## Software Technical Assessment
Software Technical Assessment
Software Description | WAMRF-010 (Rev. 07), Calibration of Weight Carts, SOP 33 | Approval Date | 2014-11-06
Assessed By | Dan Wright | Approved By | Dan Wright
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | All worksheet are appropriately named and easy to understand
| There are instructions for use | Pass | Instructions worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | The workbook opens to the instruction worksheet
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank, light blue for non-blank cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Specified on Uncertainty Analysis worksheet
| The number of digits to be rounded to is specified | Pass | Mass corrections and uncertainties are rounded according to NISTIR 6969, GLP 9
| The user is warned/notified whenever data-entry fields are left blank | Pass | In data validation cells only
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry fields on the workbook template are blank. Instructions tell user to break links before saving after a calibration is performed
| The software opens at the right location within the file | Pass | Opens to the Instruction worksheet
| Unused fields/cells are locked | Pass | Unused cells are locked and worksheets are protected
| Rows/columns that the operator need not see are hidden | N/A | No hidden cells or worksheets
| Unused sheets are removed | Pass | No unused worksheets
| Worksheets are named appropriately | Pass | Worksheets are appropriately named
B. Mathematical Specification | The correct SOP is used | Pass | SOP 33 is embedded
| The formulae and methods chosen from that SOP are specified | Pass | Formulae are show in the SOP 33 and Uncertainty Analysis worksheets
| Sources and references for formulae are specified | Pass | As specified in SOP 33
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | Pass | For Weight Cart Mass Echelon III calibrations
C. Code review | The formulae in the fields exactly match the SOP | Pass | Formulae on calculation worksheet match SOP
| Repeated calculations appropriately reference the correct cells | Pass | No issues encountered
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | Pass | Tested using data set
| Rounding is done at the appropriate locations in the file | Pass | Specified on Calculations worksheet
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | Pass | No issues encountered
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | Pass | Formatted correctly throughout workbook
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | Pass | Where appropriate
E. Component Testing | Each macro used is functional | Pass | FORMAT_CX and ROUND_UNC
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | Pass | No issues encountered
| Plotted graphs are accurate | Pass | Graphs on uncertainty worksheet display properly
| Worksheets/reports print properly, if needed to | Pass | Worksheets formatted for printing as necessary
| Conditional (color and non-color) formatting is functional | Pass | On data entry and calculation worksheets
F. Numerical Reference | Look-up tables and lists match the latest calibration report. | Pass | Linked tables are validated and up to date
| Uncertainties match the latest Scope | Pass | Test on uncertainty worksheet will not allow an uncertainty that is better than the latest published scope
| Values that reference another workbook or spreadsheet are dated | Pass | Linked references have been validated and are up to date
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct | Pass | All embedded data has been validated
| The evaluation of the embedded data is dated and documented | Pass | Records keep in Software Validation folder on computer hard drive
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | Pass | Evaluated and filed
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | Pass | Hand calculations match
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Cells are locked and worksheets are password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | State Metrologist and Program Manager only have access
| Files are backed up automatically | Pass | Backed up to agency servers weekly
| Additional back-up is available at alternate facilities | Pass | Agency IT backs up servers weekly
| Files on network drives cannot be accidentally deleted | Pass | Agency IT keeps an additional back up at a remote location

## Documentation
Documentation
Date | Description | Initials
2012-08-15 | This spreadsheet was developed by Dan Wright, WA Dept of Agriculture. This workbook follows the procedure documented in NISTIR 6969, SOP 33. | DAW
2012-08-15 | Validated ranges. | DAW
2012-08-15 | Validated workbook, file name "WA2012-08-15 WAMRF-010 (Rev. 03), SOP 33 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-02-17 | Removed GLP Rounding macros | DAW
2013-02-17 | Entered new rounding technique and corrected formulas | DAW
2013-02-17 | Validated range names. | DAW
2013-02-17 | Validated workbook, file name "WA2013-02-17 WAMRF-010 (Rev. 04), SOP 33 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-07-29 | Updated input for and added the Welch-Satterthwaite formula for effective degrees of freedom. | DAW
2013-07-29 | Revised all ISBLANK() statements. | DAW
2013-07-29 | Modified Uncertainty worksheet | DAW
2013-07-29 | Validated range names and spell checked. | DAW
2013-07-29 | Validated workbook, file name "WA2013-07-29 WAMRF-010 (Rev. 05), SOP 33 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-08-28 | Fixed standard uncertainty formulas in cells D39 & D40 on the Uncertainty Analysis worksheet | DAW
2013-08-28 | Fixed % Contibution formulas in cells I38 to I41 on the Uncertainty Analysis worksheet | DAW
2013-08-28 | Fixed Final U formula in cell I51 on the Uncertainty Analysis worksheet | DAW
2013-08-28 | Changed mg to g in the chart legend on the Uncertainty Analysis worksheet | DAW
2013-08-28 | Validated range names and spell checked. | DAW
2013-08-28 | Validated workbook, file name "WA2013-08-28 WAMRF-010 (Rev. 06), SOP 33 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-11-22 | Changed Start to Finish in cell A110 on Data Entry sheet | DAW
2014-04-10 | Referenced As Left Conventional Mass and Conventional Mass Correction to the As Found value when no adjustment is required on the Calculations worksheet. | DAW
2014-10-29 | Changed statistical df from 100 to 10000. | DAW
2014-11-04 | Embedded 2014 edition of SOP 33.
2014-11-04 | Added note 5 in Instructions concerning dependent and independent standards.
2014-11-04 | Created an evaluation of the standard deviation vs the balance resolution per NISTIR 6969, SOP 4, paragraph 5.2.1. | DAW
2014-11-06 | Moved uncertainty evaluation to Data Entry worksheet and added Precision Test. | DAW
2014-11-06 | CIPM_Pa was defaulted to nominal (0.0012 g/cm³) so the uncertainty evaluation calculate before measurement are made. | DAW
2014-11-06 | Validated range names. | DAW
2014-11-06 | Validated workbook, file name "WA2014-11-06 WAMRF-010 (Rev. 07), SOP 33 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW

## SOP 33
Note on A1: Unprotect the worksheet, then double click the document to open in Adobe Acrobat
Unprotect the worksheet, then double click the document to open in Adobe Acrobat

## Instructions
Instructions
No. | Description
1 | Validate this workbook after any changes.
2 | The worksheets are password protected. Password is password. You should protect each sheet with your own unique password. Place a password hint here so you don't forget it. Password hint is: "?"
3 | This workbook has links to the Standards, Customer List, and Mass Tolerances workbooks
4 | Break links after calibration to lock in the data from linked workbooks.
5 | The calibration of the standards are treated as dependent in the Data Entry worksheet. If you are using independent calibrated standards you need to change the formula in cell F56 on the Data Entry worksheet to a root sum square calculation. The same for cell F60 for the sensitivity weights.
Range Names
Name | Range
AF_CMx | =Calculations!$H$9
AF_Cx | =Calculations!$H$10
AL_CMx | =Calculations!$H$17
AL_Cx | =Calculations!$H$18
AVG_P | ='CIPM Air Density'!$D$16
AVG_t | ='CIPM Air Density'!$D$17
AVG_U | ='CIPM Air Density'!$D$18
Balance | ='Data Entry'!$E$16
BalID | ='Data Entry'!$C$16
BalUnit | ='Data Entry'!$J$16
Best_df | ='Data Entry'!$H$19
Best_k | ='Data Entry'!$J$19
Best_U | ='Data Entry'!$E$19
Cal_Date | ='Data Entry'!$C$21
CIPM_Pa | ='CIPM Air Density'!$D$19
Condition | ='Data Entry'!$C$14
Cs | ='Data Entry'!$D$56
Csw | ='Data Entry'!$D$60
CustomerID | ='Data Entry'!$J$4
Description | ='Data Entry'!$C$10
Final_k | ='Uncertainty Analysis'!$I$50
Final_U | ='Uncertainty Analysis'!$I$51
Final_Veff | ='Uncertainty Analysis'!$I$49
ID | ='Data Entry'!$C$13
Interval | ='Data Entry'!$J$20
IntervalQ | ='Data Entry'!$F$20
JobNo | ='Data Entry'!$I$12
k | ='Uncertainty Analysis'!$I$44
k_S | ='Data Entry'!$H$56
k_sw | ='Data Entry'!$H$60
MeasuredU | ='Uncertainty Analysis'!$I$45
Mfg | ='Data Entry'!$C$11
Mfg_Date | ='Data Entry'!$I$11
Nx | ='Data Entry'!$I$13
Obs1_1 | ='Data Entry'!$G$71
Obs1_2 | ='Data Entry'!$G$72
Obs1_3 | ='Data Entry'!$G$73
Obs1_4 | ='Data Entry'!$G$74
Obs2_1 | ='Data Entry'!$G$77
Obs2_2 | ='Data Entry'!$G$78
Obs2_3 | ='Data Entry'!$G$79
Obs2_4 | ='Data Entry'!$G$80
Obs3_1 | ='Data Entry'!$G$83
Obs3_2 | ='Data Entry'!$G$84
Obs3_3 | ='Data Entry'!$G$85
Obs3_4 | ='Data Entry'!$G$86
Obs4_1 | ='Data Entry'!$G$94
Obs4_2 | ='Data Entry'!$G$95
Obs4_3 | ='Data Entry'!$G$96
Obs4_4 | ='Data Entry'!$G$97
Obs5_1 | ='Data Entry'!$G$100
Obs5_2 | ='Data Entry'!$G$101
Obs5_3 | ='Data Entry'!$G$102
Obs5_4 | ='Data Entry'!$G$103
Obs6_1 | ='Data Entry'!$G$106
Obs6_2 | ='Data Entry'!$G$107
Obs6_3 | ='Data Entry'!$G$108
Obs6_4 | ='Data Entry'!$G$109
P | ='CIPM Air Density'!$D$4
P_1 | ='Data Entry'!$G$69
P_1_left | ='Data Entry'!$G$92
P_2 | ='Data Entry'!$G$87
P_2_left | ='Data Entry'!$G$110
P_corr | ='Data Entry'!$D$65
P_left | ='CIPM Air Density'!$D$10
PO_No | ='Data Entry'!$H$5
POC | ='Data Entry'!$H$6
POC_Phone | ='Data Entry'!$H$7
Print_Titles | =Instructions!$7:$8
Ps | ='Data Entry'!$J$56
Psw | ='Data Entry'!$J$60
Px | ='Data Entry'!$I$14
R_mean | ='Data Entry'!$I$18
Receive_Date | ='Data Entry'!$I$10
Resolution | ='Data Entry'!$H$16
RptNo | ='Data Entry'!$H$1
S | ='Data Entry'!$B$56
S_df | ='Data Entry'!$I$56
SealNo | ='Data Entry'!$I$112
SN | ='Data Entry'!$C$12
sp | ='Data Entry'!$F$17
sp_df | ='Data Entry'!$J$17
sp_unit | ='Data Entry'!$A$17
StUnit | ='Data Entry'!$A$17
sw | ='Data Entry'!$B$60
sw_df | ='Data Entry'!$I$60
t | ='CIPM Air Density'!$D$5
t_1 | ='Data Entry'!$E$69
t_1_left | ='Data Entry'!$E$92
t_2 | ='Data Entry'!$E$87
t_2_left | ='Data Entry'!$E$110
t_corr | ='Data Entry'!$D$64
t_left | ='CIPM Air Density'!$D$11
Tech | ='Data Entry'!$F$21
Tech_Int | ='Data Entry'!$J$21
Tol | ='Tables and Lists'!$A$4:$B$8
Tol_Class | ='Data Entry'!$C$15
Tol_List | ='Tables and Lists'!$A$4:$A$22
Tolerance | ='Data Entry'!$I$15
U | ='CIPM Air Density'!$D$6
U_1 | ='Data Entry'!$I$69
U_1_left | ='Data Entry'!$I$92
U_2 | ='Data Entry'!$I$87
U_2_left | ='Data Entry'!$I$110
U_corr | ='Data Entry'!$D$66
U_left | ='CIPM Air Density'!$D$12
U_P | ='Data Entry'!$F$65
U_P_k | ='Data Entry'!$H$65
U_t | ='Data Entry'!$F$64
U_t_k | ='Data Entry'!$H$64
U_U | ='Data Entry'!$F$66
U_U_k | ='Data Entry'!$H$66
UCL | =Calculations!$D$29
Us | ='Data Entry'!$F$56
Usw | ='Data Entry'!$F$60
Veff | ='Uncertainty Analysis'!$I$43
WL | =Calculations!$D$30
Xnom | ='Data Entry'!$F$13

## Report Pages 1-3
REPORT OF CALIBRATION
Issued To:
Point of Contact:
Purchase Order Number:
Report Number:
Calibration Date:
Calibration Due Date:
This is to certify that the information contained in this report is true and correct as of the date of calibration.
Dan Wright, Program Specialist 3 |  |  |  |  | Date of Issue
Accreditation LOGO |  | Accredited by the Accreditation Authority for the specific scope of accreditation under lab code XXXXXX-X. This report may not be used to claim product endorsement by Accreditation Authority or any other government agency, and may not be reproduced, except in full, without written approval from the laboratory.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date:
Artifact(s) Description
Test Item: | lb Weight Cart |  |  | Date Received:
Serial Number: |  |  |  | Model Number:
Manufacture: |  |  |  | Weight Cart Manufacture Date:
Condition: | See Inspection Checklist |  |  | Tolerance Specification: | NIST HB 105-8
Calibration Information
Job Order #: |  |  |  | Temperature:
Metrologist: |  |  |  | Pressure:
Procedure: | NIST IR 6969, SOP 33 |  |  | Humidity:
Laboratory Reference Standards Used
Description |  | Serial Number | Cert. Number |  | Cal Date | Cal Due
Traceability Statement
The artifact(s) described in this report have been compared to the Standards of the State of Washington. The Standards of the State of Washington are traceable to the National Institute of Standards and Technology (NIST) and are part of a comprehensive measurement assurance program for ensuring continued accuracy and measurement traceability within the level of uncertainty reported by this laboratory. The International System of Units (SI) for mass is the kilogram (kg) (see Conversion Factors on page 3). The report number for this report is the only unique report number to be used in referencing measurement traceability for the artifact(s) described in this report.
Uncertainty Statement
The combined standard uncertainty includes uncertainties reported for the standard, uncertainties associated with the measurement process, uncertainties for any observed deviations from reference values which are less than surveillance limits, and other uncertainties associated with the particular artifact (i.e., material density, air buoyancy corrections, etc.). The combined standard uncertainty is multiplied by a coverage factor, k, to give the expanded uncertainty, which defines an interval with a 95.45 % level of confidence. The expanded uncertainty presented in this report is consistent with NIST Technical Note 1297.
Accreditation Statement
Accredited by the Accreditation Authority for the specified scope of accreditation under lab code XXXXXX-X. This laboratory meets the requirements of ISO/IEC 17025 and ANSI/NCSL Z540-1.
Conversion Factors
From NIST Special Publication 811, Guide for the Use of the International System of Units (SI)
Factors in boldface are exact
To convert from |  |  | to |  | multiply by
pound (avoirdupois) (lb) |  |  | to kilogram (kg) |  | 4.535 923 7 E-01
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date:
Pertinent Information
● In-accordance-with ISO/IEC FDIS 17025, General Requirements for the Competence of Testing and Calibration Laboratories, paragraph 5.10.4.4 ‘A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the client. This requirement may be superseded by legal regulations.’
● In-accordance-with Washington Administrative Code (WAC) Chapter 16-663, Service Agents -- Reporting, Test Procedures, Standards And Calibration Of Weighing And Measuring Devices, Section 16-663-130, Adequacy of standards and submission of standards for certification, paragraph 2, ‘… All standards used for servicing, repairing and/or calibrating commercial weighing and measuring devices must be submitted at least every two years for examination and certification…’
● The weight cart listed above has been found and/or left within the tolerances for the specification stated above, and as close to zero error as possible, except as noted. An artifact is considered in-tolerance when the correction plus the measurement uncertainty is equal to or less than the specified tolerance. RED print indicates an out-of-tolerance reading.
● The adjusting cavity was sealed after adjustment with Seal Number
● All corrections stated in this report correlate to a “Conventional Mass” (CM), also known as ‘apparent mass”, scale verses 8.0 g/cm³ reference mass density and an air density of 1.2 mg/cm3 at 20 ºC.
● The results listed in this report relate only to the artifacts described and extent of calibrations performed.
Inspection Checklist
The weight cart was allowed to come to thermal and environmental equilibrium in the laboratory prior to calibration. The weight cart was adjusted (as needed and noted above) as close as possible to zero error. All fluid levels stated in the Inspection Checklist on page 4 were adjusted as close as possible to the full/reference marks and sealed. Liquid levels must be maintained as close to reference levels as possible during use.
The attached Inspection Checklist is an integral component of this Calibration Report and a copy must be maintained with the cart and reviewed prior to use.
Any maintenance, repairs, replacement of parts, or damage to weight cart or its components will likely result in an out-of-tolerance condition; therefore, maintenance or replacement of components such as batteries, tires, filters, or other items listed on the checklist, require calibration of the weight cart prior to subsequent use.
As Found / Left Calibration Results
Nominal Mass (lb) | As Found CM Correction (lb) | As Left CM Correction (lb) | Uncertainty ± (lb) | k-factor | Degrees of Freedom | NIST HB 105-8 Tolerance ± (lb)

## Inspection Checklist
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Inspection Date:
Inspection Checklist for Weight Cart
| Manufacturer: |  |  |  | Date of Manufacture:
| Model Number: |  |  |  | ID/SN Number:
| Nominal Mass of Weight Cart: | lb |  |  | Nominal Mass Suitably Marked:
Power Source: |  |  |  | Power Source Reference Level:
Engine Oil: |  |  |  | Engine Oil Reference Mark:
Hydraulic Oil: |  |  |  | Hydraulic Oil Reference Mark:
No. of Axles: |  | No. of Tires: |  |  | Sealed Wheel Bearings?:
Tire Sizes: | Front: |  | Middle: |  | Rear:
Fluid drain tubes extend beyond the body of the cart:
Drain holes present in locations where water may accumulate:
Weight restraint railing permanently fixed and solid:
Adjusting Cavity Accessible:
Adjusting Cavity Sealed:
|  |  | Adjusting cavity approximate capacity: | lb
Service brakes functioning properly:
Parking brakes functioning properly:
Remote control functioning properly:
General condition at time of calibration (note any accumulated dirt/debris, damage, loose parts, or evidence of tampering or unauthorized entry of seals).
List and report any repair and maintenance performed, parts replaced, etc., Leaks repaired, new battery, carburetor, exhaust system, wheels changed, welding performed, etc.. Include any comments or changes since the last calibration.

## Control Data
Control Chart Information
Note: Copy and Paste Appropriate Row(s) to Control Chart
| Technician Initials | Date | Report Number | Cx vs 8.0 (MAX) (g) | Cx vs 8.0 (MIN) (g) | Balance Description | Start Air Temperature (ºC) | Finish Air Temperature (ºC) | Start Pressure (g/cm³) | Finish Pressure (g/cm³) | Start Humidity (% RH) | Finish Humidity (% RH)
As Found
As Left

## Data Entry
Data Entry | [5 empty] | Report Number:
Customer Information
Customer ID
PO#
POC
Phone
Calibration Item & Measurement Information
Item Description |  | Weight Cart |  |  | Date Received
Manufacture |  |  |  |  | Weight Cart Manufacture Date
Serial Number |  |  |  |  | Job Order Number
ID/Model Number |  |  |  |  | Nominal Value (avdp. pound)
Artifact Condition |  | See Inspection Checklist |  |  | Assumed Density (g/cm³)
Tol. Specification |  | NIST HB 105-8 |  |  | Tolerance (± lb)
Balance ID |  |  | Balance |  |  | Resolution |  | Bal Unit
Process Standard Deviation for Balance (g) | [6 empty] | Degrees of Freedom
Range Mean from Control Chart for Balance at lb load (g)
Accredited Unc (e.g, NVLAP,A2LA, etc.) | [5 empty] | df |  | k factor
Was a Calibration Interval Requested or is one Required? | [5 empty] | What is the Interval (months)?
Calibration Date |  |  |  | Technician |  |  |  | Initials
Inspection Checklist Information
Power Source |  |  |  | Power Source Reference Level
Engine Oil |  |  |  | Engine Oil Reference Mark
Hydraulic Oil |  |  |  | Hydraulic Oil Reference Mark
Number of Axles |  |  | Number of Tires |  |  | Sealed Wheel Bearings
Tire Size: | Front |  |  | Middle |  |  | Rear
Nominal mass of the weight cart is suitably marked
Fluid drain tubes extend beyond the body of the cart
Drain holes are present in locations where water may accumulate
Weight restraint railing is permanently fixed and solid
Adjusting cavity is: |  | Accessible |  | Sealed |  | Approximate capacity (lb)
Service brakes are functioning properly
Parking brakes are functioning properly
Remote control is functioning properly
General condition at time of calibration (note any accumulated dirt/debris, damage, loose parts, or evidence of tampering or unauthorized entry of seals)
List and report any repair and maintenance performed since the last calibration (i.e., parts replaced, leaks repaired, wheels changed, welding performed, etc.)
Standards Used
Note: If the 'Calibration Due' date turns RED the standard calibration is void, calibrate before use.
Description |  |  | Serial Number |  | Report Number |  | Calibration Date | Calibration Due
Mass Standard(s) Data
Enter information for the reference weight(s) (S)
ID Code(s) for S | Nominal Mass (g) |  | CM Correction (g) |  | Us (g) |  | k factor | Degrees of Freedom (df) | Density (g/cm³)
S1
S2
S3
S4
S5
S =
Enter information for the sensitivity weight(s) (sw)
sw1
sw2
sw =
Environmental Standard(s) Data (used for MABC calculation)
Parameter |  | ID Code | Correction |  | Uncertainty |  | k factor | Unit
Temperature, Air |  | t | [5 empty] | ºC
Barometric Pressure |  | P | [5 empty] | mm Hg
Humidity |  | U | [5 empty] | % RH
Uncertainty Evaluation
Criteria: NIST HB 143, Table 4, "The expanded uncertainty must be less than 1/3 of the maximum permissible error."
|  | 1/3 of NIST HB 105-8 Tolerence = |  |  | lb
|  | Expanded Uncertainty = |  |  | lb
Precision Test (Pn) =
As Found Measurements
Start Time |  |  | t (ºC) |  | P (mm Hg) |  | U (% RH) |  | SOP 4, 2.5.2.5 Check
Observations Run One |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
ROTATE WEIGHT CART 180º
Observations Run Two |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
ROTATE WEIGHT CART 180º
Observations Run Three |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
Finish Time |  |  | t (ºC) |  | P (mm Hg) |  | U (% RH)
Adjust As Close To Zero Error As Possible And Re-Test
As Left Measurements
Start Time |  |  | t (ºC) |  | P (mm Hg) |  | U (% RH) |  | SOP 4, 2.5.2.5 Check
Observations Run One |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
ROTATE WEIGHT CART 180º
Observations Run Two |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
ROTATE WEIGHT CART 180º
Observations Run Three |  |  | Weights |  |  | Balance Readings ()
1 |  |  | X + tx
2 |  |  | S + ts
3 |  |  | S + sw
4 |  |  | X + sw
Finish Time |  |  | t (ºC) |  | P (mm Hg) |  | U (% RH)
Adjusting Cavity Calibration Seal Number:

## Calculations
Calculations
Measurement Determinations
As Found Measurement Observations Option B (XSSX) (g)
|  |  |  | As Found Weight Cart Conventional Mass, CM vs 8.0 g/cm³, First Double Substitution = | lb
|  |  |  | As Found Weight Cart Conventional Mass, CM vs 8.0 g/cm³, Second Double Substitution = | lb
|  |  |  | As Found Weight Cart Conventional Mass, CM vs 8.0 g/cm³, Third Double Substitution = | lb
|  |  |  | Average As Found Weight Cart Conventional Mass, CMx vs 8.0 g/cm³ = | lb
|  |  |  | As Found Weight Cart Conventional Mass Correction, Cx vs 8.0 g/cm³ = | lb
As Left Measurement Observations Option B (XSSX) (g)
|  |  |  | As Left Weight Cart Conventional Mass, CM vs 8.0 g/cm³, First Double Substitution = | lb
|  |  |  | As Left Weight Cart Conventional Mass, CM vs 8.0 g/cm³, Second Double Substitution = | lb
|  |  |  | As Left Weight Cart Conventional Mass, CM vs 8.0 g/cm³, Third Double Substitution = | lb
|  |  |  | Average As Left Weight Cart Conventional Mass, CMx vs 8.0 g/cm³ = | lb
|  |  |  | As Left Weight Cart Conventional Mass Correction, Cx vs 8.0 g/cm³ = | lb
Tolerance Compliance Evaluation
Criteria: The conventional mass shall not differ more than the maximum permissible error minus the expanded uncertainty from the nominal value of weight. (ASTM E 617, NIST HB 105-1 & OIML R 111)
| NIST HB 105-8 Tolerence (±) = |  | lb
| Expanded Uncertainty (±) = |  | lb
| As Found Cx vs 8.0 = |  | lb
Measurement Assurance Evaluations
Criteria: The observed range must be less than the Control Chart Control Limit to be considered in control.
| Upper Control Limit (Rmean*3.267) = |  | lb
| Warning Limit (Rmean*2.512) = |  | lb
| As Found Range of the Measurements = |  | lb
Report Conventional Mass Correction and Uncertainty (uc*k) As Follows:
Mass corrections and uncertainties are rounded according to NISTIR 6969, GLP 9
|  | Nominal Value = |  | lb
|  | As Found Conventional Mass = |  | lb
|  | As Found Conventional Mass Correction = |  | lb
|  | As Left Conventional Mass = |  | lb
|  | As Left Conventional Mass Correction = |  | lb
Effective Degrees of Freedom (veff) =
k factor =
|  | Uncertainty (uc*k) = |  | lb
|  | NIST HB 105-8 Tolerence (±) = |  | lb
|  | Assumed Density = |  | g/cm³

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NISTIR 6969, SOP 33, Calibration of Weight Carts using SOP 4 (Modified for triple double substitution measurements)
Specify Equation(s)
Double Substitution Optional Sequence B (XSSX) w/o Air Buoyancy Correction
where
S = | standard weight
X = | weight calibrated
t = | small calibrated tare weight, A subscript s or x is used to indicate the larger weight with which it is associated
sw = | small calibrated weight used to evaluate the sensitivity of the balance
N = | the nominal value of a specific weight. Subscripts s, x, are used to identify the weight.
C = | the correction for a specific weight. Subscripts s, x, are used to identify the weight.
CM = | the conventional mass of a specific weight. Subscripts s, x, t, sw are used to identify the weight.
ρa = | density of air at time of calibration
rn = | density of normal air (1.2 kg/m³)
ρ = | density of masses; subscripts s, x, ts, tx, sw are used to identify the weight
Estimating the Magnitude of the Air Buoyancy Correction
where
m0 = | nominal mass in grams
ρa = | density of air at time of measurement or normally found in the laboratory in g/cm³
ρn = | density of "normal" air, i.e., 0.0012 g/cm³
ρX = | density of the unknown mass
ρS = | density of the mass standard
Identify Components
Component | Description |  |  |  | Reference
s(p) | Uncertainty associated with the measurement process from the average range of several sets of measurements |  |  |  | NISTIR 6969, SOP 29, Para. 3.2.1.2. Standard deviation from a set of replicate measurements and NISTIR 6969, Statistical Techniques, paragraphs 8.3 & 8.4, Pooled standard deviation from standard deviation for each set of replicates from a control chart
u(s) | Uncertainty associated with the reference weight(s) and sensitivity weight |  |  |  | Calibration Report divided by the k factor
u(MABC) | Uncertainty associated with the magnitude of the air buoyancy correction where air buoyancy is not used |  |  |  | NISTIR 6969, SOP 2, Paragraph 2.3.2, Estimating the Magnitude of the Air Buoyancy Correction
Quantify, Convert, Combined & Expand
Component Description |  | Component Symbol | Reported Unc (g) | Type | Distribution | Factor to Normalize | Standard Unc (g) | % Contribution | Degrees of Freedom
Standard Deviation from CC |  | s(p) | 0 | A | Normal | 1 | 0 |  | 1
Report U divided by k-factor |  | u(s) | 0 | B | Normal | 1 | 0 |  | 1
Report U divided by k-factor |  | u(sw) | 0 | B | Normal | 1 | 0 |  | 1
Magnitude of ABC |  | u(MABC) | 0 | B | Rectangular | 1.7321 | 0 |  | 1
[7 empty] | Combined U (uc) = |  | g
Effective Degrees of Freedom, Welch-Satterthwaite formula, NIST Technical Note 1297, (B-1) (veff) =
k factor =
[7 empty] | Expanded U (uc*k) = |  | g
NVLAP Effective Degrees of Freedom (veff) =
NVLAP k factor =
[7 empty] | NVLAP U = |  | g
Final Effective Degrees of Freedom (veff) =
Final k factor =
[7 empty] | Final U (uc*k) = |  | g

## CIPM Air Density
CIPM Air Density
Air Density, SOP 2, Option B, 2012 (CIPM 2007)
Air Density For As Found Measurement Observations
Barometric Pressure (mm Hg)
Temperature (ºC)
Relative Humidity (%)
Calculated Air Density (mg/cm³)
Air Density For As Left Measurement Observations
Barometric Pressure (mm Hg)
Temperature (ºC)
Relative Humidity (%)
Calculated Air Density (mg/cm³)
Average Air Density Note: Air density results are rounded to 9 digits (g/cm³) and used for MABC calculation.
Average Barometric Pressure (mm Hg)
Average Temperature (ºC)
Average Relative Humidity (%)
Average Calculated Air Density (mg/cm³) | 0.0012

## Tables and Lists
Tables & Lists
Tolerance Lookup Table (Tol) (lb)
2000 | 0.5
3000 | 1
4000 | 1.25
5000 | 1.5
6000 | 2