## CIPM Air Density 1981_91
Password to unlock: | metrology
Do not delete these cells
Air Density Calculation |  |  |  | They are used to calculate air density.
Barometric Pressure: | 760 | mm Hg |  | 101325 | Pa
Temperature: | 20 | °C |  | 293.15 | °K
Relative Humidity: | 40 | % |  | 40 | %
Calculated Air Density: | 1.2003 | mg/cm3 | 0.0012003 | g/cm3
NOTE: When copying these cells into a spreadsheet, they must be pasted into exactly
the same cells on a new worksheet (i.e., the block from A3 to G7).
Cell references in the equations are "absolute" and fixed to the reference cells in each equation.
Once the cells are pasted, the entire block may be dragged/dropped.

## Air Density - Validation
Air Density Calculations
Password to unlock: |  | metrology |  |  | Data Entry and Data Set Comparisons |  |  |  | Equations and Variables
Active Fortran Program | [7 empty] | SOP 2, Option A |  | Reference Equations
[5 empty] | Pressure: | 760 | mm Hg | ra | 1.2001 | ((0.4646*(H4-(0.003796*H6*L5)))/(273.15+H5))
[5 empty] | Temperature: | 20 | C | es | 17.5357 | 1314600000*EXP(-5315.56/(H5+273.15))
[5 empty] | Relative humidity: | 42 | %
SOP 2, Option B (preferred for all precision work)
[5 empty] | Pressure (converted): | 101325 | Pa | ra | 1.2001 | (L10*(L11/(L12*L13*L14))*(1-0.378*L15)), Giacomo version of base equation (1982)
[5 empty] | Temperature (converted): | 293.15 | K | Ma | 0.028963 | 0.0289635 (28.96386 in 1981 version)
[8 empty] | p | 101325 | Pa
[8 empty] | Z | 0.99962 | 1-((L11/L14)*(L24+L25*H5+L26*H5^2+(L27+L28*H5)*L15+(L29+L30*H5)*L15^2))+(L11^2/L14^2)*(L31+L32*L15^2)
CIPM 81/91 Paper Data Sets, Table 2 |  |  |  |  | Reference Conversions |  |  | R | 8.3145 | 8.3145
| Pa | t | RH |  | mm Hg | Pa |  | T | 293.15 | K
Data Set 1 | 100000 | 20 | 50 |  | 750.0617 | 100000 |  | xv | 0.009735 | (L16/100)*L17*(L18/L19)
Data Set 2 | 110000 | 20 | 10 |  | 825.0679 | 110000 |  | h | 42 | %
Data Set 3 | 100000 | 15 | 90 |  | 750.0617 | 100000 |  | f | 1.004 | 1.00062+(0.0000000314*L19)+(0.00000056*H5^2)
Data Set 4 | 60000 | 25 | 50 |  | 450.037 | 60000 |  | rsv | 2339.1632 | 1*EXP(L20*L14^2+L21*L14+L22+L23/L14)
Uncertainty in Air Density Equation: |  |  |  | 0.0012 | Pa = mm Hg * 133.322368421053 |  |  | p | 101325 | Pa
[8 empty] | A | 1.2379e-05 | 1.2379e-05
Comparison of Calculation Results | [7 empty] | B | -0.019121 | -0.019121
Set 1 | [7 empty] | C | 33.9371 | 33.9371
|  |  |  | SOP 2, A | 1.1835 |  |  | D | -6343.1645 | -6343.1645
|  |  |  | SOP 2, B | 1.1835 |  |  | a0 | 1.5812e-06 | 1.5812e-06
|  |  |  | air91co24.exe | 1.1835 |  |  | a1 | -2.9331e-08 | -2.9331e-08
[8 empty] | a2 | 1.1043e-10 | 1.1043e-10
|  |  |  | CIPM Paper | 1.1835 |  |  | b0 | 5.707e-06 | 5.707e-06
[8 empty] | b1 | -2.051e-08 | -2.051e-08
[8 empty] | c0 | 0.00019898 | 0.00019898
Set 2 | [7 empty] | c1 | -2.376e-06 | -2.376e-06
|  |  |  | SOP 2, A | 1.3066 |  |  | d | 1.83e-11 | 1.83e-11
|  |  |  | SOP 2, B | 1.3066 |  |  | e | -7.65e-09 | -7.65e-09
|  |  |  | air91co24.exe | 1.3066
References:
|  |  |  | CIPM Paper | 1.3066 |  |  | Equation for the Determination of the Density of Moist Air (1981)
Metrologia 1982, 18, 33-40
P. Giacomo
Set 3 | [7 empty] | Equation for the Determination of the Density of Moist Air (1981/91)
|  |  |  | SOP 2, A | 1.2023 |  |  | Metrologia, 1992, 29, 67-70
|  |  |  | SOP 2, B | 1.2024 |  |  | R. S. Davis
|  |  |  | air91co24.exe | 1.2024
Handbook 145, SOP 2
|  |  |  | CIPM Paper | 1.2024 |  |  | R. S. Davis
Air91co24.exe
Created in Fortran by Zeina J. Jabbour
Set 4 | [7 empty] | Based on 4 % CO2 concentration.
|  |  |  | SOP 2, A | 0.69425
|  |  |  | SOP 2, B | 0.69416 |  |  | air91co24.xls
|  |  |  | air91co24.exe | 0.69416 |  |  | Created by Georgia L. Harris, Feb 2003
|  |  |  | CIMP Paper | 0.69416 |  |  | NIST SP 811; conversion factors for Pa and mm Hg

## Kragten - Air Density Unc
| CIMP 1981/91 Formula, R. S. Davis | [5 empty] | Air Density Calculations
| Kragten Spreadsheet |  | A | B | C | D
|  |  | P (mm Hg) | T (°C) | RH (%) | Equation
|  | 1 | 2 | 3 | 4 | 6 | Pressure: | 760 | mm Hg | Pressure: | 760.5 | mm Hg | Pressure: | 760 | mm Hg | Pressure: | 760 | mm Hg
| 1 | Value | 760 | 20 | 42 | 0 | Temperature: | 20 | C | Temperature: | 20 | C | Temperature: | 20.1 | C | Temperature: | 20 | C
| 2 | Uncertainty | 0.5 | 0.1 | 10 | 0.0012 | Relative humidity: | 42 | % | Relative humidity: | 42 | % | Relative humidity: | 42 | % | Relative humidity: | 52 | %
P (mm Hg) | 3 | 760 | 760.5 | 760 | 760 | 760
T (°C) | 4 | 20 | 20 | 20.1 | 20 | 20 | 133.3224
RH (%) | 5 | 42 | 42 | 42 | 52 | 42 | ra | 1.2001 |  | ra | 1.2009 |  | ra | 1.1996 |  | ra | 1.199
Equation | 6 | 0 | 0 | 0 | 0 | 0.0012 | Ma | 0.028963 |  | Ma | 0.028963 |  | Ma | 0.028963 |  | Ma | 0.028963
Corr (x) | 7 | 1.2001 | 1.2009 | 1.1996 | 1.199 | 1.2013 | p | 101325 |  | p | 101391.6612 |  | p | 101325 |  | p | 101325
u(y,xi) | 8 |  | 0.0007927 | -0.00043776 | -0.001047 | 0.0012 | Z | 0.99962 |  | Z | 0.99962 |  | Z | 0.99962 |  | Z | 0.99961
u(y)2, u(y,xi)2 | 9 | 3.3562e-06 | 6.2837e-07 | 1.9163e-07 | 1.0962e-06 | 1.44e-06 | R | 8.3145 |  | R | 8.3145 |  | R | 8.3145 |  | R | 8.3145
u(c) | 10 | 0.001832 | 43.27% | 23.9% | 57.15% | 65.5% | T | 293.15 |  | T | 293.15 |  | T | 293.25 |  | T | 293.15
k | 11 | 2 | 18.72% | 5.71% | 32.66% | 42.91% | xv | 0.009735 |  | xv | 0.0097287 |  | xv | 0.0097955 |  | xv | 0.012053
k*u | 12 | 0.003664 |  |  |  | 100% | h | 42 |  | h | 42 |  | h | 42 |  | h | 52
100*ku/c(x) | 13 | 0.15266 | < relative unc. (%) at k=1 |  |  |  | f | 1.004 |  | f | 1.004 |  | f | 1.004 |  | f | 1.004
[7 empty] | rsv | 2339.1632 |  | rsv | 2339.1632 |  | rsv | 2353.6933 |  | rsv | 2339.1632
Green Cells for P, T, RH, and their uncertainties may be changed. | [6 empty] | p | 101325 |  | p | 101391.6612 |  | p | 101325 |  | p | 101325
[7 empty] | A | 1.2379e-05 |  | A | 1.2379e-05 |  | A | 1.2379e-05 |  | A | 1.2379e-05
[7 empty] | B | -0.019121 |  | B | -0.019121 |  | B | -0.019121 |  | B | -0.019121
[7 empty] | C | 33.9371 |  | C | 33.9371 |  | C | 33.9371 |  | C | 33.9371
[7 empty] | D | -6343.1645 |  | D | -6343.1645 |  | D | -6343.1645 |  | D | -6343.1645
[7 empty] | a0 | 1.5812e-06 |  | a0 | 1.5812e-06 |  | a0 | 1.5812e-06 |  | a0 | 1.5812e-06
[7 empty] | a1 | -2.9331e-08 |  | a1 | -2.9331e-08 |  | a1 | -2.9331e-08 |  | a1 | -2.9331e-08
[7 empty] | a2 | 1.1043e-10 |  | a2 | 1.1043e-10 |  | a2 | 1.1043e-10 |  | a2 | 1.1043e-10
[7 empty] | b0 | 5.707e-06 |  | b0 | 5.707e-06 |  | b0 | 5.707e-06 |  | b0 | 5.707e-06
[7 empty] | b1 | -2.051e-08 |  | b1 | -2.051e-08 |  | b1 | -2.051e-08 |  | b1 | -2.051e-08
[7 empty] | c0 | 0.00019898 |  | c0 | 0.00019898 |  | c0 | 0.00019898 |  | c0 | 0.00019898
[7 empty] | c1 | -2.376e-06 |  | c1 | -2.376e-06 |  | c1 | -2.376e-06 |  | c1 | -2.376e-06
[7 empty] | d | 1.83e-11 |  | d | 1.83e-11 |  | d | 1.83e-11 |  | d | 1.83e-11
[7 empty] | e | -7.65e-09 |  | e | -7.65e-09 |  | e | -7.65e-09 |  | e | -7.65e-09
Password to unlock: |  | metrology

## MABC & MABC Uncertainty
SOP 2, MABC
| Kragten Spreadsheet |  | A | B | C | D | E | Entered Values |  |  | Uncertainties |  | Calculated Values
|  |  | Mass | ra | rn | rx | rs | Mass | 100 | g | 0.025 |  | Delta rho | -0.05
|  | 1 | 2 | 3 | 4 | 6 | 7 | ra | 1.15 | mg/cm3 | 0.0012 |  | Vx | 12.7551
| 1 | Value | 100 | 1.15 | 1.2 | 7.84 | 8 | rn | 1.2 | mg/cm3 | 0 |  | Vs | 12.5
| 2 | Uncertainty | 0.025 | 0.0012 | 0 | 0.05 | 0.008 | rx | 7.84 | g/cm3 | 0.05 |  | MABC | -0.012755
Mass | 3 | 100 | 100.025 | 100 | 100 | 100 | 100 | rs | 8 | g/cm3 | 0.008
ra | 4 | 1.15 | 1.15 | 1.1512 | 1.15 | 1.15 | 1.15
rn | 5 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
rx | 6 | 7.84 | 7.84 | 7.84 | 7.84 | 7.89 | 7.84 | R 111 Uncertainty for the buoyancy correction.
rs |  | 8 | 8 | 8 | 8 | 8 | 8.008
Corr (x) | 7 | -0.012755 | -0.012758 | -0.012449 | -0.012755 | -0.0087136 | -0.013379
u(y,xi) | 8 |  | -3.1888e-06 | 0.00030612 | 0 | 0.0040415 | -0.00062438
u(y)2, u(y,xi)2 | 9 | 1.6818e-05 | 1.0168e-11 | 9.3711e-08 | 0 | 1.6334e-05 | 3.8984e-07
u(c) | 10 | 0.0041009 | 0.07776% | 7.465% | 0% | 98.55% | 15.23% | ubc | 0.00030612
k | 11 | 2 | 6.046e-05% | 0.5572% | 0% | 97.12% | 2.318%
k*u | 12 | 1: 0.0082019; 7: 100%
100*ku/c(x) | 13 | 1: -32.1513; 2: < relative unc. (%) at k=1
Green Cells for Mass, air density, normal air density, density of the standard and unknown and their uncertainties
may be changed.
[9 empty] | m0 | 1 | 1 | 1 | 1 | 1
[9 empty] | rho a | 0.9 | 1 | 1.1 | 1.2 | 1.3
[9 empty] | rho n | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
[9 empty] | rho x | 7 | 7.5 | 7.996 | 8 | 8.4
[9 empty] | rho s | 8 | 8 | 8 | 8 | 8
[10 empty] | -0.0053571 | -0.0016667 | -6.2531e-06 | 0 | -0.00059524
Password to unlock: |  | 1: metrology
Note on D5: Georgia L Harris: