## Documentation
Documentation
Date | Description | Initials
2004-01-30 | Created workbook. | DAW
2004-03-03 | Updated water density formula to the Patterson/Morris equation. | DAW
2004-03-05 | Updated water density formula to the Patterson/Morris equation on the 'Uncertainty Analysis' worksheet. | DAW
2004-03-11 | Made suggested changed from NIST, WMD. Revised pressure measurements. | DAW
2004-07-06 | Updated uncertainty analysis to new format. | DAW
2004-09-10 | Updated uncertainty analysis. | DAW
2005-03-22 | Corrected Z60 formula on 'Uncertainty Analysis' workdheet. | DAW
2005-10-16 | Removed unnecessary information on 'NSCV Data Entry' worksheet. | DAW
2006-09-01 | Complete revision. | DAW
2006-11-14 | Changed reference to ASTM 674 to ASTM 694 on 'Uncertainty Analysis' worksheet. | DAW
2006-11-14 | Deleted uncertainty for viscosity and added statement to report of calibration. | DAW
2006-11-14 | Deleted note about standard reference temperature. | DAW
2007-03-16 | Major revision. Updated to Kragten uncertainty spreadsheet, added tolerances, added metric vessel capability, etc. | DAW
2007-03-30 | Reformatted standards tables, corrected errors in 'LPG Data Entry' worksheet. | DAW
2007-04-05 | Fixed conditional formating in uncertainty test on the 'Calculations' worksheet. | DAW
2007-09-24 | Rounded water density formulas to 8 places, revised tolerance evaluation, and revised 'NSCV Data Entry' worksheet. | DAW
2008-01-10 | Revised 'NSCV Data Entry' worksheet. | DAW
2008-02-07 | Removed 'NSCV Data Entry' worksheet. | DAW
2008-06-10 | Added 'Control Data' worksheet, removed 'Standards' worksheet, reformated vlookup formulas to global standards workbooks, and revised 'Documentation' worksheet. | DAW
2009-01-21 | Color coded data entry cells. Edited all comments. | DAW
2009-01-21 | Added Standards Used table on Data Entry worksheet. | DAW
2009-10-30 | Added Information worksheet. | DAW
2009-10-30 | Validated links | DAW
2009-10-30 | Validated workbook, file name "WA2009-10-30 WAMRF-014 (Rev. 20), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2010-03-26 | Updated criteria for tolerance and uncertainty evaluations on the "Calculations' sheet to NIST HB 105-4(2010) specifications. | DAW
2010-03-26 | Removed NSCV value. The neck scale either meets the new specification tolerances or it is rejected. | DAW
2010-03-26 | Validated workbook, file name "WA2010-03-26 WAMRF-014 (Rev. 21), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW

## Information
Workbook Information
No. | Description
1 | This workbook was developed by Dan Wright, Washington State Department of Agriculture Metrology Laboratory. It follows the procedure documented in NISTIR 7383, SOP 21. NIST/WMD has evaluated this spreadsheet and provided useful feedback in its development.
2 | This workbook must be validated after any changes or modifications.
3 | The worksheets are password protected. Password is metrology. You should protect each sheet with your own unique password. Place a password hint here so you don't forget it. Password hint is: "?"
4 | Validate all links and ranges after any changes and during annual reviews.
5 | Modify the report pages with your laboratory and accredition authority information.
6 | Modify the footer information on the report sheets and the header information on all other sheets if you want to use your own form number for the template.
Range Names
Date | Description
_unc1 | ='LPG Prover Data Entry'!$I$35
_unc2 | ='LPG Prover Data Entry'!$I$36
_unc3 | ='LPG Prover Data Entry'!$I$37
_unc4 | ='LPG Prover Data Entry'!$I$38
a_1 | ='LPG Prover Data Entry'!$H$35
a_2 | ='LPG Prover Data Entry'!$H$36
a_3 | ='LPG Prover Data Entry'!$H$37
a_4 | ='LPG Prover Data Entry'!$H$38
Address_1 | ='LPG Prover Data Entry'!$C$5
AirTemp | ='LPG Prover Data Entry'!$D$46
AirTemp_F | ='LPG Prover Data Entry'!$D$99
answer.list | ='Tables & Lists'!$G$14:$G$15
AsFound | =Calculations!$E$64
AsLeft | =Calculations!$E$65
B | ='LPG Prover Data Entry'!$I$13
Cal_Date | ='LPG Prover Data Entry'!$C$20
CCE.Table | ='Tables & Lists'!$D$5:$E$7
Condition | ='LPG Prover Data Entry'!$C$14
CustomerID | ='LPG Prover Data Entry'!$J$4
d2.Table | =Calculations!$A$77:$C$106
Description | ='LPG Prover Data Entry'!$C$10
drops1 | ='LPG Prover Data Entry'!$D$35
drops2 | ='LPG Prover Data Entry'!$D$36
drops3 | ='LPG Prover Data Entry'!$D$37
drops4 | ='LPG Prover Data Entry'!$D$38
grads1 | ='LPG Prover Data Entry'!$G$35
grads2 | ='LPG Prover Data Entry'!$G$36
grads3 | ='LPG Prover Data Entry'!$G$37
grads4 | ='LPG Prover Data Entry'!$G$38
Humidity | ='LPG Prover Data Entry'!$G$46
Humidity_F | ='LPG Prover Data Entry'!$G$99
increment | ='LPG Prover Data Entry'!$J$12
interval | ='LPG Prover Data Entry'!$J$19
intervalQ | ='LPG Prover Data Entry'!$F$19
k_1 | ='LPG Prover Data Entry'!$J$35
k_2 | ='LPG Prover Data Entry'!$J$36
k_3 | ='LPG Prover Data Entry'!$J$37
k_4 | ='LPG Prover Data Entry'!$J$38
Material | ='LPG Prover Data Entry'!$C$13
Material.list | ='Tables & Lists'!$D$5:$D$7
MFG | ='LPG Prover Data Entry'!$C$11
mmHg | ='LPG Prover Data Entry'!$J$16
Nom_Val | ='LPG Prover Data Entry'!$J$11
NomVal.list | ='Tables & Lists'!$G$17:$G$18
NomValUnit | ='LPG Prover Data Entry'!$G$11
NVLAP | ='LPG Prover Data Entry'!$I$18
PO.No | ='LPG Prover Data Entry'!$H$5
POC.Name | ='LPG Prover Data Entry'!$H$6
POC.Phone | ='LPG Prover Data Entry'!$H$7
PooledSD | ='LPG Prover Data Entry'!$I$17
Print_Titles | =Information!$6:$7
ProverTempCorr | ='LPG Prover Data Entry'!$E$43
Range.Unit | ='LPG Prover Data Entry'!$A$17
RangeMean.list | ='Tables & Lists'!$G$9:$G$12
RefT | ='LPG Prover Data Entry'!$J$15
RefTemp.list | ='Tables & Lists'!$G$20:$G$21
RefTempUnit | ='LPG Prover Data Entry'!$A$15
ReportedUnc | ='Uncertainty Analysis'!$F$91
Rmean | ='LPG Prover Data Entry'!$I$17
Rnd.Factor | =Calculations!$E$61
Rnd.Table | ='Tables & Lists'!$A$4:$B$32
RptNo | ='LPG Prover Data Entry'!$H$1
Scale.list | ='Tables & Lists'!$G$4:$G$7
Scale.Unit | ='LPG Prover Data Entry'!$G$12
SealNoBottom | ='LPG Prover Data Entry'!$I$96
SealNoTop | ='LPG Prover Data Entry'!$D$96
SN | ='LPG Prover Data Entry'!$C$12
Standards.Table | ='LPG Prover Data Entry'!$A$34:$J$38
StdTempCorr | ='LPG Prover Data Entry'!$E$42
t_1 | ='LPG Prover Data Entry'!$E$73
t_2 | ='LPG Prover Data Entry'!$J$73
Tech | ='LPG Prover Data Entry'!$F$20
Tolerance | =Calculations!$E$63
U60.1 | =Calculations!$E$21
U60.2 | =Calculations!$J$21
Unc | =Calculations!$E$66
V_1 | ='LPG Prover Data Entry'!$E$35
V_2 | ='LPG Prover Data Entry'!$E$36
V_3 | ='LPG Prover Data Entry'!$E$37
V_4 | ='LPG Prover Data Entry'!$E$38
WODate | ='LPG Prover Data Entry'!$I$10
Z60.1 | =Calculations!$E$27
Z60.2 | =Calculations!$E$28

## Report Pages 1-3
STATE OF WASHINGTON
DEPARTMENT OF AGRICULTURE
Weights and Measures Program Metrology Laboratory
2747 29th Avenue Southwest • Tumwater, Washington 98512-6104
Ph (360) 753-5043 • Fax (360) 586-4728 • e-mail dwright@agr.wa.gov
REPORT OF CALIBRATION
Issued To:
Point of Contact:
Purchase Order Number:
Report Number:
Calibration Date
Calibration Due Date
This is to certify that the information contained in this report is true and correct as of the date of calibration.
Dan Wright, State Metrologist |  |  |  |  | Date of Issue
Accreditation LOGO |  | Accredited by the Accreditation Authority for the specific scope of accreditation under lab code XXXXXX-X. This report may not be used to claim product endorsement by Accreditation Authority or any other government agency, and may not be reproduced, except in full, without written approval from the laboratory.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Artifact(s) Description
Test Item: | gal LPG Prover |  |  | Date Received:
Serial Number: |  |  |  | Manufacture:
Material: |  |  |  | Material CCE:
Condition: |  |  |  | Specification: | NIST HB 105-4
Calibration Information
Job Order #: |  |  |  | Temperature:
Metrologist: |  |  |  | Humidity:
Procedure: | NISTIR 7383, SOP 21 |  |  | Water Temperature:
Laboratory Reference Standards Used
Description |  | Serial Number | Cert. Number |  | Cal Date | Cal Due
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
Traceability Statement
The artifact(s) described in this report have been compared to the Standards of the State of Washington. The Standards of the State of Washington are traceable to the National Institute of Standards and Technology (NIST) and are part of a comprehensive measurement assurance program for ensuring continued accuracy and measurement traceability within the level of uncertainty reported by this laboratory. The report number for this report is the only unique report number to be used in referencing measurement traceability for the artifact(s) described in this report.
Uncertainty Statement
The combined standard uncertainty includes uncertainties reported for the standard, uncertainties associated with the measurement process, uncertainties for any observed deviations from reference values which are less than surveillance limits, and other uncertainties associated with the particular artifact (i.e., material cubical coefficient of expansion, reading meniscus, etc.). The combined standard uncertainty is multiplied by k, a coverage factor of 2, to give the expanded uncertainty (which defines an interval with an approximate 95 percent level of confidence). The expanded uncertainty presented in this report is consistent with NIST Technical Note 1297.
Certification Statement
Accredited by the Accredition Authority for the specified scope of accreditation under lab code XXXXXX-X. This laboratory meets the requirements of ISO/IEC 17025 and ANSI/NCSL Z540-1.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Pertinent Information
● In-accordance-with ISO/IEC FDIS 17025, General Requirements for the Competence of Testing and Calibration Laboratories, paragraph 5.10.4.4 ‘A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the client. This requirement may be superseded by legal regulations.’
● In-accordance-with Washington Administrative Code (WAC) Chapter 16-663, Service Agents -- Reporting, Test Procedures, Standards And Calibration Of Weighing And Measuring Devices, Section 16-663-130, Adequacy of standards and submission of standards for certification, paragraph 2, ‘… All standards used for servicing, repairing and/or calibrating commercial weighing and measuring devices must be submitted at least every two years for examination and certification…’
● The artifact is considered in-tolerance when the error is equal to or less than the specified tolerance minus the measurement uncertainty. Bold Italic print indicates an out-of-tolerance reading.
● The LPG Prover 'As Left' volume is 'In-Tolerance', the prover may be used in meter testing without a correction.
● Enter the Pressure Correction from Table 1 that corresponds with the pressure being tested on your LPG Meter Test form.
● The calibration item was calibrated in a "wet down" condition using water. The calibration data above applies when the prover is drained for a 30 (± 5) second period after cessation of the main flow.
● The drain time (using ) to the bottom zero was approximately minute(s) seconds.
● The Top Securty Seal Number is and the Bottom Security Seal Number is .
● The results listed in this report relate only to the artifacts described and extent of calibrations performed.

## Report Pages 4-5
WSDA Weights and Measures Metrology Laboratory
Report Number: |  |  |  | Calibration Date
Calibration Results
Nominal Volume (at zero mark on gauge) | Prover Volume As Found @ ºF and 100 psig (gal) | Prover Volume As Left @ ºF and 100 psig (gal) | NIST HB 105-4 Specification Tolerance ± (gal) | Uncertainty k=2 ± (gal)
Table 1 - LPG Prover Corrections @ ºF
psig | Prover Scale Reading (gal) | Pressure Correction (Pcorr) (in³)¹ | Prover Error (gal) | Prover Volume (gal)
¹Gauge scale was adjusted for nominal volume at 100 psig.
²Gauge scale could not be adjusted or did not need to be adjusted.
Chart 1 - LPG Pressure Corrections (gal) @ ºF

## Control Data
Tech Initials | Date | Test Number | Volume (Trial 1) @ ºF (gal) | Volume (Trial 2) @ ºF (gal) | Start Humidity (% RH) | Finish Humidity (% RH) | Start Air Temp. (ºC) | Finish Air Temp. (ºC) | Water Temp. (Trial 1) (ºC) | Water Temp. (Trial 2) (ºC)
Note on B1: Copy and paste values in control chart.

## LPG Prover Data Entry
Prover Data Entry | [5 empty] | Report Number:
Customer Information
Customer ID
PO#
POC
Phone
Artifact and Measurement Information
Description |  | LPG Prover |  |  |  | Date Received
Manufacture | [5 empty] | Nominal Volume (gal)
Serial Number | [5 empty] | Scale Graduations (gal)
Prover Material | [5 empty] | CCE (β) (/ ºF) =
Artifact Condition
Designated Reference Temperature For This Calibration (ºF)
Normal Barometric Pressure for Location of Calibration (mm Hg)
Process Standard Deviation of the Trials from Control Chart (gal)
Accredited (e.g. NVLAP, A2LA, etc.) Best Uncertainty for this Calibration (gal)
Was a Calibration Interval Requested or is one Required? | [5 empty] | What is the Interval (months)?
Date of Calibration |  |  |  | Technician |  |  |  | Tech Initials
Standards Used
Standard Item ID | Description |  | Serial Number |  | Report Number |  | Calibration Date | Calibration Due
Volume Standard(s) Data
ID Code | Description |  | Drops | Volume (Vs @ trefs) (gal) | Std. Ref. Temp. (trefS ºF) | Scale Graduations (gal) | CCE (α) (/ ºF) | Unc (gal) | k-factor
|  |  |  | 0 | 0 | 0 | 0 | 0 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0 | 1
Temperature Standard(s) Data
Description |  |  | ID Code | Correction (ºC) |  | Accuracy (ºC)
Used For Volume Standand(s)
Used For LPG Prover
Laboratory Test Conditions
Start | Temperature (ºC) |  |  | Humidity (% RH)
Standard(s) Observations
Trial 1 |  |  |  |  | Trial 2
Drop Number | ID Code | Volume (Vs @ tref) (gal) | Water Temp. (ºC) | Delta (Δ) (in³) | Drop Number | ID Code | Volume (Vs @ tref) (gal) | Water Temp. (ºC) | Delta (Δ) (in³)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
11 |  |  |  |  | 11
12 |  |  |  |  | 12
13 |  |  |  |  | 13
14 |  |  |  |  | 14
15 |  |  |  |  | 15
Artifact Observations
Trial 1 |  |  |  |  | Trial 2
Neck Scale Reading @ 0 psig (gal) |  |  |  |  | Neck Scale Reading @ 0 psig (gal)
Water Temperature, Bottom (ºC) |  |  |  |  | Water Temperature, Bottom (ºC)
Water Temperature, Middle (ºC) |  |  |  |  | Water Temperature, Middle (ºC)
Water Temperature, Top (ºC) |  |  |  |  | Water Temperature, Top (ºC)
Water Temperature, Average (ºC) |  |  |  |  | Water Temperature, Average (ºC)
Pressure Measurements
Adjust prover for Nominal Volume at ºF and 100 psig
If bottom gauge scale plate is adjustable, enter inside diameter of the lower neck (in)
Neck Scale Reading As Found @ 100 psig (gal)
Desired Neck Scale Reading at ºF and 100 psig is | [5 empty] | TBD
If the upper and/or lower gauge scale plate is adjustable perform only ONE of the following adjustments:
A. | If the top scale is adjustable, then adjust the upper scale to read the Desired Gauge Reading at ºF and 100 psig.
B. | If the bottom scale is adjustable, adjust the bottom scale: | [6 empty] | N/A
Note: If lower scale was adjusted, print this form, including calculations sheet, drain, rezero, and retest before the Increasing / Decreasing Pressure Test.
i | It is recommended that LPG provers be adjusted, as close as possible, to the nominal volume at 100 psig. If neither scale is adjustable, a prover volume at 100 psig must be calculated and reported.
i | It is best to report the actual volume at 100 psig to ensure that the owner/user applies the correction properly. (A high neck scale reading at 100 psig means the prover is small and the volume reported at ºF and 100 psig should be less than the nominal value of the prover.) The prover error is added to the nominal volume to determine the prover volume (see the Pressure Corrections worksheet).
i | Include a Pressure Correction Table 1 and Chart with your Report of Calibration for the owner/user to use during their testing of LPG Meters (see the Report spreadsheet).
Neck Scale Reading As Left @ 100 psig (gal)
Was upper or lower scale adjusted for nominal volume @ 100 psig during this calibration?
Increasing / Decreasing Pressure Test.
Increasing Pressure |  |  |  |  | Decreasing Pressure
Neck Scale Reading @ 0 psig (gal) |  |  |  |  | Neck Scale Reading @ 150 psig (gal)
Neck Scale Reading @ 50 psig (gal) |  |  |  |  | Neck Scale Reading @ 100 psig (gal)
Neck Scale Reading @ 100 psig (gal) |  |  |  |  | Neck Scale Reading @ 50 psig (gal)
Neck Scale Reading @ 150 psig (gal) |  |  |  |  | Neck Scale Reading @ 0 psig (gal)
Neck Scale Reading @ 200 psig (gal)
Top Scale Security Seal # |  |  |  |  | Bottom Scale Security Seal #
Laboratory Test Conditions
Finish | Temperature (ºC) |  |  | Humidity (% RH)
Prover Drain Time/Method
The drain time to bottom zero was: |  |  |  |  | minute(s) |  | seconds(s)
The prover was drained using:

## Calculations
Calculations
Measurement Trial Calculations
Volume Standard(s) Trial 1 |  |  |  |  | Volume Standard(s) Trial 2
Drop # | Std. Ref. Temp. (trefS ºF) | Water Temp. (ºF) | Water Density (g/cm³) | Delivered Volume (gal) | Drop # | Std. Ref. Temp. (trefS ºF) | Water Temp. (ºF) | Water Density (g/cm³) | Delivered Volume (gal)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
11 |  |  |  |  | 11
12 |  |  |  |  | 12
13 |  |  |  |  | 13
14 |  |  |  |  | 14
15 |  |  |  |  | 15
Trial 1 Delivered Volume (gal) = |  |  |  |  | Trial 2 Delivered Volume (gal) =
LPG Prover Trial 1 |  |  |  |  | LPG Prover Trial 2
Water Density (g/cm³) = |  |  |  |  | Water Density (g/cm³) =
Water Temperature (ºF) = |  |  |  |  | Water Temperature (ºF) =
LPG Prover Volume Determinations At The Designated Reference Temperature
Trial 1 VXtrefX =
Trial 2 VXtrefX =
Trial 1 LPG Prover Volume =
Trial 2 LPG Prover Volume =
LPG Volume Determinations At The Designated Reference Temperature And 100 psig
Prover Volume As Found =
Prover Volume As Left =
Uncertainty Evaluation
NIST HB 105-4 Criteria: The expanded uncertainty for calibration must be less than 0.2 % of the measured volume.
0.2 % of Measured Volume =
Expanded Uncertainty =
Tolerance Evaluation
NIST HB 105-4 Criteria: The maximum permissible errors shall be ± 0.2 % of the nominal capacity.
Tolerance =
Volume Error As Found =
Volume Error As Left =
Measurement Evaluation
NIST HB 105-4 Criteria: Must be capable of repeating within 0.02 % of the test volume.
0.02 % of measure Volume =
Measurement Range =
Measurement Control Calculation
Plot on Control Chart
Std. Dev. of the Trials =
Increasing / Decreasing Presure Test
Criteria: Must agree within 0.02 % of prover nominal volume.
Results Formatted For Calibration Report
Rounding factor =
LPG Prover Nominal Value =
NIST HB 105-4 Tolerance =
LPG Prover Volume As Found =
LPG Prover Volume As Left =
Uncertainty (k=2) =
Note on A57: Metric Prover Pressure Test Formula
Note on J57: Customary Prover Pressure Test Formula

## Pressure Corrections
Pressure Corrections
Pressure Correction Formulas
Prover Error at Reference Temperature and Selected Pressure
where
| VXtrefX = | Prover volume at reference temperature and 0 psig
| VNominal = | Nominal volume of prover
| Wcf = | Water compressibility correction factor (See SOP 21, Appendix A)
| Sr = | Scale reading
| Npsig = | Selected applied pressure
Pressure Correction
where
| Pcorr @ Npsig= | Prover correction at selected psig
| Wcv = | Water compressibility value per gallon 100 psig
Pressure Correction Calculations
psig | Prover Scale Reading (gal) | Pressure Correction (Pcorr) (in³)¹ | Prover Error (gal) | Prover Volume (gal)
0
10
20
30
40
50
60
70
80
90
100
110
120
130
140
150
160
170
180
190
200
* Gauge scale was adjusted to nominal volume at 100 psig.
** Gauge scale could not be adjusted or did not need to be adjusted.

## Water Compressability
Compressability of Water
Compressability of Water Equations and Calculations
The Thermal Compressability of Water (kT) equation is from the article by Frank E. Jones and Georgia L. Harris, "ITS-90 Density of Water Formulation for Volumetric Standards Calibration", as published in the Journal of Research of the National Institution of Standards and Technology, Vol. 97, No. 3, 1992.
kT (kPa)-1 = | 50.831 01x10-8 - 3.682 93x10-9t + 7.263 725x10-11t2 - 6.597 702x10-13t3 + 2.877 67x10-15t4
kT (kPa)-1 =
where
t = temperature of water
Thermal Compressibility of Water Based on Calibration Location Atmompheric Pressure
kT (atm) =
Note: This worksheet takes into account the typical atmospheric pressure at the location of the calibration.
Water Compressibility Value
Wcv = ( Vnom ) ( kT ) ( P )
Water Compressibility Factor
Wcf = 1 + ( kT ) ( P )
where
Vnom = | Nominal volume of the prover in gallons or liters
kT = | Thermal compressability of water in atmospheres (atm)
P = | Pressure in atm
Conversion Factors ( from NIST SP 811)
To Convert From |  | To | Multiply By
atmosphere (standard) |  | pascal (Pa) (exactly) | 101325
pounds per square inch (psi) |  | pascal Pa | 6894.757
millimeter of mercury (0 ºC) (mmHg) |  | pascal Pa | 133.322
Calibration Data
Nominal Volume (gal) =
Designated Reference Temperature (ºC) =
Typical Barometric Pressure at Your Location (mmHg) =
or pascal (Pa) =
Compressability of Water Verus Pressure Calculations
Compressability Value per gallon at mmHg =
Pressure (psig) | Pressure (atm) | Wcf | Wcv (in³) | Wcv (gal)
0 | 0
10 | 0.68046
20 | 1.3609
30 | 2.0414
40 | 2.7218
50 | 3.4023
60 | 4.0828
70 | 4.7632
80 | 5.4437
90 | 6.1241
100 | 6.8046
110 | 7.4851
120 | 8.1655
130 | 8.846
140 | 9.5264
150 | 10.2069
160 | 10.8874
170 | 11.5678
180 | 12.2483
190 | 12.9287
200 | 13.6092

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NIST HB 145, SOP 21, Calibration of LPG Provers
Specify Equation(s)
Volume at the Reference Temperature
Prover Volume = VXtrefX - gauge reading
where
| VXtrefX = | Volume of the unknown vessel, VX at its designated reference temperature, trefX
| VStrefS = | Volume of the standard vessel, VS at its designated reference temperature, trefS
| ρ1, ρ2, …, ρN = | Density of the water in the standard where ρ1 is the density of the water for the first delivery, ρ2 is the density of the water for the second delivery, and so on until all N deliveries are completed
| Δ1, Δ2,…,ΔN = | Volume differences between the water level and the reference mark on the standard where the subscripts 1, 2,…,N, represent each delivery as above. If the water level is below the reference line, Δ is negative. If the water level is above the reference lin
| t1, t2, …, tN = | Temperature of water for each delivery with the subscripts as above
| α = | Coefficient of cubical expansion for the standard in its designated units
| β = | Coefficient of cubical expansion for the unknown vessel in its designated units
| tx = | Temperature of water in the filled unknown vessel in designated units
| ρx = | Density of water in the prover in g/cm3
Identify Components Their Uncertainties
Data Values (based on trial 1) | [5 empty] | Uncertainties
| gal | (VS) Delivered Volume of the standard(s) |  |  |  |  | gal | (uS) Uncertainty of standard(s), VS, from calibration report, normal distribution, k =1
| gal | (Δ) Volume differences between the water level and the reference mark on the standard |  |  |  |  | gal | (usp) Uncertainty associated with the measurement process an R-Chart, normal distribution
| /ºF | (α) Cubic Coefficient of Expansion for the stand material |  |  |  |  | /ºF | (uα) NBS Report 10 081, Uncertainties in the calibration of large vessels, p. 10 (0.000009 /ºC or 0.000005 /ºF), uniform distribution
| /ºF | (β) Coefficient of cubical expansion for the unknown vessel material |  |  |  |  | /ºF | (uβ) NBS Report 10 081, Uncertainties in the calibration of large vessels, p. 10 (0.000009 /ºC or 0.000005 /ºF), uniform distribution
| ºF | (tS) Temperature of water in the standard |  |  |  |  | ºF | (utS) Accuracy for thermometer used for measuring water temperature (NIST IR 7383, SOP 19, accurate to ± 0.1 ºC), uniform distribution
| ºF | (tX) Temperature of water in the filled unknown vessel |  |  |  |  | ºF | (utX) Accuracy for thermometer used for measuring water temperature (NIST IR 7383, SOP 19, accurate to ± 0.1 ºC), uniform distribution
| g/cm³ | (ρwS) Density of water in the standard in g/cm3 |  |  |  |  | g/cm³ | (uρwS) Combined standard uncertainty of the Patterson Morris equation, [J. Res., NIST 97, 335 (1992), 10ppm (0.0000 10 g/cm³)], and possible uncorrected systematic error, k = 1
| g/cm³ | (ρwX) Density of water in the unknown vessel in g/cm3 |  |  |  |  | g/cm³ | (uρwX) Combined standard uncertainty of the Patterson Morris equation, [J. Res., NIST 97, 335 (1992), 10ppm (0.0000 10 g/cm³)], and possible uncorrected systematic error, k = 1
| gal | (Sr) Unknown vessel scale reading at 100 psig |  |  |  |  | gal | (uSr) Uncertainty associated with reading the meniscus of the standard(s), for each transfer, and unknown vessel, 1/4 of smallest division of neck scale, NIST SP 250-72, March 24, 2006, uniform distribution
| ºF | (tref) Reference Temperature for this calibration |  |  |  |  | ºF | Treated as constant, no uncertainty
Input Table (Convert Component Uncertainties For Use In Kragten Spreadsheet)
Input description |  | Variable name, symbol | Value | units | Reported uncer. | Type | Distribution | Factor to normalize | Standard uncer., ui | Rel. ui (%) | df
Volume Standards |  | VStrefS |  | gal |  | B | uc | 1 | #VALUE! | #VALUE!
delta |  | Δ (sp) |  | gal |  | A | Normal | 1 | #VALUE! | #VALUE!
Standard Cubic Coefficient of Expansion |  | α |  | /ºF |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Unknown Vessel Cubic Coefficient of Expansion |  | β |  | /ºF |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Standard Water Temperature |  | tS |  | ºF |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Unknown Vessel Water Temperature |  | tX |  | ºF |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Standard Water Density |  | ρwS |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Unknown Vessel Water Density |  | ρwX |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Unknown vessel scale reading @ 100 psig |  | Sr |  | gal |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Reference Temperature |  | trefX |  | ºF |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Kragten Spreadsheet (Calculate Each Component Uncertainty Impact, Combine and Expand)
|  | u(VStrefS) | u(Δ (sp)) | u(α) | u(β) | u(tS) | u(tX) | u(ρwS) | u(ρwX) | u(Sr) | u(trefX)
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
df
Measurement Equation Inputs | Value
VStrefS |  | #VALUE!
Δ (sp) |  |  | #VALUE!
α |  |  |  | #VALUE!
β |  |  |  |  | #VALUE!
tS | [5 empty] | #VALUE!
tX | [6 empty] | #VALUE!
ρwS | [7 empty] | #VALUE!
ρwX | [8 empty] | #VALUE!
Sr | [9 empty] | #VALUE!
trefX | [10 empty] | #VALUE!
Result | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ciui
uc | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | (ciui)2
df | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | <--rel (ciui)2 | #VALUE!
k | 2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ci | ∑ (ciui)2
U | #VALUE! | gal | #VALUE! | (Urelative, %)
Weight of Each Component Uncertainty
Determine Expanded Uncertainty For This Calibration
Expand uncertainty for this calibration (k=2) = | [5 empty] | gal
Accredited best uncertainty (e.g. NVLAP, A2LA, etc.) (k=2) = |  |  |  |  | 0 | gal
Uncertainty to be reported for this calibration (k=2) = | [5 empty] | gal

## Tables & Lists
Tables & Lists
Rounding Table (Rnd.Table) |  | Cubical Coefficient of Expansion Table (CCE.Table) & Material List (Material.List) |  | Pick Lists for Data Validation Cells
0 | TableError | Material | cm³/ºF | Scale Graduations (in³)
1e-12 | 13 | Steel, Pressure Vessel, Low Carbon | 1.6e-05 | Scale Graduations (gal)
1e-11 | 12 | Steel, Prover, Low Carbon | 1.86e-05 | Scale Graduations (L)
1e-10 | 11 | Steel, Stainless | 2.65e-05 | Scale Graduations (mL)
1e-09 | 10
1e-08 | 9 |  |  | Range Mean from CC (in³)
1e-07 | 8 |  |  | Range Mean from CC (gal)
1e-06 | 7 |  |  | Range Mean from CC (L)
1e-05 | 6 |  |  | Range Mean from CC (mL)
0.0001 | 5
0.001 | 4 |  |  | Yes
0.01 | 3 |  |  | No
0.1 | 2
1 | 1 |  |  | Nominal Volume (gal)
10 | 0 |  |  | Nominal Volume (L)
100 | -1
1000 | -2 |  |  | Designated Reference Temperature For This Calibration (ºF)
10000 | -3 |  |  | Designated Reference Temperature For This Calibration (ºC)
100000 | -4
1000000 | -5
10000000 | -6
100000000 | -7
1000000000 | -8
10000000000 | -9
100000000000 | -10
1000000000000 | -11
10000000000000 | -12
100000000000000 | -13
1000000000000000 | TableError