## Disclaimer & Cover
Disclaimer | You must validate all spreadsheets for use in your laboratory!!!
2013-01-15 | These spreadsheets are provided as a useful approach for metrology calculations according to the SOP represented here. However, NIST OWM or any other organization involved in their development, collectively and individually, do not warrant this spreadsheets for any specific purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each user agrees to decide if, when and how to use the spreadsheets, does so at his or her sole risk, and is responsible to suitable and applicable verification prior to use. When using the tools provided on the NIST OWM website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST OWM, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST OWM, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST OWM website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## Revisions
Acceptable Units
Modifications |  |  | in³ | cubic inch
Date | Initials | Modification | pt | pint
2013-01-16 | GH | Development and modification after limited input. | qt | quart
|  |  | gal | gallon
2014-01-10 | VM | Inserted Disclaimer & Cover, Revisons, Software V&V tabs to workbook and distributed information from the original Instructions tab to those tabs | mL | milliliter
|  | Inserted line below "Ave Process Std Dev:" | L | liter
|  | Added "Ref. Standard ID:" and data entry box to cells B9:C9 and "Chart maintained by:" and data entry box to cells I9:k9. | cm³ | cubic centimeter
|  | Inserted three columns between "Lower Control Limit" and Reference Value" columns. Entered headings for Upper Tolerance(X̅+x %*tol), Lower Tolerance(X̅-x %*tol) and Upper Tolerance Limit(± x %*tol) and equations to calculate the upper and lower tolerance limits, and tolerance to be displayed as a series on the graphs. | m³ | cubic meter
|  | Inserted "Bias (X̅ - ref):" in cell K9 and calculation for bias in L9. | fl dr | fluid drams
|  | Moved equation in N1 to Q1 and edited the equation to show only the Chart Title, Laboratory, Check ID, SOP and Nominal value for title block of Vol Control Chart. | fl oz | Fluid ounce(s)
|  | Edited the equation in Vol Data:Q2 to display "Balance ID: , Accepted Ref Vol: ± U, Mean:, Bias (X̅ - ref):, sp/st:, df & Ref. source:" | ft³/min | cubic feet per minute
|  | Edited the equation in Vol Data:Q3 to display "Reference Standard(s) ID:, Dates:, Chart maintained by:" | in³/min | cubic inches per minute
Added text boxes to Volume Control Chart to read and display the content of Q2 and Q3.
|  | Edited equation in AI1 to create content for Title box on Vol Sp chart showing "Standard Deviation Chart, Flask ID:, SOP: & Nominal Vol:" | End of list
Added equation to AI2 to create content for text box on Vol Sp Chart showing "Balance ID:, Accepted Ref Vol: ± U, Mean:, Bias (X̅ - ref):, Avg Process Std Dev:, df, & Ref. source:"
Added equation to AI3 to create content for text box on Vol Sp chart showing "Reference Standard(s) ID: , Dates:, & Chart maintained by:"
Edited equation in AK4 to create content for Title box on Vol R bar chart showing "Standard Deviation Chart, Flask ID:, SOP: & Nominal Vol:"
Added equation to Ak5 to create content for text box on Vol R bar Chart showing "Balance ID:, Accepted Ref Vol: ± U, Mean:, Bias (X̅ - ref):, Avg Process Std Dev:, df, & Ref. source:"
Added equation to Ak6 to create content for text box on Vol R bar chart showing "Reference Standard(s) ID: , Dates:, & Chart maintained by:"
Added text boxes to Vol Sp Chart to read and display the content of AI2 and AI3.
Added text boxes to Vol R bar chart to read and display the content of AK 5 and AK 6.
Edited text box on Vol Sp Chart to read and display contents of AI1
Edited text box on Vol R bar Chart to read and display contents of Ak4.
2018-06-26 | GH | Added En, k value, comments, clarification of cells. Added Temperature and Humidity columns for evaluating correlation of environmental conditions. Deleted RANGE charts.
2018-11-30 | GH | Updated estimates of standard deviation to represent the "pooled standard deviation" according to the SOPs and NISTIR 6969 Section 8 instead of average standard deviation. Chart limits set back to Auto calculate. Updated chart legends and titles.

## Sofware V&V
Appendix A - Checklist for Creation of Control Charts |  | Technical Assessment
This form is completed for each software file used in the laboratory. The form may be added as a tab/worsheet to a file/workbook or may be maintained in a separate electronic/paper file/location in the laboratory.
Control Charts cover entire scope (are available for each measurement parameter).
Control charts have titles (or otherwise include)
| Laboratory name or other identifying information | Codes | Assessment | Pass/Fail | Evidence
| SOP(s) | A. Software Inspection | Spreadsheet is clear and makes sense | N/A
| Balance/Other equipment |  | There are instructions for use
| Standard ID, Check standard ID |  | Instructions and data input appear on the visible portion of the first worksheet
| Nominal load/value |  | Data-entry fields are labeled and color coded (it is recommended to avoid red and green)
Dates/chronology/time periods if not used on the x axis
| Legends if multiple series and extra information are plotted; to avoid confusion |  | The Standard Operating Procedure (SOP) used is clearly specified
The number of digits to be rounded to is specified
Control charts have x and y axis with labels and |  |  | The user is warned/notified whenever data-entry fields are left blank
| All measurement values have units of measure associated with them |  | Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations
Control charts have |  |  | The software opens at the right location within the file
| Mean value (and units) |  | Unused fields/cells are locked
| Standard deviation (and units) |  | Rows/columns that the operator need not see are hidden
| Degrees of freedom or number of points noted (if not obvious or if small number) |  | Unused sheets are removed
| Alternative summaries of this information and suitable references (e.g., tables) |  | Worksheets are named appropriately
| Printed summary reports of control data with the charts | B. Mathematical Specification | The correct SOP is used
| Reference values and source and bias if appropriate/available |  | The formulae and methods chosen from that SOP are specified
Sources and references for formulae are specified
Control charts have limits that are based on |  |  | The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty
Statistical controls of
| Warning limits (i.e., two standard deviations) and | C. Code review | The formulae in the fields exactly match the SOP
| Action/control limits (i.e., three standard deviations) |  | Repeated calculations appropriately reference the correct cells
| Or specification limits (e.g., tolerances or smaller ratios of tolerances) |  | Calculations, when tested using standard data or reference test data, show appropriate accuracy
Rounding is done at the appropriate locations in the file
Good Items (on chart or in spreadsheet or database table summaries). Control charts have (when applicable and meaningful if not otherwise noted, e.g., in a table) |  | D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences
Tolerances: when applicable
| Uncertainties: for reference value, check standard, and the process output |  | Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa
Equipment information: device readability, configuration (stability settings/timing)
| Standard information: calibration date and interval information |  | “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used
Responsible staff: need on chart or in database
| Status of control: in control, out of control with latest date of review | E. Component Testing | Each macro used is functional
| History: previous limits and history of the chart/data with F-test and/or t-test results |  | Each command/button is functional
Combinations of interdependent macros are functional
SOP 9 | Page 7 of 7 |  | Plotted graphs are accurate
Worksheets/reports print properly, if needed to
Conditional (color and non-color) formatting is functional
|  | F. Numerical Reference | Look-up tables and lists match the latest calibration report.
Uncertainties match the latest Scope
Values that reference another workbook or spreadsheet are dated
When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user
|  | G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct
The evaluation of the embedded data is dated and documented
|  | H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented
|  | I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty
|  | J. Security | Equation and calculation cells are protected against inadvertent editing
Cells are locked in place; they cannot be moved/dragged
Confidentiality of passwords is appropriate
Files are backed up automatically
Additional back-up is available at alternate facilities
Files on network drives cannot be accidentally deleted
Cells P69 through P72 are necessary for the drop-down menu in cells J8 through J60
Pass
Fail
N/A

## Instructions
File Purpose: Control Chart, Standard Deviation Chart, Range Chart
Instructions
This file was created as a template for volume calibrations, but assessment is required before use.
Enter all values in the light yellow colored cells.
Incorrect calculations may result if ALL cells are not completely filled in.
It is recommended that all header information cells be filled in then those cells locked to prevent accidental alteration.
It is also recommended that as cells are filled with measurement data they be locked to prevent alteration as well.
New values will be entered at the bottom of the list so that there is a chronology to values recorded to the control chart.
WARNING: If you copy and paste to another spreadsheet, be sure to double check all reference cells, data validation for the LISTS and LOOKUP tables!!!
Password is "metrology".