Specifications, Tolerances,

and Other Technical

Requirements for

Weighing and Measuring Devices

as adopted by the 98th National Conference on Weights and Measures 2013

Editors:

Tina Butcher

Linda Crown

Rick Harshman

Juana Williams

Carol Hockert, Chief

Office of Weights and Measures

Carol Hockert, Chief

Office of Weights and Measures

Physical Measurement Laboratory

# U. S. Department of Commerce

Penny Pritzker, Secretary

### National Institute of

### Standards and Technology

Dr. Patrick D. Gallagher, Under Secretary of Commerce

for Standards and Technology and Director

| NIST Handbook | 44 |
| --- | --- |

2014 Edition

Supersedes NIST Handbook 44, 2013 Edition

Certain commercial entities, equipment, or materials may be identified in this document in order to describe an experimental procedure or concept adequately.  Such identification is not intended to imply recommendation or endorsement by the National Institute of Standards and Technology, nor is it intended to imply that the entities, materials, or equipment are necessarily the best available for the purpose.

National Institute of Standards and Technology Handbook 44, 2014 Edition

Natl. Inst. Stand. Technol. Handb. 44, 2014 Ed., 484 pages (Oct. 2013)

CODEN:  NIHAE2