Table of Contents

Section 2.23.	Weights	2-85

A.	Application	2-85

A.1.	General.	2-85

A.2.	Exceptions.	2-85

A.3.	Additional Code Requirements.	2-85

S.	Specifications	2-85

S.1.	Material.	2-85

S.2.	Design.	2-85

S.2.1.	Surface.	2-85

S.2.2.	Ring.	2-85

S.3.	Adjusting Material.	2-85

S.4.	Marking Requirements.	2-85

S.4.1.	General.	2-85

S.4.2.	Apothecaries’ Weights.	2-85

S.4.3.	Troy Weights.	2-85

S.4.4.	Metric Weights.	2-85

S.4.5.	Carat Weights.	2-85

S.4.6.	Counterpoise Weight.	2-85

N.	Notes	2-86

N.1.	Testing Procedures.	2-86

T.	Tolerances	2-86

T.1.	In Excess and in Deficiency.	2-86

T.2.	On Avoirdupois Weights.	2-86

T.3.	On Metric Weights.	2-88

T.4.	On Carat Weights.	2-88

T.5.	On Apothecaries and Troy Weights.	2-88

THIS PAGE INTENTIONALLY LEFT BLANK

# Section 2.23.	Weights

## A.	Application

A.1.	General. – This code applies to commercial weights; that is, weights used in connection with commercial weighing devices.

A.2.	Exceptions. – This code does not apply to test weights or to other “standards” of mass.

A.3.	Additional Code Requirements. – In addition to the requirements of this code, Weights shall meet the requirements of Section 1.10. General Code.

## S.	Specifications

S.1.	Material. – The material used for weights shall be as follows:

(a)	Weights of 6 g or 100 gr and larger shall be made of a metal, or a metal alloy, not softer than brass.

(b)	Weights of less than 6 g or 100 gr may be made of aluminum, but shall not be made of iron or of unplated steel, except stainless steel.

### S.2.	Design.

S.2.1.	Surface. – The surface of a weight shall be smooth and shall not be coated with thick, soft, or brittle material.  A weight of more than 2 g or 30 gr or shall not have sharp edges, points, or corners.

S.2.2.	Ring. – A ring on a weight shall not be split or removable.

S.3.	Adjusting Material. – Adjusting material shall be securely positioned and shall not project beyond the surface of the weight.

### S.4.	Marking Requirements.

S.4.1.	General. – A weight shall be marked to show clearly its nominal value, which shall include identification of the unit; however, the nominal value of a weight of 30 gr or 2 g, or less, may be designated by dots, lines, figures, distinctive shape, or other appropriate means.

S.4.2.	Apothecaries’ Weights. – On apothecaries’ dram, ounce, and pound weights, the letters “ap” shall be used in combination with the nominal value and the appropriate abbreviation of or symbol for the unit.

S.4.3.	Troy Weights. – On troy ounce and pound weights, the letter “t” shall be used in combination with the nominal value and the appropriate symbol of the unit.

S.4.4.	Metric Weights. – On metric weights, the symbols “kg,” “g,” and “mg” shall be used in combination with the nominal value of kilograms, grams, and milligrams, respectively.

S.4.5.	Carat Weights. – On carat weights, the letter “c” shall be used in combination with the nominal value.

S.4.6.	Counterpoise Weight. – A counterpoise weight shall be marked to show clearly both its nominal value and the value it represents when used on the multiplying-lever scale for which it is intended.

## N.	Notes

N.1.	Testing Procedures. – Commercial weights should be tested on a precision balance using standard weights, the errors of which, when used without correction, do not exceed 1/3 of the smallest tolerance to be applied.  (Also see Appendix A, Fundamental Considerations, paragraphs 3.2. Tolerance for Standards and 3.3. Accuracy of Standards.)

## T.	Tolerances

T.1.	In Excess and In Deficiency. – The tolerances hereinafter prescribed shall be applied equally to errors in excess and errors in deficiency.

T.2.	On Avoirdupois Weights. – The maintenance tolerances shall be as shown in Table 1. Maintenance Tolerance for Avoirdupois Weights.  Acceptance tolerances shall be one-half the maintenance tolerances.

| Table 1. Maintenance Tolerance for Avoirdupois Weights |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Maintenance Tolerance |  |  |  |  |  |  |  |
| Nominal Value | Equal-Arm Weights |  | Counterpoise Weights |  |  |  |  |
|  |  |  | For scales with multiples of less than 1000 |  | For scales with multiples of 1000 or over |  |  |
| oz | grains | mg | grains | mg | grains |  | mg |
| 1/64 | 0.1 | 6 |  |  |  |  |  |
| 1/32 | 0.3 | 19 |  |  |  |  |  |
| 1/16 | 0.4 | 26 |  |  |  |  |  |
| 1/8 | 0.5 | 32 |  |  |  |  |  |
| ¼ | 1.0 | 65 |  |  |  |  |  |
| ½ | 1.5 | 97 | 1.0 | 65 |  |  |  |
| 1 | 1.7 | 110 | 1.0 | 65 |  |  |  |
| 2 | 2.0 | 130 | 1.0 | 65 |  |  |  |
| 3 | 2.0 | 130 | 1.5 | 97 |  |  |  |
| 4 | 3.0 | 190 | 1.5 | 97 |  | 1.0 | 65 |
| 5 | 3.5 | 230 | 1.5 | 97 |  | 1.0 | 65 |
| 6 | 3.5 | 230 | 1.5 | 97 |  |  |  |
| 8 | 4.0 | 260 | 2.0 | 130 |  | 1.5 | 97 |
| 10 | 4.0 | 260 | 2.5 | 160 |  | 2.0 | 130 |
| 12 | 5.0 | 320 | 2.5 | 160 |  | 2.0 | 130 |
| Lb | Grains | mg | grains | mg |  | grains | mg |
| 1 | 5.0 | 320 | 3.0 | 190 |  | 2.5 | 160 |
| 2 | 7.0 | 450 | 6.0 | 390 |  | 4.0 | 260 |
| 3 | 9.0 | 580 | 9.0 | 580 |  | 5.0 | 320 |
| 4 | 11.0 | 710 | 11.0 | 710 |  | 6.0 | 390 |
| 5 | 15.0 | 970 | 12.0 | 780 |  | 6.5 | 420 |
| 6 | 17.0 | 1190 |  |  |  |  |  |
| 7 | 19.0 | 1200 |  |  |  |  |  |
| 8 | 21.0 | 1400 | 15.0 | 970 |  | 9.0 | 580 |
| 9 | 23.0 | 1500 |  |  |  |  |  |
| 10 | 25.0 | 1600 | 18.0 | 1160 |  | 10.0 | 650 |
| 15 | 28.0 | 1800 |  |  |  |  |  |
| 20 | 30.0 | 1900 |  |  |  |  |  |
| 25 | 35.0 | 2300 |  |  |  |  |  |
| 30 | 40.0 | 2600 |  |  |  |  |  |
| 40 | 45.0 | 2900 |  |  |  |  |  |
| 50 | 50.0 | 3200 |  |  |  |  |  |

T.3.	On Metric Weights. – The maintenance tolerances shall be as shown in Table 2. Maintenance Tolerances for Metric Weights.  Acceptance tolerances shall be one-half the maintenance tolerances.

T.4.	On Carat Weights. – The maintenance tolerances shall be as shown in Table 2. Maintenance Tolerances for Metric Weights.  Acceptance tolerances shall be one-half the maintenance tolerances.

| Table 2. Maintenance Tolerances for Metric Weights |  |  |  |
| --- | --- | --- | --- |
| Nominal Value (mg) | Maintenance Tolerance (mg) | Nominal Value (g) | Maintenance Tolerance (mg) |
| 5 or less 10 20 30 50 100 200 300 500 | 0.1 0.3 0.4 0.6 0.8 1.0 1.5 2.0 3.0 | 1 2 3 5 10 20 30 50 100 200 300 500 | 4 6 8 10 15 20 30 40 70 100 150 175 |
| Nominal Value (kg) | Maintenance Tolerance (mg) | Nominal Value (carats) | Maintenance Tolerance (mg) |
| 1 2 3 5 10 20 | 250 400 500 800 1000 1500 | 0.25* 0.5** 1.0 2.0 3.0 5.0 10.0 20.0 30.0 50.0 100.0 | 0.6 1.0 1.5 2.0 3.0 4.0 6.0 10.0 12.0 15.0 25.0 |
|  |  | *25 points or less **50 points |  |

T.5.	On Apothecaries and Troy Weights. – The maintenance tolerances shall be as shown in Table 3. Maintenance Tolerances for Apothecaries’ and Troy Weights.  Acceptance tolerances shall be one-half the maintenance tolerances.

| Table 3. Maintenance Tolerances for Apothecaries’ and Troy Weights |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Nominal Value | Maintenance Tolerance |  | Nominal Value | Maintenance Tolerance |  |
| grains | grains | mg | oz | grains | mg |
| 1 2 3 5 10 20 | 0.01 0.02 0.03 0.03 0.04 0.06 | 0.6 1.3 2.0 2.0 2.5 4.0 | 1 2 3 4 5 | 0.4 0.6 1.0 1.5 1.6 | 25.0 40.0 65.0 100.0 105.0 |
| scruples | grains | mg | oz | grains | mg |
| 1 2 | 0.06 0.10 | 4.0 6.5 | 6 7 8 9 10 | 1.8 1.9 2.0 2.1 2.2 | 115.0 125.0 130.0 135.0 145.0 |
| dr | grains | mg | oz | grains | mg |
| 0.5 1.0 2.0 3.0 4.0 5.0 6.0 | 0.07 0.10 0.20 0.30 0.40 0.50 0.60 | 4.5 6.5 13.0 20.0 25.0 30.0 40.0 | 11 12 20 30 50 | 2.4 2.5 2.9 3.7 5.4 | 155.0 160.0 190.0 240.0 350.0 |
| dwt | grains | mg | oz | grains | mg |
| 1 2 3 4 5 10 | 0.06 0.10 0.15 0.20 0.30 0.40 | 4.0 6.5 10.0 13.0 20.0 25.0 | 100 200 300 500 1 000 | 7.7 12.3 15.4 23.1 38.6 | 500.0 800.0 1 000.0 1 500.0 2 500.0 |

THIS PAGE INTENTIONALLY LEFT BLANK