## START HERE Cover Sheet
OSAC Program Office Use Only |  |  | Comment Adjudication_TO SC
This Comment Adjudication cover sheet is to be completed by the SC Chair or Technical Point of Contact during the Comment Adjudication Process.
Document Title: | 2022-S-0040 Standard for Internal Validation of DNA Extraction Methods
OSAC Unit | Human Forensic Biology
Unit Chair
Name: | Beth Ordeman
Affiliation:
Unit Technical Contact
Name: | Susan Greenspoon
Affiliation:
SC or FSSB Quality Task Group (QTG) Review Completion Date
Notes made by the QTG, if applicable:
Full Subcommittee Resolution Vote Outcome.
Google Form URL To Edit and Update with Vote Outcome. Click this link to edit the original form submission.
Categories for adjudication of ALL COMMENTS.
Resolution / Disposition to be used in the Adjudication | Definition | Rationale
Not germane The comment is not relevant to the document or its scope. | The comment is not relevant to the document or its scope. | No additional rationale from the SC
No response needed The comment either supports the standard as written or does not seek any change to it. | The comment either supports the standard as written or does not seek any change to it. | No additional rationale from the SC
No change The OSAC Proposed standard will not be revised based on this comment. | The OSAC Proposed standard will not be revised based on this comment. | SC must state WHY the standard will not be revised.
Revision was made The OSAC Proposed standard was revised based on this comment. | The OSAC Proposed standard was revised based on this comment. | SC must state WHAT revisions were made to address the comment.
Withdrawn The commenter withdrew the comment. | The commenter withdrew the comment. | No additional rationale from the SC
*If multiple comments are similar, the adjudication should be consistent across all comments. OPO will review the adjudication to ensure similarity in the Type of Adjudication (e.g., No change, Revision was made, etc.) and similarity in the justification wording. This check is to ensure that the adjudication documents published to the OSAC Public Document Repository provides viewers a clear understanding of the unit's position.
Key
You can use this language in the Rationale
No Response is Needed
Must provide a Justification for Resolution Rationale

## Comments
|  |  |  | OSAC Proposed Standard Number | 2022-S-0040
|  |  |  | Standard Title | Standard for Internal Validation of DNA Extraction Methods
|  |  |  | Date of Open Comment Submission | 2023-01-03
|  |  |  | Name of Commenter(s) | Various commenters below (see contact info on "Commenter Info" sheet)
|  |  |  | Organization of Commenter (if any) | Various commenters below (see contact info on "Commenter Info" sheet)
|  |  |  | Email Address of Commenter (will not be published) | Various commenters below (see contact info on "Commenter Info" sheet)
TABLE OF COMMENTS | Comment # | Document Line Number | Name of Commenter | Current Language | Suggested Language or Revision | Rationale | Resolution /Disposition | Subcommittee Response/Rationale | Notes (optional field) | QTG Adjudication Review of Public Comments - Summary comments
Public Comment |  | No Public Comments Received
Task Group Comment | 1 | 4.1.2 | Andrea Roth, LTG chair | Discordant data... (define discordant in section 3. Concordant is defined but not in a way that inherently provides a clear and uniform definition of discordant) | none provided | none provided | Revision was made | Verbiage added to concordance definition: Antonym: discordant. |  | moved resolution to response and added correct resolution
Task Group Comment | 2 | 4.2 | Andrea Roth, LTG chair | "known reference or casework-like samples and substrates representative of those typically analyzed by the testing laboratory" - should be and not or; and perhaps clarify that should it should include the range of samples that the lab intends to analyze. Somewhat vague description through out 4.2, 4.2.1, 4.2.2, 4.2.3 regarding the the different samples, but not sure how to clarify- perhaps adding that the range of samples used most reflect all of the factor spaces in which the tool will be used? | none provided | none provided | No change | This standard is written such that a databasing operation can follow it (they will not be testing casework-like samples). Further detail is provided in the BPR. |  | CH-ok
Task Group Comment | 3 | 4.3 | Andrea Roth, LTG chair | how will this be determined? how will the susceptibility be defined? Are there uniform standards/ thresholds of acceptable results or must this conform with ranges/ standards set by the lab? | none provided | none provided | No change | More details are in the BPR guidance. Namely, degradation and inhibition introduced by the extraction process will be assessed by DNA quantitation and STR typing, as described. |  | without her knowing DNA habits. I suggest. As described in 4.2.1, this will be determined by the loss of DNA profiles or DNA template when degraded or inihibited with items commonly seen in casework.
Task Group Comment | 4 | 4.4 | Andrea Roth, LTG chair | more than one study should be required, and criterion for the appropriateness of the study design should be included; is this used to set the analytical threshold for the lab in terms of quant? specify how this determination relates to the lab's validation/ use of the new method | none provided | none provided | No change | This is a separate issue from quantitation, but the quantitation system used/validated is what the lab uses to define the limit of template needed to generate a profile. Analytical threshold is not set using this standard. There are multiple criteria described in the standard for determining suitability of a DNA extraction method. It is also dsicussed in more detail in the BPR document. For clarification, the term, "study", does not refer to a single experiment, but can include multiple experiments. | Q | I suggest. The term, "study", does not refer to a single experiment, but a set of experiments designed to accomplish one purpose. This validation document only discusses validating and determinig sufficiency of an extraction method. It is only one part of the DNA testing methods. Q- it seems you could define what sensitivty means to extraction. Is there a upper or lower bound limit to the recovery of DNA?
Task Group Comment | 5 | 4.5 | Andrea Roth, LTG chair | what about reliability? how should this be demonstrated? this doesn't provide and guidance for uniformity or appropriateness of how this is tested or demonstrated | none provided | none provided | No change | Reliability is demonstrated by fulfilling all of the validation requirements. Guidance is provided in the best practive recommendations accompanying document. |  | Suggest: The document as a whole defines what is needed to determine fit for purpose which includes reliability ofan extraction procedure for DNA
Task Group Comment | 6 | 4.6 | Andrea Roth, LTG chair | same concern as in 4.4 and 4.5 5- including data... and retain/ make available? | none provided | none provided | No change | 4.6, 4.7 and entire document. The BPR provides the detailed recommendations, not the standard document. | Q | I don't understand the requirement to use the BPR to understand how to use the standard. The minimum req and the best practice should be different.
Task Group Comment | 7 | page 4, line 27 | Human Factors Task Group consensus OSAC/HFTG | Accumulation of test data within the laboratory for developing the laboratory standard operating procedures and demonstrating that the established protocols for the technical steps of the test and for data interpretation perform as expected in the laboratory. | Systematic collection of test data within the laboratory for developing the laboratory standard operating procedures and demonstrating that the established protocols for the technical steps of the test and for data interpretation perform as expected in the laboratory. | The word accumulation suggests a random process. | No change | Lexicon: The accumulation of test data within the laboratory for developing the laboratory standard operating procedures and determining the limits of the method(s). Internal validation demonstrates that the established protocols for the technical steps of the test and for data interpretation perform as expected in the laboratory | Definition of internal validation changed to match Lexicon. | Suggest: Revised definition to match ASB 038 and ASB 077 which are published parent documents in this series.
STR Comment | 1 | General Comment | Inman | Quantitation is likely a separate standard, but is inextricably linked to extraction. It seems like it’s a least worth discussing the relationship somehow. This is particularly problematic when talking about reproducibility, comparability, accuracy, and precision. Unless those are measured without using quantitation, one can't easily deconvolve the effect of extraction vs. quantitation, without adding a section on replicate testing during the both the extraction and quantitation steps. |  |  | No change | It is understood that questions of reproducibility, yield, etc. are defined at least in part by the quantitation system employed. We fhave a separate standard/BPR for quantitation. |  | CH-would this make the quantitation documents Normative References?
STR Comment | 2 | Title | John Butler | Standard for Internal Validation of DNA Extraction Methods | Standard for Studies to Perform When Conducting Internal Validation of DNA Extraction Methods | The document as written is limited with what is shared about how to conduct an internal validaiton and appears to be more focused on general studies that should be performed | No change | The standard is written as general standards that must be met for the internal validation of an extraction system. The details are provided in the companion BPR document. |  | Suggest: This standard is written as minimum standards that must be met. Additional recommendations are provided in the BPR.
STR Comment | 3 | 3-4 | TLV | …in conjunction with ASB 038 Standard for Internal Validation of Forensic DNA Analysis. | Should there be mention of the standard for automated platform validation standards? | Not sure if it should be within scope here or within the best practices document, but most often the extraction platforms currently involve some type of liquid handler. | No change | The automation standard is a stand-alone standard and applies to a variety of methodologies (quantitation, DNA extraction, STR setup). This standard applies to both manual and automated DNA extraction. |  | CH-ok
STR Comment | 4 | 20 | Inman | A technique used to release DNA from cells in a biological sample | A technique used to release and purify DNA from biological cells | Not to be really picky, but extraction is more complex than this. Because DNA in cells is wrapped around a protein scaffold, DNA extraction methods as practiced in forensic science laboratories is both the release of DNA from the cells AND removal of the DNA from the protein scaffold, resulting in unprotected naked ds DNA. Perhaps this can be modified to include release and purify DNA from cells, or something similar . | Revision was made | A technique used to release DNA from cells in a biological sample that may or may not be followed by purification. | We have added additional verbiage, but the standard is written to accommodate extraction procedures that simply release the DNA from the cells and disable proteins, such as Chelex, and Casework Direct. | corrected resolution
STR Comment | 5 | 26-29 | John Butler | Internal Validation (OSAC 2022-S-0040 3.6 and 2022-S-0041 3.8): Accumulation of test data within the laboratory for developing the laboratory standard operating procedures and demonstrating that the established protocols for the technical steps of the test and for data interpretation perform as expected in the laboratory. | Internal Validation (ASB 038 3.3): The accumulation of test data within the laboratory for developing the laboratory standard operating procedures and determining the limits of the method(s). Internal validation demonstrates that the established protocols for the technical steps of the test and for data interpretation perform as expected in the laboratory. | definition does not match the normative reference (umbrella document) ASB 038 Standards for Internal Validation of Forensic DNA Analysis Methods | Revision was made | See response to comment #7 by Human Factors. The defintion was changed to match the Lexicon. | Q | don't reference a response, just use the exact same response, ie paste the final resolution of 7 here as well
STR Comment | 6 | 35-58 (General) | John Butler | shall demonstrate, shall be compared, shall be documented, shall evaluate, shall assess, shall conduct, shall compare, shall determine, shall complete and document | more details on how to demonstrate, compare, document, assess, determine, and complete the requests being made | The focus is generally on what should/could be done surround DNA extraction methods yet there is not enough information to be helpful, even with the best practice recommendations. There needs to be more on how to assess, conduct, compare, document, etc. to help meet the FSSB requirements for method validation standards. This concern is a much wider issue than just this one document. |  | We are aware of this concern and will address numbers of samples and the most appropriate statistical assessments, where applicable, to perform either in the standards doc |  | What about more language on considerations for each one? I think discussion, limitations, etc would be very helpful. Particularly for non-DNA experts
STR Comment | 7 | 37 | Inman | DNA typing results obtained from the new DNA extraction chemistry shall be compared to results obtained from a previously validated DNA extraction method or published results. | none provided | The purpose of some extraction methods may be to improve the extraction efficiency (get more DNA out). So you don’t want concordance all of the time; you may want better performan. If by concordance you mean the same types, then this is out of place. | No change. | Concordance is defined as: Agreement between DNA typing results. Theoretically, a DNA extraction system could cause problems with downstream DNA typing results. |  | ok
STR Comment | 8 | 40-41 | TLV | …and a potential explanation for any discordant results shall be documented in the final validation document. | No revision, but inconsistent with best practices document, see comment. | See comment for best practices document. | Revision was made | The following verbiage has been added to the standards doc: Discordant data observed when comparing identical genetic markers using sample(s) from the same individual shall be documented in the final validation document. |  | ok
STR Comment | 9 | 44 | Inman | The laboratory shall assess the performance of the DNA extraction method using different substrates, sample types, and sample conditions (e.g. inhibition, degraded samples, etc.) | none provided | Shouldn’t there be some definition somewhere of what “assess” means? For example, establishing either a minimum acceptable outcome, or defining some other criteria that says “adapt” or “reject” the new method? Without definition, it's difficult to 'assess' adherence to this standard by an auditor | No change | Details for assessment can be found in the companion bpr. | Q | JRB-then this is not a standalone document CH-agree, what defines acceptability after assessment? Minimum requirements should be in this document
STR Comment | 10 | 45 | TLV | ….sample conditions (e.g. inhibition, degraded samples, etc.) | …sample conditions (e.g. inhibited or degraded samples, etc.) | Grammar/tense change suggestions | Revision was made | (e.g. inhibited or degraded samples, etc.) |  | ok
STR Comment | 11 | 51 | John Butler | …to determine the limits of the of the DNA extraction method… | …to determine the limits of the DNA extraction method… | remove duplicate "of the" | Revision was made | Corrected. |  | ok
STR Comment | 12 | 53 | John Butler | precision and accuracy | percent recovery? (will require DNA quantitation not simply consistent genotypes from extracted samples) | It would be helpful to provide more detail on how this goal of measuring precision and accuracy could be accomplished; this section seems to correspond to the BPR 5.5 Reproducibility and Repeatability | Revision was made | Statement now reads: The laboratory shall determine the precision of the DNA extraction method. Precision estimates can be performed using reproducibility/repeatability studies. | Given comments by others, the word accuracy was removed. | red language is no longer in the document
STR Comment | 13 | 53 | A Rasmussen | The laboratory shall demonstrate the precision and accuracy of the DNA extraction method. | Remove "and accuracy" | Accuracy (other than concordance) does not seem applicable to extraction methods. | Revision was made | Statement now reads: The laboratory shall determine the precision of the DNA extraction method. Precision estimates can be performed using reproducibility/repeatability studies. |  | red language is no longer in the document
STR Comment | 14 | 53 | Inman | The laboratory shall demonstrate the precision and accuracy of the DNA extraction method. | "determine" in place of "demonstrate" | I’m not sure this is actually possible; it depends on the quant method, inextricably bound to the extraction method. Further, I think that one determines precion and accuracy, rather than demonstrating it. | Revision was made | Statement now reads: The laboratory shall determine the precision of the DNA extraction method. |  | ok
STR Comment | 15 | 57-58 | John Butler | complete and document all applicable validation studies | include value of publicly sharing this validation data in an open access format | to permit independent assessment by others outside of the laboratory as well as enabling everyone in the laboratory to see the rationale for capabilities and limitations with the methods in use | Not germane | That is beyond the scope of this document and veers into inidividual lab policy and state practices and/or laws. |  | updated resolution
STR Comment | 16 | 61 | John Butler | The Butler Methodology book is listed in the bibliography without any explanation of why or what is being specifically cited | Provide a specific chapter or set of pages that pertain to the requirements listed in this standard and/or include additional appropriate references | It is unclear how the provided reference applies to the material discussed in this document; Chapter 2 describes different types of DNA extraction methods but not validation experiments to perform | Revision was made | reference removed |  | updated resolution and response
STR Comment | 17 | 61 | Inman | Butler, John. Advanced Topics in Forensic DNA Typing: Methodology. Academic Press, 2012. | none provided | The single reference to Butler is, I think, completely inadequate for a standard (it's like saying I used the library to create my work). Specific references should be pulled out. With no offense meant to John. | Revision was made | reference removed |  | updated resolution and response

## Commenter Info
Commenter # | Commenter's Name | DO NOT PUBLISH | Commenter's Position/Title | Commenter's Org/Affiliation | DO NOT PUBLISH | Commenter's Email | Date of Open Comment Submission
1 | Andrea Roth |  | OSAC LTG Chair |  |  |  | 2023-01-03
2 | Pate Skene |  | OSAC HFTG Chair |  |  |  | 2023-01-03
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25