## START HERE Cover Sheet
Comment Adjudication_TO SC
This Comment Adjudication cover sheet is to be completed by the SC Chair or Technical Point of Contact during the Comment Adjudication Process.
Document Title: | 2024-S-0012 Standard Practice for the Forensic Analysis of Geological Materials by Scanning Electron Microscopy and Energy Dispersive X-Ray Spectrometry
OSAC Unit | Trace Materials
Unit Chair
Name: | Sandy Koch
Affiliation: | NIST
Unit Technical Contact
Name: | Libby Stern & Jack Hieptas
Affiliation: | FBI & John Jay
SC or FSSB Quality Task Group (QTG) Review Completion Date | 2024-10-28
Notes made by the QTG, if applicable: | See notes in column L
Full Subcommittee Resolution Vote Outcome. | 25 affirmative, 1 abstention
Google Form URL To Edit and Update with Vote Outcome. Click this link to edit the original form submission.
Categories for adjudication of ALL COMMENTS.
Resolution / Disposition to be used in the Adjudication | Definition | Rationale
Not germane The comment is not relevant to the document or its scope. | The comment is not relevant to the document or its scope. | No additional rationale from the SC
No response needed The comment either supports the standard as written or does not seek any change to it. | The comment either supports the standard as written or does not seek any change to it. | No additional rationale from the SC
No change The OSAC Proposed standard will not be revised based on this comment. | The OSAC Proposed standard will not be revised based on this comment. | SC must state WHY the standard will not be revised.
Revision was made The OSAC Proposed standard was revised based on this comment. | The OSAC Proposed standard was revised based on this comment. | SC must state WHAT revisions were made to address the comment.
Withdrawn The commenter withdrew the comment. | The commenter withdrew the comment. | No additional rationale from the SC
*If multiple comments are similar, the adjudication should be consistent across all comments. OPO will review the adjudication to ensure similarity in the Type of Adjudication (e.g., No change, Revision was made, etc.) and similarity in the justification wording. This check is to ensure that the adjudication documents published to the OSAC Public Document Repository provides viewers a clear understanding of the unit's position.
Key
You can use this language in the Rationale
No Response is Needed
Must provide a Justification for Resolution Rationale

## Comments
|  |  |  | changed to " | OSAC Proposed Standard Number | 2024-S-0012
[5 empty] | Standard Title | Standard Practice for the Forensic Analysis of Geological Materials by Scanning Electron Microscopy and Energy Dispersive X-Ray Spectrometry
[5 empty] | Date of Open Comment Submission | 2024-06-04
[5 empty] | Name of Commenter(s) | Various commenters below (see contact info on "Commenter Info" sheet)
[5 empty] | Organization of Commenter (if any) | Various commenters below (see contact info on "Commenter Info" sheet)
[5 empty] | Email Address of Commenter (will not be published) | Various commenters below (see contact info on "Commenter Info" sheet)
TABLE OF COMMENTS | Comment # | Document Line Number | Should be | Name of Commenter | Current Language | Suggested Language or Revision | Rationale | Resolution /Disposition | Subcommittee Response/Rationale | Notes (optional field) | QTG Adjudication Review of Comments - Summary comments | Post ballot changes
Public Comment | 1 |  |  |  | None Received
Task Group Comment | 1 |  |  |  | None Received
STR Comment | 1 | Overall |  | Chris Palenik | None Provided | None Provided | The overall section organization of the document is fine, but I think that the content of individual sections could be better organizized. | Revision was made The OSAC Proposed standard was revised based on this comment. | There was significant reorganization of the sample preparation section and significance and use section.
STR Comment | 2 | Overall |  | Chris Palenik | None Provided | None Provided | There is, in my opinion, an imbalance in the emphases given to various topics throughout the work: (a) SEM 101 seems overemphasized in certain areas, (b) several highly detailed scenarios that are touched upon, and (d) arguably too little specificity is given in critical areas (like performance checks, sample prep and analytical conditions, and mineral identifications (provisional or otherwise), and (e) too much discussion of trace elements (e.g., ~6.2 area of the document), when trace elements are not as germaine to classificaiton/provisional id/distinction of samples in the way this document is intended to be used. This is not incorrect, but I think further refinement of the document is warranted. | Revision was made The OSAC Proposed standard was revised based on this comment. | This document has been refined. (a) The audiance for this standard includes both expert SEM users less familiar with forensic geology and forensic geologist less familiar with SEM/EDS, so the level of detail is appropriate. (b) Some of the scenarios are specific to forensic geology so they add value. (c) Sample preparation was expanded and reorganized. Detailed performance checks are laboratory and instrument specific. The approach to mineral indentification/classification is explained in 7.2. An exhaustive table of identification criteria would never be complete and is captured in mineralogy texts. (d) This section is devoted to spectral interpretation, specifically the assignment of elements to peaks. The discussion is general and not specific to trace element analysis.
STR Comment | 3 | Overall |  | Chris Palenik |  | None Provided | Scope: The scope of this document is appropriately broad (i.e., geological materials), but the text is fairly specific to only grain mounts and clay. I think some text that explains this in the scope and some text in the various preparation sections that discusses this are appropriate. To that end, the document seems to try to touch on everything though, mentioning water extracts, in situ analysis on clothing, and surface texture. I think the document would be better if it scaled back and focused, with more detail and parameters guidance, on a few main topics: the analysis of grains (polished and unpolished) and powders by SEM EDS). | Revision was made The OSAC Proposed standard was revised based on this comment. | The sample preparation section was significantly structured and expanded
STR Comment | 4 |  |  | Sara Walker | None Provided | There is no requirement for what controls should be run when performing this analysis | The minimum requirements for controls to ensure the veracity of the result being determined is essential | No change The OSAC Proposed standard will not be revised based on this comment. | Samples are so variable that there is no single methodological standard. When appropriate, standard procedures are included throughout, such as calibration of the energy levels by metal standard and options for monitoring airborne contamination.
STR Comment | 5 | 72-95 |  | RV | Scope | None Provided | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | We are unaware of any current conflicts between this document and any cited or other OSAC standards.
STR Comment | 6 | 77 | 86 | NWMR | kilo electron volts | kiloelectronvolts | It really is just one long word. | Revision was made The OSAC Proposed standard was revised based on this comment. | kilo electron volts replaced by kiloelectron volts
STR Comment | 7 | 77 |  | David Stoney | This document describes sample handling .... | This document describes aspects of sample handling .... | Edit is recommended so that the document does not appear to be all-inclusive with respect to the topic areas. | Revision was made The OSAC Proposed standard was revised based on this comment. | The theoretical foundation of SEM/EDS is covered in numerous texts such as Scanning Electron Microscopy and X-ray Microanalysis. This document describes some options for sample handling and preparation, instrument operating conditions, spectral data collection, evaluation of EDS data quality, interpretation of EDS spectra for the identification of inorganic geological materials, documentation of morphology by SEM imaging, and criteria for sample comparison. |  | I'm not seeing this change in the document | Corrected, this change was inadvertantly omitted from the document.
STR Comment | 8 | 81 | 90 | NWMR | "These are provided in the following (…)" | Eliminate | Helpful maybe but not relevant | Revision was made The OSAC Proposed standard was revised based on this comment. | Kept citations but eliminated the text.
STR Comment | 9 | 87 |  | David Stoney | ... and data are reported in elemental weight percent (wt. %) | ... and compositional data based on EDS are reported in elemental weight percent (wt. %) | Edit is recommended to specify the type of data. | Revision was made The OSAC Proposed standard was revised based on this comment. | Change made as suggested.
STR Comment | 10 | 90 |  | Sara Walker | These are provided in the following (C1723, | remove extra space between following and the parentheses | Typographical error | Revision was made The OSAC Proposed standard was revised based on this comment. | removed extra space between "following" and the left parentheses | Editorial
STR Comment | 11 | 109 |  | Chris Palenik | E766 Practice for Calibrating the Magnification of a Scanning Electron Microscope | None Provided | I realize this is potentially relevant, but it's not clear to me that anyone is actually using this, or how this should be specifically imprelmented here. | Revision was made The OSAC Proposed standard was revised based on this comment. | now referenced in 6.1.4.1
STR Comment | 12 | 115 | 123? | NWMR | "bulk analysis" | None Provided | Likely to conflict with other standards use of this term. | No change The OSAC Proposed standard will not be revised based on this comment. | See modification to 6.2.3.4 recorded on comment 92: Section 6.2.3.4 was stregthened and now reads "Overscanning (see 3.1.6) large portions of a specimen when there are multiple phases/compositional domains present should not be used for quantitative analysis but can be used for detecting the elements present." |  | I'm not sure this response goes with this line item? I agree that you don't need to revise based on the comment. That is why you have defined bulk analysis - so that folks won't be confused. | Note, comments from NWMR had incorrect line noted. column D is our best guess of correct line item
STR Comment | 13 | 115 | 123? | NWMR | "analysis of a material either by a large single area" | Eliminate "by a large single area" | In this comment, I'll address the issue of overscanning which is discussed in many places in this document. I won't claim that overscanning doesn't have its uses. As a qualitative technique to say something about a heterogeneous region, overscanning can be used. However, the results are exceedingly difficult to interprete because they depend on the fractional area represented by each different type present in the region. Furthermore, quantitative techniques like matrix correction cannot be applied to the resulting spectra because the assumption of homogeneity is violated. If bulk analysis by larger area scales (or overscans) are to be mentioned in the document, I'd like to see significant caveats describing what can and cannot be done with the spectra. | No change The OSAC Proposed standard will not be revised based on this comment. | See modification to 6.2.3.4 recorded on comment 92: Section 6.2.3.4 was stregthened and now reads "Overscanning (see 3.1.6) large portions of a specimen when there are multiple phases/compositional domains present should not be used for quantitative analysis but can be used for for detecting the elements present." |  | I would change the resolution to revised. They also don't like when you reference another comment line item. Just reference the change and copy what you had in the other comment (which you did).
STR Comment | 14 | 115-116 |  | Sara Walker | ISO/IEC 1702:2017 General requirements for the competence of testing and calibration laboratories. | ISO/IEC 17025:2017 General requirements for the competence of testing and calibration laboratories [remove the period] | The standard number has a typographical error and a title doesn't need a period at the end | Revision was made The OSAC Proposed standard was revised based on this comment. | 1702 changed to 17025, and period removed.
STR Comment | 15 | 117-118 |  | Sara Walker | ISO 22309:2011 Microbeam analysis — Quantitative analysis using energy-dispersive spectrometry (EDS) for elements with an atomic number of 11 (Na) or above. | None Provided | A title doesn't need a period at the end | Revision was made The OSAC Proposed standard was revised based on this comment. | period removed
STR Comment | 16 | 120 |  | Chris Palenik | The term “Provisional Identification” is used throughout this doc. I think it should be defined here. | None Provided | I have a sense of what you mean. This would probably be akin to what our lab would call a tentative ID by GC/MS. Here, provisional ID may mean anything from a grouping based on elements present, to a likely ID if you see Na, Si, Al, and O for a feldspar, to a probable ID such as Ba and S together or Si and O as the only elements detected. But given the possible range of weight given to a provisional ID, I think that this terms requires some explicit definition or explanaiton, either in the terms section or elsewhere. | No change The OSAC Proposed standard will not be revised based on this comment. | The explanation of "provisional" is clear from the context.
STR Comment | 17 | 123 | 131? | NWMR | "overscan" | None Provided | See comment on line 115 | No change The OSAC Proposed standard will not be revised based on this comment. | See modification to 6.2.3.4 recorded on comment 92: Section 6.2.3.4 was stregthened and now reads "Overscanning (see 3.1.6) large portions of a specimen when there are multiple phases/compositional domains present should not be used for quantitative analysis but can be used for for detecting the elements present." |  | --also, fix spelling of modification --I would change the resolution to revised. They also don't like when you reference another comment line item. Just reference the change and copy what you had in the other comment (which you did).
STR Comment | 18 | 129 | 138? | NWMR | "Discussion: …." | None Provided | I'm not sure what is being said here. | Revision was made The OSAC Proposed standard was revised based on this comment. | Section 2.1.6.1 was eliminated. |  | should be 3.1.6.1 eliminated?
STR Comment | 19 | 132 | 140 | NWMR | "standard: ..." | "a material with a known composition. The best standards have a simple flat polished geometry which eliminates ambiguities resulting for non-ideal geometries." | Its better to compare flat, polished to particles than particles to particles. | Revision was made The OSAC Proposed standard was revised based on this comment. | 3.1.7.1 added as "Discussion: The best standards have a simple flat geometry."
STR Comment | 20 | 135 |  | David Stoney | Spot analysis | spot analysis | Edit recommended as the other terms are not capitalized. | Revision was made The OSAC Proposed standard was revised based on this comment. | S changed to s.
STR Comment | 21 | 143 |  | Chris Palenik | 3.1.8.1 Discussion: There are several alternative terms and …. | 3.1.8.1 Discussion: There are several alternative or closely related terms and | add "or closely related" terms as these are not exactly synonymns | Revision was made The OSAC Proposed standard was revised based on this comment. | 3.1.8.1 changed as suggested: There are several alternative or closely related terms and |  | the change you have listed is not what is in the document.
STR Comment | 22 | 149 |  | Chris Palenik | None Provided | None Provided | The topics in section 4 are important, but I don't follow their organization….I think the section flow should be refined. | Revision was made The OSAC Proposed standard was revised based on this comment. | Significance and Use section was reorganized, moving limations to the end and moved the goals statement to the front.
STR Comment | 23 | 150 |  | Chris Palenik | 4.1 This guide is intended to advise and assist analysts in the effective application of SEM/EDS to the analysis of geological materials. | None Provided | Do you want to mention comparison? Since the dominant application of forensic geology in the US is comparision, it seems worth mentioning in this section. | Revision was made The OSAC Proposed standard was revised based on this comment. | Comparison was covered in what was 4.8, but to emphasized this point. What was 4.8 was moved to now be 4.2
STR Comment | 24 | 150 | 160 | NWMR | "semi-quantiative" | None Provided | What is the definition of "semi-quantitative?" If you provide numbers, it is quantitative (albeit maybe with large error bars.) What exactly would a semi-quantitative analysis look like? | Revision was made The OSAC Proposed standard was revised based on this comment. | In 4.6 this text was added "In this document semi-quantitative EDS analysis consists of attributing peaks to elements and observing the peak height and area"
STR Comment | 25 | 155 |  | Chris Palenik | elemental composition of the clay-sized fraction from soils submitted for forensic | None Provided | The term “elemental composition” as used in this guide is potentially misleading. To many, I think that term would imply quantitative, but for all practial purposes in this guide, it is really the “qualitative elemental composition" or at best something that looks at relative peak ratios. Perhaps defining "Elemental composition" as it is intended to be used in this guide would be helpful. | Revision was made The OSAC Proposed standard was revised based on this comment. | SEM/EDS can be used to determine the elemental composition of geologically-derived materials individual component particles with high spatial resolution (sub-micrometer). and the identification of elements in bulk elemental composition ofthe clay-sized fraction from soils submitted for forensic examination. Some references use |  | The second half of this statement is not in the document. Document ends with (sub-micrometer). Probably was accidentally removed because of next line item. | Now table accurately docuements change
STR Comment | 26 | 155-156 |  | David Stoney | Some references use SEM/EDX, SEM/EDXA, or SEM/EDAX as synonyms of SEM/EDS. | Addition of SEM/EDS to Section 3 (between Lines 126 and 127) with the following text. "SEM/EDS, n - abbreviation of Scanning Electron Microscopy and Energy Dispersive X-Ray Spectrometry. Some references use SEM/EDX, SEM/EDXA, or SEM/EDAX as synonyms of SEM/EDS. | Edit recommended as the last line of paragraph 4.2 seems out of place and fits better with the terminology section. | Revision was made The OSAC Proposed standard was revised based on this comment. | This sentence was removed. |  | No need to reference the comment number in the adjudication. | removed refernce to comment number
STR Comment | 27 | 156 |  | Chris Palenik | Some references use SEM/EDX, SEM/EDXA, or SEM/EDAX as synonyms of SEM/EDS. | Relocate or remove | I don't think this is needed. Regardless, this sentence does not belong in this section. | Revision was made The OSAC Proposed standard was revised based on this comment. | This sentence was removed. |  | No need to reference the comment number in the adjudication. | removed refernce to comment number
STR Comment | 28 | 160 |  | David Stoney | Qualitative or semi-quantitative EDS analysis can be used to confirm provisional mineral identifications.... | Qualitative or semi-quantitative EDS analysis can be used to provide data to test and provide support for provisional mineral identifications ... | Edit recommended as it makes it clear that the data resulting from EDS interact with a provisional identification by providing a means to test this identification, which may lead to support. | Revision was made The OSAC Proposed standard was revised based on this comment. | Qualitative or semi-quantitative EDS analysis data can be used to confirm provide additional information to test provisional mineral identifications made using other techniques (e.g., polarized light microscopy [PLM], Raman spectroscopy).
STR Comment | 29 | 163 |  | David Stoney | Imaging by SEM can provide additional information enabling confirmation of provisional mineral identifications. | Imaging by SEM can provide additional information to test and provide support for provisional mineral identifications. | Edit recommended as it makes it clear that the data resulting from imaging interact with a provisional identification by providing a means to test this identification, which may lead to support. | Revision was made The OSAC Proposed standard was revised based on this comment. | change to: Imaging by SEM can provide additional information enabling confirmation to test and provide support for of provisional mineral identifications and morphometric evaluation. (see comment 30)
STR Comment | 30 | 165 |  | Chris Palenik | 4.5 Imaging by SEM can provide additional information enabling confirmation of provisional mineral identifications. | 4.5 Imaging by SEM can provide additional information enabling confirmation of provisional mineral identifications and morphometric evaluation. | Add "and morphometric evaluation" or "through morphometric evaluation" | Revision was made The OSAC Proposed standard was revised based on this comment. | change to: Imaging by SEM can provide additional information enabling confirmation to test and provide support for of provisional mineral identifications and morphometric evaluation. (see comment 29)
STR Comment | 31 | 166 |  | Chris Palenik | 4.6 Limitations to SEM/EDS: | Relocate this subsection | It seems odd to place limitations as 4.6, right in the middle of 4.5 and 4.7, which are both advantages. | Revision was made The OSAC Proposed standard was revised based on this comment. | Moved Limitations to end of Significance and Use
STR Comment | 32 | 166-167, 345 |  | RV | typically around 0.1 weight % | Can this be more definitive? "typically around" seems to be loose language. | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | It would be rare to need to know the precise LOD of an element. It will vary among elements, sample geometry, and matrix. If needed, the analyst could determine the true LOD, but for context of expected LOD, this statement was included.
STR Comment | 33 | 169 |  | David Stoney | Alternative analytical techniques, such as inductively coupled plasma mass spectrometry, provide lower detection limits. | Alternative analytical techniques, such as inductively coupled plasma mass spectrometry, can provide lower detection limits. | Edit recommended as the technique has the ability to provide lower detection limits. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to: Alternative analytical techniques, such as inductively coupled plasma mass spectrometry or X-ray fluorescence, can provide lower detection limits. |  | rationale has x-ray fluorescence. Document has micro x-ray fluorescence. Clarify one. | should be micro Xray. omitted from this table on ballot
STR Comment | 34 | 169 |  | Chris Palenik | Alternative analytical techniques, such as inductively coupled plasma mass spectrometry, provide lower detection limits. | None Provided | To me, this discussion gets too deep in the weeds. ICP-MS is a method with minimal, practical, established application in forenisc geology. If you're going to keep it, I'd add micro-XRF, but I would strongly consider removing this altogether. I don't think lower detection limits are important at the level this guide is intended for. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to: Alternative analytical techniques, such as inductively coupled plasma mass spectrometry or X-ray fluorescence, can provide lower detection limits. |  | rationale has x-ray fluorescence. Document has micro x-ray fluorescence. Clarify one. | should be micro Xray. omitted from this table on ballot
STR Comment | 35 | 171 |  | David Stoney | Some samples and SEM systems require covering the sample with a conductive coating. | Some specimens and SEM systems require covering the specimen with a conductive coating. | Comment: Some analysts strongly prefer the term "specimen" to be used when referring generally to the item under examination, reserving the term "sample" to refer to a portion of a larger specimen that has been isolated (as later, in your section 5.1.2 where the term "subsamples" is currently used). The authors may wish to consider adopting this usage, which would mean additional similar edits throughout the document. I don't feel that this change essential, as the term "sample" is widely used as a synonym for specimen, but it is worth considering. | Revision was made The OSAC Proposed standard was revised based on this comment. | "sample" was replaced with "specimen" in numerous sections of this document when referring to the material under analysis. The term sample is retained when referring the evidentiary material, as this a common usage in forensic science. Note, this usage of "sample" does not align with geological "assays" (e.g. mining) in which a "sample" is representative and "specimen" is not; in contrast an evidentiary sample is almost never representative of the source.
STR Comment | 36 | 173 |  | David Stoney | Discoloration of materials can be caused by electron beam irradiation. | Electron beam irradiation can effect some materials, e.g. discoloration or heat-induced alterations such as loss of water from hydrated minerals. | Edit recommended to generalize the point as the electron beam can have other effects. | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed to: Electron beam irradiation can irrevocably change some specimens (e.g discoloration, loss of water, or migration of light elements.) |  | Just a note - it should be "e.g.," that being said, in the document, the ASTM editor will expand it to for example. | comma after e.g. was added to docuemnt
STR Comment | 37 | 174 |  | David Stoney | EDS provides elemental information, not structural information, preventing identification of polymorphs (e.g., CaCO3, Al2SiO5, TiO2, and KAlSi3O8). | EDS provides elemental information, not structural information, and so will not distinguish among polymorphs (e.g., as seen in CaCO3, Al2SiO5, TiO2, and KAlSi3O8). | Edit recommended as focuses on the inability to distinguish these materials. | Revision was made The OSAC Proposed standard was revised based on this comment. | Now reads 4.8.4 EDS only provides elemental information, which may limit its ability to identify certain minerals on its own.
STR Comment | 38 | 175 |  | Chris Palenik | 4.6.4 EDS provides elemental information, not structural information preventing identification of polymorphs (e.g., CaCO3, Al2SiO5, TiO2, and KAlSi3O8). | 4.6.4 EDS provides elemental information, not structural or chemical information, which can limit the identification of a specific phase and or preventing identification of polymorphs (e.g., CaCO3, Al2SiO5, TiO2, and KAlSi3O8). | add: or chemical information, which can limit the identification of a specific phase and or | Revision was made The OSAC Proposed standard was revised based on this comment. | Now reads 4.8.4 EDS only provides elemental information, which may limit its ability to identify certain minerals on its own.
STR Comment | 39 | 175 |  | Chris Palenik | Preventing polymorphs | rephrase | The method doesn't "prevent" polymorphs from being Ided or distinguished, it simply can't ID or distinguish polymorphs. | Revision was made The OSAC Proposed standard was revised based on this comment. | Now reads 4.8.4 EDS only provides elemental information, which may limit its ability to identify certain minerals on its own. |  | I liked the polymorph reference. I think you lose something with the current revision. Perhaps ... may limit its ability to identify certain materials (e.g., those with polymorphs) on its own. Just a suggestion. | polymorphs are more specifically discussed in note 28
STR Comment | 40 | 180 |  | Chris Palenik | 4.6.6 Procedures for quantitative methods using EDS are not covered in this guide; they are available elsewhere (ASTM E1508). | None Provided | I realize there is a fine line to walk between qual and quant, but to me, the document requires some explanation that faces the qual/quant issue more directly. You might look to XRF of glass for some help (though it's not perfectly translatable). | No change The OSAC Proposed standard will not be revised based on this comment. | In numerous places in this document, it states that elemental quantitation is not within its scope, but there are references that prescribe how perform quantiation, if the user wishes to do so. Therefore this topic will not be expanded on within the document.
STR Comment | 41 | 180 | 190,? | NWMR | "relative abundances" | see previous comment - Line 460 | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | See response to comment 115 - elemental quantitation is not within the scope. |  | OSAC doesn't want you to refer to another comment. Just copy your response from that comment into this line item. | Response to comment 115"Elemental quantitation is not within the scope of the document."
STR Comment | 42 | 186-192, 417-418, 427-428 |  | RV | SEM/EDS analysis is a component of an overall examination scheme and is not typically used on its own (7.1.1). The presence, absence, and relative abundance of mineral species or morphotypes determined by SEM/EDS in two samples can be used to evaluate whether there are exclusionary differences during a forensic comparison. | Should this be incorporated in some way into the Scope? | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | in 1.2 of the scope..."This document describes sample handling and preparation, instrument operating conditions, spectral data collection, evaluation of EDS data quality, interpretation of EDS spectra for the identification of inorganic geological materials, documentation of morphology by SEM imaging, and criteria for sample comparison." In the Significance and Use it states " SEM/EDS analysis is a component of an overall examination scheme and is not typically used on its own."
STR Comment | 43 | 196 |  | Chris Palenik | 5. Sample Preparation | None Provided | To me, the current sample preparation is patchy and could be better organized. This document scope covers soils, sediment, rocks and geologically derived samples. I realize that isolation and fractionation of samples is beyond the sope of this document, but I think some acknowledgement and better organization of this section would be helpful: acknowledging (without going into every detail) (a) the need to consider the format of the original sample matrix, (b) the potential need to isolate, separate, fractionate, polish a sample (again, based upon its original matrix (soil, sand, rock)), and (c) then how you mount each of the sample types you might present to the SEM (grains, water extracts, powders, polished samples). | Revision was made The OSAC Proposed standard was revised based on this comment. | The sample preparation section was reorganized
STR Comment | 44 | 196-197 |  | Sara Walker | 5.1.1 Use sample preparation practices that minimize the possibility of cross-contamination. | Reference applicable documents for the practices to sample preparation practices that minimize the possibility of cross contamination on sample preparation practices. If there are no existing documents to reference then add the considerations for minimizing cross contamination in more detail than the current list in the notes | Some guidance is necessary on how to minimize cross contamination, whether its in another existing document or put into this one. There is a note that lists possible methods but has no detail or prioritization of options | Revision was made The OSAC Proposed standard was revised based on this comment. | There is guidance on means to minimize cross contamination specific to this document in Note 1. However we have added a citation to the draft "WK89493 Standard Guide for the Detection and Preservation of Forensic Trace Evidence" currently being developed at ASTM and originating from the OSAC Trace Materials subcommittee. It is assumed that WK89493 will be published before 2024-S-0012 is published by the SDO. | Addition of a citation. | This reference should be up in section 2.1 since it (will be) an ASTM standard, not in the citations at the end. | Moved reference of wk89493 to terminlogy
STR Comment | 45 | 201 |  | Chris Palenik | Note 1 - Approaches for minimizing cross contamination include: preparing questioned and known | None Provided | This is a mixture of several ideas, of varying levels of specificity. For example, the last sentence of 5.1.1, seems like it should be part of a note. I think these ideas might be better in their own section with a bit more organized discussion. | Revision was made The OSAC Proposed standard was revised based on this comment. | The statement "Exposure of a sticky mount during specimen preparation can serve as an air blank." was moved to its own section and modified as shown in comment 46. The rest of Note was modified as: Note 1 - The nature of the samples affects precautionary measures recommended to prevent cross contamination. Approaches for minimizing cross contamination include: preparing questioned and known samples in separate locations, storing prepared specimens samples within covered containers to minimize dust accumulation, and placing questioned and known items within the SEM/EDS instrument at separate times. Specimens of limited quantity and small particle size not embedded in mounting material require the greatest measures to mitigate potential cross contamination (22).
STR Comment | 46 | 202 |  | David Stoney | Exposure of a sticky mount during sample preparation can serve as an air blank. | Exposure of a sticky mount during sample preparation can provide a control specimen allowing tests for possible airborne contamination. | Edit recommended as providing additional description. | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed to "Exposure of a sticky mount during sample preparation can provide a control specimen allowing tests for possible airborne contamination. "
STR Comment | 47 | 203 |  | Chris Palenik | minimize dust accumulation, and placing questioned and known items within the SEM/EDS instrument at separate times . Exposure of a sticky mount during sample preparation can serve as an air blank. | None Provided | Is this a requirement? GSR community has monitored for cross contamination, and many labs analyze stubs from multiple or different individuals together. | No change The OSAC Proposed standard will not be revised based on this comment. | No, this is not a requirement. The Note is written as guidance and notes that the precautions needed for each case may be different.
STR Comment | 48 | 206 |  | Caitlin Berry | When preparing subsamples of particulate material, use procedures that create representative subsamples. | When preparing subsamples of particulate material, use procedures that create representative subsamples and document all potential contributions to uncertainty of measurements (e.g. technician, instrument used, sample bottle location, etc). | In this section you describe a couple of methods for getting representative subsamples. I might sugeest just some verbiage to include thinking about how the preparation of the samples might introduce uncertainty in the measurements and therefore the final result. | Revision was made The OSAC Proposed standard was revised based on this comment. | Note added: "Non-representative subsampling might introduce bias in the modal abundance of particles, affecting downstream analyses."
STR Comment | 49 | 208-213 |  | Sara Walker | 5.1.2.1 Sample splitters, or cone and quartering (4) can be used to create representative subsamples, but both can be impractical for samples of limited quantity. 5.1.2.2 An alternative method of representative sub-sampling appropriate for small quantities of powder, is first to mix the particles, moistening (with an appropriate liquid such as water) to cause particle adhesion, and then selecting one or more sub-samples for analysis. | 5.1.2.1 Sample splitters, or cone and quartering (4) should be used to create representative subsamples for samples that are not of limited quantity. 5.1.2.2 For samples of limited quantity an alternative method of first mixing and moistening the particles (with an appropriate liquid such as water) to cause particle adhesion, and then selecting one or more sub-samples for analysis from this mixture. | More definitive language on what the standard practices are is necessary in a standards document | Revision was made The OSAC Proposed standard was revised based on this comment. | 5.1.2.1 (now 5.3.1.6) was not changed. Sample splitters vs cone & quartering are two distinct methods so use of SHOULD is not appropriate. 5.1.2.2 (now 5.3.1.7) added "from this mixture" to the statement.
STR Comment | 50 | 210 |  | David Stoney | An alternative method of representative sub-sampling appropriate for small quantities of powder, | An alternative method of representative sub-sampling appropriate for small quantities of powder, [add a citation] | Adding a citation for this method would be useful. | No change The OSAC Proposed standard will not be revised based on this comment. | We are unaware of a formal reference for this informal method, but inclusion within this document (and in E3294) will allow citations of this method.
STR Comment | 51 | 212 | 223 | NWMR | "Uncoated samples …. voltage" | None Provided | Use of variable pressure mode to collect EDS spectra is fraught and should not be recommented. Depending on the chamber pressure, the incident beam may contain a small fraction of the original beam, the remainder being scattered by as much as millimeters. | Revision was made The OSAC Proposed standard was revised based on this comment. | NOTE modified as: "Variable pressure mode increases the width of the spot size and can reduce the EDS signal. The beam scatter could result in contributions to the EDS spectrum from the mounting material or other material that is in proximity. The wider spot size could also worsen image quality. The analyst must balance the benefits of retaining uncoated samples against the reduced data quality and reduced spatial resolution." |  | The text in the response is not the same as the text in the document. | Table corrected to match rerised text
STR Comment | 52 | 214 | 225 | NWMR | "The width of the spot… conditions" | None Provided | Yes, but the image quality is not a good estimator of the spectrum quality. You can get a nice image from the 50% of the beam that remains in the beam but a lousy spectrum from the 50% of the beam that is scattered. It is just too easy to make an egrigious error in VP. | Revision was made The OSAC Proposed standard was revised based on this comment. | NOTE modified as: "Variable pressure mode increases the width of the spot size and can reduce the EDS signal. The beam scatter could result in contributions to the EDS spectrum from the mounting material or other material that is in proximity. The wider spot size could also worsen image quality. The analyst must balance the benefits of retaining uncoated samples against the reduced data quality and reduced spatial resolution." |  | Don't refer to previous note. Just repeat the explanation. | Table corrected to match rerised text
STR Comment | 53 | 219 | 230 | NWMR | "Metal coatings…." | None Provided | So long as you know which metal you are using, metal coatings are usually thinner and can work just fine. | No change The OSAC Proposed standard will not be revised based on this comment. | The line offers the use of metal or carbon coating. Carbon is routinely used to coat geological materials for quantiative microbeam analysis.
STR Comment | 54 | 220 |  | Chris Palenik | 5.1.4 Sample preparation is commonly carried out with the aid of a stereomicroscope. Once the SEM mounts are prepared, protect them from surface abrasion and atmospheric dust deposition with a protective lid. Storing the prepared samples in a vacuum chamber or desiccator reduces SEM chamber evacuation time . | None Provided | This point has three separate and, really, independent thoughts that probably don't all belong in the same sub-section. | Revision was made The OSAC Proposed standard was revised based on this comment. | The three sentences are now separate sections.
STR Comment | 55 | 223 |  | Caitlin Berry | Uncoated samples can be analyzed using variable pressure mode (3.1.11) or using low accelerating voltage. | perhaps remove/revise reference to 3.1.11 or add in a section 3.1.11 | I don't see a section 3.1.11 in the document, perhaps it is supposed to be 3.1.8? | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed 3.1.11 to 3.1.8.
STR Comment | 56 | 225 |  | Chris Palenik | Uncoated samples can be analyzed using variable pressure mode (3.1.11) or using low accelerating voltage . | None Provided | I would suggest mentioning that low voltage limits the elements that can be addressed/accessed. | Revision was made The OSAC Proposed standard was revised based on this comment. | An new note acknowledging these effects was added: "Lower beam current can reduce detection limits and lower accelerating voltage can affect the excitation of elements. "
STR Comment | 57 | 229 |  | Chris Palenik | The wider spot size could also worsen image quality. | remove | While techincally true, I don't think this is important here. Generally, if you need to use variable pressure, the VP is actually going to improve imaging relative to an uncoated, secondary image. Also, at the magnification levels typically used for grain analysisi, broadened spot size is not going to be a major factor in imaging. | Revision was made The OSAC Proposed standard was revised based on this comment. | NOTE modified as: "Variable pressure mode increases the width of the spot size and can reduce the EDS signal. The beam scatter could result in contributions to the EDS spectrum from the mounting material or other material that is in proximity. The wider spot size could also worsen image quality. The analyst must balance the benefits of retaining uncoated samples against the reduced data quality and reduced spatial resolution. |  | wording in response does not match wording in document. | The text corrected as noted in table
STR Comment | 58 | 233 |  | Chris Palenik | (e.g., determination of phosphorus and zirconium) and ZAF correction (see E1508). If coating is | remove/revise | Seems odd to mention ZAF here, since this is not a document on quantitative analysis (which is the only place ZAF is used. Also, there are bigger impacts on ZAF than a metal coating (sample prep, matrix differences). | Revision was made The OSAC Proposed standard was revised based on this comment. | Note modified to: Metal (e.g., gold, palladium, platinum) coatings can interfere with X-ray peak assignments (e.g., determination of phosphorus and zirconium) and ZAF correction (see E1508). If coating is required, carbon is preferred to metals when EDS analysis is performed. |  | Zirconium is still listed in the standard but is crossed out in the rationale. | Zirconium should not have been strike through in table. corrected.
STR Comment | 59 | 234 |  | David Stoney | ... use the same sample preparation technique for all compared samples. | ... use the same sample preparation technique for all samples to be compared. | Edit recommended for clarity. | Revision was made The OSAC Proposed standard was revised based on this comment. | modified to "use the same specimen preparation technique for all samples to be compared." |  | Note... you are mixing sample and specimen in the statement in the document. I recommend re-reading the statement to make sure it is what you are meaning it to say. | In comment 24, the use of sample and specemnt are explained.
STR Comment | 60 | 235 | 245? | NWMR | "remove extraneous material" | None Provided | Low temperature ashing is often a good way to burn off organics from soil while retaining inorganics. | Revision was made The OSAC Proposed standard was revised based on this comment. | extraneous materials was removed and grain coatings were replaced by "adhering materials)" |  | The document wasn't revised based on this comment, as there is no reference to ashing and the only method of removing adhering material is via washing. ashing is also destructive if you are wanting to retain the adhering material.
STR Comment | 61 | 237 |  | Chris Palenik | 5.2 Unpolished particles | None Provided | Somewhere, this section on grain mounts should really discuss the benefit of mounting grains from the same size fraction. | Revision was made The OSAC Proposed standard was revised based on this comment. | The sample preparation section was reorgianized. Selection of particle sizes is now in 5.3.1.3, 5.3.1.5, and note 5. |  | re-read 5.3.1.1 - it reads a little strangely to me. Should it just be particle separation and not "the particle separation"? Do you mean 5.3.1.3 instead of 5.3.1.5? 5.3.1.5 doesn't talk about particle sizes. | 5.3.1.1 corrected to 5.3.1.3. and "the" was revomved.
STR Comment | 62 | 248 |  | Chris Palenik | extraneous materials, it is recommended to retain the extraneous material for potential subsequent analysis . Grains can be either individually mounted or applied as a dispersion. | None Provided | The meaning of "extraneous materials" needs to be better specified. Do you mean, by using EDS to analyze dried residue from a water extract? If so, there should be a section on sample prep for collecting, preparing and analyzing the water soluble fraction of the sample. I think this idea could be put into a better organized framework. | Revision was made The OSAC Proposed standard was revised based on this comment. | extraneous materials was removed and grain coatings were replaced by "adhering materials"
STR Comment | 63 | 254 |  | David Stoney | ... enables sufficient spacing between the grains to minimize mixing of EDS spectra from distinct grains and minimizes the geometric effect of adjacent grains blocking signal from reaching the detector. | ... allows grains to be well-spaced on the stub, avoiding the mixing of EDS spectra or blocking of EDS signals that can occur when grains are very closely spaced. | Edit recommended to avoid "sufficient spacing to minimize...", the meaning of which is unclear. | Revision was made The OSAC Proposed standard was revised based on this comment. | now reads Note 7 - Back sieving of sand- or silt-sized grains (4) onto a sticky SEM stub allows grains to be well-spaced on the stub, avoiding the mixing of EDS spectra or blocking of EDS signal that can occur when grains are very closely spaced.
STR Comment | 64 | 255 |  | Chris Palenik | Note 5 - Back sieving of sand- or silt-sized grains (4) onto a sticky SEM stub enables sufficient spacing | None Provided | “back sieving” should be defined for the reader. Also, I think that this method should be specifically called out in its own point in this sub-section on preparing unpolished particles. | No change The OSAC Proposed standard will not be revised based on this comment. | now reads "Note 7 - Back sieving of sand- or silt-sized grains (4) onto a sticky SEM stub allows grains to be well-spaced on the stub, avoiding the mixing of EDS spectra or blocking of EDS signal that can occur when grains are very closely spaced." The reference provided allows the reader to learn what back sieving is if they do not know.
STR Comment | 65 | 257 | 267? | NWMR | None Provided | None Provided | Even with conductive coatings, highly textured samples (like spread clays) often charge. It might be worth mentioning the use of the Duane-Hunt limit as a tool to evaluate sample charging. | No change The OSAC Proposed standard will not be revised based on this comment. | This issue is too detailed for this document.
STR Comment | 66 | 257 | 267? | NWMR | overscanning | None Provided | see above | No change The OSAC Proposed standard will not be revised based on this comment. | Caveats for overscanning of clays are described in 6.2.3.3 and 6.2.3.4. This section is about sample preparation for overscanning.
STR Comment | 67 | 269-270 |  | David Stoney | Note 6 - The larger crystal size of diagenetic clay minerals allows identification by SEM/EDS, whereas the small mixed crystals found in soil clay minerals can prevent identification by SEM/EDS. | Note 6 - Some larger, minimally weatherd crystals may be individually characterized using SEM/EDS. | Edit recommended to remove terms 'diagenetic" and "identification". Identification seems inappropriate as it is emphasized in this document that identification requires additional complementary analyses. | Revision was made The OSAC Proposed standard was revised based on this comment. | This note was eliminated as it is too detailed for the document.
STR Comment | 68 | 271 |  | Chris Palenik | Note 6 - The larger crystal size of diagenetic clay minerals allows identification by SEM/EDS, whereas | None Provided | This is getting a bit in the weeds relative to other discussions in this document. If you're going to keep this, 1) I think this term should be defined for the reader. (2) This comment does not belong in the sample prep section. | Revision was made The OSAC Proposed standard was revised based on this comment. | This note was eliminated as it is too detailed for the document.
STR Comment | 69 | 273 |  | Chris Palenik | Note 7 - EDS data from the clay-sized fraction can be used together with other methods (e.g., X-ray diffraction, FT-IR, or staining) for identification of mineral phases present. | None Provided | This comment doesn’t belong in the sample prep section. | Revision was made The OSAC Proposed standard was revised based on this comment. | text was moved to note 20 in section 6.2.3.4.
STR Comment | 70 | 274 |  | David Stoney | When isolation of particles from their substrate is not feasible due to risk of sample loss, .... | When isolation of particles from their substrate risks their loss or alteration, .... | Edit recommended to avoid confusion of feasibility and risk. | Revision was made The OSAC Proposed standard was revised based on this comment. | suggested language was accepted.
STR Comment | 71 | 278 |  | Chris Palenik | in situ analysis can be conducted (e.g., soil embedded in fabrics). | None Provided | This seems really specific compared to other parts of the document that discuss sample prep, particularly when there are good ways to get soil off of a clothing item. | Revision was made The OSAC Proposed standard was revised based on this comment. | For clarity, example of in situ analysis of projectiles was added. and in section 5.7.2, many was changed to some |  | in rationale, should it reference 5.7.2 instead of 5.7.1 | Table corrected 5.7.2 from 5.7.1
STR Comment | 72 | 278 |  | Sara Walker | placement in the SEM chamber. | remove extra space between placement and in | Typographical error | Revision was made The OSAC Proposed standard was revised based on this comment. | Removed extra space between placement and in | Editorial
STR Comment | 73 | 284 | 289 | NWMR | "(5)" | None Provided | Is the reference right? | Revision was made The OSAC Proposed standard was revised based on this comment. | This reference was removed as the section was simplified to: "Use laboratory-established protocols for instrument operation and maintenance"
STR Comment | 74 | 286 |  | Chris Palenik | 6.1.1 SEM calibration | None Provided | Calibration and performance monitoring are two separate ideas and need to be treated separatly in this document. | Revision was made The OSAC Proposed standard was revised based on this comment. | new section (now 6.1.2.2) and associated note : Ensure electron gun and aperture alignment. Note 8: Significant changes in accelearating voltage could require adjustments to the aperture alignment. |  | don't reference comment 73. copy here. Also note that the section was changed from calibration to quality assurance.
STR Comment | 75 | 297 | 302 | NWMR | "Set the accelerating …." | Set the accelerating voltage to best interogate the sample characteristic you wish to observe. Low voltages provide more surface sensitivity while higher voltages can see through surface contamination." | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | partially incoporated. Now says: Set the accelerating voltage to best observe the sample characteristics under investigation. Note 11: Lower voltages provide more surface sensitivity while higher voltages penetrated deeper into the sample.
STR Comment | 76 | 298-299 |  | David Stoney | The grayscale intensity value of BSE images is proportional to the specimen average atomic number and can be used to assess compositional information. | The grayscale intensity value of BSE images is proportional to the specimen average atomic number and can be used to for assessment of this compositional character of the specimen. | Edit recommended to clarify the nature of the information. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "The grayscale intensity value of BSE images is generally proportional to the specimen average atomic number and can be used for assessment of this compositional character of the specimen."
STR Comment | 77 | 300 |  | Chris Palenik | Note 9 - The grayscale intensity value of BSE images is proportional to the specimen average atomic number and can be used to assess compositional information. | None Provided | Generally proportional to the atomic number but there are exceptions (e.g. CL-luminescent materials can give an an omalously high BSE signal relative to its average atomic number. There is an article on this - I think it's in the microscope, by Robert Carlton. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "The grayscale intensity value of BSE images is generally proportional to the specimen average atomic number and can be used for assessment of this compositional character of the specimen."
STR Comment | 78 | 305 | 311 | NWMR | "semi-quantiative" | None Provided | If you are making a length measurement, it is quantitative. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "For morphometric assessments, confirm the instrument scale calibration"
STR Comment | 79 | 305 | 311? | NWMR | None Provided | None Provided | There is a standard for length calibration in an SEM. | Revision was made The OSAC Proposed standard was revised based on this comment. | now referenced E766
STR Comment | 80 | 308 |  | Chris Palenik | 6.1.2.6 For non-conductive materials, use low kV, variable pressure, a conductive coating, or a combination of these techniques to mitigate charging. | None Provided | I believe this is redundant with an earlier section. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "If you observe sample charging, use a mitigation method as described in section 5.2.7 and note 2." to minimize redundancy. The early section focuses on sample preparation and this section on analysis.
STR Comment | 81 | 309 |  | RV | sufficient | Can this be defined? If not, I think a reasonable examiner will determine sufficiency according to FSSP policy. | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | Because this standard describes applications of SEM EDS for distinct applications (ID of unknown, confirmation of ID of unknown previously characterized for another method, comparison two or more samples, and provenance) what is sufficient will be distinct for each end goal.
STR Comment | 82 | 313 |  | Chris Palenik | 6.1.4.1 For semi-quantitative morphometrics, confirm the instrument scale calibration. | None Provided | This topic was touched on in the calibration/performance check session and should have already been completed before getting to this stage. | Revision was made The OSAC Proposed standard was revised based on this comment. | 6.1.4.1 reads "For morphometric assessment, confirm the instrument scale calibration (E766)."
STR Comment | 83 | 314 |  | Chris Palenik | 6.1.4.2 Particle surface texture examination - Examine and document the surface | None Provided | Given note 12, I question whether this topic should be discussed at all in this document. | No change The OSAC Proposed standard will not be revised based on this comment. | Particle surface texture analysis is a highly specialized method but it does have forensic relevance. It is important for the reader to be aware of the method. Specialized references are provided as additional resources.
STR Comment | 84 | 320 |  | Chris Palenik | Note 11 - The charging of particles due to limited contact to a conductive substrate can be reduced by using a conductive paint, commonly carbon or silver. However, this option can be supplanted using low vacuum (3.1.9) or low voltage imaging mode in systems with this capability. | None Provided | Redundant | Revision was made The OSAC Proposed standard was revised based on this comment. | This note was removed from this section but the additional detail was added to the section on sample prepratation (5.2.7. to 5.2.7.1).
STR Comment | 85 | 330-331 |  | David Stoney | If during sample preparation the sample was size-fractionated then assessing particle size and sorting is not appropriate. | If during sample preparation the sample was size-fractionated then any measurements or comparisons of particle size and sorting data must allow for this preparation. | Edit recommended as allowing for comparison within size fractions. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed as suggested with minor modification" If during sample preparation the sample was size-fractionated then any measurements or comparisons of particle size and sorting data should account for this preparation." |  | Reword document without using "must" | changed "must alow " to " should account"
STR Comment | 86 | 338 |  | Chris Palenik | 6.2.1 EDS Instrument Calibration | None Provided | What is being discussed in this section is not a calibration but a performance check. | Revision was made The OSAC Proposed standard was revised based on this comment. | calibration changed to performance verification.
STR Comment | 87 | 342 | 347 | NWMR | "reduce the analytical…" | "reduce the analytical volume, increase light element sensitivity and…" | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | changed as suggested.
STR Comment | 88 | 346 | 351 | NWMR | "aiding in" | "confirming" | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "When L or M X-ray family lines are provisionally identified, increasing the accelerating voltage can confirm the presence of L or K X-ray family lines of that element"
STR Comment | 89 | 350 |  | Chris Palenik | Note 16 - If the SEM/EDS has a beryllium window, detection of light elements (i.e., fluorine and below) is not possible. | None Provided | This seems out of place (and probably almost irrelevant in the year 2024). | Revision was made The OSAC Proposed standard was revised based on this comment. | Note was removed.
STR Comment | 90 | 354 | 359 | NWMR | "The working …" | "The optimal working distance for X-ray analysis is instrument dependent. Correctly adjusted it …. by the EDS detector The working distance is configured using tbe objective lens, then the specimen should be brought into focus using the stage Z motion. The specimen should always be mounted at the optimal working distance when collecting X-ray spectra." | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | partially accepted. Now reads: "The optimal working distance for the X-ray analysis is instrument dependent Correctly adjusted, the working distance achieves a take-off angle that optimizes X-ray collection by the EDS detector.
STR Comment | 91 | 372 |  | Chris Palenik | to survey the elements present in a sample, or mapping the distribution of the elements within a defined area. | None Provided | This is the first mention of mapping. If you're going to include it, I think it deserves its own section and a bit more discussion about how that could be used. | Revision was made The OSAC Proposed standard was revised based on this comment. | Now reads : "EDS spectra can be collected for a variety of purposes, such as: characterization of individual particles, detection of compositional domains within particles, or overscanning large areas to survey the elements present in a specimen"
STR Comment | 92 | 374 | 378? | NWMR | "bulk analysis…" | None Provided | While I agree that the backscatter image of an imaged area is a good way to visualize heterogeneity, I have qualms about overscanned spectra. These qualms could be addressed by a discussion of how these spectra can be interpreted. 6.2.3.4 does provide some appropriate caveats. | Revision was made The OSAC Proposed standard was revised based on this comment. | Section 6.2.3.4 was stregthened and now reads "Overscanning (see 3.1.6) large portions of a specimen when there are multiple phases/compositional domains present should not be used for quantitative analysis but can be used for for detecting the elements present."
STR Comment | 93 | 374 |  | Chris Palenik | Note 19 - Deflecting the beam from the center of the field of view can result in non-linear X-ray counts. This effect is greater at low magnification (10). | None Provided | What do you mean by this "non-linear x-ray counts" term? | Revision was made The OSAC Proposed standard was revised based on this comment. | changed "non-linear" to "reduced"
STR Comment | 94 | 374 |  | David Stoney | Point analysis of a material is achieved by deflecting a static beam to a spot ... | Point analysis of a material is achieved by positioning a static beam to a spot ... | Edit recommended as the instruments means of positioning the beam is not necessary. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed deflecting to positioning |  | (1) in note 19, should deflecting be changed to "positioning the beam away from the center..." (2) in 6.2.3.2 consider changing the word "to" to something else. I think it worked for deflecting, but it seems like a strange word choice for positioning... | In note 19 deflecting is an appropriate word.
STR Comment | 95 | 387 |  | Chris Palenik | phases/compositional domains present should only be used for qualitative elemental determination. | None Provided | Isn’t everything in this document limited to qualitative analysis? | Revision was made The OSAC Proposed standard was revised based on this comment. | 6.2.3.1 Now reads : "EDS spectra can be collected for a variety of purposes, such as: characterization of individual particles, detection of compositional domains within particles, or overscanning large areas to survey the elements present in a specimen" |  | add that the response was made to 6.2.3.1
STR Comment | 96 | 387 | 391? | NWMR | None Provided | "The low energy end of the continuum is particularly sensitive to problems with a spectrum. If the low energy continuum is suppressed, this often suggest oclusion of the X-ray signal." | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | the end of 6.2.4.1. now reads "The low energy end of the continuum is particularly sensitive to problems. If the low energy continuum is suppressed, this often suggests occlusion of the X-ray signal. Reject spectra of poor quality and recollect"
STR Comment | 97 | 389 |  | Chris Palenik | Note 20 - Overscanning is only recommended for fine-grained clay preparations. | None Provided | Overscan can be used analysis of individual grains (polished or grain mounts). For instance, automated particle analysis can either conduct a point analysis at the center of mass or the option can be selected to analyze the entire grain (which is simply a small raster). There is certainly some reason to rater over an entire grain. Was there any discussion as to overscanning an entire sample (pro or con)? Are there non-fine grained clays? If you are suggesting that overscan should not be used to analyze a mass of grains - you should state that explicitly. | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed to to "Overscanning is typically used for fine-grained (clay-sized) preparations."
STR Comment | 98 | 389-393 |  | Sara Walker | 6.2.4.1 During and upon completion of the acquisition of an EDS spectrum, assess the data quality for: the shape, intensity, and continuity of the background; peaks; and spectral artifacts (e.g., sum peaks, escape peaks). Gaps in the background or sudden breaks in the continuum are an indication of a poor-quality spectrum resulting from sample geometry (see 6.2.3.2). | Provide direction on what to do if the spectrum does not provide suitable data quality. | The way this section currently reads, the quality assessment is performed and notes are made on observations but there is nothing to indicate how it should be handled if a poor quality ES spectra is obtained. Does this data get rejected and a new analysis performed? Does the analyst proceed forward? What is the threshold of a good quality spectra based upon? | Revision was made The OSAC Proposed standard was revised based on this comment. | "Reject spectra of poor quality and recollect." was appended to 6.2.4.1
STR Comment | 99 | 392 | 396? | NWMR | None Provided | "When an element is identified in a spectrum, label all peaks associated with the element." | Qual is critical. Many, if not most, blunders in X-ray analysis can be traced to an error in qual. | Revision was made The OSAC Proposed standard was revised based on this comment. | This was captured in 8.1.3 by appending "Within representative spectra, annotate all peaks associated with elements that are being used for mineral identification"
STR Comment | 100 | 392 | 396? | NWMR | None Provided | "Vendors software usually provides "auto-qual" routines. These should only be used to confirm a manual qualitative analysis as many error modes have been idenfified in these routines." | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | The notes that follow this section discuss how to use software based element id and its limations.
STR Comment | 101 | 392 | 396? | NWMR | None Provided | "Use care when excluding an element whose peaks are occuluded by another element. Performing a spectrum fit may be necessary to determine the presence of the element." | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | added as new note after what was note 21
STR Comment | 102 | 392 | 396? | NWMR | None Provided | "Peak interferences (like Pb-S-Mo) are best resolved by peak fitting as more than one element may be present." | None Provided | Revision was made The OSAC Proposed standard was revised based on this comment. | appending to what was note 22 (now 23)
STR Comment | 103 | 392 | 396? | NWMR | None Provided | "Multiple beam energies may be necessary to identify both low Z and high Z elements." | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | This concept is already described in note 21.
STR Comment | 104 | 401 |  | Chris Palenik | Note 21 - Modern instrument manufacturer software packages provide auto-identification of elements; if using auto-identification, confirm all peak assignments (11). | None Provided | Redundant with end of 6.2.4.3 | No change The OSAC Proposed standard will not be revised based on this comment. | 6.2.4.3 is the procedural discription, and note 21 provides information to aid the analyst regaring that procedure.
STR Comment | 105 | 404 |  | Chris Palenik | If only a single peak with low counts is assigned to an element, the element identification should be considered provisional. | None Provided | The presence/absence of a peak can be defined quantitatively based on S/N. For instance, Oxford EDS software permits a threshold to be set for auto-ID based upon S/N ratio (e.g., S/N >2, 3, 4), which eliminates the need to decide if an ID is provisional. Also, it would be a good place to add some discussion about renalaysis/replicate analysis/longer count times to clarify this point. Notwithstanding these comments, I am not sure how often trace elements come into play for general mineral grouping/provisional IDing of a mineral, which is the focus of this document. | Revision was made The OSAC Proposed standard was revised based on this comment. | New text "Provisional identifications can be confirmed by collection of replicate spectra, those with a higher number of counts, or by complementary techniques (e.g., X-ray fluorescence, inductively coupled plasma- mass spectrometry, electron microprobe)." Peak detection thresholds are often built into software so we do not recommend specifying. While trace element analysis may be less common, providing examples with real minerals is germane.
STR Comment | 106 | 406 |  | Chris Palenik | higher number of counts or by complementary techniques (e.g., X-ray fluorescence, inductively coupled plasma- mass spectrometry, electron microprobe). | None Provided | Probably ICP-MS is not particularly relevant since this procedure is largely referencing the study of individual grains. I suppose if you specified LA-ICP-MS. Also, this focus on trace levels seems a bit beyond the real focus of the document. | Revision was made The OSAC Proposed standard was revised based on this comment. | ICP-MS was removed as and example method.
STR Comment | 107 | 407 |  | Chris Palenik | electron microprobe). | None Provided | More specifically WDS. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed to "SEM or electron microprobe equipped with wavelength-dispersive X-ray spectrometers"
STR Comment | 108 | 408 |  | Chris Palenik | Note 22 - Due to the energy resolution limits of EDS detectors (approximately 130 eV), there is the potential for overlap of X-ray peaks from multiple elements. Examples | None Provided | This section is stating a fact, but it is not discussing how to treat peak overlap. For instance, peak overlap, in some cases, can be treated with spectral simulations in DTSA or in vendor-based software packages. In other cases, it’s not possible to parse out a particular overlap (e.g., small amount of lead relative to a larger sulfur peak) - and here, another method would be needed (e.g., WDS). | Revision was made The OSAC Proposed standard was revised based on this comment. | the following was added to what is now note 23: "Peak interferences (like Pb-S-Mo) are best resolved by peak fitting as more than one element may be present."
STR Comment | 109 | 417 |  | Chris Palenik | 6.2.4.4 The use of EDS spectral databases or software for the synthesis of spectra is recommended to confirm peak identification. | None Provided | (a) Not sure how spectral databases inform peak IDs, and (b) this comment should really be tied to the note above regarding trace peak ID and overlapping peaks. That said, modern, auto-IDs made using vendor-based software are already based upon spectral fitting (even if the user doesn’t see it). | Revision was made The OSAC Proposed standard was revised based on this comment. | note 22 added: "Use care when excluding an element whose peaks are occluded by another element. Performing a spectrum fit may be necessary to determine the presence of the element."
STR Comment | 110 | 419 |  | Chris Palenik | 6.2.5 Element quantitation is beyond the scope of this document. For guidance on quantitation, see Goldstein et al. (1). | None Provided | This is redundant...was stated earlier in this document. | Revision was made The OSAC Proposed standard was revised based on this comment. | This section was changed to a note. This topic is covered in significance and use, but remains relevant for the procedural section of the docuement
STR Comment | 111 | 421 |  | Chris Palenik | 6.2.6 Automated EDS-based particle categorization or identification | None Provided | Consider making this into its own section, (1) as it is a hybrid of imaging and elemental analysis. As imaging can be used to assess grain size, shape, and other imaging metrics. (2) Also, this approach is nuanced enough that it really deserves its own section to treat it properly (see next two comments) | Revision was made The OSAC Proposed standard was revised based on this comment. | Automated morphometrics is now covered in "Note 13 - Particle analysis can use a range of morphometrics and terminology which can be carried out by automated software."
STR Comment | 112 | 422 |  | Chris Palenik | 6.2.6.1 The automated detection of particles and their elemental compositions can be used to categorize and count particles. | None Provided | What is meant by categorization: Elemental composition? Particle metrics (dimensions, shape, circularity)? This, along with my prior comment is why you should consider expanding this section. | No change The OSAC Proposed standard will not be revised based on this comment. | This is covered in what was 6.2.6.1 as "The automated detection of particles and their elemental compositions can be used to categorize and count particles. The relative abundance of general groups defined by composition can be used in the forensic characterization of geological materials (12, 13)."
STR Comment | 113 | 429 |  | Chris Palenik | 6.2.6.4 Procedures for the use of automated mineral identification are beyond the scope of this document. | None Provided | Two points to consider about this statement (1) the term used here is "mineral identification" and not "provisional," which seems to break with the rest of the document. | Revision was made The OSAC Proposed standard was revised based on this comment. | Statment changed to "Procedures for automated mineral analysis are beyond the scope of this document."
STR Comment | 114 | 457 |  | Chris Palenik | Note 26 - Some compounds occur as polymorphs (minerals having the same chemical composition but different crystal structures). The identification of a particular polymorph (e.g., SiO2, CaCO3, Al2SiO5, TiO2, and KAlSi3O8) is not possible by EDS alone. Differentiation of polymorphs can be achieved using PLM, X-ray diffraction (XRD), or Raman spectroscopy. | None Provided | (1) This was discussed earlier and seems redundant. (2) the discussion of polymorphs here seeems to be an extremely detailed point, particularly when little to no guidance is given on how one might identify or provisonally identify a mineral, and (3) crtically: if this statement is saying that one can't identify a particular polymorph of SiO2 by EDS, then this implies that one shouldn't use EDS to identify (or provisionally identify) quartz or calcite (since we can't be sure of the polymorph). To avoid this, I'd not mention polymorphs at this level. | Revision was made The OSAC Proposed standard was revised based on this comment. | The integration of morphology for mineral identification is now in what was Note 26 (now 28) - Some compounds occur as polymorphs (minerals having the same chemical composition but different crystal structures). The identification of a particular polymorph (e.g., SiO2, CaCO3, Al2SiO5, TiO2, and KAlSi3O8) is not possible by EDS alone. Differentiation of polymorphs can be achieved using morphology (see section 7.3), PLM, X-ray diffraction (XRD), or Raman spectroscopy. |  | You might make an additional note in your response that this is more of a cautionary note (similar to note 27), not an instructional note.
STR Comment | 115 | 460 | 461? | NWMR | "3.1.8.1: …." | "Elemental abundances calculated from flat, polished samples using a standards-based quantitative technique can be as accurate as ±1% (relative) for major elements and ±10% (relative) for minor elements. Each deviation from this ideal measurement protocol will reduce the expected accuracy. The accuracy of standardless analysis is very vendor specific and maybe an order of magnitude worse than standards-based analysis. Quantitative algorithms designed for bulk samples should not be assumed to provide accurate results on particles particualrly for light elements." | Somewhere there needs to be caveats about the uses and expected accuracy of quantiative analysis on the types of samples discussed in this standard. Without an appreciation of the accuracy of the technique, it is not possible to state whther two samples are the similar or different. I'd prefer to see this discussion early in the document - maybe near 4.6.6 | Not germane The comment is not relevant to the document or its scope. | Elemental quantitation is not within the scope of the document.
STR Comment | 116 | 461 |  | Chris Palenik | Note 27 - Some mineral varieties can be confidently identified by their EDS spectra alone (e.g., zircon, ZrSiO4), but for other minerals EDS analysis alone is limited to provisional identification. | None Provided | To me, this is an important point, and one that the authors should consider moving from a note to the main body and providing some more detailed guidance on this classification. | Revision was made The OSAC Proposed standard was revised based on this comment. | Note 27 was appended to 7.2.1
STR Comment | 117 | 475 | 420? | NWMR | None Provided | None Provided | I don't feel that there is enough discussion about combining observations of multiple measurements (spectra) to say something about a heterogeneous sample. In the call, there was a discussion of comparison by populations of particles grouped by spectral similarity. I don't see this discussed in the document. | No change The OSAC Proposed standard will not be revised based on this comment. | This is covered in what was 6.2.5.1 as "The automated detection of particles and their elemental compositions can be used to categorize and count particles. The relative abundance of general groups defined by composition can be used in the forensic characterization of geological materials (12, 13)." There is a need for a document on criteria for soil comparisons, regardless of analytical method. This is an ambitous effort that will need to address the range of soil types and the nature of the samples.
STR Comment | 118 | 479 |  | David Stoney | Samples with minimal differences - | Comparisons showing minimal difference - | Edit recommended as comparisons show differences. | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed to "Samples showing minimal differences - Two samples of geological material...."
STR Comment | 119 | 483 |  | Caitlin Berry | ... in similar relative abundances, then no exclusionary differences are detected. | include maybe a range if possible, or light specifics as to what might be considered similar --is there a reference that has this kind of information you could point to? | I recognize some of this is qualitative and will depend on the specific usage of this standard, but just from a statistical perspective "similar" can mean a lot of things so I'm just wondering if there is any more room for specificity around this? | Revision was made The OSAC Proposed standard was revised based on this comment. | added 7.4.2: The processes of transfer and persistence of particle assemblages can cause changes in relative proportions of particle types or sizes from the source. Knowledge of the history of the sample should influence the interpretation of similarities and differences. This can be partially mitigated by comparing similar fractions (size, or grain type). Because of this the criterial for similar abundance will vary for examination to examination.
STR Comment | 120 | 490 |  | David Stoney | The identification of uncommon features, ... | The identification of less common features, ... | Edit recommended as descriptive and avoids referral to a vague category of "uncommon." | Revision was made The OSAC Proposed standard was revised based on this comment. | Chnaged to "The identification of less common features, including minerals,..."
STR Comment | 121 | 492 |  | David Stoney | ... the probative value of the evidence. | ... the strength of association provided by the comparison. | Edit recommended to avoid reference to probative value (which includes contextual and legal case aspects unrelated to the SEM/EDS examination) and to add specificity to what is meant by evidence. | Revision was made The OSAC Proposed standard was revised based on this comment. | Changed to "... the strength of association provided by the comparison."
STR Comment | 122 | 496-498 |  | David Stoney | The local characteristics can be confirmed with more expansive known exemplars, consulting published data, or contacting experts with localized knowledge (e.g., regional geological surveys, industry, academic literature, and persons with local expertise). | The local characteristics can be ascertained by using more expansive known exemplars, consulting published data, or making use of experts with localized knowledge (arising, for example, from regional geological surveys or industrial experience, academic study, or other related local expertise). | Edits recommended for clarity and to avoid the use of the word "confirmed". | Revision was made The OSAC Proposed standard was revised based on this comment. | only changed the first phase: The local characteristics can be ascertained by using ... but reject the other changes as they do not accurately capture the intent.
STR Comment | 123 | 499 |  | David Stoney | Samples with differences - | Comparisons showing differences - | Edit recommended as comparisons show differences. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed as suggested
STR Comment | 124 | 504 |  | David Stoney | Samples with explainable differences - | Comparisons showing explainable differences - | Edit recommended as comparisons show differences. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed as suggested |  | 7.4.4, 7.4.4.1, and 7.4.4.2 all still have "samples showing..." Since you changed 7.4.3 to comparisons, you should change these as well. | also changed 7.4.4, 7.4.4.1, and 7.4.4.2 to "comparisons showing..."
STR Comment | 125 | 510 |  | Caitlin Berry | Document the justification for the explainable differences. | Document the justification for the explainable differences and the uncertainty around this classification. | Given you are describing a process for assessing similarity/categorization (do the samples have explainable differences or exclusionary differences, etc.) it would be nice to also offer guidance on how to estimate the accuracy/uncertainty of these classifications. Or at least suggest that it maybe should be done. What is the risk of error in each case? I'm sure it depends greatly, but I might suggest pointing to a reference of some kind or at least including a statement addressing that it should be considered. | No change The OSAC Proposed standard will not be revised based on this comment. | Cases are too variable to be able to specify the accuracy or uncertainty.
STR Comment | 126 | 510-511 |  | David Stoney | Document the justification for the explainable differences. | The justification for explainable differences should be documented. | Edit recommended to avoid imperative sense. | No change The OSAC Proposed standard will not be revised based on this comment. | This section should be in the affirmative because if the report writer is claiming the differences are due to some factor they must describe what those factors are.
STR Comment | 127 | 512 |  | David Stoney | Samples with exclusionary differences - | Comparisons showing exclusionary differences - | Edit recommended as comparisons show differences. | Revision was made The OSAC Proposed standard was revised based on this comment. | changed as suggested. |  | Change was not made as suggested. 7.4.4, 7.4.4.1, and 7.4.4.2 all still have "samples showing..." Since you changed 7.4.3 to comparisons, you should change these as well. | also changed 7.4.4, 7.4.4.1, and 7.4.4.2 to "comparisons showing..."
STR Comment | 128 | 521-537 |  | Sara Walker | None Provided | There is no requirement for documentation of results from controls or impacting environmental conditions, if they exist | The results of controls and the impact of environmental considerations should be documented so that the information is readily available to another person reviewing the results. | Revision was made The OSAC Proposed standard was revised based on this comment. | Included "Relevant controls" as section 8.1.2.3
STR Comment | 129 | 529 | around 529 | NWMR | None Provided | add "beam energy" | None Provided | No change The OSAC Proposed standard will not be revised based on this comment. | 8.1.2 requires documentaion of accelerating voltage
STR Comment | 130 | 530 | ? | NWMR | None Provided | None Provided | Should the spectra be in the report or just descriptions? What about spectrum comparisons? | No change The OSAC Proposed standard will not be revised based on this comment. | Report format and scope are dictated by individual laboratory procedures. As stated in 8.3, anything that is needed for a secondary analyst to review to make an independent conclusion should be included; this may include spectra and spectral overlays.

## Commenter Info
Commenter # | Commenter's Name | DO NOT PUBLISH | Commenter's Position/Title | Commenter's Org/Affiliation | DO NOT PUBLISH | Commenter's Email | Date of Open Comment Submission
1 | None Received
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25