## Slide 1
Biological and Non-biological Reference Materials for Flow Cytometry
Lili Wang
Biosytems and Biomaterials Division

## Slide 2: Understanding and Managing Sources of Variability in Measurement Process
Cell Gene Therapy Insights 2016;2(6),663-673.

## Slide 3: Use of Reference Materials for Quantitative Flow Cytometry
Cell processing/staining
Flow cytometer QC/calibration
Quantitative measurement
Sample Collection
Fixed vs. fresh
Anticoagulant
Cell count& Viability
Cell debris detection
Antibody quality
Cell debris detection
Linearity/sensitivity/resolution
Instrument settings
Sources of variability
QC & calibration beads
Compensation bead/
reference cell controls
Reference cell materials
Reference materials
controls
Flow cytometry compensation
Cell reference
materials
Data analysis/reporting
Reference cell/sample FMO controls
Linearity scale/range
Labeling fluorophore choice
Reference cell materials
Population gating
Notes: The major point is that reference materials can be used for a lot of things, consider what evidence they are providing or what aspect of measurement they are being used for.

## Slide 4
Benchmarking Flow Cytometric Scale
Cytometry Part A ● 73A: 279-288, 2008; Flow Cytometry Protocols: Third Edition, p53-65, 2011
Current Protocols in Cytometry, 75:1.29.1-14, 2016; Flow Cytometry Protocols: Fourth Edition (in press)
Two step process:
Benchmark scale
Aim: Provides evidence of linear range/proportionality and resolution, provides evidence of comparability within experiment and between experiments on single instrument
Aim: Transforms fluorescence scale to ABC scale, provides reasonable instrument independent transferable scale
1. Establish linear range in fluorescence scale using beads assigned with “Equivalent number of Reference Fluorophores (ERF) values
2. Anchor the fluorescence scale (FS) to a benchmark cell material with a known protein expression in the unit of Antibodies Bound per Cell (ABC)
Notes: This is a generalized scheme that enables clear understanding what the reference materials are giving you in terms of measurement assurance evidence. There are ways to do this in a one step process, but in that case more specialized ERF beads and more characterization would be required.

## Slide 5
Traceable ERF Value Assignment to Calibration Beads
Flow Cytometry Quantitation Consortium 81 Federal Register 136 (15 July 2016), pp. 46054-46055
ERF Value Assignment to Cytometer Calibration Beads Submitted by Consortium Members
Light obscuration and flow cytometry methods
for accurate microsphere concentration, fraction of fluorescence signal in single spheres
Different manufacturers’
relative intensity beads
NIST SRM 1934
Fluorescein
Nile red
Allophycocyanin (APC)
Coumarin 30
Calibrated fluorometer
Calibrated fluorometer
Emission of RF
Emission of bead
ERF/bead

## Slide 6
Traceable ERF Value Assignment to Commercial Microparticles
FC Bead | Ex Laser (nm) | SRM 1934
FITC | 488 | Fluorescein
PE | 488 | Fluorescein
BB515 | 488 | Fluorescein
PerCP | 488 | Nile Red
PerCP-Cy5.5 | 488 | Nile Red
PE-Cy7 | 488 | Nile Red
APC | 633 | APC
APC-R700 | 633 | APC
APC-H7 | 633 | APC
APC-Cy7 | 633 | APC
V450 | 405 | Coumarin 30
BV421 | 405 | Coumarin 30
V500-C | 405 | Coumarin 30
BV510 | 405 | Coumarin 30
BV605 | 405 | Coumarin 30
Six Peak Hard Dyed Micro-particles | Ex Laser (nm) | SRM 1934
Intensity 2-6 | 488 | Fluorescein
Intensity 2-6 | 488 | Nile Red
Intensity 2-6 | 633 | APC
Intensity 2-6 | 405 | Coumarin 30

## Slide 7
Flow Cytometer Performance Characterization
Cytometry 33A:260-266, 1998
Cytometry 33A:267-279, 1998
Current Protocols in Cytometry, 40:1.20.1-1.20.18, 2007
Single Cell Analysis (edited book), p171-199, 2017
Flow Cytometry Protocols: Fourth Edition (in press)

## Slide 8
Cellular Assays on Flow Cytometry Devices Cleared by CDRH
Purpose | Measurand | ‘quantitative’, ‘qualitative’ or ‘semi-quantitative’ | Standard used
Assessment of CMV-specific immune status and risk of CMV reactivation in immunosuppressed stem cell transplant recipients. | CMV specific CD8 MHC tetramer or dextramers | Quantitative | Beads
Immunologic assessment of patients having or suspected of having immune deficiency. | CD3+, CD3+CD4+, CD3+CD8+ | Quantitative and Qualitative | Beads
monitor forms of immunodeficiency | CD3+CD4+, CD3+CD8+ | Quantitative | Beads
Immunologic assessment of patients having or suspected of having immune deficiency. | CD3+CD4+, CD3+CD8+, CD3+, CD19+ and CD3-CD56+ | Quantitative | Beads or on a dual platform
Immunologic assessment of patients having, or suspected of having, immune deficiency. | CD3+, CD3+CD4+, CD3+CD8+, CD3-CD19+, CD3-CD56+ and/or CD16+, CD45+ Low SS, and CD45+ | Qualitative and quantitative | Beads or on a dual platform
HIV positive patients | CD4/Hgb | Quantitative | CD4 count/volume of the imaged field of views
multiparameter immunophenotyping aid in the differential diagnosis of hematopoietic neoplasms | T1: CD2, CD56, CD7, CD5, CD45 T2: CD8, CD4, CD3, CD45 B1: Kappa, Lambda, CD19, CD5, CD45 B2: CD20, CD10, CD19, CD38, CD45 M: CD7, CD13, CD34, CD33, CD45 | Qualitative | None

## Slide 9
Cytometry Part A ● 87A: 244-253, 2015
Cytometry Part A ● 87A: 254-261, 2015
CD4+ Cell Counting Standard
(WHO BS/10.2153)
(http://www.nibsc.org/documents/ifu/SS-319-20.pdf)
The first international reference standard for CD4+ cell counting for HIV/AIDS diagnosis and restricting access to anti-retroviral treatment (ART).
A round-robin international comparative study has demonstrated equivalence in capability in flow cytometric measurement of CD4+ cell concentrations and documented the associated measurement uncertainty.
Mean CD4+ lymphocyte concentration:
300.4 ± 28.3 µL-1 for 1 to 5 dilution
302.9 ± 28.3 µL-1 for 1 to 20 dilution

## Slide 10
CD34+ Cell Counting Reference Material
CD45+ cells → CD34+ cells → CD45dim cells → Blast morphology
Reference lymphocyte plot
Internal reference counting beads
http://www.nibsc.org/science_and_research/biotherapeutics/flow_cytometry_reference_materials.aspx

## Slide 11
Source of Variability for Cell Counting and Expression Analysis Using Flow Cytometry
Source of
Variability
Sample Collection
Antibody Reagents
Data Analysis
Flow Cytometer Platforms
TruCount Bead Standard
Different
Clones and
Binding Affinities
Different
Manufacturers
Reagent
Titration
Tube-to-Tube
Variability
Anticoagulant
Cytometer
Compensation
Linearity/sensitivity/
resolution calibration
Threshold and
PMT setting
Population
Gating
Volume Calibration
Dot or Density
Plots
Cell count and
viability
Cell debris
Effect of
Temperature
Number of Events
Collected
Channel for
Bead Detection
Quantitation
Choice of reference
cell materials
Single tube/separate
tube assay

## Slide 12: Need for Biological Cell Reference Materials
Stabilized Samples Diagnostic for ALL, AML, CLL,
MM, and MRD CVID: B cell panel; better standardization of memory B cells (both switched and unswitched). IgG expressing B cells - both percentage and expression. PNH (Paroxysmal Nocturnal Hemoglobinuria), CAR T Cell Manufacturing

## Slide 13
Obtaining ABC (antibodies bound per cell) values for CD4 receptors on
T-cells:
    Flow cytometry using QuantiBrite PE beads and unimolar CD4
    PE conjugate
    ii) Mass cytometry (CyTOFTM)
    iii) Single photon counting based technique (under development)
Quantifying CD4 receptor density using multiple reaction monitoring mass
spectrometry (MRM MS) combined with N15, C13 labeled recombinant CD4
protein as an internal standard
Quantifying RNA of CD4 per cell (under development)
Characterization of Human CD4+ Lymphocyte
Candidate Reference Materials
Cytometry Part A ● 81A: 567-575, 2012;
Analytical Chemistry ● 85: 1773-1777, 2013;
Clinical Proteomics ● 11: 43, 2014

## Slide 14
Procedure for Obtaining ABC Values of Biomarkers Based on Human PBMC Reference Controls
Cyto-Trol showing a consistent ABC
value of CD4 reference marker on T
cells using quantitative flow cytometry (41000, CV of 3.1%), CyTOF (38000, CV of 4.9%), and quantitative mass spectrometry (42500, CV of 12.8%)
Current Protocols in Cytometry, 75:1.29.1-14, 2016
Cytometry Part B (Clinical Cytometry) 90B:159-167, 2016
Flow Cytometry Protocols: Fourth Edition (in press)

## Slide 15
Quantifying CD20 Expression Based on CD4 Reference

## Slide 16
Biolegend’s Veri-CellsTM PBMC
BD Biosciences’ PBMC Dried Down CTT
Beckman Coulter’s Cyto-Trol Control CellsTM
CYTO 2016 Workshop on Biological
Cell Reference Materials
NIST, FDA, and a company are jointly characterizing these human blood cell preparations using conventional flow cytometry and CyTOF.
A round-robin study afterward with user communities to test their utilities/usefulness. The study allows for assay protocol development and serves a training purpose.

## Slide 17
Characterization of Human CD4+ and CD19+ Lymphocytes from Different Cell Preparations as Candidate Biological Cell Reference Materials for Quantitative Multiparameter Flow Cytometry and CyTOF
Three commercially available human PBMCs:
Veri-Cells PBMC
Cyto-Trol Control Cells
PBMC Dried Down CTT
Activity by using flow cytometry and CyTOF:
Quantifying CD4 expression on T lymphocytes
Quantifying CD19 expression on B lymphocytes
Characterize PBMC cell subsets
Characterize differential states of T lymphocytes
Apply an agreed gating strategy and analyze data
Anonymized whole blood samples

## Slide 18: Building Measurement Assurance with the Use of Reference Materials
Cell processing/staining
Flow cytometer QC/calibration
Quantitative measurement
Sample Collection
Fixed vs. fresh
Anticoagulant
Cell count& Viability
Cell debris detection
Antibody quality
Cell debris detection
Linearity/sensitivity/resolution
Instrument settings
Sources of variability
QC & calibration beads
Compensation bead/
reference cell controls
Reference cell materials
Reference materials
controls
Flow cytometry compensation
Cell reference
materials
Data analysis/reporting
Reference cell/sample FMO controls
Linearity scale/range
Labeling fluorophore choice
Reference cell materials
Population gating
Notes: The major point is that reference materials can be used for a lot of things, consider what evidence they are providing or what aspect of measurement they are being used for.

## Slide 19
Acknowledgements
NIST: P. DeRose, E. Stein, A.K. Gaigalas, H. He, K.D. Cole, M.
Misakian, P. Bajcsy, M. Wang, I.V. Turko, A. Plant
: : R.A. Hoffman, M. Bigos, J.P. Nolan and the Standardization
Task Force and Study Participants
: H. Degheidy, H. Mostowski, F. Abbasi, G.E. Marti and S. Bauer
Univ. of Toronto/DVS Sciences/Fluidigm: O. Ornatsky and S. Tanner
Flow Cytometry Quantitation Consortium:
CCQM Comparative Study Participants: 20 organizations across 10 countries