## Software V&V
Disclaimer
These spreadsheets are provided as a useful approach for metrology calculations according to the SOP represented here. However, NIST OWM or any other organization involved in their development, collectively and individually, do not warrant this spreadsheets for any specific purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each user agrees to decide if, when and how to use the spreadsheets, does so at his or her sole risk, and is responsible to suitable and applicable verification prior to use. When using the tools provided on the NIST OWM website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST OWM, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST OWM, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST WMD website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## Revision Control
Revision Control
6/?/2015 | V. Miller | This file was created to complete the uncertainty analysis of the results produced by the NIST Mass Code.
2018-08-22 | V. Miller | The file was modified to correct the calculations of the errors caused by the variability in the sensitivity factors during measurements. The equations that calculate the impact were modified to correct sensitivity values <> 1 to an equivalent value near unity.

## Instructions
The uncertainty tab contains calculations that are believed to provide a good estimate of the uncertainty of mass disseminations using weighing designs. The yellow cells in the Blue Restraint section (row 3) have comments describing the source of the information. The headings of all other columns are similarly commented to guide in locating the information.
The yellow cells are unprotected and provide input of values that may contribute to uncertainty in the measurement.
The information entered at the top of the page for the restraint propogates proportionately to all other nominal values in the table.
The uncertainty of the restraint calculated in cell O3 is the value that would be entered in the Mass Code data file for the uncertainty of the starting restraint. This will differ from the uncertainty from the calibration certificate because of the additional components related to the restraint that are included from the conditions at the time of use as the restraint.
Each line of data relates to a specific nominal value, though the nominal column could be edited to fit a specific situation.
Some values to be entered must come from the Mass Code output report, thus the calcuation of uncertainty must be iterative to utilize data from the measurements, unless average values are established by laboratory policy that set bounds limiting the range of values within specific identified, documented criteria.
The tab containing the MPE table is provided as a resource but must be validated.
The Uncertainty Calculations and MPE Table tabs are protected using the password 'metrology'.

## Uncertainty Calculations
Volume of Restraint | Nominal of Restraint | Cal Cert U of Starting Restraint (k=1) | Sensitivity used for Series 1 | Maximum Sensitivity from Series 1 | Minimum Sensitivity from Series 1 | Max A(I) MC Data Series 1 | Min A(I) MC Data Series 1 | Total U (Restraint) to MC | U²(ABC Restraint) (Output from C.6.3.1) | Uncertainty due to u(CCE) of Restraint (10 % @ 2 °C) | Rho(R) | u(Rho(R)) | u(t) | u(RH) | u(P) |  | Rho(a) Series 1 (establishes ABC of Restraint) | Rho(a) at Restraint Cal | Max Deviation of T from 20 °C | CCE of Restraint
(cm³) | (g) | (mg) | (mg/div) | (mg/div) | (mg/div) | (mg) | (mg) | (mg) | (mg²) | (mg) | (g/cm³) | (g/cm³) | (°C) | (% RH) | (mmHg) |  | (g/cm³) | (g/cm³) | (°C) | ( /°C)
250 | 2000 | 0.04654 | 1 | 1 | 1 | 1 | 1 | 0.04654 | 0.0041963 | 1.4827e-09 | 8 | 0.025 | 0.2 | 1 | 0.11 |  | 0.0011414 | 0.0012 | 2 | 4.5e-05
|  | From MC report | [10 empty] | From MC report
Unknown (X) |  | Restraint |  |  |  |  | Unknown (X) | [19 empty] | % contribution to uncertainty
Nominal (g) | Tolerance (mg) ASTM 0 for this exercise | Type B for the restraint Calculated from Cell O3 | Type A of series | Type B due to series 1 sensitivity error | Mass Code Output Uncertainty (k=1) Includes Type A and B of process, restraint and sensitivity | u²(ABC) from restraint (Proportional) (mg²) | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Average Sensitivity for Series X (mg/div) | Maximum Sensitivity from Series X (mg/div) | Minimum Sensitivity from Series X (mg/div) | Mass error from Sensitivity error (mg) | Assumed/Reported Density Rho(x) (g/cm³) | uRho x (g/cm³) (k=1) | u(t) (°C) (k=1) | u(RH) (% RH) (k=1) | u(P) (mmHg) (k=1) | Rho(a) during this cal (g/cm³) | Reference Air density for Stds cal | Uncertainty of Mole Fraction of CO₂ | CCE (/°C) | Uncertainty due to u(CCE) (10 % @ 2 °C) (mg) | u(c) (k=1) (mg) | k | Uk=2 (mg) | Pn | Nominal (g) | Total | Type B for the restraint Calculated from Cell N3 | Type A of series | u²(ABC) from restraint (Proportional) (mg²) | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Mass error from Sensitivity error (mg) | Uncertainty due to u(CCE) (10 % @ °C) (mg)
30000 | 38 | 0.6981 | 0.145 | 0 | 0.713 | 0.062944 | 4.1732 | 1 | 1 | 1 | 0 | 8 | 0.07 | 2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 2.224e-08 | 2.1782 | 2 | 4.4 | 0.34737 | 30000 | 100% | 10.27% | 0.4431% | 1.327% | 87.96% | 0% | 1.043e-14%
20000 | 25 | 0.4654 | 0.118 | 0 | 0.48013 | 0.041963 | 1.8548 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-08 | 1.4585 | 2 | 2.9 | 0.348 | 20000 | 100% | 10.18% | 0.6546% | 1.973% | 87.19% | 0% | 1.033e-14%
10000 | 13 | 0.2327 | 0.094 | 0 | 0.25097 | 0.020981 | 0.46369 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 7.4134e-09 | 0.74004 | 2 | 1.5 | 0.34615 | 10000 | 100% | 9.887% | 1.613% | 3.831% | 84.67% | 0% | 1.004e-14%
5000 | 6 | 0.11635 | 0.112 | 0 | 0.1615 | 0.010491 | 0.11592 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 3.7067e-09 | 0.39051 | 2 | 0.78 | 0.39 | 5000 | 100% | 8.877% | 8.226% | 6.879% | 76.02% | 0% | 9.01e-15%
3000 | 3.8 | 0.06981 | 0.109 | 0 | 0.12944 | 0.0062944 | 0.041732 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 2.224e-09 | 0.25452 | 2 | 0.51 | 0.40263 | 3000 | 100% | 7.523% | 18.34% | 9.716% | 64.42% | 0% | 7.635e-15%
2000 | 2.5 | 0.04654 | 0.112 | 0 | 0.12128 | 0.0041963 | 0.018548 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-09 | 0.19353 | 2 | 0.39 | 0.468 | 2000 | 100% | 5.783% | 33.49% | 11.2% | 49.52% | 0% | 5.869e-15%
1000 | 1.3 | 0.02327 | 0.01395 | 0 | 0.027131 | 0.0020981 | 0.0046369 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 7.4134e-10 | 0.086436 | 2 | 0.17 | 0.39231 | 1000 | 100% | 7.248% | 2.605% | 28.08% | 62.06% | 0% | 7.356e-15%
500 | 0.6 | 0.011635 | 0.01458 | 0 | 0.018653 | 0.0010491 | 0.0011592 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 3.7067e-10 | 0.050559 | 2 | 0.10 | 0.5 | 500 | 100% | 5.296% | 8.316% | 41.04% | 45.35% | 0% | 5.375e-15%
300 | 0.38 | 0.006981 | 0.01272 | 0 | 0.01451 | 0.00062944 | 0.00041732 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 2.224e-10 | 0.035458 | 2 | 0.071 | 0.56053 | 300 | 100% | 3.876% | 12.87% | 50.06% | 33.19% | 0% | 3.934e-15%
200 | 0.25 | 0.004654 | 0.01042 | 0 | 0.011412 | 0.00041963 | 0.00018548 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-10 | 0.027117 | 2 | 0.054 | 0.648 | 200 | 100% | 2.946% | 14.77% | 57.07% | 25.22% | 0% | 2.99e-15%
100 | 0.13 | 0.002327 | 0.01266 | 0 | 0.012872 | 0.00020981 | 4.6369e-05 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 7.4134e-11 | 0.02054 | 2 | 0.041 | 0.94615 | 100 | 100% | 1.284% | 37.99% | 49.73% | 10.99% | 0% | 1.303e-15%
50 | 0.06 | 0.0011635 | 0.00735 | 0 | 0.0074415 | 0.00010491 | 1.1592e-05 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 3.7067e-11 | 0.01311 | 2 | 0.026 | 1.3 | 50 | 100% | 0.7876% | 31.43% | 61.04% | 6.745% | 0% | 7.994e-16%
20 | 0.037 | 0.0004654 | 0.0061 | 0 | 0.0061177 | 4.1963e-05 | 1.8548e-06 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-11 | 0.0090136 | 2 | 0.018 | 1.4595 | 20 | 100% | 0.2666% | 45.8% | 51.65% | 2.283% | 0% | 2.706e-16%
20 | 0.037 | 0.0004654 | 0.00492 | 0 | 0.004942 | 4.1963e-05 | 1.8548e-06 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-11 | 0.0082608 | 2 | 0.017 | 1.3784 | 20 | 100% | 0.3174% | 35.47% | 61.49% | 2.718% | 0% | 3.221e-16%
10 | 0.025 | 0.0002327 | 0.00582 | 0 | 0.0058247 | 2.0981e-05 | 4.6369e-07 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 7.4134e-12 | 0.0074412 | 2 | 0.015 | 1.8 | 10 | 100% | 0.09779% | 61.17% | 37.89% | 0.8374% | 0% | 9.925e-17%
5 | 0.017 | 0.00011635 | 0.00295 | 0 | 0.0029523 | 1.0491e-05 | 1.1592e-07 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 3.7067e-12 | 0.0043958 | 2 | 0.0088 | 1.5529 | 5 | 100% | 0.07006% | 45.04% | 54.29% | 0.5999% | 0% | 7.111e-17%
3 | 0.017 | 6.981e-05 | 0.00185 | 0 | 0.0018513 | 6.2944e-06 | 4.1732e-08 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 2.224e-12 | 0.0031247 | 2 | 0.0062 | 1.0941 | 3 | 100% | 0.04991% | 35.05% | 64.47% | 0.4274% | 0% | 5.066e-17%
2 | 0.017 | 4.654e-05 | 0.00128 | 0 | 0.0012808 | 4.1963e-06 | 1.8548e-08 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 1.4827e-12 | 0.0024198 | 2 | 0.0048 | 0.84706 | 2 | 100% | 0.03699% | 27.98% | 71.67% | 0.3168% | 0% | 3.754e-17%
1 | 0.017 | 2.327e-05 | 0.00092 | 0 | 0.00092029 | 2.0981e-06 | 4.6369e-09 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 4.5e-05 | 7.4134e-13 | 0.0017175 | 2 | 0.0034 | 0.6 | 1 | 100% | 0.01836% | 28.69% | 71.13% | 0.1572% | 0% | 1.863e-17%
0.5 | 0.005 | 1.1635e-05 | 0.00048 | 0 | 0.00048014 | 1.0491e-06 | 1.1592e-09 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 3.2125e-13 | 0.0011317 | 2 | 0.0023 | 1.38 | 0.5 | 100% | 0.01057% | 17.99% | 81.91% | 0.09051% | 0% | 8.058e-18%
0.3 | 0.005 | 6.981e-06 | 0.00032 | 0 | 0.00032008 | 6.2944e-07 | 4.1732e-10 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 1.9275e-13 | 0.00085575 | 2 | 0.0017 | 1.02 | 0.3 | 100% | 0.006655% | 13.98% | 85.95% | 0.05699% | 0% | 5.073e-18%
0.2 | 0.005 | 4.654e-06 | 0.00024 | 0 | 0.00024005 | 4.1963e-07 | 1.8548e-10 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 1.285e-13 | 0.00069097 | 2 | 0.0014 | 0.84 | 0.2 | 100% | 0.004537% | 12.06% | 87.89% | 0.03885% | 0% | 3.458e-18%
0.1 | 0.005 | 2.327e-06 | 0.00022 | 0 | 0.00022001 | 2.0981e-07 | 4.6369e-11 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 6.425e-14 | 0.0005082 | 2 | 0.0010 | 0.6 | 0.1 | 100% | 0.002097% | 18.74% | 81.24% | 0.01795% | 0% | 1.598e-18%
0.05 | 0.005 | 1.1635e-06 | 0.00017 | 0 | 0.00017 | 1.0491e-07 | 1.1592e-11 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 3.2125e-14 | 0.00036581 | 2 | 0.00073 | 0.438 | 0.05 | 100% | 0.001012% | 21.6% | 78.39% | 0.008663% | 0% | 7.712e-19%
0.03 | 0.005 | 6.981e-07 | 0.00018 | 0 | 0.00018 | 6.2944e-08 | 4.1732e-12 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 1.9275e-14 | 0.00030879 | 2 | 0.00062 | 0.372 | 0.03 | 100% | 0.0005111% | 33.98% | 66.01% | 0.004377% | 0% | 3.896e-19%
0.02 | 0.005 | 4.654e-07 | 0.00015 | 0 | 0.00015 | 4.1963e-08 | 1.8548e-12 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 1.285e-14 | 0.0002539 | 2 | 0.00051 | 0.306 | 0.02 | 100% | 0.000336% | 34.9% | 65.09% | 0.002877% | 0% | 2.561e-19%
0.01 | 0.005 | 2.327e-07 | 0.0002 | 0 | 0.0002 | 2.0981e-08 | 4.6369e-13 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 6.425e-15 | 0.00024695 | 2 | 0.00049 | 0.294 | 0.01 | 100% | 8.88e-05% | 65.59% | 34.41% | 0.0007604% | 0% | 6.769e-20%
0.005 | 0.005 | 1.1635e-07 | 0.00016 | 0 | 0.00016 | 1.0491e-08 | 1.1592e-13 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 3.9e-05 | 3.2125e-15 | 0.00018998 | 2 | 0.00038 | 0.228 | 0.005 | 100% | 3.751e-05% | 70.93% | 29.07% | 0.0003212% | 0% | 2.859e-20%
0.003 | 0.005 | 6.981e-08 | 0.00018 | 0 | 0.00018 | 6.2944e-09 | 4.1732e-14 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 6.96e-05 | 3.4398e-15 | 0.00019671 | 2 | 0.00039 | 0.234 | 0.003 | 100% | 1.259e-05% | 83.73% | 16.27% | 0.0001079% | 0% | 3.058e-20%
0.002 | 0.005 | 4.654e-08 | 0.00015 | 0 | 0.00015 | 4.1963e-09 | 1.8548e-14 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 6.96e-05 | 2.2932e-15 | 0.00016339 | 2 | 0.00033 | 0.198 | 0.002 | 100% | 8.113e-06% | 84.28% | 15.72% | 6.948e-05% | 0% | 1.97e-20%
0.001 | 0.005 | 2.327e-08 | 0.0002 | 0 | 0.0002 | 2.0981e-09 | 4.6369e-15 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 0.0011414 | 0.0012 | 0 | 6.96e-05 | 1.1466e-15 | 0.00020518 | 2 | 0.00041 | 0.246 | 0.001 | 100% | 1.286e-06% | 95.02% | 4.984% | 1.101e-05% | 0% | 3.123e-21%
C.6.3-1 calcs
[20 empty] | Here | =((((C6)*((1-0.0012/8)/0.99985))*(8-H6)/(8*H6)*SQRT((0.00001*N6*L6/760*101325)^2+(0.000022*N6)^2+(-0.0034*N6*J6)^2+(-0.01*K6/100*N6)^2))^2+(((C6)*((1-0.0012/8)/0.99985))*(N6-0.0012)/H6)^2*(I6^2/H6^2)+(((C6)*((1-0.0012/8)/0.99985))^2*(N6-0.0012))*((N6-0.0012)-2*(M6-0.0012))*(0^2/8^4))*1000000
[20 empty] | Demo | =(((($D$13+$F$13)*((1-0.0012/$J$13)/0.99985))*($J$13-J14)/($J$13*J14)*SQRT((0.00001*$L$6*$I$4/760*101325)^2+(0.000022*$L$6)^2+(-0.0034*$L$6*I3)^2+(-0.01*$I$5/100*$L$6)^2))^2+((($D$13+$F$13)*((1-0.0012/$J$13)/0.99985))*($L$6-0.0012)/J14)^2*(L14^2/J14^2)+((($D$13+$F$13)*((1-0.0012/$J$13)/0.99985))^2*($L$6-0.0012))*(($L$6-0.0012)-2*($L$8-0.0012))*($L$13^2/$J$13^4))*1000000
[21 empty] | $D$13+$F$13 | Mass+Corr |  |  | C3 Nominal
[21 empty] | $J$13 | Den STD |  |  | 8 g/cm³
[21 empty] | J14 | Den X |  |  | L3
[21 empty] | $L$6 | Air Den now |  |  | M3
[21 empty] | $I$4 | U Press |  |  | I3
[21 empty] | I3 | U temp |  |  | G3
[21 empty] | $I$5 | U RH |  |  | H3
[21 empty] | L14 | U den X |  |  | L3
[21 empty] | $L$8 | Den Air at Std Cal |  |  | n3
|  |  | Uncertainty of CIPM 2007 air density equation Above coefficients (yellow highlighting) updated from R111 (2004) values to match CIPM 2007 air density paper. | [17 empty] | $L$13 | U den S |  |  | 0 g/cm³
By using 8 as the density of the standard and 0 as the uncertainty of the density of the standard I believe we can calculate the C.6.3-1 ABC uncertainty of the starting restraints at the time of use.
The one question I have is whether the restraint is 1 kg or 2 kg. I am guessing it is 2 kg for these purposes.
Note on A3: "Volume of Restraint" is the sum of all artifact volumes in the restraint.
Note on B3: The "Nominal of Restraint" is the combined nominal of the restraint, as established in the design.
Note on D3: "Cal Cert U of Starting Restraint" is typically the sum of the k=1 uncertainties of the restraint, though in rare cases the RSS value may be appropriate.
Note on E3: The "Sensitivity used for Series 1" will typically be the average sensitivity for Series 1 from the Mass Code report, but when no sensitivity weight is used, e.g., automated comparators, will be the sensitivity value entered for the comparator used.
Note on F3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on G3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on H3: From Mass Code report for Series 1.
Note on I3: From Mass Code report for Series 1.
Note on M3: Effective density of the restraint.
Note on N3: Uncertainty of the density of the restraint.
Note on O3: Uncertainty of the ability to measure the air temperature. Though not necessarily correct because of the difficulties accurately measuring air temperature, this value is typically the uncertainty of the thermometer used for that purpose during Series 1.
Note on P3: The uncertianty of the Relative Humidity measurement during Series 1.
Note on Q3: Uncertainty of the barometer used to measure the barometric pressure. This assumes that there is no elevation offset in the pressure measurement; the barometer is at the height of the measurements performed.
Note on S3: Enter the actual calculated air density present during Series 1.
Note on T3: This value comes from the calibration certificate for the restraint artifacts. If the restraints were not calibrated together so that the pressures were equal, use the BP from the certificates that deviates the most from Normal air density.
Note on U3: This deviation value is based on the temperature of the mass standards during Series 1 measurements.
Note on V3: Based on the material of the restraint. And effective CCE may be used if materials differ.
Note on B7: Enter the tolerance for the weight set being calibrated.
Note on D7: This uncertainty component is the value of the uncertainty of the restraint applied proportionately to the nominal value of the unknown. HINT: For nomnal X values less than the restraint, the value will be proportinately less than the uncertainty of the restraint. For nominal X values larger than the restraint, this value will be proportionaltely larger than the uncertainty of the restraint.
Note on E7: The Type A values will come from the mass code report for each nomnal mass.
Note on F7: This component will typically be Zero unless there is a known uncorrected sensitivity error caused by the use of a constant sensitivity factor when no sensitivity weight is used.
Note on G7: This value is the RSS of columns D, E, F for each nominal X value.
Note on H7: The U²(ABC) is calculated as a variance because under some conditions it may actually decrease the uncertainty of the restraint slightly and would thus be a negative value. As the square root of a negative number is an Imaginary number, the only way to utilize this component in normal mathematical equations is to keep it as a variance. This component is also distributed proportionately to the X values.
Note on I7: The U²(ABC) is calculated as a variance because under some conditions it may actually decrease the uncertainty of the measurent result slightly and would thus be a negative value. As the square root of a negative number is an Imaginary number, the only way to utilize this component in normal mathematical equations is to keep it as a variance. It will be combined with variances of the rest of the uncertainty components associated with X.
Note on J7: The average sensitivity for each X will come from the mass code report and will be the calculated average of all the MC calculated sensitivity values of each series. It may be different than the average sensitivity used for the mass code solutions.
Note on K7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on L7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on M7: The calculated sensitivity error uncertainty contribution for X. Calculations are based on the deviation of the sensitivity during the series from the average, multiplied by the difference limits set by the tolerance.
Note on N7: The density value for X entered in the input file for the comparisons.
Note on O7: The estimated uncertainty of the entered density value for X. The default 0.07 g/cm³ value is the k=1 portion of the density uncertainty given in OIML R111. If the density is known to a better uncertainty use that uncertiainty value. Most U.S. weight manufacturers believe that the uncertainty of thier assumed density is 0.025 g/cm³ (k=1)
Note on P7: The uncertainty of the air temperature in which the measurements are performed. Though not completely correct because of the difficulty of measuring air temperature, the value used is tyically the calibration uncertainty of the thermometer used to measure air temperature.
Note on Q7: Unertainty of the relative humidity measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the hygrometer.
Note on R7: Unertainty of the barometric pressure measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the barometer.
Note on S7: Enter the value of the actual calculated air density during the series for that nominal value.
Note on T7: This value will be the air density present during the comparisons of the previous series in the overall design, as the mass value of the standards used in this series come from the previous series.
Note on U7: The effect of this component on the mass uncertainty is negligible, but is shown here to demonstrate that the effect was considered. Follow the guidance for this component given in the 2007 Air Density paper.
Note on V7: Enter the CCE provided the mass code for each nominal X value.
Note on W7: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6.
Note on X7: This column shows the calculated k=1 combined uncertianty for all components included in this uncertainty budget. NOTE: Variances are already squared and are included as shown.
Note on Y7: Enter the appropriate k value based on the degrees of freedom associated with this nomal X value.
Note on Z7: k=2 is shown in the heading but each u(c) is expanded by multiplying the entered k value for that nominal X . (There is no way to easily show the actual k value for each nominal.)
Note on AA7: The Pn value is calculated based on the tolerance entered in column B. If the Pn cell background is RED, the uncertainty exceeds the Tolerance/3 limit.
Note on AB7: The Nominal value is shown here for convenience.
Note on T8: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T9: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T10: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T11: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T12: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T13: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T14: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T15: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T16: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T17: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T18: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T19: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T20: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T21: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T22: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T23: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T24: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T25: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T26: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T27: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T28: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T29: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T30: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T31: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T32: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T33: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T34: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T35: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T36: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T37: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.
Note on T38: This value will typically be left as 0.0012 g/cm³ (normal air) due to the output of the mass code being corrected to Conventional conditions.

## MPE Table
Maximum Permissible Errors for Weights (+/- in mg)
|  |  | OIML | [8 empty] | ASTM | [9 empty] | NIST | User Tolerances
Nominal | Units | Nominal (g) | E1 | E2 | F1 | F2 | M1 | M1-2 | M2 | M2-3 | M3 | 000 | 00 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | F | Other 1 | Other 2 | Other 3 | Other 4 | Other 5 | Not Specified
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
5 000 | kg | 5000000 | 3: 25000; 4: 80000; 5: 250000; 6: 500000; 7: 800000; 8: 1600000; 9: 2500000; 14: 25000; 15: 50000; 16: 100000; 17: 250000; 18: 500000; 19: 750000; 20: 500000
3 000 | kg | 3000000 | 14: 15000; 15: 30000; 16: 60000; 17: 150000; 18: 300000; 19: 450000; 20: 300000
2 000 | kg | 2000000 | 3: 10000; 4: 30000; 5: 100000; 6: 200000; 7: 300000; 8: 600000; 9: 1000000; 14: 10000; 15: 20000; 16: 40000; 17: 100000; 18: 200000; 19: 300000; 20: 200000
1 000 | kg | 1000000 | 2: 1600; 3: 5000; 4: 16000; 5: 50000; 6: 100000; 7: 160000; 8: 300000; 9: 500000; 14: 5000; 15: 10000; 16: 20000; 17: 50000; 18: 100000; 19: 150000; 20: 100000
500 | kg | 500000 | 2: 800; 3: 2500; 4: 8000; 5: 25000; 6: 50000; 7: 80000; 8: 160000; 9: 250000; 14: 2500; 15: 5000; 16: 10000; 17: 25000; 18: 50000; 19: 75000; 20: 50000
300 | kg | 300000 | 14: 1500; 15: 3000; 16: 6000; 17: 15000; 18: 30000; 19: 45000; 20: 30000
200 | kg | 200000 | 2: 300; 3: 1000; 4: 3000; 5: 10000; 6: 20000; 7: 30000; 8: 60000; 9: 100000; 14: 1000; 15: 2000; 16: 4000; 17: 10000; 18: 20000; 19: 30000; 20: 20000
100 | kg | 100000 | 2: 160; 3: 500; 4: 1600; 5: 5000; 6: 10000; 7: 16000; 8: 30000; 9: 50000; 14: 500; 15: 1000; 16: 2000; 17: 5000; 18: 10000; 19: 15000; 20: 10000
50 | kg | 50000 | 25 | 80 | 250 | 800 | 2500 | 5000 | 8000 | 16000 | 25000 | 13 | 25 | 63 | 125 | 250 | 500 | 1000 | 2500 | 5000 | 7500 | 5000
30 | kg | 30000 | 10: 7.5; 11: 15; 12: 38; 13: 75; 14: 150; 15: 300; 16: 600; 17: 1500; 18: 3000; 19: 4500; 20: 3000
25 | kg | 25000 | 10: 6.3; 11: 13; 12: 31; 13: 62; 14: 125; 15: 250; 16: 500; 17: 1200; 18: 2500; 19: 3800; 20: 2500
20 | kg | 20000 | 10 | 30 | 100 | 300 | 1000 |  | 3000 |  | 10000 | 5 | 10 | 25 | 50 | 100 | 200 | 400 | 1000 | 2000 | 3800 | 2000
10 | kg | 10000 | 5 | 16 | 50 | 160 | 500 |  | 1600 |  | 5000 | 2.5 | 5 | 13 | 25 | 50 | 100 | 200 | 500 | 1000 | 2200 | 1000
5 | kg | 5000 | 2.5 | 8 | 25 | 80 | 250 |  | 800 |  | 2500 | 1.3 | 2.5 | 6 | 12 | 25 | 50 | 100 | 250 | 500 | 1400 | 500
3 | kg | 3000 | 10: 0.75; 11: 1.5; 12: 3.8; 13: 7.5; 14: 15; 15: 30; 16: 60; 17: 150; 18: 300; 19: 1000; 20: 300
2 | kg | 2000 | 1 | 3 | 10 | 30 | 100 |  | 300 |  | 1000 | 0.5 | 1 | 2.5 | 5 | 10 | 20 | 40 | 100 | 200 | 750 | 200
1 | kg | 1000 | 0.5 | 1.6 | 5 | 16 | 50 |  | 160 |  | 500 | 0.25 | 0.5 | 1.3 | 2.5 | 5 | 10 | 20 | 50 | 100 | 470 | 100
500 | g | 500 | 0.25 | 0.8 | 2.5 | 8 | 25 |  | 80 |  | 250 | 0.13 | 0.25 | 0.6 | 1.2 | 2.5 | 5 | 10 | 30 | 50 | 300 | 70
300 | g | 300 | 10: 0.075; 11: 0.15; 12: 0.38; 13: 0.75; 14: 1.5; 15: 3; 16: 6; 17: 20; 18: 30; 19: 210; 20: 60
200 | g | 200 | 0.1 | 0.3 | 1 | 3 | 10 |  | 30 |  | 100 | 0.05 | 0.1 | 0.25 | 0.5 | 1 | 2 | 4 | 15 | 20 | 160 | 40
100 | g | 100 | 0.05 | 0.16 | 0.5 | 1.6 | 5 |  | 16 |  | 50 | 0.025 | 0.05 | 0.13 | 0.25 | 0.5 | 1 | 2 | 9 | 10 | 100 | 20
50 | g | 50 | 0.03 | 0.1 | 0.3 | 1 | 3 |  | 10 |  | 30 | 0.015 | 0.03 | 0.06 | 0.12 | 0.25 | 0.6 | 1.2 | 5.6 | 7 | 62 | 10
30 | g | 30 | 10: 0.013; 11: 0.026; 12: 0.037; 13: 0.074; 14: 0.15; 15: 0.45; 16: 0.9; 17: 4; 18: 5; 19: 44; 20: 6
20 | g | 20 | 0.025 | 0.08 | 0.25 | 0.8 | 2.5 |  | 8 |  | 25 | 0.013 | 0.025 | 0.037 | 0.074 | 0.1 | 0.35 | 0.7 | 3 | 3 | 33 | 4
10 | g | 10 | 0.02 | 0.06 | 0.2 | 0.6 | 2 |  | 6 |  | 20 | 0.01 | 0.02 | 0.025 | 0.05 | 0.074 | 0.25 | 0.5 | 2 | 2 | 21 | 2
5 | g | 5 | 0.016 | 0.05 | 0.16 | 0.5 | 1.6 |  | 5 |  | 16 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.18 | 0.36 | 1.3 | 2 | 13 | 1.5
3 | g | 3 | 10: 0.005; 11: 0.01; 12: 0.017; 13: 0.034; 14: 0.054; 15: 0.15; 16: 0.3; 17: 0.95; 18: 2; 19: 9.4; 20: 1.3
2 | g | 2 | 0.012 | 0.04 | 0.12 | 0.4 | 1.2 |  | 4 |  | 12 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.13 | 0.26 | 0.75 | 2 | 7 | 1.1
1 | g | 1 | 0.01 | 0.03 | 0.1 | 0.3 | 1 |  | 3 |  | 10 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.1 | 0.2 | 0.5 | 2 | 4.5 | 0.9
500 | mg | 0.5 | 0.008 | 0.025 | 0.08 | 0.25 | 0.8 |  | 2.5 |  |  | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.08 | 0.16 | 0.38 | 1 | 3 | 0.72
300 | mg | 0.3 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.025; 15: 0.07; 16: 0.14; 17: 0.3; 18: 1; 19: 2.2; 20: 0.61
200 | mg | 0.2 | 0.006 | 0.02 | 0.06 | 0.2 | 0.6 |  | 2 |  |  | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.06 | 0.12 | 0.26 | 1 | 1.8 | 0.54
100 | mg | 0.1 | 0.005 | 0.016 | 0.05 | 0.16 | 0.5 |  | 1.6 |  |  | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.05 | 0.1 | 0.2 | 1 | 1.2 | 0.43
50 | mg | 0.05 | 1: 0.004; 2: 0.012; 3: 0.04; 4: 0.12; 5: 0.4; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.042; 16: 0.085; 17: 0.16; 19: 0.88; 20: 0.35
30 | mg | 0.03 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.038; 16: 0.075; 17: 0.14; 19: 0.68; 20: 0.3
20 | mg | 0.02 | 1: 0.003; 2: 0.01; 3: 0.03; 4: 0.1; 5: 0.3; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.035; 16: 0.07; 17: 0.12; 19: 0.56; 20: 0.26
10 | mg | 0.01 | 1: 0.003; 2: 0.008; 3: 0.025; 4: 0.08; 5: 0.25; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.03; 16: 0.06; 17: 0.1; 19: 0.4; 20: 0.21
5 | mg | 0.005 | 1: 0.003; 2: 0.006; 3: 0.02; 4: 0.06; 5: 0.2; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.028; 16: 0.055; 17: 0.08; 20: 0.17
3 | mg | 0.003 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.026; 16: 0.052; 17: 0.07; 20: 0.14
2 | mg | 0.002 | 1: 0.003; 2: 0.006; 3: 0.02; 4: 0.06; 5: 0.2; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.025; 16: 0.05; 17: 0.06; 20: 0.12
1 | mg | 0.001 | 1: 0.003; 2: 0.006; 3: 0.02; 4: 0.06; 5: 0.2; 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.025; 16: 0.05; 17: 0.05; 20: 0.1
500 | µg | 0.0005 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.025; 16: 0.05; 17: 0.05
300 | µg | 0.0003 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014; 15: 0.025
200 | µg | 0.0002 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01; 14: 0.014
100 | µg | 0.0001 | 10: 0.002; 11: 0.003; 12: 0.005; 13: 0.01
50 | µg | 5e-05 | 10: 0.002; 11: 0.003; 12: 0.005
The values in the above table are correct as of 20180806. Verified by comparison to current documentary standards.