## Software V&V
Disclaimer
These spreadsheets are provided as a useful approach for metrology calculations according to the SOP represented here. However, NIST OWM or any other organization involved in their development, collectively and individually, do not warrant this spreadsheets for any specific purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each user agrees to decide if, when and how to use the spreadsheets, does so at his or her sole risk, and is responsible to suitable and applicable verification prior to use. When using the tools provided on the NIST OWM website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST OWM, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST OWM, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST WMD website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## Revision Control
Revision Control
6/?/2015 | V. Miller | This file was created to complete the uncertainty analysis of the results produced by the NIST Mass Code.
2018-08-22 | V. Miller | The file was modified to correct the calculations of the errors caused by the variability in the sensitivity factors during measurements. The equations that calculate the impact were modified to correct sensitivity values <> 1 to an equivalent value near unity.
2020-03-31 | V. Miller | The air density values were changed to mg/cm³, with the associated edits required for the ub² equations. The mass code output files provide the air density values in mg/cm³. All equations for 'Uncertainty due to the uncertainty of CCE' were edited to result in the correct units. A tab was created for 5321 weight sets and one for 5221 weight sets so that it is not necessary to edit the various nominal values.
2020-04-09 | V. Miller | The uncertainty calculation in L3 was edited to provide the correct units. The user can now select from a dropdown in A5:B5 the tolerance against which the calculated uncertainties are to be evaluated. This required editing the heading of the MPE tables, entering a "-" for any tolerance cell having no MPE value, adding data validation to cell A5:B5, editing the formulas in B8:B38, applying an error function to any calculations based on the MPE value to display "-", and setting conditional formatting for the Pn value column to NOT turn red if the resulting U is "-". Added lines for one additional decade.
2020-04-17 | V. Miller | Added a column for entering the temperature deviation of each series and edited the u due to CCE of the test weight to use that temperature deviation in the calculation. Added some error checking and associated responses for the uncertainty calculations.
2020-04-28 | V. Miller | Converted many cells to Named ranges and edited many equations to use the named range to better document the equations. Added a tab for lb weights, and lb weight tolerances to the MPE table so that the lookup functions on the lb Uncertainty Calculations tab. Named ranges were set to be tab specific to prevent unintentional calculations carrying over between tabs.
2020-05-05 | V. Miller | Moved u² calculation for the restraint into the uncertainty of the restraint that gets passed to the unknowns and deleted the columns related to that component in the uncertainty of the unknowns section; edited all named ranges to fit the changes.
2020-05-18 | V. Miller | Created Sensitivity Calculations Sheet and added fields in which to enter sensitivity values to calculate the Mean of the Average Sensitivity values and the max and min Observed Sensitivity Values. The output of those three calculations can be copied and pasted into the fields for each series as appropriate.
2020-05-19 | V. Miller | Removed content of Column F which had contained a term that was a proportional distribution of the sensitivity error in Series 1, after realizing that it was already included in the uncertainty of the restraint that is being proportionately being passed down from Series 1 already. Thus that error was being double counted. After removing the content and editing the subseqent equation in Column G that content was moved to Column F and column G made blank, indicating the break between the Uncertainties coming from Series 1 and those associated directly with the unknown weights.
2020-05-26 | V. Miller | Created in X3 the calculation of u_Sens_R. Edited J3 to use that named range. Changed named range CalU_R to u_Cal_R to use lower case representing the k=1 uncertainty value of the restraint.
2020-05-27 | V. Miller | Entered an equation to set the CCE of unknowns based on the entered material density. Added note to column heading indicating that Brass cannot be used with the Mass Code Uncertainty Calculator. The CCE for Brass WILL NOT display correctly and errors will result. Brass at 8.3909 is in the density range of Nichrome and only the CCE of Nichrome will display. To allow for use of measured density values, values between 2.6 and 2.8 return the CCE of Aluminum(0.000069/°C), values between 7.8 and 8.05 return the CCE of SS(0.000045/°C), values between 8.3 and 8.6 return the CCE of Nichrome(0.000039/°C), and values between 16.6 and 16.8 return a CCE for Tantalum(0.00002/°C).
2020-06-01 | V. Miller | Added a line at each decade for calculation of uncertainties related to the check standard used in each series. While this may be an unnecessary calculation as this uncertainty will not be reported on a calibration certificate, it will provide an uncertainty value to be entered in associated control charts for the measurements, as part of the overall measurement uncertainty calculations.

## Instructions
The uncertainty tab contains calculations that are believed to provide a good estimate of the uncertainty of mass disseminations using weighing designs. The yellow cells in the Blue Restraint section (row 3) have comments describing the source of the information. The headings of all other columns are similarly commented to guide in locating the information.
The yellow cells are unprotected and provide input of values that may contribute to uncertainty in the measurement.
The information entered at the top of the page for the restraint propogates proportionately to all other nominal values in the table.
The uncertainty of the restraint calculated in cell O3 is the value that would be entered in the Mass Code data file for the uncertainty of the starting restraint. This will differ from the uncertainty from the calibration certificate because of the additional components related to the restraint that are included from the conditions at the time of use as the restraint.
Each line of data relates to a specific nominal value, though the nominal column could be edited to fit a specific situation. The worksheet tables are marked with lines to delineate the table by decades, which is the typical use of designs, to make it easier to follow the table across, but is not intended to indicate that decade based designs must be used. If alternative designs are used, the originator recommends that the tables be reformatted to reflect the designs used by editing the line placement. This is not a requirement but does clarify the designs used and makes data entry by 'series' more apparent to the casual observer.
Some values to be entered must come from the Mass Code output report, thus the calcuation of uncertainty must be iterative to utilize data from the measurements, unless average values are established by laboratory policy that set bounds limiting the range of values within specific identified, documented criteria.
The tab containing the MPE table is provided as a resource but must be validated.
The Uncertainty Calculations and MPE Table tabs are protected using the password 'metrology'.

## 5221 Uncertainty Calculations
Volume of Restraint | Nominal of Restraint | Cal Cert U of Starting Restraint (k=1) | Sensitivity used for Series 1 | Maximum Sensitivity from Series 1 | Minimum Sensitivity from Series 1 | Max A(I) MC Data Series 1 | Min A(I) MC Data Series 1 | Total U (Restraint) to MC | U²(ABC Restraint) (Output from C.6.3.1) | Uncertainty due to u(CCE) of Restraint (10 % @ 2.17 °C) | Rho(R) | u(Rho(R)) | u(t) | u(RH) | u(P) |  | Rho(a) Series 1 (establishes ABC of Restraint) | Rho(a) at Restraint Cal | Max Deviation of T from 20 °C | CCE of Restraint |  | Uncertainty due to Sensitivity Error of Series 1 | [9 empty] | Maximum Permissible Errors for Weights (+/- in mg)
(cm³) | (g) | (mg) | (mg/div) | (mg/div) | (mg/div) | (mg) | (mg) | (mg) | (mg²) | (mg) | (g/cm³) | (g/cm³) | (°C) | (% RH) | (mmHg) |  | (mg/cm³) | (mg/cm³) | (°C) | ( /°C) |  | (mg) | [12 empty] | OIML | [18 empty] | NIST | User Tolerances
250 | 2000 | 0.06 | 0.9999 | 1.0002 | 0.99971 | 0.46995 | -31.4419 | 0.065093 | 0.00059738 | 0.0016602 | 7.9924 | 0.025 | 0.2 | 1 | 0.11 |  | 1.1779 | 1.2 | 2.17 | 4.5e-05 |  | 0.00608 | [9 empty] | Nominal | Units | Nominal (g) | OIML E1 | OIML E2 | OIML F1 | OIML F2 | OIML M1 | OIML M1-2 | OIML M2 | OIML M2-3 | OIML M3 | ASTM 000 | ASTM 00 | ASTM 0 | ASTM 1 | ASTM 2 | ASTM 3 | ASTM 4 | ASTM 5 | ASTM 6 | ASTM 7 | NIST F | User 1 | User 2 | User 3 | User 4 | User 5 | Not Specified
Tolerance to be evaluated: | [34 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
ASTM 0 |  | From MC report | [10 empty] | From MC report
Unknown (X) |  | Sourced from Restraint |  |  |  | Applicable to Unknown (X) | [19 empty] | % contribution to uncertainty
Nominal (g) | MPE (mg) ASTM 0 for this exercise | Type B from the restraint Calculated from Cell J3 | Type A of series From Output Report | Mass Code Output Uncertainty (k=1) Includes Type A of this process and Type B from the restraint |  | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Average Sensitivity for Series X (mg/div) | Maximum Sensitivity from Series X (mg/div) | Minimum Sensitivity from Series X (mg/div) | Mass error from Sensitivity error of Series (mg) | Assumed/Reported Density Rho(x) (g/cm³) | uRho x (g/cm³) (k=1) | u(t) (°C) (k=1) | u(RH) (% RH) (k=1) | u(P) (mmHg) (k=1) | Rho(a) during this cal (mg/cm³) | Uncertainty of Mole Fraction of CO₂ | CCE of X (/°C) | Max Deviation of T for Series from 20 °C (°C) | Uncertainty due to u(CCE) of X (10 % @ 1.55 °C) (mg) | u(c) (k=1) (mg) | k | Uk=95.45 % CI (mg) | Pn | Nominal (g) | Total | Type B for the restraint Calculated from Cell N3 | Type A of series | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Mass error from Sensitivity error (mg) | Uncertainty due to u(CCE) (10 % @ °C) (mg)
20000 | 25 | 0.65093 |  | 0.65093 |  | 1.6449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 20000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
20000 | 25 | 0.65093 |  | 0.65093 |  | 1.6449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 20000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10000 | 13 | 0.32547 |  | 0.32547 |  | 0.41124 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 10000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10000 | 13 | 0.32547 |  | 0.32547 |  | 0.41124 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 10000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
5000 | 6 | 0.16273 |  | 0.16273 |  | 0.10281 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 5000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2000 | 2.5 | 0.065093 |  | 0.065093 |  | 0.016449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 2000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2000 | 2.5 | 0.065093 |  | 0.065093 |  | 0.016449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 2000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1000 | 1.3 | 0.032547 | 0.02535 | 0.041254 |  | 0.00054031 | 0.9999 | 1.0002 | 0.99971 | 0.00024768 | 7.9435 | 0.025 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 2.17 | 0.00081008 | 0.04736 | 2 | 0.095 | 0.21923 | 1000 | 100% | 47.23% | 28.65% | 24.09% | 0.002735% | 0.02926%
1000 | 1.3 | 0.032547 | 0.02535 | 0.041254 |  | 0.00054031 | 0.9999 | 1.0002 | 0.99971 | 0.00024768 | 7.9435 | 0.025 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 2.17 | 0.00081008 | 0.04736 | 2 | 0.095 | 0.21923 | 1000 | 100% | 47.23% | 28.65% | 24.09% | 0.002735% | 0.02926%
500 | 0.6 | 0.016273 | 0.0141 | 0.021532 |  | 2.4076e-05 | 0.99891 | 1.0015 | 0.99454 | 0.0015141 | 7.982 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00042644 | 0.02214 | 2 | 0.044 | 0.22 | 500 | 100% | 54.02% | 40.56% | 4.912% | 0.4677% | 0.0371%
200 | 0.25 | 0.0065093 | 0.01097 | 0.012756 |  | 3.8733e-06 | 0.99891 | 1.0015 | 0.99454 | 0.00063089 | 7.973 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00017077 | 0.012923 | 2 | 0.026 | 0.312 | 200 | 100% | 25.37% | 72.06% | 2.319% | 0.2383% | 0.01746%
200 | 0.25 | 0.0065093 | 0.00863 | 0.01081 |  | 3.8433e-06 | 0.99891 | 1.0015 | 0.99454 | 0.00063089 | 7.986 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00017049 | 0.011005 | 2 | 0.022 | 0.264 | 200 | 100% | 34.98% | 61.49% | 3.173% | 0.3286% | 0.024%
100 | 0.13 | 0.0032547 | 0.00974 | 0.010269 |  | 5.379e-06 | 0.99891 | 1.0015 | 0.99454 | 0.00032806 | 7.9517 | 0.025 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 2.23 | 8.3162e-05 | 0.010533 | 2 | 0.021 | 0.48462 | 100 | 100% | 9.547% | 85.5% | 4.848% | 0.097% | 0.006233%
100 | 0.13 | 0.0032547 | 0.00974 | 0.010269 |  | 9.8239e-07 | 0.99891 | 1.0015 | 0.99454 | 0.00032806 | 7.9517 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 8.5612e-05 | 0.010323 | 2 | 0.021 | 0.48462 | 100 | 100% | 9.941% | 89.03% | 0.9219% | 0.101% | 0.006878%
50 | 0.06 | 0.0016273 | 0.00501 | 0.0052677 |  | 2.7882e-07 | 1 | 1.0003 | 0.9998 | 8.6917e-06 | 7.969 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 4.3986e-05 | 0.0052943 | 2 | 0.011 | 0.55 | 50 | 100% | 9.448% | 89.55% | 0.9948% | 0.0002695% | 0.006903%
20 | 0.037 | 0.00065093 | 0.00315 | 0.0032166 |  | 4.4528e-08 | 1 | 1.0003 | 0.9998 | 5.3599e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 1.7588e-05 | 0.0032235 | 2 | 0.0064 | 0.51892 | 20 | 100% | 4.078% | 95.49% | 0.4285% | 0.0002765% | 0.002977%
20 | 0.037 | 0.00065093 | 0.00228 | 0.0023711 |  | 4.4315e-08 | 1 | 1.0003 | 0.9998 | 5.3599e-06 | 7.98 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 1.757e-05 | 0.0023805 | 2 | 0.0048 | 0.38919 | 20 | 100% | 7.477% | 91.73% | 0.782% | 0.000507% | 0.005448%
10 | 0.025 | 0.00032547 | 0.00175 | 0.00178 |  | 9.6893e-09 | 1 | 1.0003 | 0.9998 | 3.6216e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.3 | 8.8075e-06 | 0.0017828 | 2 | 0.0036 | 0.432 | 10 | 100% | 3.333% | 96.36% | 0.3049% | 0.0004127% | 0.002441%
10 | 0.025 | 0.00032547 | 0.00175 | 0.00178 |  | 1.1132e-08 | 1 | 1.0003 | 0.9998 | 3.6216e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 8.794e-06 | 0.0017832 | 2 | 0.0036 | 0.432 | 10 | 100% | 3.331% | 96.32% | 0.3501% | 0.0004125% | 0.002432%
5 | 0.017 | 0.00016273 | 0.00101 | 0.001023 |  | 2.6617e-09 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.971 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 4.2085e-06 | 0.0010243 | 2 | 0.0020 | 0.35294 | 5 | 100% | 2.524% | 97.22% | 0.2537% | 0.001523% | 0.001688%
2 | 0.017 | 6.5093e-05 | 0.00073 | 0.0007329 |  | 4.8376e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.83 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 1.7137e-06 | 0.00073324 | 2 | 0.0015 | 0.26471 | 2 | 100% | 0.7881% | 99.12% | 0.08998% | 0.002972% | 0.0005462%
2 | 0.017 | 6.5093e-05 | 0.00062 | 0.00062341 |  | 4.2209e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.986 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 1.6802e-06 | 0.00062376 | 2 | 0.0012 | 0.21176 | 2 | 100% | 1.089% | 98.8% | 0.1085% | 0.004107% | 0.0007256%
1 | 0.017 | 3.2547e-05 | 0.00065 | 0.00065081 |  | 1.1053e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.984 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.2 | 8.399e-07 | 0.00065091 | 2 | 0.0013 | 0.22941 | 1 | 100% | 0.25% | 99.72% | 0.02609% | 0.003771% | 0.0001665%
1 | 0.017 | 3.2547e-05 | 0.00065 | 0.00065081 |  | 1.0564e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.984 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 8.4033e-07 | 0.00065091 | 2 | 0.0013 | 0.22941 | 1 | 100% | 0.25% | 99.72% | 0.02493% | 0.003771% | 0.0001667%
0.5 | 0.005 | 1.6273e-05 | 0.00037 | 0.00037036 |  | 2.4939e-11 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 8.005 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 3.9641e-07 | 0.00037039 | 2 | 0.00074 | 0.444 | 0.5 | 100% | 0.193% | 99.79% | 0.01818% | 0.0004135% | 0.0001145%
0.2 | 0.005 | 6.5093e-06 | 0.00028 | 0.00028008 |  | 3.996e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 8.002 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 1.5862e-07 | 0.00028008 | 2 | 0.00056 | 0.336 | 0.2 | 100% | 0.05401% | 99.94% | 0.005094% | 0.0007231% | 3.207e-05%
0.2 | 0.005 | 6.5093e-06 | 0.00023 | 0.00023009 |  | 4.0166e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.992 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 1.5882e-07 | 0.0002301 | 2 | 0.00046 | 0.276 | 0.2 | 100% | 0.08003% | 99.91% | 0.007586% | 0.001071% | 4.764e-05%
0.1 | 0.005 | 3.2547e-06 | 0.00025 | 0.00025002 |  | 1.0535e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.989 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.08 | 7.9399e-08 | 0.00025002 | 2 | 0.00050 | 0.3 | 0.1 | 100% | 0.01695% | 99.98% | 0.001685% | 0.0009075% | 1.008e-05%
0.1 | 0.005 | 3.2547e-06 | 0.00025 | 0.00025002 |  | 1.0058e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.989 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 7.944e-08 | 0.00025002 | 2 | 0.00050 | 0.3 | 0.1 | 100% | 0.01695% | 99.98% | 0.001609% | 0.0009075% | 1.01e-05%
0.05 | 0.005 | 1.6273e-06 | 0.00025 | 0.00025001 |  | 2.4021e-13 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 8.001 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 4.5e-05 | 2.09 | 3.9868e-08 | 0.00025001 | 2 | 0.00050 | 0.3 | 0.05 | 100% | 0.004237% | 99.99% | 0.0003843% | 0.001164% | 2.543e-06%
0.02 | 0.005 | 6.5093e-07 | 0.00018 | 0.00018 |  | 2.5152e-11 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.703 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 7.238e-08 | 0.00018007 | 2 | 0.00036 | 0.216 | 0.02 | 100% | 0.001307% | 99.92% | 0.07757% | 0.002243% | 1.616e-05%
0.02 | 0.005 | 6.5093e-07 | 0.00015 | 0.00015 |  | 2.492e-11 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.711 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 7.2167e-08 | 0.00015009 | 2 | 0.00030 | 0.18 | 0.02 | 100% | 0.001881% | 99.88% | 0.1106% | 0.003229% | 2.312e-05%
0.01 | 0.005 | 3.2547e-07 | 0.0002 | 0.0002 |  | 6.3572e-12 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.697 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 6.9e-05 | 2.09 | 3.6255e-08 | 0.00020002 | 2 | 0.00040 | 0.24 | 0.01 | 100% | 0.0002648% | 99.98% | 0.01589% | 0.001818% | 3.286e-06%
0.01 | 0.005 | 3.2547e-07 | 0.0002 | 0.0002 |  | 6.332e-12 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.697 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 3.6271e-08 | 0.00020002 | 2 | 0.00040 | 0.24 | 0.01 | 100% | 0.0002648% | 99.98% | 0.01583% | 0.001818% | 3.288e-06%
0.005 | 0.005 | 1.6273e-07 | 0.00017 | 0.00017 |  | 1.8919e-12 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.689 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 1.8531e-08 | 0.00017001 | 2 | 0.00034 | 0.204 | 0.005 | 100% | 9.163e-05% | 99.99% | 0.006546% | 0.0009532% | 1.188e-06%
0.002 | 0.005 | 6.5093e-08 | 0.00018 | 0.00018 |  | 2.9613e-13 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.707 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 7.3632e-09 | 0.00018 | 2 | 0.00036 | 0.216 | 0.002 | 100% | 1.308e-05% | 100% | 0.000914% | 0.0008502% | 1.673e-07%
0.002 | 0.005 | 6.5093e-08 | 0.00015 | 0.00015 |  | 2.9758e-13 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.703 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 7.3741e-09 | 0.00015 | 2 | 0.00030 | 0.18 | 0.002 | 100% | 1.883e-05% | 100% | 0.001323% | 0.001224% | 2.417e-07%
0.001 | 0.005 | 3.2547e-08 | 0.0002 | 0.0002 |  | 6.2444e-14 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.709 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.16 | 3.7319e-09 | 0.0002 | 2 | 0.00040 | 0.24 | 0.001 | 100% | 2.648e-06% | 100% | 0.0001561% | 0.0006887% | 3.482e-08%
0.001 | 0.005 | 3.2547e-08 | 0.0002 | 0.0002 |  | 7.3854e-14 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.709 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 3.6789e-09 | 0.0002 | 2 | 0.00040 | 0.24 | 0.001 | 100% | 2.648e-06% | 100% | 0.0001846% | 0.0006887% | 3.384e-08%
0.0005 | 0.005 | 1.6273e-08 | 0 | 1.6273e-08 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 1.6273e-08 | 2 | 0.000000033 | 1.98e-05 | 0.0005 | 100% | 100% | 0% | 0% | 0% | 0%
0.0002 | 0.005 | 6.5093e-09 | 0 | 6.5093e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 6.5093e-09 | 2 | 0.000000013 | 7.8e-06 | 0.0002 | 100% | 100% | 0% | 0% | 0% | 0%
0.0002 | 0.005 | 6.5093e-09 | 0 | 6.5093e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 6.5093e-09 | 2 | 0.000000013 | 7.8e-06 | 0.0002 | 100% | 100% | 0% | 0% | 0% | 0%
0.0001 | 0.005 | 3.2547e-09 | 0 | 3.2547e-09 |  | 2.0902e-17 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 4.5e-05 | 0 | 0 | 5.612e-09 | 2 | 0.000000011 | 6.6e-06 | 0.0001 | 100% | 33.63% | 0% | 66.37% | 0% | 0%
0.0001 | 0.005 | 3.2547e-09 | 0 | 3.2547e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 3.2547e-09 | 2 | 0.0000000065 | 3.9e-06 | 0.0001 | 100% | 100% | 0% | 0% | 0% | 0%
5e-05 | 0.005 | 1.6273e-09 | 0 | 1.6273e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 1.6273e-09 | 2 | 0.0000000033 | 1.98e-06 | 5e-05 | 100% | 100% | 0% | 0% | 0% | 0%
Uncertainty of CIPM 2007 air density equation Above coefficients (yellow highlighting) updated from R111 (2004) values to match CIPM 2007 air density paper.
Note on A3: "Volume of Restraint" is the sum of all artifact volumes in the restraint.
Note on B3: The "Nominal of Restraint" is the combined nominal of the restraint, as established in the design.
Note on D3: "Cal Cert U of Starting Restraint" is typically the sum of the k=1 uncertainties of the restraint, though in rare cases the RSS value may be appropriate.
Note on E3: The "Sensitivity used for Series 1" will typically be the average sensitivity for Series 1 from the Mass Code report, but when no sensitivity weight is used, e.g., automated comparators, will be the sensitivity value entered for the comparator used.
Note on F3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on G3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on H3: From Mass Code report for Series 1.
Note on I3: From Mass Code report for Series 1.
Note on L3: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in the cell above is the temperature entered in U3.
Note on M3: Effective density of the restraint.
Note on N3: Uncertainty of the density of the restraint.
Note on O3: Uncertainty of the ability to measure the air temperature. Though not necessarily correct because of the difficulties accurately measuring air temperature, this value is typically the uncertainty of the thermometer used for that purpose during Series 1.
Note on P3: The uncertianty of the Relative Humidity measurement during Series 1.
Note on Q3: Uncertainty of the barometer used to measure the barometric pressure. This assumes that there is no elevation offset in the pressure measurement; the barometer is at the height of the measurements performed.
Note on S3: Enter the actual calculated air density present during Series 1.
Note on T3: This value comes from the calibration certificate for the restraint artifacts. If the restraints were not calibrated together so that the pressures were equal, use the BP from the certificates that deviates the most from Normal air density.
Note on U3: This deviation value is based on the temperature of the mass standards during Series 1 measurements.
Note on V3: Based on the material of the restraint. The effective CCE may be used if materials differ.
Note on X3: u_Sens_R: Calculated sensitivity uncertainty to be included in the combined uncetainty. (MAX(ABS(Max_Sens_R-Avg_Sens_R),ABS(Min_Sens_R-Avg_Sens_R)))/(SQRT(3))*(MaxAi_R-MinAi_R)
Note on B7: Enter the tolerance for the weight set being calibrated.
Note on D7: This uncertainty component is the value of the uncertainty of the restraint applied proportionately to the nominal value of the unknown. HINT: For nomnal X values less than the restraint, the value will be proportionately less than the uncertainty of the restraint. For nominal X values larger than the restraint, this value will be proportionately larger than the uncertainty of the restraint.
Note on E7: The Type A values will come from the mass code report for each nomnal mass.
Note on F7: This value is the RSS of columns D and E for each nominal X value.
Note on H7: The U²(ABC) is calculated as a variance because under some conditions it may actually decrease the uncertainty of the measurent result slightly and would thus be a negative value. As the square root of a negative number is an Imaginary number, the only way to utilize this component in normal mathematical equations is to keep it as a variance. It will be combined with variances of the rest of the uncertainty components associated with X. The u²(ABC) value calculated for the unknown weight is based on the standard conditions established by OIML D28 as the internal calculations of the Mass Code software do all calculations of the mass values relative to standard conditions: Rho X: 8000 kg/m³, Temperature: 20 °C, Rho Air: 1.2 kg/m³.
Note on I7: The average sensitivity for each X will come from the mass code report and will be the calculated average of all the MC calculated sensitivity values of each series. It may be different than the average sensitivity used for the mass code solutions.
Note on J7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on K7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on L7: The calculated sensitivity error uncertainty contribution for X. Calculations are based on the deviation of the sensitivity during the series from the average, multiplied by the difference limits set by the tolerance.
Note on M7: Enter the density value for X entered in the MC input file for the comparisons. Entered values must be between: 2.6 and 2.8 g/cm³, or 7.8 and 8.05 g/cm³, or 8.3 and 8.6 g/cm³, or 16.5 and 16.8 g/cm³ to correctly generate the CCE value.
Note on N7: The estimated uncertainty of the entered density value for X. The default 0.07 g/cm³ value is the k=1 portion of the density uncertainty given in OIML R111. If the density is known to a better uncertainty use that uncertiainty value. Most U.S. weight manufacturers believe that the uncertainty of thier assumed density is 0.025 g/cm³ (k=1)
Note on O7: The uncertainty of the air temperature in which the measurements are performed. Though not completely correct because of the difficulty of measuring air temperature, the value used is tyically the calibration uncertainty of the thermometer used to measure air temperature.
Note on P7: Unertainty of the relative humidity measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the hygrometer.
Note on Q7: Unertainty of the barometric pressure measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the barometer.
Note on R7: Enter the value of the actual calculated air density during the series for that nominal value.
Note on S7: The effect of this component on the mass uncertainty is negligible, but is shown here to demonstrate that the effect was considered. Follow the guidance for this component given in the 2007 Air Density paper.
Note on T7: The CCE shown is based on the density entered in column M. It is assumed that Mass Code will NOT be used to calibrate Brass weights which have a density of 8.3909 g/cm³ and CCE of 0.000054/°C. Entering 8.3909 will return the CCE of Nichrome.
Note on U7: These values will generally be the maximum recorded deviation of the temperature from 20 °C during a series. If the maximum temperature measures 22.1 °C the deviation will be 2.1. If the minimum temperature measures 18.9 °C the deviation will be 1.1 °C.
Note on V7: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in this cell will be the average of all temperature deviations in column W.
Note on W7: This column shows the calculated k=1 combined uncertianty for all components included in this uncertainty budget. NOTE: Variances are already squared and are included as shown.
Note on X7: Enter the appropriate k value based on the degrees of freedom associated with this nomal X value.
Note on Y7: k=2 is shown in the heading but each u(c) is expanded by multiplying the entered k value for that nominal X . (There is no way to easily show the actual k value for each nominal.)
Note on Z7: The Pn value is calculated based on the tolerance entered in column B. If the Pn cell background is RED, the uncertainty exceeds the Tolerance/3 limit.
Note on AA7: The Nominal value is shown in this column for convenience.

## 5321 Uncertainty Calculations
Volume of Restraint | Nominal of Restraint | Cal Cert U of Starting Restraint (k=1) | Sensitivity used for Series 1 | Maximum Sensitivity from Series 1 | Minimum Sensitivity from Series 1 | Max A(I) MC Data Series 1 | Min A(I) MC Data Series 1 | Total U (Restraint) to MC | U²(ABC Restraint) (Output from C.6.3.1) | Uncertainty due to u(CCE) of Restraint (10 % @ 2.17 °C) | Rho(R) | u(Rho(R)) | u(t) | u(RH) | u(P) |  | Rho(a) Series 1 (establishes ABC of Restraint) | Rho(a) at Restraint Cal | Max Deviation of T from 20 °C | CCE of Restraint |  | Uncertainty due to Sensitivity Error of Series 1 | [9 empty] | Maximum Permissible Errors for Weights (+/- in mg)
(cm³) | (g) | (mg) | (mg/div) | (mg/div) | (mg/div) | (mg) | (mg) | (mg) | (mg²) | (mg) | (g/cm³) | (g/cm³) | (°C) | (% RH) | (mmHg) |  | (mg/cm³) | (mg/cm³) | (°C) | ( /°C) |  | (mg) | [12 empty] | OIML | [18 empty] | NIST | User Tolerances
250 | 2000 | 0.06 | 0.9999 | 1.0002 | 0.99971 | 0.46995 | -31.4419 | 0.065093 | 0.00059738 | 0.0016602 | 7.9924 | 0.025 | 0.2 | 1 | 0.11 |  | 1.1779 | 1.2 | 2.17 | 4.5e-05 |  | 0.00608 | [9 empty] | Nominal | Units | Nominal (g) | OIML E1 | OIML E2 | OIML F1 | OIML F2 | OIML M1 | OIML M1-2 | OIML M2 | OIML M2-3 | OIML M3 | ASTM 000 | ASTM 00 | ASTM 0 | ASTM 1 | ASTM 2 | ASTM 3 | ASTM 4 | ASTM 5 | ASTM 6 | ASTM 7 | NIST F | User 1 | User 2 | User 3 | User 4 | User 5 | Not Specified
Tolerance to be evaluated: | [34 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
ASTM 0 |  | From MC report | [10 empty] | From MC report
Unknown (X) |  | Sourced from Restraint |  |  |  | Applicable to Unknown (X) | [19 empty] | % contribution to uncertainty
Nominal (g) | MPE (mg) ASTM 0 for this exercise | Type B from the restraint Calculated from Cell J3 | Type A of series From Output Report | Mass Code Output Uncertainty (k=1) Includes Type A and B of process from the restraint |  | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Average Sensitivity for Series X (mg/div) | Maximum Sensitivity from Series X (mg/div) | Minimum Sensitivity from Series X (mg/div) | Mass error from Sensitivity error of Series (mg) | Assumed/Reported Density Rho(x) (g/cm³) | uRho x (g/cm³) (k=1) | u(t) (°C) (k=1) | u(RH) (% RH) (k=1) | u(P) (mmHg) (k=1) | Rho(a) during this cal (mg/cm³) | Uncertainty of Mole Fraction of CO₂ in the air | CCE of X (/°C) | Max Deviation of T for Series from 20 °C (°C) | Uncertainty due to u(CCE) of X (10 % @ 1.55 °C) (mg) | u(c) (k=1) (mg) | k | Uk=95.45 % CI (mg) | Pn | Nominal (g) | Total | Type B for the restraint Calculated from Cell N3 | Type A of series | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Mass error from Sensitivity error (mg) | Uncertainty due to u(CCE) (10 % @ °C) (mg)
30000 | 38 | 0.9764 |  | 0.9764 |  | 3.7011 | 1 | 1 | 1 | 0 | 8 | 0.07 | 2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 30000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
20000 | 25 | 0.65093 |  | 0.65093 |  | 1.6449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 20000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10000 | 13 | 0.32547 |  | 0.32547 |  | 0.41124 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 10000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10000 | 13 | 0.32547 |  | 0.32547 |  | 0.41124 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 10000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
5000 | 6 | 0.16273 |  | 0.16273 |  | 0.10281 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 5000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
3000 | 3.8 | 0.09764 |  | 0.09764 |  | 0.037011 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 3000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2000 | 2.5 | 0.065093 |  | 0.065093 |  | 0.016449 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 0 | 0 | - | 2 | - | - | 2000 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1000 | 1.3 | 0.032547 | 0.02535 | 0.041254 |  | 0.00054031 | 0.9999 | 1.0002 | 0.99971 | 0.00024768 | 7.9435 | 0.025 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 2.17 | 0.00081008 | 0.04736 | 2 | 0.095 | 0.21923 | 1000 | 100% | 47.23% | 28.65% | 24.09% | 0.002735% | 0.02926%
1000 | 1.3 | 0.032547 | 0.02535 | 0.041254 |  | 0.00054031 | 0.9999 | 1.0002 | 0.99971 | 0.00024768 | 7.9435 | 0.025 | 0.2 | 1 | 0.11 | 1.1414 | 0 | 4.5e-05 | 2.17 | 0.00081008 | 0.04736 | 2 | 0.095 | 0.21923 | 1000 | 100% | 47.23% | 28.65% | 24.09% | 0.002735% | 0.02926%
500 | 0.6 | 0.016273 | 0.0141 | 0.021532 |  | 2.4076e-05 | 0.99891 | 1.0015 | 0.99454 | 0.0015141 | 7.982 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00042644 | 0.02214 | 2 | 0.044 | 0.22 | 500 | 100% | 54.02% | 40.56% | 4.912% | 0.4677% | 0.0371%
300 | 0.38 | 0.009764 | 0.01097 | 0.014686 |  | 8.7148e-06 | 0.99891 | 1.0015 | 0.99454 | 0.00095895 | 7.973 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00025615 | 0.015013 | 2 | 0.030 | 0.23684 | 300 | 100% | 42.3% | 53.4% | 3.867% | 0.408% | 0.02911%
200 | 0.25 | 0.0065093 | 0.00863 | 0.01081 |  | 3.8433e-06 | 0.99891 | 1.0015 | 0.99454 | 0.00063089 | 7.986 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 0.00017049 | 0.011005 | 2 | 0.022 | 0.264 | 200 | 100% | 34.98% | 61.49% | 3.173% | 0.3286% | 0.024%
100 | 0.13 | 0.0032547 | 0.00974 | 0.010269 |  | 9.8239e-07 | 0.99891 | 1.0015 | 0.99454 | 0.00032806 | 7.9517 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 8.5612e-05 | 0.010323 | 2 | 0.021 | 0.48462 | 100 | 100% | 9.941% | 89.03% | 0.9219% | 0.101% | 0.006878%
100 | 0.13 | 0.0032547 | 0.00974 | 0.010269 |  | 9.8239e-07 | 0.99891 | 1.0015 | 0.99454 | 0.00032806 | 7.9517 | 0.025 | 0.2 | 1 | 0.11 | 1.175 | 0 | 4.5e-05 | 2.23 | 8.5612e-05 | 0.010323 | 2 | 0.021 | 0.48462 | 100 | 100% | 9.941% | 89.03% | 0.9219% | 0.101% | 0.006878%
50 | 0.06 | 0.0016273 | 0.00501 | 0.0052677 |  | 2.7882e-07 | 1 | 1.0003 | 0.9998 | 8.6917e-06 | 7.969 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 4.3986e-05 | 0.0052943 | 2 | 0.011 | 0.55 | 50 | 100% | 9.448% | 89.55% | 0.9948% | 0.0002695% | 0.006903%
30 | 0.037 | 0.0009764 | 0.00315 | 0.0032979 |  | 1.0019e-07 | 1 | 1.0003 | 0.9998 | 5.3599e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 2.6382e-05 | 0.0033131 | 2 | 0.0066 | 0.53514 | 30 | 100% | 8.685% | 90.4% | 0.9127% | 0.0002617% | 0.006341%
20 | 0.037 | 0.00065093 | 0.00228 | 0.0023711 |  | 4.4315e-08 | 1 | 1.0003 | 0.9998 | 5.3599e-06 | 7.98 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 1.757e-05 | 0.0023805 | 2 | 0.0048 | 0.38919 | 20 | 100% | 7.477% | 91.73% | 0.782% | 0.000507% | 0.005448%
10 | 0.025 | 0.00032547 | 0.00175 | 0.00178 |  | 1.1132e-08 | 1 | 1.0003 | 0.9998 | 3.6216e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 8.794e-06 | 0.0017832 | 2 | 0.0036 | 0.432 | 10 | 100% | 3.331% | 96.32% | 0.3501% | 0.0004125% | 0.002432%
10 | 0.025 | 0.00032547 | 0.00175 | 0.00178 |  | 1.1132e-08 | 1 | 1.0003 | 0.9998 | 3.6216e-06 | 7.972 | 0.025 | 0.2 | 1 | 0.11 | 1.1732 | 0 | 4.5e-05 | 2.3 | 8.794e-06 | 0.0017832 | 2 | 0.0036 | 0.432 | 10 | 100% | 3.331% | 96.32% | 0.3501% | 0.0004125% | 0.002432%
5 | 0.017 | 0.00016273 | 0.00101 | 0.001023 |  | 2.6617e-09 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.971 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 4.2085e-06 | 0.0010243 | 2 | 0.0020 | 0.35294 | 5 | 100% | 2.524% | 97.22% | 0.2537% | 0.001523% | 0.001688%
3 | 0.017 | 9.764e-05 | 0.00073 | 0.0007365 |  | 1.0885e-09 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.83 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 2.5706e-06 | 0.00073725 | 2 | 0.0015 | 0.26471 | 3 | 100% | 1.754% | 98.04% | 0.2003% | 0.00294% | 0.001216%
2 | 0.017 | 6.5093e-05 | 0.00062 | 0.00062341 |  | 4.2209e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.986 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 1.6802e-06 | 0.00062376 | 2 | 0.0012 | 0.21176 | 2 | 100% | 1.089% | 98.8% | 0.1085% | 0.004107% | 0.0007256%
1 | 0.017 | 3.2547e-05 | 0.00065 | 0.00065081 |  | 1.0564e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.984 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 8.4033e-07 | 0.00065091 | 2 | 0.0013 | 0.22941 | 1 | 100% | 0.25% | 99.72% | 0.02493% | 0.003771% | 0.0001667%
1 | 0.017 | 3.2547e-05 | 0.00065 | 0.00065081 |  | 1.0564e-10 | 1.0003 | 1.0007 | 0.99988 | 3.9974e-06 | 7.984 | 0.025 | 0.2 | 1 | 0.11 | 1.1738 | 0 | 4.5e-05 | 2.2 | 8.4033e-07 | 0.00065091 | 2 | 0.0013 | 0.22941 | 1 | 100% | 0.25% | 99.72% | 0.02493% | 0.003771% | 0.0001667%
0.5 | 0.005 | 1.6273e-05 | 0.00037 | 0.00037036 |  | 2.4939e-11 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 8.005 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 3.9641e-07 | 0.00037039 | 2 | 0.00074 | 0.444 | 0.5 | 100% | 0.193% | 99.79% | 0.01818% | 0.0004135% | 0.0001145%
0.3 | 0.005 | 9.764e-06 | 0.00028 | 0.00028017 |  | 8.9911e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 8.002 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 2.3793e-07 | 0.00028019 | 2 | 0.00056 | 0.336 | 0.3 | 100% | 0.1214% | 99.87% | 0.01145% | 0.0007226% | 7.211e-05%
0.2 | 0.005 | 6.5093e-06 | 0.00023 | 0.00023009 |  | 4.0166e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.992 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 1.5882e-07 | 0.0002301 | 2 | 0.00046 | 0.276 | 0.2 | 100% | 0.08003% | 99.91% | 0.007586% | 0.001071% | 4.764e-05%
0.1 | 0.005 | 3.2547e-06 | 0.00025 | 0.00025002 |  | 1.0058e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.989 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 7.944e-08 | 0.00025002 | 2 | 0.00050 | 0.3 | 0.1 | 100% | 0.01695% | 99.98% | 0.001609% | 0.0009075% | 1.01e-05%
0.1 | 0.005 | 3.2547e-06 | 0.00025 | 0.00025002 |  | 1.0058e-12 | 0.99999 | 1.0001 | 0.99973 | 7.5318e-07 | 7.989 | 0.025 | 0.2 | 1 | 0.11 | 1.1744 | 0 | 4.5e-05 | 2.08 | 7.944e-08 | 0.00025002 | 2 | 0.00050 | 0.3 | 0.1 | 100% | 0.01695% | 99.98% | 0.001609% | 0.0009075% | 1.01e-05%
0.05 | 0.005 | 1.6273e-06 | 0.00025 | 0.00025001 |  | 2.4021e-13 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 8.001 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 4.5e-05 | 2.09 | 3.9868e-08 | 0.00025001 | 2 | 0.00050 | 0.3 | 0.05 | 100% | 0.004237% | 99.99% | 0.0003843% | 0.001164% | 2.543e-06%
0.03 | 0.005 | 9.764e-07 | 0.00018 | 0.00018 |  | 5.6592e-11 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.703 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 1.0857e-07 | 0.00018016 | 2 | 0.00036 | 0.216 | 0.03 | 100% | 0.002937% | 99.82% | 0.1744% | 0.002241% | 3.632e-05%
0.02 | 0.005 | 6.5093e-07 | 0.00015 | 0.00015 |  | 2.492e-11 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.711 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 7.2167e-08 | 0.00015009 | 2 | 0.00030 | 0.18 | 0.02 | 100% | 0.001881% | 99.88% | 0.1106% | 0.003229% | 2.312e-05%
0.01 | 0.005 | 3.2547e-07 | 0.0002 | 0.0002 |  | 6.332e-12 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.697 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 3.6271e-08 | 0.00020002 | 2 | 0.00040 | 0.24 | 0.01 | 100% | 0.0002648% | 99.98% | 0.01583% | 0.001818% | 3.288e-06%
0.01 | 0.005 | 3.2547e-07 | 0.0002 | 0.0002 |  | 6.332e-12 | 1.0001 | 1.0004 | 0.99988 | 8.529e-07 | 2.697 | 0.025 | 0.2 | 1 | 0.11 | 1.1749 | 0 | 6.9e-05 | 2.09 | 3.6271e-08 | 0.00020002 | 2 | 0.00040 | 0.24 | 0.01 | 100% | 0.0002648% | 99.98% | 0.01583% | 0.001818% | 3.288e-06%
0.005 | 0.005 | 1.6273e-07 | 0.00017 | 0.00017 |  | 1.8919e-12 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.689 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 1.8531e-08 | 0.00017001 | 2 | 0.00034 | 0.204 | 0.005 | 100% | 9.163e-05% | 99.99% | 0.006546% | 0.0009532% | 1.188e-06%
0.003 | 0.005 | 9.764e-08 | 0.00018 | 0.00018 |  | 6.663e-13 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.707 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 1.1045e-08 | 0.00018 | 2 | 0.00036 | 0.216 | 0.003 | 100% | 2.942e-05% | 100% | 0.002056% | 0.0008502% | 3.765e-07%
0.002 | 0.005 | 6.5093e-08 | 0.00015 | 0.00015 |  | 2.9758e-13 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.703 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 7.3741e-09 | 0.00015 | 2 | 0.00030 | 0.18 | 0.002 | 100% | 1.883e-05% | 100% | 0.001323% | 0.001224% | 2.417e-07%
0.001 | 0.005 | 3.2547e-08 | 0.0002 | 0.0002 |  | 7.3854e-14 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.709 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 3.6789e-09 | 0.0002 | 2 | 0.00040 | 0.24 | 0.001 | 100% | 2.648e-06% | 100% | 0.0001846% | 0.0006887% | 3.384e-08%
0.001 | 0.005 | 3.2547e-08 | 0.0002 | 0.0002 |  | 7.3854e-14 | 0.99982 | 1 | 0.99964 | 5.2486e-07 | 2.709 | 0.025 | 0.2 | 1 | 0.11 | 1.1582 | 0 | 6.9e-05 | 2.16 | 3.6789e-09 | 0.0002 | 2 | 0.00040 | 0.24 | 0.001 | 100% | 2.648e-06% | 100% | 0.0001846% | 0.0006887% | 3.384e-08%
0.0005 | 0.005 | 1.6273e-08 | 0 | 1.6273e-08 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 1.6273e-08 | 2 | 0.000000033 | 1.98e-05 | 0.0005 | 100% | 100% | 0% | 0% | 0% | 0%
0.0003 | 0.005 | 9.764e-09 | 0 | 9.764e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 9.764e-09 | 2 | 0.000000020 | 1.2e-05 | 0.0003 | 100% | 100% | 0% | 0% | 0% | 0%
0.0002 | 0.005 | 6.5093e-09 | 0 | 6.5093e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 6.5093e-09 | 2 | 0.000000013 | 7.8e-06 | 0.0002 | 100% | 100% | 0% | 0% | 0% | 0%
0.0001 | 0.005 | 3.2547e-09 | 0 | 3.2547e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 3.2547e-09 | 2 | 0.0000000065 | 3.9e-06 | 0.0001 | 100% | 100% | 0% | 0% | 0% | 0%
0.0001 | 0.005 | 3.2547e-09 | 0 | 3.2547e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 3.2547e-09 | 2 | 0.0000000065 | 3.9e-06 | 0.0001 | 100% | 100% | 0% | 0% | 0% | 0%
5e-05 | 0.005 | 1.6273e-09 | 0 | 1.6273e-09 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 0 | 0 | 1.6273e-09 | 2 | 0.0000000033 | 1.98e-06 | 5e-05 | 100% | 100% | 0% | 0% | 0% | 0%
Uncertainty of CIPM 2007 air density equation Above coefficients (yellow highlighting) updated from R111 (2004) values to match CIPM 2007 air density paper.
Note on A3: "Volume of Restraint" is the sum of all artifact volumes in the restraint.
Note on B3: The "Nominal of Restraint" is the combined nominal of the restraint, as established in the design.
Note on D3: "Cal Cert U of Starting Restraint" is typically the sum of the k=1 uncertainties of the restraint, though in rare cases the RSS value may be appropriate.
Note on E3: The "Sensitivity used for Series 1" will typically be the average sensitivity for Series 1 from the Mass Code report, but when no sensitivity weight is used, e.g., automated comparators, will be the sensitivity value entered for the comparator used.
Note on F3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on G3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on H3: From Mass Code report for Series 1.
Note on I3: From Mass Code report for Series 1.
Note on J3: This value is the combined uncertainty (k=1) for Series 1 and includes the uncertainty of the: restraints, sensitivity error, CCE of the restraint and u²abc.
Note on L3: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in the cell above is the temperature entered in U3.
Note on M3: Effective density of the restraint.
Note on N3: Uncertainty of the density of the restraint.
Note on O3: Uncertainty of the ability to measure the air temperature. Though not necessarily correct because of the difficulties accurately measuring air temperature, this value is typically the uncertainty of the thermometer used for that purpose during Series 1.
Note on P3: The uncertianty of the Relative Humidity measurement during Series 1.
Note on Q3: Uncertainty of the barometer used to measure the barometric pressure. This assumes that there is no elevation offset in the pressure measurement; the barometer is at the height of the measurements performed.
Note on S3: Enter the actual calculated air density present during Series 1.
Note on T3: This value comes from the calibration certificate for the restraint artifacts. If the restraints were not calibrated together so that the pressures were equal, use the BP from the certificates that deviates the most from Normal air density.
Note on U3: This deviation value is based on the temperature of the mass standards during Series 1 measurements.
Note on V3: Based on the material of the restraint. The effective CCE may be used if materials differ.
Note on X3: u_Sens_R: Calculated sensitivity uncertainty to be included in the combined uncetainty. (MAX(ABS(Max_Sens_R-Avg_Sens_R),ABS(Min_Sens_R-Avg_Sens_R)))/(SQRT(3))*(MaxAi_R-MinAi_R)
Note on A4: Select from the drop down list below the tolerance classification to which this weight set is to be evaluated.
Note on B7: Enter the tolerance for the weight set being calibrated.
Note on D7: This uncertainty component is the value of the uncertainty of the restraint applied proportionately to the nominal value of the unknown. HINT: For nomnal X values less than the restraint, the value will be proportionately less than the uncertainty of the restraint. For nominal X values larger than the restraint, this value will be proportionately larger than the uncertainty of the restraint.
Note on E7: The Type A values will come from the mass code report for each nomnal mass.
Note on F7: This value is the RSS of columns D and E for each nominal X value.
Note on H7: The U²(ABC) is calculated as a variance because under some conditions it may actually decrease the uncertainty of the measurent result slightly and would thus be a negative value. As the square root of a negative number is an Imaginary number, the only way to utilize this component in normal mathematical equations is to keep it as a variance. It will be combined with variances of the rest of the uncertainty components associated with X. The u²(ABC) value calculated for the unknown weight is based on the standard conditions established by OIML D28 as the internal calculations of the Mass Code software do all calculations of the mass values relative to standard conditions: Rho X: 8000 kg/m³, Temperature: 20 °C, Rho Air: 1.2 kg/m³.
Note on I7: The average sensitivity for each X will come from the mass code report and will be the calculated average of all the MC calculated sensitivity values of each series. It may be different than the average sensitivity used for the mass code solutions.
Note on J7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on K7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on L7: The calculated sensitivity error uncertainty contribution for X. Calculations are based on the deviation of the sensitivity during the series from the average, multiplied by the difference limits set by the tolerance.
Note on M7: Enter the density value for X entered in the MC input file for the comparisons. Entered values must be between: 2.6 and 2.8 g/cm³, or 7.8 and 8.05 g/cm³, or 8.3 and 8.6 g/cm³, or 16.5 and 16.8 g/cm³ to correctly generate the CCE value.
Note on N7: The estimated uncertainty of the entered density value for X. The default 0.07 g/cm³ value is the k=1 portion of the density uncertainty given in OIML R111. If the density is known to a better uncertainty use that uncertiainty value. Most U.S. weight manufacturers believe that the uncertainty of thier assumed density is 0.025 g/cm³ (k=1)
Note on O7: The uncertainty of the air temperature in which the measurements are performed. Though not completely correct because of the difficulty of measuring air temperature, the value used is tyically the calibration uncertainty of the thermometer used to measure air temperature.
Note on P7: Unertainty of the relative humidity measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the hygrometer.
Note on Q7: Unertainty of the barometric pressure measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the barometer.
Note on R7: Enter the value of the actual calculated air density during the series for that nominal value.
Note on S7: The effect of this component on the mass uncertainty is negligible, but is shown here to demonstrate that the effect was considered. Follow the guidance for this component given in the 2007 Air Density paper.
Note on T7: The CCE shown is based on the density entered in column M. It is assumed that Mass Code will NOT be used to calibrate Brass weights which have a density of 8.3909 g/cm³ and CCE of 0.000054/°C. Entering 8.3909 will return the CCE of Nichrome.
Note on V7: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in this cell will be the average of all temperature deviations in cells U8:U43.
Note on W7: This column shows the calculated k=1 combined uncertianty for all components included in this uncertainty budget. NOTE: Variances are already squared and are included as shown. Included are: Proportional Uncertainty from Series 1, Uncertainty due to Sensitivity errors, Output uncertainty from this Series, Uncertainty due to Temperature Deviation from 20 °C reference, Variance of u²abc.
Note on X7: Enter the appropriate k value based on the degrees of freedom associated with this nomal X value.
Note on Y7: k=2 is shown in the heading but each u(c) is expanded by multiplying the entered k value for that nominal X . (There is no way to easily show the actual k value for each nominal.)
Note on Z7: The Pn value is calculated based on the tolerance entered in column B. If the Pn cell background is RED, the uncertainty exceeds the Tolerance/3 limit.
Note on AA7: The Nominal value is shown in this column for convenience.

## lb Uncertainty Calculations
Volume of Restraint | Nominal of Restraint | Cal Cert U of Starting Restraint (k=1) | Sensitivity used for Series 1 | Maximum Sensitivity from Series 1 | Minimum Sensitivity from Series 1 | Max A(I) MC Data Series 1 | Min A(I) MC Data Series 1 | Total U (Restraint) to MC | U²(ABC Restraint) (Output from C.6.3.1) | Uncertainty due to u(CCE) of Restraint (10 % @ 1.07 °C) | Rho(R) | u(Rho(R)) | u(t) | u(RH) | u(P) |  | Rho(a) Series 1 (establishes ABC of Restraint) | Rho(a) at Restraint Cal | Max Deviation of T from 20 °C | CCE of Restraint |  | Uncertainty due to Sensitivity Error of Series 1 | [9 empty] | Maximum Permissible Errors for Weights (+/- in mg)
(cm³) | (lb) | (mg) | (mg/div) | (mg/div) | (mg/div) | (mg) | (mg) | (mg) | (mg²) | (mg) | (g/cm³) | (g/cm³) | (°C) | (% RH) | (mmHg) |  | (mg/cm³) | (mg/cm³) | (°C) | ( /°C) |  | (mg) | [12 empty] | OIML | [18 empty] | NIST | User Tolerances
28480.1986 | 504 | 160.5 | 999.6 | 1001.4 | 996.4 | 2748 | 449.8 | 160.548 | -2.6334 | 0.094319 | 8.027 | 0.025 | 0.2 | 1 | 0.11 |  | 1.1913 | 1.172 | 1.07 | 4.5e-05 |  | 4.246 | [9 empty] | Nominal | Units | Nominal (g) | OIML E1 | OIML E2 | OIML F1 | OIML F2 | OIML M1 | OIML M1-2 | OIML M2 | OIML M2-3 | OIML M3 | ASTM 000 | ASTM 00 | ASTM 0 | ASTM 1 | ASTM 2 | ASTM 3 | ASTM 4 | ASTM 5 | ASTM 6 | ASTM 7 | NIST F | User 1 | User 2 | User 3 | User 4 | User 5 | Not Specified
Tolerance to be evaluated: | [34 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
ASTM 3 |  | From MC report | [10 empty] | From MC report
Unknown (X) |  | Sourced from Restraint |  |  |  | Applicable to Unknown (X) | [19 empty] | % contribution to uncertainty
Nominal (lb) | MPE (mg) ASTM 3 for this exercise | Type B from the restraint Calculated from Cell J3 (mg) | Type A of series From Output Report (mg) | Mass Code Output Uncertainty (k=1) Includes Type A and B of process from the restraint |  | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Average Sensitivity for Series X (mg/div) | Maximum Sensitivity from Series X (mg/div) | Minimum Sensitivity from Series X (mg/div) | Mass error from Sensitivity error of Series (mg) | Assumed/Reported Density Rho(x) (g/cm³) | uRho x (g/cm³) (k=1) | u(t) (°C) (k=1) | u(RH) (% RH) (k=1) | u(P) (mmHg) (k=1) | Rho(a) during this cal (mg/cm³) | Uncertainty of Mole Fraction of CO₂ | CCE of X (/°C) | Max Deviation of T for Series from 20 °C (°C) | Uncertainty due to u(CCE) of X (10 % @ 1.07 °C) (mg) | u(c) (k=1) (mg) | k | Uk=95.45 % CI (mg) | Pn | Nominal (lb) | Total | Type B for the restraint Calculated from Cell N3 | Type A of series | U²(ABC) for unknown (Output from Simplified C.6.3.1) (mg²) | Mass error from Sensitivity error (mg) | Uncertainty due to u(CCE) (10 % @ °C) (mg)
500 | 2300 | 159.2738 | 91.756 | 183.8132 |  | 0 | 999.6 | 1001.4 | 996.4 | 4.2493 | 8 | 0.025 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 | 1.07 | 0.094572 | 183.8624 | 2 | 370 | 0.48261 | 500 | 100% | 75.04% | 24.9% | 0% | 0.05341% | 2.646e-05%
50 | 230 | 15.9274 |  | 15.9274 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 50 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
30 | 140 | 9.5564 |  | 9.5564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 30 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
20 | 91 | 6.371 |  | 6.371 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 20 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10 | 45 | 3.1855 |  | 3.1855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 10 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10 | 45 | 3.1855 |  | 3.1855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 10 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
5 | 23 | 1.5927 |  | 1.5927 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 5 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
3 | 14 | 0.95564 |  | 0.95564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 3 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
3 | 14 | 0.95564 |  | 0.95564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 3 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2 | 9.1 | 0.6371 |  | 0.6371 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1 | 4.5 | 0.31855 |  | 0.31855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1 | 4.5 | 0.31855 |  | 0.31855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.5 | 2.3 | 0.15927 |  | 0.15927 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.5 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.3 | 1.4 | 0.095564 |  | 0.095564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.3 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.2 | 0.91 | 0.06371 |  | 0.06371 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.1 | 0.45 | 0.031855 |  | 0.031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.1 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.1 | 0.45 | 0.031855 |  | 0.031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.1 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.05 | 0.36 | 0.015927 |  | 0.015927 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.03 | 0.32 | 0.0095564 |  | 0.0095564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.03 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.02 | 0.23 | 0.006371 |  | 0.006371 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.02 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.01 | 0.16 | 0.0031855 |  | 0.0031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.01 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.01 | 0.16 | 0.0031855 |  | 0.0031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.01 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.005 | 0.14 | 0.0015927 |  | 0.0015927 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.005 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.003 | 0.11 | 0.00095564 |  | 0.00095564 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.003 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.002 | 0.091 | 0.0006371 |  | 0.0006371 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.002 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.001 | 0.068 | 0.00031855 |  | 0.00031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.001 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.001 | 0.068 | 0.00031855 |  | 0.00031855 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.001 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.0005 | 0.064 | 0.00015927 |  | 0.00015927 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.0005 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.0003 | 0.054 | 9.5564e-05 |  | 9.5564e-05 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.0003 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.0002 | 0.045 | 6.371e-05 |  | 6.371e-05 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.0002 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.0001 | 0.041 | 3.1855e-05 |  | 3.1855e-05 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.0001 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0.0001 | 0.041 | 3.1855e-05 |  | 3.1855e-05 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 0.0001 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
5e-05 | 0.036 | 1.5927e-05 |  | 1.5927e-05 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 5e-05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
3e-05 | 0.032 | 9.5564e-06 |  | 9.5564e-06 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 3e-05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2e-05 | 0.029 | 6.371e-06 |  | 6.371e-06 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 2e-05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1e-05 | 0.027 | 3.1855e-06 |  | 3.1855e-06 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1e-05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1e-05 | 0.027 | 3.1855e-06 |  | 3.1855e-06 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1e-05 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
5e-06 | 0.023 | 1.5927e-06 |  | 1.5927e-06 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 5e-06 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
3e-06 | 0.023 | 9.5564e-07 |  | 9.5564e-07 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 3e-06 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2e-06 | 0.023 | 6.371e-07 |  | 6.371e-07 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 2e-06 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1e-06 | 0.023 | 3.1855e-07 |  | 3.1855e-07 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1e-06 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
1e-06 | 0.023 | 3.1855e-07 |  | 3.1855e-07 |  | 0 | 1 | 1 | 1 | 0 | 8 | 0.07 | 0.2 | 1 | 0.11 | 1.2 | 0 | 4.5e-05 |  | 0 | - | 2 | - | - | 1e-06 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Uncertainty of CIPM 2007 air density equation Above coefficients (yellow highlighting) updated from R111 (2004) values to match CIPM 2007 air density paper.
Note on A3: "Volume of Restraint" is the sum of all artifact volumes in the restraint.
Note on B3: The "Nominal of Restraint" is the combined nominal of the restraint, as established in the design.
Note on D3: "Cal Cert U of Starting Restraint" is typically the sum of the k=1 uncertainties of the restraint, though in rare cases the RSS value may be appropriate.
Note on E3: The "Sensitivity used for Series 1" will typically be the average sensitivity for Series 1 from the Mass Code report, but when no sensitivity weight is used, e.g., automated comparators, will be the sensitivity value entered for the comparator used.
Note on F3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on G3: This value will typically be equal to the sensitivity used unless there is a known uncorrected sensitivity bias. Variations in sensitivity are measured in the within process standard deviation.
Note on H3: From Mass Code report for Series 1.
Note on I3: From Mass Code report for Series 1.
Note on L3: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in the cell above is the temperature entered in U3.
Note on M3: Effective density of the restraint.
Note on N3: Uncertainty of the density of the restraint.
Note on O3: Uncertainty of the ability to measure the air temperature. Though not necessarily correct because of the difficulties accurately measuring air temperature, this value is typically the uncertainty of the thermometer used for that purpose during Series 1.
Note on P3: The uncertianty of the Relative Humidity measurement during Series 1.
Note on Q3: Uncertainty of the barometer used to measure the barometric pressure. This assumes that there is no elevation offset in the pressure measurement; the barometer is at the height of the measurements performed.
Note on S3: Enter the actual calculated air density present during Series 1.
Note on T3: This value comes from the calibration certificate for the restraint artifacts. If the restraints were not calibrated together so that the pressures were equal, use the BP from the certificates that deviates the most from Normal air density.
Note on U3: This deviation value is based on the temperature of the mass standards during Series 1 measurements.
Note on V3: Based on the material of the restraint. The effective CCE may be used if materials differ.
Note on X3: u_Sens_R: Calculated sensitivity uncertainty to be included in the combined uncetainty. (MAX(ABS(Max_Sens_R-Avg_Sens_R),ABS(Min_Sens_R-Avg_Sens_R)))/(SQRT(3))*(MaxAi_R-MinAi_R)
Note on B7: Enter the tolerance for the weight set being calibrated.
Note on D7: This uncertainty component is the value of the uncertainty of the restraint applied proportionately to the nominal value of the unknown. HINT: For nomnal X values less than the restraint, the value will be proportionately less than the uncertainty of the restraint. For nominal X values larger than the restraint, this value will be proportionately larger than the uncertainty of the restraint.
Note on E7: The Type A values will come from the mass code report for each nomnal mass.
Note on F7: This value is the RSS of columns D and E for each nominal X value.
Note on H7: The U²(ABC) is calculated as a variance because under some conditions it may actually decrease the uncertainty of the measurent result slightly and would thus be a negative value. As the square root of a negative number is an Imaginary number, the only way to utilize this component in normal mathematical equations is to keep it as a variance. It will be combined with variances of the rest of the uncertainty components associated with X. The u²(ABC) value calculated for the unknown weight is based on the standard conditions established by OIML D28 as the internal calculations of the Mass Code software do all calculations of the mass values relative to standard conditions: Rho X: 8000 kg/m³, Temperature: 20 °C, Rho Air: 1.2 kg/m³.
Note on I7: The average sensitivity for each X will come from the mass code report and will be the calculated average of all the MC calculated sensitivity values of each series. It may be different than the average sensitivity used for the mass code solutions.
Note on J7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on K7: From the mass code report for the series in which the artifact is assigned a value. Variations in sensitivity may be also measured in the within process standard deviation.
Note on L7: The calculated sensitivity error uncertainty contribution for X. Calculations are based on the deviation of the sensitivity during the series from the average, multiplied by the difference limits set by the tolerance.
Note on M7: Enter the density value for X entered in the MC input file for the comparisons. Entered values must be between: 2.6 and 2.8 g/cm³, or 7.8 and 8.05 g/cm³, or 8.3 and 8.6 g/cm³, or 16.5 and 16.8 g/cm³ to correctly generate the CCE value.
Note on N7: The estimated uncertainty of the entered density value for X. The default 0.07 g/cm³ value is the k=1 portion of the density uncertainty given in OIML R111. If the density is known to a better uncertainty use that uncertiainty value. Most U.S. weight manufacturers believe that the uncertainty of thier assumed density is 0.025 g/cm³ (k=1)
Note on O7: The uncertainty of the air temperature in which the measurements are performed. Though not completely correct because of the difficulty of measuring air temperature, the value used is tyically the calibration uncertainty of the thermometer used to measure air temperature.
Note on P7: Unertainty of the relative humidity measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the hygrometer.
Note on Q7: Unertainty of the barometric pressure measurement of the air surrounding the comparison measurements. Typically the value used will be the uncertainty of the calibration of the barometer.
Note on R7: Enter the value of the actual calculated air density during the series for that nominal value.
Note on S7: The effect of this component on the mass uncertainty is negligible, but is shown here to demonstrate that the effect was considered. Follow the guidance for this component given in the 2007 Air Density paper.
Note on T7: The CCE shown is based on the density entered in column M. It is assumed that Mass Code will NOT be used to calibrate Brass weights which have a density of 8.3909 g/cm³ and CCE of 0.000054/°C. Entering 8.3909 will return the CCE of Nichrome.
Note on V7: 10 % of the CCE is used as described in guidance from EURAMET cg-19, Guidelines on the determination of uncertainty in gravimetric volume calibration version 2.1, paragraph 5.3.6. The temperature displayed in this cell will be the average of all temperature deviations in cells U8:U43.
Note on W7: This column shows the calculated k=1 combined uncertianty for all components included in this uncertainty budget. NOTE: Variances are already squared and are included as shown.
Note on X7: Enter the appropriate k value based on the degrees of freedom associated with this nomal X value.
Note on Y7: k=2 is shown in the heading but each u(c) is expanded by multiplying the entered k value for that nominal X . (There is no way to easily show the actual k value for each nominal.)
Note on Z7: The Pn value is calculated based on the tolerance entered in column B. If the Pn cell background is RED, the uncertainty exceeds the Tolerance/3 limit.
Note on AA7: The Nominal value is shown in this column for convenience.

## Sensitivity Calculation Sheet
Copy and Paste the image from the Output Text file here. |  | Enter the values from the output text file here.
AVERAGE OBSERVED
A(I) DELTA(I) SENSITIVITY DRIFT(I) SENSITIVITY
[mg] [mg] [mg/div] [mg] [mg/div] |  |  | Observed Sensitivity Values
A 1 -31.44190 -0.03200 0.99990 0.00500 0.99996 | A1 |  | 999.8517
A 2 -30.94195 -0.01101 0.99990 0.00500 0.99991 | A2 |  | 1000.8525
A 3 -31.28691 0.02800 0.99990 0.04000 1.00023 | A3 |  | 998.3541
A 4 -31.10193 0.01500 0.99990 0.00500 0.99986 | A4 |  | 1001.3537
A 5 0.46995 -0.00900 0.99990 0.02000 0.99998 | A5 |  | 996.3644
A 6 0.07999 -0.01499 0.99990 0.01000 0.99983 | A6 |  | 1000.8525
A 7 0.28497 -0.00799 0.99990 0.01500 0.99996 | A7
A 8 -0.40496 -0.02099 0.99990 -0.02500 0.99971 | A8
A 9 -0.18498 0.00101 0.99990 -0.00500 0.99971 | A9
A 10 0.18998 -0.00800 0.99990 0.00000 0.99988 | A10
A 11 -0.32969 -0.00029 1.00011 0.00055 1.00031 | A11
A12
A13
A14
A15
A16
A17
|  | Mean of Avg Sensitivities | Max Observed Sensitivity | Min Observed Sensitivity
|  | 999.6048 | 1001.3537 | 996.3644

## MPE Table
Maximum Permissible Errors for Weights (+/- in mg)
|  |  | OIML R111 (2004) | [8 empty] | ASTM E617 (2018) | [9 empty] | NIST HB 105-1 (1990) | User Tolerances
Nominal | Units | Nominal (g) | OIML E1 | OIML E2 | OIML F1 | OIML F2 | OIML M1 | OIML M1-2 | OIML M2 | OIML M2-3 | OIML M3 | ASTM 000 | ASTM 00 | ASTM 0 | ASTM 1 | ASTM 2 | ASTM 3 | ASTM 4 | ASTM 5 | ASTM 6 | ASTM 7 | NIST F | User 1 | User 2 | User 3 | User 4 | User 5 | Not Specified
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
5 000 | kg | 5000000 | - | - | 25000 | 80000 | 250000 | 500000 | 800000 | 1600000 | 2500000 | - | - | - | - | 25000 | 50000 | 100000 | 250000 | 500000 | 750000 | 500000 | - | - | - | - | - | -
3 000 | kg | 3000000 | - | - | - | - | - | - | - | - | - | - | - | - | - | 15000 | 30000 | 60000 | 150000 | 300000 | 450000 | 300000 | - | - | - | - | - | -
2 000 | kg | 2000000 | - | - | 10000 | 30000 | 100000 | 200000 | 300000 | 600000 | 1000000 | - | - | - | - | 10000 | 20000 | 40000 | 100000 | 200000 | 300000 | 200000 | - | - | - | - | - | -
1 000 | kg | 1000000 | - | 1600 | 5000 | 16000 | 50000 | 100000 | 160000 | 300000 | 500000 | - | - | - | - | 5000 | 10000 | 20000 | 50000 | 100000 | 150000 | 100000 | - | - | - | - | - | -
500 | kg | 500000 | - | 800 | 2500 | 8000 | 25000 | 50000 | 80000 | 160000 | 250000 | - | - | - | - | 2500 | 5000 | 10000 | 25000 | 50000 | 75000 | 50000 | - | - | - | - | - | -
300 | kg | 300000 | - | - | - | - | - | - | - | - | - | - | - | - | - | 1500 | 3000 | 6000 | 15000 | 30000 | 45000 | 30000 | - | - | - | - | - | -
200 | kg | 200000 | - | 300 | 1000 | 3000 | 10000 | 20000 | 30000 | 60000 | 100000 | - | - | - | - | 1000 | 2000 | 4000 | 10000 | 20000 | 30000 | 20000 | - | - | - | - | - | -
100 | kg | 100000 | - | 160 | 500 | 1600 | 5000 | 10000 | 16000 | 30000 | 50000 | - | - | - | - | 500 | 1000 | 2000 | 5000 | 10000 | 15000 | 10000 | - | - | - | - | - | -
50 | kg | 50000 | 25 | 80 | 250 | 800 | 2500 | 5000 | 8000 | 16000 | 25000 | 13 | 25 | 63 | 125 | 250 | 500 | 1000 | 2500 | 5000 | 7500 | 5000 | - | - | - | - | - | -
30 | kg | 30000 | - | - | - | - | - | - | - | - | - | 7.5 | 15 | 38 | 75 | 150 | 300 | 600 | 1500 | 3000 | 4500 | 3000 | - | - | - | - | - | -
25 | kg | 25000 | - | - | - | - | - | - | - | - | - | 6.3 | 13 | 31 | 62 | 125 | 250 | 500 | 1200 | 2500 | 3800 | 2500 | - | - | - | - | - | -
20 | kg | 20000 | 10 | 30 | 100 | 300 | 1000 | - | 3000 | - | 10000 | 5 | 10 | 25 | 50 | 100 | 200 | 400 | 1000 | 2000 | 3800 | 2000 | - | - | - | - | - | -
10 | kg | 10000 | 5 | 16 | 50 | 160 | 500 | - | 1600 | - | 5000 | 2.5 | 5 | 13 | 25 | 50 | 100 | 200 | 500 | 1000 | 2200 | 1000 | - | - | - | - | - | -
5 | kg | 5000 | 2.5 | 8 | 25 | 80 | 250 | - | 800 | - | 2500 | 1.3 | 2.5 | 6 | 12 | 25 | 50 | 100 | 250 | 500 | 1400 | 500 | - | - | - | - | - | -
3 | kg | 3000 | - | - | - | - | - | - | - | - | - | 0.75 | 1.5 | 3.8 | 7.5 | 15 | 30 | 60 | 150 | 300 | 1000 | 300 | - | - | - | - | - | -
2 | kg | 2000 | 1 | 3 | 10 | 30 | 100 | - | 300 | - | 1000 | 0.5 | 1 | 2.5 | 5 | 10 | 20 | 40 | 100 | 200 | 750 | 200 | - | - | - | - | - | -
1 | kg | 1000 | 0.5 | 1.6 | 5 | 16 | 50 | - | 160 | - | 500 | 0.25 | 0.5 | 1.3 | 2.5 | 5 | 10 | 20 | 50 | 100 | 470 | 100 | - | - | - | - | - | -
500 | g | 500 | 0.25 | 0.8 | 2.5 | 8 | 25 | - | 80 | - | 250 | 0.13 | 0.25 | 0.6 | 1.2 | 2.5 | 5 | 10 | 30 | 50 | 300 | 70 | - | - | - | - | - | -
300 | g | 300 | - | - | - | - | - | - | - | - | - | 0.075 | 0.15 | 0.38 | 0.75 | 1.5 | 3 | 6 | 20 | 30 | 210 | 60 | - | - | - | - | - | -
200 | g | 200 | 0.1 | 0.3 | 1 | 3 | 10 | - | 30 | - | 100 | 0.05 | 0.1 | 0.25 | 0.5 | 1 | 2 | 4 | 15 | 20 | 160 | 40 | - | - | - | - | - | -
100 | g | 100 | 0.05 | 0.16 | 0.5 | 1.6 | 5 | - | 16 | - | 50 | 0.025 | 0.05 | 0.13 | 0.25 | 0.5 | 1 | 2 | 9 | 10 | 100 | 20 | - | - | - | - | - | -
50 | g | 50 | 0.03 | 0.1 | 0.3 | 1 | 3 | - | 10 | - | 30 | 0.015 | 0.03 | 0.06 | 0.12 | 0.25 | 0.6 | 1.2 | 5.6 | 7 | 62 | 10 | - | - | - | - | - | -
30 | g | 30 | - | - | - | - | - | - | - | - | - | 0.014 | 0.026 | 0.037 | 0.074 | 0.15 | 0.45 | 0.9 | 4 | 5 | 44 | 6 | - | - | - | - | - | -
20 | g | 20 | 0.025 | 0.08 | 0.25 | 0.8 | 2.5 | - | 8 | - | 25 | 0.013 | 0.025 | 0.037 | 0.074 | 0.1 | 0.35 | 0.7 | 3 | 3 | 33 | 4 | - | - | - | - | - | -
10 | g | 10 | 0.02 | 0.06 | 0.2 | 0.6 | 2 | - | 6 | - | 20 | 0.01 | 0.02 | 0.025 | 0.05 | 0.074 | 0.25 | 0.5 | 2 | 2 | 21 | 2 | - | - | - | - | - | -
5 | g | 5 | 0.016 | 0.05 | 0.16 | 0.5 | 1.6 | - | 5 | - | 16 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.18 | 0.36 | 1.3 | 2 | 13 | 1.5 | - | - | - | - | - | -
3 | g | 3 | - | - | - | - | - | - | - | - | - | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.15 | 0.3 | 0.95 | 2 | 9.4 | 1.3 | - | - | - | - | - | -
2 | g | 2 | 0.012 | 0.04 | 0.12 | 0.4 | 1.2 | - | 4 | - | 12 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.13 | 0.26 | 0.75 | 2 | 7 | 1.1 | - | - | - | - | - | -
1 | g | 1 | 0.01 | 0.03 | 0.1 | 0.3 | 1 | - | 3 | - | 10 | 0.005 | 0.01 | 0.017 | 0.034 | 0.054 | 0.1 | 0.2 | 0.5 | 2 | 4.5 | 0.9 | - | - | - | - | - | -
500 | mg | 0.5 | 0.008 | 0.025 | 0.08 | 0.25 | 0.8 | - | 2.5 | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.08 | 0.16 | 0.38 | 1 | 3 | 0.72 | - | - | - | - | - | -
300 | mg | 0.3 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.07 | 0.14 | 0.3 | 1 | 2.2 | 0.61 | - | - | - | - | - | -
200 | mg | 0.2 | 0.006 | 0.02 | 0.06 | 0.2 | 0.6 | - | 2 | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.06 | 0.12 | 0.26 | 1 | 1.8 | 0.54 | - | - | - | - | - | -
100 | mg | 0.1 | 0.005 | 0.016 | 0.05 | 0.16 | 0.5 | - | 1.6 | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.025 | 0.05 | 0.1 | 0.2 | 1 | 1.2 | 0.43 | - | - | - | - | - | -
50 | mg | 0.05 | 0.004 | 0.012 | 0.04 | 0.12 | 0.4 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.042 | 0.085 | 0.16 | - | 0.88 | 0.35 | - | - | - | - | - | -
30 | mg | 0.03 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.038 | 0.075 | 0.14 | - | 0.68 | 0.3 | - | - | - | - | - | -
20 | mg | 0.02 | 0.003 | 0.01 | 0.03 | 0.1 | 0.3 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.035 | 0.07 | 0.12 | - | 0.56 | 0.26 | - | - | - | - | - | -
10 | mg | 0.01 | 0.003 | 0.008 | 0.025 | 0.08 | 0.25 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.03 | 0.06 | 0.1 | - | 0.4 | 0.21 | - | - | - | - | - | -
5 | mg | 0.005 | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.028 | 0.055 | 0.08 | - | - | 0.17 | - | - | - | - | - | -
3 | mg | 0.003 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.026 | 0.052 | 0.07 | - | - | 0.14 | - | - | - | - | - | -
2 | mg | 0.002 | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.025 | 0.05 | 0.06 | - | - | 0.12 | - | - | - | - | - | -
1 | mg | 0.001 | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.025 | 0.05 | 0.05 | - | - | 0.1 | - | - | - | - | - | -
500 | µg | 0.0005 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.025 | 0.05 | 0.05 | - | - | - | - | - | - | - | - | -
300 | µg | 0.0003 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | 0.025 | - | - | - | - | - | - | - | - | - | - | -
200 | µg | 0.0002 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | 0.014 | - | - | - | - | - | - | - | - | - | - | - | -
100 | µg | 0.0001 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | 0.01 | - | - | - | - | - | - | - | - | - | - | - | - | -
50 | µg | 5e-05 | - | - | - | - | - | - | - | - | - | 0.002 | 0.003 | 0.005 | - | - | - | - | - | - | - | - | - | - | - | - | - | -
The values in the above table are correct as of 20200505. Verified by comparison to current documentary standards.
Maximum Permissible Errors for Weights (+/- in mg)
|  |  | OIML R111 (2004) | [8 empty] | ASTM E617 (2018) | [9 empty] | NIST HB 105-1 (1990) | User Tolerances
Nominal | Units | Nominal (lb) | OIML E1 | OIML E2 | OIML F1 | OIML F2 | OIML M1 | OIML M1-2 | OIML M2 | OIML M2-3 | OIML M3 | ASTM 000 | ASTM 00 | ASTM 0 | ASTM 1 | ASTM 2 | ASTM 3 | ASTM 4 | ASTM 5 | ASTM 6 | ASTM 7 | NIST F | User 1 | User 2 | User 3 | User 4 | User 5 | Not Specified
| lb |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
10000 | lb | 10000 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 45000 | 91000 | 230000 | 450000 | 680000 | 450000 | - | - | - | - | - | -
5000 | lb | 5000 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 23000 | 45000 | 110000 | 230000 | 340000 | 230000 | - | - | - | - | - | -
3000 | lb | 3000 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 14000 | 27000 | 68000 | 140000 | 200000 | 140000 | - | - | - | - | - | -
2500 | lb | 2500 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 11000 | 23000 | 57000 | 110000 | 170000 | 110000 | - | - | - | - | - | -
2000 | lb | 2000 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 9100 | 18000 | 45000 | 91000 | 140000 | 91000 | - | - | - | - | - | -
1000 | lb | 1000 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 4500 | 9100 | 23000 | 45000 | 68000 | 45000 | - | - | - | - | - | -
500 | lb | 500 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 2300 | 4500 | 11000 | 23000 | 34000 | 23000 | - | - | - | - | - | -
300 | lb | 300 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 1400 | 2700 | 6800 | 14000 | 20000 | 14000 | - | - | - | - | - | -
200 | lb | 200 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | 910 | 1800 | 4500 | 9100 | 14000 | 9100 | - | - | - | - | - | -
100 | lb | 100 | - | - | - | - | - | - | - | - | - | - | - | 57 | 110 | 230 | 450 | 910 | 2300 | 4500 | 6800 | 4500 | - | - | - | - | - | -
50 | lb | 50 | - | - | - | - | - | - | - | - | - | - | - | 29 | 57 | 110 | 230 | 450 | 1100 | 2300 | 4100 | 2300 | - | - | - | - | - | -
30 | lb | 30 | - | - | - | - | - | - | - | - | - | - | - | 17 | 32 | 68 | 140 | 270 | 680 | 1400 | 2700 | 1400 | - | - | - | - | - | -
25 | lb | 25 | - | - | - | - | - | - | - | - | - | - | - | 14 | 28 | 57 | 110 | 230 | 570 | 1100 | 2400 | 1100 | - | - | - | - | - | -
20 | lb | 20 | - | - | - | - | - | - | - | - | - | - | - | 12 | 23 | 45 | 91 | 180 | 450 | 910 | 2000 | 910 | - | - | - | - | - | -
10 | lb | 10 | - | - | - | - | - | - | - | - | - | - | - | 5.5 | 11 | 23 | 45 | 91 | 230 | 450 | 1300 | 450 | - | - | - | - | - | -
5 | lb | 5 | - | - | - | - | - | - | - | - | - | - | - | 2.7 | 5.4 | 11 | 23 | 45 | 110 | 230 | 780 | 230 | - | - | - | - | - | -
3 | lb | 3 | - | - | - | - | - | - | - | - | - | - | - | 1.7 | 3.4 | 6.8 | 14 | 27 | 68 | 140 | 580 | 140 | - | - | - | - | - | -
2 | lb | 2 | - | - | - | - | - | - | - | - | - | - | - | 1.2 | 2.3 | 4.5 | 9.1 | 18 | 45 | 91 | 440 | 91 | - | - | - | - | - | -
1 | lb | 1 | - | - | - | - | - | - | - | - | - | - | - | 0.55 | 1.1 | 2.3 | 4.5 | 9.1 | 23 | 45 | 270 | 70 | - | - | - | - | - | -
0.5 | lb | 0.5 | - | - | - | - | - | - | - | - | - | - | - | 0.27 | 0.54 | 1.1 | 2.3 | 4.5 | 16 | 23 | 170 | 45 | - | - | - | - | - | -
0.3 | lb | 0.3 | - | - | - | - | - | - | - | - | - | - | - | 0.17 | 0.34 | 0.68 | 1.4 | 2.7 | 11 | 14 | 120 | 27 | - | - | - | - | - | -
0.2 | lb | 0.2 | - | - | - | - | - | - | - | - | - | - | - | 0.12 | 0.23 | 0.45 | 0.91 | 1.8 | 8.2 | 9.1 | 97 | 18 | - | - | - | - | - | -
0.1 | lb | 0.1 | - | - | - | - | - | - | - | - | - | - | - | 0.055 | 0.11 | 0.23 | 0.45 | 1.1 | 5.4 | 6.8 | 59 | 9.1 | - | - | - | - | - | -
0.05 | lb | 0.05 | - | - | - | - | - | - | - | - | - | - | - | 0.027 | 0.054 | 0.11 | 0.36 | 0.77 | 3.3 | 4.5 | 37 | 4.5 | - | - | - | - | - | -
0.03 | lb | 0.03 | - | - | - | - | - | - | - | - | - | - | - | 0.017 | 0.034 | 0.068 | 0.32 | 0.59 | 2.4 | 3.2 | 26 | 2.7 | - | - | - | - | - | -
0.02 | lb | 0.02 | - | - | - | - | - | - | - | - | - | - | - | 0.017 | 0.034 | 0.045 | 0.23 | 0.45 | 1.9 | 2.3 | 20 | 1.8 | - | - | - | - | - | -
0.01 | lb | 0.01 | - | - | - | - | - | - | - | - | - | - | - | 0.012 | 0.023 | 0.034 | 0.16 | 0.34 | 1.2 | 1.4 | 12 | 1.5 | - | - | - | - | - | -
0.005 | lb | 0.005 | - | - | - | - | - | - | - | - | - | - | - | 0.0075 | 0.015 | 0.024 | 0.14 | 0.27 | 0.86 | 0.91 | 7.8 | 1.2 | - | - | - | - | - | -
0.003 | lb | 0.003 | - | - | - | - | - | - | - | - | - | - | - | 0.0075 | 0.015 | 0.024 | 0.11 | 0.22 | 0.64 | 0.91 | 5.4 | 0.99 | - | - | - | - | - | -
0.002 | lb | 0.002 | - | - | - | - | - | - | - | - | - | - | - | 0.0075 | 0.015 | 0.024 | 0.091 | 0.19 | 0.5 | 0.91 | 4.2 | 0.87 | - | - | - | - | - | -
0.001 | lb | 0.001 | - | - | - | - | - | - | - | - | - | - | - | 0.0075 | 0.015 | 0.024 | 0.068 | 0.15 | 0.36 | 0.91 | 2.9 | 0.7 | - | - | - | - | - | -
0.0005 | lb | 0.0005 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.023 | 0.064 | 0.13 | 0.27 | - | 2.1 | - | - | - | - | - | - | -
0.0003 | lb | 0.0003 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.023 | 0.054 | 0.11 | 0.23 | - | 1.5 | - | - | - | - | - | - | -
0.0002 | lb | 0.0002 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.023 | 0.045 | 0.095 | 0.2 | - | 1.2 | - | - | - | - | - | - | -
0.0001 | lb | 0.0001 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.023 | 0.041 | 0.086 | 0.16 | - | 0.84 | - | - | - | - | - | - | -
5e-05 | lb | 5e-05 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.036 | 0.073 | 0.14 | - | - | - | - | - | - | - | - | -
3e-05 | lb | 3e-05 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.032 | 0.064 | 0.11 | - | - | - | - | - | - | - | - | -
2e-05 | lb | 2e-05 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.029 | 0.059 | 0.091 | - | - | - | - | - | - | - | - | -
1e-05 | lb | 1e-05 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.027 | 0.054 | 0.091 | - | - | - | - | - | - | - | - | -
5e-06 | lb | 5e-06 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.023 | 0.054 | 0.091 | - | - | - | - | - | - | - | - | -
3e-06 | lb | 3e-06 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.023 | 0.054 | 0.091 | - | - | - | - | - | - | - | - | -
2e-06 | lb | 2e-06 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.023 | 0.054 | 0.091 | - | - | - | - | - | - | - | - | -
1e-06 | lb | 1e-06 | - | - | - | - | - | - | - | - | - | - | - | 0.0045 | 0.0091 | 0.014 | 0.023 | 0.054 | 0.091 | - | - | - | - | - | - | - | - | -
The values in the above table are correct as of 20200505. Verified by comparison to current documentary standards.