## Software Technical Assessment
Software Technical Assessment
Software Description | WAQCF-021 (Rev. 01), Advanced Mass Control Chart.xlt | Approval Date | 2015-10-09
Assessed By | Dan Wright | Approved By | Dan Wright
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | Very well laid out
| There are instructions for use | Pass | Instruction worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | Clearly marked worksheets for instructions and data input
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank data entry cells, light blue for non-blank data entry cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Chart Procedure (SOP 9) and Measurement Procedure (SOP 28) are both specified
| The number of digits to be rounded to is specified | N/A | No rounding
| The user is warned/notified whenever data-entry fields are left blank | Pass | Visual, blank data-entry cells are color coded
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry cells are blank on the template
| The software opens at the right location within the file | Pass | Opens to the instruction worksheet
| Unused fields/cells are locked | Pass | All unused cells are locked
| Rows/columns that the operator need not see are hidden | Pass | Yes
| Unused sheets are removed | Pass | Yes
| Worksheets are named appropriately | Pass | Yes
B. Mathematical Specification | The correct SOP is used | Pass | SOP 9
| The formulae and methods chosen from that SOP are specified | Pass | SOP 9 method as explained on Instructions worksheet
| Sources and references for formulae are specified | Pass | SOP 9
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | Pass
C. Code review | The formulae in the fields exactly match the SOP | Pass | Upper / Lower control limits
| Repeated calculations appropriately reference the correct cells | Pass | Uses absolute cell references
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | N/A
| Rounding is done at the appropriate locations in the file | N/A | No rounding
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | N/A
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | N/A
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | N/A
E. Component Testing | Each macro used is functional | N/A | No macros
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | N/A
| Plotted graphs are accurate | Pass | Charts depict correct information
| Worksheets/reports print properly, if needed to | Pass | All worksheets and charts are formated for printing
| Conditional (color and non-color) formatting is functional | Pass | For blank and nonblank cells
F. Numerical Reference | Look-up tables and lists match the latest calibration report | Pass | Standards workbook is validated to latest calibration reports and Mass Tolerances workbook is validated to latest references
| Uncertainties match the latest Scope | N/A
| Values that reference another workbook or spreadsheet are dated | N/A
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct | N/A | No embedded data
| The evaluation of the embedded data is dated and documented | N/A
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | N/A
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | Pass | Hand calculation did not reveal any significant differences
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Locked and password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | Password hint in instructions
| Files are backed up automatically | Pass | Manual backup
| Additional back-up is available at alternate facilities | Pass | Backed-up to agency network drives
| Files on network drives cannot be accidentally deleted | Pass

## Documentation
Documentation
Date | Description | Initials
2010-08-31 | Developed by Dan Wright, State Metrologist, Washington Metrology Laboratory | DAW
2010-08-31 | Validated workbook, file name 'WA2010-08-31 WAQCF-021 (Rev. 00), Advanced Mass Control Chart Workbook V&V'.pdf on Laboratory Computer C:\ drive and Agency Server H:\ drive and a paper copy in laboratory files. | DAW
2015-10-09 | Added MPE Finder worksheet. | DAW
2015-10-09 | Added precision test to data entry sheet, | DAW
2015-10-09 | Validated links and range names. | DAW
2015-10-09 | Validated workbook, file name 'WA2015-10-09 WAQCF-021 (Rev. 01), Advanced Mass Control Chart Workbook V&V'.pdf on Laboratory Computer C:\ drive and Agency Server H:\ drive and a paper copy in laboratory files. | DAW

## Instructions
Instructions
No. | Description
1 | Enter requested information in the light yellow cells rows 1 through 10 on the Data Entry worksheet.
2 | On the Data Entry spreadsheet enter the best uncertainty from youe CMC tables in rows 12 down, column B. These uncertainties are used for the error bars on the s(t) Chart. To add an additional weighing design, unprotect sheet and place cursor in column A where it states Insert Row Here and insert row. This will be the next row you will enter new data.
3 | Enter data from your measurements in the light yellow cells in rows 2 down column A through P on the LAP Data Entry worksheet. After entering data highlight columns Q through AD on the row above the one your entering data and copy down.
4 | On the LAP Data Entry worksheet place the cursor in column A where it states Insert Row Here and insert a row. This will be the next row you will enter new data.
5 | Modify the 'CC MPE %' on the Data Entry worksheet to adjust the limits just outside the UCL and LCL on the s(t) Chart. This gives a better visual of your data points.
6 | Modify horizontal chart axis on the temperature and humidity charts to match the requirements.
7 | Do not change the formulas in L1 through L10 on the Data Entry spreadsheet, these are used for the control chart titles.
8 | Password protected. Hint: password
9 | You must link this workbook to the Mass MPE workbook for the lookups on the MPE Finder worksheet to work.
10 | The MPE Finder worksheet is linked to the Mass MPE workbook.
12 | Name files as follows; Balance (ID), Load (ID), SOP, e.g., CC30002 (S5), 10 kg (C55), SOP 28.
Range Names
Name | Range
ASTM_ap | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 ap'!$B$3:$F$18
ASTM_dwt | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 dwt'!$B$3:$E$19
ASTM_g | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 g'!$B$3:$L$48
ASTM_gr | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 gr'!$B$3:$E$27
ASTM_lb | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 lb'!$B$3:$J$39
ASTM_oz | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 oz'!$B$3:$J$31
ASTM_ozt | ='C:\LINKED\TOL\[Mass MPE.xls]ASTM E 617 oz t'!$B$3:$G$31
ChkID | ='Data Entry'!$C$2
Design.Table | ='Data Entry'!$A$13:$G$43
NIST_ap | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-1 ap'!$B$3:$I$21
NIST_g | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-1 g'!$B$3:$K$43
NIST_gr | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-1 gr'!$B$3:$I$27
NIST_lb | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-1 lb'!$B$3:$K$31
NIST_oz | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-1 oz'!$B$3:$I$31
OIML | ='C:\LINKED\TOL\[Mass MPE.xls]OIML R 111'!$B$3:$K$32
PooledSwTable | ='LAP Data Entry'!$U:$V
Spec | ='Data Entry'!$G$9
Unit | ='MPE Finder'!$G$4:$G$15
Weightcart | ='C:\LINKED\TOL\[Mass MPE.xls]NIST HB 105-8'!$B$3:$C$11

## Data Entry
Laboratory: |  |  |  |  | Pooled sw = |  | #N/A | d.f = | 0 | Do Not Delete, Used for Chart Titles and Legends
Check ID Code: |  |  |  |  | Standard Deviation (st) = |  | #DIV/0! | d.f = | -1 | #DIV/0!
Check Description: |  |  |  |  | Check Standard Average = |  | #DIV/0! | n = | 0 | #N/A
Report True Mass Correction: | [9 empty] | ● ● F Ratio Chart () Check Standard Balance: ● Measurement Procedure: SOP 28
Calibration Report U (k=2): | [9 empty] | ● ● T Value Chart () Check Standard Balance: ● Measurement Procedure: SOP 28
Balance ID Code: | [9 empty] | ● ● Check Standard vs Humidity Chart () Check Standard Balance: ● Measurement Procedure: SOP 28
Balance Description: | [9 empty] | ● ● Check Standard vs Pressure Chart () Check Standard Balance: ● Measurement Procedure: SOP 28
Procedure: |  | SOP 28 |  |  | Nominal Value: |  | Nominal Value Unit |  |  | ● ● Check Standard vs Temperature Chart () Check Standard Balance: ● Measurement Procedure: SOP 28
Parameter: |  |  |  |  | Specification: |  |  | Class: |  | milligrams | Control Limits (Chk Mean ± 3 s) | Warning Limits (Chk Mean ± 2 s)
Control Chart Unit: |  | mg |  |  | MPE (mg) = |  | CC MPE %: |  |  | Check Mean ± 0 % MPE |  |  | 0
Weighing Design | Best Uncertainty | sb | K1 | K2 | Restraint Position | Check Position | Design Matrix | Precision Test (Pn)
A.1.1 |  | #DIV/0! | 0.8165 | 1.4142 | 1 1 1 | 1 1 1 | See Comment | ---
A.1.2 |  | #DIV/0! | 0.70711 | 1.4142 | 1 1 1 1 | 1 -1 1 1 | See Comment | ---
A.1.4 |  | #DIV/0! | 0.63246 | 1.4142 | 1 1 1 1 1 | 1 -1 1 1 1 | See Comment | ---
A.1.4 |  | #DIV/0! | 0.54772 | 1.2247 | 1 1 1 1 1 | 1 1 1 1 1 | See Comment | ---
A.2.1 |  | #DIV/0! | 0.70711 | 1.4142 | 1 1 1 1 1 1 | 1 -1 1 1 1 1 | See Comment | ---
B.1a |  | #DIV/0! | 0.70711 | 1.4142 | 2 1 1 1 | 2 1 1 -1 | See Comment | ---
B.3a |  | #DIV/0! | 0.57735 | 1.4142 | 2 2 1 1 | 2 2 1 -1 | See Comment | ---
B.4 |  | #DIV/0! | 0.43818 | 1.0583 | 2 2 1 1 1 | 2 2 1 1 1 | See Comment | ---
B.8 |  | #DIV/0! | 0.5069 | 1.0607 | 2 2 2 1 1 1 | 2 2 2 1 1 1 | See Comment | ---
C.1 |  | #DIV/0! | 0.41748 | 1.0149 | 5 3 2 1 1 | 5 3 2 1 1 | See Comment | ---
C.1a |  | #DIV/0! | 0.53452 | 1.4142 | 5 3 2 1 1 | 5 3 2 1 -1 | See Comment | ---
C.2 |  | #DIV/0! | 0.35509 | 1.0149 | 5 3 2 1 1 1 | 5 3 2 1 1 1 | See Comment | ---
C.2a |  | #DIV/0! | 0.5 | 1.4142 | 5 3 2 1 1 1 | 5 3 2 1 1 1 | See Comment | ---
C.9 |  | #DIV/0! | 0.41404 | 1.0198 | 5 2 2 1 1 | 5 2 2 1 1 | See Comment | ---
C.9a |  | #DIV/0! | 0.53452 | 1.4142 | 5 2 2 1 1 | 5 2 2 1 -1 | See Comment | ---
C.10 |  | #DIV/0! | 0.46445 | 1.0198 | 5 2 2 1 1 1 | 5 2 2 1 1 1 | See Comment | ---
C.10a |  | #DIV/0! | 0.53452 | 1.4142 | 5 2 2 1 1 1 | 5 2 2 1 1 -1 | See Comment | ---
D.1a |  | #DIV/0! | 0.5 | 1.4142 | 4 2 2 1 1 | 4 2 2 1 -1 | See Comment | ---
D.1a |  | #DIV/0! | 0.5 | 1.4142 | 4 2 2 1 1 | 4 2 2 1 1 | See Comment | ---
D.14 |  | #DIV/0! | 0.47746 | 1.0008 | 5 4 3 2 1 | 5 4 3 2 1 | See Comment | ---
WA.C.1a |  | #DIV/0! | 1.0954 | 1.5232 | 5 3 2 1 1 | 5 3 2 1 -1 | See Comment | ---
WA.C.2a |  | #DIV/0! | 0.8165 | 1.4142 | 5 3 2 1 1 1 | 5 3 2 1 1 -1 | See Comment | ---
WA.C.2a |  | #DIV/0! | 0.70711 | 1.2247 | 5 3 2 1 1 1 | 5 3 2 1 1 1 | See Comment | ---
WA.C.5 |  | #DIV/0! | 0.4916 | 1.01 | 5 5 3 2 1 1 | 5 5 3 2 1 1 | See Comment | ---
WA.C.10 |  | #DIV/0! | 0.52599 | 1.0198 | 5 2 2 1 1 1 | 5 2 2 1 1 1 | See Comment | ---
WA.C.10a |  | #DIV/0! | 0.8165 | 1.4142 | 5 2 2 1 1 1 | 5 2 2 1 1 -1 | See Comment | ---
WA.C.10a |  | #DIV/0! | 0.70711 | 1.2247 | 5 2 2 1 1 1 | 5 2 2 1 1 1 | See Comment | ---
2.3.4.1.9 |  | #DIV/0! | 0.44721 | 1.0198 | 5 2 2 1 1 1 1 | 5 2 2 1 1 1 1 | See Comment | ---
2.3.4.1.17 |  | #DIV/0! | 0.41952 | 1.0392 | 3 2 1 1 1 | 3 2 1 1 1 | See Comment | ---
2.3.4.1.17a |  | #DIV/0! | 0.63246 | 1.4142 | 3 2 1 1 1 | 3 2 1 1 1 | See Comment | ---
Insert Row Here
Note on C6: M5 - XP2003KL M6 - AX1005 S4 - CC30002 S5 - CC5001 S6 - CC111 S7 - CCE6
Note on H13: Design Matrix 1 1 1 Y1 + - Y2 + - Y3 + -
Note on H14: Design Matrix 1 1 1 1 Y1 + - Y2 + - Y3 + - Y4 + - Y5 + - Y6 + -
Note on H15: Design Matrix 1 1 1 1 1 Y1 + - Y2 + - Y3 + - Y4 + - Y5 + - Y6 + - Y7 + - Y8 + - Y9 + - Y10 + -
Note on H16: Design Matrix 1 1 1 1 1 Y1 + - Y2 + - Y3 + - Y4 + - Y5 + - Y6 + - Y7 + - Y8 + - Y9 + - Y10 + -
Note on H17: Design Matrix 1 1 1 1 1 1 Y1 + - Y2 + - Y3 + - Y4 + - Y5 + - Y6 + - Y7 + - Y8 + -
Note on H18: Design Matrix 2 1 1 1 Y1 + - - Y2 + - - Y3 + - - Y4 + - Y5 + - Y6 + -
Note on H19: Design Matrix 2 2 1 1 Y1 + - + - Y2 + - - + Y3 + - Y4 + - - Y5 + - - Y6 + -
Note on H20: Design Matrix 2 2 1 1 1 Y1 + - - + Y2 + - - + Y3 + - + - Y4 + - Y5 + - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - -
Note on H21: Design Matrix 2 2 2 1 1 1 Y1 + - Y2 + - Y3 + - - Y4 + - - Y5 + - - Y6 + - Y7 + - - Y8 + - - Y9 + - - Y10 + - - Y11 + - - Y12 + - -
Note on H22: Design Matrix 5 3 2 1 1 Y1 + - - + - Y2 + - - - + Y3 + - - Y4 + - - - Y5 + - - Y6 + - - Y7 + - - Y8 + -
Note on H23: Design Matrix 5 3 2 1 1 Y1 + - - + - Y2 + - - - + Y3 + - - Y4 + - - - Y5 + - - Y6 + - - Y7 + - - Y8 + -
Note on H24: Design Matrix 5 3 2 1 1 1 Y1 + - - + - Y2 + - - + - Y3 + - - - + Y4 + - - Y5 + - - - - Y6 + - + - - Y7 + - - + - Y8 + - - - + Y9 + - - Y10 + - - Y11 + - -
Note on H25: Design Matrix 5 3 2 1 1 1 Y1 + - - + - Y2 + - - + - Y3 + - - - + Y4 + - - Y5 + - - - - Y6 + - + - - Y7 + - - + - Y8 + - - - + Y9 + - - Y10 + - - Y11 + - -
Note on H26: Design Matrix 5 2 2 1 1 Y1 + - - - Y2 + - - - Y3 + - + - Y4 + - - + Y5 + - Y6 + - - Y7 + - - Y8 + -
Note on H27: Design Matrix 5 2 2 1 1 Y1 + - - - Y2 + - - - Y3 + - + - Y4 + - - + Y5 + - Y6 + - - Y7 + - - Y8 + -
Note on H28: Design Matrix 5 2 2 1 1 1 Y1 + - - - - + Y2 + - - - + - Y3 + - - + - - Y4 + - - - - Y5 + - - - - Y6 + - + - Y7 + - - + Y8 + - + -
Note on H29: Design Matrix 5 2 2 1 1 1 Y1 + - - - - + Y2 + - - - + - Y3 + - - + - - Y4 + - - - - Y5 + - - - - Y6 + - + - Y7 + - - + Y8 + - + -
Note on H30: Design Matrix 4 2 2 1 1 Y1 + - - + - Y2 + - - - + Y3 + - - - Y4 + - - - Y5 + - + - Y6 + - - +
Note on H31: Design Matrix 4 2 2 1 1 Y1 + - - + - Y2 + - - - + Y3 + - - - Y4 + - - - Y5 + - + - Y6 + - - +
Note on H32: Design Matrix 5 4 3 2 1 Y1 + - - + Y2 + - - + Y3 + - - Y4 + - - Y5 + - - + Y6 + - - Y7 + - -
Note on H33: Design Matrix 5 3 2 1 1 Y1 + - - Y2 + - - - Y3 + - - Y4 + - - Y5 + - - Y6 + -
Note on H34: Design Matrix 5 3 2 1 1 1 Y1 + - - Y2 + - - - Y3 + - - - Y4 + - - - Y5 + - - - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - - Y11 + - -
Note on H35: Design Matrix 5 3 2 1 1 1 Y1 + - - Y2 + - - - Y3 + - - - Y4 + - - - Y5 + - - - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - - Y11 + - -
Note on H36: Design Matrix 5 5 3 2 1 1 Y1 + - - Y2 + - - Y3 + - - - Y4 + - - - Y5 + - - Y6 + - - Y7 + - - Y8 + -
Note on H37: Design Matrix 5 2 2 1 1 1 Y1 + - - - Y2 + - - - Y3 + - - - Y4 + - Y5 + - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - -
Note on H38: Design Matrix 5 2 2 1 1 1 Y1 + - - - Y2 + - - - Y3 + - - - Y4 + - Y5 + - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - -
Note on H39: Design Matrix 5 2 2 1 1 1 Y1 + - - - Y2 + - - - Y3 + - - - Y4 + - Y5 + - - Y6 + - - Y7 + - - Y8 + - - Y9 + - - Y10 + - -
Note on H40: Design Matrix 5 2 2 1 1 1 1 Y1 + - - - Y2 + - - - Y3 + - - - Y4 + - - - Y5 + + - - - Y6 + + - - - Y7 + + - - - - Y8 + - Y9 + - Y10 + -
Note on H41: Design Matrix 3 2 1 1 1 Y1 + - - Y2 + - - Y3 + - - Y4 + - - - Y5 + - - Y6 + - - Y7 + - - Y8 + - Y9 + - Y10 + -
Note on H42: Design Matrix 3 2 1 1 1 Y1 + - - Y2 + - - Y3 + - - Y4 + - - - Y5 + - - Y6 + - - Y7 + - - Y8 + - Y9 + - Y10 + -

## MPE Finder
Maximum Permissible Error (MPE) Finder |  |  |  | Do Not Delete | [23 empty] | Do Not Delete
Standard Specification: |  |  | 0 |  | MPE Lookup Table | [7 empty] | Class Table Column Lookup
ASTM E 617 Specification Class: |  |  | --- |  | Unit | gram | ASTM E 617 | OIML R 111 | NIST HB 105-1 | NIST HB 105-1 F | NIST HB 105-8 |  | Class | ASTM g | ASTM lb | ASTM oz | ASTM oz t | ASTM dwt | ASTM gr | ASTM carat | ASTM ap | Class | OIML R 111 | Class | 105-1 g | 105-1 lb | 105-1 oz | 105-1 ap | 105-1 gr
NIST HB 105-1 (draft) & OIML R 111 Spec. Class: |  |  | --- |  | kg | --- | --- | --- | --- | --- | --- |  | 000 | 2 | [7 empty] | E1 | 2 | E1 | 2 | 2 | 2 | 2 | 2
Mass Nominal Value: |  |  | 0 |  | g | --- | --- | --- | --- | --- | --- |  | 00 | 3 | [7 empty] | E2 | 3 | E2 | 3 | 3 | 3 | 3 | 3
Nominal Value Unit: |  |  | 0% |  | mg | --- | --- | --- | --- | --- | --- |  | 0 | 4 | 2 | 2 | [5 empty] | F1 | 4 | F1 | 4 | 4 | 4 | 4 | 4
| MPE (± mg) = | FALSE |  |  | lb | --- | --- | --- | --- | --- | --- |  | 1 | 5 | 3 | 3 | [5 empty] | F2 | 5 | F2 | 5 | 5 | 5 | 5 | 5
[5 empty] | oz | --- | --- | --- | --- | --- | --- |  | 2 | 6 | 4 | 4 | 2 |  |  |  |  | M1 | 6 | M1 | 6 | 6 | 6 | 6 | 6
Interpolation of Intermediate Values |  |  |  |  | oz t | --- | --- | --- | --- | --- | --- |  | 3 | 7 | 5 | 5 | 3 |  |  | 2 |  | M1-2 | 7 | M1-2 | 7 | 7
For ASTM E 617 and NIST HB 105-1 (draft) Only |  |  |  |  | dwt | --- | --- | --- | --- | --- | --- |  | 4 | 8 | 6 | 6 | 4 | 2 | 2 | 3 | 2 | M2 | 8 | M2 | 8 | 8 | 7 | 7 | 7
nominal value to be interpolated (x) g: |  |  | #N/A |  | gr | --- | --- | --- | --- | --- | --- |  | 5 | 9 | 7 | 7 | 5 | 3 | 3 | 4 | 3 | M2-3 | 9 | M2-3 | 9 | 9
lower nominal value from table (x1) 0: |  |  |  |  | carat | --- | --- | --- | --- | --- | --- |  | 6 | 10 | 8 | 8 | 6 | 4 | 4 | 5 | 4 | M3 | 10 | M3 | 10 | 10 | 8 | 8 | 8
lower nominal value from table (x1) g: |  |  | #N/A |  | oz ap | --- | --- | --- | --- | --- | --- |  | 7 | 11 | 9 | 9 |  |  |  | 6 | 5
upper nominal value from table (x2) 0: |  |  |  |  | dr ap | --- | --- | --- | --- | --- | ---
upper nominal value from table (x2) g: |  |  | #N/A |  | s ap | --- | --- | --- | --- | --- | ---
lower table value cooresponding with x1 (y1) mg: |  |  | #N/A |  | Class F formated to two significant digits |  |  |  |  | #N/A
upper table value cooresponding with x2 (y2) mg: |  |  | #N/A
interpolated table value (y) mg: |  |  | #N/A |  | Interpolation of Intermediate Values Lower Value Lookup |  |  |  |  | Interpolation of Intermediate Values Upper Value Lookup
MPE (± mg) = |  | --- |  |  | Unit | gram | ASTM E 617 | NIST HB 105-1 |  | Unit | gram | ASTM E 617 | NIST HB 105-1
Formula |  |  |  |  | kg | --- | --- | --- |  | kg | --- | --- | ---
[5 empty] | g | --- | --- | --- |  | g | --- | --- | ---
[5 empty] | mg | --- | --- | --- |  | mg | --- | --- | ---
[5 empty] | lb | --- | --- | --- |  | lb | --- | --- | ---
[5 empty] | oz | --- | --- | --- |  | oz | --- | --- | ---
[5 empty] | oz t | --- | --- | --- |  | oz t | --- | --- | ---
[5 empty] | dwt | --- | --- | --- |  | dwt | --- | --- | ---
[5 empty] | gr | --- | --- | --- |  | gr | --- | --- | ---
[5 empty] | carat | --- | --- | --- |  | carat | --- | --- | ---
[5 empty] | oz ap | --- | --- | --- |  | oz ap | --- | --- | ---
[5 empty] | dr ap | --- | --- | --- |  | dr ap | --- | --- | ---
[5 empty] | s ap | --- | --- | --- |  | s ap | --- | --- | ---
|  |  |  | Do Not Delete | [23 empty] | Do Not Delete

## LAP Data Entry
Test Number | Operator ID | Date | Starting Restraint ID | Check Standard ID | Check Observed Value | Balance ID | Observed sw | Degrees of Freedom | Weighing Design ID | Average Temp | Average Pressure | Average Humidity | Average Air Density | F Ratio | T Value | Check Mean Value | U (k=2) | Accepted sw | d.f.*sw2 | Pooled d.f. | Pooled sw | UWL Limit | LWL Limit | UCL Limit | LCL Limit | Calibrated Value | T Value Mean | Mean + 0 % of MPE | Mean - 0 % of MPE
[16 empty] | #DIV/0! | #N/A | #N/A |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | #N/A | #DIV/0! | #DIV/0!
Insert Row Here