## Data
UserName | EntryDate | SerialNumber | Overall | Barrel | Serial Number | Overall Stats |  |  | Barrel Stats
All dimensions in inches
Al | 2015-01-01 | 11 | 20.4375 | 9.6875 |  | Means | Stdevs |  | Means | Stdevs
Al | 2015-01-01 | 22 | 28.6875 | 11.875
Al | 2015-01-01 | 33 | 18.5 | 12.25 | #11 | 20.4344 | 0.024627 |  | 9.7344 | 0.028932
Al | 2015-01-01 | 44 | 20.1875 | 13.625
Betty | 2015-01-01 | 11 | 20.4375 | 9.75 | #22 | 28.6812 | 0.027951 | * | 11.0281 | 0.20122
Betty | 2015-01-01 | 22 | 28.6875 | 11
Betty | 2015-01-01 | 33 | 18.5 | 12.25 | #33 | 18.7125 | 0.076948 |  | 12.25 | 0.020278
Betty | 2015-01-01 | 44 | 20.25 | 13.5625
Carl | 2015-01-01 | 11 | 20.4375 | 9.6875 | #44 | 20.2336 | 0.027766 |  | 13.6184 | 0.019237
Carl | 2015-01-01 | 22 | 28.6875 | 11
Carl | 2015-01-01 | 33 | 18.6875 | 12.25 | #22* | *The first value for SN 22 widely |  |  | 10.9836 | 0.028276
Carl | 2015-01-01 | 44 | 20.1875 | 13.625 |  | differs from the mean. It is removed
Denise | 2015-01-01 | 11 | 20.4375 | 9.75 |  | here as an example of a possible
Denise | 2015-01-01 | 22 | 28.5625 | 10.9375 |  | action for dealing with outliers.
Denise | 2015-01-01 | 33 | 18.6875 | 12.25
Denise | 2015-01-01 | 44 | 20.25 | 13.625 |  |  | Pooled Stdev - OL |  |  | Pooled Stdev - BL
Ed | 2015-01-01 | 11 | 20.4375 | 9.75 |  |  | 0.044944 |  |  | 0.024586
Ed | 2015-01-01 | 22 | 28.6875 | 11
Ed | 2015-01-01 | 33 | 18.6875 | 12.25 |  | Total Number of Observations for OL or BL (N):
Ed | 2015-01-01 | 44 | 20.1875 | 13.625 |  |  |  | 80
Betty | 2015-01-02 | 44 | 20.25 | 13.625
Betty | 2015-01-02 | 11 | 20.4375 | 9.6875 |  | Number of Repetitions of the Same Quantity by Each User (R):
Betty | 2015-01-02 | 22 | 28.6875 | 11 |  |  |  | 4
Betty | 2015-01-02 | 33 | 18.6875 | 12.25
Carl | 2015-01-03 | 44 | 20.25 | 13.625 |  | Corrected, Conservative Estmate of Standard Deviation
Carl | 2015-01-03 | 11 | 20.4375 | 9.75 |  | = Pooled Stdev × √[(N-1)/(N-R)]
Carl | 2015-01-03 | 22 | 28.6875 | 11 |  |  | 0.045822 |  |  | 0.025066 | †
Carl | 2015-01-03 | 33 | 18.75 | 12.25 |  | Enter this value on page 2, cell C7. |  |  | Enter this value on page 3, cell C7.
Denise | 2015-01-04 | 44 | 20.25 | 13.625
Denise | 2015-01-04 | 11 | 20.5 | 9.75 |  | †Although the mean and stdev for BL are calculated from only
Denise | 2015-01-04 | 22 | 28.6875 | 11 |  | 79 measurements here instead of 80,
Denise | 2015-01-04 | 33 | 18.75 | 12.25 |  | the error in the correction factor is negligible.
Ed | 2015-01-05 | 44 | 20.25 | 13.625
Ed | 2015-01-05 | 11 | 20.4375 | 9.6875 |  | Data and calculations adapted from results by:
Ed | 2015-01-05 | 22 | 28.6875 | 10.9375 |  | Johnson County Sheriff's Office,
Ed | 2015-01-05 | 33 | 18.75 | 12.25 |  | Virginia Dept. of Forensic Science
Al | 2015-01-06 | 44 | 20.25 | 13.625
Al | 2015-01-06 | 11 | 20.4375 | 9.75 |  | Note: The variance was not decomposed into person and gage components.
Al | 2015-01-06 | 22 | 28.6875 | 10.9375
Al | 2015-01-06 | 33 | 18.75 | 12.25
Al | 2015-01-08 | 33 | 18.75 | 12.25
Al | 2015-01-09 | 44 | 20.25 | 13.625
Al | 2015-01-10 | 11 | 20.4375 | 9.75
Al | 2015-01-11 | 22 | 28.6875 | 11
Ed | 2015-01-08 | 33 | 18.75 | 12.25
Ed | 2015-01-09 | 44 | 20.25 | 13.625
Ed | 2015-01-10 | 11 | 20.4375 | 9.75
Ed | 2015-01-11 | 22 | 28.6875 | 11
Denise | 2015-01-08 | 33 | 18.75 | 12.25
Denise | 2015-01-09 | 44 | 20.25 | 13.625
Denise | 2015-01-10 | 11 | 20.4375 | 9.75
Denise | 2015-01-11 | 22 | 28.6875 | 11
Carl | 2015-01-08 | 33 | 18.75 | 12.3125
Carl | 2015-01-09 | 44 | 20.25 | 13.625
Carl | 2015-01-10 | 11 | 20.4375 | 9.75
Carl | 2015-01-11 | 22 | 28.6875 | 11
Betty | 2015-01-08 | 33 | 18.75 | 12.25
Betty | 2015-01-09 | 44 | 20.25 | 13.625
Betty | 2015-01-10 | 11 | 20.375 | 9.6875
Betty | 2015-01-11 | 22 | 28.6875 | 10.9375
Ed | 2015-01-14 | 22 | 28.6875 | 11
Ed | 2015-01-14 | 33 | 18.75 | 12.25
Ed | 2015-01-14 | 44 | 20.25 | 13.625
Ed | 2015-01-14 | 11 | 20.375 | 9.75
Denise | 2015-01-18 | 22 | 28.6875 | 10.9375
Denise | 2015-01-18 | 33 | 18.75 | 12.25
Denise | 2015-01-18 | 44 | 20.25 | 13.5625
Denise | 2015-01-18 | 11 | 20.4375 | 9.75
Carl | 2015-01-22 | 22 | 28.6875 | 11
Carl | 2015-01-22 | 33 | 18.75 | 12.1875
Carl | 2015-01-22 | 44 | 20.1875 | 13.625
Carl | 2015-01-22 | 11 | 20.4375 | 9.75
Betty | 2015-01-26 | 22 | 28.6875 | 11
Betty | 2015-01-26 | 33 | 18.75 | 12.25
Betty | 2015-01-26 | 44 | 20.1875 | 13.625
Betty | 2015-01-26 | 11 | 20.4375 | 9.75
Al | 2015-01-31 | 22 | 28.6875 | 11
Al | 2015-01-31 | 33 | 18.75 | 12.25
Al | 2015-01-31 | 44 | 20.25 | 13.625
Al | 2015-01-31 | 11 | 20.4375 | 9.75
Example Spreadsheet for Calculating Uncertainty for Measurement of Overall Length and Barrel Length
November 14, 2017
The spreadsheet begins on Sheet 1 with simulated data involving 80 measurements each for barrel length and overall length, shown in columns D and E. The 80 measurements include five observers taking four measurements each on four firearms. The calculations here yield a mean value for the overall length and barrel length of each firearm and a Type A standard uncertainty (Stdev) calculated from 80 measurements by all observers. The four values of standard uncertainty are pooled to yield a value of statistical reproducibility for OL and BL measurements, shown in yellow. This value is carried onto Sheet 2 as one component of a summary uncertainty budget tabulated in ASCLD/LAB recommended format.

## OL Summary ASCLD-LAB format rev
Measurement Uncertainty Estimation Form
Measurement: |  |  | Overall Length of a Firearm
Range of measurement values: |  |  | 19 inch to 28 inch
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2017-11-14
Attention! For explanatory notes see rows 25-39 on the BL Summary (next sheet).
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Standard Uncertainty | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 0.045822 | inch | A | normal | 1 | >60 | 0.045822 | 86.4282
2 | Length scale readability (1/16 inch) | 0.0625 | inch | B | rectangular | 3.46 | ∞ | 0.018064 | 13.4312
3 | Measuring scale calibration uncertainty | 0.0014 | inch | A | normal | 1 | 19 | 0.0014 | 0.08068
4 | Reference length standard calibration uncertainty (obtained from vendor) | 0.0002 | inch | B | normal | 2 | ∞ | 0.0001 | 0.00041163
5 | Thermal expansion of measuring tool for ± 5 °C (conservative choices) | 0.00208 | inch | B | rectangular | 1.73 | ∞ | 0.0012023 | 0.059503
6
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.049289
| Expanded Unc | U (k=2) | [5 empty] | 0.098577
| Expanded Unc | U (k=3) | [5 empty] | 0.14787
| Reported Uncertainty: | 0.098577 |  | k=2
| Reported Uncertainty: | 0.14787 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
The basis for the data above:
1 | See explanatory notes on the BL spreadsheet.
2
3
4
5
6
7
8
9
10
Revision: | 2017-11-14
Note on C7: Standard deviation or outside limit Cells formatted to display 8 decimal places
Note on D7: All units must be the same or all must be a percentage
Note on E7: Type A Evaluation - by statistical means of data from the measurement process. Type B Evaluation - by other means.
Note on F7: Normal and Rectangular are the most common Distribution types. If unknown, Rectangular Distribution is the conservative choice.
Note on G7: The Divisor that converts the value of the uncertainty component to a standard uncertainty (equivalent to 1 standard deviation)
Note on I7: Equal or equivalent to one standard deviation
Note on C19: Limit to at most two significant figures unless documented rationale for reporting additional significant figures. The laboratory procedure for estimating MU must include the process for rounding the expanded uncertainty

## BL Summary ASCLD-LAB format rev
Measurement Uncertainty Estimation Form
Measurement: |  |  | Barrel Length of a Firearm
Range of measurement values: |  |  | 9.7 inch to 13.5 inch
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2017-11-14
Attention! See explanatory notes on rows 25-39.
Line Item | Uncertainty Component (Notes: see row 32) | Value (Notes: row 33) | Units (row 34) | Type (row 35) | Statistical Distribution (row 36) | Divisor (row 36) | Degrees Freedom (n-1) (row 37) | Standard Uncertainty (row 38) | Component Contribution % (row 39)
1 (Notes: row 25) | Measurement Process Reproducibility (from page 1) | 0.025066 | inch | A | normal | 1 | >60 | 0.025066 | 65.637
2 (Notes: row 26) | Length scale readability (1/16 inch) | 0.0625 | inch | B | rectangular | 3.4641 | ∞ | 0.018042 | 34.0062
3 (Notes: row 27) | Measuring scale calibration uncertainty | 0.0014 | inch | A | normal | 1 | 19 | 0.0014 | 0.20476
4 (Notes: row 28) | Reference length standard calibration uncertainty (obtained from vendor) | 0.0002 | inch | B | normal | 2 | ∞ | 0.0001 | 0.0010447
5 (Notes: row 29) | Thermal expansion of measuring tool for ± 5 °C (conservative choices) | 0.00208 | inch | B | rectangular | 1.73 | ∞ | 0.0012023 | 0.15101
6 (Notes: row 30) | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.030939
| Expanded Unc | U (k=2) | [5 empty] | 0.061879 |  |  | 0.143
| Expanded Unc | U (k=3) | [5 empty] | 0.092818
| Reported Uncertainty: | 0.061879 |  | k=2
| Reported Uncertainty: | 0.092818 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
The basis for the data above:
1 | The first component of uncertainty is an estimate of the standard deviation of the repeated measurements by the observers in the study. It is copied from cell K39 shown in yellow on sheet one.
2 | This line shows the estimated uncertainty based on scale resolution. For 1/16 inch gradations, the value that is read could range over the full 1/16 inch space between gradations. For 1/8 inch gradations, the estimate should be larger; for electronic readout, the estimated value here could be much smaller.
3 | The measuring scale, be it a ruler or digital gauge, could have an uncertainty obtainable from the vendor or from a calibration lab. In this example, the uncertainty is obtained by comparing a length on the measuring scale with the length of a reference standard 20 times. The calculated standard deviation is 0.00140 and is shown in Column C.
4 | The reference length standard should have an uncertainty obtainable from the vendor or from a calibration lab. In this case, the value in Column C has a stated uncertainty from the vendor of 0.00020 inch. That represents a 2 standard deviation or 95 % confidence interval, and the result seems to be derived from statistical considerations and from multiple sources of uncertainty. So a normal distribution of expected values is assumed.
5 | Lengths are defined at the standard temperature of 20°C. You can make a small error in a length measurement if the measuring tool has a different coefficient of thermal expanision than the object being measured, and the ambient temperature is not 20°C. The measuring tool and the object will expand differently. We make a worst case estimate of the possible error by assuming a stainless steel measuring scale 26 inches long, a barrel whose length does not change with temperature, and a temperature anywhere between 15°C and 25°C. The thermal expansion coefficient for stainless steel is approximately 0.0016 %/°C. The possible error is then ± (0.0016)(%/°C)x(5°C)x(26 inch) = ± 0.00208 inch.
6 | other sources, etc.
Column B shows the different sources of uncertainty that are considered significant.
Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation.
Column D: The units of the source components in this template are all assumed to be the same units of length. If one source uncertainty component has different units from the others, such as % or meters, the values need to be converted into equivalent values in inches in this case, and the converted values need to be recorded in column C for combination with the other uncertainty components.
Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B.
Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C as in line 8, then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in line11, then the divisor is √3. If the uncertainty value came from someone else like the vendor in line 10, then the vendor's 2 standard deviation estimate must be divided by 2.
Column H: The more data points you have, the more confident you can be that your statistical calculations of 95 % confidence intervals are accurate. The degrees of freedom is used for this. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f.
Column I = Column C / Column G. All the stated uncertainty values in Column C have been converted into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I16 here.
Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty.
Revision: | 2017-11-14
Note on C7: Standard deviation or outside limit Cells formatted to display 8 decimal places
Note on D7: All units must be the same or all must be a percentage
Note on E7: Type A Evaluation - by statistical means of data from the measurement process. Type B Evaluation - by other means.
Note on F7: Normal and Rectangular are the most common Distribution types. If unknown, Rectangular Distribution is the conservative choice.
Note on G7: The Divisor that converts the value of the uncertainty component to a standard uncertainty (equivalent to 1 standard deviation)
Note on I7: Equal or equivalent to one standard deviation
Note on C19: Limit to at most two significant figures unless documented rationale for reporting additional significant figures. The laboratory procedure for estimating MU must include the process for rounding the expanded uncertainty