## Slide 1: Calibration
Lane C. Sander, Ph.D.
Chemical Sciences Division
Material Measurement Laboratory
National Institute of Standards and Technology1
1Contribution of the National Institute of Standards and Technology. Not subject to copyright. Certain commercial equipment, instruments, or materials are identified to specify adequately the experimental procedure. Such identification does not imply recommendation or endorsement by the National Institute of Standards and Technology, nor does it imply that the materials or equipment identified are the best available for the purpose.
Notes: In the context of chemical metrology, calibration is the process of relating a known quantity of an analyte to the corresponding measured instrumental response through a mathematical relationship. Calibration permits the assignment of analyte levels in unknown samples based on the known levels of the calibrants.
This presentation will provide an overview and application of various calibration models, with recommendations of ways to improve measurement accuracy.

## Slide 2: Approaches to Quantitation
External Standard Calibration Approach
Calibrants
Unknown
Calibration Curve
amount
response
R
A
X
[
]
A
X
amount
response
Notes: The basis for external standard calibration is conceptually very simple. Instrumental response can be related to the quantity of analyte present. Usually this response is linearly related to the mass of the analyte. If one or more samples are prepared with known composition, the response of the instrument can be calibrated.
The figure on the left shows chromatograms for four calibrants with different concentrations. Chromatographic response is characterized by the area of the peak corresponding to the analyte, which is plotted as a function of the quantity of the analyte, as is shown on the figure at the right. Quantity is commonly reported in units of mass fraction, mole fraction, mass concentration, or amount of substance concentration, but for simplicity, the generalized term “concentration” will be used to encompass all of these units. Thus, the area response for the analyte in an unknown sample can be related to the concentration through the calibration curve.

## Slide 3: Approaches to Calibration
Considerations for linear and nonlinear instrumental response
  Linear models can be used for most instrumentation
  Nonlinear response can often be modeled as linear, for sufficiently small ranges
Mathematical models
  Zero-Intercept Model
  Fixed-Intercept Model
  Unconstrained (Fitted) Model
  Bracketed Calibration
Notes: Several common approaches are used in calibrating methods. Most frequently, calibration models are based on linear instrumental response. However, when instrumental response is nonlinear, a linear model may still provide accurate results if the calibration range is sufficiently limited.
Details of the calibration model are important to achieve accurate results. Knowledge of the applicability and limitations of the various approaches can inform the selection of the most appropriate model for specific applications.
1) Four models can be discussed. If no instrumental response is observed for a blank calibrant, then the mathematical relation can be described as a zero-intercept model. However, if a response is observed for the blank calibrant, the mathematical relation can be described as a fixed-intercept model. If both the slope and the intercept are determined based on the calibration data, the mathematical relation can be described as having an calculated slope and intercept. A subset of these models can be described when the level of the unknown sample is known within a defined interval. For such cases, a linear relation is defined by calibrants placed at the upper and lower bounds of the interval.
In the next series of slides, these closely related models will be discussed in detail.

## Slide 4: Calibration
cal
cal
cal
cal
unknown
Zero-Intercept
Model
Recommended
when no response is
observed for a zero-
level (blank) calibrant
Fixed-Intercept
Model
cal
cal
cal
unknown
Recommended when
the response for a zero-
level calibrant can be
attributed to a constant
interference
Calibrants clustered near unknown
Bracketed Calibration
Best accuracy is provided
for all calibration models
Knowledge of unknown is
required
Unconstrained
(Fitted) Model
cal
cal
cal
cal
Slope and intercept are calcu-
lated to give best fit to data
Low-levels may exhibit poor
accuracy
unknown
Calibration should be guided by knowledge of unknown samples
Calibrant levels should be similar to unknowns
Notes: In planning the preparation of calibration solutions, the concentration range to be targeted is guided by knowledge or assumptions about the concentration of the unknown samples. Ideally the levels in the calibrants should be similar to the levels in the unknowns. This minimizes any issues of linearity of instrumental response.
1) In the first example, the intercept is fixed at zero. This is the best approach if no response is observed for the blank calibrant, and linear response has been established.
2) In the next example, a fixed-intercept is determined based on a non-zero response for a blank calibrant. This type of model can result when the solvent or other matrix used to prepare the calibrants is contaminated with the analyte.
3) In the third example, the calibration data is fit to a line, typically by using simple linear regression with the ordinary least squares method. This method minimizes the sum of the squared deviations between the data points and the fitted line; however, other more complex methods exist. Both the slope and the intercept are determined by this approach.
4) Finally, in the last example, calibration solutions are prepared that approximate the level estimated for the unknown. Bracketing often provides the best accuracy, but this approach requires some foreknowledge of the properties of the unknown.

## Slide 5: Zero-Intercept Model: y = mx
Appropriate when the zero-level calibrant (blank) does not produce an instrumental response, and the overall response is linear
Intercept is fixed at zero; the slope is varied to fit data
Fixed (zero) intercept promotes accuracy of low-level analytes
Approaches
  Averaged response factors
  Linear regression
amount
response
variations
in slope
Notes: The zero-intercept model is appropriate when little or no response is observed for zero-level (blank) calibrants, and the overall instrumental response is linear
1) In this calibration model, the intercept is fixed at zero, and the slope is varied to fit the calibration data.
2) A significant advantage is provided for low-level analytes by constraining the intercept to be zero. This effectively changes the weighting of the model to favor the accuracy of low-level analytes.
3) Different mathematical approaches can be utilized to develop the calibration model; namely, averaged response factors and linear regression.

## Slide 6: Calibration Models
Zero-intercept
linear regression
slope of regression line
Cal 1
Cal 2
Cal 3
Cal 4
Unweighted: dilute calibrants are less important than concentrated calibrants
Linear response is assumed
Fixed (zero) intercept
y = mx
averaged
response factors
Avg RF = (RF1 + RF2 + RF3 + RF4) / 4
RF2
RF3
RF4
RF1
Cal 1
Cal 2
Cal 3
Cal 4
All calibrants are weighted equally
Linear response is assumed
Fixed (zero) intercept
Notes: In the figure at the left, 4 calibration solutions are shown. A response factor can be calculated for each calibrant by dividing the analyte response by the level of that analyte. These response factors can be averaged to give an average response factor. This approach weights all calibrants equally. It also assumes a linear instrumental response, and a zero-intercept.
Another common approach is to use a linear regression model with a zero-intercept as shown at the right. For the four calibrants levels shown in the first example, a line can be calculated that passes through zero. Dilute calibrants are less important than concentrated calibrants, so if one of the calibrants is very concentrated compared with the other calibrants, the concentrated calibrant will dominate the calculation of the linear regression more than the dilute calibrants. As with averaged response factors, linear response is assumed, and a zero-intercept is fixed.

## Slide 7: Fixed-Intercept Model: y = mx + b
Appropriate when the zero-level calibrant (blank) produces a nonzero instrumental response, and this bias is present in all calibrants. The overall instrumental response must be linear.
Intercept is based on blank response; slope is varied to fit data
Constrained (non-zero) intercept favors accuracy of low-level analytes
Approaches
  Averaged response factors
  (equal weighting for low-level and high-level calibrants)
  Linear regression
  (greater weighting for high-level calibrants)
amount
response
variations
in slope
Notes: If a blank level calibrant produces a nonzero response and the source of the response can be identified, the use of a different mathematical model y=mx+b may be more appropriate. When the bias is constant, the calibration curves are offset by a known amount “b”. In the calibration model, the value for b should be constrained so that the response curve always passes through the y axis at this fixed level. This approach is analogous to the zero-intercept model, when the blank response has been subtracted from each of the calibrants. Thus, the corrected response can be written as (Y-b) = mX.
1) As an example, suppose that the solvent used to prepare the calibrant is contaminated with the analyte to be measured. In this case, each calibrant would produce a constant biased response due to the presence of the contaminant.
2) As with the zero-intercept model, the accuracy of the low-level analytes is promoted by constraining the intercept.
3) The same approaches described for the zero-intercept model can also be used for the fixed-intercept model. Thus, averaged response factors and linear regression also apply to the constrained-intercept model.
If linear regression is used to determine the slope of the line (m), the value of b should be specified explicitly rather than determined as a consequence of the fitting algorithm.

## Slide 8: Unconstrained (Fitted) Model: y = mx + b
Both the slope and intercept are unconstrained to provide the best fit to the data.
The calibration model is valid only between the upper and lower bounds of the calibrants.
High-level calibrants are weighted more strongly than low-level calibrants.
amount
response
calibrants
variations
in slope
variations
in intercept
Notes: In some cases, a calibration model may be warranted that is based on a calculated intercept. This approach emphasizes a limited range of concentrations to achieve the best fit to the data. Typically, this approach gives more weight to higher concentration data, and may result in poor accuracy for low-level analytes.
When this approach is used, it is essential to evaluate the validity of the model for different concentration ranges.
For best results, the calibration model should be limited for use only between the upper and lower bounds of the calibrants.

## Slide 9: Bracketed Calibration
Calibrants are prepared within a narrow interval to enclose the expected levels of the unknowns
The model should not be used (extrapolated) outside of the (yellow) calibrated interval
Bracketed calibration provides the most reliable and accurate results
Multiple calibration solutions may be required to cover extended concentration ranges
amount
response
calibrants
variations
in slope
variations
in intercept
Notes: When the levels of the unknowns are expected within a known interval, calibrants can be prepared to bracket this interval.
Zero-intercept, fixed-intercept, and unconstrained (fitted) approaches can be used with bracketed calibrants; however, as will all calibration approaches, the model should not be extrapolated for use outside of the calibrated interval.
This approach provides the most robust calibration model; however, knowledge of the unknown is required, and additional calibrants are required when unknowns fall outside of the narrow calibration window.
Because the calibration interval is limited, additional calibration solutions may be required to cover the desired range of unknowns

## Slide 10: Extended Calibration Range Issues
(Generated data shown to illustrate potential error sources)
Linear Regression Model
y = mx + b
using points 0 to 1000
y = 0.9755x + 3.6453
This example:
Fitted line is above the responses
for the low-concentration calibrants
Measurements are biased low
0
200
400
600
800
1000
1200
1400
0
200
400
600
800
1000
R² = 0.9999
R
E
S
P
O
N
S
E
CONCENTRATION
Plot enlarged to show low-concentration data
CONCENTRATION
0
2
4
6
8
10
12
14
0
2
4
6
8
10
R
E
S
P
O
N
S
E
All data fitted
x
y
0
1.2
1
2.3
2
3.4
5
6.6
7
8.8
10
12
20
19
50
53
70
75
100
107
200
207
500
493
700
687
1000
976
100 ´ magnification
Notes: In the following series of slides, an artificial data set was produced to illustrate potential error sources.
In this example, data is plotted that spans three orders of magnitude, from 0 to 1000. The linear regression coefficient for this data was determined to be R squared equal to 0.9999. Based solely on this regression coefficient, it might be thought that the resulting linear model would provide accurate results.
1) However, if the range shown in yellow from 0 to 10 is magnified, the model is seen to be significantly biased for low-level data. This is a consequence of the application of linear regression to data for an unconstrained model. This approach weights high-level concentrations with greater importance than low-level data. In this example, low-level measurements are biased low, but other data might provide different biases.

## Slide 11: Extrapolation Issues
(Generated data shown to illustrate potential error sources)
This example:
Dashed line is extrapolated from
low-concentration data to higher concentrations
Linear Regression Model
y = mx + b
using points 0 to 10
y = 1.0909x + 1.189
CONCENTRATION
0
2
4
6
8
10
12
14
0
2
4
6
8
10
100 ´ zoom-out
R² = 1.0000
Low-conc. data
0
1.2
1
2.3
2
3.4
5
6.6
7
8.8
10
12
20
19
50
53
70
75
100
107
200
207
500
493
700
687
1000
976
x
y
Notes: It is useful to examine this data set from another perspective. A subset of the data, shown in yellow, is plotted over the interval from 0 to 10. If this data is modeled by linear regression, a perfect regression coefficient of R squared = 1 is determined.
1) If the linear relation is plotted for the interval 0 to 1000, shown in blue, it is evident that the extrapolated model is biased for the high-level data.
These examples illustrate the importance of evaluating the calibration data over the intervals that the data will be used.

## Slide 12
Influence of the Calibration Model
y = mx+b vs y = mx
Data fit over 1 order of magnitude
0
2
4
6
8
10
12
14
0
2
4
6
8
10
y = 1.091x + 1.189
y = 1.257x
y=mx+b model 0-10
y=mx model 0-10
R2 = 0.9999
R2 = 0.9876
Relative intercept is large
y = mx model is not appropriate
Data fit over 3 orders of magnitude
200
400
600
800
1000
0
400
600
800
1000
0
200
y = 0.976x + 3.645
y = 0.981x
y=mx+b model 0-1000
y=mx model 0-1000
R2 = 1.0000
R2 = 0.9998
Relative intercept is small
Model choice is less important
Data fit over 2 orders of magnitude
0
20
40
60
80
100
0
20
40
60
80
100
y = 1.056x + 0.889
y = 1.069x
y=mx+b model 0-100
y=mx model 0-100
R2 = 0.9991
R2 = 0.9992
Relative intercept is small
Model choice is less important
Notes: The question can be asked “How does the choice of the calibration model affect potential biases for different sets of calibration data?”
In the following figures, the data from the previous slide is fit to zero-intercept and unconstrained data models, shown as dashed and solid lines, respectively.
1) For the interval 0 to 1000, regression coefficients are approximately 1, and the corresponding lines are nearly identical.
2) If a subset of the data is considered over the interval 0 to 100, the regression coefficients are only slightly decreased, and again the corresponding lines are nearly identical.
3) For the interval 0 to 10, the regression coefficients differed significantly, and it is apparent that the zero-intercept model is not appropriate for this data.

## Slide 13: Recommendations
Verify response linearity
If a range of unknown samples are to be determined, utilize the fixed-intercept model
  Determine the intercept with a zero-level (blank) calibrant
  Determine the slope with one or more calibrants at the maximum level expected for unknown samples
  Utilize linear regression or averaged response factors based on desired weighting of data
If unknown sample levels are expected within a defined interval, bracket calibrants to enclose this interval.
Notes: A few recommendations can be made based on the potential advantages and limitations of the various calibration models.
1) These recommendations assume linear instrument response over the relevant measurement range.
2) In most cases, good accuracy at low and high levels can be achieved by using the fixed-intercept calibration model.
Determine the intercept with a zero-level blank calibrant. For calibrants prepared in pure solvents, the instrumental response will likely be negligible, and the calibration model will effectively be a zero-intercept model.
Estimate the upper limit for the levels in the unknown samples, but don’t significantly exceed this upper limit.
If intermediate calibrants are prepared, carefully consider how to utilize the data. The averaged response factor approach weights low-level and high-level calibrants equally; whereas linear regression gives more importance to high-level calibrants, especially with the fixed-intercept model.
Finally, if unknown sample levels are expected to occur within a defined interval, bracket the calibrants to enclose this interval.
To conclude, aspects of the different calibration approaches can have significant impact on measurement accuracy. It is important consider the measurement goals so that the results will be fit for purpose.

## Slide 14: www.nist.gov
Notes: To learn more about the National Institute of Standards and Technology, visit www.nist.gov