## Software Technical Assessment
Software Technical Assessment
Software Description | WAQCF-022 (Rev. 01), Mass Control Chart.xlt | Approval Date | 2011-06-20
Assessed By | Dan Wright | Approved By | Dan Wright
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | Very well laid out
| There are instructions for use | Pass | Instruction worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | Clearly marked worksheets for instructions and data input
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank data entry cells, light blue for non-blank data entry cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Chart Procedure (SOP 9) and Measurement Procedure are both specified
| The number of digits to be rounded to is specified | N/A | No rounding
| The user is warned/notified whenever data-entry fields are left blank | Pass | Visual, blank data-entry cells are color coded
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry cells are blank on the template, except percent of tolerance which has a default of 33%
| The software opens at the right location within the file | Pass | Opens to the instruction worksheet
| Unused fields/cells are locked | Pass | All unused cells are locked
| Rows/columns that the operator need not see are hidden | Pass | Yes
| Unused sheets are removed | Pass | Yes
| Worksheets are named appropriately | Pass | Yes
B. Mathematical Specification | The correct SOP is used | Pass | SOP 9 and appropriate measurement SOP
| The formulae and methods chosen from that SOP are specified | Pass | SOP 9 method as explained on Instructions worksheet
| Sources and references for formulae are specified | Pass | SOP 9
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | N/A
C. Code review | The formulae in the fields exactly match the SOP | Pass | Upper / Lower control limits
| Repeated calculations appropriately reference the correct cells | Pass | Uses absolute cell references
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | N/A
| Rounding is done at the appropriate locations in the file | N/A | No rounding
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | N/A
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | N/A
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | N/A
E. Component Testing | Each macro used is functional | N/A | No macros
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | N/A
| Plotted graphs are accurate | Pass | Charts depict correct information
| Worksheets/reports print properly, if needed to | Pass | All worksheets and charts are formated for printing
| Conditional (color and non-color) formatting is functional | Pass | For blank and nonblank cells and Precision Test
F. Numerical Reference | Look-up tables and lists match the latest calibration report | Pass | Standards workbook is validated to latest calibration reports and Mass Tolerances workbook is validated to latest references
| Uncertainties match the latest Scope | N/A
| Values that reference another workbook or spreadsheet are dated | N/A
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct | N/A | No embedded data
| The evaluation of the embedded data is dated and documented | N/A
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | N/A
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | Pass | Basic arithmetic calculations
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Locked and password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | Password hint in instructions
| Files are backed up automatically | Pass | Manual backup
| Additional back-up is available at alternate facilities | Pass | Backed-up to agency network drives
| Files on network drives cannot be accidentally deleted | Pass

## Documentation
Documentation
Date | Description | Initials
2010-08-31 | Developed by Dan Wright, State Metrologist, Washington Metrology Laboratory | DAW
2010-08-31 | Validated workbook, file name 'WA010-08-31 WAQCF-022 (Rev. 00), Mass Control Chart Workbook V&V'.pdf on Laboratory Computer C:\ drive and Agency Server H:\ drive and a paper copy in laboratory files. | DAW
2011-06-20 | Added Best Uncertainty error bars and precision test. | DAW
2011-06-20 | Validated links and range names. | DAW
2011-06-20 | Validated workbook, file name 'WA011-06-20 WAQCF-022 (Rev. 01), Mass Control Chart Workbook V&V'.pdf on Laboratory Computer C:\ drive and Agency Server H:\ drive and a paper copy in laboratory files. | DAW

## Instructions
Instructions
No. | Description
1 | Enter requested information in the light yellow cells rows 1 through 8 on the Data Entry worksheet.
2 | Enter data from your measurements in the light yellow cells in rows 11 down column A through K. After entering data highlight columns L through Z on the row above the one your entering data and copy down.
3 | Place the cursor in column A where it states Insert Row Here and insert a row. This will be the next row you will enter new data.
4 | Modify the 'Percent of Tolerance for Control Chart' to adjust the limits just outside the UCL and LCL. This gives a better visual of your data points.
5 | Modify horizontal chart axis on the temperature and humidity charts to match the requirements.
6 | Do not change the formulas in M1 through M4, these are used for the control chart titles.
7 | Password protected. Password: password
8 | The formulas in cells C3, C5, and I4 on the Data Entry worksheet is linked to the Standards workbook.
9 | The formulas in cells D2 through D20 on the Tolerance Lookup worksheet are linked to the Mass Tolerances workbook.
10 | The control chart parameters for Mass Echelon III are based on the tolerance and Mass Echelons I & II are based on the standard deviation as described in SOP 9.
11 | Name files as follows; Balance (ID), Load (ID), SOP, e.g., C5000 (S8), 10 lb (CA18), SOP 8.
Range Names
Name | Range
ChkID | ='Data Entry'!$C$2
Spec | ='Data Entry'!$I$1
Spec.list | ='Tolerance Lookup'!$A$2:$A$20

## Data Entry
Laboratory: |  | Washington |  | Tolerance Class: |  |  |  | NIST HB 105-1, Class F |  |  |  | Do Not Delete, Used for Chart Titles and Legends
Check ID Code: |  | CA17 |  | Tolerance (g) = |  |  | 0.23 | % of Tol. for CC = |  | 33% |  | Washington ● Mass Echelon III ● X-Chart (SOP 9) 5 lb (CA17) Check Standard Balance: Sartorius C5000 ● Measurement Procedure: SOP 8 Standard Deviation = 0.001000 g ● NIST HB 105-1, Class F tolerance = ± 0.23 g
Check Description: |  | 5 lb |  | Best Unc. (k=2) (g) = |  |  | 0.012 | Precision Test (Pn) = |  | 0.15652 |  | Washington ● Mass Echelon III ● Check Standard vs Humidity Chart 5 lb (CA17) Check Standard Balance: Sartorius C5000 ● Measurement Procedure: SOP 8 Standard Deviation = 0.001000 g
Balance ID Code: |  | S8 |  | Check Standard Correction (g) = |  |  |  | 0.04706 |  |  |  | Washington ● Mass Echelon III ● Check Standard vs Temperature Chart 5 lb (CA17) Check Standard Balance: Sartorius C5000 ● Measurement Procedure: SOP 8 Standard Deviation = 0.001000 g
Balance Description: |  | Sartorius C5000 |  | Check Standard Mean = |  |  |  | 0.0466 |  |  |  | grams
Procedure: |  | SOP 8 |  | Number of Values = |  |  |  | 3 |  |  |  | Control Limits (mean ± 1/4 tolerance) | Warning Limits (mean ± 1/10 tolerance) |  |  | Check Observed Value ± Best U (k=2)
Parameter: |  | Mass Echelon III |  | Standard Deviation = |  |  |  | 0.001 |  |  |  | Check Mean ± 33 % tolerance
Control Chart Unit: |  | g |  | Standard Bias / Drift = |  |  |  | -0.00045995
Tech. Initials | Date | Report Number | Check Observed Value | Balance ID | Start Temp. (º C) | Finish Temp. (º C) | Start Pressure (g/cm³) | Finish Pressure (g/cm³) | Start Humidity (% RH) | Finish Humidity (% RH) | Average Air Temp. (º C) | Range (º C) | Average Pressure (g/cm³) | Range (g/cm³) | Average Humidity (% RH) | Range (%RH) | LCL (mean - 1/4 tolerance) | LWL (mean - 1/10 tolerance) | UWL (mean + 1/10 tolerance) | UCL (mean + 1/4 tolerance) | Check Mean | Mean + 33 % of tol | Mean - 33 % of tol | Calibrated Value | Best U (k=2)
DW | 2012-03-27 | 200446-0-L3237 | 0.0466 | S8 | 22.6 | 22.6 | 748.9 | 748.9 | 42 | 42 | 22.6 | 0 | 748.9 | 0 | 42 | 0 | -0.0109 | 0.0236 | 0.0696 | 0.1041 | 0.0466 | 0.1225 | -0.0293 | 0.04706 | 0.012
DW | 2012-04-11 | 200446-0-L3240 | 0.0476 | S8 | 22.3 | 22.3 | 749.9 | 749.9 | 43 | 43 | 22.3 | 0 | 749.9 | 0 | 43 | 0 | -0.0109 | 0.0236 | 0.0696 | 0.1041 | 0.0466 | 0.1225 | -0.0293 | 0.04706 | 0.012
DW | 2012-04-12 | 200446-0-L3249 | 0.0456 | S8 | 22.5 | 22.5 | 749.8 | 749.8 | 40 | 40 | 22.5 | 0 | 749.8 | 0 | 40 | 0 | -0.0109 | 0.0236 | 0.0696 | 0.1041 | 0.0466 | 0.1225 | -0.0293 | 0.04706 | 0.012
Insert Row Here

## Tolerance Lookup
Specification List (Spec.list)
ASTM E 617, Class 0
ASTM E 617, Class 1
ASTM E 617, Class 2
ASTM E 617, Class 3
ASTM E 617, Class 4
ASTM E 617, Class 5
ASTM E 617, Class 6
ASTM E 617, Class 7
NIST HB 105-1, Class F
NIST HB 105-8
Not Applicable
OIML R 111, Class E1
OIML R 111, Class E2
OIML R 111, Class F1
OIML R 111, Class F2
OIML R 111, Class M1
OIML R 111, Class M2
OIML R 111, Class M3
User/Manufacture