## Instructions and Disclaimer
Disclaimer
This spreadsheets is provided as a template to create metrology calculations according to the SOP represented here in the NIST OWM seminars. All calculations are entered and the validation the responsibility of the user. Nothing is protected or secure in this workbook.

## Change and Rev Control
Owner | NIST OWM
Tab | Change | Date | Initials
All | Creation of Templates base on SOPs from NISTIR 6969 in place in 3/2017. | 2017-03-23 | GH
All | Updated based on use for Validation exercise/demo in May 2017 Mass seminar. Nothing is protected! | 2017-05-24 | GH
All | Evaluated against other spreadsheets; no changes. | 2018-04-17 | GH
All | Updated based on class and TA inputs in May 2018. | 2018-06-05 | GH

## Instructions
This spreadsheet is set up for people to use the values on each tab and create equations for each of the SOPs on the worksheets.
| Worksheet | Applicable SOP | Instructions
1 | MABC = Magnitude of the Air Buoyancy Correction | 2 | Fill in the yellow cells on each worksheeet with values/equations of intermediate calculation values as requested.
2 | Effective Density | 2, 28
3 | Air Density | 2
4 | Mass and Conventional Mass Conversions | 2
5 | SOP 4 , Option A | 4
6 | SOP 5 | 5
7 | SOP 7, Option A | 7
8 | SOP 8 | 8
9 | Uncertainties | 29
*ABC = Air Buoyancy Correction

## Software V&V
Technical Assessment
This form is completed for each software file used in the laboratory. The form may be added as a tab/worksheet to a file/workbook or may be maintained in a separate electronic/paper file/location in the laboratory.
Codes | Assessment | Pass/Fail | Evidence Reviewed and Observed
A. Software Inspection | Spreadsheet is clear and makes sense
There are instructions for use
Instructions and data input appear on the visible portion of the first worksheet
Data-entry fields are labeled and color coded (it is recommended to avoid red and green)
The Standard Operating Procedure (SOP) used is clearly specified
The number of digits to be rounded to is specified
The user is warned/notified whenever data-entry fields are left blank
Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations
The software opens at the right location within the file
Unused fields/cells are locked
Rows/columns that the operator need not see are hidden
Unused sheets are removed
Worksheets are named appropriately
B. Mathematical Specification | The correct SOP is used
The formulae and methods chosen from that SOP are specified
Sources and references for formulae are specified
The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty
C. Code review | The formulae in the fields exactly match the SOP
Repeated calculations appropriately reference the correct cells
Calculations, when tested using standard data or reference test data, show appropriate accuracy
Rounding is done at the appropriate locations in the file
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences
Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa
“Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used
E. Component Testing | Each macro used is functional
Each command/button is functional
Combinations of interdependent macros are functional
Plotted graphs are accurate
Worksheets/reports print properly, if needed to
Conditional (color and non-color) formatting is functional
F. Numerical Reference | Look-up tables and lists match the latest calibration report.
Uncertainties match the latest Scope
Values that reference another workbook or spreadsheet are dated
When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct
The evaluation of the embedded data is dated and documented
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty
J. Security | Equation and calculation cells are protected against inadvertent editing
Cells are locked in place; they cannot be moved/dragged
Confidentiality of passwords is appropriate
Files are backed up automatically
Additional back-up is available at alternate facilities
Files on network drives cannot be accidentally deleted

## MABC
MABC - Magnitude of the Air Buoyancy Correction. | [6 empty] | SOP 2
Fill in the yellow cells with values/equations as appropriate.
Air density, ρa (mg/cm3) | 1.1795
Normal air density, ρn (mg/cm3) |  | look up this value in SOP 2
X Mass (g) | X Density (g/cm3) | X Volume (M/ρ) or cm3 | S Mass (g) | S Density (g/cm3) | S Volume (M/ρ) or cm3 | MABC (mg)
1000 | 7.85 |  | 1000 | 8.03
0.2 | 8 |  | 0.2 | 8
22679.6185 | 7.2 |  | 22679.6185 | 8

## Effective Density
Effective Density - used for when weights are used together as a summation. |  |  |  | SOP 2 (also SOP 28)
Fill in the yellow cells with values/equations as appropriate.
Mass (g) | Density (g/cm3) | Calculated Volume (M/ρ)
0.1 | 16.6
0.05 | 16.6
0.03 | 2.7
0.02 | 2.7
SUMS
Effective Density

## Air Density
Air Density Calculation. | EXTRA CREDIT IF YOU GET BORED OR WANT A CHALLENGE!!! |  | SOP 2
Fill in the yellow cells with values/equations as appropriate.
Pressure (mmHg) | Temp (°C) | RH (%)
760 | 20 | 42
Pressure (Pa) | Temp (K)
Look up the Exact Conversions in SP 811 for mmHg to Pa.
Look up the Exact Conversions in SP 811 for °C to K.
Variables | Calculations
ρa
xv
psv
f
Z
Variable | Description | Verify these variables!
Ma | 0.028965
Mv | 0.018015
p | Pa
T | K
R | 8.3145
h | %
t | °C
A | 1.2379e-05
B | -0.019121
C | 33.9371
D | -6343.1645
a0 | 1.5812e-06
a1 | -2.9331e-08
a2 | 1.1043e-10
b0 | 5.707e-06
b1 | -2.051e-08
c0 | 0.00019898
c1 | -2.376e-06
d | 1.83e-11
e | -7.65e-09
Note on B5: Enter values
Note on B8: Problem with conversion factors. Nothing exact in SP 811, Circ 1038, or HB 44. Doesn't match the Mass seminar spreadsheet.

## Mx and CMx Conversions
Converting between CMx and Mx. |  |  | SOP 2
Fill in the yellow cells with values/equations as appropriate.
Given True Mass, Find the Conventional Mass.
Mx (g) | 1000.0045
ρx (g/cm3) | 7.85
Numerator Buoyancy Corr |  | A
Numerator (Mass X with buoyancy corr) |  | B
Denominator (conventional buoy corr) |  | C
CMx Calculated (g) |  | 3.2 Equation
Given the Conventional Mass, Find the True Mass.
CMx (g) | 1000.0016
ρx (g/cm3) | 7.85 |  |  | Rewrite and Insert the equation (using "insert equation") to solve for Mx.
Numerator Buoyancy Corr
Numerator |  |  |  | Solution:
Denominator
Mx Calculated (g)
𝑀_𝑥=(〖𝐶𝑀〗_𝑥 (1−0.0012/8))/((1−0.0012/𝜌_𝑥 ))
𝐵: 𝑀𝑥 (1−0.0012/𝜌_𝑥 )
𝐴: (1−0.0012/𝜌_𝑥 )
𝐶: (1−0.0012/8)

## SOP 7 Solution
Calculate dx, Mx and CMx using SOP 7, Option A with ABC, Equal Nominals and no Tare Weights |  |  |  | SOP 7
Fill in the yellow cells with values/equations as appropriate.
Givens
ρn | 0.0012 | g/cm3
Sample Values
ρa | 0.0011848 | g/cm3
Ms | 2000.0058 | g
Msw | 0.20001 | g
ρs | 7.85 | g/cm3
ρsw | 16.6 | g/cm3
ρx | 7.95 | g/cm3
Observations
O1 | 0
O2 | 0.0044
O3 | 0.2046
Intermediate Calculations |  |  | Equation
Measured difference, dx without ABC |  | g | A
Buoyancy on Sensitivity weight |  |  | B
Sensitivity Factor |  |  | C
dx with ABC |  | g | D (or A x C)
Mass of Standard, Ms with ABC |  | g | E
Ms + dx, Numerator |  | g | F (or E + D)
Buoyancy Correction for X, Denominator |  |  | G
Final Calculations
Mass of X, Mx |  | g | Eqn 3.2.2.1 (also: (E+D)/G)
Conventional Mass of X, CMx |  | g | Convert Mass to Conventional Mass
𝐴: (𝑂_2−𝑂_1)
𝐶: (𝑀_𝑠𝑤 (1−𝜌_𝑎/𝜌_𝑠𝑤 ))/((𝑂_3−𝑂_2))
𝐷: 𝑑_𝑥= (𝑂_2−𝑂_1)(𝑀_𝑠𝑤 (1−𝜌_𝑎/𝜌_𝑠𝑤 ))/((𝑂_3−𝑂_2))
𝐸: 𝑀_𝑠 (1−𝜌_𝑎/𝜌_𝑠 )
Calculation of Conventional Mass from Mass
𝐹: 𝑀_𝑠 (1−𝜌_𝑎/𝜌_𝑠 )+(𝑂_2−𝑂_1)(𝑀_𝑠𝑤 (1−𝜌_𝑎/𝜌_𝑠𝑤 ))/((𝑂_3−𝑂_2))
𝐺: (1−𝜌_𝑎/𝜌_𝑥 )
𝐵: (1−𝜌_𝑎/𝜌_𝑠𝑤 )

## SOP 4 Solution
Calculate dx, Mx and CMx using SOP 4, Option A with ABC, Equal Nominals and no Tare Weights |  |  | SOP 4, Sub X for Sc variables in 3.2.2.1
Fill in the yellow cells with values/equations as appropriate.
Givens
ρn | 0.0012 | g/cm3
Sample Values
ρa | 0.0011848 | g/cm3
Ms | 2000.0058 | g
Msw | 0.20001 | g
ρs | 7.85 | g/cm3
ρsw | 16.6 | g/cm3
ρx | 7.95 | g/cm3
Observations
O1 | 0
O2 | 0.0044
O3 | 0.2046
O4 | 0.2006
Intermediate Calculations
Measured difference, dx without ABC |  | g
ABC for Sensitivity weight
Sensitivity Factor
dx with ABC |  | g
Mass of Standard, Ms with ABC |  | g
Ms + dx, Numerator |  | g
Buoyancy Correction for X, Denominator
Final Calculations
Mass of X, Mx |  | g
Conventional Mass of X, CMx |  | g
Note from 3.1 Calculate the value for the test weight correction using the same equations, but substitute X for the Sc subscript for the unknown in equations 3.1.1. and 3.1.2. as appropriate.

## SOP 5 Solution
Calculate dx, Mx and CMx using SOP 5,using ABC, Equal Nominals and no Tare Weights |  |  |  |  | SOP 5
Fill in the yellow cells with values/equations as appropriate.
Givens
ρn | 0.0012 | g/cm3
Enter Values
ρa | 0.0011812 | g/cm3
Ms | 10 | g
Msw | 0.030007 | g
ρs | 7.95 | g/cm3
ρsw | 7.95 | g/cm3
ρx | 7.95 | g/cm3
Observations | S vs X | S vs Sc | X vs Sc
O1 | 0 | 0 | 0.002
O2 | 0.003 | -0.002 | -0.007
O3 | 30.009 | 30.005 | 30
O4 | 30.006 | 30.007 | 30.004
| a1 | a2 | a3
Eqn 3.2
within process standard deviation, sw |  | 0.577 (a1 - a2 + a3)
Measured and calculated difference, dx |  | Equation 3.8
Mass of Standard, Ms with ABC
Ms with ABC + dx (Numerator)
Buoyancy Correction for X (Denominator)
Calculated Mx |  | Equation 3.9
Calculated CMx
Calculation of Conventional Mass from Mass

## SOP 8 Solution
Calculate conventional mass corrections (Cx) for 5 weights using SOP 8 |  |  |  | SOP 8
Fill in the yellow cells with values/equations as appropriate.
Sample Values
Ms | 22679.6685 | g
Nominal | 50 | lb
Nominal |  | g
Cs |  | g
Observations and Corrections (g)
S | 0 | Cx Corrections (g) | Cx Corrections (mg)
X1 | -0.073
X2 | 0.258
X3 | 1.851
X4 | 2.439
X5 | -0.476 |  |  | NOTE: DO NOT use an Excel CONVERT function without Validating against SP 811 and providing objective evidence!!!
S | 0.001 | drift

## Uncertainty Solution
Calculate the standard uncertainty, the combined and expanded uncertainties for a mass calibration and round the final uncertainty. | [8 empty] | SOP 29 and SOP 4 and 5
Fill in the yellow cells with values/equations as appropriate.
Uncertainty Budget (CMx)
Variable | Uncertainty Component for X | Value | Unit | Reported Value | Unit | Distribution | Divisor | df | Standard uncertainty | Unit
s(p) | Process Uncertainty | 0.00488 | mg | 0.00488 | mg | Normal, 1s | 1 | 9 | 0.00488 | mg
u(s) | Uncertainty of Standard | 0.022 | mg | 0.022 | mg | Normal, 2s | 2 | 10000 | 0.011 | mg
u(sw) | Uncertainty of sw | 0.000385 | mg | 0.000385 | mg | Normal, 1s | 1 | 10000 |  | mg
u(abc) | Uncertainty of the buoyancy correction | [7 empty] | -1.955e-12 | mg2
| uc |  |  |  | mg | <-- Enter the RSS Equation, Note that the value in K9 is already squared!
| Use a k value for 9 degrees of freedom (95.45 %) - for now. |  |  | 2.32
| Expanded Uncertainty |  |  |  | mg | <-- Multiply the RSS Solution by the k value.
| Rounded Value |  |  |  | mg | Round the value per GLP 9.
Effective degrees of freedom will be covered during seminars.
Alternatively, you can read NISTIR 6969, Section 8 (2018) or the Guide to the Expression of Uncertainty in Measurement to discover the Welch-Satterthwaite equation.
SOP 29, 3.4 (Rectangular Distribution)
SOP 29, 3.5 (Combined Uncertainty)
SOP 29, 3.6 (Expanded)

## Linear Interpolations
To be completed in class, for interpolating table values.
|  |  |  | df |  | t
|  |  | X1 | 17 | y1 | 2.16
|  |  | x | 17.45 | y
|  |  | x2 | 18 | y2 | 2.15
| Example with t-Table | Example with Mass (in class)
𝑦=𝑦_1+(𝑥−𝑥_1 )((𝑦_2−𝑦_1)/(𝑥_2−𝑥_1 ))
y = the t value to be determined
x = the df of the y to be determined or given
y1 = the t value associated with x1
y2 = the t value of associated with x2
x1 = the df for the value less than x
x2 = the df for the value greater than x