## Cone Summary
Date: | 2019-03-26
Time: | 16:05
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B1S-Test2
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0413 |  | Start Test | 0 | 2019-03-26 16:13
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 6 | 2019-03-26 16:13
Sample Area | 1 | m^2 | Wetting | 217 | 2019-03-26 16:17
Cone Heat Flux | 75 | kW/m^2 | Dripping | 239 | 2019-03-26 16:17
Cone Orientation |  |  | Bottom Ignition | 352 | 2019-03-26 16:19
Sample Grid |  |  | Flame Out | 703 | 2019-03-26 16:25
Sample Thickness | 106 | mm
Sample Initial Mass | 51 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 24.2921 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.587 | kW/m2 | Average Heat Release Rate | 1.6982 | kW/m2
Total Heat Release | 1.0586 | MJ/m2 | Average Mass Loss Rate | 0.07915 | g/m2s
Peak Mass Loss Rate | 0.20329 | g/m2s | Effective Heat of Combustion | 21.5225 | MJ/kg
Total Mass Lost | 47.1457 | g | Average Specific Extinction Area | 103.9506 | m2/kg
Time to Ignition | 6 | s | Peak CO Yield | 0.075547 | kg/kg
Time to Peak Heat Release | 414 | s | Peak CO2 Yield | 6.4277 | kg/kg
Time to Flame Out | 703 | s | Average CO Yield | 0.026283 | kg/kg
Average Heat Release Rate | 1.5166 | kW/m2 | Average CO2 Yield | 1.9844 | kg/kg
Average Mass Loss Rate | 0.066493 | g/m2s
Effective Heat of Combustion | 22.4535 | MJ/kg
Dripping | 239 | s | Time to dripping
HRRD | 0.64081 | kW | HRR at dripping time
THRD | 0.30859 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.3244 | kW | AHRR from Ignition to Dripping
Wetting | 217 | s | Time to wetting
HRRW | 0.64193 | kW | HRR at wetting time
THRW | 0.29462 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.3672 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-26 16:05 | 1553630725.6857 | -454 | 0.20955 | 0.000444 | 0 | 93.1256 | 92.2422 | 129.224 | 21.2612 | 0.001822 | 0.001657 | -152.838 | 0.21999 | 8.4024 | 0.003262 | 0.004261 | 0.004088 | 5.4e-05 | 0.00122 | 0.29966 | 1.3674 | 2.5473 | 100.79 | 21.07 | 16.578 | 65.6592 | 100.3228 | 61.6915 | 0.004113 | 1.0544 | 0.025657 | 24.3338 | 0.085253 | 0.13087 | -7e-06 | -0.000583 | -0.000583 | -9.2642
2019-03-26 16:05 | 1553630726.6871 | -453 | 0.20955 | 0.000444 | 0 | 93.1256 | 92.2422 | 129.224 | 21.2612 | 0.001822 | 0.001657 | -152.838 | 0.21999 | 8.4024 | 0.003262 | 0.004261 | 0.004088 | 5.4e-05 | 0.00122 | 0.29966 | 1.3674 | 2.5473 | 100.79 | 21.07 | 16.582 | 65.6592 | 100.3228 | 61.6915 | 0.004114 | 1.0544 | 0.025657 | 24.3338 | 0.085253 | 0.13087 | -7e-06 | -0.000583 | -0.000583 | -9.2684
2019-03-26 16:05 | 1553630727.6868 | -452 | 0.20953 | 0.000439 | 0 | 93.192 | 92.3145 | 129.224 | 21.2612 | 0.001823 | 0.00176 | -152.838 | 0.22016 | 8.4004 | 0.003289 | 0.004443 | 0.004289 | 5.3e-05 | 0.001223 | 0.3314 | 1.385 | 2.5473 | 100.79 | 21.07 | 16.589 | 65.6953 | 100.9742 | 64.1776 | 0.004116 | 1.0466 | 0.025562 | 24.424 | 0.084624 | 0.13098 | 0 | -5.4e-05 | -5.4e-05 | -101.0758
2019-03-26 16:05 | 1553630728.6858 | -451 | 0.2095 | 0.000443 | 0 | 93.2053 | 92.3031 | 129.224 | 21.2612 | 0.001806 | 0.001744 | -152.838 | 0.21953 | 8.4037 | 0.003297 | 0.004333 | 0.004077 | 5.5e-05 | 0.001219 | 0.32244 | 1.3474 | 2.5474 | 100.79 | 21.07 | 16.592 | 65.2902 | 101.1792 | 63.7921 | 0.004116 | 1.0478 | 0.025577 | 24.4101 | 0.087006 | 0.1306 | -5e-06 | -0.000402 | -0.000402 | -13.4004
2019-03-26 16:05 | 1553630729.6863 | -450 | 0.20947 | 0.000443 | 0 | 93.2638 | 92.1451 | 129.224 | 21.2612 | 0.001807 | 0.001698 | -152.838 | 0.21966 | 8.4034 | 0.003314 | 0.00461 | 0.00442 | 5.7e-05 | 0.001229 | 0.35474 | 1.2966 | 2.5474 | 100.79 | 21.06 | 16.595 | 65.3166 | 101.5912 | 62.6737 | 0.004115 | 1.0513 | 0.025619 | 24.3695 | 0.10794 | 0.13068 | -5e-06 | -0.000361 | -0.000361 | -14.9458
2019-03-26 16:05 | 1553630730.6858 | -449 | 0.20948 | 0.000445 | 1e-06 | 93.3075 | 92.4685 | 129.224 | 21.2612 | 0.001837 | 0.001721 | -152.838 | 0.22031 | 8.4033 | 0.003271 | 0.004538 | 0.004362 | 5.5e-05 | 0.001228 | 0.30431 | 1.3118 | 2.5473 | 100.79 | 21.06 | 16.595 | 66.0253 | 100.5521 | 63.2469 | 0.004115 | 1.0495 | 0.025597 | 24.3903 | 0.08079 | 0.13106 | -8e-06 | -0.000648 | -0.000648 | -8.3521
2019-03-26 16:06 | 1553630731.6874 | -448 | 0.20946 | 0.000439 | 1e-06 | 93.4429 | 92.4312 | 129.224 | 21.2612 | 0.001851 | 0.001701 | -152.838 | 0.21982 | 8.4036 | 0.003266 | 0.004721 | 0.004492 | 5.5e-05 | 0.001232 | 0.33196 | 1.3156 | 2.5473 | 100.79 | 21.06 | 16.595 | 66.3702 | 100.4256 | 62.7649 | 0.004115 | 1.051 | 0.025616 | 24.3728 | 0.097379 | 0.13077 | -1e-06 | -0.000117 | -0.000117 | -46.3582
2019-03-26 16:06 | 1553630732.687 | -447 | 0.20951 | 0.000441 | 1e-06 | 93.4539 | 92.4873 | 129.224 | 21.2612 | 0.001814 | 0.001704 | -152.838 | 0.21988 | 8.4039 | 0.003357 | 0.004749 | 0.004565 | 5.3e-05 | 0.001241 | 0.35706 | 1.3632 | 2.5474 | 100.79 | 21.06 | 16.595 | 65.4805 | 102.6242 | 62.8266 | 0.004115 | 1.0508 | 0.025613 | 24.3751 | 0.092993 | 0.13081 | -4e-06 | -0.000319 | -0.000319 | -16.9328
2019-03-26 16:06 | 1553630733.6861 | -446 | 0.2095 | 0.000435 | 1e-06 | 93.4825 | 92.6372 | 129.224 | 21.2612 | 0.001812 | 0.001671 | -152.838 | 0.21918 | 8.4041 | 0.003235 | 0.004601 | 0.004475 | 5.3e-05 | 0.001243 | 0.29966 | 1.3982 | 2.5475 | 100.79 | 21.06 | 16.594 | 65.4165 | 99.6733 | 62.0438 | 0.004114 | 1.0533 | 0.025643 | 24.3466 | 0.081245 | 0.13039 | 5e-06 | 0.000334 | 0.000334 | 16.1455
2019-03-26 16:06 | 1553630734.6865 | -445 | 0.20946 | 0.00044 | 0 | 93.6277 | 92.8164 | 129.224 | 21.2612 | 0.001832 | 0.00168 | -152.838 | 0.21968 | 8.4023 | 0.003294 | 0.004874 | 0.004711 | 5.3e-05 | 0.001246 | 0.34971 | 1.425 | 2.5476 | 100.79 | 21.07 | 16.589 | 65.9074 | 101.1092 | 62.2401 | 0.004116 | 1.0526 | 0.025636 | 24.3538 | 0.077847 | 0.13069 | -1e-06 | -8.5e-05 | -8.5e-05 | -63.2078
2019-03-26 16:06 | 1553630735.6873 | -444 | 0.20949 | 0.000438 | 0 | 93.6744 | 92.8134 | 129.224 | 21.2612 | 0.001842 | 0.001685 | -152.838 | 0.21928 | 8.4025 | 0.00336 | 0.004886 | 0.00469 | 5.3e-05 | 0.001249 | 0.3246 | 1.4047 | 2.5476 | 100.79 | 21.07 | 16.585 | 66.1455 | 102.6987 | 62.3776 | 0.004115 | 1.0522 | 0.025631 | 24.3588 | 0.082593 | 0.13045 | 1e-06 | 7.5e-05 | 7.5e-05 | 71.743
2019-03-26 16:06 | 1553630736.6866 | -443 | 0.20947 | 0.000438 | 0 | 93.8361 | 92.9636 | 129.5162 | 21.2612 | 0.001832 | 0.001734 | -152.8154 | 0.22024 | 8.4071 | 0.003363 | 0.004923 | 0.004778 | 5.4e-05 | 0.001253 | 0.33787 | 1.3555 | 2.5475 | 100.79 | 21.07 | 16.579 | 65.9081 | 102.7711 | 63.5524 | 0.004113 | 1.0485 | 0.025615 | 24.4289 | 0.083556 | 0.13102 | 2e-06 | 0.000108 | 0.000108 | 50.2289
2019-03-26 16:06 | 1553630737.6855 | -442 | 0.20949 | 0.000435 | 0 | 93.8987 | 92.9687 | 129.4801 | 21.2612 | 0.001794 | 0.001715 | -152.8208 | 0.21956 | 8.4079 | 0.003354 | 0.004497 | 0.004289 | 5.6e-05 | 0.001257 | 0.33876 | 1.3935 | 2.5476 | 100.79 | 21.07 | 16.566 | 64.9959 | 102.5542 | 63.0969 | 0.00411 | 1.05 | 0.025628 | 24.409 | 0.089034 | 0.13062 | 5e-06 | 0.000355 | 0.000355 | 15.2004
2019-03-26 16:06 | 1553630738.6869 | -441 | 0.20949 | 0.000434 | 1e-06 | 94.0233 | 93.2057 | 129.4843 | 21.2612 | 0.001846 | 0.001726 | -152.8093 | 0.21966 | 8.4005 | 0.003336 | 0.005375 | 0.005206 | 5.5e-05 | 0.001254 | 0.34811 | 1.3752 | 2.5475 | 100.79 | 21.07 | 16.558 | 66.2352 | 102.1128 | 63.3547 | 0.004108 | 1.0492 | 0.025619 | 24.4188 | 0.07812 | 0.13068 | 6e-06 | 0.00044 | 0.00044 | 12.2561
2019-03-26 16:06 | 1553630739.6864 | -440 | 0.2095 | 0.000442 | 0 | 93.9911 | 93.2393 | 129.3611 | 21.2612 | 0.001831 | 0.001721 | -152.7901 | 0.21899 | 8.4016 | 0.003285 | 0.005061 | 0.004887 | 5.5e-05 | 0.001255 | 0.30143 | 1.4179 | 2.5476 | 100.79 | 21.07 | 16.538 | 65.8761 | 100.8857 | 63.2406 | 0.004103 | 1.0495 | 0.025611 | 24.403 | 0.071833 | 0.13028 | -4e-06 | -0.000284 | -0.000284 | -18.9261
2019-03-26 16:06 | 1553630740.6864 | -439 | 0.2095 | 0.000442 | 0 | 94.0439 | 93.2944 | 129.5629 | 21.2612 | 0.001824 | 0.00171 | -152.7749 | 0.21981 | 8.4032 | 0.003324 | 0.004818 | 0.004633 | 5.6e-05 | 0.001249 | 0.34058 | 1.3763 | 2.5475 | 100.79 | 21.07 | 16.527 | 65.7147 | 101.8158 | 62.9782 | 0.0041 | 1.0503 | 0.025641 | 24.4125 | 0.071573 | 0.13076 | -4e-06 | -0.000316 | -0.000316 | -17.0946
2019-03-26 16:06 | 1553630741.6868 | -438 | 0.20951 | 0.000433 | 0 | 94.2028 | 93.3169 | 129.7631 | 21.2612 | 0.001828 | 0.001687 | -152.7693 | 0.21899 | 8.4045 | 0.003339 | 0.004358 | 0.004204 | 5.4e-05 | 0.001244 | 0.35994 | 1.3638 | 2.5475 | 100.79 | 21.07 | 16.517 | 65.8163 | 102.1968 | 62.4096 | 0.004098 | 1.0521 | 0.025683 | 24.4107 | 0.084505 | 0.13028 | 8e-06 | 0.000594 | 0.000594 | 9.0513
2019-03-26 16:06 | 1553630742.6855 | -437 | 0.2095 | 0.000435 | 0 | 94.2659 | 93.7361 | 129.6108 | 21.2612 | 0.00179 | 0.001731 | -152.7657 | 0.22011 | 8.4011 | 0.003372 | 0.004493 | 0.004376 | 5.4e-05 | 0.001243 | 0.32565 | 1.3166 | 2.5474 | 100.79 | 21.07 | 16.496 | 64.8889 | 102.9792 | 63.4797 | 0.004093 | 1.0488 | 0.025627 | 24.4352 | 0.050411 | 0.13095 | 6e-06 | 0.0004 | 0.0004 | 13.5081
2019-03-26 16:06 | 1553630743.686 | -436 | 0.20949 | 0.000437 | 0 | 94.3099 | 93.5055 | 129.4714 | 21.2612 | 0.001784 | 0.001699 | -152.3708 | 0.21944 | 8.4044 | 0.003424 | 0.004483 | 0.00422 | 5.4e-05 | 0.001235 | 0.31885 | 1.324 | 2.5474 | 100.79 | 21.07 | 16.487 | 64.7563 | 104.2541 | 62.7149 | 0.00409 | 1.0512 | 0.025642 | 24.3943 | 0.076616 | 0.13054 | 3e-06 | 0.000226 | 0.000226 | 23.8874
2019-03-26 16:06 | 1553630744.6855 | -435 | 0.2095 | 0.000439 | 0 | 94.4947 | 93.5375 | 129.2726 | 21.2612 | 0.001835 | 0.001688 | -112.0487 | 0.22014 | 8.4028 | 0.003347 | 0.00424 | 0.004111 | 5.8e-05 | 0.00123 | 0.32908 | 1.3418 | 2.5475 | 100.79 | 21.07 | 16.479 | 65.9887 | 102.3806 | 62.4385 | 0.004088 | 1.052 | 0.025633 | 24.3655 | 0.09107 | 0.13096 | 0 | 4.8e-05 | 4.8e-05 | 111.5775
2019-03-26 16:06 | 1553630745.6856 | -434 | 0.20952 | 0.000437 | 0 | 94.5652 | 93.8642 | 129.4765 | 21.2612 | 0.001835 | 0.001654 | -69.9639 | 0.21966 | 8.408 | 0.003391 | 0.004811 | 0.004684 | 5.9e-05 | 0.001222 | 0.3553 | 1.3487 | 2.5475 | 100.79 | 21.07 | 16.464 | 65.9887 | 103.443 | 61.6298 | 0.004085 | 1.0546 | 0.025684 | 24.3554 | 0.066554 | 0.13068 | 2e-06 | 0.000141 | 0.000141 | 38.2024
2019-03-26 16:06 | 1553630746.6856 | -433 | 0.20949 | 0.000438 | 0 | 94.6069 | 93.7153 | 129.4094 | 21.2612 | 0.001799 | 0.001729 | -27.9027 | 0.2192 | 8.4031 | 0.003371 | 0.004836 | 0.004543 | 5.7e-05 | 0.00122 | 0.35778 | 1.3343 | 2.5474 | 100.79 | 21.07 | 16.459 | 65.1239 | 102.9683 | 63.4265 | 0.004083 | 1.0489 | 0.025609 | 24.4143 | 0.084692 | 0.1304 | 1e-06 | 6.2e-05 | 6.2e-05 | 87.0743
2019-03-26 16:06 | 1553630747.6867 | -432 | 0.20949 | 0.000441 | 0 | 94.7199 | 93.9715 | 129.4604 | 21.2612 | 0.001791 | 0.001738 | 14.1863 | 0.21982 | 8.4028 | 0.003399 | 0.004316 | 0.004199 | 5.8e-05 | 0.001221 | 0.32693 | 1.3683 | 2.5475 | 100.79 | 21.07 | 16.452 | 64.9334 | 103.6346 | 63.6443 | 0.004082 | 1.0483 | 0.025606 | 24.427 | 0.07096 | 0.13077 | -2e-06 | -0.000156 | -0.000156 | -34.6779
2019-03-26 16:06 | 1553630748.6868 | -431 | 0.20948 | 0.000439 | 0 | 94.7568 | 94.0074 | 129.7286 | 21.2612 | 0.001825 | 0.001637 | 42.8644 | 0.22033 | 8.4012 | 0.003427 | 0.0043 | 0.004034 | 5.8e-05 | 0.001225 | 0.33539 | 1.3682 | 2.5474 | 100.79 | 21.07 | 16.45 | 65.7412 | 104.3225 | 61.2165 | 0.004081 | 1.0559 | 0.025725 | 24.364 | 0.07102 | 0.13107 | 1e-06 | 0.000101 | 0.000101 | 53.6682
2019-03-26 16:06 | 1553630749.6869 | -430 | 0.20946 | 0.000446 | 0 | 94.838 | 94.2969 | 129.9807 | 21.2612 | 0.001788 | 0.001682 | 42.8727 | 0.24983 | 8.4008 | 0.003498 | 0.005121 | 0.005105 | 6.2e-05 | 0.001229 | 0.33052 | 1.3508 | 2.5475 | 100.79 | 21.07 | 16.449 | 64.8569 | 106.0475 | 62.2916 | 0.004081 | 1.0525 | 0.025709 | 24.4269 | 0.051183 | 0.14863 | -9e-06 | -0.000662 | -0.000662 | -9.2654
2019-03-26 16:06 | 1553630750.686 | -429 | 0.20946 | 0.000441 | 0 | 94.8418 | 94.3274 | 129.7913 | 21.2612 | 0.001757 | 0.001621 | 42.8674 | 0.2516 | 8.4045 | 0.003517 | 0.004685 | 0.004663 | 6.1e-05 | 0.001223 | 0.3413 | 1.3542 | 2.5475 | 100.79 | 21.07 | 16.451 | 64.1059 | 106.5072 | 60.8327 | 0.004082 | 1.0571 | 0.025746 | 24.3559 | 0.048643 | 0.14968 | -1e-06 | -9.2e-05 | -9.2e-05 | -66.9319
2019-03-26 16:06 | 1553630751.686 | -428 | 0.20947 | 0.00045 | 0 | 95.0154 | 94.2979 | 129.7976 | 21.2612 | 0.001789 | 0.001685 | 42.8921 | 0.25244 | 8.4013 | 0.003551 | 0.004846 | 0.0048 | 5.7e-05 | 0.001221 | 0.32853 | 1.3258 | 2.5474 | 100.79 | 21.07 | 16.453 | 64.8802 | 107.3396 | 62.3713 | 0.004082 | 1.0522 | 0.025688 | 24.4125 | 0.067803 | 0.15018 | -1.3e-05 | -0.001006 | -0.001006 | -6.1661
2019-03-26 16:06 | 1553630752.6862 | -427 | 0.20947 | 0.000452 | 0 | 95.0919 | 94.473 | 129.7505 | 21.2612 | 0.001786 | 0.001679 | 42.888 | 0.25149 | 8.4019 | 0.003548 | 0.005272 | 0.005184 | 5.9e-05 | 0.001218 | 0.33682 | 1.3886 | 2.5474 | 100.79 | 21.07 | 16.454 | 64.7905 | 107.2656 | 62.2259 | 0.004082 | 1.0527 | 0.025688 | 24.4028 | 0.0584 | 0.14961 | -1.6e-05 | -0.001264 | -0.001264 | -4.8886
2019-03-26 16:06 | 1553630753.686 | -426 | 0.20944 | 0.000456 | 0 | 95.1072 | 94.4481 | 129.9697 | 21.2612 | 0.001829 | 0.001682 | 42.5127 | 0.2414 | 8.403 | 0.003561 | 0.004657 | 0.004486 | 6e-05 | 0.001213 | 0.32068 | 1.3836 | 2.5475 | 100.79 | 21.07 | 16.455 | 65.8302 | 107.5678 | 62.2892 | 0.004083 | 1.0525 | 0.025708 | 24.4257 | 0.0622 | 0.14361 | -2.1e-05 | -0.001681 | -0.001681 | -3.5275
2019-03-26 16:06 | 1553630754.6851 | -425 | 0.20944 | 0.000462 | 0 | 95.2203 | 94.6606 | 130.0922 | 21.2612 | 0.001788 | 0.001683 | 2.1959 | 0.21959 | 8.4019 | 0.003581 | 0.00473 | 0.004528 | 5.9e-05 | 0.001209 | 0.29535 | 1.3748 | 2.5475 | 100.79 | 21.07 | 16.456 | 64.8429 | 108.0618 | 62.3236 | 0.004083 | 1.0524 | 0.025719 | 24.4385 | 0.052728 | 0.13064 | -2.9e-05 | -0.002242 | -0.002242 | -2.4064
2019-03-26 16:06 | 1553630755.6862 | -424 | 0.2094 | 0.000464 | 0 | 95.2831 | 94.7513 | 129.9828 | 21.2612 | 0.001834 | 0.001673 | -39.8849 | 0.21979 | 8.4025 | 0.003528 | 0.004867 | 0.004702 | 5.6e-05 | 0.001199 | 0.31996 | 1.3741 | 2.5474 | 100.79 | 21.07 | 16.455 | 65.9526 | 106.7748 | 62.0821 | 0.004083 | 1.0531 | 0.025717 | 24.4194 | 0.05006 | 0.13075 | -3.1e-05 | -0.002436 | -0.002436 | -2.2165
2019-03-26 16:06 | 1553630756.6862 | -423 | 0.2095 | 0.000463 | 0 | 95.3342 | 94.6508 | 130.2349 | 21.2612 | 0.001752 | 0.001655 | -81.9663 | 0.21977 | 8.4035 | 0.003544 | 0.004742 | 0.004534 | 5.9e-05 | 0.001203 | 0.32244 | 1.3843 | 2.5475 | 100.79 | 21.07 | 16.452 | 63.9904 | 107.1537 | 61.6587 | 0.004082 | 1.0545 | 0.025758 | 24.4276 | 0.064346 | 0.13074 | -3e-05 | -0.002355 | -0.002355 | -2.293
2019-03-26 16:06 | 1553630757.6853 | -422 | 0.20947 | 0.000459 | 0 | 95.3711 | 94.7897 | 130.0641 | 21.2612 | 0.001782 | 0.001714 | -124.0506 | 0.21965 | 8.4019 | 0.00354 | 0.004563 | 0.004362 | 5.7e-05 | 0.001202 | 0.34164 | 1.3835 | 2.5475 | 100.79 | 21.08 | 16.448 | 64.6962 | 107.0685 | 63.0625 | 0.004083 | 1.0501 | 0.025687 | 24.4627 | 0.054697 | 0.13067 | -2.6e-05 | -0.002013 | -0.002013 | -2.6808
2019-03-26 16:06 | 1553630758.6862 | -421 | 0.20949 | 0.000464 | 0 | 95.3643 | 94.7154 | 130.0702 | 21.2612 | 0.001813 | 0.001646 | -152.7228 | 0.21944 | 8.4007 | 0.003598 | 0.004737 | 0.004542 | 5.7e-05 | 0.001206 | 0.31725 | 1.375 | 2.5475 | 100.79 | 21.07 | 16.445 | 65.4477 | 108.4723 | 61.4407 | 0.00408 | 1.0552 | 0.02575 | 24.4042 | 0.061073 | 0.13055 | -3.1e-05 | -0.002417 | -0.002417 | -2.2311
2019-03-26 16:06 | 1553630759.6856 | -420 | 0.20949 | 0.000468 | 0 | 95.3708 | 94.7308 | 130.0735 | 21.2612 | 0.001785 | 0.001681 | -152.7306 | 0.2196 | 8.4024 | 0.003653 | 0.005708 | 0.005498 | 5.8e-05 | 0.00121 | 0.34595 | 1.3259 | 2.5474 | 100.79 | 21.07 | 16.443 | 64.7664 | 109.7942 | 62.2658 | 0.00408 | 1.0526 | 0.025719 | 24.4346 | 0.060222 | 0.13064 | -3.6e-05 | -0.002829 | -0.002829 | -1.9069
2019-03-26 16:06 | 1553630760.6864 | -419 | 0.20948 | 0.000461 | 0 | 95.5175 | 94.7715 | 129.9584 | 21.2612 | 0.00176 | 0.001677 | -152.7232 | 0.21899 | 8.401 | 0.003559 | 0.005099 | 0.004964 | 5.7e-05 | 0.001206 | 0.32405 | 1.3501 | 2.5475 | 100.79 | 21.07 | 16.439 | 64.18 | 107.519 | 62.1712 | 0.004079 | 1.0529 | 0.025711 | 24.4204 | 0.070131 | 0.13028 | -2.8e-05 | -0.002196 | -0.002196 | -2.4501
2019-03-26 16:06 | 1553630761.6855 | -418 | 0.20949 | 0.000462 | 1e-06 | 95.5603 | 95.0199 | 129.7324 | 21.2612 | 0.001779 | 0.001606 | -152.7555 | 0.21877 | 8.404 | 0.003553 | 0.005205 | 0.005087 | 5.8e-05 | 0.001205 | 0.36066 | 1.3588 | 2.5474 | 100.79 | 21.07 | 16.436 | 64.636 | 107.3694 | 60.4768 | 0.004078 | 1.0582 | 0.025754 | 24.3374 | 0.050723 | 0.13015 | -2.9e-05 | -0.00234 | -0.00234 | -2.2975
2019-03-26 16:06 | 1553630762.6855 | -417 | 0.2095 | 0.000462 | 0 | 95.6686 | 94.8825 | 129.4939 | 21.2612 | 0.001794 | 0.001678 | -152.7656 | 0.22 | 8.402 | 0.003544 | 0.004648 | 0.004493 | 5.7e-05 | 0.0012 | 0.34164 | 1.381 | 2.5476 | 100.79 | 21.07 | 16.43 | 64.9904 | 107.1537 | 62.1994 | 0.004076 | 1.0528 | 0.025664 | 24.3777 | 0.073809 | 0.13088 | -2.8e-05 | -0.002116 | -0.002116 | -2.5546
2019-03-26 16:06 | 1553630763.6865 | -416 | 0.20951 | 0.000463 | 0 | 95.6855 | 95.0586 | 129.3967 | 21.2612 | 0.001796 | 0.00169 | -152.7723 | 0.21903 | 8.4054 | 0.003544 | 0.005473 | 0.005244 | 5.4e-05 | 0.001199 | 0.31924 | 1.3764 | 2.5474 | 100.79 | 21.07 | 16.428 | 65.0325 | 107.1466 | 62.4917 | 0.004076 | 1.0519 | 0.025643 | 24.3792 | 0.058792 | 0.1303 | -3.1e-05 | -0.002401 | -0.002401 | -2.2413
2019-03-26 16:06 | 1553630764.6855 | -415 | 0.2095 | 0.000451 | 0 | 95.7956 | 95.1598 | 129.2343 | 21.2612 | 0.001798 | 0.001669 | -152.8002 | 0.22032 | 8.4033 | 0.003484 | 0.005261 | 0.005076 | 5.3e-05 | 0.001213 | 0.33483 | 1.3766 | 2.5475 | 100.79 | 21.07 | 16.427 | 65.0982 | 105.7046 | 61.986 | 0.004076 | 1.0534 | 0.025647 | 24.3455 | 0.059562 | 0.13107 | -1.5e-05 | -0.001174 | -0.001174 | -4.6112
2019-03-26 16:06 | 1553630765.6866 | -414 | 0.20953 | 0.000457 | 0 | 95.885 | 95.2112 | 129.4134 | 21.2612 | 0.001727 | 0.001641 | -152.8052 | 0.22007 | 8.4064 | 0.003467 | 0.004739 | 0.004575 | 5.6e-05 | 0.001214 | 0.3115 | 1.3572 | 2.5475 | 100.79 | 21.07 | 16.424 | 63.3718 | 105.2851 | 61.3179 | 0.004075 | 1.0555 | 0.02569 | 24.3381 | 0.063076 | 0.13092 | -2.3e-05 | -0.001784 | -0.001784 | -3.0308
2019-03-26 16:06 | 1553630766.6872 | -413 | 0.20953 | 0.00046 | 0 | 95.7976 | 95.3948 | 129.1735 | 21.2612 | 0.001783 | 0.001641 | -152.7856 | 0.21836 | 8.4003 | 0.003423 | 0.005085 | 0.004936 | 5.4e-05 | 0.001213 | 0.30807 | 1.3867 | 2.5476 | 100.79 | 21.07 | 16.422 | 64.7366 | 104.2265 | 61.3022 | 0.004074 | 1.0556 | 0.025667 | 24.3149 | 0.037696 | 0.1299 | -2.6e-05 | -0.002072 | -0.002072 | -2.5887
2019-03-26 16:06 | 1553630767.6858 | -412 | 0.20953 | 0.000459 | 0 | 95.8703 | 95.2401 | 129.3374 | 21.2612 | 0.001788 | 0.001631 | -152.7766 | 0.22011 | 8.4041 | 0.00357 | 0.004406 | 0.004278 | 5.4e-05 | 0.001214 | 0.34147 | 1.398 | 2.5476 | 100.79 | 21.07 | 16.42 | 64.86 | 107.7884 | 61.0601 | 0.004074 | 1.0564 | 0.025692 | 24.3216 | 0.058984 | 0.13095 | -2.5e-05 | -0.001987 | -0.001987 | -2.7211
2019-03-26 16:06 | 1553630768.6858 | -411 | 0.20956 | 0.000454 | 1e-06 | 95.96 | 95.2623 | 129.2191 | 21.2612 | 0.001756 | 0.001658 | -152.7855 | 0.21931 | 8.4016 | 0.003433 | 0.003949 | 0.003743 | 5.6e-05 | 0.001219 | 0.32853 | 1.4094 | 2.5477 | 100.79 | 21.07 | 16.42 | 64.0715 | 104.4672 | 61.7119 | 0.004074 | 1.0543 | 0.025655 | 24.3341 | 0.065269 | 0.13046 | -1.9e-05 | -0.001508 | -0.001508 | -3.5725
2019-03-26 16:06 | 1553630769.6849 | -410 | 0.20949 | 0.000453 | 0 | 95.9505 | 95.466 | 129.2275 | 21.2612 | 0.001765 | 0.001652 | -152.8059 | 0.21944 | 8.4034 | 0.003476 | 0.004978 | 0.004855 | 5.1e-05 | 0.00123 | 0.35817 | 1.3714 | 2.5475 | 100.79 | 21.07 | 16.422 | 64.2847 | 105.4992 | 61.5672 | 0.004074 | 1.0548 | 0.025662 | 24.3297 | 0.045281 | 0.13055 | -1.8e-05 | -0.001413 | -0.001413 | -3.8166
2019-03-26 16:06 | 1553630770.686 | -409 | 0.20949 | 0.00045 | 1e-06 | 95.9638 | 95.3739 | 129.1782 | 21.2612 | 0.001751 | 0.001664 | -152.8067 | 0.21977 | 8.403 | 0.00342 | 0.004821 | 0.004633 | 5.2e-05 | 0.001239 | 0.3298 | 1.3608 | 2.5474 | 100.79 | 21.07 | 16.426 | 63.949 | 104.1495 | 61.862 | 0.004075 | 1.0538 | 0.025646 | 24.3357 | 0.055154 | 0.13074 | -1.4e-05 | -0.001112 | -0.001112 | -4.8577
2019-03-26 16:06 | 1553630771.686 | -408 | 0.20949 | 0.000448 | 0 | 96.0441 | 95.6136 | 129.4125 | 21.2612 | 0.001754 | 0.001675 | -152.7994 | 0.21938 | 8.4028 | 0.003444 | 0.00486 | 0.004637 | 5.1e-05 | 0.001239 | 0.32244 | 1.3765 | 2.5475 | 100.79 | 21.07 | 16.428 | 64.0231 | 104.726 | 62.1353 | 0.004076 | 1.053 | 0.025658 | 24.3677 | 0.040184 | 0.13051 | -1.2e-05 | -0.000947 | -0.000947 | -5.6923
2019-03-26 16:06 | 1553630772.686 | -407 | 0.20951 | 0.00045 | 1e-06 | 96.068 | 95.4889 | 129.4944 | 21.2612 | 0.001767 | 0.00168 | -152.7886 | 0.21944 | 8.4033 | 0.003367 | 0.005165 | 0.004952 | 5.1e-05 | 0.001235 | 0.35043 | 1.3546 | 2.5473 | 100.79 | 21.07 | 16.431 | 64.3424 | 102.8787 | 62.2495 | 0.004077 | 1.0526 | 0.025662 | 24.3796 | 0.054085 | 0.13054 | -1.4e-05 | -0.001134 | -0.001134 | -4.7536
2019-03-26 16:06 | 1553630773.686 | -406 | 0.20957 | 0.000447 | 0 | 96.1643 | 95.5353 | 129.324 | 21.2612 | 0.001771 | 0.001682 | -152.7787 | 0.22005 | 8.4026 | 0.003394 | 0.00487 | 0.004719 | 5.3e-05 | 0.001231 | 0.335 | 1.2943 | 2.5475 | 100.79 | 21.07 | 16.44 | 64.4322 | 103.5222 | 62.3017 | 0.004079 | 1.0524 | 0.025643 | 24.3654 | 0.058696 | 0.13091 | -9e-06 | -0.000739 | -0.000739 | -7.3118
2019-03-26 16:06 | 1553630774.6851 | -405 | 0.20949 | 0.000451 | 0 | 96.277 | 95.5729 | 129.3512 | 21.2612 | 0.001776 | 0.001672 | -152.7664 | 0.21968 | 8.4003 | 0.003373 | 0.004744 | 0.004516 | 5.5e-05 | 0.001226 | 0.30596 | 1.3255 | 2.5475 | 100.79 | 21.07 | 16.446 | 64.5494 | 103.0051 | 62.0626 | 0.00408 | 1.0532 | 0.025655 | 24.3593 | 0.065654 | 0.13069 | -1.5e-05 | -0.00122 | -0.00122 | -4.4232
2019-03-26 16:06 | 1553630775.6851 | -404 | 0.20956 | 0.000448 | 1e-06 | 96.1775 | 95.7733 | 129.1848 | 21.2612 | 0.001798 | 0.001688 | -152.7566 | 0.22031 | 8.4021 | 0.003388 | 0.004749 | 0.004627 | 5.4e-05 | 0.001222 | 0.31188 | 1.3854 | 2.5475 | 100.79 | 21.07 | 16.451 | 65.0826 | 103.3667 | 62.4377 | 0.004082 | 1.052 | 0.025624 | 24.3573 | 0.037676 | 0.13106 | -1.2e-05 | -0.000932 | -0.000932 | -5.8087
2019-03-26 16:06 | 1553630776.6851 | -403 | 0.20952 | 0.000448 | 0 | 96.2682 | 95.702 | 129.0977 | 21.2612 | 0.001821 | 0.001674 | -152.7768 | 0.21961 | 8.4002 | 0.003442 | 0.004914 | 0.004799 | 5.2e-05 | 0.001221 | 0.33019 | 1.4011 | 2.5476 | 100.79 | 21.07 | 16.462 | 65.6327 | 104.6771 | 62.1009 | 0.004084 | 1.0531 | 0.025629 | 24.3368 | 0.052758 | 0.13065 | -1.1e-05 | -0.000878 | -0.000878 | -6.1452
2019-03-26 16:06 | 1553630777.6867 | -402 | 0.20952 | 0.000447 | 0 | 96.2924 | 95.9346 | 128.8672 | 21.2612 | 0.001759 | 0.00165 | -152.767 | 0.21998 | 8.404 | 0.003411 | 0.004631 | 0.004409 | 5.1e-05 | 0.001223 | 0.37017 | 1.3873 | 2.5475 | 100.79 | 21.07 | 16.467 | 64.1504 | 103.934 | 61.5282 | 0.004085 | 1.0549 | 0.025628 | 24.2943 | 0.033292 | 0.13086 | -1e-05 | -0.000797 | -0.000797 | -6.782
2019-03-26 16:06 | 1553630778.6853 | -401 | 0.2095 | 0.000447 | 0 | 96.3198 | 95.6718 | 128.9456 | 21.2612 | 0.001815 | 0.001682 | -152.7874 | 0.21888 | 8.4024 | 0.003447 | 0.004563 | 0.004358 | 5.4e-05 | 0.001213 | 0.33539 | 1.3653 | 2.5474 | 100.79 | 21.07 | 16.479 | 65.4985 | 104.7999 | 62.2954 | 0.004088 | 1.0525 | 0.025606 | 24.3295 | 0.060385 | 0.13021 | -9e-06 | -0.000693 | -0.000693 | -7.7623
2019-03-26 16:06 | 1553630779.6854 | -400 | 0.20955 | 0.000446 | -1e-06 | 96.3393 | 95.7619 | 128.9512 | 21.2612 | 0.001751 | 0.00174 | -152.759 | 0.21833 | 8.4041 | 0.003339 | 0.005074 | 0.004979 | 5.3e-05 | 0.001205 | 0.31492 | 1.3705 | 2.5476 | 100.79 | 21.07 | 16.483 | 63.9528 | 102.1787 | 63.6936 | 0.004089 | 1.0481 | 0.025553 | 24.3807 | 0.053775 | 0.12989 | -8e-06 | -0.000577 | -0.000577 | -9.3027
2019-03-26 16:06 | 1553630780.6854 | -399 | 0.20953 | 0.000446 | 0 | 96.4335 | 95.9669 | 128.9573 | 21.2612 | 0.001781 | 0.001677 | -152.7858 | 0.21957 | 8.4012 | 0.003351 | 0.00388 | 0.003681 | 5.1e-05 | 0.001194 | 0.3298 | 1.3489 | 2.5474 | 100.79 | 21.07 | 16.487 | 64.6868 | 102.4716 | 62.1861 | 0.00409 | 1.0528 | 0.025611 | 24.3267 | 0.04338 | 0.13062 | -8e-06 | -0.000626 | -0.000626 | -8.6132
2019-03-26 16:06 | 1553630781.6864 | -398 | 0.20949 | 0.00044 | 0 | 96.4524 | 95.9656 | 128.7383 | 21.2612 | 0.001833 | 0.001679 | -152.7955 | 0.22001 | 8.4023 | 0.003348 | 0.004694 | 0.004518 | 5.2e-05 | 0.00119 | 0.32693 | 1.4042 | 2.5475 | 100.79 | 21.07 | 16.493 | 65.93 | 102.4002 | 62.2307 | 0.004092 | 1.0527 | 0.025588 | 24.3076 | 0.045258 | 0.13088 | 0 | 2.9e-05 | 2.9e-05 | 187.3106
2019-03-26 16:06 | 1553630782.6865 | -397 | 0.20955 | 0.00045 | 0 | 96.4984 | 95.9289 | 129.0592 | 21.2612 | 0.001816 | 0.001737 | -152.807 | 0.21914 | 8.4046 | 0.003403 | 0.004651 | 0.004541 | 5e-05 | 0.00118 | 0.35723 | 1.331 | 2.5475 | 100.79 | 21.07 | 16.495 | 65.5242 | 103.7407 | 63.614 | 0.004092 | 1.0483 | 0.025567 | 24.388 | 0.052943 | 0.13037 | -1.3e-05 | -0.000999 | -0.000999 | -5.392
2019-03-26 16:06 | 1553630783.6846 | -396 | 0.2095 | 0.000448 | 0 | 96.5542 | 95.9951 | 129.3205 | 21.2612 | 0.00176 | 0.001704 | -152.8185 | 0.2198 | 8.4015 | 0.003499 | 0.004395 | 0.004162 | 5.1e-05 | 0.001164 | 0.32947 | 1.3474 | 2.5475 | 100.79 | 21.07 | 16.498 | 64.1754 | 106.0703 | 62.8196 | 0.004093 | 1.0508 | 0.025623 | 24.3839 | 0.05194 | 0.13076 | -1.1e-05 | -0.000809 | -0.000809 | -6.6737
2019-03-26 16:06 | 1553630784.6855 | -395 | 0.20951 | 0.000447 | -1e-06 | 96.5341 | 96.018 | 129.4071 | 21.2612 | 0.001871 | 0.001695 | -152.8254 | 0.21987 | 8.4034 | 0.00348 | 0.00486 | 0.004697 | 5.1e-05 | 0.001161 | 0.35994 | 1.3933 | 2.5474 | 100.79 | 21.07 | 16.499 | 66.843 | 105.5959 | 62.6181 | 0.004093 | 1.0515 | 0.025639 | 24.3847 | 0.04795 | 0.1308 | -9e-06 | -0.000641 | -0.000641 | -8.4236
2019-03-26 16:06 | 1553630785.6847 | -394 | 0.20948 | 0.000449 | 0 | 96.5158 | 96.0607 | 129.3909 | 21.2612 | 0.001772 | 0.001653 | -152.8356 | 0.21921 | 8.4027 | 0.003382 | 0.004929 | 0.004716 | 5.1e-05 | 0.001157 | 0.29999 | 1.3815 | 2.5474 | 100.79 | 21.07 | 16.5 | 64.4673 | 103.2424 | 61.5978 | 0.004094 | 1.0547 | 0.025677 | 24.3461 | 0.042279 | 0.13041 | -1.3e-05 | -0.001037 | -0.001037 | -5.1959
2019-03-26 16:06 | 1553630786.6858 | -393 | 0.20948 | 0.000452 | 0 | 96.6507 | 96.1868 | 129.3822 | 21.2612 | 0.001827 | 0.001721 | -152.8265 | 0.21958 | 8.399 | 0.003401 | 0.00474 | 0.004612 | 4.8e-05 | 0.001159 | 0.33268 | 1.3604 | 2.5474 | 100.79 | 21.07 | 16.5 | 65.7896 | 103.6881 | 63.2414 | 0.004094 | 1.0495 | 0.025613 | 24.405 | 0.043041 | 0.13063 | -1.5e-05 | -0.001203 | -0.001203 | -4.4839
2019-03-26 16:06 | 1553630787.6861 | -392 | 0.20945 | 0.000454 | 1e-06 | 96.698 | 96.3522 | 129.4226 | 21.2612 | 0.001796 | 0.001651 | -152.8283 | 0.21859 | 8.4054 | 0.00352 | 0.004508 | 0.004287 | 4.9e-05 | 0.001159 | 0.30862 | 1.3689 | 2.5475 | 100.79 | 21.07 | 16.508 | 65.0402 | 106.5766 | 61.5539 | 0.004096 | 1.0548 | 0.025682 | 24.3475 | 0.032044 | 0.13004 | -1.9e-05 | -0.001477 | -0.001477 | -3.6367
2019-03-26 16:06 | 1553630788.6848 | -391 | 0.20947 | 0.000451 | 0 | 96.8048 | 96.3468 | 129.2169 | 21.2612 | 0.001842 | 0.001687 | -152.8156 | 0.21838 | 8.4048 | 0.003444 | 0.004922 | 0.004759 | 5.1e-05 | 0.001162 | 0.33516 | 1.383 | 2.5475 | 100.79 | 21.07 | 16.511 | 66.1487 | 104.7267 | 62.4228 | 0.004096 | 1.0521 | 0.025628 | 24.3597 | 0.042422 | 0.12991 | -1.6e-05 | -0.001233 | -0.001233 | -4.352
2019-03-26 16:06 | 1553630789.6858 | -390 | 0.20951 | 0.000449 | 0 | 96.8264 | 96.492 | 128.8224 | 21.2612 | 0.001795 | 0.001737 | -152.8253 | 0.21959 | 8.4026 | 0.003443 | 0.004348 | 0.004168 | 5e-05 | 0.001167 | 0.35834 | 1.3667 | 2.5474 | 100.79 | 21.07 | 16.52 | 65.0224 | 104.6969 | 63.6154 | 0.004099 | 1.0483 | 0.025544 | 24.3657 | 0.030941 | 0.13063 | -1.2e-05 | -0.000961 | -0.000961 | -5.6128
2019-03-26 16:06 | 1553630790.685 | -389 | 0.20951 | 0.000445 | 0 | 96.9855 | 96.5717 | 128.8698 | 21.2612 | 0.001824 | 0.001725 | -152.8369 | 0.21829 | 8.4027 | 0.003469 | 0.005312 | 0.00515 | 5.3e-05 | 0.001172 | 0.32731 | 1.3507 | 2.5475 | 100.79 | 21.07 | 16.524 | 65.7068 | 105.3457 | 63.3343 | 0.0041 | 1.0492 | 0.025559 | 24.36 | 0.038245 | 0.12986 | -8e-06 | -0.000641 | -0.000641 | -8.3633
2019-03-26 16:06 | 1553630791.685 | -388 | 0.20953 | 0.000445 | 0 | 97.0838 | 96.7014 | 128.6881 | 21.2612 | 0.001848 | 0.001673 | -152.8245 | 0.21815 | 8.401 | 0.003476 | 0.004537 | 0.004405 | 4.8e-05 | 0.001178 | 0.34147 | 1.3235 | 2.5475 | 100.79 | 21.07 | 16.531 | 66.2829 | 105.5118 | 62.0831 | 0.004101 | 1.0531 | 0.025589 | 24.2975 | 0.035295 | 0.12978 | -7e-06 | -0.000568 | -0.000568 | -9.4355
2019-03-26 16:07 | 1553630792.6854 | -387 | 0.2095 | 0.000446 | 1e-06 | 97.1511 | 96.7871 | 128.6595 | 21.2612 | 0.001783 | 0.001695 | -152.8231 | 0.21825 | 8.4015 | 0.003449 | 0.004274 | 0.004034 | 4.9e-05 | 0.001183 | 0.36282 | 1.3229 | 2.5474 | 100.79 | 21.07 | 16.546 | 64.725 | 104.8589 | 62.6003 | 0.004105 | 1.0515 | 0.025566 | 24.3136 | 0.033569 | 0.12983 | -8e-06 | -0.000683 | -0.000683 | -7.8555
2019-03-26 16:07 | 1553630793.686 | -386 | 0.2095 | 0.000446 | 0 | 97.2928 | 96.9616 | 128.5367 | 21.2612 | 0.001798 | 0.001668 | -152.822 | 0.21921 | 8.4039 | 0.003446 | 0.005315 | 0.005126 | 5.1e-05 | 0.001184 | 0.34219 | 1.3192 | 2.5474 | 100.79 | 21.07 | 16.553 | 65.0871 | 104.7709 | 61.9728 | 0.004107 | 1.0535 | 0.025578 | 24.2792 | 0.030495 | 0.13041 | -8e-06 | -0.000625 | -0.000625 | -8.618
2019-03-26 16:07 | 1553630794.6855 | -385 | 0.20951 | 0.000447 | 0 | 97.3111 | 96.9055 | 128.3097 | 21.2612 | 0.001843 | 0.001669 | -152.7998 | 0.21974 | 8.4026 | 0.00344 | 0.005214 | 0.005022 | 4.9e-05 | 0.001183 | 0.3079 | 1.3317 | 2.5475 | 100.79 | 21.07 | 16.559 | 66.1751 | 104.6417 | 61.979 | 0.004108 | 1.0535 | 0.025555 | 24.258 | 0.037358 | 0.13072 | -9e-06 | -0.000728 | -0.000728 | -7.4211
2019-03-26 16:07 | 1553630795.6851 | -384 | 0.20949 | 0.000448 | 0 | 97.4524 | 97.3371 | 128.3032 | 21.2612 | 0.00182 | 0.001713 | -152.7928 | 0.21907 | 8.4021 | 0.003406 | 0.004829 | 0.004594 | 4.8e-05 | 0.001183 | 0.33572 | 1.3587 | 2.5475 | 100.79 | 21.07 | 16.574 | 65.6226 | 103.8013 | 63.0555 | 0.004112 | 1.0501 | 0.025513 | 24.2963 | 0.010593 | 0.13033 | -1.1e-05 | -0.000852 | -0.000852 | -6.3159
2019-03-26 16:07 | 1553630796.6865 | -383 | 0.20951 | 0.000444 | 0 | 97.5339 | 97.2971 | 128.3839 | 21.2612 | 0.001788 | 0.001672 | -152.8036 | 0.21955 | 8.4019 | 0.003407 | 0.004715 | 0.004577 | 5e-05 | 0.001171 | 0.31094 | 1.3474 | 2.5474 | 100.79 | 21.07 | 16.581 | 64.8492 | 103.8405 | 62.0611 | 0.004114 | 1.0532 | 0.025559 | 24.268 | 0.021744 | 0.13061 | -6e-06 | -0.00047 | -0.00047 | -11.4873
2019-03-26 16:07 | 1553630797.6853 | -382 | 0.20951 | 0.000444 | 0 | 97.5543 | 97.2719 | 128.6273 | 21.2612 | 0.001796 | 0.001679 | -152.8245 | 0.2186 | 8.4055 | 0.003521 | 0.004646 | 0.004496 | 5e-05 | 0.001163 | 0.32261 | 1.3473 | 2.5475 | 100.79 | 21.08 | 16.597 | 65.0513 | 106.5953 | 62.2251 | 0.00412 | 1.0527 | 0.025577 | 24.2969 | 0.02593 | 0.13004 | -5e-06 | -0.000401 | -0.000401 | -13.3842
2019-03-26 16:07 | 1553630798.6853 | -381 | 0.20951 | 0.000448 | 0 | 97.6812 | 97.3206 | 128.9369 | 21.2612 | 0.001804 | 0.001721 | -152.8219 | 0.21849 | 8.4058 | 0.003424 | 0.004965 | 0.004802 | 4.5e-05 | 0.001159 | 0.32244 | 1.3932 | 2.5477 | 100.79 | 21.08 | 16.603 | 65.2457 | 104.2423 | 63.2414 | 0.004122 | 1.0495 | 0.025569 | 24.363 | 0.033079 | 0.12998 | -1e-05 | -0.000783 | -0.000783 | -6.8598
2019-03-26 16:07 | 1553630799.6857 | -380 | 0.20949 | 0.00044 | 0 | 97.7039 | 97.4581 | 129.2214 | 21.2612 | 0.001841 | 0.0017 | -152.8333 | 0.21886 | 8.4028 | 0.003388 | 0.004931 | 0.004778 | 4.8e-05 | 0.001156 | 0.34058 | 1.3648 | 2.5475 | 100.79 | 21.08 | 16.609 | 66.1167 | 103.3761 | 62.7385 | 0.004123 | 1.0511 | 0.025616 | 24.3716 | 0.022533 | 0.1302 | 0 | 2.4e-05 | 2.4e-05 | 224.4965
2019-03-26 16:07 | 1553630800.6845 | -379 | 0.20953 | 0.000439 | 0 | 97.8253 | 97.4067 | 129.1496 | 21.2612 | 0.001767 | 0.001705 | -152.8073 | 0.21936 | 8.4008 | 0.003375 | 0.004895 | 0.004629 | 4.7e-05 | 0.001158 | 0.30127 | 1.3569 | 2.5475 | 100.79 | 21.08 | 16.622 | 64.3393 | 103.0508 | 62.8579 | 0.004126 | 1.0507 | 0.025605 | 24.3692 | 0.038354 | 0.1305 | 1e-06 | 0.000112 | 0.000112 | 48.0589
2019-03-26 16:07 | 1553630801.6854 | -378 | 0.20952 | 0.000445 | 0 | 97.8805 | 97.5923 | 129.2526 | 21.2612 | 0.00177 | 0.00171 | -152.8084 | 0.21943 | 8.4042 | 0.003388 | 0.004579 | 0.004449 | 5.1e-05 | 0.001165 | 0.33196 | 1.3749 | 2.5475 | 100.79 | 21.08 | 16.629 | 64.4103 | 103.3722 | 62.972 | 0.004128 | 1.0503 | 0.025611 | 24.383 | 0.026376 | 0.13054 | -8e-06 | -0.000587 | -0.000587 | -9.1824
2019-03-26 16:07 | 1553630802.6855 | -377 | 0.20949 | 0.000445 | 0 | 98.0009 | 97.7324 | 129.1665 | 21.2612 | 0.001765 | 0.001727 | -152.8051 | 0.21924 | 8.4025 | 0.003432 | 0.00472 | 0.004532 | 4.8e-05 | 0.001172 | 0.33395 | 1.3678 | 2.5476 | 100.79 | 21.08 | 16.636 | 64.2994 | 104.4303 | 63.3827 | 0.00413 | 1.0491 | 0.025587 | 24.3898 | 0.024537 | 0.13043 | -8e-06 | -0.000614 | -0.000614 | -8.7663
2019-03-26 16:07 | 1553630803.6855 | -376 | 0.20949 | 0.000444 | 0 | 97.9891 | 97.7338 | 129.3052 | 21.2612 | 0.001767 | 0.001714 | -152.7956 | 0.2196 | 8.4035 | 0.003516 | 0.004641 | 0.004438 | 4.6e-05 | 0.001184 | 0.32621 | 1.3543 | 2.5475 | 100.79 | 21.07 | 16.647 | 64.3487 | 106.4679 | 63.0656 | 0.00413 | 1.0501 | 0.025612 | 24.3914 | 0.023336 | 0.13064 | -6e-06 | -0.000497 | -0.000497 | -10.852
2019-03-26 16:07 | 1553630804.6861 | -375 | 0.20953 | 0.000441 | 0 | 98.0402 | 98.1368 | 129.3594 | 21.2612 | 0.001828 | 0.001693 | -152.7953 | 0.21959 | 8.4023 | 0.003518 | 0.005346 | 0.005168 | 4.9e-05 | 0.001187 | 0.29502 | 1.3361 | 2.5475 | 100.79 | 21.07 | 16.653 | 65.8153 | 106.5206 | 62.5589 | 0.004132 | 1.0516 | 0.025637 | 24.3781 | -0.008815 | 0.13063 | -2e-06 | -0.000165 | -0.000165 | -32.7233
2019-03-26 16:07 | 1553630805.6851 | -374 | 0.20955 | 0.000444 | 1e-06 | 98.1022 | 97.7758 | 129.3144 | 21.2612 | 0.001812 | 0.001683 | -152.8009 | 0.21931 | 8.403 | 0.00349 | 0.004974 | 0.004807 | 4.9e-05 | 0.001188 | 0.33323 | 1.3421 | 2.5475 | 100.79 | 21.07 | 16.665 | 65.4338 | 105.8603 | 62.3292 | 0.004135 | 1.0524 | 0.025641 | 24.3655 | 0.029805 | 0.13047 | -7e-06 | -0.000599 | -0.000599 | -9.002
2019-03-26 16:07 | 1553630806.6847 | -373 | 0.2095 | 0.000438 | 0 | 98.1706 | 97.9916 | 129.1585 | 21.2612 | 0.001801 | 0.001702 | -152.807 | 0.21903 | 8.4021 | 0.003487 | 0.005352 | 0.005177 | 4.6e-05 | 0.001196 | 0.35419 | 1.3604 | 2.5475 | 100.79 | 21.07 | 16.672 | 65.1518 | 105.7722 | 62.7705 | 0.004136 | 1.051 | 0.025609 | 24.3668 | 0.016326 | 0.1303 | 1e-06 | 4.6e-05 | 4.6e-05 | 116.9835
2019-03-26 16:07 | 1553630807.6848 | -372 | 0.2095 | 0.000444 | 0 | 98.2096 | 97.8141 | 128.8646 | 21.2612 | 0.001793 | 0.001744 | -152.7954 | 0.21896 | 8.4051 | 0.003474 | 0.004893 | 0.004651 | 4.9e-05 | 0.001192 | 0.36193 | 1.3635 | 2.5475 | 100.79 | 21.07 | 16.677 | 64.9615 | 105.4715 | 63.7991 | 0.004138 | 1.0478 | 0.025541 | 24.3763 | 0.036091 | 0.13026 | -6e-06 | -0.000441 | -0.000441 | -12.189
2019-03-26 16:07 | 1553630808.6848 | -371 | 0.20949 | 0.000446 | 0 | 98.24 | 98.1577 | 128.67 | 21.2612 | 0.001815 | 0.001731 | -152.7943 | 0.21963 | 8.4029 | 0.003442 | 0.004778 | 0.004726 | 5.4e-05 | 0.001191 | 0.35026 | 1.3419 | 2.5475 | 100.79 | 21.07 | 16.688 | 65.4877 | 104.6938 | 63.4671 | 0.00414 | 1.0488 | 0.025534 | 24.3459 | 0.007499 | 0.13066 | -9e-06 | -0.000697 | -0.000697 | -7.7386
2019-03-26 16:07 | 1553630809.6849 | -370 | 0.20949 | 0.00044 | 0 | 98.361 | 98.1187 | 128.5055 | 21.2612 | 0.001815 | 0.001697 | -152.7838 | 0.21939 | 8.4043 | 0.003436 | 0.004292 | 0.004175 | 4.9e-05 | 0.001193 | 0.32516 | 1.3486 | 2.5473 | 100.79 | 21.07 | 16.694 | 65.4971 | 104.545 | 62.6674 | 0.004142 | 1.0513 | 0.025548 | 24.3014 | 0.022065 | 0.13051 | -1e-06 | -2.3e-05 | -2.3e-05 | -234.3554
2019-03-26 16:07 | 1553630810.6852 | -369 | 0.20951 | 0.000444 | 0 | 98.3197 | 98.0746 | 128.6637 | 21.2612 | 0.001771 | 0.001713 | -152.7766 | 0.21983 | 8.398 | 0.003403 | 0.005453 | 0.005221 | 4.8e-05 | 0.001188 | 0.31974 | 1.3349 | 2.5475 | 100.79 | 21.07 | 16.703 | 64.4524 | 103.747 | 63.0478 | 0.004144 | 1.0501 | 0.025549 | 24.3302 | 0.022322 | 0.13078 | -5e-06 | -0.000362 | -0.000362 | -14.9025
2019-03-26 16:07 | 1553630811.686 | -368 | 0.20947 | 0.000435 | 0 | 98.4499 | 98.2027 | 128.7466 | 21.2612 | 0.001851 | 0.00167 | -152.7618 | 0.21935 | 8.4019 | 0.003468 | 0.004869 | 0.004663 | 4.9e-05 | 0.001186 | 0.28926 | 1.3507 | 2.5475 | 100.79 | 21.07 | 16.708 | 66.374 | 105.3236 | 62.0079 | 0.004145 | 1.0534 | 0.025597 | 24.3003 | 0.022483 | 0.13049 | 6e-06 | 0.000502 | 0.000502 | 10.7292
2019-03-26 16:07 | 1553630812.686 | -367 | 0.20949 | 0.000439 | 0 | 98.4705 | 98.3722 | 128.6672 | 21.2612 | 0.001855 | 0.00169 | -152.7525 | 0.21865 | 8.4039 | 0.003487 | 0.004988 | 0.004842 | 4.7e-05 | 0.001182 | 0.33716 | 1.3617 | 2.5475 | 100.79 | 21.07 | 16.713 | 66.456 | 105.7793 | 62.483 | 0.004147 | 1.0519 | 0.025571 | 24.31 | 0.008942 | 0.13008 | 0 | -9e-06 | -9e-06 | -627.1151
2019-03-26 16:07 | 1553630813.6851 | -366 | 0.20952 | 0.000435 | 1e-06 | 98.5727 | 98.3967 | 128.4233 | 21.2612 | 0.00176 | 0.001659 | -152.7675 | 0.21888 | 8.4039 | 0.003486 | 0.004727 | 0.004539 | 4.6e-05 | 0.001183 | 0.35081 | 1.3802 | 2.5475 | 100.79 | 21.08 | 16.72 | 64.1848 | 105.7414 | 61.7555 | 0.004151 | 1.0542 | 0.025575 | 24.2607 | 0.015984 | 0.13021 | -0.000111 | -0.008616 | -0.008616 | -0.62415
2019-03-26 16:07 | 1553630814.6854 | -365 | 0.20948 | 0.000437 | 0 | 98.6621 | 98.3631 | 128.2337 | 21.2612 | 0.001831 | 0.00169 | -152.7732 | 0.21819 | 8.4035 | 0.003456 | 0.004828 | 0.004686 | 4.7e-05 | 0.001179 | 0.32372 | 1.3361 | 2.5473 | 100.79 | 21.08 | 16.724 | 65.8911 | 105.0232 | 62.4994 | 0.004152 | 1.0518 | 0.025527 | 24.2696 | 0.027152 | 0.1298 | 9.6e-05 | 0.007389 | 0.007389 | 0.7255
2019-03-26 16:07 | 1553630815.6856 | -364 | 0.2095 | 0.000442 | 0 | 98.7312 | 98.4189 | 128.079 | 21.2612 | 0.001812 | 0.001706 | -152.7591 | 0.21938 | 8.4031 | 0.003443 | 0.005145 | 0.004897 | 4.7e-05 | 0.00118 | 0.32012 | 1.3144 | 2.5474 | 100.79 | 21.08 | 16.73 | 65.4213 | 104.7197 | 62.8665 | 0.004153 | 1.0507 | 0.025498 | 24.2683 | 0.028339 | 0.13051 | -3.8e-05 | -0.00291 | -0.00291 | -1.8522
2019-03-26 16:07 | 1553630816.6843 | -363 | 0.20951 | 0.000438 | 0 | 98.7347 | 98.4555 | 128.0635 | 21.2612 | 0.001819 | 0.00167 | -152.7585 | 0.21875 | 8.4026 | 0.00342 | 0.005224 | 0.00512 | 4.6e-05 | 0.001185 | 0.33373 | 1.322 | 2.5474 | 100.79 | 21.08 | 16.733 | 65.5962 | 104.1526 | 62.0118 | 0.004154 | 1.0534 | 0.025529 | 24.2359 | 0.025331 | 0.13014 | -7.9e-05 | -0.006059 | -0.006059 | -0.88698
2019-03-26 16:07 | 1553630817.6847 | -362 | 0.20948 | 0.000437 | 0 | 98.7955 | 98.6643 | 127.8889 | 21.2612 | 0.001865 | 0.001697 | -152.7768 | 0.21857 | 8.4039 | 0.003405 | 0.004655 | 0.004417 | 4.9e-05 | 0.001186 | 0.32156 | 1.3192 | 2.5475 | 100.79 | 21.08 | 16.736 | 66.7049 | 103.7872 | 62.6581 | 0.004155 | 1.0513 | 0.025487 | 24.2427 | 0.011892 | 0.13003 | 0.000115 | 0.008856 | 0.008856 | 0.60637
2019-03-26 16:07 | 1553630818.6853 | -361 | 0.20949 | 0.00044 | 0 | 98.8419 | 98.7433 | 127.8957 | 21.2612 | 0.001777 | 0.001713 | -152.7832 | 0.21926 | 8.4033 | 0.003389 | 0.004801 | 0.004677 | 4.9e-05 | 0.001191 | 0.32372 | 1.3247 | 2.5475 | 100.79 | 21.08 | 16.739 | 64.5884 | 103.4035 | 63.0548 | 0.004156 | 1.0501 | 0.025473 | 24.2577 | 0.008928 | 0.13044 | 8.3e-05 | 0.006386 | 0.006386 | 0.84359
2019-03-26 16:07 | 1553630819.6853 | -360 | 0.20949 | 0.00044 | 0 | 98.8865 | 98.7876 | 127.9562 | 21.2612 | 0.001779 | 0.001685 | -152.7784 | 0.21966 | 8.4059 | 0.003407 | 0.004947 | 0.004766 | 4.7e-05 | 0.00119 | 0.31022 | 1.3551 | 2.5476 | 100.79 | 21.08 | 16.741 | 64.6242 | 103.8422 | 62.379 | 0.004156 | 1.0522 | 0.025504 | 24.239 | 0.008943 | 0.13067 | 7.8e-05 | 0.005999 | 0.005999 | 0.89967
2019-03-26 16:07 | 1553630820.6844 | -359 | 0.20948 | 0.000441 | 0 | 98.9027 | 98.8031 | 128.0602 | 21.2612 | 0.001795 | 0.001776 | -152.78 | 0.21949 | 8.4025 | 0.003349 | 0.00446 | 0.004277 | 4.5e-05 | 0.001193 | 0.3517 | 1.3533 | 2.5474 | 100.79 | 21.08 | 16.741 | 65.0217 | 102.4293 | 64.5682 | 0.004156 | 1.0454 | 0.025432 | 24.3279 | 0.009011 | 0.13057 | 0.000114 | 0.008741 | 0.008741 | 0.61692
2019-03-26 16:07 | 1553630821.6844 | -358 | 0.2095 | 0.000437 | 0 | 98.9558 | 98.7433 | 128.2241 | 21.2612 | 0.001774 | 0.001638 | -152.788 | 0.21907 | 8.404 | 0.003276 | 0.004487 | 0.004243 | 4.8e-05 | 0.001196 | 0.36569 | 1.3459 | 2.5474 | 100.79 | 21.08 | 16.742 | 64.5111 | 100.6745 | 61.236 | 0.004156 | 1.0558 | 0.025575 | 24.223 | 0.019233 | 0.13033 | -3.4e-05 | -0.002602 | -0.002602 | -2.0686
2019-03-26 16:07 | 1553630822.6847 | -357 | 0.20952 | 0.000437 | 0 | 99.0243 | 98.8431 | 128.1832 | 21.2612 | 0.001817 | 0.00169 | -152.7895 | 0.21933 | 8.4034 | 0.003364 | 0.004238 | 0.004052 | 4.8e-05 | 0.001197 | 0.32781 | 1.3503 | 2.5475 | 100.79 | 21.08 | 16.743 | 65.5507 | 102.7922 | 62.4806 | 0.004157 | 1.0519 | 0.025523 | 24.2642 | 0.016382 | 0.13048 | -0.000109 | -0.008375 | -0.008375 | -0.64341
2019-03-26 16:07 | 1553630823.6855 | -356 | 0.20949 | 0.000443 | 0 | 99.0665 | 98.8993 | 128.245 | 21.2612 | 0.001794 | 0.001705 | -152.7793 | 0.21903 | 8.4039 | 0.00331 | 0.004385 | 0.004232 | 4.6e-05 | 0.001191 | 0.33964 | 1.3598 | 2.5474 | 100.79 | 21.08 | 16.743 | 64.9834 | 101.4761 | 62.8641 | 0.004157 | 1.0507 | 0.025515 | 24.2839 | 0.015115 | 0.1303 | 2.7e-05 | 0.002065 | 0.002065 | 2.6056
2019-03-26 16:07 | 1553630824.6855 | -355 | 0.20949 | 0.000438 | 0 | 99.0952 | 98.8381 | 128.3762 | 21.2612 | 0.001763 | 0.0017 | -152.7919 | 0.21922 | 8.4046 | 0.003343 | 0.004825 | 0.004543 | 5e-05 | 0.001185 | 0.32045 | 1.3188 | 2.5475 | 100.79 | 21.08 | 16.744 | 64.2402 | 102.2816 | 62.7337 | 0.004157 | 1.0511 | 0.025533 | 24.2916 | 0.023235 | 0.13042 | 4.2e-05 | 0.003285 | 0.003285 | 1.6396
2019-03-26 16:07 | 1553630825.6861 | -354 | 0.20952 | 0.00044 | -1e-06 | 99.0783 | 99.0425 | 128.4494 | 21.2612 | 0.001797 | 0.001688 | -152.8069 | 0.21983 | 8.404 | 0.003349 | 0.004629 | 0.004443 | 5.1e-05 | 0.001186 | 0.32875 | 1.3476 | 2.5475 | 100.79 | 21.08 | 16.745 | 65.0777 | 102.4317 | 62.4368 | 0.004157 | 1.052 | 0.025551 | 24.2878 | 0.003239 | 0.13078 | -0.000121 | -0.009299 | -0.009299 | -0.58081
2019-03-26 16:07 | 1553630826.6847 | -353 | 0.20949 | 0.00044 | 0 | 99.1465 | 99.1386 | 128.6738 | 21.2612 | 0.001757 | 0.001699 | -152.7993 | 0.21853 | 8.4063 | 0.003414 | 0.00458 | 0.004427 | 5e-05 | 0.001189 | 0.35347 | 1.3824 | 2.5474 | 100.79 | 21.08 | 16.746 | 64.1097 | 103.9939 | 62.7127 | 0.004157 | 1.0512 | 0.025563 | 24.319 | 0.000715 | 0.13001 | 5.3e-05 | 0.004099 | 0.004099 | 1.3098
2019-03-26 16:07 | 1553630827.6848 | -352 | 0.2095 | 0.000438 | 1e-06 | 99.2543 | 99.0804 | 128.9193 | 21.2612 | 0.001807 | 0.001721 | -152.7747 | 0.21951 | 8.4035 | 0.003356 | 0.004433 | 0.004275 | 4.9e-05 | 0.001194 | 0.33157 | 1.3408 | 2.5475 | 100.79 | 21.08 | 16.747 | 65.3065 | 102.5975 | 63.2257 | 0.004158 | 1.0496 | 0.025568 | 24.3608 | 0.015678 | 0.13059 | 9e-06 | 0.000679 | 0.000679 | 7.9429
2019-03-26 16:07 | 1553630828.6863 | -351 | 0.20951 | 0.00044 | 0 | 99.2888 | 99.0092 | 128.6306 | 21.2612 | 0.001834 | 0.001669 | -152.7764 | 0.21859 | 8.4051 | 0.003336 | 0.004934 | 0.004753 | 4.6e-05 | 0.001196 | 0.34468 | 1.391 | 2.5475 | 100.79 | 21.08 | 16.747 | 65.9481 | 102.1269 | 61.987 | 0.004158 | 1.0534 | 0.025587 | 24.2886 | 0.025225 | 0.13004 | -6.7e-05 | -0.005181 | -0.005181 | -1.0366
2019-03-26 16:07 | 1553630829.6839 | -350 | 0.20952 | 0.000441 | 0 | 99.3679 | 99.2095 | 128.5996 | 21.2612 | 0.001849 | 0.001699 | -152.7749 | 0.21901 | 8.4011 | 0.003376 | 0.004435 | 0.004225 | 4.7e-05 | 0.001198 | 0.2975 | 1.3848 | 2.5476 | 100.79 | 21.09 | 16.747 | 66.3242 | 103.0915 | 62.7173 | 0.00416 | 1.0511 | 0.025556 | 24.3121 | 0.014268 | 0.13029 | -0.000102 | -0.007853 | -0.007853 | -0.68523
2019-03-26 16:07 | 1553630830.6843 | -349 | 0.20951 | 0.000441 | 1e-06 | 99.2625 | 99.2865 | 128.4203 | 21.2612 | 0.001852 | 0.001725 | -152.777 | 0.21922 | 8.4008 | 0.003383 | 0.004505 | 0.004348 | 5e-05 | 0.001197 | 0.32781 | 1.3803 | 2.5475 | 100.79 | 21.09 | 16.747 | 66.4007 | 103.2637 | 63.3321 | 0.00416 | 1.0492 | 0.025514 | 24.3174 | -0.002161 | 0.13042 | -7.9e-05 | -0.006121 | -0.006121 | -0.87988
2019-03-26 16:07 | 1553630831.684 | -348 | 0.20951 | 0.00044 | 0 | 99.4193 | 99.2317 | 128.0569 | 21.2612 | 0.001829 | 0.001689 | -152.7874 | 0.21948 | 8.4019 | 0.003478 | 0.004999 | 0.004849 | 4.6e-05 | 0.00119 | 0.32173 | 1.3405 | 2.5476 | 100.79 | 21.09 | 16.747 | 65.8326 | 105.5644 | 62.4775 | 0.00416 | 1.0519 | 0.025511 | 24.2521 | 0.01689 | 0.13057 | -3.8e-05 | -0.002931 | -0.002931 | -1.8395
2019-03-26 16:07 | 1553630832.6854 | -347 | 0.2095 | 0.000442 | 0 | 99.3638 | 99.345 | 127.7506 | 21.2612 | 0.001837 | 0.001703 | -152.7796 | 0.21906 | 8.4056 | 0.003492 | 0.005047 | 0.004866 | 4.6e-05 | 0.001184 | 0.33196 | 1.3537 | 2.5474 | 100.79 | 21.09 | 16.747 | 66.0315 | 105.8886 | 62.8064 | 0.00416 | 1.0509 | 0.025468 | 24.235 | 0.00169 | 0.13032 | -1.5e-05 | -0.001179 | -0.001179 | -4.5632
2019-03-26 16:07 | 1553630833.6841 | -346 | 0.20948 | 0.000442 | 1e-06 | 99.4258 | 99.3584 | 127.6171 | 21.2612 | 0.001811 | 0.001747 | -152.776 | 0.21876 | 8.4019 | 0.003523 | 0.005272 | 0.005108 | 4.7e-05 | 0.001181 | 0.35098 | 1.3573 | 2.5474 | 100.79 | 21.09 | 16.745 | 65.3946 | 106.6433 | 63.8537 | 0.00416 | 1.0476 | 0.025415 | 24.26 | 0.006058 | 0.13014 | 0.00011 | 0.008421 | 0.008421 | 0.63824
2019-03-26 16:07 | 1553630834.6851 | -345 | 0.20951 | 0.000437 | 1e-06 | 99.4393 | 99.3487 | 127.6736 | 21.2612 | 0.001778 | 0.001698 | -152.7724 | 0.21882 | 8.4058 | 0.003512 | 0.004955 | 0.004863 | 4.5e-05 | 0.001181 | 0.33052 | 1.3409 | 2.5475 | 100.79 | 21.09 | 16.744 | 64.6134 | 106.3837 | 62.6877 | 0.004159 | 1.0512 | 0.025465 | 24.2234 | 0.008157 | 0.13018 | -8.4e-05 | -0.006523 | -0.006523 | -0.82417
2019-03-26 16:07 | 1553630835.6845 | -344 | 0.20951 | 0.000439 | 0 | 99.5255 | 99.4576 | 127.8138 | 21.2612 | 0.00183 | 0.00166 | -152.7559 | 0.21901 | 8.4042 | 0.003525 | 0.004397 | 0.004204 | 4.5e-05 | 0.001183 | 0.34058 | 1.367 | 2.5475 | 100.79 | 21.09 | 16.743 | 65.8559 | 106.6977 | 61.7705 | 0.004159 | 1.0541 | 0.025513 | 24.2036 | 0.006106 | 0.13029 | -5.5e-05 | -0.004281 | -0.004281 | -1.257
2019-03-26 16:07 | 1553630836.6843 | -343 | 0.20951 | 0.000441 | 0 | 99.6117 | 99.4946 | 127.9933 | 21.2612 | 0.00176 | 0.001736 | -152.7454 | 0.21854 | 8.4036 | 0.003523 | 0.004765 | 0.004502 | 4.9e-05 | 0.001186 | 0.34811 | 1.3172 | 2.5474 | 100.79 | 21.09 | 16.739 | 64.1745 | 106.6426 | 63.599 | 0.004158 | 1.0484 | 0.025462 | 24.2866 | 0.010524 | 0.13001 | -4.1e-05 | -0.003163 | -0.003163 | -1.6973
2019-03-26 16:07 | 1553630837.6857 | -342 | 0.20952 | 0.000437 | 0 | 99.6472 | 99.4768 | 128.2088 | 21.2612 | 0.001772 | 0.001716 | -152.7541 | 0.21965 | 8.4013 | 0.003478 | 0.004855 | 0.004736 | 5e-05 | 0.00119 | 0.32781 | 1.3515 | 2.5475 | 100.79 | 21.09 | 16.737 | 64.4666 | 105.5549 | 63.1133 | 0.004158 | 1.0499 | 0.025502 | 24.2895 | 0.015306 | 0.13067 | -0.000142 | -0.01097 | -0.01097 | -0.49194
2019-03-26 16:07 | 1553630838.6844 | -341 | 0.2095 | 0.000442 | 0 | 99.6693 | 99.6714 | 128.2682 | 21.2612 | 0.00182 | 0.001697 | -152.7538 | 0.22002 | 8.4007 | 0.003513 | 0.004697 | 0.00453 | 5e-05 | 0.001195 | 0.33699 | 1.3632 | 2.5474 | 100.79 | 21.09 | 16.734 | 65.6164 | 106.4058 | 62.6487 | 0.004157 | 1.0514 | 0.025525 | 24.2783 | -0.000191 | 0.13089 | 3e-06 | 0.000238 | 0.000238 | 22.6939
2019-03-26 16:07 | 1553630839.6854 | -340 | 0.20948 | 0.000441 | 0 | 99.7691 | 99.7615 | 128.3083 | 21.2612 | 0.001816 | 0.0017 | -152.7687 | 0.219 | 8.4032 | 0.003521 | 0.004594 | 0.004384 | 4.7e-05 | 0.001199 | 0.33395 | 1.3301 | 2.5473 | 100.79 | 21.09 | 16.725 | 65.5343 | 106.5892 | 62.7346 | 0.004155 | 1.0511 | 0.025526 | 24.2852 | 0.000677 | 0.13028 | 0.000105 | 0.008112 | 0.008112 | 0.66332
2019-03-26 16:07 | 1553630840.6848 | -339 | 0.20947 | 0.000438 | 0 | 99.8128 | 99.7171 | 128.1055 | 21.2612 | 0.001807 | 0.001708 | -152.7444 | 0.22052 | 8.4052 | 0.003593 | 0.004901 | 0.00471 | 5.2e-05 | 0.001204 | 0.34987 | 1.3456 | 2.5474 | 100.79 | 21.09 | 16.72 | 65.3183 | 108.336 | 62.9282 | 0.004153 | 1.0505 | 0.025498 | 24.273 | 0.008573 | 0.13119 | 0.00019 | 0.01456 | 0.01456 | 0.37212
2019-03-26 16:07 | 1553630841.6856 | -338 | 0.20948 | 0.000437 | 1e-06 | 99.8789 | 99.8513 | 128.1384 | 21.2612 | 0.001799 | 0.001752 | -152.7432 | 0.22028 | 8.4038 | 0.003446 | 0.005333 | 0.005142 | 5e-05 | 0.001205 | 0.3804 | 1.3435 | 2.5474 | 100.79 | 21.09 | 16.71 | 65.1097 | 104.7723 | 63.9803 | 0.004151 | 1.0472 | 0.025462 | 24.3141 | 0.002473 | 0.13105 | 9.6e-05 | 0.007352 | 0.007352 | 0.73611
2019-03-26 16:07 | 1553630842.6848 | -337 | 0.20948 | 0.00044 | 0 | 99.8919 | 99.9588 | 128.4318 | 21.2612 | 0.001874 | 0.001686 | -152.7376 | 0.21996 | 8.4035 | 0.003477 | 0.00449 | 0.004348 | 4.6e-05 | 0.001194 | 0.31349 | 1.3135 | 2.5475 | 100.79 | 21.09 | 16.704 | 66.9281 | 105.5392 | 62.3925 | 0.004149 | 1.0522 | 0.025551 | 24.2845 | -0.005995 | 0.13086 | 0.000116 | 0.008913 | 0.008913 | 0.60636
2019-03-26 16:07 | 1553630843.6846 | -336 | 0.20947 | 0.000439 | 1e-06 | 99.9084 | 99.8661 | 128.5851 | 21.2612 | 0.001823 | 0.001696 | -152.7327 | 0.22003 | 8.4025 | 0.003447 | 0.004624 | 0.004489 | 5.1e-05 | 0.001195 | 0.32587 | 1.3329 | 2.5476 | 100.79 | 21.09 | 16.696 | 65.6866 | 104.8053 | 62.6438 | 0.004147 | 1.0514 | 0.025557 | 24.3081 | 0.003791 | 0.1309 | 0.000203 | 0.015663 | 0.015663 | 0.34513
2019-03-26 16:07 | 1553630844.6846 | -335 | 0.20951 | 0.000442 | 0 | 99.9512 | 99.7595 | 128.7599 | 21.2612 | 0.001807 | 0.00165 | -152.7343 | 0.22048 | 8.4047 | 0.003434 | 0.004721 | 0.004522 | 4.9e-05 | 0.001192 | 0.33572 | 1.3147 | 2.5475 | 100.79 | 21.09 | 16.678 | 65.3027 | 104.4979 | 61.5219 | 0.004143 | 1.0549 | 0.025617 | 24.284 | 0.017172 | 0.13116 | -8.9e-05 | -0.00689 | -0.00689 | -0.78621
2019-03-26 16:07 | 1553630845.6847 | -334 | 0.20953 | 0.000442 | 0 | 99.9518 | 99.9561 | 128.8313 | 21.2612 | 0.001791 | 0.001659 | -152.7685 | 0.21905 | 8.4005 | 0.003411 | 0.004798 | 0.004668 | 4.9e-05 | 0.001187 | 0.35579 | 1.3408 | 2.5476 | 100.79 | 21.09 | 16.67 | 64.9247 | 103.9217 | 61.7493 | 0.004141 | 1.0542 | 0.025616 | 24.2989 | -0.00039 | 0.13031 | -0.000217 | -0.016785 | -0.016785 | -0.32064
2019-03-26 16:07 | 1553630846.6837 | -333 | 0.20953 | 0.000437 | 0 | 100.0099 | 100.0106 | 128.5032 | 21.2612 | 0.001787 | 0.001654 | -152.7911 | 0.21981 | 8.4013 | 0.003495 | 0.004435 | 0.004264 | 4.5e-05 | 0.001188 | 0.3194 | 1.289 | 2.5473 | 100.79 | 21.1 | 16.65 | 64.8367 | 105.976 | 61.6127 | 0.004139 | 1.0546 | 0.025588 | 24.263 | -6e-05 | 0.13077 | -0.000169 | -0.013063 | -0.013063 | -0.41343
2019-03-26 16:07 | 1553630847.6848 | -332 | 0.20951 | 0.000438 | 0 | 100.0365 | 99.8405 | 128.0614 | 21.2612 | 0.00179 | 0.001707 | -152.7693 | 0.21936 | 8.4037 | 0.00343 | 0.004798 | 0.004668 | 5.1e-05 | 0.001192 | 0.29623 | 1.3421 | 2.5476 | 100.79 | 21.1 | 16.64 | 64.8927 | 104.3879 | 62.8892 | 0.004136 | 1.0506 | 0.025496 | 24.2674 | 0.017542 | 0.1305 | -6.7e-05 | -0.005173 | -0.005173 | -1.0419
2019-03-26 16:07 | 1553630848.6837 | -331 | 0.20948 | 0.000441 | 0 | 99.9905 | 100.0842 | 128.1384 | 21.2612 | 0.001806 | 0.001724 | -152.7466 | 0.21932 | 8.4018 | 0.003493 | 0.004502 | 0.004275 | 4.4e-05 | 0.0012 | 0.34827 | 1.3733 | 2.5476 | 100.79 | 21.1 | 16.629 | 65.2846 | 105.9153 | 63.3124 | 0.004133 | 1.0493 | 0.025487 | 24.29 | -0.008384 | 0.13047 | 0.000136 | 0.010446 | 0.010446 | 0.51585
2019-03-26 16:07 | 1553630849.6848 | -330 | 0.20949 | 0.00044 | 0 | 100.0519 | 99.8462 | 128.3393 | 21.2612 | 0.00178 | 0.001695 | -152.7261 | 0.21994 | 8.4025 | 0.003512 | 0.004373 | 0.004201 | 4.6e-05 | 0.001207 | 0.34811 | 1.3291 | 2.5475 | 100.79 | 21.1 | 16.603 | 64.6478 | 106.3726 | 62.6188 | 0.004127 | 1.0515 | 0.025533 | 24.284 | 0.018402 | 0.13084 | 3.9e-05 | 0.003014 | 0.003014 | 1.793
2019-03-26 16:07 | 1553630850.6844 | -329 | 0.20952 | 0.000437 | 0 | 100.051 | 100.0889 | 128.5938 | 21.2612 | 0.001852 | 0.001682 | -152.7469 | 0.21969 | 8.403 | 0.003596 | 0.005069 | 0.004901 | 4.4e-05 | 0.001206 | 0.3147 | 1.3547 | 2.5476 | 100.79 | 21.1 | 16.59 | 66.3834 | 108.4123 | 62.3024 | 0.004124 | 1.0524 | 0.025571 | 24.2966 | -0.003393 | 0.13069 | -0.000129 | -0.009977 | -0.009977 | -0.541
2019-03-26 16:08 | 1553630851.6853 | -328 | 0.2095 | 0.000437 | 0 | 100.0817 | 100.0936 | 129.1972 | 21.2612 | 0.001799 | 0.0017 | -152.7626 | 0.21987 | 8.4052 | 0.003568 | 0.00547 | 0.005265 | 5e-05 | 0.001218 | 0.33644 | 1.3971 | 2.5476 | 100.79 | 21.1 | 16.562 | 65.1198 | 107.7481 | 62.7315 | 0.004117 | 1.0511 | 0.025614 | 24.3691 | -0.00107 | 0.1308 | 2.7e-05 | 0.002087 | 0.002087 | 2.5886
2019-03-26 16:08 | 1553630852.6839 | -327 | 0.2095 | 0.00044 | 0 | 100.1277 | 99.9037 | 129.2385 | 21.2612 | 0.001769 | 0.001718 | -152.7828 | 0.22001 | 8.4042 | 0.00358 | 0.004821 | 0.004631 | 5e-05 | 0.001213 | 0.31404 | 1.3514 | 2.5476 | 100.79 | 21.09 | 16.547 | 64.3901 | 108.0444 | 63.1672 | 0.00411 | 1.0497 | 0.025602 | 24.3888 | 0.020034 | 0.13089 | -2.9e-05 | -0.002228 | -0.002228 | -2.4267
2019-03-26 16:08 | 1553630853.6845 | -326 | 0.20949 | 0.000437 | 0 | 100.149 | 100.0301 | 128.9264 | 21.2612 | 0.001845 | 0.001736 | -152.7863 | 0.21976 | 8.4037 | 0.00368 | 0.005041 | 0.004884 | 4.9e-05 | 0.001217 | 0.33516 | 1.3687 | 2.5475 | 100.79 | 21.09 | 16.531 | 66.2196 | 110.4584 | 63.5945 | 0.004106 | 1.0484 | 0.025555 | 24.3748 | 0.010623 | 0.13074 | 4.5e-05 | 0.003413 | 0.003413 | 1.5819
2019-03-26 16:08 | 1553630854.6843 | -325 | 0.20949 | 0.000438 | 0 | 100.1109 | 100.1104 | 128.7954 | 21.2612 | 0.00179 | 0.001715 | -152.7872 | 0.2197 | 8.4054 | 0.003608 | 0.005124 | 0.004921 | 5e-05 | 0.001213 | 0.32444 | 1.3734 | 2.5476 | 100.79 | 21.09 | 16.5 | 64.9028 | 108.7079 | 63.1039 | 0.004099 | 1.0499 | 0.02556 | 24.3446 | 4e-05 | 0.1307 | 4.3e-05 | 0.003347 | 0.003347 | 1.6129
2019-03-26 16:08 | 1553630855.6841 | -324 | 0.20952 | 0.000437 | 0 | 100.1971 | 100.053 | 128.5717 | 21.2612 | 0.001847 | 0.00171 | -152.7553 | 0.21986 | 8.4027 | 0.003585 | 0.005079 | 0.004954 | 5e-05 | 0.001208 | 0.34036 | 1.346 | 2.5475 | 100.79 | 21.09 | 16.485 | 66.2634 | 108.157 | 62.9619 | 0.004095 | 1.0504 | 0.025544 | 24.3184 | 0.012875 | 0.13079 | -8.9e-05 | -0.00683 | -0.00683 | -0.79088
2019-03-26 16:08 | 1553630856.6855 | -323 | 0.20951 | 0.00044 | -1e-06 | 100.1256 | 100.2046 | 128.5161 | 21.2612 | 0.00179 | 0.00169 | -152.7269 | 0.21989 | 8.4074 | 0.003504 | 0.005218 | 0.004981 | 4.8e-05 | 0.001206 | 0.31078 | 1.3273 | 2.5475 | 100.79 | 21.09 | 16.47 | 64.9052 | 106.1838 | 62.4886 | 0.004091 | 1.0519 | 0.025556 | 24.296 | -0.007047 | 0.13081 | -4.3e-05 | -0.003284 | -0.003284 | -1.6453
2019-03-26 16:08 | 1553630857.6842 | -322 | 0.20953 | 0.00044 | -1e-06 | 100.2053 | 100.0405 | 128.6435 | 21.2612 | 0.001815 | 0.001681 | -152.7228 | 0.21957 | 8.4019 | 0.003577 | 0.004814 | 0.004626 | 5e-05 | 0.001207 | 0.33892 | 1.3254 | 2.5474 | 100.79 | 21.09 | 16.44 | 65.5103 | 107.9703 | 62.2745 | 0.004084 | 1.0525 | 0.025577 | 24.3003 | 0.014725 | 0.13062 | -0.000172 | -0.01324 | -0.01324 | -0.40745
2019-03-26 16:08 | 1553630858.6852 | -321 | 0.20953 | 0.000439 | 0 | 100.1876 | 100.0899 | 128.6684 | 21.2612 | 0.001844 | 0.001666 | -152.7248 | 0.21989 | 8.4014 | 0.003734 | 0.005137 | 0.005013 | 5e-05 | 0.001206 | 0.29142 | 1.3434 | 2.5475 | 100.79 | 21.09 | 16.426 | 66.1915 | 111.7801 | 61.9251 | 0.00408 | 1.0536 | 0.025593 | 24.2899 | 0.008727 | 0.13081 | -0.000168 | -0.013012 | -0.013012 | -0.4152
2019-03-26 16:08 | 1553630859.6843 | -320 | 0.20952 | 0.000432 | 0 | 100.1711 | 100.3031 | 128.5182 | 21.2612 | 0.001789 | 0.001713 | -152.7337 | 0.22029 | 8.402 | 0.00368 | 0.005009 | 0.004795 | 5.2e-05 | 0.001212 | 0.35795 | 1.3644 | 2.5476 | 100.79 | 21.09 | 16.402 | 64.8732 | 110.4734 | 63.0562 | 0.004074 | 1.0501 | 0.025535 | 24.3167 | -0.011775 | 0.13105 | -0.000139 | -0.010697 | -0.010697 | -0.50595
2019-03-26 16:08 | 1553630860.6834 | -319 | 0.20953 | 0.00044 | -1e-06 | 100.1578 | 100.1585 | 128.7491 | 21.2612 | 0.001836 | 0.001726 | -152.719 | 0.21955 | 8.4013 | 0.003641 | 0.005111 | 0.004933 | 5e-05 | 0.001216 | 0.33445 | 1.363 | 2.5475 | 100.79 | 21.09 | 16.392 | 66.0065 | 109.5198 | 63.361 | 0.004072 | 1.0491 | 0.025546 | 24.3496 | -6.2e-05 | 0.13061 | -0.000189 | -0.014564 | -0.014564 | -0.37039
2019-03-26 16:08 | 1553630861.6844 | -318 | 0.20953 | 0.000443 | 0 | 100.1056 | 100.0183 | 128.4384 | 21.2612 | 0.001792 | 0.001722 | -152.7107 | 0.21892 | 8.4035 | 0.003613 | 0.005234 | 0.005041 | 5.4e-05 | 0.001218 | 0.34058 | 1.3637 | 2.5475 | 100.79 | 21.1 | 16.384 | 64.9428 | 108.841 | 63.2594 | 0.004072 | 1.0495 | 0.025519 | 24.3165 | 0.007798 | 0.13024 | -0.000174 | -0.013392 | -0.013392 | -0.40163
2019-03-26 16:08 | 1553630862.6844 | -317 | 0.20952 | 0.000439 | 0 | 100.2242 | 100.2409 | 128.5463 | 21.2612 | 0.001841 | 0.001718 | -152.7017 | 0.21968 | 8.4024 | 0.003643 | 0.005065 | 0.004871 | 5.1e-05 | 0.001214 | 0.31686 | 1.3825 | 2.5475 | 100.79 | 21.1 | 16.369 | 66.1323 | 109.5553 | 63.1641 | 0.004069 | 1.0497 | 0.025533 | 24.3233 | -0.001484 | 0.13069 | -0.000111 | -0.008497 | -0.008497 | -0.6352
2019-03-26 16:08 | 1553630863.6838 | -316 | 0.20952 | 0.00044 | 0 | 100.211 | 100.1854 | 129.1334 | 21.2612 | 0.001829 | 0.00172 | -152.7198 | 0.21936 | 8.4025 | 0.003524 | 0.004892 | 0.00475 | 5.3e-05 | 0.001211 | 0.35098 | 1.3555 | 2.5475 | 100.79 | 21.1 | 16.363 | 65.8333 | 106.6766 | 63.2125 | 0.004067 | 1.0496 | 0.02559 | 24.3805 | 0.002281 | 0.1305 | -0.000119 | -0.00918 | -0.00918 | -0.5871
2019-03-26 16:08 | 1553630864.6835 | -315 | 0.20952 | 0.000437 | 0 | 100.2074 | 100.1666 | 129.3261 | 21.2612 | 0.001778 | 0.001696 | -152.7386 | 0.22045 | 8.4022 | 0.00356 | 0.005669 | 0.005485 | 5e-05 | 0.001199 | 0.31006 | 1.3407 | 2.5474 | 100.79 | 21.1 | 16.354 | 64.6211 | 107.5364 | 62.6292 | 0.004065 | 1.0514 | 0.025631 | 24.3775 | 0.003646 | 0.13115 | -0.000136 | -0.010519 | -0.010519 | -0.51492
2019-03-26 16:08 | 1553630865.6846 | -314 | 0.20953 | 0.000441 | 0 | 100.1401 | 100.0469 | 129.5686 | 21.2612 | 0.001827 | 0.001752 | -152.7786 | 0.21924 | 8.4016 | 0.003544 | 0.005003 | 0.004834 | 5.2e-05 | 0.001199 | 0.33212 | 1.3524 | 2.5474 | 100.79 | 21.1 | 16.35 | 65.7795 | 107.1488 | 63.9762 | 0.004064 | 1.0472 | 0.025604 | 24.4493 | 0.008329 | 0.13043 | -0.000213 | -0.016458 | -0.016458 | -0.32729
2019-03-26 16:08 | 1553630866.6846 | -313 | 0.20953 | 0.00044 | 0 | 100.234 | 100.1995 | 129.715 | 21.2612 | 0.001773 | 0.001694 | -152.7749 | 0.21918 | 8.4015 | 0.003518 | 0.004828 | 0.004615 | 4.7e-05 | 0.001195 | 0.31421 | 1.3578 | 2.5475 | 100.79 | 21.1 | 16.347 | 64.4986 | 106.5261 | 62.593 | 0.004063 | 1.0515 | 0.025671 | 24.4128 | 0.003075 | 0.13039 | -0.000202 | -0.015616 | -0.015616 | -0.34485
2019-03-26 16:08 | 1553630867.6837 | -312 | 0.20952 | 0.000442 | 0 | 100.2597 | 100.4217 | 129.8437 | 21.2612 | 0.001825 | 0.001748 | -152.7912 | 0.2202 | 8.403 | 0.003599 | 0.004778 | 0.004618 | 5e-05 | 0.001195 | 0.30198 | 1.3469 | 2.5474 | 100.79 | 21.1 | 16.342 | 65.752 | 108.4912 | 63.8777 | 0.004062 | 1.0475 | 0.025635 | 24.4716 | -0.014447 | 0.131 | -0.000131 | -0.010142 | -0.010142 | -0.53343
2019-03-26 16:08 | 1553630868.6847 | -311 | 0.20953 | 0.000439 | 0 | 100.3343 | 100.3118 | 129.777 | 21.2612 | 0.001795 | 0.001678 | -152.7947 | 0.21934 | 8.4041 | 0.003589 | 0.00495 | 0.004789 | 5.1e-05 | 0.001195 | 0.31205 | 1.3947 | 2.5475 | 100.79 | 21.1 | 16.341 | 65.0294 | 108.2423 | 62.2104 | 0.004062 | 1.0527 | 0.025692 | 24.4047 | 0.00201 | 0.13048 | -0.000172 | -0.013363 | -0.013363 | -0.40328
2019-03-26 16:08 | 1553630869.6848 | -310 | 0.20952 | 0.000438 | 0 | 100.2608 | 100.172 | 129.7678 | 21.2612 | 0.001838 | 0.00174 | -152.7785 | 0.21987 | 8.4021 | 0.003581 | 0.004779 | 0.004494 | 5e-05 | 0.001201 | 0.32156 | 1.3906 | 2.5476 | 100.79 | 21.1 | 16.339 | 66.0573 | 108.0485 | 63.7013 | 0.004061 | 1.0481 | 0.025634 | 24.4581 | 0.007933 | 0.1308 | -0.000123 | -0.009504 | -0.009504 | -0.56838
2019-03-26 16:08 | 1553630870.6849 | -309 | 0.20953 | 0.000437 | 0 | 100.2381 | 100.4392 | 129.4728 | 21.2612 | 0.001807 | 0.001735 | -152.7943 | 0.22027 | 8.4014 | 0.003574 | 0.004956 | 0.004763 | 5.2e-05 | 0.001205 | 0.28584 | 1.3098 | 2.5475 | 100.79 | 21.1 | 16.337 | 65.2986 | 107.8773 | 63.5827 | 0.004061 | 1.0484 | 0.025609 | 24.4259 | -0.017926 | 0.13104 | -0.000193 | -0.014971 | -0.014971 | -0.36149
2019-03-26 16:08 | 1553630871.6849 | -308 | 0.20952 | 0.000442 | 0 | 100.1661 | 100.0846 | 129.3104 | 21.2612 | 0.001808 | 0.001698 | -152.763 | 0.2191 | 8.4023 | 0.003581 | 0.004556 | 0.00439 | 5.1e-05 | 0.001203 | 0.32189 | 1.3528 | 2.5476 | 100.79 | 21.1 | 16.334 | 65.3409 | 108.0672 | 62.6783 | 0.00406 | 1.0513 | 0.025628 | 24.3778 | 0.007283 | 0.13034 | -9.9e-05 | -0.007625 | -0.007625 | -0.70605
2019-03-26 16:08 | 1553630872.6849 | -307 | 0.20953 | 0.000442 | 0 | 100.2092 | 100.218 | 129.2245 | 21.2612 | 0.00183 | 0.001623 | -152.7465 | 0.22049 | 8.4027 | 0.00356 | 0.004827 | 0.004663 | 5.2e-05 | 0.001196 | 0.31996 | 1.3943 | 2.5476 | 100.79 | 21.09 | 16.329 | 65.8699 | 107.5537 | 60.871 | 0.004056 | 1.057 | 0.025688 | 24.3041 | -0.000788 | 0.13117 | -0.000192 | -0.014909 | -0.014909 | -0.36337
2019-03-26 16:08 | 1553630873.6843 | -306 | 0.20953 | 0.000436 | 0 | 100.19 | 100.1229 | 128.9897 | 21.2612 | 0.001796 | 0.0017 | -152.7565 | 0.2184 | 8.4014 | 0.003613 | 0.004382 | 0.004177 | 5.3e-05 | 0.001192 | 0.31365 | 1.381 | 2.5476 | 100.79 | 21.09 | 16.326 | 65.0311 | 108.8276 | 62.7399 | 0.004055 | 1.0511 | 0.025593 | 24.3498 | 0.005994 | 0.12992 | -0.000157 | -0.012175 | -0.012175 | -0.44074
2019-03-26 16:08 | 1553630874.6843 | -305 | 0.20951 | 0.000441 | 1e-06 | 100.1932 | 100.3367 | 128.7538 | 21.2612 | 0.001785 | 0.001702 | -152.7391 | 0.21917 | 8.4006 | 0.003597 | 0.004644 | 0.004477 | 5.1e-05 | 0.001187 | 0.3246 | 1.3656 | 2.5476 | 100.79 | 21.09 | 16.325 | 64.7664 | 108.4455 | 62.7782 | 0.004055 | 1.051 | 0.025569 | 24.3289 | -0.012796 | 0.13038 | -7.3e-05 | -0.005681 | -0.005681 | -0.94788
2019-03-26 16:08 | 1553630875.6831 | -304 | 0.20951 | 0.000437 | 0 | 100.2056 | 100.1871 | 128.6222 | 21.2612 | 0.001779 | 0.001683 | -152.6988 | 0.21893 | 8.4011 | 0.003577 | 0.004967 | 0.004752 | 4.9e-05 | 0.00118 | 0.32836 | 1.4267 | 2.5476 | 100.79 | 21.09 | 16.324 | 64.6312 | 107.964 | 62.3292 | 0.004055 | 1.0524 | 0.025573 | 24.3002 | 0.001657 | 0.13024 | -2.8e-05 | -0.002165 | -0.002165 | -2.4845
2019-03-26 16:08 | 1553630876.6843 | -303 | 0.2095 | 0.000439 | 0 | 100.2487 | 100.4046 | 128.4405 | 21.2612 | 0.00181 | 0.001707 | -152.7176 | 0.21899 | 8.4001 | 0.003627 | 0.004677 | 0.004527 | 5.4e-05 | 0.001182 | 0.3637 | 1.3761 | 2.5477 | 100.79 | 21.09 | 16.324 | 65.3792 | 109.177 | 62.9048 | 0.004055 | 1.0506 | 0.025533 | 24.3039 | -0.013894 | 0.13028 | -1.3e-05 | -0.001033 | -0.001033 | -5.2064
2019-03-26 16:08 | 1553630877.6841 | -302 | 0.20951 | 0.000437 | 0 | 100.1903 | 100.0906 | 128.4844 | 21.2612 | 0.001782 | 0.001711 | -152.7266 | 0.21967 | 8.4019 | 0.003564 | 0.004877 | 0.004617 | 5.2e-05 | 0.001182 | 0.33301 | 1.3742 | 2.5475 | 100.79 | 21.09 | 16.324 | 64.6952 | 107.653 | 62.9953 | 0.004055 | 1.0503 | 0.025534 | 24.3113 | 0.008904 | 0.13068 | -6.4e-05 | -0.004889 | -0.004889 | -1.1038
2019-03-26 16:08 | 1553630878.6832 | -301 | 0.20949 | 0.00044 | 0 | 100.1885 | 100.2947 | 128.4926 | 21.2612 | 0.001852 | 0.001755 | -152.7312 | 0.21976 | 8.4031 | 0.003608 | 0.005115 | 0.004924 | 5.1e-05 | 0.001187 | 0.32516 | 1.4159 | 2.5475 | 100.79 | 21.09 | 16.325 | 66.3997 | 108.7117 | 64.0481 | 0.004055 | 1.047 | 0.025495 | 24.3501 | -0.009471 | 0.13073 | 8.3e-05 | 0.006429 | 0.006429 | 0.83983
2019-03-26 16:08 | 1553630879.6843 | -300 | 0.20954 | 0.000439 | 0 | 100.1962 | 100.2328 | 128.6224 | 21.2612 | 0.001803 | 0.001717 | -152.7562 | 0.21974 | 8.4017 | 0.003562 | 0.005332 | 0.005232 | 5e-05 | 0.001192 | 0.35817 | 1.3851 | 2.5476 | 100.79 | 21.09 | 16.327 | 65.2214 | 107.5877 | 63.1383 | 0.004056 | 1.0498 | 0.025542 | 24.3295 | -0.003267 | 0.13072 | -0.00024 | -0.018541 | -0.018541 | -0.29119
2019-03-26 16:08 | 1553630880.6846 | -299 | 0.20954 | 0.000438 | 0 | 100.182 | 100.2499 | 128.5691 | 21.2612 | 0.001849 | 0.001697 | -152.7621 | 0.21889 | 8.4023 | 0.003551 | 0.003837 | 0.003598 | 5.1e-05 | 0.0012 | 0.31116 | 1.3845 | 2.5477 | 100.79 | 21.08 | 16.328 | 66.3069 | 107.3364 | 62.6682 | 0.004053 | 1.0513 | 0.025554 | 24.3075 | -0.006062 | 0.13022 | -0.000244 | -0.018818 | -0.018818 | -0.28578
2019-03-26 16:08 | 1553630881.6844 | -298 | 0.20951 | 0.000439 | 0 | 100.1953 | 100.2819 | 128.3806 | 21.2612 | 0.001865 | 0.001673 | -152.7949 | 0.21919 | 8.4025 | 0.003554 | 0.004915 | 0.004777 | 5.3e-05 | 0.001202 | 0.33428 | 1.393 | 2.5476 | 100.79 | 21.09 | 16.328 | 66.7011 | 107.4056 | 62.0758 | 0.004056 | 1.0532 | 0.025558 | 24.2682 | -0.007725 | 0.1304 | -3.7e-05 | -0.00288 | -0.00288 | -1.8696
2019-03-26 16:08 | 1553630882.6844 | -297 | 0.20948 | 0.00044 | 0 | 100.1788 | 100.1723 | 128.0539 | 21.2612 | 0.00183 | 0.001725 | -152.808 | 0.21849 | 8.4 | 0.003517 | 0.004888 | 0.0047 | 5.3e-05 | 0.001205 | 0.32947 | 1.3197 | 2.5474 | 100.79 | 21.09 | 16.33 | 65.8552 | 106.5096 | 63.3321 | 0.004056 | 1.0492 | 0.025478 | 24.2827 | 0.000579 | 0.12998 | 0.000128 | 0.009823 | 0.009823 | 0.54647
2019-03-26 16:08 | 1553630883.6848 | -296 | 0.20951 | 0.000437 | 0 | 100.2487 | 100.0977 | 127.8553 | 21.2612 | 0.00183 | 0.001691 | -152.7905 | 0.22005 | 8.4043 | 0.003574 | 0.004801 | 0.004627 | 5.2e-05 | 0.001212 | 0.34003 | 1.3519 | 2.5476 | 100.79 | 21.09 | 16.332 | 65.8574 | 107.8924 | 62.5126 | 0.004057 | 1.0518 | 0.025489 | 24.2343 | 0.013489 | 0.13091 | -6.7e-05 | -0.005158 | -0.005158 | -1.0481
2019-03-26 16:08 | 1553630884.6837 | -295 | 0.20955 | 0.000442 | 0 | 100.19 | 100.1464 | 127.9093 | 21.2612 | 0.001838 | 0.001692 | -152.7904 | 0.21932 | 8.4036 | 0.003565 | 0.005085 | 0.004941 | 5.1e-05 | 0.001216 | 0.33301 | 1.3561 | 2.5476 | 100.79 | 21.09 | 16.332 | 66.0611 | 107.6576 | 62.529 | 0.004057 | 1.0517 | 0.025494 | 24.24 | 0.003892 | 0.13047 | -0.000287 | -0.022054 | -0.022054 | -0.24433
2019-03-26 16:08 | 1553630885.6839 | -294 | 0.20947 | 0.000434 | 0 | 100.1209 | 100.1676 | 127.814 | 21.2612 | 0.00179 | 0.00173 | -152.7908 | 0.21928 | 8.4028 | 0.003556 | 0.005374 | 0.005102 | 5.5e-05 | 0.001219 | 0.30525 | 1.3832 | 2.5476 | 100.79 | 21.09 | 16.329 | 64.9038 | 107.4583 | 63.464 | 0.004056 | 1.0488 | 0.025449 | 24.2647 | -0.004168 | 0.13045 | 0.000173 | 0.013248 | 0.013248 | 0.40667
2019-03-26 16:08 | 1553630886.6846 | -293 | 0.20953 | 0.000436 | -1e-06 | 100.2455 | 100.2553 | 128.1102 | 21.2612 | 0.001802 | 0.001668 | -152.7948 | 0.21913 | 8.4019 | 0.003594 | 0.004744 | 0.004634 | 5.3e-05 | 0.001221 | 0.30182 | 1.3476 | 2.5474 | 100.79 | 21.09 | 16.324 | 65.1911 | 108.3613 | 61.9557 | 0.004055 | 1.0535 | 0.025536 | 24.2383 | -0.000878 | 0.13036 | -0.000177 | -0.013635 | -0.013635 | -0.39487
2019-03-26 16:08 | 1553630887.6846 | -292 | 0.20951 | 0.000437 | 0 | 100.102 | 100.135 | 128.0997 | 21.2612 | 0.00183 | 0.001716 | -152.7835 | 0.21881 | 8.4041 | 0.003579 | 0.005586 | 0.00532 | 5.1e-05 | 0.001223 | 0.34595 | 1.3827 | 2.5475 | 100.79 | 21.09 | 16.321 | 65.8709 | 108.0089 | 63.1172 | 0.004054 | 1.0499 | 0.025491 | 24.2793 | -0.002944 | 0.13017 | -3.8e-05 | -0.002913 | -0.002913 | -1.8456
2019-03-26 16:08 | 1553630888.6848 | -291 | 0.20951 | 0.00044 | 0 | 100.1256 | 100.1199 | 128.1339 | 21.2612 | 0.001811 | 0.001689 | -152.7879 | 0.21913 | 8.4042 | 0.003535 | 0.004704 | 0.004586 | 5.6e-05 | 0.001226 | 0.32565 | 1.3869 | 2.5475 | 100.79 | 21.09 | 16.319 | 65.4088 | 106.952 | 62.4792 | 0.004054 | 1.0519 | 0.025518 | 24.2595 | 0.000518 | 0.13036 | -7.6e-05 | -0.005817 | -0.005817 | -0.92549
2019-03-26 16:08 | 1553630889.6848 | -290 | 0.2095 | 0.000439 | 0 | 100.2033 | 100.2268 | 127.8539 | 21.2612 | 0.001813 | 0.001686 | -152.7887 | 0.21948 | 8.4003 | 0.003536 | 0.004579 | 0.004364 | 5.3e-05 | 0.001226 | 0.29806 | 1.3395 | 2.5474 | 100.79 | 21.09 | 16.313 | 65.4533 | 106.9646 | 62.3985 | 0.004052 | 1.0521 | 0.025493 | 24.23 | -0.002095 | 0.13057 | 3.1e-05 | 0.0024 | 0.0024 | 2.2472
2019-03-26 16:08 | 1553630890.6852 | -289 | 0.2095 | 0.00044 | 0 | 100.0716 | 100.22 | 127.8032 | 21.2612 | 0.001792 | 0.001692 | -152.7849 | 0.21949 | 8.4024 | 0.003561 | 0.004954 | 0.004768 | 5.6e-05 | 0.001228 | 0.3246 | 1.3447 | 2.5473 | 100.79 | 21.09 | 16.311 | 64.9396 | 107.5789 | 62.5478 | 0.004052 | 1.0517 | 0.025483 | 24.2306 | -0.013254 | 0.13058 | -1.5e-05 | -0.001099 | -0.001099 | -4.9079
2019-03-26 16:08 | 1553630891.6839 | -288 | 0.2095 | 0.000436 | 0 | 100.1206 | 100.2385 | 128.0567 | 21.2612 | 0.00181 | 0.001678 | -152.7742 | 0.22018 | 8.404 | 0.003513 | 0.005019 | 0.004809 | 5.3e-05 | 0.001227 | 0.28871 | 1.3449 | 2.5475 | 100.79 | 21.09 | 16.308 | 65.3744 | 106.4065 | 62.1924 | 0.004051 | 1.0528 | 0.025522 | 24.2418 | -0.010526 | 0.13098 | -2e-05 | -0.001528 | -0.001528 | -3.5412
2019-03-26 16:08 | 1553630892.6833 | -287 | 0.20948 | 0.000439 | 0 | 100.1292 | 100.1827 | 128.4149 | 21.2612 | 0.001804 | 0.0017 | -152.7752 | 0.21995 | 8.4037 | 0.003562 | 0.004807 | 0.004659 | 5.4e-05 | 0.001231 | 0.33212 | 1.3601 | 2.5474 | 100.79 | 21.09 | 16.303 | 65.2433 | 107.6025 | 62.7211 | 0.00405 | 1.0511 | 0.025537 | 24.2948 | -0.00478 | 0.13085 | 0.000128 | 0.009921 | 0.009921 | 0.54471
2019-03-26 16:08 | 1553630893.684 | -286 | 0.20951 | 0.000441 | 0 | 100.138 | 100.0926 | 128.4475 | 21.2612 | 0.001821 | 0.001715 | -152.7746 | 0.21955 | 8.4021 | 0.003482 | 0.005046 | 0.004812 | 5.3e-05 | 0.001232 | 0.33301 | 1.3564 | 2.5475 | 100.79 | 21.09 | 16.302 | 65.6421 | 105.6432 | 63.1015 | 0.00405 | 1.0499 | 0.025526 | 24.3117 | 0.004058 | 0.13061 | -6e-05 | -0.004604 | -0.004604 | -1.1716
2019-03-26 16:08 | 1553630894.683 | -285 | 0.20953 | 0.000434 | 0 | 100.1776 | 100.4076 | 128.5517 | 21.2612 | 0.001834 | 0.001727 | -152.7716 | 0.22026 | 8.4002 | 0.003526 | 0.005072 | 0.005008 | 5.2e-05 | 0.001228 | 0.2844 | 1.3291 | 2.5474 | 100.79 | 21.09 | 16.3 | 65.9495 | 106.7142 | 63.3788 | 0.004049 | 1.0491 | 0.025526 | 24.3315 | -0.020513 | 0.13103 | -0.000162 | -0.012442 | -0.012442 | -0.43494
2019-03-26 16:08 | 1553630895.6847 | -284 | 0.2095 | 0.000439 | -1e-06 | 100.136 | 100.0496 | 128.7177 | 21.2612 | 0.001778 | 0.001675 | -152.7882 | 0.21978 | 8.4037 | 0.003499 | 0.00458 | 0.004403 | 5.3e-05 | 0.001224 | 0.32244 | 1.3578 | 2.5475 | 100.79 | 21.09 | 16.299 | 64.6033 | 106.0735 | 62.1315 | 0.004049 | 1.053 | 0.02559 | 24.3021 | 0.007719 | 0.13075 | 2.4e-05 | 0.001924 | 0.001924 | 2.8066
2019-03-26 16:08 | 1553630896.6831 | -283 | 0.20951 | 0.00044 | 0 | 100.0899 | 100.0035 | 128.7508 | 21.2612 | 0.001842 | 0.001667 | -152.7636 | 0.22009 | 8.4045 | 0.003503 | 0.004781 | 0.004642 | 5.3e-05 | 0.001208 | 0.3063 | 1.3657 | 2.5474 | 100.79 | 21.09 | 16.3 | 66.1556 | 106.1516 | 61.9424 | 0.004049 | 1.0536 | 0.0256 | 24.2983 | 0.007723 | 0.13093 | -5.1e-05 | -0.00392 | -0.00392 | -1.3795
2019-03-26 16:08 | 1553630897.6845 | -282 | 0.2095 | 0.000438 | 0 | 100.1428 | 100.3518 | 128.8078 | 21.2612 | 0.001822 | 0.001665 | -152.764 | 0.21957 | 8.3993 | 0.003426 | 0.004445 | 0.0042 | 5e-05 | 0.001204 | 0.3402 | 1.3526 | 2.5476 | 100.79 | 21.09 | 16.308 | 65.6695 | 104.3067 | 61.8774 | 0.004051 | 1.0538 | 0.025608 | 24.3014 | -0.018651 | 0.13062 | -2e-05 | -0.001516 | -0.001516 | -3.5577
2019-03-26 16:08 | 1553630898.6834 | -281 | 0.2095 | 0.000437 | 0 | 100.1546 | 100.133 | 129.1498 | 21.2612 | 0.001819 | 0.001792 | -152.7576 | 0.22008 | 8.4026 | 0.003473 | 0.00506 | 0.004843 | 4.8e-05 | 0.001197 | 0.32549 | 1.3883 | 2.5475 | 100.79 | 21.09 | 16.311 | 65.6024 | 105.4361 | 64.9442 | 0.004052 | 1.0442 | 0.025526 | 24.4447 | 0.00193 | 0.13093 | -1.7e-05 | -0.001307 | -0.001307 | -4.1361
2019-03-26 16:08 | 1553630899.6832 | -280 | 0.20951 | 0.000441 | 0 | 100.1991 | 100.2728 | 129.3751 | 21.2612 | 0.001815 | 0.001759 | -152.7525 | 0.21912 | 8.4044 | 0.003448 | 0.004541 | 0.004376 | 4.7e-05 | 0.001197 | 0.33212 | 1.3746 | 2.5475 | 100.79 | 21.09 | 16.317 | 65.5023 | 104.8329 | 64.155 | 0.004053 | 1.0467 | 0.025578 | 24.4375 | -0.006573 | 0.13036 | -5.4e-05 | -0.004167 | -0.004167 | -1.292
2019-03-26 16:08 | 1553630900.6833 | -279 | 0.20949 | 0.000433 | 0 | 100.1735 | 100.2244 | 129.6329 | 21.2612 | 0.001857 | 0.001675 | -152.7606 | 0.21951 | 8.4027 | 0.00339 | 0.004855 | 0.004761 | 5.3e-05 | 0.001195 | 0.31636 | 1.3595 | 2.5475 | 100.79 | 21.09 | 16.32 | 66.5061 | 103.4217 | 62.1259 | 0.004054 | 1.053 | 0.025681 | 24.3881 | -0.004547 | 0.13059 | 9.1e-05 | 0.007047 | 0.007047 | 0.7653
2019-03-26 16:08 | 1553630901.6833 | -278 | 0.2095 | 0.000441 | 0 | 100.1504 | 100.0251 | 129.588 | 21.2612 | 0.001887 | 0.00176 | -152.7594 | 0.21961 | 8.4025 | 0.00345 | 0.004607 | 0.004316 | 5.2e-05 | 0.001197 | 0.31957 | 1.3696 | 2.5474 | 100.79 | 21.1 | 16.323 | 67.2215 | 104.8785 | 64.1713 | 0.004057 | 1.0466 | 0.025598 | 24.4582 | 0.011205 | 0.13065 | -1.1e-05 | -0.000887 | -0.000887 | -6.0814
2019-03-26 16:08 | 1553630902.6844 | -277 | 0.20948 | 0.000437 | 0 | 100.1543 | 100.1474 | 129.1782 | 21.2612 | 0.001875 | 0.001728 | -152.7664 | 0.22014 | 8.4035 | 0.003451 | 0.005178 | 0.004989 | 5.5e-05 | 0.001199 | 0.32997 | 1.3416 | 2.5476 | 100.79 | 21.1 | 16.326 | 66.9358 | 104.9115 | 63.3976 | 0.004058 | 1.049 | 0.025587 | 24.3914 | 0.000613 | 0.13096 | 0.000154 | 0.011873 | 0.011873 | 0.45554
2019-03-26 16:08 | 1553630903.6837 | -276 | 0.20953 | 0.000442 | 0 | 100.211 | 100.2486 | 129.3021 | 21.2612 | 0.001856 | 0.001687 | -152.7676 | 0.21995 | 8.4021 | 0.003488 | 0.004629 | 0.00449 | 5.1e-05 | 0.001195 | 0.3282 | 1.3906 | 2.5474 | 100.79 | 21.1 | 16.328 | 66.4943 | 105.788 | 62.4235 | 0.004058 | 1.0521 | 0.025637 | 24.3678 | -0.00336 | 0.13085 | -0.000182 | -0.0141 | -0.0141 | -0.38326
2019-03-26 16:08 | 1553630904.6849 | -275 | 0.20949 | 0.000441 | 0 | 100.1767 | 100.0271 | 129.4888 | 21.2612 | 0.001778 | 0.001669 | -152.7711 | 0.21956 | 8.4017 | 0.003473 | 0.004541 | 0.004316 | 5.2e-05 | 0.001199 | 0.31868 | 1.3452 | 2.5476 | 100.79 | 21.1 | 16.329 | 64.6204 | 105.4276 | 61.9964 | 0.004059 | 1.0534 | 0.025671 | 24.3698 | 0.013371 | 0.13062 | 9.7e-05 | 0.007547 | 0.007547 | 0.71473
2019-03-26 16:08 | 1553630905.6846 | -274 | 0.20949 | 0.000438 | -1e-06 | 100.1915 | 100.1498 | 129.4503 | 21.2612 | 0.001783 | 0.00171 | -152.7526 | 0.21972 | 8.4037 | 0.003486 | 0.004511 | 0.00435 | 5.2e-05 | 0.001203 | 0.29336 | 1.3515 | 2.5476 | 100.79 | 21.1 | 16.329 | 64.7226 | 105.7516 | 62.9734 | 0.004059 | 1.0503 | 0.02563 | 24.4017 | 0.003723 | 0.13071 | 7.2e-05 | 0.005583 | 0.005583 | 0.96702
2019-03-26 16:08 | 1553630906.6836 | -273 | 0.20952 | 0.000437 | 0 | 100.2257 | 100.2052 | 129.201 | 21.2612 | 0.001806 | 0.00172 | -152.7682 | 0.21956 | 8.4034 | 0.00349 | 0.004936 | 0.004813 | 5e-05 | 0.001203 | 0.33395 | 1.3338 | 2.5475 | 100.79 | 21.1 | 16.328 | 65.2854 | 105.8532 | 63.2204 | 0.004058 | 1.0496 | 0.025596 | 24.3872 | 0.001827 | 0.13062 | -8.4e-05 | -0.006537 | -0.006537 | -0.82522
2019-03-26 16:08 | 1553630907.6826 | -272 | 0.20952 | 0.00044 | 0 | 100.2189 | 100.217 | 128.874 | 21.2612 | 0.001782 | 0.001695 | -152.7894 | 0.21994 | 8.4026 | 0.003434 | 0.004356 | 0.004176 | 5.3e-05 | 0.001208 | 0.2891 | 1.3612 | 2.5476 | 100.79 | 21.1 | 16.327 | 64.6952 | 104.4846 | 62.6017 | 0.004058 | 1.0515 | 0.025587 | 24.3339 | 0.000171 | 0.13084 | -0.000124 | -0.00956 | -0.00956 | -0.56524
2019-03-26 16:08 | 1553630908.683 | -271 | 0.20947 | 0.00044 | -1e-06 | 100.1876 | 100.2731 | 128.2795 | 21.2612 | 0.001805 | 0.001735 | -152.7932 | 0.21913 | 8.4036 | 0.003474 | 0.005027 | 0.004841 | 4.9e-05 | 0.001208 | 0.3131 | 1.3559 | 2.5476 | 100.79 | 21.1 | 16.325 | 65.262 | 105.4662 | 63.5656 | 0.004058 | 1.0485 | 0.025492 | 24.3125 | -0.007631 | 0.13036 | 0.00018 | 0.01387 | 0.01387 | 0.38817
2019-03-26 16:08 | 1553630909.6831 | -270 | 0.20952 | 0.000439 | -1e-06 | 100.2694 | 100.3525 | 128.1738 | 21.2612 | 0.001845 | 0.00174 | -152.781 | 0.22022 | 8.4033 | 0.003412 | 0.005141 | 0.00499 | 5.2e-05 | 0.001205 | 0.31598 | 1.3262 | 2.5475 | 100.79 | 21.1 | 16.323 | 66.2211 | 103.9476 | 63.6929 | 0.004057 | 1.0481 | 0.025476 | 24.3071 | -0.007408 | 0.13101 | -9e-05 | -0.006912 | -0.006912 | -0.78282
2019-03-26 16:08 | 1553630910.6828 | -269 | 0.20948 | 0.000434 | 0 | 100.213 | 100.1501 | 127.7919 | 21.2612 | 0.001854 | 0.001725 | -152.7803 | 0.2191 | 8.4056 | 0.003443 | 0.00424 | 0.004068 | 5.3e-05 | 0.001207 | 0.33627 | 1.3294 | 2.5475 | 100.79 | 21.1 | 16.319 | 66.4334 | 104.7134 | 63.3391 | 0.004056 | 1.0492 | 0.025452 | 24.2581 | 0.005617 | 0.13034 | 0.000136 | 0.010481 | 0.010481 | 0.51361
2019-03-26 16:08 | 1553630911.6829 | -268 | 0.2095 | 0.000438 | 0 | 100.2942 | 100.3478 | 127.9398 | 21.2612 | 0.001849 | 0.001717 | -152.7758 | 0.21959 | 8.402 | 0.003434 | 0.005177 | 0.004989 | 5.1e-05 | 0.001202 | 0.28727 | 1.3485 | 2.5475 | 100.79 | 21.1 | 16.317 | 66.3163 | 104.4948 | 63.1516 | 0.004056 | 1.0498 | 0.025473 | 24.2654 | -0.004777 | 0.13064 | -5e-06 | -0.000351 | -0.000351 | -15.3574
2019-03-26 16:09 | 1553630912.6829 | -267 | 0.20951 | 0.000439 | 0 | 100.2083 | 100.1952 | 127.9863 | 21.2612 | 0.001863 | 0.001695 | -152.776 | 0.21927 | 8.4039 | 0.00344 | 0.004803 | 0.004666 | 5.2e-05 | 0.001206 | 0.32637 | 1.4131 | 2.5477 | 100.79 | 21.1 | 16.313 | 66.6496 | 104.6262 | 62.6219 | 0.004055 | 1.0514 | 0.025498 | 24.2507 | 0.001173 | 0.13045 | -5.6e-05 | -0.004256 | -0.004256 | -1.2657
2019-03-26 16:09 | 1553630913.6829 | -266 | 0.2095 | 0.000443 | 0 | 100.2154 | 100.2163 | 127.6286 | 21.2612 | 0.001827 | 0.001702 | -152.7824 | 0.21962 | 8.4028 | 0.003488 | 0.004948 | 0.004773 | 5.3e-05 | 0.0012 | 0.32228 | 1.3935 | 2.5476 | 100.79 | 21.1 | 16.311 | 65.778 | 105.8006 | 62.7844 | 0.004054 | 1.0509 | 0.025456 | 24.2226 | -8.5e-05 | 0.13065 | -2.5e-05 | -0.001866 | -0.001866 | -2.8924
2019-03-26 16:09 | 1553630914.684 | -265 | 0.20952 | 0.000437 | 0 | 100.2408 | 100.1746 | 127.4018 | 21.2612 | 0.001829 | 0.001745 | -152.7703 | 0.21994 | 8.4035 | 0.003448 | 0.005086 | 0.004841 | 5e-05 | 0.0012 | 0.3126 | 1.3731 | 2.5475 | 100.79 | 21.1 | 16.309 | 65.8264 | 104.8312 | 63.8255 | 0.004054 | 1.0477 | 0.025394 | 24.2385 | 0.005902 | 0.13084 | -8.7e-05 | -0.006673 | -0.006673 | -0.80976
2019-03-26 16:09 | 1553630915.683 | -264 | 0.20952 | 0.000436 | 0 | 100.2594 | 100.3068 | 127.1589 | 21.2612 | 0.001809 | 0.001743 | -152.7953 | 0.21983 | 8.4029 | 0.003395 | 0.004316 | 0.004166 | 5.2e-05 | 0.001201 | 0.31885 | 1.3553 | 2.5475 | 100.79 | 21.1 | 16.304 | 65.3611 | 103.5474 | 63.774 | 0.004053 | 1.0478 | 0.025372 | 24.2136 | -0.004227 | 0.13078 | -0.000114 | -0.008676 | -0.008676 | -0.62252
2019-03-26 16:09 | 1553630916.6845 | -263 | 0.20954 | 0.000441 | 0 | 100.3019 | 100.3172 | 127.0943 | 21.2612 | 0.001769 | 0.001708 | -152.7902 | 0.2199 | 8.4002 | 0.003459 | 0.004861 | 0.004642 | 5.2e-05 | 0.001196 | 0.33085 | 1.3685 | 2.5475 | 100.79 | 21.1 | 16.302 | 64.3901 | 105.0947 | 62.9267 | 0.004052 | 1.0505 | 0.025398 | 24.177 | -0.001365 | 0.13082 | -0.000217 | -0.016613 | -0.016613 | -0.32521
2019-03-26 16:09 | 1553630917.6841 | -262 | 0.20951 | 0.000438 | 0 | 100.3063 | 100.3605 | 127.6225 | 21.2612 | 0.001758 | 0.001723 | -152.7681 | 0.21872 | 8.402 | 0.003366 | 0.004581 | 0.004327 | 5.3e-05 | 0.001187 | 0.31404 | 1.3643 | 2.5476 | 100.79 | 21.1 | 16.296 | 64.1237 | 102.8544 | 63.2782 | 0.004051 | 1.0494 | 0.025437 | 24.2398 | -0.004836 | 0.13012 | -6.4e-05 | -0.004875 | -0.004875 | -1.1023
2019-03-26 16:09 | 1553630918.6842 | -261 | 0.20954 | 0.000442 | 0 | 100.2848 | 100.3804 | 128.2553 | 21.2612 | 0.001797 | 0.001703 | -152.7523 | 0.21951 | 8.4021 | 0.003372 | 0.004716 | 0.004495 | 5.2e-05 | 0.001188 | 0.31758 | 1.3499 | 2.5474 | 100.79 | 21.1 | 16.293 | 65.0631 | 102.9855 | 62.8049 | 0.00405 | 1.0509 | 0.025518 | 24.2827 | -0.008525 | 0.13059 | -0.000251 | -0.01929 | -0.01929 | -0.27958
2019-03-26 16:09 | 1553630919.6832 | -260 | 0.20951 | 0.000436 | 0 | 100.3293 | 100.4799 | 128.36 | 21.2612 | 0.001823 | 0.001719 | -152.7494 | 0.21934 | 8.404 | 0.003458 | 0.004655 | 0.004521 | 5.4e-05 | 0.001186 | 0.33826 | 1.3388 | 2.5474 | 100.79 | 21.1 | 16.29 | 65.6897 | 105.0603 | 63.2009 | 0.004049 | 1.0496 | 0.025513 | 24.307 | -0.013413 | 0.13048 | -3.2e-05 | -0.002478 | -0.002478 | -2.1749
2019-03-26 16:09 | 1553630920.6846 | -259 | 0.20951 | 0.000437 | 0 | 100.3432 | 100.2701 | 128.3752 | 21.2612 | 0.001823 | 0.001765 | -152.7453 | 0.21933 | 8.406 | 0.003463 | 0.005021 | 0.004856 | 5.3e-05 | 0.001189 | 0.30502 | 1.3315 | 2.5474 | 100.79 | 21.1 | 16.283 | 65.6828 | 105.1852 | 64.2909 | 0.004047 | 1.0462 | 0.025474 | 24.3478 | 0.006517 | 0.13048 | -4e-05 | -0.003027 | -0.003027 | -1.7803
2019-03-26 16:09 | 1553630921.6843 | -258 | 0.20954 | 0.000436 | 0 | 100.2989 | 100.2157 | 128.0666 | 21.2612 | 0.00184 | 0.001663 | -152.7431 | 0.21995 | 8.4051 | 0.003469 | 0.004723 | 0.004558 | 5.4e-05 | 0.001194 | 0.335 | 1.3099 | 2.5474 | 100.79 | 21.1 | 16.279 | 66.108 | 105.3299 | 61.8422 | 0.004046 | 1.0539 | 0.025536 | 24.2301 | 0.007428 | 0.13085 | -0.000223 | -0.017146 | -0.017146 | -0.31517
2019-03-26 16:09 | 1553630922.6844 | -257 | 0.2095 | 0.000435 | 0 | 100.3532 | 100.2762 | 128.2203 | 21.2612 | 0.001874 | 0.001684 | -152.7254 | 0.21944 | 8.4043 | 0.003482 | 0.004152 | 0.004005 | 5.1e-05 | 0.001194 | 0.35258 | 1.3168 | 2.5474 | 100.79 | 21.1 | 16.273 | 66.9149 | 105.6517 | 62.3438 | 0.004045 | 1.0523 | 0.025532 | 24.2628 | 0.006872 | 0.13054 | 3.9e-05 | 0.003038 | 0.003038 | 1.7747
2019-03-26 16:09 | 1553630923.6824 | -256 | 0.20953 | 0.000439 | 0 | 100.3786 | 100.3975 | 128.5186 | 21.2612 | 0.001831 | 0.001705 | -152.7161 | 0.21896 | 8.4031 | 0.003483 | 0.004905 | 0.004723 | 5.4e-05 | 0.001196 | 0.32045 | 1.3253 | 2.5475 | 100.79 | 21.1 | 16.264 | 65.8956 | 105.6713 | 62.8627 | 0.004043 | 1.0507 | 0.025542 | 24.3097 | -0.001684 | 0.13026 | -0.0002 | -0.015411 | -0.015411 | -0.34908
2019-03-26 16:09 | 1553630924.6835 | -255 | 0.20954 | 0.00044 | 0 | 100.311 | 100.4278 | 128.5801 | 21.2612 | 0.001836 | 0.001722 | -152.7244 | 0.21898 | 8.4047 | 0.003481 | 0.004512 | 0.004295 | 5.1e-05 | 0.0012 | 0.30934 | 1.3314 | 2.5475 | 100.79 | 21.1 | 16.261 | 66.012 | 105.6243 | 63.2577 | 0.004042 | 1.0495 | 0.025533 | 24.3299 | -0.010405 | 0.13027 | -0.000239 | -0.018437 | -0.018437 | -0.29181
2019-03-26 16:09 | 1553630925.6829 | -254 | 0.2095 | 0.000439 | 0 | 100.3414 | 100.4264 | 128.8116 | 21.2612 | 0.001773 | 0.001725 | -152.713 | 0.21957 | 8.4028 | 0.003489 | 0.00465 | 0.004513 | 5.2e-05 | 0.001201 | 0.35906 | 1.3375 | 2.5475 | 100.79 | 21.1 | 16.253 | 64.4791 | 105.8217 | 63.3321 | 0.00404 | 1.0492 | 0.025553 | 24.3544 | -0.007574 | 0.13062 | -8e-06 | -0.00056 | -0.00056 | -9.628
2019-03-26 16:09 | 1553630926.6826 | -253 | 0.20954 | 0.000437 | 0 | 100.3757 | 100.3054 | 128.9965 | 21.2612 | 0.00185 | 0.001699 | -152.6955 | 0.21957 | 8.4037 | 0.003467 | 0.00491 | 0.004613 | 5.3e-05 | 0.001206 | 0.33179 | 1.3425 | 2.5475 | 100.79 | 21.1 | 16.251 | 66.342 | 105.3018 | 62.7033 | 0.004039 | 1.0512 | 0.025596 | 24.3491 | 0.006262 | 0.13062 | -0.000225 | -0.017391 | -0.017391 | -0.31021
2019-03-26 16:09 | 1553630927.6826 | -252 | 0.20952 | 0.000437 | -1e-06 | 100.4406 | 100.4106 | 128.5301 | 21.2612 | 0.001823 | 0.00175 | -152.7059 | 0.21996 | 8.4023 | 0.003519 | 0.004834 | 0.004689 | 5.1e-05 | 0.001207 | 0.35994 | 1.3568 | 2.5475 | 100.79 | 21.1 | 16.248 | 65.6991 | 106.5506 | 63.9348 | 0.004039 | 1.0473 | 0.025503 | 24.3496 | 0.00267 | 0.13086 | -9.2e-05 | -0.007063 | -0.007063 | -0.7652
2019-03-26 16:09 | 1553630928.6827 | -251 | 0.20949 | 0.00044 | -1e-06 | 100.3532 | 100.4735 | 128.2607 | 21.2612 | 0.001827 | 0.001714 | -152.7123 | 0.21988 | 8.3996 | 0.003526 | 0.005048 | 0.004784 | 5e-05 | 0.001201 | 0.3063 | 1.395 | 2.5475 | 100.79 | 21.1 | 16.246 | 65.7982 | 106.7278 | 63.0572 | 0.004038 | 1.0501 | 0.025509 | 24.2924 | -0.010712 | 0.13081 | 6.9e-05 | 0.005368 | 0.005368 | 1.0064
2019-03-26 16:09 | 1553630929.6827 | -250 | 0.20949 | 0.000442 | -1e-06 | 100.3447 | 100.369 | 128.1144 | 21.2612 | 0.001831 | 0.001701 | -152.7422 | 0.21923 | 8.4005 | 0.00343 | 0.005215 | 0.005094 | 5e-05 | 0.001199 | 0.34844 | 1.4471 | 2.5476 | 100.79 | 21.1 | 16.246 | 65.8761 | 104.3879 | 62.7479 | 0.004038 | 1.0511 | 0.025506 | 24.2673 | -0.002164 | 0.13042 | 8.2e-05 | 0.006389 | 0.006389 | 0.84307
2019-03-26 16:09 | 1553630930.6838 | -249 | 0.20951 | 0.000444 | -1e-06 | 100.3559 | 100.4379 | 128.0856 | 21.2612 | 0.001875 | 0.001697 | -152.7363 | 0.22026 | 8.4039 | 0.003429 | 0.005002 | 0.004769 | 5.4e-05 | 0.001208 | 0.33019 | 1.3921 | 2.5476 | 100.79 | 21.11 | 16.249 | 66.9521 | 104.3617 | 62.6633 | 0.004041 | 1.0513 | 0.025507 | 24.2616 | -0.007303 | 0.13103 | -8e-05 | -0.006111 | -0.006111 | -0.88559
2019-03-26 16:09 | 1553630931.6828 | -248 | 0.20949 | 0.000438 | -1e-06 | 100.4084 | 100.3673 | 128.0222 | 21.2612 | 0.001826 | 0.001707 | -152.7324 | 0.21977 | 8.4019 | 0.003442 | 0.004733 | 0.004523 | 5.2e-05 | 0.001209 | 0.29463 | 1.3401 | 2.5475 | 100.79 | 21.11 | 16.252 | 65.753 | 104.6826 | 62.8978 | 0.004042 | 1.0506 | 0.025491 | 24.264 | 0.003668 | 0.13074 | 9.8e-05 | 0.007589 | 0.007589 | 0.71151
2019-03-26 16:09 | 1553630932.684 | -247 | 0.20949 | 0.00044 | -1e-06 | 100.3887 | 100.3115 | 127.9853 | 21.2612 | 0.001822 | 0.001696 | -152.7486 | 0.2201 | 8.4043 | 0.003514 | 0.004458 | 0.004274 | 5.4e-05 | 0.001211 | 0.3298 | 1.352 | 2.5475 | 100.79 | 21.11 | 16.256 | 65.6609 | 106.4301 | 62.6243 | 0.004043 | 1.0514 | 0.025498 | 24.2507 | 0.00688 | 0.13094 | 6.8e-05 | 0.005247 | 0.005247 | 1.0306
2019-03-26 16:09 | 1553630933.684 | -246 | 0.20952 | 0.000442 | -1e-06 | 100.3267 | 100.2419 | 127.791 | 21.2612 | 0.001802 | 0.00172 | -152.7567 | 0.22032 | 8.4034 | 0.003522 | 0.004915 | 0.004735 | 5.1e-05 | 0.00121 | 0.33035 | 1.3464 | 2.5474 | 100.79 | 21.11 | 16.264 | 65.1831 | 106.6278 | 63.2132 | 0.004045 | 1.0496 | 0.025456 | 24.2535 | 0.007562 | 0.13107 | -0.00011 | -0.008446 | -0.008446 | -0.64093
2019-03-26 16:09 | 1553630934.684 | -245 | 0.20952 | 0.000442 | -1e-06 | 100.3391 | 100.4214 | 127.7976 | 21.2612 | 0.001752 | 0.001706 | -152.7663 | 0.22016 | 8.4012 | 0.003441 | 0.004886 | 0.004712 | 4.8e-05 | 0.001218 | 0.31116 | 1.3595 | 2.5474 | 100.79 | 21.11 | 16.268 | 63.9848 | 104.6637 | 62.8742 | 0.004046 | 1.0507 | 0.02547 | 24.2419 | -0.007336 | 0.13097 | -0.000138 | -0.010558 | -0.010558 | -0.51231
2019-03-26 16:09 | 1553630935.683 | -244 | 0.20951 | 0.00044 | -1e-06 | 100.3981 | 100.4762 | 128.1525 | 21.2612 | 0.001819 | 0.001694 | -152.7714 | 0.21933 | 8.402 | 0.003543 | 0.004241 | 0.004058 | 5e-05 | 0.001221 | 0.30525 | 1.3888 | 2.5476 | 100.79 | 21.11 | 16.277 | 65.6014 | 107.1386 | 62.5815 | 0.004048 | 1.0516 | 0.025516 | 24.2649 | -0.006954 | 0.13048 | -3.2e-05 | -0.002394 | -0.002394 | -2.2507
2019-03-26 16:09 | 1553630936.6845 | -243 | 0.20948 | 0.000442 | -1e-06 | 100.34 | 100.3841 | 128.2154 | 21.2612 | 0.001841 | 0.001728 | -152.79 | 0.21952 | 8.4029 | 0.00354 | 0.0052 | 0.005015 | 4.8e-05 | 0.001224 | 0.31813 | 1.3684 | 2.5477 | 100.79 | 21.11 | 16.283 | 66.1299 | 107.0653 | 63.3961 | 0.00405 | 1.049 | 0.025492 | 24.3003 | -0.003933 | 0.13059 | 0.000106 | 0.008251 | 0.008251 | 0.6537
2019-03-26 16:09 | 1553630937.6834 | -242 | 0.20949 | 0.000441 | -1e-06 | 100.3757 | 100.3921 | 128.5806 | 21.2612 | 0.001793 | 0.001682 | -152.7854 | 0.21971 | 8.4026 | 0.003453 | 0.004954 | 0.004754 | 4.8e-05 | 0.001216 | 0.3095 | 1.3873 | 2.5475 | 100.79 | 21.11 | 16.287 | 64.9599 | 104.9405 | 62.2923 | 0.004051 | 1.0525 | 0.02557 | 24.295 | -0.001468 | 0.1307 | 9.2e-05 | 0.007157 | 0.007157 | 0.75428
2019-03-26 16:09 | 1553630938.6822 | -241 | 0.20952 | 0.000437 | 0 | 100.4442 | 100.4893 | 128.633 | 21.2612 | 0.001849 | 0.001747 | -152.792 | 0.22045 | 8.4034 | 0.003509 | 0.004824 | 0.004576 | 4.9e-05 | 0.001208 | 0.32875 | 1.3841 | 2.5477 | 100.79 | 21.11 | 16.296 | 66.3124 | 106.2971 | 63.8638 | 0.004053 | 1.0476 | 0.025515 | 24.3568 | -0.004019 | 0.13115 | -8e-05 | -0.006162 | -0.006162 | -0.87901
2019-03-26 16:09 | 1553630939.6825 | -240 | 0.20951 | 0.000437 | -1e-06 | 100.4524 | 100.497 | 128.4447 | 21.2612 | 0.001839 | 0.001723 | -152.7833 | 0.22022 | 8.4016 | 0.00347 | 0.004582 | 0.004432 | 4.9e-05 | 0.001205 | 0.3314 | 1.3641 | 2.5476 | 100.79 | 21.11 | 16.301 | 66.0854 | 105.3544 | 63.2859 | 0.004054 | 1.0494 | 0.025519 | 24.3181 | -0.003971 | 0.13101 | -5.6e-05 | -0.004238 | -0.004238 | -1.2766
2019-03-26 16:09 | 1553630940.6832 | -239 | 0.20953 | 0.000438 | 0 | 100.4247 | 100.4594 | 128.7055 | 21.2612 | 0.001827 | 0.00175 | -152.7792 | 0.21962 | 8.4055 | 0.003385 | 0.004395 | 0.00424 | 4.8e-05 | 0.001197 | 0.34667 | 1.3569 | 2.5474 | 100.79 | 21.11 | 16.304 | 65.7896 | 103.2974 | 63.9247 | 0.004055 | 1.0474 | 0.02552 | 24.3658 | -0.003091 | 0.13065 | -0.000184 | -0.014148 | -0.014148 | -0.38138
2019-03-26 16:09 | 1553630941.6833 | -238 | 0.2095 | 0.00044 | -1e-06 | 100.4164 | 100.2318 | 129.0451 | 21.2612 | 0.001776 | 0.001697 | -152.7949 | 0.22015 | 8.4062 | 0.003441 | 0.004877 | 0.004728 | 4.7e-05 | 0.001199 | 0.34324 | 1.3719 | 2.5476 | 100.79 | 21.11 | 16.308 | 64.5602 | 104.6489 | 62.6595 | 0.004056 | 1.0513 | 0.025602 | 24.3521 | 0.016459 | 0.13097 | -1e-05 | -0.000665 | -0.000665 | -8.133
2019-03-26 16:09 | 1553630942.6834 | -237 | 0.20949 | 0.00044 | -1e-06 | 100.4433 | 100.4759 | 129.3369 | 21.2612 | 0.001807 | 0.001645 | -152.813 | 0.2197 | 8.4063 | 0.003423 | 0.005072 | 0.004816 | 5.1e-05 | 0.001201 | 0.34451 | 1.3662 | 2.5477 | 100.79 | 21.11 | 16.307 | 65.301 | 104.2321 | 61.4126 | 0.004056 | 1.0552 | 0.025679 | 24.3343 | -0.002901 | 0.1307 | 4.8e-05 | 0.003785 | 0.003785 | 1.426
2019-03-26 16:09 | 1553630943.6825 | -236 | 0.20952 | 0.000437 | -1e-06 | 100.34 | 100.4439 | 129.3064 | 21.2612 | 0.001821 | 0.001733 | -152.8212 | 0.21928 | 8.4018 | 0.003469 | 0.00525 | 0.005063 | 5e-05 | 0.001204 | 0.28926 | 1.3664 | 2.5475 | 100.79 | 21.11 | 16.305 | 65.6515 | 105.3379 | 63.5319 | 0.004055 | 1.0486 | 0.025595 | 24.4084 | -0.009263 | 0.13045 | -0.0001 | -0.007636 | -0.007636 | -0.70556
2019-03-26 16:09 | 1553630944.6838 | -235 | 0.20952 | 0.000433 | 0 | 100.4199 | 100.3881 | 129.0017 | 21.2612 | 0.001766 | 0.001728 | -152.799 | 0.21894 | 8.4016 | 0.003458 | 0.004741 | 0.004578 | 4.8e-05 | 0.001212 | 0.34595 | 1.3669 | 2.5475 | 100.79 | 21.11 | 16.303 | 64.323 | 105.0664 | 63.4055 | 0.004055 | 1.049 | 0.025569 | 24.3751 | 0.002836 | 0.13025 | -0.000118 | -0.009065 | -0.009065 | -0.59342
2019-03-26 16:09 | 1553630945.6834 | -234 | 0.20948 | 0.000438 | -1e-06 | 100.4424 | 100.4291 | 128.713 | 21.2612 | 0.001825 | 0.001725 | -152.7854 | 0.22061 | 8.4021 | 0.003455 | 0.005395 | 0.005205 | 4.5e-05 | 0.001207 | 0.33069 | 1.3707 | 2.5476 | 100.79 | 21.11 | 16.301 | 65.7513 | 104.9917 | 63.3227 | 0.004054 | 1.0493 | 0.025544 | 24.3448 | 0.001181 | 0.13124 | 0.000127 | 0.009829 | 0.009829 | 0.55144
2019-03-26 16:09 | 1553630946.6836 | -233 | 0.20953 | 0.00044 | 0 | 100.4208 | 100.4584 | 128.4963 | 21.2612 | 0.001874 | 0.001679 | -152.7898 | 0.21977 | 8.401 | 0.003423 | 0.004743 | 0.004586 | 5e-05 | 0.001203 | 0.32444 | 1.3028 | 2.5474 | 100.79 | 21.11 | 16.295 | 66.9079 | 104.218 | 62.2181 | 0.004053 | 1.0527 | 0.025564 | 24.2843 | -0.003343 | 0.13074 | -0.000162 | -0.012514 | -0.012514 | -0.43148
2019-03-26 16:09 | 1553630947.6836 | -232 | 0.20952 | 0.000439 | 0 | 100.4155 | 100.38 | 128.2724 | 21.2612 | 0.001852 | 0.001713 | -152.7944 | 0.21999 | 8.4036 | 0.003448 | 0.004249 | 0.004052 | 4.6e-05 | 0.001206 | 0.3011 | 1.3375 | 2.5476 | 100.79 | 21.11 | 16.291 | 66.3834 | 104.8227 | 63.0478 | 0.004052 | 1.0501 | 0.025511 | 24.2931 | 0.003161 | 0.13087 | -9.7e-05 | -0.007518 | -0.007518 | -0.71893
2019-03-26 16:09 | 1553630948.6827 | -231 | 0.20952 | 0.000437 | 0 | 100.4161 | 100.378 | 128.0804 | 21.2612 | 0.001767 | 0.001692 | -152.8049 | 0.21966 | 8.4021 | 0.003318 | 0.005069 | 0.004826 | 4.8e-05 | 0.001206 | 0.32516 | 1.3134 | 2.5476 | 100.79 | 21.11 | 16.28 | 64.3564 | 101.6919 | 62.5369 | 0.004049 | 1.0517 | 0.025511 | 24.2565 | 0.003393 | 0.13067 | -0.000129 | -0.009935 | -0.009935 | -0.54323
2019-03-26 16:09 | 1553630949.6838 | -230 | 0.2095 | 0.000441 | 0 | 100.3866 | 100.3105 | 128.4034 | 21.2612 | 0.001806 | 0.001666 | -152.8078 | 0.21957 | 8.4005 | 0.003396 | 0.004675 | 0.004501 | 4.8e-05 | 0.00121 | 0.3366 | 1.3481 | 2.5474 | 100.79 | 21.11 | 16.274 | 65.2769 | 103.5789 | 61.9142 | 0.004048 | 1.0537 | 0.025567 | 24.2645 | 0.006786 | 0.13062 | 1e-06 | 5.2e-05 | 5.2e-05 | 103.7712
2019-03-26 16:09 | 1553630950.6828 | -229 | 0.20946 | 0.000438 | 0 | 100.3907 | 100.4358 | 128.3095 | 21.2612 | 0.001794 | 0.001724 | -152.8181 | 0.21929 | 8.4017 | 0.003367 | 0.004824 | 0.00458 | 4.8e-05 | 0.001205 | 0.33019 | 1.3472 | 2.5476 | 100.79 | 21.11 | 16.269 | 64.9967 | 102.8598 | 63.3172 | 0.004046 | 1.0493 | 0.025504 | 24.3064 | -0.004019 | 0.13046 | 0.000268 | 0.020614 | 0.020614 | 0.26137
2019-03-26 16:09 | 1553630951.6838 | -228 | 0.20949 | 0.000441 | 0 | 100.3957 | 100.2856 | 128.2978 | 21.2612 | 0.001796 | 0.001684 | -152.8024 | 0.21951 | 8.4014 | 0.003417 | 0.004359 | 0.004214 | 4.9e-05 | 0.001209 | 0.3131 | 1.3536 | 2.5476 | 100.79 | 21.12 | 16.255 | 65.0489 | 104.0741 | 62.3479 | 0.004045 | 1.0523 | 0.02554 | 24.2702 | 0.00982 | 0.13059 | 5e-05 | 0.003842 | 0.003842 | 1.4039
2019-03-26 16:09 | 1553630952.6828 | -227 | 0.20951 | 0.000437 | 0 | 100.3783 | 100.3972 | 128.1508 | 21.2612 | 0.001822 | 0.001708 | -152.7745 | 0.21994 | 8.4043 | 0.003384 | 0.004801 | 0.004574 | 5.2e-05 | 0.001215 | 0.29015 | 1.3469 | 2.5475 | 100.79 | 21.12 | 16.249 | 65.6772 | 103.2834 | 62.9267 | 0.004044 | 1.0505 | 0.025503 | 24.2772 | -0.001681 | 0.13084 | -5.1e-05 | -0.003925 | -0.003925 | -1.3768
2019-03-26 16:09 | 1553630953.6829 | -226 | 0.20951 | 0.000441 | 0 | 100.3919 | 100.3837 | 128.0464 | 21.2612 | 0.001872 | 0.001712 | -152.7495 | 0.21998 | 8.4075 | 0.003357 | 0.004469 | 0.00431 | 5e-05 | 0.001215 | 0.32117 | 1.3496 | 2.5474 | 100.79 | 21.12 | 16.242 | 66.861 | 102.6154 | 63.0228 | 0.004042 | 1.0502 | 0.025489 | 24.2708 | 0.000727 | 0.13087 | -4.1e-05 | -0.003171 | -0.003171 | -1.7043
2019-03-26 16:09 | 1553630954.7243 | -225 | 0.20952 | 0.000439 | 0 | 100.4217 | 100.3538 | 128.1952 | 21.2612 | 0.00187 | 0.001768 | -152.7738 | 0.21951 | 8.4031 | 0.003256 | 0.004988 | 0.004789 | 5.2e-05 | 0.001215 | 0.32388 | 1.3644 | 2.5475 | 100.79 | 21.12 | 16.232 | 66.829 | 100.1714 | 64.3807 | 0.00404 | 1.046 | 0.025452 | 24.3339 | 0.006049 | 0.13059 | -0.000119 | -0.009116 | -0.009116 | -0.59164
2019-03-26 16:09 | 1553630955.6689 | -224 | 0.20947 | 0.000441 | 0 | 100.3733 | 100.4795 | 128.2485 | 21.2612 | 0.001868 | 0.001726 | -152.7823 | 0.21942 | 8.4011 | 0.003351 | 0.004541 | 0.00436 | 4.8e-05 | 0.001216 | 0.31437 | 1.3729 | 2.5474 | 100.79 | 21.12 | 16.228 | 66.7682 | 102.4883 | 63.3461 | 0.004039 | 1.0492 | 0.025497 | 24.3016 | -0.009462 | 0.13054 | 0.000176 | 0.01351 | 0.01351 | 0.39904
2019-03-26 16:09 | 1553630956.6683 | -223 | 0.2095 | 0.000438 | 0 | 100.3828 | 100.4927 | 128.3459 | 21.2612 | 0.001865 | 0.001708 | -152.7869 | 0.21972 | 8.4032 | 0.003359 | 0.005121 | 0.004916 | 4.9e-05 | 0.001213 | 0.323 | 1.3563 | 2.5475 | 100.79 | 21.12 | 16.226 | 66.6941 | 102.6767 | 62.9313 | 0.004038 | 1.0505 | 0.025522 | 24.2959 | -0.009788 | 0.13071 | 1.6e-05 | 0.001226 | 0.001226 | 4.4029
2019-03-26 16:09 | 1553630957.6681 | -222 | 0.20953 | 0.000437 | 0 | 100.4034 | 100.2103 | 128.4623 | 21.2612 | 0.001833 | 0.001708 | -152.7892 | 0.21964 | 8.4029 | 0.003437 | 0.004628 | 0.004488 | 4.8e-05 | 0.001218 | 0.3465 | 1.3313 | 2.5475 | 100.79 | 21.12 | 16.227 | 65.9401 | 104.575 | 62.9267 | 0.004038 | 1.0505 | 0.025534 | 24.3067 | 0.017222 | 0.13067 | -0.000141 | -0.010869 | -0.010869 | -0.4965
2019-03-26 16:09 | 1553630958.6692 | -221 | 0.20949 | 0.00044 | 0 | 100.3314 | 100.1972 | 128.7048 | 21.2612 | 0.001798 | 0.001779 | -152.769 | 0.21859 | 8.402 | 0.003317 | 0.004398 | 0.0042 | 4.5e-05 | 0.00122 | 0.31222 | 1.364 | 2.5475 | 100.79 | 21.11 | 16.23 | 65.0996 | 101.6675 | 64.6312 | 0.004037 | 1.0452 | 0.025494 | 24.3913 | 0.011974 | 0.13004 | 9.6e-05 | 0.007393 | 0.007393 | 0.72647
2019-03-26 16:09 | 1553630959.6681 | -220 | 0.20956 | 0.00044 | 0 | 100.3518 | 100.4241 | 128.7203 | 21.2612 | 0.001851 | 0.001698 | -152.7628 | 0.21886 | 8.403 | 0.003364 | 0.004739 | 0.004621 | 4.5e-05 | 0.001219 | 0.31636 | 1.3787 | 2.5476 | 100.79 | 21.11 | 16.235 | 66.3702 | 102.79 | 62.6776 | 0.004038 | 1.0513 | 0.025569 | 24.3221 | -0.006444 | 0.1302 | -0.000355 | -0.027387 | -0.027387 | -0.19634
2019-03-26 16:09 | 1553630960.6682 | -219 | 0.20952 | 0.000442 | 0 | 100.3863 | 100.5007 | 128.7846 | 21.2612 | 0.001867 | 0.001717 | -152.7666 | 0.21965 | 8.4027 | 0.003395 | 0.00416 | 0.00405 | 4.9e-05 | 0.001218 | 0.34954 | 1.3546 | 2.5475 | 100.79 | 21.11 | 16.238 | 66.7581 | 103.5428 | 63.1516 | 0.004039 | 1.0498 | 0.025557 | 24.3453 | -0.01019 | 0.13067 | -0.000109 | -0.008373 | -0.008373 | -0.64454
2019-03-26 16:09 | 1553630961.6692 | -218 | 0.20955 | 0.000439 | 0 | 100.4202 | 100.3659 | 128.9038 | 21.2612 | 0.001806 | 0.001735 | -152.7645 | 0.22035 | 8.402 | 0.003428 | 0.004399 | 0.004159 | 4.7e-05 | 0.001214 | 0.35601 | 1.3551 | 2.5475 | 100.79 | 21.12 | 16.244 | 65.2892 | 104.3547 | 63.5851 | 0.004043 | 1.0484 | 0.025553 | 24.3723 | 0.004839 | 0.13109 | -0.000292 | -0.022546 | -0.022546 | -0.24013
2019-03-26 16:09 | 1553630962.6693 | -217 | 0.20947 | 0.00044 | 0 | 100.3485 | 100.3199 | 128.8083 | 21.2612 | 0.001808 | 0.001717 | -152.7771 | 0.21937 | 8.4021 | 0.003381 | 0.004693 | 0.004502 | 4.8e-05 | 0.001214 | 0.31852 | 1.3718 | 2.5476 | 100.79 | 21.12 | 16.248 | 65.3323 | 103.2032 | 63.1509 | 0.004044 | 1.0498 | 0.02556 | 24.3476 | 0.002553 | 0.1305 | 0.000209 | 0.016113 | 0.016113 | 0.33451
2019-03-26 16:09 | 1553630963.6688 | -216 | 0.2095 | 0.000443 | 0 | 100.3934 | 100.374 | 128.8665 | 21.2612 | 0.001822 | 0.001693 | -152.7829 | 0.22018 | 8.4015 | 0.003318 | 0.004892 | 0.004646 | 4.7e-05 | 0.001217 | 0.30198 | 1.3649 | 2.5476 | 100.79 | 21.12 | 16.252 | 65.6678 | 101.6699 | 62.5697 | 0.004045 | 1.0516 | 0.025588 | 24.332 | 0.001728 | 0.13098 | 2.7e-05 | 0.002115 | 0.002115 | 2.5578
2019-03-26 16:09 | 1553630964.67 | -215 | 0.20948 | 0.000437 | 0 | 100.3267 | 100.4288 | 128.8879 | 21.2612 | 0.001794 | 0.00166 | -152.7614 | 0.21888 | 8.402 | 0.003231 | 0.00441 | 0.004251 | 4.7e-05 | 0.00122 | 0.34202 | 1.3868 | 2.5474 | 100.79 | 21.12 | 16.255 | 65.0036 | 99.5688 | 61.7806 | 0.004045 | 1.0541 | 0.02562 | 24.3054 | -0.0091 | 0.13021 | 0.000151 | 0.011693 | 0.011693 | 0.45992
2019-03-26 16:09 | 1553630965.6685 | -214 | 0.2095 | 0.000434 | 0 | 100.3798 | 100.5165 | 128.5245 | 21.2612 | 0.001801 | 0.00172 | -152.7623 | 0.21944 | 8.4031 | 0.003286 | 0.004097 | 0.003931 | 4.9e-05 | 0.001217 | 0.2792 | 1.3674 | 2.5475 | 100.79 | 21.12 | 16.263 | 65.1629 | 100.8958 | 63.218 | 0.004047 | 1.0496 | 0.025529 | 24.3232 | -0.012175 | 0.13055 | 3.9e-05 | 0.002976 | 0.002976 | 1.8118
2019-03-26 16:09 | 1553630966.6688 | -213 | 0.20947 | 0.000439 | 0 | 100.3698 | 100.2718 | 128.6379 | 21.2612 | 0.001834 | 0.001699 | -152.7677 | 0.22087 | 8.404 | 0.003368 | 0.004422 | 0.004207 | 5e-05 | 0.001214 | 0.31974 | 1.3932 | 2.5474 | 100.79 | 21.12 | 16.266 | 65.962 | 102.8937 | 62.6971 | 0.004048 | 1.0512 | 0.02556 | 24.315 | 0.008735 | 0.1314 | 0.000182 | 0.014051 | 0.014051 | 0.38623
2019-03-26 16:09 | 1553630967.6699 | -212 | 0.2095 | 0.000442 | 0 | 100.3825 | 100.2657 | 128.6038 | 21.2612 | 0.001856 | 0.00173 | -152.7692 | 0.21917 | 8.4043 | 0.003324 | 0.004712 | 0.004539 | 4.7e-05 | 0.001217 | 0.34235 | 1.3383 | 2.5475 | 100.79 | 21.12 | 16.269 | 66.4786 | 101.8231 | 63.4484 | 0.004049 | 1.0489 | 0.025528 | 24.339 | 0.010406 | 0.13038 | 2.8e-05 | 0.002149 | 0.002149 | 2.506
2019-03-26 16:09 | 1553630968.6685 | -211 | 0.20953 | 0.000438 | 0 | 100.3335 | 100.3145 | 128.4909 | 21.2612 | 0.001809 | 0.001706 | -152.7847 | 0.22013 | 8.4031 | 0.003363 | 0.004544 | 0.004408 | 4.9e-05 | 0.001222 | 0.33356 | 1.3271 | 2.5474 | 100.79 | 21.12 | 16.271 | 65.3525 | 102.7774 | 62.8767 | 0.004049 | 1.0506 | 0.025539 | 24.3076 | 0.001691 | 0.13095 | -0.000138 | -0.01064 | -0.01064 | -0.50831
2019-03-26 16:09 | 1553630969.668 | -210 | 0.2095 | 0.000437 | 0 | 100.332 | 100.2419 | 128.2811 | 21.2612 | 0.001856 | 0.001734 | -152.7822 | 0.21994 | 8.4018 | 0.003386 | 0.004731 | 0.004577 | 4.8e-05 | 0.001227 | 0.32532 | 1.3767 | 2.5475 | 100.79 | 21.12 | 16.27 | 66.4762 | 103.3313 | 63.5608 | 0.004049 | 1.0485 | 0.025492 | 24.3125 | 0.008036 | 0.13084 | 3.2e-05 | 0.002458 | 0.002458 | 2.1984
2019-03-26 16:09 | 1553630970.6676 | -209 | 0.20955 | 0.000439 | 0 | 100.3281 | 100.371 | 128.2022 | 21.2612 | 0.001791 | 0.0017 | -152.7727 | 0.21967 | 8.4023 | 0.003411 | 0.004717 | 0.004569 | 4.8e-05 | 0.001233 | 0.32693 | 1.3818 | 2.5477 | 100.79 | 21.12 | 16.268 | 64.9153 | 103.9389 | 62.7298 | 0.004049 | 1.0511 | 0.025516 | 24.275 | -0.003817 | 0.13068 | -0.000309 | -0.023832 | -0.023832 | -0.22646
2019-03-26 16:10 | 1553630971.6691 | -208 | 0.20955 | 0.000437 | 0 | 100.3155 | 100.3377 | 128.0682 | 21.2612 | 0.001836 | 0.001721 | -152.8007 | 0.2196 | 8.4053 | 0.003324 | 0.004277 | 0.004069 | 5e-05 | 0.001223 | 0.33539 | 1.3153 | 2.5475 | 100.79 | 21.12 | 16.265 | 66.0027 | 101.8316 | 63.2375 | 0.004048 | 1.0495 | 0.025483 | 24.2806 | -0.001982 | 0.13064 | -0.000273 | -0.020987 | -0.020987 | -0.25709
2019-03-26 16:10 | 1553630972.6692 | -207 | 0.20952 | 0.000444 | 0 | 100.314 | 100.2281 | 127.7833 | 21.2612 | 0.001807 | 0.001695 | -152.8077 | 0.22018 | 8.404 | 0.003377 | 0.004595 | 0.004432 | 4.9e-05 | 0.00123 | 0.34468 | 1.3753 | 2.5476 | 100.79 | 21.12 | 16.264 | 65.3041 | 103.1041 | 62.6073 | 0.004048 | 1.0515 | 0.025479 | 24.2309 | 0.007661 | 0.13099 | -0.000127 | -0.009805 | -0.009805 | -0.55176
2019-03-26 16:10 | 1553630973.6688 | -206 | 0.20955 | 0.000439 | 0 | 100.2505 | 100.3323 | 127.8943 | 21.2612 | 0.00186 | 0.001734 | -152.8175 | 0.21982 | 8.4041 | 0.003436 | 0.004514 | 0.00431 | 5.4e-05 | 0.001232 | 0.33445 | 1.3513 | 2.5476 | 100.79 | 21.12 | 16.263 | 66.5871 | 104.5277 | 63.5615 | 0.004047 | 1.0485 | 0.025453 | 24.2758 | -0.007295 | 0.13077 | -0.000278 | -0.021323 | -0.021323 | -0.25329
2019-03-26 16:10 | 1553630974.6692 | -205 | 0.20951 | 0.000439 | 0 | 100.2045 | 100.1051 | 128.159 | 21.2612 | 0.001813 | 0.001706 | -152.8116 | 0.21982 | 8.4049 | 0.003314 | 0.004182 | 0.003974 | 5.1e-05 | 0.001241 | 0.34219 | 1.3495 | 2.5475 | 100.79 | 21.12 | 16.262 | 65.4533 | 101.5765 | 62.8868 | 0.004047 | 1.0506 | 0.025505 | 24.2766 | 0.008877 | 0.13077 | -1.6e-05 | -0.001254 | -0.001254 | -4.3071
2019-03-26 16:10 | 1553630975.6689 | -204 | 0.20951 | 0.000436 | 0 | 100.2213 | 100.2614 | 128.6449 | 21.2612 | 0.00183 | 0.001738 | -152.7822 | 0.21909 | 8.4024 | 0.003328 | 0.004666 | 0.004505 | 4.9e-05 | 0.001248 | 0.34451 | 1.384 | 2.5475 | 100.79 | 21.12 | 16.261 | 65.8646 | 101.9344 | 63.6467 | 0.004047 | 1.0482 | 0.025525 | 24.35 | -0.003577 | 0.13034 | -2.3e-05 | -0.001733 | -0.001733 | -3.1065
2019-03-26 16:10 | 1553630976.6683 | -203 | 0.20951 | 0.000443 | 0 | 100.2417 | 100.1619 | 128.49 | 21.2612 | 0.001772 | 0.001727 | -152.7669 | 0.22016 | 8.4003 | 0.003357 | 0.004973 | 0.004738 | 5.1e-05 | 0.001252 | 0.32372 | 1.3534 | 2.5476 | 100.79 | 21.12 | 16.258 | 64.469 | 102.6154 | 63.3868 | 0.004046 | 1.0491 | 0.025519 | 24.326 | 0.007122 | 0.13098 | -3.7e-05 | -0.002886 | -0.002886 | -1.8744
2019-03-26 16:10 | 1553630977.668 | -202 | 0.20948 | 0.000444 | 0 | 100.2275 | 100.3027 | 128.4156 | 21.2612 | 0.001841 | 0.001716 | -152.7452 | 0.21923 | 8.4 | 0.003325 | 0.004901 | 0.004772 | 4.8e-05 | 0.001255 | 0.32278 | 1.4016 | 2.5474 | 100.79 | 21.12 | 16.255 | 66.1268 | 101.8512 | 63.1172 | 0.004045 | 1.0499 | 0.025522 | 24.3092 | -0.006712 | 0.13042 | 0.000133 | 0.010255 | 0.010255 | 0.52527
2019-03-26 16:10 | 1553630978.6681 | -201 | 0.2095 | 0.000444 | 0 | 100.2653 | 100.2614 | 128.263 | 21.2612 | 0.001861 | 0.001718 | -152.7284 | 0.22085 | 8.4027 | 0.003279 | 0.004678 | 0.004481 | 4.9e-05 | 0.001256 | 0.33069 | 1.389 | 2.5475 | 100.79 | 21.11 | 16.246 | 66.5958 | 100.7287 | 63.1711 | 0.004041 | 1.0497 | 0.025505 | 24.2967 | 0.000347 | 0.13138 | 3.6e-05 | 0.002779 | 0.002779 | 1.9527
2019-03-26 16:10 | 1553630979.669 | -200 | 0.20948 | 0.000444 | -1e-06 | 100.1555 | 100.254 | 128.5362 | 21.2612 | 0.001805 | 0.001761 | -152.7171 | 0.21989 | 8.4031 | 0.003404 | 0.004209 | 0.003987 | 5e-05 | 0.001256 | 0.32853 | 1.406 | 2.5476 | 100.79 | 21.11 | 16.241 | 65.2502 | 103.7722 | 64.1964 | 0.004039 | 1.0465 | 0.025493 | 24.3596 | -0.008794 | 0.13081 | 0.000145 | 0.011185 | 0.011185 | 0.48302
2019-03-26 16:10 | 1553630980.6691 | -199 | 0.20952 | 0.000441 | 0 | 100.1983 | 100.2832 | 128.5823 | 21.2612 | 0.001825 | 0.001723 | -152.722 | 0.22066 | 8.402 | 0.003356 | 0.00495 | 0.004779 | 5.5e-05 | 0.001258 | 0.33268 | 1.3776 | 2.5475 | 100.79 | 21.11 | 16.236 | 65.7325 | 102.607 | 63.2977 | 0.004038 | 1.0493 | 0.025532 | 24.3315 | -0.007582 | 0.13127 | -8.5e-05 | -0.006495 | -0.006495 | -0.83475
2019-03-26 16:10 | 1553630981.6686 | -198 | 0.20948 | 0.000439 | 0 | 100.2257 | 100.0526 | 128.3794 | 21.2612 | 0.001859 | 0.001725 | -152.6935 | 0.21895 | 8.4035 | 0.003461 | 0.004733 | 0.004569 | 5e-05 | 0.001259 | 0.33555 | 1.3652 | 2.5475 | 100.79 | 21.11 | 16.223 | 66.5606 | 105.1413 | 63.3374 | 0.004035 | 1.0492 | 0.02551 | 24.3137 | 0.015461 | 0.13025 | 0.000159 | 0.012219 | 0.012219 | 0.44026
2019-03-26 16:10 | 1553630982.6682 | -197 | 0.20946 | 0.000439 | 0 | 100.1723 | 100.2187 | 128.7928 | 21.2612 | 0.001809 | 0.001725 | -152.674 | 0.21968 | 8.4036 | 0.003417 | 0.005201 | 0.004991 | 5e-05 | 0.001256 | 0.32637 | 1.3935 | 2.5475 | 100.79 | 21.11 | 16.218 | 65.3573 | 104.0874 | 63.3311 | 0.004034 | 1.0492 | 0.025551 | 24.3526 | -0.004142 | 0.13069 | 0.000262 | 0.020208 | 0.020208 | 0.2671
2019-03-26 16:10 | 1553630983.6683 | -196 | 0.20946 | 0.000442 | 0 | 100.2198 | 100.2315 | 129.1867 | 21.2612 | 0.001859 | 0.001736 | -152.6707 | 0.21972 | 8.4032 | 0.003416 | 0.00418 | 0.003977 | 5.2e-05 | 0.001249 | 0.32228 | 1.3773 | 2.5477 | 100.79 | 21.11 | 16.205 | 66.5575 | 104.0608 | 63.5998 | 0.00403 | 1.0484 | 0.02558 | 24.3996 | -0.001039 | 0.13071 | 0.000273 | 0.021072 | 0.021072 | 0.25618
2019-03-26 16:10 | 1553630984.6693 | -195 | 0.20949 | 0.000437 | 0 | 100.1853 | 100.1508 | 129.0078 | 21.2612 | 0.001853 | 0.001743 | -152.6705 | 0.22005 | 8.4043 | 0.003484 | 0.004948 | 0.004777 | 5e-05 | 0.001245 | 0.29789 | 1.3344 | 2.5475 | 100.79 | 21.11 | 16.198 | 66.4115 | 105.7007 | 63.7733 | 0.004029 | 1.0479 | 0.025556 | 24.3889 | 0.00308 | 0.13091 | 9.6e-05 | 0.007405 | 0.007405 | 0.73008
2019-03-26 16:10 | 1553630985.6684 | -194 | 0.2095 | 0.000439 | 0 | 100.1555 | 100.0835 | 128.7376 | 21.2612 | 0.001875 | 0.001724 | -152.6959 | 0.21974 | 8.4016 | 0.003421 | 0.004445 | 0.00426 | 5.4e-05 | 0.001241 | 0.33754 | 1.313 | 2.5474 | 100.79 | 21.11 | 16.191 | 66.9476 | 104.1778 | 63.3203 | 0.004027 | 1.0493 | 0.025546 | 24.347 | 0.006425 | 0.13072 | -1e-05 | -0.000784 | -0.000784 | -6.8878
2019-03-26 16:10 | 1553630986.6685 | -193 | 0.20947 | 0.000438 | 0 | 100.1203 | 100.1209 | 128.6815 | 21.2612 | 0.00183 | 0.001719 | -152.6848 | 0.21929 | 8.4035 | 0.003406 | 0.004754 | 0.004553 | 5.6e-05 | 0.001244 | 0.28888 | 1.3347 | 2.5476 | 100.79 | 21.11 | 16.177 | 65.8709 | 103.8156 | 63.1843 | 0.004023 | 1.0497 | 0.025546 | 24.3368 | -4.7e-05 | 0.13045 | 0.000183 | 0.014166 | 0.014166 | 0.38033
2019-03-26 16:10 | 1553630987.6686 | -192 | 0.20949 | 0.000439 | 0 | 100.1451 | 100.256 | 128.7214 | 21.2612 | 0.001861 | 0.001707 | -152.7203 | 0.22013 | 8.3988 | 0.00344 | 0.004853 | 0.00468 | 5.5e-05 | 0.001244 | 0.31813 | 1.3318 | 2.5475 | 100.79 | 21.11 | 16.171 | 66.5958 | 104.6283 | 62.9094 | 0.004022 | 1.0505 | 0.02556 | 24.3306 | -0.009897 | 0.13095 | 9.2e-05 | 0.007124 | 0.007124 | 0.75917
2019-03-26 16:10 | 1553630988.6679 | -191 | 0.20949 | 0.000434 | 0 | 100.1437 | 100.2009 | 128.8609 | 21.2612 | 0.001855 | 0.001772 | -152.7309 | 0.22085 | 8.4029 | 0.003515 | 0.005494 | 0.005325 | 5.5e-05 | 0.001246 | 0.323 | 1.3374 | 2.5476 | 100.79 | 21.11 | 16.165 | 66.4567 | 106.4427 | 64.4565 | 0.00402 | 1.0457 | 0.025516 | 24.3998 | -0.005109 | 0.13139 | 8.3e-05 | 0.006393 | 0.006393 | 0.8488
2019-03-26 16:10 | 1553630989.6676 | -190 | 0.20951 | 0.000436 | 0 | 100.1377 | 100.0133 | 128.6062 | 21.2612 | 0.001809 | 0.001759 | -152.7322 | 0.21937 | 8.4057 | 0.003469 | 0.005152 | 0.004939 | 5.6e-05 | 0.001249 | 0.34346 | 1.3144 | 2.5475 | 100.79 | 21.11 | 16.158 | 65.3479 | 105.3292 | 64.148 | 0.004019 | 1.0467 | 0.025502 | 24.3645 | 0.011124 | 0.1305 | -3e-05 | -0.002295 | -0.002295 | -2.3482
2019-03-26 16:10 | 1553630990.6676 | -189 | 0.20948 | 0.000441 | 0 | 100.1366 | 100.0973 | 128.8958 | 21.2612 | 0.001856 | 0.001694 | -152.7243 | 0.22037 | 8.4005 | 0.003448 | 0.004883 | 0.004715 | 5.6e-05 | 0.001244 | 0.34435 | 1.3353 | 2.5475 | 100.79 | 21.11 | 16.155 | 66.4755 | 104.8281 | 62.5815 | 0.004018 | 1.0516 | 0.02559 | 24.3352 | 0.003506 | 0.1311 | 0.000134 | 0.010376 | 0.010376 | 0.5218
2019-03-26 16:10 | 1553630991.6687 | -188 | 0.20948 | 0.000441 | 0 | 100.0826 | 100.0002 | 128.9369 | 21.2612 | 0.001829 | 0.001779 | -152.7478 | 0.21925 | 8.4059 | 0.003474 | 0.004787 | 0.004661 | 5.5e-05 | 0.001241 | 0.31492 | 1.3345 | 2.5474 | 100.79 | 21.11 | 16.15 | 65.8389 | 105.463 | 64.6374 | 0.004017 | 1.0452 | 0.025516 | 24.4135 | 0.007365 | 0.13043 | 0.000147 | 0.011322 | 0.011322 | 0.47578
2019-03-26 16:10 | 1553630992.6687 | -187 | 0.20947 | 0.000442 | 0 | 100.1649 | 100.2133 | 128.9745 | 21.2612 | 0.001817 | 0.001761 | -152.7486 | 0.21979 | 8.4038 | 0.00345 | 0.004824 | 0.004666 | 5.2e-05 | 0.001235 | 0.29966 | 1.3768 | 2.5475 | 100.79 | 21.11 | 16.148 | 65.5444 | 104.8792 | 64.1995 | 0.004016 | 1.0465 | 0.025537 | 24.4012 | -0.004321 | 0.13075 | 0.000187 | 0.014441 | 0.014441 | 0.37394
2019-03-26 16:10 | 1553630993.6679 | -186 | 0.20949 | 0.000443 | 0 | 100.211 | 100.1417 | 128.4032 | 21.2612 | 0.001865 | 0.00174 | -152.7385 | 0.22026 | 8.4029 | 0.003413 | 0.004367 | 0.004221 | 5.5e-05 | 0.001236 | 0.34938 | 1.3623 | 2.5476 | 100.79 | 21.11 | 16.145 | 66.7064 | 103.9711 | 63.6951 | 0.004015 | 1.0481 | 0.025499 | 24.3289 | 0.006183 | 0.13103 | 5e-05 | 0.003812 | 0.003812 | 1.4197
2019-03-26 16:10 | 1553630994.6688 | -185 | 0.20949 | 0.000439 | 0 | 100.1144 | 100.0247 | 128.2248 | 21.2612 | 0.001897 | 0.001718 | -152.7504 | 0.21909 | 8.4028 | 0.003358 | 0.005127 | 0.004924 | 5.7e-05 | 0.001236 | 0.30829 | 1.3955 | 2.5476 | 100.79 | 21.11 | 16.14 | 67.4756 | 102.6501 | 63.1617 | 0.004014 | 1.0498 | 0.025501 | 24.2927 | 0.008019 | 0.13033 | 0.000113 | 0.008754 | 0.008754 | 0.61489
2019-03-26 16:10 | 1553630995.6683 | -184 | 0.20948 | 0.000441 | -1e-06 | 100.1345 | 100.0758 | 128.3365 | 21.2612 | 0.001862 | 0.001705 | -152.7412 | 0.21878 | 8.4014 | 0.003329 | 0.004792 | 0.004578 | 5.4e-05 | 0.001227 | 0.31382 | 1.3469 | 2.5474 | 100.79 | 21.11 | 16.137 | 66.6402 | 101.9582 | 62.861 | 0.004013 | 1.0507 | 0.025524 | 24.2925 | 0.005244 | 0.13015 | 0.000154 | 0.011936 | 0.011936 | 0.45036
2019-03-26 16:10 | 1553630996.6679 | -183 | 0.20947 | 0.000442 | -1e-06 | 100.1014 | 100.3427 | 128.4529 | 21.2612 | 0.001848 | 0.001784 | -152.7553 | 0.21987 | 8.4017 | 0.003317 | 0.004417 | 0.004246 | 5.8e-05 | 0.001227 | 0.30879 | 1.3695 | 2.5475 | 100.79 | 21.11 | 16.133 | 66.3031 | 101.6597 | 64.746 | 0.004012 | 1.0448 | 0.025464 | 24.3715 | -0.021534 | 0.1308 | 0.000173 | 0.013342 | 0.013342 | 0.4049
2019-03-26 16:10 | 1553630997.6684 | -182 | 0.20947 | 0.000442 | -1e-06 | 100.0752 | 100.1397 | 128.4607 | 21.2612 | 0.001805 | 0.001662 | -152.7432 | 0.21953 | 8.3999 | 0.003355 | 0.004547 | 0.004402 | 5.1e-05 | 0.001232 | 0.3282 | 1.3973 | 2.5474 | 100.79 | 21.11 | 16.131 | 65.2644 | 102.5748 | 61.8268 | 0.004012 | 1.0539 | 0.025576 | 24.2668 | -0.005765 | 0.1306 | 0.000217 | 0.016809 | 0.016809 | 0.3209
2019-03-26 16:10 | 1553630998.6681 | -181 | 0.20947 | 0.000441 | 0 | 100.0902 | 100.0385 | 128.3398 | 21.2612 | 0.001911 | 0.001704 | -152.7399 | 0.21922 | 8.4024 | 0.003388 | 0.004744 | 0.004571 | 5.3e-05 | 0.001225 | 0.28998 | 1.4365 | 2.5476 | 100.79 | 21.11 | 16.13 | 67.8026 | 103.365 | 62.8251 | 0.004012 | 1.0508 | 0.025526 | 24.2915 | 0.004623 | 0.13042 | 0.00017 | 0.013111 | 0.013111 | 0.41082
2019-03-26 16:10 | 1553630999.6691 | -180 | 0.20947 | 0.000443 | -1e-06 | 100.1448 | 100.1209 | 128.6919 | 21.2612 | 0.001839 | 0.001717 | -152.7423 | 0.21923 | 8.4011 | 0.003451 | 0.004801 | 0.004618 | 5.4e-05 | 0.001225 | 0.34235 | 1.3712 | 2.5474 | 100.79 | 21.11 | 16.129 | 66.0799 | 104.9124 | 63.1429 | 0.004011 | 1.0498 | 0.025549 | 24.3363 | 0.002141 | 0.13042 | 0.000185 | 0.014287 | 0.014287 | 0.37701
2019-03-26 16:10 | 1553631000.6681 | -179 | 0.2095 | 0.000446 | -1e-06 | 100.0734 | 100.0469 | 128.3135 | 21.2612 | 0.001806 | 0.001739 | -152.7626 | 0.21951 | 8.4028 | 0.003488 | 0.004981 | 0.004819 | 5.2e-05 | 0.001223 | 0.33931 | 1.3493 | 2.5475 | 100.79 | 21.11 | 16.129 | 65.2902 | 105.8069 | 63.6717 | 0.004011 | 1.0482 | 0.025491 | 24.3196 | 0.002368 | 0.13059 | 3.2e-05 | 0.002542 | 0.002542 | 2.122
2019-03-26 16:10 | 1553631001.6672 | -178 | 0.20951 | 0.000445 | -1e-06 | 100.1398 | 99.9914 | 128.1659 | 21.2612 | 0.001839 | 0.00176 | -152.7615 | 0.21917 | 8.4057 | 0.003512 | 0.004451 | 0.004234 | 5e-05 | 0.001217 | 0.31741 | 1.3485 | 2.5476 | 100.79 | 21.11 | 16.129 | 66.0698 | 106.3757 | 64.188 | 0.004011 | 1.0466 | 0.025457 | 24.3242 | 0.013263 | 0.13038 | -3e-05 | -0.002285 | -0.002285 | -2.357
2019-03-26 16:10 | 1553631002.6686 | -177 | 0.2095 | 0.000443 | 0 | 100.0693 | 100.0325 | 128.0189 | 21.2612 | 0.001809 | 0.001701 | -152.7568 | 0.21938 | 8.4039 | 0.003461 | 0.004828 | 0.004647 | 5.4e-05 | 0.001216 | 0.31653 | 1.3747 | 2.5475 | 100.79 | 21.11 | 16.128 | 65.3667 | 105.1427 | 62.7493 | 0.004011 | 1.051 | 0.025497 | 24.2583 | 0.003291 | 0.13051 | 2.1e-05 | 0.001633 | 0.001633 | 3.3008
2019-03-26 16:10 | 1553631003.6683 | -176 | 0.2095 | 0.000444 | 0 | 100.1218 | 100.0412 | 128.1694 | 21.2612 | 0.001863 | 0.00177 | -152.7549 | 0.22086 | 8.4014 | 0.003454 | 0.004195 | 0.003965 | 5.3e-05 | 0.001218 | 0.31924 | 1.344 | 2.5474 | 100.79 | 21.11 | 16.127 | 66.6487 | 104.9808 | 64.412 | 0.004011 | 1.0459 | 0.025449 | 24.3326 | 0.007205 | 0.13139 | 2.9e-05 | 0.002261 | 0.002261 | 2.3999
2019-03-26 16:10 | 1553631004.6687 | -175 | 0.20951 | 0.000446 | -1e-06 | 100.1803 | 100.1646 | 128.3482 | 21.2612 | 0.001838 | 0.001794 | -152.7284 | 0.22024 | 8.4029 | 0.003448 | 0.004881 | 0.004672 | 5.1e-05 | 0.001219 | 0.33339 | 1.3465 | 2.5475 | 100.79 | 21.11 | 16.125 | 66.0448 | 104.8416 | 65.0012 | 0.00401 | 1.044 | 0.025444 | 24.3708 | 0.001401 | 0.13102 | -2.4e-05 | -0.001803 | -0.001803 | -3.0015
2019-03-26 16:10 | 1553631005.6684 | -174 | 0.2095 | 0.000444 | -1e-06 | 100.1135 | 100.3162 | 128.2837 | 21.2612 | 0.001854 | 0.001711 | -152.7302 | 0.2205 | 8.404 | 0.003465 | 0.004773 | 0.004595 | 5.3e-05 | 0.00122 | 0.30525 | 1.3527 | 2.5475 | 100.79 | 21.11 | 16.126 | 66.4373 | 105.2332 | 62.9869 | 0.004011 | 1.0503 | 0.025514 | 24.292 | -0.018085 | 0.13118 | 5e-06 | 0.000403 | 0.000403 | 13.4591
2019-03-26 16:10 | 1553631006.6677 | -173 | 0.2095 | 0.000439 | -1e-06 | 100.1242 | 100.1252 | 128.1863 | 21.2612 | 0.001879 | 0.001745 | -152.7318 | 0.21973 | 8.4022 | 0.003473 | 0.004874 | 0.004675 | 5.2e-05 | 0.001221 | 0.35601 | 1.3473 | 2.5475 | 100.79 | 21.11 | 16.126 | 67.0483 | 105.4441 | 63.8075 | 0.004011 | 1.0477 | 0.025473 | 24.3124 | -9.5e-05 | 0.13072 | 4.3e-05 | 0.003323 | 0.003323 | 1.6248
2019-03-26 16:10 | 1553631007.6679 | -172 | 0.20948 | 0.000441 | 0 | 100.1091 | 100.0506 | 128.1206 | 21.2612 | 0.001826 | 0.00178 | -152.7269 | 0.21942 | 8.4034 | 0.003415 | 0.005567 | 0.005385 | 5.1e-05 | 0.001223 | 0.33411 | 1.351 | 2.5474 | 100.79 | 21.11 | 16.126 | 65.7662 | 104.03 | 64.6509 | 0.004011 | 1.0451 | 0.025435 | 24.3366 | 0.00523 | 0.13054 | 0.000168 | 0.012952 | 0.012952 | 0.41625
2019-03-26 16:10 | 1553631008.6686 | -171 | 0.20947 | 0.000441 | 0 | 100.177 | 100.0365 | 128.1367 | 21.2612 | 0.001821 | 0.00177 | -152.7301 | 0.22056 | 8.4027 | 0.003475 | 0.004374 | 0.004236 | 5.1e-05 | 0.001223 | 0.33843 | 1.3854 | 2.5476 | 100.79 | 21.11 | 16.126 | 65.6375 | 105.4943 | 64.4158 | 0.004011 | 1.0459 | 0.025445 | 24.3296 | 0.012556 | 0.13121 | 0.000206 | 0.015871 | 0.015871 | 0.34143
2019-03-26 16:10 | 1553631009.668 | -170 | 0.20946 | 0.000441 | 0 | 100.1974 | 100.0459 | 127.83 | 21.2612 | 0.001844 | 0.001771 | -152.7383 | 0.2192 | 8.4011 | 0.00343 | 0.003866 | 0.00369 | 5.3e-05 | 0.001217 | 0.35153 | 1.3902 | 2.5477 | 100.79 | 21.11 | 16.126 | 66.2078 | 104.3981 | 64.4346 | 0.004011 | 1.0458 | 0.025414 | 24.3012 | 0.013533 | 0.1304 | 0.000251 | 0.019229 | 0.019229 | 0.28007
2019-03-26 16:10 | 1553631010.6687 | -169 | 0.2095 | 0.00044 | 0 | 100.1994 | 100.0953 | 128.0555 | 21.2612 | 0.001901 | 0.001727 | -152.7162 | 0.21959 | 8.4051 | 0.003452 | 0.004727 | 0.004576 | 5.4e-05 | 0.001215 | 0.32261 | 1.3774 | 2.5474 | 100.79 | 21.11 | 16.123 | 67.5663 | 104.9279 | 63.3882 | 0.00401 | 1.0491 | 0.025476 | 24.2849 | 0.0093 | 0.13063 | 5.1e-05 | 0.003931 | 0.003931 | 1.3726
2019-03-26 16:10 | 1553631011.6687 | -168 | 0.20949 | 0.000441 | 0 | 100.1844 | 100.132 | 128.1893 | 21.2612 | 0.001824 | 0.001729 | -152.6853 | 0.22045 | 8.4005 | 0.003464 | 0.005169 | 0.005 | 5.4e-05 | 0.001209 | 0.3178 | 1.3563 | 2.5475 | 100.79 | 21.11 | 16.121 | 65.7107 | 105.2152 | 63.425 | 0.004009 | 1.0489 | 0.025488 | 24.2989 | 0.004682 | 0.13115 | 0.000101 | 0.007778 | 0.007778 | 0.69633
2019-03-26 16:10 | 1553631012.6677 | -167 | 0.2095 | 0.000444 | 0 | 100.2222 | 100.1195 | 128.2318 | 21.2612 | 0.001834 | 0.001751 | -152.7086 | 0.21978 | 8.402 | 0.003453 | 0.004985 | 0.004783 | 5.4e-05 | 0.001202 | 0.31852 | 1.3723 | 2.5474 | 100.79 | 21.11 | 16.117 | 65.9505 | 104.9437 | 63.9543 | 0.004008 | 1.0473 | 0.025472 | 24.322 | 0.009166 | 0.13075 | 3.5e-05 | 0.002707 | 0.002707 | 1.9945
2019-03-26 16:10 | 1553631013.6677 | -166 | 0.20949 | 0.000442 | -1e-06 | 100.182 | 100.2863 | 128.2853 | 21.2612 | 0.001883 | 0.001747 | -152.7075 | 0.21981 | 8.4055 | 0.003438 | 0.004577 | 0.004446 | 5.2e-05 | 0.001205 | 0.30989 | 1.3891 | 2.5475 | 100.79 | 21.11 | 16.115 | 67.1286 | 104.5939 | 63.8693 | 0.004008 | 1.0476 | 0.025481 | 24.324 | -0.009301 | 0.13077 | 0.00011 | 0.008477 | 0.008477 | 0.63713
2019-03-26 16:10 | 1553631014.6682 | -165 | 0.20947 | 0.000441 | -1e-06 | 100.2195 | 100.4264 | 128.014 | 21.2612 | 0.001884 | 0.001758 | -152.7373 | 0.21898 | 8.4003 | 0.00347 | 0.004322 | 0.004152 | 5.4e-05 | 0.001206 | 0.3298 | 1.3622 | 2.5475 | 100.79 | 21.11 | 16.109 | 67.1597 | 105.367 | 64.1167 | 0.004007 | 1.0468 | 0.025444 | 24.3072 | -0.018448 | 0.13027 | 0.000225 | 0.017319 | 0.017319 | 0.31066
2019-03-26 16:10 | 1553631015.6679 | -164 | 0.20949 | 0.000433 | 0 | 100.2697 | 100.1162 | 127.8598 | 21.2612 | 0.001826 | 0.001718 | -152.7454 | 0.21931 | 8.4025 | 0.003538 | 0.004625 | 0.004515 | 5.5e-05 | 0.001216 | 0.33516 | 1.3522 | 2.5474 | 100.79 | 21.11 | 16.107 | 65.7756 | 107.0144 | 63.1689 | 0.004006 | 1.0497 | 0.025465 | 24.2584 | 0.013707 | 0.13046 | 0.000113 | 0.008698 | 0.008698 | 0.61945
2019-03-26 16:10 | 1553631016.6671 | -163 | 0.20949 | 0.00044 | 0 | 100.2532 | 100.1034 | 127.729 | 21.2612 | 0.001866 | 0.001732 | -152.7262 | 0.21944 | 8.403 | 0.00352 | 0.004793 | 0.004571 | 5.4e-05 | 0.001222 | 0.32925 | 1.3173 | 2.5474 | 100.79 | 21.11 | 16.105 | 66.7352 | 106.5693 | 63.4953 | 0.004006 | 1.0487 | 0.02544 | 24.2578 | 0.013374 | 0.13055 | 0.000104 | 0.008019 | 0.008019 | 0.67239
2019-03-26 16:10 | 1553631017.668 | -162 | 0.20947 | 0.00044 | 0 | 100.177 | 100.3377 | 127.8682 | 21.2612 | 0.001906 | 0.001768 | -152.7201 | 0.21962 | 8.4012 | 0.003465 | 0.004501 | 0.004322 | 5e-05 | 0.001232 | 0.34971 | 1.3297 | 2.5474 | 100.79 | 21.11 | 16.098 | 67.6997 | 105.2356 | 64.3728 | 0.004004 | 1.046 | 0.02542 | 24.3026 | -0.014335 | 0.13065 | 0.00018 | 0.013858 | 0.013858 | 0.38938
2019-03-26 16:10 | 1553631018.667 | -161 | 0.20948 | 0.000439 | 0 | 100.2585 | 100.2278 | 127.9213 | 21.2612 | 0.001853 | 0.001742 | -152.7074 | 0.22034 | 8.4007 | 0.003433 | 0.005227 | 0.005094 | 5e-05 | 0.001228 | 0.34108 | 1.3532 | 2.5474 | 100.79 | 21.11 | 16.095 | 66.4101 | 104.4563 | 63.7396 | 0.004003 | 1.048 | 0.025449 | 24.2848 | 0.002741 | 0.13108 | 0.000124 | 0.009538 | 0.009538 | 0.56759
2019-03-26 16:10 | 1553631019.6671 | -160 | 0.20948 | 0.00044 | 0 | 100.27 | 100.1313 | 127.7598 | 21.2612 | 0.001875 | 0.001713 | -152.7143 | 0.21981 | 8.4027 | 0.003511 | 0.004535 | 0.004289 | 5.3e-05 | 0.001236 | 0.32587 | 1.332 | 2.5474 | 100.79 | 21.11 | 16.087 | 66.9437 | 106.3568 | 63.0517 | 0.004001 | 1.0501 | 0.025459 | 24.2447 | 0.012382 | 0.13077 | 0.000156 | 0.012007 | 0.012007 | 0.44981
2019-03-26 16:10 | 1553631020.6672 | -159 | 0.20949 | 0.000441 | -1e-06 | 100.2874 | 100.101 | 127.6138 | 21.2612 | 0.001837 | 0.001753 | -152.7096 | 0.21938 | 8.4016 | 0.003483 | 0.005466 | 0.005332 | 5.3e-05 | 0.001238 | 0.30342 | 1.3893 | 2.5475 | 100.79 | 21.11 | 16.082 | 66.0176 | 105.6864 | 64.0051 | 0.004 | 1.0471 | 0.025409 | 24.2652 | 0.016638 | 0.13051 | 9.6e-05 | 0.007395 | 0.007395 | 0.72885
2019-03-26 16:10 | 1553631021.6682 | -158 | 0.20953 | 0.00044 | -1e-06 | 100.2349 | 100.092 | 127.606 | 21.2612 | 0.001862 | 0.001717 | -152.7191 | 0.21986 | 8.3998 | 0.003478 | 0.004298 | 0.004142 | 5.4e-05 | 0.001238 | 0.3246 | 1.3826 | 2.5475 | 100.79 | 21.12 | 16.077 | 66.6215 | 105.5591 | 63.1407 | 0.004001 | 1.0498 | 0.025441 | 24.2333 | 0.012762 | 0.13079 | -0.00014 | -0.010754 | -0.010754 | -0.5023
2019-03-26 16:10 | 1553631022.6683 | -157 | 0.20951 | 0.000439 | -1e-06 | 100.306 | 100.3353 | 127.5948 | 21.2612 | 0.00186 | 0.001775 | -152.7012 | 0.22031 | 8.4015 | 0.003547 | 0.004963 | 0.004718 | 5.2e-05 | 0.001229 | 0.31222 | 1.3585 | 2.5476 | 100.79 | 21.12 | 16.072 | 66.5784 | 107.2362 | 64.5345 | 0.004 | 1.0455 | 0.025387 | 24.2824 | -0.002615 | 0.13106 | -1.6e-05 | -0.001204 | -0.001204 | -4.4973
2019-03-26 16:10 | 1553631023.6683 | -156 | 0.2095 | 0.000441 | -1e-06 | 100.3054 | 100.2842 | 127.9551 | 21.2612 | 0.001872 | 0.001771 | -152.6992 | 0.2193 | 8.401 | 0.003478 | 0.004917 | 0.004681 | 5.1e-05 | 0.00122 | 0.31045 | 1.4033 | 2.5475 | 100.79 | 21.12 | 16.07 | 66.8673 | 105.5464 | 64.443 | 0.003999 | 1.0458 | 0.025426 | 24.3134 | 0.001889 | 0.13046 | 4.5e-05 | 0.003466 | 0.003466 | 1.5545
2019-03-26 16:10 | 1553631024.6687 | -155 | 0.20951 | 0.000438 | 0 | 100.2877 | 100.3565 | 128.0314 | 21.2612 | 0.0019 | 0.00177 | -152.6762 | 0.21963 | 8.4004 | 0.003536 | 0.004973 | 0.004841 | 5.2e-05 | 0.001224 | 0.33229 | 1.378 | 2.5475 | 100.79 | 21.12 | 16.066 | 67.5482 | 106.9622 | 64.4158 | 0.003998 | 1.0459 | 0.025435 | 24.3196 | -0.006135 | 0.13066 | -4.7e-05 | -0.00361 | -0.00361 | -1.495
2019-03-26 16:10 | 1553631025.6674 | -154 | 0.20952 | 0.00044 | -1e-06 | 100.3645 | 100.3915 | 128.2579 | 21.2612 | 0.001896 | 0.001775 | -152.6835 | 0.21989 | 8.4012 | 0.003521 | 0.005148 | 0.004925 | 5.1e-05 | 0.00122 | 0.31205 | 1.3729 | 2.5476 | 100.79 | 21.12 | 16.066 | 67.446 | 106.6033 | 64.5376 | 0.003998 | 1.0455 | 0.025453 | 24.3455 | -0.002408 | 0.13081 | -9.4e-05 | -0.007147 | -0.007147 | -0.75587
2019-03-26 16:10 | 1553631026.6684 | -153 | 0.20953 | 0.00044 | -1e-06 | 100.3261 | 100.3155 | 128.4245 | 21.2612 | 0.001823 | 0.001712 | -152.7036 | 0.21964 | 8.4029 | 0.003518 | 0.004931 | 0.004781 | 5.2e-05 | 0.001221 | 0.31437 | 1.3666 | 2.5475 | 100.79 | 21.12 | 16.067 | 65.6936 | 106.5215 | 63.0298 | 0.003999 | 1.0502 | 0.025526 | 24.3069 | 0.000943 | 0.13066 | -0.000167 | -0.012824 | -0.012824 | -0.4208
2019-03-26 16:10 | 1553631027.6675 | -152 | 0.20954 | 0.000436 | -1e-06 | 100.3113 | 100.3888 | 128.3691 | 21.2612 | 0.001928 | 0.00174 | -152.707 | 0.21944 | 8.4004 | 0.003532 | 0.004674 | 0.004488 | 5.1e-05 | 0.001223 | 0.32516 | 1.379 | 2.5475 | 100.79 | 21.12 | 16.069 | 68.2262 | 106.8726 | 63.685 | 0.003999 | 1.0481 | 0.025496 | 24.3253 | -0.006905 | 0.13055 | -0.000227 | -0.017466 | -0.017466 | -0.30868
2019-03-26 16:10 | 1553631028.6685 | -151 | 0.2095 | 0.000438 | -1e-06 | 100.3205 | 100.3168 | 128.2365 | 21.2612 | 0.001855 | 0.001746 | -152.7154 | 0.21924 | 8.4028 | 0.003462 | 0.005099 | 0.004826 | 5e-05 | 0.001229 | 0.32173 | 1.3313 | 2.5474 | 100.79 | 21.12 | 16.073 | 66.4522 | 105.1757 | 63.8373 | 0.004 | 1.0477 | 0.025477 | 24.3182 | 0.000323 | 0.13043 | 4.4e-05 | 0.003395 | 0.003395 | 1.5867
2019-03-26 16:10 | 1553631029.6676 | -150 | 0.20951 | 0.000439 | -1e-06 | 100.3155 | 100.1794 | 128.5496 | 21.2612 | 0.001857 | 0.001759 | -152.7076 | 0.21901 | 8.4018 | 0.003527 | 0.005414 | 0.005277 | 5e-05 | 0.001232 | 0.33627 | 1.3025 | 2.5474 | 100.79 | 21.12 | 16.077 | 66.5012 | 106.7545 | 64.1511 | 0.004001 | 1.0467 | 0.025496 | 24.3592 | 0.012143 | 0.13029 | -2.7e-05 | -0.002004 | -0.002004 | -2.6855
2019-03-26 16:10 | 1553631030.6675 | -149 | 0.20954 | 0.000435 | 0 | 100.3671 | 100.3367 | 128.4931 | 21.2612 | 0.001899 | 0.001758 | -152.7133 | 0.21899 | 8.4052 | 0.00349 | 0.005057 | 0.004889 | 5.1e-05 | 0.001233 | 0.33826 | 1.3759 | 2.5476 | 100.79 | 21.12 | 16.086 | 67.5187 | 105.8476 | 64.1208 | 0.004003 | 1.0468 | 0.025492 | 24.3528 | 0.002712 | 0.13028 | -0.000181 | -0.013897 | -0.013897 | -0.38717
2019-03-26 16:10 | 1553631031.668 | -148 | 0.20954 | 0.000438 | -1e-06 | 100.288 | 100.2442 | 128.552 | 21.2612 | 0.00182 | 0.001787 | -152.7108 | 0.21822 | 8.4079 | 0.003471 | 0.00499 | 0.004827 | 4.9e-05 | 0.001232 | 0.32964 | 1.4275 | 2.5477 | 100.79 | 21.12 | 16.093 | 65.6202 | 105.393 | 64.8162 | 0.004005 | 1.0446 | 0.025471 | 24.3835 | 0.003904 | 0.12982 | -0.000187 | -0.01433 | -0.01433 | -0.37415
2019-03-26 16:11 | 1553631032.6667 | -147 | 0.20951 | 0.000441 | 0 | 100.3467 | 100.4658 | 128.4844 | 21.2612 | 0.001851 | 0.001765 | -152.689 | 0.21947 | 8.4049 | 0.003382 | 0.004856 | 0.004697 | 4.9e-05 | 0.001228 | 0.32261 | 1.4012 | 2.5476 | 100.79 | 21.12 | 16.11 | 66.3656 | 103.2378 | 64.308 | 0.004009 | 1.0462 | 0.025484 | 24.3587 | -0.010603 | 0.13056 | -2.7e-05 | -0.002087 | -0.002087 | -2.5834
2019-03-26 16:11 | 1553631033.6668 | -146 | 0.20949 | 0.000441 | -1e-06 | 100.3273 | 100.3972 | 128.0041 | 21.2612 | 0.001905 | 0.001741 | -152.6978 | 0.21928 | 8.4011 | 0.003539 | 0.004558 | 0.004389 | 5.4e-05 | 0.001231 | 0.31188 | 1.3203 | 2.5475 | 100.79 | 21.11 | 16.118 | 67.6684 | 107.0253 | 63.7083 | 0.004009 | 1.0481 | 0.025459 | 24.2915 | -0.006232 | 0.13045 | 8.3e-05 | 0.006459 | 0.006459 | 0.83412
2019-03-26 16:11 | 1553631034.6679 | -145 | 0.20955 | 0.000442 | -1e-06 | 100.3069 | 100.4093 | 127.8985 | 21.2612 | 0.001844 | 0.001737 | -152.697 | 0.21854 | 8.4034 | 0.003493 | 0.004197 | 0.00398 | 5e-05 | 0.001231 | 0.34202 | 1.3447 | 2.5475 | 100.79 | 21.11 | 16.126 | 66.1953 | 105.9185 | 63.6132 | 0.004011 | 1.0483 | 0.025452 | 24.2781 | -0.009127 | 0.13001 | -0.000259 | -0.019811 | -0.019811 | -0.27103
2019-03-26 16:11 | 1553631035.6679 | -144 | 0.20953 | 0.000438 | -1e-06 | 100.3999 | 100.4049 | 127.9023 | 21.2612 | 0.001859 | 0.001772 | -152.6877 | 0.21993 | 8.4038 | 0.003574 | 0.00461 | 0.004401 | 4.9e-05 | 0.001226 | 0.35026 | 1.3671 | 2.5475 | 100.79 | 21.12 | 16.141 | 66.5465 | 107.886 | 64.4618 | 0.004017 | 1.0457 | 0.02542 | 24.309 | -0.00045 | 0.13084 | -0.000127 | -0.00969 | -0.00969 | -0.55765
2019-03-26 16:11 | 1553631036.6683 | -143 | 0.20952 | 0.000441 | -1e-06 | 100.2744 | 100.102 | 127.7466 | 21.2612 | 0.001877 | 0.001745 | -152.685 | 0.21935 | 8.4049 | 0.003545 | 0.005317 | 0.005108 | 4.8e-05 | 0.001226 | 0.33892 | 1.3384 | 2.5475 | 100.79 | 21.11 | 16.148 | 66.9889 | 107.1828 | 63.8092 | 0.004016 | 1.0477 | 0.025429 | 24.2707 | 0.01539 | 0.13049 | -9.5e-05 | -0.007238 | -0.007238 | -0.74456
2019-03-26 16:11 | 1553631037.6673 | -142 | 0.20954 | 0.000437 | -1e-06 | 100.3863 | 100.3236 | 127.8182 | 21.2612 | 0.001908 | 0.001717 | -152.6896 | 0.21932 | 8.4027 | 0.003492 | 0.005027 | 0.004891 | 5e-05 | 0.001229 | 0.29717 | 1.3714 | 2.5476 | 100.79 | 21.12 | 16.154 | 67.7269 | 105.9068 | 63.1391 | 0.00402 | 1.0498 | 0.025462 | 24.2534 | 0.005591 | 0.13048 | -0.000181 | -0.013852 | -0.013852 | -0.38901
2019-03-26 16:11 | 1553631038.668 | -141 | 0.20953 | 0.00044 | 0 | 100.4161 | 100.4876 | 127.9276 | 21.2612 | 0.001872 | 0.001719 | -152.7069 | 0.21959 | 8.4039 | 0.00354 | 0.004809 | 0.004599 | 5e-05 | 0.00123 | 0.30718 | 1.3201 | 2.5474 | 100.79 | 21.12 | 16.164 | 66.8624 | 107.0639 | 63.1906 | 0.004023 | 1.0497 | 0.025471 | 24.2656 | -0.006367 | 0.13063 | -0.000141 | -0.010834 | -0.010834 | -0.49799
2019-03-26 16:11 | 1553631039.6681 | -140 | 0.20953 | 0.000437 | -1e-06 | 100.3134 | 100.3653 | 127.8269 | 21.2612 | 0.001867 | 0.001791 | -152.7116 | 0.21994 | 8.4007 | 0.003493 | 0.005185 | 0.004968 | 4.9e-05 | 0.001233 | 0.30845 | 1.3329 | 2.5474 | 100.79 | 21.12 | 16.168 | 66.747 | 105.9202 | 64.9334 | 0.004024 | 1.0443 | 0.025395 | 24.3188 | -0.004624 | 0.13084 | -0.000172 | -0.013155 | -0.013155 | -0.41078
2019-03-26 16:11 | 1553631040.6683 | -139 | 0.20955 | 0.000441 | -1e-06 | 100.3901 | 100.3303 | 128.0084 | 21.2612 | 0.00187 | 0.001758 | -152.7223 | 0.2194 | 8.4051 | 0.003526 | 0.004577 | 0.004413 | 5e-05 | 0.00124 | 0.35026 | 1.34 | 2.5475 | 100.79 | 21.12 | 16.174 | 66.8134 | 106.7331 | 64.1379 | 0.004025 | 1.0467 | 0.025443 | 24.3074 | 0.005333 | 0.13052 | -0.000245 | -0.01877 | -0.01877 | -0.28718
2019-03-26 16:11 | 1553631041.6672 | -138 | 0.20951 | 0.000439 | -1e-06 | 100.3228 | 100.5202 | 127.8837 | 21.2612 | 0.001861 | 0.001744 | -152.7281 | 0.21949 | 8.3996 | 0.003473 | 0.004616 | 0.004437 | 5e-05 | 0.00124 | 0.31382 | 1.3719 | 2.5474 | 100.79 | 21.12 | 16.174 | 66.6073 | 105.4267 | 63.7873 | 0.004025 | 1.0478 | 0.025444 | 24.283 | -0.017582 | 0.13058 | -3.8e-05 | -0.002855 | -0.002855 | -1.8891
2019-03-26 16:11 | 1553631042.6675 | -137 | 0.20954 | 0.00044 | 0 | 100.3615 | 100.4422 | 127.6783 | 21.2612 | 0.001904 | 0.001769 | -152.7513 | 0.21909 | 8.4005 | 0.003521 | 0.004117 | 0.003949 | 5.4e-05 | 0.00124 | 0.33107 | 1.3845 | 2.5474 | 100.79 | 21.12 | 16.173 | 67.6389 | 106.604 | 64.3831 | 0.004025 | 1.046 | 0.025401 | 24.2849 | -0.007192 | 0.13034 | -0.000216 | -0.016528 | -0.016528 | -0.32569
2019-03-26 16:11 | 1553631043.6673 | -136 | 0.20952 | 0.000438 | -1e-06 | 100.4126 | 100.3165 | 127.7225 | 21.2612 | 0.001853 | 0.001755 | -152.741 | 0.21901 | 8.4015 | 0.003554 | 0.004407 | 0.004218 | 5.3e-05 | 0.001238 | 0.34075 | 1.3409 | 2.5474 | 100.79 | 21.12 | 16.166 | 66.4171 | 107.4024 | 64.0481 | 0.004023 | 1.047 | 0.025418 | 24.277 | 0.008561 | 0.13029 | -6.6e-05 | -0.005026 | -0.005026 | -1.0706
2019-03-26 16:11 | 1553631044.6673 | -135 | 0.20954 | 0.00044 | -1e-06 | 100.3739 | 100.7596 | 128.0384 | 21.2612 | 0.001889 | 0.001741 | -152.7411 | 0.21939 | 8.4027 | 0.00361 | 0.004696 | 0.004517 | 5.1e-05 | 0.001236 | 0.33229 | 1.2956 | 2.5473 | 100.79 | 21.12 | 16.161 | 67.2736 | 108.7647 | 63.7256 | 0.004022 | 1.048 | 0.025462 | 24.2954 | -0.034302 | 0.13051 | -0.000196 | -0.014999 | -0.014999 | -0.35938
2019-03-26 16:11 | 1553631045.6673 | -134 | 0.2095 | 0.000438 | -1e-06 | 100.3928 | 100.4406 | 127.9116 | 21.2612 | 0.001875 | 0.001747 | -152.7465 | 0.21948 | 8.4026 | 0.003542 | 0.00492 | 0.004751 | 5.3e-05 | 0.001237 | 0.32444 | 1.3301 | 2.5475 | 100.79 | 21.12 | 16.146 | 66.9406 | 107.1126 | 63.8631 | 0.004018 | 1.0476 | 0.025444 | 24.2883 | -0.004255 | 0.13057 | 2e-05 | 0.001623 | 0.001623 | 3.3223
2019-03-26 16:11 | 1553631046.6684 | -133 | 0.20951 | 0.000441 | -1e-06 | 100.3639 | 100.292 | 128.1534 | 21.2612 | 0.001772 | 0.001753 | -152.7448 | 0.21915 | 8.4047 | 0.003533 | 0.004913 | 0.004693 | 5.3e-05 | 0.00124 | 0.33284 | 1.3922 | 2.5476 | 100.79 | 21.12 | 16.138 | 64.4579 | 106.8882 | 64.0161 | 0.004016 | 1.0471 | 0.025462 | 24.3168 | 0.006409 | 0.13037 | -4.6e-05 | -0.003465 | -0.003465 | -1.5539
2019-03-26 16:11 | 1553631047.6676 | -132 | 0.20953 | 0.00044 | -1e-06 | 100.4577 | 100.4463 | 127.9731 | 21.2612 | 0.001847 | 0.001742 | -152.7346 | 0.21949 | 8.4027 | 0.003553 | 0.004915 | 0.004768 | 5.6e-05 | 0.001246 | 0.32156 | 1.3884 | 2.5477 | 100.79 | 21.12 | 16.129 | 66.2781 | 107.3709 | 63.749 | 0.004014 | 1.0479 | 0.025454 | 24.2901 | 0.001021 | 0.13058 | -0.000181 | -0.013834 | -0.013834 | -0.38981
2019-03-26 16:11 | 1553631048.6679 | -131 | 0.20956 | 0.000438 | 0 | 100.3949 | 100.3582 | 127.902 | 21.2612 | 0.001856 | 0.001731 | -152.7205 | 0.21976 | 8.4005 | 0.003506 | 0.004829 | 0.00464 | 5.6e-05 | 0.00125 | 0.33212 | 1.3441 | 2.5475 | 100.79 | 21.12 | 16.109 | 66.4794 | 106.227 | 63.4703 | 0.004009 | 1.0488 | 0.025458 | 24.2733 | 0.003267 | 0.13073 | -0.000319 | -0.024526 | -0.024526 | -0.22014
2019-03-26 16:11 | 1553631049.6676 | -130 | 0.20957 | 0.000438 | -1e-06 | 100.3733 | 100.2335 | 127.8518 | 21.2612 | 0.001855 | 0.001739 | -152.7048 | 0.2193 | 8.4023 | 0.003552 | 0.004897 | 0.004804 | 5.2e-05 | 0.001244 | 0.33229 | 1.3402 | 2.5475 | 100.79 | 21.12 | 16.099 | 66.4529 | 107.3481 | 63.6686 | 0.004006 | 1.0482 | 0.025445 | 24.2757 | 0.012469 | 0.13046 | -0.000379 | -0.029055 | -0.029055 | -0.18544
2019-03-26 16:11 | 1553631050.6666 | -129 | 0.2095 | 0.000439 | -1e-06 | 100.3441 | 100.5105 | 127.6396 | 21.2612 | 0.001866 | 0.001774 | -152.6945 | 0.22014 | 8.4021 | 0.003498 | 0.004437 | 0.004234 | 5.1e-05 | 0.001243 | 0.33892 | 1.3858 | 2.5474 | 100.79 | 21.12 | 16.079 | 66.7213 | 106.0507 | 64.5087 | 0.004002 | 1.0456 | 0.025392 | 24.2857 | -0.01482 | 0.13096 | 1.3e-05 | 0.001048 | 0.001048 | 5.1625
2019-03-26 16:11 | 1553631051.6677 | -128 | 0.20951 | 0.000441 | -1e-06 | 100.4019 | 100.532 | 127.7027 | 21.2612 | 0.001851 | 0.001741 | -152.7056 | 0.2208 | 8.4033 | 0.003539 | 0.004426 | 0.004331 | 5.4e-05 | 0.001242 | 0.32405 | 1.3579 | 2.5475 | 100.79 | 21.12 | 16.068 | 66.3584 | 107.0457 | 63.7155 | 0.003999 | 1.048 | 0.025429 | 24.2632 | -0.011578 | 0.13135 | -1.4e-05 | -0.001018 | -0.001018 | -5.3307
2019-03-26 16:11 | 1553631052.6667 | -127 | 0.20955 | 0.000437 | 0 | 100.3352 | 100.4446 | 127.9809 | 21.2612 | 0.001828 | 0.001779 | -152.7136 | 0.22009 | 8.4025 | 0.003543 | 0.004616 | 0.004488 | 5.2e-05 | 0.001247 | 0.3361 | 1.3666 | 2.5475 | 100.79 | 21.12 | 16.058 | 65.8045 | 107.126 | 64.6384 | 0.003996 | 1.0452 | 0.025421 | 24.3228 | -0.009744 | 0.13093 | -0.000276 | -0.021176 | -0.021176 | -0.25535
2019-03-26 16:11 | 1553631053.6678 | -126 | 0.2095 | 0.000434 | -1e-06 | 100.4143 | 100.3962 | 128.2466 | 21.2612 | 0.00184 | 0.001698 | -152.7239 | 0.21936 | 8.4043 | 0.003583 | 0.005145 | 0.00488 | 5.3e-05 | 0.001241 | 0.33644 | 1.3509 | 2.5474 | 100.79 | 21.12 | 16.04 | 66.1087 | 108.1051 | 62.6776 | 0.003992 | 1.0513 | 0.025522 | 24.2773 | 0.001618 | 0.1305 | 5e-05 | 0.003921 | 0.003921 | 1.3745
2019-03-26 16:11 | 1553631054.6668 | -125 | 0.20953 | 0.000437 | -1e-06 | 100.2886 | 100.5475 | 128.0985 | 21.2612 | 0.001928 | 0.00176 | -152.7523 | 0.21977 | 8.401 | 0.003569 | 0.005217 | 0.005046 | 5.6e-05 | 0.001241 | 0.33588 | 1.325 | 2.5474 | 100.79 | 21.12 | 16.032 | 68.2262 | 107.7688 | 64.1738 | 0.00399 | 1.0466 | 0.025451 | 24.3173 | -0.023058 | 0.13074 | -0.000149 | -0.011441 | -0.011441 | -0.47196
2019-03-26 16:11 | 1553631055.6682 | -124 | 0.2095 | 0.000437 | -1e-06 | 100.3689 | 100.4217 | 128.1459 | 21.2612 | 0.001802 | 0.001773 | -152.7356 | 0.21987 | 8.401 | 0.003509 | 0.00471 | 0.004467 | 5.5e-05 | 0.001244 | 0.33157 | 1.3309 | 2.5475 | 100.79 | 21.12 | 16.025 | 65.1925 | 106.301 | 64.4948 | 0.003988 | 1.0456 | 0.025443 | 24.3333 | -0.004708 | 0.1308 | 1.4e-05 | 0.001109 | 0.001109 | 4.8704
2019-03-26 16:11 | 1553631056.668 | -123 | 0.20951 | 0.000439 | 0 | 100.3771 | 100.4395 | 127.8368 | 21.2612 | 0.001897 | 0.001745 | -152.7426 | 0.21949 | 8.4016 | 0.003577 | 0.004831 | 0.004689 | 5.4e-05 | 0.001244 | 0.35961 | 1.369 | 2.5475 | 100.79 | 21.12 | 16.019 | 67.4787 | 107.9577 | 63.8068 | 0.003987 | 1.0477 | 0.025438 | 24.2792 | -0.005559 | 0.13057 | -2.8e-05 | -0.002119 | -0.002119 | -2.5447
2019-03-26 16:11 | 1553631057.6683 | -122 | 0.20955 | 0.000441 | 0 | 100.4052 | 100.4291 | 127.9391 | 21.2612 | 0.001929 | 0.001791 | -152.7263 | 0.21965 | 8.4031 | 0.003626 | 0.004665 | 0.004497 | 5.7e-05 | 0.001248 | 0.32892 | 1.3088 | 2.5474 | 100.79 | 21.12 | 16.011 | 68.2503 | 109.1517 | 64.9288 | 0.003985 | 1.0443 | 0.025406 | 24.3293 | -0.002132 | 0.13067 | -0.000288 | -0.022079 | -0.022079 | -0.24443
2019-03-26 16:11 | 1553631058.667 | -121 | 0.2095 | 0.000448 | -1e-06 | 100.3771 | 100.3723 | 128.0037 | 21.2612 | 0.001853 | 0.0018 | -152.726 | 0.21926 | 8.4026 | 0.003617 | 0.004448 | 0.004246 | 5.7e-05 | 0.001248 | 0.31996 | 1.3415 | 2.5475 | 100.79 | 21.13 | 16.01 | 66.4192 | 108.9214 | 65.1427 | 0.003987 | 1.0436 | 0.025405 | 24.3432 | 0.000431 | 0.13044 | 2.2e-05 | 0.001705 | 0.001705 | 3.1602
2019-03-26 16:11 | 1553631059.668 | -120 | 0.2095 | 0.000448 | -1e-06 | 100.3848 | 100.4086 | 127.8952 | 21.2612 | 0.001912 | 0.001737 | -152.7398 | 0.21972 | 8.4015 | 0.003626 | 0.004713 | 0.004505 | 5.6e-05 | 0.00125 | 0.37144 | 1.3335 | 2.5475 | 100.79 | 21.13 | 16.009 | 67.8432 | 109.1437 | 63.6342 | 0.003987 | 1.0483 | 0.025451 | 24.2785 | -0.00212 | 0.13071 | 3.9e-05 | 0.00305 | 0.00305 | 1.7697
2019-03-26 16:11 | 1553631060.6681 | -119 | 0.20953 | 0.000443 | -1e-06 | 100.386 | 100.4301 | 128.0593 | 21.2612 | 0.001894 | 0.001755 | -152.732 | 0.2195 | 8.4031 | 0.003672 | 0.005016 | 0.004835 | 5.7e-05 | 0.001254 | 0.32709 | 1.3432 | 2.5476 | 100.79 | 21.13 | 16.01 | 67.4095 | 110.2729 | 64.0474 | 0.003987 | 1.047 | 0.025452 | 24.309 | -0.003931 | 0.13058 | -0.00013 | -0.009912 | -0.009912 | -0.5441
2019-03-26 16:11 | 1553631061.6687 | -118 | 0.20954 | 0.000443 | -1e-06 | 100.4076 | 100.3908 | 128.1785 | 21.2612 | 0.001883 | 0.001768 | -152.7208 | 0.22003 | 8.4062 | 0.003642 | 0.004749 | 0.004614 | 5.4e-05 | 0.001254 | 0.33516 | 1.3634 | 2.5476 | 100.79 | 21.13 | 16.013 | 67.1325 | 109.5332 | 64.3636 | 0.003988 | 1.046 | 0.025451 | 24.3317 | 0.001492 | 0.1309 | -0.000233 | -0.017808 | -0.017808 | -0.30356
2019-03-26 16:11 | 1553631062.6682 | -117 | 0.20953 | 0.000447 | 0 | 100.3816 | 100.4392 | 128.2954 | 21.2612 | 0.001891 | 0.001801 | -152.7196 | 0.22027 | 8.4 | 0.003591 | 0.00498 | 0.004749 | 5.4e-05 | 0.001256 | 0.32925 | 1.3521 | 2.5476 | 100.79 | 21.13 | 16.017 | 67.3237 | 108.3115 | 65.1528 | 0.003989 | 1.0436 | 0.025433 | 24.3713 | -0.005134 | 0.13104 | -0.000182 | -0.013947 | -0.013947 | -0.38802
2019-03-26 16:11 | 1553631063.6672 | -116 | 0.20952 | 0.000449 | -1e-06 | 100.3925 | 100.377 | 128.1356 | 21.2612 | 0.001905 | 0.001775 | -152.7066 | 0.21969 | 8.4069 | 0.003568 | 0.004611 | 0.004451 | 5.7e-05 | 0.001256 | 0.30165 | 1.3431 | 2.5475 | 100.79 | 21.13 | 16.021 | 67.6636 | 107.7348 | 64.533 | 0.00399 | 1.0455 | 0.025441 | 24.3337 | 0.001379 | 0.13069 | -9.6e-05 | -0.007311 | -0.007311 | -0.73831
2019-03-26 16:11 | 1553631064.6673 | -115 | 0.20954 | 0.000447 | -1e-06 | 100.3627 | 100.3447 | 128.0821 | 21.2612 | 0.001874 | 0.001753 | -152.6977 | 0.21887 | 8.4022 | 0.003648 | 0.004824 | 0.004686 | 5.3e-05 | 0.001255 | 0.32693 | 1.3443 | 2.5474 | 100.79 | 21.13 | 16.026 | 66.9194 | 109.679 | 64.0082 | 0.003991 | 1.0471 | 0.025455 | 24.3098 | 0.001598 | 0.13021 | -0.00019 | -0.014563 | -0.014563 | -0.36926
2019-03-26 16:11 | 1553631065.6664 | -114 | 0.20953 | 0.000448 | -1e-06 | 100.3624 | 100.4032 | 128.0602 | 21.2612 | 0.001843 | 0.001737 | -152.7155 | 0.21944 | 8.4037 | 0.003628 | 0.00518 | 0.00497 | 5.5e-05 | 0.001259 | 0.33948 | 1.3221 | 2.5475 | 100.79 | 21.13 | 16.035 | 66.1612 | 109.1886 | 63.6178 | 0.003993 | 1.0483 | 0.025468 | 24.2936 | -0.00364 | 0.13054 | -0.000161 | -0.012343 | -0.012343 | -0.4368
2019-03-26 16:11 | 1553631066.6667 | -113 | 0.20951 | 0.000449 | -1e-06 | 100.3597 | 100.2893 | 128.1128 | 21.2612 | 0.001905 | 0.001765 | -152.7041 | 0.22042 | 8.4071 | 0.003709 | 0.005141 | 0.005015 | 5.4e-05 | 0.001258 | 0.32803 | 1.3434 | 2.5475 | 100.79 | 21.13 | 16.039 | 67.6552 | 111.1773 | 64.2917 | 0.003994 | 1.0462 | 0.025448 | 24.3229 | 0.006281 | 0.13113 | -2.3e-05 | -0.001745 | -0.001745 | -3.1043
2019-03-26 16:11 | 1553631067.6665 | -112 | 0.20951 | 0.00045 | -1e-06 | 100.4073 | 100.5001 | 128.0182 | 21.2612 | 0.001866 | 0.001727 | -152.7283 | 0.21974 | 8.4017 | 0.003705 | 0.004807 | 0.004648 | 5.3e-05 | 0.001263 | 0.29502 | 1.3428 | 2.5475 | 100.79 | 21.13 | 16.048 | 66.7189 | 111.0612 | 63.3798 | 0.003996 | 1.0491 | 0.025473 | 24.281 | -0.008263 | 0.13072 | -2.4e-05 | -0.001796 | -0.001796 | -3.0061
2019-03-26 16:11 | 1553631068.6666 | -111 | 0.2095 | 0.000451 | -1e-06 | 100.3166 | 100.2755 | 128.1419 | 21.2612 | 0.001881 | 0.001739 | -152.7304 | 0.21868 | 8.4007 | 0.003619 | 0.005188 | 0.005031 | 5.1e-05 | 0.001261 | 0.34418 | 1.3345 | 2.5474 | 100.79 | 21.13 | 16.051 | 67.0849 | 108.9695 | 63.678 | 0.003997 | 1.0481 | 0.025474 | 24.3035 | 0.003668 | 0.13009 | 1.1e-05 | 0.000905 | 0.000905 | 5.9399
2019-03-26 16:11 | 1553631069.6675 | -110 | 0.20956 | 0.000451 | -1e-06 | 100.4019 | 100.3518 | 128.2994 | 21.2612 | 0.001896 | 0.001765 | -152.7309 | 0.21959 | 8.4037 | 0.003683 | 0.004782 | 0.004627 | 5.2e-05 | 0.001262 | 0.32549 | 1.3734 | 2.5475 | 100.79 | 21.13 | 16.054 | 67.4376 | 110.5357 | 64.2986 | 0.003998 | 1.0462 | 0.025466 | 24.3408 | 0.004467 | 0.13063 | -0.000332 | -0.025472 | -0.025472 | -0.2118
2019-03-26 16:11 | 1553631070.6676 | -109 | 0.20954 | 0.000453 | -1e-06 | 100.319 | 100.3693 | 128.1914 | 21.2612 | 0.001911 | 0.001741 | -152.7461 | 0.21942 | 8.4015 | 0.003548 | 0.004705 | 0.004441 | 5.2e-05 | 0.00126 | 0.30989 | 1.3854 | 2.5475 | 100.79 | 21.13 | 16.058 | 67.8151 | 107.2575 | 63.7124 | 0.003999 | 1.048 | 0.025477 | 24.3095 | -0.004483 | 0.13053 | -0.000215 | -0.016443 | -0.016443 | -0.32786
2019-03-26 16:11 | 1553631071.6666 | -108 | 0.20948 | 0.000451 | 0 | 100.3405 | 100.3078 | 127.8811 | 21.2612 | 0.001946 | 0.001759 | -152.7528 | 0.21953 | 8.403 | 0.003633 | 0.00487 | 0.00475 | 5.4e-05 | 0.001265 | 0.33069 | 1.3744 | 2.5476 | 100.79 | 21.13 | 16.06 | 68.6575 | 109.3147 | 64.1528 | 0.003999 | 1.0467 | 0.02543 | 24.2959 | 0.002922 | 0.1306 | 0.000146 | 0.011228 | 0.011228 | 0.48036
2019-03-26 16:11 | 1553631072.6668 | -107 | 0.20948 | 0.000449 | -1e-06 | 100.306 | 100.4338 | 127.3901 | 21.2612 | 0.001918 | 0.001746 | -152.7529 | 0.21968 | 8.4035 | 0.003627 | 0.00524 | 0.005025 | 5.4e-05 | 0.001263 | 0.31669 | 1.3362 | 2.5475 | 100.79 | 21.13 | 16.062 | 67.9827 | 109.176 | 63.8388 | 0.004 | 1.0476 | 0.025393 | 24.2379 | -0.011391 | 0.13069 | 0.000119 | 0.009181 | 0.009181 | 0.58787
2019-03-26 16:11 | 1553631073.667 | -106 | 0.20953 | 0.000453 | -1e-06 | 100.3715 | 100.3222 | 127.311 | 21.2612 | 0.00186 | 0.001742 | -152.7651 | 0.21985 | 8.4064 | 0.00368 | 0.004439 | 0.004246 | 5.2e-05 | 0.001268 | 0.30143 | 1.3489 | 2.5475 | 100.79 | 21.13 | 16.064 | 66.5832 | 110.4671 | 63.7343 | 0.004 | 1.048 | 0.025389 | 24.2266 | 0.004396 | 0.13079 | -0.00014 | -0.010721 | -0.010721 | -0.50383
2019-03-26 16:11 | 1553631074.6678 | -105 | 0.20953 | 0.00045 | -1e-06 | 100.3166 | 100.3363 | 127.5133 | 21.2612 | 0.001886 | 0.001775 | -152.7641 | 0.21993 | 8.4043 | 0.003677 | 0.004283 | 0.00409 | 5.1e-05 | 0.001268 | 0.29214 | 1.34 | 2.5474 | 100.79 | 21.13 | 16.064 | 67.2135 | 110.3898 | 64.5431 | 0.004 | 1.0455 | 0.025379 | 24.275 | -0.001757 | 0.13084 | -0.000146 | -0.01114 | -0.01114 | -0.48505
2019-03-26 16:11 | 1553631075.6668 | -104 | 0.20949 | 0.000448 | -1e-06 | 100.267 | 100.4483 | 127.4626 | 21.2612 | 0.001842 | 0.001741 | -152.7459 | 0.21906 | 8.4036 | 0.003629 | 0.004912 | 0.004776 | 5.2e-05 | 0.001273 | 0.34108 | 1.3692 | 2.5474 | 100.79 | 21.14 | 16.061 | 66.1588 | 109.2265 | 63.7232 | 0.004002 | 1.048 | 0.025404 | 24.2406 | -0.016154 | 0.13032 | 0.000101 | 0.007807 | 0.007807 | 0.68942
2019-03-26 16:11 | 1553631076.6683 | -103 | 0.20948 | 0.000448 | -1e-06 | 100.309 | 100.3992 | 127.6145 | 21.2612 | 0.001938 | 0.001727 | -152.7532 | 0.21983 | 8.4031 | 0.003666 | 0.004767 | 0.004599 | 5e-05 | 0.001274 | 0.33284 | 1.3449 | 2.5473 | 100.79 | 21.14 | 16.059 | 68.4625 | 110.1231 | 63.3836 | 0.004001 | 1.0491 | 0.025432 | 24.2429 | -0.008044 | 0.13078 | 0.000151 | 0.011682 | 0.011682 | 0.46233
2019-03-26 16:11 | 1553631077.667 | -102 | 0.20947 | 0.000451 | -1e-06 | 100.3577 | 100.338 | 127.469 | 21.2612 | 0.001876 | 0.001807 | -152.754 | 0.21977 | 8.4025 | 0.003557 | 0.004879 | 0.004667 | 4.9e-05 | 0.001272 | 0.323 | 1.3587 | 2.5476 | 100.79 | 21.14 | 16.056 | 66.9781 | 107.4656 | 65.2986 | 0.004001 | 1.0431 | 0.025346 | 24.2979 | 0.001751 | 0.13074 | 0.000189 | 0.014515 | 0.014515 | 0.37202
2019-03-26 16:11 | 1553631078.6664 | -101 | 0.20953 | 0.000448 | -1e-06 | 100.2414 | 100.2792 | 127.4307 | 21.2612 | 0.001845 | 0.001751 | -152.7329 | 0.21995 | 8.4024 | 0.003683 | 0.005395 | 0.005227 | 5e-05 | 0.001272 | 0.33035 | 1.3377 | 2.5476 | 100.79 | 21.14 | 16.048 | 66.2266 | 110.5389 | 63.9709 | 0.003999 | 1.0472 | 0.025392 | 24.2465 | -0.003375 | 0.13085 | -0.000129 | -0.009886 | -0.009886 | -0.54662
2019-03-26 16:11 | 1553631079.6671 | -100 | 0.20946 | 0.000446 | -1e-06 | 100.3237 | 100.261 | 127.1587 | 21.2612 | 0.001962 | 0.001798 | -152.7204 | 0.21956 | 8.4016 | 0.003602 | 0.004407 | 0.004242 | 5.4e-05 | 0.001277 | 0.33019 | 1.3424 | 2.5475 | 100.79 | 21.14 | 16.043 | 69.0286 | 108.5573 | 65.0888 | 0.003997 | 1.0438 | 0.025323 | 24.2607 | 0.00559 | 0.13062 | 0.000276 | 0.02117 | 0.02117 | 0.25481
2019-03-26 16:11 | 1553631080.6661 | -99 | 0.20951 | 0.000448 | 0 | 100.3196 | 100.2631 | 127.1514 | 21.2612 | 0.00183 | 0.0018 | -152.7103 | 0.2199 | 8.4038 | 0.00362 | 0.00506 | 0.004822 | 5e-05 | 0.001277 | 0.35889 | 1.3534 | 2.5475 | 100.79 | 21.14 | 16.032 | 65.8677 | 108.9972 | 65.1487 | 0.003995 | 1.0436 | 0.02532 | 24.2622 | 0.005041 | 0.13082 | -1.1e-05 | -0.000792 | -0.000792 | -6.8199
2019-03-26 16:11 | 1553631081.6662 | -98 | 0.20943 | 0.000444 | -1e-06 | 100.3426 | 100.6009 | 127.5401 | 21.2612 | 0.001874 | 0.001748 | -152.6985 | 0.21987 | 8.4038 | 0.003558 | 0.004681 | 0.004598 | 5.5e-05 | 0.001274 | 0.33251 | 1.3221 | 2.5475 | 100.79 | 21.14 | 16.025 | 66.9288 | 107.5088 | 63.8818 | 0.003993 | 1.0475 | 0.025406 | 24.2537 | -0.022995 | 0.1308 | 0.000441 | 0.03389 | 0.03389 | 0.1594
2019-03-26 16:11 | 1553631082.6672 | -97 | 0.20947 | 0.000446 | -1e-06 | 100.3095 | 100.4197 | 127.7628 | 21.2612 | 0.00197 | 0.001773 | -152.7005 | 0.21993 | 8.4031 | 0.003625 | 0.004391 | 0.004202 | 4.9e-05 | 0.00127 | 0.34274 | 1.347 | 2.5476 | 100.79 | 21.14 | 16.017 | 69.2291 | 109.1366 | 64.4986 | 0.003991 | 1.0456 | 0.025405 | 24.2971 | -0.009818 | 0.13084 | 0.000222 | 0.017054 | 0.017054 | 0.31684
2019-03-26 16:11 | 1553631083.6666 | -96 | 0.20948 | 0.00045 | -1e-06 | 100.2157 | 100.3333 | 127.6518 | 21.2612 | 0.001859 | 0.001762 | -152.7101 | 0.21903 | 8.4019 | 0.003588 | 0.004948 | 0.004786 | 5.4e-05 | 0.001272 | 0.32101 | 1.3774 | 2.5477 | 100.79 | 21.14 | 16.001 | 66.5544 | 108.2194 | 64.2339 | 0.003987 | 1.0464 | 0.025404 | 24.277 | -0.010493 | 0.1303 | 0.000169 | 0.013012 | 0.013012 | 0.41358
2019-03-26 16:11 | 1553631084.6663 | -95 | 0.20945 | 0.000452 | -1e-06 | 100.3228 | 100.4409 | 127.365 | 21.2612 | 0.001894 | 0.00175 | -152.7105 | 0.2202 | 8.4018 | 0.003537 | 0.005039 | 0.004862 | 5.1e-05 | 0.001275 | 0.29159 | 1.3441 | 2.5476 | 100.79 | 21.14 | 15.993 | 67.4064 | 106.9954 | 63.9247 | 0.003985 | 1.0474 | 0.025387 | 24.2386 | -0.01052 | 0.131 | 0.000335 | 0.025703 | 0.025703 | 0.21049
2019-03-26 16:11 | 1553631085.6673 | -94 | 0.20949 | 0.000446 | -1e-06 | 100.2718 | 100.2442 | 127.5086 | 21.2612 | 0.001924 | 0.001741 | -152.7195 | 0.21938 | 8.4055 | 0.003517 | 0.004912 | 0.004714 | 5.3e-05 | 0.001279 | 0.31349 | 1.3653 | 2.5476 | 100.79 | 21.14 | 15.977 | 68.1248 | 106.5121 | 63.717 | 0.003981 | 1.048 | 0.025409 | 24.2448 | 0.002456 | 0.13051 | 0.000105 | 0.008138 | 0.008138 | 0.6623
2019-03-26 16:11 | 1553631086.6665 | -93 | 0.20948 | 0.000453 | -1e-06 | 100.3093 | 100.4718 | 127.418 | 21.2612 | 0.001918 | 0.001792 | -152.7119 | 0.21969 | 8.4033 | 0.003597 | 0.00485 | 0.004576 | 5.6e-05 | 0.001279 | 0.31094 | 1.3939 | 2.5477 | 100.79 | 21.14 | 15.97 | 67.9733 | 108.4448 | 64.9497 | 0.003979 | 1.0442 | 0.025354 | 24.2805 | -0.014484 | 0.13069 | 0.000168 | 0.012909 | 0.012909 | 0.41813
2019-03-26 16:11 | 1553631087.6675 | -92 | 0.20942 | 0.000457 | -1e-06 | 100.2992 | 100.2399 | 127.3631 | 21.2612 | 0.001882 | 0.001737 | -152.7037 | 0.21974 | 8.403 | 0.003643 | 0.004406 | 0.004275 | 5.5e-05 | 0.001283 | 0.33107 | 1.4277 | 2.5476 | 100.79 | 21.14 | 15.963 | 67.1145 | 109.5522 | 63.6303 | 0.003978 | 1.0483 | 0.025398 | 24.2278 | 0.005294 | 0.13072 | 0.000477 | 0.036603 | 0.036603 | 0.1475
2019-03-26 16:11 | 1553631088.6677 | -91 | 0.20949 | 0.000457 | -1e-06 | 100.2428 | 100.2866 | 127.1732 | 21.2612 | 0.001897 | 0.001813 | -152.7224 | 0.21842 | 8.4033 | 0.003716 | 0.00535 | 0.005106 | 5.9e-05 | 0.001288 | 0.32189 | 1.3847 | 2.5476 | 100.79 | 21.14 | 15.953 | 67.478 | 111.3475 | 65.4446 | 0.003975 | 1.0427 | 0.025311 | 24.2749 | -0.003903 | 0.12994 | 8.1e-05 | 0.006209 | 0.006209 | 0.86431
2019-03-26 16:11 | 1553631089.6665 | -90 | 0.20946 | 0.000458 | -1e-06 | 100.2452 | 100.254 | 127.3748 | 21.2612 | 0.00186 | 0.001786 | -152.7444 | 0.21936 | 8.4025 | 0.003706 | 0.005028 | 0.004876 | 5.7e-05 | 0.00129 | 0.35187 | 1.3428 | 2.5475 | 100.79 | 21.14 | 15.95 | 66.5832 | 111.088 | 64.7936 | 0.003974 | 1.0447 | 0.025355 | 24.2708 | -0.000784 | 0.1305 | 0.000237 | 0.018171 | 0.018171 | 0.2966
2019-03-26 16:11 | 1553631090.6666 | -89 | 0.20945 | 0.000459 | -1e-06 | 100.3163 | 100.2315 | 127.6013 | 21.2612 | 0.00193 | 0.00179 | -152.7267 | 0.21985 | 8.404 | 0.003736 | 0.005684 | 0.005528 | 5.8e-05 | 0.001285 | 0.33179 | 1.3792 | 2.5475 | 100.79 | 21.14 | 15.946 | 68.2683 | 111.8261 | 64.9021 | 0.003973 | 1.0444 | 0.025374 | 24.2962 | 0.007571 | 0.13079 | 0.000323 | 0.024776 | 0.024776 | 0.21801
2019-03-26 16:12 | 1553631091.6667 | -88 | 0.20947 | 0.000463 | -1e-06 | 100.283 | 100.2052 | 127.7415 | 21.2612 | 0.001893 | 0.001797 | -152.7225 | 0.21975 | 8.4029 | 0.003758 | 0.004875 | 0.004638 | 5.6e-05 | 0.001288 | 0.3413 | 1.3739 | 2.5476 | 100.79 | 21.14 | 15.946 | 67.3821 | 112.351 | 65.0638 | 0.003973 | 1.0439 | 0.025382 | 24.3154 | 0.006937 | 0.13073 | 0.000193 | 0.01487 | 0.01487 | 0.36309
2019-03-26 16:12 | 1553631092.6679 | -87 | 0.20945 | 0.000459 | 0 | 100.2269 | 100.2479 | 128.0034 | 21.2612 | 0.001932 | 0.001772 | -152.7194 | 0.21824 | 8.4023 | 0.003782 | 0.004455 | 0.004351 | 5.8e-05 | 0.001296 | 0.34219 | 1.3453 | 2.5475 | 100.79 | 21.14 | 15.948 | 68.3048 | 112.9402 | 64.4642 | 0.003974 | 1.0457 | 0.02543 | 24.3187 | -0.001877 | 0.12983 | 0.000303 | 0.023285 | 0.023285 | 0.23028
2019-03-26 16:12 | 1553631093.6667 | -86 | 0.20949 | 0.000458 | -1e-06 | 100.2753 | 100.2903 | 128.1135 | 21.2612 | 0.001868 | 0.001807 | -152.6867 | 0.21963 | 8.4033 | 0.003757 | 0.004526 | 0.00431 | 6.3e-05 | 0.001291 | 0.29159 | 1.3557 | 2.5475 | 100.79 | 21.14 | 15.952 | 66.7665 | 112.3282 | 65.3048 | 0.003975 | 1.0431 | 0.02541 | 24.3594 | -0.001336 | 0.13066 | 8.3e-05 | 0.006402 | 0.006402 | 0.84285
2019-03-26 16:12 | 1553631094.6668 | -85 | 0.20951 | 0.00046 | -1e-06 | 100.2496 | 100.2496 | 128.183 | 21.2612 | 0.001874 | 0.001759 | -152.6894 | 0.21929 | 8.401 | 0.003776 | 0.00527 | 0.005072 | 6e-05 | 0.001293 | 0.29142 | 1.3568 | 2.5476 | 100.79 | 21.14 | 15.955 | 66.9187 | 112.8068 | 64.141 | 0.003976 | 1.0467 | 0.02546 | 24.3241 | 1e-06 | 0.13045 | -5.4e-05 | -0.004114 | -0.004114 | -1.3096
2019-03-26 16:12 | 1553631095.6662 | -84 | 0.2095 | 0.000457 | -1e-06 | 100.1903 | 100.2099 | 127.9196 | 21.2612 | 0.001921 | 0.001791 | -152.6828 | 0.21889 | 8.4036 | 0.003688 | 0.004868 | 0.004664 | 6.2e-05 | 0.001292 | 0.3402 | 1.3273 | 2.5475 | 100.79 | 21.14 | 15.958 | 68.0522 | 110.657 | 64.9139 | 0.003976 | 1.0443 | 0.025405 | 24.3269 | -0.001754 | 0.13022 | -6e-06 | -0.000356 | -0.000356 | -15.0956
2019-03-26 16:12 | 1553631096.6669 | -83 | 0.20953 | 0.000457 | -1e-06 | 100.2242 | 100.1693 | 127.8027 | 21.2612 | 0.001889 | 0.001753 | -152.7014 | 0.21844 | 8.4013 | 0.003704 | 0.004913 | 0.004694 | 5.9e-05 | 0.001298 | 0.33124 | 1.3459 | 2.5475 | 100.79 | 21.15 | 15.964 | 67.2736 | 111.043 | 64.0113 | 0.00398 | 1.0471 | 0.025427 | 24.2833 | 0.004907 | 0.12995 | -0.000138 | -0.01057 | -0.01057 | -0.50778
2019-03-26 16:12 | 1553631097.666 | -82 | 0.20947 | 0.000454 | 0 | 100.213 | 100.5458 | 128.0555 | 21.2612 | 0.001909 | 0.001745 | -152.6729 | 0.21992 | 8.4021 | 0.003706 | 0.005328 | 0.005114 | 6.1e-05 | 0.001301 | 0.31708 | 1.3198 | 2.5475 | 100.79 | 21.15 | 15.967 | 67.7605 | 111.0894 | 63.8036 | 0.003981 | 1.0478 | 0.02546 | 24.2999 | -0.02965 | 0.13083 | 0.00018 | 0.013868 | 0.013868 | 0.38961
2019-03-26 16:12 | 1553631098.667 | -81 | 0.20954 | 0.000451 | -1e-06 | 100.2316 | 100.3394 | 127.9086 | 21.2612 | 0.001963 | 0.001774 | -152.6701 | 0.21983 | 8.4008 | 0.003738 | 0.004789 | 0.004609 | 6e-05 | 0.001305 | 0.31996 | 1.3325 | 2.5476 | 100.79 | 21.15 | 15.976 | 69.0645 | 111.8709 | 64.5049 | 0.003983 | 1.0456 | 0.025419 | 24.3112 | -0.00961 | 0.13078 | -0.000233 | -0.017835 | -0.017835 | -0.30285
2019-03-26 16:12 | 1553631099.666 | -80 | 0.20951 | 0.000452 | -1e-06 | 100.2762 | 100.06 | 127.8926 | 21.2612 | 0.001885 | 0.001777 | -152.6534 | 0.21818 | 8.4045 | 0.003752 | 0.004479 | 0.004291 | 6.1e-05 | 0.001296 | 0.3361 | 1.3648 | 2.5475 | 100.79 | 21.15 | 15.982 | 67.1864 | 112.2128 | 64.5734 | 0.003985 | 1.0454 | 0.025415 | 24.3121 | 0.019303 | 0.1298 | -2.4e-05 | -0.001764 | -0.001764 | -3.0381
2019-03-26 16:12 | 1553631100.667 | -79 | 0.20955 | 0.000448 | 0 | 100.177 | 100.1851 | 127.5218 | 21.2612 | 0.001911 | 0.001811 | -152.6882 | 0.21916 | 8.4 | 0.003735 | 0.005229 | 0.005077 | 5.7e-05 | 0.00129 | 0.34058 | 1.3454 | 2.5475 | 100.79 | 21.15 | 15.989 | 67.8019 | 111.811 | 65.3924 | 0.003986 | 1.0428 | 0.025348 | 24.3063 | -0.00072 | 0.13038 | -0.00026 | -0.01985 | -0.01985 | -0.27126
2019-03-26 16:12 | 1553631101.6661 | -78 | 0.20951 | 0.000449 | 0 | 100.2169 | 100.1807 | 127.4098 | 21.2612 | 0.001844 | 0.001786 | -152.6909 | 0.21842 | 8.4037 | 0.003755 | 0.005266 | 0.005144 | 5.5e-05 | 0.001279 | 0.34363 | 1.3533 | 2.5475 | 100.79 | 21.15 | 16.005 | 66.1891 | 112.2792 | 64.7974 | 0.00399 | 1.0447 | 0.025359 | 24.2742 | 0.003228 | 0.12994 | -2.2e-05 | -0.001627 | -0.001627 | -3.298
2019-03-26 16:12 | 1553631102.6661 | -77 | 0.20951 | 0.00045 | -1e-06 | 100.1847 | 100.0869 | 127.3748 | 21.2612 | 0.001906 | 0.001761 | -152.6916 | 0.21926 | 8.4021 | 0.003651 | 0.005191 | 0.005022 | 6.1e-05 | 0.001271 | 0.33859 | 1.3673 | 2.5475 | 100.79 | 21.15 | 16.014 | 67.6761 | 109.7563 | 64.194 | 0.003993 | 1.0465 | 0.025378 | 24.2492 | 0.008733 | 0.13044 | -4.7e-05 | -0.003582 | -0.003582 | -1.5038
2019-03-26 16:12 | 1553631103.6689 | -76 | 0.20953 | 0.00045 | -1e-06 | 100.2603 | 100.1921 | 127.6856 | 21.2612 | 0.001875 | 0.001801 | -152.7131 | 0.21932 | 8.4001 | 0.00366 | 0.004679 | 0.004461 | 5.9e-05 | 0.001272 | 0.3131 | 1.3583 | 2.5475 | 100.79 | 21.14 | 16.032 | 66.9327 | 109.9858 | 65.1518 | 0.003995 | 1.0436 | 0.025373 | 24.3132 | 0.00608 | 0.13047 | -0.000131 | -0.009944 | -0.009944 | -0.54189
2019-03-26 16:12 | 1553631104.6676 | -75 | 0.20952 | 0.000445 | -1e-06 | 100.2481 | 100.2829 | 127.7922 | 21.2612 | 0.00189 | 0.001755 | -152.7096 | 0.21923 | 8.4005 | 0.003634 | 0.005188 | 0.004968 | 5.8e-05 | 0.001278 | 0.35867 | 1.323 | 2.5474 | 100.79 | 21.15 | 16.041 | 67.3126 | 109.3407 | 64.0527 | 0.003999 | 1.047 | 0.025425 | 24.2838 | -0.003099 | 0.13042 | -5.5e-05 | -0.004181 | -0.004181 | -1.2883
2019-03-26 16:12 | 1553631105.6673 | -74 | 0.20949 | 0.000443 | 0 | 100.2414 | 100.1726 | 127.8567 | 21.2612 | 0.001854 | 0.001798 | -152.7392 | 0.21924 | 8.4044 | 0.003742 | 0.005036 | 0.004913 | 5.8e-05 | 0.001278 | 0.33035 | 1.3377 | 2.5475 | 100.79 | 21.15 | 16.05 | 66.4466 | 111.9619 | 65.091 | 0.004002 | 1.0438 | 0.025392 | 24.3273 | 0.006135 | 0.13043 | 0.000109 | 0.00839 | 0.00839 | 0.64204
2019-03-26 16:12 | 1553631106.6663 | -73 | 0.20952 | 0.000445 | 0 | 100.2375 | 100.133 | 128.0128 | 21.2612 | 0.001856 | 0.00176 | -152.7174 | 0.21794 | 8.4011 | 0.003644 | 0.005791 | 0.005591 | 5.9e-05 | 0.00128 | 0.3178 | 1.3232 | 2.5475 | 100.79 | 21.14 | 16.067 | 66.4794 | 109.5799 | 64.1675 | 0.004003 | 1.0466 | 0.025442 | 24.3089 | 0.009335 | 0.12965 | -6.3e-05 | -0.004832 | -0.004832 | -1.1081
2019-03-26 16:12 | 1553631107.6674 | -72 | 0.20946 | 0.000448 | -1e-06 | 100.2508 | 100.2705 | 127.9377 | 21.2612 | 0.001869 | 0.001771 | -152.7431 | 0.21879 | 8.4033 | 0.003702 | 0.004696 | 0.004536 | 5.9e-05 | 0.001279 | 0.30901 | 1.3393 | 2.5475 | 100.79 | 21.14 | 16.074 | 66.7939 | 111.0082 | 64.451 | 0.004005 | 1.0457 | 0.025424 | 24.312 | -0.001753 | 0.13016 | 0.000305 | 0.023442 | 0.023442 | 0.22931
2019-03-26 16:12 | 1553631108.6675 | -71 | 0.20948 | 0.000448 | -1e-06 | 100.2703 | 100.1316 | 128.1429 | 21.2612 | 0.001894 | 0.001794 | -152.7546 | 0.21935 | 8.4021 | 0.003726 | 0.005452 | 0.005271 | 5.9e-05 | 0.001279 | 0.32925 | 1.3239 | 2.5475 | 100.79 | 21.15 | 16.082 | 67.4078 | 111.5835 | 64.9935 | 0.00401 | 1.0441 | 0.025424 | 24.351 | 0.012378 | 0.13049 | 0.000155 | 0.011963 | 0.011963 | 0.45051
2019-03-26 16:12 | 1553631109.6665 | -70 | 0.20951 | 0.000447 | 0 | 100.177 | 100.2923 | 128.2731 | 21.2612 | 0.001882 | 0.001735 | -152.7601 | 0.21928 | 8.403 | 0.003743 | 0.004769 | 0.004621 | 5.8e-05 | 0.001284 | 0.31382 | 1.3911 | 2.5476 | 100.79 | 21.15 | 16.099 | 67.1185 | 111.9831 | 63.5827 | 0.004014 | 1.0484 | 0.02549 | 24.3125 | -0.010289 | 0.13045 | -4.3e-05 | -0.003309 | -0.003309 | -1.6279
2019-03-26 16:12 | 1553631110.6666 | -69 | 0.20948 | 0.000444 | 0 | 100.1808 | 100.1814 | 128.4503 | 21.2612 | 0.001867 | 0.001799 | -152.7591 | 0.21931 | 8.4035 | 0.003744 | 0.00468 | 0.004532 | 5.7e-05 | 0.001286 | 0.3465 | 1.373 | 2.5476 | 100.79 | 21.15 | 16.106 | 66.7384 | 112.021 | 65.1215 | 0.004016 | 1.0437 | 0.02545 | 24.3848 | -4.7e-05 | 0.13047 | 0.000149 | 0.011484 | 0.011484 | 0.4692
2019-03-26 16:12 | 1553631111.6656 | -68 | 0.20946 | 0.000449 | 0 | 100.1741 | 99.9928 | 128.1114 | 21.2612 | 0.001954 | 0.001802 | -152.7521 | 0.21888 | 8.402 | 0.003709 | 0.005267 | 0.005121 | 5.7e-05 | 0.001288 | 0.32964 | 1.382 | 2.5475 | 100.79 | 21.15 | 16.121 | 68.8306 | 111.1636 | 65.1809 | 0.004019 | 1.0435 | 0.025414 | 24.3548 | 0.0162 | 0.13021 | 0.000275 | 0.02115 | 0.02115 | 0.25427
2019-03-26 16:12 | 1553631112.6657 | -67 | 0.20947 | 0.000453 | -1e-06 | 100.1038 | 100.1266 | 127.9356 | 21.2612 | 0.001962 | 0.001771 | -152.7373 | 0.21918 | 8.4019 | 0.003779 | 0.005425 | 0.005208 | 6e-05 | 0.001285 | 0.31094 | 1.4251 | 2.5476 | 100.79 | 21.15 | 16.126 | 69.0262 | 112.8722 | 64.4329 | 0.004021 | 1.0458 | 0.025425 | 24.3111 | -0.002035 | 0.13039 | 0.000198 | 0.015217 | 0.015217 | 0.35388
2019-03-26 16:12 | 1553631113.6668 | -66 | 0.20947 | 0.000456 | 0 | 100.1478 | 100.2378 | 127.6511 | 21.2612 | 0.00193 | 0.001826 | -152.7227 | 0.21985 | 8.4025 | 0.003792 | 0.005758 | 0.005526 | 5.8e-05 | 0.001276 | 0.32748 | 1.4043 | 2.5476 | 100.79 | 21.15 | 16.131 | 68.2659 | 113.1788 | 65.7739 | 0.004022 | 1.0417 | 0.025346 | 24.3323 | -0.00804 | 0.13079 | 0.00018 | 0.013768 | 0.013768 | 0.39233
2019-03-26 16:12 | 1553631114.6667 | -65 | 0.2095 | 0.000452 | 0 | 100.061 | 100.1394 | 127.5664 | 21.2612 | 0.001906 | 0.001763 | -152.7132 | 0.21959 | 8.4047 | 0.0038 | 0.00436 | 0.004206 | 5.8e-05 | 0.001275 | 0.32731 | 1.3758 | 2.5474 | 100.79 | 21.15 | 16.14 | 67.6903 | 113.3756 | 64.2416 | 0.004024 | 1.0464 | 0.025395 | 24.2692 | -0.007001 | 0.13064 | 5.3e-05 | 0.004038 | 0.004038 | 1.3362
2019-03-26 16:12 | 1553631115.6662 | -64 | 0.20952 | 0.000453 | -1e-06 | 100.1041 | 100.1669 | 127.6861 | 21.2612 | 0.001951 | 0.001767 | -152.6952 | 0.22005 | 8.4 | 0.003773 | 0.004598 | 0.004489 | 6e-05 | 0.001276 | 0.30596 | 1.3817 | 2.5474 | 100.79 | 21.15 | 16.144 | 68.7683 | 112.7364 | 64.3525 | 0.004025 | 1.0461 | 0.025403 | 24.2845 | -0.005611 | 0.13091 | -7.8e-05 | -0.005954 | -0.005954 | -0.90807
2019-03-26 16:12 | 1553631116.6672 | -63 | 0.2095 | 0.000454 | -1e-06 | 100.038 | 100.0325 | 127.8703 | 21.2612 | 0.001862 | 0.00183 | -152.7 | 0.21909 | 8.3993 | 0.003738 | 0.005305 | 0.005083 | 5.9e-05 | 0.001277 | 0.32421 | 1.4357 | 2.5475 | 100.79 | 21.15 | 16.152 | 66.6417 | 111.8757 | 65.8699 | 0.004027 | 1.0414 | 0.025364 | 24.3566 | 0.000494 | 0.13034 | 3.2e-05 | 0.002469 | 0.002469 | 2.1806
2019-03-26 16:12 | 1553631117.6663 | -62 | 0.2095 | 0.000453 | -1e-06 | 100.0902 | 100.022 | 128.1403 | 21.2612 | 0.001878 | 0.001783 | -152.7037 | 0.21955 | 8.4007 | 0.003739 | 0.005194 | 0.005052 | 5.9e-05 | 0.00128 | 0.32012 | 1.3819 | 2.5475 | 100.79 | 21.15 | 16.156 | 67.0185 | 111.9081 | 64.7313 | 0.004028 | 1.0449 | 0.025434 | 24.3413 | 0.006096 | 0.13061 | 4.4e-05 | 0.003449 | 0.003449 | 1.5638
2019-03-26 16:12 | 1553631118.667 | -61 | 0.20948 | 0.00045 | 0 | 100.0442 | 100.1474 | 128.3938 | 21.2612 | 0.001911 | 0.001796 | -152.7013 | 0.21814 | 8.4027 | 0.003725 | 0.005221 | 0.005049 | 5.9e-05 | 0.001279 | 0.29789 | 1.322 | 2.5474 | 100.79 | 21.15 | 16.159 | 67.8134 | 111.5529 | 65.0373 | 0.004029 | 1.0439 | 0.025447 | 24.3764 | -0.009226 | 0.12977 | 0.000141 | 0.010846 | 0.010846 | 0.49416
2019-03-26 16:12 | 1553631119.667 | -60 | 0.20949 | 0.000455 | 0 | 100.0141 | 99.905 | 128.3414 | 21.2612 | 0.00191 | 0.001751 | -152.7045 | 0.21865 | 8.4037 | 0.003725 | 0.004567 | 0.0044 | 5.6e-05 | 0.001273 | 0.31957 | 1.3248 | 2.5474 | 100.79 | 21.15 | 16.165 | 67.787 | 111.548 | 63.9716 | 0.00403 | 1.0472 | 0.025482 | 24.333 | 0.009755 | 0.13007 | 5.7e-05 | 0.004416 | 0.004416 | 1.2166
2019-03-26 16:12 | 1553631120.666 | -59 | 0.2095 | 0.000451 | 0 | 100.0899 | 99.9071 | 128.2325 | 21.2612 | 0.001906 | 0.0018 | -152.682 | 0.21908 | 8.4026 | 0.003717 | 0.004217 | 0.004081 | 5.8e-05 | 0.001275 | 0.32908 | 1.3332 | 2.5476 | 100.79 | 21.15 | 16.166 | 67.6949 | 111.3524 | 65.1504 | 0.004031 | 1.0436 | 0.025427 | 24.3652 | 0.016357 | 0.13033 | 5e-06 | 0.000409 | 0.000409 | 13.1702
2019-03-26 16:12 | 1553631121.6671 | -58 | 0.20951 | 0.000451 | 0 | 100.0613 | 99.935 | 128.4841 | 21.2612 | 0.001929 | 0.001749 | -152.6977 | 0.21907 | 8.4016 | 0.0037 | 0.004745 | 0.004534 | 5.8e-05 | 0.001259 | 0.33876 | 1.3768 | 2.5474 | 100.79 | 21.15 | 16.163 | 68.2471 | 110.9411 | 63.9059 | 0.00403 | 1.0474 | 0.025499 | 24.3442 | 0.0113 | 0.13032 | -1.3e-05 | -0.001002 | -0.001002 | -5.3723
2019-03-26 16:12 | 1553631122.6662 | -57 | 0.2095 | 0.000446 | -1e-06 | 100.0126 | 99.8869 | 128.5264 | 21.2612 | 0.0019 | 0.001775 | -152.715 | 0.21832 | 8.4012 | 0.003574 | 0.005178 | 0.004991 | 5.9e-05 | 0.001243 | 0.30702 | 1.3578 | 2.5476 | 100.79 | 21.15 | 16.158 | 67.5357 | 107.8846 | 64.533 | 0.004029 | 1.0455 | 0.02548 | 24.3708 | 0.011249 | 0.12988 | 5.1e-05 | 0.003932 | 0.003932 | 1.3643
2019-03-26 16:12 | 1553631123.6662 | -56 | 0.20949 | 0.000448 | -1e-06 | 99.9993 | 99.8933 | 128.4414 | 21.2612 | 0.001881 | 0.001796 | -152.7019 | 0.21935 | 8.4029 | 0.003588 | 0.005569 | 0.00544 | 5.8e-05 | 0.001234 | 0.34274 | 1.3402 | 2.5476 | 100.79 | 21.15 | 16.152 | 67.0818 | 108.2177 | 65.0388 | 0.004027 | 1.0439 | 0.025452 | 24.381 | 0.009489 | 0.13049 | 9.9e-05 | 0.007667 | 0.007667 | 0.70293
2019-03-26 16:12 | 1553631124.6676 | -55 | 0.20949 | 0.000449 | -1e-06 | 100.0108 | 100.0187 | 128.5238 | 21.2612 | 0.001927 | 0.001749 | -152.7146 | 0.21914 | 8.4016 | 0.003506 | 0.005125 | 0.004969 | 5.7e-05 | 0.001227 | 0.33754 | 1.376 | 2.5475 | 100.79 | 21.16 | 16.136 | 68.1942 | 106.2287 | 63.9114 | 0.004026 | 1.0474 | 0.025503 | 24.3481 | -0.000702 | 0.13037 | 0.000125 | 0.009659 | 0.009659 | 0.55741
2019-03-26 16:12 | 1553631125.6667 | -54 | 0.20949 | 0.000442 | 0 | 100.0164 | 99.821 | 128.4201 | 21.2612 | 0.001904 | 0.001828 | -152.6956 | 0.21931 | 8.4032 | 0.003575 | 0.005245 | 0.00508 | 5.8e-05 | 0.001228 | 0.33323 | 1.3663 | 2.5475 | 100.79 | 21.16 | 16.127 | 67.6466 | 107.9067 | 65.8122 | 0.004023 | 1.0415 | 0.025421 | 24.4068 | 0.017494 | 0.13046 | 0.00012 | 0.009262 | 0.009262 | 0.58179
2019-03-26 16:12 | 1553631126.6667 | -53 | 0.20953 | 0.000443 | -1e-06 | 99.9724 | 99.9017 | 128.0539 | 21.2612 | 0.001877 | 0.001774 | -152.7062 | 0.21911 | 8.4041 | 0.003593 | 0.004891 | 0.004656 | 5.7e-05 | 0.001226 | 0.32892 | 1.3403 | 2.5475 | 100.79 | 21.16 | 16.106 | 66.9788 | 108.3596 | 64.5251 | 0.004018 | 1.0455 | 0.025433 | 24.3257 | 0.006333 | 0.13035 | -0.000149 | -0.011363 | -0.011363 | -0.47377
2019-03-26 16:12 | 1553631127.6664 | -52 | 0.20951 | 0.000445 | -1e-06 | 99.9843 | 100.0399 | 128.0091 | 21.2612 | 0.001914 | 0.001764 | -152.69 | 0.21957 | 8.4026 | 0.003612 | 0.005078 | 0.004956 | 5.8e-05 | 0.001234 | 0.31365 | 1.3796 | 2.5476 | 100.79 | 21.16 | 16.095 | 67.8899 | 108.8181 | 64.2777 | 0.004015 | 1.0463 | 0.025438 | 24.3125 | -0.004972 | 0.13062 | -3.3e-05 | -0.002496 | -0.002496 | -2.1612
2019-03-26 16:12 | 1553631128.6665 | -51 | 0.2095 | 0.000443 | 0 | 99.994 | 99.8519 | 127.8741 | 21.2612 | 0.001902 | 0.001825 | -152.6923 | 0.21883 | 8.4024 | 0.003564 | 0.005176 | 0.004915 | 5.8e-05 | 0.001237 | 0.32101 | 1.3732 | 2.5475 | 100.79 | 21.16 | 16.083 | 67.592 | 107.6474 | 65.7489 | 0.004012 | 1.0417 | 0.025369 | 24.3526 | 0.012716 | 0.13018 | 4.8e-05 | 0.003735 | 0.003735 | 1.4394
2019-03-26 16:12 | 1553631129.6659 | -50 | 0.20951 | 0.000444 | -1e-06 | 99.9603 | 100.0556 | 127.9473 | 21.2612 | 0.001879 | 0.001782 | -152.6916 | 0.2189 | 8.4017 | 0.003563 | 0.004203 | 0.004032 | 5.8e-05 | 0.001236 | 0.31542 | 1.3335 | 2.5475 | 100.79 | 21.17 | 16.061 | 67.0428 | 107.6316 | 64.707 | 0.004009 | 1.045 | 0.025416 | 24.3221 | -0.008524 | 0.13023 | -2.6e-05 | -0.001936 | -0.001936 | -2.778
2019-03-26 16:12 | 1553631130.6666 | -49 | 0.20953 | 0.000443 | -1e-06 | 99.9955 | 100.0109 | 128.0039 | 21.2612 | 0.001882 | 0.001777 | -152.6966 | 0.21949 | 8.4008 | 0.003467 | 0.004744 | 0.004559 | 5.8e-05 | 0.001239 | 0.33892 | 1.3635 | 2.5475 | 100.79 | 21.17 | 16.05 | 67.1137 | 105.2812 | 64.5744 | 0.004007 | 1.0454 | 0.025426 | 24.3227 | -0.001383 | 0.13058 | -0.000126 | -0.009642 | -0.009642 | -0.55932
2019-03-26 16:12 | 1553631131.6655 | -48 | 0.20951 | 0.000443 | -1e-06 | 100.002 | 100.0593 | 128.1114 | 21.2612 | 0.001897 | 0.00179 | -152.7146 | 0.21877 | 8.4041 | 0.003463 | 0.004769 | 0.004573 | 5.9e-05 | 0.001235 | 0.36066 | 1.3858 | 2.5476 | 100.79 | 21.17 | 16.038 | 67.4836 | 105.1869 | 64.9076 | 0.004004 | 1.0443 | 0.025424 | 24.3449 | -0.005131 | 0.13014 | 6e-06 | 0.000521 | 0.000521 | 10.3224
2019-03-26 16:12 | 1553631132.6655 | -47 | 0.20948 | 0.00044 | 0 | 99.9376 | 99.8418 | 128.1431 | 21.2612 | 0.001895 | 0.001774 | -152.7215 | 0.21965 | 8.4042 | 0.003473 | 0.004918 | 0.004713 | 5.6e-05 | 0.001226 | 0.30663 | 1.4128 | 2.5476 | 100.79 | 21.17 | 16.017 | 67.4179 | 105.4245 | 64.5041 | 0.003998 | 1.0456 | 0.025443 | 24.3334 | 0.008575 | 0.13067 | 0.000156 | 0.011978 | 0.011978 | 0.45054
2019-03-26 16:12 | 1553631133.6661 | -46 | 0.2095 | 0.000441 | -1e-06 | 99.8963 | 100.0163 | 128.3231 | 21.2612 | 0.001871 | 0.001795 | -152.7497 | 0.21925 | 8.4008 | 0.003479 | 0.005003 | 0.004799 | 5.9e-05 | 0.001219 | 0.29391 | 1.3892 | 2.5476 | 100.79 | 21.17 | 16.006 | 66.8509 | 105.5903 | 65.0207 | 0.003996 | 1.044 | 0.025441 | 24.3691 | -0.010741 | 0.13043 | 1.4e-05 | 0.001117 | 0.001117 | 4.8214
2019-03-26 16:12 | 1553631134.6667 | -45 | 0.20951 | 0.000447 | -1e-06 | 99.8313 | 99.8519 | 128.4546 | 21.2612 | 0.001923 | 0.00182 | -152.7434 | 0.21853 | 8.4041 | 0.003511 | 0.005355 | 0.005138 | 6.3e-05 | 0.001225 | 0.32875 | 1.3432 | 2.5476 | 100.79 | 21.17 | 15.985 | 68.1067 | 106.3553 | 65.6219 | 0.00399 | 1.0421 | 0.025431 | 24.4033 | -0.001844 | 0.13001 | -2.1e-05 | -0.001607 | -0.001607 | -3.3422
2019-03-26 16:12 | 1553631135.6672 | -44 | 0.2095 | 0.000444 | -1e-06 | 99.8925 | 99.8872 | 128.5703 | 21.2612 | 0.001927 | 0.001818 | -152.7524 | 0.21909 | 8.4017 | 0.003571 | 0.005085 | 0.004944 | 5.6e-05 | 0.001222 | 0.31421 | 1.3705 | 2.5476 | 100.79 | 21.17 | 15.975 | 68.1856 | 107.8253 | 65.575 | 0.003988 | 1.0423 | 0.025445 | 24.4126 | 0.000467 | 0.13034 | 3.1e-05 | 0.002456 | 0.002456 | 2.1915
2019-03-26 16:12 | 1553631136.6659 | -43 | 0.20948 | 0.00044 | 0 | 99.9066 | 99.8096 | 128.8057 | 21.2612 | 0.001916 | 0.001784 | -152.7464 | 0.21965 | 8.4002 | 0.003508 | 0.00482 | 0.004585 | 5.8e-05 | 0.001224 | 0.33035 | 1.3443 | 2.5473 | 100.79 | 21.17 | 15.965 | 67.9228 | 106.2908 | 64.7585 | 0.003985 | 1.0448 | 0.025499 | 24.4054 | 0.008692 | 0.13067 | 0.000151 | 0.01166 | 0.01166 | 0.46284
2019-03-26 16:12 | 1553631137.666 | -42 | 0.20949 | 0.000444 | -1e-06 | 99.9105 | 99.8345 | 128.7125 | 21.2612 | 0.00188 | 0.001807 | -152.7455 | 0.21948 | 8.4032 | 0.003526 | 0.004568 | 0.004441 | 5.7e-05 | 0.00122 | 0.30879 | 1.3524 | 2.5475 | 100.79 | 21.17 | 15.946 | 67.0647 | 106.7183 | 65.3159 | 0.003981 | 1.0431 | 0.025468 | 24.4167 | 0.006807 | 0.13057 | 0.000109 | 0.008391 | 0.008391 | 0.64268
2019-03-26 16:12 | 1553631138.6669 | -41 | 0.20952 | 0.000447 | -1e-06 | 99.893 | 99.9444 | 128.4492 | 21.2612 | 0.00189 | 0.001802 | -152.7382 | 0.21902 | 8.4023 | 0.003616 | 0.005595 | 0.005457 | 5.6e-05 | 0.001219 | 0.32925 | 1.3515 | 2.5474 | 100.79 | 21.17 | 15.935 | 67.2948 | 108.9182 | 65.1817 | 0.003978 | 1.0435 | 0.025447 | 24.3869 | -0.004596 | 0.13029 | -8.7e-05 | -0.006657 | -0.006657 | -0.80834
2019-03-26 16:12 | 1553631139.667 | -40 | 0.20952 | 0.000443 | 0 | 99.9592 | 100.0049 | 127.9537 | 21.2612 | 0.001919 | 0.001743 | -152.729 | 0.21988 | 8.4032 | 0.003608 | 0.004966 | 0.004732 | 5.8e-05 | 0.001222 | 0.33682 | 1.4168 | 2.5476 | 100.79 | 21.18 | 15.913 | 68.0108 | 108.7064 | 63.7569 | 0.003975 | 1.0479 | 0.025452 | 24.2885 | -0.004091 | 0.13081 | -9.2e-05 | -0.007068 | -0.007068 | -0.76431
2019-03-26 16:12 | 1553631140.6654 | -39 | 0.20952 | 0.000441 | -1e-06 | 99.9423 | 99.8345 | 127.9321 | 21.2612 | 0.001899 | 0.001781 | -152.7264 | 0.21865 | 8.4031 | 0.003559 | 0.004724 | 0.004524 | 6e-05 | 0.001227 | 0.32997 | 1.4389 | 2.5475 | 100.79 | 21.17 | 15.9 | 67.5148 | 107.5105 | 64.6735 | 0.003969 | 1.0451 | 0.025415 | 24.3195 | 0.009661 | 0.13008 | -5.7e-05 | -0.004337 | -0.004337 | -1.2386
2019-03-26 16:12 | 1553631141.6651 | -38 | 0.20953 | 0.000445 | -1e-06 | 99.9648 | 99.9414 | 127.8778 | 21.2612 | 0.001865 | 0.001786 | -152.7252 | 0.21983 | 8.4025 | 0.003601 | 0.004736 | 0.004582 | 5.7e-05 | 0.001228 | 0.30702 | 1.3811 | 2.5476 | 100.79 | 21.18 | 15.886 | 66.6963 | 108.5495 | 64.7991 | 0.003968 | 1.0447 | 0.025405 | 24.3188 | 0.002096 | 0.13078 | -0.000133 | -0.010161 | -0.010161 | -0.53156
2019-03-26 16:12 | 1553631142.6655 | -37 | 0.20952 | 0.00044 | 0 | 99.9031 | 100.0009 | 128.0771 | 21.2612 | 0.001881 | 0.001833 | -152.7165 | 0.21949 | 8.4047 | 0.003577 | 0.004975 | 0.004805 | 5.8e-05 | 0.001232 | 0.34954 | 1.4006 | 2.5476 | 100.79 | 21.18 | 15.86 | 67.0755 | 107.9671 | 65.9276 | 0.003962 | 1.0412 | 0.025383 | 24.3784 | -0.00875 | 0.13057 | -0.000103 | -0.007871 | -0.007871 | -0.68511
2019-03-26 16:12 | 1553631143.6652 | -36 | 0.20952 | 0.000444 | -1e-06 | 99.9465 | 99.7192 | 127.9687 | 21.2612 | 0.001889 | 0.001755 | -152.7063 | 0.21899 | 8.4037 | 0.003613 | 0.004975 | 0.004858 | 5.6e-05 | 0.001229 | 0.33483 | 1.3897 | 2.5476 | 100.79 | 21.18 | 15.846 | 67.2847 | 108.8441 | 64.0551 | 0.003958 | 1.047 | 0.025442 | 24.3007 | 0.020367 | 0.13028 | -0.000103 | -0.007845 | -0.007845 | -0.68588
2019-03-26 16:12 | 1553631144.6652 | -35 | 0.20952 | 0.000444 | 0 | 99.9007 | 99.8368 | 127.6056 | 21.2612 | 0.001933 | 0.001804 | -152.7028 | 0.22004 | 8.4024 | 0.003599 | 0.004596 | 0.004401 | 5.9e-05 | 0.001228 | 0.30287 | 1.3693 | 2.5476 | 100.79 | 21.18 | 15.833 | 68.3354 | 108.4999 | 65.2394 | 0.003955 | 1.0433 | 0.025362 | 24.3088 | 0.005724 | 0.1309 | -6.2e-05 | -0.004753 | -0.004753 | -1.1373
2019-03-26 16:12 | 1553631145.6653 | -34 | 0.20951 | 0.000439 | -1e-06 | 99.9391 | 100.0536 | 127.4833 | 21.2612 | 0.001886 | 0.001755 | -152.7062 | 0.22022 | 8.4028 | 0.003687 | 0.005217 | 0.005088 | 6e-05 | 0.001231 | 0.33268 | 1.3495 | 2.5476 | 100.79 | 21.17 | 15.821 | 67.209 | 110.6225 | 64.0628 | 0.003949 | 1.047 | 0.025394 | 24.2548 | -0.010245 | 0.13101 | -3.7e-05 | -0.002793 | -0.002793 | -1.9369
2019-03-26 16:12 | 1553631146.6667 | -33 | 0.20953 | 0.000444 | 0 | 99.9639 | 99.736 | 127.3983 | 21.2612 | 0.001907 | 0.001802 | -152.7056 | 0.21913 | 8.3994 | 0.003659 | 0.005052 | 0.004836 | 5.2e-05 | 0.001228 | 0.35059 | 1.3414 | 2.5474 | 100.79 | 21.18 | 15.809 | 67.7081 | 109.9527 | 65.1793 | 0.003949 | 1.0435 | 0.025343 | 24.2868 | 0.020418 | 0.13036 | -0.00014 | -0.010671 | -0.010671 | -0.50451
2019-03-26 16:12 | 1553631147.6658 | -32 | 0.20952 | 0.00044 | 0 | 99.9828 | 99.8825 | 127.2556 | 21.2612 | 0.001877 | 0.001802 | -152.7096 | 0.21937 | 8.4045 | 0.003668 | 0.00529 | 0.005194 | 5.5e-05 | 0.001225 | 0.32516 | 1.3587 | 2.5475 | 100.79 | 21.18 | 15.788 | 67.0022 | 110.1657 | 65.1831 | 0.003944 | 1.0435 | 0.025329 | 24.2734 | 0.008973 | 0.1305 | -9.3e-05 | -0.007123 | -0.007123 | -0.75671
2019-03-26 16:12 | 1553631148.6655 | -31 | 0.20954 | 0.000447 | 0 | 99.9686 | 99.7813 | 127.5863 | 21.2612 | 0.001915 | 0.001779 | -152.697 | 0.2179 | 8.4023 | 0.003718 | 0.004651 | 0.00444 | 5.6e-05 | 0.001224 | 0.32637 | 1.3302 | 2.5475 | 100.79 | 21.18 | 15.778 | 67.8954 | 111.3967 | 64.636 | 0.003941 | 1.0452 | 0.025382 | 24.2852 | 0.016771 | 0.12963 | -0.000185 | -0.014122 | -0.014122 | -0.37911
2019-03-26 16:12 | 1553631149.6665 | -30 | 0.20948 | 0.000445 | 0 | 100.0383 | 99.8553 | 127.8673 | 21.2612 | 0.001894 | 0.001808 | -152.7008 | 0.2193 | 8.4048 | 0.003661 | 0.005375 | 0.005185 | 5.8e-05 | 0.001224 | 0.31222 | 1.3269 | 2.5474 | 100.79 | 21.18 | 15.764 | 67.397 | 109.9931 | 65.3323 | 0.003938 | 1.043 | 0.025384 | 24.337 | 0.016375 | 0.13046 | 0.000188 | 0.014443 | 0.014443 | 0.37305
2019-03-26 16:12 | 1553631150.6669 | -29 | 0.20952 | 0.000443 | 0 | 100.0011 | 99.9488 | 128.1346 | 21.2612 | 0.001913 | 0.001744 | -152.7222 | 0.21983 | 8.4034 | 0.003668 | 0.005362 | 0.005187 | 5.9e-05 | 0.001221 | 0.347 | 1.3267 | 2.5475 | 100.79 | 21.18 | 15.758 | 67.8565 | 110.1712 | 63.7897 | 0.003936 | 1.0478 | 0.025469 | 24.3069 | 0.004682 | 0.13078 | -4.7e-05 | -0.003631 | -0.003631 | -1.4875
2019-03-26 16:12 | 1553631151.6656 | -28 | 0.20954 | 0.000445 | 0 | 99.9881 | 99.9608 | 128.2583 | 21.2612 | 0.001926 | 0.001754 | -152.7018 | 0.21949 | 8.4005 | 0.003717 | 0.005056 | 0.004875 | 5.5e-05 | 0.001219 | 0.30525 | 1.3275 | 2.5474 | 100.79 | 21.18 | 15.754 | 68.1808 | 111.3675 | 64.0286 | 0.003935 | 1.0471 | 0.025472 | 24.3272 | 0.002437 | 0.13057 | -0.000186 | -0.014314 | -0.014314 | -0.37673
2019-03-26 16:13 | 1553631152.6667 | -27 | 0.20951 | 0.000446 | 0 | 100.0034 | 99.7723 | 127.9837 | 21.2612 | 0.001857 | 0.001791 | -152.688 | 0.21873 | 8.4015 | 0.00372 | 0.005121 | 0.004921 | 5.5e-05 | 0.001218 | 0.32764 | 1.3389 | 2.5475 | 100.79 | 21.18 | 15.749 | 66.5176 | 111.4234 | 64.924 | 0.003934 | 1.0443 | 0.025411 | 24.3334 | 0.020701 | 0.13012 | -1.9e-05 | -0.001468 | -0.001468 | -3.661
2019-03-26 16:13 | 1553631153.6667 | -26 | 0.20954 | 0.000447 | -1e-06 | 99.9751 | 99.939 | 127.9471 | 21.2612 | 0.001888 | 0.001831 | -152.7018 | 0.21943 | 8.4036 | 0.003642 | 0.004912 | 0.004655 | 5.6e-05 | 0.001218 | 0.32549 | 1.3964 | 2.5475 | 100.79 | 21.18 | 15.749 | 67.265 | 109.5286 | 65.8824 | 0.003934 | 1.0413 | 0.025371 | 24.3644 | 0.003231 | 0.13054 | -0.000222 | -0.017008 | -0.017008 | -0.31698
2019-03-26 16:13 | 1553631154.6661 | -25 | 0.20953 | 0.000448 | -1e-06 | 100.0386 | 99.9733 | 128.2827 | 21.2612 | 0.001903 | 0.001831 | -152.6977 | 0.21955 | 8.3993 | 0.003718 | 0.005401 | 0.005201 | 5.7e-05 | 0.001214 | 0.29087 | 1.346 | 2.5475 | 100.79 | 21.18 | 15.752 | 67.6254 | 111.3745 | 65.8894 | 0.003935 | 1.0413 | 0.025404 | 24.3966 | 0.005839 | 0.13061 | -0.000139 | -0.010623 | -0.010623 | -0.50778
2019-03-26 16:13 | 1553631155.6672 | -24 | 0.20952 | 0.000446 | -1e-06 | 100.0232 | 100.0399 | 128.2166 | 21.2612 | 0.001931 | 0.001786 | -152.6956 | 0.21987 | 8.3977 | 0.003709 | 0.005102 | 0.004952 | 5.6e-05 | 0.001218 | 0.33212 | 1.339 | 2.5477 | 100.79 | 21.18 | 15.754 | 68.2861 | 111.1661 | 64.7943 | 0.003935 | 1.0447 | 0.025439 | 24.3509 | -0.001487 | 0.1308 | -5.7e-05 | -0.004294 | -0.004294 | -1.2581
2019-03-26 16:13 | 1553631156.6667 | -23 | 0.20952 | 0.000443 | 0 | 100.0492 | 100.2627 | 128.2137 | 21.2612 | 0.001944 | 0.00178 | -152.7034 | 0.21994 | 8.404 | 0.003651 | 0.004931 | 0.004769 | 5.6e-05 | 0.00122 | 0.31277 | 1.383 | 2.5476 | 100.79 | 21.18 | 15.759 | 68.6113 | 109.7485 | 64.6492 | 0.003936 | 1.0451 | 0.025444 | 24.3454 | -0.019069 | 0.13084 | -4.8e-05 | -0.003653 | -0.003653 | -1.4792
2019-03-26 16:13 | 1553631157.6658 | -22 | 0.20953 | 0.000443 | 0 | 100.0253 | 99.8734 | 128.256 | 21.2612 | 0.001906 | 0.001774 | -152.7015 | 0.21993 | 8.4034 | 0.003644 | 0.004931 | 0.004783 | 5.4e-05 | 0.001227 | 0.33539 | 1.4217 | 2.5477 | 100.79 | 21.18 | 15.771 | 67.6802 | 109.5964 | 64.5205 | 0.003939 | 1.0455 | 0.025453 | 24.3447 | 0.013588 | 0.13084 | -0.000168 | -0.012917 | -0.012917 | -0.41835
2019-03-26 16:13 | 1553631158.6649 | -21 | 0.20953 | 0.000445 | -1e-06 | 99.9925 | 100.0177 | 128.3928 | 21.2612 | 0.00194 | 0.001716 | -152.7177 | 0.21922 | 8.4035 | 0.003622 | 0.004994 | 0.004738 | 5.5e-05 | 0.001232 | 0.31669 | 1.4406 | 2.5477 | 100.79 | 21.18 | 15.779 | 68.5133 | 109.0492 | 63.122 | 0.003941 | 1.0499 | 0.02552 | 24.3072 | -0.002249 | 0.13042 | -0.000136 | -0.010437 | -0.010437 | -0.51606
2019-03-26 16:13 | 1553631159.666 | -20 | 0.20951 | 0.000442 | 0 | 100.0554 | 99.9659 | 128.8156 | 21.2612 | 0.001935 | 0.001807 | -152.7118 | 0.21852 | 8.4027 | 0.003596 | 0.004062 | 0.003901 | 5.1e-05 | 0.001232 | 0.323 | 1.3714 | 2.5475 | 100.79 | 21.18 | 15.793 | 68.3743 | 108.4108 | 65.3128 | 0.003945 | 1.0431 | 0.025479 | 24.4264 | 0.008005 | 0.13 | -2.5e-05 | -0.001923 | -0.001923 | -2.7918
2019-03-26 16:13 | 1553631160.6651 | -19 | 0.20953 | 0.000443 | -1e-06 | 100.0197 | 99.8876 | 129.0179 | 21.2612 | 0.001859 | 0.00175 | -152.7036 | 0.21939 | 8.4021 | 0.003549 | 0.004526 | 0.004307 | 5.1e-05 | 0.001236 | 0.33229 | 1.3733 | 2.5476 | 100.79 | 21.18 | 15.8 | 66.5527 | 107.2796 | 63.9341 | 0.003947 | 1.0474 | 0.025551 | 24.3957 | 0.011822 | 0.13051 | -0.000108 | -0.008327 | -0.008327 | -0.64736
2019-03-26 16:13 | 1553631161.6661 | -18 | 0.20952 | 0.000448 | -1e-06 | 100.0303 | 99.9998 | 129.1404 | 21.2612 | 0.001922 | 0.001754 | -152.708 | 0.21985 | 8.4001 | 0.003578 | 0.004826 | 0.004635 | 5.5e-05 | 0.001237 | 0.32659 | 1.4184 | 2.5474 | 100.79 | 21.18 | 15.805 | 68.0817 | 107.9798 | 64.0363 | 0.003948 | 1.047 | 0.025559 | 24.411 | 0.002724 | 0.13079 | -5.8e-05 | -0.004423 | -0.004423 | -1.2213
2019-03-26 16:13 | 1553631162.6655 | -17 | 0.20953 | 0.000443 | 0 | 100.0427 | 100.1511 | 129.4017 | 21.2612 | 0.001934 | 0.001742 | -152.7183 | 0.21929 | 8.4023 | 0.003573 | 0.005588 | 0.005388 | 5.7e-05 | 0.001237 | 0.30453 | 1.3971 | 2.5475 | 100.79 | 21.17 | 15.81 | 68.3587 | 107.8664 | 63.7427 | 0.003947 | 1.0479 | 0.025596 | 24.425 | -0.009688 | 0.13045 | -0.000116 | -0.0089 | -0.0089 | -0.60539
2019-03-26 16:13 | 1553631163.6662 | -16 | 0.20951 | 0.000445 | 0 | 100.0377 | 100.0099 | 129.9155 | 21.2612 | 0.001874 | 0.00173 | -152.7191 | 0.21798 | 8.4037 | 0.003611 | 0.004826 | 0.004695 | 5.6e-05 | 0.00124 | 0.34811 | 1.3527 | 2.5475 | 100.79 | 21.18 | 15.81 | 66.9086 | 108.7827 | 63.4484 | 0.003949 | 1.0489 | 0.025658 | 24.4628 | 0.002482 | 0.12968 | -1.1e-05 | -0.000894 | -0.000894 | -5.9922
2019-03-26 16:13 | 1553631164.6666 | -15 | 0.20952 | 0.00044 | -1e-06 | 100.0805 | 100.0671 | 129.9725 | 21.2612 | 0.001859 | 0.001773 | -152.727 | 0.21899 | 8.4 | 0.003616 | 0.005188 | 0.004929 | 5.3e-05 | 0.001244 | 0.32029 | 1.3596 | 2.5475 | 100.79 | 21.17 | 15.807 | 66.5513 | 108.9049 | 64.4774 | 0.003946 | 1.0457 | 0.025625 | 24.5055 | 0.001198 | 0.13028 | -7.6e-05 | -0.005852 | -0.005852 | -0.91943
2019-03-26 16:13 | 1553631165.6663 | -14 | 0.20952 | 0.000447 | 0 | 100.0713 | 99.9017 | 130.5671 | 21.2612 | 0.001839 | 0.001767 | -152.7352 | 0.21847 | 8.4044 | 0.003582 | 0.004877 | 0.004739 | 5.5e-05 | 0.001249 | 0.35513 | 1.3082 | 2.5474 | 100.79 | 21.17 | 15.805 | 66.0767 | 108.0766 | 64.3448 | 0.003945 | 1.0461 | 0.025688 | 24.5567 | 0.015176 | 0.12997 | -7.7e-05 | -0.005958 | -0.005958 | -0.90092
2019-03-26 16:13 | 1553631166.6654 | -13 | 0.20952 | 0.000439 | -1e-06 | 100.0058 | 99.9098 | 131.0734 | 21.2612 | 0.001901 | 0.001758 | -152.7532 | 0.21907 | 8.4029 | 0.003559 | 0.005347 | 0.00516 | 5.2e-05 | 0.001243 | 0.3246 | 1.3285 | 2.5475 | 100.79 | 21.17 | 15.801 | 67.5615 | 107.534 | 64.1386 | 0.003944 | 1.0467 | 0.025746 | 24.5967 | 0.008595 | 0.13033 | -3.9e-05 | -0.002998 | -0.002998 | -1.7953
2019-03-26 16:13 | 1553631167.6667 | -12 | 0.20955 | 0.000445 | -1e-06 | 100.0244 | 100.06 | 131.1311 | 21.2612 | 0.001893 | 0.00171 | -138.2533 | 0.21966 | 8.4036 | 0.003592 | 0.004461 | 0.004303 | 5.7e-05 | 0.001245 | 0.29894 | 1.3633 | 2.5474 | 100.79 | 21.17 | 15.791 | 67.3703 | 108.3154 | 62.9751 | 0.003942 | 1.0503 | 0.025796 | 24.5597 | -0.003185 | 0.13067 | -0.000265 | -0.020629 | -0.020629 | -0.26162
2019-03-26 16:13 | 1553631168.6656 | -11 | 0.20952 | 0.000439 | -1e-06 | 100.0294 | 100.139 | 130.883 | 21.2612 | 0.001846 | 0.001712 | -96.2031 | 0.21989 | 8.4046 | 0.003583 | 0.005173 | 0.004946 | 5.5e-05 | 0.001241 | 0.30215 | 1.3625 | 2.5475 | 100.79 | 21.17 | 15.785 | 66.2571 | 108.1089 | 63.0103 | 0.00394 | 1.0502 | 0.02577 | 24.5377 | -0.009795 | 0.13081 | -7.8e-05 | -0.006064 | -0.006064 | -0.89094
2019-03-26 16:13 | 1553631169.6665 | -10 | 0.20953 | 0.000442 | -1e-06 | 100.0793 | 100.0291 | 130.7161 | 21.2612 | 0.001928 | 0.001702 | -54.1124 | 0.21959 | 8.4029 | 0.003682 | 0.005547 | 0.005359 | 5.2e-05 | 0.001244 | 0.3178 | 1.3719 | 2.5475 | 100.79 | 21.17 | 15.778 | 68.2075 | 110.5143 | 62.7914 | 0.003939 | 1.0509 | 0.025762 | 24.5141 | 0.004489 | 0.13064 | -0.000116 | -0.009014 | -0.009014 | -0.59852
2019-03-26 16:13 | 1553631170.6659 | -9 | 0.20951 | 0.000439 | -1e-06 | 100.0516 | 99.7981 | 130.5307 | 21.2612 | 0.001847 | 0.001734 | -12.0392 | 0.21884 | 8.4037 | 0.003626 | 0.004666 | 0.004544 | 5.1e-05 | 0.001242 | 0.31636 | 1.3711 | 2.5475 | 100.79 | 21.17 | 15.762 | 66.2624 | 109.143 | 63.5461 | 0.003935 | 1.0486 | 0.025715 | 24.5242 | 0.022683 | 0.13019 | -1.7e-05 | -0.001285 | -0.001285 | -4.184
2019-03-26 16:13 | 1553631171.6656 | -8 | 0.20953 | 0.000441 | 0 | 100.0651 | 99.9824 | 130.2983 | 21.2612 | 0.001873 | 0.001713 | 30.0517 | 0.21928 | 8.4037 | 0.003673 | 0.004691 | 0.004527 | 5.1e-05 | 0.001238 | 0.34108 | 1.3739 | 2.5475 | 100.79 | 21.17 | 15.753 | 66.8891 | 110.2911 | 63.0478 | 0.003932 | 1.0501 | 0.025711 | 24.4842 | 0.007402 | 0.13045 | -0.000155 | -0.012056 | -0.012056 | -0.44689
2019-03-26 16:13 | 1553631172.666 | -7 | 0.20954 | 0.000432 | 0 | 100.0262 | 100.1309 | 130.1868 | 21.2612 | 0.001892 | 0.001702 | 72.1329 | 0.23034 | 8.4023 | 0.003707 | 0.004959 | 0.004798 | 5e-05 | 0.001228 | 0.31205 | 1.3731 | 2.5475 | 100.79 | 21.17 | 15.735 | 67.3508 | 111.1116 | 62.7844 | 0.003928 | 1.0509 | 0.02571 | 24.4642 | -0.009365 | 0.13703 | -0.000159 | -0.012313 | -0.012313 | -0.45963
2019-03-26 16:13 | 1553631173.6647 | -6 | 0.20954 | 0.000436 | -1e-06 | 100.0338 | 99.7871 | 129.9195 | 21.2612 | 0.00181 | 0.001718 | 114.2081 | 0.25049 | 8.4028 | 0.003835 | 0.005246 | 0.005173 | 5.1e-05 | 0.001231 | 0.33339 | 1.409 | 2.5476 | 100.79 | 21.17 | 15.727 | 65.3862 | 114.2253 | 63.1554 | 0.003926 | 1.0498 | 0.02567 | 24.4525 | 0.022094 | 0.14901 | -0.000202 | -0.015593 | -0.015593 | -0.39469
2019-03-26 16:13 | 1553631174.6647 | -5 | 0.20953 | 0.000441 | -1e-06 | 100.1372 | 100.0392 | 129.8906 | 21.2612 | 0.00183 | 0.001684 | 268.2392 | 0.25155 | 8.4029 | 0.003834 | 0.005068 | 0.004944 | 5.2e-05 | 0.001232 | 0.32372 | 1.3295 | 2.5475 | 100.79 | 21.17 | 15.72 | 65.8677 | 114.2133 | 62.3564 | 0.003924 | 1.0523 | 0.025697 | 24.4207 | 0.008757 | 0.14965 | -0.000128 | -0.009882 | -0.009882 | -0.62544
2019-03-26 16:13 | 1553631175.6651 | -4 | 0.2095 | 0.000439 | -1e-06 | 100.0613 | 100.0607 | 129.3777 | 21.2612 | 0.00184 | 0.001723 | 267.9632 | 0.25135 | 8.4032 | 0.003967 | 0.005119 | 0.00503 | 5.4e-05 | 0.001239 | 0.31669 | 1.3479 | 2.5476 | 100.79 | 21.17 | 15.711 | 66.1025 | 117.4451 | 63.2897 | 0.003922 | 1.0494 | 0.025611 | 24.4064 | 5.4e-05 | 0.14953 | 2.1e-05 | 0.001695 | 0.001695 | 3.6428
2019-03-26 16:13 | 1553631176.6658 | -3 | 0.20954 | 0.000443 | 0 | 100.0214 | 99.8271 | 129.0181 | 21.2612 | 0.001797 | 0.001731 | 268.1744 | 0.25182 | 8.4065 | 0.003966 | 0.005203 | 0.005209 | 5.3e-05 | 0.001242 | 0.30342 | 1.3484 | 2.5475 | 100.79 | 21.17 | 15.707 | 65.0669 | 117.43 | 63.478 | 0.003921 | 1.0488 | 0.025568 | 24.3792 | 0.0174 | 0.14981 | -0.00019 | -0.014638 | -0.014638 | -0.42266
2019-03-26 16:13 | 1553631177.6651 | -2 | 0.20955 | 0.000439 | 0 | 100.0702 | 100.2627 | 128.7419 | 21.2612 | 0.001827 | 0.001688 | 268.0364 | 0.25212 | 8.4009 | 0.004021 | 0.005441 | 0.005324 | 5.2e-05 | 0.001244 | 0.3361 | 1.3428 | 2.5476 | 100.79 | 21.17 | 15.699 | 65.7826 | 118.7739 | 62.4377 | 0.003919 | 1.052 | 0.02558 | 24.3155 | -0.017196 | 0.14999 | -0.000268 | -0.020656 | -0.020656 | -0.29988
2019-03-26 16:13 | 1553631178.6659 | -1 | 0.20951 | 0.000441 | -1e-06 | 100.0737 | 100.0946 | 127.7603 | 21.2612 | 0.001907 | 0.001758 | 267.9829 | 0.25151 | 8.4034 | 0.004179 | 0.005654 | 0.005616 | 5.3e-05 | 0.001241 | 0.28639 | 1.3255 | 2.5476 | 100.79 | 21.17 | 15.696 | 67.7161 | 122.6266 | 64.1268 | 0.003918 | 1.0468 | 0.025419 | 24.2835 | -0.001872 | 0.14962 | 1.2e-05 | 0.000985 | 0.000985 | 6.2739
2019-03-26 16:13 | 1553631179.6659 | 0 | 0.20954 | 0.000442 | 0 | 100.1259 | 99.7188 | 126.5111 | 21.2612 | 0.001815 | 0.001722 | 268.1102 | 0.25191 | 8.4033 | 0.004181 | 0.005282 | 0.005192 | 5.3e-05 | 0.001232 | 0.33124 | 1.3581 | 2.8454 | 100.79 | 21.17 | 15.694 | 65.5079 | 122.6822 | 63.2556 | 0.003918 | 1.0495 | 0.025327 | 24.1332 | 0.036444 | 0.14986 | -0.000198 | -0.015133 | -0.015133 | -0.40899
2019-03-26 16:13 | 1553631180.6663 | 1 | 0.20953 | 0.00044 | 2e-06 | 98.1258 | 92.4991 | 124.8539 | 21.2612 | 0.001861 | 0.001722 | 268.2535 | 0.25222 | 8.402 | 0.004202 | 0.006036 | 0.005484 | 5.4e-05 | 0.001226 | 0.34882 | 1.3552 | 2.5478 | 100.79 | 21.16 | 15.689 | 66.5972 | 123.2097 | 63.2727 | 0.003914 | 1.0494 | 0.02516 | 23.9753 | 0.52819 | 0.15005 | -0.000156 | -0.011988 | -0.011988 | -0.51692
2019-03-26 16:13 | 1553631181.6651 | 2 | 0.20949 | 0.000435 | 4e-06 | 97.3866 | 88.3878 | 123.0052 | 21.2612 | 0.001836 | 0.001781 | 268.3235 | 0.25146 | 8.4037 | 0.004175 | 0.005207 | 0.005521 | 5.5e-05 | 0.001225 | 0.35601 | 1.3774 | 2.5477 | 100.79 | 21.16 | 15.687 | 65.9947 | 122.5455 | 64.6829 | 0.003914 | 1.045 | 0.024921 | 23.8469 | 0.86722 | 0.14959 | 0.000146 | 0.010759 | 0.010759 | 0.57422
2019-03-26 16:13 | 1553631182.6654 | 3 | 0.20939 | 0.000443 | 8e-06 | 98.3882 | 92.0604 | 121.295 | 21.2612 | 0.001861 | 0.001868 | 268.0891 | 0.25099 | 8.406 | 0.004272 | 0.006349 | 0.007021 | 5.3e-05 | 0.001222 | 0.34435 | 1.363 | 2.5474 | 100.79 | 21.16 | 15.685 | 66.6003 | 124.9161 | 66.7735 | 0.003913 | 1.0386 | 0.024671 | 23.7537 | 0.5946 | 0.14931 | 0.000683 | 0.050355 | 0.050355 | 0.12246
2019-03-26 16:13 | 1553631183.6651 | 4 | 0.20869 | 0.000446 | 1.4e-05 | 99.2056 | 95.9064 | 119.1452 | 21.2612 | 0.001936 | 0.001998 | 268.1428 | 0.25175 | 8.3983 | 0.004172 | 0.006237 | 0.00725 | 5.5e-05 | 0.001216 | 0.33251 | 1.3601 | 2.5476 | 100.79 | 21.16 | 15.681 | 68.4094 | 122.4652 | 69.9098 | 0.003912 | 1.0291 | 0.024339 | 23.6506 | 0.30252 | 0.14977 | 0.004939 | 0.36196 | 0.36196 | 0.017089
2019-03-26 16:13 | 1553631184.6666 | 5 | 0.20713 | 0.000467 | 2.3e-05 | 99.6944 | 95.3638 | 116.5226 | 21.2612 | 0.002005 | 0.002014 | 268.3298 | 0.25251 | 8.4009 | 0.00414 | 0.006597 | 0.006589 | 5.2e-05 | 0.00121 | 0.31094 | 1.3801 | 2.5345 | 100.79 | 21.16 | 15.679 | 70.0744 | 121.691 | 70.2725 | 0.003912 | 1.028 | 0.024057 | 23.4012 | 0.39723 | 0.15022 | 0.014247 | 1.032 | 1.032 | 0.006011
2019-03-26 16:13 | 1553631185.6656 | 6 | 0.20599 | 0.000549 | 3.4e-05 | 99.8561 | 93.6773 | 113.701 | 21.2612 | 0.002088 | 0.002067 | 267.9427 | 0.25229 | 8.4055 | 0.004061 | 0.005254 | 0.005244 | 4.9e-05 | 0.001201 | 0.31653 | 1.3694 | 1.8894 | 100.79 | 21.15 | 15.674 | 72.0649 | 119.7404 | 71.5591 | 0.003908 | 1.0242 | 0.02372 | 23.1594 | 0.57133 | 0.15009 | 0.020978 | 1.4973 | 1.4973 | 0.00414
2019-03-26 16:13 | 1553631186.6667 | 7 | 0.20491 | 0.000972 | 4.7e-05 | 99.9996 | 91.3208 | 110.6811 | 21.2612 | 0.00214 | 0.002178 | 267.8419 | 0.25116 | 8.4005 | 0.00418 | 0.005896 | 0.005833 | 5e-05 | 0.001204 | 0.29878 | 1.3864 | 2.5477 | 100.79 | 21.15 | 15.672 | 73.3121 | 122.6726 | 74.2296 | 0.003907 | 1.0163 | 0.023312 | 22.9381 | 0.81205 | 0.14941 | 0.026923 | 1.8872 | 1.8872 | 0.00327
2019-03-26 16:13 | 1553631187.6658 | 8 | 0.20399 | 0.001662 | 6.3e-05 | 100.0451 | 87.0351 | 108.1266 | 21.2612 | 0.002227 | 0.002292 | 267.7378 | 0.25183 | 8.4042 | 0.004181 | 0.005451 | 0.005412 | 5.1e-05 | 0.001206 | 0.30613 | 1.396 | 2.5475 | 100.79 | 21.15 | 15.67 | 75.4083 | 122.6829 | 76.9548 | 0.003907 | 1.0084 | 0.022952 | 22.7607 | 1.2461 | 0.14982 | 0.031558 | 2.1764 | 2.1764 | 0.002843
2019-03-26 16:13 | 1553631188.6654 | 9 | 0.2033 | 0.00247 | 8.1e-05 | 100.0684 | 82.1634 | 106.5739 | 21.2612 | 0.002305 | 0.002371 | 267.6081 | 0.25158 | 8.4025 | 0.004107 | 0.005217 | 0.005115 | 5.1e-05 | 0.001209 | 0.31564 | 1.3676 | 2.5476 | 100.79 | 21.15 | 15.669 | 77.2744 | 120.8821 | 78.8669 | 0.003907 | 1.0029 | 0.022724 | 22.6583 | 1.7634 | 0.14966 | 0.034627 | 2.363 | 2.363 | 0.002616
2019-03-26 16:13 | 1553631189.6658 | 10 | 0.20259 | 0.003361 | 9.7e-05 | 100.0923 | 77.7933 | 105.1214 | 21.2612 | 0.002406 | 0.002462 | 267.4443 | 0.251 | 8.401 | 0.004124 | 0.005312 | 0.005227 | 5.5e-05 | 0.00121 | 0.31045 | 1.3761 | 2.5476 | 100.79 | 21.15 | 15.667 | 79.7043 | 121.2987 | 81.0564 | 0.003906 | 0.99672 | 0.022499 | 22.5732 | 2.2544 | 0.14932 | 0.037769 | 2.5505 | 2.5505 | 0.002418
2019-03-26 16:13 | 1553631190.6649 | 11 | 0.20034 | 0.00432 | 0.000112 | 100.0702 | 74.3349 | 104.0696 | 21.2612 | 0.002591 | 0.002551 | 267.2625 | 0.25045 | 8.4016 | 0.004152 | 0.005885 | 0.005733 | 5.2e-05 | 0.00121 | 0.31542 | 1.3275 | 2.5475 | 100.79 | 21.15 | 15.663 | 84.1538 | 121.9815 | 83.1949 | 0.003905 | 0.99074 | 0.022319 | 22.5277 | 2.6591 | 0.14899 | 0.050001 | 3.3462 | 3.3462 | 0.001839
2019-03-26 16:13 | 1553631191.6649 | 12 | 0.19879 | 0.005361 | 0.000124 | 100.0495 | 71.7932 | 103.7041 | 21.2612 | 0.002597 | 0.002656 | 267.0851 | 0.25154 | 8.4029 | 0.004249 | 0.005935 | 0.005877 | 5e-05 | 0.001208 | 0.31222 | 1.3713 | 2.5476 | 100.79 | 21.15 | 15.661 | 84.2995 | 124.3593 | 85.7027 | 0.003905 | 0.98382 | 0.022202 | 22.5671 | 2.9685 | 0.14964 | 0.057954 | 3.8552 | 3.8552 | 0.001603
2019-03-26 16:13 | 1553631192.665 | 13 | 0.19763 | 0.006466 | 0.000133 | 100.1088 | 70.1248 | 103.4537 | 21.2612 | 0.002696 | 0.002643 | 266.9249 | 0.2509 | 8.402 | 0.004225 | 0.005735 | 0.005612 | 5.3e-05 | 0.001206 | 0.27544 | 1.3645 | 2.5475 | 100.79 | 21.14 | 15.659 | 86.68 | 123.7533 | 85.3999 | 0.003902 | 0.98465 | 0.022184 | 22.5303 | 3.1841 | 0.14926 | 0.063489 | 4.2178 | 4.2178 | 0.001462
2019-03-26 16:13 | 1553631193.6646 | 14 | 0.19688 | 0.007607 | 0.000138 | 100.0929 | 69.9036 | 103.5232 | 21.2612 | 0.002789 | 0.002701 | 266.7594 | 0.25093 | 8.4028 | 0.004223 | 0.006039 | 0.005938 | 5.6e-05 | 0.001207 | 0.30071 | 1.3517 | 2.5475 | 100.79 | 21.14 | 15.659 | 88.9049 | 123.7161 | 86.7861 | 0.003902 | 0.98086 | 0.022149 | 22.5814 | 3.2109 | 0.14928 | 0.066557 | 4.4132 | 4.4132 | 0.001397
2019-03-26 16:13 | 1553631194.6648 | 15 | 0.19681 | 0.008647 | 0.000142 | 100.0982 | 70.7299 | 103.8929 | 21.2612 | 0.002866 | 0.002757 | 266.5592 | 0.25125 | 8.4033 | 0.004271 | 0.00542 | 0.005418 | 5.3e-05 | 0.001202 | 0.29645 | 1.3175 | 2.5473 | 100.79 | 21.14 | 15.659 | 90.7794 | 124.8906 | 88.1471 | 0.003902 | 0.97716 | 0.022147 | 22.6644 | 3.1063 | 0.14947 | 0.065738 | 4.3585 | 4.3585 | 0.001416
2019-03-26 16:13 | 1553631195.6658 | 16 | 0.19689 | 0.009262 | 0.000144 | 100.1239 | 70.3104 | 104.8691 | 21.2612 | 0.002859 | 0.002788 | 266.367 | 0.25189 | 8.4015 | 0.004277 | 0.005257 | 0.005219 | 5.2e-05 | 0.001194 | 0.30973 | 1.3693 | 2.5476 | 100.79 | 21.14 | 15.659 | 90.6074 | 125.0433 | 88.8876 | 0.003902 | 0.97516 | 0.022228 | 22.794 | 3.1618 | 0.14985 | 0.06457 | 4.297 | 4.297 | 0.00144
2019-03-26 16:13 | 1553631196.6648 | 17 | 0.1971 | 0.009529 | 0.000145 | 100.102 | 70.6932 | 106.1354 | 21.2612 | 0.002938 | 0.002823 | 266.2073 | 0.25125 | 8.4028 | 0.004327 | 0.005971 | 0.005859 | 5.6e-05 | 0.001188 | 0.2975 | 1.376 | 2.5475 | 100.79 | 21.14 | 15.659 | 92.4964 | 126.273 | 89.7345 | 0.003902 | 0.97289 | 0.022336 | 22.958 | 3.1113 | 0.14947 | 0.062994 | 4.2129 | 4.2129 | 0.001465
2019-03-26 16:13 | 1553631197.6659 | 18 | 0.1974 | 0.009747 | 0.000144 | 100.0105 | 71.2493 | 107.1724 | 21.2612 | 0.002916 | 0.002863 | 266.0515 | 0.2506 | 8.4035 | 0.004364 | 0.005848 | 0.005806 | 5.3e-05 | 0.001176 | 0.27959 | 1.349 | 2.5475 | 100.79 | 21.14 | 15.658 | 91.9794 | 127.1735 | 90.6864 | 0.003902 | 0.97034 | 0.022415 | 23.1001 | 3.033 | 0.14908 | 0.06094 | 4.0907 | 4.0907 | 0.001505
2019-03-26 16:13 | 1553631198.6653 | 19 | 0.19781 | 0.0099 | 0.000144 | 100.0235 | 72.6117 | 107.8559 | 21.2612 | 0.002992 | 0.002929 | 265.9106 | 0.25144 | 8.4003 | 0.004398 | 0.005719 | 0.005665 | 5.1e-05 | 0.00117 | 0.31492 | 1.3351 | 2.5473 | 100.79 | 21.15 | 15.658 | 93.8129 | 128.0187 | 92.2877 | 0.003904 | 0.96609 | 0.022437 | 23.2246 | 2.8647 | 0.14958 | 0.058366 | 3.9224 | 3.9224 | 0.001575
2019-03-26 16:13 | 1553631199.6653 | 20 | 0.19818 | 0.009962 | 0.000143 | 100.1469 | 71.8184 | 108.5282 | 21.2612 | 0.003033 | 0.002951 | 265.7449 | 0.25078 | 8.4047 | 0.004431 | 0.005895 | 0.005853 | 5.1e-05 | 0.001161 | 0.32748 | 1.3728 | 2.5475 | 100.79 | 21.14 | 15.657 | 94.7858 | 128.8295 | 92.825 | 0.003901 | 0.96467 | 0.02249 | 23.314 | 2.974 | 0.14919 | 0.056062 | 3.7772 | 3.7772 | 0.001631
2019-03-26 16:13 | 1553631200.666 | 21 | 0.19858 | 0.009951 | 0.000141 | 100.1141 | 72.1566 | 109.0558 | 21.2612 | 0.003059 | 0.002958 | 265.6319 | 0.25019 | 8.4025 | 0.004327 | 0.006788 | 0.006644 | 5.1e-05 | 0.001164 | 0.30901 | 1.4146 | 2.5475 | 100.79 | 21.14 | 15.658 | 95.431 | 126.2705 | 92.9798 | 0.003902 | 0.96426 | 0.02254 | 23.3755 | 2.9291 | 0.14883 | 0.053672 | 3.6248 | 3.6248 | 0.001696
2019-03-26 16:13 | 1553631201.6664 | 22 | 0.1991 | 0.009872 | 0.000138 | 100.164 | 72.1633 | 109.6119 | 21.2612 | 0.003143 | 0.002958 | 265.5282 | 0.25069 | 8.404 | 0.004401 | 0.00568 | 0.005636 | 5.2e-05 | 0.001165 | 0.33428 | 1.3724 | 2.5475 | 100.79 | 21.15 | 15.659 | 97.4543 | 128.0826 | 92.9931 | 0.003904 | 0.96423 | 0.022597 | 23.4355 | 2.9327 | 0.14914 | 0.050708 | 3.4341 | 3.4341 | 0.001794
2019-03-26 16:13 | 1553631202.6661 | 23 | 0.19945 | 0.009729 | 0.000135 | 100.1773 | 71.9935 | 109.8095 | 21.2612 | 0.003103 | 0.003011 | 265.4279 | 0.25113 | 8.4036 | 0.004416 | 0.006043 | 0.005994 | 5.3e-05 | 0.001165 | 0.31166 | 1.3696 | 2.5476 | 100.79 | 21.15 | 15.66 | 96.4915 | 128.4476 | 94.2543 | 0.003904 | 0.96092 | 0.022579 | 23.497 | 2.955 | 0.1494 | 0.048785 | 3.3017 | 3.3017 | 0.001869
2019-03-26 16:13 | 1553631203.6646 | 24 | 0.19985 | 0.009545 | 0.000131 | 100.2086 | 72.1434 | 110.1459 | 21.2612 | 0.003154 | 0.002979 | 265.3261 | 0.25041 | 8.4033 | 0.004393 | 0.005527 | 0.005459 | 5.2e-05 | 0.001174 | 0.29502 | 1.3729 | 2.5475 | 100.79 | 21.15 | 15.662 | 97.7113 | 127.888 | 93.4946 | 0.003905 | 0.96291 | 0.022637 | 23.5086 | 2.9392 | 0.14897 | 0.046601 | 3.1626 | 3.1626 | 0.001945
2019-03-26 16:13 | 1553631204.6656 | 25 | 0.2 | 0.009324 | 0.000127 | 100.1398 | 73.9281 | 110.7632 | 21.2612 | 0.003168 | 0.003039 | 265.2407 | 0.25052 | 8.4006 | 0.004462 | 0.005733 | 0.00566 | 5e-05 | 0.001173 | 0.28119 | 1.4101 | 2.5476 | 100.79 | 21.14 | 15.662 | 98.0524 | 129.5695 | 94.9284 | 0.003903 | 0.95916 | 0.022656 | 23.6204 | 2.7144 | 0.14904 | 0.045962 | 3.1222 | 3.1222 | 0.001971
2019-03-26 16:13 | 1553631205.6642 | 26 | 0.20054 | 0.009088 | 0.000122 | 100.162 | 74.9406 | 111.2538 | 21.2612 | 0.003177 | 0.002997 | 265.1448 | 0.25037 | 8.4034 | 0.004369 | 0.005963 | 0.005823 | 5e-05 | 0.001177 | 0.30862 | 1.3992 | 2.5476 | 100.79 | 21.14 | 15.662 | 98.2735 | 127.2867 | 93.9241 | 0.003903 | 0.96178 | 0.022737 | 23.6403 | 2.5947 | 0.14895 | 0.043036 | 2.9346 | 2.9346 | 0.002096
2019-03-26 16:13 | 1553631206.6658 | 27 | 0.20041 | 0.008841 | 0.000118 | 100.2015 | 75.0516 | 111.5512 | 21.2612 | 0.003245 | 0.003059 | 265.0417 | 0.25069 | 8.4041 | 0.004388 | 0.005709 | 0.005559 | 5.1e-05 | 0.001171 | 0.28567 | 1.3674 | 2.5475 | 100.79 | 21.14 | 15.66 | 99.9155 | 127.7662 | 95.4255 | 0.003902 | 0.95786 | 0.022721 | 23.7203 | 2.585 | 0.14914 | 0.044155 | 3.0089 | 3.0089 | 0.002047
2019-03-26 16:13 | 1553631207.6643 | 28 | 0.20089 | 0.00859 | 0.000112 | 100.1499 | 76.2819 | 111.6622 | 21.2612 | 0.003239 | 0.003014 | 264.9464 | 0.25046 | 8.4047 | 0.004446 | 0.006318 | 0.006258 | 5.2e-05 | 0.001167 | 0.32731 | 1.376 | 2.5476 | 100.79 | 21.14 | 15.66 | 99.7596 | 129.1875 | 94.3397 | 0.003902 | 0.96069 | 0.022766 | 23.6971 | 2.435 | 0.149 | 0.041598 | 2.8409 | 2.8409 | 0.002166
2019-03-26 16:13 | 1553631208.6644 | 29 | 0.20166 | 0.008258 | 0.000108 | 100.1853 | 77.5661 | 111.962 | 21.2612 | 0.003224 | 0.003026 | 264.8339 | 0.24982 | 8.4051 | 0.004447 | 0.006274 | 0.006256 | 5.2e-05 | 0.001171 | 0.29856 | 1.3064 | 2.5475 | 100.79 | 21.14 | 15.662 | 99.4104 | 129.2099 | 94.6316 | 0.003903 | 0.95993 | 0.022787 | 23.7383 | 2.2888 | 0.14862 | 0.03736 | 2.5546 | 2.5546 | 0.002403
2019-03-26 16:13 | 1553631209.6644 | 30 | 0.20209 | 0.007923 | 0.000104 | 100.2688 | 78.1574 | 112.4234 | 21.2612 | 0.003209 | 0.003046 | 264.7754 | 0.25071 | 8.4036 | 0.004392 | 0.006327 | 0.006232 | 5.2e-05 | 0.001173 | 0.32189 | 1.3198 | 2.5475 | 100.79 | 21.14 | 15.662 | 99.0418 | 127.8539 | 95.1161 | 0.003903 | 0.95867 | 0.022819 | 23.8028 | 2.2284 | 0.14914 | 0.035219 | 2.412 | 2.412 | 0.002554
2019-03-26 16:13 | 1553631210.6659 | 31 | 0.20227 | 0.007678 | 0.0001 | 100.278 | 79.7014 | 112.8652 | 21.2612 | 0.003246 | 0.003094 | 264.6915 | 0.25057 | 8.4002 | 0.00435 | 0.006428 | 0.006387 | 5.3e-05 | 0.001178 | 0.31078 | 1.3683 | 2.5475 | 100.79 | 21.14 | 15.661 | 99.9462 | 126.8198 | 96.2689 | 0.003902 | 0.95568 | 0.022828 | 23.8869 | 2.0542 | 0.14906 | 0.034409 | 2.3578 | 2.3578 | 0.002611
2019-03-26 16:14 | 1553631211.6646 | 32 | 0.20236 | 0.007379 | 9.7e-05 | 100.2945 | 80.6554 | 112.945 | 21.2612 | 0.003255 | 0.003077 | 264.6004 | 0.25083 | 8.4037 | 0.004339 | 0.005784 | 0.005707 | 5.3e-05 | 0.001183 | 0.30309 | 1.3502 | 2.5475 | 100.79 | 21.14 | 15.658 | 100.1462 | 126.5635 | 95.8476 | 0.003902 | 0.95677 | 0.022849 | 23.8817 | 1.9492 | 0.14922 | 0.034284 | 2.3515 | 2.3515 | 0.002621
2019-03-26 16:14 | 1553631212.666 | 33 | 0.20262 | 0.007044 | 9.4e-05 | 100.3326 | 80.3633 | 113.1257 | 21.2612 | 0.00332 | 0.003053 | 264.5011 | 0.25089 | 8.4025 | 0.004365 | 0.005597 | 0.005536 | 5.5e-05 | 0.001188 | 0.32604 | 1.3425 | 2.5475 | 100.79 | 21.14 | 15.654 | 101.7367 | 127.1973 | 95.2727 | 0.003901 | 0.95826 | 0.022885 | 23.8821 | 1.9851 | 0.14925 | 0.033107 | 2.2746 | 2.2746 | 0.00271
2019-03-26 16:14 | 1553631213.6651 | 34 | 0.20293 | 0.006814 | 9.2e-05 | 100.3261 | 80.6564 | 113.2872 | 21.2612 | 0.003293 | 0.003102 | 264.3913 | 0.25093 | 8.4051 | 0.00435 | 0.005363 | 0.005246 | 5.2e-05 | 0.001188 | 0.3235 | 1.3955 | 2.5475 | 100.79 | 21.14 | 15.647 | 101.0872 | 126.8247 | 96.4625 | 0.003899 | 0.95518 | 0.022865 | 23.9377 | 1.9519 | 0.14928 | 0.031551 | 2.166 | 2.166 | 0.002846
2019-03-26 16:14 | 1553631214.6647 | 35 | 0.2031 | 0.006587 | 9e-05 | 100.3411 | 80.0106 | 113.3632 | 21.2612 | 0.003263 | 0.003053 | 264.3078 | 0.25109 | 8.4004 | 0.004469 | 0.005702 | 0.005652 | 5.2e-05 | 0.001186 | 0.3178 | 1.3844 | 2.5475 | 100.79 | 21.14 | 15.642 | 100.3402 | 129.7473 | 95.2822 | 0.003898 | 0.95824 | 0.022909 | 23.9075 | 2.0252 | 0.14937 | 0.030801 | 2.1187 | 2.1187 | 0.002912
2019-03-26 16:14 | 1553631215.6651 | 36 | 0.2032 | 0.006378 | 8.9e-05 | 100.4096 | 80.7411 | 113.3222 | 21.2612 | 0.003316 | 0.003112 | 264.2254 | 0.25082 | 8.4027 | 0.004382 | 0.005804 | 0.005819 | 5.5e-05 | 0.001182 | 0.34324 | 1.3663 | 2.5475 | 100.79 | 21.14 | 15.638 | 101.6276 | 127.6221 | 96.6936 | 0.003897 | 0.95458 | 0.022861 | 23.9489 | 1.95 | 0.14921 | 0.030438 | 2.0895 | 2.0895 | 0.002949
2019-03-26 16:14 | 1553631216.6658 | 37 | 0.20328 | 0.00618 | 8.7e-05 | 100.3712 | 81.3429 | 113.1363 | 21.2612 | 0.003313 | 0.003079 | 264.1288 | 0.25071 | 8.4019 | 0.004331 | 0.005714 | 0.005637 | 5.5e-05 | 0.001182 | 0.29839 | 1.3687 | 2.5476 | 100.79 | 21.14 | 15.63 | 101.5622 | 126.3699 | 95.9039 | 0.003895 | 0.95662 | 0.022867 | 23.9037 | 1.8802 | 0.14915 | 0.03025 | 2.0772 | 2.0772 | 0.002965
2019-03-26 16:14 | 1553631217.6648 | 38 | 0.20331 | 0.005999 | 8.5e-05 | 100.3633 | 81.321 | 112.9534 | 21.2612 | 0.003367 | 0.003111 | 264.0531 | 0.24992 | 8.4017 | 0.004345 | 0.005518 | 0.005465 | 5.5e-05 | 0.001189 | 0.33395 | 1.3726 | 2.5475 | 100.79 | 21.14 | 15.628 | 102.8685 | 126.698 | 96.6747 | 0.003894 | 0.95463 | 0.022825 | 23.9093 | 1.8819 | 0.14868 | 0.03028 | 2.0756 | 2.0756 | 0.002958
2019-03-26 16:14 | 1553631218.6659 | 39 | 0.20328 | 0.005837 | 8.4e-05 | 100.3949 | 81.7614 | 112.7891 | 21.2612 | 0.003317 | 0.003114 | 263.9877 | 0.25106 | 8.4016 | 0.004378 | 0.005575 | 0.005496 | 5.8e-05 | 0.001194 | 0.31885 | 1.3178 | 2.5475 | 100.79 | 21.14 | 15.628 | 101.6629 | 127.5224 | 96.7491 | 0.003894 | 0.95444 | 0.022806 | 23.8943 | 1.8364 | 0.14936 | 0.030629 | 2.0978 | 2.0978 | 0.00294
2019-03-26 16:14 | 1553631219.664 | 40 | 0.20323 | 0.005704 | 8.3e-05 | 100.3594 | 82.2723 | 112.4729 | 21.2612 | 0.003373 | 0.003087 | 263.8784 | 0.251 | 8.4003 | 0.004438 | 0.00613 | 0.006009 | 5.6e-05 | 0.001197 | 0.33964 | 1.3659 | 2.5474 | 100.79 | 21.14 | 15.629 | 103.0082 | 128.9889 | 96.1037 | 0.003894 | 0.9561 | 0.022794 | 23.84 | 1.7775 | 0.14932 | 0.031111 | 2.1297 | 2.1297 | 0.002896
2019-03-26 16:14 | 1553631220.665 | 41 | 0.20314 | 0.005601 | 8.2e-05 | 100.3727 | 83.2997 | 112.2375 | 21.2612 | 0.003389 | 0.0031 | 263.7869 | 0.2503 | 8.4051 | 0.004425 | 0.005924 | 0.005976 | 5.5e-05 | 0.001204 | 0.34435 | 1.342 | 2.5475 | 100.79 | 21.14 | 15.63 | 103.4013 | 128.6749 | 96.4123 | 0.003895 | 0.95531 | 0.02276 | 23.825 | 1.6677 | 0.14891 | 0.031803 | 2.1739 | 2.1739 | 0.002829
2019-03-26 16:14 | 1553631221.665 | 42 | 0.20308 | 0.005518 | 8.1e-05 | 100.3281 | 84.4507 | 111.9676 | 21.2612 | 0.003318 | 0.003116 | 263.7057 | 0.25061 | 8.4053 | 0.004431 | 0.005677 | 0.005642 | 5.6e-05 | 0.001196 | 0.33539 | 1.3697 | 2.5476 | 100.79 | 21.14 | 15.631 | 101.6856 | 128.8096 | 96.7906 | 0.003895 | 0.95433 | 0.022721 | 23.8085 | 1.541 | 0.14909 | 0.032274 | 2.2023 | 2.2023 | 0.002796
2019-03-26 16:14 | 1553631222.664 | 43 | 0.20279 | 0.005462 | 8.1e-05 | 100.4391 | 83.8106 | 111.6371 | 21.2612 | 0.003448 | 0.003164 | 263.6004 | 0.25031 | 8.3999 | 0.004372 | 0.005933 | 0.005813 | 5.7e-05 | 0.001194 | 0.33069 | 1.3491 | 2.5475 | 100.79 | 21.14 | 15.631 | 104.8242 | 127.3638 | 97.956 | 0.003895 | 0.95133 | 0.022652 | 23.8107 | 1.6189 | 0.14891 | 0.034034 | 2.3152 | 2.3152 | 0.002656
2019-03-26 16:14 | 1553631223.6651 | 44 | 0.20266 | 0.00543 | 8.1e-05 | 100.3491 | 82.5934 | 111.5451 | 21.2612 | 0.003418 | 0.003185 | 263.5231 | 0.25045 | 8.4033 | 0.004387 | 0.006014 | 0.005958 | 6e-05 | 0.001191 | 0.32532 | 1.3345 | 2.5475 | 100.79 | 21.14 | 15.631 | 104.0993 | 127.7471 | 98.4741 | 0.003895 | 0.95001 | 0.022627 | 23.8175 | 1.7417 | 0.14899 | 0.034865 | 2.369 | 2.369 | 0.002597
2019-03-26 16:14 | 1553631224.6652 | 45 | 0.20259 | 0.005425 | 8e-05 | 100.3884 | 83.424 | 111.5709 | 21.2612 | 0.003405 | 0.003167 | 263.4186 | 0.25032 | 8.4005 | 0.004396 | 0.005787 | 0.005711 | 5.9e-05 | 0.00119 | 0.33644 | 1.3744 | 2.5475 | 100.79 | 21.14 | 15.634 | 103.7894 | 127.9615 | 98.0319 | 0.003896 | 0.95114 | 0.022643 | 23.8061 | 1.6557 | 0.14892 | 0.035283 | 2.3991 | 2.3991 | 0.002564
2019-03-26 16:14 | 1553631225.6641 | 46 | 0.20258 | 0.005446 | 7.9e-05 | 100.4037 | 84.4991 | 111.7404 | 21.2612 | 0.00345 | 0.003123 | 263.3221 | 0.25118 | 8.4009 | 0.00436 | 0.006188 | 0.006088 | 5.7e-05 | 0.001189 | 0.32532 | 1.3168 | 2.5475 | 100.79 | 21.14 | 15.637 | 104.8816 | 127.0674 | 96.9701 | 0.003896 | 0.95387 | 0.022693 | 23.7901 | 1.5426 | 0.14943 | 0.03531 | 2.4063 | 2.4063 | 0.002565
2019-03-26 16:14 | 1553631226.6642 | 47 | 0.20261 | 0.005479 | 7.9e-05 | 100.3653 | 84.8813 | 112.0467 | 21.2612 | 0.003497 | 0.00319 | 263.2386 | 0.25047 | 8.4032 | 0.004389 | 0.005662 | 0.005661 | 5.5e-05 | 0.001189 | 0.34108 | 1.33 | 2.5475 | 100.79 | 21.13 | 15.641 | 106.0114 | 127.7989 | 98.5956 | 0.003895 | 0.9497 | 0.022674 | 23.8749 | 1.4988 | 0.14901 | 0.035082 | 2.3888 | 2.3888 | 0.002576
2019-03-26 16:14 | 1553631227.6653 | 48 | 0.2027 | 0.005517 | 7.8e-05 | 100.3574 | 84.9855 | 112.5161 | 21.2612 | 0.003488 | 0.003138 | 263.1276 | 0.25078 | 8.4 | 0.0044 | 0.006468 | 0.006343 | 5.6e-05 | 0.001194 | 0.33682 | 1.3431 | 2.5475 | 100.79 | 21.13 | 15.641 | 105.8052 | 128.0492 | 97.3234 | 0.003895 | 0.95296 | 0.02276 | 23.8839 | 1.4871 | 0.14919 | 0.034512 | 2.3591 | 2.3591 | 0.002612
2019-03-26 16:14 | 1553631228.6653 | 49 | 0.20282 | 0.005561 | 7.8e-05 | 100.4146 | 85.0299 | 112.914 | 21.2612 | 0.003468 | 0.003201 | 263.0369 | 0.25072 | 8.3996 | 0.004433 | 0.006025 | 0.005954 | 5.9e-05 | 0.001197 | 0.31166 | 1.359 | 2.5476 | 100.79 | 21.13 | 15.641 | 105.326 | 128.87 | 98.8583 | 0.003895 | 0.94903 | 0.022753 | 23.9756 | 1.4875 | 0.14915 | 0.033783 | 2.3086 | 2.3086 | 0.002668
2019-03-26 16:14 | 1553631229.6644 | 50 | 0.20284 | 0.005583 | 7.7e-05 | 100.4081 | 84.5176 | 113.3693 | 21.2612 | 0.003466 | 0.003164 | 262.966 | 0.25068 | 8.4011 | 0.004412 | 0.00635 | 0.006258 | 5.8e-05 | 0.001192 | 0.32997 | 1.3479 | 2.5474 | 100.79 | 21.13 | 15.637 | 105.2773 | 128.3584 | 97.9551 | 0.003894 | 0.95134 | 0.022827 | 23.9947 | 1.541 | 0.14913 | 0.033625 | 2.3053 | 2.3053 | 0.002672
2019-03-26 16:14 | 1553631230.6644 | 51 | 0.203 | 0.005602 | 7.6e-05 | 100.345 | 84.208 | 113.8451 | 21.2612 | 0.003445 | 0.003213 | 262.8601 | 0.24959 | 8.404 | 0.0045 | 0.005913 | 0.005786 | 5.7e-05 | 0.00119 | 0.31924 | 1.3753 | 2.5475 | 100.79 | 21.13 | 15.633 | 104.7541 | 130.5246 | 99.1494 | 0.003893 | 0.94828 | 0.022838 | 24.0837 | 1.5682 | 0.14848 | 0.032635 | 2.2387 | 2.2387 | 0.002739
2019-03-26 16:14 | 1553631231.6655 | 52 | 0.20304 | 0.005596 | 7.5e-05 | 100.324 | 83.7955 | 114.3118 | 21.2612 | 0.00345 | 0.003162 | 262.7601 | 0.25015 | 8.4011 | 0.004406 | 0.006476 | 0.006456 | 5.4e-05 | 0.001186 | 0.31166 | 1.3737 | 2.5476 | 100.79 | 21.13 | 15.626 | 104.869 | 128.2054 | 97.9152 | 0.003891 | 0.95144 | 0.022923 | 24.093 | 1.6102 | 0.14881 | 0.03239 | 2.2302 | 2.2302 | 0.002756
2019-03-26 16:14 | 1553631232.6645 | 53 | 0.20322 | 0.005581 | 7.4e-05 | 100.2986 | 83.9239 | 115.1345 | 21.2612 | 0.003511 | 0.003123 | 262.69 | 0.25006 | 8.4018 | 0.004465 | 0.005939 | 0.005817 | 5.2e-05 | 0.001184 | 0.30055 | 1.37 | 2.5475 | 100.79 | 21.13 | 15.624 | 106.3641 | 129.658 | 96.9763 | 0.003891 | 0.95385 | 0.023034 | 24.1489 | 1.5943 | 0.14876 | 0.031328 | 2.1677 | 2.1677 | 0.002834
2019-03-26 16:14 | 1553631233.6655 | 54 | 0.2033 | 0.005556 | 7.4e-05 | 100.3054 | 84.2547 | 115.5504 | 21.2612 | 0.003464 | 0.003092 | 262.6024 | 0.24993 | 8.404 | 0.004408 | 0.006182 | 0.006174 | 5.3e-05 | 0.001182 | 0.31421 | 1.3417 | 2.5475 | 100.79 | 21.13 | 15.622 | 105.2174 | 128.2555 | 96.2179 | 0.00389 | 0.95581 | 0.0231 | 24.1677 | 1.5597 | 0.14869 | 0.030902 | 2.1444 | 2.1444 | 0.002864
2019-03-26 16:14 | 1553631234.6636 | 55 | 0.20349 | 0.005522 | 7.4e-05 | 100.2505 | 85.1701 | 115.9175 | 21.2612 | 0.00348 | 0.003152 | 262.55 | 0.25084 | 8.403 | 0.004324 | 0.005952 | 0.005851 | 5.7e-05 | 0.001184 | 0.31116 | 1.3334 | 2.5475 | 100.79 | 21.13 | 15.62 | 105.6068 | 126.2003 | 97.6628 | 0.00389 | 0.95208 | 0.023091 | 24.2533 | 1.4582 | 0.14923 | 0.029785 | 2.0662 | 2.0662 | 0.002983
2019-03-26 16:14 | 1553631235.6637 | 56 | 0.20365 | 0.005477 | 7.4e-05 | 100.2886 | 85.1664 | 116.1926 | 21.2612 | 0.003461 | 0.003138 | 262.4852 | 0.2507 | 8.4052 | 0.004333 | 0.005422 | 0.005409 | 5.4e-05 | 0.001181 | 0.31614 | 1.3499 | 2.5475 | 100.79 | 21.13 | 15.613 | 105.1476 | 126.4146 | 97.3275 | 0.003888 | 0.95295 | 0.023129 | 24.2711 | 1.4619 | 0.14914 | 0.028916 | 2.0093 | 2.0093 | 0.003066
2019-03-26 16:14 | 1553631236.6642 | 57 | 0.20387 | 0.005411 | 7.4e-05 | 100.2062 | 84.7784 | 116.4339 | 21.2612 | 0.00351 | 0.003142 | 262.3914 | 0.25141 | 8.4034 | 0.004314 | 0.00532 | 0.005212 | 5.3e-05 | 0.001184 | 0.3115 | 1.3672 | 2.5475 | 100.79 | 21.13 | 15.607 | 106.334 | 125.9457 | 97.4159 | 0.003886 | 0.95272 | 0.02315 | 24.2992 | 1.4954 | 0.14957 | 0.027679 | 1.9252 | 1.9252 | 0.003209
2019-03-26 16:14 | 1553631237.6651 | 58 | 0.20414 | 0.005335 | 7.3e-05 | 100.1133 | 84.2648 | 116.8116 | 21.2612 | 0.00344 | 0.003154 | 262.333 | 0.2513 | 8.4024 | 0.004293 | 0.005735 | 0.005687 | 5.2e-05 | 0.001189 | 0.31725 | 1.3559 | 2.5475 | 100.79 | 21.13 | 15.598 | 104.6363 | 125.4428 | 97.706 | 0.003884 | 0.95197 | 0.023179 | 24.3481 | 1.5415 | 0.1495 | 0.026108 | 1.8182 | 1.8182 | 0.003396
2019-03-26 16:14 | 1553631238.6648 | 59 | 0.20434 | 0.005238 | 7.2e-05 | 100.1071 | 85.4524 | 117.0501 | 21.2612 | 0.003491 | 0.003092 | 262.2561 | 0.25044 | 8.4038 | 0.004299 | 0.005992 | 0.005899 | 5.1e-05 | 0.001192 | 0.32244 | 1.3982 | 2.5477 | 100.79 | 21.14 | 15.579 | 105.8831 | 125.59 | 96.2252 | 0.003882 | 0.95579 | 0.023249 | 24.3242 | 1.4157 | 0.14899 | 0.025029 | 1.7485 | 1.7485 | 0.003519
2019-03-26 16:14 | 1553631239.6649 | 60 | 0.20449 | 0.005128 | 7e-05 | 100.084 | 86.1032 | 117.1026 | 21.2612 | 0.003488 | 0.003054 | 262.1718 | 0.2508 | 8.4033 | 0.004207 | 0.006246 | 0.006136 | 5.2e-05 | 0.00119 | 0.27328 | 1.3938 | 2.5476 | 100.79 | 21.14 | 15.569 | 105.7894 | 123.3226 | 95.2901 | 0.003879 | 0.95822 | 0.023284 | 24.2989 | 1.3458 | 0.1492 | 0.024304 | 1.7004 | 1.7004 | 0.003624
2019-03-26 16:14 | 1553631240.6649 | 61 | 0.20457 | 0.005013 | 6.9e-05 | 99.991 | 86.0394 | 117.2108 | 21.2612 | 0.003462 | 0.003139 | 262.1168 | 0.25127 | 8.4009 | 0.004253 | 0.005636 | 0.005603 | 5e-05 | 0.00119 | 0.30757 | 1.3543 | 2.5475 | 100.79 | 21.13 | 15.556 | 105.1655 | 124.4625 | 97.3415 | 0.003874 | 0.95291 | 0.02323 | 24.3777 | 1.3441 | 0.14948 | 0.023962 | 1.6727 | 1.6727 | 0.003691
2019-03-26 16:14 | 1553631241.6639 | 62 | 0.20483 | 0.004888 | 6.8e-05 | 99.9435 | 86.3577 | 117.6085 | 21.2612 | 0.003416 | 0.003084 | 262.0617 | 0.25084 | 8.4043 | 0.004201 | 0.005853 | 0.005679 | 5.1e-05 | 0.001188 | 0.32853 | 1.3572 | 2.5474 | 100.79 | 21.13 | 15.532 | 104.0552 | 123.17 | 96.0246 | 0.003868 | 0.95631 | 0.023311 | 24.3756 | 1.3069 | 0.14923 | 0.02253 | 1.5783 | 1.5783 | 0.003905
2019-03-26 16:14 | 1553631242.6641 | 63 | 0.20491 | 0.004763 | 6.6e-05 | 99.917 | 86.2374 | 117.3972 | 21.2612 | 0.003395 | 0.003062 | 261.9928 | 0.25099 | 8.4031 | 0.004165 | 0.005781 | 0.005714 | 4.7e-05 | 0.001194 | 0.31614 | 1.3294 | 2.5474 | 100.79 | 21.13 | 15.52 | 103.5396 | 122.288 | 95.5006 | 0.003865 | 0.95767 | 0.023306 | 24.3364 | 1.317 | 0.14931 | 0.022199 | 1.555 | 1.555 | 0.003966
2019-03-26 16:14 | 1553631243.6641 | 64 | 0.20503 | 0.004647 | 6.5e-05 | 99.826 | 86.549 | 117.4848 | 21.2612 | 0.003379 | 0.003083 | 261.935 | 0.25063 | 8.4029 | 0.004249 | 0.005316 | 0.005285 | 5.3e-05 | 0.001192 | 0.32068 | 1.367 | 2.5475 | 100.79 | 21.13 | 15.494 | 103.1576 | 124.3593 | 96.005 | 0.003858 | 0.95636 | 0.023299 | 24.3621 | 1.2766 | 0.1491 | 0.021625 | 1.5143 | 1.5143 | 0.004066
2019-03-26 16:14 | 1553631244.6641 | 65 | 0.20516 | 0.004531 | 6.4e-05 | 99.7345 | 86.6256 | 117.5855 | 21.2612 | 0.003388 | 0.003099 | 261.8665 | 0.24987 | 8.4031 | 0.004114 | 0.005162 | 0.005019 | 5e-05 | 0.001188 | 0.32029 | 1.3964 | 2.5477 | 100.79 | 21.13 | 15.48 | 103.3824 | 121.0399 | 96.384 | 0.003855 | 0.95538 | 0.023297 | 24.385 | 1.2604 | 0.14865 | 0.021018 | 1.4717 | 1.4717 | 0.004172
2019-03-26 16:14 | 1553631245.6643 | 66 | 0.2052 | 0.004428 | 6.3e-05 | 99.7071 | 87.1393 | 117.4425 | 21.2612 | 0.003456 | 0.003095 | 261.8083 | 0.2514 | 8.4019 | 0.004155 | 0.006085 | 0.005992 | 4.8e-05 | 0.001183 | 0.31277 | 1.406 | 2.5475 | 100.79 | 21.13 | 15.468 | 105.0271 | 122.0593 | 96.2838 | 0.003852 | 0.95564 | 0.023286 | 24.3669 | 1.2051 | 0.14956 | 0.020865 | 1.4603 | 1.4603 | 0.00423
2019-03-26 16:14 | 1553631246.6642 | 67 | 0.20525 | 0.004325 | 6.2e-05 | 99.6442 | 87.3726 | 117.5218 | 21.2612 | 0.003383 | 0.003017 | 261.7601 | 0.25063 | 8.4035 | 0.004189 | 0.005498 | 0.005432 | 5e-05 | 0.001176 | 0.32659 | 1.4064 | 2.5475 | 100.79 | 21.13 | 15.445 | 103.2497 | 122.8942 | 94.4085 | 0.003846 | 0.96052 | 0.023353 | 24.3132 | 1.1755 | 0.1491 | 0.02072 | 1.4545 | 1.4545 | 0.004234
2019-03-26 16:14 | 1553631247.6657 | 68 | 0.20528 | 0.004232 | 6.2e-05 | 99.5574 | 87.025 | 117.6096 | 21.2612 | 0.003357 | 0.00304 | 261.6626 | 0.25179 | 8.4041 | 0.004213 | 0.006313 | 0.006237 | 5e-05 | 0.00117 | 0.29789 | 1.4437 | 2.5477 | 100.79 | 21.13 | 15.435 | 102.628 | 123.4703 | 94.9574 | 0.003844 | 0.95908 | 0.023344 | 24.3404 | 1.2034 | 0.14979 | 0.020676 | 1.4509 | 1.4509 | 0.004264
2019-03-26 16:14 | 1553631248.6648 | 69 | 0.20541 | 0.004148 | 6.2e-05 | 99.4423 | 87.1722 | 117.7239 | 21.2612 | 0.003393 | 0.003045 | 261.6025 | 0.25025 | 8.403 | 0.004148 | 0.005967 | 0.00587 | 4.8e-05 | 0.001171 | 0.29302 | 1.3692 | 2.5476 | 100.79 | 21.13 | 15.424 | 103.5064 | 121.8838 | 95.0903 | 0.003841 | 0.95874 | 0.023352 | 24.3567 | 1.1779 | 0.14887 | 0.019965 | 1.4014 | 1.4014 | 0.004387
2019-03-26 16:14 | 1553631249.6645 | 70 | 0.2054 | 0.00407 | 6.2e-05 | 99.3959 | 87.2563 | 117.8216 | 21.2612 | 0.003379 | 0.003014 | 261.5338 | 0.25028 | 8.4023 | 0.004155 | 0.005811 | 0.005772 | 5e-05 | 0.001171 | 0.29927 | 1.3562 | 2.5476 | 100.79 | 21.13 | 15.42 | 103.1695 | 122.0442 | 94.3436 | 0.00384 | 0.96068 | 0.023385 | 24.342 | 1.1651 | 0.14889 | 0.020162 | 1.4173 | 1.4173 | 0.004339
2019-03-26 16:14 | 1553631250.6634 | 71 | 0.20543 | 0.003999 | 6.2e-05 | 99.3779 | 87.4082 | 117.7594 | 21.2612 | 0.003386 | 0.003021 | 261.4564 | 0.2506 | 8.4041 | 0.004126 | 0.005578 | 0.005477 | 4.9e-05 | 0.001165 | 0.27759 | 1.3475 | 2.5476 | 100.79 | 21.13 | 15.417 | 103.3345 | 121.3296 | 94.5026 | 0.003839 | 0.96027 | 0.023374 | 24.3409 | 1.148 | 0.14908 | 0.020053 | 1.4089 | 1.4089 | 0.00437
2019-03-26 16:14 | 1553631251.6636 | 72 | 0.20554 | 0.003934 | 6.1e-05 | 99.1489 | 87.381 | 117.4878 | 21.2612 | 0.003406 | 0.002957 | 261.3814 | 0.25099 | 8.4025 | 0.004079 | 0.005463 | 0.005411 | 5.2e-05 | 0.001159 | 0.30613 | 1.3556 | 2.5476 | 100.79 | 21.13 | 15.415 | 103.8051 | 120.1846 | 92.9627 | 0.003839 | 0.96431 | 0.023396 | 24.2618 | 1.1301 | 0.14931 | 0.019478 | 1.3699 | 1.3699 | 0.004502
2019-03-26 16:14 | 1553631252.6636 | 73 | 0.20559 | 0.003866 | 6.1e-05 | 99.0665 | 87.496 | 117.6512 | 21.2612 | 0.003409 | 0.002975 | 261.322 | 0.25066 | 8.4017 | 0.004118 | 0.005904 | 0.005857 | 4.9e-05 | 0.001158 | 0.32621 | 1.3913 | 2.5475 | 100.79 | 21.13 | 15.414 | 103.8839 | 121.1336 | 93.403 | 0.003838 | 0.96315 | 0.023398 | 24.2933 | 1.1109 | 0.14912 | 0.019222 | 1.3519 | 1.3519 | 0.004555
2019-03-26 16:14 | 1553631253.665 | 74 | 0.2057 | 0.00381 | 6e-05 | 98.998 | 87.3686 | 117.4634 | 21.2612 | 0.003335 | 0.003029 | 261.2182 | 0.25094 | 8.4034 | 0.004181 | 0.005781 | 0.005742 | 5e-05 | 0.00116 | 0.31758 | 1.3735 | 2.5474 | 100.79 | 21.13 | 15.414 | 102.0821 | 122.6783 | 94.7091 | 0.003838 | 0.95973 | 0.023338 | 24.3171 | 1.1178 | 0.14928 | 0.01865 | 1.3083 | 1.3083 | 0.004712
2019-03-26 16:14 | 1553631254.6637 | 75 | 0.20571 | 0.003751 | 5.9e-05 | 98.9054 | 88.015 | 117.1667 | 21.2612 | 0.003318 | 0.002981 | 261.144 | 0.25012 | 8.4029 | 0.004154 | 0.005691 | 0.005638 | 5e-05 | 0.001158 | 0.33339 | 1.3487 | 2.5475 | 100.79 | 21.13 | 15.415 | 101.6801 | 122.0356 | 93.5305 | 0.003839 | 0.96281 | 0.023346 | 24.2474 | 1.0434 | 0.1488 | 0.018698 | 1.3123 | 1.3123 | 0.004683
2019-03-26 16:14 | 1553631255.7219 | 76 | 0.20579 | 0.003694 | 5.9e-05 | 98.7439 | 87.9508 | 117.5186 | 21.2612 | 0.003372 | 0.002967 | 261.0708 | 0.25141 | 8.3995 | 0.004108 | 0.005673 | 0.005535 | 4.7e-05 | 0.001159 | 0.27936 | 1.3535 | 2.5475 | 100.79 | 21.13 | 15.415 | 102.9997 | 120.9041 | 93.2004 | 0.003839 | 0.96368 | 0.023391 | 24.2729 | 1.0353 | 0.14956 | 0.018293 | 1.2863 | 1.2863 | 0.004802
2019-03-26 16:14 | 1553631256.6702 | 77 | 0.20575 | 0.003642 | 6e-05 | 98.7017 | 87.9841 | 117.3453 | 21.2612 | 0.003344 | 0.002962 | 261.0081 | 0.2505 | 8.4051 | 0.004142 | 0.005346 | 0.005358 | 4.7e-05 | 0.001156 | 0.32925 | 1.342 | 2.5475 | 100.79 | 21.13 | 15.419 | 102.3218 | 121.7289 | 93.0924 | 0.00384 | 0.96397 | 0.023378 | 24.2514 | 1.0281 | 0.14902 | 0.018541 | 1.3029 | 1.3029 | 0.004724
2019-03-26 16:14 | 1553631257.6704 | 78 | 0.2058 | 0.003586 | 6.1e-05 | 98.6028 | 87.8815 | 117.0963 | 21.2612 | 0.003332 | 0.00297 | 260.9798 | 0.25091 | 8.4033 | 0.004081 | 0.005457 | 0.005378 | 5e-05 | 0.001157 | 0.33516 | 1.3297 | 2.5475 | 100.79 | 21.13 | 15.421 | 102.031 | 120.2384 | 93.266 | 0.00384 | 0.96351 | 0.023347 | 24.2314 | 1.0296 | 0.14927 | 0.018314 | 1.2852 | 1.2852 | 0.004797
2019-03-26 16:14 | 1553631258.6707 | 79 | 0.20583 | 0.003549 | 6.1e-05 | 98.5219 | 88.2019 | 117.1038 | 21.2612 | 0.003285 | 0.002983 | 260.8996 | 0.25076 | 8.4036 | 0.004183 | 0.00571 | 0.005605 | 4.8e-05 | 0.001159 | 0.33339 | 1.3421 | 2.5475 | 100.79 | 21.13 | 15.422 | 100.8801 | 122.7243 | 93.5814 | 0.00384 | 0.96268 | 0.023338 | 24.2426 | 0.98972 | 0.14918 | 0.018224 | 1.2784 | 1.2784 | 0.004819
2019-03-26 16:14 | 1553631259.6705 | 80 | 0.20585 | 0.003514 | 6.1e-05 | 98.4691 | 87.6439 | 117.3383 | 21.2612 | 0.003312 | 0.002978 | 260.8432 | 0.25112 | 8.4037 | 0.00412 | 0.005812 | 0.005702 | 5e-05 | 0.001157 | 0.32278 | 1.326 | 2.5475 | 100.79 | 21.12 | 15.426 | 101.5324 | 121.1845 | 93.4758 | 0.003839 | 0.96296 | 0.023365 | 24.2634 | 1.0417 | 0.14939 | 0.018147 | 1.2745 | 1.2745 | 0.004841
2019-03-26 16:14 | 1553631260.6705 | 81 | 0.20584 | 0.003476 | 6.1e-05 | 98.4614 | 87.892 | 117.5606 | 21.2612 | 0.003315 | 0.002931 | 260.7971 | 0.24998 | 8.4017 | 0.00408 | 0.005248 | 0.00522 | 5e-05 | 0.00116 | 0.31509 | 1.3308 | 2.5474 | 100.79 | 21.12 | 15.428 | 101.6087 | 120.2154 | 92.3313 | 0.003839 | 0.96597 | 0.023423 | 24.2484 | 1.0157 | 0.14871 | 0.018256 | 1.2853 | 1.2853 | 0.004778
2019-03-26 16:14 | 1553631261.6707 | 82 | 0.20583 | 0.003437 | 6.2e-05 | 98.2937 | 87.7723 | 117.6979 | 21.2612 | 0.003272 | 0.002973 | 260.76 | 0.25037 | 8.4032 | 0.004134 | 0.005536 | 0.005436 | 4.8e-05 | 0.001164 | 0.33516 | 1.3485 | 2.5474 | 100.79 | 21.12 | 15.436 | 100.5772 | 121.5471 | 93.3427 | 0.003841 | 0.96331 | 0.023405 | 24.2961 | 1.0127 | 0.14895 | 0.018372 | 1.2924 | 1.2924 | 0.00476
2019-03-26 16:14 | 1553631262.6706 | 83 | 0.20588 | 0.003413 | 6.2e-05 | 98.2111 | 87.5935 | 117.7167 | 21.2612 | 0.003269 | 0.002917 | 260.7015 | 0.2506 | 8.4064 | 0.004104 | 0.005327 | 0.005231 | 4.8e-05 | 0.001159 | 0.30038 | 1.3585 | 2.5476 | 100.79 | 21.12 | 15.441 | 100.4868 | 120.7924 | 91.9905 | 0.003843 | 0.96688 | 0.02345 | 24.2532 | 1.0234 | 0.14908 | 0.018061 | 1.2729 | 1.2729 | 0.004837
2019-03-26 16:14 | 1553631263.6716 | 84 | 0.2059 | 0.003386 | 6.2e-05 | 98.1134 | 87.7904 | 118.004 | 21.2612 | 0.003291 | 0.002961 | 260.6914 | 0.25079 | 8.4022 | 0.004001 | 0.005886 | 0.005837 | 5.1e-05 | 0.001157 | 0.29374 | 1.3546 | 2.5475 | 100.79 | 21.12 | 15.446 | 101.0364 | 118.288 | 93.0642 | 0.003844 | 0.96404 | 0.023444 | 24.3184 | 0.99437 | 0.14919 | 0.017998 | 1.2682 | 1.2682 | 0.004859
2019-03-26 16:14 | 1553631264.6706 | 85 | 0.20594 | 0.003356 | 6.3e-05 | 98.0697 | 88.0039 | 118.4974 | 21.2612 | 0.003275 | 0.002909 | 260.6249 | 0.25113 | 8.4026 | 0.00399 | 0.005819 | 0.005878 | 4.8e-05 | 0.001156 | 0.31437 | 1.3255 | 2.5475 | 100.79 | 21.12 | 15.456 | 100.6312 | 118.0187 | 91.8066 | 0.003846 | 0.96736 | 0.023533 | 24.3273 | 0.96867 | 0.1494 | 0.01781 | 1.2597 | 1.2597 | 0.004898
2019-03-26 16:14 | 1553631265.6707 | 86 | 0.20598 | 0.003326 | 6.2e-05 | 97.967 | 87.8136 | 118.3446 | 21.2612 | 0.003289 | 0.002939 | 260.5583 | 0.25054 | 8.4002 | 0.004047 | 0.005587 | 0.005495 | 4.8e-05 | 0.001158 | 0.30182 | 1.3399 | 2.5475 | 100.79 | 21.12 | 15.461 | 100.9689 | 119.4183 | 92.5268 | 0.003848 | 0.96546 | 0.023495 | 24.3356 | 0.97865 | 0.14905 | 0.017566 | 1.2404 | 1.2404 | 0.004963
2019-03-26 16:14 | 1553631266.6711 | 87 | 0.20599 | 0.003302 | 6.2e-05 | 97.848 | 87.5087 | 118.4781 | 21.2612 | 0.003241 | 0.002888 | 260.509 | 0.25071 | 8.4033 | 0.004051 | 0.005582 | 0.005369 | 4.8e-05 | 0.001163 | 0.31725 | 1.3399 | 2.5475 | 100.79 | 21.12 | 15.467 | 99.8246 | 119.5119 | 91.3032 | 0.003849 | 0.9687 | 0.023548 | 24.3086 | 0.9989 | 0.14914 | 0.017541 | 1.2415 | 1.2415 | 0.004962
2019-03-26 16:14 | 1553631267.6708 | 88 | 0.20601 | 0.003281 | 6.2e-05 | 97.7426 | 87.8836 | 118.6959 | 21.2612 | 0.00329 | 0.002899 | 260.4482 | 0.25014 | 8.4018 | 0.003984 | 0.005044 | 0.004898 | 5e-05 | 0.001165 | 0.31813 | 1.348 | 2.5475 | 100.79 | 21.12 | 15.478 | 101.0074 | 117.8602 | 91.5753 | 0.003852 | 0.96798 | 0.023561 | 24.34 | 0.95103 | 0.14881 | 0.017459 | 1.2363 | 1.2363 | 0.004971
2019-03-26 16:14 | 1553631268.6709 | 89 | 0.206 | 0.00325 | 6.1e-05 | 97.5983 | 87.4939 | 118.4509 | 21.2612 | 0.003244 | 0.002887 | 260.4069 | 0.25097 | 8.4026 | 0.003999 | 0.006055 | 0.006052 | 4.9e-05 | 0.00117 | 0.33395 | 1.3074 | 2.5475 | 100.79 | 21.12 | 15.48 | 99.8826 | 118.2436 | 91.2625 | 0.003852 | 0.96881 | 0.023546 | 24.3044 | 0.97755 | 0.1493 | 0.017531 | 1.2407 | 1.2407 | 0.00497
2019-03-26 16:14 | 1553631269.6699 | 90 | 0.20604 | 0.003237 | 6e-05 | 97.582 | 87.5198 | 118.3953 | 21.2612 | 0.003249 | 0.002899 | 260.3604 | 0.25042 | 8.4024 | 0.003999 | 0.005772 | 0.005667 | 4.7e-05 | 0.001165 | 0.28888 | 1.3328 | 2.5473 | 100.79 | 21.12 | 15.485 | 100.0074 | 118.2326 | 91.569 | 0.003854 | 0.96799 | 0.023531 | 24.3089 | 0.97342 | 0.14897 | 0.017318 | 1.2249 | 1.2249 | 0.005023
2019-03-26 16:14 | 1553631270.6699 | 91 | 0.20606 | 0.003212 | 5.9e-05 | 97.4796 | 87.7128 | 118.6469 | 21.2612 | 0.003255 | 0.002916 | 260.2896 | 0.25051 | 8.4013 | 0.003981 | 0.0061 | 0.006038 | 4.7e-05 | 0.001168 | 0.32388 | 1.401 | 2.5473 | 100.79 | 21.12 | 15.484 | 100.1549 | 117.7945 | 91.9678 | 0.003853 | 0.96693 | 0.023543 | 24.3481 | 0.94432 | 0.14903 | 0.017253 | 1.2209 | 1.2209 | 0.005041
2019-03-26 16:14 | 1553631271.671 | 92 | 0.20614 | 0.003195 | 5.8e-05 | 97.4516 | 88.0422 | 119.0741 | 21.2612 | 0.003254 | 0.002891 | 260.2013 | 0.25045 | 8.4029 | 0.004044 | 0.005293 | 0.00525 | 5.1e-05 | 0.001174 | 0.32316 | 1.3454 | 2.5473 | 100.79 | 21.12 | 15.482 | 100.1221 | 119.3335 | 91.3625 | 0.003853 | 0.96854 | 0.023605 | 24.3716 | 0.90822 | 0.14899 | 0.016814 | 1.193 | 1.193 | 0.005158
2019-03-26 16:15 | 1553631272.67 | 93 | 0.20615 | 0.003182 | 5.8e-05 | 97.3146 | 87.9958 | 119.2496 | 21.2612 | 0.003251 | 0.002908 | 260.15 | 0.25151 | 8.4046 | 0.004105 | 0.005448 | 0.005373 | 5.1e-05 | 0.001173 | 0.31813 | 1.3898 | 2.5472 | 100.79 | 21.12 | 15.474 | 100.057 | 120.8273 | 91.7825 | 0.003851 | 0.96743 | 0.023609 | 24.4036 | 0.90035 | 0.14963 | 0.01677 | 1.1901 | 1.1901 | 0.005192
2019-03-26 16:15 | 1553631273.6711 | 94 | 0.20614 | 0.003158 | 5.8e-05 | 97.2299 | 88.0816 | 119.4804 | 21.2612 | 0.00323 | 0.00287 | 260.0785 | 0.25095 | 8.4025 | 0.004043 | 0.005366 | 0.005347 | 4.9e-05 | 0.001172 | 0.32156 | 1.3849 | 2.5474 | 100.79 | 21.12 | 15.466 | 99.5642 | 119.3066 | 90.8637 | 0.003849 | 0.96987 | 0.023661 | 24.3965 | 0.88385 | 0.14929 | 0.016866 | 1.1996 | 1.1996 | 0.00514
2019-03-26 16:15 | 1553631274.6711 | 95 | 0.2062 | 0.003137 | 5.8e-05 | 97.1865 | 87.9864 | 119.3905 | 21.2612 | 0.003225 | 0.002862 | 260.0461 | 0.25073 | 8.4017 | 0.003983 | 0.0051 | 0.004965 | 4.9e-05 | 0.001169 | 0.31238 | 1.3414 | 2.5474 | 100.79 | 21.12 | 15.451 | 99.437 | 117.8507 | 90.6785 | 0.003845 | 0.97036 | 0.023658 | 24.3811 | 0.88953 | 0.14916 | 0.016486 | 1.1725 | 1.1725 | 0.005254
2019-03-26 16:15 | 1553631275.6706 | 96 | 0.20621 | 0.003122 | 5.8e-05 | 97.1127 | 88.3051 | 119.5743 | 21.2612 | 0.003235 | 0.00289 | 259.9774 | 0.25107 | 8.4031 | 0.004021 | 0.005831 | 0.005817 | 5e-05 | 0.001168 | 0.31116 | 1.3466 | 2.5475 | 100.79 | 21.12 | 15.442 | 99.6708 | 118.7817 | 91.3422 | 0.003843 | 0.96859 | 0.023655 | 24.4221 | 0.8504 | 0.14936 | 0.016472 | 1.1712 | 1.1712 | 0.005267
2019-03-26 16:15 | 1553631276.6718 | 97 | 0.20626 | 0.003098 | 5.8e-05 | 97.0917 | 88.1643 | 119.6839 | 21.2612 | 0.003216 | 0.002863 | 259.9168 | 0.25116 | 8.4029 | 0.00397 | 0.0066 | 0.006541 | 4.9e-05 | 0.001166 | 0.30574 | 1.3301 | 2.5475 | 100.79 | 21.12 | 15.435 | 99.2064 | 117.5377 | 90.6997 | 0.003841 | 0.97031 | 0.023687 | 24.4117 | 0.86274 | 0.14941 | 0.016204 | 1.1537 | 1.1537 | 0.005349
2019-03-26 16:15 | 1553631277.6717 | 98 | 0.20627 | 0.003078 | 5.7e-05 | 96.9872 | 88.3757 | 119.6639 | 21.2612 | 0.003194 | 0.002843 | 259.8781 | 0.25033 | 8.4 | 0.003933 | 0.005662 | 0.005645 | 4.6e-05 | 0.001167 | 0.32211 | 1.4019 | 2.5473 | 100.79 | 21.12 | 15.418 | 98.6756 | 116.6366 | 90.2058 | 0.003837 | 0.97162 | 0.023701 | 24.3931 | 0.83169 | 0.14892 | 0.01615 | 1.1506 | 1.1506 | 0.005345
2019-03-26 16:15 | 1553631278.6713 | 99 | 0.20629 | 0.003055 | 5.7e-05 | 96.9391 | 88.2661 | 119.7848 | 21.2612 | 0.003151 | 0.002866 | 259.8363 | 0.2507 | 8.4031 | 0.004002 | 0.005303 | 0.005204 | 4.6e-05 | 0.001165 | 0.32587 | 1.3705 | 2.5472 | 100.79 | 21.12 | 15.409 | 97.6338 | 118.3022 | 90.7669 | 0.003835 | 0.97013 | 0.023695 | 24.4243 | 0.83835 | 0.14914 | 0.016086 | 1.1457 | 1.1457 | 0.005376
2019-03-26 16:15 | 1553631279.6708 | 100 | 0.20628 | 0.003038 | 5.7e-05 | 96.8249 | 88.4097 | 119.8536 | 21.2612 | 0.003251 | 0.002833 | 259.7857 | 0.25097 | 8.4019 | 0.003925 | 0.005241 | 0.005246 | 5e-05 | 0.001164 | 0.30359 | 1.346 | 2.5474 | 100.79 | 21.12 | 15.395 | 100.0616 | 116.4316 | 89.965 | 0.003831 | 0.97227 | 0.023728 | 24.4043 | 0.81326 | 0.1493 | 0.01613 | 1.1505 | 1.1505 | 0.00536
2019-03-26 16:15 | 1553631280.6695 | 101 | 0.20636 | 0.003021 | 5.7e-05 | 96.8305 | 88.4291 | 119.8041 | 21.2612 | 0.003177 | 0.002833 | 259.7594 | 0.25095 | 8.4045 | 0.003952 | 0.005617 | 0.005526 | 4.4e-05 | 0.001166 | 0.31974 | 1.302 | 2.5474 | 100.79 | 21.12 | 15.389 | 98.2767 | 117.0878 | 89.9838 | 0.00383 | 0.97222 | 0.023722 | 24.3999 | 0.81181 | 0.14929 | 0.015655 | 1.1163 | 1.1163 | 0.005523
2019-03-26 16:15 | 1553631281.6695 | 102 | 0.2064 | 0.003002 | 5.7e-05 | 96.7806 | 88.6201 | 119.5092 | 21.2612 | 0.003115 | 0.002892 | 259.7101 | 0.25061 | 8.4055 | 0.003884 | 0.005815 | 0.005742 | 4.8e-05 | 0.001169 | 0.31708 | 1.3357 | 2.5472 | 100.79 | 21.12 | 15.385 | 96.7749 | 115.4179 | 91.4063 | 0.003829 | 0.96842 | 0.023647 | 24.4176 | 0.78791 | 0.14909 | 0.015477 | 1.1001 | 1.1001 | 0.005597
2019-03-26 16:15 | 1553631282.6695 | 103 | 0.20637 | 0.002987 | 5.8e-05 | 96.772 | 88.5092 | 119.5653 | 21.2612 | 0.003181 | 0.002821 | 259.6474 | 0.25043 | 8.4036 | 0.003908 | 0.005882 | 0.005864 | 4.8e-05 | 0.001167 | 0.31349 | 1.3151 | 2.5472 | 100.79 | 21.12 | 15.38 | 98.3675 | 116.0215 | 89.6861 | 0.003828 | 0.97302 | 0.023708 | 24.3656 | 0.79832 | 0.14898 | 0.015638 | 1.1144 | 1.1144 | 0.005521
2019-03-26 16:15 | 1553631283.6695 | 104 | 0.20639 | 0.00297 | 5.8e-05 | 96.6165 | 88.7102 | 119.4816 | 21.2612 | 0.003156 | 0.002855 | 259.5936 | 0.25091 | 8.4045 | 0.003818 | 0.005063 | 0.004932 | 5.1e-05 | 0.001168 | 0.33356 | 1.3297 | 2.5474 | 100.79 | 21.12 | 15.38 | 97.764 | 113.8135 | 90.5089 | 0.003828 | 0.97081 | 0.023673 | 24.3847 | 0.76364 | 0.14927 | 0.015588 | 1.1091 | 1.1091 | 0.005558
2019-03-26 16:15 | 1553631284.67 | 105 | 0.20641 | 0.002952 | 5.7e-05 | 96.6469 | 88.5021 | 119.6163 | 21.2612 | 0.003137 | 0.00281 | 259.5437 | 0.2502 | 8.4061 | 0.003894 | 0.005062 | 0.005043 | 4.9e-05 | 0.001168 | 0.30718 | 1.361 | 2.5476 | 100.79 | 21.12 | 15.381 | 97.2992 | 115.6711 | 89.4142 | 0.003828 | 0.97375 | 0.023722 | 24.3617 | 0.78746 | 0.14884 | 0.015491 | 1.1046 | 1.1046 | 0.005565
2019-03-26 16:15 | 1553631285.6697 | 106 | 0.20641 | 0.002934 | 5.6e-05 | 96.5657 | 88.9556 | 119.5299 | 21.2612 | 0.003216 | 0.002787 | 259.5213 | 0.25155 | 8.406 | 0.003867 | 0.005038 | 0.005015 | 5.4e-05 | 0.001171 | 0.33052 | 1.3869 | 2.5475 | 100.79 | 21.12 | 15.389 | 99.2151 | 115.0208 | 88.8604 | 0.00383 | 0.97523 | 0.023732 | 24.3343 | 0.73423 | 0.14965 | 0.015478 | 1.1041 | 1.1041 | 0.005598
2019-03-26 16:15 | 1553631286.6711 | 107 | 0.20644 | 0.002919 | 5.6e-05 | 96.5899 | 88.9162 | 119.6684 | 21.2612 | 0.00314 | 0.002855 | 259.4692 | 0.25086 | 8.4006 | 0.003874 | 0.005571 | 0.005487 | 5.5e-05 | 0.001164 | 0.28871 | 1.4074 | 2.5475 | 100.79 | 21.12 | 15.393 | 97.3736 | 115.1956 | 90.5144 | 0.003831 | 0.9708 | 0.023691 | 24.4039 | 0.74042 | 0.14923 | 0.015316 | 1.0907 | 1.0907 | 0.005651
2019-03-26 16:15 | 1553631287.6708 | 108 | 0.20648 | 0.002898 | 5.5e-05 | 96.5182 | 89.2847 | 119.8156 | 21.2612 | 0.003148 | 0.002838 | 259.4052 | 0.25071 | 8.4033 | 0.003841 | 0.005605 | 0.005533 | 5.3e-05 | 0.00116 | 0.34036 | 1.3618 | 2.5475 | 100.79 | 21.12 | 15.406 | 97.5695 | 114.37 | 90.0924 | 0.003834 | 0.97193 | 0.02372 | 24.4048 | 0.69679 | 0.14914 | 0.015114 | 1.0777 | 1.0777 | 0.005716
2019-03-26 16:15 | 1553631288.6698 | 109 | 0.20658 | 0.002883 | 5.4e-05 | 96.4426 | 89.1885 | 119.9322 | 21.2612 | 0.003105 | 0.002785 | 259.3665 | 0.25003 | 8.4027 | 0.003831 | 0.005955 | 0.005904 | 5.1e-05 | 0.001162 | 0.32731 | 1.4173 | 2.5477 | 100.79 | 21.12 | 15.414 | 96.54 | 114.1288 | 88.8159 | 0.003836 | 0.97535 | 0.023773 | 24.3737 | 0.69942 | 0.14874 | 0.014548 | 1.0396 | 1.0396 | 0.005909
2019-03-26 16:15 | 1553631289.6699 | 110 | 0.20655 | 0.002863 | 5.4e-05 | 96.4479 | 89.4245 | 120.2439 | 21.2612 | 0.00311 | 0.002801 | 259.3094 | 0.25056 | 8.4018 | 0.003751 | 0.005487 | 0.005371 | 5.6e-05 | 0.001162 | 0.30359 | 1.4425 | 2.5476 | 100.79 | 21.12 | 15.422 | 96.6419 | 112.1829 | 89.2095 | 0.003838 | 0.9743 | 0.023791 | 24.4186 | 0.67628 | 0.14906 | 0.01473 | 1.0535 | 1.0535 | 0.005844
2019-03-26 16:15 | 1553631290.6713 | 111 | 0.20655 | 0.002847 | 5.3e-05 | 96.3095 | 89.278 | 120.1996 | 21.2612 | 0.003101 | 0.002765 | 259.2642 | 0.25111 | 8.4029 | 0.003696 | 0.005567 | 0.005468 | 5.5e-05 | 0.001159 | 0.33843 | 1.4057 | 2.5476 | 100.79 | 21.12 | 15.438 | 96.4374 | 110.8513 | 88.3385 | 0.003842 | 0.97664 | 0.023815 | 24.3847 | 0.67811 | 0.14938 | 0.014752 | 1.0562 | 1.0562 | 0.005841
2019-03-26 16:15 | 1553631291.6714 | 112 | 0.20664 | 0.002831 | 5.2e-05 | 96.3107 | 89.4363 | 120.0341 | 21.2612 | 0.003093 | 0.002765 | 259.2299 | 0.25078 | 8.4012 | 0.003706 | 0.005587 | 0.00549 | 5.7e-05 | 0.001159 | 0.27328 | 1.3603 | 2.5475 | 100.79 | 21.13 | 15.444 | 96.2438 | 111.1045 | 88.3409 | 0.003846 | 0.97664 | 0.023799 | 24.368 | 0.66237 | 0.14919 | 0.014268 | 1.0208 | 1.0208 | 0.006036
2019-03-26 16:15 | 1553631292.67 | 113 | 0.20662 | 0.00281 | 5.2e-05 | 96.2584 | 89.5667 | 120.012 | 21.2612 | 0.003108 | 0.002743 | 259.1863 | 0.25097 | 8.4038 | 0.00368 | 0.005296 | 0.005191 | 5.8e-05 | 0.001162 | 0.29927 | 1.3654 | 2.5475 | 100.79 | 21.13 | 15.459 | 96.5987 | 110.4671 | 87.806 | 0.00385 | 0.97808 | 0.023814 | 24.3478 | 0.64448 | 0.1493 | 0.014387 | 1.0301 | 1.0301 | 0.005986
2019-03-26 16:15 | 1553631293.6715 | 114 | 0.20664 | 0.002787 | 5.1e-05 | 96.1687 | 89.8219 | 120.1439 | 21.2612 | 0.003115 | 0.002769 | 259.1388 | 0.24961 | 8.4032 | 0.003671 | 0.005468 | 0.005391 | 5.6e-05 | 0.001159 | 0.31852 | 1.3341 | 2.5474 | 100.79 | 21.13 | 15.466 | 96.7718 | 110.243 | 88.4393 | 0.003851 | 0.97637 | 0.023806 | 24.3825 | 0.61069 | 0.14849 | 0.014296 | 1.0232 | 1.0232 | 0.005994
2019-03-26 16:15 | 1553631294.6701 | 115 | 0.20668 | 0.002768 | 5e-05 | 96.179 | 89.8242 | 120.1064 | 21.2612 | 0.003112 | 0.002756 | 259.0993 | 0.25089 | 8.4018 | 0.00374 | 0.005506 | 0.005434 | 5.5e-05 | 0.001163 | 0.31868 | 1.3727 | 2.5475 | 100.79 | 21.13 | 15.473 | 96.71 | 111.9152 | 88.1168 | 0.003853 | 0.97724 | 0.023813 | 24.3678 | 0.61142 | 0.14925 | 0.014084 | 1.0084 | 1.0084 | 0.006113
2019-03-26 16:15 | 1553631295.6706 | 116 | 0.20672 | 0.002753 | 4.8e-05 | 96.2283 | 89.8864 | 120.1076 | 21.2612 | 0.003134 | 0.002724 | 259.0553 | 0.2506 | 8.403 | 0.003646 | 0.005459 | 0.005383 | 5.7e-05 | 0.001158 | 0.32189 | 1.3384 | 2.5474 | 100.79 | 21.13 | 15.485 | 97.2215 | 109.6445 | 87.343 | 0.003856 | 0.97934 | 0.023839 | 24.3418 | 0.60981 | 0.14908 | 0.013857 | 0.99327 | 0.99327 | 0.006199
2019-03-26 16:15 | 1553631296.6702 | 117 | 0.20671 | 0.002743 | 4.6e-05 | 96.1719 | 90.2912 | 119.9552 | 21.2612 | 0.00309 | 0.002782 | 259.0012 | 0.25068 | 8.4041 | 0.00369 | 0.005304 | 0.005218 | 5.8e-05 | 0.001149 | 0.3235 | 1.3111 | 2.5474 | 100.79 | 21.13 | 15.491 | 96.1812 | 110.7164 | 88.7463 | 0.003858 | 0.97554 | 0.023778 | 24.3737 | 0.56438 | 0.14913 | 0.013924 | 0.99559 | 0.99559 | 0.006186
2019-03-26 16:15 | 1553631297.6703 | 118 | 0.20675 | 0.002721 | 4.4e-05 | 96.1244 | 90.5392 | 120.1177 | 21.2612 | 0.003036 | 0.002758 | 258.9524 | 0.25077 | 8.4013 | 0.003672 | 0.00523 | 0.005172 | 5.5e-05 | 0.001149 | 0.3235 | 1.3512 | 2.5475 | 100.79 | 21.13 | 15.498 | 94.8594 | 110.2605 | 88.1613 | 0.003859 | 0.97712 | 0.023813 | 24.3705 | 0.53542 | 0.14918 | 0.013702 | 0.98137 | 0.98137 | 0.006278
2019-03-26 16:15 | 1553631298.6693 | 119 | 0.2068 | 0.002702 | 4.3e-05 | 96.1392 | 90.6424 | 120.3465 | 21.2612 | 0.003112 | 0.002714 | 258.8874 | 0.25073 | 8.4013 | 0.003741 | 0.005202 | 0.005139 | 5.5e-05 | 0.001148 | 0.27721 | 1.3404 | 2.5475 | 100.79 | 21.13 | 15.501 | 96.6943 | 111.9451 | 87.1009 | 0.00386 | 0.98 | 0.023871 | 24.3578 | 0.5266 | 0.14916 | 0.013471 | 0.96713 | 0.96713 | 0.00637
2019-03-26 16:15 | 1553631299.6694 | 120 | 0.20682 | 0.002684 | 4.2e-05 | 96.0462 | 90.823 | 119.9207 | 21.2612 | 0.003064 | 0.002738 | 258.8621 | 0.25089 | 8.401 | 0.003745 | 0.005748 | 0.005649 | 5.4e-05 | 0.001146 | 0.33555 | 1.3414 | 2.5474 | 100.79 | 21.13 | 15.504 | 95.5368 | 112.0407 | 87.6871 | 0.003861 | 0.97841 | 0.023809 | 24.3345 | 0.50015 | 0.14925 | 0.013333 | 0.95484 | 0.95484 | 0.006456
2019-03-26 16:15 | 1553631300.6698 | 121 | 0.20683 | 0.002667 | 4.2e-05 | 96.0081 | 90.988 | 119.8348 | 21.2612 | 0.003122 | 0.002745 | 258.8334 | 0.25061 | 8.4016 | 0.0037 | 0.004805 | 0.00464 | 5.5e-05 | 0.00114 | 0.29679 | 1.3838 | 2.5474 | 100.79 | 21.13 | 15.507 | 96.9512 | 110.9547 | 87.8652 | 0.003862 | 0.97792 | 0.023795 | 24.3318 | 0.48036 | 0.14909 | 0.013336 | 0.95442 | 0.95442 | 0.006452
2019-03-26 16:15 | 1553631301.6695 | 122 | 0.20684 | 0.002646 | 4.2e-05 | 96.0057 | 91.0741 | 120.1451 | 21.2612 | 0.00308 | 0.002734 | 258.7917 | 0.25015 | 8.4035 | 0.003731 | 0.00497 | 0.00492 | 5.6e-05 | 0.001136 | 0.29264 | 1.3859 | 2.5475 | 100.79 | 21.13 | 15.509 | 95.9297 | 111.6988 | 87.5896 | 0.003862 | 0.97867 | 0.023835 | 24.354 | 0.47169 | 0.14882 | 0.013304 | 0.95372 | 0.95372 | 0.006444
2019-03-26 16:15 | 1553631302.671 | 123 | 0.2069 | 0.002628 | 4.3e-05 | 95.8756 | 90.8704 | 120.1303 | 21.2612 | 0.00302 | 0.002751 | 258.7496 | 0.25103 | 8.4044 | 0.003854 | 0.005039 | 0.004971 | 5.4e-05 | 0.00114 | 0.30829 | 1.3763 | 2.5476 | 100.79 | 21.13 | 15.512 | 94.4712 | 114.6895 | 88.0096 | 0.003863 | 0.97753 | 0.023819 | 24.3666 | 0.47958 | 0.14934 | 0.012958 | 0.92827 | 0.92827 | 0.006644
2019-03-26 16:15 | 1553631303.6695 | 124 | 0.20693 | 0.002606 | 4.2e-05 | 95.8517 | 91.0146 | 120.2237 | 21.2612 | 0.003078 | 0.002769 | 258.7091 | 0.25075 | 8.4049 | 0.003715 | 0.005547 | 0.005441 | 5.3e-05 | 0.00114 | 0.32372 | 1.3805 | 2.5476 | 100.79 | 21.13 | 15.512 | 95.8899 | 111.3113 | 88.4283 | 0.003863 | 0.9764 | 0.023815 | 24.3902 | 0.46317 | 0.14917 | 0.012802 | 0.91695 | 0.91695 | 0.006719
2019-03-26 16:15 | 1553631304.671 | 125 | 0.20692 | 0.002589 | 4.2e-05 | 95.7253 | 91.3685 | 120.2916 | 21.2612 | 0.003034 | 0.002793 | 258.6613 | 0.25043 | 8.4015 | 0.003734 | 0.005623 | 0.005561 | 5.5e-05 | 0.001136 | 0.31974 | 1.3592 | 2.5475 | 100.79 | 21.13 | 15.512 | 94.8186 | 111.7643 | 89.0056 | 0.003863 | 0.97484 | 0.023802 | 24.4166 | 0.41665 | 0.14898 | 0.012891 | 0.92289 | 0.92289 | 0.006667
2019-03-26 16:15 | 1553631305.6707 | 126 | 0.20693 | 0.002571 | 4.1e-05 | 95.6648 | 91.459 | 120.6948 | 21.2612 | 0.003068 | 0.002739 | 258.6191 | 0.2503 | 8.4024 | 0.00374 | 0.005698 | 0.00564 | 5.6e-05 | 0.001137 | 0.32372 | 1.3811 | 2.5475 | 100.79 | 21.14 | 15.51 | 95.6282 | 111.9254 | 87.7129 | 0.003865 | 0.97834 | 0.023885 | 24.4138 | 0.40215 | 0.1489 | 0.012826 | 0.92148 | 0.92148 | 0.006674
2019-03-26 16:15 | 1553631306.6697 | 127 | 0.20694 | 0.002554 | 4e-05 | 95.7274 | 91.6462 | 121.1213 | 21.2612 | 0.003085 | 0.002749 | 258.607 | 0.25073 | 8.4051 | 0.003711 | 0.005687 | 0.005611 | 5.2e-05 | 0.001133 | 0.30525 | 1.366 | 2.5476 | 100.79 | 21.14 | 15.509 | 96.0504 | 111.2174 | 87.948 | 0.003864 | 0.9777 | 0.023919 | 24.4648 | 0.3897 | 0.14916 | 0.012771 | 0.91887 | 0.91887 | 0.006704
2019-03-26 16:15 | 1553631307.6698 | 128 | 0.20693 | 0.002539 | 4e-05 | 95.7469 | 91.6812 | 121.0434 | 21.2612 | 0.0031 | 0.002716 | 258.5755 | 0.25015 | 8.4017 | 0.003661 | 0.005054 | 0.004985 | 5.2e-05 | 0.00113 | 0.33892 | 1.3457 | 2.5475 | 100.79 | 21.14 | 15.509 | 96.421 | 110.0025 | 87.1454 | 0.003864 | 0.97988 | 0.023938 | 24.4298 | 0.38811 | 0.14882 | 0.012881 | 0.92755 | 0.92755 | 0.006626
2019-03-26 16:15 | 1553631308.6698 | 129 | 0.20699 | 0.002525 | 4e-05 | 95.7103 | 92.0459 | 121.1265 | 21.2612 | 0.003006 | 0.002705 | 258.5215 | 0.25147 | 8.4063 | 0.003678 | 0.00492 | 0.004846 | 5.3e-05 | 0.00113 | 0.30702 | 1.3423 | 2.5475 | 100.79 | 21.14 | 15.506 | 94.1402 | 110.4039 | 86.904 | 0.003864 | 0.98054 | 0.023954 | 24.43 | 0.34918 | 0.1496 | 0.012521 | 0.90223 | 0.90223 | 0.006848
2019-03-26 16:15 | 1553631309.6699 | 130 | 0.20705 | 0.002509 | 4e-05 | 95.6468 | 92.1827 | 121.6107 | 21.2612 | 0.003003 | 0.002703 | 258.4633 | 0.25152 | 8.4004 | 0.003716 | 0.004966 | 0.004878 | 5.6e-05 | 0.001126 | 0.31437 | 1.347 | 2.5475 | 100.79 | 21.14 | 15.505 | 94.0657 | 111.3264 | 86.8549 | 0.003863 | 0.98067 | 0.024004 | 24.4771 | 0.32996 | 0.14963 | 0.012217 | 0.88211 | 0.88211 | 0.007006
2019-03-26 16:15 | 1553631310.6699 | 131 | 0.20706 | 0.00249 | 4e-05 | 95.6099 | 92.2446 | 121.5776 | 21.2612 | 0.003024 | 0.002687 | 258.3946 | 0.25054 | 8.4011 | 0.003687 | 0.005198 | 0.005202 | 5.8e-05 | 0.001126 | 0.33179 | 1.3381 | 2.5475 | 100.79 | 21.14 | 15.503 | 94.5887 | 110.6383 | 86.4661 | 0.003863 | 0.98173 | 0.024014 | 24.4605 | 0.32051 | 0.14905 | 0.012155 | 0.87793 | 0.87793 | 0.007012
2019-03-26 16:15 | 1553631311.67 | 132 | 0.20709 | 0.00248 | 4e-05 | 95.6453 | 92.3575 | 121.6283 | 21.2612 | 0.002955 | 0.002697 | 258.3776 | 0.25116 | 8.403 | 0.003725 | 0.004866 | 0.004815 | 5.5e-05 | 0.001117 | 0.30397 | 1.38 | 2.5476 | 100.79 | 21.14 | 15.501 | 92.9118 | 111.5585 | 86.7026 | 0.003862 | 0.98108 | 0.024011 | 24.4737 | 0.31288 | 0.14941 | 0.012006 | 0.86701 | 0.86701 | 0.007117
2019-03-26 16:15 | 1553631312.67 | 133 | 0.20707 | 0.00246 | 4e-05 | 95.6527 | 92.5424 | 121.4718 | 21.2612 | 0.002989 | 0.002751 | 258.3462 | 0.25077 | 8.4025 | 0.003705 | 0.004974 | 0.004904 | 5.4e-05 | 0.001117 | 0.30646 | 1.3434 | 2.5476 | 100.79 | 21.14 | 15.499 | 93.7317 | 111.077 | 88.0082 | 0.003862 | 0.97754 | 0.023952 | 24.5022 | 0.29568 | 0.14918 | 0.012146 | 0.87501 | 0.87501 | 0.007041
2019-03-26 16:15 | 1553631313.6691 | 134 | 0.2071 | 0.002447 | 4e-05 | 95.6285 | 92.5834 | 121.2744 | 21.2612 | 0.002984 | 0.002689 | 258.2989 | 0.25034 | 8.4034 | 0.003692 | 0.00572 | 0.00557 | 5.3e-05 | 0.001117 | 0.32693 | 1.3395 | 2.5474 | 100.79 | 21.14 | 15.491 | 93.6149 | 110.7558 | 86.5067 | 0.00386 | 0.98162 | 0.023982 | 24.4314 | 0.28945 | 0.14893 | 0.011977 | 0.86394 | 0.86394 | 0.007119
2019-03-26 16:15 | 1553631314.6691 | 135 | 0.20718 | 0.002433 | 4e-05 | 95.6382 | 92.8174 | 121.0972 | 21.2612 | 0.003039 | 0.002702 | 258.2503 | 0.24986 | 8.4018 | 0.003584 | 0.005314 | 0.005253 | 5.4e-05 | 0.001125 | 0.30845 | 1.3469 | 2.5475 | 100.79 | 21.14 | 15.487 | 94.9511 | 108.1405 | 86.8167 | 0.003859 | 0.98077 | 0.023954 | 24.424 | 0.26779 | 0.14864 | 0.011502 | 0.82863 | 0.82863 | 0.007408
2019-03-26 16:15 | 1553631315.6694 | 136 | 0.20717 | 0.002421 | 4e-05 | 95.7165 | 93.1055 | 120.7843 | 21.2612 | 0.002978 | 0.002707 | 258.1992 | 0.25018 | 8.4023 | 0.003613 | 0.00528 | 0.005212 | 5.6e-05 | 0.00113 | 0.32908 | 1.368 | 2.5475 | 100.79 | 21.14 | 15.483 | 93.457 | 108.83 | 86.944 | 0.003858 | 0.98043 | 0.023919 | 24.3968 | 0.24738 | 0.14883 | 0.011564 | 0.83192 | 0.83192 | 0.007389
2019-03-26 16:15 | 1553631316.6693 | 137 | 0.20725 | 0.002414 | 3.8e-05 | 95.7312 | 93.2037 | 120.5296 | 21.2612 | 0.003023 | 0.002735 | 258.1485 | 0.25125 | 8.4007 | 0.003611 | 0.005201 | 0.005093 | 5.6e-05 | 0.001135 | 0.31598 | 1.3718 | 2.5475 | 100.79 | 21.14 | 15.474 | 94.5486 | 108.7946 | 87.606 | 0.003856 | 0.97863 | 0.023872 | 24.3934 | 0.23933 | 0.14947 | 0.011143 | 0.8001 | 0.8001 | 0.007715
2019-03-26 16:15 | 1553631317.6707 | 138 | 0.20717 | 0.002399 | 3.9e-05 | 95.6379 | 93.0584 | 120.3003 | 21.2612 | 0.002955 | 0.002672 | 258.1182 | 0.24991 | 8.4049 | 0.003538 | 0.005145 | 0.005054 | 5.6e-05 | 0.001136 | 0.29789 | 1.348 | 2.5474 | 100.79 | 21.14 | 15.47 | 92.9017 | 107.0071 | 86.0992 | 0.003855 | 0.98273 | 0.023899 | 24.3193 | 0.24456 | 0.14867 | 0.011633 | 0.83625 | 0.83625 | 0.007342
2019-03-26 16:15 | 1553631318.6693 | 139 | 0.20719 | 0.002386 | 3.9e-05 | 95.7005 | 93.318 | 120.0601 | 21.2612 | 0.003072 | 0.002703 | 258.0918 | 0.2505 | 8.4014 | 0.003509 | 0.004975 | 0.004891 | 5.7e-05 | 0.001144 | 0.32045 | 1.353 | 2.5474 | 100.79 | 21.14 | 15.459 | 95.7277 | 106.3136 | 86.8542 | 0.003852 | 0.98067 | 0.02385 | 24.3205 | 0.2255 | 0.14902 | 0.011509 | 0.8256 | 0.8256 | 0.007455
2019-03-26 16:15 | 1553631319.6694 | 140 | 0.20721 | 0.002371 | 3.9e-05 | 95.6371 | 93.3448 | 119.7435 | 21.2612 | 0.002974 | 0.002691 | 258.0401 | 0.25073 | 8.4031 | 0.003475 | 0.005143 | 0.005056 | 5.5e-05 | 0.001144 | 0.29645 | 1.3631 | 2.5474 | 100.79 | 21.14 | 15.454 | 93.3716 | 105.481 | 86.552 | 0.003851 | 0.98149 | 0.023829 | 24.2782 | 0.21699 | 0.14916 | 0.011391 | 0.81641 | 0.81641 | 0.007546
2019-03-26 16:15 | 1553631320.6698 | 141 | 0.20717 | 0.002362 | 3.9e-05 | 95.612 | 93.2739 | 119.9228 | 21.2612 | 0.002995 | 0.002667 | 258.0101 | 0.24999 | 8.4016 | 0.003484 | 0.005489 | 0.005447 | 5.8e-05 | 0.001142 | 0.34379 | 1.3644 | 2.5475 | 100.79 | 21.14 | 15.448 | 93.8718 | 105.7067 | 85.975 | 0.003849 | 0.98307 | 0.023866 | 24.2769 | 0.22145 | 0.14872 | 0.011663 | 0.83728 | 0.83728 | 0.007336
2019-03-26 16:15 | 1553631321.6705 | 142 | 0.20721 | 0.002351 | 4e-05 | 95.6819 | 93.408 | 119.8695 | 21.2612 | 0.002978 | 0.00264 | 257.9762 | 0.25063 | 8.4035 | 0.003483 | 0.004841 | 0.004817 | 5.6e-05 | 0.001137 | 0.28998 | 1.357 | 2.5475 | 100.79 | 21.14 | 15.438 | 93.4587 | 105.6667 | 85.3241 | 0.003847 | 0.98486 | 0.023882 | 24.2495 | 0.21513 | 0.1491 | 0.011464 | 0.82342 | 0.82342 | 0.007478
2019-03-26 16:15 | 1553631322.6705 | 143 | 0.20721 | 0.002347 | 3.9e-05 | 95.6143 | 93.1442 | 119.9864 | 21.2612 | 0.002989 | 0.002661 | 257.9469 | 0.25116 | 8.4018 | 0.003474 | 0.005338 | 0.005259 | 5.5e-05 | 0.001141 | 0.28854 | 1.344 | 2.5474 | 100.79 | 21.14 | 15.433 | 93.73 | 105.4715 | 85.8425 | 0.003845 | 0.98343 | 0.023877 | 24.2789 | 0.23412 | 0.14942 | 0.011464 | 0.82331 | 0.82331 | 0.007495
2019-03-26 16:15 | 1553631323.6696 | 144 | 0.20729 | 0.002337 | 4e-05 | 95.6146 | 93.3294 | 120.004 | 21.2612 | 0.002904 | 0.002632 | 257.9138 | 0.25073 | 8.404 | 0.00356 | 0.004682 | 0.004687 | 5.7e-05 | 0.001137 | 0.29247 | 1.3182 | 2.5475 | 100.79 | 21.14 | 15.425 | 91.6871 | 107.5428 | 85.129 | 0.003843 | 0.98539 | 0.023902 | 24.2565 | 0.21638 | 0.14916 | 0.010972 | 0.7887 | 0.7887 | 0.007811
2019-03-26 16:15 | 1553631324.6706 | 145 | 0.20717 | 0.002333 | 4e-05 | 95.6374 | 93.2857 | 120.1503 | 21.2612 | 0.002984 | 0.002634 | 257.8657 | 0.24992 | 8.4025 | 0.003512 | 0.004857 | 0.004773 | 5.6e-05 | 0.001136 | 0.31741 | 1.273 | 2.5475 | 100.79 | 21.14 | 15.421 | 93.6065 | 106.3767 | 85.1908 | 0.003842 | 0.98522 | 0.023915 | 24.2734 | 0.22269 | 0.14868 | 0.011718 | 0.84286 | 0.84286 | 0.007285
2019-03-26 16:15 | 1553631325.6711 | 146 | 0.20723 | 0.002328 | 3.8e-05 | 95.5866 | 93.2974 | 120.3277 | 21.2612 | 0.002918 | 0.002691 | 257.8197 | 0.25052 | 8.4052 | 0.003514 | 0.004454 | 0.004449 | 5.5e-05 | 0.001128 | 0.29087 | 1.3336 | 2.5475 | 100.79 | 21.14 | 15.419 | 92.0194 | 106.4318 | 86.5472 | 0.003842 | 0.98151 | 0.023887 | 24.3372 | 0.21681 | 0.14903 | 0.011357 | 0.81601 | 0.81601 | 0.007543
2019-03-26 16:15 | 1553631326.6696 | 147 | 0.20718 | 0.002326 | 3.8e-05 | 95.6459 | 93.4322 | 120.5287 | 21.2612 | 0.00292 | 0.002652 | 257.7657 | 0.25104 | 8.4015 | 0.003474 | 0.00546 | 0.005402 | 5.6e-05 | 0.001121 | 0.32709 | 1.3042 | 2.5476 | 100.79 | 21.14 | 15.418 | 92.063 | 105.4613 | 85.613 | 0.003842 | 0.98406 | 0.023938 | 24.3259 | 0.20945 | 0.14935 | 0.011633 | 0.83766 | 0.83766 | 0.007363
2019-03-26 16:15 | 1553631327.6711 | 148 | 0.20726 | 0.00232 | 3.7e-05 | 95.5715 | 93.4988 | 120.972 | 21.2612 | 0.002954 | 0.002651 | 257.7101 | 0.2509 | 8.4045 | 0.003451 | 0.004937 | 0.004837 | 5.6e-05 | 0.001118 | 0.31725 | 1.3432 | 2.5476 | 100.79 | 21.14 | 15.42 | 92.8797 | 104.9052 | 85.6012 | 0.003842 | 0.9841 | 0.023983 | 24.3702 | 0.19612 | 0.14926 | 0.011198 | 0.80782 | 0.80782 | 0.007631
2019-03-26 16:15 | 1553631328.6698 | 149 | 0.2073 | 0.002315 | 3.7e-05 | 95.4532 | 93.7029 | 120.9828 | 21.2612 | 0.002946 | 0.00265 | 257.6704 | 0.24957 | 8.4011 | 0.003442 | 0.004866 | 0.004752 | 5.6e-05 | 0.001111 | 0.3079 | 1.353 | 2.5476 | 100.79 | 21.14 | 15.422 | 92.6912 | 104.6812 | 85.5678 | 0.003843 | 0.98419 | 0.023985 | 24.3701 | 0.16554 | 0.14847 | 0.010953 | 0.79022 | 0.79022 | 0.00776
2019-03-26 16:15 | 1553631329.6688 | 150 | 0.20732 | 0.002309 | 3.7e-05 | 95.4552 | 93.6208 | 121.1469 | 21.2612 | 0.002968 | 0.002659 | 257.6482 | 0.25144 | 8.4043 | 0.003506 | 0.004745 | 0.004672 | 5.6e-05 | 0.00111 | 0.30359 | 1.348 | 2.5475 | 100.79 | 21.14 | 15.427 | 93.2216 | 106.2287 | 85.7785 | 0.003844 | 0.98361 | 0.023994 | 24.3938 | 0.17356 | 0.14958 | 0.01085 | 0.78314 | 0.78314 | 0.007888
2019-03-26 16:15 | 1553631330.6699 | 151 | 0.20732 | 0.002302 | 3.7e-05 | 95.3374 | 93.5092 | 121.2833 | 21.2612 | 0.002941 | 0.002621 | 257.5911 | 0.25041 | 8.4061 | 0.003537 | 0.005037 | 0.004994 | 5.5e-05 | 0.001112 | 0.29231 | 1.3489 | 2.5474 | 100.79 | 21.14 | 15.429 | 92.5849 | 106.9811 | 84.87 | 0.003844 | 0.98611 | 0.024038 | 24.3766 | 0.17319 | 0.14897 | 0.010852 | 0.78466 | 0.78466 | 0.007841
2019-03-26 16:16 | 1553631331.6693 | 152 | 0.20729 | 0.00229 | 3.7e-05 | 95.222 | 93.5755 | 121.4105 | 21.2612 | 0.002915 | 0.002681 | 257.5071 | 0.25046 | 8.4029 | 0.003424 | 0.004618 | 0.004493 | 5.6e-05 | 0.00111 | 0.31492 | 1.3617 | 2.5476 | 100.79 | 21.14 | 15.431 | 91.9396 | 104.2503 | 86.3075 | 0.003845 | 0.98216 | 0.024002 | 24.4383 | 0.15602 | 0.149 | 0.011039 | 0.79701 | 0.79701 | 0.007721
2019-03-26 16:16 | 1553631332.67 | 153 | 0.20729 | 0.00228 | 3.8e-05 | 95.1595 | 93.5284 | 121.4931 | 21.2612 | 0.002909 | 0.002645 | 257.4612 | 0.25065 | 8.4016 | 0.00344 | 0.005071 | 0.004975 | 5.6e-05 | 0.001107 | 0.30646 | 1.3114 | 2.5476 | 100.79 | 21.14 | 15.434 | 91.8035 | 104.6269 | 85.4521 | 0.003846 | 0.9845 | 0.024039 | 24.4175 | 0.15464 | 0.14911 | 0.011044 | 0.79861 | 0.79861 | 0.007711
2019-03-26 16:16 | 1553631333.6694 | 154 | 0.20729 | 0.00227 | 3.8e-05 | 95.0175 | 93.4638 | 121.5664 | 21.2612 | 0.002912 | 0.00263 | 257.4268 | 0.25037 | 8.4027 | 0.003451 | 0.004413 | 0.004335 | 5.6e-05 | 0.001105 | 0.29142 | 1.3539 | 2.5476 | 100.79 | 21.14 | 15.433 | 91.8738 | 104.8935 | 85.0761 | 0.003845 | 0.98554 | 0.024059 | 24.4121 | 0.14746 | 0.14895 | 0.011074 | 0.80135 | 0.80135 | 0.007676
2019-03-26 16:16 | 1553631334.6694 | 155 | 0.20726 | 0.002265 | 3.8e-05 | 94.9142 | 93.2181 | 121.8269 | 21.2612 | 0.002885 | 0.002664 | 257.377 | 0.25049 | 8.4027 | 0.003511 | 0.005039 | 0.005018 | 5.5e-05 | 0.001107 | 0.33268 | 1.3262 | 2.5474 | 100.79 | 21.14 | 15.431 | 91.2297 | 106.3609 | 85.9072 | 0.003845 | 0.98326 | 0.024057 | 24.4666 | 0.16128 | 0.14901 | 0.011229 | 0.81254 | 0.81254 | 0.007574
2019-03-26 16:16 | 1553631335.6702 | 156 | 0.20731 | 0.00226 | 3.9e-05 | 94.8592 | 93.1062 | 121.6283 | 21.2612 | 0.002896 | 0.002609 | 257.3335 | 0.25091 | 8.4052 | 0.003561 | 0.004904 | 0.004817 | 5.4e-05 | 0.001095 | 0.30757 | 1.3104 | 2.5475 | 100.79 | 21.14 | 15.425 | 91.4837 | 107.5607 | 84.5898 | 0.003843 | 0.98688 | 0.024082 | 24.4017 | 0.16685 | 0.14927 | 0.010957 | 0.79363 | 0.79363 | 0.007768
2019-03-26 16:16 | 1553631336.6695 | 157 | 0.20736 | 0.002251 | 3.8e-05 | 94.701 | 93.1179 | 121.6422 | 21.2612 | 0.002913 | 0.002626 | 257.2812 | 0.24996 | 8.4022 | 0.003423 | 0.005114 | 0.005035 | 5.1e-05 | 0.001096 | 0.30431 | 1.3098 | 2.5474 | 100.79 | 21.14 | 15.423 | 91.8975 | 104.2289 | 84.9955 | 0.003843 | 0.98576 | 0.024069 | 24.4169 | 0.15079 | 0.1487 | 0.010678 | 0.77299 | 0.77299 | 0.007945
2019-03-26 16:16 | 1553631337.6692 | 158 | 0.20731 | 0.002243 | 3.8e-05 | 94.6204 | 93.1045 | 121.3577 | 21.2612 | 0.002897 | 0.002635 | 257.2397 | 0.24972 | 8.4047 | 0.003521 | 0.004832 | 0.004727 | 5e-05 | 0.001091 | 0.3131 | 1.347 | 2.5477 | 100.79 | 21.14 | 15.414 | 91.5237 | 106.6057 | 85.1971 | 0.003841 | 0.9852 | 0.024034 | 24.3952 | 0.14446 | 0.14856 | 0.010958 | 0.79217 | 0.79217 | 0.007745
2019-03-26 16:16 | 1553631338.6706 | 159 | 0.20736 | 0.002242 | 3.8e-05 | 94.4489 | 93.0685 | 121.4718 | 21.2612 | 0.002923 | 0.002605 | 257.1877 | 0.25026 | 8.4035 | 0.003455 | 0.005008 | 0.004922 | 5e-05 | 0.001092 | 0.30541 | 1.3652 | 2.5474 | 100.79 | 21.14 | 15.408 | 92.1501 | 104.9948 | 84.4806 | 0.003839 | 0.98718 | 0.02407 | 24.3823 | 0.13169 | 0.14888 | 0.010656 | 0.77147 | 0.77147 | 0.00797
2019-03-26 16:16 | 1553631339.6693 | 160 | 0.20735 | 0.002233 | 3.7e-05 | 94.3261 | 92.8463 | 121.5617 | 21.2612 | 0.002873 | 0.002605 | 257.1366 | 0.25102 | 8.4034 | 0.003467 | 0.005098 | 0.005047 | 4.9e-05 | 0.001097 | 0.27632 | 1.3451 | 2.5474 | 100.79 | 21.14 | 15.402 | 90.9396 | 105.2882 | 84.4885 | 0.003838 | 0.98716 | 0.024078 | 24.3916 | 0.14144 | 0.14933 | 0.01075 | 0.77859 | 0.77859 | 0.007921
2019-03-26 16:16 | 1553631340.6707 | 161 | 0.20736 | 0.002228 | 3.6e-05 | 94.1986 | 92.6419 | 121.2807 | 21.2612 | 0.00287 | 0.002631 | 257.1169 | 0.25006 | 8.4016 | 0.003502 | 0.005133 | 0.005021 | 5.2e-05 | 0.001103 | 0.29502 | 1.3282 | 2.5475 | 100.79 | 21.14 | 15.391 | 90.8553 | 106.1348 | 85.1151 | 0.003835 | 0.98543 | 0.024029 | 24.3847 | 0.14905 | 0.14876 | 0.010678 | 0.77187 | 0.77187 | 0.00796
2019-03-26 16:16 | 1553631341.6704 | 162 | 0.20737 | 0.00222 | 3.6e-05 | 94.0357 | 92.4507 | 121.1756 | 21.2612 | 0.002919 | 0.002633 | 257.0814 | 0.25012 | 8.4039 | 0.003476 | 0.005079 | 0.005069 | 5.1e-05 | 0.001111 | 0.31293 | 1.3478 | 2.5475 | 100.79 | 21.14 | 15.387 | 92.0459 | 105.5086 | 85.1579 | 0.003834 | 0.98531 | 0.024018 | 24.3756 | 0.15205 | 0.1488 | 0.01063 | 0.76797 | 0.76797 | 0.008002
2019-03-26 16:16 | 1553631342.6704 | 163 | 0.20736 | 0.002206 | 3.7e-05 | 93.9699 | 92.2752 | 121.2239 | 21.2612 | 0.002926 | 0.002578 | 257.032 | 0.25069 | 8.402 | 0.003527 | 0.005476 | 0.005299 | 5.1e-05 | 0.001115 | 0.32892 | 1.3877 | 2.5476 | 100.79 | 21.14 | 15.385 | 92.2101 | 106.7559 | 83.8253 | 0.003834 | 0.98899 | 0.024067 | 24.3351 | 0.16278 | 0.14913 | 0.010694 | 0.77417 | 0.77417 | 0.007956
2019-03-26 16:16 | 1553631343.6689 | 164 | 0.20735 | 0.002201 | 3.7e-05 | 93.8016 | 92.3673 | 120.9061 | 21.2612 | 0.002886 | 0.002649 | 256.9956 | 0.25095 | 8.403 | 0.003449 | 0.004946 | 0.004941 | 5.3e-05 | 0.00112 | 0.30934 | 1.3944 | 2.5475 | 100.79 | 21.14 | 15.382 | 91.2516 | 104.8462 | 85.5341 | 0.003833 | 0.98428 | 0.023978 | 24.3613 | 0.13783 | 0.14929 | 0.010792 | 0.77838 | 0.77838 | 0.007921
2019-03-26 16:16 | 1553631344.669 | 165 | 0.2074 | 0.002189 | 3.7e-05 | 93.7753 | 92.4016 | 120.5948 | 21.2612 | 0.002862 | 0.002574 | 256.9639 | 0.25023 | 8.3992 | 0.003535 | 0.004678 | 0.004576 | 5e-05 | 0.001122 | 0.34147 | 1.3828 | 2.5476 | 100.79 | 21.14 | 15.382 | 90.6763 | 106.9449 | 83.7378 | 0.003833 | 0.98923 | 0.024008 | 24.2689 | 0.132 | 0.14886 | 0.010506 | 0.75867 | 0.75867 | 0.008104
2019-03-26 16:16 | 1553631345.6686 | 166 | 0.20738 | 0.00219 | 3.7e-05 | 93.6608 | 92.3293 | 120.761 | 21.2612 | 0.002921 | 0.002548 | 256.9491 | 0.25048 | 8.403 | 0.003467 | 0.005257 | 0.005199 | 5.3e-05 | 0.001123 | 0.28926 | 1.3505 | 2.5476 | 100.79 | 21.14 | 15.387 | 92.0905 | 105.2955 | 83.1225 | 0.003834 | 0.99094 | 0.024045 | 24.2647 | 0.12807 | 0.14901 | 0.010592 | 0.76607 | 0.76607 | 0.008033
2019-03-26 16:16 | 1553631346.6697 | 167 | 0.20738 | 0.002186 | 3.7e-05 | 93.6516 | 92.3528 | 120.4972 | 21.2612 | 0.002873 | 0.002621 | 256.9168 | 0.25059 | 8.4034 | 0.003505 | 0.005102 | 0.004982 | 5.5e-05 | 0.001133 | 0.30182 | 1.332 | 2.5474 | 100.79 | 21.14 | 15.391 | 90.9428 | 106.2057 | 84.8731 | 0.003835 | 0.9861 | 0.02396 | 24.2976 | 0.12492 | 0.14908 | 0.010597 | 0.76375 | 0.76375 | 0.008061
2019-03-26 16:16 | 1553631347.6701 | 168 | 0.20737 | 0.002184 | 3.7e-05 | 93.556 | 92.1888 | 120.8793 | 21.2612 | 0.002821 | 0.002659 | 256.8711 | 0.25013 | 8.4031 | 0.003518 | 0.005368 | 0.005336 | 5.1e-05 | 0.001126 | 0.33339 | 1.3653 | 2.5475 | 100.79 | 21.14 | 15.394 | 89.6955 | 106.5286 | 85.7886 | 0.003836 | 0.98358 | 0.023967 | 24.3672 | 0.13168 | 0.1488 | 0.010658 | 0.76835 | 0.76835 | 0.007998
2019-03-26 16:16 | 1553631348.6688 | 169 | 0.20737 | 0.002181 | 3.8e-05 | 93.5229 | 92.1747 | 120.9131 | 21.2612 | 0.002846 | 0.002602 | 256.8405 | 0.25035 | 8.4018 | 0.003602 | 0.005506 | 0.00538 | 5.5e-05 | 0.001121 | 0.28279 | 1.363 | 2.5475 | 100.79 | 21.14 | 15.404 | 90.2885 | 108.5677 | 84.408 | 0.003838 | 0.98738 | 0.024017 | 24.3237 | 0.12989 | 0.14894 | 0.010668 | 0.77066 | 0.77066 | 0.007982
2019-03-26 16:16 | 1553631349.6688 | 170 | 0.20737 | 0.002177 | 3.8e-05 | 93.4249 | 92.0355 | 120.492 | 21.2612 | 0.002925 | 0.00261 | 256.7916 | 0.24956 | 8.402 | 0.003511 | 0.005186 | 0.005112 | 5.3e-05 | 0.001122 | 0.323 | 1.3662 | 2.5475 | 100.79 | 21.14 | 15.409 | 92.1828 | 106.3616 | 84.6148 | 0.003839 | 0.98681 | 0.023968 | 24.2883 | 0.13403 | 0.14846 | 0.010713 | 0.77226 | 0.77226 | 0.00794
2019-03-26 16:16 | 1553631350.6698 | 171 | 0.20737 | 0.002182 | 3.8e-05 | 93.3441 | 91.7461 | 120.5848 | 21.2612 | 0.002899 | 0.002623 | 256.7521 | 0.25026 | 8.4036 | 0.003557 | 0.005232 | 0.00516 | 5.5e-05 | 0.001128 | 0.28296 | 1.3752 | 2.5475 | 100.79 | 21.14 | 15.42 | 91.5745 | 107.4772 | 84.9224 | 0.003842 | 0.98596 | 0.023967 | 24.3081 | 0.15445 | 0.14888 | 0.010665 | 0.76877 | 0.76877 | 0.007998
2019-03-26 16:16 | 1553631351.6693 | 172 | 0.20737 | 0.002183 | 3.9e-05 | 93.2579 | 91.8193 | 120.4458 | 21.2612 | 0.002902 | 0.002573 | 256.7067 | 0.24913 | 8.4018 | 0.003535 | 0.005026 | 0.004901 | 5.5e-05 | 0.001136 | 0.32068 | 1.3811 | 2.5474 | 100.79 | 21.14 | 15.424 | 91.6369 | 106.9513 | 83.7106 | 0.003843 | 0.98931 | 0.023994 | 24.253 | 0.13905 | 0.14821 | 0.010667 | 0.76976 | 0.76976 | 0.007952
2019-03-26 16:16 | 1553631352.669 | 173 | 0.20735 | 0.002179 | 4e-05 | 93.1835 | 91.6792 | 120.3296 | 21.2612 | 0.002868 | 0.002662 | 256.6423 | 0.24965 | 8.4033 | 0.003585 | 0.004995 | 0.004974 | 5.3e-05 | 0.001139 | 0.31598 | 1.3954 | 2.5476 | 100.79 | 21.14 | 15.429 | 90.8115 | 108.1437 | 85.8627 | 0.003844 | 0.98338 | 0.02391 | 24.3142 | 0.14558 | 0.14852 | 0.010789 | 0.77582 | 0.77582 | 0.007906
2019-03-26 16:16 | 1553631353.6691 | 174 | 0.20735 | 0.002184 | 4.1e-05 | 93.0521 | 91.7094 | 120.6115 | 21.2612 | 0.002862 | 0.002663 | 256.5837 | 0.24984 | 8.402 | 0.003492 | 0.005254 | 0.005182 | 5.8e-05 | 0.001145 | 0.29264 | 1.3497 | 2.5474 | 100.79 | 21.14 | 15.436 | 90.6621 | 105.902 | 85.8798 | 0.003846 | 0.98333 | 0.023938 | 24.3433 | 0.13001 | 0.14863 | 0.010783 | 0.77616 | 0.77616 | 0.007909
2019-03-26 16:16 | 1553631354.6694 | 175 | 0.20736 | 0.002183 | 4.1e-05 | 92.9435 | 91.7619 | 120.7106 | 21.2612 | 0.002909 | 0.002603 | 256.5329 | 0.25066 | 8.4021 | 0.003655 | 0.005028 | 0.004969 | 5.4e-05 | 0.001147 | 0.32477 | 1.3556 | 2.5474 | 100.79 | 21.14 | 15.438 | 91.7972 | 109.8501 | 84.4337 | 0.003847 | 0.98731 | 0.023996 | 24.3042 | 0.11445 | 0.14912 | 0.010769 | 0.77706 | 0.77706 | 0.007925
2019-03-26 16:16 | 1553631355.6691 | 176 | 0.20736 | 0.002186 | 4.1e-05 | 92.8558 | 91.5091 | 120.8054 | 21.2612 | 0.002915 | 0.002626 | 256.4646 | 0.24941 | 8.4007 | 0.003655 | 0.005621 | 0.005598 | 5.8e-05 | 0.001148 | 0.30198 | 1.3473 | 2.5475 | 100.79 | 21.14 | 15.44 | 91.9474 | 109.8645 | 84.9934 | 0.003847 | 0.98577 | 0.023986 | 24.3327 | 0.13068 | 0.14838 | 0.010713 | 0.77265 | 0.77265 | 0.007931
2019-03-26 16:16 | 1553631356.6695 | 177 | 0.20736 | 0.002182 | 4.1e-05 | 92.6454 | 91.4953 | 120.8495 | 21.2612 | 0.002864 | 0.002581 | 256.4138 | 0.2498 | 8.4018 | 0.003614 | 0.005196 | 0.005103 | 5.7e-05 | 0.001146 | 0.32709 | 1.3464 | 2.5475 | 100.79 | 21.15 | 15.442 | 90.7083 | 108.8687 | 83.9095 | 0.00385 | 0.98876 | 0.024027 | 24.3003 | 0.11173 | 0.14861 | 0.010757 | 0.77716 | 0.77716 | 0.007897
2019-03-26 16:16 | 1553631357.6692 | 178 | 0.20736 | 0.002185 | 4.1e-05 | 92.6377 | 91.2677 | 120.4216 | 21.2612 | 0.002936 | 0.00263 | 256.3492 | 0.24987 | 8.4046 | 0.003632 | 0.004821 | 0.00476 | 6e-05 | 0.001149 | 0.31564 | 1.3366 | 2.5476 | 100.79 | 21.15 | 15.442 | 92.4596 | 109.2873 | 85.0845 | 0.00385 | 0.98552 | 0.023945 | 24.2972 | 0.13326 | 0.14865 | 0.010715 | 0.77153 | 0.77153 | 0.007957
2019-03-26 16:16 | 1553631358.6692 | 179 | 0.20741 | 0.00219 | 4.1e-05 | 92.5229 | 91.0653 | 120.089 | 21.2612 | 0.002806 | 0.002648 | 256.293 | 0.24864 | 8.4033 | 0.003615 | 0.004956 | 0.004822 | 5.4e-05 | 0.001147 | 0.35762 | 1.329 | 2.5476 | 100.79 | 21.15 | 15.443 | 89.3133 | 108.8781 | 85.5278 | 0.00385 | 0.9843 | 0.023897 | 24.2786 | 0.14203 | 0.14792 | 0.010455 | 0.75123 | 0.75123 | 0.008132
2019-03-26 16:16 | 1553631359.6693 | 180 | 0.2074 | 0.002189 | 4.1e-05 | 92.3074 | 91.0321 | 120.2749 | 21.2612 | 0.002868 | 0.002602 | 256.2645 | 0.24948 | 8.4046 | 0.003554 | 0.005456 | 0.005422 | 5.8e-05 | 0.001147 | 0.34036 | 1.3617 | 2.5476 | 100.79 | 21.15 | 15.445 | 90.8153 | 107.3947 | 84.4135 | 0.003851 | 0.98736 | 0.023953 | 24.2596 | 0.12444 | 0.14842 | 0.010507 | 0.75678 | 0.75678 | 0.0081
2019-03-26 16:16 | 1553631360.6687 | 181 | 0.20742 | 0.002189 | 4.1e-05 | 92.2188 | 91.0868 | 120.6502 | 21.2612 | 0.002899 | 0.002707 | 256.2275 | 0.25035 | 8.4034 | 0.003523 | 0.00557 | 0.005487 | 5.8e-05 | 0.001142 | 0.32012 | 1.3459 | 2.5475 | 100.79 | 21.15 | 15.446 | 91.5541 | 106.6475 | 86.9478 | 0.003851 | 0.98042 | 0.023906 | 24.3834 | 0.11047 | 0.14894 | 0.010363 | 0.74488 | 0.74488 | 0.008258
2019-03-26 16:16 | 1553631361.6694 | 182 | 0.20743 | 0.002182 | 4.1e-05 | 92.158 | 90.8388 | 120.9631 | 21.2612 | 0.002835 | 0.002617 | 256.1912 | 0.25064 | 8.4051 | 0.00361 | 0.004806 | 0.004762 | 5.9e-05 | 0.001136 | 0.30198 | 1.3336 | 2.5475 | 100.79 | 21.15 | 15.451 | 90.0276 | 108.756 | 84.7771 | 0.003852 | 0.98636 | 0.024009 | 24.3413 | 0.12897 | 0.14911 | 0.010323 | 0.74525 | 0.74525 | 0.008263
2019-03-26 16:16 | 1553631362.6695 | 183 | 0.20744 | 0.002178 | 4.1e-05 | 92.1887 | 90.6317 | 121.0885 | 21.2612 | 0.002917 | 0.002626 | 256.1482 | 0.25085 | 8.4034 | 0.0036 | 0.004986 | 0.00484 | 5.8e-05 | 0.00114 | 0.28672 | 1.402 | 2.5477 | 100.79 | 21.15 | 15.453 | 91.9895 | 108.5063 | 85.0035 | 0.003853 | 0.98574 | 0.024014 | 24.3616 | 0.15236 | 0.14923 | 0.010245 | 0.73969 | 0.73969 | 0.008332
2019-03-26 16:16 | 1553631363.6685 | 184 | 0.20742 | 0.002178 | 4.2e-05 | 92.0275 | 90.6599 | 121.0854 | 21.2612 | 0.002883 | 0.002597 | 256.0891 | 0.24972 | 8.4013 | 0.00357 | 0.00595 | 0.005768 | 5.7e-05 | 0.001141 | 0.32853 | 1.3779 | 2.5475 | 100.79 | 21.15 | 15.459 | 91.1733 | 107.7829 | 84.2925 | 0.003854 | 0.9877 | 0.024038 | 24.3371 | 0.13392 | 0.14856 | 0.0104 | 0.75158 | 0.75158 | 0.008164
2019-03-26 16:16 | 1553631364.6699 | 185 | 0.20746 | 0.002168 | 4.4e-05 | 92.0086 | 90.6384 | 120.7484 | 21.2612 | 0.0028 | 0.002605 | 256.0625 | 0.24908 | 8.401 | 0.00355 | 0.004876 | 0.004872 | 5.7e-05 | 0.001137 | 0.29479 | 1.379 | 2.5476 | 100.79 | 21.15 | 15.462 | 89.1681 | 107.3095 | 84.4768 | 0.003855 | 0.98719 | 0.023998 | 24.3094 | 0.13421 | 0.14818 | 0.01017 | 0.73362 | 0.73362 | 0.008342
2019-03-26 16:16 | 1553631365.6697 | 186 | 0.20744 | 0.00216 | 4.5e-05 | 91.9281 | 90.5513 | 120.5176 | 21.2612 | 0.002838 | 0.002588 | 256.0183 | 0.24988 | 8.4047 | 0.003582 | 0.004721 | 0.004624 | 5.6e-05 | 0.00114 | 0.31365 | 1.3895 | 2.5476 | 100.79 | 21.15 | 15.467 | 90.0987 | 108.0713 | 84.0843 | 0.003856 | 0.98827 | 0.023988 | 24.2729 | 0.13497 | 0.14865 | 0.010289 | 0.74189 | 0.74189 | 0.008275
2019-03-26 16:16 | 1553631366.6697 | 187 | 0.20746 | 0.002159 | 4.6e-05 | 91.8442 | 90.5258 | 120.8906 | 21.2612 | 0.00283 | 0.002617 | 255.9862 | 0.24947 | 8.4017 | 0.003582 | 0.005135 | 0.005098 | 5.6e-05 | 0.001131 | 0.32189 | 1.3458 | 2.5476 | 100.79 | 21.15 | 15.478 | 89.8901 | 108.0815 | 84.7779 | 0.003859 | 0.98636 | 0.024002 | 24.334 | 0.12933 | 0.14841 | 0.010143 | 0.73166 | 0.73166 | 0.008377
2019-03-26 16:16 | 1553631367.669 | 188 | 0.20748 | 0.002151 | 4.5e-05 | 91.8903 | 90.4286 | 121.1349 | 21.2612 | 0.002832 | 0.002613 | 255.9508 | 0.25086 | 8.4043 | 0.003556 | 0.004854 | 0.004729 | 5.4e-05 | 0.001129 | 0.29374 | 1.3187 | 2.5474 | 100.79 | 21.15 | 15.486 | 89.9399 | 107.4481 | 84.6704 | 0.003861 | 0.98665 | 0.02403 | 24.3549 | 0.14342 | 0.14923 | 0.010058 | 0.72641 | 0.72641 | 0.008485
2019-03-26 16:16 | 1553631368.6698 | 189 | 0.20748 | 0.002153 | 4.6e-05 | 91.7616 | 90.2074 | 121.6854 | 21.2612 | 0.002848 | 0.00254 | 255.879 | 0.25033 | 8.4043 | 0.003539 | 0.004973 | 0.004942 | 5.5e-05 | 0.001138 | 0.3131 | 1.3172 | 2.5476 | 100.79 | 21.15 | 15.499 | 90.3393 | 107.0347 | 82.9227 | 0.003864 | 0.9915 | 0.024143 | 24.3505 | 0.15279 | 0.14892 | 0.010037 | 0.72825 | 0.72825 | 0.008445
2019-03-26 16:16 | 1553631369.6702 | 190 | 0.20752 | 0.002146 | 4.6e-05 | 91.6639 | 90.4411 | 121.6182 | 21.2612 | 0.002848 | 0.00261 | 255.8161 | 0.25082 | 8.4052 | 0.003538 | 0.004815 | 0.004749 | 5.9e-05 | 0.001136 | 0.29087 | 1.3068 | 2.5475 | 100.79 | 21.15 | 15.505 | 90.3434 | 107.0197 | 84.6016 | 0.003866 | 0.98684 | 0.02408 | 24.4011 | 0.12012 | 0.14921 | 0.009819 | 0.71051 | 0.71051 | 0.008673
2019-03-26 16:16 | 1553631370.669 | 191 | 0.20748 | 0.002141 | 4.5e-05 | 91.6718 | 90.4471 | 121.0263 | 21.2612 | 0.002845 | 0.002586 | 255.7513 | 0.24973 | 8.4017 | 0.003567 | 0.005131 | 0.005051 | 5.6e-05 | 0.001139 | 0.32421 | 1.3427 | 2.5475 | 100.79 | 21.15 | 15.511 | 90.2581 | 107.7127 | 84.0405 | 0.003867 | 0.9884 | 0.02404 | 24.3226 | 0.1203 | 0.14857 | 0.010053 | 0.72631 | 0.72631 | 0.008448
2019-03-26 16:16 | 1553631371.6693 | 192 | 0.20752 | 0.002139 | 4.5e-05 | 91.7081 | 90.5275 | 120.6179 | 21.2612 | 0.002811 | 0.002595 | 255.7219 | 0.25013 | 8.4016 | 0.003574 | 0.004766 | 0.004768 | 5.6e-05 | 0.00114 | 0.32388 | 1.3454 | 2.5473 | 100.79 | 21.15 | 15.52 | 89.4361 | 107.8853 | 84.238 | 0.00387 | 0.98785 | 0.023993 | 24.2882 | 0.1159 | 0.1488 | 0.009802 | 0.70678 | 0.70678 | 0.008695
2019-03-26 16:16 | 1553631372.6702 | 193 | 0.20751 | 0.002137 | 4.5e-05 | 91.6795 | 90.4797 | 120.3195 | 21.2612 | 0.002816 | 0.002638 | 255.6803 | 0.24957 | 8.3999 | 0.003554 | 0.00527 | 0.005076 | 5.5e-05 | 0.001141 | 0.30663 | 1.3774 | 2.5476 | 100.79 | 21.15 | 15.524 | 89.5611 | 107.4049 | 85.2743 | 0.003871 | 0.98499 | 0.023929 | 24.2933 | 0.11783 | 0.14847 | 0.009893 | 0.71146 | 0.71146 | 0.008618
2019-03-26 16:16 | 1553631373.6703 | 194 | 0.20752 | 0.002124 | 4.5e-05 | 91.6984 | 90.4844 | 120.2918 | 21.2612 | 0.002813 | 0.00262 | 255.6458 | 0.25032 | 8.4058 | 0.003601 | 0.005146 | 0.005075 | 5.8e-05 | 0.001141 | 0.29983 | 1.3977 | 2.5475 | 100.79 | 21.15 | 15.53 | 89.4985 | 108.5313 | 84.8397 | 0.003872 | 0.98619 | 0.02394 | 24.2758 | 0.1192 | 0.14891 | 0.009873 | 0.71035 | 0.71035 | 0.008658
2019-03-26 16:16 | 1553631374.6695 | 195 | 0.20756 | 0.00212 | 4.6e-05 | 91.7108 | 90.3271 | 120.5122 | 21.2612 | 0.002854 | 0.002606 | 255.5618 | 0.25025 | 8.4036 | 0.003552 | 0.004714 | 0.004745 | 5.6e-05 | 0.001143 | 0.32676 | 1.3533 | 2.5474 | 100.79 | 21.15 | 15.533 | 90.48 | 107.3568 | 84.5212 | 0.003873 | 0.98707 | 0.023973 | 24.2872 | 0.13598 | 0.14887 | 0.009612 | 0.69239 | 0.69239 | 0.00888
2019-03-26 16:16 | 1553631375.6681 | 196 | 0.20756 | 0.002111 | 4.6e-05 | 91.5738 | 90.2159 | 120.7146 | 21.2612 | 0.002871 | 0.002589 | 255.5344 | 0.2503 | 8.4035 | 0.003621 | 0.005392 | 0.005298 | 6e-05 | 0.001143 | 0.31332 | 1.4051 | 2.5476 | 100.79 | 21.15 | 15.535 | 90.8811 | 109.0225 | 84.0975 | 0.003873 | 0.98824 | 0.024007 | 24.2932 | 0.13364 | 0.1489 | 0.009645 | 0.69578 | 0.69578 | 0.008839
2019-03-26 16:16 | 1553631376.6685 | 197 | 0.20755 | 0.002103 | 4.8e-05 | 91.5358 | 90.1822 | 120.5477 | 21.2612 | 0.002877 | 0.002653 | 255.4941 | 0.25052 | 8.4047 | 0.003557 | 0.004836 | 0.004701 | 5.8e-05 | 0.001139 | 0.32997 | 1.358 | 2.5474 | 100.79 | 21.15 | 15.534 | 91.0264 | 107.4656 | 85.652 | 0.003873 | 0.98396 | 0.023939 | 24.3291 | 0.13325 | 0.14904 | 0.009678 | 0.69608 | 0.69608 | 0.008843
2019-03-26 16:16 | 1553631377.6693 | 198 | 0.20755 | 0.002092 | 4.8e-05 | 91.4428 | 90.2713 | 120.4843 | 21.2612 | 0.002855 | 0.002605 | 255.4566 | 0.24991 | 8.404 | 0.003507 | 0.00465 | 0.004621 | 6e-05 | 0.001137 | 0.32853 | 1.3728 | 2.5475 | 100.79 | 21.15 | 15.534 | 90.505 | 106.2695 | 84.4885 | 0.003873 | 0.98716 | 0.023971 | 24.2833 | 0.11533 | 0.14867 | 0.009714 | 0.69959 | 0.69959 | 0.008777
2019-03-26 16:16 | 1553631378.6698 | 199 | 0.20755 | 0.002089 | 4.8e-05 | 91.3681 | 90.2488 | 120.4719 | 21.2612 | 0.002865 | 0.002553 | 255.4323 | 0.24983 | 8.4005 | 0.003487 | 0.004547 | 0.004463 | 6e-05 | 0.001137 | 0.323 | 1.3721 | 2.5475 | 100.79 | 21.15 | 15.532 | 90.7505 | 105.7807 | 83.2307 | 0.003873 | 0.99064 | 0.024012 | 24.2393 | 0.11025 | 0.14863 | 0.009733 | 0.7022 | 0.7022 | 0.008742
2019-03-26 16:16 | 1553631379.6693 | 200 | 0.20757 | 0.002082 | 4.8e-05 | 91.2125 | 89.9748 | 121.0404 | 21.2612 | 0.002828 | 0.002574 | 255.3851 | 0.2502 | 8.4028 | 0.003563 | 0.00513 | 0.005042 | 5.9e-05 | 0.001138 | 0.28512 | 1.3445 | 2.5476 | 100.79 | 21.15 | 15.524 | 89.8417 | 107.6207 | 83.7496 | 0.003871 | 0.9892 | 0.024051 | 24.3141 | 0.1222 | 0.14884 | 0.009581 | 0.69229 | 0.69229 | 0.00888
2019-03-26 16:16 | 1553631380.6688 | 201 | 0.20755 | 0.002077 | 4.8e-05 | 91.0823 | 89.7819 | 121.5736 | 21.2612 | 0.002795 | 0.002579 | 255.318 | 0.24989 | 8.4041 | 0.003495 | 0.005204 | 0.005112 | 5.8e-05 | 0.001132 | 0.33644 | 1.3334 | 2.5474 | 100.79 | 21.15 | 15.522 | 89.0627 | 105.9791 | 83.8667 | 0.00387 | 0.98888 | 0.0241 | 24.3716 | 0.12863 | 0.14866 | 0.00973 | 0.7045 | 0.7045 | 0.008715
2019-03-26 16:16 | 1553631381.6695 | 202 | 0.20757 | 0.002072 | 4.9e-05 | 90.977 | 89.5556 | 121.8187 | 21.2612 | 0.002821 | 0.002531 | 255.2462 | 0.25037 | 8.4065 | 0.003531 | 0.00488 | 0.004806 | 5.9e-05 | 0.001136 | 0.34667 | 1.3681 | 2.5475 | 100.79 | 21.15 | 15.519 | 89.6774 | 106.8527 | 82.7191 | 0.003869 | 0.99206 | 0.024164 | 24.3569 | 0.14084 | 0.14895 | 0.009623 | 0.69848 | 0.69848 | 0.008807
2019-03-26 16:16 | 1553631382.6695 | 203 | 0.20758 | 0.002066 | 5e-05 | 90.8061 | 89.3173 | 122.1713 | 21.2612 | 0.002831 | 0.00257 | 255.2139 | 0.24964 | 8.4051 | 0.00356 | 0.005114 | 0.005054 | 6.1e-05 | 0.001133 | 0.32853 | 1.3814 | 2.5476 | 100.79 | 21.15 | 15.517 | 89.9277 | 107.5435 | 83.6474 | 0.003869 | 0.98948 | 0.024167 | 24.4239 | 0.14786 | 0.14851 | 0.009548 | 0.69311 | 0.69311 | 0.008849
2019-03-26 16:16 | 1553631383.6686 | 204 | 0.20761 | 0.002058 | 5.1e-05 | 90.7122 | 89.2669 | 122.314 | 21.2612 | 0.002795 | 0.002589 | 255.1634 | 0.2496 | 8.4008 | 0.003524 | 0.005106 | 0.005069 | 5.7e-05 | 0.001123 | 0.31885 | 1.3225 | 2.5474 | 100.79 | 21.15 | 15.516 | 89.0578 | 106.6812 | 84.1006 | 0.003869 | 0.98823 | 0.024166 | 24.4537 | 0.14366 | 0.14849 | 0.009412 | 0.68314 | 0.68314 | 0.008977
2019-03-26 16:16 | 1553631384.6686 | 205 | 0.20761 | 0.002055 | 5.1e-05 | 90.6163 | 89.2625 | 122.3637 | 21.2612 | 0.002859 | 0.002535 | 255.1457 | 0.25099 | 8.4047 | 0.003525 | 0.005184 | 0.00505 | 6.1e-05 | 0.001125 | 0.33539 | 1.3327 | 2.5477 | 100.79 | 21.15 | 15.515 | 90.6113 | 106.6914 | 82.7962 | 0.003868 | 0.99185 | 0.024215 | 24.414 | 0.13464 | 0.14931 | 0.009403 | 0.68381 | 0.68381 | 0.009018
2019-03-26 16:16 | 1553631385.6696 | 206 | 0.20762 | 0.002046 | 5.2e-05 | 90.5345 | 89.0894 | 122.1234 | 21.2612 | 0.002823 | 0.002576 | 255.0674 | 0.25037 | 8.4013 | 0.003486 | 0.004997 | 0.004826 | 5.8e-05 | 0.001115 | 0.33212 | 1.3322 | 2.5474 | 100.79 | 21.15 | 15.515 | 89.7362 | 105.7463 | 83.7864 | 0.003868 | 0.9891 | 0.024158 | 24.4239 | 0.14392 | 0.14894 | 0.009354 | 0.67863 | 0.67863 | 0.009064
2019-03-26 16:16 | 1553631386.6686 | 207 | 0.20763 | 0.002043 | 5.1e-05 | 90.3462 | 88.732 | 122.1722 | 21.2612 | 0.002833 | 0.002628 | 255.0054 | 0.25101 | 8.4054 | 0.003507 | 0.004671 | 0.004752 | 6e-05 | 0.001114 | 0.33732 | 1.3238 | 2.5474 | 100.79 | 21.15 | 15.514 | 89.9705 | 106.2522 | 85.0379 | 0.003868 | 0.98564 | 0.02412 | 24.4715 | 0.16125 | 0.14933 | 0.009277 | 0.672 | 0.672 | 0.009177
2019-03-26 16:16 | 1553631387.6702 | 208 | 0.20762 | 0.002034 | 5e-05 | 90.2299 | 88.8107 | 122.1703 | 21.2612 | 0.0028 | 0.002573 | 254.9725 | 0.25052 | 8.4008 | 0.003459 | 0.004839 | 0.004653 | 5.3e-05 | 0.00111 | 0.31614 | 1.3765 | 2.5476 | 100.79 | 21.15 | 15.515 | 89.1712 | 105.1073 | 83.713 | 0.003868 | 0.9893 | 0.024165 | 24.426 | 0.1418 | 0.14904 | 0.009381 | 0.68084 | 0.68084 | 0.009041
2019-03-26 16:16 | 1553631388.6688 | 209 | 0.20763 | 0.002024 | 5e-05 | 89.9412 | 88.6789 | 122.0356 | 21.2612 | 0.002837 | 0.002572 | 254.9255 | 0.25092 | 8.4019 | 0.0034 | 0.004911 | 0.00486 | 5.3e-05 | 0.001105 | 0.32101 | 1.332 | 2.5475 | 100.79 | 21.15 | 15.516 | 90.0746 | 103.67 | 83.6957 | 0.003869 | 0.98935 | 0.024152 | 24.412 | 0.12642 | 0.14927 | 0.009295 | 0.67428 | 0.67428 | 0.009143
2019-03-26 16:16 | 1553631389.6702 | 210 | 0.20764 | 0.002019 | 4.8e-05 | 89.8086 | 88.3125 | 121.6718 | 21.2612 | 0.002813 | 0.002625 | 254.8814 | 0.25045 | 8.4044 | 0.003392 | 0.005277 | 0.005149 | 5.5e-05 | 0.0011 | 0.29662 | 1.3499 | 2.5475 | 100.79 | 21.15 | 15.516 | 89.4877 | 103.4766 | 84.9715 | 0.003869 | 0.98583 | 0.024073 | 24.4191 | 0.15026 | 0.14899 | 0.009267 | 0.67015 | 0.67015 | 0.009182
2019-03-26 16:16 | 1553631390.6683 | 211 | 0.20769 | 0.002016 | 4.6e-05 | 89.6129 | 88.1142 | 121.4553 | 21.2612 | 0.002815 | 0.002576 | 254.8378 | 0.2507 | 8.4045 | 0.003399 | 0.004606 | 0.004557 | 5.4e-05 | 0.001104 | 0.32925 | 1.3605 | 2.5476 | 100.79 | 21.15 | 15.518 | 89.544 | 103.6457 | 83.7832 | 0.003869 | 0.98911 | 0.024092 | 24.3569 | 0.15086 | 0.14914 | 0.008979 | 0.64987 | 0.64987 | 0.009478
2019-03-26 16:16 | 1553631391.6689 | 212 | 0.20772 | 0.002006 | 4.5e-05 | 89.4039 | 87.8278 | 121.7272 | 21.2612 | 0.002872 | 0.002553 | 254.8049 | 0.24974 | 8.4025 | 0.003454 | 0.004647 | 0.004649 | 5.5e-05 | 0.001103 | 0.29839 | 1.3584 | 2.5476 | 100.79 | 21.16 | 15.519 | 90.9044 | 104.9808 | 83.2268 | 0.003872 | 0.99065 | 0.024137 | 24.3651 | 0.1591 | 0.14857 | 0.008798 | 0.63797 | 0.63797 | 0.009618
2019-03-26 16:17 | 1553631392.668 | 213 | 0.20772 | 0.001995 | 4.4e-05 | 89.1719 | 87.538 | 122.0452 | 21.2612 | 0.002823 | 0.002585 | 254.743 | 0.25014 | 8.4024 | 0.003364 | 0.00465 | 0.004482 | 4.9e-05 | 0.001098 | 0.3533 | 1.3699 | 2.5475 | 100.79 | 21.16 | 15.519 | 89.7439 | 102.7852 | 83.9946 | 0.003872 | 0.98852 | 0.024143 | 24.4232 | 0.16541 | 0.14881 | 0.008818 | 0.63968 | 0.63968 | 0.009608
2019-03-26 16:17 | 1553631393.668 | 214 | 0.20772 | 0.001982 | 4.3e-05 | 89.1247 | 87.6116 | 122.0506 | 21.2612 | 0.002801 | 0.002554 | 254.6911 | 0.25005 | 8.4025 | 0.003408 | 0.004618 | 0.004538 | 5.5e-05 | 0.001101 | 0.32836 | 1.3877 | 2.5476 | 100.79 | 21.16 | 15.518 | 89.2047 | 103.8633 | 83.2566 | 0.003871 | 0.99057 | 0.024168 | 24.3985 | 0.15316 | 0.14875 | 0.008853 | 0.64297 | 0.64297 | 0.009555
2019-03-26 16:17 | 1553631394.668 | 215 | 0.2077 | 0.001973 | 4.2e-05 | 89.047 | 87.4869 | 122.402 | 21.2612 | 0.002809 | 0.002576 | 254.6513 | 0.2497 | 8.4009 | 0.003408 | 0.004777 | 0.004686 | 5.3e-05 | 0.001103 | 0.3298 | 1.3535 | 2.5475 | 100.79 | 21.16 | 15.516 | 89.3892 | 103.8674 | 83.7965 | 0.003871 | 0.98907 | 0.024185 | 24.452 | 0.1581 | 0.14855 | 0.008963 | 0.65148 | 0.65148 | 0.009417
2019-03-26 16:17 | 1553631395.6692 | 216 | 0.20771 | 0.001967 | 4.1e-05 | 89.0777 | 87.3699 | 122.6541 | 21.2612 | 0.002739 | 0.002552 | 254.5906 | 0.2501 | 8.4019 | 0.003481 | 0.005093 | 0.005024 | 5.5e-05 | 0.001103 | 0.32211 | 1.3527 | 2.5475 | 100.79 | 21.16 | 15.514 | 87.7146 | 105.6257 | 83.2153 | 0.00387 | 0.99068 | 0.024229 | 24.4573 | 0.17316 | 0.14879 | 0.008909 | 0.64877 | 0.64877 | 0.009471
2019-03-26 16:17 | 1553631396.6695 | 217 | 0.20773 | 0.001955 | 4e-05 | 89.0066 | 87.4499 | 122.5369 | 21.2612 | 0.002828 | 0.00252 | 254.5369 | 0.2494 | 8.4025 | 0.0035 | 0.00488 | 0.00477 | 5.3e-05 | 0.00111 | 0.31404 | 1.3156 | 2.5474 | 100.79 | 21.16 | 15.511 | 89.8588 | 106.0861 | 82.4313 | 0.00387 | 0.99287 | 0.024245 | 24.4187 | 0.15782 | 0.14837 | 0.008809 | 0.64193 | 0.64193 | 0.009546
2019-03-26 16:17 | 1553631397.6685 | 218 | 0.20773 | 0.001942 | 3.9e-05 | 88.9924 | 87.6052 | 122.4398 | 21.2612 | 0.00278 | 0.002548 | 254.4957 | 0.25102 | 8.402 | 0.003505 | 0.005539 | 0.005444 | 5.3e-05 | 0.00111 | 0.29214 | 1.2937 | 2.5475 | 100.79 | 21.16 | 15.502 | 88.6916 | 106.2214 | 83.1114 | 0.003867 | 0.99097 | 0.024212 | 24.4323 | 0.14053 | 0.14933 | 0.008814 | 0.64153 | 0.64153 | 0.009614
2019-03-26 16:17 | 1553631398.6697 | 219 | 0.20774 | 0.001942 | 3.8e-05 | 89.0417 | 87.5871 | 122.2182 | 21.2612 | 0.002717 | 0.002512 | 254.4648 | 0.25043 | 8.4018 | 0.003532 | 0.004775 | 0.004735 | 5.3e-05 | 0.001119 | 0.2959 | 1.3014 | 2.5476 | 100.79 | 21.17 | 15.496 | 87.1827 | 106.8685 | 82.2565 | 0.003868 | 0.99336 | 0.024219 | 24.3809 | 0.14733 | 0.14898 | 0.008753 | 0.6373 | 0.6373 | 0.009655
2019-03-26 16:17 | 1553631399.6683 | 220 | 0.20779 | 0.001931 | 3.7e-05 | 89.0822 | 87.7535 | 122.4438 | 21.2612 | 0.002729 | 0.00256 | 254.4502 | 0.25037 | 8.4033 | 0.003427 | 0.005736 | 0.005558 | 5.8e-05 | 0.001126 | 0.31957 | 1.3153 | 2.5474 | 100.79 | 21.17 | 15.483 | 87.4763 | 104.3327 | 83.3954 | 0.003865 | 0.99018 | 0.024203 | 24.4425 | 0.13442 | 0.14895 | 0.008502 | 0.61861 | 0.61861 | 0.009944
2019-03-26 16:17 | 1553631400.6687 | 221 | 0.20783 | 0.001918 | 3.6e-05 | 89.1973 | 87.8936 | 122.5024 | 21.2612 | 0.002764 | 0.002507 | 254.392 | 0.24972 | 8.4036 | 0.003459 | 0.005718 | 0.005635 | 5.6e-05 | 0.001129 | 0.33572 | 1.3262 | 2.5475 | 100.79 | 21.17 | 15.475 | 88.312 | 105.1073 | 82.138 | 0.003863 | 0.99369 | 0.024251 | 24.4052 | 0.13169 | 0.14856 | 0.008262 | 0.60236 | 0.60236 | 0.010186
2019-03-26 16:17 | 1553631401.6684 | 222 | 0.2078 | 0.00191 | 3.5e-05 | 89.3195 | 87.892 | 122.4398 | 21.2612 | 0.002801 | 0.002515 | 254.323 | 0.24956 | 8.4002 | 0.00338 | 0.005059 | 0.005004 | 5.8e-05 | 0.001125 | 0.32676 | 1.3268 | 2.5559 | 100.79 | 21.17 | 15.467 | 89.2103 | 103.185 | 82.3222 | 0.003861 | 0.99317 | 0.024239 | 24.4053 | 0.14411 | 0.14846 | 0.00847 | 0.61735 | 0.61735 | 0.009932
2019-03-26 16:17 | 1553631402.6685 | 223 | 0.2078 | 0.001901 | 3.5e-05 | 89.3877 | 87.9582 | 122.0903 | 21.2612 | 0.00277 | 0.002499 | 254.2873 | 0.2505 | 8.404 | 0.003452 | 0.004714 | 0.004691 | 6.1e-05 | 0.001127 | 0.26842 | 1.3253 | 2.858 | 100.79 | 21.17 | 15.449 | 88.4564 | 104.9335 | 81.9368 | 0.003857 | 0.99425 | 0.024217 | 24.3572 | 0.1442 | 0.14902 | 0.008435 | 0.61423 | 0.61423 | 0.01002
2019-03-26 16:17 | 1553631403.6685 | 224 | 0.20778 | 0.001895 | 3.6e-05 | 89.447 | 87.97 | 122.2503 | 21.2612 | 0.002806 | 0.002577 | 254.263 | 0.25034 | 8.4048 | 0.003434 | 0.004797 | 0.004728 | 5.5e-05 | 0.001131 | 0.32477 | 1.3181 | 2.5475 | 100.79 | 21.17 | 15.441 | 89.3119 | 104.4979 | 83.8066 | 0.003855 | 0.98904 | 0.024169 | 24.4372 | 0.14894 | 0.14893 | 0.008574 | 0.6231 | 0.6231 | 0.009871
2019-03-26 16:17 | 1553631404.6685 | 225 | 0.20776 | 0.001885 | 3.6e-05 | 89.5046 | 88.1303 | 121.7884 | 21.2612 | 0.002772 | 0.002571 | 254.2079 | 0.24999 | 8.4006 | 0.00341 | 0.005247 | 0.00523 | 5.6e-05 | 0.001125 | 0.30735 | 1.3126 | 2.5475 | 100.79 | 21.17 | 15.423 | 88.5089 | 103.9081 | 83.6748 | 0.00385 | 0.98941 | 0.024128 | 24.3865 | 0.1384 | 0.14872 | 0.008739 | 0.63399 | 0.63399 | 0.009688
2019-03-26 16:17 | 1553631405.6686 | 226 | 0.20775 | 0.001886 | 3.7e-05 | 89.5477 | 88.2685 | 121.4741 | 21.2612 | 0.002786 | 0.002571 | 254.1947 | 0.24976 | 8.4016 | 0.003459 | 0.004911 | 0.004796 | 5.8e-05 | 0.001127 | 0.31116 | 1.3179 | 2.5475 | 100.79 | 21.17 | 15.414 | 88.8448 | 105.1018 | 83.663 | 0.003848 | 0.98944 | 0.024097 | 24.3546 | 0.1287 | 0.14858 | 0.008751 | 0.63407 | 0.63407 | 0.009678
2019-03-26 16:17 | 1553631406.6676 | 227 | 0.20773 | 0.001883 | 3.7e-05 | 89.6141 | 88.2012 | 121.4291 | 21.2612 | 0.002741 | 0.002589 | 254.1658 | 0.25086 | 8.4015 | 0.003462 | 0.005412 | 0.00535 | 6e-05 | 0.001131 | 0.30807 | 1.3737 | 2.5475 | 100.79 | 21.17 | 15.407 | 87.7543 | 105.1726 | 84.1014 | 0.003846 | 0.98823 | 0.024078 | 24.3651 | 0.14214 | 0.14923 | 0.008912 | 0.64524 | 0.64524 | 0.009552
2019-03-26 16:17 | 1553631407.6686 | 228 | 0.20771 | 0.001879 | 3.7e-05 | 89.749 | 88.3525 | 121.8009 | 21.2612 | 0.002795 | 0.002587 | 254.0979 | 0.24995 | 8.4037 | 0.003418 | 0.004828 | 0.004773 | 5.7e-05 | 0.00114 | 0.34307 | 1.3664 | 2.5475 | 100.79 | 21.17 | 15.393 | 89.0627 | 104.1102 | 84.0593 | 0.003843 | 0.98834 | 0.024116 | 24.4009 | 0.14027 | 0.1487 | 0.009006 | 0.65311 | 0.65311 | 0.009403
2019-03-26 16:17 | 1553631408.6677 | 229 | 0.20772 | 0.001889 | 3.6e-05 | 89.8464 | 88.5337 | 122.1858 | 21.2612 | 0.002766 | 0.00248 | 254.0563 | 0.25026 | 8.4019 | 0.003528 | 0.005199 | 0.005091 | 6.1e-05 | 0.001142 | 0.29894 | 1.3724 | 2.5476 | 100.79 | 21.17 | 15.388 | 88.3495 | 106.7671 | 81.4837 | 0.003841 | 0.99552 | 0.024242 | 24.3512 | 0.13165 | 0.14888 | 0.008946 | 0.65216 | 0.65216 | 0.009428
2019-03-26 16:17 | 1553631409.6688 | 230 | 0.20773 | 0.001886 | 3.6e-05 | 89.9736 | 88.5532 | 122.2323 | 21.2612 | 0.00283 | 0.002524 | 253.9984 | 0.2499 | 8.4046 | 0.003488 | 0.005154 | 0.005123 | 5.8e-05 | 0.001139 | 0.33179 | 1.3664 | 2.5474 | 100.79 | 21.17 | 15.383 | 89.9072 | 105.7894 | 82.5366 | 0.00384 | 0.99257 | 0.024211 | 24.3919 | 0.14234 | 0.14866 | 0.00891 | 0.64868 | 0.64868 | 0.009465
2019-03-26 16:17 | 1553631410.6678 | 231 | 0.20772 | 0.001892 | 3.5e-05 | 90.0976 | 88.49 | 122.2201 | 21.2612 | 0.002748 | 0.002575 | 253.978 | 0.25017 | 8.4008 | 0.003465 | 0.004825 | 0.004742 | 6e-05 | 0.001134 | 0.32565 | 1.3544 | 2.5474 | 100.79 | 21.17 | 15.381 | 87.9292 | 105.2538 | 83.7575 | 0.00384 | 0.98918 | 0.024168 | 24.4325 | 0.16104 | 0.14883 | 0.008982 | 0.65289 | 0.65289 | 0.009414
2019-03-26 16:17 | 1553631411.6689 | 232 | 0.20771 | 0.001899 | 3.4e-05 | 89.9831 | 88.5394 | 121.8898 | 21.2612 | 0.002748 | 0.002548 | 253.9583 | 0.25021 | 8.4041 | 0.003592 | 0.004824 | 0.004682 | 5.7e-05 | 0.001132 | 0.29535 | 1.3405 | 2.5475 | 100.79 | 21.17 | 15.381 | 87.916 | 108.3146 | 83.1247 | 0.00384 | 0.99094 | 0.024157 | 24.3779 | 0.14467 | 0.14885 | 0.008984 | 0.6527 | 0.6527 | 0.009418
2019-03-26 16:17 | 1553631412.6694 | 233 | 0.20771 | 0.001895 | 3.3e-05 | 90.0418 | 88.6117 | 121.5936 | 21.2612 | 0.002767 | 0.002582 | 253.9029 | 0.25022 | 8.4019 | 0.003524 | 0.005209 | 0.005194 | 5.6e-05 | 0.001129 | 0.32803 | 1.3297 | 2.5475 | 100.79 | 21.17 | 15.384 | 88.3962 | 106.6788 | 83.94 | 0.00384 | 0.98867 | 0.0241 | 24.3761 | 0.14321 | 0.14886 | 0.008999 | 0.65234 | 0.65234 | 0.009424
2019-03-26 16:17 | 1553631413.668 | 234 | 0.20774 | 0.001896 | 3.4e-05 | 90.1254 | 88.6534 | 121.3899 | 21.2612 | 0.002753 | 0.002585 | 253.8552 | 0.24974 | 8.4002 | 0.003496 | 0.004903 | 0.004852 | 5.6e-05 | 0.001132 | 0.33157 | 1.3144 | 2.5475 | 100.79 | 21.17 | 15.387 | 88.0441 | 106.0043 | 84.0001 | 0.003841 | 0.98851 | 0.024078 | 24.3577 | 0.1473 | 0.14857 | 0.008825 | 0.63904 | 0.63904 | 0.009602
2019-03-26 16:17 | 1553631414.668 | 235 | 0.20776 | 0.001897 | 3.4e-05 | 90.2564 | 89.0168 | 121.4927 | 21.2612 | 0.002735 | 0.002511 | 253.8289 | 0.25037 | 8.402 | 0.003493 | 0.00471 | 0.004648 | 5.7e-05 | 0.00113 | 0.30359 | 1.3175 | 2.5474 | 100.79 | 21.17 | 15.389 | 87.6147 | 105.9122 | 82.2371 | 0.003842 | 0.99341 | 0.024148 | 24.3078 | 0.1237 | 0.14895 | 0.008727 | 0.63377 | 0.63377 | 0.009706
2019-03-26 16:17 | 1553631415.669 | 236 | 0.20777 | 0.001894 | 3.4e-05 | 90.3013 | 88.9939 | 121.5591 | 21.2612 | 0.002815 | 0.002498 | 253.7828 | 0.25068 | 8.4009 | 0.003512 | 0.004953 | 0.004927 | 5.2e-05 | 0.001132 | 0.30541 | 1.3433 | 2.5474 | 100.79 | 21.17 | 15.392 | 89.5469 | 106.3743 | 81.9104 | 0.003842 | 0.99433 | 0.024165 | 24.3032 | 0.13045 | 0.14913 | 0.008678 | 0.63072 | 0.63072 | 0.009765
2019-03-26 16:17 | 1553631416.6691 | 237 | 0.20778 | 0.001891 | 3.4e-05 | 90.4105 | 88.9452 | 121.7178 | 21.2612 | 0.00272 | 0.002536 | 253.7436 | 0.25033 | 8.4042 | 0.003547 | 0.005173 | 0.005101 | 5.6e-05 | 0.001129 | 0.32731 | 1.3764 | 2.5475 | 100.79 | 21.17 | 15.399 | 87.2522 | 107.2228 | 82.8345 | 0.003844 | 0.99174 | 0.02415 | 24.3507 | 0.14616 | 0.14893 | 0.008622 | 0.62619 | 0.62619 | 0.009822
2019-03-26 16:17 | 1553631417.6681 | 238 | 0.20778 | 0.001889 | 3.4e-05 | 90.4938 | 89.127 | 121.3964 | 21.2612 | 0.002708 | 0.002545 | 253.7005 | 0.24974 | 8.4028 | 0.00351 | 0.004343 | 0.004365 | 5.8e-05 | 0.001131 | 0.3115 | 1.3528 | 2.5477 | 100.79 | 21.17 | 15.401 | 86.9743 | 106.3231 | 83.0523 | 0.003845 | 0.99114 | 0.02411 | 24.326 | 0.13612 | 0.14857 | 0.00862 | 0.62503 | 0.62503 | 0.009817
2019-03-26 16:17 | 1553631418.6692 | 239 | 0.20774 | 0.001881 | 3.4e-05 | 90.5537 | 89.3173 | 121.3103 | 21.2612 | 0.002769 | 0.002531 | 253.6611 | 0.24998 | 8.4027 | 0.003499 | 0.004472 | 0.004352 | 5.5e-05 | 0.001134 | 0.31669 | 1.3989 | 2.7858 | 100.79 | 21.17 | 15.402 | 88.4362 | 106.0727 | 82.7145 | 0.003845 | 0.99208 | 0.024113 | 24.3058 | 0.12297 | 0.14871 | 0.008836 | 0.64081 | 0.64081 | 0.009584
2019-03-26 16:17 | 1553631419.6692 | 240 | 0.20774 | 0.001878 | 3.6e-05 | 90.7083 | 89.3055 | 121.191 | 21.2612 | 0.002728 | 0.002508 | 253.6071 | 0.24987 | 8.4009 | 0.003429 | 0.005066 | 0.004984 | 5.5e-05 | 0.001132 | 0.32139 | 1.3487 | 2.5475 | 100.79 | 21.17 | 15.405 | 87.4366 | 104.3642 | 82.1551 | 0.003846 | 0.99364 | 0.02412 | 24.2748 | 0.13941 | 0.14865 | 0.008836 | 0.64092 | 0.64092 | 0.009579
2019-03-26 16:17 | 1553631420.6683 | 241 | 0.20774 | 0.001877 | 3.6e-05 | 90.8524 | 89.5987 | 121.076 | 21.2612 | 0.002729 | 0.002544 | 253.5715 | 0.25052 | 8.4021 | 0.00351 | 0.004983 | 0.004937 | 5.9e-05 | 0.00113 | 0.34235 | 1.3631 | 2.5475 | 100.79 | 21.17 | 15.406 | 87.4669 | 106.3333 | 83.028 | 0.003846 | 0.99121 | 0.024079 | 24.2931 | 0.12429 | 0.14904 | 0.008883 | 0.64323 | 0.64323 | 0.009569
2019-03-26 16:17 | 1553631421.6674 | 242 | 0.20775 | 0.00188 | 3.4e-05 | 90.892 | 89.6464 | 121.0338 | 21.2612 | 0.002668 | 0.002503 | 253.5432 | 0.25013 | 8.4023 | 0.003597 | 0.004939 | 0.004878 | 5.3e-05 | 0.001129 | 0.32388 | 1.3308 | 2.5475 | 100.79 | 21.17 | 15.408 | 86.0025 | 108.4353 | 82.0359 | 0.003846 | 0.99397 | 0.024109 | 24.255 | 0.12342 | 0.1488 | 0.008819 | 0.63946 | 0.63946 | 0.00961
2019-03-26 16:17 | 1553631422.6674 | 243 | 0.20777 | 0.001881 | 3.5e-05 | 90.9197 | 89.6383 | 121.3338 | 21.2612 | 0.002762 | 0.002526 | 253.5019 | 0.25013 | 8.4023 | 0.003616 | 0.005096 | 0.004989 | 5.3e-05 | 0.001129 | 0.31022 | 1.3596 | 2.5475 | 100.79 | 21.17 | 15.409 | 88.2549 | 108.9073 | 82.5936 | 0.003847 | 0.99242 | 0.02412 | 24.3041 | 0.12696 | 0.1488 | 0.008691 | 0.63037 | 0.63037 | 0.009749
2019-03-26 16:17 | 1553631423.6685 | 244 | 0.20776 | 0.001881 | 3.5e-05 | 90.9841 | 89.7331 | 121.3075 | 21.2612 | 0.002714 | 0.002546 | 253.4554 | 0.24868 | 8.4016 | 0.003564 | 0.004981 | 0.004953 | 5.4e-05 | 0.001126 | 0.32947 | 1.3418 | 2.5475 | 100.79 | 21.17 | 15.411 | 87.111 | 107.6348 | 83.0624 | 0.003847 | 0.99111 | 0.024101 | 24.3174 | 0.12383 | 0.14794 | 0.008751 | 0.63423 | 0.63423 | 0.009634
2019-03-26 16:17 | 1553631424.6685 | 245 | 0.20777 | 0.001886 | 3.5e-05 | 91.0077 | 89.8024 | 121.61 | 21.2612 | 0.002787 | 0.002532 | 253.4167 | 0.25002 | 8.4045 | 0.003564 | 0.004987 | 0.004958 | 5.4e-05 | 0.001131 | 0.30702 | 1.3365 | 2.5475 | 100.79 | 21.17 | 15.413 | 88.8604 | 107.6365 | 82.7292 | 0.003848 | 0.99204 | 0.024143 | 24.3364 | 0.11925 | 0.14874 | 0.008669 | 0.62941 | 0.62941 | 0.00976
2019-03-26 16:17 | 1553631425.6686 | 246 | 0.20777 | 0.001884 | 3.4e-05 | 91.0298 | 89.8104 | 121.6868 | 21.2612 | 0.002746 | 0.002522 | 253.4028 | 0.25061 | 8.4023 | 0.003502 | 0.005172 | 0.005062 | 5e-05 | 0.001132 | 0.3465 | 1.3495 | 2.5475 | 100.79 | 21.17 | 15.416 | 87.8715 | 106.1404 | 82.4827 | 0.003848 | 0.99272 | 0.024159 | 24.3356 | 0.12062 | 0.14909 | 0.008674 | 0.63018 | 0.63018 | 0.009771
2019-03-26 16:17 | 1553631426.6676 | 247 | 0.20776 | 0.001886 | 3.4e-05 | 91.1269 | 89.9402 | 121.5915 | 21.2612 | 0.002742 | 0.00251 | 253.3396 | 0.24951 | 8.4025 | 0.003481 | 0.004406 | 0.00437 | 5.3e-05 | 0.00114 | 0.33467 | 1.345 | 2.5476 | 100.79 | 21.17 | 15.423 | 87.7925 | 105.6274 | 82.1957 | 0.00385 | 0.99353 | 0.024159 | 24.3163 | 0.11725 | 0.14843 | 0.008731 | 0.63434 | 0.63434 | 0.009664
2019-03-26 16:17 | 1553631427.6676 | 248 | 0.20778 | 0.001884 | 3.3e-05 | 91.2447 | 89.8521 | 121.503 | 21.2612 | 0.002696 | 0.002529 | 253.3001 | 0.25031 | 8.4041 | 0.003506 | 0.004917 | 0.004838 | 5.2e-05 | 0.00114 | 0.33229 | 1.345 | 2.5475 | 100.79 | 21.17 | 15.43 | 86.687 | 106.2372 | 82.659 | 0.003852 | 0.99223 | 0.024134 | 24.3232 | 0.13756 | 0.14891 | 0.008628 | 0.62624 | 0.62624 | 0.00982
2019-03-26 16:17 | 1553631428.6677 | 249 | 0.20778 | 0.001883 | 3.3e-05 | 91.2972 | 89.8656 | 121.2497 | 21.2612 | 0.002748 | 0.002518 | 253.2565 | 0.25037 | 8.4006 | 0.003477 | 0.00481 | 0.00474 | 5.5e-05 | 0.001148 | 0.33339 | 1.3715 | 2.5474 | 100.79 | 21.17 | 15.438 | 87.9292 | 105.5346 | 82.4017 | 0.003854 | 0.99295 | 0.024118 | 24.2891 | 0.14138 | 0.14895 | 0.008632 | 0.62615 | 0.62615 | 0.009824
2019-03-26 16:17 | 1553631429.6687 | 250 | 0.20777 | 0.001886 | 3.3e-05 | 91.3011 | 90.2508 | 120.8723 | 21.2612 | 0.002731 | 0.002546 | 253.238 | 0.24996 | 8.4011 | 0.003414 | 0.004606 | 0.004593 | 5.2e-05 | 0.001148 | 0.31116 | 1.3688 | 2.5475 | 100.8 | 21.17 | 15.458 | 87.517 | 104.0142 | 83.0763 | 0.003858 | 0.99107 | 0.024057 | 24.2743 | 0.10349 | 0.1487 | 0.008636 | 0.62489 | 0.62489 | 0.009828
2019-03-26 16:17 | 1553631430.6678 | 251 | 0.20779 | 0.001886 | 3.3e-05 | 91.3265 | 90.1345 | 120.9709 | 21.2612 | 0.002715 | 0.002526 | 253.2124 | 0.25038 | 8.4015 | 0.003462 | 0.00515 | 0.005007 | 5.6e-05 | 0.001149 | 0.34772 | 1.365 | 2.5476 | 100.8 | 21.17 | 15.469 | 87.1406 | 105.1757 | 82.5936 | 0.003861 | 0.99242 | 0.024084 | 24.2677 | 0.11751 | 0.14895 | 0.008522 | 0.61728 | 0.61728 | 0.009966
2019-03-26 16:17 | 1553631431.6689 | 252 | 0.20779 | 0.001885 | 3.4e-05 | 91.4378 | 90.3076 | 121.3277 | 21.2612 | 0.002773 | 0.002531 | 253.1842 | 0.2505 | 8.4036 | 0.003443 | 0.005474 | 0.005403 | 5.2e-05 | 0.001143 | 0.31868 | 1.3489 | 2.5475 | 100.8 | 21.17 | 15.48 | 88.526 | 104.7103 | 82.7027 | 0.003864 | 0.99211 | 0.024115 | 24.3072 | 0.11124 | 0.14902 | 0.008558 | 0.62068 | 0.62068 | 0.009916
2019-03-26 16:17 | 1553631432.6689 | 253 | 0.20777 | 0.001881 | 3.5e-05 | 91.5361 | 90.2696 | 121.4401 | 21.2612 | 0.002712 | 0.002564 | 253.168 | 0.24973 | 8.4017 | 0.003423 | 0.004399 | 0.004357 | 5.5e-05 | 0.00114 | 0.30093 | 1.3718 | 2.5475 | 100.8 | 21.17 | 15.5 | 87.0494 | 104.2156 | 83.4952 | 0.003869 | 0.98991 | 0.0241 | 24.3455 | 0.12461 | 0.14857 | 0.008675 | 0.6287 | 0.6287 | 0.009759
2019-03-26 16:17 | 1553631433.668 | 254 | 0.2078 | 0.001885 | 3.5e-05 | 91.6211 | 90.388 | 121.5096 | 21.2612 | 0.002793 | 0.002522 | 253.1468 | 0.24971 | 8.4025 | 0.003479 | 0.004788 | 0.004691 | 5.4e-05 | 0.001133 | 0.29159 | 1.3731 | 2.5476 | 100.8 | 21.17 | 15.509 | 89.015 | 105.584 | 82.4936 | 0.003871 | 0.99269 | 0.024141 | 24.3182 | 0.1212 | 0.14855 | 0.008464 | 0.61435 | 0.61435 | 0.009987
2019-03-26 16:17 | 1553631434.668 | 255 | 0.20779 | 0.001893 | 3.5e-05 | 91.7318 | 90.4414 | 121.2356 | 21.2612 | 0.002777 | 0.002527 | 253.0837 | 0.24985 | 8.4011 | 0.003423 | 0.005426 | 0.00532 | 5.6e-05 | 0.001121 | 0.32748 | 1.3942 | 2.5475 | 100.79 | 21.17 | 15.527 | 88.6221 | 104.2139 | 82.6145 | 0.003876 | 0.99236 | 0.024109 | 24.2949 | 0.12671 | 0.14864 | 0.008548 | 0.61968 | 0.61968 | 0.009906
2019-03-26 16:17 | 1553631435.668 | 256 | 0.20779 | 0.001895 | 3.5e-05 | 91.8056 | 90.4713 | 121.2798 | 21.2612 | 0.002776 | 0.002533 | 253.0302 | 0.24982 | 8.4047 | 0.003559 | 0.005138 | 0.005122 | 5.5e-05 | 0.001112 | 0.31614 | 1.3318 | 2.5475 | 100.79 | 21.17 | 15.534 | 88.6026 | 107.5105 | 82.7496 | 0.003878 | 0.99198 | 0.024109 | 24.304 | 0.13095 | 0.14862 | 0.008535 | 0.61874 | 0.61874 | 0.00992
2019-03-26 16:17 | 1553631436.6681 | 257 | 0.20778 | 0.001895 | 3.5e-05 | 91.9558 | 90.8209 | 121.0666 | 21.2612 | 0.002736 | 0.002571 | 253.0188 | 0.25043 | 8.4028 | 0.003529 | 0.00444 | 0.004436 | 5.8e-05 | 0.001109 | 0.30453 | 1.3764 | 2.5475 | 100.79 | 21.17 | 15.541 | 87.6435 | 106.8049 | 83.6717 | 0.00388 | 0.98942 | 0.024057 | 24.3141 | 0.11107 | 0.14898 | 0.008591 | 0.62146 | 0.62146 | 0.009901
2019-03-26 16:17 | 1553631437.6682 | 258 | 0.2078 | 0.001895 | 3.4e-05 | 92.0331 | 90.754 | 120.9556 | 21.2612 | 0.002694 | 0.002553 | 252.9969 | 0.24961 | 8.4019 | 0.003437 | 0.004996 | 0.004811 | 5.4e-05 | 0.001097 | 0.31437 | 1.3567 | 2.5474 | 100.79 | 21.17 | 15.551 | 86.616 | 104.5702 | 83.2355 | 0.003882 | 0.99063 | 0.02406 | 24.288 | 0.12518 | 0.14849 | 0.008503 | 0.61524 | 0.61524 | 0.009968
2019-03-26 16:17 | 1553631438.6682 | 259 | 0.20777 | 0.001893 | 3.4e-05 | 92.119 | 90.9154 | 120.8704 | 21.2612 | 0.002757 | 0.002508 | 252.9471 | 0.24998 | 8.4045 | 0.003461 | 0.005315 | 0.00521 | 5.9e-05 | 0.001096 | 0.33124 | 1.3429 | 2.5474 | 100.79 | 21.17 | 15.556 | 88.1488 | 105.1396 | 82.1537 | 0.003883 | 0.99364 | 0.024089 | 24.2426 | 0.11764 | 0.14871 | 0.008656 | 0.62708 | 0.62708 | 0.009794
2019-03-26 16:17 | 1553631439.6683 | 260 | 0.20778 | 0.001887 | 3.4e-05 | 92.2229 | 91.0479 | 121.1533 | 21.2612 | 0.002687 | 0.002536 | 252.8942 | 0.24898 | 8.4041 | 0.003465 | 0.00503 | 0.005005 | 5.2e-05 | 0.001095 | 0.32117 | 1.3516 | 2.5475 | 100.79 | 21.17 | 15.557 | 86.4497 | 105.2521 | 82.8251 | 0.003884 | 0.99177 | 0.024094 | 24.2939 | 0.1147 | 0.14812 | 0.008582 | 0.62182 | 0.62182 | 0.009837
2019-03-26 16:17 | 1553631440.6682 | 261 | 0.20777 | 0.001886 | 3.3e-05 | 92.3342 | 91.3396 | 121.2626 | 21.2612 | 0.002783 | 0.00255 | 252.8486 | 0.24975 | 8.3998 | 0.00337 | 0.004362 | 0.004344 | 5.5e-05 | 0.001101 | 0.34899 | 1.3635 | 2.5474 | 100.79 | 21.17 | 15.557 | 88.7735 | 102.9423 | 83.166 | 0.003884 | 0.99082 | 0.024093 | 24.3165 | 0.09687 | 0.14858 | 0.008649 | 0.62677 | 0.62677 | 0.00979
2019-03-26 16:17 | 1553631441.6684 | 262 | 0.20781 | 0.001891 | 3.2e-05 | 92.3933 | 91.3282 | 121.2164 | 21.2612 | 0.002775 | 0.002555 | 252.7954 | 0.24954 | 8.4036 | 0.003332 | 0.004008 | 0.003868 | 5.4e-05 | 0.001104 | 0.29856 | 1.3528 | 2.5475 | 100.79 | 21.17 | 15.555 | 88.5688 | 102.0216 | 83.2737 | 0.003883 | 0.99052 | 0.024085 | 24.3155 | 0.10371 | 0.14845 | 0.00841 | 0.60924 | 0.60924 | 0.010064
2019-03-26 16:17 | 1553631442.6684 | 263 | 0.20779 | 0.001889 | 3.1e-05 | 92.5016 | 91.2835 | 121.0467 | 21.2612 | 0.002737 | 0.002537 | 252.7422 | 0.24905 | 8.4015 | 0.003388 | 0.004713 | 0.004652 | 5.3e-05 | 0.001109 | 0.32261 | 1.3336 | 2.5475 | 100.79 | 21.17 | 15.547 | 87.6614 | 103.3776 | 82.8424 | 0.003881 | 0.99172 | 0.024083 | 24.2838 | 0.11857 | 0.14816 | 0.008527 | 0.61773 | 0.61773 | 0.009906
2019-03-26 16:17 | 1553631443.6674 | 264 | 0.20782 | 0.001888 | 3.1e-05 | 92.4753 | 91.2139 | 120.8814 | 21.2612 | 0.002714 | 0.002524 | 252.6941 | 0.24939 | 8.4027 | 0.003392 | 0.004949 | 0.004901 | 5.2e-05 | 0.00111 | 0.3235 | 1.3397 | 2.5475 | 100.79 | 21.17 | 15.541 | 87.1165 | 103.4766 | 82.5397 | 0.00388 | 0.99257 | 0.024077 | 24.2569 | 0.12285 | 0.14836 | 0.008392 | 0.60776 | 0.60776 | 0.010082
2019-03-26 16:17 | 1553631444.6675 | 265 | 0.20781 | 0.00188 | 3e-05 | 92.5964 | 91.498 | 120.9742 | 21.2612 | 0.002773 | 0.00253 | 252.6557 | 0.25011 | 8.4 | 0.00342 | 0.005161 | 0.005036 | 5.4e-05 | 0.001113 | 0.33395 | 1.3372 | 2.5475 | 100.79 | 21.17 | 15.535 | 88.5214 | 104.1463 | 82.6909 | 0.003878 | 0.99214 | 0.024081 | 24.2713 | 0.10673 | 0.14879 | 0.008454 | 0.61245 | 0.61245 | 0.010034
2019-03-26 16:17 | 1553631445.6675 | 266 | 0.2078 | 0.001878 | 3e-05 | 92.7797 | 91.7269 | 120.9444 | 21.2612 | 0.002736 | 0.002547 | 252.602 | 0.2495 | 8.4025 | 0.003433 | 0.004729 | 0.004676 | 5.5e-05 | 0.001111 | 0.28263 | 1.3976 | 2.5476 | 100.79 | 21.17 | 15.519 | 87.6457 | 104.4773 | 83.0819 | 0.003874 | 0.99106 | 0.024064 | 24.2817 | 0.10207 | 0.14843 | 0.008529 | 0.61748 | 0.61748 | 0.009928
2019-03-26 16:17 | 1553631446.6686 | 267 | 0.20779 | 0.001874 | 3e-05 | 92.7941 | 91.5746 | 120.9324 | 21.2612 | 0.002732 | 0.002514 | 252.5576 | 0.25041 | 8.4049 | 0.003431 | 0.004378 | 0.004249 | 5.7e-05 | 0.001112 | 0.30326 | 1.4073 | 2.5476 | 100.79 | 21.17 | 15.509 | 87.542 | 104.4107 | 82.308 | 0.003872 | 0.99321 | 0.024089 | 24.2541 | 0.11833 | 0.14897 | 0.008583 | 0.622 | 0.622 | 0.009891
2019-03-26 16:17 | 1553631447.6676 | 268 | 0.20783 | 0.001874 | 3.1e-05 | 92.8476 | 91.6872 | 121.1584 | 21.2612 | 0.002779 | 0.00251 | 252.4975 | 0.25084 | 8.4032 | 0.003478 | 0.005666 | 0.005596 | 5.4e-05 | 0.001116 | 0.30182 | 1.3673 | 2.5474 | 100.79 | 21.17 | 15.492 | 88.6666 | 105.5503 | 82.2012 | 0.003867 | 0.99351 | 0.024116 | 24.2731 | 0.11249 | 0.14923 | 0.008326 | 0.60396 | 0.60396 | 0.010204
2019-03-26 16:17 | 1553631448.6677 | 269 | 0.20779 | 0.001866 | 3e-05 | 92.893 | 91.6045 | 121.2955 | 21.2612 | 0.0027 | 0.002517 | 252.4523 | 0.25008 | 8.4027 | 0.003506 | 0.005415 | 0.005368 | 5.5e-05 | 0.001114 | 0.30127 | 1.3422 | 2.5475 | 100.79 | 21.17 | 15.483 | 86.7666 | 106.2452 | 82.3823 | 0.003865 | 0.99301 | 0.024123 | 24.293 | 0.12494 | 0.14877 | 0.008576 | 0.62234 | 0.62234 | 0.009873
2019-03-26 16:17 | 1553631449.6687 | 270 | 0.20779 | 0.001872 | 3.1e-05 | 92.9721 | 91.9168 | 121.1251 | 21.2612 | 0.002739 | 0.00256 | 252.4056 | 0.24982 | 8.4035 | 0.003496 | 0.004698 | 0.004657 | 5.6e-05 | 0.001121 | 0.29358 | 1.3899 | 2.5477 | 100.79 | 21.17 | 15.474 | 87.7107 | 106.0005 | 83.4079 | 0.003863 | 0.99015 | 0.024071 | 24.3109 | 0.10211 | 0.14861 | 0.008573 | 0.62077 | 0.62077 | 0.009887
2019-03-26 16:17 | 1553631450.6688 | 271 | 0.20778 | 0.001866 | 3.1e-05 | 93.0002 | 91.7299 | 120.7763 | 21.2612 | 0.002764 | 0.002521 | 252.3543 | 0.25022 | 8.4023 | 0.003484 | 0.004737 | 0.004601 | 5.5e-05 | 0.001121 | 0.2975 | 1.3886 | 2.5476 | 100.79 | 21.17 | 15.453 | 88.3064 | 105.7038 | 82.4553 | 0.003858 | 0.9928 | 0.024069 | 24.2435 | 0.12301 | 0.14886 | 0.008641 | 0.62569 | 0.62569 | 0.009826
2019-03-26 16:18 | 1553631451.6688 | 272 | 0.20773 | 0.001868 | 3.1e-05 | 93.0595 | 92.1222 | 120.7683 | 21.2612 | 0.002774 | 0.002557 | 252.2949 | 0.25072 | 8.3999 | 0.003474 | 0.005017 | 0.004865 | 5.5e-05 | 0.00112 | 0.33467 | 1.3958 | 2.5476 | 100.79 | 21.17 | 15.443 | 88.5518 | 105.4504 | 83.3269 | 0.003855 | 0.99037 | 0.024039 | 24.2723 | 0.090544 | 0.14915 | 0.008956 | 0.64765 | 0.64765 | 0.009511
2019-03-26 16:18 | 1553631452.6679 | 273 | 0.20774 | 0.001867 | 3.2e-05 | 93.0872 | 91.9686 | 121.0544 | 21.2612 | 0.00275 | 0.002513 | 252.2725 | 0.2505 | 8.4006 | 0.003471 | 0.004963 | 0.00492 | 5.5e-05 | 0.001126 | 0.30773 | 1.3494 | 2.5474 | 100.79 | 21.17 | 15.427 | 87.9832 | 105.3937 | 82.2652 | 0.003851 | 0.99333 | 0.024103 | 24.2649 | 0.10814 | 0.14902 | 0.008867 | 0.64291 | 0.64291 | 0.009573
2019-03-26 16:18 | 1553631453.6679 | 274 | 0.20776 | 0.00187 | 3.3e-05 | 93.0843 | 92.0863 | 121.2056 | 21.2612 | 0.002754 | 0.002561 | 252.2379 | 0.25074 | 8.4019 | 0.003452 | 0.004967 | 0.004897 | 5.7e-05 | 0.001132 | 0.30071 | 1.3418 | 2.5476 | 100.79 | 21.17 | 15.421 | 88.0691 | 104.9359 | 83.4305 | 0.00385 | 0.99009 | 0.024079 | 24.3198 | 0.09642 | 0.14916 | 0.00873 | 0.63223 | 0.63223 | 0.009744
2019-03-26 16:18 | 1553631454.668 | 275 | 0.20771 | 0.00187 | 3.5e-05 | 93.1605 | 92.0916 | 121.1509 | 21.2612 | 0.002745 | 0.002599 | 252.1969 | 0.25028 | 8.401 | 0.003528 | 0.005097 | 0.005143 | 5.9e-05 | 0.001133 | 0.31509 | 1.4074 | 2.5476 | 100.79 | 21.17 | 15.416 | 87.852 | 106.7671 | 84.3325 | 0.003848 | 0.98759 | 0.024043 | 24.345 | 0.10321 | 0.14889 | 0.009058 | 0.65494 | 0.65494 | 0.009389
2019-03-26 16:18 | 1553631455.667 | 276 | 0.20767 | 0.001872 | 3.6e-05 | 93.1404 | 92.0059 | 120.9167 | 21.2612 | 0.002746 | 0.002535 | 252.1819 | 0.24973 | 8.3994 | 0.003479 | 0.004552 | 0.004435 | 5.4e-05 | 0.001133 | 0.29856 | 1.3886 | 2.5475 | 100.79 | 21.17 | 15.413 | 87.8833 | 105.5802 | 82.7917 | 0.003848 | 0.99186 | 0.024072 | 24.269 | 0.10962 | 0.14857 | 0.009309 | 0.67386 | 0.67386 | 0.009105
2019-03-26 16:18 | 1553631456.6671 | 277 | 0.20768 | 0.001874 | 3.8e-05 | 93.2151 | 92.1921 | 121.0312 | 21.2612 | 0.002811 | 0.002517 | 252.1269 | 0.24916 | 8.4005 | 0.003532 | 0.005033 | 0.004952 | 5.5e-05 | 0.001134 | 0.32549 | 1.4233 | 2.5476 | 100.79 | 21.17 | 15.415 | 89.4323 | 106.8615 | 82.3635 | 0.003848 | 0.99306 | 0.024097 | 24.2659 | 0.098698 | 0.14823 | 0.009246 | 0.66993 | 0.66993 | 0.009138
2019-03-26 16:18 | 1553631457.6681 | 278 | 0.20767 | 0.001881 | 3.8e-05 | 93.2827 | 92.179 | 121.0575 | 21.2612 | 0.002767 | 0.002556 | 252.0939 | 0.24959 | 8.4013 | 0.003553 | 0.005445 | 0.005372 | 5.2e-05 | 0.001124 | 0.33069 | 1.4352 | 2.5476 | 100.79 | 21.17 | 15.419 | 88.3729 | 107.3765 | 83.305 | 0.003849 | 0.99043 | 0.024068 | 24.3006 | 0.10645 | 0.14848 | 0.009294 | 0.67261 | 0.67261 | 0.009117
2019-03-26 16:18 | 1553631458.6676 | 279 | 0.2077 | 0.001894 | 3.9e-05 | 93.2148 | 91.8936 | 121.1605 | 21.2612 | 0.002739 | 0.002589 | 252.0628 | 0.24919 | 8.4033 | 0.003484 | 0.005293 | 0.005273 | 5.6e-05 | 0.001124 | 0.28982 | 1.4223 | 2.5476 | 100.79 | 21.17 | 15.422 | 87.7114 | 105.7077 | 84.0951 | 0.00385 | 0.98824 | 0.024052 | 24.3379 | 0.12768 | 0.14824 | 0.009102 | 0.65818 | 0.65818 | 0.009302
2019-03-26 16:18 | 1553631459.6683 | 280 | 0.20769 | 0.001899 | 3.9e-05 | 93.3624 | 92.295 | 121.0655 | 21.2612 | 0.00275 | 0.002551 | 252.0096 | 0.25007 | 8.4014 | 0.003481 | 0.004503 | 0.004414 | 5.6e-05 | 0.001123 | 0.3246 | 1.3816 | 2.5476 | 100.79 | 21.17 | 15.426 | 87.9692 | 105.6384 | 83.1997 | 0.003851 | 0.99073 | 0.024073 | 24.2978 | 0.10285 | 0.14877 | 0.00913 | 0.66077 | 0.66077 | 0.009299
2019-03-26 16:18 | 1553631460.6672 | 281 | 0.20768 | 0.001905 | 3.8e-05 | 93.3635 | 92.2533 | 121.4337 | 21.2612 | 0.002803 | 0.002555 | 251.9585 | 0.24984 | 8.4004 | 0.003553 | 0.00496 | 0.004944 | 5.1e-05 | 0.001118 | 0.32477 | 1.3615 | 2.5475 | 100.79 | 21.17 | 15.435 | 89.2392 | 107.3884 | 83.2886 | 0.003853 | 0.99048 | 0.024106 | 24.3378 | 0.107 | 0.14863 | 0.009156 | 0.66364 | 0.66364 | 0.00925
2019-03-26 16:18 | 1553631461.6687 | 282 | 0.20769 | 0.001915 | 3.8e-05 | 93.4928 | 92.6429 | 121.311 | 21.2612 | 0.002714 | 0.002552 | 251.894 | 0.25006 | 8.4017 | 0.003559 | 0.004534 | 0.004464 | 5.4e-05 | 0.001113 | 0.32244 | 1.3896 | 2.5477 | 100.79 | 21.17 | 15.441 | 87.1054 | 107.5182 | 83.223 | 0.003855 | 0.99066 | 0.024096 | 24.3233 | 0.081682 | 0.14876 | 0.00909 | 0.65854 | 0.65854 | 0.009329
2019-03-26 16:18 | 1553631462.6683 | 283 | 0.2077 | 0.001919 | 3.8e-05 | 93.5722 | 92.5367 | 121.0446 | 21.2612 | 0.002791 | 0.002541 | 251.8118 | 0.24934 | 8.4021 | 0.003537 | 0.004778 | 0.004702 | 5.6e-05 | 0.001111 | 0.31868 | 1.381 | 2.5475 | 100.79 | 21.17 | 15.447 | 88.9728 | 106.9986 | 82.9491 | 0.003856 | 0.99142 | 0.024079 | 24.2872 | 0.099537 | 0.14833 | 0.009052 | 0.65536 | 0.65536 | 0.009348
2019-03-26 16:18 | 1553631463.6674 | 284 | 0.20771 | 0.001927 | 3.9e-05 | 93.7092 | 92.6023 | 120.6737 | 21.2612 | 0.002685 | 0.002545 | 251.7599 | 0.24988 | 8.398 | 0.00357 | 0.004701 | 0.004601 | 5.5e-05 | 0.001108 | 0.29927 | 1.3347 | 2.5474 | 100.79 | 21.17 | 15.459 | 86.4209 | 107.7846 | 83.0444 | 0.003859 | 0.99116 | 0.024039 | 24.2532 | 0.10629 | 0.14865 | 0.009008 | 0.65103 | 0.65103 | 0.00943
2019-03-26 16:18 | 1553631464.6675 | 285 | 0.2077 | 0.001917 | 3.8e-05 | 93.8122 | 92.6725 | 120.6737 | 21.2612 | 0.002784 | 0.002553 | 251.7318 | 0.24959 | 8.402 | 0.003529 | 0.005031 | 0.004935 | 5.3e-05 | 0.00111 | 0.31526 | 1.3129 | 2.5475 | 100.79 | 21.17 | 15.464 | 88.7995 | 106.7828 | 83.241 | 0.00386 | 0.99061 | 0.024032 | 24.2599 | 0.10933 | 0.14848 | 0.009027 | 0.65223 | 0.65223 | 0.009402
2019-03-26 16:18 | 1553631465.6675 | 286 | 0.20769 | 0.001924 | 3.8e-05 | 93.9164 | 92.7371 | 121.0188 | 21.2612 | 0.002769 | 0.002543 | 251.676 | 0.24965 | 8.4032 | 0.003561 | 0.005077 | 0.005021 | 5.1e-05 | 0.001104 | 0.3246 | 1.3639 | 2.5476 | 100.8 | 21.17 | 15.47 | 88.4345 | 107.5741 | 83.0068 | 0.003861 | 0.99126 | 0.024074 | 24.2866 | 0.11303 | 0.14851 | 0.009126 | 0.66055 | 0.66055 | 0.009286
2019-03-26 16:18 | 1553631466.6686 | 287 | 0.20773 | 0.001917 | 3.7e-05 | 94.0419 | 92.9609 | 121.1849 | 21.2612 | 0.002723 | 0.00254 | 251.6635 | 0.25004 | 8.4018 | 0.003602 | 0.004974 | 0.004895 | 5.5e-05 | 0.001099 | 0.31205 | 1.4018 | 2.5475 | 100.8 | 21.18 | 15.472 | 87.3141 | 108.5565 | 82.9126 | 0.003864 | 0.99153 | 0.024094 | 24.3 | 0.1034 | 0.14875 | 0.008877 | 0.64309 | 0.64309 | 0.009553
2019-03-26 16:18 | 1553631467.6686 | 288 | 0.20772 | 0.001917 | 3.7e-05 | 94.0856 | 93.1334 | 121.1608 | 21.2612 | 0.002755 | 0.002521 | 251.6149 | 0.2501 | 8.4035 | 0.003595 | 0.004862 | 0.004777 | 5.5e-05 | 0.001095 | 0.32068 | 1.3489 | 2.5474 | 100.8 | 21.18 | 15.474 | 88.091 | 108.3894 | 82.4751 | 0.003865 | 0.99275 | 0.024107 | 24.2827 | 0.090982 | 0.14879 | 0.008942 | 0.64812 | 0.64812 | 0.009481
2019-03-26 16:18 | 1553631468.6677 | 289 | 0.2077 | 0.001917 | 3.8e-05 | 94.1452 | 93.1885 | 121.1286 | 21.2612 | 0.002731 | 0.002499 | 251.5726 | 0.25028 | 8.4062 | 0.003655 | 0.005225 | 0.005136 | 5.6e-05 | 0.001091 | 0.29264 | 1.3542 | 2.5474 | 100.8 | 21.17 | 15.475 | 87.5155 | 109.8606 | 81.9361 | 0.003863 | 0.99425 | 0.024122 | 24.2611 | 0.091355 | 0.14889 | 0.009031 | 0.65498 | 0.65498 | 0.009388
2019-03-26 16:18 | 1553631469.6667 | 290 | 0.20773 | 0.001916 | 3.8e-05 | 94.2766 | 93.1969 | 121.2777 | 21.2612 | 0.002744 | 0.002542 | 251.5444 | 0.24963 | 8.4038 | 0.003671 | 0.005219 | 0.005187 | 5.2e-05 | 0.001097 | 0.32444 | 1.3283 | 2.5475 | 100.8 | 21.18 | 15.476 | 87.8412 | 110.2517 | 82.9616 | 0.003865 | 0.99139 | 0.024102 | 24.311 | 0.10302 | 0.1485 | 0.00889 | 0.64422 | 0.64422 | 0.00952
2019-03-26 16:18 | 1553631470.6678 | 291 | 0.20772 | 0.001914 | 3.9e-05 | 94.335 | 93.1932 | 121.1504 | 21.2612 | 0.002744 | 0.002539 | 251.4964 | 0.25078 | 8.4052 | 0.00366 | 0.005206 | 0.005175 | 5.6e-05 | 0.0011 | 0.32029 | 1.369 | 2.5476 | 100.8 | 21.17 | 15.478 | 87.8363 | 109.967 | 82.9088 | 0.003863 | 0.99154 | 0.024091 | 24.2965 | 0.10892 | 0.14919 | 0.008901 | 0.64464 | 0.64464 | 0.009558
2019-03-26 16:18 | 1553631471.6669 | 292 | 0.20769 | 0.001905 | 3.9e-05 | 94.3967 | 93.4709 | 120.8314 | 21.2612 | 0.002711 | 0.002503 | 251.4861 | 0.25019 | 8.403 | 0.003618 | 0.006013 | 0.005919 | 5.6e-05 | 0.001103 | 0.27959 | 1.3314 | 2.5475 | 100.8 | 21.17 | 15.479 | 87.0352 | 108.9467 | 82.0428 | 0.003864 | 0.99395 | 0.024088 | 24.2349 | 0.088155 | 0.14883 | 0.009127 | 0.66094 | 0.66094 | 0.0093
2019-03-26 16:18 | 1553631472.6669 | 293 | 0.20774 | 0.001904 | 4.1e-05 | 94.5945 | 93.562 | 120.9274 | 21.2612 | 0.00273 | 0.002577 | 251.4547 | 0.25049 | 8.3992 | 0.003615 | 0.005172 | 0.005115 | 5.3e-05 | 0.001099 | 0.32803 | 1.3827 | 2.5475 | 100.8 | 21.17 | 15.479 | 87.4881 | 108.8869 | 83.8058 | 0.003864 | 0.98904 | 0.024038 | 24.3046 | 0.098163 | 0.14901 | 0.008835 | 0.63831 | 0.63831 | 0.009642
2019-03-26 16:18 | 1553631473.6673 | 294 | 0.20777 | 0.001896 | 4.2e-05 | 94.7007 | 93.9419 | 121.3866 | 21.2612 | 0.002716 | 0.002565 | 251.3991 | 0.24939 | 8.4007 | 0.003594 | 0.005887 | 0.005815 | 5.3e-05 | 0.001097 | 0.31686 | 1.3611 | 2.5476 | 100.79 | 21.17 | 15.478 | 87.1454 | 108.3785 | 83.5334 | 0.003864 | 0.9898 | 0.024093 | 24.3414 | 0.071964 | 0.14836 | 0.00862 | 0.62409 | 0.62409 | 0.009818
2019-03-26 16:18 | 1553631474.667 | 295 | 0.20773 | 0.001899 | 4.2e-05 | 94.8504 | 93.9731 | 121.556 | 21.2612 | 0.002764 | 0.002515 | 251.3496 | 0.25024 | 8.4045 | 0.00361 | 0.005301 | 0.005216 | 5.2e-05 | 0.001096 | 0.29662 | 1.3408 | 2.5474 | 100.79 | 21.17 | 15.477 | 88.319 | 108.7725 | 82.3198 | 0.003864 | 0.99318 | 0.024151 | 24.317 | 0.083112 | 0.14887 | 0.008897 | 0.64571 | 0.64571 | 0.009522
2019-03-26 16:18 | 1553631475.668 | 296 | 0.20778 | 0.001897 | 4.3e-05 | 94.9289 | 93.9889 | 121.6258 | 21.2612 | 0.002724 | 0.002517 | 251.3247 | 0.24955 | 8.405 | 0.003656 | 0.005293 | 0.005216 | 5.5e-05 | 0.001099 | 0.30469 | 1.3772 | 2.5475 | 100.79 | 21.17 | 15.476 | 87.3461 | 109.888 | 82.376 | 0.003863 | 0.99302 | 0.024156 | 24.3258 | 0.08901 | 0.14846 | 0.008601 | 0.62428 | 0.62428 | 0.009821
2019-03-26 16:18 | 1553631476.6681 | 297 | 0.20775 | 0.00189 | 4.4e-05 | 95.0848 | 94.1557 | 121.4734 | 21.2612 | 0.002736 | 0.002547 | 251.2602 | 0.24947 | 8.4036 | 0.003577 | 0.005216 | 0.005111 | 5.6e-05 | 0.001101 | 0.31022 | 1.3486 | 2.5476 | 100.79 | 21.17 | 15.477 | 87.6467 | 107.9562 | 83.0996 | 0.003864 | 0.99101 | 0.024116 | 24.3353 | 0.087829 | 0.14841 | 0.008756 | 0.63444 | 0.63444 | 0.009661
2019-03-26 16:18 | 1553631477.6671 | 298 | 0.20776 | 0.001889 | 4.5e-05 | 95.1426 | 94.2606 | 121.4129 | 21.2612 | 0.002705 | 0.002581 | 251.2147 | 0.25011 | 8.4037 | 0.00361 | 0.004918 | 0.004885 | 5.7e-05 | 0.001109 | 0.29927 | 1.354 | 2.5475 | 100.79 | 21.17 | 15.478 | 86.9016 | 108.7496 | 83.9063 | 0.003864 | 0.98877 | 0.024083 | 24.3568 | 0.083312 | 0.14879 | 0.008738 | 0.63218 | 0.63218 | 0.009721
2019-03-26 16:18 | 1553631478.6672 | 299 | 0.20773 | 0.001893 | 4.6e-05 | 95.2025 | 94.4169 | 121.1842 | 21.2612 | 0.002777 | 0.002497 | 251.1629 | 0.25 | 8.4022 | 0.003497 | 0.005216 | 0.005065 | 5.6e-05 | 0.001107 | 0.30215 | 1.376 | 2.5475 | 100.79 | 21.17 | 15.479 | 88.6158 | 106.0075 | 81.883 | 0.003864 | 0.9944 | 0.024129 | 24.2648 | 0.074123 | 0.14873 | 0.008897 | 0.64492 | 0.64492 | 0.009524
2019-03-26 16:18 | 1553631479.6674 | 300 | 0.20775 | 0.001891 | 4.7e-05 | 95.1837 | 94.3123 | 121.2173 | 21.2612 | 0.002728 | 0.002545 | 251.1285 | 0.24963 | 8.4045 | 0.003596 | 0.005164 | 0.005188 | 5.7e-05 | 0.001105 | 0.31022 | 1.3151 | 2.5474 | 100.79 | 21.17 | 15.48 | 87.4491 | 108.4154 | 83.0388 | 0.003864 | 0.99118 | 0.024093 | 24.3076 | 0.082258 | 0.14851 | 0.008783 | 0.63563 | 0.63563 | 0.009649
2019-03-26 16:18 | 1553631480.6676 | 301 | 0.20775 | 0.001886 | 4.9e-05 | 95.2365 | 94.3201 | 121.2401 | 21.2612 | 0.002723 | 0.00255 | 251.1003 | 0.24968 | 8.4042 | 0.003583 | 0.005509 | 0.005455 | 5.6e-05 | 0.001104 | 0.29856 | 1.358 | 2.5474 | 100.79 | 21.17 | 15.48 | 87.3141 | 108.1026 | 83.1598 | 0.003864 | 0.99084 | 0.024091 | 24.314 | 0.086488 | 0.14853 | 0.008789 | 0.63592 | 0.63592 | 0.009647
2019-03-26 16:18 | 1553631481.6683 | 302 | 0.20776 | 0.001889 | 5e-05 | 95.3271 | 94.7292 | 121.3471 | 21.2612 | 0.002702 | 0.00253 | 251.0358 | 0.25033 | 8.4002 | 0.003529 | 0.004875 | 0.004717 | 5.6e-05 | 0.001103 | 0.30486 | 1.3107 | 2.5474 | 100.79 | 21.17 | 15.481 | 86.8229 | 106.7969 | 82.6895 | 0.003865 | 0.99215 | 0.024118 | 24.3087 | 0.056283 | 0.14893 | 0.008698 | 0.62993 | 0.62993 | 0.009764
2019-03-26 16:18 | 1553631482.6684 | 303 | 0.20776 | 0.001887 | 5e-05 | 95.3312 | 94.4626 | 121.1152 | 21.2612 | 0.002774 | 0.002526 | 250.9932 | 0.24987 | 8.4018 | 0.003511 | 0.005356 | 0.005405 | 5.5e-05 | 0.001098 | 0.32405 | 1.3208 | 2.5474 | 100.79 | 21.17 | 15.482 | 88.5518 | 106.3459 | 82.588 | 0.003865 | 0.99243 | 0.024098 | 24.282 | 0.081877 | 0.14865 | 0.008726 | 0.63145 | 0.63145 | 0.009722
2019-03-26 16:18 | 1553631483.6664 | 304 | 0.20775 | 0.001887 | 5e-05 | 95.3782 | 94.4838 | 120.9411 | 21.2612 | 0.002783 | 0.00253 | 250.946 | 0.24988 | 8.4022 | 0.00353 | 0.005238 | 0.005184 | 6e-05 | 0.001098 | 0.30934 | 1.3386 | 2.5475 | 100.79 | 21.17 | 15.485 | 88.7675 | 106.8103 | 82.695 | 0.003866 | 0.99213 | 0.024077 | 24.2682 | 0.084274 | 0.14865 | 0.008783 | 0.63503 | 0.63503 | 0.009668
2019-03-26 16:18 | 1553631484.6665 | 305 | 0.20771 | 0.001884 | 5e-05 | 95.4487 | 94.7409 | 120.6587 | 21.2612 | 0.002736 | 0.002544 | 250.9005 | 0.24995 | 8.4009 | 0.00356 | 0.004754 | 0.004652 | 5.5e-05 | 0.001095 | 0.31094 | 1.4142 | 2.5477 | 100.79 | 21.17 | 15.487 | 87.6481 | 107.5551 | 83.0179 | 0.003866 | 0.99123 | 0.024038 | 24.2508 | 0.066576 | 0.1487 | 0.009007 | 0.65022 | 0.65022 | 0.009445
2019-03-26 16:18 | 1553631485.6665 | 306 | 0.20769 | 0.001886 | 5e-05 | 95.5751 | 94.6804 | 120.3106 | 21.2612 | 0.002752 | 0.002529 | 250.8243 | 0.25015 | 8.4027 | 0.003529 | 0.005114 | 0.005087 | 5.7e-05 | 0.001092 | 0.30973 | 1.3593 | 2.5476 | 100.79 | 21.17 | 15.49 | 88.0113 | 106.7986 | 82.6568 | 0.003867 | 0.99224 | 0.024016 | 24.2035 | 0.084122 | 0.14882 | 0.009167 | 0.66119 | 0.66119 | 0.009295
2019-03-26 16:18 | 1553631486.6676 | 307 | 0.2077 | 0.001884 | 5.1e-05 | 95.5355 | 94.6539 | 120.3547 | 21.2612 | 0.00275 | 0.00252 | 250.7771 | 0.24972 | 8.4015 | 0.003563 | 0.004673 | 0.004577 | 5.7e-05 | 0.001092 | 0.33124 | 1.4133 | 2.5476 | 100.79 | 21.17 | 15.496 | 87.987 | 107.6207 | 82.4421 | 0.003868 | 0.99284 | 0.024027 | 24.2007 | 0.082928 | 0.14856 | 0.009083 | 0.65537 | 0.65537 | 0.009362
2019-03-26 16:18 | 1553631487.667 | 308 | 0.20773 | 0.001887 | 5.2e-05 | 95.5556 | 94.794 | 120.462 | 21.2612 | 0.002728 | 0.002557 | 250.7415 | 0.25086 | 8.4028 | 0.003524 | 0.005272 | 0.005078 | 5.5e-05 | 0.001092 | 0.32693 | 1.3981 | 2.5475 | 100.79 | 21.17 | 15.499 | 87.4546 | 106.6773 | 83.3321 | 0.003869 | 0.99036 | 0.024008 | 24.2417 | 0.071571 | 0.14924 | 0.008914 | 0.64257 | 0.64257 | 0.009592
2019-03-26 16:18 | 1553631488.6676 | 309 | 0.20775 | 0.001896 | 5.3e-05 | 95.6406 | 94.8885 | 120.9695 | 21.2612 | 0.002758 | 0.002501 | 250.6961 | 0.25001 | 8.4016 | 0.003465 | 0.005281 | 0.005307 | 5.7e-05 | 0.00109 | 0.29446 | 1.3794 | 2.5474 | 100.79 | 21.17 | 15.504 | 88.1613 | 105.2458 | 81.9976 | 0.00387 | 0.99408 | 0.024104 | 24.2472 | 0.070616 | 0.14873 | 0.008729 | 0.63156 | 0.63156 | 0.009726
2019-03-26 16:18 | 1553631489.6667 | 310 | 0.20774 | 0.001893 | 5.4e-05 | 95.6388 | 94.8586 | 120.8965 | 21.2612 | 0.002733 | 0.002557 | 250.6289 | 0.24966 | 8.4026 | 0.00347 | 0.0053 | 0.0052 | 5.6e-05 | 0.001086 | 0.30901 | 1.4133 | 2.5476 | 100.79 | 21.17 | 15.506 | 87.5685 | 105.3559 | 83.3353 | 0.003871 | 0.99035 | 0.024051 | 24.2855 | 0.073271 | 0.14852 | 0.008828 | 0.63736 | 0.63736 | 0.009624
2019-03-26 16:18 | 1553631490.6667 | 311 | 0.20775 | 0.001891 | 5.4e-05 | 95.6465 | 94.7752 | 120.8681 | 21.2612 | 0.002688 | 0.002533 | 250.5713 | 0.24931 | 8.4023 | 0.003483 | 0.005037 | 0.004904 | 5.3e-05 | 0.001087 | 0.29535 | 1.3759 | 2.5475 | 100.79 | 21.17 | 15.509 | 86.4841 | 105.6818 | 82.759 | 0.003872 | 0.99195 | 0.024068 | 24.263 | 0.081854 | 0.14832 | 0.008762 | 0.63296 | 0.63296 | 0.009677
2019-03-26 16:18 | 1553631491.6678 | 312 | 0.20771 | 0.001891 | 5.6e-05 | 95.6238 | 94.9238 | 121.2122 | 21.2612 | 0.002742 | 0.002512 | 250.5254 | 0.25049 | 8.4036 | 0.003485 | 0.00508 | 0.005069 | 5.5e-05 | 0.001083 | 0.31222 | 1.3494 | 2.5475 | 100.79 | 21.17 | 15.512 | 87.7786 | 105.7186 | 82.2409 | 0.003872 | 0.9934 | 0.02412 | 24.2798 | 0.065716 | 0.14902 | 0.008994 | 0.65114 | 0.65114 | 0.009452
2019-03-26 16:18 | 1553631492.6668 | 313 | 0.20771 | 0.001888 | 5.6e-05 | 95.6533 | 94.7312 | 121.506 | 21.2612 | 0.002704 | 0.002557 | 250.4972 | 0.24974 | 8.4052 | 0.003391 | 0.004656 | 0.004566 | 5.9e-05 | 0.001081 | 0.31094 | 1.3349 | 2.5475 | 100.79 | 21.16 | 15.514 | 86.8658 | 103.4578 | 83.3237 | 0.00387 | 0.99038 | 0.024112 | 24.3463 | 0.086646 | 0.14857 | 0.009039 | 0.65417 | 0.65417 | 0.00938
2019-03-26 16:18 | 1553631493.6679 | 314 | 0.20771 | 0.001891 | 5.6e-05 | 95.7572 | 94.7981 | 121.7286 | 21.2612 | 0.002705 | 0.002573 | 250.4318 | 0.25046 | 8.405 | 0.003404 | 0.005098 | 0.004921 | 5.7e-05 | 0.001082 | 0.32117 | 1.3269 | 2.5474 | 100.79 | 21.16 | 15.513 | 86.8908 | 103.7518 | 83.7246 | 0.00387 | 0.98927 | 0.024121 | 24.3822 | 0.090042 | 0.149 | 0.008981 | 0.65016 | 0.65016 | 0.009465
2019-03-26 16:18 | 1553631494.6679 | 315 | 0.20771 | 0.00189 | 5.7e-05 | 95.8038 | 94.8687 | 121.9605 | 21.2612 | 0.002747 | 0.002544 | 250.3974 | 0.25035 | 8.4011 | 0.003444 | 0.005608 | 0.005606 | 5.8e-05 | 0.001084 | 0.33035 | 1.3771 | 2.5476 | 100.79 | 21.17 | 15.511 | 87.898 | 104.726 | 83.0172 | 0.003872 | 0.99123 | 0.024168 | 24.3813 | 0.087739 | 0.14893 | 0.008994 | 0.65235 | 0.65235 | 0.009429
2019-03-26 16:18 | 1553631495.667 | 316 | 0.20775 | 0.001889 | 5.7e-05 | 95.8531 | 94.9685 | 122.2154 | 21.2612 | 0.002779 | 0.002505 | 250.3456 | 0.24948 | 8.4024 | 0.003525 | 0.004565 | 0.004584 | 5.5e-05 | 0.001083 | 0.31669 | 1.3467 | 2.5474 | 100.79 | 21.17 | 15.509 | 88.669 | 106.6962 | 82.0866 | 0.003872 | 0.99383 | 0.024224 | 24.3748 | 0.082932 | 0.14842 | 0.00875 | 0.63605 | 0.63605 | 0.009637
2019-03-26 16:18 | 1553631496.667 | 317 | 0.20771 | 0.001888 | 5.8e-05 | 95.9488 | 94.8448 | 122.0454 | 21.2612 | 0.002731 | 0.002527 | 250.3108 | 0.25063 | 8.4024 | 0.003432 | 0.004132 | 0.004025 | 5.7e-05 | 0.001076 | 0.3194 | 1.336 | 2.5475 | 100.79 | 21.17 | 15.503 | 87.5239 | 104.4451 | 82.6005 | 0.00387 | 0.9924 | 0.02419 | 24.3755 | 0.10351 | 0.1491 | 0.009034 | 0.65578 | 0.65578 | 0.00939
2019-03-26 16:18 | 1553631497.6685 | 318 | 0.20772 | 0.00189 | 5.9e-05 | 95.8953 | 95.0273 | 121.8818 | 21.2612 | 0.002716 | 0.002548 | 250.2673 | 0.25047 | 8.4021 | 0.00344 | 0.004747 | 0.004681 | 5.6e-05 | 0.001068 | 0.28871 | 1.3294 | 2.5475 | 100.79 | 21.17 | 15.499 | 87.168 | 104.6325 | 83.1107 | 0.003869 | 0.99098 | 0.024157 | 24.3766 | 0.081331 | 0.14901 | 0.008962 | 0.64959 | 0.64959 | 0.009474
2019-03-26 16:18 | 1553631498.6671 | 319 | 0.2078 | 0.001893 | 5.9e-05 | 95.8729 | 95.14 | 121.392 | 21.2612 | 0.002654 | 0.002526 | 250.2139 | 0.24957 | 8.4028 | 0.003393 | 0.004592 | 0.004538 | 5.5e-05 | 0.001068 | 0.30431 | 1.3423 | 2.5474 | 100.79 | 21.16 | 15.495 | 85.6676 | 103.5064 | 82.5779 | 0.003866 | 0.99246 | 0.024126 | 24.3094 | 0.068645 | 0.14847 | 0.008436 | 0.61053 | 0.61053 | 0.010043
2019-03-26 16:18 | 1553631499.6672 | 320 | 0.2077 | 0.001888 | 6e-05 | 95.9467 | 94.8777 | 121.2486 | 21.2612 | 0.002814 | 0.00252 | 250.1675 | 0.25017 | 8.4025 | 0.003461 | 0.004733 | 0.004646 | 5.6e-05 | 0.00107 | 0.30342 | 1.3689 | 2.5476 | 100.79 | 21.16 | 15.489 | 89.5243 | 105.1342 | 82.4376 | 0.003864 | 0.99285 | 0.024117 | 24.2902 | 0.10021 | 0.14883 | 0.009054 | 0.65516 | 0.65516 | 0.009382
2019-03-26 16:18 | 1553631500.6673 | 321 | 0.2077 | 0.001894 | 6.1e-05 | 95.9742 | 95.0015 | 121.1939 | 21.2612 | 0.002763 | 0.002518 | 250.1002 | 0.2498 | 8.4024 | 0.003457 | 0.005184 | 0.005077 | 5.1e-05 | 0.001073 | 0.28313 | 1.3529 | 2.5474 | 100.79 | 21.16 | 15.488 | 88.2918 | 105.0499 | 82.3883 | 0.003864 | 0.99299 | 0.024113 | 24.283 | 0.091116 | 0.14861 | 0.009039 | 0.6539 | 0.6539 | 0.009386
2019-03-26 16:18 | 1553631501.6673 | 322 | 0.20768 | 0.001889 | 6.3e-05 | 95.9928 | 95.1077 | 121.161 | 21.2612 | 0.002757 | 0.002485 | 250.0599 | 0.25045 | 8.4049 | 0.003432 | 0.005049 | 0.004968 | 5.7e-05 | 0.001076 | 0.30574 | 1.3585 | 2.5475 | 100.79 | 21.16 | 15.485 | 88.1488 | 104.4405 | 81.6046 | 0.003863 | 0.99518 | 0.024136 | 24.253 | 0.082854 | 0.14899 | 0.009189 | 0.66534 | 0.66534 | 0.009248
2019-03-26 16:18 | 1553631502.6674 | 323 | 0.20774 | 0.001893 | 6.4e-05 | 96.0217 | 95.2153 | 121.0493 | 21.2612 | 0.002707 | 0.002555 | 250.0089 | 0.25017 | 8.4031 | 0.003474 | 0.004789 | 0.004705 | 5.6e-05 | 0.001082 | 0.29048 | 1.3947 | 2.5476 | 100.79 | 21.16 | 15.484 | 86.943 | 105.4497 | 83.2744 | 0.003863 | 0.99052 | 0.024068 | 24.2988 | 0.075438 | 0.14883 | 0.008801 | 0.63526 | 0.63526 | 0.009676
2019-03-26 16:18 | 1553631503.6674 | 324 | 0.20772 | 0.001895 | 6.4e-05 | 96.0332 | 95.2398 | 121.0103 | 21.2612 | 0.002727 | 0.002556 | 249.9864 | 0.24966 | 8.4039 | 0.003447 | 0.004581 | 0.004511 | 5.6e-05 | 0.001084 | 0.32803 | 1.41 | 2.5476 | 100.79 | 21.16 | 15.484 | 87.4195 | 104.8172 | 83.3009 | 0.003863 | 0.99045 | 0.024064 | 24.2958 | 0.074205 | 0.14853 | 0.008936 | 0.64491 | 0.64491 | 0.009512
2019-03-26 16:18 | 1553631504.6664 | 325 | 0.20769 | 0.001896 | 6.3e-05 | 96.122 | 95.1053 | 120.9584 | 21.2612 | 0.002793 | 0.00254 | 249.9228 | 0.25095 | 8.4019 | 0.00342 | 0.005152 | 0.005113 | 5.5e-05 | 0.001084 | 0.31116 | 1.3598 | 2.5474 | 100.79 | 21.16 | 15.483 | 89.015 | 104.1504 | 82.9265 | 0.003863 | 0.99149 | 0.024071 | 24.2778 | 0.095114 | 0.14929 | 0.009131 | 0.65929 | 0.65929 | 0.009352
2019-03-26 16:18 | 1553631505.6665 | 326 | 0.20774 | 0.001901 | 6.2e-05 | 96.1463 | 95.3316 | 120.9012 | 21.2612 | 0.002736 | 0.002588 | 249.8859 | 0.25043 | 8.4015 | 0.003514 | 0.005002 | 0.004973 | 5.7e-05 | 0.001085 | 0.32333 | 1.3668 | 2.5476 | 100.79 | 21.16 | 15.483 | 87.6279 | 106.4371 | 84.0881 | 0.003863 | 0.98826 | 0.024026 | 24.3116 | 0.076114 | 0.14898 | 0.008837 | 0.63685 | 0.63685 | 0.009662
2019-03-26 16:18 | 1553631506.6666 | 327 | 0.20773 | 0.0019 | 6.1e-05 | 96.1805 | 95.2697 | 121.0073 | 21.2612 | 0.002743 | 0.002547 | 249.8176 | 0.25041 | 8.4028 | 0.003589 | 0.005777 | 0.005726 | 5.4e-05 | 0.001083 | 0.28174 | 1.3643 | 2.5475 | 100.79 | 21.16 | 15.485 | 87.7966 | 108.2627 | 83.0881 | 0.003863 | 0.99104 | 0.024071 | 24.2882 | 0.085104 | 0.14897 | 0.008884 | 0.64156 | 0.64156 | 0.00959
2019-03-26 16:18 | 1553631507.6666 | 328 | 0.20774 | 0.001904 | 5.9e-05 | 96.1761 | 95.2516 | 121.2018 | 21.2612 | 0.002751 | 0.002534 | 249.7414 | 0.24986 | 8.4026 | 0.003561 | 0.005222 | 0.005157 | 5.5e-05 | 0.001088 | 0.31205 | 1.3766 | 2.5475 | 100.79 | 21.16 | 15.486 | 88.0012 | 107.575 | 82.7847 | 0.003863 | 0.99188 | 0.0241 | 24.2974 | 0.086397 | 0.14864 | 0.008798 | 0.63617 | 0.63617 | 0.00965
2019-03-26 16:18 | 1553631508.6667 | 329 | 0.20777 | 0.001898 | 5.9e-05 | 96.1271 | 95.5531 | 121.3258 | 21.2612 | 0.002726 | 0.002536 | 249.7016 | 0.25028 | 8.4013 | 0.003504 | 0.004733 | 0.004654 | 5.6e-05 | 0.001096 | 0.31813 | 1.3478 | 2.5475 | 100.79 | 21.16 | 15.489 | 87.3966 | 106.186 | 82.8205 | 0.003864 | 0.99178 | 0.024111 | 24.311 | 0.053568 | 0.14889 | 0.008648 | 0.6256 | 0.6256 | 0.009829
2019-03-26 16:18 | 1553631509.6667 | 330 | 0.20774 | 0.001884 | 5.9e-05 | 96.0934 | 95.2731 | 121.2373 | 21.2612 | 0.002725 | 0.002579 | 249.6445 | 0.25037 | 8.4053 | 0.003501 | 0.004809 | 0.004755 | 5.7e-05 | 0.001102 | 0.3115 | 1.3519 | 2.5475 | 100.79 | 21.16 | 15.49 | 87.3843 | 106.1208 | 83.8628 | 0.003864 | 0.98889 | 0.024067 | 24.3377 | 0.076687 | 0.14894 | 0.008807 | 0.63593 | 0.63593 | 0.009673
2019-03-26 16:18 | 1553631510.6677 | 331 | 0.20777 | 0.001882 | 6.1e-05 | 96.2097 | 95.4963 | 121.1516 | 21.2612 | 0.002706 | 0.002532 | 249.6111 | 0.24907 | 8.4015 | 0.003503 | 0.005001 | 0.004885 | 5.7e-05 | 0.001106 | 0.31885 | 1.3048 | 2.5473 | 100.79 | 21.16 | 15.491 | 86.9228 | 106.1554 | 82.7361 | 0.003865 | 0.99202 | 0.024097 | 24.2907 | 0.066575 | 0.14817 | 0.00863 | 0.62378 | 0.62378 | 0.00981
2019-03-26 16:19 | 1553631511.6678 | 332 | 0.20773 | 0.001869 | 6.3e-05 | 96.1156 | 95.2075 | 120.9159 | 21.2612 | 0.002737 | 0.002551 | 249.5656 | 0.24956 | 8.4004 | 0.00339 | 0.004583 | 0.004538 | 5.3e-05 | 0.001099 | 0.32333 | 1.3195 | 2.5474 | 100.79 | 21.17 | 15.495 | 87.6513 | 103.4294 | 83.1802 | 0.003868 | 0.99078 | 0.024058 | 24.2822 | 0.084904 | 0.14846 | 0.008932 | 0.64451 | 0.64451 | 0.009514
2019-03-26 16:19 | 1553631512.6669 | 333 | 0.20771 | 0.001861 | 6.5e-05 | 96.1598 | 95.1638 | 121.0814 | 21.2612 | 0.002731 | 0.002551 | 249.4881 | 0.24993 | 8.4038 | 0.003481 | 0.005217 | 0.005171 | 5.3e-05 | 0.001096 | 0.30143 | 1.3141 | 2.5473 | 100.79 | 21.17 | 15.496 | 87.5155 | 105.6369 | 83.1973 | 0.003868 | 0.99073 | 0.024074 | 24.2994 | 0.09313 | 0.14869 | 0.009054 | 0.65372 | 0.65372 | 0.009393
2019-03-26 16:19 | 1553631513.6678 | 334 | 0.20772 | 0.001858 | 6.7e-05 | 96.1519 | 95.0573 | 121.169 | 21.2612 | 0.002714 | 0.002556 | 249.4438 | 0.25007 | 8.4058 | 0.003507 | 0.005182 | 0.005079 | 5.7e-05 | 0.001094 | 0.32925 | 1.3197 | 2.5474 | 100.79 | 21.16 | 15.5 | 87.104 | 106.2491 | 83.2987 | 0.003867 | 0.99045 | 0.024079 | 24.3116 | 0.10241 | 0.14877 | 0.008951 | 0.64629 | 0.64629 | 0.009507
2019-03-26 16:19 | 1553631514.6679 | 335 | 0.2077 | 0.001852 | 6.8e-05 | 96.1826 | 95.265 | 121.1619 | 21.2612 | 0.00275 | 0.002527 | 249.3837 | 0.24992 | 8.4031 | 0.003507 | 0.004707 | 0.004683 | 5.5e-05 | 0.001092 | 0.30486 | 1.3815 | 2.5476 | 100.79 | 21.17 | 15.503 | 87.9675 | 106.2505 | 82.6051 | 0.00387 | 0.99238 | 0.024102 | 24.2872 | 0.085738 | 0.14868 | 0.0091 | 0.65763 | 0.65763 | 0.009337
2019-03-26 16:19 | 1553631515.666 | 336 | 0.20773 | 0.001861 | 6.8e-05 | 96.1448 | 95.1426 | 121.3199 | 21.2612 | 0.00283 | 0.002565 | 249.3199 | 0.24954 | 8.4028 | 0.003415 | 0.004693 | 0.004629 | 5.5e-05 | 0.001092 | 0.32659 | 1.3587 | 2.5477 | 100.79 | 21.16 | 15.505 | 89.9048 | 104.0215 | 83.5358 | 0.003868 | 0.98979 | 0.024086 | 24.3348 | 0.09372 | 0.14845 | 0.008904 | 0.64294 | 0.64294 | 0.009536
2019-03-26 16:19 | 1553631516.666 | 337 | 0.20774 | 0.001861 | 7e-05 | 96.1687 | 95.0176 | 121.2537 | 21.2612 | 0.002726 | 0.002533 | 249.2513 | 0.24966 | 8.404 | 0.003431 | 0.004826 | 0.004726 | 5.4e-05 | 0.00109 | 0.33196 | 1.3606 | 2.5475 | 100.79 | 21.17 | 15.513 | 87.4029 | 104.4058 | 82.7558 | 0.003873 | 0.99196 | 0.024106 | 24.3016 | 0.10771 | 0.14852 | 0.008864 | 0.64043 | 0.64043 | 0.009578
2019-03-26 16:19 | 1553631517.6671 | 338 | 0.20771 | 0.001864 | 7.1e-05 | 96.1108 | 95.1053 | 121.0636 | 21.2612 | 0.002731 | 0.002529 | 249.2081 | 0.25006 | 8.4044 | 0.003389 | 0.005305 | 0.005232 | 5.9e-05 | 0.001088 | 0.31006 | 1.3417 | 2.5476 | 100.79 | 21.16 | 15.516 | 87.5107 | 103.4108 | 82.6707 | 0.003871 | 0.9922 | 0.02409 | 24.2796 | 0.09407 | 0.14876 | 0.009039 | 0.6527 | 0.6527 | 0.009413
2019-03-26 16:19 | 1553631518.6665 | 339 | 0.20768 | 0.001868 | 7.2e-05 | 96.1126 | 94.9658 | 121.0934 | 21.2612 | 0.002749 | 0.002623 | 249.1449 | 0.25046 | 8.4028 | 0.00346 | 0.005039 | 0.004934 | 5.4e-05 | 0.001085 | 0.33251 | 1.3879 | 2.5475 | 100.79 | 21.17 | 15.521 | 87.956 | 105.1207 | 84.9308 | 0.003875 | 0.98594 | 0.024017 | 24.3596 | 0.10737 | 0.149 | 0.009179 | 0.66076 | 0.66076 | 0.009313
2019-03-26 16:19 | 1553631519.6666 | 340 | 0.20771 | 0.001876 | 7.4e-05 | 96.1073 | 94.9493 | 121.3652 | 21.2612 | 0.002748 | 0.002564 | 249.1014 | 0.24981 | 8.403 | 0.003454 | 0.004989 | 0.004944 | 5.2e-05 | 0.00108 | 0.31924 | 1.3775 | 2.5476 | 100.79 | 21.17 | 15.522 | 87.9362 | 104.9752 | 83.4937 | 0.003875 | 0.98991 | 0.024092 | 24.3379 | 0.10842 | 0.14861 | 0.009029 | 0.65179 | 0.65179 | 0.009417
2019-03-26 16:19 | 1553631520.6662 | 341 | 0.20765 | 0.001869 | 7.6e-05 | 96.1749 | 95.2092 | 121.541 | 21.2612 | 0.002787 | 0.002552 | 249.0587 | 0.24998 | 8.3994 | 0.00343 | 0.004992 | 0.004886 | 5.4e-05 | 0.001083 | 0.28335 | 1.3773 | 2.5475 | 100.79 | 21.17 | 15.523 | 88.8751 | 104.3816 | 83.2184 | 0.003875 | 0.99067 | 0.024119 | 24.3462 | 0.090266 | 0.14871 | 0.00935 | 0.67571 | 0.67571 | 0.009089
2019-03-26 16:19 | 1553631521.6663 | 342 | 0.20765 | 0.001868 | 7.8e-05 | 96.2342 | 95.1023 | 121.7422 | 21.2612 | 0.002741 | 0.00255 | 248.9863 | 0.24964 | 8.4012 | 0.003471 | 0.004961 | 0.004867 | 5.1e-05 | 0.001083 | 0.30646 | 1.3789 | 2.5476 | 100.79 | 21.17 | 15.523 | 87.7552 | 105.3849 | 83.1629 | 0.003875 | 0.99083 | 0.024141 | 24.3644 | 0.10583 | 0.14851 | 0.00935 | 0.67627 | 0.67627 | 0.00907
2019-03-26 16:19 | 1553631522.6663 | 343 | 0.20768 | 0.001869 | 7.9e-05 | 96.1852 | 94.9967 | 121.834 | 21.2612 | 0.002742 | 0.002539 | 248.9586 | 0.24985 | 8.4009 | 0.003498 | 0.004829 | 0.004797 | 5.2e-05 | 0.00108 | 0.29911 | 1.3305 | 2.5473 | 100.79 | 21.17 | 15.522 | 87.7911 | 106.0429 | 82.8963 | 0.003875 | 0.99157 | 0.024159 | 24.3645 | 0.11121 | 0.14864 | 0.009165 | 0.66321 | 0.66321 | 0.009256
2019-03-26 16:19 | 1553631523.6678 | 344 | 0.20765 | 0.001871 | 8e-05 | 96.0952 | 94.9759 | 121.5105 | 21.2612 | 0.00277 | 0.002499 | 248.9014 | 0.24957 | 8.4033 | 0.003504 | 0.004967 | 0.004867 | 5.2e-05 | 0.00108 | 0.32477 | 1.3921 | 2.5475 | 100.79 | 21.17 | 15.516 | 88.4682 | 106.1947 | 81.9407 | 0.003873 | 0.99424 | 0.024159 | 24.2994 | 0.10479 | 0.14847 | 0.009356 | 0.67709 | 0.67709 | 0.009056
2019-03-26 16:19 | 1553631524.6675 | 345 | 0.20765 | 0.001875 | 8.2e-05 | 96.1247 | 95.0334 | 121.2551 | 21.2612 | 0.002784 | 0.002547 | 248.8618 | 0.2487 | 8.4032 | 0.003509 | 0.004865 | 0.004792 | 5.1e-05 | 0.001072 | 0.27975 | 1.3127 | 2.5475 | 100.79 | 21.16 | 15.511 | 88.7995 | 106.3206 | 83.0833 | 0.00387 | 0.99105 | 0.024095 | 24.3129 | 0.10213 | 0.14795 | 0.009359 | 0.67541 | 0.67541 | 0.009047
2019-03-26 16:19 | 1553631525.6675 | 346 | 0.20764 | 0.001877 | 8.4e-05 | 96.1288 | 94.8105 | 120.6873 | 21.2612 | 0.002731 | 0.002505 | 248.8208 | 0.24949 | 8.403 | 0.003549 | 0.004802 | 0.004778 | 5.3e-05 | 0.001069 | 0.30829 | 1.3416 | 2.5475 | 100.79 | 21.16 | 15.506 | 87.5107 | 107.2821 | 82.0873 | 0.003868 | 0.99383 | 0.024073 | 24.222 | 0.12352 | 0.14842 | 0.009403 | 0.67782 | 0.67782 | 0.009043
2019-03-26 16:19 | 1553631526.6676 | 347 | 0.20766 | 0.001878 | 8.5e-05 | 96.102 | 94.6771 | 120.1097 | 21.2612 | 0.002741 | 0.002578 | 248.7594 | 0.24939 | 8.4052 | 0.00353 | 0.005125 | 0.005075 | 5.4e-05 | 0.001067 | 0.29717 | 1.326 | 2.5474 | 100.79 | 21.16 | 15.493 | 87.7574 | 106.8197 | 83.8306 | 0.003865 | 0.98898 | 0.023956 | 24.2232 | 0.13362 | 0.14836 | 0.009309 | 0.66768 | 0.66768 | 0.009177
2019-03-26 16:19 | 1553631527.6666 | 348 | 0.20764 | 0.001876 | 8.9e-05 | 96.0397 | 94.8781 | 119.8536 | 21.2612 | 0.002706 | 0.002525 | 248.6901 | 0.24978 | 8.4017 | 0.003515 | 0.004794 | 0.004714 | 5.1e-05 | 0.001067 | 0.3282 | 1.3142 | 2.5474 | 100.79 | 21.16 | 15.486 | 86.9197 | 106.4567 | 82.5561 | 0.003863 | 0.99252 | 0.023973 | 24.1541 | 0.10884 | 0.1486 | 0.009405 | 0.67493 | 0.67493 | 0.009093
2019-03-26 16:19 | 1553631528.6666 | 349 | 0.20765 | 0.001878 | 9.2e-05 | 96.0987 | 94.795 | 119.4928 | 21.2612 | 0.002803 | 0.002625 | 248.6129 | 0.2498 | 8.4026 | 0.003543 | 0.004873 | 0.00482 | 5.5e-05 | 0.001071 | 0.3298 | 1.3571 | 2.5475 | 100.79 | 21.16 | 15.478 | 89.2627 | 107.1228 | 84.9785 | 0.003861 | 0.98581 | 0.023856 | 24.1997 | 0.12217 | 0.14861 | 0.009365 | 0.66855 | 0.66855 | 0.00918
2019-03-26 16:19 | 1553631529.6667 | 350 | 0.20731 | 0.001879 | 9.7e-05 | 96.1135 | 94.8122 | 119.0668 | 21.2612 | 0.002742 | 0.00256 | 248.5381 | 0.24978 | 8.4026 | 0.003591 | 0.005172 | 0.005077 | 5.4e-05 | 0.001067 | 0.3314 | 1.3753 | 2.5476 | 100.79 | 21.16 | 15.461 | 87.7841 | 108.3044 | 83.4024 | 0.003857 | 0.99016 | 0.023866 | 24.1033 | 0.12193 | 0.14859 | 0.011373 | 0.81294 | 0.81294 | 0.007549
2019-03-26 16:19 | 1553631530.6658 | 351 | 0.20664 | 0.001882 | 0.0001 | 96.053 | 94.6666 | 118.4431 | 21.2612 | 0.002802 | 0.002668 | 248.4687 | 0.24994 | 8.4024 | 0.003625 | 0.005259 | 0.005276 | 5.1e-05 | 0.001065 | 0.32139 | 1.4055 | 2.5476 | 100.79 | 21.17 | 15.454 | 89.2377 | 109.1311 | 86.0025 | 0.003858 | 0.983 | 0.023717 | 24.1276 | 0.13004 | 0.14869 | 0.015429 | 1.0974 | 1.0974 | 0.005596
2019-03-26 16:19 | 1553631531.6668 | 352 | 0.20604 | 0.001912 | 0.000104 | 96.0943 | 94.7762 | 117.7476 | 21.2612 | 0.00283 | 0.002659 | 248.4089 | 0.25003 | 8.4015 | 0.003582 | 0.005385 | 0.005338 | 5.4e-05 | 0.001057 | 0.34435 | 1.3842 | 2.726 | 100.79 | 21.17 | 15.438 | 89.9103 | 108.0759 | 85.7949 | 0.003854 | 0.98356 | 0.023654 | 24.0497 | 0.12354 | 0.14874 | 0.01898 | 1.347 | 1.347 | 0.004561
2019-03-26 16:19 | 1553631532.6658 | 353 | 0.20566 | 0.001981 | 0.000107 | 96.1164 | 94.9275 | 117.2162 | 21.2612 | 0.002873 | 0.002635 | 248.3146 | 0.24998 | 8.4026 | 0.003573 | 0.005182 | 0.005166 | 5e-05 | 0.001054 | 0.32565 | 1.372 | 2.6879 | 100.79 | 21.17 | 15.431 | 90.9326 | 107.8734 | 85.2197 | 0.003852 | 0.98514 | 0.02362 | 23.9761 | 0.11133 | 0.14871 | 0.021135 | 1.498 | 1.498 | 0.0041
2019-03-26 16:19 | 1553631533.6663 | 354 | 0.20522 | 0.002105 | 0.000109 | 96.1746 | 94.8176 | 116.8001 | 21.2612 | 0.002842 | 0.002656 | 248.194 | 0.25028 | 8.4027 | 0.003642 | 0.004765 | 0.004649 | 5.2e-05 | 0.001053 | 0.33035 | 1.4006 | 2.5475 | 100.79 | 21.17 | 15.425 | 90.2003 | 109.5505 | 85.7135 | 0.003851 | 0.98379 | 0.023562 | 23.95 | 0.12711 | 0.14889 | 0.023607 | 1.6692 | 1.6692 | 0.003684
2019-03-26 16:19 | 1553631534.6669 | 355 | 0.20487 | 0.002393 | 0.000111 | 96.182 | 95.0159 | 116.4811 | 21.2612 | 0.002862 | 0.002689 | 248.0987 | 0.24903 | 8.4029 | 0.003683 | 0.004766 | 0.004659 | 5e-05 | 0.001048 | 0.31636 | 1.4206 | 2.5475 | 100.79 | 21.17 | 15.411 | 90.6833 | 110.5466 | 86.5067 | 0.003847 | 0.98162 | 0.023504 | 23.9437 | 0.1091 | 0.14815 | 0.025371 | 1.7896 | 1.7896 | 0.003419
2019-03-26 16:19 | 1553631535.667 | 356 | 0.20469 | 0.002754 | 0.000112 | 96.2056 | 95.1006 | 116.3189 | 21.2612 | 0.002884 | 0.002721 | 248.0049 | 0.2496 | 8.4003 | 0.003689 | 0.005189 | 0.005108 | 4.8e-05 | 0.001048 | 0.323 | 1.3843 | 2.5476 | 100.79 | 21.17 | 15.404 | 91.196 | 110.6728 | 87.2758 | 0.003845 | 0.97952 | 0.023462 | 23.9526 | 0.10333 | 0.14849 | 0.026037 | 1.8333 | 1.8333 | 0.003345
2019-03-26 16:19 | 1553631536.666 | 357 | 0.20445 | 0.003066 | 0.000112 | 96.2422 | 95.2243 | 116.4705 | 21.2612 | 0.002864 | 0.002716 | 247.9367 | 0.25001 | 8.4015 | 0.003543 | 0.005486 | 0.005444 | 4.9e-05 | 0.001046 | 0.30342 | 1.3717 | 2.5475 | 100.79 | 21.17 | 15.391 | 90.7185 | 107.1403 | 87.1593 | 0.003842 | 0.97984 | 0.023481 | 23.9643 | 0.095102 | 0.14873 | 0.027039 | 1.9055 | 1.9055 | 0.003224
2019-03-26 16:19 | 1553631537.6671 | 358 | 0.20444 | 0.003301 | 0.000112 | 96.2283 | 95.1618 | 116.6034 | 21.2612 | 0.002897 | 0.00271 | 247.8404 | 0.24966 | 8.4038 | 0.003605 | 0.005112 | 0.005038 | 5.2e-05 | 0.001042 | 0.33682 | 1.3906 | 2.5476 | 100.79 | 21.17 | 15.385 | 91.5049 | 108.6377 | 87.0133 | 0.003841 | 0.98024 | 0.023499 | 23.9731 | 0.099687 | 0.14853 | 0.026812 | 1.8909 | 1.8909 | 0.003244
2019-03-26 16:19 | 1553631538.6667 | 359 | 0.20445 | 0.003484 | 0.000113 | 96.3806 | 95.3917 | 116.7212 | 21.2612 | 0.002897 | 0.002758 | 247.7522 | 0.25112 | 8.403 | 0.003586 | 0.005321 | 0.005252 | 4.9e-05 | 0.001043 | 0.32549 | 1.3283 | 2.5475 | 100.79 | 21.17 | 15.379 | 91.5206 | 108.1808 | 88.1659 | 0.003839 | 0.97711 | 0.023474 | 24.0236 | 0.092248 | 0.14939 | 0.026557 | 1.8708 | 1.8708 | 0.003298
2019-03-26 16:19 | 1553631539.6674 | 360 | 0.20447 | 0.003657 | 0.000113 | 96.4302 | 95.4603 | 116.985 | 21.2612 | 0.002965 | 0.002682 | 247.6711 | 0.24933 | 8.4027 | 0.003543 | 0.004744 | 0.004676 | 4.6e-05 | 0.001042 | 0.33179 | 1.3399 | 2.5475 | 100.79 | 21.17 | 15.371 | 93.1621 | 107.1355 | 86.3318 | 0.003837 | 0.9821 | 0.02356 | 23.9896 | 0.090421 | 0.14833 | 0.026226 | 1.8542 | 1.8542 | 0.003304
2019-03-26 16:19 | 1553631540.6662 | 361 | 0.20447 | 0.003806 | 0.000113 | 96.4305 | 95.6082 | 117.2735 | 21.2612 | 0.002907 | 0.00273 | 247.5624 | 0.24931 | 8.4019 | 0.003589 | 0.005254 | 0.005185 | 4.9e-05 | 0.001032 | 0.3011 | 1.3635 | 2.5476 | 100.79 | 21.17 | 15.37 | 91.7473 | 108.2556 | 87.502 | 0.003837 | 0.97891 | 0.023551 | 24.0582 | 0.0766 | 0.14832 | 0.026042 | 1.8405 | 1.8405 | 0.003328
2019-03-26 16:19 | 1553631541.6667 | 362 | 0.20446 | 0.003935 | 0.000114 | 96.4704 | 95.655 | 117.6 | 21.2612 | 0.00293 | 0.002758 | 247.4625 | 0.24986 | 8.4022 | 0.003621 | 0.00534 | 0.005259 | 4.6e-05 | 0.001032 | 0.31885 | 1.3493 | 2.5473 | 100.79 | 21.17 | 15.371 | 92.3024 | 109.0397 | 88.1574 | 0.003837 | 0.97713 | 0.023562 | 24.1136 | 0.075925 | 0.14864 | 0.025909 | 1.8319 | 1.8319 | 0.003351
2019-03-26 16:19 | 1553631542.6677 | 363 | 0.20452 | 0.004042 | 0.000114 | 96.594 | 95.6933 | 117.6134 | 21.2612 | 0.002914 | 0.002749 | 247.3609 | 0.24971 | 8.4031 | 0.003712 | 0.004656 | 0.004579 | 4.7e-05 | 0.001034 | 0.31741 | 1.3715 | 2.5475 | 100.79 | 21.17 | 15.38 | 91.924 | 111.2427 | 87.9613 | 0.003839 | 0.97766 | 0.02357 | 24.1084 | 0.083803 | 0.14855 | 0.025433 | 1.7989 | 1.7989 | 0.003411
2019-03-26 16:19 | 1553631543.6664 | 364 | 0.20455 | 0.004127 | 0.000114 | 96.6342 | 95.8224 | 117.7347 | 21.2612 | 0.002974 | 0.002707 | 247.3126 | 0.25043 | 8.4042 | 0.0037 | 0.005166 | 0.005119 | 4.7e-05 | 0.001033 | 0.31885 | 1.3779 | 2.5476 | 100.79 | 21.17 | 15.386 | 93.3803 | 110.9498 | 86.9485 | 0.003841 | 0.98041 | 0.023615 | 24.087 | 0.075461 | 0.14898 | 0.025135 | 1.7812 | 1.7812 | 0.003454
2019-03-26 16:19 | 1553631544.6678 | 365 | 0.20463 | 0.004188 | 0.000113 | 96.6867 | 95.8365 | 117.877 | 21.2612 | 0.003025 | 0.002746 | 247.2089 | 0.25007 | 8.4047 | 0.003659 | 0.005555 | 0.005557 | 4.8e-05 | 0.001029 | 0.34794 | 1.382 | 2.5476 | 100.79 | 21.17 | 15.4 | 94.5964 | 109.9471 | 87.8746 | 0.003844 | 0.9779 | 0.023599 | 24.1325 | 0.079006 | 0.14877 | 0.024623 | 1.7437 | 1.7437 | 0.003524
2019-03-26 16:19 | 1553631545.6664 | 366 | 0.20468 | 0.004239 | 0.000113 | 96.7815 | 96.0815 | 117.9082 | 21.2612 | 0.002974 | 0.002765 | 247.1257 | 0.24948 | 8.4033 | 0.00363 | 0.004596 | 0.004553 | 4.6e-05 | 0.001024 | 0.31581 | 1.3779 | 2.5476 | 100.79 | 21.17 | 15.408 | 93.3827 | 109.2533 | 88.3426 | 0.003846 | 0.97663 | 0.023587 | 24.1513 | 0.064925 | 0.14841 | 0.024256 | 1.7168 | 1.7168 | 0.00357
2019-03-26 16:19 | 1553631546.6666 | 367 | 0.20472 | 0.004263 | 0.000113 | 96.8435 | 96.098 | 118.0021 | 21.2612 | 0.002933 | 0.002704 | 247.0401 | 0.24989 | 8.4038 | 0.003594 | 0.005568 | 0.005446 | 4.6e-05 | 0.001029 | 0.29822 | 1.3788 | 2.5476 | 100.79 | 21.16 | 15.428 | 92.3839 | 108.3659 | 86.8573 | 0.003849 | 0.98066 | 0.023645 | 24.1113 | 0.069119 | 0.14866 | 0.023943 | 1.6988 | 1.6988 | 0.003614
2019-03-26 16:19 | 1553631547.666 | 368 | 0.20473 | 0.00428 | 0.000113 | 96.8246 | 96.1747 | 117.7847 | 21.2612 | 0.002971 | 0.002796 | 246.9802 | 0.25022 | 8.4011 | 0.003458 | 0.004355 | 0.004316 | 4.5e-05 | 0.001033 | 0.33035 | 1.383 | 2.5474 | 100.79 | 21.16 | 15.441 | 93.2928 | 105.0719 | 89.0846 | 0.003852 | 0.97463 | 0.02355 | 24.1635 | 0.060242 | 0.14886 | 0.023871 | 1.6869 | 1.6869 | 0.003645
2019-03-26 16:19 | 1553631548.6666 | 369 | 0.20472 | 0.00429 | 0.000113 | 96.9232 | 96.3357 | 117.481 | 21.2612 | 0.002941 | 0.002745 | 246.8927 | 0.24943 | 8.4011 | 0.003558 | 0.004609 | 0.004568 | 4.3e-05 | 0.001036 | 0.33052 | 1.3509 | 2.5475 | 100.79 | 21.16 | 15.454 | 92.5731 | 107.4906 | 87.8659 | 0.003855 | 0.97792 | 0.02356 | 24.0916 | 0.054383 | 0.14838 | 0.023958 | 1.6937 | 1.6937 | 0.003618
2019-03-26 16:19 | 1553631549.6661 | 370 | 0.2047 | 0.004284 | 0.000113 | 96.9725 | 96.4022 | 117.0432 | 21.2612 | 0.002961 | 0.002729 | 246.7984 | 0.25036 | 8.3983 | 0.003524 | 0.005435 | 0.005286 | 4.5e-05 | 0.001035 | 0.31581 | 1.3667 | 2.5474 | 100.79 | 21.16 | 15.485 | 93.0627 | 106.6797 | 87.4756 | 0.003863 | 0.97898 | 0.023529 | 24.0337 | 0.052753 | 0.14894 | 0.02404 | 1.6972 | 1.6972 | 0.003624
2019-03-26 16:19 | 1553631550.6672 | 371 | 0.2047 | 0.004285 | 0.000113 | 97.046 | 96.376 | 117.0672 | 21.2612 | 0.003002 | 0.002796 | 246.709 | 0.2494 | 8.4055 | 0.003549 | 0.005833 | 0.005793 | 4.4e-05 | 0.001036 | 0.31437 | 1.3662 | 2.5475 | 100.79 | 21.16 | 15.502 | 94.0541 | 107.2874 | 89.0752 | 0.003867 | 0.97466 | 0.023479 | 24.0894 | 0.061963 | 0.14837 | 0.024035 | 1.6933 | 1.6933 | 0.003619
2019-03-26 16:19 | 1553631551.6659 | 372 | 0.20462 | 0.004292 | 0.000113 | 97.1868 | 96.5071 | 116.839 | 21.2612 | 0.003008 | 0.002786 | 246.64 | 0.24921 | 8.4026 | 0.003515 | 0.005222 | 0.005152 | 4.6e-05 | 0.001037 | 0.31188 | 1.3915 | 2.5475 | 100.79 | 21.16 | 15.518 | 94.1972 | 106.4662 | 88.833 | 0.003871 | 0.97531 | 0.023464 | 24.0579 | 0.062772 | 0.14825 | 0.024519 | 1.7264 | 1.7264 | 0.003547
2019-03-26 16:19 | 1553631552.6662 | 373 | 0.20453 | 0.0043 | 0.000113 | 97.2387 | 96.6547 | 116.5266 | 21.2612 | 0.003054 | 0.002814 | 246.5776 | 0.2501 | 8.4011 | 0.003558 | 0.004547 | 0.004481 | 4.3e-05 | 0.001031 | 0.33892 | 1.3711 | 2.5474 | 100.79 | 21.16 | 15.547 | 95.2908 | 107.508 | 89.5158 | 0.003879 | 0.97347 | 0.02341 | 24.0484 | 0.053885 | 0.14878 | 0.025059 | 1.7604 | 1.7604 | 0.003491
2019-03-26 16:19 | 1553631553.6663 | 374 | 0.2045 | 0.004308 | 0.000115 | 97.2517 | 96.6833 | 115.9663 | 21.2612 | 0.002988 | 0.00285 | 246.4595 | 0.24944 | 8.4031 | 0.003483 | 0.004402 | 0.004363 | 4.5e-05 | 0.001032 | 0.36497 | 1.3372 | 2.5476 | 100.79 | 21.15 | 15.559 | 93.7073 | 105.6723 | 90.388 | 0.003879 | 0.97114 | 0.023326 | 24.0193 | 0.052436 | 0.14839 | 0.025261 | 1.7681 | 1.7681 | 0.003466
2019-03-26 16:19 | 1553631554.666 | 375 | 0.20434 | 0.004322 | 0.000117 | 97.3789 | 96.5902 | 115.4844 | 21.2612 | 0.002989 | 0.002842 | 246.3679 | 0.24972 | 8.4014 | 0.003502 | 0.005315 | 0.005254 | 4.6e-05 | 0.001034 | 0.32621 | 1.3801 | 2.5476 | 100.79 | 21.15 | 15.582 | 93.7348 | 106.1453 | 90.201 | 0.003885 | 0.97164 | 0.023284 | 23.9632 | 0.072748 | 0.14856 | 0.026181 | 1.8291 | 1.8291 | 0.003354
2019-03-26 16:19 | 1553631555.6664 | 376 | 0.20411 | 0.004345 | 0.00012 | 97.4061 | 96.7165 | 115.2091 | 21.2612 | 0.00304 | 0.002886 | 246.2362 | 0.24987 | 8.4004 | 0.003534 | 0.004837 | 0.004737 | 4.8e-05 | 0.001039 | 0.31868 | 1.3578 | 2.5475 | 100.79 | 21.15 | 15.592 | 94.9651 | 106.9267 | 91.2601 | 0.003887 | 0.96881 | 0.023222 | 23.9695 | 0.063545 | 0.14865 | 0.027506 | 1.9165 | 1.9165 | 0.003203
2019-03-26 16:19 | 1553631556.679 | 377 | 0.20391 | 0.004379 | 0.000123 | 97.4291 | 96.3585 | 114.7092 | 21.2612 | 0.003099 | 0.002824 | 246.1311 | 0.24905 | 8.4047 | 0.003607 | 0.004562 | 0.004508 | 4.6e-05 | 0.001039 | 0.31957 | 1.3276 | 2.5475 | 100.79 | 21.16 | 15.603 | 96.3778 | 108.696 | 89.7564 | 0.003893 | 0.97283 | 0.023219 | 23.868 | 0.09883 | 0.14816 | 0.028672 | 1.9973 | 1.9973 | 0.003064
2019-03-26 16:19 | 1553631557.674 | 378 | 0.2036 | 0.004434 | 0.00013 | 97.5413 | 96.5989 | 114.1153 | 21.2612 | 0.003085 | 0.002915 | 246.0132 | 0.24917 | 8.3993 | 0.00356 | 0.005192 | 0.005093 | 4.7e-05 | 0.001041 | 0.32388 | 1.3612 | 2.5474 | 100.79 | 21.16 | 15.621 | 96.0403 | 107.5449 | 91.9553 | 0.003897 | 0.96697 | 0.023089 | 23.8781 | 0.086839 | 0.14823 | 0.030441 | 2.1083 | 2.1083 | 0.002904
2019-03-26 16:19 | 1553631558.6737 | 379 | 0.20313 | 0.004507 | 0.000136 | 97.5915 | 96.6584 | 113.7304 | 21.2612 | 0.003049 | 0.002986 | 245.9147 | 0.2494 | 8.4059 | 0.00365 | 0.004731 | 0.004697 | 4.5e-05 | 0.001036 | 0.30254 | 1.3645 | 2.5475 | 100.79 | 21.15 | 15.629 | 95.1844 | 109.7281 | 93.6651 | 0.003897 | 0.96246 | 0.022997 | 23.8936 | 0.085932 | 0.14837 | 0.033176 | 2.2882 | 2.2882 | 0.002678
2019-03-26 16:19 | 1553631559.6751 | 380 | 0.20281 | 0.0046 | 0.000144 | 97.6653 | 96.4372 | 113.2381 | 21.2612 | 0.003137 | 0.002959 | 245.7922 | 0.24969 | 8.4036 | 0.00368 | 0.004688 | 0.004635 | 4.7e-05 | 0.001038 | 0.29502 | 1.3274 | 2.5474 | 100.79 | 21.15 | 15.642 | 97.3038 | 110.4742 | 93.0087 | 0.0039 | 0.96419 | 0.022967 | 23.8205 | 0.11318 | 0.14854 | 0.034965 | 2.408 | 2.408 | 0.002548
2019-03-26 16:19 | 1553631560.6737 | 381 | 0.2024 | 0.00472 | 0.000152 | 97.6948 | 96.5643 | 112.4875 | 21.2612 | 0.003121 | 0.002939 | 245.6618 | 0.25017 | 8.4047 | 0.003728 | 0.005607 | 0.005494 | 4.7e-05 | 0.001041 | 0.32228 | 1.3804 | 2.5476 | 100.79 | 21.16 | 15.647 | 96.9128 | 111.6278 | 92.5299 | 0.003904 | 0.96545 | 0.022906 | 23.7259 | 0.10411 | 0.14883 | 0.037274 | 2.5596 | 2.5596 | 0.002401
2019-03-26 16:19 | 1553631561.6743 | 382 | 0.20218 | 0.004858 | 0.00016 | 97.8179 | 96.0825 | 111.927 | 21.2612 | 0.003152 | 0.003022 | 245.4991 | 0.24963 | 8.4043 | 0.00377 | 0.005381 | 0.005311 | 4.6e-05 | 0.001039 | 0.31686 | 1.373 | 2.5475 | 100.79 | 21.16 | 15.651 | 97.6683 | 112.6631 | 94.5284 | 0.003905 | 0.9602 | 0.022787 | 23.7313 | 0.1601 | 0.1485 | 0.038411 | 2.6234 | 2.6234 | 0.002338
2019-03-26 16:19 | 1553631562.6738 | 383 | 0.20179 | 0.005031 | 0.000169 | 97.856 | 95.9568 | 111.4895 | 21.2612 | 0.003186 | 0.003073 | 245.3484 | 0.24913 | 8.4014 | 0.003734 | 0.005395 | 0.005318 | 4.9e-05 | 0.001045 | 0.31166 | 1.3417 | 2.5474 | 100.79 | 21.15 | 15.656 | 98.4999 | 111.7723 | 95.7598 | 0.003903 | 0.957 | 0.022704 | 23.7245 | 0.1753 | 0.14821 | 0.040501 | 2.7556 | 2.7556 | 0.002221
2019-03-26 16:19 | 1553631563.674 | 384 | 0.2015 | 0.005223 | 0.000176 | 97.8991 | 95.7326 | 111.3977 | 21.2612 | 0.003209 | 0.003027 | 245.2253 | 0.25039 | 8.4024 | 0.003729 | 0.005053 | 0.00499 | 4.4e-05 | 0.001046 | 0.32029 | 1.3923 | 2.5475 | 100.79 | 21.15 | 15.658 | 99.0568 | 111.657 | 94.6599 | 0.003904 | 0.95986 | 0.022729 | 23.6793 | 0.20016 | 0.14896 | 0.041963 | 2.8577 | 2.8577 | 0.002153
2019-03-26 16:19 | 1553631564.6734 | 385 | 0.20114 | 0.005429 | 0.000182 | 98.0077 | 95.8012 | 111.1538 | 21.2612 | 0.003275 | 0.003042 | 245.0903 | 0.24897 | 8.4038 | 0.003761 | 0.004937 | 0.004774 | 4.6e-05 | 0.001046 | 0.32444 | 1.3931 | 2.5475 | 100.79 | 21.15 | 15.661 | 100.6438 | 112.426 | 95.009 | 0.003905 | 0.95895 | 0.022693 | 23.6646 | 0.20368 | 0.14811 | 0.043886 | 2.9835 | 2.9835 | 0.00205
2019-03-26 16:19 | 1553631565.6741 | 386 | 0.20078 | 0.005652 | 0.000189 | 98.0292 | 96.0015 | 110.8041 | 21.2612 | 0.003282 | 0.003113 | 244.9639 | 0.24927 | 8.404 | 0.003722 | 0.004984 | 0.005008 | 4.7e-05 | 0.00105 | 0.32853 | 1.3679 | 2.5477 | 100.79 | 21.15 | 15.664 | 100.7986 | 111.4818 | 96.7341 | 0.003905 | 0.95448 | 0.022605 | 23.6826 | 0.18696 | 0.14829 | 0.045763 | 3.0984 | 3.0984 | 0.001977
2019-03-26 16:19 | 1553631566.6731 | 387 | 0.20061 | 0.005879 | 0.000194 | 98.1119 | 96.0903 | 110.6611 | 21.2612 | 0.003317 | 0.003142 | 244.792 | 0.25004 | 8.4024 | 0.003694 | 0.006121 | 0.005953 | 5e-05 | 0.001051 | 0.31349 | 1.34 | 2.5475 | 100.79 | 21.15 | 15.668 | 101.6486 | 110.8071 | 97.4222 | 0.003906 | 0.9527 | 0.022569 | 23.6894 | 0.18623 | 0.14875 | 0.046472 | 3.141 | 3.141 | 0.001956
2019-03-26 16:19 | 1553631567.6735 | 388 | 0.20036 | 0.006112 | 0.000197 | 98.204 | 95.7373 | 110.3086 | 21.2612 | 0.00335 | 0.00315 | 244.65 | 0.24918 | 8.4037 | 0.003765 | 0.005193 | 0.005104 | 4.6e-05 | 0.001049 | 0.30254 | 1.4102 | 2.5476 | 100.79 | 21.15 | 15.677 | 102.4615 | 112.5287 | 97.6096 | 0.003909 | 0.95222 | 0.022527 | 23.6576 | 0.22754 | 0.14824 | 0.047654 | 3.2147 | 3.2147 | 0.001904
2019-03-26 16:19 | 1553631568.6732 | 389 | 0.20022 | 0.00634 | 0.0002 | 98.253 | 96.0304 | 109.9713 | 21.2612 | 0.003418 | 0.003178 | 244.5004 | 0.24904 | 8.4013 | 0.003779 | 0.004884 | 0.004862 | 4.8e-05 | 0.001055 | 0.32853 | 1.3973 | 2.5476 | 100.79 | 21.15 | 15.683 | 104.0969 | 112.8708 | 98.3016 | 0.00391 | 0.95045 | 0.022472 | 23.6434 | 0.20466 | 0.14815 | 0.048261 | 3.2474 | 3.2474 | 0.001884
2019-03-26 16:19 | 1553631569.6743 | 390 | 0.19987 | 0.006558 | 0.000201 | 98.2719 | 95.9693 | 109.8886 | 21.2612 | 0.00341 | 0.003205 | 244.3536 | 0.24927 | 8.406 | 0.003745 | 0.005098 | 0.004955 | 4.6e-05 | 0.001055 | 0.31758 | 1.3719 | 2.5476 | 100.79 | 21.15 | 15.691 | 103.9207 | 112.0383 | 98.9516 | 0.003912 | 0.94879 | 0.022444 | 23.6552 | 0.21208 | 0.14829 | 0.050043 | 3.3628 | 3.3628 | 0.001821
2019-03-26 16:19 | 1553631570.6747 | 391 | 0.19943 | 0.00678 | 0.000203 | 98.4091 | 96.2496 | 109.3967 | 21.2612 | 0.003471 | 0.003279 | 244.2268 | 0.24924 | 8.4037 | 0.003853 | 0.005484 | 0.00541 | 4.9e-05 | 0.001056 | 0.28711 | 1.3833 | 2.5476 | 100.79 | 21.15 | 15.707 | 105.3828 | 114.6649 | 100.7294 | 0.003916 | 0.94428 | 0.02234 | 23.6585 | 0.19847 | 0.14827 | 0.052394 | 3.5041 | 3.5041 | 0.001748
2019-03-26 16:20 | 1553631571.6737 | 392 | 0.19914 | 0.007009 | 0.000203 | 98.4661 | 96.2963 | 108.9664 | 21.2612 | 0.003469 | 0.003249 | 244.0904 | 0.24957 | 8.4029 | 0.00392 | 0.005198 | 0.005174 | 4.8e-05 | 0.001059 | 0.31188 | 1.3561 | 2.5475 | 100.79 | 21.14 | 15.715 | 105.3277 | 116.2979 | 100.0019 | 0.003916 | 0.94612 | 0.022318 | 23.589 | 0.1993 | 0.14847 | 0.053873 | 3.5992 | 3.5992 | 0.001704
2019-03-26 16:20 | 1553631572.6744 | 393 | 0.19892 | 0.007243 | 0.000203 | 98.495 | 96.0624 | 108.7714 | 21.2612 | 0.003547 | 0.003274 | 243.9372 | 0.24849 | 8.4036 | 0.003916 | 0.005439 | 0.005319 | 4.6e-05 | 0.001056 | 0.31669 | 1.3292 | 2.5474 | 100.79 | 21.14 | 15.732 | 107.2323 | 116.2179 | 100.6133 | 0.00392 | 0.94457 | 0.02228 | 23.5871 | 0.22369 | 0.14783 | 0.054891 | 3.6607 | 3.6607 | 0.001668
2019-03-26 16:20 | 1553631573.6735 | 394 | 0.19869 | 0.007472 | 0.000204 | 98.5753 | 96.2762 | 108.5387 | 21.2612 | 0.003516 | 0.003339 | 243.7695 | 0.24877 | 8.4069 | 0.003904 | 0.005462 | 0.005386 | 4.7e-05 | 0.001058 | 0.32244 | 1.3515 | 2.5475 | 100.79 | 21.14 | 15.74 | 106.4837 | 115.9242 | 102.192 | 0.003922 | 0.9406 | 0.022209 | 23.6116 | 0.21109 | 0.148 | 0.056006 | 3.723 | 3.723 | 0.001642
2019-03-26 16:20 | 1553631574.6746 | 395 | 0.19852 | 0.007696 | 0.000206 | 98.5969 | 96.3202 | 108.3831 | 21.2612 | 0.003525 | 0.003334 | 243.6343 | 0.24959 | 8.4008 | 0.003952 | 0.005332 | 0.005183 | 4.6e-05 | 0.001061 | 0.33284 | 1.3658 | 2.5475 | 100.79 | 21.14 | 15.748 | 106.697 | 117.0849 | 102.0775 | 0.003924 | 0.94088 | 0.022196 | 23.5911 | 0.20896 | 0.14848 | 0.056737 | 3.7692 | 3.7692 | 0.001627
2019-03-26 16:20 | 1553631575.6735 | 396 | 0.19851 | 0.007911 | 0.00021 | 98.7052 | 96.3612 | 108.2085 | 21.2612 | 0.003591 | 0.003358 | 243.4782 | 0.24892 | 8.4019 | 0.004036 | 0.005386 | 0.005294 | 4.8e-05 | 0.001066 | 0.30093 | 1.3853 | 2.5477 | 100.79 | 21.14 | 15.763 | 108.3027 | 119.1345 | 102.6382 | 0.003928 | 0.93948 | 0.022162 | 23.5897 | 0.21497 | 0.14808 | 0.056534 | 3.7496 | 3.7496 | 0.001631
2019-03-26 16:20 | 1553631576.674 | 397 | 0.19848 | 0.008108 | 0.000213 | 98.7179 | 96.5491 | 107.8754 | 21.2612 | 0.003589 | 0.003402 | 243.3297 | 0.24937 | 8.4046 | 0.004007 | 0.005055 | 0.005004 | 4.5e-05 | 0.001058 | 0.31045 | 1.351 | 2.5474 | 100.79 | 21.14 | 15.769 | 108.2571 | 118.4385 | 103.7032 | 0.003929 | 0.93683 | 0.022097 | 23.5867 | 0.1987 | 0.14835 | 0.05647 | 3.7342 | 3.7342 | 0.001641
2019-03-26 16:20 | 1553631577.6727 | 398 | 0.19844 | 0.00829 | 0.000217 | 98.7696 | 96.3713 | 107.7055 | 21.2612 | 0.003606 | 0.003449 | 243.1611 | 0.24907 | 8.4033 | 0.004008 | 0.005209 | 0.005188 | 4.5e-05 | 0.001052 | 0.33157 | 1.3311 | 2.5476 | 100.79 | 21.14 | 15.774 | 108.6527 | 118.4529 | 104.8462 | 0.00393 | 0.93399 | 0.022046 | 23.6038 | 0.21987 | 0.14817 | 0.056474 | 3.7256 | 3.7256 | 0.001643
2019-03-26 16:20 | 1553631578.6741 | 399 | 0.19826 | 0.008448 | 0.000218 | 98.7707 | 96.218 | 107.2593 | 21.2612 | 0.003679 | 0.003485 | 243.0003 | 0.24898 | 8.4031 | 0.004068 | 0.005649 | 0.005555 | 4.3e-05 | 0.001055 | 0.30829 | 1.3565 | 2.5474 | 100.79 | 21.14 | 15.783 | 110.4496 | 119.9316 | 105.7281 | 0.003933 | 0.93182 | 0.021974 | 23.5823 | 0.23421 | 0.14812 | 0.057339 | 3.7703 | 3.7703 | 0.001622
2019-03-26 16:20 | 1553631579.6738 | 400 | 0.19798 | 0.00859 | 0.000219 | 98.8502 | 96.2728 | 106.7469 | 21.2612 | 0.003689 | 0.003499 | 242.8507 | 0.24957 | 8.4013 | 0.004022 | 0.005779 | 0.005688 | 4.6e-05 | 0.00105 | 0.30038 | 1.3374 | 2.5475 | 100.79 | 21.14 | 15.786 | 110.6784 | 118.7944 | 106.0608 | 0.003933 | 0.931 | 0.021912 | 23.5363 | 0.23631 | 0.14847 | 0.058835 | 3.8573 | 3.8573 | 0.00159
2019-03-26 16:20 | 1553631580.6742 | 401 | 0.19771 | 0.008719 | 0.00022 | 98.9963 | 96.2597 | 106.4819 | 21.2612 | 0.003745 | 0.003498 | 242.6689 | 0.24922 | 8.4024 | 0.004162 | 0.005284 | 0.005225 | 4.5e-05 | 0.001055 | 0.31797 | 1.3491 | 2.5476 | 100.79 | 21.14 | 15.792 | 112.0471 | 122.2285 | 106.0397 | 0.003935 | 0.93105 | 0.021886 | 23.5064 | 0.25074 | 0.14826 | 0.060321 | 3.9496 | 3.9496 | 0.00155
2019-03-26 16:20 | 1553631581.6739 | 402 | 0.19746 | 0.008848 | 0.000219 | 99.1586 | 96.5327 | 106.4819 | 21.2612 | 0.00376 | 0.00354 | 242.5219 | 0.24887 | 8.4022 | 0.004181 | 0.004937 | 0.004849 | 4.6e-05 | 0.001055 | 0.30071 | 1.3404 | 2.5475 | 100.79 | 21.14 | 15.794 | 112.4197 | 122.6917 | 107.0709 | 0.003935 | 0.92853 | 0.021856 | 23.5383 | 0.24007 | 0.14805 | 0.061613 | 4.0285 | 4.0285 | 0.001518
2019-03-26 16:20 | 1553631582.6729 | 403 | 0.19722 | 0.00898 | 0.000217 | 99.0996 | 96.6792 | 106.418 | 21.2612 | 0.003819 | 0.003533 | 242.3391 | 0.24868 | 8.4035 | 0.004132 | 0.005615 | 0.005534 | 4.3e-05 | 0.001058 | 0.31725 | 1.3378 | 2.5474 | 100.79 | 21.14 | 15.796 | 113.8364 | 121.494 | 106.8843 | 0.003936 | 0.92898 | 0.021855 | 23.5255 | 0.22117 | 0.14794 | 0.062906 | 4.1126 | 4.1126 | 0.001486
2019-03-26 16:20 | 1553631583.673 | 404 | 0.19682 | 0.009123 | 0.000214 | 99.1669 | 96.6742 | 106.0439 | 21.2612 | 0.003783 | 0.003599 | 242.1693 | 0.24954 | 8.4027 | 0.00424 | 0.006038 | 0.005966 | 4.6e-05 | 0.001056 | 0.28047 | 1.3632 | 2.5476 | 100.79 | 21.14 | 15.8 | 112.9592 | 124.1199 | 108.4912 | 0.003937 | 0.92507 | 0.02177 | 23.5337 | 0.22771 | 0.14845 | 0.065087 | 4.2384 | 4.2384 | 0.001447
2019-03-26 16:20 | 1553631584.6732 | 405 | 0.19656 | 0.009275 | 0.000211 | 99.1288 | 96.806 | 105.6197 | 21.2612 | 0.003873 | 0.003617 | 241.9875 | 0.24963 | 8.4031 | 0.004236 | 0.005589 | 0.005521 | 4.5e-05 | 0.001059 | 0.3235 | 1.3544 | 2.5475 | 100.79 | 21.14 | 15.802 | 115.1678 | 124.0229 | 108.9318 | 0.003937 | 0.924 | 0.021714 | 23.5001 | 0.21209 | 0.14851 | 0.06649 | 4.3183 | 4.3183 | 0.00142
2019-03-26 16:20 | 1553631585.6741 | 406 | 0.19632 | 0.00944 | 0.000209 | 99.1778 | 96.8655 | 105.1867 | 21.2612 | 0.003921 | 0.003631 | 241.8315 | 0.24951 | 8.4023 | 0.004274 | 0.005977 | 0.00593 | 4.1e-05 | 0.001052 | 0.3178 | 1.3319 | 2.5474 | 100.79 | 21.14 | 15.806 | 116.3199 | 124.9717 | 109.2652 | 0.003938 | 0.9232 | 0.02166 | 23.4621 | 0.21101 | 0.14843 | 0.0677 | 4.3857 | 4.3857 | 0.001398
2019-03-26 16:20 | 1553631586.6731 | 407 | 0.19606 | 0.009621 | 0.000207 | 99.3148 | 96.9146 | 104.7865 | 21.2612 | 0.003937 | 0.003716 | 241.6753 | 0.24922 | 8.4019 | 0.004282 | 0.006201 | 0.006109 | 4.4e-05 | 0.001053 | 0.32333 | 1.3663 | 2.5476 | 100.79 | 21.14 | 15.809 | 116.7222 | 125.1618 | 111.3295 | 0.003939 | 0.91824 | 0.021561 | 23.4806 | 0.21882 | 0.14826 | 0.069027 | 4.4509 | 4.4509 | 0.001376
2019-03-26 16:20 | 1553631587.6731 | 408 | 0.19597 | 0.00981 | 0.000206 | 99.3756 | 97.1304 | 104.5196 | 21.2612 | 0.003979 | 0.003716 | 241.4928 | 0.25002 | 8.4056 | 0.004345 | 0.005407 | 0.005364 | 4.2e-05 | 0.001051 | 0.27616 | 1.3808 | 2.5474 | 100.79 | 21.14 | 15.81 | 117.7462 | 126.6987 | 111.3342 | 0.003939 | 0.91823 | 0.021533 | 23.4508 | 0.2044 | 0.14874 | 0.069351 | 4.466 | 4.466 | 0.001375
2019-03-26 16:20 | 1553631588.6732 | 409 | 0.1959 | 0.010009 | 0.000207 | 99.3449 | 96.9209 | 104.3745 | 21.2612 | 0.003989 | 0.003766 | 241.2981 | 0.24976 | 8.403 | 0.004376 | 0.005566 | 0.005404 | 4.5e-05 | 0.00105 | 0.31006 | 1.3508 | 2.5474 | 100.79 | 21.14 | 15.812 | 117.9964 | 127.477 | 112.5618 | 0.00394 | 0.91531 | 0.021484 | 23.4719 | 0.22095 | 0.14858 | 0.069565 | 4.4694 | 4.4694 | 0.001373
2019-03-26 16:20 | 1553631589.6732 | 410 | 0.19592 | 0.010186 | 0.000208 | 99.4435 | 97.0295 | 104.4067 | 21.2612 | 0.004004 | 0.003761 | 241.1145 | 0.24953 | 8.3996 | 0.004474 | 0.006227 | 0.006227 | 4.4e-05 | 0.00105 | 0.32659 | 1.3576 | 2.5475 | 100.79 | 21.14 | 15.813 | 118.3545 | 129.8869 | 112.4387 | 0.00394 | 0.9156 | 0.021491 | 23.4718 | 0.21981 | 0.14844 | 0.069207 | 4.4479 | 4.4479 | 0.001378
2019-03-26 16:20 | 1553631590.6746 | 411 | 0.19591 | 0.010361 | 0.000208 | 99.4677 | 96.9821 | 104.6329 | 21.2612 | 0.004075 | 0.003807 | 240.9462 | 0.24892 | 8.4028 | 0.004505 | 0.005503 | 0.005353 | 4.1e-05 | 0.001055 | 0.35297 | 1.3537 | 2.5476 | 100.79 | 21.14 | 15.811 | 120.0861 | 130.637 | 113.5534 | 0.00394 | 0.91296 | 0.021483 | 23.5311 | 0.22635 | 0.14808 | 0.069034 | 4.4352 | 4.4352 | 0.001379
2019-03-26 16:20 | 1553631591.6746 | 412 | 0.19585 | 0.010506 | 0.000208 | 99.4854 | 97.0087 | 104.5428 | 21.2612 | 0.004055 | 0.003797 | 240.7449 | 0.24905 | 8.4053 | 0.004557 | 0.006438 | 0.006328 | 4.1e-05 | 0.001046 | 0.31454 | 1.3477 | 2.5475 | 100.79 | 21.14 | 15.808 | 119.6126 | 131.9297 | 113.2974 | 0.003939 | 0.91357 | 0.021481 | 23.5132 | 0.2255 | 0.14816 | 0.069223 | 4.4468 | 4.4468 | 0.001376
2019-03-26 16:20 | 1553631592.6744 | 413 | 0.19562 | 0.010639 | 0.000208 | 99.5306 | 97.0406 | 104.1842 | 21.2612 | 0.00413 | 0.003837 | 240.5704 | 0.24853 | 8.4018 | 0.00452 | 0.005893 | 0.005868 | 3.7e-05 | 0.001049 | 0.33411 | 1.342 | 2.5474 | 100.79 | 21.14 | 15.803 | 121.4441 | 131.0051 | 114.2813 | 0.003938 | 0.91125 | 0.021417 | 23.5027 | 0.22661 | 0.14785 | 0.070437 | 4.5109 | 4.5109 | 0.001354
2019-03-26 16:20 | 1553631593.6724 | 414 | 0.19539 | 0.010765 | 0.000208 | 99.592 | 97.0846 | 104.044 | 21.2612 | 0.004105 | 0.003829 | 240.4183 | 0.24887 | 8.4016 | 0.004448 | 0.005629 | 0.005516 | 3.9e-05 | 0.001045 | 0.30541 | 1.3426 | 2.5475 | 100.79 | 21.14 | 15.791 | 120.8288 | 129.2441 | 114.0939 | 0.003935 | 0.91169 | 0.021408 | 23.4812 | 0.22807 | 0.14805 | 0.071661 | 4.587 | 4.587 | 0.001333
2019-03-26 16:20 | 1553631594.6729 | 415 | 0.19539 | 0.010898 | 0.000208 | 99.5642 | 97.2047 | 104.0321 | 21.2612 | 0.004147 | 0.003903 | 240.2359 | 0.2491 | 8.4024 | 0.004451 | 0.005777 | 0.005717 | 4.3e-05 | 0.00105 | 0.31708 | 1.3528 | 2.5477 | 100.79 | 21.14 | 15.785 | 121.8506 | 129.3055 | 115.8879 | 0.003933 | 0.90748 | 0.021357 | 23.5342 | 0.21453 | 0.14819 | 0.071501 | 4.566 | 4.566 | 0.00134
2019-03-26 16:20 | 1553631595.6735 | 416 | 0.19555 | 0.011023 | 0.000208 | 99.6345 | 97.1593 | 104.0417 | 21.2612 | 0.004175 | 0.003883 | 240.0404 | 0.24974 | 8.4033 | 0.004486 | 0.006103 | 0.006005 | 4.1e-05 | 0.001049 | 0.30901 | 1.333 | 2.5476 | 100.79 | 21.13 | 15.779 | 122.5311 | 130.1773 | 115.4162 | 0.003929 | 0.90858 | 0.021371 | 23.521 | 0.22501 | 0.14857 | 0.070411 | 4.4997 | 4.4997 | 0.001364
2019-03-26 16:20 | 1553631596.6735 | 417 | 0.1957 | 0.011128 | 0.00021 | 99.6353 | 97.2151 | 104.0527 | 21.2612 | 0.004267 | 0.003909 | 239.8632 | 0.24869 | 8.4018 | 0.004473 | 0.005861 | 0.00581 | 4.2e-05 | 0.001051 | 0.32781 | 1.3485 | 2.5474 | 100.79 | 21.14 | 15.773 | 124.7888 | 129.8471 | 116.0303 | 0.00393 | 0.90715 | 0.021355 | 23.5408 | 0.21996 | 0.14794 | 0.069371 | 4.4302 | 4.4302 | 0.001379
2019-03-26 16:20 | 1553631597.6737 | 418 | 0.19572 | 0.011206 | 0.000211 | 99.6374 | 97.3442 | 104.1393 | 21.2612 | 0.004248 | 0.003907 | 239.7088 | 0.24768 | 8.4036 | 0.004409 | 0.005806 | 0.005677 | 3.8e-05 | 0.001048 | 0.31725 | 1.3654 | 2.5476 | 100.79 | 21.13 | 15.758 | 124.3321 | 128.2867 | 115.9876 | 0.003924 | 0.90725 | 0.021365 | 23.5493 | 0.20827 | 0.14734 | 0.069185 | 4.4204 | 4.4204 | 0.001377
2019-03-26 16:20 | 1553631598.6737 | 419 | 0.19571 | 0.011263 | 0.000215 | 99.7 | 97.4487 | 104.4703 | 21.2612 | 0.004268 | 0.003982 | 239.5408 | 0.24835 | 8.4024 | 0.004513 | 0.00567 | 0.005622 | 4e-05 | 0.001053 | 0.32278 | 1.3487 | 2.5474 | 100.79 | 21.13 | 15.751 | 124.8268 | 130.8469 | 117.8206 | 0.003922 | 0.903 | 0.021349 | 23.6422 | 0.20429 | 0.14774 | 0.06917 | 4.4158 | 4.4158 | 0.001382
2019-03-26 16:20 | 1553631599.6741 | 420 | 0.19565 | 0.011287 | 0.000217 | 99.713 | 97.1589 | 104.3463 | 21.2612 | 0.004347 | 0.003962 | 239.3367 | 0.24851 | 8.4048 | 0.004503 | 0.005927 | 0.005951 | 3.9e-05 | 0.001054 | 0.35491 | 1.3893 | 2.5476 | 100.79 | 21.13 | 15.744 | 126.7513 | 130.5859 | 117.3421 | 0.003921 | 0.9041 | 0.021349 | 23.6137 | 0.23209 | 0.14784 | 0.069523 | 4.4382 | 4.4382 | 0.001376
2019-03-26 16:20 | 1553631600.6732 | 421 | 0.19562 | 0.011307 | 0.000219 | 99.7738 | 97.1963 | 104.0926 | 21.2612 | 0.004317 | 0.003952 | 239.145 | 0.24848 | 8.402 | 0.004451 | 0.005358 | 0.005193 | 3.9e-05 | 0.00106 | 0.31885 | 1.3702 | 2.5474 | 100.79 | 21.13 | 15.732 | 126.0293 | 129.3158 | 117.0983 | 0.003918 | 0.90467 | 0.02133 | 23.5776 | 0.23411 | 0.14782 | 0.069655 | 4.4425 | 4.4425 | 0.001374
2019-03-26 16:20 | 1553631601.6729 | 422 | 0.19557 | 0.01132 | 0.00022 | 99.785 | 97.0389 | 104.0485 | 21.2612 | 0.004333 | 0.003972 | 238.9825 | 0.24854 | 8.4023 | 0.004497 | 0.006035 | 0.005812 | 4e-05 | 0.00106 | 0.32731 | 1.35 | 2.5474 | 100.79 | 21.13 | 15.727 | 126.4146 | 130.4439 | 117.5662 | 0.003916 | 0.90358 | 0.021313 | 23.5867 | 0.2496 | 0.14786 | 0.069911 | 4.4552 | 4.4552 | 0.001371
2019-03-26 16:20 | 1553631602.6739 | 423 | 0.19556 | 0.011335 | 0.00022 | 99.7708 | 96.7532 | 103.9745 | 21.2612 | 0.004348 | 0.004018 | 238.8142 | 0.24825 | 8.4034 | 0.004517 | 0.006278 | 0.006275 | 4.3e-05 | 0.001057 | 0.31564 | 1.3476 | 2.5475 | 100.79 | 21.13 | 15.719 | 126.7704 | 130.9283 | 118.693 | 0.003914 | 0.90099 | 0.021274 | 23.6123 | 0.27471 | 0.14768 | 0.069956 | 4.45 | 4.45 | 0.001371
2019-03-26 16:20 | 1553631603.6731 | 424 | 0.19557 | 0.011353 | 0.000219 | 99.8024 | 96.6977 | 103.8102 | 21.2612 | 0.004394 | 0.003968 | 238.6273 | 0.24837 | 8.4002 | 0.0046 | 0.005868 | 0.005799 | 4e-05 | 0.001056 | 0.33196 | 1.3202 | 2.5474 | 100.79 | 21.13 | 15.716 | 127.9089 | 132.9779 | 117.4815 | 0.003914 | 0.90378 | 0.021291 | 23.5572 | 0.28267 | 0.14776 | 0.069906 | 4.4503 | 4.4503 | 0.001371
2019-03-26 16:20 | 1553631604.6741 | 425 | 0.19566 | 0.011373 | 0.000218 | 99.7915 | 96.7972 | 103.6408 | 21.2612 | 0.004414 | 0.004007 | 238.425 | 0.24855 | 8.4025 | 0.004611 | 0.006015 | 0.005911 | 3.8e-05 | 0.001058 | 0.32278 | 1.3424 | 2.5474 | 100.79 | 21.13 | 15.714 | 128.3943 | 133.2528 | 118.437 | 0.003913 | 0.90158 | 0.021247 | 23.5667 | 0.27249 | 0.14786 | 0.069319 | 4.4041 | 4.4041 | 0.001387
2019-03-26 16:20 | 1553631605.6744 | 426 | 0.1957 | 0.011389 | 0.000217 | 99.8473 | 96.9599 | 103.5319 | 21.2612 | 0.00441 | 0.004086 | 238.25 | 0.24915 | 8.404 | 0.004622 | 0.00697 | 0.006863 | 4e-05 | 0.001062 | 0.32997 | 1.391 | 2.5475 | 100.79 | 21.14 | 15.711 | 128.2978 | 133.5078 | 120.36 | 0.003915 | 0.89717 | 0.021184 | 23.6121 | 0.26247 | 0.14822 | 0.069101 | 4.3774 | 4.3774 | 0.001398
2019-03-26 16:20 | 1553631606.6731 | 427 | 0.19573 | 0.011399 | 0.000217 | 99.8437 | 96.7542 | 103.5004 | 21.2612 | 0.004419 | 0.004078 | 238.0629 | 0.24864 | 8.4045 | 0.00471 | 0.006312 | 0.006239 | 4.1e-05 | 0.001064 | 0.30989 | 1.3766 | 2.5475 | 100.79 | 21.14 | 15.708 | 128.5258 | 135.6919 | 120.1695 | 0.003914 | 0.8976 | 0.021186 | 23.6028 | 0.28115 | 0.14792 | 0.06887 | 4.3632 | 4.3632 | 0.0014
2019-03-26 16:20 | 1553631607.6741 | 428 | 0.19587 | 0.011397 | 0.000217 | 99.8346 | 96.735 | 103.6077 | 21.2612 | 0.004409 | 0.004086 | 237.8596 | 0.24852 | 8.4024 | 0.004644 | 0.00587 | 0.005728 | 4e-05 | 0.001066 | 0.29087 | 1.3729 | 2.5475 | 100.79 | 21.14 | 15.703 | 128.2843 | 134.0706 | 120.359 | 0.003913 | 0.89717 | 0.021192 | 23.6207 | 0.2821 | 0.14785 | 0.068054 | 4.3129 | 4.3129 | 0.001416
2019-03-26 16:20 | 1553631608.6722 | 429 | 0.19586 | 0.011384 | 0.000216 | 99.8228 | 96.8046 | 103.6415 | 21.2612 | 0.004484 | 0.00409 | 237.693 | 0.24846 | 8.4039 | 0.004651 | 0.006102 | 0.006028 | 4.2e-05 | 0.001069 | 0.3314 | 1.3313 | 2.5475 | 100.79 | 21.14 | 15.701 | 130.1136 | 134.2352 | 120.4614 | 0.003912 | 0.89694 | 0.021193 | 23.6276 | 0.27461 | 0.14781 | 0.068123 | 4.3175 | 4.3175 | 0.001414
2019-03-26 16:20 | 1553631609.6726 | 430 | 0.19596 | 0.011371 | 0.000217 | 99.8827 | 96.7962 | 103.9713 | 21.2612 | 0.004481 | 0.004061 | 237.5059 | 0.24839 | 8.4014 | 0.004676 | 0.006011 | 0.005922 | 4.1e-05 | 0.001061 | 0.31238 | 1.3794 | 2.5475 | 100.79 | 21.14 | 15.697 | 130.0441 | 134.852 | 119.7412 | 0.003911 | 0.89858 | 0.021246 | 23.6435 | 0.28076 | 0.14777 | 0.067581 | 4.294 | 4.294 | 0.001421
2019-03-26 16:20 | 1553631610.6723 | 431 | 0.19598 | 0.011345 | 0.000217 | 99.7897 | 97.0245 | 103.9839 | 21.2612 | 0.004428 | 0.004088 | 237.3396 | 0.24815 | 8.4013 | 0.004657 | 0.006345 | 0.00627 | 4.1e-05 | 0.001062 | 0.30055 | 1.368 | 2.5475 | 100.79 | 21.14 | 15.691 | 128.7474 | 134.3777 | 120.4138 | 0.00391 | 0.89705 | 0.021229 | 23.6652 | 0.25136 | 0.14762 | 0.06749 | 4.2848 | 4.2848 | 0.001423
2019-03-26 16:20 | 1553631611.6727 | 432 | 0.19602 | 0.011323 | 0.000216 | 99.9013 | 96.7942 | 103.987 | 21.2612 | 0.00451 | 0.004079 | 237.1546 | 0.2484 | 8.4037 | 0.00466 | 0.006406 | 0.006283 | 4.5e-05 | 0.001065 | 0.30469 | 1.3583 | 2.5475 | 100.79 | 21.14 | 15.687 | 130.7672 | 134.4607 | 120.1956 | 0.003909 | 0.89755 | 0.021235 | 23.659 | 0.28261 | 0.14777 | 0.067287 | 4.2732 | 4.2732 | 0.001428
2019-03-26 16:20 | 1553631612.6734 | 433 | 0.19607 | 0.011294 | 0.000215 | 99.8919 | 96.3054 | 104.0886 | 21.2612 | 0.004509 | 0.004075 | 236.9844 | 0.24833 | 8.4023 | 0.00461 | 0.005865 | 0.005785 | 4.2e-05 | 0.001076 | 0.33987 | 1.3355 | 2.5475 | 100.79 | 21.14 | 15.684 | 130.7416 | 133.2232 | 120.0885 | 0.003908 | 0.89779 | 0.021248 | 23.6673 | 0.32704 | 0.14773 | 0.067001 | 4.2579 | 4.2579 | 0.001433
2019-03-26 16:20 | 1553631613.6733 | 434 | 0.19614 | 0.011273 | 0.000213 | 99.868 | 96.1622 | 104.4074 | 21.2612 | 0.00448 | 0.004178 | 236.7839 | 0.24868 | 8.4001 | 0.004553 | 0.006343 | 0.006287 | 4.3e-05 | 0.001075 | 0.33787 | 1.3083 | 2.5475 | 100.79 | 21.14 | 15.677 | 130.0305 | 131.8226 | 122.6224 | 0.003906 | 0.89204 | 0.021213 | 23.7798 | 0.33821 | 0.14794 | 0.066645 | 4.2284 | 4.2284 | 0.001445
2019-03-26 16:20 | 1553631614.6724 | 435 | 0.19617 | 0.011241 | 0.000212 | 99.8883 | 96.093 | 104.7076 | 21.2612 | 0.004523 | 0.004121 | 236.6194 | 0.24831 | 8.4014 | 0.004523 | 0.005649 | 0.00554 | 4.4e-05 | 0.001078 | 0.28943 | 1.2951 | 2.5474 | 100.79 | 21.14 | 15.674 | 131.0823 | 131.0912 | 121.217 | 0.003906 | 0.89522 | 0.021281 | 23.7716 | 0.34648 | 0.14772 | 0.066469 | 4.2309 | 4.2309 | 0.001442
2019-03-26 16:20 | 1553631615.6735 | 436 | 0.19614 | 0.011211 | 0.000211 | 99.8839 | 96.1663 | 105.1273 | 21.2612 | 0.004539 | 0.004144 | 236.4301 | 0.24876 | 8.4043 | 0.004556 | 0.006291 | 0.006175 | 4.2e-05 | 0.001075 | 0.33085 | 1.3343 | 2.5475 | 100.79 | 21.14 | 15.668 | 131.4647 | 131.8881 | 121.7926 | 0.003904 | 0.89392 | 0.021308 | 23.8366 | 0.33927 | 0.14798 | 0.066718 | 4.2521 | 4.2521 | 0.001437
2019-03-26 16:20 | 1553631616.6735 | 437 | 0.19615 | 0.011178 | 0.000211 | 99.9178 | 96.5428 | 105.4364 | 21.2612 | 0.004569 | 0.004164 | 236.252 | 0.24874 | 8.4025 | 0.004475 | 0.006112 | 0.006057 | 4.4e-05 | 0.001078 | 0.3178 | 1.3664 | 2.5475 | 100.79 | 21.14 | 15.666 | 132.2229 | 129.8933 | 122.2603 | 0.003904 | 0.89286 | 0.021327 | 23.8857 | 0.30735 | 0.14798 | 0.066707 | 4.2551 | 4.2551 | 0.001436
2019-03-26 16:20 | 1553631617.6736 | 438 | 0.19618 | 0.011145 | 0.000213 | 99.9069 | 96.4503 | 105.3834 | 21.2612 | 0.0046 | 0.004136 | 236.0956 | 0.24887 | 8.403 | 0.004517 | 0.006409 | 0.006315 | 4.5e-05 | 0.001079 | 0.29789 | 1.3041 | 2.5475 | 100.79 | 21.13 | 15.663 | 132.9779 | 130.9315 | 121.5743 | 0.0039 | 0.89441 | 0.02134 | 23.859 | 0.31495 | 0.14805 | 0.066512 | 4.2453 | 4.2453 | 0.00144
2019-03-26 16:20 | 1553631618.673 | 439 | 0.19625 | 0.011121 | 0.000215 | 99.927 | 96.6023 | 105.2613 | 21.2612 | 0.004582 | 0.004106 | 235.9142 | 0.24894 | 8.4018 | 0.004604 | 0.006165 | 0.006134 | 4.6e-05 | 0.001081 | 0.31332 | 1.2949 | 2.5474 | 100.79 | 21.13 | 15.66 | 132.5326 | 133.0656 | 120.8486 | 0.0039 | 0.89606 | 0.021347 | 23.8233 | 0.30266 | 0.1481 | 0.066128 | 4.2222 | 4.2222 | 0.001449
2019-03-26 16:20 | 1553631619.6728 | 440 | 0.19613 | 0.0111 | 0.000216 | 99.8954 | 96.5219 | 105.0475 | 21.2612 | 0.00459 | 0.00416 | 235.7728 | 0.249 | 8.4023 | 0.004603 | 0.006704 | 0.006567 | 4.4e-05 | 0.00108 | 0.31868 | 1.3149 | 2.5473 | 100.79 | 21.13 | 15.658 | 132.7291 | 133.0616 | 122.1793 | 0.003899 | 0.89304 | 0.021289 | 23.8392 | 0.30728 | 0.14813 | 0.066882 | 4.2586 | 4.2586 | 0.001437
2019-03-26 16:20 | 1553631620.6728 | 441 | 0.19639 | 0.011075 | 0.000215 | 99.8969 | 95.9995 | 105.4202 | 21.2612 | 0.00459 | 0.004125 | 235.6014 | 0.2491 | 8.4021 | 0.004588 | 0.005999 | 0.005925 | 4.6e-05 | 0.001082 | 0.33196 | 1.3335 | 2.5474 | 100.79 | 21.13 | 15.658 | 132.7348 | 132.6789 | 121.3161 | 0.003899 | 0.89499 | 0.02135 | 23.8554 | 0.35595 | 0.14819 | 0.065364 | 4.1743 | 4.1743 | 0.001466
2019-03-26 16:20 | 1553631621.6742 | 442 | 0.19651 | 0.011056 | 0.000213 | 99.906 | 95.6798 | 105.4089 | 21.2612 | 0.004627 | 0.00418 | 235.4417 | 0.24846 | 8.4024 | 0.004599 | 0.005838 | 0.00581 | 4.7e-05 | 0.001078 | 0.29502 | 1.3478 | 2.5476 | 100.79 | 21.13 | 15.658 | 133.6467 | 132.9633 | 122.6592 | 0.003899 | 0.89196 | 0.021313 | 23.8947 | 0.38661 | 0.14781 | 0.064671 | 4.1231 | 4.1231 | 0.001481
2019-03-26 16:20 | 1553631622.6742 | 443 | 0.19662 | 0.011023 | 0.000211 | 99.8358 | 95.8267 | 105.481 | 21.2612 | 0.00468 | 0.004185 | 235.2717 | 0.24976 | 8.4039 | 0.004507 | 0.006201 | 0.006067 | 4.6e-05 | 0.00108 | 0.29231 | 1.4099 | 2.5476 | 100.79 | 21.13 | 15.658 | 134.9462 | 130.6961 | 122.7941 | 0.003899 | 0.89165 | 0.021317 | 23.9069 | 0.36659 | 0.14858 | 0.064086 | 4.0867 | 4.0867 | 0.001502
2019-03-26 16:20 | 1553631623.6723 | 444 | 0.19673 | 0.010984 | 0.00021 | 99.9022 | 96.4745 | 105.7946 | 21.2612 | 0.00465 | 0.004162 | 235.1308 | 0.24885 | 8.4041 | 0.004532 | 0.005838 | 0.005778 | 4.8e-05 | 0.001077 | 0.29927 | 1.3671 | 2.5474 | 100.79 | 21.13 | 15.659 | 134.2135 | 131.2947 | 122.2133 | 0.003899 | 0.89296 | 0.021364 | 23.9249 | 0.31228 | 0.14804 | 0.063443 | 4.0549 | 4.0549 | 0.001508
2019-03-26 16:20 | 1553631624.672 | 445 | 0.1969 | 0.010942 | 0.000211 | 99.922 | 96.537 | 105.8218 | 21.2612 | 0.004667 | 0.004219 | 234.9791 | 0.24892 | 8.4049 | 0.004612 | 0.006592 | 0.006544 | 4.4e-05 | 0.001078 | 0.28926 | 1.3144 | 2.5476 | 100.79 | 21.13 | 15.66 | 134.6232 | 133.2831 | 123.6102 | 0.0039 | 0.88982 | 0.021329 | 23.9702 | 0.30825 | 0.14808 | 0.062502 | 3.9885 | 3.9885 | 0.001533
2019-03-26 16:20 | 1553631625.672 | 446 | 0.19704 | 0.010892 | 0.00021 | 99.9394 | 96.3414 | 105.9481 | 21.2612 | 0.004605 | 0.00415 | 234.8266 | 0.24838 | 8.4012 | 0.004559 | 0.005886 | 0.005765 | 4.6e-05 | 0.00108 | 0.33699 | 1.3305 | 2.5473 | 100.79 | 21.13 | 15.659 | 133.1 | 131.9759 | 121.9196 | 0.003899 | 0.89363 | 0.021387 | 23.9333 | 0.32796 | 0.14776 | 0.061711 | 3.949 | 3.949 | 0.001545
2019-03-26 16:20 | 1553631626.6734 | 447 | 0.19721 | 0.010836 | 0.000209 | 99.955 | 96.4886 | 105.9704 | 21.2612 | 0.004644 | 0.004164 | 234.6805 | 0.24854 | 8.4028 | 0.004605 | 0.005827 | 0.005737 | 5e-05 | 0.00108 | 0.31404 | 1.3395 | 2.5474 | 100.79 | 21.13 | 15.659 | 134.0496 | 133.104 | 122.277 | 0.003899 | 0.89282 | 0.02138 | 23.9467 | 0.3157 | 0.14786 | 0.060805 | 3.8899 | 3.8899 | 0.00157
2019-03-26 16:20 | 1553631627.6721 | 448 | 0.19723 | 0.010763 | 0.000206 | 99.8659 | 96.1659 | 106.1551 | 21.2612 | 0.004655 | 0.004191 | 234.5249 | 0.24843 | 8.4009 | 0.004598 | 0.006524 | 0.006421 | 5e-05 | 0.001083 | 0.30735 | 1.3574 | 2.5475 | 100.79 | 21.13 | 15.66 | 134.34 | 132.9281 | 122.9395 | 0.0039 | 0.89133 | 0.021381 | 23.9876 | 0.33769 | 0.14779 | 0.060759 | 3.8872 | 3.8872 | 0.00157
2019-03-26 16:20 | 1553631628.6721 | 449 | 0.1974 | 0.01069 | 0.000203 | 99.8721 | 95.9417 | 106.637 | 21.2612 | 0.004658 | 0.004159 | 234.3733 | 0.2486 | 8.4011 | 0.004601 | 0.005964 | 0.005874 | 5e-05 | 0.001082 | 0.29247 | 1.3417 | 2.5475 | 100.79 | 21.13 | 15.66 | 134.4169 | 132.9914 | 122.1389 | 0.0039 | 0.89313 | 0.021451 | 24.0177 | 0.35912 | 0.14789 | 0.059849 | 3.842 | 3.842 | 0.00159
2019-03-26 16:20 | 1553631629.6726 | 450 | 0.19758 | 0.010607 | 0.0002 | 99.8963 | 95.9746 | 107.1046 | 21.2612 | 0.004667 | 0.004138 | 234.2141 | 0.24838 | 8.4042 | 0.004552 | 0.005798 | 0.005751 | 4.6e-05 | 0.001081 | 0.28191 | 1.331 | 2.5475 | 100.79 | 21.13 | 15.66 | 134.6286 | 131.793 | 121.6305 | 0.0039 | 0.89428 | 0.021512 | 24.0548 | 0.35822 | 0.14776 | 0.058896 | 3.7919 | 3.7919 | 0.001609
2019-03-26 16:20 | 1553631630.6733 | 451 | 0.19772 | 0.010514 | 0.000198 | 99.9246 | 96.0009 | 107.3144 | 21.2612 | 0.004674 | 0.004149 | 234.0649 | 0.24777 | 8.4024 | 0.004494 | 0.006762 | 0.006645 | 4.8e-05 | 0.001084 | 0.31277 | 1.342 | 2.5475 | 100.79 | 21.13 | 15.662 | 134.7935 | 130.3608 | 121.9061 | 0.0039 | 0.89366 | 0.021525 | 24.0867 | 0.35831 | 0.1474 | 0.058129 | 3.7451 | 3.7451 | 0.001625
2019-03-26 16:20 | 1553631631.6734 | 452 | 0.19789 | 0.010418 | 0.000197 | 99.8455 | 95.9804 | 107.5412 | 21.2612 | 0.004644 | 0.004176 | 233.9187 | 0.24835 | 8.4041 | 0.004509 | 0.006444 | 0.006395 | 4.6e-05 | 0.001078 | 0.30917 | 1.3731 | 2.5475 | 100.79 | 21.13 | 15.662 | 134.0688 | 130.7273 | 122.5605 | 0.0039 | 0.89218 | 0.02153 | 24.1321 | 0.35314 | 0.14774 | 0.057256 | 3.69 | 3.69 | 0.001654
2019-03-26 16:21 | 1553631632.6724 | 453 | 0.1981 | 0.010314 | 0.000198 | 99.9335 | 95.7699 | 107.7475 | 21.2612 | 0.004612 | 0.004119 | 233.7701 | 0.24784 | 8.4036 | 0.004423 | 0.005415 | 0.005384 | 4.7e-05 | 0.001084 | 0.31797 | 1.3597 | 2.5476 | 100.79 | 21.13 | 15.663 | 133.2838 | 128.6302 | 121.1669 | 0.0039 | 0.89533 | 0.021589 | 24.1127 | 0.38065 | 0.14744 | 0.05614 | 3.6281 | 3.6281 | 0.001678
2019-03-26 16:21 | 1553631633.6726 | 454 | 0.19843 | 0.010205 | 0.000199 | 99.8913 | 95.7155 | 107.6806 | 21.2612 | 0.004677 | 0.004164 | 233.6265 | 0.2481 | 8.4022 | 0.004498 | 0.005937 | 0.005793 | 4.9e-05 | 0.001078 | 0.33395 | 1.3234 | 2.5475 | 100.79 | 21.13 | 15.666 | 134.8663 | 130.4727 | 122.2603 | 0.003901 | 0.89286 | 0.021552 | 24.1386 | 0.38195 | 0.14759 | 0.054273 | 3.5017 | 3.5017 | 0.001741
2019-03-26 16:21 | 1553631634.6725 | 455 | 0.19856 | 0.010088 | 0.000199 | 99.9051 | 95.6247 | 108.1019 | 21.2612 | 0.004578 | 0.004135 | 233.49 | 0.24842 | 8.4039 | 0.004459 | 0.005987 | 0.005902 | 4.8e-05 | 0.001079 | 0.31526 | 1.3641 | 2.5474 | 100.79 | 21.13 | 15.668 | 132.4297 | 129.5081 | 121.5701 | 0.003902 | 0.89442 | 0.021613 | 24.1647 | 0.39168 | 0.14779 | 0.053627 | 3.4699 | 3.4699 | 0.001759
2019-03-26 16:21 | 1553631635.6735 | 456 | 0.19862 | 0.009972 | 0.0002 | 99.9736 | 95.1954 | 108.3028 | 21.2612 | 0.004622 | 0.004105 | 233.3388 | 0.24885 | 8.4009 | 0.004497 | 0.006086 | 0.006003 | 4.7e-05 | 0.001077 | 0.2896 | 1.355 | 2.5473 | 100.79 | 21.13 | 15.67 | 133.5228 | 130.4326 | 120.8178 | 0.003902 | 0.89613 | 0.021654 | 24.1641 | 0.43806 | 0.14804 | 0.053444 | 3.4645 | 3.4645 | 0.001765
2019-03-26 16:21 | 1553631636.673 | 457 | 0.19877 | 0.009852 | 0.000201 | 99.9795 | 94.9692 | 108.535 | 21.2612 | 0.004627 | 0.004081 | 233.1612 | 0.2487 | 8.401 | 0.004513 | 0.006235 | 0.006176 | 4.8e-05 | 0.00108 | 0.32477 | 1.3993 | 2.5476 | 100.79 | 21.13 | 15.676 | 133.6475 | 130.8351 | 120.233 | 0.003904 | 0.89746 | 0.021693 | 24.172 | 0.45987 | 0.14795 | 0.052704 | 3.4229 | 3.4229 | 0.001785
2019-03-26 16:21 | 1553631637.6726 | 458 | 0.19911 | 0.009737 | 0.000203 | 99.9196 | 94.791 | 108.8298 | 21.2612 | 0.004569 | 0.004101 | 233.0148 | 0.2478 | 8.4033 | 0.004427 | 0.006061 | 0.005994 | 5e-05 | 0.001075 | 0.28639 | 1.388 | 2.5475 | 100.79 | 21.12 | 15.68 | 132.2211 | 128.7154 | 120.7202 | 0.003902 | 0.89635 | 0.021709 | 24.2198 | 0.4713 | 0.14742 | 0.050793 | 3.3014 | 3.3014 | 0.001844
2019-03-26 16:21 | 1553631638.674 | 459 | 0.19932 | 0.009618 | 0.000205 | 99.9249 | 94.6616 | 108.8052 | 21.2612 | 0.004598 | 0.004046 | 232.8969 | 0.24781 | 8.4026 | 0.004519 | 0.00613 | 0.006054 | 4.7e-05 | 0.001077 | 0.31653 | 1.4033 | 2.5475 | 100.79 | 21.12 | 15.685 | 132.9195 | 130.9763 | 119.3914 | 0.003903 | 0.89938 | 0.021744 | 24.1761 | 0.484 | 0.14742 | 0.049698 | 3.2353 | 3.2353 | 0.001882
2019-03-26 16:21 | 1553631639.6731 | 460 | 0.19935 | 0.009497 | 0.000209 | 99.9317 | 94.6088 | 108.8566 | 21.2612 | 0.004597 | 0.004115 | 232.7757 | 0.24806 | 8.403 | 0.004505 | 0.006054 | 0.005948 | 4.6e-05 | 0.001077 | 0.30414 | 1.3665 | 2.5475 | 100.79 | 21.12 | 15.689 | 132.9042 | 130.6353 | 121.0732 | 0.003904 | 0.89555 | 0.021702 | 24.2336 | 0.48959 | 0.14757 | 0.04965 | 3.2258 | 3.2258 | 0.001889
2019-03-26 16:21 | 1553631640.6727 | 461 | 0.19949 | 0.009367 | 0.000211 | 99.9267 | 94.5456 | 108.703 | 21.2612 | 0.00456 | 0.004063 | 232.639 | 0.24741 | 8.4012 | 0.00446 | 0.005991 | 0.005948 | 4.5e-05 | 0.001081 | 0.32117 | 1.3502 | 2.5475 | 100.79 | 21.12 | 15.691 | 131.9919 | 129.5383 | 119.7903 | 0.003905 | 0.89847 | 0.021722 | 24.1771 | 0.49512 | 0.14718 | 0.048946 | 3.1829 | 3.1829 | 0.00191
2019-03-26 16:21 | 1553631641.6732 | 462 | 0.19956 | 0.009236 | 0.000212 | 99.9406 | 94.2256 | 108.8197 | 21.2612 | 0.004623 | 0.004023 | 232.4884 | 0.24743 | 8.4034 | 0.004428 | 0.005778 | 0.005683 | 4.7e-05 | 0.001087 | 0.34827 | 1.314 | 2.5476 | 100.79 | 21.12 | 15.696 | 133.5437 | 128.7521 | 118.8245 | 0.003906 | 0.90068 | 0.021761 | 24.1603 | 0.52669 | 0.1472 | 0.048698 | 3.1724 | 3.1724 | 0.001916
2019-03-26 16:21 | 1553631642.6719 | 463 | 0.19962 | 0.0091 | 0.000212 | 99.8609 | 94.0185 | 108.6629 | 21.2612 | 0.004565 | 0.004062 | 232.3502 | 0.24758 | 8.4012 | 0.004319 | 0.005481 | 0.005405 | 4.8e-05 | 0.001088 | 0.3011 | 1.3217 | 2.5475 | 100.79 | 21.12 | 15.699 | 132.1197 | 126.0651 | 119.769 | 0.003907 | 0.89852 | 0.021719 | 24.172 | 0.53923 | 0.14728 | 0.048482 | 3.1523 | 3.1523 | 0.00193
2019-03-26 16:21 | 1553631643.6729 | 464 | 0.19965 | 0.008971 | 0.000212 | 99.9069 | 93.7486 | 108.9152 | 21.2612 | 0.004549 | 0.004052 | 232.1968 | 0.24791 | 8.4037 | 0.004327 | 0.006 | 0.005939 | 5e-05 | 0.001092 | 0.31813 | 1.2924 | 2.5474 | 100.79 | 21.13 | 15.702 | 131.7258 | 126.2538 | 119.5373 | 0.00391 | 0.89905 | 0.021751 | 24.1929 | 0.56907 | 0.14749 | 0.048442 | 3.1543 | 3.1543 | 0.001931
2019-03-26 16:21 | 1553631644.6725 | 465 | 0.19971 | 0.008844 | 0.00021 | 99.971 | 93.7741 | 108.5071 | 21.2612 | 0.004559 | 0.004025 | 232.0709 | 0.2486 | 8.4024 | 0.004342 | 0.005978 | 0.005915 | 4.9e-05 | 0.001095 | 0.31885 | 1.305 | 2.5475 | 100.79 | 21.13 | 15.702 | 131.9592 | 126.6327 | 118.8643 | 0.00391 | 0.90059 | 0.021728 | 24.1268 | 0.57237 | 0.14789 | 0.048262 | 3.1395 | 3.1395 | 0.001946
2019-03-26 16:21 | 1553631645.6721 | 466 | 0.19972 | 0.008741 | 0.000208 | 99.9353 | 94.0313 | 108.1716 | 21.2612 | 0.004551 | 0.004112 | 231.9648 | 0.24842 | 8.4043 | 0.004374 | 0.005381 | 0.005307 | 5.2e-05 | 0.001097 | 0.28943 | 1.3413 | 2.5476 | 100.79 | 21.13 | 15.701 | 131.7642 | 127.4301 | 120.9964 | 0.00391 | 0.89572 | 0.021636 | 24.1549 | 0.54468 | 0.14779 | 0.048358 | 3.1324 | 3.1324 | 0.001949
2019-03-26 16:21 | 1553631646.6731 | 467 | 0.19963 | 0.008638 | 0.000205 | 99.9332 | 94.1573 | 108.255 | 21.2612 | 0.004601 | 0.004068 | 231.8478 | 0.24828 | 8.4034 | 0.004358 | 0.005694 | 0.005542 | 4.8e-05 | 0.001097 | 0.28711 | 1.3586 | 2.5476 | 100.79 | 21.13 | 15.693 | 133.0057 | 127.0213 | 119.9338 | 0.003908 | 0.89814 | 0.021674 | 24.1316 | 0.53251 | 0.1477 | 0.048992 | 3.1791 | 3.1791 | 0.001919
2019-03-26 16:21 | 1553631647.6735 | 468 | 0.19963 | 0.008558 | 0.000203 | 99.8987 | 94.7433 | 108.4803 | 21.2612 | 0.004618 | 0.00407 | 231.6972 | 0.24944 | 8.3997 | 0.004355 | 0.006054 | 0.005967 | 4.9e-05 | 0.001103 | 0.31797 | 1.358 | 2.5476 | 100.79 | 21.13 | 15.687 | 133.4284 | 126.941 | 119.98 | 0.003906 | 0.89804 | 0.021695 | 24.1581 | 0.47393 | 0.14839 | 0.049098 | 3.1893 | 3.1893 | 0.001922
2019-03-26 16:21 | 1553631648.6722 | 469 | 0.1996 | 0.008489 | 0.000201 | 99.8939 | 94.7762 | 108.1245 | 21.2612 | 0.004567 | 0.004073 | 231.5468 | 0.24936 | 8.402 | 0.004357 | 0.005761 | 0.005785 | 5.2e-05 | 0.001101 | 0.31293 | 1.3826 | 2.5477 | 100.79 | 21.13 | 15.681 | 132.1763 | 126.9896 | 120.044 | 0.003905 | 0.89789 | 0.021658 | 24.1204 | 0.4704 | 0.14834 | 0.049344 | 3.1998 | 3.1998 | 0.001915
2019-03-26 16:21 | 1553631649.6722 | 470 | 0.19961 | 0.008438 | 0.000199 | 99.9037 | 94.715 | 108.0174 | 21.2612 | 0.004561 | 0.004087 | 231.4041 | 0.24835 | 8.4021 | 0.004399 | 0.006087 | 0.006002 | 5.2e-05 | 0.001099 | 0.31116 | 1.3944 | 2.5475 | 100.79 | 21.13 | 15.666 | 132.0239 | 128.0315 | 120.3971 | 0.003901 | 0.89709 | 0.021637 | 24.1193 | 0.47704 | 0.14774 | 0.049396 | 3.2002 | 3.2002 | 0.001907
2019-03-26 16:21 | 1553631650.6722 | 471 | 0.19969 | 0.008405 | 0.000197 | 99.8842 | 94.5379 | 108.1942 | 21.2612 | 0.004602 | 0.004077 | 231.2731 | 0.2489 | 8.4012 | 0.004433 | 0.006161 | 0.006022 | 5.1e-05 | 0.001101 | 0.33124 | 1.3702 | 2.5474 | 100.79 | 21.13 | 15.658 | 133.0161 | 128.8661 | 120.1497 | 0.003899 | 0.89765 | 0.021662 | 24.1314 | 0.49205 | 0.14807 | 0.048943 | 3.1747 | 3.1747 | 0.001926
2019-03-26 16:21 | 1553631651.6734 | 472 | 0.19969 | 0.008377 | 0.000195 | 99.8907 | 94.6186 | 108.4998 | 21.2612 | 0.004557 | 0.004103 | 231.1194 | 0.24825 | 8.401 | 0.004468 | 0.005915 | 0.005884 | 4.9e-05 | 0.001092 | 0.32853 | 1.4156 | 2.5476 | 100.79 | 21.13 | 15.64 | 131.9247 | 129.7401 | 120.7821 | 0.003895 | 0.89621 | 0.021675 | 24.1849 | 0.485 | 0.14768 | 0.048967 | 3.1783 | 3.1783 | 0.001919
2019-03-26 16:21 | 1553631652.6729 | 473 | 0.1997 | 0.00835 | 0.000192 | 99.8833 | 94.7339 | 108.8751 | 21.2612 | 0.004522 | 0.004051 | 231.0016 | 0.24799 | 8.4028 | 0.004447 | 0.006386 | 0.00634 | 4.8e-05 | 0.001086 | 0.32637 | 1.3699 | 2.5475 | 100.79 | 21.13 | 15.629 | 131.0464 | 129.201 | 119.5033 | 0.003892 | 0.89913 | 0.021748 | 24.1874 | 0.47344 | 0.14753 | 0.04897 | 3.1893 | 3.1893 | 0.00191
2019-03-26 16:21 | 1553631653.6735 | 474 | 0.19968 | 0.008336 | 0.00019 | 99.8449 | 94.9225 | 108.6336 | 21.2612 | 0.004605 | 0.004046 | 230.8676 | 0.24859 | 8.4004 | 0.004445 | 0.00638 | 0.006272 | 5e-05 | 0.00109 | 0.32532 | 1.3581 | 2.5476 | 100.79 | 21.13 | 15.618 | 133.0919 | 129.1691 | 119.385 | 0.003889 | 0.8994 | 0.021727 | 24.1569 | 0.45222 | 0.14789 | 0.049091 | 3.1943 | 3.1943 | 0.001912
2019-03-26 16:21 | 1553631654.6725 | 475 | 0.19968 | 0.008304 | 0.000189 | 99.9025 | 95.1423 | 108.8167 | 21.2612 | 0.004563 | 0.004126 | 230.7088 | 0.24792 | 8.4017 | 0.004438 | 0.00641 | 0.006385 | 5.1e-05 | 0.001089 | 0.31996 | 1.3528 | 2.5475 | 100.79 | 21.13 | 15.595 | 132.0749 | 129 | 121.3281 | 0.003883 | 0.89497 | 0.021691 | 24.237 | 0.43668 | 0.14749 | 0.049133 | 3.1919 | 3.1919 | 0.001908
2019-03-26 16:21 | 1553631655.6715 | 476 | 0.1997 | 0.008284 | 0.000187 | 99.8508 | 95.3736 | 109.053 | 21.2612 | 0.004507 | 0.004055 | 230.5772 | 0.24837 | 8.4045 | 0.004389 | 0.006657 | 0.006494 | 5.2e-05 | 0.001091 | 0.32731 | 1.3273 | 2.5474 | 100.79 | 21.13 | 15.584 | 130.6865 | 127.795 | 119.596 | 0.003881 | 0.89892 | 0.021763 | 24.21 | 0.41034 | 0.14776 | 0.049048 | 3.1969 | 3.1969 | 0.001909
2019-03-26 16:21 | 1553631656.6716 | 477 | 0.19979 | 0.008266 | 0.000186 | 99.862 | 95.1312 | 109.1655 | 21.2612 | 0.004558 | 0.004033 | 230.4899 | 0.24874 | 8.4047 | 0.00449 | 0.005482 | 0.005446 | 5.1e-05 | 0.001086 | 0.32012 | 1.3343 | 2.5474 | 100.79 | 21.12 | 15.565 | 131.9425 | 130.2605 | 119.0664 | 0.003874 | 0.90013 | 0.021789 | 24.2061 | 0.4341 | 0.14798 | 0.048542 | 3.1678 | 3.1678 | 0.001929
2019-03-26 16:21 | 1553631657.6716 | 478 | 0.1999 | 0.008256 | 0.000186 | 99.7944 | 94.8424 | 109.054 | 21.2612 | 0.004503 | 0.004031 | 230.3742 | 0.24765 | 8.4003 | 0.004499 | 0.006507 | 0.006383 | 5.2e-05 | 0.001087 | 0.3413 | 1.3346 | 2.5475 | 100.79 | 21.12 | 15.557 | 130.5923 | 130.499 | 119.029 | 0.003872 | 0.90021 | 0.021779 | 24.1926 | 0.45524 | 0.14733 | 0.047889 | 3.1239 | 3.1239 | 0.001948
2019-03-26 16:21 | 1553631658.6731 | 479 | 0.20003 | 0.00823 | 0.000186 | 99.8222 | 94.7655 | 109.3852 | 21.2612 | 0.004539 | 0.004012 | 230.2558 | 0.24809 | 8.4036 | 0.004524 | 0.005946 | 0.005857 | 5.1e-05 | 0.001087 | 0.32012 | 1.2911 | 2.5474 | 100.79 | 21.12 | 15.55 | 131.4849 | 131.0951 | 118.555 | 0.00387 | 0.9013 | 0.021825 | 24.2147 | 0.46499 | 0.14759 | 0.047147 | 3.0821 | 3.0821 | 0.001978
2019-03-26 16:21 | 1553631659.6717 | 480 | 0.2001 | 0.008209 | 0.000186 | 99.8561 | 95.0277 | 109.2133 | 21.2612 | 0.004545 | 0.004056 | 230.129 | 0.24885 | 8.4019 | 0.00451 | 0.005593 | 0.005467 | 5e-05 | 0.001092 | 0.29302 | 1.3361 | 2.5475 | 100.79 | 21.12 | 15.54 | 131.6198 | 130.7718 | 119.638 | 0.003867 | 0.89882 | 0.021778 | 24.2291 | 0.44331 | 0.14804 | 0.046761 | 3.0503 | 3.0503 | 0.002004
2019-03-26 16:21 | 1553631660.6727 | 481 | 0.20011 | 0.008168 | 0.000186 | 99.8809 | 95.0099 | 109.1462 | 21.2612 | 0.00453 | 0.004072 | 230.0067 | 0.24777 | 8.4022 | 0.004417 | 0.006127 | 0.006007 | 4.8e-05 | 0.001093 | 0.3298 | 1.313 | 2.5475 | 100.79 | 21.12 | 15.537 | 131.2587 | 128.4651 | 120.0093 | 0.003867 | 0.89797 | 0.021761 | 24.2331 | 0.44721 | 0.1474 | 0.046733 | 3.0461 | 3.0461 | 0.001998
2019-03-26 16:21 | 1553631661.6719 | 482 | 0.20014 | 0.00813 | 0.000185 | 99.9081 | 95.0603 | 109.008 | 21.2612 | 0.004541 | 0.004014 | 229.9165 | 0.24795 | 8.4026 | 0.004471 | 0.006081 | 0.005984 | 4.9e-05 | 0.001093 | 0.33588 | 1.371 | 2.5476 | 100.79 | 21.12 | 15.532 | 131.5344 | 129.8144 | 118.6122 | 0.003865 | 0.90117 | 0.021786 | 24.1746 | 0.4449 | 0.1475 | 0.046602 | 3.0411 | 3.0411 | 0.002003
2019-03-26 16:21 | 1553631662.6722 | 483 | 0.20012 | 0.008091 | 0.000184 | 99.9122 | 94.8028 | 109.0504 | 21.2612 | 0.004483 | 0.004052 | 229.7777 | 0.24717 | 8.4029 | 0.004481 | 0.005431 | 0.005412 | 5.3e-05 | 0.001096 | 0.32372 | 1.3544 | 2.5475 | 100.79 | 21.12 | 15.531 | 130.0967 | 130.0568 | 119.5341 | 0.003865 | 0.89906 | 0.021764 | 24.2078 | 0.46953 | 0.14704 | 0.046787 | 3.0502 | 3.0502 | 0.001991
2019-03-26 16:21 | 1553631663.6728 | 484 | 0.20013 | 0.008047 | 0.000184 | 99.8033 | 94.7624 | 109.3312 | 21.2612 | 0.004493 | 0.00402 | 229.6573 | 0.24842 | 8.4015 | 0.004497 | 0.005756 | 0.005635 | 5e-05 | 0.001096 | 0.32836 | 1.3303 | 2.5475 | 100.79 | 21.12 | 15.53 | 130.3394 | 130.44 | 118.7446 | 0.003865 | 0.90087 | 0.021814 | 24.2145 | 0.46358 | 0.14778 | 0.046747 | 3.0546 | 3.0546 | 0.001998
2019-03-26 16:21 | 1553631664.6719 | 485 | 0.20019 | 0.008008 | 0.000183 | 99.837 | 94.9857 | 109.5584 | 21.2612 | 0.004525 | 0.004007 | 229.5487 | 0.24874 | 8.405 | 0.004448 | 0.006102 | 0.006021 | 5.1e-05 | 0.001099 | 0.3282 | 1.3259 | 2.5475 | 100.79 | 21.12 | 15.531 | 131.1207 | 129.2362 | 118.4243 | 0.003865 | 0.9016 | 0.021846 | 24.2298 | 0.44555 | 0.14797 | 0.04648 | 3.0416 | 3.0416 | 0.002009
2019-03-26 16:21 | 1553631665.673 | 486 | 0.20023 | 0.007967 | 0.000182 | 99.8364 | 95.0004 | 109.5887 | 21.2612 | 0.004487 | 0.003972 | 229.4247 | 0.24872 | 8.4017 | 0.004442 | 0.005997 | 0.00589 | 5e-05 | 0.001094 | 0.28224 | 1.3475 | 2.5475 | 100.79 | 21.12 | 15.531 | 130.1887 | 129.087 | 117.5679 | 0.003865 | 0.90358 | 0.021873 | 24.2066 | 0.44411 | 0.14796 | 0.046263 | 3.0312 | 3.0312 | 0.002016
2019-03-26 16:21 | 1553631666.6724 | 487 | 0.20037 | 0.007926 | 0.000182 | 99.8455 | 95.0099 | 109.5251 | 21.2612 | 0.004544 | 0.00398 | 229.3096 | 0.24859 | 8.4018 | 0.004404 | 0.006407 | 0.006342 | 4.6e-05 | 0.001098 | 0.32189 | 1.3265 | 2.5473 | 100.79 | 21.12 | 15.531 | 131.5878 | 128.1464 | 117.7808 | 0.003865 | 0.90309 | 0.02186 | 24.2062 | 0.44404 | 0.14788 | 0.045475 | 2.978 | 2.978 | 0.002051
2019-03-26 16:21 | 1553631667.6721 | 488 | 0.20046 | 0.00789 | 0.000181 | 99.8479 | 95.3632 | 109.2093 | 21.2612 | 0.004516 | 0.003952 | 229.2009 | 0.24718 | 8.4044 | 0.004382 | 0.005691 | 0.005664 | 4.7e-05 | 0.001095 | 0.30021 | 1.3266 | 2.5475 | 100.79 | 21.12 | 15.53 | 130.9005 | 127.6037 | 117.0966 | 0.003865 | 0.90467 | 0.021848 | 24.1501 | 0.41105 | 0.14705 | 0.04496 | 2.9427 | 2.9427 | 0.002064
2019-03-26 16:21 | 1553631668.6735 | 489 | 0.20053 | 0.007847 | 0.00018 | 99.8768 | 95.1655 | 109.1809 | 21.2612 | 0.004498 | 0.004012 | 229.0621 | 0.2477 | 8.4046 | 0.004378 | 0.005618 | 0.005463 | 5.1e-05 | 0.001099 | 0.30541 | 1.3883 | 2.5475 | 100.79 | 21.12 | 15.53 | 130.4759 | 127.5248 | 118.5518 | 0.003865 | 0.90131 | 0.021804 | 24.1919 | 0.4322 | 0.14736 | 0.04462 | 2.9148 | 2.9148 | 0.002088
2019-03-26 16:21 | 1553631669.6723 | 490 | 0.20054 | 0.007807 | 0.000177 | 99.7714 | 94.9436 | 109.483 | 21.2612 | 0.004451 | 0.003992 | 228.9399 | 0.24851 | 8.4028 | 0.004447 | 0.006544 | 0.006468 | 4.9e-05 | 0.001095 | 0.30773 | 1.3851 | 2.5476 | 100.79 | 21.12 | 15.529 | 129.3222 | 129.2202 | 118.0741 | 0.003865 | 0.90241 | 0.021848 | 24.2106 | 0.44364 | 0.14784 | 0.04459 | 2.9187 | 2.9187 | 0.002092
2019-03-26 16:21 | 1553631670.6723 | 491 | 0.20059 | 0.007759 | 0.000176 | 99.8083 | 94.7823 | 110.0149 | 21.2612 | 0.004468 | 0.004001 | 228.808 | 0.24712 | 8.4037 | 0.004424 | 0.006129 | 0.006021 | 4.9e-05 | 0.001094 | 0.3147 | 1.3658 | 2.5476 | 100.79 | 21.12 | 15.527 | 129.7362 | 128.6412 | 118.2888 | 0.003864 | 0.90192 | 0.021895 | 24.276 | 0.46216 | 0.14701 | 0.044354 | 2.9096 | 2.9096 | 0.002087
2019-03-26 16:21 | 1553631671.6722 | 492 | 0.20064 | 0.007717 | 0.000175 | 99.708 | 95.0236 | 110.1466 | 21.2612 | 0.004493 | 0.003943 | 228.6949 | 0.24804 | 8.4035 | 0.004466 | 0.005992 | 0.005919 | 5e-05 | 0.001093 | 0.29789 | 1.3295 | 2.5475 | 100.79 | 21.12 | 15.526 | 130.3473 | 129.6915 | 116.8733 | 0.003864 | 0.90519 | 0.021948 | 24.2466 | 0.43041 | 0.14756 | 0.044123 | 2.9016 | 2.9016 | 0.0021
2019-03-26 16:21 | 1553631672.6723 | 493 | 0.20072 | 0.007673 | 0.000173 | 99.8119 | 95.6788 | 110.3041 | 21.2612 | 0.004487 | 0.004024 | 228.612 | 0.24745 | 8.4024 | 0.004437 | 0.004962 | 0.004841 | 4.8e-05 | 0.001088 | 0.32444 | 1.333 | 2.5475 | 100.79 | 21.11 | 15.523 | 130.195 | 128.9754 | 118.8443 | 0.003861 | 0.90064 | 0.021908 | 24.3251 | 0.37827 | 0.14721 | 0.04369 | 2.868 | 2.868 | 0.00212
2019-03-26 16:21 | 1553631673.6724 | 494 | 0.20073 | 0.007635 | 0.000171 | 99.7971 | 95.3867 | 110.1602 | 21.2612 | 0.004457 | 0.00391 | 228.5046 | 0.24941 | 8.4039 | 0.004443 | 0.00572 | 0.005644 | 4.8e-05 | 0.001085 | 0.31669 | 1.3399 | 2.5475 | 100.79 | 21.11 | 15.521 | 129.4562 | 129.1133 | 116.0557 | 0.00386 | 0.90709 | 0.021972 | 24.2226 | 0.40429 | 0.14838 | 0.043692 | 2.8767 | 2.8767 | 0.00213
2019-03-26 16:21 | 1553631674.6718 | 495 | 0.20084 | 0.007595 | 0.000168 | 99.7629 | 95.0778 | 109.8344 | 21.2612 | 0.004498 | 0.003924 | 228.3987 | 0.24802 | 8.4024 | 0.004473 | 0.006039 | 0.005958 | 5.2e-05 | 0.001081 | 0.29214 | 1.3295 | 2.5475 | 100.79 | 21.11 | 15.516 | 130.4781 | 129.8446 | 116.4038 | 0.003859 | 0.90628 | 0.02193 | 24.1976 | 0.43024 | 0.14755 | 0.043113 | 2.8333 | 2.8333 | 0.002151
2019-03-26 16:21 | 1553631675.6725 | 496 | 0.20098 | 0.00755 | 0.000166 | 99.7177 | 95.0589 | 109.8684 | 21.2612 | 0.004427 | 0.003929 | 228.2653 | 0.24802 | 8.4004 | 0.00446 | 0.005816 | 0.00562 | 4.8e-05 | 0.001085 | 0.33085 | 1.4117 | 2.5475 | 100.79 | 21.11 | 15.514 | 128.7147 | 129.524 | 116.5297 | 0.003859 | 0.90599 | 0.02193 | 24.2053 | 0.42796 | 0.14755 | 0.042299 | 2.78 | 2.78 | 0.002192
2019-03-26 16:21 | 1553631676.6715 | 497 | 0.20108 | 0.007508 | 0.000163 | 99.7608 | 95.1036 | 109.9689 | 21.2612 | 0.004442 | 0.003917 | 228.1259 | 0.24834 | 8.4008 | 0.004418 | 0.006362 | 0.00637 | 4.6e-05 | 0.001079 | 0.32211 | 1.3873 | 2.5475 | 100.79 | 21.11 | 15.512 | 129.0838 | 128.5105 | 116.2242 | 0.003858 | 0.9067 | 0.021948 | 24.2068 | 0.42762 | 0.14774 | 0.04178 | 2.7484 | 2.7484 | 0.00222
2019-03-26 16:21 | 1553631677.6719 | 498 | 0.20109 | 0.00746 | 0.000162 | 99.7162 | 95.5191 | 110.4172 | 21.2612 | 0.004424 | 0.003952 | 228.0113 | 0.24872 | 8.4014 | 0.004466 | 0.005485 | 0.005331 | 4.9e-05 | 0.001081 | 0.32925 | 1.4432 | 2.5477 | 100.79 | 21.11 | 15.509 | 128.6412 | 129.6851 | 117.0888 | 0.003857 | 0.90469 | 0.021969 | 24.283 | 0.38463 | 0.14796 | 0.041756 | 2.7494 | 2.7494 | 0.002223
2019-03-26 16:21 | 1553631678.6716 | 499 | 0.20111 | 0.007401 | 0.000161 | 99.7859 | 95.6028 | 110.6733 | 21.2612 | 0.004398 | 0.003909 | 227.9413 | 0.24715 | 8.4001 | 0.004413 | 0.006349 | 0.006277 | 4.7e-05 | 0.001083 | 0.33948 | 1.4054 | 2.5475 | 100.79 | 21.11 | 15.509 | 128.002 | 128.3656 | 116.0359 | 0.003857 | 0.90714 | 0.022024 | 24.2784 | 0.38304 | 0.14703 | 0.041736 | 2.755 | 2.755 | 0.002204
2019-03-26 16:21 | 1553631679.6731 | 500 | 0.20119 | 0.00734 | 0.00016 | 99.7354 | 95.8062 | 110.8487 | 21.2612 | 0.004452 | 0.003915 | 227.8631 | 0.24813 | 8.4002 | 0.004408 | 0.006173 | 0.006121 | 4.9e-05 | 0.001078 | 0.31188 | 1.3819 | 2.5476 | 100.79 | 21.11 | 15.509 | 129.323 | 128.2452 | 116.1879 | 0.003857 | 0.90678 | 0.022037 | 24.3023 | 0.35951 | 0.14761 | 0.04133 | 2.73 | 2.73 | 0.002233
2019-03-26 16:21 | 1553631680.6731 | 501 | 0.2012 | 0.007288 | 0.000159 | 99.7682 | 95.8516 | 110.282 | 21.2612 | 0.004474 | 0.003925 | 227.7638 | 0.24801 | 8.4053 | 0.004371 | 0.005664 | 0.005566 | 4.8e-05 | 0.001088 | 0.29518 | 1.3572 | 2.5475 | 100.79 | 21.11 | 15.508 | 129.8655 | 127.3552 | 116.426 | 0.003857 | 0.90623 | 0.021974 | 24.2476 | 0.35821 | 0.14754 | 0.041286 | 2.7193 | 2.7193 | 0.002241
2019-03-26 16:21 | 1553631681.6718 | 502 | 0.20129 | 0.007241 | 0.000156 | 99.6938 | 95.4371 | 110.3128 | 21.2612 | 0.004357 | 0.003859 | 227.656 | 0.24838 | 8.4035 | 0.00432 | 0.006172 | 0.006058 | 4.8e-05 | 0.001088 | 0.30558 | 1.3255 | 2.5474 | 100.79 | 21.11 | 15.508 | 126.9928 | 126.0992 | 114.8168 | 0.003857 | 0.90999 | 0.022022 | 24.2008 | 0.3903 | 0.14776 | 0.040838 | 2.696 | 2.696 | 0.002264
2019-03-26 16:21 | 1553631682.6728 | 503 | 0.20137 | 0.0072 | 0.000153 | 99.7242 | 95.2963 | 110.1335 | 21.2612 | 0.004383 | 0.00392 | 227.5219 | 0.24848 | 8.402 | 0.004329 | 0.0058 | 0.005731 | 4.9e-05 | 0.001092 | 0.31061 | 1.3314 | 2.5474 | 100.79 | 21.11 | 15.509 | 127.6506 | 126.3223 | 116.3025 | 0.003857 | 0.90652 | 0.021963 | 24.2274 | 0.40624 | 0.14782 | 0.040406 | 2.6603 | 2.6603 | 0.002295
2019-03-26 16:21 | 1553631683.6718 | 504 | 0.20145 | 0.007157 | 0.000149 | 99.6941 | 95.2331 | 110.1957 | 21.2612 | 0.004367 | 0.003919 | 227.4429 | 0.24648 | 8.4022 | 0.004287 | 0.005806 | 0.005698 | 4.7e-05 | 0.001081 | 0.32372 | 1.2968 | 2.5474 | 100.79 | 21.11 | 15.51 | 127.2498 | 125.289 | 116.2757 | 0.003858 | 0.90658 | 0.021969 | 24.2334 | 0.40947 | 0.14663 | 0.039976 | 2.6331 | 2.6331 | 0.0023
2019-03-26 16:21 | 1553631684.6719 | 505 | 0.20157 | 0.007121 | 0.000147 | 99.6781 | 95.3484 | 110.5269 | 21.2612 | 0.004422 | 0.003894 | 227.3255 | 0.24744 | 8.4037 | 0.004336 | 0.005457 | 0.005344 | 4.9e-05 | 0.001075 | 0.3235 | 1.3266 | 2.5476 | 100.79 | 21.11 | 15.511 | 128.6014 | 126.4901 | 115.6623 | 0.003858 | 0.90801 | 0.02202 | 24.2507 | 0.39722 | 0.1472 | 0.039299 | 2.5946 | 2.5946 | 0.002343
2019-03-26 16:21 | 1553631685.6723 | 506 | 0.20167 | 0.007063 | 0.000143 | 99.7688 | 95.5988 | 110.5233 | 21.2612 | 0.004423 | 0.00388 | 227.241 | 0.24726 | 8.4033 | 0.00432 | 0.005922 | 0.005908 | 4.8e-05 | 0.00107 | 0.2891 | 1.3237 | 2.5475 | 100.79 | 21.11 | 15.514 | 128.6213 | 126.0921 | 115.3419 | 0.003859 | 0.90876 | 0.022029 | 24.2403 | 0.38189 | 0.14709 | 0.038811 | 2.5637 | 2.5637 | 0.00237
2019-03-26 16:21 | 1553631686.6721 | 507 | 0.20181 | 0.007008 | 0.000139 | 99.682 | 95.5544 | 110.609 | 21.2612 | 0.004341 | 0.003876 | 227.1451 | 0.24722 | 8.4054 | 0.004289 | 0.005682 | 0.005568 | 4.6e-05 | 0.001071 | 0.29717 | 1.3407 | 2.5475 | 100.79 | 21.11 | 15.522 | 126.6072 | 125.328 | 115.2446 | 0.003861 | 0.90899 | 0.02204 | 24.2466 | 0.37825 | 0.14707 | 0.038054 | 2.5152 | 2.5152 | 0.002415
2019-03-26 16:21 | 1553631687.6721 | 508 | 0.20187 | 0.00694 | 0.000135 | 99.69 | 95.6671 | 110.6275 | 21.2612 | 0.004402 | 0.003892 | 227.0324 | 0.24772 | 8.4027 | 0.00426 | 0.00594 | 0.005828 | 5.1e-05 | 0.00107 | 0.31797 | 1.3518 | 2.5475 | 100.79 | 21.11 | 15.527 | 128.1177 | 124.6209 | 115.614 | 0.003862 | 0.90812 | 0.022031 | 24.2602 | 0.36843 | 0.14737 | 0.037732 | 2.4932 | 2.4932 | 0.002441
2019-03-26 16:21 | 1553631688.671 | 509 | 0.20202 | 0.006873 | 0.000132 | 99.6902 | 95.3823 | 110.9256 | 21.2612 | 0.004393 | 0.003867 | 226.9179 | 0.24721 | 8.4028 | 0.004289 | 0.005148 | 0.005181 | 4.6e-05 | 0.001069 | 0.36282 | 1.3321 | 2.5474 | 100.79 | 21.11 | 15.532 | 127.8777 | 125.3354 | 115.0183 | 0.003863 | 0.90952 | 0.022078 | 24.2742 | 0.39512 | 0.14707 | 0.036937 | 2.4461 | 2.4461 | 0.002483
2019-03-26 16:21 | 1553631689.6722 | 510 | 0.2022 | 0.006799 | 0.000127 | 99.6542 | 95.7682 | 110.959 | 21.2612 | 0.004377 | 0.003858 | 226.8018 | 0.24779 | 8.4011 | 0.004235 | 0.005729 | 0.005611 | 4.6e-05 | 0.001067 | 0.33157 | 1.4107 | 2.5475 | 100.79 | 21.11 | 15.544 | 127.4914 | 124.0055 | 114.8033 | 0.003866 | 0.91002 | 0.022087 | 24.2712 | 0.35577 | 0.14741 | 0.035997 | 2.3851 | 2.3851 | 0.002553
2019-03-26 16:21 | 1553631690.6712 | 511 | 0.20222 | 0.006719 | 0.000123 | 99.6767 | 95.9578 | 111.588 | 21.2612 | 0.004348 | 0.003852 | 226.7161 | 0.24793 | 8.4027 | 0.004252 | 0.004694 | 0.004548 | 4.8e-05 | 0.001066 | 0.31365 | 1.3919 | 2.5475 | 100.79 | 21.1 | 15.551 | 126.7746 | 124.4214 | 114.6444 | 0.003865 | 0.91039 | 0.022154 | 24.3349 | 0.3401 | 0.1475 | 0.035933 | 2.3883 | 2.3883 | 0.002551
2019-03-26 16:22 | 1553631691.6714 | 512 | 0.20231 | 0.006643 | 0.000118 | 99.6557 | 96.0883 | 111.7554 | 21.2612 | 0.004323 | 0.003821 | 226.6111 | 0.24784 | 8.4027 | 0.004322 | 0.006036 | 0.005953 | 4.8e-05 | 0.001066 | 0.32012 | 1.3847 | 2.5475 | 100.79 | 21.1 | 15.567 | 126.1606 | 126.1351 | 113.8942 | 0.003869 | 0.91216 | 0.022192 | 24.3295 | 0.32607 | 0.14744 | 0.035539 | 2.3665 | 2.3665 | 0.002573
2019-03-26 16:22 | 1553631692.6723 | 513 | 0.20247 | 0.006565 | 0.000114 | 99.6162 | 96.1481 | 111.7807 | 21.2612 | 0.004291 | 0.003788 | 226.5293 | 0.24737 | 8.4026 | 0.004167 | 0.005929 | 0.005839 | 4.5e-05 | 0.001064 | 0.33229 | 1.3803 | 2.5475 | 100.79 | 21.1 | 15.576 | 125.3734 | 122.3509 | 113.0928 | 0.003872 | 0.91405 | 0.022218 | 24.3071 | 0.31695 | 0.14716 | 0.034643 | 2.3098 | 2.3098 | 0.002631
2019-03-26 16:22 | 1553631693.6723 | 514 | 0.20248 | 0.006497 | 0.000109 | 99.6194 | 96.172 | 111.7803 | 21.2612 | 0.004304 | 0.00381 | 226.436 | 0.24726 | 8.4046 | 0.004271 | 0.006297 | 0.006297 | 4.7e-05 | 0.001061 | 0.31741 | 1.3373 | 2.5475 | 100.79 | 21.1 | 15.585 | 125.6893 | 124.8938 | 113.6245 | 0.003874 | 0.91279 | 0.022203 | 24.3238 | 0.31502 | 0.14709 | 0.034694 | 2.3118 | 2.3118 | 0.002628
2019-03-26 16:22 | 1553631694.6725 | 515 | 0.20253 | 0.006426 | 0.000105 | 99.5855 | 96.2083 | 111.8324 | 21.2612 | 0.004301 | 0.00384 | 226.3435 | 0.24787 | 8.4009 | 0.004198 | 0.005593 | 0.005581 | 4.8e-05 | 0.001063 | 0.32477 | 1.3322 | 2.5474 | 100.79 | 21.1 | 15.603 | 125.6376 | 123.1031 | 114.3517 | 0.003878 | 0.91108 | 0.022187 | 24.3523 | 0.30859 | 0.14746 | 0.034493 | 2.297 | 2.297 | 0.002651
2019-03-26 16:22 | 1553631695.6714 | 516 | 0.20256 | 0.006355 | 0.000102 | 99.6082 | 96.5881 | 111.9005 | 21.2612 | 0.004302 | 0.003868 | 226.2517 | 0.24713 | 8.4038 | 0.004231 | 0.005801 | 0.005742 | 4.7e-05 | 0.001055 | 0.30541 | 1.335 | 2.5475 | 100.79 | 21.1 | 15.613 | 125.6552 | 123.9045 | 115.0461 | 0.003881 | 0.90945 | 0.022174 | 24.3815 | 0.27539 | 0.14702 | 0.034367 | 2.2875 | 2.2875 | 0.002654
2019-03-26 16:22 | 1553631696.6718 | 517 | 0.20259 | 0.00629 | 0.0001 | 99.6176 | 96.6584 | 111.8451 | 21.2612 | 0.004293 | 0.003818 | 226.1456 | 0.2477 | 8.4012 | 0.004187 | 0.005679 | 0.005549 | 4.8e-05 | 0.001054 | 0.28296 | 1.3366 | 2.5474 | 100.79 | 21.1 | 15.63 | 125.4205 | 122.8379 | 113.8316 | 0.003885 | 0.91231 | 0.022203 | 24.3373 | 0.26973 | 0.14736 | 0.03431 | 2.2868 | 2.2868 | 0.002661
2019-03-26 16:22 | 1553631697.6726 | 518 | 0.20255 | 0.006236 | 9.8e-05 | 99.6286 | 96.3259 | 111.8681 | 21.2612 | 0.004304 | 0.003806 | 226.0566 | 0.2474 | 8.4021 | 0.00421 | 0.005734 | 0.005639 | 4.4e-05 | 0.001056 | 0.31421 | 1.3354 | 2.5474 | 100.79 | 21.1 | 15.638 | 125.695 | 123.387 | 113.5234 | 0.003887 | 0.91303 | 0.022214 | 24.3301 | 0.30153 | 0.14718 | 0.034579 | 2.306 | 2.306 | 0.002636
2019-03-26 16:22 | 1553631698.6717 | 519 | 0.2026 | 0.006185 | 9.7e-05 | 99.6153 | 96.5101 | 111.6697 | 21.2612 | 0.004249 | 0.003773 | 225.9594 | 0.24741 | 8.4016 | 0.004133 | 0.006235 | 0.006197 | 4.6e-05 | 0.001061 | 0.35275 | 1.3424 | 2.5475 | 100.79 | 21.1 | 15.645 | 124.3488 | 121.5075 | 112.7215 | 0.003889 | 0.91493 | 0.022218 | 24.2833 | 0.28325 | 0.14718 | 0.03438 | 2.2931 | 2.2931 | 0.002651
2019-03-26 16:22 | 1553631699.6727 | 520 | 0.20264 | 0.00614 | 9.7e-05 | 99.522 | 96.4755 | 111.4458 | 21.2612 | 0.004319 | 0.003811 | 225.8766 | 0.24644 | 8.4039 | 0.004228 | 0.006121 | 0.006066 | 4.4e-05 | 0.001059 | 0.30807 | 1.3437 | 2.5475 | 100.79 | 21.1 | 15.664 | 126.0691 | 123.8489 | 113.6442 | 0.003893 | 0.91275 | 0.022169 | 24.288 | 0.27808 | 0.14661 | 0.034209 | 2.2767 | 2.2767 | 0.00266
2019-03-26 16:22 | 1553631700.6719 | 521 | 0.20267 | 0.006108 | 9.6e-05 | 99.5598 | 96.289 | 111.0719 | 21.2612 | 0.004243 | 0.003817 | 225.7855 | 0.246 | 8.4025 | 0.004195 | 0.006033 | 0.005917 | 4.5e-05 | 0.001064 | 0.31454 | 1.3602 | 2.5474 | 100.79 | 21.1 | 15.672 | 124.2067 | 123.0213 | 113.7984 | 0.003895 | 0.91238 | 0.022127 | 24.252 | 0.29879 | 0.14634 | 0.034047 | 2.2617 | 2.2617 | 0.002672
2019-03-26 16:22 | 1553631701.6712 | 522 | 0.2027 | 0.006084 | 9.5e-05 | 99.5565 | 96.411 | 110.8491 | 21.2612 | 0.004279 | 0.003812 | 225.697 | 0.24879 | 8.3999 | 0.004171 | 0.005797 | 0.005705 | 4.4e-05 | 0.001066 | 0.30862 | 1.3485 | 2.5474 | 100.79 | 21.1 | 15.68 | 125.0965 | 122.4454 | 113.6671 | 0.003897 | 0.91269 | 0.022109 | 24.2236 | 0.28717 | 0.148 | 0.033864 | 2.2478 | 2.2478 | 0.002719
2019-03-26 16:22 | 1553631702.6719 | 523 | 0.20276 | 0.006052 | 9.4e-05 | 99.509 | 96.1773 | 110.8137 | 21.2612 | 0.004292 | 0.003795 | 225.5777 | 0.24819 | 8.4022 | 0.004221 | 0.006207 | 0.006122 | 4.4e-05 | 0.001068 | 0.3126 | 1.3985 | 2.5476 | 100.79 | 21.1 | 15.696 | 125.4156 | 123.6764 | 113.2485 | 0.003901 | 0.91368 | 0.022117 | 24.2066 | 0.3046 | 0.14765 | 0.033555 | 2.2282 | 2.2282 | 0.002737
2019-03-26 16:22 | 1553631703.6719 | 524 | 0.20277 | 0.006028 | 9.3e-05 | 99.5146 | 96.3155 | 111.0644 | 21.2612 | 0.004258 | 0.003881 | 225.4687 | 0.24691 | 8.4067 | 0.004261 | 0.00495 | 0.004967 | 4.6e-05 | 0.001069 | 0.26432 | 1.3754 | 2.5476 | 100.79 | 21.1 | 15.704 | 124.5635 | 124.6528 | 115.3658 | 0.003903 | 0.9087 | 0.022082 | 24.3003 | 0.29226 | 0.14689 | 0.033564 | 2.2252 | 2.2252 | 0.002726
2019-03-26 16:22 | 1553631704.6723 | 525 | 0.20277 | 0.005993 | 9.3e-05 | 99.5199 | 96.4722 | 111.2423 | 21.2612 | 0.004236 | 0.003778 | 225.4006 | 0.24784 | 8.4014 | 0.00411 | 0.005655 | 0.005507 | 4.5e-05 | 0.001075 | 0.33467 | 1.3512 | 2.5474 | 100.79 | 21.1 | 15.72 | 124.026 | 120.944 | 112.8525 | 0.003907 | 0.91462 | 0.022171 | 24.2409 | 0.27821 | 0.14744 | 0.033567 | 2.2345 | 2.2345 | 0.002725
2019-03-26 16:22 | 1553631705.6724 | 526 | 0.20287 | 0.005967 | 9.3e-05 | 99.5418 | 96.4577 | 111.4841 | 21.2612 | 0.004307 | 0.003842 | 225.3112 | 0.24825 | 8.4019 | 0.004139 | 0.005243 | 0.005175 | 4.5e-05 | 0.001073 | 0.31293 | 1.3472 | 2.5476 | 100.79 | 21.1 | 15.727 | 125.7673 | 121.6631 | 114.4166 | 0.003909 | 0.91093 | 0.02215 | 24.3164 | 0.28151 | 0.14769 | 0.033036 | 2.1971 | 2.1971 | 0.002776
2019-03-26 16:22 | 1553631706.671 | 527 | 0.20287 | 0.005929 | 9.4e-05 | 99.5037 | 96.4547 | 111.7805 | 21.2612 | 0.004217 | 0.003778 | 225.2376 | 0.24629 | 8.4032 | 0.004079 | 0.005415 | 0.005278 | 4.4e-05 | 0.00107 | 0.30718 | 1.3979 | 2.5475 | 100.79 | 21.1 | 15.735 | 123.5595 | 120.1917 | 112.8455 | 0.003911 | 0.91464 | 0.022225 | 24.2993 | 0.27837 | 0.14652 | 0.033043 | 2.2049 | 2.2049 | 0.002744
2019-03-26 16:22 | 1553631707.672 | 528 | 0.20287 | 0.005896 | 9.4e-05 | 99.5497 | 96.692 | 111.8436 | 21.2612 | 0.004224 | 0.003763 | 225.1492 | 0.24825 | 8.4004 | 0.004081 | 0.005861 | 0.005728 | 4.3e-05 | 0.00107 | 0.31006 | 1.4174 | 2.5477 | 100.79 | 21.1 | 15.75 | 123.7438 | 120.2394 | 112.4854 | 0.003915 | 0.91549 | 0.022242 | 24.2948 | 0.26053 | 0.14768 | 0.033096 | 2.2101 | 2.2101 | 0.00276
2019-03-26 16:22 | 1553631708.6711 | 529 | 0.20288 | 0.005877 | 9.4e-05 | 99.4243 | 96.1531 | 112.0251 | 21.2612 | 0.004199 | 0.003754 | 225.0522 | 0.24775 | 8.4029 | 0.004086 | 0.005178 | 0.005112 | 4.4e-05 | 0.001076 | 0.32709 | 1.3686 | 2.5476 | 100.79 | 21.1 | 15.758 | 123.1159 | 120.3526 | 112.2641 | 0.003917 | 0.91602 | 0.022266 | 24.3075 | 0.29923 | 0.14739 | 0.033093 | 2.2123 | 2.2123 | 0.002752
2019-03-26 16:22 | 1553631709.6722 | 530 | 0.20303 | 0.00585 | 9.4e-05 | 99.3824 | 96.1861 | 112.1861 | 21.2612 | 0.004167 | 0.003718 | 224.9598 | 0.24697 | 8.4003 | 0.003984 | 0.00543 | 0.005374 | 4.6e-05 | 0.001071 | 0.31404 | 1.3353 | 2.5475 | 100.79 | 21.1 | 15.776 | 122.3382 | 117.8704 | 111.3974 | 0.003921 | 0.91808 | 0.022307 | 24.2976 | 0.2924 | 0.14692 | 0.032235 | 2.1589 | 2.1589 | 0.002811
2019-03-26 16:22 | 1553631710.6725 | 531 | 0.20314 | 0.005824 | 9.3e-05 | 99.4547 | 96.1629 | 112.4985 | 21.2612 | 0.004229 | 0.003762 | 224.8674 | 0.24727 | 8.4001 | 0.003999 | 0.005604 | 0.005488 | 4.4e-05 | 0.001066 | 0.32676 | 1.3611 | 2.5474 | 100.79 | 21.1 | 15.784 | 123.8702 | 118.2453 | 112.4616 | 0.003923 | 0.91555 | 0.022307 | 24.3651 | 0.30106 | 0.1471 | 0.031556 | 2.1136 | 2.1136 | 0.002874
2019-03-26 16:22 | 1553631711.6713 | 532 | 0.20324 | 0.005786 | 9.3e-05 | 99.442 | 96.4849 | 112.6772 | 21.2612 | 0.004252 | 0.003693 | 224.8057 | 0.24765 | 8.4026 | 0.003947 | 0.005243 | 0.005139 | 4.9e-05 | 0.00107 | 0.29551 | 1.3576 | 2.5476 | 100.79 | 21.1 | 15.792 | 124.4363 | 116.9604 | 110.7747 | 0.003925 | 0.91957 | 0.022374 | 24.331 | 0.27002 | 0.14733 | 0.031039 | 2.0852 | 2.0852 | 0.002918
2019-03-26 16:22 | 1553631712.6723 | 533 | 0.20334 | 0.005738 | 9.2e-05 | 99.4104 | 96.2426 | 113.059 | 21.2612 | 0.004214 | 0.003753 | 224.7439 | 0.24733 | 8.4018 | 0.004025 | 0.005385 | 0.005333 | 4.7e-05 | 0.001067 | 0.3095 | 1.3492 | 2.5474 | 100.8 | 21.1 | 15.81 | 123.4983 | 118.8658 | 112.2476 | 0.003929 | 0.91605 | 0.022369 | 24.4189 | 0.28967 | 0.14714 | 0.030506 | 2.049 | 2.049 | 0.002966
2019-03-26 16:22 | 1553631713.6723 | 534 | 0.20356 | 0.00567 | 9.1e-05 | 99.3945 | 96.7226 | 113.018 | 21.2612 | 0.004112 | 0.003731 | 224.6689 | 0.24827 | 8.4003 | 0.003986 | 0.004955 | 0.004926 | 4.6e-05 | 0.00107 | 0.3079 | 1.3286 | 2.5472 | 100.8 | 21.1 | 15.817 | 121.0098 | 117.9276 | 111.7115 | 0.003931 | 0.91733 | 0.022381 | 24.3975 | 0.24373 | 0.1477 | 0.029229 | 1.9644 | 1.9644 | 0.003105
2019-03-26 16:22 | 1553631714.6727 | 535 | 0.20366 | 0.0056 | 9e-05 | 99.3555 | 96.7226 | 112.8452 | 21.2612 | 0.004203 | 0.003701 | 224.595 | 0.24836 | 8.4036 | 0.00389 | 0.006025 | 0.005889 | 4.8e-05 | 0.001075 | 0.30735 | 1.3566 | 2.5475 | 100.8 | 21.1 | 15.829 | 123.2232 | 115.5872 | 110.9656 | 0.003934 | 0.91911 | 0.022385 | 24.3552 | 0.24023 | 0.14775 | 0.028737 | 1.9317 | 1.9317 | 0.003159
2019-03-26 16:22 | 1553631715.6724 | 536 | 0.20379 | 0.005516 | 9e-05 | 99.3201 | 96.8715 | 112.8199 | 21.2612 | 0.004131 | 0.003732 | 224.5152 | 0.24817 | 8.402 | 0.003947 | 0.00582 | 0.005801 | 5e-05 | 0.001076 | 0.33467 | 1.3497 | 2.5475 | 100.8 | 21.1 | 15.834 | 121.4725 | 116.9667 | 111.7329 | 0.003935 | 0.91728 | 0.02236 | 24.3768 | 0.22328 | 0.14764 | 0.028056 | 1.8839 | 1.8839 | 0.003237
2019-03-26 16:22 | 1553631716.6715 | 537 | 0.20391 | 0.005427 | 8.9e-05 | 99.3239 | 96.6836 | 112.6342 | 21.2612 | 0.004183 | 0.003692 | 224.4447 | 0.24764 | 8.406 | 0.003907 | 0.0058 | 0.005604 | 4.7e-05 | 0.00108 | 0.31061 | 1.368 | 2.5476 | 100.8 | 21.1 | 15.839 | 122.7424 | 115.9986 | 110.7424 | 0.003937 | 0.91965 | 0.022371 | 24.3254 | 0.24099 | 0.14732 | 0.027463 | 1.8449 | 1.8449 | 0.003298
2019-03-26 16:22 | 1553631717.6719 | 538 | 0.20405 | 0.005344 | 8.9e-05 | 99.3118 | 96.8248 | 112.8339 | 21.2612 | 0.004156 | 0.003659 | 224.3915 | 0.24842 | 8.4035 | 0.003938 | 0.005449 | 0.005403 | 5e-05 | 0.001082 | 0.32228 | 1.3488 | 2.5476 | 100.8 | 21.1 | 15.846 | 122.0816 | 116.7356 | 109.9503 | 0.003938 | 0.92155 | 0.022414 | 24.3218 | 0.22685 | 0.14779 | 0.026701 | 1.7972 | 1.7972 | 0.003396
2019-03-26 16:22 | 1553631718.6706 | 539 | 0.20417 | 0.005255 | 8.8e-05 | 99.3278 | 96.8288 | 113.0504 | 21.2612 | 0.004161 | 0.003683 | 224.3254 | 0.24582 | 8.4009 | 0.003898 | 0.005277 | 0.005269 | 4.8e-05 | 0.001079 | 0.3314 | 1.3566 | 2.5475 | 100.8 | 21.11 | 15.849 | 122.1896 | 115.7779 | 110.5466 | 0.003941 | 0.92012 | 0.022418 | 24.364 | 0.22791 | 0.14624 | 0.026106 | 1.7575 | 1.7575 | 0.003436
2019-03-26 16:22 | 1553631719.671 | 540 | 0.20419 | 0.005168 | 8.9e-05 | 99.2619 | 96.7088 | 113.3999 | 21.2612 | 0.00406 | 0.003677 | 224.2704 | 0.24871 | 8.4025 | 0.00385 | 0.005419 | 0.005311 | 4.7e-05 | 0.001081 | 0.33035 | 1.3198 | 2.5475 | 100.8 | 21.11 | 15.851 | 119.7324 | 114.5945 | 110.3888 | 0.003942 | 0.92049 | 0.022457 | 24.3967 | 0.23307 | 0.14796 | 0.026098 | 1.7601 | 1.7601 | 0.003472
2019-03-26 16:22 | 1553631720.6721 | 541 | 0.20423 | 0.005088 | 8.8e-05 | 99.3068 | 96.6191 | 113.305 | 21.2612 | 0.004096 | 0.003687 | 224.2117 | 0.24674 | 8.4045 | 0.003862 | 0.005174 | 0.00509 | 4.8e-05 | 0.00108 | 0.33323 | 1.3487 | 2.5475 | 100.79 | 21.11 | 15.853 | 120.6011 | 114.8911 | 110.6422 | 0.003943 | 0.91989 | 0.02244 | 24.3945 | 0.24542 | 0.14679 | 0.025957 | 1.7492 | 1.7492 | 0.003466
2019-03-26 16:22 | 1553631721.6717 | 542 | 0.20446 | 0.005011 | 8.7e-05 | 99.2454 | 96.493 | 113.7085 | 21.2612 | 0.004082 | 0.00363 | 224.1653 | 0.24624 | 8.4031 | 0.003862 | 0.004864 | 0.004806 | 4.8e-05 | 0.001078 | 0.30237 | 1.3194 | 2.5476 | 100.79 | 21.11 | 15.852 | 120.2687 | 114.8862 | 109.2367 | 0.003943 | 0.92327 | 0.022521 | 24.3931 | 0.25157 | 0.14649 | 0.024683 | 1.6694 | 1.6694 | 0.003624
2019-03-26 16:22 | 1553631722.6718 | 543 | 0.20434 | 0.004937 | 8.6e-05 | 99.2478 | 96.211 | 113.8888 | 21.2612 | 0.0041 | 0.0036 | 224.1259 | 0.24806 | 8.4051 | 0.003826 | 0.00504 | 0.004916 | 4.6e-05 | 0.001083 | 0.31116 | 1.3456 | 2.5474 | 100.79 | 21.11 | 15.848 | 120.705 | 114.0221 | 108.5179 | 0.003942 | 0.92501 | 0.02256 | 24.3895 | 0.27796 | 0.14757 | 0.02548 | 1.7265 | 1.7265 | 0.00353
2019-03-26 16:22 | 1553631723.6718 | 544 | 0.20453 | 0.004856 | 8.4e-05 | 99.2513 | 96.253 | 114.1277 | 21.2612 | 0.004046 | 0.003624 | 224.091 | 0.24718 | 8.4031 | 0.003786 | 0.005931 | 0.005903 | 4.8e-05 | 0.001083 | 0.32228 | 1.3489 | 2.5475 | 100.79 | 21.11 | 15.843 | 119.3938 | 113.0334 | 109.1114 | 0.00394 | 0.92357 | 0.022567 | 24.434 | 0.27437 | 0.14705 | 0.024442 | 1.6566 | 1.6566 | 0.003666
2019-03-26 16:22 | 1553631724.6719 | 545 | 0.20497 | 0.004784 | 8.2e-05 | 99.2401 | 96.8177 | 114.5855 | 21.2612 | 0.004067 | 0.003578 | 224.0153 | 0.2467 | 8.4038 | 0.003786 | 0.005536 | 0.005462 | 4.8e-05 | 0.001085 | 0.32516 | 1.3597 | 2.5476 | 100.79 | 21.11 | 15.839 | 119.9005 | 113.0383 | 107.9829 | 0.003939 | 0.92631 | 0.022645 | 24.4468 | 0.22104 | 0.14676 | 0.021897 | 1.4893 | 1.4893 | 0.00407
2019-03-26 16:22 | 1553631725.6709 | 546 | 0.20511 | 0.004696 | 8e-05 | 99.236 | 97.1552 | 114.9974 | 21.2612 | 0.003955 | 0.003585 | 223.9764 | 0.24851 | 8.4055 | 0.003806 | 0.00573 | 0.005623 | 4.8e-05 | 0.00108 | 0.30038 | 1.3181 | 2.5475 | 100.79 | 21.11 | 15.83 | 117.1664 | 113.5249 | 108.1437 | 0.003937 | 0.92592 | 0.022681 | 24.4959 | 0.18954 | 0.14784 | 0.021177 | 1.4427 | 1.4427 | 0.004232
2019-03-26 16:22 | 1553631726.671 | 547 | 0.20521 | 0.004595 | 7.8e-05 | 99.2135 | 97.2631 | 115.3699 | 21.2612 | 0.004017 | 0.00356 | 223.923 | 0.24746 | 8.4062 | 0.003765 | 0.00444 | 0.004284 | 4.9e-05 | 0.001087 | 0.30574 | 1.3352 | 2.5474 | 100.79 | 21.11 | 15.824 | 118.6867 | 112.5185 | 107.5435 | 0.003936 | 0.92737 | 0.022736 | 24.5162 | 0.17759 | 0.14721 | 0.020683 | 1.4125 | 1.4125 | 0.004304
2019-03-26 16:22 | 1553631727.6714 | 548 | 0.20523 | 0.004495 | 7.6e-05 | 99.2613 | 97.5694 | 115.6053 | 21.2612 | 0.004046 | 0.003533 | 223.8597 | 0.24865 | 8.4019 | 0.00373 | 0.005194 | 0.005179 | 5.1e-05 | 0.001091 | 0.347 | 1.3197 | 2.5474 | 100.79 | 21.11 | 15.813 | 119.3914 | 111.6752 | 106.9008 | 0.003933 | 0.92894 | 0.022778 | 24.5205 | 0.15378 | 0.14792 | 0.020671 | 1.4145 | 1.4145 | 0.004319
2019-03-26 16:22 | 1553631728.6713 | 549 | 0.20531 | 0.004389 | 7.5e-05 | 99.1979 | 97.6397 | 115.6661 | 21.2612 | 0.003998 | 0.003585 | 223.822 | 0.24766 | 8.4029 | 0.003781 | 0.004773 | 0.00466 | 4.7e-05 | 0.001091 | 0.32101 | 1.3313 | 2.5474 | 100.79 | 21.11 | 15.808 | 118.2087 | 112.9246 | 108.1626 | 0.003932 | 0.92587 | 0.022746 | 24.5676 | 0.14162 | 0.14733 | 0.020329 | 1.3892 | 1.3892 | 0.00438
2019-03-26 16:22 | 1553631729.6721 | 550 | 0.20541 | 0.004289 | 7.4e-05 | 99.0984 | 97.682 | 115.6384 | 21.2612 | 0.003988 | 0.003541 | 223.7937 | 0.24842 | 8.4029 | 0.00375 | 0.004757 | 0.004712 | 4.8e-05 | 0.001096 | 0.31542 | 1.3123 | 2.5473 | 100.79 | 21.11 | 15.802 | 117.9735 | 112.1678 | 107.0748 | 0.00393 | 0.92852 | 0.022776 | 24.5296 | 0.12876 | 0.14778 | 0.019901 | 1.3618 | 1.3618 | 0.004482
2019-03-26 16:22 | 1553631730.6726 | 551 | 0.20553 | 0.004193 | 7.3e-05 | 99.1447 | 97.7392 | 116.1224 | 21.2612 | 0.004021 | 0.00353 | 223.7606 | 0.24764 | 8.4026 | 0.003732 | 0.005427 | 0.005303 | 4.9e-05 | 0.001096 | 0.33754 | 1.3382 | 2.5475 | 100.79 | 21.11 | 15.791 | 118.7627 | 111.7234 | 106.819 | 0.003927 | 0.92914 | 0.022831 | 24.5726 | 0.12772 | 0.14732 | 0.019275 | 1.3221 | 1.3221 | 0.004602
2019-03-26 16:22 | 1553631731.6716 | 552 | 0.2057 | 0.004099 | 7.1e-05 | 99.2041 | 97.6524 | 116.1583 | 21.2612 | 0.00397 | 0.003501 | 223.6834 | 0.2476 | 8.402 | 0.003712 | 0.005135 | 0.005056 | 5.1e-05 | 0.001096 | 0.34899 | 1.3502 | 2.5475 | 100.79 | 21.11 | 15.786 | 117.5218 | 111.2466 | 106.1246 | 0.003926 | 0.93084 | 0.022856 | 24.554 | 0.14101 | 0.1473 | 0.018391 | 1.2629 | 1.2629 | 0.004817
2019-03-26 16:22 | 1553631732.6714 | 553 | 0.20587 | 0.004008 | 7e-05 | 99.0178 | 97.6363 | 116.3982 | 21.2612 | 0.003925 | 0.003533 | 223.6327 | 0.24853 | 8.405 | 0.00371 | 0.005433 | 0.005434 | 5.1e-05 | 0.001094 | 0.32493 | 1.3448 | 2.5474 | 100.79 | 21.11 | 15.775 | 116.4355 | 111.2023 | 106.8882 | 0.003923 | 0.92897 | 0.022856 | 24.604 | 0.12568 | 0.14785 | 0.017486 | 1.2007 | 1.2007 | 0.005086
2019-03-26 16:22 | 1553631733.6717 | 554 | 0.20587 | 0.003919 | 7e-05 | 99.058 | 97.6581 | 116.6792 | 21.2612 | 0.003927 | 0.003495 | 223.5832 | 0.24784 | 8.4017 | 0.003645 | 0.004679 | 0.004559 | 4.8e-05 | 0.0011 | 0.28783 | 1.3492 | 2.5476 | 100.79 | 21.11 | 15.771 | 116.4743 | 109.601 | 105.9619 | 0.003922 | 0.93124 | 0.022912 | 24.6037 | 0.1273 | 0.14744 | 0.017564 | 1.2091 | 1.2091 | 0.005036
2019-03-26 16:22 | 1553631734.6714 | 555 | 0.20596 | 0.003826 | 7e-05 | 99.0346 | 97.756 | 116.9005 | 21.2612 | 0.003927 | 0.003518 | 223.5399 | 0.24702 | 8.4025 | 0.003722 | 0.004633 | 0.004506 | 5e-05 | 0.001096 | 0.33644 | 1.3591 | 2.5474 | 100.79 | 21.11 | 15.766 | 116.4665 | 111.485 | 106.5174 | 0.003921 | 0.92988 | 0.022917 | 24.645 | 0.11624 | 0.14695 | 0.017142 | 1.1801 | 1.1801 | 0.005143
2019-03-26 16:22 | 1553631735.6708 | 556 | 0.20596 | 0.003741 | 6.9e-05 | 99.0022 | 97.804 | 117.1402 | 21.2612 | 0.003895 | 0.003467 | 223.5106 | 0.24724 | 8.4017 | 0.003649 | 0.005291 | 0.005186 | 5.3e-05 | 0.001096 | 0.32477 | 1.3414 | 2.5475 | 100.79 | 21.11 | 15.761 | 115.7091 | 109.698 | 105.2843 | 0.00392 | 0.93291 | 0.022978 | 24.6302 | 0.10891 | 0.14709 | 0.01723 | 1.1894 | 1.1894 | 0.005107
2019-03-26 16:22 | 1553631736.6719 | 557 | 0.20603 | 0.003661 | 6.9e-05 | 99.0341 | 97.6776 | 117.2585 | 21.2612 | 0.003925 | 0.003495 | 223.475 | 0.24655 | 8.3999 | 0.003667 | 0.00492 | 0.004835 | 5.4e-05 | 0.001096 | 0.32211 | 1.3777 | 2.5476 | 100.79 | 21.11 | 15.757 | 116.426 | 110.1365 | 105.9721 | 0.003919 | 0.93122 | 0.022969 | 24.665 | 0.12336 | 0.14667 | 0.016936 | 1.1686 | 1.1686 | 0.005183
2019-03-26 16:22 | 1553631737.6709 | 558 | 0.20613 | 0.003579 | 6.9e-05 | 98.9892 | 97.8733 | 117.2993 | 21.2612 | 0.003871 | 0.003458 | 223.4561 | 0.24827 | 8.4039 | 0.003685 | 0.005252 | 0.005233 | 5e-05 | 0.001091 | 0.29048 | 1.3446 | 2.5475 | 100.79 | 21.11 | 15.755 | 115.1078 | 110.5768 | 105.0821 | 0.003918 | 0.93341 | 0.023 | 24.6403 | 0.1014 | 0.14769 | 0.016454 | 1.1369 | 1.1369 | 0.005365
2019-03-26 16:22 | 1553631738.6716 | 559 | 0.20618 | 0.003503 | 6.8e-05 | 98.9606 | 97.8642 | 117.5226 | 21.2612 | 0.003873 | 0.003426 | 223.4236 | 0.24782 | 8.401 | 0.00368 | 0.005874 | 0.005715 | 5.5e-05 | 0.001095 | 0.27527 | 1.3547 | 2.5474 | 100.79 | 21.11 | 15.752 | 115.1688 | 110.4584 | 104.3067 | 0.003918 | 0.93533 | 0.023045 | 24.6385 | 0.099646 | 0.14743 | 0.016243 | 1.1246 | 1.1246 | 0.005414
2019-03-26 16:22 | 1553631739.6706 | 560 | 0.20632 | 0.00343 | 6.8e-05 | 98.978 | 98.0659 | 118.0547 | 21.2612 | 0.003878 | 0.003446 | 223.3531 | 0.24828 | 8.402 | 0.003764 | 0.005518 | 0.005479 | 4.9e-05 | 0.001102 | 0.32964 | 1.3655 | 2.5475 | 100.79 | 21.11 | 15.752 | 115.2746 | 112.4995 | 104.7881 | 0.003918 | 0.93414 | 0.023082 | 24.7099 | 0.082804 | 0.1477 | 0.015453 | 1.0714 | 1.0714 | 0.005693
2019-03-26 16:22 | 1553631740.6711 | 561 | 0.2064 | 0.003354 | 6.8e-05 | 98.9892 | 98.2199 | 117.9228 | 21.2612 | 0.003829 | 0.003444 | 223.3024 | 0.24765 | 8.4014 | 0.003694 | 0.005184 | 0.005122 | 5.2e-05 | 0.001106 | 0.33124 | 1.3474 | 2.5475 | 100.79 | 21.11 | 15.75 | 114.0805 | 110.8 | 104.7393 | 0.003917 | 0.93426 | 0.023071 | 24.6945 | 0.069786 | 0.14733 | 0.015094 | 1.046 | 1.046 | 0.005817
2019-03-26 16:22 | 1553631741.6708 | 562 | 0.20651 | 0.003281 | 6.8e-05 | 98.9809 | 98.0952 | 118.0352 | 21.2612 | 0.003857 | 0.00343 | 223.2666 | 0.24731 | 8.4023 | 0.003673 | 0.004911 | 0.004925 | 5.3e-05 | 0.001111 | 0.32587 | 1.3299 | 2.5475 | 100.79 | 21.11 | 15.748 | 114.7773 | 110.2967 | 104.3879 | 0.003917 | 0.93513 | 0.023093 | 24.6948 | 0.080404 | 0.14713 | 0.014504 | 1.006 | 1.006 | 0.00604
2019-03-26 16:22 | 1553631742.6712 | 563 | 0.20657 | 0.003211 | 6.9e-05 | 98.9461 | 98.1059 | 117.8474 | 21.2612 | 0.00379 | 0.00335 | 223.2323 | 0.24765 | 8.4016 | 0.003704 | 0.004639 | 0.004578 | 5.4e-05 | 0.001114 | 0.30879 | 1.3338 | 2.5475 | 100.79 | 21.11 | 15.747 | 113.1274 | 111.0476 | 102.4498 | 0.003916 | 0.93995 | 0.023134 | 24.6117 | 0.076275 | 0.14733 | 0.014218 | 0.98784 | 0.98784 | 0.006159
2019-03-26 16:22 | 1553631743.6718 | 564 | 0.20658 | 0.00314 | 7e-05 | 98.9626 | 98.0915 | 117.854 | 21.2612 | 0.003827 | 0.003347 | 223.2085 | 0.24917 | 8.4039 | 0.003689 | 0.005281 | 0.00517 | 5.8e-05 | 0.001122 | 0.31669 | 1.3212 | 2.5475 | 100.79 | 21.11 | 15.744 | 114.0301 | 110.6903 | 102.375 | 0.003916 | 0.94014 | 0.023137 | 24.61 | 0.079087 | 0.14823 | 0.014271 | 0.9916 | 0.9916 | 0.006174
2019-03-26 16:22 | 1553631744.6719 | 565 | 0.20668 | 0.003065 | 7.1e-05 | 98.8785 | 98.1802 | 117.8831 | 21.2612 | 0.003726 | 0.003366 | 223.1925 | 0.24737 | 8.4052 | 0.003586 | 0.005306 | 0.005242 | 5.4e-05 | 0.001123 | 0.3501 | 1.3224 | 2.5473 | 100.79 | 21.11 | 15.743 | 111.5845 | 108.184 | 102.8356 | 0.003915 | 0.93899 | 0.023125 | 24.6281 | 0.063391 | 0.14716 | 0.013779 | 0.95676 | 0.95676 | 0.006352
2019-03-26 16:22 | 1553631745.672 | 566 | 0.20671 | 0.003004 | 7.2e-05 | 98.9263 | 98.1815 | 117.6871 | 21.2612 | 0.003767 | 0.00336 | 223.1442 | 0.24848 | 8.4032 | 0.003638 | 0.00526 | 0.005232 | 5.5e-05 | 0.001124 | 0.3194 | 1.3238 | 2.5474 | 100.79 | 21.11 | 15.741 | 112.5879 | 109.4387 | 102.7089 | 0.003915 | 0.9393 | 0.02311 | 24.6035 | 0.067594 | 0.14782 | 0.013674 | 0.94881 | 0.94881 | 0.006434
2019-03-26 16:22 | 1553631746.672 | 567 | 0.20671 | 0.002943 | 7.3e-05 | 98.9732 | 98.2918 | 117.5735 | 21.2612 | 0.003767 | 0.003369 | 223.1158 | 0.24798 | 8.4048 | 0.003594 | 0.005111 | 0.005006 | 5.4e-05 | 0.001125 | 0.31205 | 1.3457 | 2.5476 | 100.79 | 21.11 | 15.738 | 112.5786 | 108.3676 | 102.9227 | 0.003914 | 0.93877 | 0.023092 | 24.5986 | 0.061797 | 0.14752 | 0.013704 | 0.95012 | 0.95012 | 0.006413
2019-03-26 16:22 | 1553631747.6721 | 568 | 0.20673 | 0.002886 | 7.4e-05 | 98.8564 | 98.1328 | 117.4204 | 21.2612 | 0.003761 | 0.003345 | 223.104 | 0.24777 | 8.4025 | 0.003518 | 0.004919 | 0.004761 | 5.4e-05 | 0.001122 | 0.32189 | 1.3561 | 2.5475 | 100.79 | 21.11 | 15.736 | 112.4214 | 106.5206 | 102.3375 | 0.003914 | 0.94023 | 0.023095 | 24.5634 | 0.065708 | 0.1474 | 0.013677 | 0.94825 | 0.94825 | 0.00642
2019-03-26 16:22 | 1553631748.6721 | 569 | 0.20682 | 0.002836 | 7.5e-05 | 98.8732 | 98.3187 | 117.3845 | 21.2612 | 0.003694 | 0.003348 | 223.081 | 0.24864 | 8.4014 | 0.003563 | 0.004424 | 0.004356 | 5.2e-05 | 0.001124 | 0.33771 | 1.3895 | 2.5475 | 100.79 | 21.11 | 15.732 | 110.8141 | 107.6207 | 102.4041 | 0.003913 | 0.94007 | 0.02309 | 24.5619 | 0.050302 | 0.14792 | 0.013209 | 0.91545 | 0.91545 | 0.006673
2019-03-26 16:22 | 1553631749.6701 | 570 | 0.2069 | 0.002783 | 7.6e-05 | 98.8398 | 98.2609 | 117.0134 | 21.2612 | 0.003733 | 0.003347 | 223.0579 | 0.24675 | 8.4016 | 0.003615 | 0.005703 | 0.005586 | 5.5e-05 | 0.001126 | 0.32693 | 1.3924 | 2.5475 | 100.79 | 21.11 | 15.728 | 111.7431 | 108.8837 | 102.3854 | 0.003912 | 0.94011 | 0.023054 | 24.5224 | 0.052545 | 0.14679 | 0.012783 | 0.88437 | 0.88437 | 0.006855
2019-03-26 16:22 | 1553631750.6702 | 571 | 0.20702 | 0.002735 | 7.7e-05 | 98.8404 | 98.2589 | 117.0378 | 21.2612 | 0.003688 | 0.003351 | 223.0346 | 0.24894 | 8.4001 | 0.003545 | 0.0041 | 0.004101 | 5.4e-05 | 0.001122 | 0.33964 | 1.3578 | 2.5475 | 100.79 | 21.11 | 15.724 | 110.6634 | 107.1716 | 102.4891 | 0.003911 | 0.93985 | 0.023053 | 24.5283 | 0.052783 | 0.14809 | 0.012113 | 0.83772 | 0.83772 | 0.007301
2019-03-26 16:22 | 1553631751.6706 | 572 | 0.20708 | 0.002688 | 7.8e-05 | 98.8327 | 98.2935 | 116.9827 | 21.2612 | 0.003716 | 0.003352 | 223.0108 | 0.2473 | 8.4013 | 0.003586 | 0.004996 | 0.004863 | 5.2e-05 | 0.001125 | 0.32029 | 1.3611 | 2.5475 | 100.79 | 21.11 | 15.718 | 111.3475 | 108.176 | 102.5055 | 0.003909 | 0.93981 | 0.023047 | 24.5231 | 0.048937 | 0.14712 | 0.011805 | 0.81614 | 0.81614 | 0.007445
2019-03-26 16:23 | 1553631752.6716 | 573 | 0.20714 | 0.002641 | 7.8e-05 | 98.8254 | 98.3564 | 116.9695 | 21.2612 | 0.00374 | 0.003314 | 222.9953 | 0.24715 | 8.3995 | 0.003582 | 0.005103 | 0.004982 | 5.3e-05 | 0.001129 | 0.32731 | 1.3705 | 2.5475 | 100.79 | 21.11 | 15.715 | 111.9176 | 108.0766 | 101.5758 | 0.003909 | 0.94214 | 0.023074 | 24.4913 | 0.04255 | 0.14703 | 0.011486 | 0.79488 | 0.79488 | 0.007639
2019-03-26 16:23 | 1553631753.6704 | 574 | 0.20716 | 0.00259 | 7.9e-05 | 98.8227 | 98.2316 | 116.9543 | 21.2612 | 0.003586 | 0.003302 | 222.9538 | 0.24729 | 8.3988 | 0.003572 | 0.005108 | 0.005043 | 5.3e-05 | 0.001126 | 0.29983 | 1.3698 | 2.5476 | 100.79 | 21.11 | 15.71 | 108.1816 | 107.8419 | 101.2921 | 0.003907 | 0.94286 | 0.023082 | 24.4805 | 0.053659 | 0.14711 | 0.011468 | 0.79386 | 0.79386 | 0.007654
2019-03-26 16:23 | 1553631754.6714 | 575 | 0.20719 | 0.002534 | 8e-05 | 98.8301 | 98.1876 | 116.7876 | 21.2612 | 0.003651 | 0.003279 | 222.9301 | 0.24888 | 8.4023 | 0.003555 | 0.005212 | 0.005079 | 5.5e-05 | 0.001121 | 0.31902 | 1.3592 | 2.5475 | 100.79 | 21.11 | 15.707 | 109.7563 | 107.4238 | 100.727 | 0.003907 | 0.94428 | 0.023083 | 24.4446 | 0.058337 | 0.14806 | 0.011335 | 0.78458 | 0.78458 | 0.007794
2019-03-26 16:23 | 1553631755.6718 | 576 | 0.20725 | 0.002481 | 8.1e-05 | 98.7994 | 98.3086 | 116.5841 | 21.2612 | 0.003649 | 0.003315 | 222.9084 | 0.24794 | 8.4027 | 0.003541 | 0.005569 | 0.005577 | 5.3e-05 | 0.001119 | 0.3246 | 1.3488 | 2.5476 | 100.79 | 21.11 | 15.705 | 109.7177 | 107.0852 | 101.6164 | 0.003906 | 0.94204 | 0.023035 | 24.4523 | 0.044541 | 0.1475 | 0.011047 | 0.76291 | 0.76291 | 0.007985
2019-03-26 16:23 | 1553631756.6705 | 577 | 0.20728 | 0.002435 | 8.2e-05 | 98.7746 | 98.1614 | 116.7292 | 21.2612 | 0.003677 | 0.00327 | 222.877 | 0.24801 | 8.4025 | 0.003468 | 0.005187 | 0.005054 | 5.4e-05 | 0.001117 | 0.3282 | 1.3124 | 2.5475 | 100.79 | 21.11 | 15.7 | 110.3811 | 105.3134 | 100.5098 | 0.003905 | 0.94483 | 0.023083 | 24.4313 | 0.055702 | 0.14754 | 0.010941 | 0.75707 | 0.75707 | 0.008049
2019-03-26 16:23 | 1553631757.6718 | 578 | 0.20732 | 0.002389 | 8.2e-05 | 98.782 | 98.0824 | 116.9655 | 21.2612 | 0.003652 | 0.003233 | 222.8224 | 0.24823 | 8.4025 | 0.003493 | 0.00541 | 0.005336 | 5.2e-05 | 0.001114 | 0.32173 | 1.3189 | 2.5474 | 100.79 | 21.11 | 15.696 | 109.7714 | 105.9192 | 99.6278 | 0.003904 | 0.94707 | 0.023134 | 24.4272 | 0.063572 | 0.14767 | 0.010722 | 0.74347 | 0.74347 | 0.008203
2019-03-26 16:23 | 1553631758.6706 | 579 | 0.20738 | 0.002347 | 8.4e-05 | 98.8265 | 98.275 | 116.6475 | 21.2612 | 0.003634 | 0.003172 | 222.8009 | 0.24769 | 8.4026 | 0.003573 | 0.005019 | 0.004931 | 4.9e-05 | 0.001117 | 0.33035 | 1.3745 | 2.5475 | 100.79 | 21.11 | 15.689 | 109.3347 | 107.8744 | 98.1394 | 0.003902 | 0.95086 | 0.023149 | 24.3452 | 0.050058 | 0.14735 | 0.010435 | 0.72382 | 0.72382 | 0.008408
2019-03-26 16:23 | 1553631759.6709 | 580 | 0.2074 | 0.002307 | 8.5e-05 | 98.8286 | 98.3241 | 116.8233 | 21.2612 | 0.003613 | 0.003259 | 222.7849 | 0.2464 | 8.4018 | 0.003499 | 0.004764 | 0.004697 | 5.1e-05 | 0.001112 | 0.33699 | 1.3778 | 2.5474 | 100.79 | 21.11 | 15.687 | 108.8325 | 106.0562 | 100.2626 | 0.003902 | 0.94546 | 0.0231 | 24.4331 | 0.045779 | 0.14658 | 0.01035 | 0.71631 | 0.71631 | 0.008452
2019-03-26 16:23 | 1553631760.6717 | 581 | 0.20742 | 0.002272 | 8.6e-05 | 98.7796 | 98.3779 | 116.6696 | 21.2612 | 0.003546 | 0.003233 | 222.7593 | 0.24948 | 8.4046 | 0.003508 | 0.004806 | 0.004704 | 5.5e-05 | 0.001117 | 0.33411 | 1.4295 | 2.5475 | 100.79 | 21.11 | 15.683 | 107.2063 | 106.2869 | 99.6215 | 0.003901 | 0.94708 | 0.023105 | 24.3961 | 0.036452 | 0.14842 | 0.010248 | 0.70928 | 0.70928 | 0.008642
2019-03-26 16:23 | 1553631761.6712 | 582 | 0.20746 | 0.002233 | 8.7e-05 | 98.722 | 98.1647 | 116.7571 | 21.2612 | 0.003554 | 0.00318 | 222.7312 | 0.24666 | 8.4025 | 0.003488 | 0.005394 | 0.005389 | 5.4e-05 | 0.001111 | 0.32604 | 1.3992 | 2.5475 | 100.79 | 21.11 | 15.677 | 107.4105 | 105.8052 | 98.3347 | 0.003899 | 0.95036 | 0.023154 | 24.363 | 0.050637 | 0.14674 | 0.010096 | 0.70011 | 0.70011 | 0.008656
2019-03-26 16:23 | 1553631762.6709 | 583 | 0.20747 | 0.002201 | 8.9e-05 | 98.7675 | 98.2535 | 116.8975 | 21.2612 | 0.00353 | 0.003204 | 222.7018 | 0.24763 | 8.4019 | 0.003551 | 0.00466 | 0.004497 | 5.1e-05 | 0.001105 | 0.31924 | 1.3918 | 2.5475 | 100.79 | 21.11 | 15.674 | 106.8197 | 107.3253 | 98.9212 | 0.003898 | 0.94886 | 0.023149 | 24.3969 | 0.046672 | 0.14732 | 0.010085 | 0.69915 | 0.69915 | 0.008702
2019-03-26 16:23 | 1553631763.6718 | 584 | 0.20747 | 0.002163 | 9e-05 | 98.7073 | 98.1876 | 116.9369 | 21.2612 | 0.00351 | 0.003122 | 222.6869 | 0.24943 | 8.405 | 0.003569 | 0.004418 | 0.004373 | 5.1e-05 | 0.001101 | 0.32908 | 1.3774 | 2.5475 | 100.79 | 21.11 | 15.668 | 106.3381 | 107.7656 | 96.9324 | 0.003897 | 0.95396 | 0.023215 | 24.3357 | 0.047217 | 0.14839 | 0.01008 | 0.70073 | 0.70073 | 0.008746
2019-03-26 16:23 | 1553631764.6698 | 585 | 0.20749 | 0.002135 | 9.2e-05 | 98.7522 | 98.154 | 117.3104 | 21.2612 | 0.003521 | 0.003193 | 222.6769 | 0.24675 | 8.4043 | 0.003597 | 0.004614 | 0.004506 | 4.8e-05 | 0.001099 | 0.33323 | 1.327 | 2.5474 | 100.79 | 21.11 | 15.664 | 106.6064 | 108.4392 | 98.6669 | 0.003896 | 0.94951 | 0.023198 | 24.4316 | 0.054345 | 0.14679 | 0.010027 | 0.69641 | 0.69641 | 0.008705
2019-03-26 16:23 | 1553631765.6703 | 586 | 0.20751 | 0.002106 | 9.3e-05 | 98.7235 | 98.2716 | 117.854 | 21.2612 | 0.003492 | 0.003169 | 222.6467 | 0.24847 | 8.4025 | 0.003546 | 0.005263 | 0.005128 | 5.1e-05 | 0.001092 | 0.34683 | 1.3627 | 2.5474 | 100.79 | 21.11 | 15.661 | 105.8964 | 107.1969 | 98.0735 | 0.003895 | 0.95103 | 0.02327 | 24.4686 | 0.041035 | 0.14781 | 0.009915 | 0.69062 | 0.69062 | 0.008839
2019-03-26 16:23 | 1553631766.6711 | 587 | 0.20754 | 0.002078 | 9.5e-05 | 98.7191 | 98.1462 | 118.1108 | 21.2612 | 0.003542 | 0.00317 | 222.6071 | 0.24645 | 8.4049 | 0.003572 | 0.005018 | 0.004878 | 5.6e-05 | 0.00109 | 0.29734 | 1.3646 | 2.5476 | 100.8 | 21.11 | 15.654 | 107.1165 | 107.8263 | 98.0923 | 0.003893 | 0.95098 | 0.023295 | 24.4959 | 0.052054 | 0.14661 | 0.009765 | 0.6807 | 0.6807 | 0.008896
2019-03-26 16:23 | 1553631767.6704 | 588 | 0.20758 | 0.002053 | 9.6e-05 | 98.7409 | 98.2286 | 118.1554 | 21.2612 | 0.003476 | 0.003154 | 222.5878 | 0.24857 | 8.4043 | 0.003563 | 0.005873 | 0.005863 | 5.1e-05 | 0.001087 | 0.33771 | 1.3709 | 2.5475 | 100.8 | 21.11 | 15.651 | 105.4992 | 107.6168 | 97.7123 | 0.003892 | 0.95196 | 0.023312 | 24.488 | 0.04653 | 0.14788 | 0.009595 | 0.66913 | 0.66913 | 0.009127
2019-03-26 16:23 | 1553631768.67 | 589 | 0.2076 | 0.002041 | 9.8e-05 | 98.7129 | 98.192 | 118.8041 | 21.2612 | 0.00349 | 0.003135 | 222.5412 | 0.24713 | 8.4043 | 0.003527 | 0.005498 | 0.005418 | 5.1e-05 | 0.00108 | 0.30757 | 1.3515 | 2.5474 | 100.8 | 21.11 | 15.649 | 105.854 | 106.7418 | 97.2521 | 0.003892 | 0.95314 | 0.02339 | 24.5399 | 0.047327 | 0.14702 | 0.009503 | 0.66488 | 0.66488 | 0.009132
2019-03-26 16:23 | 1553631769.67 | 590 | 0.20766 | 0.002011 | 0.0001 | 98.6704 | 98.1116 | 119.0274 | 21.2612 | 0.003457 | 0.003158 | 222.5222 | 0.24943 | 8.4029 | 0.003627 | 0.004757 | 0.004715 | 5.3e-05 | 0.00108 | 0.3183 | 1.3716 | 2.5474 | 100.8 | 21.12 | 15.643 | 105.0477 | 109.1619 | 97.8077 | 0.003893 | 0.95171 | 0.023394 | 24.5813 | 0.050796 | 0.14838 | 0.009156 | 0.64037 | 0.64037 | 0.00957
2019-03-26 16:23 | 1553631770.6714 | 591 | 0.2077 | 0.00199 | 0.000101 | 98.6784 | 98.3073 | 119.5217 | 21.2612 | 0.003472 | 0.003127 | 222.5239 | 0.24799 | 8.4027 | 0.003533 | 0.005295 | 0.005188 | 5.4e-05 | 0.001076 | 0.31133 | 1.3531 | 2.5474 | 100.8 | 21.12 | 15.641 | 105.408 | 106.8819 | 97.0681 | 0.003892 | 0.95361 | 0.023466 | 24.6077 | 0.0337 | 0.14753 | 0.00891 | 0.6249 | 0.6249 | 0.00975
2019-03-26 16:23 | 1553631771.6712 | 592 | 0.20772 | 0.001967 | 0.000103 | 98.6143 | 98.2276 | 119.551 | 21.2612 | 0.003417 | 0.003119 | 222.5021 | 0.24812 | 8.4 | 0.003559 | 0.004936 | 0.004722 | 4.9e-05 | 0.001076 | 0.36497 | 1.357 | 2.5474 | 100.8 | 21.12 | 15.635 | 104.0867 | 107.5105 | 96.865 | 0.003891 | 0.95414 | 0.023476 | 24.604 | 0.035143 | 0.1476 | 0.008862 | 0.62166 | 0.62166 | 0.009806
2019-03-26 16:23 | 1553631772.6702 | 593 | 0.20773 | 0.00194 | 0.000104 | 98.733 | 98.1657 | 119.697 | 21.2612 | 0.003412 | 0.003121 | 222.4655 | 0.24799 | 8.4041 | 0.003646 | 0.00484 | 0.004844 | 5e-05 | 0.00108 | 0.30093 | 1.389 | 2.5475 | 100.8 | 21.12 | 15.633 | 103.968 | 109.6256 | 96.9082 | 0.00389 | 0.95403 | 0.023489 | 24.6205 | 0.051534 | 0.14753 | 0.008795 | 0.61711 | 0.61711 | 0.009874
2019-03-26 16:23 | 1553631773.6714 | 594 | 0.20777 | 0.001914 | 0.000106 | 98.7114 | 98.3005 | 119.6956 | 21.2612 | 0.00338 | 0.003112 | 222.4274 | 0.24853 | 8.4026 | 0.003595 | 0.005055 | 0.004998 | 5.1e-05 | 0.00108 | 0.3569 | 1.3455 | 2.5474 | 100.8 | 21.12 | 15.63 | 103.1923 | 108.3926 | 96.6897 | 0.003889 | 0.95459 | 0.023495 | 24.613 | 0.037308 | 0.14785 | 0.008621 | 0.60491 | 0.60491 | 0.010094
2019-03-26 16:23 | 1553631774.6703 | 595 | 0.2078 | 0.001888 | 0.000108 | 98.6792 | 98.1136 | 119.5172 | 21.2612 | 0.003423 | 0.003091 | 222.4085 | 0.24829 | 8.4047 | 0.003557 | 0.004492 | 0.004369 | 5.4e-05 | 0.001075 | 0.3298 | 1.3611 | 2.5474 | 100.8 | 21.12 | 15.626 | 104.2202 | 107.4687 | 96.1938 | 0.003888 | 0.95587 | 0.023494 | 24.5782 | 0.051415 | 0.14771 | 0.008417 | 0.5903 | 0.5903 | 0.010334
2019-03-26 16:23 | 1553631775.6703 | 596 | 0.20784 | 0.001861 | 0.000109 | 98.6789 | 98.1066 | 119.1698 | 21.2612 | 0.003398 | 0.003054 | 222.4009 | 0.24799 | 8.4024 | 0.003562 | 0.005244 | 0.005112 | 5.1e-05 | 0.00107 | 0.30718 | 1.374 | 2.5474 | 100.8 | 21.12 | 15.623 | 103.6268 | 107.5835 | 95.3065 | 0.003888 | 0.95817 | 0.023488 | 24.513 | 0.052032 | 0.14753 | 0.00824 | 0.57752 | 0.57752 | 0.01055
2019-03-26 16:23 | 1553631776.6706 | 597 | 0.20786 | 0.001839 | 0.000112 | 98.6621 | 98.1261 | 119.0701 | 21.2612 | 0.003383 | 0.003007 | 222.4061 | 0.24831 | 8.4025 | 0.003585 | 0.005081 | 0.004989 | 5.1e-05 | 0.001072 | 0.30613 | 1.3879 | 2.5472 | 100.8 | 21.12 | 15.621 | 103.2645 | 108.1454 | 94.1651 | 0.003887 | 0.96115 | 0.023514 | 24.4647 | 0.048729 | 0.14772 | 0.008161 | 0.57242 | 0.57242 | 0.010658
2019-03-26 16:23 | 1553631777.6715 | 598 | 0.20786 | 0.001817 | 0.000113 | 98.6108 | 98.273 | 119.1508 | 21.2612 | 0.003361 | 0.003126 | 222.3879 | 0.24771 | 8.4042 | 0.003643 | 0.004567 | 0.004534 | 5.1e-05 | 0.001074 | 0.34075 | 1.3834 | 2.5474 | 100.8 | 21.12 | 15.618 | 102.7263 | 109.5743 | 97.0469 | 0.003886 | 0.95367 | 0.023431 | 24.5688 | 0.030691 | 0.14736 | 0.008151 | 0.56958 | 0.56958 | 0.010685
2019-03-26 16:23 | 1553631778.6707 | 599 | 0.20786 | 0.001795 | 0.000116 | 98.7073 | 97.8659 | 119.1156 | 21.2612 | 0.003321 | 0.003068 | 222.382 | 0.24907 | 8.4051 | 0.003718 | 0.005119 | 0.005031 | 4.9e-05 | 0.001073 | 0.31813 | 1.345 | 2.5475 | 100.8 | 21.12 | 15.612 | 101.7493 | 111.3784 | 95.6425 | 0.003885 | 0.9573 | 0.023472 | 24.5186 | 0.076571 | 0.14817 | 0.008203 | 0.57415 | 0.57415 | 0.010658
2019-03-26 16:23 | 1553631779.6706 | 600 | 0.20788 | 0.001774 | 0.000118 | 98.6801 | 98.0128 | 119.1985 | 21.2612 | 0.003369 | 0.003062 | 222.3517 | 0.24753 | 8.4039 | 0.003687 | 0.004162 | 0.004039 | 4.8e-05 | 0.001067 | 0.34307 | 1.3616 | 2.5475 | 100.79 | 21.12 | 15.609 | 102.9094 | 110.6429 | 95.4851 | 0.003885 | 0.95771 | 0.023485 | 24.5218 | 0.060694 | 0.14726 | 0.008105 | 0.56737 | 0.56737 | 0.010719
2019-03-26 16:23 | 1553631780.6696 | 601 | 0.20794 | 0.001753 | 0.00012 | 98.5954 | 98.1859 | 119.4578 | 21.2612 | 0.003346 | 0.003061 | 222.318 | 0.24846 | 8.4025 | 0.003711 | 0.005877 | 0.005824 | 4.9e-05 | 0.001066 | 0.33771 | 1.3436 | 2.5474 | 100.79 | 21.12 | 15.604 | 102.3554 | 111.2244 | 95.4747 | 0.003883 | 0.95774 | 0.023511 | 24.5482 | 0.037226 | 0.14781 | 0.007737 | 0.54179 | 0.54179 | 0.011267
2019-03-26 16:23 | 1553631781.6706 | 602 | 0.20795 | 0.001735 | 0.000122 | 98.5857 | 98.2495 | 119.5501 | 21.2612 | 0.003385 | 0.003052 | 222.3053 | 0.24903 | 8.4024 | 0.003738 | 0.00501 | 0.004968 | 5e-05 | 0.001066 | 0.32211 | 1.3579 | 2.5473 | 100.79 | 21.12 | 15.602 | 103.3037 | 111.8828 | 95.2539 | 0.003883 | 0.95831 | 0.023527 | 24.5503 | 0.030556 | 0.14815 | 0.007733 | 0.54178 | 0.54178 | 0.011294
2019-03-26 16:23 | 1553631782.6698 | 603 | 0.208 | 0.001714 | 0.000124 | 98.6465 | 98.3012 | 119.6494 | 21.2612 | 0.003302 | 0.003035 | 222.3008 | 0.24845 | 8.4016 | 0.003702 | 0.004646 | 0.004484 | 5e-05 | 0.001059 | 0.33572 | 1.3337 | 2.5473 | 100.79 | 21.12 | 15.599 | 101.3009 | 110.9987 | 94.8353 | 0.003882 | 0.9594 | 0.02355 | 24.5465 | 0.03136 | 0.1478 | 0.007412 | 0.51936 | 0.51936 | 0.011753
2019-03-26 16:23 | 1553631783.6709 | 604 | 0.20804 | 0.001689 | 0.000127 | 98.5759 | 98.0037 | 119.8003 | 21.2612 | 0.003363 | 0.003025 | 222.2888 | 0.24819 | 8.4029 | 0.003675 | 0.005214 | 0.005157 | 4.7e-05 | 0.001054 | 0.30757 | 1.3302 | 2.5474 | 100.79 | 21.12 | 15.595 | 102.775 | 110.3344 | 94.5995 | 0.003881 | 0.96002 | 0.023572 | 24.5541 | 0.052072 | 0.14765 | 0.007222 | 0.5062 | 0.5062 | 0.012047
2019-03-26 16:23 | 1553631784.6702 | 605 | 0.20803 | 0.001667 | 0.000129 | 98.6143 | 98.1503 | 119.7536 | 21.2612 | 0.003302 | 0.003003 | 222.2476 | 0.24944 | 8.4043 | 0.003697 | 0.004609 | 0.004447 | 5.1e-05 | 0.001045 | 0.31365 | 1.3838 | 2.5473 | 100.79 | 21.12 | 15.595 | 101.3023 | 110.8732 | 94.0768 | 0.003881 | 0.96138 | 0.023585 | 24.5319 | 0.042187 | 0.14839 | 0.007297 | 0.51163 | 0.51163 | 0.011979
2019-03-26 16:23 | 1553631785.6709 | 606 | 0.20803 | 0.001642 | 0.000132 | 98.5851 | 98.2051 | 119.7545 | 21.2612 | 0.003333 | 0.003023 | 222.2343 | 0.24826 | 8.4018 | 0.003647 | 0.005465 | 0.005477 | 5.1e-05 | 0.001044 | 0.34036 | 1.3624 | 2.5475 | 100.79 | 21.12 | 15.597 | 102.0327 | 109.6649 | 94.5503 | 0.003882 | 0.96014 | 0.023569 | 24.5478 | 0.034544 | 0.14769 | 0.007363 | 0.51588 | 0.51588 | 0.011824
2019-03-26 16:23 | 1553631786.6713 | 607 | 0.20804 | 0.001617 | 0.000135 | 98.568 | 98.1089 | 119.7106 | 21.2612 | 0.003316 | 0.003044 | 222.204 | 0.24854 | 8.4036 | 0.003596 | 0.005442 | 0.00534 | 4.7e-05 | 0.001045 | 0.33588 | 1.4267 | 2.5474 | 100.79 | 21.13 | 15.599 | 101.6321 | 108.4329 | 95.0621 | 0.003884 | 0.95881 | 0.023549 | 24.5604 | 0.041751 | 0.14786 | 0.007299 | 0.51073 | 0.51073 | 0.011956
2019-03-26 16:23 | 1553631787.6709 | 608 | 0.20807 | 0.001593 | 0.000138 | 98.5266 | 98.1375 | 119.6604 | 21.2612 | 0.003286 | 0.00303 | 222.204 | 0.24838 | 8.4037 | 0.003622 | 0.005447 | 0.005335 | 5e-05 | 0.001048 | 0.32532 | 1.369 | 2.5474 | 100.79 | 21.13 | 15.604 | 100.8983 | 109.0579 | 94.7335 | 0.003886 | 0.95967 | 0.023554 | 24.5443 | 0.035396 | 0.14776 | 0.007157 | 0.50055 | 0.50055 | 0.012192
2019-03-26 16:23 | 1553631788.6701 | 609 | 0.20806 | 0.001578 | 0.00014 | 98.5145 | 98.0515 | 119.6665 | 21.2612 | 0.003308 | 0.003012 | 222.2007 | 0.24761 | 8.4021 | 0.003572 | 0.004813 | 0.004804 | 4.9e-05 | 0.001049 | 0.32908 | 1.3718 | 2.5473 | 100.79 | 21.12 | 15.612 | 101.4407 | 107.846 | 94.2888 | 0.003885 | 0.96083 | 0.023569 | 24.53 | 0.042144 | 0.1473 | 0.007222 | 0.50536 | 0.50536 | 0.012038
2019-03-26 16:23 | 1553631789.6701 | 610 | 0.20811 | 0.001552 | 0.000142 | 98.485 | 98.0037 | 119.6747 | 21.2612 | 0.003259 | 0.002936 | 222.2126 | 0.24854 | 8.4048 | 0.003608 | 0.00487 | 0.004719 | 4.4e-05 | 0.001046 | 0.33179 | 1.3955 | 2.5474 | 100.79 | 21.12 | 15.617 | 100.2553 | 108.7173 | 92.4627 | 0.003887 | 0.96563 | 0.023629 | 24.4699 | 0.043819 | 0.14786 | 0.006986 | 0.48969 | 0.48969 | 0.01247
2019-03-26 16:23 | 1553631790.6701 | 611 | 0.2081 | 0.001537 | 0.000143 | 98.5201 | 97.8531 | 119.5959 | 21.2612 | 0.003269 | 0.002997 | 222.2174 | 0.24925 | 8.4033 | 0.003589 | 0.004701 | 0.004704 | 4.8e-05 | 0.00105 | 0.34844 | 1.3677 | 2.5474 | 100.79 | 21.12 | 15.622 | 100.4892 | 108.2444 | 93.9311 | 0.003888 | 0.96176 | 0.023574 | 24.5109 | 0.060764 | 0.14828 | 0.007033 | 0.49181 | 0.49181 | 0.012452
2019-03-26 16:23 | 1553631791.6711 | 612 | 0.20808 | 0.001522 | 0.000144 | 98.5231 | 97.8904 | 119.6351 | 21.2612 | 0.003253 | 0.002963 | 222.2115 | 0.24815 | 8.402 | 0.003581 | 0.005446 | 0.005368 | 4.8e-05 | 0.00105 | 0.31996 | 1.3548 | 2.5474 | 100.79 | 21.12 | 15.632 | 100.1015 | 108.0538 | 93.1025 | 0.00389 | 0.96394 | 0.023604 | 24.4872 | 0.057622 | 0.14762 | 0.007153 | 0.50094 | 0.50094 | 0.012171
2019-03-26 16:23 | 1553631792.6703 | 613 | 0.2081 | 0.001509 | 0.000144 | 98.4756 | 98.112 | 119.816 | 21.2612 | 0.003226 | 0.00296 | 222.205 | 0.24747 | 8.4017 | 0.003542 | 0.004313 | 0.004223 | 4.5e-05 | 0.001049 | 0.33826 | 1.3459 | 2.5475 | 100.79 | 21.12 | 15.638 | 99.4646 | 107.1134 | 93.0377 | 0.003892 | 0.96411 | 0.023624 | 24.5035 | 0.033087 | 0.14722 | 0.007046 | 0.49371 | 0.49371 | 0.012315
2019-03-26 16:23 | 1553631793.6702 | 614 | 0.20814 | 0.001497 | 0.000144 | 98.5325 | 97.922 | 119.7475 | 21.2612 | 0.00326 | 0.002964 | 222.1898 | 0.24863 | 8.4041 | 0.003553 | 0.005034 | 0.004925 | 4.5e-05 | 0.001047 | 0.33107 | 1.3554 | 2.5474 | 100.79 | 21.12 | 15.647 | 100.2759 | 107.3656 | 93.1199 | 0.003894 | 0.96389 | 0.023615 | 24.4993 | 0.055593 | 0.14791 | 0.006869 | 0.48092 | 0.48092 | 0.012702
2019-03-26 16:23 | 1553631794.6714 | 615 | 0.20814 | 0.001487 | 0.000144 | 98.4803 | 97.8084 | 120.0059 | 21.2612 | 0.003242 | 0.002956 | 222.1917 | 0.24743 | 8.4004 | 0.003619 | 0.005337 | 0.005289 | 4.9e-05 | 0.001047 | 0.36193 | 1.3202 | 2.5474 | 100.79 | 21.12 | 15.652 | 99.849 | 108.9719 | 92.9289 | 0.003895 | 0.9644 | 0.023646 | 24.5193 | 0.061233 | 0.1472 | 0.006857 | 0.4807 | 0.4807 | 0.012647
2019-03-26 16:23 | 1553631795.6703 | 616 | 0.20813 | 0.001476 | 0.000144 | 98.4522 | 97.9284 | 120.2841 | 21.2612 | 0.003191 | 0.002915 | 222.1799 | 0.24868 | 8.4024 | 0.003534 | 0.004752 | 0.004657 | 4.8e-05 | 0.001043 | 0.34612 | 1.3343 | 2.5473 | 100.79 | 21.13 | 15.656 | 98.6082 | 106.9229 | 91.9442 | 0.003899 | 0.967 | 0.023706 | 24.5147 | 0.047718 | 0.14794 | 0.006936 | 0.48763 | 0.48763 | 0.01253
2019-03-26 16:23 | 1553631796.6704 | 617 | 0.20819 | 0.001468 | 0.000143 | 98.4623 | 97.8501 | 120.3134 | 21.2612 | 0.0032 | 0.002937 | 222.1871 | 0.24881 | 8.4028 | 0.003616 | 0.004806 | 0.004719 | 4.6e-05 | 0.001044 | 0.31006 | 1.3349 | 2.5474 | 100.79 | 21.12 | 15.657 | 98.8269 | 108.9151 | 92.4706 | 0.003896 | 0.96561 | 0.023691 | 24.5353 | 0.055786 | 0.14802 | 0.006599 | 0.46325 | 0.46325 | 0.013196
2019-03-26 16:23 | 1553631797.6705 | 618 | 0.20813 | 0.001461 | 0.000144 | 98.4643 | 97.9567 | 120.1965 | 21.2612 | 0.003177 | 0.002956 | 222.1607 | 0.24827 | 8.4028 | 0.003491 | 0.005569 | 0.005468 | 4.7e-05 | 0.001045 | 0.32372 | 1.3132 | 2.5474 | 100.79 | 21.12 | 15.654 | 98.2805 | 105.8809 | 92.9492 | 0.003896 | 0.96434 | 0.023664 | 24.5395 | 0.046238 | 0.14769 | 0.006931 | 0.48641 | 0.48641 | 0.01254
2019-03-26 16:23 | 1553631798.6696 | 619 | 0.2082 | 0.001447 | 0.000144 | 98.454 | 97.9637 | 120.3017 | 21.2612 | 0.003157 | 0.002908 | 222.11 | 0.24822 | 8.4039 | 0.003546 | 0.005208 | 0.005121 | 4.1e-05 | 0.001044 | 0.30038 | 1.3316 | 2.5474 | 100.79 | 21.13 | 15.645 | 97.7787 | 107.2134 | 91.7832 | 0.003896 | 0.96742 | 0.023713 | 24.5111 | 0.044655 | 0.14767 | 0.006545 | 0.45976 | 0.45976 | 0.013265
2019-03-26 16:23 | 1553631799.6706 | 620 | 0.20819 | 0.001437 | 0.000144 | 98.4888 | 98.078 | 120.2662 | 21.2612 | 0.003243 | 0.002928 | 222.0484 | 0.24845 | 8.404 | 0.003563 | 0.00491 | 0.004866 | 4.3e-05 | 0.001043 | 0.32947 | 1.3394 | 2.5473 | 100.79 | 21.13 | 15.639 | 99.8662 | 107.6231 | 92.272 | 0.003894 | 0.96613 | 0.023693 | 24.5239 | 0.037389 | 0.1478 | 0.006599 | 0.46327 | 0.46327 | 0.013177
2019-03-26 16:23 | 1553631800.6696 | 621 | 0.20816 | 0.001428 | 0.000145 | 98.4351 | 97.9109 | 120.1275 | 21.2612 | 0.003164 | 0.002907 | 222.0248 | 0.24844 | 8.4049 | 0.003583 | 0.005394 | 0.005333 | 4.5e-05 | 0.001044 | 0.33052 | 1.379 | 2.5475 | 100.79 | 21.13 | 15.635 | 97.9488 | 108.1036 | 91.7599 | 0.003893 | 0.96749 | 0.023696 | 24.4925 | 0.047759 | 0.14779 | 0.006787 | 0.47669 | 0.47669 | 0.012805
2019-03-26 16:23 | 1553631801.6698 | 622 | 0.20817 | 0.001421 | 0.000145 | 98.4525 | 98.0904 | 120.1529 | 21.2612 | 0.003191 | 0.002912 | 222.0183 | 0.24795 | 8.4011 | 0.003483 | 0.00532 | 0.005188 | 4.4e-05 | 0.00104 | 0.34036 | 1.3684 | 2.5474 | 100.79 | 21.13 | 15.624 | 98.6051 | 105.6849 | 91.8888 | 0.003891 | 0.96715 | 0.023695 | 24.4994 | 0.032957 | 0.14751 | 0.00674 | 0.47328 | 0.47328 | 0.012872
2019-03-26 16:23 | 1553631802.6707 | 623 | 0.20818 | 0.001414 | 0.000147 | 98.4694 | 98.1741 | 119.8334 | 21.2612 | 0.003185 | 0.002881 | 222.0058 | 0.24803 | 8.4029 | 0.003575 | 0.005593 | 0.005601 | 4.5e-05 | 0.001035 | 0.30453 | 1.3571 | 2.5475 | 100.79 | 21.13 | 15.618 | 98.4569 | 107.9067 | 91.1304 | 0.003889 | 0.96916 | 0.023688 | 24.4414 | 0.026856 | 0.14755 | 0.006728 | 0.47221 | 0.47221 | 0.012905
2019-03-26 16:23 | 1553631803.6708 | 624 | 0.20818 | 0.001408 | 0.000148 | 98.4218 | 97.9422 | 120.0012 | 21.2612 | 0.003165 | 0.002913 | 221.9936 | 0.2481 | 8.4035 | 0.003471 | 0.004716 | 0.004623 | 4.6e-05 | 0.001036 | 0.33229 | 1.3668 | 2.5475 | 100.79 | 21.13 | 15.611 | 97.985 | 105.3866 | 91.9083 | 0.003887 | 0.96709 | 0.023679 | 24.4846 | 0.043696 | 0.14759 | 0.006731 | 0.47217 | 0.47217 | 0.01291
2019-03-26 16:23 | 1553631804.6699 | 625 | 0.20821 | 0.001404 | 0.000149 | 98.4387 | 98.1009 | 119.8632 | 21.2612 | 0.003111 | 0.00289 | 221.9508 | 0.24825 | 8.4011 | 0.003521 | 0.004448 | 0.004382 | 4.7e-05 | 0.001039 | 0.33251 | 1.3814 | 2.5474 | 100.79 | 21.13 | 15.599 | 96.6793 | 106.6033 | 91.3524 | 0.003884 | 0.96857 | 0.023683 | 24.4519 | 0.030747 | 0.14769 | 0.006538 | 0.45842 | 0.45842 | 0.013305
2019-03-26 16:23 | 1553631805.6709 | 626 | 0.20821 | 0.001401 | 0.000151 | 98.4062 | 98.0928 | 120.0019 | 21.2612 | 0.003124 | 0.002924 | 221.8966 | 0.24823 | 8.4039 | 0.003475 | 0.004983 | 0.004853 | 4.7e-05 | 0.001038 | 0.32604 | 1.3321 | 2.5474 | 100.79 | 21.13 | 15.594 | 96.9865 | 105.4788 | 92.1648 | 0.003883 | 0.96641 | 0.023671 | 24.4933 | 0.028531 | 0.14768 | 0.00656 | 0.45963 | 0.45963 | 0.013269
2019-03-26 16:23 | 1553631806.6711 | 627 | 0.20822 | 0.001395 | 0.000152 | 98.3994 | 97.9385 | 120.2829 | 21.2612 | 0.003153 | 0.002876 | 221.856 | 0.24864 | 8.4039 | 0.003517 | 0.005014 | 0.005044 | 4.9e-05 | 0.001038 | 0.30381 | 1.3578 | 2.5473 | 100.79 | 21.13 | 15.582 | 97.6925 | 106.5121 | 91.0061 | 0.00388 | 0.96949 | 0.023736 | 24.483 | 0.041995 | 0.14791 | 0.006469 | 0.45433 | 0.45433 | 0.013446
2019-03-26 16:23 | 1553631807.671 | 628 | 0.20821 | 0.001386 | 0.000154 | 98.3994 | 97.8743 | 120.5756 | 21.2612 | 0.003154 | 0.002885 | 221.851 | 0.24849 | 8.401 | 0.003521 | 0.004895 | 0.00483 | 4.5e-05 | 0.001037 | 0.31813 | 1.3774 | 2.5474 | 100.79 | 21.13 | 15.578 | 97.7277 | 106.612 | 91.2203 | 0.003879 | 0.96892 | 0.023758 | 24.52 | 0.047861 | 0.14783 | 0.006554 | 0.46071 | 0.46071 | 0.013252
2019-03-26 16:23 | 1553631808.6714 | 629 | 0.20825 | 0.001384 | 0.000156 | 98.4012 | 97.9351 | 120.3216 | 21.2612 | 0.00313 | 0.002843 | 221.8377 | 0.24748 | 8.403 | 0.003497 | 0.004962 | 0.004925 | 4.2e-05 | 0.001037 | 0.33627 | 1.3486 | 2.5473 | 100.79 | 21.13 | 15.574 | 97.1432 | 106.0114 | 90.2181 | 0.003878 | 0.97159 | 0.023766 | 24.4605 | 0.042463 | 0.14723 | 0.006327 | 0.44453 | 0.44453 | 0.013678
2019-03-26 16:23 | 1553631809.6701 | 630 | 0.20825 | 0.00138 | 0.000158 | 98.3501 | 97.9123 | 120.228 | 21.2612 | 0.003121 | 0.002851 | 221.8332 | 0.24846 | 8.4049 | 0.003514 | 0.00485 | 0.004767 | 4.6e-05 | 0.001044 | 0.31614 | 1.3286 | 2.5473 | 100.79 | 21.13 | 15.566 | 96.9114 | 106.4318 | 90.4065 | 0.003876 | 0.97109 | 0.02375 | 24.4573 | 0.039908 | 0.14781 | 0.006287 | 0.44121 | 0.44121 | 0.013836
2019-03-26 16:23 | 1553631810.6701 | 631 | 0.20824 | 0.001375 | 0.000159 | 98.4269 | 97.9926 | 120.1036 | 21.2612 | 0.003151 | 0.002836 | 221.8087 | 0.249 | 8.3998 | 0.003516 | 0.004976 | 0.004925 | 4.3e-05 | 0.001043 | 0.30917 | 1.2933 | 2.5472 | 100.79 | 21.13 | 15.562 | 97.6393 | 106.4781 | 90.0384 | 0.003875 | 0.97207 | 0.02375 | 24.4323 | 0.039548 | 0.14813 | 0.006406 | 0.44963 | 0.44963 | 0.013606
2019-03-26 16:24 | 1553631811.6692 | 632 | 0.20825 | 0.001366 | 0.000161 | 98.3852 | 97.9926 | 119.9768 | 21.2612 | 0.003135 | 0.002865 | 221.7995 | 0.24816 | 8.4016 | 0.003506 | 0.005245 | 0.005215 | 4.3e-05 | 0.00105 | 0.32875 | 1.3546 | 2.5474 | 100.79 | 21.13 | 15.555 | 97.2608 | 106.2396 | 90.7411 | 0.003873 | 0.9702 | 0.023714 | 24.443 | 0.035765 | 0.14763 | 0.006344 | 0.44445 | 0.44445 | 0.013719
2019-03-26 16:24 | 1553631812.6697 | 633 | 0.20833 | 0.001358 | 0.000164 | 98.3548 | 97.9372 | 120.0761 | 21.2612 | 0.003127 | 0.002852 | 221.783 | 0.24803 | 8.4028 | 0.003456 | 0.004587 | 0.004438 | 4.2e-05 | 0.00105 | 0.3194 | 1.3235 | 2.5473 | 100.79 | 21.13 | 15.551 | 97.0524 | 105.0218 | 90.4385 | 0.003872 | 0.971 | 0.023734 | 24.4429 | 0.038065 | 0.14755 | 0.005846 | 0.40894 | 0.40894 | 0.014902
2019-03-26 16:24 | 1553631813.6703 | 634 | 0.20831 | 0.001347 | 0.000166 | 98.3811 | 97.958 | 120.1198 | 21.2612 | 0.003093 | 0.002874 | 221.7786 | 0.24879 | 8.4005 | 0.003469 | 0.00524 | 0.005139 | 4.3e-05 | 0.001051 | 0.31509 | 1.4022 | 2.5473 | 100.79 | 21.13 | 15.549 | 96.2532 | 105.3435 | 90.963 | 0.003872 | 0.9696 | 0.023721 | 24.465 | 0.038551 | 0.148 | 0.006001 | 0.41971 | 0.41971 | 0.014564
2019-03-26 16:24 | 1553631814.6702 | 635 | 0.2083 | 0.001349 | 0.000169 | 98.4024 | 98.033 | 120.3287 | 21.2612 | 0.00307 | 0.002855 | 221.7922 | 0.24909 | 8.4018 | 0.003478 | 0.004695 | 0.004617 | 4.6e-05 | 0.001055 | 0.36387 | 1.3658 | 2.5473 | 100.79 | 21.13 | 15.544 | 95.6903 | 105.5549 | 90.5004 | 0.003871 | 0.97084 | 0.023757 | 24.4707 | 0.03364 | 0.14818 | 0.006021 | 0.42162 | 0.42162 | 0.014515
2019-03-26 16:24 | 1553631815.6703 | 636 | 0.20825 | 0.001344 | 0.000169 | 98.3439 | 97.8245 | 120.4418 | 21.2612 | 0.003103 | 0.00286 | 221.7996 | 0.24745 | 8.4021 | 0.003478 | 0.004973 | 0.004866 | 4.2e-05 | 0.001057 | 0.34578 | 1.3443 | 2.5473 | 100.79 | 21.12 | 15.541 | 96.4946 | 105.5503 | 90.6207 | 0.003868 | 0.97052 | 0.023764 | 24.4862 | 0.047363 | 0.14721 | 0.00634 | 0.44457 | 0.44457 | 0.013676
2019-03-26 16:24 | 1553631816.6709 | 637 | 0.20837 | 0.001338 | 0.00017 | 98.3044 | 98.0703 | 120.4662 | 21.2612 | 0.003074 | 0.002869 | 221.8044 | 0.24802 | 8.4008 | 0.00344 | 0.005174 | 0.005052 | 4.6e-05 | 0.001058 | 0.33035 | 1.3779 | 2.5474 | 100.79 | 21.12 | 15.536 | 95.7928 | 104.6252 | 90.8473 | 0.003866 | 0.96991 | 0.023759 | 24.4963 | 0.021324 | 0.14755 | 0.005621 | 0.3929 | 0.3929 | 0.015509
2019-03-26 16:24 | 1553631817.6706 | 638 | 0.20827 | 0.001332 | 0.00017 | 98.343 | 98.0545 | 120.1524 | 21.2612 | 0.003073 | 0.002799 | 221.7888 | 0.24859 | 8.4046 | 0.003497 | 0.005343 | 0.005251 | 4.5e-05 | 0.001053 | 0.32709 | 1.3924 | 2.5474 | 100.79 | 21.12 | 15.535 | 95.7629 | 106.0247 | 89.1486 | 0.003866 | 0.97446 | 0.023784 | 24.4073 | 0.026283 | 0.14788 | 0.006212 | 0.43573 | 0.43573 | 0.014017
2019-03-26 16:24 | 1553631818.6697 | 639 | 0.20838 | 0.00133 | 0.00017 | 98.4189 | 98.0861 | 120.058 | 21.2612 | 0.00309 | 0.002807 | 221.81 | 0.24827 | 8.403 | 0.003485 | 0.004698 | 0.004648 | 4.5e-05 | 0.001052 | 0.31974 | 1.3607 | 2.5476 | 100.79 | 21.12 | 15.535 | 96.1764 | 105.732 | 89.3526 | 0.003866 | 0.97391 | 0.023768 | 24.4045 | 0.030298 | 0.14769 | 0.005576 | 0.38984 | 0.38984 | 0.015647
2019-03-26 16:24 | 1553631819.671 | 640 | 0.20809 | 0.001324 | 0.00017 | 98.3052 | 98.0481 | 120.2022 | 21.2612 | 0.00305 | 0.002823 | 221.8054 | 0.24881 | 8.3987 | 0.003357 | 0.004973 | 0.004762 | 4.5e-05 | 0.001041 | 0.33251 | 1.3928 | 2.5473 | 100.79 | 21.12 | 15.532 | 95.1929 | 102.6305 | 89.7362 | 0.003865 | 0.97288 | 0.02377 | 24.4321 | 0.023428 | 0.14802 | 0.007335 | 0.51596 | 0.51596 | 0.011848
2019-03-26 16:24 | 1553631820.6697 | 641 | 0.20848 | 0.001322 | 0.000171 | 98.4313 | 97.9338 | 120.3665 | 21.2612 | 0.003018 | 0.002815 | 221.7759 | 0.24848 | 8.4036 | 0.003443 | 0.005119 | 0.005082 | 4.7e-05 | 0.001046 | 0.34324 | 1.3863 | 2.5475 | 100.79 | 21.12 | 15.532 | 94.4374 | 104.7124 | 89.5322 | 0.003865 | 0.97343 | 0.023792 | 24.4419 | 0.045322 | 0.14782 | 0.004962 | 0.34608 | 0.34608 | 0.01764
2019-03-26 16:24 | 1553631821.6697 | 642 | 0.20849 | 0.001315 | 0.000173 | 98.3944 | 98.0148 | 120.4998 | 21.2612 | 0.003006 | 0.002811 | 221.7494 | 0.24836 | 8.4032 | 0.003387 | 0.005157 | 0.00505 | 4.5e-05 | 0.001046 | 0.35026 | 1.3895 | 2.5475 | 100.79 | 21.12 | 15.531 | 94.1487 | 103.3432 | 89.447 | 0.003865 | 0.97366 | 0.023808 | 24.4526 | 0.034572 | 0.14775 | 0.004928 | 0.34378 | 0.34378 | 0.01775
2019-03-26 16:24 | 1553631822.6698 | 643 | 0.20836 | 0.001311 | 0.000175 | 98.3525 | 98.12 | 120.8552 | 21.2612 | 0.003052 | 0.002814 | 221.7273 | 0.24788 | 8.4022 | 0.00339 | 0.005279 | 0.005202 | 4.7e-05 | 0.001045 | 0.33445 | 1.36 | 2.5473 | 100.79 | 21.11 | 15.53 | 95.2455 | 103.4294 | 89.5064 | 0.003863 | 0.9735 | 0.023842 | 24.4906 | 0.021164 | 0.14747 | 0.00573 | 0.4018 | 0.4018 | 0.015158
2019-03-26 16:24 | 1553631823.6699 | 644 | 0.20838 | 0.001303 | 0.000177 | 98.428 | 98.0541 | 120.6723 | 21.2612 | 0.00305 | 0.002837 | 221.7061 | 0.24792 | 8.4025 | 0.003378 | 0.004769 | 0.004653 | 4.6e-05 | 0.001045 | 0.33467 | 1.3368 | 2.5474 | 100.79 | 21.11 | 15.529 | 95.2008 | 103.1269 | 90.0746 | 0.003862 | 0.97198 | 0.023805 | 24.4912 | 0.034042 | 0.14749 | 0.005615 | 0.3928 | 0.3928 | 0.015507
2019-03-26 16:24 | 1553631824.6709 | 645 | 0.20836 | 0.001299 | 0.000178 | 98.4044 | 98.0999 | 120.5686 | 21.2612 | 0.003068 | 0.00278 | 221.685 | 0.24799 | 8.3996 | 0.003418 | 0.004962 | 0.004897 | 4.6e-05 | 0.001044 | 0.31741 | 1.3241 | 2.5473 | 100.79 | 21.11 | 15.527 | 95.6369 | 104.1046 | 88.6916 | 0.003862 | 0.97569 | 0.02384 | 24.4341 | 0.027727 | 0.14753 | 0.005724 | 0.40115 | 0.40115 | 0.015188
2019-03-26 16:24 | 1553631825.671 | 646 | 0.20835 | 0.001293 | 0.00018 | 98.4449 | 97.7849 | 120.3 | 21.2612 | 0.003045 | 0.002767 | 221.6843 | 0.24873 | 8.4029 | 0.003413 | 0.005198 | 0.00511 | 4.7e-05 | 0.001048 | 0.34252 | 1.3451 | 2.5474 | 100.79 | 21.11 | 15.526 | 95.0896 | 103.9861 | 88.39 | 0.003862 | 0.9765 | 0.023823 | 24.3967 | 0.060167 | 0.14797 | 0.005778 | 0.40469 | 0.40469 | 0.015101
2019-03-26 16:24 | 1553631826.6709 | 647 | 0.20836 | 0.001289 | 0.000179 | 98.3377 | 97.961 | 120.105 | 21.2612 | 0.003006 | 0.002804 | 221.6639 | 0.24887 | 8.4048 | 0.003465 | 0.004389 | 0.004338 | 4.7e-05 | 0.001046 | 0.30934 | 1.3654 | 2.5474 | 100.79 | 21.11 | 15.526 | 94.148 | 105.2465 | 89.2791 | 0.003862 | 0.97411 | 0.023775 | 24.4068 | 0.034328 | 0.14805 | 0.005772 | 0.40351 | 0.40351 | 0.015154
2019-03-26 16:24 | 1553631827.669 | 648 | 0.20843 | 0.001285 | 0.000178 | 98.3911 | 97.836 | 120.5676 | 21.2612 | 0.003017 | 0.002798 | 221.6429 | 0.24846 | 8.4036 | 0.003482 | 0.004496 | 0.004394 | 4.5e-05 | 0.001051 | 0.34755 | 1.3366 | 2.5473 | 100.79 | 21.11 | 15.524 | 94.4164 | 105.6495 | 89.125 | 0.003861 | 0.97452 | 0.023826 | 24.4486 | 0.050612 | 0.14781 | 0.005328 | 0.37245 | 0.37245 | 0.01639
2019-03-26 16:24 | 1553631828.669 | 649 | 0.20842 | 0.00128 | 0.000178 | 98.4239 | 97.9587 | 120.8275 | 21.2612 | 0.003013 | 0.002819 | 221.618 | 0.24878 | 8.4066 | 0.003469 | 0.005155 | 0.005025 | 4.8e-05 | 0.001055 | 0.33356 | 1.3505 | 2.5474 | 100.79 | 21.11 | 15.523 | 94.3163 | 105.3277 | 89.6423 | 0.003861 | 0.97313 | 0.023834 | 24.4924 | 0.04238 | 0.148 | 0.005374 | 0.37592 | 0.37592 | 0.01626
2019-03-26 16:24 | 1553631829.6701 | 650 | 0.20844 | 0.001272 | 0.000177 | 98.3722 | 97.7936 | 120.9864 | 21.2612 | 0.002962 | 0.002775 | 221.6036 | 0.24898 | 8.4059 | 0.003462 | 0.004892 | 0.00484 | 4.7e-05 | 0.001064 | 0.3298 | 1.3705 | 2.5473 | 100.79 | 21.11 | 15.52 | 93.0837 | 105.175 | 88.5869 | 0.00386 | 0.97597 | 0.023885 | 24.4728 | 0.052768 | 0.14812 | 0.00526 | 0.36852 | 0.36852 | 0.016599
2019-03-26 16:24 | 1553631830.669 | 651 | 0.20839 | 0.001267 | 0.000177 | 98.4124 | 98.0575 | 120.8744 | 21.2612 | 0.003084 | 0.002835 | 221.6096 | 0.24812 | 8.404 | 0.00348 | 0.004873 | 0.004814 | 4.6e-05 | 0.001063 | 0.3454 | 1.3629 | 2.5474 | 100.79 | 21.11 | 15.519 | 96.0292 | 105.6015 | 90.0175 | 0.00386 | 0.97213 | 0.023827 | 24.5098 | 0.032314 | 0.14761 | 0.005579 | 0.3906 | 0.3906 | 0.015607
2019-03-26 16:24 | 1553631831.6702 | 652 | 0.20843 | 0.001265 | 0.000177 | 98.4667 | 98.1926 | 120.9167 | 21.2612 | 0.003028 | 0.002752 | 221.5854 | 0.2485 | 8.4004 | 0.003473 | 0.004575 | 0.004482 | 4.7e-05 | 0.001066 | 0.3361 | 1.3384 | 2.5474 | 100.79 | 21.11 | 15.518 | 94.6801 | 105.4322 | 88.0214 | 0.00386 | 0.9775 | 0.023897 | 24.4466 | 0.024931 | 0.14783 | 0.00533 | 0.37376 | 0.37376 | 0.016335
2019-03-26 16:24 | 1553631832.6692 | 653 | 0.20842 | 0.00126 | 0.00018 | 98.4714 | 98.1335 | 120.5573 | 21.2612 | 0.003044 | 0.002797 | 221.5851 | 0.24902 | 8.4028 | 0.003398 | 0.004684 | 0.004588 | 4.3e-05 | 0.001064 | 0.34468 | 1.3381 | 2.5473 | 100.79 | 21.1 | 15.518 | 95.07 | 103.6181 | 89.0947 | 0.003857 | 0.97461 | 0.023826 | 24.4465 | 0.030751 | 0.14814 | 0.005399 | 0.37744 | 0.37744 | 0.01621
2019-03-26 16:24 | 1553631833.6706 | 654 | 0.2084 | 0.00125 | 0.000182 | 98.4821 | 97.9647 | 120.5751 | 21.2612 | 0.003002 | 0.002777 | 221.5688 | 0.24851 | 8.4036 | 0.00349 | 0.005065 | 0.004995 | 4.4e-05 | 0.001069 | 0.35961 | 1.3468 | 2.5475 | 100.79 | 21.1 | 15.518 | 94.051 | 105.8493 | 88.6189 | 0.003857 | 0.97589 | 0.023843 | 24.4323 | 0.04711 | 0.14784 | 0.005527 | 0.38687 | 0.38687 | 0.015783
2019-03-26 16:24 | 1553631834.6693 | 655 | 0.20839 | 0.00125 | 0.000185 | 98.4688 | 98.116 | 120.5216 | 21.2612 | 0.003012 | 0.00275 | 221.5599 | 0.24752 | 8.4038 | 0.003427 | 0.005288 | 0.005186 | 4.3e-05 | 0.001065 | 0.31813 | 1.3413 | 2.5474 | 100.8 | 21.1 | 15.519 | 94.285 | 104.3264 | 87.9849 | 0.003857 | 0.9776 | 0.023859 | 24.4054 | 0.032103 | 0.14725 | 0.005575 | 0.39038 | 0.39038 | 0.015579
2019-03-26 16:24 | 1553631835.6704 | 656 | 0.20844 | 0.001242 | 0.000187 | 98.3861 | 97.9748 | 120.769 | 21.2612 | 0.002981 | 0.002799 | 221.5254 | 0.24815 | 8.4016 | 0.003514 | 0.004744 | 0.004664 | 4.3e-05 | 0.001071 | 0.32709 | 1.3455 | 2.5473 | 100.8 | 21.1 | 15.518 | 93.5508 | 106.4293 | 89.1542 | 0.003857 | 0.97445 | 0.023845 | 24.47 | 0.037472 | 0.14762 | 0.005316 | 0.37136 | 0.37136 | 0.016418
2019-03-26 16:24 | 1553631836.6695 | 657 | 0.20839 | 0.001242 | 0.000189 | 98.4646 | 98.2437 | 120.8706 | 21.2612 | 0.002956 | 0.00278 | 221.4947 | 0.24774 | 8.4036 | 0.003556 | 0.005208 | 0.005118 | 4.5e-05 | 0.001066 | 0.35187 | 1.3484 | 2.5473 | 100.8 | 21.1 | 15.518 | 92.9446 | 107.4396 | 88.7087 | 0.003857 | 0.97564 | 0.023869 | 24.4652 | 0.02009 | 0.14738 | 0.005596 | 0.39181 | 0.39181 | 0.015535
2019-03-26 16:24 | 1553631837.6695 | 658 | 0.20836 | 0.001236 | 0.000191 | 98.392 | 98.1567 | 120.7965 | 21.2612 | 0.00297 | 0.002766 | 221.494 | 0.24823 | 8.4033 | 0.003491 | 0.005215 | 0.00506 | 4.5e-05 | 0.00107 | 0.34092 | 1.3417 | 2.5473 | 100.8 | 21.1 | 15.515 | 93.2865 | 105.876 | 88.3621 | 0.003856 | 0.97658 | 0.023873 | 24.446 | 0.021421 | 0.14768 | 0.00577 | 0.40429 | 0.40429 | 0.015086
2019-03-26 16:24 | 1553631838.6695 | 659 | 0.20843 | 0.001231 | 0.000192 | 98.4558 | 98.0713 | 120.7756 | 21.2612 | 0.002963 | 0.002726 | 221.4799 | 0.24739 | 8.4041 | 0.00351 | 0.005039 | 0.004929 | 4.5e-05 | 0.001064 | 0.30845 | 1.3411 | 2.5473 | 100.8 | 21.1 | 15.513 | 93.1129 | 106.3357 | 87.3858 | 0.003856 | 0.97922 | 0.023904 | 24.4109 | 0.035 | 0.14717 | 0.005368 | 0.37576 | 0.37576 | 0.016176
2019-03-26 16:24 | 1553631839.6705 | 660 | 0.20843 | 0.001226 | 0.000194 | 98.4514 | 98.0861 | 120.5275 | 21.2612 | 0.002966 | 0.002773 | 221.4765 | 0.24806 | 8.4022 | 0.003563 | 0.004645 | 0.004588 | 4.2e-05 | 0.001068 | 0.3011 | 1.3502 | 2.5473 | 100.8 | 21.1 | 15.51 | 93.1833 | 107.6284 | 88.5229 | 0.003855 | 0.97615 | 0.023842 | 24.4242 | 0.033248 | 0.14757 | 0.005404 | 0.37722 | 0.37722 | 0.016157
2019-03-26 16:24 | 1553631840.6696 | 661 | 0.20846 | 0.001221 | 0.000195 | 98.4705 | 98.1069 | 120.6746 | 21.2612 | 0.002939 | 0.00272 | 221.4477 | 0.24838 | 8.4023 | 0.003545 | 0.004873 | 0.004792 | 4.7e-05 | 0.001075 | 0.33732 | 1.3515 | 2.5475 | 100.8 | 21.1 | 15.504 | 92.534 | 107.1937 | 87.2452 | 0.003853 | 0.97961 | 0.023898 | 24.3959 | 0.033091 | 0.14776 | 0.005229 | 0.36545 | 0.36545 | 0.016699
2019-03-26 16:24 | 1553631841.6696 | 662 | 0.20846 | 0.001223 | 0.000196 | 98.5131 | 98.0511 | 120.5056 | 21.2612 | 0.002997 | 0.002788 | 221.4357 | 0.24809 | 8.4019 | 0.003567 | 0.004642 | 0.004524 | 4.3e-05 | 0.00107 | 0.33627 | 1.3945 | 2.5474 | 100.8 | 21.09 | 15.501 | 93.9343 | 107.7144 | 88.9018 | 0.00385 | 0.97512 | 0.023827 | 24.4348 | 0.04204 | 0.14759 | 0.005203 | 0.36245 | 0.36245 | 0.016817
2019-03-26 16:24 | 1553631842.6696 | 663 | 0.20847 | 0.001223 | 0.000197 | 98.4652 | 98.0908 | 120.5744 | 21.2612 | 0.002925 | 0.002739 | 221.4078 | 0.24723 | 8.4029 | 0.003533 | 0.004993 | 0.004897 | 4.5e-05 | 0.001071 | 0.31238 | 1.3879 | 2.5475 | 100.8 | 21.09 | 15.497 | 92.1843 | 106.8945 | 87.7191 | 0.003849 | 0.97832 | 0.023873 | 24.4018 | 0.03408 | 0.14708 | 0.00514 | 0.35852 | 0.35852 | 0.016943
2019-03-26 16:24 | 1553631843.6699 | 664 | 0.20847 | 0.001214 | 0.000198 | 98.5024 | 97.995 | 120.7552 | 21.2612 | 0.002946 | 0.002721 | 221.3958 | 0.24812 | 8.4031 | 0.003538 | 0.005363 | 0.005238 | 4.5e-05 | 0.001071 | 0.31061 | 1.3613 | 2.5473 | 100.8 | 21.09 | 15.495 | 92.7021 | 107.0236 | 87.2852 | 0.003849 | 0.9795 | 0.023905 | 24.4054 | 0.046198 | 0.1476 | 0.005149 | 0.35959 | 0.35959 | 0.016953
2019-03-26 16:24 | 1553631844.6688 | 665 | 0.20849 | 0.001211 | 0.000199 | 98.5001 | 98.2696 | 120.8833 | 21.2612 | 0.002885 | 0.002787 | 221.3746 | 0.24811 | 8.4016 | 0.003508 | 0.00495 | 0.004907 | 4.4e-05 | 0.001071 | 0.32565 | 1.3372 | 2.5473 | 100.8 | 21.09 | 15.491 | 91.2353 | 106.2775 | 88.8667 | 0.003848 | 0.97522 | 0.023865 | 24.4719 | 0.020951 | 0.1476 | 0.005052 | 0.35202 | 0.35202 | 0.017317
2019-03-26 16:24 | 1553631845.6691 | 666 | 0.20848 | 0.001203 | 0.000199 | 98.5889 | 98.1268 | 120.6798 | 21.2612 | 0.002944 | 0.00272 | 221.3622 | 0.24799 | 8.402 | 0.003494 | 0.004994 | 0.004904 | 4.6e-05 | 0.001065 | 0.32781 | 1.3517 | 2.5473 | 100.8 | 21.09 | 15.484 | 92.6529 | 105.9476 | 87.2602 | 0.003846 | 0.97957 | 0.023898 | 24.3969 | 0.042028 | 0.14753 | 0.005113 | 0.35684 | 0.35684 | 0.017075
2019-03-26 16:24 | 1553631846.6689 | 667 | 0.20847 | 0.001198 | 0.000199 | 98.5564 | 98.161 | 120.688 | 21.2612 | 0.002998 | 0.002688 | 221.3482 | 0.24701 | 8.4035 | 0.00349 | 0.005166 | 0.005072 | 4.7e-05 | 0.001065 | 0.32493 | 1.3298 | 2.5473 | 100.8 | 21.09 | 15.481 | 93.9555 | 105.8399 | 86.4762 | 0.003845 | 0.9817 | 0.023925 | 24.3712 | 0.035957 | 0.14694 | 0.005189 | 0.36275 | 0.36275 | 0.01673
2019-03-26 16:24 | 1553631847.6689 | 668 | 0.20849 | 0.001193 | 0.000199 | 98.5443 | 98.1694 | 120.5507 | 21.2612 | 0.00298 | 0.002697 | 221.3207 | 0.24881 | 8.4034 | 0.003511 | 0.004859 | 0.004796 | 4.6e-05 | 0.001069 | 0.3361 | 1.349 | 2.5473 | 100.8 | 21.09 | 15.475 | 93.5156 | 106.3599 | 86.698 | 0.003844 | 0.9811 | 0.023904 | 24.3649 | 0.034093 | 0.14802 | 0.005037 | 0.35146 | 0.35146 | 0.017394
2019-03-26 16:24 | 1553631848.67 | 669 | 0.20847 | 0.001192 | 0.000198 | 98.5626 | 98.3449 | 120.551 | 21.2612 | 0.002921 | 0.002691 | 221.3003 | 0.24817 | 8.4032 | 0.003477 | 0.005344 | 0.005298 | 4.4e-05 | 0.00106 | 0.3183 | 1.3303 | 2.5474 | 100.8 | 21.09 | 15.472 | 92.0951 | 105.5244 | 86.5433 | 0.003843 | 0.98152 | 0.023909 | 24.3596 | 0.01978 | 0.14764 | 0.005188 | 0.36249 | 0.36249 | 0.016821
2019-03-26 16:24 | 1553631849.669 | 670 | 0.20855 | 0.001186 | 0.000198 | 98.5458 | 98.2131 | 120.6063 | 21.2612 | 0.002895 | 0.002743 | 221.2911 | 0.24869 | 8.4025 | 0.003499 | 0.005172 | 0.005085 | 4.6e-05 | 0.001057 | 0.33859 | 1.3681 | 2.5473 | 100.8 | 21.09 | 15.468 | 91.4736 | 106.0689 | 87.8043 | 0.003842 | 0.97809 | 0.023873 | 24.4079 | 0.030246 | 0.14794 | 0.004703 | 0.32703 | 0.32703 | 0.018684
2019-03-26 16:24 | 1553631850.6701 | 671 | 0.2085 | 0.001182 | 0.000198 | 98.5048 | 98.279 | 120.6085 | 21.2612 | 0.002929 | 0.00272 | 221.2777 | 0.2473 | 8.4034 | 0.003552 | 0.005168 | 0.005076 | 4.8e-05 | 0.001057 | 0.33555 | 1.3455 | 2.5472 | 100.8 | 21.09 | 15.462 | 92.2797 | 107.3537 | 87.2469 | 0.00384 | 0.9796 | 0.023892 | 24.3893 | 0.020523 | 0.14712 | 0.005 | 0.34868 | 0.34868 | 0.017426
2019-03-26 16:24 | 1553631851.6702 | 672 | 0.20855 | 0.001179 | 0.000198 | 98.5939 | 98.2686 | 120.6939 | 21.2612 | 0.00291 | 0.002705 | 221.2739 | 0.24806 | 8.4015 | 0.003606 | 0.005052 | 0.005004 | 4.7e-05 | 0.00105 | 0.29358 | 1.3352 | 2.5474 | 100.8 | 21.09 | 15.459 | 91.8372 | 108.6692 | 86.8978 | 0.00384 | 0.98055 | 0.023912 | 24.3861 | 0.029562 | 0.14757 | 0.004736 | 0.32993 | 0.32993 | 0.018472
2019-03-26 16:24 | 1553631852.671 | 673 | 0.20854 | 0.001175 | 0.0002 | 98.5697 | 98.3261 | 120.704 | 21.2612 | 0.002887 | 0.002697 | 221.2536 | 0.24856 | 8.4034 | 0.003523 | 0.005228 | 0.005088 | 4.6e-05 | 0.001054 | 0.34202 | 1.3521 | 2.5473 | 100.8 | 21.09 | 15.451 | 91.2774 | 106.6601 | 86.7074 | 0.003838 | 0.98107 | 0.023919 | 24.3807 | 0.022135 | 0.14787 | 0.004805 | 0.33488 | 0.33488 | 0.018236
2019-03-26 16:24 | 1553631853.6703 | 674 | 0.20856 | 0.001172 | 0.000202 | 98.6294 | 98.3106 | 120.9265 | 21.2612 | 0.002881 | 0.002714 | 221.235 | 0.24868 | 8.4012 | 0.00353 | 0.00534 | 0.005217 | 4.9e-05 | 0.001057 | 0.30237 | 1.3134 | 2.5473 | 100.8 | 21.09 | 15.446 | 91.1336 | 106.8173 | 87.0992 | 0.003837 | 0.98 | 0.023928 | 24.4164 | 0.028951 | 0.14794 | 0.004643 | 0.3232 | 0.3232 | 0.018905
2019-03-26 16:24 | 1553631854.6694 | 675 | 0.20857 | 0.001167 | 0.000204 | 98.5576 | 98.2609 | 121.1422 | 21.2612 | 0.002923 | 0.002739 | 221.2222 | 0.24857 | 8.4017 | 0.003554 | 0.00504 | 0.00506 | 4.6e-05 | 0.001058 | 0.33804 | 1.2988 | 2.5474 | 100.8 | 21.09 | 15.441 | 92.1334 | 107.3986 | 87.7107 | 0.003835 | 0.97834 | 0.023929 | 24.4589 | 0.026971 | 0.14787 | 0.004629 | 0.32206 | 0.32206 | 0.018963
2019-03-26 16:24 | 1553631855.6693 | 676 | 0.20858 | 0.001164 | 0.000205 | 98.6361 | 98.3369 | 121.3816 | 21.2612 | 0.002875 | 0.002667 | 221.2133 | 0.24777 | 8.4018 | 0.003502 | 0.005277 | 0.005237 | 4.5e-05 | 0.001058 | 0.34324 | 1.3287 | 2.5474 | 100.8 | 21.09 | 15.431 | 90.995 | 106.1365 | 85.9861 | 0.003833 | 0.98304 | 0.02401 | 24.4245 | 0.027182 | 0.1474 | 0.004554 | 0.31762 | 0.31762 | 0.019166
2019-03-26 16:24 | 1553631856.6706 | 677 | 0.2086 | 0.001155 | 0.000206 | 98.5736 | 98.1977 | 121.4537 | 21.2612 | 0.002923 | 0.002697 | 221.2023 | 0.24732 | 8.4022 | 0.003507 | 0.004473 | 0.004354 | 4.6e-05 | 0.001059 | 0.35314 | 1.3159 | 2.5474 | 100.8 | 21.09 | 15.427 | 92.1313 | 106.253 | 86.6988 | 0.003832 | 0.98109 | 0.023994 | 24.456 | 0.034173 | 0.14713 | 0.004426 | 0.30812 | 0.30812 | 0.019721
2019-03-26 16:24 | 1553631857.7312 | 678 | 0.20858 | 0.001146 | 0.000206 | 98.5969 | 98.2458 | 121.5398 | 21.2612 | 0.002907 | 0.002693 | 221.1911 | 0.24805 | 8.4034 | 0.003544 | 0.004591 | 0.004553 | 4.5e-05 | 0.001058 | 0.3178 | 1.3135 | 2.5476 | 100.8 | 21.09 | 15.423 | 91.762 | 107.1653 | 86.6066 | 0.003831 | 0.98135 | 0.024005 | 24.4615 | 0.031911 | 0.14757 | 0.004563 | 0.31817 | 0.31817 | 0.019155
2019-03-26 16:24 | 1553631858.6788 | 679 | 0.20857 | 0.001143 | 0.000206 | 98.596 | 98.1637 | 121.622 | 21.2612 | 0.002891 | 0.002712 | 221.1742 | 0.24795 | 8.4018 | 0.003462 | 0.004769 | 0.004728 | 4.5e-05 | 0.001061 | 0.35026 | 1.3335 | 2.5473 | 100.8 | 21.09 | 15.414 | 91.3634 | 105.1726 | 87.0703 | 0.003829 | 0.98008 | 0.023998 | 24.4855 | 0.039301 | 0.14751 | 0.004661 | 0.32516 | 0.32516 | 0.018736
2019-03-26 16:24 | 1553631859.6789 | 680 | 0.20855 | 0.001137 | 0.000207 | 98.5299 | 98.2733 | 121.807 | 21.2612 | 0.002888 | 0.002696 | 221.1469 | 0.24819 | 8.4052 | 0.003518 | 0.005101 | 0.005025 | 4.4e-05 | 0.001064 | 0.34396 | 1.3647 | 2.5474 | 100.8 | 21.09 | 15.411 | 91.2945 | 106.523 | 86.6769 | 0.003828 | 0.98115 | 0.024029 | 24.4908 | 0.023321 | 0.14765 | 0.00478 | 0.3342 | 0.3342 | 0.018246
2019-03-26 16:24 | 1553631860.6782 | 681 | 0.20856 | 0.001135 | 0.000208 | 98.6007 | 98.2958 | 121.9595 | 21.2612 | 0.002884 | 0.002657 | 221.1502 | 0.24709 | 8.4015 | 0.003484 | 0.00506 | 0.005008 | 4.6e-05 | 0.001061 | 0.34235 | 1.3497 | 2.5474 | 100.8 | 21.09 | 15.407 | 91.2133 | 105.6934 | 85.7292 | 0.003827 | 0.98374 | 0.024076 | 24.4738 | 0.0277 | 0.14699 | 0.004709 | 0.32958 | 0.32958 | 0.01842
2019-03-26 16:24 | 1553631861.6789 | 682 | 0.20854 | 0.001129 | 0.000208 | 98.6043 | 98.3621 | 121.9795 | 21.2612 | 0.002858 | 0.002675 | 221.1337 | 0.24826 | 8.4037 | 0.003475 | 0.004265 | 0.004136 | 4.7e-05 | 0.001061 | 0.34058 | 1.3795 | 2.5474 | 100.8 | 21.09 | 15.406 | 90.5669 | 105.4873 | 86.1687 | 0.003827 | 0.98254 | 0.024063 | 24.4908 | 0.021997 | 0.14769 | 0.004828 | 0.33809 | 0.33809 | 0.018041
2019-03-26 16:24 | 1553631862.6783 | 683 | 0.20855 | 0.00113 | 0.000208 | 98.5438 | 98.277 | 122.179 | 21.2612 | 0.002856 | 0.002701 | 221.1317 | 0.24722 | 8.4039 | 0.003479 | 0.004489 | 0.004388 | 4.6e-05 | 0.001059 | 0.36177 | 1.4067 | 2.5473 | 100.8 | 21.09 | 15.405 | 90.5286 | 105.5731 | 86.8073 | 0.003826 | 0.9808 | 0.024062 | 24.5326 | 0.024244 | 0.14707 | 0.004796 | 0.3357 | 0.3357 | 0.018093
2019-03-26 16:24 | 1553631863.6781 | 684 | 0.20854 | 0.001121 | 0.00021 | 98.6267 | 98.3002 | 122.1008 | 21.2612 | 0.002849 | 0.002679 | 221.1351 | 0.2485 | 8.4041 | 0.003463 | 0.005136 | 0.00501 | 4.9e-05 | 0.001059 | 0.35242 | 1.4107 | 2.5475 | 100.8 | 21.09 | 15.406 | 90.3496 | 105.1956 | 86.271 | 0.003827 | 0.98226 | 0.024072 | 24.5065 | 0.029658 | 0.14783 | 0.004819 | 0.33737 | 0.33737 | 0.018097
2019-03-26 16:24 | 1553631864.6791 | 685 | 0.20858 | 0.001119 | 0.000213 | 98.666 | 98.2767 | 121.7544 | 21.2612 | 0.002845 | 0.002723 | 221.1132 | 0.2485 | 8.4017 | 0.003518 | 0.004667 | 0.004567 | 4.5e-05 | 0.001057 | 0.32875 | 1.4099 | 2.5474 | 100.8 | 21.09 | 15.407 | 90.2699 | 106.5388 | 87.3242 | 0.003827 | 0.97939 | 0.024002 | 24.5075 | 0.03536 | 0.14783 | 0.004593 | 0.31988 | 0.31988 | 0.019087
2019-03-26 16:24 | 1553631865.6791 | 686 | 0.20854 | 0.001113 | 0.000217 | 98.6031 | 98.0783 | 121.4769 | 21.2612 | 0.002818 | 0.002764 | 221.1146 | 0.24822 | 8.4015 | 0.003511 | 0.004831 | 0.00474 | 4.4e-05 | 0.001062 | 0.31741 | 1.4008 | 2.5475 | 100.8 | 21.09 | 15.408 | 89.6033 | 106.3672 | 88.3089 | 0.003827 | 0.97672 | 0.023942 | 24.513 | 0.047728 | 0.14767 | 0.004851 | 0.3375 | 0.3375 | 0.01807
2019-03-26 16:24 | 1553631866.6782 | 687 | 0.20859 | 0.001112 | 0.00022 | 98.5765 | 98.2397 | 121.5246 | 21.2612 | 0.002901 | 0.002679 | 221.1191 | 0.2475 | 8.4032 | 0.003604 | 0.005064 | 0.004978 | 4.7e-05 | 0.001064 | 0.33069 | 1.3848 | 2.5472 | 100.8 | 21.09 | 15.409 | 91.6025 | 108.6039 | 86.2779 | 0.003827 | 0.98224 | 0.024015 | 24.4488 | 0.030614 | 0.14724 | 0.004513 | 0.31386 | 0.31386 | 0.019375
2019-03-26 16:24 | 1553631867.6796 | 688 | 0.20858 | 0.001107 | 0.000221 | 98.7191 | 98.3362 | 121.7194 | 21.2612 | 0.002917 | 0.002652 | 221.1129 | 0.248 | 8.4002 | 0.003661 | 0.0049 | 0.004781 | 4.2e-05 | 0.001069 | 0.32029 | 1.3583 | 2.5472 | 100.8 | 21.09 | 15.408 | 92.0069 | 109.9921 | 85.62 | 0.003827 | 0.98404 | 0.024056 | 24.446 | 0.034761 | 0.14753 | 0.004571 | 0.31854 | 0.31854 | 0.019129
2019-03-26 16:24 | 1553631868.6792 | 689 | 0.20857 | 0.001103 | 0.000221 | 98.6229 | 98.3987 | 121.6373 | 21.2612 | 0.002809 | 0.002672 | 221.1123 | 0.24901 | 8.4038 | 0.003651 | 0.004941 | 0.004856 | 4.8e-05 | 0.001068 | 0.32781 | 1.3332 | 2.5473 | 100.8 | 21.09 | 15.408 | 89.3868 | 109.7556 | 86.0946 | 0.003827 | 0.98274 | 0.024032 | 24.4539 | 0.020353 | 0.14813 | 0.004645 | 0.32357 | 0.32357 | 0.018908
2019-03-26 16:24 | 1553631869.6783 | 690 | 0.20859 | 0.001095 | 0.00022 | 98.6651 | 98.2468 | 121.9102 | 21.2612 | 0.002891 | 0.002683 | 221.1227 | 0.24881 | 8.4017 | 0.003686 | 0.005111 | 0.005063 | 4.7e-05 | 0.00107 | 0.34036 | 1.3446 | 2.5474 | 100.8 | 21.09 | 15.408 | 91.3806 | 110.6091 | 86.3568 | 0.003827 | 0.98203 | 0.02405 | 24.4902 | 0.038003 | 0.14802 | 0.004568 | 0.31834 | 0.31834 | 0.019203
2019-03-26 16:24 | 1553631870.6783 | 691 | 0.20861 | 0.001086 | 0.000218 | 98.5789 | 98.2689 | 121.7891 | 21.2612 | 0.002856 | 0.002682 | 221.1185 | 0.24848 | 8.4035 | 0.003627 | 0.004974 | 0.004937 | 4.7e-05 | 0.001074 | 0.35059 | 1.3568 | 2.5474 | 100.8 | 21.09 | 15.41 | 90.5183 | 109.1816 | 86.3311 | 0.003828 | 0.9821 | 0.024039 | 24.4772 | 0.028166 | 0.14782 | 0.004443 | 0.30919 | 0.30919 | 0.019745
2019-03-26 16:24 | 1553631871.6783 | 692 | 0.2086 | 0.001081 | 0.000217 | 98.6801 | 98.4962 | 121.7701 | 21.2612 | 0.002834 | 0.002675 | 221.0943 | 0.24773 | 8.4003 | 0.00357 | 0.004637 | 0.004438 | 4.9e-05 | 0.001069 | 0.34092 | 1.3565 | 2.5475 | 100.8 | 21.09 | 15.411 | 89.998 | 107.7923 | 86.1812 | 0.003828 | 0.98251 | 0.024042 | 24.4702 | 0.016688 | 0.14738 | 0.004491 | 0.31275 | 0.31275 | 0.019462
2019-03-26 16:25 | 1553631872.6785 | 693 | 0.20863 | 0.001078 | 0.000216 | 98.6155 | 98.3231 | 121.4173 | 21.2612 | 0.002848 | 0.002663 | 221.0916 | 0.24855 | 8.4043 | 0.003656 | 0.00507 | 0.005036 | 4.5e-05 | 0.001066 | 0.32189 | 1.3514 | 2.5473 | 100.8 | 21.08 | 15.415 | 90.334 | 109.8873 | 85.879 | 0.003826 | 0.98333 | 0.024017 | 24.4244 | 0.026562 | 0.14786 | 0.004329 | 0.30076 | 0.30076 | 0.020304
2019-03-26 16:25 | 1553631873.6776 | 694 | 0.20863 | 0.001072 | 0.000216 | 98.6542 | 98.2131 | 121.2204 | 21.2612 | 0.002823 | 0.002589 | 221.0591 | 0.24782 | 8.4011 | 0.003591 | 0.004913 | 0.004854 | 4.3e-05 | 0.00107 | 0.31293 | 1.348 | 2.5474 | 100.8 | 21.09 | 15.416 | 89.7432 | 108.2967 | 84.0929 | 0.003829 | 0.98825 | 0.024058 | 24.3438 | 0.040074 | 0.14743 | 0.004348 | 0.30262 | 0.30262 | 0.02012
2019-03-26 16:25 | 1553631874.6775 | 695 | 0.20864 | 0.00107 | 0.000216 | 98.6961 | 98.4122 | 121.4295 | 21.2612 | 0.002765 | 0.002618 | 221.0483 | 0.24786 | 8.404 | 0.003504 | 0.005487 | 0.005345 | 4.4e-05 | 0.001066 | 0.33323 | 1.3785 | 2.5474 | 100.8 | 21.09 | 15.418 | 88.344 | 106.179 | 84.799 | 0.00383 | 0.9863 | 0.024055 | 24.3889 | 0.025767 | 0.14745 | 0.00428 | 0.29771 | 0.29771 | 0.020456
2019-03-26 16:25 | 1553631875.6777 | 696 | 0.20865 | 0.001056 | 0.000215 | 98.6778 | 98.1641 | 121.3669 | 21.2612 | 0.00277 | 0.002636 | 221.054 | 0.24785 | 8.4012 | 0.00352 | 0.005251 | 0.00515 | 4.1e-05 | 0.001068 | 0.33771 | 1.351 | 2.5474 | 100.8 | 21.09 | 15.422 | 88.4557 | 106.5654 | 85.2305 | 0.003831 | 0.98511 | 0.024034 | 24.3973 | 0.046685 | 0.14745 | 0.004223 | 0.29341 | 0.29341 | 0.020755
2019-03-26 16:25 | 1553631876.6787 | 697 | 0.20863 | 0.001053 | 0.000214 | 98.7099 | 98.3839 | 121.1011 | 21.2612 | 0.002833 | 0.002603 | 221.0469 | 0.24817 | 8.4016 | 0.003491 | 0.005586 | 0.005478 | 4.6e-05 | 0.001066 | 0.28528 | 1.4014 | 2.5474 | 100.8 | 21.09 | 15.424 | 89.972 | 105.8753 | 84.4486 | 0.003831 | 0.98727 | 0.024034 | 24.344 | 0.029591 | 0.14764 | 0.004403 | 0.30644 | 0.30644 | 0.019898
2019-03-26 16:25 | 1553631877.6781 | 698 | 0.20864 | 0.001046 | 0.000213 | 98.7114 | 98.4044 | 120.9657 | 21.2612 | 0.002808 | 0.00263 | 221.0172 | 0.24785 | 8.4023 | 0.003554 | 0.004606 | 0.004516 | 4.5e-05 | 0.001072 | 0.34562 | 1.3815 | 2.5474 | 100.8 | 21.08 | 15.427 | 89.3596 | 107.4112 | 85.0792 | 0.003829 | 0.98553 | 0.023999 | 24.3518 | 0.027861 | 0.14745 | 0.004333 | 0.30102 | 0.30102 | 0.02023
2019-03-26 16:25 | 1553631878.6787 | 699 | 0.20865 | 0.001044 | 0.000211 | 98.7392 | 98.3305 | 120.9343 | 21.2612 | 0.00278 | 0.002598 | 221.0062 | 0.24842 | 8.4046 | 0.003589 | 0.004959 | 0.004973 | 4.6e-05 | 0.001071 | 0.32908 | 1.3612 | 2.5474 | 100.79 | 21.09 | 15.435 | 88.694 | 108.2619 | 84.313 | 0.003834 | 0.98764 | 0.024022 | 24.3226 | 0.0371 | 0.14778 | 0.004261 | 0.29618 | 0.29618 | 0.020607
2019-03-26 16:25 | 1553631879.6787 | 700 | 0.20864 | 0.001035 | 0.000208 | 98.7533 | 98.4511 | 120.6531 | 21.2612 | 0.002763 | 0.002606 | 220.9828 | 0.24791 | 8.4024 | 0.003538 | 0.005147 | 0.005076 | 4.7e-05 | 0.001068 | 0.30757 | 1.3444 | 2.5474 | 100.79 | 21.09 | 15.44 | 88.2783 | 107.0236 | 84.5008 | 0.003835 | 0.98712 | 0.023988 | 24.3007 | 0.027412 | 0.14749 | 0.004327 | 0.30073 | 0.30073 | 0.020254
2019-03-26 16:25 | 1553631880.6778 | 701 | 0.2087 | 0.001037 | 0.000206 | 98.7197 | 98.4064 | 120.7885 | 21.2612 | 0.002829 | 0.002595 | 220.9781 | 0.24821 | 8.4035 | 0.003534 | 0.005469 | 0.005388 | 4.3e-05 | 0.001066 | 0.30215 | 1.3524 | 2.5474 | 100.79 | 21.09 | 15.448 | 89.8877 | 106.9253 | 84.2449 | 0.003837 | 0.98783 | 0.02401 | 24.3056 | 0.028427 | 0.14766 | 0.003976 | 0.2757 | 0.2757 | 0.022119
2019-03-26 16:25 | 1553631881.6788 | 702 | 0.2087 | 0.001037 | 0.000204 | 98.6887 | 98.3348 | 121.0631 | 21.2612 | 0.002799 | 0.002607 | 220.9963 | 0.24801 | 8.4046 | 0.003457 | 0.004394 | 0.004268 | 4.8e-05 | 0.001071 | 0.32173 | 1.3283 | 2.5474 | 100.79 | 21.09 | 15.452 | 89.1532 | 105.0373 | 84.5289 | 0.003838 | 0.98704 | 0.024028 | 24.3429 | 0.032129 | 0.14754 | 0.003984 | 0.27667 | 0.27667 | 0.022025
2019-03-26 16:25 | 1553631882.6789 | 703 | 0.20867 | 0.001031 | 0.000201 | 98.6996 | 98.5308 | 121.4276 | 21.2612 | 0.002764 | 0.002685 | 220.9645 | 0.24771 | 8.4045 | 0.003539 | 0.005406 | 0.005284 | 4.4e-05 | 0.001078 | 0.29623 | 1.3641 | 2.5473 | 100.79 | 21.09 | 15.463 | 88.3182 | 107.033 | 86.4115 | 0.003841 | 0.98188 | 0.024001 | 24.4436 | 0.01531 | 0.14736 | 0.004182 | 0.29085 | 0.29085 | 0.020925
2019-03-26 16:25 | 1553631883.6779 | 704 | 0.20869 | 0.001025 | 0.000199 | 98.7238 | 98.3658 | 121.7793 | 21.2612 | 0.002815 | 0.002643 | 220.9722 | 0.24765 | 8.4028 | 0.003515 | 0.005321 | 0.00527 | 4.4e-05 | 0.001081 | 0.33356 | 1.3577 | 2.5474 | 100.79 | 21.09 | 15.467 | 89.5353 | 106.464 | 85.4015 | 0.003842 | 0.98464 | 0.024069 | 24.4445 | 0.032499 | 0.14733 | 0.004097 | 0.2856 | 0.2856 | 0.021305
2019-03-26 16:25 | 1553631884.679 | 705 | 0.20867 | 0.001026 | 0.000197 | 98.7129 | 98.4599 | 121.9041 | 21.2612 | 0.002829 | 0.00264 | 220.9798 | 0.2486 | 8.4019 | 0.003445 | 0.004221 | 0.004087 | 4.5e-05 | 0.001083 | 0.31116 | 1.3541 | 2.5473 | 100.79 | 21.09 | 15.472 | 89.8807 | 104.7677 | 85.3352 | 0.003843 | 0.98483 | 0.024084 | 24.4548 | 0.022955 | 0.14789 | 0.00416 | 0.29047 | 0.29047 | 0.021027
2019-03-26 16:25 | 1553631885.6783 | 706 | 0.20868 | 0.001021 | 0.000195 | 98.7507 | 98.4596 | 122.1036 | 21.2612 | 0.002822 | 0.002563 | 220.9506 | 0.24794 | 8.3999 | 0.003511 | 0.005417 | 0.005269 | 4.5e-05 | 0.001084 | 0.33411 | 1.3937 | 2.5474 | 100.79 | 21.08 | 15.482 | 89.7032 | 106.3498 | 83.4735 | 0.003843 | 0.98997 | 0.024166 | 24.4112 | 0.026408 | 0.1475 | 0.004147 | 0.29061 | 0.29061 | 0.020962
2019-03-26 16:25 | 1553631886.6795 | 707 | 0.20867 | 0.00102 | 0.000193 | 98.717 | 98.3187 | 122.4541 | 21.2612 | 0.002777 | 0.002646 | 220.9533 | 0.24728 | 8.4034 | 0.00342 | 0.005117 | 0.004986 | 4.6e-05 | 0.001088 | 0.31741 | 1.401 | 2.5474 | 100.79 | 21.08 | 15.487 | 88.6245 | 104.1613 | 85.4607 | 0.003845 | 0.98448 | 0.024134 | 24.5142 | 0.036164 | 0.14711 | 0.004171 | 0.29216 | 0.29216 | 0.020796
2019-03-26 16:25 | 1553631887.6782 | 708 | 0.20867 | 0.001017 | 0.000192 | 98.6875 | 98.4155 | 122.544 | 21.2612 | 0.002752 | 0.002606 | 220.9583 | 0.24756 | 8.403 | 0.003459 | 0.004969 | 0.004874 | 4.8e-05 | 0.001088 | 0.33301 | 1.3572 | 2.5475 | 100.79 | 21.08 | 15.493 | 88.0308 | 105.087 | 84.5174 | 0.003846 | 0.98708 | 0.024174 | 24.4909 | 0.024686 | 0.14727 | 0.004195 | 0.29444 | 0.29444 | 0.020657
2019-03-26 16:25 | 1553631888.6792 | 709 | 0.20865 | 0.00102 | 0.000189 | 98.7058 | 98.6078 | 122.5377 | 21.2612 | 0.002787 | 0.002617 | 220.9508 | 0.24821 | 8.4005 | 0.003452 | 0.005015 | 0.004898 | 4.4e-05 | 0.001093 | 0.37 | 1.4023 | 2.5474 | 100.79 | 21.08 | 15.505 | 88.8688 | 104.9272 | 84.774 | 0.003849 | 0.98637 | 0.024165 | 24.4991 | 0.008886 | 0.14766 | 0.004319 | 0.30355 | 0.30355 | 0.02009
2019-03-26 16:25 | 1553631889.6772 | 710 | 0.20868 | 0.001028 | 0.000186 | 98.6539 | 98.2989 | 122.4431 | 21.2612 | 0.002787 | 0.002607 | 220.9439 | 0.24853 | 8.406 | 0.003464 | 0.00499 | 0.004925 | 4.7e-05 | 0.001096 | 0.33284 | 1.4257 | 2.5474 | 100.79 | 21.08 | 15.511 | 88.8573 | 105.223 | 84.5313 | 0.003851 | 0.98704 | 0.024164 | 24.4813 | 0.032243 | 0.14785 | 0.004163 | 0.29234 | 0.29234 | 0.020887
2019-03-26 16:25 | 1553631890.6778 | 711 | 0.20866 | 0.001026 | 0.000183 | 98.7598 | 98.508 | 122.552 | 21.2612 | 0.002834 | 0.002629 | 220.9289 | 0.24931 | 8.4043 | 0.00347 | 0.004978 | 0.004914 | 4.8e-05 | 0.001091 | 0.3131 | 1.4415 | 2.5474 | 100.79 | 21.08 | 15.523 | 89.9908 | 105.3694 | 85.072 | 0.003854 | 0.98555 | 0.024157 | 24.5107 | 0.02284 | 0.14831 | 0.004249 | 0.29868 | 0.29868 | 0.020508
2019-03-26 16:25 | 1553631891.6774 | 712 | 0.20866 | 0.001023 | 0.000182 | 98.7309 | 98.5285 | 122.4663 | 21.2612 | 0.002774 | 0.002624 | 220.9147 | 0.24829 | 8.4045 | 0.003382 | 0.005215 | 0.005132 | 4.9e-05 | 0.001094 | 0.3131 | 1.363 | 2.5475 | 100.79 | 21.08 | 15.529 | 88.5611 | 103.2245 | 84.9347 | 0.003855 | 0.98593 | 0.024153 | 24.4974 | 0.018358 | 0.14771 | 0.004307 | 0.30297 | 0.30297 | 0.020136
2019-03-26 16:25 | 1553631892.6785 | 713 | 0.20864 | 0.00103 | 0.000179 | 98.7268 | 98.6387 | 122.4975 | 21.2612 | 0.002758 | 0.002652 | 220.9335 | 0.24838 | 8.4006 | 0.003363 | 0.0048 | 0.004717 | 5e-05 | 0.001091 | 0.3418 | 1.316 | 2.5473 | 100.79 | 21.08 | 15.534 | 88.1728 | 102.7607 | 85.6099 | 0.003856 | 0.98407 | 0.024133 | 24.5236 | 0.00798 | 0.14776 | 0.004396 | 0.30939 | 0.30939 | 0.019724
2019-03-26 16:25 | 1553631893.6784 | 714 | 0.20867 | 0.001026 | 0.000176 | 98.6993 | 98.4582 | 122.2731 | 21.2612 | 0.00278 | 0.00268 | 220.9211 | 0.24808 | 8.4023 | 0.00335 | 0.004794 | 0.0047 | 5e-05 | 0.001097 | 0.34418 | 1.315 | 2.5474 | 100.79 | 21.08 | 15.543 | 88.6892 | 102.467 | 86.296 | 0.003859 | 0.98219 | 0.024088 | 24.5246 | 0.021877 | 0.14758 | 0.00421 | 0.29544 | 0.29544 | 0.020631
2019-03-26 16:25 | 1553631894.6779 | 715 | 0.20866 | 0.001032 | 0.000174 | 98.8572 | 98.7187 | 122.2313 | 21.2612 | 0.002785 | 0.002594 | 220.8936 | 0.24788 | 8.4022 | 0.003329 | 0.004853 | 0.00473 | 4.7e-05 | 0.001099 | 0.33467 | 1.3067 | 2.5473 | 100.79 | 21.07 | 15.546 | 88.8096 | 101.955 | 84.2185 | 0.003857 | 0.9879 | 0.024154 | 24.4494 | 0.012541 | 0.14747 | 0.004293 | 0.30246 | 0.30246 | 0.020136
2019-03-26 16:25 | 1553631895.6776 | 716 | 0.20866 | 0.001033 | 0.000172 | 98.8106 | 98.6828 | 122.3534 | 21.2612 | 0.002792 | 0.002575 | 220.8924 | 0.24894 | 8.401 | 0.003361 | 0.004946 | 0.00493 | 5.2e-05 | 0.001098 | 0.33445 | 1.3182 | 2.5474 | 100.79 | 21.07 | 15.549 | 88.9798 | 102.7215 | 83.7597 | 0.003858 | 0.98917 | 0.024181 | 24.4459 | 0.011579 | 0.14809 | 0.004288 | 0.30257 | 0.30257 | 0.020214
2019-03-26 16:25 | 1553631896.6786 | 717 | 0.20863 | 0.001031 | 0.00017 | 98.7728 | 98.3523 | 122.2431 | 21.2612 | 0.002792 | 0.002624 | 220.868 | 0.24778 | 8.4052 | 0.003291 | 0.005424 | 0.005293 | 5e-05 | 0.001099 | 0.33987 | 1.3309 | 2.5474 | 100.79 | 21.07 | 15.55 | 88.983 | 101.0371 | 84.9316 | 0.003858 | 0.98594 | 0.024131 | 24.475 | 0.03816 | 0.1474 | 0.004434 | 0.31264 | 0.31264 | 0.019472
2019-03-26 16:25 | 1553631897.6787 | 718 | 0.20866 | 0.001036 | 0.000168 | 98.7899 | 98.2814 | 122.0614 | 21.2612 | 0.002735 | 0.002611 | 220.854 | 0.24818 | 8.4049 | 0.003297 | 0.004596 | 0.00466 | 5e-05 | 0.001097 | 0.35817 | 1.3489 | 2.5475 | 100.79 | 21.07 | 15.551 | 87.6168 | 101.1823 | 84.6196 | 0.003858 | 0.98679 | 0.024123 | 24.4461 | 0.046163 | 0.14764 | 0.00429 | 0.30222 | 0.30222 | 0.020175
2019-03-26 16:25 | 1553631898.6787 | 719 | 0.20867 | 0.001036 | 0.000165 | 98.7923 | 98.713 | 122.2276 | 21.2612 | 0.002796 | 0.002614 | 220.856 | 0.24768 | 8.4038 | 0.003376 | 0.00444 | 0.004363 | 5.2e-05 | 0.001092 | 0.31492 | 1.3472 | 2.5473 | 100.79 | 21.07 | 15.552 | 89.0728 | 103.0782 | 84.7045 | 0.003858 | 0.98656 | 0.024137 | 24.4657 | 0.007181 | 0.14735 | 0.004216 | 0.29716 | 0.29716 | 0.020479
2019-03-26 16:25 | 1553631899.6789 | 720 | 0.20863 | 0.001038 | 0.000163 | 98.7489 | 98.626 | 122.2257 | 21.2612 | 0.002798 | 0.002637 | 220.8518 | 0.24836 | 8.3996 | 0.003352 | 0.004435 | 0.004336 | 5.5e-05 | 0.001088 | 0.31222 | 1.3323 | 2.5474 | 100.79 | 21.07 | 15.552 | 89.1274 | 102.5041 | 85.2579 | 0.003858 | 0.98504 | 0.024118 | 24.4844 | 0.011144 | 0.14775 | 0.004448 | 0.31389 | 0.31389 | 0.01944
2019-03-26 16:25 | 1553631900.6788 | 721 | 0.20867 | 0.001037 | 0.000161 | 98.6928 | 98.5419 | 121.9842 | 21.2612 | 0.002792 | 0.002591 | 220.8256 | 0.24748 | 8.4022 | 0.003329 | 0.004976 | 0.004866 | 4.9e-05 | 0.001088 | 0.33107 | 1.334 | 2.5472 | 100.79 | 21.07 | 15.552 | 88.983 | 101.9354 | 84.1576 | 0.003858 | 0.98807 | 0.024131 | 24.4226 | 0.013687 | 0.14723 | 0.004195 | 0.29582 | 0.29582 | 0.020555
2019-03-26 16:25 | 1553631901.6789 | 722 | 0.20868 | 0.001041 | 0.000159 | 98.7817 | 98.4612 | 121.794 | 21.2612 | 0.002755 | 0.00261 | 220.8014 | 0.24813 | 8.4024 | 0.003353 | 0.005088 | 0.004969 | 5e-05 | 0.001081 | 0.3454 | 1.3614 | 2.5474 | 100.79 | 21.07 | 15.552 | 88.1004 | 102.5384 | 84.596 | 0.003858 | 0.98686 | 0.024098 | 24.4186 | 0.029062 | 0.14761 | 0.004178 | 0.29425 | 0.29425 | 0.020718
2019-03-26 16:25 | 1553631902.6789 | 723 | 0.20869 | 0.001041 | 0.000158 | 98.8059 | 98.6236 | 121.6905 | 21.2612 | 0.002774 | 0.002604 | 220.7789 | 0.24846 | 8.4032 | 0.003275 | 0.004735 | 0.004687 | 4.8e-05 | 0.00108 | 0.34987 | 1.3601 | 2.5475 | 100.79 | 21.07 | 15.552 | 88.5448 | 100.6517 | 84.4597 | 0.003858 | 0.98724 | 0.024092 | 24.4035 | 0.016516 | 0.14781 | 0.004072 | 0.28652 | 0.28652 | 0.021306
2019-03-26 16:25 | 1553631903.678 | 724 | 0.20865 | 0.001045 | 0.000156 | 98.7105 | 98.5436 | 121.5133 | 21.2612 | 0.002835 | 0.002596 | 220.7684 | 0.24863 | 8.4017 | 0.003291 | 0.005287 | 0.005118 | 5.1e-05 | 0.00108 | 0.35275 | 1.3175 | 2.5473 | 100.79 | 21.07 | 15.554 | 90.0276 | 101.0166 | 84.2668 | 0.003859 | 0.98777 | 0.024081 | 24.3792 | 0.01514 | 0.14791 | 0.004326 | 0.30495 | 0.30495 | 0.020032
2019-03-26 16:25 | 1553631904.679 | 725 | 0.20868 | 0.00104 | 0.000155 | 98.7831 | 98.5691 | 121.5218 | 21.2612 | 0.00271 | 0.002584 | 220.7711 | 0.24868 | 8.4022 | 0.003311 | 0.005019 | 0.004999 | 4.9e-05 | 0.001079 | 0.30431 | 1.316 | 2.5474 | 100.79 | 21.07 | 15.555 | 87.0087 | 101.5223 | 83.9867 | 0.003859 | 0.98854 | 0.024091 | 24.3705 | 0.019399 | 0.14794 | 0.004189 | 0.29522 | 0.29522 | 0.020696
2019-03-26 16:25 | 1553631905.6782 | 726 | 0.20866 | 0.001044 | 0.000154 | 98.8174 | 98.6105 | 121.5549 | 21.2612 | 0.002742 | 0.002529 | 220.7475 | 0.24838 | 8.4025 | 0.003395 | 0.00471 | 0.004582 | 5.1e-05 | 0.001067 | 0.31669 | 1.3526 | 2.5474 | 100.79 | 21.07 | 15.556 | 87.7786 | 103.5387 | 82.6707 | 0.003859 | 0.9922 | 0.024139 | 24.3288 | 0.018747 | 0.14776 | 0.00427 | 0.30181 | 0.30181 | 0.02022
2019-03-26 16:25 | 1553631906.6771 | 727 | 0.20865 | 0.001039 | 0.000153 | 98.869 | 98.6122 | 121.745 | 21.2612 | 0.002777 | 0.002563 | 220.7427 | 0.24805 | 8.4029 | 0.003386 | 0.005431 | 0.005295 | 4.8e-05 | 0.001061 | 0.31974 | 1.4266 | 2.5476 | 100.8 | 21.07 | 15.56 | 88.6134 | 103.3313 | 83.4673 | 0.00386 | 0.98998 | 0.024131 | 24.3751 | 0.023269 | 0.14757 | 0.004367 | 0.3088 | 0.3088 | 0.019736
2019-03-26 16:25 | 1553631907.6772 | 728 | 0.20866 | 0.001044 | 0.000151 | 98.8313 | 98.6364 | 121.7354 | 21.2612 | 0.002721 | 0.002634 | 220.7494 | 0.24781 | 8.4047 | 0.003383 | 0.005073 | 0.004999 | 4.8e-05 | 0.00106 | 0.30989 | 1.3581 | 2.5474 | 100.8 | 21.07 | 15.561 | 87.2679 | 103.2473 | 85.1721 | 0.00386 | 0.98527 | 0.024073 | 24.4323 | 0.017655 | 0.14742 | 0.004264 | 0.3007 | 0.3007 | 0.020248
2019-03-26 16:25 | 1553631908.6772 | 729 | 0.20866 | 0.001045 | 0.00015 | 98.8764 | 98.6908 | 121.4197 | 21.2612 | 0.002662 | 0.002588 | 220.7235 | 0.24832 | 8.4027 | 0.003411 | 0.004908 | 0.004814 | 4.4e-05 | 0.001062 | 0.3282 | 1.3534 | 2.5474 | 100.8 | 21.07 | 15.563 | 85.8689 | 103.9302 | 84.078 | 0.003861 | 0.98829 | 0.024078 | 24.3633 | 0.016805 | 0.14773 | 0.004278 | 0.30184 | 0.30184 | 0.020213
2019-03-26 16:25 | 1553631909.6773 | 730 | 0.20864 | 0.001046 | 0.000149 | 98.8702 | 98.4999 | 121.323 | 21.2612 | 0.002754 | 0.002574 | 220.7092 | 0.24793 | 8.4033 | 0.003337 | 0.004274 | 0.004213 | 4.5e-05 | 0.001062 | 0.31061 | 1.4008 | 2.5474 | 100.8 | 21.07 | 15.564 | 88.0799 | 102.1433 | 83.7433 | 0.003861 | 0.98922 | 0.02408 | 24.3422 | 0.033566 | 0.1475 | 0.004399 | 0.31072 | 0.31072 | 0.019605
2019-03-26 16:25 | 1553631910.6783 | 731 | 0.20866 | 0.001046 | 0.000148 | 98.9009 | 98.5856 | 121.1655 | 21.2612 | 0.002727 | 0.002635 | 220.7056 | 0.24824 | 8.4015 | 0.003334 | 0.004625 | 0.004526 | 4.8e-05 | 0.001057 | 0.31094 | 1.3892 | 2.5475 | 100.8 | 21.07 | 15.564 | 87.4327 | 102.0593 | 85.1961 | 0.003861 | 0.98521 | 0.024015 | 24.3759 | 0.028562 | 0.14768 | 0.004275 | 0.30094 | 0.30094 | 0.020267
2019-03-26 16:25 | 1553631911.6784 | 732 | 0.20865 | 0.00105 | 0.000146 | 98.8871 | 98.8041 | 121.76 | 21.2612 | 0.002727 | 0.002587 | 220.7048 | 0.24773 | 8.4028 | 0.003389 | 0.004661 | 0.004534 | 4.6e-05 | 0.001058 | 0.32421 | 1.3897 | 2.5475 | 100.8 | 21.07 | 15.563 | 87.4318 | 103.3919 | 84.0523 | 0.003861 | 0.98836 | 0.024113 | 24.3966 | 0.007505 | 0.14738 | 0.004366 | 0.30889 | 0.30889 | 0.019705
2019-03-26 16:25 | 1553631912.6774 | 733 | 0.20865 | 0.001053 | 0.000145 | 98.9166 | 98.6703 | 121.5366 | 21.2612 | 0.00274 | 0.002597 | 220.6954 | 0.24875 | 8.4025 | 0.003352 | 0.004996 | 0.005031 | 4.8e-05 | 0.001058 | 0.32012 | 1.3731 | 2.5473 | 100.79 | 21.07 | 15.561 | 87.7379 | 102.5024 | 84.2901 | 0.003861 | 0.98771 | 0.024083 | 24.3823 | 0.022295 | 0.14798 | 0.004348 | 0.30725 | 0.30725 | 0.019891
2019-03-26 16:25 | 1553631913.6784 | 734 | 0.20863 | 0.00106 | 0.000143 | 98.8897 | 98.7678 | 121.5117 | 21.2612 | 0.002777 | 0.002591 | 220.6901 | 0.24828 | 8.4059 | 0.003321 | 0.005132 | 0.004996 | 4.6e-05 | 0.001061 | 0.3454 | 1.3595 | 2.5475 | 100.79 | 21.07 | 15.56 | 88.6276 | 101.7515 | 84.1576 | 0.00386 | 0.98807 | 0.024085 | 24.3753 | 0.011032 | 0.1477 | 0.004492 | 0.31788 | 0.31788 | 0.01919
2019-03-26 16:25 | 1553631914.6785 | 735 | 0.20862 | 0.001058 | 0.000142 | 98.8788 | 98.6976 | 121.0946 | 21.2612 | 0.002729 | 0.002616 | 220.678 | 0.24879 | 8.4034 | 0.003336 | 0.005276 | 0.005278 | 5e-05 | 0.001061 | 0.31885 | 1.4457 | 2.5475 | 100.79 | 21.07 | 15.557 | 87.4577 | 102.1191 | 84.7545 | 0.00386 | 0.98642 | 0.024023 | 24.3537 | 0.016409 | 0.14801 | 0.004521 | 0.31923 | 0.31923 | 0.019148
2019-03-26 16:25 | 1553631915.6786 | 736 | 0.20861 | 0.001067 | 0.00014 | 98.9594 | 98.6908 | 120.926 | 21.2612 | 0.002718 | 0.002562 | 220.6755 | 0.24786 | 8.4042 | 0.003396 | 0.00548 | 0.005365 | 4.8e-05 | 0.001064 | 0.33826 | 1.4287 | 2.5476 | 100.79 | 21.07 | 15.555 | 87.2087 | 103.5772 | 83.4562 | 0.003859 | 0.99001 | 0.02405 | 24.2926 | 0.024305 | 0.14745 | 0.004596 | 0.32518 | 0.32518 | 0.018727
2019-03-26 16:25 | 1553631916.6776 | 737 | 0.20865 | 0.001065 | 0.000139 | 98.9012 | 98.756 | 120.6974 | 21.2612 | 0.00281 | 0.002548 | 220.6781 | 0.24807 | 8.4044 | 0.0034 | 0.005195 | 0.005123 | 4.8e-05 | 0.001059 | 0.29192 | 1.4289 | 2.5475 | 100.79 | 21.07 | 15.549 | 89.4149 | 103.6748 | 83.1184 | 0.003858 | 0.99095 | 0.024039 | 24.2581 | 0.013139 | 0.14758 | 0.004359 | 0.30787 | 0.30787 | 0.019797
2019-03-26 16:25 | 1553631917.6776 | 738 | 0.20864 | 0.001062 | 0.000137 | 98.9186 | 98.8276 | 120.908 | 21.2612 | 0.00278 | 0.002551 | 220.6725 | 0.24843 | 8.4015 | 0.003331 | 0.004914 | 0.00483 | 4.8e-05 | 0.001062 | 0.32781 | 1.4156 | 2.5475 | 100.79 | 21.07 | 15.546 | 88.708 | 101.9998 | 83.1973 | 0.003857 | 0.99073 | 0.024057 | 24.282 | 0.008231 | 0.14779 | 0.004395 | 0.31084 | 0.31084 | 0.019636
2019-03-26 16:25 | 1553631918.6787 | 739 | 0.20866 | 0.001069 | 0.000136 | 98.9405 | 98.6041 | 121.2511 | 21.2612 | 0.002676 | 0.00257 | 220.66 | 0.24893 | 8.401 | 0.003357 | 0.005082 | 0.005027 | 4.5e-05 | 0.001061 | 0.34562 | 1.3499 | 2.5475 | 100.79 | 21.07 | 15.54 | 86.1827 | 102.6154 | 83.6505 | 0.003855 | 0.98948 | 0.024076 | 24.3319 | 0.030462 | 0.14809 | 0.004292 | 0.30372 | 0.30372 | 0.020138
2019-03-26 16:25 | 1553631919.6787 | 740 | 0.20864 | 0.001067 | 0.000135 | 98.9145 | 98.5816 | 121.1802 | 21.2612 | 0.002744 | 0.002571 | 220.6469 | 0.24813 | 8.4021 | 0.003447 | 0.0048 | 0.004723 | 4.8e-05 | 0.001064 | 0.34578 | 1.372 | 2.5475 | 100.79 | 21.07 | 15.537 | 87.8412 | 104.7951 | 83.6748 | 0.003855 | 0.98941 | 0.024068 | 24.3256 | 0.030156 | 0.14761 | 0.004395 | 0.31111 | 0.31111 | 0.019596
2019-03-26 16:25 | 1553631920.6778 | 741 | 0.20867 | 0.001069 | 0.000135 | 98.9245 | 98.8542 | 121.2983 | 21.2612 | 0.002757 | 0.002586 | 220.6316 | 0.24813 | 8.4022 | 0.003472 | 0.004943 | 0.00485 | 5e-05 | 0.001061 | 0.30596 | 1.395 | 2.5476 | 100.79 | 21.07 | 15.534 | 88.1502 | 105.4102 | 84.0304 | 0.003854 | 0.98842 | 0.024068 | 24.3495 | 0.006362 | 0.14761 | 0.004245 | 0.30026 | 0.30026 | 0.020304
2019-03-26 16:25 | 1553631921.6779 | 742 | 0.20863 | 0.001075 | 0.000135 | 98.983 | 98.9006 | 120.8821 | 21.2612 | 0.002786 | 0.002529 | 220.6303 | 0.24835 | 8.4015 | 0.003486 | 0.00516 | 0.0051 | 4.6e-05 | 0.001066 | 0.335 | 1.3884 | 2.5476 | 100.79 | 21.07 | 15.528 | 88.833 | 105.747 | 82.6659 | 0.003853 | 0.99221 | 0.024072 | 24.2613 | 0.007448 | 0.14774 | 0.004456 | 0.31566 | 0.31566 | 0.01933
2019-03-26 16:25 | 1553631922.6771 | 743 | 0.20859 | 0.001079 | 0.000134 | 98.9431 | 98.7581 | 121.0016 | 21.2612 | 0.002715 | 0.00256 | 220.6385 | 0.24842 | 8.4009 | 0.003427 | 0.004776 | 0.004703 | 4.5e-05 | 0.001071 | 0.31636 | 1.3493 | 2.5474 | 100.79 | 21.07 | 15.525 | 87.129 | 104.313 | 83.4125 | 0.003852 | 0.99014 | 0.024059 | 24.2987 | 0.016746 | 0.14778 | 0.00468 | 0.33178 | 0.33178 | 0.018396
2019-03-26 16:25 | 1553631923.6769 | 744 | 0.20863 | 0.001078 | 0.000133 | 98.9786 | 98.8034 | 121.2638 | 21.2612 | 0.002714 | 0.002615 | 220.6381 | 0.24809 | 8.4036 | 0.003416 | 0.004558 | 0.004427 | 5e-05 | 0.001074 | 0.30702 | 1.3636 | 2.5475 | 100.79 | 21.07 | 15.522 | 87.1203 | 104.0489 | 84.7187 | 0.003851 | 0.98652 | 0.024041 | 24.3695 | 0.015839 | 0.14759 | 0.00443 | 0.31346 | 0.31346 | 0.019446
2019-03-26 16:25 | 1553631924.6771 | 745 | 0.20861 | 0.001083 | 0.000132 | 98.9827 | 98.8327 | 122.0431 | 21.2612 | 0.002703 | 0.002569 | 220.6192 | 0.24822 | 8.4039 | 0.003428 | 0.005283 | 0.005205 | 4.8e-05 | 0.001068 | 0.33035 | 1.3771 | 2.5475 | 100.79 | 21.07 | 15.515 | 86.8369 | 104.3351 | 83.6255 | 0.003849 | 0.98955 | 0.024155 | 24.4103 | 0.013565 | 0.14767 | 0.004575 | 0.32559 | 0.32559 | 0.018731
2019-03-26 16:25 | 1553631925.678 | 746 | 0.20862 | 0.001085 | 0.000131 | 98.9904 | 98.8979 | 122.3372 | 21.2612 | 0.002745 | 0.002612 | 220.6119 | 0.24803 | 8.4016 | 0.00347 | 0.005232 | 0.00517 | 5e-05 | 0.001066 | 0.31653 | 1.3596 | 2.5476 | 100.79 | 21.07 | 15.511 | 87.8505 | 105.3622 | 84.6468 | 0.003848 | 0.98672 | 0.02415 | 24.4747 | 0.008358 | 0.14755 | 0.004489 | 0.31931 | 0.31931 | 0.019085
2019-03-26 16:25 | 1553631926.6781 | 747 | 0.20862 | 0.00109 | 0.00013 | 99.0016 | 98.8929 | 122.4144 | 21.2612 | 0.002758 | 0.002589 | 220.5973 | 0.24864 | 8.4013 | 0.0034 | 0.004193 | 0.004136 | 5.1e-05 | 0.001071 | 0.32211 | 1.3664 | 2.5475 | 100.79 | 21.07 | 15.504 | 88.1777 | 103.6787 | 84.103 | 0.003847 | 0.98822 | 0.024176 | 24.4638 | 0.009828 | 0.14791 | 0.004521 | 0.32207 | 0.32207 | 0.018967
2019-03-26 16:25 | 1553631927.6781 | 748 | 0.20865 | 0.001092 | 0.000128 | 99.1055 | 98.7631 | 122.159 | 21.2612 | 0.002727 | 0.002574 | 220.5754 | 0.24871 | 8.4031 | 0.003472 | 0.005283 | 0.00522 | 4.9e-05 | 0.001064 | 0.31564 | 1.3571 | 2.5474 | 100.79 | 21.07 | 15.5 | 87.4272 | 105.4087 | 83.7419 | 0.003846 | 0.98922 | 0.024163 | 24.4259 | 0.030955 | 0.14796 | 0.004296 | 0.30556 | 0.30556 | 0.019998
2019-03-26 16:25 | 1553631928.6772 | 749 | 0.20864 | 0.001092 | 0.000126 | 99.0382 | 98.826 | 122.0806 | 21.2612 | 0.002705 | 0.002589 | 220.5745 | 0.24751 | 8.4048 | 0.003504 | 0.004825 | 0.004764 | 4.8e-05 | 0.001067 | 0.30525 | 1.3423 | 2.5476 | 100.79 | 21.07 | 15.497 | 86.9009 | 106.1758 | 84.0913 | 0.003845 | 0.98825 | 0.024143 | 24.43 | 0.019187 | 0.14724 | 0.004407 | 0.31358 | 0.31358 | 0.019393
2019-03-26 16:25 | 1553631929.6782 | 750 | 0.20863 | 0.001096 | 0.000123 | 99.0893 | 98.9688 | 122.5271 | 21.2612 | 0.00279 | 0.002566 | 220.5702 | 0.24821 | 8.4034 | 0.003423 | 0.004775 | 0.004684 | 5e-05 | 0.001067 | 0.32372 | 1.3426 | 2.5474 | 100.79 | 21.07 | 15.49 | 88.9478 | 104.236 | 83.5445 | 0.003843 | 0.98977 | 0.024206 | 24.4559 | 0.010876 | 0.14766 | 0.004434 | 0.3165 | 0.3165 | 0.019268
2019-03-26 16:25 | 1553631930.6783 | 751 | 0.20864 | 0.001104 | 0.000121 | 99.04 | 99.0939 | 122.8902 | 21.2612 | 0.00273 | 0.002538 | 220.5643 | 0.24779 | 8.4017 | 0.003479 | 0.004641 | 0.004565 | 5e-05 | 0.001066 | 0.31045 | 1.3234 | 2.5474 | 100.79 | 21.07 | 15.486 | 87.4929 | 105.585 | 82.8758 | 0.003842 | 0.99163 | 0.024264 | 24.4692 | -0.00487 | 0.14741 | 0.004348 | 0.31109 | 0.31109 | 0.01957
2019-03-26 16:25 | 1553631931.6774 | 752 | 0.20865 | 0.0011 | 0.00012 | 99.081 | 99.0065 | 123.0613 | 21.2612 | 0.0027 | 0.002535 | 220.5537 | 0.24827 | 8.4042 | 0.003406 | 0.005089 | 0.005086 | 4.7e-05 | 0.001069 | 0.3465 | 1.33 | 2.5474 | 100.8 | 21.07 | 15.48 | 86.7714 | 103.824 | 82.8095 | 0.00384 | 0.99181 | 0.024283 | 24.4839 | 0.006728 | 0.14769 | 0.004297 | 0.30763 | 0.30763 | 0.019828
2019-03-26 16:26 | 1553631932.6785 | 753 | 0.20863 | 0.001096 | 0.000119 | 99.089 | 98.9218 | 123.235 | 21.2612 | 0.002702 | 0.002592 | 220.5449 | 0.24929 | 8.4032 | 0.003401 | 0.004561 | 0.004404 | 4.9e-05 | 0.001065 | 0.32029 | 1.3445 | 2.5475 | 100.8 | 21.07 | 15.476 | 86.8159 | 103.6937 | 84.1764 | 0.003839 | 0.98802 | 0.024254 | 24.5482 | 0.015104 | 0.1483 | 0.004464 | 0.31958 | 0.31958 | 0.019166
2019-03-26 16:26 | 1553631933.6775 | 754 | 0.20863 | 0.001098 | 0.000117 | 99.091 | 98.9335 | 122.9259 | 21.2612 | 0.002701 | 0.00253 | 220.5427 | 0.24875 | 8.4026 | 0.003377 | 0.00504 | 0.004905 | 5.1e-05 | 0.001069 | 0.31526 | 1.3254 | 2.5474 | 100.8 | 21.07 | 15.473 | 86.7962 | 103.1206 | 82.6871 | 0.003838 | 0.99216 | 0.024274 | 24.4662 | 0.014227 | 0.14798 | 0.00442 | 0.31673 | 0.31673 | 0.019296
2019-03-26 16:26 | 1553631934.6779 | 755 | 0.20865 | 0.001093 | 0.000116 | 99.0653 | 98.8781 | 122.4893 | 21.2612 | 0.002715 | 0.002528 | 220.5514 | 0.24769 | 8.4038 | 0.003339 | 0.004332 | 0.004269 | 5e-05 | 0.001071 | 0.3418 | 1.346 | 2.5475 | 100.79 | 21.07 | 15.468 | 87.1374 | 102.1843 | 82.6457 | 0.003838 | 0.99227 | 0.024233 | 24.4213 | 0.016924 | 0.14735 | 0.004313 | 0.30843 | 0.30843 | 0.019732
2019-03-26 16:26 | 1553631935.6765 | 756 | 0.20864 | 0.001097 | 0.000115 | 99.1722 | 98.834 | 122.1849 | 21.2612 | 0.002681 | 0.00252 | 220.5305 | 0.24765 | 8.4026 | 0.003332 | 0.004605 | 0.004514 | 4.8e-05 | 0.001075 | 0.32388 | 1.3434 | 2.5475 | 100.79 | 21.07 | 15.466 | 86.3179 | 102.0153 | 82.4421 | 0.003837 | 0.99284 | 0.024209 | 24.384 | 0.030552 | 0.14733 | 0.004384 | 0.31335 | 0.31335 | 0.019418
2019-03-26 16:26 | 1553631936.6766 | 757 | 0.20866 | 0.001097 | 0.000114 | 99.1108 | 98.8787 | 122.0609 | 21.2612 | 0.00267 | 0.002594 | 220.5128 | 0.24922 | 8.4027 | 0.003329 | 0.004828 | 0.004732 | 4.6e-05 | 0.001076 | 0.32925 | 1.3572 | 2.5474 | 100.79 | 21.07 | 15.465 | 86.0484 | 101.955 | 84.2216 | 0.003837 | 0.98789 | 0.024137 | 24.4325 | 0.020967 | 0.14826 | 0.004281 | 0.30498 | 0.30498 | 0.020077
2019-03-26 16:26 | 1553631937.6767 | 758 | 0.20864 | 0.001098 | 0.000112 | 99.1208 | 99.0804 | 122.3093 | 21.2612 | 0.002621 | 0.002544 | 220.5087 | 0.2481 | 8.4052 | 0.00338 | 0.004099 | 0.004018 | 4.8e-05 | 0.001079 | 0.34562 | 1.3777 | 2.5476 | 100.79 | 21.07 | 15.463 | 84.8738 | 103.1765 | 83.0109 | 0.003836 | 0.99125 | 0.024202 | 24.4159 | 0.003646 | 0.14759 | 0.004392 | 0.31406 | 0.31406 | 0.019409
2019-03-26 16:26 | 1553631938.6768 | 759 | 0.20862 | 0.001099 | 0.000111 | 99.1757 | 98.9026 | 122.3137 | 21.2612 | 0.002734 | 0.002488 | 220.5103 | 0.24757 | 8.4034 | 0.003363 | 0.004945 | 0.004863 | 5.1e-05 | 0.001079 | 0.31902 | 1.3262 | 2.5475 | 100.79 | 21.07 | 15.462 | 87.592 | 102.7624 | 81.6686 | 0.003836 | 0.995 | 0.024248 | 24.3703 | 0.024667 | 0.14728 | 0.004502 | 0.32278 | 0.32278 | 0.018845
2019-03-26 16:26 | 1553631939.6777 | 760 | 0.20861 | 0.001097 | 0.000109 | 99.1731 | 99.0236 | 122.2069 | 21.2612 | 0.002759 | 0.002588 | 220.4979 | 0.24867 | 8.4012 | 0.003395 | 0.004346 | 0.004322 | 5.2e-05 | 0.001079 | 0.29264 | 1.3407 | 2.5475 | 100.79 | 21.07 | 15.461 | 88.1801 | 103.5576 | 84.0703 | 0.003836 | 0.98831 | 0.024156 | 24.4419 | 0.01349 | 0.14793 | 0.004569 | 0.32656 | 0.32656 | 0.018709
2019-03-26 16:26 | 1553631940.6778 | 761 | 0.20865 | 0.001104 | 0.000106 | 99.2233 | 99.0495 | 121.8985 | 21.2612 | 0.002683 | 0.002542 | 220.4841 | 0.24811 | 8.4014 | 0.003369 | 0.004634 | 0.004479 | 5e-05 | 0.001074 | 0.30558 | 1.3723 | 2.5476 | 100.79 | 21.07 | 15.46 | 86.3537 | 102.9259 | 82.9703 | 0.003836 | 0.99137 | 0.024163 | 24.3735 | 0.015677 | 0.1476 | 0.004322 | 0.30879 | 0.30879 | 0.019741
2019-03-26 16:26 | 1553631941.6772 | 762 | 0.20863 | 0.001109 | 0.000105 | 99.3121 | 99.0798 | 121.9119 | 21.2612 | 0.002717 | 0.002523 | 220.4759 | 0.2495 | 8.4026 | 0.00336 | 0.005172 | 0.005035 | 5.1e-05 | 0.00108 | 0.29878 | 1.3992 | 2.5475 | 100.79 | 21.07 | 15.46 | 87.1788 | 102.7043 | 82.5061 | 0.003836 | 0.99266 | 0.02418 | 24.3589 | 0.020951 | 0.14843 | 0.004451 | 0.31851 | 0.31851 | 0.019246
2019-03-26 16:26 | 1553631942.6769 | 763 | 0.20861 | 0.001108 | 0.000104 | 99.2256 | 99.0441 | 121.8635 | 21.2612 | 0.002683 | 0.002586 | 220.4379 | 0.24927 | 8.4035 | 0.003376 | 0.004922 | 0.004939 | 5.1e-05 | 0.001082 | 0.34346 | 1.4137 | 2.5476 | 100.8 | 21.07 | 15.46 | 86.3708 | 103.094 | 84.0273 | 0.003835 | 0.98843 | 0.024124 | 24.4061 | 0.016375 | 0.14829 | 0.00455 | 0.32505 | 0.32505 | 0.018842
2019-03-26 16:26 | 1553631943.678 | 764 | 0.20865 | 0.00111 | 0.000103 | 99.2197 | 99.2415 | 122.1684 | 21.2612 | 0.002733 | 0.002577 | 220.4275 | 0.24844 | 8.4007 | 0.003363 | 0.00457 | 0.004501 | 4.7e-05 | 0.00108 | 0.3361 | 1.3547 | 2.5475 | 100.8 | 21.08 | 15.459 | 87.5733 | 102.7796 | 83.8152 | 0.003837 | 0.98902 | 0.024161 | 24.4294 | -0.00196 | 0.1478 | 0.004333 | 0.30979 | 0.30979 | 0.019704
2019-03-26 16:26 | 1553631944.678 | 765 | 0.20865 | 0.001116 | 0.000102 | 99.2416 | 99.0586 | 122.0666 | 21.2612 | 0.002721 | 0.002548 | 220.3955 | 0.24849 | 8.4032 | 0.003379 | 0.005109 | 0.004991 | 4.7e-05 | 0.001081 | 0.34379 | 1.3473 | 2.5474 | 100.8 | 21.08 | 15.458 | 87.2734 | 103.1601 | 83.1076 | 0.003837 | 0.99098 | 0.024175 | 24.395 | 0.016507 | 0.14783 | 0.004335 | 0.3102 | 0.3102 | 0.019681
2019-03-26 16:26 | 1553631945.677 | 766 | 0.20872 | 0.001122 | 0.0001 | 99.2979 | 98.8586 | 122.036 | 21.2612 | 0.002668 | 0.002536 | 220.3847 | 0.24799 | 8.4037 | 0.003419 | 0.004878 | 0.004787 | 5e-05 | 0.001089 | 0.34595 | 1.3311 | 2.5474 | 100.8 | 21.08 | 15.457 | 86.0017 | 104.115 | 82.8167 | 0.003837 | 0.99179 | 0.024182 | 24.3819 | 0.039666 | 0.14753 | 0.003901 | 0.2787 | 0.2787 | 0.021862
2019-03-26 16:26 | 1553631946.6771 | 767 | 0.20873 | 0.001119 | 9.9e-05 | 99.2463 | 99.1598 | 121.9527 | 21.2612 | 0.002669 | 0.002564 | 220.3838 | 0.24815 | 8.4029 | 0.003382 | 0.005076 | 0.005101 | 4.9e-05 | 0.00108 | 0.32173 | 1.3628 | 2.5474 | 100.8 | 21.08 | 15.456 | 86.0143 | 103.2197 | 83.5 | 0.003837 | 0.98989 | 0.02415 | 24.397 | 0.0078 | 0.14762 | 0.003865 | 0.27577 | 0.27577 | 0.022108
2019-03-26 16:26 | 1553631947.6781 | 768 | 0.20866 | 0.001118 | 9.6e-05 | 99.2177 | 99.3877 | 121.9511 | 21.2612 | 0.002725 | 0.002516 | 220.3849 | 0.24945 | 8.4006 | 0.003377 | 0.004971 | 0.004832 | 4.8e-05 | 0.001088 | 0.31133 | 1.3766 | 2.5475 | 100.8 | 21.07 | 15.451 | 87.3803 | 103.1121 | 82.3479 | 0.003833 | 0.9931 | 0.024189 | 24.3574 | -0.015316 | 0.1484 | 0.004275 | 0.30629 | 0.30629 | 0.02001
2019-03-26 16:26 | 1553631948.6772 | 769 | 0.20861 | 0.001126 | 9.5e-05 | 99.3351 | 99.141 | 121.7044 | 21.2612 | 0.002757 | 0.002606 | 220.3575 | 0.24876 | 8.4042 | 0.003354 | 0.005207 | 0.005075 | 4.5e-05 | 0.001091 | 0.31526 | 1.3648 | 2.5476 | 100.8 | 21.07 | 15.446 | 88.1425 | 102.5573 | 84.5188 | 0.003832 | 0.98707 | 0.024091 | 24.4069 | 0.017503 | 0.14798 | 0.004566 | 0.32629 | 0.32629 | 0.018731
2019-03-26 16:26 | 1553631949.6778 | 770 | 0.20861 | 0.001126 | 9.4e-05 | 99.267 | 98.9295 | 121.576 | 21.2612 | 0.002694 | 0.002523 | 220.3659 | 0.24831 | 8.4049 | 0.003351 | 0.005123 | 0.005123 | 4.8e-05 | 0.001087 | 0.33411 | 1.3486 | 2.5475 | 100.8 | 21.07 | 15.438 | 86.6206 | 102.4828 | 82.5133 | 0.00383 | 0.99264 | 0.024147 | 24.3256 | 0.030458 | 0.14772 | 0.004564 | 0.32694 | 0.32694 | 0.01866
2019-03-26 16:26 | 1553631950.6774 | 771 | 0.20863 | 0.001129 | 9.3e-05 | 99.2802 | 99.1352 | 121.5955 | 21.2612 | 0.002772 | 0.002577 | 220.3607 | 0.24896 | 8.4057 | 0.003363 | 0.004757 | 0.004638 | 4.8e-05 | 0.001081 | 0.3063 | 1.3448 | 2.5474 | 100.8 | 21.07 | 15.434 | 88.5135 | 102.7648 | 83.8111 | 0.003829 | 0.98903 | 0.024105 | 24.3719 | 0.013073 | 0.1481 | 0.004427 | 0.31648 | 0.31648 | 0.019327
2019-03-26 16:26 | 1553631951.6773 | 772 | 0.20859 | 0.001133 | 9.2e-05 | 99.2389 | 98.9319 | 121.5105 | 21.2612 | 0.002723 | 0.002526 | 220.35 | 0.24846 | 8.4022 | 0.00341 | 0.004736 | 0.004695 | 4.5e-05 | 0.00108 | 0.33644 | 1.3781 | 2.5475 | 100.8 | 21.07 | 15.43 | 87.3326 | 103.9185 | 82.5981 | 0.003828 | 0.9924 | 0.024137 | 24.3219 | 0.027718 | 0.14781 | 0.004675 | 0.33508 | 0.33508 | 0.018218
2019-03-26 16:26 | 1553631952.6774 | 773 | 0.20857 | 0.001133 | 9e-05 | 99.316 | 99.0801 | 121.5042 | 21.2612 | 0.002706 | 0.002539 | 220.3665 | 0.24795 | 8.4061 | 0.003386 | 0.004694 | 0.00461 | 4.8e-05 | 0.001075 | 0.3079 | 1.3765 | 2.5475 | 100.8 | 21.07 | 15.423 | 86.9079 | 103.3194 | 82.9032 | 0.003826 | 0.99155 | 0.024126 | 24.3317 | 0.021266 | 0.14751 | 0.004812 | 0.34495 | 0.34495 | 0.017661
2019-03-26 16:26 | 1553631953.6764 | 774 | 0.20855 | 0.001136 | 8.9e-05 | 99.3086 | 99.1762 | 121.7046 | 21.2612 | 0.00272 | 0.002511 | 220.3246 | 0.24906 | 8.4027 | 0.003473 | 0.004685 | 0.004575 | 4.6e-05 | 0.001073 | 0.33682 | 1.4183 | 2.5475 | 100.8 | 21.07 | 15.418 | 87.264 | 105.4456 | 82.2315 | 0.003825 | 0.99343 | 0.024169 | 24.3288 | 0.011926 | 0.14817 | 0.004939 | 0.35491 | 0.35491 | 0.017242
2019-03-26 16:26 | 1553631954.6765 | 775 | 0.2086 | 0.001143 | 8.8e-05 | 99.2569 | 98.9567 | 122.183 | 21.2612 | 0.002715 | 0.002504 | 220.3203 | 0.24832 | 8.4038 | 0.003506 | 0.004706 | 0.004632 | 4.7e-05 | 0.001072 | 0.30574 | 1.3672 | 2.5476 | 100.8 | 21.07 | 15.415 | 87.129 | 106.2372 | 82.0686 | 0.003824 | 0.99388 | 0.024222 | 24.371 | 0.027092 | 0.14773 | 0.004622 | 0.33264 | 0.33264 | 0.018342
2019-03-26 16:26 | 1553631955.6766 | 776 | 0.20856 | 0.001146 | 8.7e-05 | 99.2687 | 99.2499 | 122.4186 | 21.2612 | 0.002667 | 0.002537 | 220.298 | 0.2478 | 8.4 | 0.003566 | 0.004923 | 0.004868 | 4.7e-05 | 0.001069 | 0.31332 | 1.3859 | 2.5476 | 100.8 | 21.07 | 15.411 | 85.9892 | 107.701 | 82.8587 | 0.003823 | 0.99168 | 0.024218 | 24.4216 | 0.001699 | 0.14741 | 0.004853 | 0.34948 | 0.34948 | 0.017421
2019-03-26 16:26 | 1553631956.6776 | 777 | 0.2086 | 0.001143 | 8.5e-05 | 99.3 | 99.1557 | 122.5496 | 21.2612 | 0.002674 | 0.00253 | 220.2684 | 0.24823 | 8.4036 | 0.003548 | 0.004426 | 0.004326 | 4.8e-05 | 0.00107 | 0.29878 | 1.3749 | 2.5476 | 100.8 | 21.07 | 15.404 | 86.1492 | 107.2631 | 82.6808 | 0.003821 | 0.99217 | 0.024237 | 24.4285 | 0.013005 | 0.14768 | 0.004642 | 0.33447 | 0.33447 | 0.018235
2019-03-26 16:26 | 1553631957.6776 | 778 | 0.2086 | 0.001151 | 8.5e-05 | 99.3741 | 99.1668 | 122.3163 | 21.2612 | 0.002753 | 0.002582 | 220.2437 | 0.24823 | 8.3998 | 0.003501 | 0.004946 | 0.004907 | 4.7e-05 | 0.001073 | 0.347 | 1.3814 | 2.5475 | 100.8 | 21.07 | 15.4 | 88.0378 | 106.1176 | 83.9306 | 0.00382 | 0.9887 | 0.024172 | 24.4481 | 0.018676 | 0.14767 | 0.00464 | 0.33336 | 0.33336 | 0.018295
2019-03-26 16:26 | 1553631958.6768 | 779 | 0.20858 | 0.001145 | 8.4e-05 | 99.3623 | 99.4052 | 122.501 | 21.2612 | 0.002675 | 0.002588 | 220.2305 | 0.2489 | 8.4009 | 0.003539 | 0.004749 | 0.004511 | 4.8e-05 | 0.001066 | 0.32244 | 1.3793 | 2.5476 | 100.8 | 21.07 | 15.397 | 86.1819 | 107.0362 | 84.0867 | 0.00382 | 0.98827 | 0.024185 | 24.4719 | -0.003859 | 0.14807 | 0.004718 | 0.33936 | 0.33936 | 0.01802
2019-03-26 16:26 | 1553631959.6777 | 780 | 0.20858 | 0.001153 | 8.4e-05 | 99.3768 | 99.2065 | 122.4729 | 21.2612 | 0.002735 | 0.002501 | 220.2205 | 0.24926 | 8.4027 | 0.003506 | 0.004977 | 0.005052 | 5e-05 | 0.001064 | 0.31454 | 1.3956 | 2.5476 | 100.8 | 21.07 | 15.391 | 87.6192 | 106.2484 | 81.9952 | 0.003818 | 0.99409 | 0.024253 | 24.3973 | 0.015338 | 0.14829 | 0.004749 | 0.34258 | 0.34258 | 0.017877
2019-03-26 16:26 | 1553631960.6767 | 781 | 0.20858 | 0.001148 | 8.3e-05 | 99.4016 | 99.2431 | 122.3536 | 21.2612 | 0.002665 | 0.002562 | 220.1936 | 0.2479 | 8.4022 | 0.003515 | 0.00524 | 0.005085 | 4.8e-05 | 0.001059 | 0.33395 | 1.3542 | 2.5474 | 100.8 | 21.07 | 15.39 | 85.926 | 106.4662 | 83.4548 | 0.003818 | 0.99002 | 0.024192 | 24.4355 | 0.014266 | 0.14748 | 0.004765 | 0.34293 | 0.34293 | 0.017761
2019-03-26 16:26 | 1553631961.6768 | 782 | 0.20858 | 0.001143 | 8.3e-05 | 99.3458 | 99.2989 | 122.4255 | 21.2612 | 0.002622 | 0.002508 | 220.1862 | 0.24818 | 8.4004 | 0.003505 | 0.00473 | 0.004697 | 5e-05 | 0.001054 | 0.323 | 1.3484 | 2.5474 | 100.8 | 21.07 | 15.387 | 84.8957 | 106.2081 | 82.1488 | 0.003817 | 0.99366 | 0.024243 | 24.3979 | 0.004217 | 0.14764 | 0.004735 | 0.3414 | 0.3414 | 0.01786
2019-03-26 16:26 | 1553631962.6773 | 783 | 0.20857 | 0.001148 | 8.4e-05 | 99.4187 | 99.1319 | 122.5489 | 21.2612 | 0.002739 | 0.002513 | 220.1768 | 0.24817 | 8.4029 | 0.003512 | 0.005148 | 0.005 | 4.6e-05 | 0.001047 | 0.32244 | 1.3572 | 2.5475 | 100.8 | 21.07 | 15.386 | 87.7083 | 106.3876 | 82.2808 | 0.003817 | 0.99329 | 0.024251 | 24.4147 | 0.025841 | 0.14764 | 0.004774 | 0.34436 | 0.34436 | 0.017706
2019-03-26 16:26 | 1553631963.6779 | 784 | 0.20857 | 0.001147 | 8.4e-05 | 99.3735 | 99.3669 | 122.3013 | 21.2612 | 0.002658 | 0.002548 | 220.1955 | 0.24745 | 8.4043 | 0.003476 | 0.005235 | 0.005132 | 4.8e-05 | 0.001047 | 0.32997 | 1.3418 | 2.5474 | 100.8 | 21.07 | 15.384 | 85.7496 | 105.5038 | 83.1153 | 0.003816 | 0.99096 | 0.024198 | 24.4187 | 0.0006 | 0.14721 | 0.004823 | 0.34723 | 0.34723 | 0.017509
2019-03-26 16:26 | 1553631964.678 | 785 | 0.20858 | 0.00115 | 8.4e-05 | 99.4042 | 99.3366 | 122.0799 | 21.2612 | 0.002694 | 0.002501 | 220.2002 | 0.24844 | 8.4038 | 0.003459 | 0.005484 | 0.005404 | 4.9e-05 | 0.001045 | 0.2907 | 1.3611 | 2.5475 | 100.8 | 21.07 | 15.38 | 86.6372 | 105.0972 | 81.982 | 0.003815 | 0.99412 | 0.024215 | 24.3577 | 0.006086 | 0.1478 | 0.004739 | 0.34129 | 0.34129 | 0.017885
2019-03-26 16:26 | 1553631965.677 | 786 | 0.20855 | 0.001152 | 8.3e-05 | 99.4083 | 99.5319 | 122.0386 | 21.2612 | 0.002707 | 0.002559 | 220.2001 | 0.24942 | 8.4046 | 0.003535 | 0.005244 | 0.00516 | 4.8e-05 | 0.001042 | 0.34003 | 1.3448 | 2.5475 | 100.8 | 21.07 | 15.378 | 86.9509 | 106.9418 | 83.3814 | 0.003815 | 0.99022 | 0.024163 | 24.4015 | -0.011111 | 0.14838 | 0.004942 | 0.35541 | 0.35541 | 0.017242
2019-03-26 16:26 | 1553631966.6771 | 787 | 0.20855 | 0.001156 | 8.2e-05 | 99.3257 | 99.1799 | 122.1403 | 21.2612 | 0.002713 | 0.002557 | 220.1962 | 0.24804 | 8.4029 | 0.003426 | 0.005047 | 0.004979 | 4.6e-05 | 0.001036 | 0.34612 | 1.3518 | 2.5474 | 100.8 | 21.07 | 15.374 | 87.0929 | 104.2934 | 83.3384 | 0.003814 | 0.99034 | 0.024174 | 24.4102 | 0.013135 | 0.14756 | 0.004897 | 0.35236 | 0.35236 | 0.017296
2019-03-26 16:26 | 1553631967.6763 | 788 | 0.20859 | 0.001158 | 8.1e-05 | 99.4013 | 99.2586 | 122.5034 | 21.2612 | 0.002712 | 0.00256 | 220.1854 | 0.249 | 8.4001 | 0.003517 | 0.005498 | 0.005397 | 4.8e-05 | 0.001026 | 0.32781 | 1.3316 | 2.5474 | 100.8 | 21.08 | 15.373 | 87.0619 | 106.5128 | 83.4024 | 0.003816 | 0.99016 | 0.024208 | 24.4487 | 0.012846 | 0.14813 | 0.004691 | 0.33786 | 0.33786 | 0.018108
2019-03-26 16:26 | 1553631968.6772 | 789 | 0.20858 | 0.001161 | 8e-05 | 99.4287 | 99.185 | 122.447 | 21.2612 | 0.002734 | 0.002493 | 220.1961 | 0.24873 | 8.4007 | 0.003584 | 0.005055 | 0.004987 | 4.7e-05 | 0.001022 | 0.31614 | 1.3699 | 2.5475 | 100.8 | 21.08 | 15.372 | 87.5911 | 108.1184 | 81.7909 | 0.003816 | 0.99466 | 0.024258 | 24.3877 | 0.021952 | 0.14797 | 0.004732 | 0.3416 | 0.3416 | 0.01789
2019-03-26 16:26 | 1553631969.6762 | 790 | 0.20859 | 0.001171 | 8e-05 | 99.4724 | 99.2885 | 122.3419 | 21.2612 | 0.002694 | 0.002528 | 220.1583 | 0.24786 | 8.4031 | 0.003603 | 0.005154 | 0.005115 | 4.9e-05 | 0.001016 | 0.31332 | 1.3445 | 2.5476 | 100.8 | 21.08 | 15.368 | 86.6184 | 108.5874 | 82.6465 | 0.003815 | 0.99227 | 0.024218 | 24.4066 | 0.016549 | 0.14745 | 0.004659 | 0.33573 | 0.33573 | 0.018139
2019-03-26 16:26 | 1553631970.6773 | 791 | 0.20857 | 0.001163 | 8.1e-05 | 99.4591 | 99.5114 | 122.3668 | 21.2612 | 0.002635 | 0.002535 | 220.1652 | 0.24899 | 8.4063 | 0.003536 | 0.004635 | 0.004548 | 4.4e-05 | 0.001025 | 0.34506 | 1.3825 | 2.5475 | 100.8 | 21.07 | 15.367 | 85.2055 | 106.9748 | 82.7972 | 0.003812 | 0.99185 | 0.024215 | 24.4143 | -0.004701 | 0.14812 | 0.004767 | 0.34352 | 0.34352 | 0.017808
2019-03-26 16:26 | 1553631971.6763 | 792 | 0.20859 | 0.001171 | 8e-05 | 99.5385 | 99.4129 | 122.0332 | 21.2612 | 0.002755 | 0.00257 | 220.1367 | 0.24819 | 8.4007 | 0.003441 | 0.005072 | 0.004934 | 4.5e-05 | 0.001024 | 0.35115 | 1.3634 | 2.5475 | 100.8 | 21.07 | 15.363 | 88.0941 | 104.6567 | 83.6334 | 0.003811 | 0.98952 | 0.024154 | 24.4096 | 0.011295 | 0.14765 | 0.004676 | 0.33609 | 0.33609 | 0.018144
2019-03-26 16:26 | 1553631972.6764 | 793 | 0.20857 | 0.001171 | 8.1e-05 | 99.4916 | 99.4664 | 121.7638 | 21.2612 | 0.002748 | 0.002519 | 220.1075 | 0.2481 | 8.4037 | 0.003521 | 0.004995 | 0.004918 | 4.4e-05 | 0.001021 | 0.3418 | 1.355 | 2.5477 | 100.8 | 21.07 | 15.357 | 87.916 | 106.5924 | 82.4174 | 0.00381 | 0.99291 | 0.024168 | 24.3411 | 0.002269 | 0.14759 | 0.004803 | 0.34549 | 0.34549 | 0.017643
2019-03-26 16:26 | 1553631973.6764 | 794 | 0.20856 | 0.001167 | 8.1e-05 | 99.5249 | 99.3739 | 121.5697 | 21.2612 | 0.002718 | 0.002526 | 220.0879 | 0.2479 | 8.402 | 0.003494 | 0.005285 | 0.00518 | 4.5e-05 | 0.001023 | 0.33428 | 1.3464 | 2.5475 | 100.8 | 21.07 | 15.352 | 87.1952 | 105.9381 | 82.5919 | 0.003808 | 0.99242 | 0.024143 | 24.3276 | 0.013584 | 0.14747 | 0.004859 | 0.34915 | 0.34915 | 0.017444
2019-03-26 16:26 | 1553631974.6765 | 795 | 0.20855 | 0.001169 | 8.2e-05 | 99.6268 | 99.584 | 121.5725 | 21.2612 | 0.002665 | 0.002488 | 220.0867 | 0.24846 | 8.4037 | 0.00348 | 0.004862 | 0.00481 | 4.5e-05 | 0.001024 | 0.31708 | 1.3466 | 2.5476 | 100.8 | 21.07 | 15.346 | 85.9291 | 105.5976 | 81.6811 | 0.003807 | 0.99497 | 0.024174 | 24.2967 | 0.003841 | 0.14781 | 0.004884 | 0.35145 | 0.35145 | 0.017369
2019-03-26 16:26 | 1553631975.6765 | 796 | 0.20855 | 0.00117 | 8.1e-05 | 99.5618 | 99.6445 | 121.5542 | 21.2612 | 0.002698 | 0.002521 | 220.0797 | 0.24817 | 8.4017 | 0.003489 | 0.005205 | 0.005104 | 4.4e-05 | 0.001024 | 0.34235 | 1.3324 | 2.5474 | 100.8 | 21.07 | 15.342 | 86.7252 | 105.8336 | 82.4585 | 0.003806 | 0.99279 | 0.024146 | 24.3215 | -0.007425 | 0.14764 | 0.00492 | 0.35368 | 0.35368 | 0.01724
2019-03-26 16:26 | 1553631976.678 | 797 | 0.20854 | 0.001175 | 8.1e-05 | 99.5424 | 99.5581 | 121.4321 | 21.2612 | 0.002686 | 0.002535 | 220.0878 | 0.24883 | 8.4006 | 0.003536 | 0.005077 | 0.004995 | 4.4e-05 | 0.001031 | 0.32068 | 1.322 | 2.5473 | 100.8 | 21.07 | 15.338 | 86.4271 | 106.9726 | 82.8095 | 0.003805 | 0.99181 | 0.024122 | 24.3213 | -0.001417 | 0.14803 | 0.004933 | 0.35425 | 0.35425 | 0.017258
2019-03-26 16:26 | 1553631977.678 | 798 | 0.20855 | 0.001178 | 8e-05 | 99.5199 | 99.4603 | 121.4781 | 21.2612 | 0.002697 | 0.002484 | 220.0884 | 0.24784 | 8.4036 | 0.003514 | 0.004746 | 0.004687 | 4.7e-05 | 0.001029 | 0.35081 | 1.3578 | 2.5476 | 100.8 | 21.07 | 15.332 | 86.6963 | 106.4419 | 81.5828 | 0.003804 | 0.99524 | 0.024168 | 24.2839 | 0.00536 | 0.14744 | 0.004874 | 0.35073 | 0.35073 | 0.017362
2019-03-26 16:26 | 1553631978.6767 | 799 | 0.20853 | 0.001177 | 8e-05 | 99.6167 | 99.4452 | 121.5971 | 21.2612 | 0.002687 | 0.002526 | 220.0794 | 0.24902 | 8.4064 | 0.003616 | 0.00466 | 0.004604 | 4.3e-05 | 0.001028 | 0.31277 | 1.3252 | 2.5475 | 100.8 | 21.07 | 15.329 | 86.4584 | 108.9119 | 82.5779 | 0.003803 | 0.99246 | 0.024146 | 24.3299 | 0.015419 | 0.14814 | 0.005001 | 0.35966 | 0.35966 | 0.017012
2019-03-26 16:26 | 1553631979.6777 | 800 | 0.20854 | 0.001186 | 7.9e-05 | 99.5896 | 99.618 | 121.4948 | 21.2612 | 0.002671 | 0.002475 | 220.0805 | 0.2479 | 8.4018 | 0.003528 | 0.004567 | 0.004518 | 4.9e-05 | 0.001032 | 0.34938 | 1.348 | 2.5474 | 100.8 | 21.07 | 15.324 | 86.0742 | 106.7765 | 81.3636 | 0.003802 | 0.99586 | 0.024178 | 24.2781 | -0.002548 | 0.14748 | 0.004918 | 0.35419 | 0.35419 | 0.017196
2019-03-26 16:26 | 1553631980.6768 | 801 | 0.20851 | 0.00119 | 7.7e-05 | 99.5917 | 99.5201 | 121.4443 | 21.2612 | 0.002708 | 0.002573 | 220.0751 | 0.24831 | 8.4053 | 0.003552 | 0.004943 | 0.004699 | 4.6e-05 | 0.001033 | 0.32333 | 1.3609 | 2.5475 | 100.8 | 21.07 | 15.322 | 86.9603 | 107.3488 | 83.7121 | 0.003801 | 0.9893 | 0.024093 | 24.3533 | 0.006425 | 0.14772 | 0.005095 | 0.36585 | 0.36585 | 0.016676
2019-03-26 16:26 | 1553631981.6778 | 802 | 0.20855 | 0.001192 | 7.5e-05 | 99.5503 | 99.6227 | 121.4988 | 21.2612 | 0.002695 | 0.002494 | 220.0646 | 0.24911 | 8.4022 | 0.003593 | 0.005146 | 0.00513 | 4.7e-05 | 0.001033 | 0.33804 | 1.4211 | 2.5476 | 100.8 | 21.07 | 15.322 | 86.6449 | 108.3494 | 81.8284 | 0.003801 | 0.99455 | 0.024162 | 24.2944 | -0.006497 | 0.14819 | 0.004842 | 0.34862 | 0.34862 | 0.017556
2019-03-26 16:26 | 1553631982.6768 | 803 | 0.20856 | 0.001195 | 7.5e-05 | 99.6324 | 99.6254 | 121.3997 | 21.2612 | 0.002684 | 0.002496 | 220.0699 | 0.24816 | 8.4013 | 0.003575 | 0.005389 | 0.005372 | 4.3e-05 | 0.001034 | 0.30934 | 1.3244 | 2.5474 | 100.8 | 21.07 | 15.322 | 86.392 | 107.9127 | 81.8549 | 0.003801 | 0.99448 | 0.024151 | 24.2854 | 0.000631 | 0.14763 | 0.004808 | 0.34603 | 0.34603 | 0.017621
2019-03-26 16:26 | 1553631983.6759 | 804 | 0.20857 | 0.001191 | 7.3e-05 | 99.6835 | 99.538 | 121.5621 | 21.2612 | 0.002636 | 0.002514 | 220.0782 | 0.24925 | 8.4033 | 0.003562 | 0.005602 | 0.00552 | 4.7e-05 | 0.001031 | 0.30055 | 1.3207 | 2.5475 | 100.8 | 21.07 | 15.322 | 85.2291 | 107.6025 | 82.2972 | 0.003801 | 0.99324 | 0.024152 | 24.3168 | 0.013065 | 0.14828 | 0.004751 | 0.34194 | 0.34194 | 0.017909
2019-03-26 16:26 | 1553631984.677 | 805 | 0.20857 | 0.001189 | 7.2e-05 | 99.6598 | 99.6146 | 121.0692 | 21.2612 | 0.002668 | 0.002529 | 220.0692 | 0.24875 | 8.4039 | 0.003597 | 0.004687 | 0.00459 | 4.4e-05 | 0.001032 | 0.32156 | 1.3222 | 2.5475 | 100.8 | 21.07 | 15.322 | 86.0102 | 108.4455 | 82.6513 | 0.003801 | 0.99225 | 0.024091 | 24.2795 | 0.004061 | 0.14798 | 0.004784 | 0.34352 | 0.34352 | 0.017791
2019-03-26 16:26 | 1553631985.677 | 806 | 0.20859 | 0.001191 | 7.2e-05 | 99.6784 | 99.6166 | 120.8007 | 21.2612 | 0.002659 | 0.002501 | 220.0565 | 0.24838 | 8.3988 | 0.003618 | 0.004818 | 0.004652 | 4.5e-05 | 0.001033 | 0.34075 | 1.3313 | 2.5475 | 100.8 | 21.07 | 15.321 | 85.7893 | 108.9632 | 81.989 | 0.003801 | 0.99411 | 0.024087 | 24.23 | 0.005549 | 0.14776 | 0.004619 | 0.3315 | 0.3315 | 0.018409
2019-03-26 16:26 | 1553631986.676 | 807 | 0.20855 | 0.001191 | 7.1e-05 | 99.7142 | 99.5733 | 120.9171 | 21.2612 | 0.002691 | 0.002484 | 220.0129 | 0.24795 | 8.402 | 0.003551 | 0.005175 | 0.005135 | 4.6e-05 | 0.001044 | 0.35778 | 1.352 | 2.5476 | 100.8 | 21.07 | 15.321 | 86.5534 | 107.326 | 81.589 | 0.003801 | 0.99523 | 0.024112 | 24.228 | 0.012648 | 0.14751 | 0.004903 | 0.35262 | 0.35262 | 0.017277
2019-03-26 16:26 | 1553631987.6765 | 808 | 0.20859 | 0.001187 | 7.1e-05 | 99.661 | 99.6012 | 120.8648 | 21.2612 | 0.002707 | 0.002507 | 220.0051 | 0.24872 | 8.4042 | 0.00354 | 0.005317 | 0.005201 | 4.6e-05 | 0.001042 | 0.32444 | 1.3343 | 2.5475 | 100.8 | 21.07 | 15.318 | 86.9502 | 107.0512 | 82.1224 | 0.0038 | 0.99373 | 0.024089 | 24.241 | 0.005375 | 0.14796 | 0.004631 | 0.33249 | 0.33249 | 0.018379
2019-03-26 16:26 | 1553631988.6762 | 809 | 0.20857 | 0.001185 | 7.2e-05 | 99.726 | 99.4331 | 120.8805 | 21.2612 | 0.00266 | 0.002497 | 219.9996 | 0.24826 | 8.4018 | 0.003558 | 0.005258 | 0.005177 | 4.6e-05 | 0.001033 | 0.33052 | 1.3888 | 2.5476 | 100.8 | 21.07 | 15.316 | 85.8151 | 107.4898 | 81.8791 | 0.0038 | 0.99441 | 0.024099 | 24.2343 | 0.026308 | 0.14769 | 0.004762 | 0.34209 | 0.34209 | 0.01783
2019-03-26 16:26 | 1553631989.6762 | 810 | 0.20852 | 0.00118 | 7.2e-05 | 99.7147 | 99.5054 | 121.3626 | 21.2612 | 0.002646 | 0.002513 | 219.9962 | 0.24854 | 8.4037 | 0.003625 | 0.005115 | 0.005058 | 4.6e-05 | 0.001028 | 0.35275 | 1.3371 | 2.5475 | 100.8 | 21.07 | 15.314 | 85.4795 | 109.1303 | 82.2854 | 0.003799 | 0.99328 | 0.024133 | 24.2964 | 0.018803 | 0.14786 | 0.005071 | 0.36502 | 0.36502 | 0.016729
2019-03-26 16:26 | 1553631990.6772 | 811 | 0.20851 | 0.001186 | 7.2e-05 | 99.7127 | 99.6035 | 121.7027 | 21.2612 | 0.002688 | 0.00253 | 219.9811 | 0.248 | 8.4028 | 0.003555 | 0.004817 | 0.004751 | 4.8e-05 | 0.00103 | 0.34612 | 1.3886 | 2.5475 | 100.8 | 21.07 | 15.309 | 86.4747 | 107.427 | 82.6808 | 0.003798 | 0.99217 | 0.024153 | 24.344 | 0.009799 | 0.14753 | 0.005114 | 0.36847 | 0.36847 | 0.016536
2019-03-26 16:26 | 1553631991.6763 | 812 | 0.20855 | 0.001187 | 7.2e-05 | 99.6802 | 99.4627 | 121.6391 | 21.2612 | 0.002699 | 0.002488 | 219.9785 | 0.24859 | 8.4029 | 0.003555 | 0.004898 | 0.004841 | 4.9e-05 | 0.00103 | 0.32173 | 1.3374 | 2.5475 | 100.8 | 21.07 | 15.305 | 86.7551 | 107.4347 | 81.6794 | 0.003797 | 0.99497 | 0.024181 | 24.3033 | 0.019543 | 0.14788 | 0.004908 | 0.35386 | 0.35386 | 0.01726
2019-03-26 16:27 | 1553631992.6763 | 813 | 0.20855 | 0.001189 | 7.2e-05 | 99.7215 | 99.6395 | 121.4969 | 21.2612 | 0.002688 | 0.002509 | 219.9709 | 0.24867 | 8.4044 | 0.003591 | 0.005291 | 0.00515 | 4.9e-05 | 0.00103 | 0.35906 | 1.3103 | 2.5474 | 100.8 | 21.06 | 15.295 | 86.4723 | 108.2967 | 82.1777 | 0.003792 | 0.99358 | 0.02415 | 24.3062 | 0.007363 | 0.14793 | 0.0049 | 0.3529 | 0.3529 | 0.017313
2019-03-26 16:27 | 1553631993.6765 | 814 | 0.20853 | 0.00119 | 7.2e-05 | 99.7897 | 99.6771 | 121.4535 | 21.2612 | 0.002634 | 0.002536 | 219.9229 | 0.249 | 8.4005 | 0.0036 | 0.005332 | 0.005235 | 4.7e-05 | 0.001033 | 0.32731 | 1.3528 | 2.5474 | 100.8 | 21.06 | 15.29 | 85.193 | 108.5235 | 82.8282 | 0.003791 | 0.99176 | 0.024124 | 24.3241 | 0.010098 | 0.14813 | 0.005022 | 0.36134 | 0.36134 | 0.016931
2019-03-26 16:27 | 1553631994.6774 | 815 | 0.20855 | 0.001196 | 7.2e-05 | 99.7525 | 99.7094 | 121.7208 | 21.2612 | 0.002647 | 0.002519 | 219.9076 | 0.24808 | 8.4038 | 0.003574 | 0.005037 | 0.004972 | 4.7e-05 | 0.001038 | 0.35131 | 1.3776 | 2.5475 | 100.8 | 21.06 | 15.286 | 85.5006 | 107.8773 | 82.4157 | 0.00379 | 0.99291 | 0.024164 | 24.3367 | 0.003868 | 0.14758 | 0.004865 | 0.35055 | 0.35055 | 0.017387
2019-03-26 16:27 | 1553631995.6765 | 816 | 0.2085 | 0.0012 | 7.2e-05 | 99.7295 | 99.626 | 121.9342 | 21.2612 | 0.002681 | 0.002504 | 219.9222 | 0.24855 | 8.4015 | 0.003529 | 0.004787 | 0.004665 | 4.6e-05 | 0.00103 | 0.28567 | 1.3294 | 2.5474 | 100.8 | 21.06 | 15.28 | 86.3061 | 106.786 | 82.0522 | 0.003788 | 0.99393 | 0.024198 | 24.3456 | 0.009285 | 0.14786 | 0.00514 | 0.37113 | 0.37113 | 0.016454
2019-03-26 16:27 | 1553631996.6776 | 817 | 0.20852 | 0.001204 | 7.1e-05 | 99.7756 | 99.7921 | 121.9548 | 21.2612 | 0.002703 | 0.002503 | 219.9199 | 0.24828 | 8.4025 | 0.003522 | 0.005117 | 0.005007 | 4.5e-05 | 0.00103 | 0.33732 | 1.3188 | 2.5475 | 100.8 | 21.06 | 15.278 | 86.8354 | 106.6191 | 82.0272 | 0.003788 | 0.994 | 0.024201 | 24.3468 | -0.001483 | 0.1477 | 0.005074 | 0.36634 | 0.36634 | 0.016651
2019-03-26 16:27 | 1553631997.6757 | 818 | 0.20852 | 0.00121 | 7.1e-05 | 99.8281 | 99.6731 | 121.733 | 21.2612 | 0.002656 | 0.002532 | 219.884 | 0.24867 | 8.4008 | 0.003525 | 0.00539 | 0.005324 | 4.5e-05 | 0.001029 | 0.35131 | 1.3192 | 2.5475 | 100.8 | 21.06 | 15.275 | 85.7051 | 106.6962 | 82.7308 | 0.003787 | 0.99203 | 0.024155 | 24.3487 | 0.013899 | 0.14793 | 0.005014 | 0.36136 | 0.36136 | 0.016908
2019-03-26 16:27 | 1553631998.6756 | 819 | 0.20851 | 0.001212 | 7e-05 | 99.7936 | 99.7319 | 121.4734 | 21.2612 | 0.002645 | 0.002472 | 219.8636 | 0.24886 | 8.4017 | 0.00352 | 0.004516 | 0.004382 | 4.5e-05 | 0.001031 | 0.34971 | 1.3554 | 2.5476 | 100.8 | 21.06 | 15.275 | 85.4506 | 106.5756 | 81.2895 | 0.003787 | 0.99607 | 0.024178 | 24.2734 | 0.005526 | 0.14805 | 0.005088 | 0.36704 | 0.36704 | 0.016658
2019-03-26 16:27 | 1553631999.6767 | 820 | 0.20849 | 0.001215 | 7e-05 | 99.8243 | 99.6045 | 121.2657 | 21.2612 | 0.002676 | 0.002517 | 219.8646 | 0.24849 | 8.4025 | 0.003534 | 0.005288 | 0.005206 | 4.7e-05 | 0.001031 | 0.35347 | 1.3712 | 2.5476 | 100.8 | 21.06 | 15.275 | 86.1952 | 106.9049 | 82.3823 | 0.003787 | 0.99301 | 0.02412 | 24.29 | 0.019711 | 0.14783 | 0.005232 | 0.37671 | 0.37671 | 0.016207
2019-03-26 16:27 | 1553632000.6767 | 821 | 0.20852 | 0.001221 | 6.9e-05 | 99.8066 | 99.7827 | 120.8199 | 21.2612 | 0.002673 | 0.002505 | 219.8402 | 0.24819 | 8.4021 | 0.003521 | 0.005365 | 0.005263 | 4.6e-05 | 0.001028 | 0.32604 | 1.3608 | 2.5475 | 100.8 | 21.06 | 15.281 | 86.1148 | 106.5892 | 82.088 | 0.003789 | 0.99383 | 0.024086 | 24.2353 | 0.002139 | 0.14765 | 0.005037 | 0.36208 | 0.36208 | 0.016842
2019-03-26 16:27 | 1553632001.6772 | 822 | 0.2085 | 0.001223 | 6.9e-05 | 99.8449 | 99.8879 | 120.9446 | 21.2612 | 0.002638 | 0.002466 | 219.8269 | 0.2486 | 8.399 | 0.003476 | 0.00479 | 0.004671 | 4.6e-05 | 0.00104 | 0.33268 | 1.3972 | 2.5475 | 100.8 | 21.06 | 15.286 | 85.289 | 105.4984 | 81.1516 | 0.00379 | 0.99645 | 0.02413 | 24.2158 | -0.003849 | 0.14789 | 0.005117 | 0.36855 | 0.36855 | 0.016573
2019-03-26 16:27 | 1553632002.6758 | 823 | 0.20854 | 0.001224 | 6.8e-05 | 99.8839 | 99.9111 | 121.3047 | 21.2612 | 0.002706 | 0.002559 | 219.7945 | 0.24784 | 8.4019 | 0.003492 | 0.005522 | 0.005475 | 4.8e-05 | 0.001036 | 0.32012 | 1.3716 | 2.5475 | 100.8 | 21.05 | 15.294 | 86.9245 | 105.8974 | 83.3908 | 0.003789 | 0.9902 | 0.02409 | 24.3284 | -0.002436 | 0.14744 | 0.004902 | 0.3524 | 0.3524 | 0.017279
2019-03-26 16:27 | 1553632003.6769 | 824 | 0.20852 | 0.001228 | 6.7e-05 | 99.8184 | 99.9454 | 121.5166 | 21.2612 | 0.002658 | 0.002521 | 219.8256 | 0.24901 | 8.4047 | 0.003511 | 0.005161 | 0.005029 | 4.7e-05 | 0.001038 | 0.35922 | 1.3679 | 2.5476 | 100.8 | 21.05 | 15.314 | 85.7573 | 106.3505 | 82.4664 | 0.003794 | 0.99277 | 0.024142 | 24.318 | -0.011376 | 0.14813 | 0.005022 | 0.36189 | 0.36189 | 0.016905
2019-03-26 16:27 | 1553632004.676 | 825 | 0.20854 | 0.001229 | 6.5e-05 | 99.8715 | 99.6613 | 121.5044 | 21.2612 | 0.002727 | 0.002535 | 219.8268 | 0.24773 | 8.4041 | 0.003461 | 0.004645 | 0.004565 | 4.6e-05 | 0.001047 | 0.34202 | 1.3552 | 2.5476 | 100.8 | 21.05 | 15.325 | 87.4279 | 105.157 | 82.8087 | 0.003797 | 0.99182 | 0.024129 | 24.3285 | 0.018842 | 0.14737 | 0.004925 | 0.35477 | 0.35477 | 0.017156
2019-03-26 16:27 | 1553632005.677 | 826 | 0.20856 | 0.001235 | 6.5e-05 | 99.8588 | 99.7037 | 121.495 | 21.2612 | 0.002654 | 0.002523 | 219.8091 | 0.24892 | 8.4057 | 0.003449 | 0.005293 | 0.005182 | 4.9e-05 | 0.001054 | 0.32637 | 1.3498 | 2.5475 | 100.8 | 21.05 | 15.351 | 85.6621 | 104.862 | 82.5171 | 0.003804 | 0.99263 | 0.024138 | 24.3176 | 0.013905 | 0.14808 | 0.004777 | 0.34411 | 0.34411 | 0.017773
2019-03-26 16:27 | 1553632006.677 | 827 | 0.20855 | 0.001233 | 6.5e-05 | 99.8951 | 99.7823 | 121.4079 | 21.2612 | 0.002682 | 0.00248 | 219.809 | 0.24861 | 8.4014 | 0.003483 | 0.004901 | 0.004847 | 5e-05 | 0.001053 | 0.35347 | 1.3388 | 2.5475 | 100.8 | 21.05 | 15.365 | 86.3427 | 105.6706 | 81.483 | 0.003807 | 0.99552 | 0.024165 | 24.2735 | 0.010102 | 0.1479 | 0.004852 | 0.34997 | 0.34997 | 0.017454
2019-03-26 16:27 | 1553632007.6771 | 828 | 0.20853 | 0.00123 | 6.7e-05 | 99.8623 | 99.9575 | 121.3924 | 21.2612 | 0.002682 | 0.002567 | 219.8544 | 0.24797 | 8.4034 | 0.003467 | 0.004921 | 0.004849 | 4.5e-05 | 0.00105 | 0.33411 | 1.3617 | 2.5475 | 100.8 | 21.05 | 15.378 | 86.3436 | 105.2938 | 83.5608 | 0.00381 | 0.98972 | 0.024093 | 24.343 | -0.008518 | 0.14752 | 0.004968 | 0.35728 | 0.35728 | 0.017053
2019-03-26 16:27 | 1553632008.6771 | 829 | 0.20851 | 0.00123 | 6.7e-05 | 99.8922 | 99.7632 | 121.8778 | 21.2612 | 0.002648 | 0.002535 | 188.5029 | 0.2483 | 8.4033 | 0.00343 | 0.004345 | 0.004282 | 4.4e-05 | 0.001045 | 0.35563 | 1.3323 | 2.5476 | 100.8 | 21.04 | 15.406 | 85.5177 | 104.3918 | 82.7924 | 0.003815 | 0.99186 | 0.024167 | 24.3653 | 0.011556 | 0.14771 | 0.005064 | 0.36537 | 0.36537 | 0.016697
2019-03-26 16:27 | 1553632009.6771 | 830 | 0.20852 | 0.001231 | 6.7e-05 | 99.909 | 99.7457 | 122.452 | 21.2612 | 0.002665 | 0.00249 | 151.2258 | 0.24772 | 8.4031 | 0.003421 | 0.004825 | 0.004759 | 4.7e-05 | 0.00104 | 0.32565 | 1.3736 | 2.5476 | 100.8 | 21.05 | 15.421 | 85.919 | 104.1756 | 81.7176 | 0.003821 | 0.99486 | 0.024261 | 24.3857 | 0.01463 | 0.14737 | 0.004998 | 0.36196 | 0.36196 | 0.016815
2019-03-26 16:27 | 1553632010.6763 | 831 | 0.20848 | 0.001233 | 6.6e-05 | 99.8821 | 99.9363 | 123.2906 | 21.2612 | 0.002696 | 0.00249 | 113.9878 | 0.24886 | 8.4029 | 0.003399 | 0.004627 | 0.004567 | 4.6e-05 | 0.001042 | 0.35618 | 1.3724 | 2.5475 | 100.8 | 21.05 | 15.444 | 86.6785 | 103.6528 | 81.7145 | 0.003827 | 0.99487 | 0.024344 | 24.469 | -0.004852 | 0.14805 | 0.005286 | 0.38434 | 0.38434 | 0.015909
2019-03-26 16:27 | 1553632011.6763 | 832 | 0.20845 | 0.001239 | 6.5e-05 | 99.9379 | 99.8852 | 123.7883 | 21.2612 | 0.002657 | 0.002484 | 76.7595 | 0.24811 | 8.4049 | 0.003388 | 0.00491 | 0.004814 | 4.8e-05 | 0.001044 | 0.31725 | 1.383 | 2.5475 | 100.8 | 21.05 | 15.453 | 85.7347 | 103.3831 | 81.582 | 0.003829 | 0.99525 | 0.024397 | 24.5137 | 0.004718 | 0.1476 | 0.005421 | 0.39525 | 0.39525 | 0.015423
2019-03-26 16:27 | 1553632012.6764 | 833 | 0.20849 | 0.001247 | 6.3e-05 | 99.8322 | 99.8536 | 124.1054 | 21.2612 | 0.002648 | 0.00246 | 39.5153 | 0.24897 | 8.4019 | 0.003409 | 0.004848 | 0.004753 | 4.6e-05 | 0.001044 | 0.34379 | 1.345 | 2.5476 | 100.8 | 21.04 | 15.46 | 85.5233 | 103.8744 | 80.9956 | 0.003828 | 0.99689 | 0.024449 | 24.5248 | -0.001915 | 0.14811 | 0.005163 | 0.3772 | 0.3772 | 0.016217
2019-03-26 16:27 | 1553632013.6764 | 834 | 0.20859 | 0.001257 | 6e-05 | 99.8984 | 99.7975 | 124.2497 | 21.2612 | 0.002589 | 0.002482 | 2.2296 | 0.22394 | 8.4032 | 0.003393 | 0.005087 | 0.004906 | 4.7e-05 | 0.001045 | 0.32068 | 1.3809 | 2.5475 | 100.8 | 21.04 | 15.471 | 84.0968 | 103.5043 | 81.5195 | 0.003831 | 0.99542 | 0.024445 | 24.5572 | 0.009037 | 0.13322 | 0.004599 | 0.33569 | 0.33569 | 0.016391
2019-03-26 16:27 | 1553632014.6764 | 835 | 0.20867 | 0.00126 | 5.8e-05 | 99.8951 | 99.9535 | 124.6394 | 21.2612 | 0.002595 | 0.002444 | -35.0449 | 0.22021 | 8.403 | 0.003438 | 0.005371 | 0.005204 | 4.9e-05 | 0.001048 | 0.32405 | 1.3826 | 2.5475 | 100.8 | 21.04 | 15.475 | 84.2372 | 104.5828 | 80.6197 | 0.003832 | 0.99795 | 0.024514 | 24.5645 | -0.005223 | 0.131 | 0.004073 | 0.29786 | 0.29786 | 0.018164
2019-03-26 16:27 | 1553632015.6764 | 836 | 0.20877 | 0.001251 | 5.6e-05 | 99.9193 | 99.8082 | 124.752 | 21.2612 | 0.002601 | 0.002461 | -72.3086 | 0.21947 | 8.4037 | 0.003438 | 0.004711 | 0.00452 | 4.7e-05 | 0.001046 | 0.3194 | 1.376 | 2.5475 | 100.8 | 21.04 | 15.481 | 84.4017 | 104.5796 | 81.0143 | 0.003833 | 0.99684 | 0.024512 | 24.5893 | 0.009949 | 0.13056 | 0.003486 | 0.25454 | 0.25454 | 0.021184
2019-03-26 16:27 | 1553632016.6755 | 837 | 0.20889 | 0.001233 | 5.4e-05 | 99.9034 | 99.7592 | 125.0328 | 21.2612 | 0.002609 | 0.002403 | -109.5677 | 0.21959 | 8.4025 | 0.00345 | 0.004398 | 0.004152 | 4.9e-05 | 0.001042 | 0.33251 | 1.3801 | 2.5476 | 100.8 | 21.04 | 15.483 | 84.5782 | 104.8894 | 79.6397 | 0.003834 | 1.0007 | 0.024587 | 24.5691 | 0.012922 | 0.13063 | 0.002779 | 0.20294 | 0.20294 | 0.026585
2019-03-26 16:27 | 1553632017.6756 | 838 | 0.209 | 0.001205 | 5.2e-05 | 99.9462 | 99.9635 | 125.5947 | 21.2612 | 0.002608 | 0.002432 | -146.8545 | 0.21991 | 8.4026 | 0.003376 | 0.005055 | 0.004922 | 5.1e-05 | 0.001034 | 0.32853 | 1.3744 | 2.5476 | 100.8 | 21.04 | 15.484 | 84.5657 | 103.0743 | 80.3195 | 0.003834 | 0.9988 | 0.024618 | 24.648 | -0.001554 | 0.13082 | 0.002174 | 0.15833 | 0.15833 | 0.034125
2019-03-26 16:27 | 1553632018.6767 | 839 | 0.20907 | 0.001166 | 5.1e-05 | 100.0005 | 99.779 | 125.7752 | 21.2612 | 0.00256 | 0.002395 | -152.7759 | 0.21947 | 8.4049 | 0.003459 | 0.004145 | 0.004054 | 4.9e-05 | 0.001033 | 0.33196 | 1.3328 | 2.5474 | 100.8 | 21.04 | 15.486 | 83.3961 | 105.0979 | 79.4493 | 0.003835 | 1.0013 | 0.024666 | 24.6353 | 0.019834 | 0.13056 | 0.001829 | 0.13302 | 0.13302 | 0.040537
2019-03-26 16:27 | 1553632019.6756 | 840 | 0.20913 | 0.00112 | 5e-05 | 99.9825 | 99.9501 | 125.5555 | 21.2612 | 0.002665 | 0.002333 | -152.7707 | 0.21966 | 8.4024 | 0.003414 | 0.00452 | 0.004312 | 4.6e-05 | 0.001033 | 0.33284 | 1.3399 | 2.5475 | 100.8 | 21.04 | 15.488 | 85.9377 | 104.0167 | 77.9433 | 0.003835 | 1.0056 | 0.024698 | 24.5611 | 0.002898 | 0.13067 | 0.001498 | 0.10865 | 0.10865 | 0.049673
2019-03-26 16:27 | 1553632020.6758 | 841 | 0.20917 | 0.001064 | 4.7e-05 | 99.9243 | 99.856 | 125.5006 | 21.2612 | 0.002616 | 0.002395 | -152.768 | 0.21944 | 8.4038 | 0.003418 | 0.004699 | 0.004511 | 4.8e-05 | 0.001036 | 0.35187 | 1.3545 | 2.5476 | 100.8 | 21.04 | 15.488 | 84.7538 | 104.1007 | 79.4455 | 0.003835 | 1.0013 | 0.02464 | 24.6083 | 0.006121 | 0.13055 | 0.001312 | 0.09473 | 0.09473 | 0.056915
2019-03-26 16:27 | 1553632021.6758 | 842 | 0.20924 | 0.001012 | 4.4e-05 | 99.9465 | 100.1289 | 125.7733 | 21.2612 | 0.002509 | 0.002394 | -152.7789 | 0.21964 | 8.4022 | 0.003441 | 0.004542 | 0.004313 | 4.6e-05 | 0.001037 | 0.35906 | 1.3802 | 2.5476 | 100.8 | 21.03 | 15.489 | 82.1808 | 104.6669 | 79.415 | 0.003833 | 1.0014 | 0.024667 | 24.6339 | -0.016314 | 0.13067 | 0.001011 | 0.072551 | 0.072551 | 0.074383
2019-03-26 16:27 | 1553632022.6768 | 843 | 0.20929 | 0.000962 | 4e-05 | 99.8936 | 99.783 | 125.9766 | 21.2612 | 0.002533 | 0.00234 | -152.7769 | 0.22036 | 8.4015 | 0.003454 | 0.005109 | 0.004933 | 4.9e-05 | 0.001033 | 0.35043 | 1.3644 | 2.5476 | 100.8 | 21.04 | 15.491 | 82.7455 | 104.9822 | 78.1054 | 0.003836 | 1.0051 | 0.024733 | 24.608 | 0.00991 | 0.13109 | 0.000752 | 0.053678 | 0.053678 | 0.10086
2019-03-26 16:27 | 1553632023.6758 | 844 | 0.20929 | 0.000901 | 3.5e-05 | 99.8452 | 99.7632 | 126.2613 | 21.2612 | 0.002528 | 0.002317 | -152.7753 | 0.21964 | 8.4031 | 0.003473 | 0.00477 | 0.004613 | 5.2e-05 | 0.001033 | 0.34307 | 1.3684 | 2.5476 | 100.8 | 21.04 | 15.494 | 82.6419 | 105.4448 | 77.5619 | 0.003837 | 1.0067 | 0.02478 | 24.6167 | 0.007352 | 0.13066 | 0.000808 | 0.05822 | 0.05822 | 0.092689
2019-03-26 16:27 | 1553632024.6759 | 845 | 0.20937 | 0.000859 | 3.1e-05 | 99.9314 | 99.8788 | 126.4587 | 21.2612 | 0.00252 | 0.002373 | -152.7735 | 0.22067 | 8.4022 | 0.00346 | 0.004156 | 0.003942 | 4.9e-05 | 0.001032 | 0.34683 | 1.3508 | 2.5475 | 100.8 | 21.04 | 15.496 | 82.439 | 105.1279 | 78.9137 | 0.003837 | 1.0028 | 0.024752 | 24.6834 | 0.004708 | 0.13127 | 0.000419 | 0.029327 | 0.029327 | 0.18487
2019-03-26 16:27 | 1553632025.676 | 846 | 0.20936 | 0.000803 | 2.7e-05 | 99.9775 | 99.9124 | 127.0594 | 21.2612 | 0.002467 | 0.002333 | -152.7596 | 0.22025 | 8.4004 | 0.003427 | 0.004892 | 0.004705 | 4.6e-05 | 0.001033 | 0.3235 | 1.3577 | 2.5476 | 100.8 | 21.03 | 15.498 | 81.1648 | 104.314 | 77.9526 | 0.003835 | 1.0055 | 0.024845 | 24.7081 | 0.005819 | 0.13103 | 0.000552 | 0.039713 | 0.039713 | 0.13626
2019-03-26 16:27 | 1553632026.676 | 847 | 0.20942 | 0.00076 | 2.3e-05 | 99.9164 | 99.9578 | 127.2561 | 21.2612 | 0.002497 | 0.002291 | -152.7535 | 0.22033 | 8.4026 | 0.003447 | 0.005239 | 0.005087 | 4.6e-05 | 0.001028 | 0.32676 | 1.4021 | 2.5475 | 100.8 | 21.03 | 15.499 | 81.901 | 104.8046 | 76.9462 | 0.003835 | 1.0084 | 0.0249 | 24.6918 | -0.003711 | 0.13107 | 0.000244 | 0.01688 | 0.01688 | 0.32069
2019-03-26 16:27 | 1553632027.6771 | 848 | 0.20946 | 0.000724 | 1.9e-05 | 100.0011 | 99.8183 | 127.5826 | 21.2612 | 0.002486 | 0.002282 | -152.7703 | 0.22023 | 8.405 | 0.003367 | 0.004412 | 0.004178 | 4.4e-05 | 0.001023 | 0.32068 | 1.4027 | 2.5476 | 100.8 | 21.03 | 15.499 | 81.6318 | 102.8653 | 76.7218 | 0.003835 | 1.0091 | 0.02494 | 24.7155 | 0.016362 | 0.13101 | 5.7e-05 | 0.003099 | 0.003099 | 1.7461
2019-03-26 16:27 | 1553632028.6771 | 849 | 0.20947 | 0.000681 | 1.6e-05 | 99.9305 | 99.7639 | 127.7835 | 21.2612 | 0.002447 | 0.002314 | -152.7511 | 0.21946 | 8.4039 | 0.003391 | 0.005091 | 0.004828 | 4.8e-05 | 0.001018 | 0.32333 | 1.3607 | 2.5475 | 100.8 | 21.03 | 15.496 | 80.6766 | 103.4546 | 77.4925 | 0.003835 | 1.0069 | 0.024932 | 24.7622 | 0.014931 | 0.13055 | 3e-05 | 0.001269 | 0.001269 | 4.2502
2019-03-26 16:27 | 1553632029.6752 | 850 | 0.20947 | 0.000643 | 1.4e-05 | 99.9503 | 99.8997 | 127.8248 | 21.2612 | 0.002405 | 0.002259 | -152.7596 | 0.22174 | 8.4025 | 0.003452 | 0.005167 | 0.005079 | 4.9e-05 | 0.001017 | 0.32261 | 1.3604 | 2.5475 | 100.8 | 21.03 | 15.493 | 79.6848 | 104.9156 | 76.18 | 0.003834 | 1.0106 | 0.024983 | 24.7198 | 0.004533 | 0.13191 | 0.000122 | 0.008315 | 0.008315 | 0.65518
2019-03-26 16:27 | 1553632030.6752 | 851 | 0.20948 | 0.000621 | 1.1e-05 | 99.9214 | 99.9676 | 127.8912 | 21.2612 | 0.002507 | 0.002246 | -152.7631 | 0.21978 | 8.4007 | 0.003439 | 0.004515 | 0.00428 | 4.7e-05 | 0.001011 | 0.34235 | 1.3669 | 2.5475 | 100.8 | 21.03 | 15.49 | 82.1287 | 104.6126 | 75.8643 | 0.003833 | 1.0115 | 0.025001 | 24.7151 | -0.004134 | 0.13075 | 6.2e-05 | 0.003972 | 0.003972 | 1.3594
2019-03-26 16:27 | 1553632031.6763 | 852 | 0.20949 | 0.000594 | 9e-06 | 99.8942 | 99.8012 | 127.7682 | 21.2612 | 0.002403 | 0.002271 | -152.7667 | 0.2203 | 8.403 | 0.003388 | 0.005054 | 0.004869 | 4.6e-05 | 0.001013 | 0.33229 | 1.3446 | 2.5475 | 100.8 | 21.03 | 15.482 | 79.6404 | 103.3848 | 76.4481 | 0.003831 | 1.0099 | 0.024968 | 24.7238 | 0.008335 | 0.13106 | 6.6e-05 | 0.004372 | 0.004372 | 1.238
2019-03-26 16:27 | 1553632032.6753 | 853 | 0.20949 | 0.000572 | 7e-06 | 99.9568 | 99.8018 | 127.8349 | 21.2612 | 0.002407 | 0.00227 | -152.7711 | 0.22051 | 8.4037 | 0.003404 | 0.004762 | 0.004559 | 4.8e-05 | 0.001016 | 0.30829 | 1.3581 | 2.5475 | 100.8 | 21.03 | 15.477 | 79.7269 | 103.7627 | 76.427 | 0.00383 | 1.0099 | 0.024975 | 24.7295 | 0.013877 | 0.13118 | 6.5e-05 | 0.004442 | 0.004442 | 1.2198
2019-03-26 16:27 | 1553632033.6753 | 854 | 0.20949 | 0.000549 | 6e-06 | 99.9167 | 99.901 | 127.7652 | 21.2612 | 0.002402 | 0.002225 | -152.7562 | 0.22052 | 8.4007 | 0.003361 | 0.005117 | 0.004931 | 4.3e-05 | 0.001023 | 0.30574 | 1.3598 | 2.5474 | 100.8 | 21.03 | 15.469 | 79.614 | 102.7254 | 75.3646 | 0.003828 | 1.013 | 0.025006 | 24.6852 | 0.0014 | 0.13118 | 8.2e-05 | 0.005802 | 0.005802 | 0.93384
2019-03-26 16:27 | 1553632034.6764 | 855 | 0.20949 | 0.000532 | 6e-06 | 99.9302 | 99.8785 | 127.9142 | 21.2612 | 0.002434 | 0.002243 | -152.7655 | 0.21955 | 8.4047 | 0.003431 | 0.005088 | 0.004903 | 4.6e-05 | 0.001028 | 0.35081 | 1.3557 | 2.5475 | 100.8 | 21.03 | 15.465 | 80.3702 | 104.424 | 75.7841 | 0.003827 | 1.0118 | 0.025006 | 24.7144 | 0.004633 | 0.13061 | 0.000138 | 0.01007 | 0.01007 | 0.53565
2019-03-26 16:27 | 1553632035.6755 | 856 | 0.20952 | 0.000518 | 4e-06 | 99.9373 | 99.6136 | 127.9999 | 21.2612 | 0.002364 | 0.002242 | -152.7775 | 0.21998 | 8.4051 | 0.003399 | 0.004681 | 0.004489 | 4.6e-05 | 0.00103 | 0.31885 | 1.361 | 2.5475 | 100.8 | 21.03 | 15.462 | 78.6838 | 103.6314 | 75.7511 | 0.003826 | 1.0119 | 0.025015 | 24.7215 | 0.029021 | 0.13087 | -3.9e-05 | -0.00319 | -0.00319 | -1.6941
2019-03-26 16:27 | 1553632036.6765 | 857 | 0.20951 | 0.000509 | 4e-06 | 99.8405 | 99.6902 | 128.4008 | 21.2612 | 0.00239 | 0.002222 | -152.7815 | 0.22062 | 8.4031 | 0.003365 | 0.0047 | 0.004521 | 4.6e-05 | 0.001034 | 0.34667 | 1.3196 | 2.5476 | 100.8 | 21.03 | 15.452 | 79.3177 | 102.8237 | 75.2702 | 0.003824 | 1.0133 | 0.025072 | 24.7432 | 0.013472 | 0.13125 | 6.2e-05 | 0.004448 | 0.004448 | 1.2187
2019-03-26 16:27 | 1553632037.6766 | 858 | 0.20953 | 0.000498 | 3e-06 | 99.973 | 99.899 | 127.9358 | 21.2612 | 0.002379 | 0.002185 | -152.7661 | 0.21991 | 8.4024 | 0.003367 | 0.004679 | 0.004512 | 4.5e-05 | 0.001038 | 0.3235 | 1.3378 | 2.5475 | 100.8 | 21.03 | 15.449 | 79.0495 | 102.8692 | 74.3949 | 0.003823 | 1.0158 | 0.025058 | 24.6673 | 0.006627 | 0.13082 | -7e-05 | -0.005473 | -0.005473 | -0.98725
2019-03-26 16:27 | 1553632038.6767 | 859 | 0.20953 | 0.000488 | 2e-06 | 99.829 | 99.898 | 127.6964 | 21.2612 | 0.00238 | 0.00225 | -152.774 | 0.22046 | 8.4015 | 0.003377 | 0.004802 | 0.00464 | 4.7e-05 | 0.001042 | 0.33588 | 1.3341 | 2.5475 | 100.8 | 21.03 | 15.445 | 79.0711 | 103.116 | 75.9633 | 0.003822 | 1.0113 | 0.024978 | 24.6997 | -0.00618 | 0.13115 | -3e-05 | -0.002404 | -0.002404 | -2.2531
2019-03-26 16:27 | 1553632039.6757 | 860 | 0.20951 | 0.000481 | 2e-06 | 99.958 | 99.8462 | 127.8283 | 21.2612 | 0.002363 | 0.002211 | -152.7646 | 0.22038 | 8.4004 | 0.003389 | 0.004291 | 0.004094 | 4.6e-05 | 0.001038 | 0.3194 | 1.358 | 2.5474 | 100.8 | 21.03 | 15.438 | 78.6799 | 103.4028 | 75.0246 | 0.00382 | 1.014 | 0.025024 | 24.6792 | 0.010007 | 0.1311 | 8.6e-05 | 0.006347 | 0.006347 | 0.85315
2019-03-26 16:27 | 1553632040.6757 | 861 | 0.2095 | 0.000478 | 2e-06 | 99.9624 | 99.9165 | 127.994 | 21.2612 | 0.00235 | 0.002156 | -152.7595 | 0.22052 | 8.4033 | 0.003452 | 0.004534 | 0.004318 | 4.8e-05 | 0.001039 | 0.35723 | 1.3995 | 2.5475 | 100.8 | 21.03 | 15.434 | 78.3455 | 104.9272 | 73.7004 | 0.003819 | 1.0179 | 0.025088 | 24.6482 | 0.004111 | 0.13118 | 0.00012 | 0.008999 | 0.008999 | 0.60207
2019-03-26 16:27 | 1553632041.6757 | 862 | 0.20955 | 0.000473 | 1e-06 | 99.9884 | 99.8768 | 128.206 | 21.2612 | 0.002371 | 0.002202 | -152.7773 | 0.21994 | 8.4019 | 0.003388 | 0.00457 | 0.00443 | 5e-05 | 0.001043 | 0.32228 | 1.3543 | 2.5477 | 100.8 | 21.03 | 15.428 | 78.8623 | 103.3659 | 74.8002 | 0.003818 | 1.0146 | 0.025069 | 24.7077 | 0.009987 | 0.13084 | -0.000157 | -0.011952 | -0.011952 | -0.45214
2019-03-26 16:27 | 1553632042.6768 | 863 | 0.2095 | 0.000466 | 0 | 99.9323 | 99.8459 | 128.3431 | 21.2612 | 0.002389 | 0.002185 | -152.7709 | 0.22068 | 8.4011 | 0.003385 | 0.004938 | 0.004637 | 4.8e-05 | 0.001039 | 0.31852 | 1.3479 | 2.5474 | 100.8 | 21.03 | 15.424 | 79.3004 | 103.2928 | 74.3863 | 0.003817 | 1.0159 | 0.025098 | 24.7062 | 0.007738 | 0.13129 | 0.00018 | 0.013642 | 0.013642 | 0.39747
2019-03-26 16:27 | 1553632043.6749 | 864 | 0.20952 | 0.000463 | 1e-06 | 99.9801 | 100.0113 | 128.0654 | 21.2612 | 0.00233 | 0.002162 | -152.7656 | 0.22012 | 8.4028 | 0.003409 | 0.005222 | 0.005113 | 4.6e-05 | 0.001037 | 0.32278 | 1.4092 | 2.5477 | 100.8 | 21.03 | 15.42 | 77.8753 | 103.8807 | 73.8438 | 0.003816 | 1.0174 | 0.02509 | 24.6602 | -0.002787 | 0.13095 | 2.6e-05 | 0.001916 | 0.001916 | 2.8232
2019-03-26 16:27 | 1553632044.6749 | 865 | 0.20952 | 0.000457 | 0 | 99.8992 | 99.8116 | 128.0576 | 21.2612 | 0.002352 | 0.002215 | -152.7568 | 0.22059 | 8.4011 | 0.003386 | 0.004176 | 0.00402 | 4.7e-05 | 0.001035 | 0.3314 | 1.3971 | 2.5476 | 100.8 | 21.03 | 15.414 | 78.4139 | 103.3219 | 75.1152 | 0.003814 | 1.0137 | 0.025044 | 24.7046 | 0.007851 | 0.13123 | 3.4e-05 | 0.002545 | 0.002545 | 2.1293
2019-03-26 16:27 | 1553632045.6749 | 866 | 0.20954 | 0.000459 | 0 | 99.9412 | 99.8624 | 127.7011 | 21.2612 | 0.002304 | 0.002147 | -152.7563 | 0.2205 | 8.4009 | 0.003371 | 0.005076 | 0.004862 | 4.6e-05 | 0.001036 | 0.37 | 1.3933 | 2.5476 | 100.8 | 21.04 | 15.402 | 77.24 | 102.9658 | 73.4704 | 0.003814 | 1.0185 | 0.025068 | 24.6118 | 0.007055 | 0.13118 | -5.1e-05 | -0.003911 | -0.003911 | -1.3853
2019-03-26 16:27 | 1553632046.6761 | 867 | 0.20953 | 0.000457 | 0 | 99.9748 | 99.9007 | 127.0188 | 21.2612 | 0.002329 | 0.002148 | -152.7502 | 0.21977 | 8.403 | 0.003437 | 0.004501 | 0.00427 | 5e-05 | 0.001031 | 0.34147 | 1.3902 | 2.5477 | 100.8 | 21.04 | 15.395 | 77.8426 | 104.5656 | 73.4954 | 0.003812 | 1.0185 | 0.025 | 24.5469 | 0.006635 | 0.13074 | -2.9e-05 | -0.002143 | -0.002143 | -2.5194
2019-03-26 16:27 | 1553632047.675 | 868 | 0.20954 | 0.000454 | 0 | 99.9769 | 100.0846 | 127.0953 | 21.2612 | 0.002301 | 0.00217 | -152.77 | 0.22018 | 8.4021 | 0.003392 | 0.004955 | 0.004796 | 4.9e-05 | 0.001029 | 0.30127 | 1.3746 | 2.5477 | 100.8 | 21.04 | 15.387 | 77.1746 | 103.4696 | 74.0216 | 0.00381 | 1.0169 | 0.024989 | 24.5729 | -0.009629 | 0.13099 | -8.9e-05 | -0.006717 | -0.006717 | -0.80542
2019-03-26 16:27 | 1553632048.6751 | 869 | 0.20952 | 0.000454 | 0 | 99.9854 | 99.9884 | 127.119 | 21.2612 | 0.002311 | 0.002129 | -152.7649 | 0.21999 | 8.4015 | 0.00339 | 0.004665 | 0.00449 | 4.9e-05 | 0.00103 | 0.32532 | 1.3669 | 2.5475 | 100.8 | 21.04 | 15.374 | 77.4192 | 103.435 | 73.0495 | 0.003807 | 1.0198 | 0.025026 | 24.5408 | -0.000267 | 0.13087 | 4e-05 | 0.002986 | 0.002986 | 1.8102
2019-03-26 16:27 | 1553632049.6762 | 870 | 0.20958 | 0.000453 | 0 | 99.9887 | 99.8153 | 127.0685 | 21.2612 | 0.002301 | 0.002158 | -152.7846 | 0.21985 | 8.4043 | 0.003435 | 0.004483 | 0.004262 | 4.9e-05 | 0.001032 | 0.31725 | 1.3719 | 2.5475 | 100.8 | 21.04 | 15.368 | 77.1684 | 104.52 | 73.734 | 0.003805 | 1.0178 | 0.024996 | 24.5601 | 0.015524 | 0.13079 | -0.000336 | -0.025368 | -0.025368 | -0.21293
2019-03-26 16:27 | 1553632050.6753 | 871 | 0.20958 | 0.000452 | -1e-06 | 100.0256 | 99.981 | 126.8746 | 21.2612 | 0.002302 | 0.002078 | -152.7941 | 0.21952 | 8.4029 | 0.003246 | 0.004552 | 0.004341 | 5.2e-05 | 0.001031 | 0.335 | 1.3703 | 2.5475 | 100.8 | 21.04 | 15.361 | 77.2119 | 99.9322 | 71.8272 | 0.003804 | 1.0234 | 0.025046 | 24.4738 | 0.003986 | 0.13059 | -0.000315 | -0.023772 | -0.023772 | -0.22688
2019-03-26 16:28 | 1553632051.6763 | 872 | 0.20955 | 0.000451 | 0 | 100.0017 | 99.8254 | 126.8932 | 21.2612 | 0.002325 | 0.002139 | -152.7992 | 0.22049 | 8.4031 | 0.003327 | 0.005234 | 0.005021 | 5.1e-05 | 0.001031 | 0.34468 | 1.3324 | 2.5474 | 100.8 | 21.04 | 15.348 | 77.7568 | 101.911 | 73.2996 | 0.0038 | 1.019 | 0.024995 | 24.5278 | 0.015782 | 0.13117 | -0.000135 | -0.010193 | -0.010193 | -0.53148
2019-03-26 16:28 | 1553632052.6763 | 873 | 0.20955 | 0.000445 | 0 | 100.0548 | 99.9336 | 127.0272 | 21.2612 | 0.002266 | 0.002094 | -152.8201 | 0.22027 | 8.4025 | 0.003363 | 0.005085 | 0.004908 | 4.8e-05 | 0.00103 | 0.30973 | 1.347 | 2.5476 | 100.8 | 21.04 | 15.342 | 76.3412 | 102.7719 | 72.2083 | 0.003799 | 1.0223 | 0.025047 | 24.5021 | 0.01084 | 0.13104 | -8.7e-05 | -0.006538 | -0.006538 | -0.82769
2019-03-26 16:28 | 1553632053.6754 | 874 | 0.20955 | 0.000446 | 0 | 99.9905 | 99.9901 | 127.2091 | 21.2612 | 0.0023 | 0.002145 | -152.8378 | 0.2198 | 8.4023 | 0.003331 | 0.004651 | 0.00446 | 4.9e-05 | 0.001034 | 0.32731 | 1.3682 | 2.5476 | 100.8 | 21.04 | 15.331 | 77.1449 | 102.0044 | 73.4205 | 0.003796 | 1.0187 | 0.025022 | 24.5626 | 3.2e-05 | 0.13076 | -0.000126 | -0.00955 | -0.00955 | -0.56546
2019-03-26 16:28 | 1553632054.6754 | 875 | 0.20953 | 0.00045 | 0 | 100.0722 | 100.0788 | 126.8953 | 21.2612 | 0.002249 | 0.00207 | -152.8305 | 0.22067 | 8.4035 | 0.003262 | 0.004694 | 0.004489 | 5e-05 | 0.001039 | 0.33644 | 1.4004 | 2.5476 | 100.8 | 21.04 | 15.326 | 75.9414 | 100.3315 | 71.622 | 0.003795 | 1.024 | 0.025056 | 24.4686 | -0.000592 | 0.13128 | 1.6e-05 | 0.001244 | 0.001244 | 4.3598
2019-03-26 16:28 | 1553632055.6765 | 876 | 0.20954 | 0.000447 | 0 | 100.0058 | 99.9558 | 127.0664 | 21.2612 | 0.00226 | 0.002119 | -152.8162 | 0.22036 | 8.4038 | 0.003334 | 0.005133 | 0.004952 | 5.4e-05 | 0.001038 | 0.3314 | 1.4271 | 2.5476 | 100.8 | 21.04 | 15.323 | 76.2057 | 102.0664 | 72.8126 | 0.003794 | 1.0205 | 0.025029 | 24.5273 | 0.004473 | 0.13109 | -2.7e-05 | -0.002037 | -0.002037 | -2.6578
2019-03-26 16:28 | 1553632056.6755 | 877 | 0.20949 | 0.000447 | 0 | 100.1106 | 100.1037 | 127.337 | 21.2612 | 0.002249 | 0.002128 | -152.8205 | 0.22073 | 8.4023 | 0.003325 | 0.004557 | 0.004406 | 5.3e-05 | 0.001041 | 0.31133 | 1.3642 | 2.5475 | 100.8 | 21.04 | 15.32 | 75.9265 | 101.8473 | 73.0238 | 0.003794 | 1.0198 | 0.025048 | 24.5609 | 0.000614 | 0.13131 | 0.00024 | 0.018136 | 0.018136 | 0.29904
2019-03-26 16:28 | 1553632057.6766 | 878 | 0.20951 | 0.000449 | 0 | 100.0926 | 99.9619 | 127.4511 | 21.2612 | 0.002258 | 0.002124 | -152.7987 | 0.21955 | 8.4022 | 0.003342 | 0.004331 | 0.004086 | 5.3e-05 | 0.00104 | 0.32228 | 1.3235 | 2.5475 | 100.8 | 21.04 | 15.319 | 76.1564 | 102.2581 | 72.9178 | 0.003793 | 1.0202 | 0.025063 | 24.5681 | 0.01169 | 0.13061 | 0.000148 | 0.01119 | 0.01119 | 0.48204
2019-03-26 16:28 | 1553632058.6756 | 879 | 0.20952 | 0.000446 | 0 | 100.1233 | 100.0641 | 127.7211 | 21.2612 | 0.002231 | 0.002088 | -152.8232 | 0.22011 | 8.4026 | 0.003313 | 0.004553 | 0.004409 | 5.4e-05 | 0.001045 | 0.32947 | 1.3499 | 2.5476 | 100.8 | 21.04 | 15.322 | 75.4917 | 101.5576 | 72.0625 | 0.003794 | 1.0227 | 0.025121 | 24.5637 | 0.005293 | 0.13094 | 6.4e-05 | 0.004832 | 0.004832 | 1.1191
2019-03-26 16:28 | 1553632059.6756 | 880 | 0.20947 | 0.000446 | 0 | 100.1383 | 100.1602 | 127.8612 | 21.2612 | 0.002207 | 0.002085 | -152.812 | 0.22056 | 8.3993 | 0.00331 | 0.005064 | 0.004835 | 5.1e-05 | 0.001041 | 0.34108 | 1.3554 | 2.5476 | 100.8 | 21.04 | 15.323 | 74.9312 | 101.4768 | 72.0017 | 0.003794 | 1.0229 | 0.025137 | 24.575 | -0.001952 | 0.13121 | 0.000359 | 0.027245 | 0.027245 | 0.19889
2019-03-26 16:28 | 1553632060.6757 | 881 | 0.20956 | 0.000443 | -1e-06 | 100.1354 | 100.0516 | 127.8652 | 21.2612 | 0.002236 | 0.002052 | -152.809 | 0.22016 | 8.4041 | 0.003395 | 0.005303 | 0.005127 | 5.3e-05 | 0.00105 | 0.32101 | 1.3887 | 2.5476 | 100.8 | 21.04 | 15.324 | 75.6265 | 103.5554 | 71.2034 | 0.003795 | 1.0252 | 0.025167 | 24.547 | 0.007486 | 0.13098 | -0.000141 | -0.010636 | -0.010636 | -0.5086
2019-03-26 16:28 | 1553632061.6758 | 882 | 0.20955 | 0.000443 | 0 | 100.1404 | 100.2046 | 127.5391 | 21.2612 | 0.00227 | 0.002062 | -152.8104 | 0.21936 | 8.4013 | 0.003387 | 0.005164 | 0.004993 | 5.3e-05 | 0.001046 | 0.32045 | 1.3581 | 2.5475 | 100.8 | 21.04 | 15.322 | 76.4263 | 103.34 | 71.4404 | 0.003794 | 1.0245 | 0.025126 | 24.5241 | -0.005729 | 0.1305 | -9.8e-05 | -0.007452 | -0.007452 | -0.72323
2019-03-26 16:28 | 1553632062.6748 | 883 | 0.2095 | 0.00044 | -1e-06 | 100.1345 | 99.9619 | 127.2502 | 21.2612 | 0.0022 | 0.002087 | -152.8115 | 0.22007 | 8.4042 | 0.003373 | 0.00514 | 0.004932 | 5.8e-05 | 0.001045 | 0.3194 | 1.3788 | 2.5477 | 100.8 | 21.04 | 15.322 | 74.762 | 103.0145 | 72.0502 | 0.003794 | 1.0227 | 0.025075 | 24.518 | 0.015435 | 0.13092 | 0.000196 | 0.014875 | 0.014875 | 0.3635
2019-03-26 16:28 | 1553632063.6749 | 884 | 0.20952 | 0.000444 | 0 | 100.2216 | 100.2056 | 127.753 | 21.2612 | 0.002254 | 0.002072 | -152.8095 | 0.21935 | 8.4031 | 0.003339 | 0.004462 | 0.004275 | 5.6e-05 | 0.001041 | 0.32764 | 1.3562 | 2.5476 | 100.8 | 21.04 | 15.321 | 76.0536 | 102.1951 | 71.6689 | 0.003794 | 1.0239 | 0.025139 | 24.5528 | 0.001429 | 0.13049 | 6.1e-05 | 0.004644 | 0.004644 | 1.1606
2019-03-26 16:28 | 1553632064.6749 | 885 | 0.20951 | 0.000439 | 0 | 100.2263 | 100.2604 | 128.1536 | 21.2612 | 0.002232 | 0.002064 | -152.8049 | 0.21973 | 8.4006 | 0.003403 | 0.004573 | 0.004376 | 5.7e-05 | 0.001043 | 0.33732 | 1.356 | 2.5476 | 100.8 | 21.04 | 15.319 | 75.5135 | 103.7509 | 71.4973 | 0.003793 | 1.0244 | 0.025184 | 24.5851 | -0.003039 | 0.13072 | 0.000149 | 0.011368 | 0.011368 | 0.4749
2019-03-26 16:28 | 1553632065.676 | 886 | 0.20954 | 0.000447 | 0 | 100.2703 | 100.293 | 127.9257 | 21.2612 | 0.002168 | 0.002098 | -152.8186 | 0.21902 | 8.4028 | 0.003376 | 0.004397 | 0.004231 | 5.6e-05 | 0.001039 | 0.34252 | 1.3388 | 2.5476 | 100.8 | 21.04 | 15.318 | 73.9733 | 103.0956 | 72.2958 | 0.003793 | 1.022 | 0.025133 | 24.5917 | -0.002024 | 0.13029 | -4e-05 | -0.003075 | -0.003075 | -1.7499
2019-03-26 16:28 | 1553632066.675 | 887 | 0.20953 | 0.000443 | 0 | 100.3243 | 100.2328 | 127.9396 | 21.2612 | 0.002162 | 0.002056 | -152.8216 | 0.22085 | 8.4014 | 0.003343 | 0.004697 | 0.004466 | 5.2e-05 | 0.001047 | 0.33843 | 1.3234 | 2.5475 | 100.8 | 21.04 | 15.315 | 73.8299 | 102.2888 | 71.2962 | 0.003792 | 1.025 | 0.02517 | 24.5574 | 0.008162 | 0.13138 | 3.3e-05 | 0.002536 | 0.002536 | 2.1399
2019-03-26 16:28 | 1553632067.6764 | 888 | 0.20953 | 0.000445 | 0 | 100.2992 | 100.3266 | 128.0267 | 21.2612 | 0.00221 | 0.002088 | -152.8213 | 0.22007 | 8.4033 | 0.00338 | 0.00501 | 0.004898 | 5.6e-05 | 0.001049 | 0.32549 | 1.3741 | 2.5474 | 100.8 | 21.04 | 15.315 | 75.0006 | 103.1797 | 72.0617 | 0.003792 | 1.0227 | 0.025151 | 24.5931 | -0.002441 | 0.13092 | -1.5e-05 | -0.001141 | -0.001141 | -4.7406
2019-03-26 16:28 | 1553632068.6761 | 889 | 0.20951 | 0.000445 | -1e-06 | 100.3069 | 100.2506 | 128.0116 | 21.2612 | 0.002186 | 0.002061 | -152.7978 | 0.21952 | 8.4037 | 0.003407 | 0.005016 | 0.004872 | 5.8e-05 | 0.001057 | 0.34451 | 1.3959 | 2.5476 | 100.8 | 21.04 | 15.315 | 74.4082 | 103.8468 | 71.4053 | 0.003792 | 1.0246 | 0.025174 | 24.5682 | 0.005019 | 0.13059 | 0.000136 | 0.010346 | 0.010346 | 0.52129
2019-03-26 16:28 | 1553632069.6756 | 890 | 0.20954 | 0.000441 | 0 | 100.3113 | 100.371 | 127.9635 | 21.2612 | 0.002201 | 0.002043 | -152.7897 | 0.21995 | 8.4051 | 0.00338 | 0.004789 | 0.004621 | 5.4e-05 | 0.001053 | 0.33373 | 1.3839 | 2.5475 | 100.8 | 21.04 | 15.313 | 74.7878 | 103.1874 | 70.978 | 0.003792 | 1.0259 | 0.025184 | 24.5484 | -0.005317 | 0.13085 | -2e-05 | -0.001519 | -0.001519 | -3.5565
2019-03-26 16:28 | 1553632070.6766 | 891 | 0.20951 | 0.000442 | 0 | 100.34 | 100.4903 | 128.0011 | 21.2612 | 0.002131 | 0.002113 | -152.7907 | 0.22093 | 8.4026 | 0.003342 | 0.004619 | 0.004481 | 5.8e-05 | 0.001056 | 0.31885 | 1.3583 | 2.5475 | 100.8 | 21.04 | 15.311 | 73.0915 | 102.2675 | 72.6706 | 0.003791 | 1.0209 | 0.025126 | 24.6123 | -0.013393 | 0.13143 | 0.00013 | 0.009843 | 0.009843 | 0.5515
2019-03-26 16:28 | 1553632071.6753 | 892 | 0.20953 | 0.000445 | -1e-06 | 100.3081 | 100.2876 | 128.0706 | 21.2612 | 0.002169 | 0.002062 | -152.7761 | 0.22016 | 8.4011 | 0.003268 | 0.004412 | 0.004198 | 5.5e-05 | 0.001059 | 0.34739 | 1.3386 | 2.5475 | 100.8 | 21.04 | 15.306 | 74.0012 | 100.4807 | 71.4288 | 0.00379 | 1.0246 | 0.025179 | 24.5747 | 0.001826 | 0.13098 | -5e-06 | -0.00037 | -0.00037 | -14.6392
2019-03-26 16:28 | 1553632072.6758 | 893 | 0.20957 | 0.000445 | 0 | 100.2874 | 100.1572 | 127.8811 | 21.2612 | 0.002217 | 0.002029 | -152.7618 | 0.21987 | 8.408 | 0.003331 | 0.004436 | 0.00435 | 5.5e-05 | 0.001056 | 0.30773 | 1.3389 | 2.5474 | 100.8 | 21.04 | 15.303 | 75.1565 | 101.9841 | 70.6529 | 0.003789 | 1.0269 | 0.025188 | 24.5289 | 0.011624 | 0.1308 | -0.000254 | -0.019378 | -0.019378 | -0.27878
2019-03-26 16:28 | 1553632073.6764 | 894 | 0.20957 | 0.000445 | 0 | 100.3152 | 100.3495 | 127.7103 | 21.2612 | 0.002134 | 0.002072 | -152.7637 | 0.22083 | 8.4033 | 0.003245 | 0.0041 | 0.003799 | 5.6e-05 | 0.001059 | 0.32372 | 1.3359 | 2.5475 | 100.8 | 21.04 | 15.301 | 73.1571 | 99.9148 | 71.6814 | 0.003789 | 1.0238 | 0.025134 | 24.5491 | -0.003058 | 0.13137 | -0.000202 | -0.015382 | -0.015382 | -0.35273
2019-03-26 16:28 | 1553632074.6754 | 895 | 0.20957 | 0.000444 | 0 | 100.3907 | 100.3424 | 127.5262 | 21.2612 | 0.00216 | 0.002036 | -152.7902 | 0.22012 | 8.4035 | 0.003278 | 0.004605 | 0.004445 | 4.6e-05 | 0.001063 | 0.33212 | 1.3367 | 2.5475 | 100.8 | 21.03 | 15.294 | 73.7962 | 100.7161 | 70.8223 | 0.003785 | 1.0264 | 0.025147 | 24.5008 | 0.004307 | 0.13095 | -0.000228 | -0.017358 | -0.017358 | -0.31156
2019-03-26 16:28 | 1553632075.6755 | 896 | 0.20956 | 0.000443 | 0 | 100.3751 | 100.5216 | 127.9462 | 21.2612 | 0.002148 | 0.002059 | -152.7904 | 0.22087 | 8.399 | 0.003267 | 0.00423 | 0.004012 | 4.9e-05 | 0.001066 | 0.34595 | 1.3497 | 2.5474 | 100.8 | 21.03 | 15.291 | 73.5055 | 100.4515 | 71.3688 | 0.003784 | 1.0247 | 0.025168 | 24.5606 | -0.013044 | 0.13139 | -0.000149 | -0.011355 | -0.011355 | -0.4779
2019-03-26 16:28 | 1553632076.6744 | 897 | 0.20952 | 0.000446 | 0 | 100.3175 | 100.2772 | 128.0898 | 21.2612 | 0.002102 | 0.001965 | -152.8046 | 0.21988 | 8.4026 | 0.00319 | 0.004749 | 0.004553 | 5.4e-05 | 0.001068 | 0.31996 | 1.3256 | 2.5475 | 100.8 | 21.03 | 15.285 | 72.4087 | 98.5831 | 69.0996 | 0.003783 | 1.0315 | 0.025266 | 24.4934 | 0.003597 | 0.13081 | 0.000103 | 0.007883 | 0.007883 | 0.68531
2019-03-26 16:28 | 1553632077.6755 | 898 | 0.20952 | 0.000442 | -1e-06 | 100.3485 | 100.5061 | 127.9278 | 21.2612 | 0.002118 | 0.002052 | -152.8251 | 0.22085 | 8.4018 | 0.003283 | 0.004764 | 0.00458 | 5.2e-05 | 0.00107 | 0.3183 | 1.3424 | 2.5475 | 100.8 | 21.03 | 15.281 | 72.7751 | 100.8424 | 71.2073 | 0.003782 | 1.0252 | 0.025173 | 24.5531 | -0.014036 | 0.13138 | 9.5e-05 | 0.007254 | 0.007254 | 0.74798
2019-03-26 16:28 | 1553632078.675 | 899 | 0.20951 | 0.000439 | 0 | 100.3364 | 100.2829 | 127.5882 | 21.2612 | 0.002118 | 0.00201 | -152.8429 | 0.22011 | 8.4028 | 0.003338 | 0.004699 | 0.004529 | 5.1e-05 | 0.001068 | 0.32892 | 1.3745 | 2.5476 | 100.8 | 21.03 | 15.278 | 72.7806 | 102.1567 | 70.1797 | 0.003781 | 1.0283 | 0.025177 | 24.4839 | 0.004772 | 0.13095 | 0.000168 | 0.012787 | 0.012787 | 0.42294
2019-03-26 16:28 | 1553632079.6757 | 900 | 0.20953 | 0.000443 | 0 | 100.3057 | 100.4742 | 127.342 | 21.2612 | 0.002107 | 0.002 | -152.8376 | 0.22076 | 8.3998 | 0.003262 | 0.003994 | 0.003812 | 5.2e-05 | 0.001074 | 0.32068 | 1.3948 | 2.5476 | 100.8 | 21.03 | 15.271 | 72.5248 | 100.3221 | 69.9389 | 0.003779 | 1.029 | 0.025161 | 24.4517 | -0.015009 | 0.13133 | 8e-06 | 0.000617 | 0.000617 | 8.7935
2019-03-26 16:28 | 1553632080.6747 | 901 | 0.2095 | 0.000445 | 0 | 100.345 | 100.4453 | 127.2145 | 21.2612 | 0.00213 | 0.002 | -152.8455 | 0.22057 | 8.4017 | 0.003256 | 0.004954 | 0.004681 | 5.3e-05 | 0.001077 | 0.3366 | 1.3857 | 2.5475 | 100.8 | 21.03 | 15.269 | 73.0759 | 100.1745 | 69.9434 | 0.003779 | 1.029 | 0.025149 | 24.4396 | -0.008935 | 0.13122 | 0.000184 | 0.013953 | 0.013953 | 0.3884
2019-03-26 16:28 | 1553632081.6757 | 902 | 0.20949 | 0.000441 | 0 | 100.3535 | 100.3908 | 127.1683 | 21.2612 | 0.002141 | 0.002037 | -152.8497 | 0.22095 | 8.4027 | 0.003275 | 0.004961 | 0.004797 | 5.2e-05 | 0.001076 | 0.31382 | 1.3761 | 2.5474 | 100.8 | 21.03 | 15.267 | 73.3301 | 100.6384 | 70.8408 | 0.003778 | 1.0263 | 0.025111 | 24.4671 | -0.003321 | 0.13144 | 0.00025 | 0.018933 | 0.018933 | 0.28673
2019-03-26 16:28 | 1553632082.6758 | 903 | 0.20951 | 0.000447 | 0 | 100.2809 | 100.2136 | 127.3603 | 21.2612 | 0.002137 | 0.00205 | -152.8626 | 0.22039 | 8.4023 | 0.003306 | 0.00473 | 0.004547 | 5.1e-05 | 0.00108 | 0.3402 | 1.3652 | 2.5476 | 100.8 | 21.03 | 15.266 | 73.2295 | 101.3802 | 71.1426 | 0.003778 | 1.0254 | 0.025119 | 24.4963 | 0.006002 | 0.13111 | 0.000113 | 0.00856 | 0.00856 | 0.63257
2019-03-26 16:28 | 1553632083.6758 | 904 | 0.20952 | 0.000442 | 0 | 100.3228 | 100.2725 | 127.0432 | 21.2612 | 0.002213 | 0.002063 | -152.8696 | 0.22043 | 8.4033 | 0.003309 | 0.005219 | 0.005017 | 5.4e-05 | 0.001077 | 0.31349 | 1.3527 | 2.5475 | 100.8 | 21.03 | 15.266 | 75.0597 | 101.4548 | 71.456 | 0.003778 | 1.0245 | 0.025076 | 24.4769 | 0.004491 | 0.13113 | 6.2e-05 | 0.004722 | 0.004722 | 1.147
2019-03-26 16:28 | 1553632084.6749 | 905 | 0.20953 | 0.000441 | 0 | 100.283 | 100.2584 | 126.9084 | 21.2612 | 0.002137 | 0.001978 | -152.8567 | 0.22135 | 8.403 | 0.003292 | 0.004628 | 0.004468 | 5.7e-05 | 0.001078 | 0.31277 | 1.341 | 2.5475 | 100.8 | 21.03 | 15.269 | 73.2499 | 101.055 | 69.4209 | 0.003779 | 1.0306 | 0.025137 | 24.3916 | 0.002197 | 0.13168 | 2.2e-05 | 0.001673 | 0.001673 | 3.2502
2019-03-26 16:28 | 1553632085.675 | 906 | 0.20952 | 0.000444 | 0 | 100.262 | 100.2345 | 126.5033 | 21.2612 | 0.00213 | 0.002003 | -152.8516 | 0.21999 | 8.4021 | 0.003358 | 0.004873 | 0.004626 | 5.3e-05 | 0.001075 | 0.33964 | 1.4047 | 2.5476 | 100.8 | 21.03 | 15.273 | 73.069 | 102.6382 | 70.0206 | 0.00378 | 1.0288 | 0.025075 | 24.3739 | 0.002457 | 0.13087 | 0.000101 | 0.007655 | 0.007655 | 0.7061
2019-03-26 16:28 | 1553632086.675 | 907 | 0.20953 | 0.000448 | 0 | 100.2676 | 100.2896 | 126.539 | 21.2612 | 0.002125 | 0.002017 | -152.8183 | 0.22017 | 8.3988 | 0.003357 | 0.004878 | 0.004692 | 5.3e-05 | 0.001076 | 0.31924 | 1.3682 | 2.5474 | 100.8 | 21.03 | 15.28 | 72.9505 | 102.628 | 70.3489 | 0.003781 | 1.0278 | 0.025067 | 24.389 | -0.001961 | 0.13098 | -2e-06 | -0.000179 | -0.000179 | -30.2044
2019-03-26 16:28 | 1553632087.6751 | 908 | 0.20952 | 0.000447 | 0 | 100.1785 | 100.0429 | 126.4545 | 21.2612 | 0.002187 | 0.002034 | -152.8161 | 0.21992 | 8.4039 | 0.003395 | 0.004143 | 0.003957 | 5.1e-05 | 0.001073 | 0.32084 | 1.3735 | 2.5476 | 100.8 | 21.03 | 15.283 | 74.4519 | 103.5442 | 70.7605 | 0.003782 | 1.0266 | 0.025043 | 24.3955 | 0.012116 | 0.13083 | 5.9e-05 | 0.004474 | 0.004474 | 1.2076
2019-03-26 16:28 | 1553632088.6751 | 909 | 0.20952 | 0.000444 | 0 | 100.1451 | 100.1394 | 126.6493 | 21.2612 | 0.002118 | 0.001976 | -152.7835 | 0.22029 | 8.4038 | 0.003323 | 0.005217 | 0.005011 | 5e-05 | 0.00107 | 0.32925 | 1.4139 | 2.5476 | 100.8 | 21.03 | 15.287 | 72.7859 | 101.8081 | 69.3781 | 0.003783 | 1.0307 | 0.025113 | 24.3651 | 0.000516 | 0.13105 | 8.3e-05 | 0.006289 | 0.006289 | 0.86058
2019-03-26 16:28 | 1553632089.6761 | 910 | 0.20954 | 0.000443 | 0 | 100.1646 | 100.1246 | 126.5686 | 21.2612 | 0.002126 | 0.00203 | -152.7999 | 0.2208 | 8.4049 | 0.003357 | 0.00463 | 0.004528 | 5.6e-05 | 0.001068 | 0.33876 | 1.3448 | 2.5475 | 100.8 | 21.03 | 15.293 | 72.9693 | 102.6297 | 70.6726 | 0.003784 | 1.0268 | 0.025058 | 24.4034 | 0.003577 | 0.13135 | -5.9e-05 | -0.00448 | -0.00448 | -1.2109
2019-03-26 16:28 | 1553632090.6762 | 911 | 0.20955 | 0.000441 | 0 | 100.123 | 100.0372 | 126.42 | 21.2612 | 0.002069 | 0.001988 | -152.7955 | 0.221 | 8.4007 | 0.003386 | 0.004439 | 0.004308 | 5.2e-05 | 0.00107 | 0.33804 | 1.3671 | 2.5476 | 100.8 | 21.03 | 15.296 | 71.6151 | 103.3352 | 69.6565 | 0.003785 | 1.0299 | 0.02508 | 24.353 | 0.007671 | 0.13148 | -7.5e-05 | -0.005683 | -0.005683 | -0.95546
2019-03-26 16:28 | 1553632091.6742 | 912 | 0.20952 | 0.000447 | 0 | 100.0834 | 100.0973 | 126.5655 | 21.2612 | 0.002149 | 0.002029 | -152.7883 | 0.22074 | 8.4012 | 0.003357 | 0.004563 | 0.004334 | 5.3e-05 | 0.001064 | 0.28047 | 1.3766 | 2.5476 | 100.8 | 21.03 | 15.297 | 73.5375 | 102.6249 | 70.6553 | 0.003785 | 1.0269 | 0.025058 | 24.4025 | -0.001242 | 0.13132 | 4.3e-05 | 0.003238 | 0.003238 | 1.6749
2019-03-26 16:28 | 1553632092.6753 | 913 | 0.20953 | 0.00044 | 0 | 100.0114 | 100.2012 | 126.4315 | 21.2612 | 0.002065 | 0.002013 | -152.7924 | 0.22035 | 8.4048 | 0.003356 | 0.004703 | 0.004563 | 5.6e-05 | 0.001071 | 0.32587 | 1.3371 | 2.5476 | 100.8 | 21.03 | 15.302 | 71.5209 | 102.5983 | 70.2547 | 0.003787 | 1.0281 | 0.02506 | 24.3753 | -0.016958 | 0.13108 | 1.2e-05 | 0.00088 | 0.00088 | 6.15
2019-03-26 16:28 | 1553632093.6743 | 914 | 0.20954 | 0.000443 | 0 | 99.9949 | 99.9353 | 126.6686 | 21.2612 | 0.002139 | 0.002009 | -152.7989 | 0.21994 | 8.4015 | 0.003346 | 0.003915 | 0.003749 | 5.5e-05 | 0.001068 | 0.30829 | 1.3497 | 2.5475 | 100.8 | 21.03 | 15.304 | 73.2874 | 102.3656 | 70.1595 | 0.003787 | 1.0284 | 0.025087 | 24.3948 | 0.005331 | 0.13084 | -2.2e-05 | -0.001656 | -0.001656 | -3.2629
2019-03-26 16:28 | 1553632094.6754 | 915 | 0.20953 | 0.000443 | 0 | 100.0179 | 99.9531 | 126.858 | 21.2612 | 0.00211 | 0.00203 | -152.8078 | 0.22036 | 8.4023 | 0.003388 | 0.00464 | 0.004388 | 5.4e-05 | 0.001068 | 0.35314 | 1.366 | 2.5477 | 100.8 | 21.03 | 15.307 | 72.5918 | 103.3659 | 70.6695 | 0.003788 | 1.0268 | 0.025087 | 24.4311 | 0.005796 | 0.13109 | 3e-05 | 0.002307 | 0.002307 | 2.3468
2019-03-26 16:28 | 1553632095.6744 | 916 | 0.20952 | 0.000445 | 0 | 99.9317 | 100.0015 | 127.222 | 21.2612 | 0.002083 | 0.00206 | -152.784 | 0.22009 | 8.4001 | 0.003346 | 0.004343 | 0.004188 | 5.5e-05 | 0.001065 | 0.31924 | 1.4285 | 2.5476 | 100.8 | 21.03 | 15.308 | 71.9527 | 102.3469 | 71.3913 | 0.003788 | 1.0247 | 0.025096 | 24.4918 | -0.006247 | 0.13093 | 9.9e-05 | 0.007473 | 0.007473 | 0.72363
2019-03-26 16:28 | 1553632096.6755 | 917 | 0.20954 | 0.000439 | 0 | 99.8783 | 99.8106 | 127.1009 | 21.2612 | 0.00211 | 0.002009 | -152.8074 | 0.22048 | 8.4009 | 0.003352 | 0.00512 | 0.004804 | 5.6e-05 | 0.001064 | 0.35275 | 1.3948 | 2.5476 | 100.8 | 21.03 | 15.309 | 72.5865 | 102.5024 | 70.1696 | 0.003788 | 1.0283 | 0.025129 | 24.4367 | 0.006065 | 0.13116 | -4.2e-05 | -0.003179 | -0.003179 | -1.704
2019-03-26 16:28 | 1553632097.6755 | 918 | 0.20954 | 0.000442 | 0 | 99.8426 | 99.8546 | 127.3366 | 21.2612 | 0.002149 | 0.002028 | -152.8347 | 0.22034 | 8.4007 | 0.003356 | 0.004394 | 0.004251 | 5.4e-05 | 0.001059 | 0.30613 | 1.3888 | 2.5476 | 100.8 | 21.03 | 15.307 | 73.5267 | 102.5934 | 70.6101 | 0.003788 | 1.027 | 0.025136 | 24.4751 | -0.00108 | 0.13108 | -5e-05 | -0.003794 | -0.003794 | -1.427
2019-03-26 16:28 | 1553632098.6746 | 919 | 0.20955 | 0.000443 | 0 | 99.7968 | 99.6519 | 127.4342 | 21.2612 | 0.002084 | 0.001998 | -152.8567 | 0.22116 | 8.4061 | 0.003441 | 0.004812 | 0.004572 | 5.2e-05 | 0.001058 | 0.33699 | 1.3587 | 2.5474 | 100.8 | 21.03 | 15.306 | 71.9745 | 104.6654 | 69.9028 | 0.003788 | 1.0291 | 0.025172 | 24.4592 | 0.012996 | 0.13157 | -0.000109 | -0.00829 | -0.00829 | -0.65544
2019-03-26 16:28 | 1553632099.6756 | 920 | 0.20956 | 0.00045 | -1e-06 | 99.7514 | 99.8247 | 127.5532 | 21.2612 | 0.002079 | 0.001962 | -152.8517 | 0.22061 | 8.3988 | 0.003386 | 0.00492 | 0.004826 | 5.7e-05 | 0.00105 | 0.31293 | 1.3541 | 2.5475 | 100.8 | 21.03 | 15.305 | 71.8489 | 103.3173 | 69.0279 | 0.003787 | 1.0318 | 0.025216 | 24.4394 | -0.006575 | 0.13124 | -0.000141 | -0.01072 | -0.01072 | -0.50563
2019-03-26 16:28 | 1553632100.6747 | 921 | 0.20953 | 0.000446 | -1e-06 | 99.6793 | 99.3897 | 127.7063 | 21.2612 | 0.002147 | 0.001982 | -152.8557 | 0.22024 | 8.402 | 0.003369 | 0.004836 | 0.004623 | 5.4e-05 | 0.001051 | 0.3314 | 1.3891 | 2.5475 | 100.8 | 21.03 | 15.306 | 73.4892 | 102.9109 | 69.5154 | 0.003788 | 1.0303 | 0.025213 | 24.4715 | 0.026026 | 0.13102 | 2.1e-05 | 0.001642 | 0.001642 | 3.296
2019-03-26 16:28 | 1553632101.6757 | 922 | 0.20949 | 0.000445 | 0 | 99.5819 | 99.303 | 127.4483 | 21.2612 | 0.002111 | 0.001965 | -152.8651 | 0.21944 | 8.4026 | 0.003412 | 0.004291 | 0.004077 | 5.7e-05 | 0.001051 | 0.33323 | 1.3769 | 2.5476 | 100.8 | 21.03 | 15.308 | 72.6074 | 103.9641 | 69.1044 | 0.003788 | 1.0315 | 0.025202 | 24.4321 | 0.025089 | 0.13054 | 0.000228 | 0.017367 | 0.017367 | 0.31045
2019-03-26 16:28 | 1553632102.6757 | 923 | 0.20952 | 0.000445 | 0 | 99.4875 | 99.4038 | 127.1673 | 21.2612 | 0.002088 | 0.001988 | -152.8369 | 0.22097 | 8.3992 | 0.003428 | 0.004208 | 0.004039 | 5.6e-05 | 0.001049 | 0.30558 | 1.379 | 2.5475 | 100.8 | 21.03 | 15.313 | 72.061 | 104.3431 | 69.6673 | 0.003789 | 1.0298 | 0.025154 | 24.4252 | 0.007522 | 0.13146 | 6.1e-05 | 0.004616 | 0.004616 | 1.176
2019-03-26 16:28 | 1553632103.6748 | 924 | 0.20956 | 0.000444 | 0 | 99.3995 | 99.3477 | 127.0058 | 21.2612 | 0.002146 | 0.001985 | -152.8294 | 0.2214 | 8.4057 | 0.003379 | 0.004279 | 0.004133 | 5.8e-05 | 0.001052 | 0.34274 | 1.3558 | 2.5475 | 100.8 | 21.03 | 15.315 | 73.4596 | 103.1695 | 69.5894 | 0.00379 | 1.0301 | 0.025141 | 24.4069 | 0.004663 | 0.13171 | -0.000136 | -0.010348 | -0.010348 | -0.52569
2019-03-26 16:28 | 1553632104.6753 | 925 | 0.20953 | 0.000442 | 0 | 99.4166 | 99.2556 | 126.7129 | 21.2612 | 0.002017 | 0.001958 | -152.8314 | 0.21971 | 8.4028 | 0.003415 | 0.005026 | 0.004847 | 5.5e-05 | 0.00105 | 0.32405 | 1.3767 | 2.5475 | 100.8 | 21.03 | 15.318 | 70.3443 | 104.0411 | 68.9282 | 0.003791 | 1.0321 | 0.025136 | 24.3553 | 0.0145 | 0.1307 | 1.6e-05 | 0.001246 | 0.001246 | 4.3328
2019-03-26 16:28 | 1553632105.6759 | 926 | 0.20955 | 0.00045 | 0 | 99.2817 | 99.4371 | 126.4559 | 21.2612 | 0.002038 | 0.001945 | -152.8612 | 0.2205 | 8.3991 | 0.003376 | 0.003989 | 0.003757 | 5.6e-05 | 0.00105 | 0.29695 | 1.3662 | 2.5475 | 100.8 | 21.03 | 15.321 | 70.8665 | 103.0915 | 68.6169 | 0.003791 | 1.033 | 0.025122 | 24.3195 | -0.013989 | 0.13117 | -9.4e-05 | -0.007077 | -0.007077 | -0.7655
2019-03-26 16:28 | 1553632106.675 | 927 | 0.20952 | 0.000444 | 0 | 99.3027 | 99.2307 | 126.7455 | 21.2612 | 0.002084 | 0.001976 | -152.843 | 0.21996 | 8.4018 | 0.00345 | 0.00446 | 0.004272 | 5.8e-05 | 0.001046 | 0.33516 | 1.3886 | 2.5475 | 100.8 | 21.03 | 15.323 | 71.9721 | 104.8739 | 69.3649 | 0.003792 | 1.0307 | 0.025123 | 24.3739 | 0.006485 | 0.13085 | 6.2e-05 | 0.004708 | 0.004708 | 1.1478
2019-03-26 16:28 | 1553632107.6751 | 928 | 0.2095 | 0.000446 | 0 | 99.156 | 98.8145 | 126.6977 | 21.2612 | 0.002101 | 0.001972 | -152.8108 | 0.22001 | 8.406 | 0.003346 | 0.004603 | 0.004415 | 5.7e-05 | 0.001046 | 0.31741 | 1.4374 | 2.5476 | 100.8 | 21.03 | 15.327 | 72.3729 | 102.368 | 69.2781 | 0.003793 | 1.031 | 0.025122 | 24.3662 | 0.030852 | 0.13089 | 0.000183 | 0.013881 | 0.013881 | 0.38943
2019-03-26 16:28 | 1553632108.6751 | 929 | 0.20956 | 0.000446 | -1e-06 | 99.1205 | 99.0667 | 126.6174 | 21.2612 | 0.002083 | 0.00199 | -152.8248 | 0.22141 | 8.403 | 0.003358 | 0.004767 | 0.004656 | 5.8e-05 | 0.001051 | 0.32853 | 1.4235 | 2.5475 | 100.8 | 21.03 | 15.33 | 71.9527 | 102.6564 | 69.6985 | 0.003794 | 1.0297 | 0.025098 | 24.3735 | 0.004864 | 0.13172 | -0.000136 | -0.010303 | -0.010303 | -0.528
2019-03-26 16:28 | 1553632109.6751 | 930 | 0.20952 | 0.000442 | 0 | 99.0884 | 98.9362 | 126.7523 | 21.2612 | 0.002079 | 0.001978 | -152.8243 | 0.22059 | 8.4032 | 0.003415 | 0.004809 | 0.004708 | 5.8e-05 | 0.001052 | 0.32947 | 1.4067 | 2.5475 | 100.8 | 21.03 | 15.333 | 71.8404 | 104.0387 | 69.4195 | 0.003794 | 1.0306 | 0.025122 | 24.3765 | 0.013744 | 0.13123 | 6.9e-05 | 0.005293 | 0.005293 | 1.0239
2019-03-26 16:28 | 1553632110.6756 | 931 | 0.20957 | 0.000443 | 0 | 98.9951 | 98.6068 | 126.865 | 21.2612 | 0.002083 | 0.001981 | -152.8118 | 0.22105 | 8.4024 | 0.003463 | 0.004536 | 0.004381 | 5.7e-05 | 0.001052 | 0.31996 | 1.403 | 2.5474 | 100.8 | 21.03 | 15.34 | 71.944 | 105.2009 | 69.4796 | 0.003796 | 1.0304 | 0.025131 | 24.3895 | 0.035153 | 0.1315 | -0.000217 | -0.01647 | -0.01647 | -0.32976
2019-03-26 16:28 | 1553632111.6743 | 932 | 0.20952 | 0.000449 | 0 | 98.8319 | 98.7981 | 127.2469 | 21.2612 | 0.002035 | 0.001976 | -152.7881 | 0.22066 | 8.4034 | 0.003427 | 0.004887 | 0.004683 | 6.1e-05 | 0.001055 | 0.33395 | 1.4005 | 2.5476 | 100.8 | 21.03 | 15.342 | 70.7824 | 104.3273 | 69.3664 | 0.003797 | 1.0307 | 0.025173 | 24.4221 | 0.003058 | 0.13127 | 4.6e-05 | 0.003521 | 0.003521 | 1.5399
2019-03-26 16:29 | 1553632112.6757 | 933 | 0.20954 | 0.000453 | 0 | 98.8301 | 98.5597 | 127.8701 | 21.2612 | 0.002069 | 0.001918 | -152.7981 | 0.22024 | 8.4001 | 0.003365 | 0.00488 | 0.004677 | 5.6e-05 | 0.001064 | 0.3402 | 1.3905 | 2.5476 | 100.8 | 21.03 | 15.345 | 71.5949 | 102.8293 | 67.9858 | 0.003797 | 1.0349 | 0.025285 | 24.4325 | 0.024502 | 0.13102 | -6.6e-05 | -0.005016 | -0.005016 | -1.0787
2019-03-26 16:29 | 1553632113.6744 | 934 | 0.20956 | 0.000447 | 0 | 98.7581 | 98.551 | 128.232 | 21.2612 | 0.002037 | 0.001953 | -152.783 | 0.22071 | 8.4031 | 0.003406 | 0.005429 | 0.005252 | 5.7e-05 | 0.00106 | 0.31332 | 1.3562 | 2.5475 | 100.8 | 21.04 | 15.346 | 70.8276 | 103.8187 | 68.8236 | 0.0038 | 1.0324 | 0.02529 | 24.4971 | 0.018774 | 0.1313 | -0.000154 | -0.011748 | -0.011748 | -0.46158
2019-03-26 16:29 | 1553632114.6754 | 935 | 0.20954 | 0.000452 | -1e-06 | 98.6279 | 98.4572 | 128.7217 | 21.2612 | 0.002043 | 0.001975 | -152.7568 | 0.22091 | 8.4042 | 0.00336 | 0.004784 | 0.004625 | 5.9e-05 | 0.001058 | 0.3298 | 1.3401 | 2.5475 | 100.8 | 21.04 | 15.346 | 70.9735 | 102.7099 | 69.3414 | 0.0038 | 1.0308 | 0.025319 | 24.5624 | 0.015492 | 0.13142 | -2.1e-05 | -0.001538 | -0.001538 | -3.5291
2019-03-26 16:29 | 1553632115.6754 | 936 | 0.20956 | 0.000447 | -1e-06 | 98.64 | 98.637 | 128.6332 | 21.2612 | 0.00201 | 0.001946 | -152.7532 | 0.22106 | 8.4015 | 0.003377 | 0.005496 | 0.005259 | 5.2e-05 | 0.001053 | 0.32803 | 1.3565 | 2.5474 | 100.8 | 21.04 | 15.348 | 70.1821 | 103.0995 | 68.645 | 0.0038 | 1.0329 | 0.025336 | 24.5289 | 0.000266 | 0.13151 | -0.00018 | -0.013768 | -0.013768 | -0.39448
2019-03-26 16:29 | 1553632116.6749 | 937 | 0.20952 | 0.00045 | -1e-06 | 98.4629 | 98.2952 | 128.1497 | 21.2612 | 0.002019 | 0.001981 | -152.7634 | 0.22022 | 8.4011 | 0.003396 | 0.004608 | 0.004414 | 5.2e-05 | 0.001045 | 0.33196 | 1.3443 | 2.5474 | 100.8 | 21.04 | 15.35 | 70.3996 | 103.574 | 69.4911 | 0.003801 | 1.0304 | 0.025257 | 24.5131 | 0.015247 | 0.13101 | 0.000103 | 0.007921 | 0.007921 | 0.68305
2019-03-26 16:29 | 1553632117.6755 | 938 | 0.20957 | 0.000453 | -1e-06 | 98.467 | 98.3873 | 128.0745 | 21.2612 | 0.002039 | 0.001936 | -152.7668 | 0.22107 | 8.3997 | 0.003392 | 0.004575 | 0.004312 | 5.7e-05 | 0.001042 | 0.32211 | 1.3911 | 2.5475 | 100.8 | 21.04 | 15.352 | 70.88 | 103.4805 | 68.3976 | 0.003801 | 1.0337 | 0.02529 | 24.4668 | 0.007245 | 0.13151 | -0.000253 | -0.019322 | -0.019322 | -0.28111
2019-03-26 16:29 | 1553632118.6756 | 939 | 0.20952 | 0.000452 | 0 | 98.3876 | 98.1422 | 128.0077 | 21.2612 | 0.002043 | 0.001939 | -152.7551 | 0.22084 | 8.4019 | 0.003379 | 0.004537 | 0.004403 | 5.4e-05 | 0.001046 | 0.33876 | 1.3861 | 2.5475 | 100.8 | 21.04 | 15.358 | 70.9915 | 103.1671 | 68.4765 | 0.003803 | 1.0334 | 0.025281 | 24.4632 | 0.022336 | 0.13138 | 5.4e-05 | 0.00413 | 0.00413 | 1.3138
2019-03-26 16:29 | 1553632119.6747 | 940 | 0.20953 | 0.000448 | -1e-06 | 98.2695 | 97.954 | 128.0062 | 21.2612 | 0.002049 | 0.001907 | -152.7521 | 0.22062 | 8.4032 | 0.003374 | 0.005006 | 0.004722 | 5.2e-05 | 0.001043 | 0.33483 | 1.4021 | 2.5475 | 100.8 | 21.04 | 15.364 | 71.1349 | 103.0366 | 67.7019 | 0.003804 | 1.0358 | 0.025309 | 24.4353 | 0.028768 | 0.13125 | -7e-06 | -0.000471 | -0.000471 | -11.5096
2019-03-26 16:29 | 1553632120.6761 | 941 | 0.20952 | 0.000449 | 0 | 98.2934 | 98.001 | 127.9473 | 21.2612 | 0.002051 | 0.001936 | -152.7549 | 0.21949 | 8.4026 | 0.003432 | 0.00469 | 0.004494 | 5.7e-05 | 0.001041 | 0.30359 | 1.3869 | 2.5476 | 100.8 | 21.04 | 15.375 | 71.1676 | 104.442 | 68.4087 | 0.003807 | 1.0336 | 0.025277 | 24.455 | 0.026648 | 0.13058 | 5.5e-05 | 0.004251 | 0.004251 | 1.2686
2019-03-26 16:29 | 1553632121.6757 | 942 | 0.20952 | 0.000446 | -1e-06 | 98.1582 | 97.9482 | 127.561 | 21.2612 | 0.00202 | 0.002001 | -152.7737 | 0.22022 | 8.4033 | 0.003426 | 0.004778 | 0.00459 | 5.7e-05 | 0.001044 | 0.33157 | 1.41 | 2.5475 | 100.8 | 21.04 | 15.382 | 70.4371 | 104.2941 | 69.9778 | 0.003809 | 1.0289 | 0.025181 | 24.4741 | 0.019155 | 0.13101 | 0.00011 | 0.008447 | 0.008447 | 0.6405
2019-03-26 16:29 | 1553632122.6752 | 943 | 0.20954 | 0.000444 | -1e-06 | 98.0366 | 97.7402 | 127.1896 | 21.2612 | 0.002029 | 0.001936 | -152.7707 | 0.2204 | 8.4048 | 0.003388 | 0.004595 | 0.004443 | 5.7e-05 | 0.001049 | 0.30237 | 1.4076 | 2.5475 | 100.8 | 21.04 | 15.388 | 70.6507 | 103.3691 | 68.4008 | 0.00381 | 1.0337 | 0.025203 | 24.3822 | 0.027089 | 0.13112 | -1.1e-05 | -0.000826 | -0.000826 | -6.553
2019-03-26 16:29 | 1553632123.6753 | 944 | 0.20952 | 0.000448 | 0 | 97.9333 | 97.5499 | 126.834 | 21.2612 | 0.002073 | 0.001957 | -152.7983 | 0.22095 | 8.4028 | 0.003431 | 0.004989 | 0.004695 | 5.9e-05 | 0.001052 | 0.31686 | 1.3887 | 2.5476 | 100.8 | 21.04 | 15.403 | 71.7054 | 104.424 | 68.9202 | 0.003814 | 1.0321 | 0.025148 | 24.3666 | 0.035087 | 0.13144 | 0.000104 | 0.007885 | 0.007885 | 0.68846
2019-03-26 16:29 | 1553632124.6739 | 945 | 0.20953 | 0.000448 | -1e-06 | 97.8409 | 97.6269 | 126.6639 | 21.2612 | 0.002059 | 0.001981 | -152.8021 | 0.22074 | 8.4028 | 0.003421 | 0.004374 | 0.004249 | 5.9e-05 | 0.001055 | 0.33787 | 1.3847 | 2.5476 | 100.8 | 21.04 | 15.409 | 71.3618 | 104.1715 | 69.4796 | 0.003816 | 1.0304 | 0.025111 | 24.3702 | 0.019588 | 0.13132 | -1.1e-05 | -0.000799 | -0.000799 | -6.7866
2019-03-26 16:29 | 1553632125.6751 | 946 | 0.20952 | 0.000449 | 0 | 97.7751 | 97.4538 | 126.6678 | 21.2612 | 0.002005 | 0.001955 | -152.7968 | 0.22009 | 8.4043 | 0.003414 | 0.004239 | 0.003993 | 5.4e-05 | 0.001056 | 0.33157 | 1.4368 | 2.5476 | 100.8 | 21.04 | 15.422 | 70.0564 | 104.0096 | 68.8601 | 0.003819 | 1.0323 | 0.025134 | 24.3485 | 0.029444 | 0.13093 | 8.4e-05 | 0.006387 | 0.006387 | 0.84666
2019-03-26 16:29 | 1553632126.6744 | 947 | 0.20954 | 0.000441 | 0 | 97.6213 | 97.4064 | 127.035 | 21.2612 | 0.002114 | 0.001952 | -152.7951 | 0.22 | 8.4001 | 0.003454 | 0.004497 | 0.004388 | 5.7e-05 | 0.001056 | 0.32012 | 1.3636 | 2.5476 | 100.8 | 21.04 | 15.428 | 72.697 | 104.9832 | 68.8056 | 0.00382 | 1.0324 | 0.025173 | 24.3818 | 0.019716 | 0.13088 | -2.9e-05 | -0.002187 | -0.002187 | -2.472
2019-03-26 16:29 | 1553632127.674 | 948 | 0.20953 | 0.000442 | -1e-06 | 97.5658 | 96.9979 | 127.2521 | 21.2612 | 0.002083 | 0.001937 | -152.7996 | 0.21914 | 8.4043 | 0.003407 | 0.004263 | 0.004079 | 5.5e-05 | 0.001059 | 0.35817 | 1.3593 | 2.5473 | 100.8 | 21.04 | 15.432 | 71.9325 | 103.8429 | 68.4243 | 0.003821 | 1.0336 | 0.025208 | 24.389 | 0.052214 | 0.13037 | 4e-05 | 0.003094 | 0.003094 | 1.7403
2019-03-26 16:29 | 1553632128.6741 | 949 | 0.20956 | 0.000445 | -1e-06 | 97.4114 | 96.9919 | 127.211 | 21.2612 | 0.002009 | 0.001927 | -152.8096 | 0.22012 | 8.4019 | 0.003395 | 0.004777 | 0.004623 | 6.2e-05 | 0.001062 | 0.33754 | 1.406 | 2.5475 | 100.8 | 21.04 | 15.438 | 70.1679 | 103.5554 | 68.2027 | 0.003823 | 1.0343 | 0.025212 | 24.3772 | 0.038607 | 0.13095 | -0.000146 | -0.011068 | -0.011068 | -0.48863
2019-03-26 16:29 | 1553632129.6755 | 950 | 0.20952 | 0.000442 | 0 | 97.3376 | 96.8251 | 127.0875 | 21.2612 | 0.002006 | 0.00193 | -152.8043 | 0.22054 | 8.4056 | 0.003391 | 0.004595 | 0.00439 | 5.8e-05 | 0.001057 | 0.33588 | 1.3716 | 2.5476 | 100.8 | 21.04 | 15.44 | 70.097 | 103.4493 | 68.2596 | 0.003823 | 1.0341 | 0.025198 | 24.3674 | 0.047218 | 0.1312 | 0.0001 | 0.007621 | 0.007621 | 0.711
2019-03-26 16:29 | 1553632130.6752 | 951 | 0.20955 | 0.000439 | 0 | 97.1298 | 96.6224 | 126.9476 | 21.2612 | 0.002036 | 0.001961 | -152.8151 | 0.22035 | 8.4038 | 0.00333 | 0.004446 | 0.004261 | 5.9e-05 | 0.001054 | 0.32012 | 1.3796 | 2.5476 | 100.8 | 21.05 | 15.443 | 70.8237 | 101.9676 | 69.0099 | 0.003826 | 1.0318 | 0.025156 | 24.3807 | 0.046847 | 0.13108 | -9.9e-05 | -0.007487 | -0.007487 | -0.7231
2019-03-26 16:29 | 1553632131.6745 | 952 | 0.20953 | 0.000442 | -1e-06 | 97.0227 | 96.4281 | 127.0148 | 21.2612 | 0.002055 | 0.001941 | -152.8185 | 0.21987 | 8.4019 | 0.003416 | 0.004726 | 0.004566 | 5.6e-05 | 0.001053 | 0.34987 | 1.3782 | 2.5476 | 100.8 | 21.05 | 15.443 | 71.2697 | 104.06 | 68.5327 | 0.003826 | 1.0333 | 0.025181 | 24.3701 | 0.054979 | 0.1308 | -4e-06 | -0.00028 | -0.00028 | -19.2883
2019-03-26 16:29 | 1553632132.6756 | 953 | 0.20954 | 0.000442 | 0 | 96.8653 | 96.3182 | 127.2753 | 21.2612 | 0.002022 | 0.001968 | -152.8207 | 0.22031 | 8.4051 | 0.003406 | 0.004459 | 0.004254 | 6.2e-05 | 0.001054 | 0.33445 | 1.3885 | 2.5476 | 100.8 | 21.05 | 15.443 | 70.4798 | 103.8083 | 69.1791 | 0.003826 | 1.0313 | 0.025183 | 24.4182 | 0.050665 | 0.13106 | -1.3e-05 | -0.000947 | -0.000947 | -5.717
2019-03-26 16:29 | 1553632133.6748 | 954 | 0.20955 | 0.000441 | 0 | 96.7021 | 96.3044 | 127.3964 | 21.2612 | 0.00203 | 0.00193 | -152.8048 | 0.22159 | 8.4018 | 0.003336 | 0.004623 | 0.004479 | 5.7e-05 | 0.001052 | 0.3402 | 1.401 | 2.5475 | 100.8 | 21.05 | 15.442 | 70.6702 | 102.1269 | 68.2551 | 0.003826 | 1.0341 | 0.025229 | 24.3968 | 0.036859 | 0.13182 | -0.000107 | -0.008186 | -0.008186 | -0.66507
2019-03-26 16:29 | 1553632134.6755 | 955 | 0.20955 | 0.000441 | 0 | 96.6094 | 96.2476 | 127.2502 | 21.2612 | 0.002025 | 0.001942 | -152.8153 | 0.22007 | 8.4004 | 0.003331 | 0.004731 | 0.004524 | 5.5e-05 | 0.001043 | 0.32637 | 1.3715 | 2.5476 | 100.8 | 21.05 | 15.441 | 70.5431 | 101.9967 | 68.5546 | 0.003826 | 1.0332 | 0.025203 | 24.3935 | 0.033559 | 0.13092 | -8.2e-05 | -0.006196 | -0.006196 | -0.87269
2019-03-26 16:29 | 1553632135.6745 | 956 | 0.20955 | 0.00044 | 0 | 96.4128 | 95.8674 | 127.3279 | 21.2612 | 0.002029 | 0.001942 | -152.8358 | 0.22024 | 8.4018 | 0.003375 | 0.004904 | 0.004669 | 5.5e-05 | 0.001042 | 0.30735 | 1.3806 | 2.5476 | 100.8 | 21.05 | 15.439 | 70.6389 | 103.0642 | 68.5553 | 0.003825 | 1.0332 | 0.025211 | 24.401 | 0.050742 | 0.13102 | -9.9e-05 | -0.007583 | -0.007583 | -0.71363
2019-03-26 16:29 | 1553632136.6758 | 957 | 0.20957 | 0.000439 | 0 | 96.3178 | 95.6889 | 127.1145 | 21.2612 | 0.002018 | 0.001913 | -152.8304 | 0.22068 | 8.4068 | 0.003371 | 0.004648 | 0.004508 | 5.8e-05 | 0.001049 | 0.32875 | 1.3685 | 2.5476 | 100.8 | 21.05 | 15.437 | 70.3864 | 102.9637 | 67.862 | 0.003825 | 1.0353 | 0.025215 | 24.3558 | 0.05859 | 0.13128 | -0.000197 | -0.014963 | -0.014963 | -0.36237
2019-03-26 16:29 | 1553632137.6757 | 958 | 0.20953 | 0.000443 | 0 | 96.1725 | 95.6835 | 127.0955 | 21.2612 | 0.002038 | 0.001925 | -152.8103 | 0.22005 | 8.4013 | 0.003416 | 0.004875 | 0.004608 | 5.5e-05 | 0.001045 | 0.34058 | 1.3759 | 2.5476 | 100.8 | 21.05 | 15.434 | 70.855 | 104.0646 | 68.1373 | 0.003824 | 1.0345 | 0.025203 | 24.3638 | 0.045596 | 0.13091 | -5e-06 | -0.00034 | -0.00034 | -15.8902
2019-03-26 16:29 | 1553632138.6735 | 959 | 0.20957 | 0.000441 | 0 | 96.0654 | 95.6297 | 127.2152 | 21.2612 | 0.002015 | 0.001953 | -152.8146 | 0.2204 | 8.4035 | 0.003488 | 0.004743 | 0.004619 | 5.4e-05 | 0.001049 | 0.34147 | 1.3516 | 2.5476 | 100.8 | 21.05 | 15.433 | 70.3114 | 105.8069 | 68.8097 | 0.003824 | 1.0324 | 0.02519 | 24.3993 | 0.040654 | 0.13112 | -0.000224 | -0.016987 | -0.016987 | -0.31878
2019-03-26 16:29 | 1553632139.6737 | 960 | 0.20954 | 0.000443 | 0 | 95.8694 | 95.1752 | 127.4901 | 21.2612 | 0.002 | 0.001927 | -152.8361 | 0.22068 | 8.4042 | 0.003487 | 0.004732 | 0.004604 | 6.3e-05 | 0.001051 | 0.3063 | 1.3531 | 2.5476 | 100.8 | 21.05 | 15.432 | 69.9552 | 105.78 | 68.1894 | 0.003824 | 1.0343 | 0.02524 | 24.4034 | 0.064997 | 0.13129 | -5e-05 | -0.00378 | -0.00378 | -1.4342
2019-03-26 16:29 | 1553632140.6741 | 961 | 0.20957 | 0.000437 | -1e-06 | 95.6899 | 95.0556 | 127.5396 | 21.2612 | 0.00201 | 0.001922 | -152.8183 | 0.22074 | 8.4038 | 0.003411 | 0.004506 | 0.004302 | 6.1e-05 | 0.001052 | 0.33268 | 1.3404 | 2.5474 | 100.8 | 21.06 | 15.43 | 70.1898 | 103.9248 | 68.0702 | 0.003825 | 1.0347 | 0.02525 | 24.4039 | 0.05949 | 0.13132 | -0.000187 | -0.014186 | -0.014186 | -0.38232
2019-03-26 16:29 | 1553632141.6737 | 962 | 0.20951 | 0.000442 | 0 | 95.4909 | 94.8515 | 127.4739 | 21.2612 | 0.002047 | 0.001959 | -152.8237 | 0.21964 | 8.4058 | 0.003413 | 0.004334 | 0.004144 | 5.8e-05 | 0.001059 | 0.35778 | 1.3497 | 2.5476 | 100.8 | 21.06 | 15.429 | 71.078 | 103.9726 | 68.957 | 0.003825 | 1.032 | 0.02521 | 24.4293 | 0.060095 | 0.13067 | 0.000161 | 0.012262 | 0.012262 | 0.44011
2019-03-26 16:29 | 1553632142.6738 | 963 | 0.20949 | 0.000443 | 0 | 95.349 | 94.711 | 127.0615 | 21.2612 | 0.001993 | 0.001919 | -152.8173 | 0.22009 | 8.4049 | 0.003413 | 0.004258 | 0.00413 | 5.6e-05 | 0.001064 | 0.34235 | 1.3314 | 2.5476 | 100.8 | 21.06 | 15.422 | 69.7774 | 103.9743 | 67.9993 | 0.003823 | 1.0349 | 0.025205 | 24.3556 | 0.060046 | 0.13093 | 0.000248 | 0.018886 | 0.018886 | 0.28632
2019-03-26 16:29 | 1553632143.6748 | 964 | 0.20953 | 0.000441 | 0 | 95.1872 | 94.5248 | 127.1286 | 21.2612 | 0.001994 | 0.001949 | -152.8231 | 0.22098 | 8.4001 | 0.003428 | 0.004498 | 0.004313 | 5.9e-05 | 0.001066 | 0.31078 | 1.3811 | 2.5477 | 100.8 | 21.06 | 15.419 | 69.7985 | 104.3375 | 68.719 | 0.003823 | 1.0327 | 0.025185 | 24.3877 | 0.062464 | 0.13146 | 2.3e-05 | 0.001751 | 0.001751 | 3.101
2019-03-26 16:29 | 1553632144.6738 | 965 | 0.20953 | 0.000442 | -1e-06 | 95.0423 | 94.3224 | 127.4049 | 21.2612 | 0.001997 | 0.001924 | -152.8136 | 0.21985 | 8.4033 | 0.003347 | 0.004663 | 0.004467 | 5.9e-05 | 0.001063 | 0.31885 | 1.3742 | 2.5476 | 100.8 | 21.06 | 15.414 | 69.8819 | 102.3813 | 68.1168 | 0.003821 | 1.0345 | 0.025235 | 24.3927 | 0.068004 | 0.13079 | 3e-05 | 0.002337 | 0.002337 | 2.3115
2019-03-26 16:29 | 1553632145.6753 | 966 | 0.20953 | 0.000437 | 0 | 94.9015 | 93.9701 | 127.6063 | 21.2612 | 0.002009 | 0.001927 | -152.7934 | 0.22017 | 8.403 | 0.003365 | 0.004378 | 0.004187 | 5.9e-05 | 0.001063 | 0.34451 | 1.3355 | 2.5475 | 100.8 | 21.06 | 15.402 | 70.1571 | 102.8104 | 68.2043 | 0.003819 | 1.0342 | 0.025251 | 24.4151 | 0.088214 | 0.13098 | 5.1e-05 | 0.003917 | 0.003917 | 1.3811
2019-03-26 16:29 | 1553632146.6753 | 967 | 0.20952 | 0.000442 | 0 | 94.7069 | 93.8662 | 127.6527 | 21.2612 | 0.00197 | 0.001959 | -152.7901 | 0.22031 | 8.4057 | 0.003384 | 0.004375 | 0.004234 | 5.8e-05 | 0.001065 | 0.34147 | 1.3703 | 2.5476 | 100.8 | 21.06 | 15.396 | 69.2205 | 103.2756 | 68.9695 | 0.003817 | 1.0319 | 0.025228 | 24.4469 | 0.079754 | 0.13106 | 8.6e-05 | 0.006556 | 0.006556 | 0.82566
2019-03-26 16:29 | 1553632147.6744 | 968 | 0.20955 | 0.000445 | 0 | 94.5021 | 93.8242 | 127.7016 | 21.2612 | 0.001999 | 0.00193 | -152.8007 | 0.22029 | 8.4034 | 0.003331 | 0.004518 | 0.004333 | 5.6e-05 | 0.00107 | 0.35369 | 1.3797 | 2.5475 | 100.8 | 21.06 | 15.385 | 69.9302 | 101.995 | 68.2729 | 0.003814 | 1.034 | 0.025258 | 24.4267 | 0.06439 | 0.13105 | -7.5e-05 | -0.005674 | -0.005674 | -0.95383
2019-03-26 16:29 | 1553632148.6741 | 969 | 0.20955 | 0.00044 | 0 | 94.2993 | 93.6471 | 127.6436 | 21.2612 | 0.002003 | 0.001915 | -152.7854 | 0.22015 | 8.4027 | 0.003343 | 0.004482 | 0.004292 | 5.7e-05 | 0.001079 | 0.34772 | 1.3421 | 2.5475 | 100.8 | 21.07 | 15.381 | 70.0276 | 102.2864 | 67.8954 | 0.003816 | 1.0352 | 0.025266 | 24.4076 | 0.06208 | 0.13097 | -0.000115 | -0.008778 | -0.008778 | -0.61619
2019-03-26 16:29 | 1553632149.6745 | 970 | 0.20953 | 0.000443 | 0 | 94.2278 | 93.3418 | 127.5429 | 21.2612 | 0.00206 | 0.001911 | -152.7931 | 0.22102 | 8.4032 | 0.003312 | 0.004986 | 0.004798 | 5.8e-05 | 0.001081 | 0.31133 | 1.3415 | 2.5475 | 100.8 | 21.07 | 15.374 | 71.3976 | 101.5262 | 67.8002 | 0.003814 | 1.0355 | 0.02526 | 24.3946 | 0.084503 | 0.13148 | 3.8e-05 | 0.00288 | 0.00288 | 1.8856
2019-03-26 16:29 | 1553632150.6752 | 971 | 0.20952 | 0.000442 | 0 | 94.0785 | 93.2084 | 127.5471 | 21.2612 | 0.002037 | 0.001911 | -152.8064 | 0.22106 | 8.4031 | 0.003362 | 0.004854 | 0.004638 | 5.9e-05 | 0.001081 | 0.30879 | 1.3688 | 2.5476 | 100.8 | 21.07 | 15.357 | 70.8346 | 102.7372 | 67.8057 | 0.00381 | 1.0355 | 0.02526 | 24.3952 | 0.083112 | 0.13151 | 6e-05 | 0.004585 | 0.004585 | 1.1846
2019-03-26 16:29 | 1553632151.6746 | 972 | 0.20952 | 0.00044 | 0 | 93.9321 | 93.2198 | 127.6103 | 21.2612 | 0.001963 | 0.001864 | -152.7986 | 0.22067 | 8.4 | 0.003326 | 0.004145 | 0.004009 | 6.1e-05 | 0.001084 | 0.36248 | 1.3253 | 2.5475 | 100.8 | 21.07 | 15.348 | 69.0584 | 101.8701 | 66.6722 | 0.003807 | 1.0389 | 0.025309 | 24.3606 | 0.068084 | 0.13128 | 0.000126 | 0.009613 | 0.009613 | 0.56403
2019-03-26 16:29 | 1553632152.6753 | 973 | 0.20952 | 0.00044 | 0 | 93.8001 | 92.8826 | 128.0231 | 21.2612 | 0.001979 | 0.001871 | -152.8224 | 0.2205 | 8.4035 | 0.003306 | 0.004152 | 0.003983 | 5.7e-05 | 0.001085 | 0.33826 | 1.3482 | 2.5474 | 100.8 | 21.07 | 15.339 | 69.4327 | 101.3848 | 66.8461 | 0.003805 | 1.0384 | 0.025343 | 24.4062 | 0.087921 | 0.13118 | 9.4e-05 | 0.00719 | 0.00719 | 0.75346
2019-03-26 16:29 | 1553632153.6733 | 974 | 0.20952 | 0.000438 | 0 | 93.5879 | 92.5125 | 128.221 | 21.2612 | 0.001974 | 0.001939 | -152.8292 | 0.2206 | 8.4034 | 0.003366 | 0.004024 | 0.003799 | 6.1e-05 | 0.001084 | 0.33301 | 1.3408 | 2.5475 | 100.8 | 21.07 | 15.317 | 69.318 | 102.8496 | 68.489 | 0.0038 | 1.0334 | 0.025301 | 24.484 | 0.10337 | 0.13123 | 7.1e-05 | 0.005485 | 0.005485 | 0.98823
2019-03-26 16:29 | 1553632154.6748 | 975 | 0.20953 | 0.00044 | 0 | 93.4943 | 92.6755 | 128.1377 | 21.2612 | 0.002012 | 0.001888 | -152.8356 | 0.22005 | 8.4008 | 0.003309 | 0.005111 | 0.005001 | 5.7e-05 | 0.001085 | 0.32101 | 1.3365 | 2.5474 | 100.8 | 21.07 | 15.307 | 70.246 | 101.4719 | 67.265 | 0.003797 | 1.0371 | 0.025339 | 24.4322 | 0.078675 | 0.13091 | 3.5e-05 | 0.002697 | 0.002697 | 2.0044
2019-03-26 16:29 | 1553632155.6748 | 976 | 0.20951 | 0.000436 | 0 | 93.4202 | 92.3441 | 128.0799 | 21.2612 | 0.00198 | 0.001909 | -152.8458 | 0.22046 | 8.4065 | 0.003381 | 0.004896 | 0.00474 | 6e-05 | 0.001095 | 0.31852 | 1.3707 | 2.5475 | 100.8 | 21.07 | 15.286 | 69.4748 | 103.2032 | 67.7658 | 0.003792 | 1.0356 | 0.025314 | 24.4446 | 0.10363 | 0.13115 | 0.000177 | 0.013525 | 0.013525 | 0.40047
2019-03-26 16:29 | 1553632156.6744 | 977 | 0.20953 | 0.000435 | 0 | 93.2109 | 92.3834 | 128.4499 | 21.2612 | 0.001964 | 0.001921 | -152.8602 | 0.22038 | 8.4039 | 0.003289 | 0.004693 | 0.004458 | 5.9e-05 | 0.001086 | 0.32101 | 1.449 | 2.5474 | 100.8 | 21.07 | 15.277 | 69.0755 | 100.9853 | 68.0397 | 0.00379 | 1.0347 | 0.025341 | 24.4897 | 0.079762 | 0.1311 | 3.5e-05 | 0.002716 | 0.002716 | 1.9938
2019-03-26 16:29 | 1553632157.6745 | 978 | 0.20952 | 0.000437 | 0 | 93.0642 | 92.1814 | 128.3412 | 21.2612 | 0.001994 | 0.001898 | -152.8587 | 0.22093 | 8.4026 | 0.003294 | 0.004812 | 0.00471 | 5.6e-05 | 0.001092 | 0.31205 | 1.3441 | 2.5474 | 100.8 | 21.07 | 15.268 | 69.7954 | 101.0976 | 67.4867 | 0.003788 | 1.0364 | 0.025351 | 24.4595 | 0.085256 | 0.13143 | 0.00012 | 0.009204 | 0.009204 | 0.58973
2019-03-26 16:29 | 1553632158.7307 | 979 | 0.20953 | 0.000439 | 0 | 92.8918 | 91.8093 | 128.4302 | 21.2612 | 0.002012 | 0.001932 | -152.8626 | 0.22079 | 8.4029 | 0.003282 | 0.00424 | 0.004056 | 5.6e-05 | 0.001092 | 0.30735 | 1.3577 | 2.5475 | 100.8 | 21.07 | 15.249 | 70.235 | 100.7986 | 68.3183 | 0.003783 | 1.0339 | 0.025328 | 24.4979 | 0.10485 | 0.13135 | 6.1e-05 | 0.004637 | 0.004637 | 1.1697
2019-03-26 16:29 | 1553632159.6825 | 980 | 0.20952 | 0.000434 | 0 | 92.6817 | 91.7044 | 128.6766 | 21.2612 | 0.00193 | 0.001902 | -152.8481 | 0.22005 | 8.4041 | 0.003282 | 0.004316 | 0.004058 | 5.5e-05 | 0.001098 | 0.3449 | 1.3457 | 2.5475 | 100.8 | 21.07 | 15.24 | 68.276 | 100.8063 | 67.5881 | 0.003781 | 1.0361 | 0.02538 | 24.4951 | 0.094817 | 0.13091 | 7.2e-05 | 0.00549 | 0.00549 | 0.98487
2019-03-26 16:29 | 1553632160.6825 | 981 | 0.20951 | 0.000437 | 0 | 92.4792 | 91.5047 | 128.7456 | 21.2612 | 0.002015 | 0.001899 | -152.8323 | 0.22015 | 8.4011 | 0.003288 | 0.004295 | 0.00414 | 5.6e-05 | 0.001102 | 0.33157 | 1.3026 | 2.5474 | 100.8 | 21.07 | 15.222 | 70.3013 | 100.95 | 67.5201 | 0.003776 | 1.0363 | 0.025389 | 24.4992 | 0.09475 | 0.13096 | 0.000163 | 0.012511 | 0.012511 | 0.43232
2019-03-26 16:29 | 1553632161.6825 | 982 | 0.20953 | 0.000437 | 0 | 92.3756 | 91.4203 | 128.9118 | 21.2612 | 0.001994 | 0.001894 | -152.8226 | 0.2204 | 8.4045 | 0.003339 | 0.004519 | 0.004311 | 5.4e-05 | 0.001108 | 0.31868 | 1.3553 | 2.5474 | 100.8 | 21.07 | 15.215 | 69.8076 | 102.1828 | 67.3977 | 0.003774 | 1.0367 | 0.02541 | 24.5107 | 0.092975 | 0.13112 | 1.3e-05 | 0.001003 | 0.001003 | 5.3985
2019-03-26 16:29 | 1553632162.6838 | 983 | 0.20954 | 0.000439 | -1e-06 | 92.2061 | 91.1924 | 128.7797 | 21.2612 | 0.001977 | 0.001883 | -152.8002 | 0.21988 | 8.4029 | 0.003323 | 0.005331 | 0.00516 | 5.5e-05 | 0.001101 | 0.3298 | 1.3178 | 2.5475 | 100.8 | 21.07 | 15.209 | 69.4007 | 101.7987 | 67.1411 | 0.003773 | 1.0375 | 0.025407 | 24.4889 | 0.098881 | 0.13081 | -2.5e-05 | -0.001921 | -0.001921 | -2.812
2019-03-26 16:29 | 1553632163.6833 | 984 | 0.20952 | 0.00044 | 0 | 92.2014 | 91.0694 | 128.5355 | 21.2612 | 0.001983 | 0.001898 | -152.7891 | 0.22093 | 8.4039 | 0.003371 | 0.004524 | 0.004361 | 5.6e-05 | 0.001101 | 0.30143 | 1.2986 | 2.5474 | 100.79 | 21.07 | 15.2 | 69.5505 | 102.9532 | 67.5047 | 0.003771 | 1.0364 | 0.025369 | 24.4787 | 0.1105 | 0.13143 | 0.00011 | 0.008482 | 0.008482 | 0.63997
2019-03-26 16:29 | 1553632164.6841 | 985 | 0.20954 | 0.000435 | 0 | 92.0553 | 90.8465 | 128.6672 | 21.2612 | 0.001977 | 0.001904 | -152.8169 | 0.22067 | 8.4055 | 0.003318 | 0.004324 | 0.004203 | 5.5e-05 | 0.001098 | 0.31902 | 1.3412 | 2.5474 | 100.79 | 21.07 | 15.198 | 69.3976 | 101.6776 | 67.6504 | 0.003771 | 1.0359 | 0.025377 | 24.4965 | 0.11823 | 0.13128 | -1.3e-05 | -0.000977 | -0.000977 | -5.5521
2019-03-26 16:29 | 1553632165.6831 | 986 | 0.20952 | 0.000439 | 0 | 91.8079 | 90.6663 | 128.6712 | 21.2612 | 0.002007 | 0.001906 | -152.8079 | 0.22111 | 8.4042 | 0.003248 | 0.004398 | 0.004281 | 5.8e-05 | 0.0011 | 0.32228 | 1.3737 | 2.5477 | 100.79 | 21.07 | 15.198 | 70.1119 | 99.9924 | 67.6793 | 0.003771 | 1.0358 | 0.025376 | 24.4979 | 0.11192 | 0.13154 | 8.1e-05 | 0.006256 | 0.006256 | 0.86843
2019-03-26 16:29 | 1553632166.6828 | 987 | 0.20954 | 0.000437 | 0 | 91.7073 | 90.4787 | 128.0297 | 21.2612 | 0.001976 | 0.001885 | -152.8136 | 0.22074 | 8.4048 | 0.003236 | 0.004926 | 0.004683 | 5.3e-05 | 0.001097 | 0.32012 | 1.333 | 2.5475 | 100.79 | 21.07 | 15.196 | 69.3632 | 99.7037 | 67.1714 | 0.00377 | 1.0374 | 0.025331 | 24.4185 | 0.12063 | 0.13132 | -4e-05 | -0.003078 | -0.003078 | -1.7618
2019-03-26 16:29 | 1553632167.6839 | 988 | 0.20953 | 0.000436 | 0 | 91.5307 | 90.1903 | 127.6633 | 21.2612 | 0.002005 | 0.001863 | -152.825 | 0.2204 | 8.4023 | 0.003377 | 0.004777 | 0.004594 | 5.4e-05 | 0.001099 | 0.32173 | 1.3692 | 2.5476 | 100.79 | 21.07 | 15.196 | 70.0689 | 103.101 | 66.6455 | 0.00377 | 1.039 | 0.025315 | 24.3647 | 0.13196 | 0.13112 | 1.6e-05 | 0.001266 | 0.001266 | 4.2785
2019-03-26 16:29 | 1553632168.6832 | 989 |  | 0.00044 |  | 91.4109 | 90.3604 | 127.4751 | 21.2612 | 0.002014 | 0.001881 | -152.8136 | 0.22059 | 8.4008 | 0.003292 | 0.004742 | 0.004545 | 5.2e-05 | 0.001102 | 0.35402 | 1.3494 | 2.5474 | 100.79 | 21.07 | 15.199 | 70.2919 | 101.0582 | 67.0786 | 0.003771 | 1.0377 | 0.02528 | 24.3623 | 0.10339 | 0.13123
2019-03-26 16:29 | 1553632169.6832 | 990 |  | 0.000437 |  | 91.2686 | 90.1224 | 127.0833 | 21.2612 | 0.001981 | 0.001848 | -152.8255 | 0.22046 | 8.4007 | 0.003319 | 0.004396 | 0.004273 | 5.1e-05 | 0.001108 | 0.31902 | 1.3662 | 2.5476 | 100.79 | 21.07 | 15.202 | 69.4873 | 101.6934 | 66.2975 | 0.003772 | 1.0401 | 0.02527 | 24.2969 | 0.11304 | 0.13115
2019-03-26 16:29 | 1553632170.6832 | 991 |  |  |  | 91.2282 | 89.8962 | 127.1239 | 21.2612 | 0.001941 | 0.001894 | -152.8368 | 0.21914 | 8.4003 | 0.003223 | 0.004324 | 0.004112 | 5.3e-05 | 0.001109 | 0.31902 | 1.3791 | 2.5476 | 100.79 | 21.07 | 15.205 | 68.5296 | 99.3839 | 67.39 | 0.003772 | 1.0367 | 0.025234 | 24.3398 | 0.13156 | 0.13037
2019-03-26 16:30 | 1553632171.6832 | 992 |  |  |  | 91.1089 | 90.1147 | 127.2004 | 21.2612 | 0.00195 | 0.001925 | -152.8634 | 0.21939 | 8.4008 | 0.003252 | 0.004301 | 0.004073 | 5.1e-05 | 0.001109 | 0.33682 | 1.3682 | 2.5476 | 100.79 | 21.07 | 15.205 | 68.7339 | 100.0836 | 68.1519 | 0.003772 | 1.0344 | 0.025213 | 24.3744 | 0.098146 | 0.13051
2019-03-26 16:30 | 1553632172.6832 | 993 |  |  |  | 90.9944 | 89.8004 | 127.2863 | 21.2612 | 0.001981 | 0.001847 | -152.8804 | 0.22009 | 8.4015 | 0.00322 | 0.004395 | 0.0042 | 5.2e-05 | 0.001105 | 0.32444 | 1.3779 | 2.5475 | 100.79 | 21.07 | 15.207 | 69.4849 | 99.3203 | 66.2586 | 0.003773 | 1.0402 | 0.025292 | 24.3149 | 0.11815 | 0.13093
2019-03-26 16:30 | 1553632173.6832 | 994 |  |  |  | 90.7824 | 89.558 |  | 21.2612 | 0.002 | 0.001879 | -152.8993 | 0.22038 | 8.4031 | 0.003171 | 0.004232 | 0.004035 | 5.1e-05 | 0.001104 | 0.33516 | 1.4161 | 2.5476 | 100.79 | 21.07 | 15.209 | 69.9466 | 98.1269 | 67.0334 | 0.003773 | 1.0378 |  |  | 0.12146 | 0.1311
2019-03-26 16:30 | 1553632174.6832 | 995 |  |  |  | 90.6142 | 89.4497 |  | 21.2612 | 0.001986 |  | -152.8835 | 0.21979 | 8.3997 | 0.003158 | 0.004439 | 0.00427 | 5.4e-05 | 0.001106 | 0.32908 | 1.3895 | 2.5475 | 100.79 | 21.07 | 15.208 | 69.612 | 97.8212 |  | 0.003773 |  |  |  | 0.11569 | 0.13075
2019-03-26 16:30 | 1553632175.6832 | 996 |  |  |  | 90.4633 | 89.1889 |  | 21.2612 | 0.001961 |  |  | 0.21979 | 8.4041 | 0.003148 | 0.00396 | 0.003813 | 5.2e-05 | 0.001106 | 0.33196 | 1.3597 | 2.5476 | 100.79 | 21.07 | 15.208 | 69.0123 | 97.5601 |  | 0.003773 |  |  |  | 0.12691 | 0.13076
2019-03-26 16:30 | 1553632176.6832 | 997 |  |  |  | 90.3352 | 88.9327 |  | 21.2612 |  |  |  | 0.22078 | 8.4066 | 0.003216 | 0.004271 | 0.004032 | 5.4e-05 | 0.00111 | 0.31725 | 1.3455 | 2.5475 | 100.79 | 21.07 | 15.202 |  | 99.205 |  | 0.003772 |  |  |  | 0.13996 | 0.13134
2019-03-26 16:30 | 1553632177.6832 | 998 |  |  |  | 90.221 | 88.7899 |  | 21.2612 |  |  |  | 0.21981 | 8.3997 | 0.003199 | 0.004881 | 0.004675 | 5e-05 | 0.001112 | 0.27721 | 1.33 | 2.5474 | 100.79 | 21.07 | 15.198 |  | 98.7996 |  | 0.003771 |  |  |  | 0.14302 | 0.13077
2019-03-26 16:30 | 1553632178.6832 | 999 |  |  |  | 90.0997 | 88.9798 |  | 21.2612 |  |  |  | 0.22024 | 8.4017 | 0.003198 | 0.005256 | 0.005055 | 5e-05 | 0.001118 | 0.33284 | 1.3603 | 2.5475 | 100.79 | 21.07 | 15.191 |  | 98.7783 |  | 0.003769 |  |  |  | 0.11187 | 0.13102
2019-03-26 16:30 | 1553632179.6832 | 1000 |  |  |  | 90.0191 | 88.7435 |  | 21.2612 |  |  |  | 0.21955 | 8.402 | 0.003152 | 0.0046 | 0.004484 | 4.8e-05 | 0.00112 | 0.30326 | 1.357 | 2.5476 | 100.79 | 21.07 | 15.186 |  | 97.6582 |  | 0.003768 |  |  |  | 0.12766 | 0.13061
2019-03-26 16:30 | 1553632180.6832

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-26 16:13 | 0 | 0 | 0 | 0 | 0 | 51 | -0.015133
2019-03-26 16:13 | 1 | 0 | 0 | 0 | 0 | 51.1434 | -0.011988
2019-03-26 16:13 | 2 | 0 | 0 | 0 | 0 | 51.2133 | 0.010759
2019-03-26 16:13 | 3 | 0 | 0 | 0 | 0 | 50.9789 | 0.050355
2019-03-26 16:13 | 4 | 0 | 0 | 0 | 0 | 51.0326 | 0.36196
2019-03-26 16:13 | 5 | 0 | 0 | 0 | 0 | 51.2197 | 1.032
2019-03-26 16:13 | 6 | 0 | 0 | 0 | 0 | 50.8325 | 1.4973
2019-03-26 16:13 | 7 | 0 | 0 | 0 | 0 | 50.7318 | 1.8872
2019-03-26 16:13 | 8 | 0 | 0 | 0 | 0 | 50.6276 | 2.1764
2019-03-26 16:13 | 9 | 0 | 0 | 0 | 0 | 50.4979 | 2.363
2019-03-26 16:13 | 10 | 0 | 0 | 0 | 0 | 50.3341 | 2.5505
2019-03-26 16:13 | 11 | 0 | 0 | 0 | 0 | 50.1523 | 3.3462
2019-03-26 16:13 | 12 | 0 | 0 | 0 | 0 | 49.9749 | 3.8552
2019-03-26 16:13 | 13 | 0 | 0 | 0 | 0 | 49.8147 | 4.2178
2019-03-26 16:13 | 14 | 0.18392 | 0 | 0 | 0 | 49.6493 | 4.4132
2019-03-26 16:13 | 15 | 0.20185 | 0 | 0 | 0 | 49.4491 | 4.3585
2019-03-26 16:13 | 16 | 0.17562 | 0 | 0 | 0 | 49.2568 | 4.297
2019-03-26 16:13 | 17 | 0.15628 | 0 | 0 | 0 | 49.0971 | 4.2129
2019-03-26 16:13 | 18 | 0.146 | 0 | 0 | 0 | 48.9413 | 4.0907
2019-03-26 16:13 | 19 | 0.15644 | 0 | 0 | 0 | 48.8004 | 3.9224
2019-03-26 16:13 | 20 | 0.14215 | 0 | 0 | 0 | 48.6347 | 3.7772
2019-03-26 16:13 | 21 | 0.10423 | 0 | 0 | 0 | 48.5217 | 3.6248
2019-03-26 16:13 | 22 | 0.10113 | 0 | 0 | 0 | 48.418 | 3.4341
2019-03-26 16:13 | 23 | 0.10211 | 0 | 0 | 0 | 48.3177 | 3.3017
2019-03-26 16:13 | 24 | 0.092867 | 0 | 0 | 0 | 48.2159 | 3.1626
2019-03-26 16:13 | 25 | 0.088703 | 0 | 0 | 0 | 48.1305 | 3.1222
2019-03-26 16:13 | 26 | 0.10096 | 0 | 0 | 0 | 48.0346 | 2.9346
2019-03-26 16:13 | 27 | 0.098396 | 0 | 0 | 0 | 47.9316 | 3.0089
2019-03-26 16:13 | 28 | 0.10777 | 0 | 0 | 0 | 47.8362 | 2.8409
2019-03-26 16:13 | 29 | 0.084821 | 0 | 0 | 0 | 47.7237 | 2.5546
2019-03-26 16:13 | 30 | 0.066098 | 0 | 0 | 0 | 47.6652 | 2.412
2019-03-26 16:13 | 31 | 0.088905 | 0 | 0 | 0 | 47.5813 | 2.3578
2019-03-26 16:13 | 32 | 0.094935 | 0 | 0 | 0 | 47.4902 | 2.3515
2019-03-26 16:14 | 33 | 0.1074 | 0 | 0 | 0 | 47.3909 | 2.2746
2019-03-26 16:14 | 34 | 0.097655 | 0 | 0 | 0 | 47.2811 | 2.166
2019-03-26 16:14 | 35 | 0.079595 | 0 | 0 | 0 | 47.1976 | 2.1187
2019-03-26 16:14 | 36 | 0.09111 | 0 | 0 | 0 | 47.1152 | 2.0895
2019-03-26 16:14 | 37 | 0.088202 | 0 | 0 | 0 | 47.0186 | 2.0772
2019-03-26 16:14 | 38 | 0.065167 | 0 | 0 | 0 | 46.9429 | 2.0756
2019-03-26 16:14 | 39 | 0.087981 | 0 | 0 | 0 | 46.8775 | 2.0978
2019-03-26 16:14 | 40 | 0.1049 | 0 | 0 | 0 | 46.7682 | 2.1297
2019-03-26 16:14 | 41 | 0.082868 | 0 | 0 | 0 | 46.6767 | 2.1739
2019-03-26 16:14 | 42 | 0.094754 | 0 | 0 | 0 | 46.5955 | 2.2023
2019-03-26 16:14 | 43 | 0.090985 | 0 | 0 | 0 | 46.4902 | 2.3152
2019-03-26 16:14 | 44 | 0.089188 | 0 | 0 | 0 | 46.413 | 2.369
2019-03-26 16:14 | 45 | 0.10387 | 0.016856 | 1.7962 | 23.0979 | 46.3085 | 2.3991
2019-03-26 16:14 | 46 | 0.087075 | 0.019899 | 2.1557 | 27.635 | 46.2119 | 2.4063
2019-03-26 16:14 | 47 | 0.09787 | 0.017689 | 1.9279 | 24.4082 | 46.1284 | 2.3888
2019-03-26 16:14 | 48 | 0.1048 | 0.016372 | 1.8198 | 22.5101 | 46.0174 | 2.3591
2019-03-26 16:14 | 49 | 0.076195 | 0.022512 | 2.5222 | 30.2991 | 45.9267 | 2.3086
2019-03-26 16:14 | 50 | 0.087204 | 0.019481 | 2.2197 | 26.4362 | 45.8558 | 2.3053
2019-03-26 16:14 | 51 | 0.10835 | 0.015483 | 1.7934 | 20.662 | 45.75 | 2.2387
2019-03-26 16:14 | 52 | 0.083095 | 0.019997 | 2.3447 | 26.8395 | 45.6499 | 2.2302
2019-03-26 16:14 | 53 | 0.079293 | 0.020777 | 2.4624 | 27.3385 | 45.5799 | 2.1677
2019-03-26 16:14 | 54 | 0.070463 | 0.023447 | 2.7664 | 30.4327 | 45.4922 | 2.1444
2019-03-26 16:14 | 55 | 0.053226 | 0.031029 | 3.6385 | 38.8193 | 45.4398 | 2.0662
2019-03-26 16:14 | 56 | 0.083275 | 0.019865 | 2.3104 | 24.1279 | 45.3751 | 2.0093
2019-03-26 16:14 | 57 | 0.077007 | 0.021501 | 2.4706 | 25.0001 | 45.2812 | 1.9252
2019-03-26 16:14 | 58 | 0.064068 | 0.025526 | 2.9315 | 28.3795 | 45.2228 | 1.8182
2019-03-26 16:14 | 59 | 0.08458 | 0.019128 | 2.1868 | 20.6722 | 45.1459 | 1.7485
2019-03-26 16:14 | 60 | 0.070273 | 0.022417 | 2.5806 | 24.1973 | 45.0616 | 1.7004
2019-03-26 16:14 | 61 | 0.051479 | 0.030094 | 3.4358 | 32.4927 | 45.0066 | 1.6727
2019-03-26 16:14 | 62 | 0.062941 | 0.024341 | 2.7495 | 25.076 | 44.9515 | 1.5783
2019-03-26 16:14 | 63 | 0.063575 | 0.023385 | 2.6519 | 24.4584 | 44.8826 | 1.555
2019-03-26 16:14 | 64 | 0.06311 | 0.023193 | 2.6057 | 23.9943 | 44.8248 | 1.5143
2019-03-26 16:14 | 65 | 0.065057 | 0.022151 | 2.4644 | 22.6215 | 44.7563 | 1.4717
2019-03-26 16:14 | 66 | 0.048234 | 0.029396 | 3.2468 | 30.2762 | 44.6982 | 1.4603
2019-03-26 16:14 | 67 | 0.075128 | 0.018627 | 2.0419 | 19.3604 | 44.6499 | 1.4545
2019-03-26 16:14 | 68 | 0.082188 | 0.01702 | 1.8257 | 17.6537 | 44.5525 | 1.4509
2019-03-26 16:14 | 69 | 0.060563 | 0.023106 | 2.4292 | 23.1398 | 44.4923 | 1.4014
2019-03-26 16:14 | 70 | 0.073944 | 0.018951 | 1.9549 | 19.1669 | 44.4237 | 1.4173
2019-03-26 16:14 | 71 | 0.078263 | 0.017897 | 1.814 | 18.0024 | 44.3462 | 1.4089
2019-03-26 16:14 | 72 | 0.06331 | 0.021787 | 2.208 | 21.6371 | 44.2712 | 1.3699
2019-03-26 16:14 | 73 | 0.082748 | 0.016671 | 1.6603 | 16.3379 | 44.2118 | 1.3519
2019-03-26 16:14 | 74 | 0.092761 | 0.01459 | 1.4559 | 14.1042 | 44.108 | 1.3083
2019-03-26 16:14 | 75 | 0.072129 | 0.018457 | 1.844 | 18.1931 | 44.0338 | 1.3123
2019-03-26 16:14 | 76 | 0.070765 | 0.018849 | 1.8545 | 18.1769 | 43.9606 | 1.2863
2019-03-26 16:14 | 77 | 0.040307 | 0.033635 | 3.2083 | 32.3251 | 43.8979 | 1.3029
2019-03-26 16:14 | 78 | 0.053311 | 0.02582 | 2.3852 | 24.1082 | 43.8696 | 1.2852
2019-03-26 16:14 | 79 | 0.073452 | 0.018733 | 1.7127 | 17.4052 | 43.7895 | 1.2784
2019-03-26 16:14 | 80 | 0.050052 | 0.027522 | 2.4915 | 25.4629 | 43.7331 | 1.2745
2019-03-26 16:14 | 81 | 0.038945 | 0.035459 | 3.1752 | 33.0036 | 43.6869 | 1.2853
2019-03-26 16:14 | 82 | 0.051084 | 0.027455 | 2.3917 | 25.2997 | 43.6499 | 1.2924
2019-03-26 16:14 | 83 | 0.031418 | 0.044727 | 3.8691 | 40.5167 | 43.5913 | 1.2729
2019-03-26 16:14 | 84 | 0.034276 | 0.040986 | 3.5175 | 36.9997 | 43.5812 | 1.2682
2019-03-26 16:14 | 85 | 0.072681 | 0.019716 | 1.6504 | 17.3316 | 43.5147 | 1.2597
2019-03-26 16:14 | 86 | 0.05695 | 0.024722 | 2.084 | 21.7804 | 43.4481 | 1.2404
2019-03-26 16:14 | 87 | 0.055258 | 0.025537 | 2.1372 | 22.4666 | 43.3989 | 1.2415
2019-03-26 16:14 | 88 | 0.051614 | 0.027354 | 2.2748 | 23.9532 | 43.338 | 1.2363
2019-03-26 16:14 | 89 | 0.040235 | 0.034503 | 2.8887 | 30.8371 | 43.2967 | 1.2407
2019-03-26 16:14 | 90 | 0.057629 | 0.023679 | 2.0074 | 21.2548 | 43.2502 | 1.2249
2019-03-26 16:14 | 91 | 0.084649 | 0.01586 | 1.3568 | 14.4233 | 43.1794 | 1.2209
2019-03-26 16:14 | 92 | 0.069582 | 0.019017 | 1.6462 | 17.1456 | 43.0911 | 1.193
2019-03-26 16:14 | 93 | 0.061605 | 0.021483 | 1.8521 | 19.319 | 43.0398 | 1.1901
2019-03-26 16:15 | 94 | 0.050587 | 0.02622 | 2.2434 | 23.7142 | 42.9683 | 1.1996
2019-03-26 16:15 | 95 | 0.047938 | 0.027666 | 2.3514 | 24.4579 | 42.9359 | 1.1725
2019-03-26 16:15 | 96 | 0.06949 | 0.019083 | 1.6141 | 16.8546 | 42.8672 | 1.1712
2019-03-26 16:15 | 97 | 0.048732 | 0.027248 | 2.2871 | 23.6747 | 42.8066 | 1.1537
2019-03-26 16:15 | 98 | 0.037707 | 0.034628 | 2.9384 | 30.5143 | 42.7679 | 1.1506
2019-03-26 16:15 | 99 | 0.048464 | 0.026935 | 2.2686 | 23.6406 | 42.7261 | 1.1457
2019-03-26 16:15 | 100 | 0.037229 | 0.035113 | 2.9408 | 30.9022 | 42.6755 | 1.1505
2019-03-26 16:15 | 101 | 0.034668 | 0.037697 | 3.1396 | 32.2007 | 42.6493 | 1.1163
2019-03-26 16:15 | 102 | 0.058688 | 0.022198 | 1.8371 | 18.7449 | 42.5999 | 1.1001
2019-03-26 16:15 | 103 | 0.059709 | 0.022258 | 1.8013 | 18.6635 | 42.5372 | 1.1144
2019-03-26 16:15 | 104 | 0.053426 | 0.024839 | 1.9988 | 20.7606 | 42.4834 | 1.1091
2019-03-26 16:15 | 105 | 0.033301 | 0.039244 | 3.1938 | 33.1695 | 42.4335 | 1.1046
2019-03-26 16:15 | 106 | 0.033923 | 0.037865 | 3.1175 | 32.5483 | 42.4112 | 1.1041
2019-03-26 16:15 | 107 | 0.062648 | 0.020468 | 1.6765 | 17.4107 | 42.3591 | 1.0907
2019-03-26 16:15 | 108 | 0.050812 | 0.024815 | 2.0547 | 21.2091 | 42.2951 | 1.0777
2019-03-26 16:15 | 109 | 0.046832 | 0.026494 | 2.2228 | 22.1994 | 42.2563 | 1.0396
2019-03-26 16:15 | 110 | 0.053582 | 0.023174 | 1.9307 | 19.6611 | 42.1992 | 1.0535
2019-03-26 16:15 | 111 | 0.037983 | 0.032118 | 2.7112 | 27.8068 | 42.154 | 1.0562
2019-03-26 16:15 | 112 | 0.037758 | 0.031678 | 2.7102 | 27.0363 | 42.1197 | 1.0208
2019-03-26 16:15 | 113 | 0.046937 | 0.025499 | 2.1653 | 21.9453 | 42.0761 | 1.0301
2019-03-26 16:15 | 114 | 0.043455 | 0.027004 | 2.3189 | 23.5462 | 42.0287 | 1.0232
2019-03-26 16:15 | 115 | 0.040251 | 0.02859 | 2.4872 | 25.052 | 41.9891 | 1.0084
2019-03-26 16:15 | 116 | 0.049847 | 0.022187 | 1.9997 | 19.9264 | 41.9451 | 0.99327
2019-03-26 16:15 | 117 | 0.050972 | 0.02074 | 1.9435 | 19.5322 | 41.891 | 0.99559
2019-03-26 16:15 | 118 | 0.059784 | 0.016939 | 1.6461 | 16.4153 | 41.8422 | 0.98137
2019-03-26 16:15 | 119 | 0.046225 | 0.021462 | 2.1193 | 20.9222 | 41.7772 | 0.96713
2019-03-26 16:15 | 120 | 0.022603 | 0.04276 | 4.294 | 42.2441 | 41.7519 | 0.95484
2019-03-26 16:15 | 121 | 0.035458 | 0.027242 | 2.7183 | 26.917 | 41.7232 | 0.95442
2019-03-26 16:15 | 122 | 0.043067 | 0.022466 | 2.2242 | 22.1451 | 41.6815 | 0.95372
2019-03-26 16:15 | 123 | 0.040706 | 0.024319 | 2.3356 | 22.8041 | 41.6395 | 0.92827
2019-03-26 16:15 | 124 | 0.044541 | 0.021704 | 2.1162 | 20.5865 | 41.5989 | 0.91695
2019-03-26 16:15 | 125 | 0.048094 | 0.02009 | 1.9461 | 19.1894 | 41.5511 | 0.92289
2019-03-26 16:15 | 126 | 0.025046 | 0.03779 | 3.7238 | 36.791 | 41.5089 | 0.92148
2019-03-26 16:15 | 127 | 0.017399 | 0.053148 | 5.3327 | 52.8118 | 41.4968 | 0.91887
2019-03-26 16:15 | 128 | 0.04401 | 0.021029 | 2.0975 | 21.076 | 41.4654 | 0.92755
2019-03-26 16:15 | 129 | 0.057105 | 0.016217 | 1.6087 | 15.7994 | 41.4113 | 0.90223
2019-03-26 16:15 | 130 | 0.068093 | 0.013629 | 1.3433 | 12.9544 | 41.3532 | 0.88211
2019-03-26 16:15 | 131 | 0.04257 | 0.021809 | 2.1333 | 20.6232 | 41.2844 | 0.87793
2019-03-26 16:15 | 132 | 0.0186 | 0.049909 | 4.8625 | 46.6144 | 41.2674 | 0.86701
2019-03-26 16:15 | 133 | 0.040411 | 0.022914 | 2.2145 | 21.6526 | 41.236 | 0.87501
2019-03-26 16:15 | 134 | 0.049081 | 0.01889 | 1.816 | 17.6023 | 41.1887 | 0.86394
2019-03-26 16:15 | 135 | 0.049977 | 0.01853 | 1.7712 | 16.5804 | 41.1401 | 0.82863
2019-03-26 16:15 | 136 | 0.052758 | 0.017528 | 1.6671 | 15.7687 | 41.0891 | 0.83192
2019-03-26 16:15 | 137 | 0.040795 | 0.021492 | 2.1455 | 19.6126 | 41.0384 | 0.8001
2019-03-26 16:15 | 138 | 0.024563 | 0.036675 | 3.5452 | 34.0455 | 41.0081 | 0.83625
2019-03-26 16:15 | 139 | 0.040547 | 0.022172 | 2.1316 | 20.3616 | 40.9816 | 0.8256
2019-03-26 16:15 | 140 | 0.042641 | 0.021065 | 2.0124 | 19.1463 | 40.9299 | 0.81641
2019-03-26 16:15 | 141 | 0.030503 | 0.029493 | 2.8069 | 27.4491 | 40.8999 | 0.83728
2019-03-26 16:15 | 142 | 0.031591 | 0.029227 | 2.6994 | 26.0654 | 40.8661 | 0.82342
2019-03-26 16:15 | 143 | 0.029599 | 0.030408 | 2.8756 | 27.8158 | 40.8367 | 0.82331
2019-03-26 16:15 | 144 | 0.041094 | 0.022487 | 2.0645 | 19.1924 | 40.8036 | 0.7887
2019-03-26 16:15 | 145 | 0.047643 | 0.019406 | 1.7786 | 17.6911 | 40.7555 | 0.84286
2019-03-26 16:15 | 146 | 0.049665 | 0.017664 | 1.7006 | 16.4302 | 40.7095 | 0.81601
2019-03-26 16:15 | 147 | 0.056765 | 0.015488 | 1.4898 | 14.7564 | 40.6556 | 0.83766
2019-03-26 16:15 | 148 | 0.049269 | 0.017408 | 1.7152 | 16.3962 | 40.5999 | 0.80782
2019-03-26 16:15 | 149 | 0.026716 | 0.032106 | 3.1566 | 29.5785 | 40.5602 | 0.79022
2019-03-26 16:15 | 150 | 0.035981 | 0.023847 | 2.3386 | 21.765 | 40.538 | 0.78314
2019-03-26 16:15 | 151 | 0.07665 | 0.011215 | 1.0965 | 10.237 | 40.4809 | 0.78466
2019-03-26 16:15 | 152 | 0.068136 | 0.012597 | 1.2252 | 11.6973 | 40.3969 | 0.79701
2019-03-26 16:16 | 153 | 0.035664 | 0.024756 | 2.3341 | 22.3926 | 40.351 | 0.79861
2019-03-26 16:16 | 154 | 0.041655 | 0.021213 | 1.9913 | 19.2378 | 40.3167 | 0.80135
2019-03-26 16:16 | 155 | 0.047234 | 0.018706 | 1.7521 | 17.2024 | 40.2668 | 0.81254
2019-03-26 16:16 | 156 | 0.048247 | 0.018814 | 1.7133 | 16.4492 | 40.2233 | 0.79363
2019-03-26 16:16 | 157 | 0.046739 | 0.018913 | 1.7606 | 16.5383 | 40.1711 | 0.77299
2019-03-26 16:16 | 158 | 0.045923 | 0.019222 | 1.7829 | 17.2501 | 40.1295 | 0.79217
2019-03-26 16:16 | 159 | 0.055061 | 0.016056 | 1.4886 | 14.0113 | 40.0776 | 0.77147
2019-03-26 16:16 | 160 | 0.034062 | 0.025279 | 2.3974 | 22.8583 | 40.0264 | 0.77859
2019-03-26 16:16 | 161 | 0.023799 | 0.035131 | 3.4166 | 32.4326 | 40.0067 | 0.77187
2019-03-26 16:16 | 162 | 0.04481 | 0.01865 | 1.8073 | 17.1386 | 39.9712 | 0.76797
2019-03-26 16:16 | 163 | 0.044444 | 0.019365 | 1.8143 | 17.4189 | 39.9219 | 0.77417
2019-03-26 16:16 | 164 | 0.034421 | 0.024911 | 2.3287 | 22.6132 | 39.8855 | 0.77838
2019-03-26 16:16 | 165 | 0.021445 | 0.040036 | 3.7221 | 35.3782 | 39.8537 | 0.75867
2019-03-26 16:16 | 166 | 0.021022 | 0.040904 | 3.8046 | 36.4418 | 39.8389 | 0.76607
2019-03-26 16:16 | 167 | 0.041662 | 0.020567 | 1.9094 | 18.3321 | 39.8066 | 0.76375
2019-03-26 16:16 | 168 | 0.037737 | 0.022712 | 2.1067 | 20.3608 | 39.761 | 0.76835
2019-03-26 16:16 | 169 | 0.039297 | 0.022447 | 2.0245 | 19.6111 | 39.7303 | 0.77066
2019-03-26 16:16 | 170 | 0.045193 | 0.019478 | 1.7536 | 17.088 | 39.6814 | 0.77226
2019-03-26 16:16 | 171 | 0.040082 | 0.021961 | 1.9816 | 19.1798 | 39.642 | 0.76877
2019-03-26 16:16 | 172 | 0.055878 | 0.016186 | 1.4237 | 13.7756 | 39.5965 | 0.76976
2019-03-26 16:16 | 173 | 0.063756 | 0.014499 | 1.2411 | 12.1685 | 39.5321 | 0.77582
2019-03-26 16:16 | 174 | 0.052745 | 0.017985 | 1.5054 | 14.7153 | 39.4735 | 0.77616
2019-03-26 16:16 | 175 | 0.060373 | 0.01575 | 1.3178 | 12.8709 | 39.4228 | 0.77706
2019-03-26 16:16 | 176 | 0.059872 | 0.015876 | 1.3301 | 12.905 | 39.3544 | 0.77265
2019-03-26 16:16 | 177 | 0.056945 | 0.01672 | 1.3983 | 13.6474 | 39.3036 | 0.77716
2019-03-26 16:16 | 178 | 0.063841 | 0.014863 | 1.2447 | 12.0851 | 39.239 | 0.77153
2019-03-26 16:16 | 179 | 0.040925 | 0.023139 | 1.9422 | 18.3562 | 39.1829 | 0.75123
2019-03-26 16:16 | 180 | 0.030551 | 0.031069 | 2.6067 | 24.7712 | 39.1543 | 0.75678
2019-03-26 16:16 | 181 | 0.036782 | 0.025755 | 2.1608 | 20.2513 | 39.1173 | 0.74488
2019-03-26 16:16 | 182 | 0.038221 | 0.024892 | 2.0818 | 19.4986 | 39.081 | 0.74525
2019-03-26 16:16 | 183 | 0.054338 | 0.017513 | 1.4619 | 13.6128 | 39.038 | 0.73969
2019-03-26 16:16 | 184 | 0.042709 | 0.022847 | 1.8618 | 17.5977 | 38.9789 | 0.75158
2019-03-26 16:16 | 185 | 0.033688 | 0.030295 | 2.3457 | 21.7773 | 38.9524 | 0.73362
2019-03-26 16:16 | 186 | 0.039376 | 0.026496 | 1.9986 | 18.8412 | 38.9081 | 0.74189
2019-03-26 16:16 | 187 | 0.029669 | 0.035968 | 2.6528 | 24.6608 | 38.876 | 0.73166
2019-03-26 16:16 | 188 | 0.054631 | 0.019131 | 1.437 | 13.2966 | 38.8407 | 0.72641
2019-03-26 16:16 | 189 | 0.070253 | 0.015279 | 1.1238 | 10.3661 | 38.7688 | 0.72825
2019-03-26 16:16 | 190 | 0.066061 | 0.016206 | 1.1881 | 10.7554 | 38.7059 | 0.71051
2019-03-26 16:16 | 191 | 0.046245 | 0.022609 | 1.6904 | 15.7056 | 38.6411 | 0.72631
2019-03-26 16:16 | 192 | 0.033137 | 0.031492 | 2.3523 | 21.3293 | 38.6117 | 0.70678
2019-03-26 16:16 | 193 | 0.034956 | 0.029773 | 2.2218 | 20.353 | 38.5701 | 0.71146
2019-03-26 16:16 | 194 | 0.063334 | 0.01644 | 1.2194 | 11.2159 | 38.5356 | 0.71035
2019-03-26 16:16 | 195 | 0.058725 | 0.01815 | 1.3144 | 11.7905 | 38.4517 | 0.69239
2019-03-26 16:16 | 196 | 0.029402 | 0.036302 | 2.6179 | 23.6644 | 38.4242 | 0.69578
2019-03-26 16:16 | 197 | 0.041111 | 0.027015 | 1.8599 | 16.9316 | 38.3839 | 0.69608
2019-03-26 16:16 | 198 | 0.02874 | 0.038695 | 2.6501 | 24.3421 | 38.3464 | 0.69959
2019-03-26 16:16 | 199 | 0.03298 | 0.033778 | 2.31 | 21.2916 | 38.3221 | 0.7022
2019-03-26 16:16 | 200 | 0.058641 | 0.019028 | 1.2969 | 11.8056 | 38.2749 | 0.69229
2019-03-26 16:16 | 201 | 0.074391 | 0.01503 | 1.022 | 9.4703 | 38.2079 | 0.7045
2019-03-26 16:16 | 202 | 0.050946 | 0.022463 | 1.4926 | 13.7102 | 38.136 | 0.69848
2019-03-26 16:16 | 203 | 0.040863 | 0.028581 | 1.8558 | 16.9618 | 38.1037 | 0.69311
2019-03-26 16:16 | 204 | 0.030586 | 0.038946 | 2.4696 | 22.335 | 38.0532 | 0.68314
2019-03-26 16:16 | 205 | 0.04663 | 0.025597 | 1.6208 | 14.6645 | 38.0355 | 0.68381
2019-03-26 16:16 | 206 | 0.077606 | 0.015645 | 0.96733 | 8.7445 | 37.9572 | 0.67863
2019-03-26 16:16 | 207 | 0.044878 | 0.026492 | 1.6677 | 14.9738 | 37.8952 | 0.672
2019-03-26 16:16 | 208 | 0.037778 | 0.030912 | 1.976 | 18.0219 | 37.8623 | 0.68084
2019-03-26 16:16 | 209 | 0.046752 | 0.024965 | 1.5881 | 14.4224 | 37.8153 | 0.67428
2019-03-26 16:16 | 210 | 0.044518 | 0.025087 | 1.6582 | 15.0536 | 37.7713 | 0.67015
2019-03-26 16:16 | 211 | 0.03581 | 0.029911 | 2.06 | 18.1475 | 37.7276 | 0.64987
2019-03-26 16:16 | 212 | 0.047354 | 0.022169 | 1.553 | 13.4722 | 37.6947 | 0.63797
2019-03-26 16:16 | 213 | 0.060312 | 0.017024 | 1.2129 | 10.6062 | 37.6328 | 0.63968
2019-03-26 16:17 | 214 | 0.043244 | 0.023227 | 1.6824 | 14.8684 | 37.5809 | 0.64297
2019-03-26 16:17 | 215 | 0.049862 | 0.01969 | 1.4535 | 13.0658 | 37.5411 | 0.65148
2019-03-26 16:17 | 216 | 0.059987 | 0.016006 | 1.2067 | 10.8152 | 37.4804 | 0.64877
2019-03-26 16:17 | 217 | 0.047713 | 0.019645 | 1.5088 | 13.454 | 37.4267 | 0.64193
2019-03-26 16:17 | 218 | 0.036333 | 0.025119 | 1.9656 | 17.657 | 37.3855 | 0.64153
2019-03-26 16:17 | 219 | 0.018273 | 0.04868 | 3.9094 | 34.8772 | 37.3547 | 0.6373
2019-03-26 16:17 | 220 | 0.034177 | 0.025325 | 2.0769 | 18.1002 | 37.34 | 0.61861
2019-03-26 16:17 | 221 | 0.070004 | 0.012054 | 1.0092 | 8.6047 | 37.2818 | 0.60236
2019-03-26 16:17 | 222 | 0.054168 | 0.015137 | 1.2981 | 11.3968 | 37.2128 | 0.61735
2019-03-26 16:17 | 223 | 0.024668 | 0.03321 | 2.8345 | 24.9001 | 37.1771 | 0.61423
2019-03-26 16:17 | 224 | 0.042258 | 0.0199 | 1.6461 | 14.7449 | 37.1528 | 0.6231
2019-03-26 16:17 | 225 | 0.035415 | 0.023706 | 1.9505 | 17.902 | 37.0977 | 0.63399
2019-03-26 16:17 | 226 | 0.014309 | 0.060224 | 4.8239 | 44.313 | 37.0845 | 0.63407
2019-03-26 16:17 | 227 | 0.051845 | 0.016608 | 1.3282 | 12.4456 | 37.0556 | 0.64524
2019-03-26 16:17 | 228 | 0.056648 | 0.015224 | 1.2149 | 11.5292 | 36.9878 | 0.65311
2019-03-26 16:17 | 229 | 0.050696 | 0.016638 | 1.3719 | 12.8641 | 36.9461 | 0.65216
2019-03-26 16:17 | 230 | 0.040552 | 0.020773 | 1.7102 | 15.996 | 36.8882 | 0.64868
2019-03-26 16:17 | 231 | 0.013931 | 0.058686 | 4.9852 | 46.8659 | 36.8678 | 0.65289
2019-03-26 16:17 | 232 | 0.03818 | 0.020792 | 1.8249 | 17.0956 | 36.8482 | 0.6527
2019-03-26 16:17 | 233 | 0.056342 | 0.013643 | 1.2311 | 11.5782 | 36.7927 | 0.65234
2019-03-26 16:17 | 234 | 0.034697 | 0.022804 | 1.9983 | 18.4176 | 36.745 | 0.63904
2019-03-26 16:17 | 235 | 0.034963 | 0.022697 | 1.99 | 18.1271 | 36.7187 | 0.63377
2019-03-26 16:17 | 236 | 0.043987 | 0.018053 | 1.5803 | 14.3387 | 36.6727 | 0.63072
2019-03-26 16:17 | 237 | 0.040887 | 0.01941 | 1.6964 | 15.3152 | 36.6334 | 0.62619
2019-03-26 16:17 | 238 | 0.040338 | 0.019641 | 1.7148 | 15.4949 | 36.5903 | 0.62503
2019-03-26 16:17 | 239 | 0.047966 | 0.01652 | 1.4362 | 13.3597 | 36.5509 | 0.64081
2019-03-26 16:17 | 240 | 0.046625 | 0.018 | 1.4756 | 13.746 | 36.4969 | 0.64092
2019-03-26 16:17 | 241 | 0.029324 | 0.028571 | 2.3409 | 21.9353 | 36.4613 | 0.64323
2019-03-26 16:17 | 242 | 0.033744 | 0.023479 | 2.0401 | 18.9503 | 36.433 | 0.63946
2019-03-26 16:17 | 243 | 0.045642 | 0.017877 | 1.5098 | 13.8113 | 36.3917 | 0.63037
2019-03-26 16:17 | 244 | 0.045072 | 0.018089 | 1.5276 | 14.0716 | 36.3452 | 0.63423
2019-03-26 16:17 | 245 | 0.021554 | 0.037892 | 3.2086 | 29.2019 | 36.3066 | 0.62941
2019-03-26 16:17 | 246 | 0.038465 | 0.02064 | 1.7972 | 16.3832 | 36.2926 | 0.63018
2019-03-26 16:17 | 247 | 0.055101 | 0.014408 | 1.2559 | 11.5124 | 36.2294 | 0.63434
2019-03-26 16:17 | 248 | 0.041676 | 0.01847 | 1.657 | 15.0264 | 36.1899 | 0.62624
2019-03-26 16:17 | 249 | 0.030789 | 0.024984 | 2.2403 | 20.3367 | 36.1463 | 0.62615
2019-03-26 16:17 | 250 | 0.019736 | 0.038878 | 3.4916 | 31.6623 | 36.1278 | 0.62489
2019-03-26 16:17 | 251 | 0.028463 | 0.026988 | 2.4238 | 21.6868 | 36.1023 | 0.61728
2019-03-26 16:17 | 252 | 0.022044 | 0.035949 | 3.1319 | 28.1565 | 36.0741 | 0.62068
2019-03-26 16:17 | 253 | 0.014241 | 0.057246 | 4.8346 | 44.1459 | 36.0578 | 0.6287
2019-03-26 16:17 | 254 | 0.043332 | 0.018846 | 1.595 | 14.1776 | 36.0366 | 0.61435
2019-03-26 16:17 | 255 | 0.065322 | 0.012485 | 1.0612 | 9.4867 | 35.9735 | 0.61968
2019-03-26 16:17 | 256 | 0.030803 | 0.026477 | 2.2527 | 20.087 | 35.92 | 0.61874
2019-03-26 16:17 | 257 | 0.010772 | 0.075547 | 6.4277 | 57.6915 | 35.9086 | 0.62146
2019-03-26 16:17 | 258 | 0.036461 | 0.021685 | 1.8993 | 16.8742 | 35.8867 | 0.61524
2019-03-26 16:17 | 259 | 0.054299 | 0.014579 | 1.2755 | 11.5485 | 35.8369 | 0.62708
2019-03-26 16:17 | 260 | 0.048844 | 0.01621 | 1.4137 | 12.7306 | 35.784 | 0.62182
2019-03-26 16:17 | 261 | 0.04882 | 0.01574 | 1.4136 | 12.8382 | 35.7384 | 0.62677
2019-03-26 16:17 | 262 | 0.054291 | 0.013721 | 1.2741 | 11.2216 | 35.6852 | 0.60924
2019-03-26 16:17 | 263 | 0.051456 | 0.014023 | 1.3428 | 12.0052 | 35.632 | 0.61773
2019-03-26 16:17 | 264 | 0.04153 | 0.01737 | 1.6624 | 14.634 | 35.5839 | 0.60776
2019-03-26 16:17 | 265 | 0.045981 | 0.015186 | 1.4954 | 13.3197 | 35.5455 | 0.61245
2019-03-26 16:17 | 266 | 0.049018 | 0.014235 | 1.4003 | 12.597 | 35.4918 | 0.61748
2019-03-26 16:17 | 267 | 0.052728 | 0.013247 | 1.3003 | 11.7965 | 35.4474 | 0.622
2019-03-26 16:17 | 268 | 0.053846 | 0.013419 | 1.2748 | 11.2165 | 35.3873 | 0.60396
2019-03-26 16:17 | 269 | 0.044338 | 0.015776 | 1.542 | 14.0364 | 35.3421 | 0.62234
2019-03-26 16:17 | 270 | 0.048464 | 0.014881 | 1.4122 | 12.8088 | 35.2954 | 0.62077
2019-03-26 16:17 | 271 | 0.058802 | 0.012264 | 1.1601 | 10.6406 | 35.2441 | 0.62569
2019-03-26 16:17 | 272 | 0.040516 | 0.017777 | 1.6833 | 15.985 | 35.1847 | 0.64765
2019-03-26 16:18 | 273 | 0.024863 | 0.029984 | 2.749 | 25.8586 | 35.1624 | 0.64291
2019-03-26 16:18 | 274 | 0.041001 | 0.018731 | 1.668 | 15.4196 | 35.1278 | 0.63223
2019-03-26 16:18 | 275 | 0.025252 | 0.032209 | 2.7042 | 25.9365 | 35.0867 | 0.65494
2019-03-26 16:18 | 276 | 0.034678 | 0.024153 | 1.9736 | 19.4318 | 35.0717 | 0.67386
2019-03-26 16:18 | 277 | 0.047441 | 0.018655 | 1.4457 | 14.1213 | 35.0167 | 0.66993
2019-03-26 16:18 | 278 | 0.028334 | 0.031198 | 2.4267 | 23.7383 | 34.9838 | 0.67261
2019-03-26 16:18 | 279 | 0.042216 | 0.021476 | 1.6389 | 15.5909 | 34.9527 | 0.65818
2019-03-26 16:18 | 280 | 0.052919 | 0.017147 | 1.312 | 12.4865 | 34.8994 | 0.66077
2019-03-26 16:18 | 281 | 0.056096 | 0.015783 | 1.2433 | 11.8303 | 34.8483 | 0.66364
2019-03-26 16:18 | 282 | 0.076963 | 0.011499 | 0.91062 | 8.5566 | 34.7839 | 0.65854
2019-03-26 16:18 | 283 | 0.070519 | 0.012541 | 0.9952 | 9.2934 | 34.7017 | 0.65536
2019-03-26 16:18 | 284 | 0.03518 | 0.025757 | 1.9999 | 18.5059 | 34.6497 | 0.65103
2019-03-26 16:18 | 285 | 0.043583 | 0.020252 | 1.6055 | 14.9652 | 34.6216 | 0.65223
2019-03-26 16:18 | 286 | 0.033444 | 0.026438 | 2.1035 | 19.751 | 34.5658 | 0.66055
2019-03-26 16:18 | 287 | 0.027467 | 0.03137 | 2.5541 | 23.4134 | 34.5533 | 0.64309
2019-03-26 16:18 | 288 | 0.049662 | 0.017359 | 1.4133 | 13.0507 | 34.5047 | 0.64812
2019-03-26 16:18 | 289 | 0.0331 | 0.026766 | 2.1218 | 19.7879 | 34.4624 | 0.65498
2019-03-26 16:18 | 290 | 0.040023 | 0.022118 | 1.7525 | 16.0965 | 34.4342 | 0.64422
2019-03-26 16:18 | 291 | 0.029014 | 0.031298 | 2.4137 | 22.2181 | 34.3863 | 0.64464
2019-03-26 16:18 | 292 | 0.015688 | 0.057877 | 4.4426 | 42.1303 | 34.3759 | 0.66094
2019-03-26 16:18 | 293 | 0.045763 | 0.020815 | 1.519 | 13.948 | 34.3446 | 0.63831
2019-03-26 16:18 | 294 | 0.056652 | 0.017264 | 1.2247 | 11.0162 | 34.2889 | 0.62409
2019-03-26 16:18 | 295 | 0.033381 | 0.02937 | 2.0867 | 19.344 | 34.2394 | 0.64571
2019-03-26 16:18 | 296 | 0.044214 | 0.022706 | 1.5741 | 14.1196 | 34.2146 | 0.62428
2019-03-26 16:18 | 297 | 0.057812 | 0.01774 | 1.1974 | 10.9743 | 34.15 | 0.63444
2019-03-26 16:18 | 298 | 0.048549 | 0.021575 | 1.4232 | 13.0215 | 34.1045 | 0.63218
2019-03-26 16:18 | 299 | 0.044112 | 0.024319 | 1.5727 | 14.62 | 34.0527 | 0.64492
2019-03-26 16:18 | 300 | 0.026798 | 0.040842 | 2.5822 | 23.7198 | 34.0183 | 0.63563
2019-03-26 16:18 | 301 | 0.04771 | 0.023914 | 1.4464 | 13.3288 | 33.9901 | 0.63592
2019-03-26 16:18 | 302 | 0.056166 | 0.020751 | 1.232 | 11.2155 | 33.9256 | 0.62993
2019-03-26 16:18 | 303 | 0.043198 | 0.026959 | 1.5988 | 14.6175 | 33.8831 | 0.63145
2019-03-26 16:18 | 304 | 0.044239 | 0.026301 | 1.5598 | 14.3544 | 33.8358 | 0.63503
2019-03-26 16:18 | 305 | 0.063089 | 0.018413 | 1.0903 | 10.3065 | 33.7903 | 0.65022
2019-03-26 16:18 | 306 | 0.065213 | 0.017797 | 1.0549 | 10.139 | 33.7141 | 0.66119
2019-03-26 16:18 | 307 | 0.038189 | 0.031013 | 1.8003 | 17.1614 | 33.6669 | 0.65537
2019-03-26 16:18 | 308 | 0.037665 | 0.032035 | 1.8268 | 17.0601 | 33.6313 | 0.64257
2019-03-26 16:18 | 309 | 0.057914 | 0.02132 | 1.1985 | 10.9052 | 33.586 | 0.63156
2019-03-26 16:18 | 310 | 0.065207 | 0.019251 | 1.0605 | 9.7744 | 33.5187 | 0.63736
2019-03-26 16:18 | 311 | 0.052402 | 0.023972 | 1.3191 | 12.079 | 33.4612 | 0.63296
2019-03-26 16:18 | 312 | 0.032961 | 0.039607 | 2.1017 | 19.7549 | 33.4153 | 0.65114
2019-03-26 16:18 | 313 | 0.047937 | 0.027224 | 1.4423 | 13.6463 | 33.3871 | 0.65417
2019-03-26 16:18 | 314 | 0.051601 | 0.025301 | 1.3426 | 12.5998 | 33.3216 | 0.65016
2019-03-26 16:18 | 315 | 0.041909 | 0.03177 | 1.6554 | 15.5661 | 33.2872 | 0.65235
2019-03-26 16:18 | 316 | 0.044038 | 0.030304 | 1.5782 | 14.4432 | 33.2354 | 0.63605
2019-03-26 16:18 | 317 | 0.036913 | 0.036736 | 1.8792 | 17.7656 | 33.2006 | 0.65578
2019-03-26 16:18 | 318 | 0.049702 | 0.027716 | 1.3952 | 13.0695 | 33.1571 | 0.64959
2019-03-26 16:18 | 319 | 0.048981 | 0.028088 | 1.4162 | 12.4647 | 33.1038 | 0.61053
2019-03-26 16:18 | 320 | 0.058514 | 0.023902 | 1.1819 | 11.1966 | 33.0574 | 0.65516
2019-03-26 16:18 | 321 | 0.054674 | 0.026002 | 1.2687 | 11.9601 | 32.9901 | 0.6539
2019-03-26 16:18 | 322 | 0.045803 | 0.032086 | 1.5118 | 14.5261 | 32.9497 | 0.66534
2019-03-26 16:18 | 323 | 0.034235 | 0.043488 | 2.0213 | 18.556 | 32.8987 | 0.63526
2019-03-26 16:18 | 324 | 0.042865 | 0.034726 | 1.6158 | 15.0452 | 32.8762 | 0.64491
2019-03-26 16:18 | 325 | 0.051013 | 0.028732 | 1.3588 | 12.9239 | 32.8127 | 0.65929
2019-03-26 16:18 | 326 | 0.049779 | 0.028923 | 1.3936 | 12.7936 | 32.7758 | 0.63685
2019-03-26 16:18 | 327 | 0.0779 | 0.018218 | 0.89169 | 8.2357 | 32.7074 | 0.64156
2019-03-26 16:18 | 328 | 0.05716 | 0.024043 | 1.2193 | 11.1298 | 32.6313 | 0.63617
2019-03-26 16:18 | 329 | 0.047441 | 0.028982 | 1.4651 | 13.1869 | 32.5915 | 0.6256
2019-03-26 16:18 | 330 | 0.045695 | 0.030034 | 1.5071 | 13.9169 | 32.5343 | 0.63593
2019-03-26 16:18 | 331 | 0.034812 | 0.04081 | 1.9786 | 17.9185 | 32.5009 | 0.62378
2019-03-26 16:18 | 332 | 0.065291 | 0.022436 | 1.046 | 9.8712 | 32.4554 | 0.64451
2019-03-26 16:19 | 333 | 0.062187 | 0.024321 | 1.0942 | 10.5123 | 32.3779 | 0.65372
2019-03-26 16:19 | 334 | 0.04916 | 0.031718 | 1.3822 | 13.1467 | 32.3337 | 0.64629
2019-03-26 16:19 | 335 | 0.062862 | 0.025199 | 1.0785 | 10.4616 | 32.2735 | 0.65763
2019-03-26 16:19 | 336 | 0.068611 | 0.023072 | 0.99225 | 9.3709 | 32.2098 | 0.64294
2019-03-26 16:19 | 337 | 0.054686 | 0.029824 | 1.246 | 11.7112 | 32.1411 | 0.64043
2019-03-26 16:19 | 338 | 0.052734 | 0.031349 | 1.2933 | 12.3772 | 32.0979 | 0.6527
2019-03-26 16:19 | 339 | 0.055053 | 0.030359 | 1.2377 | 12.0023 | 32.0347 | 0.66076
2019-03-26 16:19 | 340 | 0.03893 | 0.044262 | 1.7633 | 16.7426 | 31.9912 | 0.65179
2019-03-26 16:19 | 341 | 0.061236 | 0.028932 | 1.1181 | 11.0345 | 31.9486 | 0.67571
2019-03-26 16:19 | 342 | 0.050102 | 0.036325 | 1.367 | 13.4978 | 31.8761 | 0.67627
2019-03-26 16:19 | 343 | 0.040143 | 0.045953 | 1.7084 | 16.5213 | 31.8484 | 0.66321
2019-03-26 16:19 | 344 | 0.050751 | 0.036807 | 1.3527 | 13.3413 | 31.7913 | 0.67709
2019-03-26 16:19 | 345 | 0.037172 | 0.051373 | 1.8459 | 18.1699 | 31.7516 | 0.67541
2019-03-26 16:19 | 346 | 0.05064 | 0.038595 | 1.3552 | 13.3852 | 31.7106 | 0.67782
2019-03-26 16:19 | 347 | 0.066353 | 0.029661 | 1.0298 | 10.0624 | 31.6492 | 0.66768
2019-03-26 16:19 | 348 | 0.074129 | 0.027819 | 0.92145 | 9.1048 | 31.58 | 0.67493
2019-03-26 16:19 | 349 | 0.077158 | 0.027493 | 0.8819 | 8.6647 | 31.5027 | 0.66855
2019-03-26 16:19 | 350 | 0.072668 | 0.030791 | 0.93728 | 11.187 | 31.4279 | 0.81294
2019-03-26 16:19 | 351 | 0.061259 | 0.037419 | 1.1066 | 17.9134 | 31.3585 | 1.0974
2019-03-26 16:19 | 352 | 0.074039 | 0.032113 | 0.92776 | 18.1931 | 31.2987 | 1.347
2019-03-26 16:19 | 353 | 0.11242 | 0.021728 | 0.63215 | 13.3248 | 31.2045 | 1.498
2019-03-26 16:19 | 354 | 0.11029 | 0.022506 | 0.68299 | 15.1343 | 31.0838 | 1.6692
2019-03-26 16:19 | 355 | 0.094563 | 0.026666 | 0.90337 | 18.9244 | 30.9885 | 1.7896
2019-03-26 16:19 | 356 | 0.078523 | 0.032344 | 1.2498 | 23.3474 | 30.8948 | 1.8333
2019-03-26 16:19 | 357 | 0.080786 | 0.031464 | 1.3535 | 23.587 | 30.8265 | 1.9055
2019-03-26 16:19 | 358 | 0.095195 | 0.026722 | 1.2376 | 19.8634 | 30.7303 | 1.8909
2019-03-26 16:19 | 359 | 0.081697 | 0.031381 | 1.5204 | 22.8996 | 30.642 | 1.8708
2019-03-26 16:19 | 360 | 0.095076 | 0.027064 | 1.3764 | 19.5028 | 30.5609 | 1.8542
2019-03-26 16:19 | 361 | 0.10643 | 0.024167 | 1.2791 | 17.2925 | 30.4522 | 1.8405
2019-03-26 16:19 | 362 | 0.10441 | 0.024865 | 1.3487 | 17.5456 | 30.3524 | 1.8319
2019-03-26 16:19 | 363 | 0.070512 | 0.036831 | 2.0521 | 25.512 | 30.2507 | 1.7989
2019-03-26 16:19 | 364 | 0.073251 | 0.035521 | 2.0208 | 24.3166 | 30.2024 | 1.7812
2019-03-26 16:19 | 365 | 0.097836 | 0.026344 | 1.5343 | 17.8225 | 30.0988 | 1.7437
2019-03-26 16:19 | 366 | 0.084887 | 0.030347 | 1.789 | 20.2243 | 30.0156 | 1.7168
2019-03-26 16:19 | 367 | 0.070687 | 0.036533 | 2.1658 | 24.0321 | 29.9299 | 1.6988
2019-03-26 16:19 | 368 | 0.070935 | 0.036259 | 2.1581 | 23.7803 | 29.87 | 1.6869
2019-03-26 16:19 | 369 | 0.093588 | 0.027494 | 1.6403 | 18.0974 | 29.7826 | 1.6937
2019-03-26 16:19 | 370 | 0.094132 | 0.027299 | 1.6264 | 18.0299 | 29.6882 | 1.6972
2019-03-26 16:19 | 371 | 0.079336 | 0.032322 | 1.926 | 21.3433 | 29.5989 | 1.6933
2019-03-26 16:19 | 372 | 0.059373 | 0.043162 | 2.5762 | 29.0763 | 29.5298 | 1.7264
2019-03-26 16:19 | 373 | 0.09194 | 0.027809 | 1.6629 | 19.147 | 29.4674 | 1.7604
2019-03-26 16:19 | 374 | 0.10617 | 0.02442 | 1.4376 | 16.6534 | 29.3493 | 1.7681
2019-03-26 16:19 | 375 | 0.11162 | 0.023589 | 1.3693 | 16.3871 | 29.2577 | 1.8291
2019-03-26 16:19 | 376 | 0.12066 | 0.022321 | 1.2701 | 15.8831 | 29.1261 | 1.9165
2019-03-26 16:19 | 377 | 0.11089 | 0.024892 | 1.3926 | 18.011 | 29.0209 | 1.9973
2019-03-26 16:19 | 378 | 0.10729 | 0.027039 | 1.4492 | 19.6505 | 28.9031 | 2.1083
2019-03-26 16:19 | 379 | 0.10823 | 0.027931 | 1.4546 | 21.1429 | 28.8045 | 2.2882
2019-03-26 16:19 | 380 | 0.12571 | 0.025427 | 1.2764 | 19.1551 | 28.6821 | 2.408
2019-03-26 16:19 | 381 | 0.14824 | 0.022701 | 1.1078 | 17.2672 | 28.5517 | 2.5596
2019-03-26 16:19 | 382 | 0.16174 | 0.021787 | 1.0395 | 16.2203 | 28.3889 | 2.6234
2019-03-26 16:19 | 383 | 0.13494 | 0.027483 | 1.2857 | 20.4215 | 28.2382 | 2.7556
2019-03-26 16:19 | 384 | 0.1274 | 0.030347 | 1.4152 | 22.4299 | 28.1151 | 2.8577
2019-03-26 16:19 | 385 | 0.12789 | 0.031212 | 1.4631 | 23.3277 | 27.9802 | 2.9835
2019-03-26 16:19 | 386 | 0.15095 | 0.027355 | 1.2855 | 20.5255 | 27.8537 | 3.0984
2019-03-26 16:19 | 387 | 0.16009 | 0.026434 | 1.2588 | 19.6203 | 27.6818 | 3.141
2019-03-26 16:19 | 388 | 0.14356 | 0.029877 | 1.4566 | 22.3925 | 27.5398 | 3.2147
2019-03-26 16:19 | 389 | 0.15051 | 0.028861 | 1.4377 | 21.5762 | 27.3902 | 3.2474
2019-03-26 16:19 | 390 | 0.13578 | 0.032113 | 1.6465 | 24.7673 | 27.2434 | 3.3628
2019-03-26 16:19 | 391 | 0.1285 | 0.034109 | 1.7902 | 27.2687 | 27.1166 | 3.5041
2019-03-26 16:19 | 392 | 0.14434 | 0.030337 | 1.646 | 24.9354 | 26.9803 | 3.5992
2019-03-26 16:20 | 393 | 0.16458 | 0.02656 | 1.4892 | 22.2422 | 26.8271 | 3.6607
2019-03-26 16:20 | 394 | 0.15095 | 0.02901 | 1.6697 | 24.6638 | 26.6593 | 3.723
2019-03-26 16:20 | 395 | 0.1436 | 0.030775 | 1.8067 | 26.2474 | 26.5241 | 3.7692
2019-03-26 16:20 | 396 | 0.15235 | 0.029526 | 1.7479 | 24.6123 | 26.368 | 3.7496
2019-03-26 16:20 | 397 | 0.1585 | 0.028701 | 1.7168 | 23.5593 | 26.2195 | 3.7342
2019-03-26 16:20 | 398 | 0.16736 | 0.027629 | 1.6586 | 22.2618 | 26.051 | 3.7256
2019-03-26 16:20 | 399 | 0.1519 | 0.03048 | 1.8561 | 24.8206 | 25.8901 | 3.7703
2019-03-26 16:20 | 400 | 0.16762 | 0.027669 | 1.7055 | 23.0114 | 25.7405 | 3.8573
2019-03-26 16:20 | 401 | 0.1641 | 0.028359 | 1.7661 | 24.0679 | 25.5587 | 3.9496
2019-03-26 16:20 | 402 | 0.16308 | 0.028368 | 1.8011 | 24.7032 | 25.4117 | 4.0285
2019-03-26 16:20 | 403 | 0.17829 | 0.025709 | 1.6718 | 23.0664 | 25.2289 | 4.1126
2019-03-26 16:20 | 404 | 0.17689 | 0.025455 | 1.7053 | 23.9604 | 25.0591 | 4.2384
2019-03-26 16:20 | 405 | 0.16989 | 0.026065 | 1.8004 | 25.4178 | 24.8773 | 4.3183
2019-03-26 16:20 | 406 | 0.15176 | 0.028831 | 2.0463 | 28.8993 | 24.7213 | 4.3857
2019-03-26 16:20 | 407 | 0.1683 | 0.025631 | 1.872 | 26.4465 | 24.5651 | 4.4509
2019-03-26 16:20 | 408 | 0.19174 | 0.02236 | 1.6733 | 23.2925 | 24.3827 | 4.466
2019-03-26 16:20 | 409 | 0.19148 | 0.022447 | 1.7056 | 23.341 | 24.1879 | 4.4694
2019-03-26 16:20 | 410 | 0.17224 | 0.025084 | 1.9303 | 25.8234 | 24.0043 | 4.4479
2019-03-26 16:20 | 411 | 0.18573 | 0.023253 | 1.8202 | 23.8794 | 23.836 | 4.4352
2019-03-26 16:20 | 412 | 0.19255 | 0.022427 | 1.7801 | 23.0938 | 23.6347 | 4.4468
2019-03-26 16:20 | 413 | 0.15852 | 0.027161 | 2.1831 | 28.4566 | 23.4602 | 4.5109
2019-03-26 16:20 | 414 | 0.16427 | 0.026199 | 2.1307 | 27.923 | 23.3082 | 4.587
2019-03-26 16:20 | 415 | 0.193 | 0.022246 | 1.8316 | 23.6576 | 23.1257 | 4.566
2019-03-26 16:20 | 416 | 0.18933 | 0.022692 | 1.8898 | 23.7662 | 22.9303 | 4.4997
2019-03-26 16:20 | 417 | 0.16315 | 0.026567 | 2.2123 | 27.1542 | 22.753 | 4.4302
2019-03-26 16:20 | 418 | 0.15631 | 0.027874 | 2.3263 | 28.2791 | 22.5987 | 4.4204
2019-03-26 16:20 | 419 | 0.18822 | 0.023571 | 1.9404 | 23.4617 | 22.4306 | 4.4158
2019-03-26 16:20 | 420 | 0.20329 | 0.022026 | 1.8003 | 21.832 | 22.2266 | 4.4382
2019-03-26 16:20 | 421 | 0.17561 | 0.02571 | 2.086 | 25.2982 | 22.0349 | 4.4425
2019-03-26 16:20 | 422 | 0.16144 | 0.028071 | 2.2697 | 27.5958 | 21.8723 | 4.4552
2019-03-26 16:20 | 423 | 0.1768 | 0.025586 | 2.0715 | 25.1697 | 21.704 | 4.45
2019-03-26 16:20 | 424 | 0.19842 | 0.022713 | 1.8503 | 22.4292 | 21.5171 | 4.4503
2019-03-26 16:20 | 425 | 0.18895 | 0.023693 | 1.9424 | 23.309 | 21.3148 | 4.4041
2019-03-26 16:20 | 426 | 0.1774 | 0.025045 | 2.0656 | 24.675 | 21.1398 | 4.3774
2019-03-26 16:20 | 427 | 0.19928 | 0.022298 | 1.8406 | 21.895 | 20.9528 | 4.3632
2019-03-26 16:20 | 428 | 0.18462 | 0.024074 | 1.9869 | 23.3604 | 20.7494 | 4.3129
2019-03-26 16:20 | 429 | 0.17549 | 0.025212 | 2.088 | 24.6023 | 20.5828 | 4.3175
2019-03-26 16:20 | 430 | 0.17685 | 0.025196 | 2.0748 | 24.2797 | 20.3957 | 4.294
2019-03-26 16:20 | 431 | 0.17517 | 0.025417 | 2.0882 | 24.4604 | 20.2294 | 4.2848
2019-03-26 16:20 | 432 | 0.17663 | 0.025099 | 2.0676 | 24.1937 | 20.0444 | 4.2732
2019-03-26 16:20 | 433 | 0.18709 | 0.0236 | 1.9481 | 22.7587 | 19.8742 | 4.2579
2019-03-26 16:20 | 434 | 0.18295 | 0.023871 | 1.9853 | 23.1125 | 19.6737 | 4.2284
2019-03-26 16:20 | 435 | 0.17485 | 0.024938 | 2.0779 | 24.1968 | 19.5092 | 4.2309
2019-03-26 16:20 | 436 | 0.18756 | 0.023169 | 1.9344 | 22.671 | 19.3199 | 4.2521
2019-03-26 16:20 | 437 | 0.16424 | 0.026482 | 2.2046 | 25.9083 | 19.1419 | 4.2551
2019-03-26 16:20 | 438 | 0.17043 | 0.025777 | 2.1195 | 24.9094 | 18.9854 | 4.2453
2019-03-26 16:20 | 439 | 0.16095 | 0.02756 | 2.2402 | 26.2327 | 18.8041 | 4.2222
2019-03-26 16:20 | 440 | 0.15405 | 0.028851 | 2.3299 | 27.6448 | 18.6627 | 4.2586
2019-03-26 16:20 | 441 | 0.16722 | 0.026531 | 2.1476 | 24.9628 | 18.4913 | 4.1743
2019-03-26 16:20 | 442 | 0.16632 | 0.02638 | 2.1518 | 24.7898 | 18.3315 | 4.1231
2019-03-26 16:20 | 443 | 0.15536 | 0.027982 | 2.2971 | 26.3046 | 18.1615 | 4.0867
2019-03-26 16:20 | 444 | 0.14382 | 0.030151 | 2.4782 | 28.195 | 18.0207 | 4.0549
2019-03-26 16:20 | 445 | 0.15353 | 0.028331 | 2.3087 | 25.9777 | 17.8689 | 3.9885
2019-03-26 16:20 | 446 | 0.14858 | 0.029217 | 2.3813 | 26.5787 | 17.7165 | 3.949
2019-03-26 16:20 | 447 | 0.15067 | 0.028665 | 2.3354 | 25.818 | 17.5703 | 3.8899
2019-03-26 16:20 | 448 | 0.15377 | 0.027684 | 2.2729 | 25.2792 | 17.4147 | 3.8872
2019-03-26 16:20 | 449 | 0.15594 | 0.02699 | 2.2335 | 24.6379 | 17.2631 | 3.842
2019-03-26 16:20 | 450 | 0.15505 | 0.02682 | 2.2352 | 24.4565 | 17.1039 | 3.7919
2019-03-26 16:20 | 451 | 0.14662 | 0.028095 | 2.3443 | 25.5427 | 16.9547 | 3.7451
2019-03-26 16:20 | 452 | 0.14759 | 0.027776 | 2.3082 | 25.0019 | 16.8086 | 3.69
2019-03-26 16:20 | 453 | 0.14693 | 0.028118 | 2.3017 | 24.6923 | 16.6599 | 3.6281
2019-03-26 16:21 | 454 | 0.13839 | 0.029954 | 2.4139 | 25.304 | 16.5163 | 3.5017
2019-03-26 16:21 | 455 | 0.14104 | 0.029474 | 2.348 | 24.6028 | 16.3798 | 3.4699
2019-03-26 16:21 | 456 | 0.16821 | 0.024885 | 1.9497 | 20.5968 | 16.2286 | 3.4645
2019-03-26 16:21 | 457 | 0.1666 | 0.025296 | 1.9483 | 20.5452 | 16.0511 | 3.4229
2019-03-26 16:21 | 458 | 0.12928 | 0.032947 | 2.4833 | 25.5364 | 15.9046 | 3.3014
2019-03-26 16:21 | 459 | 0.11589 | 0.037176 | 2.7408 | 27.9168 | 15.7868 | 3.2353
2019-03-26 16:21 | 460 | 0.12807 | 0.03423 | 2.4443 | 25.1881 | 15.6655 | 3.2258
2019-03-26 16:21 | 461 | 0.14592 | 0.030358 | 2.1178 | 21.813 | 15.5289 | 3.1829
2019-03-26 16:21 | 462 | 0.14431 | 0.030898 | 2.1153 | 21.9829 | 15.3783 | 3.1724
2019-03-26 16:21 | 463 | 0.14711 | 0.03025 | 2.0405 | 21.4275 | 15.24 | 3.1523
2019-03-26 16:21 | 464 | 0.14255 | 0.031264 | 2.079 | 22.1272 | 15.0866 | 3.1543
2019-03-26 16:21 | 465 | 0.11278 | 0.039105 | 2.5879 | 27.8379 | 14.9607 | 3.1395
2019-03-26 16:21 | 466 | 0.1071 | 0.040614 | 2.6821 | 29.2487 | 14.8546 | 3.1324
2019-03-26 16:21 | 467 | 0.1347 | 0.03188 | 2.1109 | 23.6008 | 14.7377 | 3.1791
2019-03-26 16:21 | 468 | 0.15398 | 0.027644 | 1.8313 | 20.7119 | 14.587 | 3.1893
2019-03-26 16:21 | 469 | 0.1475 | 0.028525 | 1.8931 | 21.6927 | 14.4366 | 3.1998
2019-03-26 16:21 | 470 | 0.13434 | 0.030979 | 2.0642 | 23.8223 | 14.2939 | 3.2002
2019-03-26 16:21 | 471 | 0.14435 | 0.028573 | 1.9157 | 21.9927 | 14.1629 | 3.1747
2019-03-26 16:21 | 472 | 0.13624 | 0.029984 | 2.0241 | 23.3278 | 14.0093 | 3.1783
2019-03-26 16:21 | 473 | 0.12085 | 0.033396 | 2.2823 | 26.3915 | 13.8914 | 3.1893
2019-03-26 16:21 | 474 | 0.15002 | 0.026595 | 1.8336 | 21.2919 | 13.7575 | 3.1943
2019-03-26 16:21 | 475 | 0.151 | 0.026241 | 1.8118 | 21.1387 | 13.5986 | 3.1919
2019-03-26 16:21 | 476 | 0.10481 | 0.037529 | 2.6125 | 30.5022 | 13.467 | 3.1969
2019-03-26 16:21 | 477 | 0.097545 | 0.040156 | 2.8044 | 32.4756 | 13.3797 | 3.1678
2019-03-26 16:21 | 478 | 0.11872 | 0.032978 | 2.3003 | 26.3123 | 13.2641 | 3.1239
2019-03-26 16:21 | 479 | 0.12323 | 0.03184 | 2.2139 | 25.0118 | 13.1456 | 3.0821
2019-03-26 16:21 | 480 | 0.12793 | 0.030604 | 2.1225 | 23.8436 | 13.0188 | 3.0503
2019-03-26 16:21 | 481 | 0.10185 | 0.038409 | 2.6505 | 29.9067 | 12.8965 | 3.0461
2019-03-26 16:21 | 482 | 0.11337 | 0.034361 | 2.3729 | 26.8255 | 12.8063 | 3.0411
2019-03-26 16:21 | 483 | 0.13458 | 0.02876 | 1.9873 | 22.665 | 12.6675 | 3.0502
2019-03-26 16:21 | 484 | 0.11171 | 0.034727 | 2.3866 | 27.3441 | 12.5472 | 3.0546
2019-03-26 16:21 | 485 | 0.11605 | 0.033295 | 2.2895 | 26.209 | 12.4385 | 3.0416
2019-03-26 16:21 | 486 | 0.12136 | 0.031704 | 2.1809 | 24.9769 | 12.3146 | 3.0312
2019-03-26 16:21 | 487 | 0.10865 | 0.03539 | 2.4219 | 27.408 | 12.1994 | 2.978
2019-03-26 16:21 | 488 | 0.12455 | 0.030687 | 2.1021 | 23.6266 | 12.0908 | 2.9427
2019-03-26 16:21 | 489 | 0.13221 | 0.028691 | 1.9655 | 22.0462 | 11.952 | 2.9148
2019-03-26 16:21 | 490 | 0.12728 | 0.029366 | 2.0354 | 22.9323 | 11.8297 | 2.9187
2019-03-26 16:21 | 491 | 0.12587 | 0.02959 | 2.0499 | 23.116 | 11.6978 | 2.9096
2019-03-26 16:21 | 492 | 0.094355 | 0.039344 | 2.7264 | 30.7516 | 11.5847 | 2.9016
2019-03-26 16:21 | 493 | 0.092721 | 0.039508 | 2.7536 | 30.9317 | 11.5018 | 2.868
2019-03-26 16:21 | 494 | 0.10638 | 0.034135 | 2.395 | 27.0402 | 11.3944 | 2.8767
2019-03-26 16:21 | 495 | 0.11905 | 0.02991 | 2.1249 | 23.7985 | 11.2886 | 2.8333
2019-03-26 16:21 | 496 | 0.14077 | 0.024994 | 1.7863 | 19.7475 | 11.1551 | 2.78
2019-03-26 16:21 | 497 | 0.13119 | 0.026357 | 1.9078 | 20.9501 | 11.0157 | 2.7484
2019-03-26 16:21 | 498 | 0.089586 | 0.038397 | 2.7785 | 30.6897 | 10.9011 | 2.7494
2019-03-26 16:21 | 499 | 0.068665 | 0.049911 | 3.6054 | 40.1226 | 10.8311 | 2.755
2019-03-26 16:21 | 500 | 0.08868 | 0.038429 | 2.7703 | 30.7844 | 10.7529 | 2.73
2019-03-26 16:21 | 501 | 0.10307 | 0.032763 | 2.3599 | 26.3833 | 10.6536 | 2.7193
2019-03-26 16:21 | 502 | 0.12627 | 0.026296 | 1.9181 | 21.3511 | 10.5459 | 2.696
2019-03-26 16:21 | 503 | 0.10559 | 0.030758 | 2.2746 | 25.1946 | 10.4117 | 2.6603
2019-03-26 16:21 | 504 | 0.096339 | 0.03284 | 2.4788 | 27.3318 | 10.3327 | 2.6331
2019-03-26 16:21 | 505 | 0.10317 | 0.030325 | 2.3084 | 25.1492 | 10.2153 | 2.5946
2019-03-26 16:21 | 506 | 0.086051 | 0.035382 | 2.7462 | 29.7928 | 10.1308 | 2.5637
2019-03-26 16:21 | 507 | 0.1051 | 0.028173 | 2.2321 | 23.9318 | 10.035 | 2.5152
2019-03-26 16:21 | 508 | 0.11487 | 0.025024 | 2.0215 | 21.7038 | 9.9222 | 2.4932
2019-03-26 16:21 | 509 | 0.11797 | 0.023877 | 1.9537 | 20.7349 | 9.8077 | 2.4461
2019-03-26 16:21 | 510 | 0.099444 | 0.027263 | 2.2935 | 23.9843 | 9.6916 | 2.3851
2019-03-26 16:21 | 511 | 0.094772 | 0.02779 | 2.3855 | 25.2007 | 9.6059 | 2.3883
2019-03-26 16:21 | 512 | 0.094028 | 0.026917 | 2.3813 | 25.1679 | 9.5009 | 2.3665
2019-03-26 16:22 | 513 | 0.085652 | 0.028581 | 2.5865 | 26.9668 | 9.4191 | 2.3098
2019-03-26 16:22 | 514 | 0.093948 | 0.024898 | 2.3321 | 24.6072 | 9.3259 | 2.3118
2019-03-26 16:22 | 515 | 0.09091 | 0.024768 | 2.3819 | 25.2668 | 9.2333 | 2.297
2019-03-26 16:22 | 516 | 0.1003 | 0.021795 | 2.1339 | 22.8066 | 9.1415 | 2.2875
2019-03-26 16:22 | 517 | 0.09809 | 0.021878 | 2.1624 | 23.3134 | 9.0354 | 2.2868
2019-03-26 16:22 | 518 | 0.092859 | 0.022659 | 2.2658 | 24.8335 | 8.9464 | 2.306
2019-03-26 16:22 | 519 | 0.089937 | 0.02316 | 2.3207 | 25.4967 | 8.8493 | 2.2931
2019-03-26 16:22 | 520 | 0.086009 | 0.024165 | 2.4037 | 26.4708 | 8.7665 | 2.2767
2019-03-26 16:22 | 521 | 0.087955 | 0.023342 | 2.3338 | 25.7142 | 8.6753 | 2.2617
2019-03-26 16:22 | 522 | 0.10455 | 0.019417 | 1.9541 | 21.4998 | 8.5868 | 2.2478
2019-03-26 16:22 | 523 | 0.12011 | 0 | 0 | 0 | 8.4675 | 2.2282
2019-03-26 16:22 | 524 | 0.085859 | 0 | 0 | 0 | 8.3586 | 2.2252
2019-03-26 16:22 | 525 | 0.076658 | 0 | 0 | 0 | 8.2905 | 2.2345
2019-03-26 16:22 | 526 | 0.082085 | 0 | 0 | 0 | 8.2011 | 2.1971
2019-03-26 16:22 | 527 | 0.079018 | 0 | 0 | 0 | 8.1274 | 2.2049
2019-03-26 16:22 | 528 | 0.094266 | 0 | 0 | 0 | 8.039 | 2.2101
2019-03-26 16:22 | 529 | 0.095389 | 0 | 0 | 0 | 7.9421 | 2.2123
2019-03-26 16:22 | 530 | 0.094595 | 0 | 0 | 0 | 7.8496 | 2.1589
2019-03-26 16:22 | 531 | 0.077066 | 0 | 0 | 0 | 7.7572 | 2.1136
2019-03-26 16:22 | 532 | 0.058103 | 0 | 0 | 0 | 7.6955 | 2.0852
2019-03-26 16:22 | 533 | 0.068448 | 0 | 0 | 0 | 7.6337 | 2.049
2019-03-26 16:22 | 534 | 0.075045 | 0 | 0 | 0 | 7.5588 | 1.9644
2019-03-26 16:22 | 535 | 0.077544 | 0 | 0 | 0 | 7.4849 | 1.9317
2019-03-26 16:22 | 536 | 0.07712 | 0 | 0 | 0 | 7.405 | 1.8839
2019-03-26 16:22 | 537 | 0.060041 | 0 | 0 | 0 | 7.3345 | 1.8449
2019-03-26 16:22 | 538 | 0.059089 | 0 | 0 | 0 | 7.2813 | 1.7972
2019-03-26 16:22 | 539 | 0.061308 | 0 | 0 | 0 | 7.2152 | 1.7575
2019-03-26 16:22 | 540 | 0.056957 | 0 | 0 | 0 | 7.1602 | 1.7601
2019-03-26 16:22 | 541 | 0.053395 | 0 | 0 | 0 | 7.1015 | 1.7492
2019-03-26 16:22 | 542 | 0.042282 | 0 | 0 | 0 | 7.0552 | 1.6694
2019-03-26 16:22 | 543 | 0.033211 | 0 | 0 | 0 | 7.0157 | 1.7265
2019-03-26 16:22 | 544 | 0.057973 | 0 | 0 | 0 | 6.9808 | 1.6566
2019-03-26 16:22 | 545 | 0.059464 | 0 | 0 | 0 | 6.9051 | 1.4893
2019-03-26 16:22 | 546 | 0.042289 | 0 | 0 | 0 | 6.8662 | 1.4427
2019-03-26 16:22 | 547 | 0.061723 | 0 | 0 | 0 | 6.8128 | 1.4125
2019-03-26 16:22 | 548 | 0.052076 | 0 | 0 | 0 | 6.7495 | 1.4145
2019-03-26 16:22 | 549 | 0.030449 | 0 | 0 | 0 | 6.7118 | 1.3892
2019-03-26 16:22 | 550 | 0.026275 | 0 | 0 | 0 | 6.6835 | 1.3618
2019-03-26 16:22 | 551 | 0.057774 | 0 | 0 | 0 | 6.6504 | 1.3221
2019-03-26 16:22 | 552 | 0.067703 | 0 | 0 | 0 | 6.5732 | 1.2629
2019-03-26 16:22 | 553 | 0.048375 | 0 | 0 | 0 | 6.5225 | 1.2007
2019-03-26 16:22 | 554 | 0.047487 | 0 | 0 | 0 | 6.473 | 1.2091
2019-03-26 16:22 | 555 | 0.035275 | 0 | 0 | 0 | 6.4297 | 1.1801
2019-03-26 16:22 | 556 | 0.032651 | 0 | 0 | 0 | 6.4004 | 1.1894
2019-03-26 16:22 | 557 | 0.026612 | 0 | 0 | 0 | 6.3648 | 1.1686
2019-03-26 16:22 | 558 | 0.021159 | 0 | 0 | 0 | 6.346 | 1.1369
2019-03-26 16:22 | 559 | 0.054315 | 0 | 0 | 0 | 6.3134 | 1.1246
2019-03-26 16:22 | 560 | 0.064961 | 0 | 0 | 0 | 6.2429 | 1.0714
2019-03-26 16:22 | 561 | 0.041743 | 0 | 0 | 0 | 6.1923 | 1.046
2019-03-26 16:22 | 562 | 0.034689 | 0 | 0 | 0 | 6.1564 | 1.006
2019-03-26 16:22 | 563 | 0.029585 | 0 | 0 | 0 | 6.1221 | 0.98784
2019-03-26 16:22 | 564 | 0.016327 | 0 | 0 | 0 | 6.0983 | 0.9916
2019-03-26 16:22 | 565 | 0.033148 | 0 | 0 | 0 | 6.0824 | 0.95676
2019-03-26 16:22 | 566 | 0.042446 | 0 | 0 | 0 | 6.034 | 0.94881
2019-03-26 16:22 | 567 | 0.017487 | 0 | 0 | 0 | 6.0056 | 0.95012
2019-03-26 16:22 | 568 | 0.016032 | 0 | 0 | 0 | 5.9938 | 0.94825
2019-03-26 16:22 | 569 | 0.023938 | 0 | 0 | 0 | 5.9708 | 0.91545
2019-03-26 16:22 | 570 | 0.023161 | 0 | 0 | 0 | 5.9478 | 0.88437
2019-03-26 16:22 | 571 | 0.024255 | 0 | 0 | 0 | 5.9244 | 0.83772
2019-03-26 16:22 | 572 | 0.01753 | 0 | 0 | 0 | 5.9007 | 0.81614
2019-03-26 16:22 | 573 | 0.029319 | 0 | 0 | 0 | 5.8851 | 0.79488
2019-03-26 16:23 | 574 | 0.034917 | 0 | 0 | 0 | 5.8436 | 0.79386
2019-03-26 16:23 | 575 | 0.020439 | 0 | 0 | 0 | 5.8199 | 0.78458
2019-03-26 16:23 | 576 | 0.024455 | 0 | 0 | 0 | 5.7982 | 0.76291
2019-03-26 16:23 | 577 | 0.046554 | 0 | 0 | 0 | 5.7668 | 0.75707
2019-03-26 16:23 | 578 | 0.040452 | 0 | 0 | 0 | 5.7122 | 0.74347
2019-03-26 16:23 | 579 | 0.01519 | 0 | 0 | 0 | 5.6907 | 0.72382
2019-03-26 16:23 | 580 | 0.020095 | 0 | 0 | 0 | 5.6747 | 0.71631
2019-03-26 16:23 | 581 | 0.027507 | 0 | 0 | 0 | 5.6492 | 0.70928
2019-03-26 16:23 | 582 | 0.030224 | 0 | 0 | 0 | 5.6211 | 0.70011
2019-03-26 16:23 | 583 | 0.02266 | 0 | 0 | 0 | 5.5916 | 0.69915
2019-03-26 16:23 | 584 | 0.0095272 | 0 | 0 | 0 | 5.5768 | 0.70073
2019-03-26 16:23 | 585 | 0.018942 | 0 | 0 | 0 | 5.5667 | 0.69641
2019-03-26 16:23 | 586 | 0.038306 | 0 | 0 | 0 | 5.5365 | 0.69062
2019-03-26 16:23 | 587 | 0.027918 | 0 | 0 | 0 | 5.4969 | 0.6807
2019-03-26 16:23 | 588 | 0.033541 | 0 | 0 | 0 | 5.4777 | 0.66913
2019-03-26 16:23 | 589 | 0.03686 | 0 | 0 | 0 | 5.431 | 0.66488
2019-03-26 16:23 | 590 | 0.0043798 | 0 | 0 | 0 | 5.412 | 0.64037
2019-03-26 16:23 | 591 | 0.0070694 | 0 | 0 | 0 | 5.4137 | 0.6249
2019-03-26 16:23 | 592 | 0.031051 | 0 | 0 | 0 | 5.3919 | 0.62166
2019-03-26 16:23 | 593 | 0.040178 | 0 | 0 | 0 | 5.3553 | 0.61711
2019-03-26 16:23 | 594 | 0.029567 | 0 | 0 | 0 | 5.3172 | 0.60491
2019-03-26 16:23 | 595 | 0.012738 | 0 | 0 | 0 | 5.2983 | 0.5903
2019-03-26 16:23 | 596 | -0.0017232 | 0 | 0 | 0 | 5.2907 | 0.57752
2019-03-26 16:23 | 597 | 0.0064677 | 0 | 0 | 0 | 5.2959 | 0.57242
2019-03-26 16:23 | 598 | 0.012003 | 0 | 0 | 0 | 5.2777 | 0.56958
2019-03-26 16:23 | 599 | 0.016754 | 0 | 0 | 0 | 5.2718 | 0.57415
2019-03-26 16:23 | 600 | 0.035787 | 0 | 0 | 0 | 5.2415 | 0.56737
2019-03-26 16:23 | 601 | 0.02417 | 0 | 0 | 0 | 5.2078 | 0.54179
2019-03-26 16:23 | 602 | 0.0062197 | 0 | 0 | 0 | 5.1951 | 0.54178
2019-03-26 16:23 | 603 | 0.0051416 | 0 | 0 | 0 | 5.1906 | 0.51936
2019-03-26 16:23 | 604 | 0.029505 | 0 | 0 | 0 | 5.1786 | 0.5062
2019-03-26 16:23 | 605 | 0.028267 | 0 | 0 | 0 | 5.1375 | 0.51163
2019-03-26 16:23 | 606 | 0.022005 | 0 | 0 | 0 | 5.1241 | 0.51588
2019-03-26 16:23 | 607 | 0.016278 | 0 | 0 | 0 | 5.0938 | 0.51073
2019-03-26 16:23 | 608 | 0.00041367 | 0 | 0 | 0 | 5.0938 | 0.50055
2019-03-26 16:23 | 609 | -0.0046003 | 0 | 0 | 0 | 5.0905 | 0.50536
2019-03-26 16:23 | 610 | -0.010494 | 0 | 0 | 0 | 5.1024 | 0.48969
2019-03-26 16:23 | 611 | 0.0010897 | 0 | 0 | 0 | 5.1072 | 0.49181
2019-03-26 16:23 | 612 | 0.0063311 | 0 | 0 | 0 | 5.1013 | 0.50094
2019-03-26 16:23 | 613 | 0.01231 | 0 | 0 | 0 | 5.0948 | 0.49371
2019-03-26 16:23 | 614 | 0.0062728 | 0 | 0 | 0 | 5.0796 | 0.48092
2019-03-26 16:23 | 615 | 0.005122 | 0 | 0 | 0 | 5.0815 | 0.4807
2019-03-26 16:23 | 616 | 0.000642 | 0 | 0 | 0 | 5.0697 | 0.48763
2019-03-26 16:23 | 617 | 0.0059727 | 0 | 0 | 0 | 5.0769 | 0.46325
2019-03-26 16:23 | 618 | 0.040427 | 0 | 0 | 0 | 5.0505 | 0.48641
2019-03-26 16:23 | 619 | 0.06133 | 0 | 0 | 0 | 4.9998 | 0.45976
2019-03-26 16:23 | 620 | 0.044939 | 0 | 0 | 0 | 4.9383 | 0.46327
2019-03-26 16:23 | 621 | 0.011444 | 0 | 0 | 0 | 4.9146 | 0.47669
2019-03-26 16:23 | 622 | 0.0080707 | 0 | 0 | 0 | 4.9081 | 0.47328
2019-03-26 16:23 | 623 | 0.01024 | 0 | 0 | 0 | 4.8956 | 0.47221
2019-03-26 16:23 | 624 | 0.02656 | 0 | 0 | 0 | 4.8835 | 0.47217
2019-03-26 16:23 | 625 | 0.052222 | 0 | 0 | 0 | 4.8406 | 0.45842
2019-03-26 16:23 | 626 | 0.051297 | 0 | 0 | 0 | 4.7864 | 0.45963
2019-03-26 16:23 | 627 | 0.02099 | 0 | 0 | 0 | 4.7458 | 0.45433
2019-03-26 16:23 | 628 | 0.0069101 | 0 | 0 | 0 | 4.7408 | 0.46071
2019-03-26 16:23 | 629 | 0.0078938 | 0 | 0 | 0 | 4.7275 | 0.44453
2019-03-26 16:23 | 630 | 0.015034 | 0 | 0 | 0 | 4.723 | 0.44121
2019-03-26 16:23 | 631 | 0.017907 | 0 | 0 | 0 | 4.6985 | 0.44963
2019-03-26 16:23 | 632 | 0.012588 | 0 | 0 | 0 | 4.6893 | 0.44445
2019-03-26 16:24 | 633 | 0.012535 | 0 | 0 | 0 | 4.6728 | 0.40894
2019-03-26 16:24 | 634 | -0.0060922 | 0 | 0 | 0 | 4.6685 | 0.41971
2019-03-26 16:24 | 635 | -0.012155 | 0 | 0 | 0 | 4.682 | 0.42162
2019-03-26 16:24 | 636 | -0.0073425 | 0 | 0 | 0 | 4.6894 | 0.44457
2019-03-26 16:24 | 637 | 0.0086407 | 0 | 0 | 0 | 4.6943 | 0.3929
2019-03-26 16:24 | 638 | -0.0032071 | 0 | 0 | 0 | 4.6786 | 0.43573
2019-03-26 16:24 | 639 | -0.013422 | 0 | 0 | 0 | 4.6998 | 0.38984
2019-03-26 16:24 | 640 | 0.019462 | 0 | 0 | 0 | 4.6952 | 0.51596
2019-03-26 16:24 | 641 | 0.030406 | 0 | 0 | 0 | 4.6657 | 0.34608
2019-03-26 16:24 | 642 | 0.024133 | 0 | 0 | 0 | 4.6393 | 0.34378
2019-03-26 16:24 | 643 | 0.021308 | 0 | 0 | 0 | 4.6171 | 0.4018
2019-03-26 16:24 | 644 | 0.022737 | 0 | 0 | 0 | 4.5959 | 0.3928
2019-03-26 16:24 | 645 | 0.0092497 | 0 | 0 | 0 | 4.5748 | 0.40115
2019-03-26 16:24 | 646 | 0.0087977 | 0 | 0 | 0 | 4.5741 | 0.40469
2019-03-26 16:24 | 647 | 0.022058 | 0 | 0 | 0 | 4.5537 | 0.40351
2019-03-26 16:24 | 648 | 0.023849 | 0 | 0 | 0 | 4.5327 | 0.37245
2019-03-26 16:24 | 649 | 0.02167 | 0 | 0 | 0 | 4.5079 | 0.37592
2019-03-26 16:24 | 650 | 0.0008325 | 0 | 0 | 0 | 4.4934 | 0.36852
2019-03-26 16:24 | 651 | 0.0093826 | 0 | 0 | 0 | 4.4994 | 0.3906
2019-03-26 16:24 | 652 | 0.013429 | 0 | 0 | 0 | 4.4752 | 0.37376
2019-03-26 16:24 | 653 | 0.0069338 | 0 | 0 | 0 | 4.4749 | 0.37744
2019-03-26 16:24 | 654 | 0.011831 | 0 | 0 | 0 | 4.4586 | 0.38687
2019-03-26 16:24 | 655 | 0.021373 | 0 | 0 | 0 | 4.4497 | 0.39038
2019-03-26 16:24 | 656 | 0.037191 | 0 | 0 | 0 | 4.4152 | 0.37136
2019-03-26 16:24 | 657 | 0.014265 | 0 | 0 | 0 | 4.3846 | 0.39181
2019-03-26 16:24 | 658 | 0.0058271 | 0 | 0 | 0 | 4.3838 | 0.40429
2019-03-26 16:24 | 659 | 0.0077564 | 0 | 0 | 0 | 4.3697 | 0.37576
2019-03-26 16:24 | 660 | 0.01662 | 0 | 0 | 0 | 4.3663 | 0.37722
2019-03-26 16:24 | 661 | 0.021198 | 0 | 0 | 0 | 4.3375 | 0.36545
2019-03-26 16:24 | 662 | 0.01984 | 0 | 0 | 0 | 4.3255 | 0.36245
2019-03-26 16:24 | 663 | 0.020516 | 0 | 0 | 0 | 4.2976 | 0.35852
2019-03-26 16:24 | 664 | 0.016023 | 0 | 0 | 0 | 4.2856 | 0.35959
2019-03-26 16:24 | 665 | 0.017397 | 0 | 0 | 0 | 4.2644 | 0.35202
2019-03-26 16:24 | 666 | 0.011315 | 0 | 0 | 0 | 4.2521 | 0.35684
2019-03-26 16:24 | 667 | 0.021483 | 0 | 0 | 0 | 4.2381 | 0.36275
2019-03-26 16:24 | 668 | 0.026052 | 0 | 0 | 0 | 4.2105 | 0.35146
2019-03-26 16:24 | 669 | 0.013901 | 0 | 0 | 0 | 4.1901 | 0.36249
2019-03-26 16:24 | 670 | 0.011152 | 0 | 0 | 0 | 4.1809 | 0.32703
2019-03-26 16:24 | 671 | 0.0075324 | 0 | 0 | 0 | 4.1675 | 0.34868
2019-03-26 16:24 | 672 | 0.011401 | 0 | 0 | 0 | 4.1637 | 0.32993
2019-03-26 16:24 | 673 | 0.021345 | 0 | 0 | 0 | 4.1434 | 0.33488
2019-03-26 16:24 | 674 | 0.015845 | 0 | 0 | 0 | 4.1248 | 0.3232
2019-03-26 16:24 | 675 | 0.010138 | 0 | 0 | 0 | 4.1121 | 0.32206
2019-03-26 16:24 | 676 | 0.0096601 | 0 | 0 | 0 | 4.1032 | 0.31762
2019-03-26 16:24 | 677 | 0.010822 | 0 | 0 | 0 | 4.0921 | 0.30812
2019-03-26 16:24 | 678 | 0.013159 | 0 | 0 | 0 | 4.0809 | 0.31817
2019-03-26 16:24 | 679 | 0.025126 | 0 | 0 | 0 | 4.0641 | 0.32516
2019-03-26 16:24 | 680 | 0.01124 | 0 | 0 | 0 | 4.0367 | 0.3342
2019-03-26 16:24 | 681 | 0.0052565 | 0 | 0 | 0 | 4.04 | 0.32958
2019-03-26 16:24 | 682 | 0.011376 | 0 | 0 | 0 | 4.0235 | 0.33809
2019-03-26 16:24 | 683 | -0.0040439 | 0 | 0 | 0 | 4.0215 | 0.3357
2019-03-26 16:24 | 684 | 0.010735 | 0 | 0 | 0 | 4.025 | 0.33737
2019-03-26 16:24 | 685 | 0.012677 | 0 | 0 | 0 | 4.003 | 0.31988
2019-03-26 16:24 | 686 | -0.0058151 | 0 | 0 | 0 | 4.0044 | 0.3375
2019-03-26 16:24 | 687 | 0.0010401 | 0 | 0 | 0 | 4.0089 | 0.31386
2019-03-26 16:24 | 688 | 0.0052148 | 0 | 0 | 0 | 4.0027 | 0.31854
2019-03-26 16:24 | 689 | -0.0065937 | 0 | 0 | 0 | 4.0021 | 0.32357
2019-03-26 16:24 | 690 | -0.0056798 | 0 | 0 | 0 | 4.0125 | 0.31834
2019-03-26 16:24 | 691 | 0.017234 | 0 | 0 | 0 | 4.0083 | 0.30919
2019-03-26 16:24 | 692 | 0.012649 | 0 | 0 | 0 | 3.9841 | 0.31275
2019-03-26 16:24 | 693 | 0.017627 | 0 | 0 | 0 | 3.9814 | 0.30076
2019-03-26 16:25 | 694 | 0.025456 | 0 | 0 | 0 | 3.9489 | 0.30262
2019-03-26 16:25 | 695 | -0.00031908 | 0 | 0 | 0 | 3.9382 | 0.29771
2019-03-26 16:25 | 696 | -0.002551 | 0 | 0 | 0 | 3.9438 | 0.29341
2019-03-26 16:25 | 697 | 0.020961 | 0 | 0 | 0 | 3.9368 | 0.30644
2019-03-26 16:25 | 698 | 0.021213 | 0 | 0 | 0 | 3.9071 | 0.30102
2019-03-26 16:25 | 699 | 0.017212 | 0 | 0 | 0 | 3.8961 | 0.29618
2019-03-26 16:25 | 700 | 0.017011 | 0 | 0 | 0 | 3.8726 | 0.30073
2019-03-26 16:25 | 701 | -0.012446 | 0 | 0 | 0 | 3.8679 | 0.2757
2019-03-26 16:25 | 702 | 0.008164 | 0 | 0 | 0 | 3.8861 | 0.27667
2019-03-26 16:25 | 703 | 0.016177 | 0 | 0 | 0 | 3.8543 | 0.29085

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062505 | 0.062705 | 0 | 0.2095 | 2019-03-26 11:46 | Group A | 0.1 | 0.22 | Do Nothing | 0.2095 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025211 | 0.02534 | 0 | 0.0261 | 2019-03-26 11:46 | Group A | -0.01 | 0.1 | Do Nothing | 0.00043979 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002731 | 0.002746 | 0 | 0.0011 | 2019-03-26 11:46 | Group A | -0.01 | 0.1 | Do Nothing | 1.2692e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.4835 | 181.973 | 0 | 100 | 2019-03-26 11:46 | None | 0 | 1000 | Do Nothing | 99.5783 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.3052 | 207.2338 | 0 | 100 | 2019-03-26 11:46 | None | 0 | 1000 | Do Nothing | 99.4819 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -15.8259 | 1447 | 0 |  | 2019-03-26 11:39 | None | -10 | 1000 | Do Nothing | 128.349 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2612 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018089 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017001 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -152.7595 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.21943 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4033 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0034835 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0048197 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0046413 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 4.8346e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011965 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.3325 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3483 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5475 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.79 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 21.0884 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 16.6555 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 65.352 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 105.6915 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 62.7336 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0041369 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0511 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025542 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.3002 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN