## Cone Summary
Date: | 2019-03-20
Time: | 17:41
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B2S-Test1
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0411 |  | Start Test | 0 | 2019-03-20 17:49
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 7 | 2019-03-20 17:49
Sample Area | 1 | m^2 | Wetting | 156 | 2019-03-20 17:52
Cone Heat Flux | 75 | kW/m^2 | Dripping | 181 | 2019-03-20 17:52
Cone Orientation |  |  | Bottom Ignition | 256 | 2019-03-20 17:53
Sample Grid |  |  | Flame Out | 524 | 2019-03-20 17:58
Sample Thickness | 106 | mm
Sample Initial Mass | 48 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 24.0253 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.5291 | kW/m2 | Average Heat Release Rate | 2.323 | kW/m2
Total Heat Release | 1.0826 | MJ/m2 | Average Mass Loss Rate | 0.098704 | g/m2s
Peak Mass Loss Rate | 0.19402 | g/m2s | Effective Heat of Combustion | 23.5952 | MJ/kg
Total Mass Lost | 44.3037 | g | Average Specific Extinction Area | 63.9251 | m2/kg
Time to Ignition | 7 | s | Peak CO Yield | 0.05084 | kg/kg
Time to Peak Heat Release | 382 | s | Peak CO2 Yield | 6.9664 | kg/kg
Time to Flame Out | 524 | s | Average CO Yield | 0.023798 | kg/kg
Average Heat Release Rate | 2.09 | kW/m2 | Average CO2 Yield | 2.0551 | kg/kg
Average Mass Loss Rate | 0.085472 | g/m2s
Effective Heat of Combustion | 24.4363 | MJ/kg
Dripping | 181 | s | Time to dripping
HRRD | 0.99356 | kW | HRR at dripping time
THRD | 0.20648 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.1867 | kW | AHRR from Ignition to Dripping
Wetting | 156 | s | Time to wetting
HRRW | 0.94838 | kW | HRR at wetting time
THRW | 0.18291 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.1884 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-20 17:41 | 1553118058.0738 | -482 | 0.20951 | 0.000451 | 0 | 109.651 | 110.8158 | 126.8437 | 21.0359 | 0.002014 | 0.001847 | -151.8798 | 0.23903 | 8.4019 | 0.003709 | 0.005621 | 0.005444 | 6.3e-05 | 0.001248 | 0.26377 | 1.3498 | 2.548 | 100.58 | 20.67 | 19.025 | 70.0702 | 110.9377 | 66.039 | 0.004615 | 1.0409 | 0.025134 | 24.1472 | -0.094519 | 0.1422 | -7e-06 | -0.000548 | -0.000548 | -10.6707
2019-03-20 17:41 | 1553118059.0738 | -481 | 0.20951 | 0.000451 | 0 | 109.651 | 110.8158 | 126.8437 | 21.0359 | 0.002014 | 0.001847 | -151.8798 | 0.23903 | 8.4019 | 0.003709 | 0.005621 | 0.005444 | 6.3e-05 | 0.001248 | 0.26377 | 1.3498 | 2.548 | 100.58 | 20.67 | 19.025 | 70.0702 | 110.9377 | 66.039 | 0.004615 | 1.0409 | 0.025134 | 24.1472 | -0.094519 | 0.1422 | -7e-06 | -0.000548 | -0.000548 | -10.6721
2019-03-20 17:41 | 1553118060.0729 | -480 | 0.20946 | 0.000453 | 1e-06 | 109.5119 | 111.2781 | 126.8437 | 21.0366 | 0.002044 | 0.001866 | -151.8798 | 0.23907 | 8.402 | 0.003599 | 0.005083 | 0.004824 | 5.9e-05 | 0.001244 | 0.2756 | 1.3478 | 2.548 | 100.58 | 20.67 | 19.024 | 70.7826 | 108.2638 | 66.5024 | 0.004614 | 1.0394 | 0.025116 | 24.1637 | -0.14311 | 0.14222 | -9e-06 | -0.000763 | -0.000763 | -7.6589
2019-03-20 17:41 | 1553118061.0743 | -479 | 0.20942 | 0.000445 | 0 | 109.562 | 111.061 | 126.8437 | 21.0374 | 0.00198 | 0.001804 | -151.8798 | 0.2378 | 8.4048 | 0.003646 | 0.005166 | 0.004911 | 6.2e-05 | 0.001247 | 0.27721 | 1.3095 | 2.5482 | 100.58 | 20.67 | 19.023 | 69.2499 | 109.4067 | 65.0126 | 0.004614 | 1.044 | 0.025172 | 24.1106 | -0.12155 | 0.14147 | 1e-06 | 6.7e-05 | 6.7e-05 | 86.8283
2019-03-20 17:41 | 1553118062.074 | -478 | 0.20949 | 0.000442 | 1e-06 | 109.5375 | 110.9945 | 126.8437 | 21.0369 | 0.002008 | 0.001846 | -151.8798 | 0.23848 | 8.4023 | 0.003648 | 0.004636 | 0.00445 | 6.1e-05 | 0.001252 | 0.27201 | 1.2962 | 2.548 | 100.58 | 20.67 | 19.021 | 69.9161 | 109.4582 | 66.0198 | 0.004614 | 1.0409 | 0.025134 | 24.1465 | -0.11819 | 0.14187 | 4e-06 | 0.000232 | 0.000232 | 25.1341
2019-03-20 17:41 | 1553118063.0745 | -477 | 0.20944 | 0.000449 | 1e-06 | 109.552 | 111.2139 | 126.8437 | 21.0356 | 0.001979 | 0.001833 | -151.8798 | 0.23845 | 8.402 | 0.003542 | 0.005004 | 0.004809 | 6.2e-05 | 0.001255 | 0.29695 | 1.3005 | 2.548 | 100.58 | 20.67 | 19.016 | 69.213 | 106.8884 | 65.7085 | 0.004613 | 1.0419 | 0.025146 | 24.1354 | -0.13467 | 0.14185 | -5e-06 | -0.000432 | -0.000432 | -13.4874
2019-03-20 17:41 | 1553118064.0731 | -476 | 0.20947 | 0.00045 | 1e-06 | 109.5048 | 111.0277 | 126.8437 | 21.0366 | 0.001941 | 0.001809 | -151.8798 | 0.23853 | 8.4026 | 0.003655 | 0.005137 | 0.004947 | 6.3e-05 | 0.001258 | 0.27848 | 1.2841 | 2.548 | 100.58 | 20.67 | 19.013 | 68.3131 | 109.6424 | 65.129 | 0.004612 | 1.0437 | 0.025167 | 24.1148 | -0.12354 | 0.14191 | -6e-06 | -0.000475 | -0.000475 | -12.2872
2019-03-20 17:41 | 1553118065.0732 | -475 | 0.20944 | 0.000445 | 0 | 109.5126 | 110.9986 | 126.8437 | 21.0361 | 0.001964 | 0.001865 | -151.8798 | 0.23805 | 8.3987 | 0.003681 | 0.004626 | 0.004396 | 6.4e-05 | 0.001265 | 0.28064 | 1.2856 | 2.5481 | 100.58 | 20.67 | 19.01 | 68.8616 | 110.2617 | 66.4887 | 0.004611 | 1.0395 | 0.025117 | 24.1632 | -0.12056 | 0.14161 | 1e-06 | 3.8e-05 | 3.8e-05 | 154.7122
2019-03-20 17:41 | 1553118066.0747 | -474 | 0.20943 | 0.000452 | 0 | 109.5261 | 110.8618 | 126.8437 | 21.0373 | 0.001967 | 0.001863 | -151.8798 | 0.23889 | 8.4046 | 0.003611 | 0.00552 | 0.005304 | 6.3e-05 | 0.00126 | 0.28694 | 1.3067 | 2.5481 | 100.58 | 20.66 | 19.005 | 68.9339 | 108.5631 | 66.4305 | 0.004607 | 1.0397 | 0.025119 | 24.1611 | -0.10842 | 0.14211 | -8e-06 | -0.000614 | -0.000614 | -9.5166
2019-03-20 17:41 | 1553118067.0733 | -473 | 0.20949 | 0.000446 | 0 | 109.4068 | 110.7601 | 126.8437 | 21.0365 | 0.001966 | 0.001845 | -151.8798 | 0.23931 | 8.4028 | 0.003643 | 0.00525 | 0.005021 | 6.3e-05 | 0.001256 | 0.29231 | 1.3022 | 2.5479 | 100.59 | 20.66 | 19.001 | 68.9228 | 109.3419 | 66.0038 | 0.004606 | 1.041 | 0.025135 | 24.1459 | -0.10996 | 0.14236 | 0 | -7e-06 | -7e-06 | -895.5946
2019-03-20 17:41 | 1553118068.0743 | -472 | 0.20946 | 0.00045 | 0 | 109.4543 | 110.975 | 126.8437 | 21.0353 | 0.001975 | 0.001843 | -151.8798 | 0.23833 | 8.4026 | 0.003681 | 0.005005 | 0.004814 | 6.3e-05 | 0.001256 | 0.30359 | 1.3136 | 2.548 | 100.59 | 20.66 | 18.994 | 69.1332 | 110.2626 | 65.948 | 0.004604 | 1.0411 | 0.025137 | 24.144 | -0.12341 | 0.14178 | -5e-06 | -0.000395 | -0.000395 | -14.7466
2019-03-20 17:41 | 1553118069.0738 | -471 | 0.20948 | 0.000453 | 1e-06 | 109.4246 | 110.9534 | 126.8709 | 21.0366 | 0.001917 | 0.00187 | -151.8745 | 0.23852 | 8.4019 | 0.003612 | 0.005276 | 0.00506 | 6.5e-05 | 0.001251 | 0.2995 | 1.3123 | 2.5482 | 100.59 | 20.66 | 18.989 | 67.7288 | 108.5852 | 66.6023 | 0.004603 | 1.0391 | 0.025115 | 24.1698 | -0.1241 | 0.1419 | -1e-05 | -0.000791 | -0.000791 | -7.3759
2019-03-20 17:41 | 1553118070.0725 | -470 | 0.20949 | 0.00045 | 1e-06 | 109.4 | 110.8326 | 126.7508 | 21.0367 | 0.001991 | 0.001843 | -151.8824 | 0.23904 | 8.4008 | 0.00366 | 0.005183 | 0.004898 | 6.6e-05 | 0.00125 | 0.28799 | 1.3269 | 2.5481 | 100.59 | 20.66 | 18.983 | 69.5018 | 109.7499 | 65.9617 | 0.004601 | 1.0411 | 0.025127 | 24.1356 | -0.11637 | 0.14221 | -5e-06 | -0.000444 | -0.000444 | -13.1526
2019-03-20 17:41 | 1553118071.0725 | -469 | 0.20948 | 0.000455 | 0 | 109.4346 | 110.8651 | 126.5902 | 21.0369 | 0.002003 | 0.00181 | -151.8864 | 0.23927 | 8.4006 | 0.003678 | 0.005413 | 0.005189 | 6.1e-05 | 0.001249 | 0.26936 | 1.3566 | 2.5481 | 100.59 | 20.66 | 18.97 | 69.7991 | 110.19 | 65.1729 | 0.004598 | 1.0435 | 0.025141 | 24.0922 | -0.11617 | 0.14234 | -1.2e-05 | -0.000901 | -0.000901 | -6.4961
2019-03-20 17:41 | 1553118072.0739 | -468 | 0.20945 | 0.000449 | 0 | 109.3913 | 110.7882 | 126.5604 | 21.0374 | 0.001978 | 0.001855 | -151.8738 | 0.23864 | 8.4046 | 0.003695 | 0.004812 | 0.00468 | 6e-05 | 0.001248 | 0.29103 | 1.3288 | 2.548 | 100.59 | 20.66 | 18.963 | 69.2085 | 110.5983 | 66.238 | 0.004596 | 1.0402 | 0.025098 | 24.1273 | -0.1135 | 0.14197 | -4e-06 | -0.00027 | -0.00027 | -21.5889
2019-03-20 17:41 | 1553118073.0726 | -467 | 0.2095 | 0.000446 | 0 | 109.3848 | 110.7605 | 126.4311 | 21.0368 | 0.00196 | 0.001823 | -151.8698 | 0.23905 | 8.4027 | 0.003659 | 0.005349 | 0.005055 | 5.8e-05 | 0.001246 | 0.29031 | 1.34 | 2.5481 | 100.59 | 20.66 | 18.958 | 68.7625 | 109.7381 | 65.4843 | 0.004595 | 1.0426 | 0.025113 | 24.0882 | -0.11179 | 0.14221 | 0 | 3e-05 | 3e-05 | 196.6851
2019-03-20 17:41 | 1553118074.0736 | -466 | 0.20948 | 0.000449 | 1e-06 | 109.3647 | 110.9492 | 126.3261 | 21.0368 | 0.002 | 0.001845 | -151.8726 | 0.23869 | 8.4015 | 0.003665 | 0.005451 | 0.005155 | 6.1e-05 | 0.001244 | 0.28816 | 1.3301 | 2.548 | 100.58 | 20.66 | 18.951 | 69.7374 | 109.8856 | 65.9915 | 0.004594 | 1.041 | 0.025084 | 24.0962 | -0.12866 | 0.14199 | -5e-06 | -0.00038 | -0.00038 | -15.3459
2019-03-20 17:41 | 1553118075.0741 | -465 | 0.20949 | 0.000447 | 1e-06 | 109.334 | 110.8637 | 126.202 | 21.0377 | 0.001957 | 0.001819 | -151.8866 | 0.23986 | 8.4048 | 0.003662 | 0.004902 | 0.004779 | 5.9e-05 | 0.001235 | 0.26322 | 1.3515 | 2.5481 | 100.58 | 20.66 | 18.948 | 68.697 | 109.7965 | 65.3853 | 0.004593 | 1.0429 | 0.025094 | 24.0628 | -0.12427 | 0.14269 | -2e-06 | -0.000194 | -0.000194 | -30.1942
2019-03-20 17:41 | 1553118076.0727 | -464 | 0.2095 | 0.000452 | 0 | 109.2991 | 110.8323 | 126.1447 | 21.0371 | 0.001983 | 0.001823 | -151.8997 | 0.23889 | 8.4042 | 0.003654 | 0.004474 | 0.004186 | 5.9e-05 | 0.001235 | 0.27815 | 1.3007 | 2.5479 | 100.58 | 20.66 | 18.946 | 69.3299 | 109.6035 | 65.4665 | 0.004593 | 1.0426 | 0.025085 | 24.0602 | -0.12459 | 0.14212 | -8e-06 | -0.000623 | -0.000623 | -9.3693
2019-03-20 17:41 | 1553118077.0729 | -463 | 0.20951 | 0.000449 | 1e-06 | 109.4 | 110.6723 | 125.7471 | 21.0371 | 0.001957 | 0.001823 | -151.9025 | 0.23975 | 8.4021 | 0.003603 | 0.004903 | 0.004668 | 6e-05 | 0.001239 | 0.26952 | 1.3638 | 2.548 | 100.58 | 20.66 | 18.943 | 68.7065 | 108.3676 | 65.4821 | 0.004592 | 1.0426 | 0.025045 | 24.0228 | -0.10342 | 0.14263 | -4e-06 | -0.000371 | -0.000371 | -15.797
2019-03-20 17:41 | 1553118078.0733 | -462 | 0.20952 | 0.000447 | 0 | 109.3369 | 110.9889 | 126.0168 | 21.0377 | 0.001975 | 0.001853 | -151.8923 | 0.23875 | 8.4022 | 0.003575 | 0.006007 | 0.005872 | 6.1e-05 | 0.001234 | 0.27367 | 1.3169 | 2.5479 | 100.58 | 20.66 | 18.942 | 69.1338 | 107.6978 | 66.1921 | 0.004592 | 1.0404 | 0.025046 | 24.0738 | -0.13413 | 0.14203 | -2e-06 | -0.000194 | -0.000194 | -30.1157
2019-03-20 17:41 | 1553118079.0739 | -461 | 0.2095 | 0.000456 | 1e-06 | 109.3356 | 110.8681 | 126.0884 | 21.0382 | 0.001967 | 0.001786 | -151.894 | 0.23825 | 8.4034 | 0.003648 | 0.005063 | 0.004793 | 5.6e-05 | 0.001225 | 0.29103 | 1.3475 | 2.5479 | 100.58 | 20.67 | 18.942 | 68.9463 | 109.4658 | 64.587 | 0.004595 | 1.0453 | 0.025113 | 24.0236 | -0.1245 | 0.14173 | -1.3e-05 | -0.001023 | -0.001023 | -5.6921
2019-03-20 17:41 | 1553118080.0734 | -460 | 0.2095 | 0.000449 | 1e-06 | 109.3483 | 110.9063 | 126.1125 | 21.0387 | 0.00193 | 0.001793 | -151.8855 | 0.23885 | 8.4018 | 0.003563 | 0.004952 | 0.004768 | 5.5e-05 | 0.001223 | 0.28279 | 1.3428 | 2.548 | 100.58 | 20.67 | 18.943 | 68.053 | 107.4103 | 64.7655 | 0.004595 | 1.0448 | 0.025108 | 24.0323 | -0.12654 | 0.14209 | -4e-06 | -0.000323 | -0.000323 | -18.1002
2019-03-20 17:41 | 1553118081.073 | -459 | 0.20951 | 0.000452 | 1e-06 | 109.2215 | 110.565 | 125.8004 | 21.0386 | 0.001979 | 0.001856 | -151.8855 | 0.23862 | 8.4056 | 0.003593 | 0.00483 | 0.004588 | 5.6e-05 | 0.001223 | 0.28655 | 1.3711 | 2.548 | 100.58 | 20.67 | 18.948 | 69.2274 | 108.1342 | 66.2804 | 0.004596 | 1.0401 | 0.025021 | 24.0562 | -0.10936 | 0.14195 | -9e-06 | -0.00073 | -0.00073 | -7.995
2019-03-20 17:41 | 1553118082.0731 | -458 | 0.20949 | 0.000451 | 0 | 109.2661 | 110.7497 | 125.445 | 21.0381 | 0.001934 | 0.001833 | -151.9017 | 0.23903 | 8.4038 | 0.003731 | 0.005715 | 0.005544 | 5.9e-05 | 0.001228 | 0.29087 | 1.3905 | 2.548 | 100.58 | 20.67 | 18.952 | 68.1327 | 111.4804 | 65.7039 | 0.004597 | 1.0419 | 0.025007 | 24.0018 | -0.12063 | 0.1422 | -7e-06 | -0.000514 | -0.000514 | -11.3675
2019-03-20 17:41 | 1553118083.0745 | -457 | 0.20953 | 0.000451 | 0 | 109.2506 | 110.6484 | 125.4891 | 21.0388 | 0.001983 | 0.001843 | -151.9103 | 0.23928 | 8.3999 | 0.003672 | 0.005215 | 0.005039 | 5.9e-05 | 0.001227 | 0.29015 | 1.3634 | 2.548 | 100.58 | 20.67 | 18.957 | 69.3283 | 110.0436 | 65.9473 | 0.004598 | 1.0411 | 0.025002 | 24.0147 | -0.11371 | 0.14235 | -7e-06 | -0.000524 | -0.000524 | -11.1659
2019-03-20 17:41 | 1553118084.0732 | -456 | 0.2095 | 0.000446 | 0 | 109.28 | 110.7811 | 125.9712 | 21.0385 | 0.002041 | 0.001805 | -151.9108 | 0.23954 | 8.4007 | 0.003657 | 0.004903 | 0.004716 | 5.8e-05 | 0.001231 | 0.29878 | 1.3568 | 2.5481 | 100.58 | 20.67 | 18.965 | 70.7192 | 109.6909 | 65.0543 | 0.0046 | 1.0439 | 0.025083 | 24.029 | -0.12202 | 0.1425 | 1e-06 | 6.2e-05 | 6.2e-05 | 95.1688
2019-03-20 17:41 | 1553118085.0722 | -455 | 0.20949 | 0.000452 | 0 | 109.2816 | 110.6218 | 126.0865 | 21.0389 | 0.001983 | 0.001827 | -151.9112 | 0.23892 | 8.4038 | 0.003664 | 0.00523 | 0.005015 | 5.8e-05 | 0.00123 | 0.2808 | 1.3809 | 2.5481 | 100.58 | 20.67 | 18.968 | 69.333 | 109.8449 | 65.5775 | 0.004601 | 1.0423 | 0.025076 | 24.0586 | -0.10903 | 0.14213 | -8e-06 | -0.000626 | -0.000626 | -9.3341
2019-03-20 17:41 | 1553118086.0737 | -454 | 0.2095 | 0.000448 | 0 | 109.2477 | 110.7041 | 125.8888 | 21.0383 | 0.001919 | 0.001854 | -151.9221 | 0.23959 | 8.4014 | 0.003672 | 0.005251 | 0.004928 | 5.9e-05 | 0.001236 | 0.27312 | 1.3646 | 2.5482 | 100.58 | 20.67 | 18.971 | 67.7961 | 110.0431 | 66.2099 | 0.004602 | 1.0403 | 0.025033 | 24.0622 | -0.11845 | 0.14253 | -3e-06 | -0.0002 | -0.0002 | -29.343
2019-03-20 17:41 | 1553118087.0734 | -453 | 0.20949 | 0.000452 | 1e-06 | 109.2519 | 110.7627 | 126.1123 | 21.0388 | 0.001992 | 0.00184 | -151.9182 | 0.23899 | 8.4018 | 0.003594 | 0.005057 | 0.004876 | 5.7e-05 | 0.001231 | 0.28313 | 1.3494 | 2.5481 | 100.58 | 20.67 | 18.974 | 69.5389 | 108.1558 | 65.878 | 0.004602 | 1.0413 | 0.025067 | 24.0718 | -0.12285 | 0.14217 | -9e-06 | -0.000691 | -0.000691 | -8.4588
2019-03-20 17:41 | 1553118088.0724 | -452 | 0.20951 | 0.000452 | 1e-06 | 109.2011 | 110.5325 | 126.0722 | 21.0394 | 0.001943 | 0.001826 | -151.9135 | 0.23907 | 8.4024 | 0.003608 | 0.005203 | 0.004978 | 5.9e-05 | 0.001235 | 0.27992 | 1.3102 | 2.548 | 100.58 | 20.67 | 18.976 | 68.3713 | 108.5015 | 65.5376 | 0.004603 | 1.0424 | 0.025076 | 24.0558 | -0.1084 | 0.14222 | -8e-06 | -0.000682 | -0.000682 | -8.5665
2019-03-20 17:41 | 1553118089.0735 | -451 | 0.2095 | 0.000449 | 0 | 109.1927 | 110.491 | 125.8335 | 21.0396 | 0.002007 | 0.001861 | -151.9283 | 0.23853 | 8.4023 | 0.00363 | 0.005155 | 0.004962 | 6e-05 | 0.001231 | 0.28838 | 1.299 | 2.5481 | 100.58 | 20.67 | 18.981 | 69.9 | 109.0281 | 66.3789 | 0.004604 | 1.0398 | 0.025021 | 24.0629 | -0.10573 | 0.1419 | -4e-06 | -0.000324 | -0.000324 | -18.0135
2019-03-20 17:41 | 1553118090.0726 | -450 | 0.2095 | 0.000448 | 0 | 109.2218 | 110.6454 | 125.8649 | 21.0393 | 0.001991 | 0.001817 | -151.9491 | 0.2385 | 8.4031 | 0.003608 | 0.00528 | 0.005038 | 5.7e-05 | 0.00123 | 0.28655 | 1.3203 | 2.548 | 100.58 | 20.67 | 18.982 | 69.5137 | 108.5068 | 65.3236 | 0.004604 | 1.0431 | 0.025063 | 24.0285 | -0.11583 | 0.14188 | -2e-06 | -0.00022 | -0.00022 | -26.4626
2019-03-20 17:42 | 1553118091.0726 | -449 | 0.2095 | 0.000449 | 0 | 109.1257 | 110.7156 | 125.8297 | 21.039 | 0.002019 | 0.001814 | -151.9554 | 0.23966 | 8.4038 | 0.003568 | 0.004802 | 0.004543 | 5.6e-05 | 0.001225 | 0.29087 | 1.3477 | 2.5482 | 100.58 | 20.67 | 18.983 | 70.1784 | 107.5303 | 65.267 | 0.004605 | 1.0432 | 0.025062 | 24.0231 | -0.12938 | 0.14257 | -4e-06 | -0.000294 | -0.000294 | -19.9097
2019-03-20 17:42 | 1553118092.0726 | -448 | 0.20951 | 0.000451 | 0 | 109.2111 | 110.5975 | 126.2226 | 21.0395 | 0.001982 | 0.00185 | -151.9363 | 0.2396 | 8.4049 | 0.003591 | 0.004881 | 0.004675 | 6.1e-05 | 0.001231 | 0.29856 | 1.322 | 2.5481 | 100.58 | 20.67 | 18.983 | 69.2963 | 108.0838 | 66.1284 | 0.004605 | 1.0406 | 0.025069 | 24.0912 | -0.11284 | 0.14253 | -7e-06 | -0.000539 | -0.000539 | -10.867
2019-03-20 17:42 | 1553118093.0727 | -447 | 0.2095 | 0.000448 | 0 | 109.2021 | 110.6469 | 126.302 | 21.0407 | 0.001944 | 0.001873 | -151.9435 | 0.23932 | 8.402 | 0.003608 | 0.005407 | 0.005161 | 5.5e-05 | 0.001228 | 0.29031 | 1.3728 | 2.5481 | 100.58 | 20.67 | 18.982 | 68.3889 | 108.5074 | 66.6788 | 0.004604 | 1.0389 | 0.025056 | 24.1183 | -0.11757 | 0.14237 | -3e-06 | -0.000243 | -0.000243 | -24.0535
2019-03-20 17:42 | 1553118094.0731 | -446 | 0.20951 | 0.00045 | 0 | 109.2842 | 110.6768 | 125.9245 | 21.0404 | 0.001999 | 0.001834 | -151.9249 | 0.23883 | 8.4033 | 0.00357 | 0.004318 | 0.004126 | 5.7e-05 | 0.001231 | 0.2995 | 1.3629 | 2.5479 | 100.58 | 20.67 | 18.977 | 69.6979 | 107.5657 | 65.7423 | 0.004603 | 1.0418 | 0.025053 | 24.049 | -0.11326 | 0.14208 | -5e-06 | -0.000431 | -0.000431 | -13.5401
2019-03-20 17:42 | 1553118095.0728 | -445 | 0.20949 | 0.000445 | 0 | 109.1691 | 110.5886 | 125.9999 | 21.0408 | 0.002016 | 0.001772 | -151.915 | 0.23904 | 8.4018 | 0.003674 | 0.004428 | 0.004134 | 6.1e-05 | 0.001235 | 0.28047 | 1.3523 | 2.5481 | 100.58 | 20.66 | 18.972 | 70.1294 | 110.1014 | 64.2548 | 0.004599 | 1.0464 | 0.025116 | 24.0034 | -0.11555 | 0.14221 | 2e-06 | 0.000103 | 0.000103 | 56.6413
2019-03-20 17:42 | 1553118096.0728 | -444 | 0.20949 | 0.000452 | 0 | 109.2305 | 110.503 | 126.0175 | 21.04 | 0.00197 | 0.00181 | -151.9127 | 0.23908 | 8.4023 | 0.003594 | 0.00475 | 0.004477 | 5.5e-05 | 0.001241 | 0.29374 | 1.3511 | 2.5479 | 100.58 | 20.66 | 18.969 | 69.0075 | 108.157 | 65.1696 | 0.004598 | 1.0435 | 0.025084 | 24.0376 | -0.1036 | 0.14223 | -8e-06 | -0.00059 | -0.00059 | -9.9098
2019-03-20 17:42 | 1553118097.0735 | -443 | 0.2095 | 0.000448 | 0 | 109.1649 | 110.6368 | 125.7839 | 21.0406 | 0.001916 | 0.00184 | -151.9146 | 0.23888 | 8.4018 | 0.00361 | 0.004877 | 0.004666 | 5.8e-05 | 0.001245 | 0.2855 | 1.3555 | 2.5481 | 100.58 | 20.66 | 18.962 | 67.7171 | 108.5371 | 65.8844 | 0.004597 | 1.0413 | 0.025034 | 24.0406 | -0.1198 | 0.14211 | -3e-06 | -0.000262 | -0.000262 | -22.2878
2019-03-20 17:42 | 1553118098.074 | -442 | 0.20946 | 0.000449 | 0 | 109.108 | 110.7332 | 125.3736 | 21.04 | 0.001957 | 0.001856 | -151.9188 | 0.23893 | 8.4039 | 0.003529 | 0.004713 | 0.004538 | 5.6e-05 | 0.001247 | 0.28982 | 1.317 | 2.5482 | 100.58 | 20.66 | 18.959 | 68.7062 | 106.5731 | 66.2686 | 0.004596 | 1.0401 | 0.024979 | 24.015 | -0.13225 | 0.14214 | -4e-06 | -0.000324 | -0.000324 | -18.0456
2019-03-20 17:42 | 1553118099.073 | -441 | 0.20949 | 0.00045 | 0 | 109.1726 | 110.7713 | 125.257 | 21.0402 | 0.001963 | 0.00184 | -151.9008 | 0.2398 | 8.4055 | 0.003538 | 0.005522 | 0.005249 | 5.8e-05 | 0.001252 | 0.27417 | 1.3148 | 2.5483 | 100.58 | 20.66 | 18.954 | 68.8391 | 106.7898 | 65.8794 | 0.004595 | 1.0413 | 0.024982 | 23.99 | -0.13003 | 0.14266 | -4e-06 | -0.000309 | -0.000309 | -18.9662
2019-03-20 17:42 | 1553118100.0724 | -440 | 0.20947 | 0.00045 | 0 | 109.1966 | 110.5953 | 125.2734 | 21.0398 | 0.001935 | 0.001777 | -151.8952 | 0.23896 | 8.4001 | 0.003509 | 0.005039 | 0.00474 | 5.5e-05 | 0.00125 | 0.29319 | 1.4039 | 2.5483 | 100.58 | 20.66 | 18.953 | 68.175 | 106.0841 | 64.3607 | 0.004594 | 1.046 | 0.02504 | 23.9378 | -0.11385 | 0.14216 | -5e-06 | -0.000371 | -0.000371 | -15.7535
2019-03-20 17:42 | 1553118101.0721 | -439 | 0.20951 | 0.000449 | 1e-06 | 109.1956 | 110.5273 | 125.5976 | 21.0401 | 0.001958 | 0.001802 | -151.8908 | 0.23857 | 8.4016 | 0.003634 | 0.005044 | 0.004957 | 5.8e-05 | 0.001251 | 0.3079 | 1.3568 | 2.5481 | 100.58 | 20.66 | 18.952 | 68.7275 | 109.1231 | 64.9715 | 0.004594 | 1.0441 | 0.025049 | 23.9904 | -0.10842 | 0.14192 | -4e-06 | -0.000366 | -0.000366 | -15.918
2019-03-20 17:42 | 1553118102.0732 | -438 | 0.20948 | 0.000446 | 1e-06 | 109.129 | 110.5445 | 125.4098 | 21.0397 | 0.001948 | 0.001829 | -151.9037 | 0.23889 | 8.4005 | 0.003691 | 0.005361 | 0.005102 | 5.5e-05 | 0.00125 | 0.28888 | 1.326 | 2.5482 | 100.58 | 20.66 | 18.953 | 68.4737 | 110.5003 | 65.6259 | 0.004594 | 1.0421 | 0.025006 | 23.9957 | -0.11527 | 0.14211 | -1e-06 | -0.000149 | -0.000149 | -39.3033
2019-03-20 17:42 | 1553118103.0726 | -437 | 0.20947 | 0.000449 | 1e-06 | 109.2134 | 110.8285 | 125.5342 | 21.0404 | 0.001976 | 0.001841 | -151.8904 | 0.23859 | 8.401 | 0.003683 | 0.005337 | 0.005132 | 5.7e-05 | 0.001251 | 0.29717 | 1.3587 | 2.5482 | 100.58 | 20.66 | 18.955 | 69.1638 | 110.3221 | 65.913 | 0.004595 | 1.0412 | 0.025008 | 24.0178 | -0.13131 | 0.14194 | -5e-06 | -0.000392 | -0.000392 | -14.8971
2019-03-20 17:42 | 1553118104.0723 | -436 | 0.2095 | 0.000449 | 0 | 109.0782 | 110.605 | 125.4915 | 21.0409 | 0.001952 | 0.001819 | -151.9097 | 0.23772 | 8.4042 | 0.003617 | 0.004658 | 0.004363 | 5.3e-05 | 0.001251 | 0.2943 | 1.3763 | 2.5483 | 100.58 | 20.66 | 18.962 | 68.5708 | 108.7062 | 65.386 | 0.004597 | 1.0429 | 0.025023 | 23.995 | -0.12433 | 0.14142 | -4e-06 | -0.000305 | -0.000305 | -19.081
2019-03-20 17:42 | 1553118105.0733 | -435 | 0.20949 | 0.000448 | 0 | 109.1309 | 110.6237 | 125.6849 | 21.0411 | 0.001997 | 0.00187 | -151.9174 | 0.23832 | 8.4005 | 0.003686 | 0.004897 | 0.004679 | 5.9e-05 | 0.001247 | 0.26792 | 1.3784 | 2.5482 | 100.58 | 20.66 | 18.967 | 69.6527 | 110.3842 | 66.6066 | 0.004598 | 1.0391 | 0.024998 | 24.0567 | -0.12152 | 0.14178 | -2e-06 | -0.000138 | -0.000138 | -42.2698
2019-03-20 17:42 | 1553118106.0738 | -434 | 0.20948 | 0.000446 | 1e-06 | 109.2402 | 110.6857 | 125.8858 | 21.0409 | 0.00195 | 0.001848 | -151.9092 | 0.23805 | 8.4009 | 0.003547 | 0.005118 | 0.004935 | 5.8e-05 | 0.001257 | 0.2912 | 1.3418 | 2.548 | 100.58 | 20.66 | 18.973 | 68.5316 | 107.0111 | 66.0821 | 0.004599 | 1.0407 | 0.025037 | 24.0574 | -0.11758 | 0.14161 | -2e-06 | -0.000201 | -0.000201 | -28.9328
2019-03-20 17:42 | 1553118107.0737 | -433 | 0.20947 | 0.000448 | 0 | 109.1972 | 110.6592 | 125.9978 | 21.0406 | 0.001957 | 0.001841 | -151.9227 | 0.23881 | 8.4021 | 0.00353 | 0.004703 | 0.004503 | 5.9e-05 | 0.001256 | 0.27201 | 1.3656 | 2.5483 | 100.58 | 20.66 | 18.987 | 68.7109 | 106.6036 | 65.9164 | 0.004603 | 1.0412 | 0.025054 | 24.0622 | -0.11896 | 0.14207 | -3e-06 | -0.000243 | -0.000243 | -23.9932
2019-03-20 17:42 | 1553118108.0725 | -432 | 0.20948 | 0.000452 | 0 | 109.1215 | 110.5209 | 126.0367 | 21.04 | 0.001984 | 0.001806 | -151.9296 | 0.23822 | 8.4033 | 0.003505 | 0.004754 | 0.004608 | 6e-05 | 0.001259 | 0.2907 | 1.3621 | 2.5481 | 100.58 | 20.66 | 18.996 | 69.3396 | 105.9946 | 65.0652 | 0.004605 | 1.0438 | 0.02509 | 24.0357 | -0.11398 | 0.14172 | -8e-06 | -0.000649 | -0.000649 | -8.9701
2019-03-20 17:42 | 1553118109.0734 | -431 | 0.20952 | 0.000447 | 1e-06 | 109.1044 | 110.8177 | 126.0877 | 21.0405 | 0.001901 | 0.001882 | -151.9411 | 0.23956 | 8.4026 | 0.003525 | 0.005316 | 0.005045 | 6e-05 | 0.00127 | 0.27632 | 1.3046 | 2.548 | 100.58 | 20.65 | 19.005 | 67.3566 | 106.4885 | 66.8963 | 0.004604 | 1.0382 | 0.025027 | 24.1055 | -0.13937 | 0.14252 | -2e-06 | -0.000223 | -0.000223 | -26.2906
2019-03-20 17:42 | 1553118110.0726 | -430 | 0.20953 | 0.000449 | 1e-06 | 109.1658 | 110.6502 | 126.1085 | 21.0407 | 0.001974 | 0.001854 | -151.9315 | 0.23872 | 8.4036 | 0.003538 | 0.005374 | 0.005184 | 5.9e-05 | 0.001272 | 0.27616 | 1.3705 | 2.5481 | 100.58 | 20.65 | 19.026 | 69.1134 | 106.7888 | 66.2287 | 0.004609 | 1.0403 | 0.025054 | 24.0838 | -0.1208 | 0.14202 | -4e-06 | -0.000357 | -0.000357 | -16.3274
2019-03-20 17:42 | 1553118111.0736 | -429 | 0.20949 | 0.000451 | 0 | 109.1287 | 110.4716 | 126.2203 | 21.0405 | 0.001985 | 0.001842 | -151.9243 | 0.23894 | 8.4015 | 0.003512 | 0.00494 | 0.004687 | 5.9e-05 | 0.00127 | 0.30453 | 1.3259 | 2.548 | 100.58 | 20.66 | 19.038 | 69.3665 | 106.1644 | 65.9263 | 0.004615 | 1.0412 | 0.025076 | 24.0838 | -0.1094 | 0.14215 | -7e-06 | -0.000534 | -0.000534 | -10.9338
2019-03-20 17:42 | 1553118112.0737 | -428 | 0.20946 | 0.000452 | 0 | 109.269 | 110.6405 | 126.2107 | 21.0415 | 0.002031 | 0.001848 | -151.9304 | 0.23832 | 8.4029 | 0.003519 | 0.004211 | 0.00403 | 6e-05 | 0.001268 | 0.26985 | 1.3246 | 2.5482 | 100.58 | 20.65 | 19.06 | 70.4851 | 106.3477 | 66.0865 | 0.004618 | 1.0407 | 0.025069 | 24.0885 | -0.11157 | 0.14178 | -7e-06 | -0.000549 | -0.000549 | -10.6137
2019-03-20 17:42 | 1553118113.0727 | -427 | 0.20949 | 0.000446 | 0 | 109.2383 | 110.6286 | 125.9769 | 21.0412 | 0.001901 | 0.001832 | -151.9148 | 0.23859 | 8.4015 | 0.003483 | 0.005402 | 0.005097 | 6e-05 | 0.001266 | 0.26465 | 1.3392 | 2.5481 | 100.58 | 20.66 | 19.071 | 67.3509 | 105.46 | 65.6882 | 0.004623 | 1.0419 | 0.025061 | 24.0521 | -0.11312 | 0.14194 | 0 | -1.2e-05 | -1.2e-05 | -477.9415
2019-03-20 17:42 | 1553118114.0738 | -426 | 0.20948 | 0.000445 | 1e-06 | 109.1506 | 110.6058 | 125.878 | 21.0412 | 0.00184 | 0.0018 | -151.898 | 0.23857 | 8.4036 | 0.003516 | 0.005163 | 0.004998 | 5.7e-05 | 0.001269 | 0.29679 | 1.3243 | 2.548 | 100.58 | 20.66 | 19.083 | 65.8958 | 106.2587 | 64.9203 | 0.004626 | 1.0443 | 0.025079 | 24.0154 | -0.11845 | 0.14192 | 0 | -3.6e-05 | -3.6e-05 | -161.0103
2019-03-20 17:42 | 1553118115.0731 | -425 | 0.2095 | 0.000451 | 1e-06 | 109.1464 | 110.6046 | 125.8243 | 21.0412 | 0.001999 | 0.001871 | -151.8914 | 0.2393 | 8.4027 | 0.003592 | 0.003983 | 0.003751 | 6.1e-05 | 0.001268 | 0.27936 | 1.3843 | 2.5482 | 100.58 | 20.66 | 19.104 | 69.7167 | 108.1091 | 66.6209 | 0.004631 | 1.0391 | 0.025011 | 24.0706 | -0.11871 | 0.14236 | -6e-06 | -0.00052 | -0.00052 | -11.247
2019-03-20 17:42 | 1553118116.0719 | -424 | 0.20951 | 0.00045 | 0 | 109.1875 | 110.5463 | 125.7682 | 21.0416 | 0.002018 | 0.00187 | -151.8733 | 0.23857 | 8.4011 | 0.003707 | 0.004653 | 0.004475 | 6.1e-05 | 0.001269 | 0.28119 | 1.3499 | 2.5481 | 100.58 | 20.66 | 19.115 | 70.1632 | 110.9015 | 66.6103 | 0.004634 | 1.0391 | 0.025006 | 24.0648 | -0.11063 | 0.14193 | -5e-06 | -0.000409 | -0.000409 | -14.2735
2019-03-20 17:42 | 1553118117.0733 | -423 | 0.20949 | 0.00045 | 0 | 109.1591 | 110.5172 | 125.9093 | 21.0411 | 0.00198 | 0.001871 | -151.8642 | 0.23882 | 8.402 | 0.003605 | 0.004766 | 0.004632 | 6.1e-05 | 0.001269 | 0.27057 | 1.3407 | 2.5481 | 100.58 | 20.66 | 19.136 | 69.2573 | 108.4313 | 66.6331 | 0.004639 | 1.039 | 0.025019 | 24.0791 | -0.1106 | 0.14208 | -6e-06 | -0.000465 | -0.000465 | -12.5616
2019-03-20 17:42 | 1553118118.0734 | -422 | 0.20948 | 0.00045 | 1e-06 | 109.1649 | 110.651 | 126.0895 | 21.042 | 0.002007 | 0.00185 | -151.8743 | 0.23838 | 8.4025 | 0.003654 | 0.004945 | 0.004653 | 6e-05 | 0.001271 | 0.2557 | 1.3498 | 2.5482 | 100.58 | 20.66 | 19.146 | 69.8944 | 109.6131 | 66.1371 | 0.004641 | 1.0406 | 0.025055 | 24.0788 | -0.12094 | 0.14181 | -5e-06 | -0.000433 | -0.000433 | -13.4631
2019-03-20 17:42 | 1553118119.0735 | -421 | 0.2095 | 0.000449 | 0 | 109.1927 | 110.5497 | 126.1569 | 21.0423 | 0.001968 | 0.001852 | -151.8733 | 0.23848 | 8.4027 | 0.003679 | 0.005308 | 0.005061 | 6e-05 | 0.001274 | 0.28136 | 1.343 | 2.5483 | 100.58 | 20.66 | 19.154 | 68.9543 | 110.2332 | 66.1669 | 0.004643 | 1.0405 | 0.025061 | 24.0863 | -0.11048 | 0.14187 | -4e-06 | -0.000345 | -0.000345 | -16.8996
2019-03-20 17:42 | 1553118120.0721 | -420 | 0.20949 | 0.000452 | 1e-06 | 109.1597 | 110.5351 | 125.9959 | 21.0414 | 0.001877 | 0.001806 | -151.8721 | 0.23797 | 8.4028 | 0.003676 | 0.004828 | 0.004614 | 5.7e-05 | 0.001271 | 0.27345 | 1.3412 | 2.548 | 100.58 | 20.66 | 19.167 | 66.7787 | 110.159 | 65.0642 | 0.004646 | 1.0439 | 0.025086 | 24.0317 | -0.112 | 0.14157 | -8e-06 | -0.00067 | -0.00067 | -8.6879
2019-03-20 17:42 | 1553118121.0722 | -419 | 0.20945 | 0.000448 | 0 | 109.1034 | 110.5624 | 125.5964 | 21.043 | 0.001981 | 0.00187 | -151.886 | 0.23877 | 8.4013 | 0.003624 | 0.005705 | 0.005492 | 5.8e-05 | 0.001263 | 0.27295 | 1.3545 | 2.548 | 100.58 | 20.66 | 19.172 | 69.2692 | 108.8752 | 66.5998 | 0.004648 | 1.0391 | 0.024989 | 24.048 | -0.11882 | 0.14204 | -3e-06 | -0.000246 | -0.000246 | -23.6937
2019-03-20 17:42 | 1553118122.0736 | -418 | 0.20947 | 0.000451 | 1e-06 | 109.1519 | 110.5721 | 125.606 | 21.0422 | 0.001972 | 0.001804 | -151.8767 | 0.23839 | 8.4014 | 0.003664 | 0.005019 | 0.004826 | 5.8e-05 | 0.001259 | 0.28191 | 1.3153 | 2.5482 | 100.58 | 20.66 | 19.177 | 69.0728 | 109.8554 | 65.0228 | 0.004649 | 1.044 | 0.025048 | 23.9931 | -0.11563 | 0.14182 | -7e-06 | -0.000611 | -0.000611 | -9.5338
2019-03-20 17:42 | 1553118123.0733 | -417 | 0.20947 | 0.000454 | 1e-06 | 109.1332 | 110.6648 | 125.5006 | 21.0426 | 0.002029 | 0.001771 | -151.8986 | 0.23824 | 8.4 | 0.003684 | 0.005526 | 0.005244 | 6e-05 | 0.001257 | 0.2625 | 1.3429 | 2.5481 | 100.58 | 20.66 | 19.179 | 70.4338 | 110.3446 | 64.2183 | 0.004649 | 1.0465 | 0.025068 | 23.9545 | -0.12466 | 0.14173 | -1.1e-05 | -0.000836 | -0.000836 | -6.9639
2019-03-20 17:42 | 1553118124.0723 | -416 | 0.20946 | 0.000451 | 0 | 109.118 | 110.577 | 125.6398 | 21.0426 | 0.001983 | 0.001869 | -151.9136 | 0.23803 | 8.4016 | 0.003779 | 0.005405 | 0.005213 | 5.8e-05 | 0.001254 | 0.28313 | 1.3373 | 2.548 | 100.58 | 20.66 | 19.18 | 69.3234 | 112.6559 | 66.5886 | 0.004649 | 1.0392 | 0.024994 | 24.0518 | -0.1188 | 0.1416 | -6e-06 | -0.000466 | -0.000466 | -12.4907
2019-03-20 17:42 | 1553118125.0728 | -415 | 0.2095 | 0.00045 | 0 | 109.0252 | 110.6387 | 125.3267 | 21.0425 | 0.001989 | 0.001817 | -151.909 | 0.2394 | 8.4045 | 0.003822 | 0.004708 | 0.004472 | 5.8e-05 | 0.001249 | 0.26465 | 1.3157 | 2.5481 | 100.58 | 20.66 | 19.179 | 69.4769 | 113.7058 | 65.3257 | 0.004649 | 1.043 | 0.025009 | 23.9771 | -0.1314 | 0.14242 | -5e-06 | -0.000392 | -0.000392 | -14.9271
2019-03-20 17:42 | 1553118126.0725 | -414 | 0.20949 | 0.000453 | 1e-06 | 109.0776 | 110.6372 | 125.1985 | 21.0427 | 0.002038 | 0.001802 | -151.9361 | 0.23815 | 8.4041 | 0.003745 | 0.005631 | 0.005441 | 6.1e-05 | 0.001253 | 0.2792 | 1.3014 | 2.5481 | 100.58 | 20.66 | 19.178 | 70.6382 | 111.837 | 64.9725 | 0.004649 | 1.0441 | 0.025009 | 23.9523 | -0.12698 | 0.14168 | -1e-05 | -0.000833 | -0.000833 | -6.9904
2019-03-20 17:42 | 1553118127.0725 | -413 | 0.20953 | 0.000448 | 1e-06 | 109.1102 | 110.4693 | 125.0779 | 21.043 | 0.001968 | 0.001779 | -151.9272 | 0.23772 | 8.4045 | 0.003729 | 0.004918 | 0.00469 | 5.9e-05 | 0.001249 | 0.29463 | 1.3311 | 2.5481 | 100.58 | 20.66 | 19.178 | 68.9752 | 111.4324 | 64.4153 | 0.004649 | 1.0459 | 0.025018 | 23.9211 | -0.11073 | 0.14142 | -3e-06 | -0.000295 | -0.000295 | -19.6946
2019-03-20 17:42 | 1553118128.0729 | -412 | 0.20952 | 0.000451 | 0 | 109.0316 | 110.4196 | 124.992 | 21.0432 | 0.00198 | 0.001851 | -151.9162 | 0.23904 | 8.4037 | 0.003822 | 0.005217 | 0.004986 | 5.9e-05 | 0.00125 | 0.26145 | 1.2915 | 2.548 | 100.58 | 20.66 | 19.173 | 69.2631 | 113.7002 | 66.1522 | 0.004648 | 1.0405 | 0.024945 | 23.9743 | -0.11315 | 0.14221 | -7e-06 | -0.000525 | -0.000525 | -11.1388
2019-03-20 17:42 | 1553118129.0726 | -411 | 0.2095 | 0.00045 | 1e-06 | 109.0672 | 110.4656 | 124.6784 | 21.0439 | 0.001982 | 0.001793 | -151.912 | 0.23871 | 8.4006 | 0.003829 | 0.005184 | 0.004918 | 5.6e-05 | 0.001254 | 0.27328 | 1.3164 | 2.5482 | 100.58 | 20.66 | 19.17 | 69.2958 | 113.8787 | 64.7682 | 0.004647 | 1.0448 | 0.024965 | 23.8953 | -0.11395 | 0.14201 | -6e-06 | -0.000457 | -0.000457 | -12.7603
2019-03-20 17:42 | 1553118130.0726 | -410 | 0.20954 | 0.000447 | 0 | 109.0565 | 110.4234 | 124.7357 | 21.0437 | 0.001961 | 0.00185 | -151.9297 | 0.23859 | 8.3995 | 0.003696 | 0.004588 | 0.004368 | 5.7e-05 | 0.001259 | 0.29087 | 1.3231 | 2.5483 | 100.58 | 20.66 | 19.166 | 68.8012 | 110.6378 | 66.1245 | 0.004646 | 1.0406 | 0.024921 | 23.9487 | -0.11141 | 0.14194 | -2e-06 | -0.000122 | -0.000122 | -47.9966
2019-03-20 17:42 | 1553118131.0721 | -409 | 0.20951 | 0.00045 | 1e-06 | 109.0892 | 110.5665 | 124.7338 | 21.0436 | 0.002024 | 0.00183 | -151.92 | 0.23949 | 8.3998 | 0.003731 | 0.004609 | 0.004323 | 5.8e-05 | 0.001256 | 0.27096 | 1.3245 | 2.548 | 100.58 | 20.66 | 19.165 | 70.3037 | 111.4969 | 65.6532 | 0.004646 | 1.042 | 0.024938 | 23.9319 | -0.12032 | 0.14247 | -6e-06 | -0.000473 | -0.000473 | -12.3796
2019-03-20 17:42 | 1553118132.0727 | -408 | 0.20951 | 0.000452 | 0 | 108.9912 | 110.522 | 124.5711 | 21.0437 | 0.001973 | 0.001835 | -151.9211 | 0.23914 | 8.4017 | 0.003782 | 0.005664 | 0.005558 | 5.6e-05 | 0.001258 | 0.27527 | 1.3189 | 2.5481 | 100.58 | 20.66 | 19.165 | 69.0961 | 112.721 | 65.7726 | 0.004646 | 1.0417 | 0.024917 | 23.9205 | -0.12476 | 0.14226 | -7e-06 | -0.000554 | -0.000554 | -10.5538
2019-03-20 17:42 | 1553118133.0718 | -407 | 0.20953 | 0.000452 | 1e-06 | 109.0588 | 110.4129 | 124.7493 | 21.0442 | 0.001941 | 0.001848 | -151.9218 | 0.2382 | 8.4027 | 0.003803 | 0.004852 | 0.004644 | 5.8e-05 | 0.001264 | 0.28335 | 1.3238 | 2.5481 | 100.58 | 20.66 | 19.164 | 68.3189 | 113.2382 | 66.0836 | 0.004646 | 1.0407 | 0.024924 | 23.9486 | -0.11038 | 0.1417 | -8e-06 | -0.000652 | -0.000652 | -8.9323
2019-03-20 17:42 | 1553118134.0718 | -406 | 0.20953 | 0.000451 | 0 | 109.0239 | 110.4488 | 124.6061 | 21.0445 | 0.001927 | 0.001788 | -151.9174 | 0.23906 | 8.4042 | 0.003799 | 0.005323 | 0.005015 | 6e-05 | 0.001266 | 0.27704 | 1.3382 | 2.5481 | 100.58 | 20.66 | 19.164 | 67.9901 | 113.1381 | 64.648 | 0.004646 | 1.0451 | 0.024962 | 23.8841 | -0.11615 | 0.14222 | -7e-06 | -0.000513 | -0.000513 | -11.386
2019-03-20 17:42 | 1553118135.0733 | -405 | 0.20948 | 0.000452 | 0 | 109.0355 | 110.4454 | 124.8178 | 21.0441 | 0.00199 | 0.00185 | -151.9217 | 0.23805 | 8.4028 | 0.003791 | 0.004657 | 0.004526 | 5.7e-05 | 0.001268 | 0.28313 | 1.3459 | 2.5482 | 100.58 | 20.66 | 19.163 | 69.4902 | 112.9599 | 66.1374 | 0.004645 | 1.0406 | 0.024929 | 23.9571 | -0.11492 | 0.14161 | -8e-06 | -0.000613 | -0.000613 | -9.4913
2019-03-20 17:42 | 1553118136.072 | -404 | 0.20953 | 0.000454 | 0 | 109.0245 | 110.5071 | 124.8014 | 21.0442 | 0.00199 | 0.001843 | -151.9189 | 0.23835 | 8.4026 | 0.003742 | 0.005301 | 0.005078 | 6e-05 | 0.001276 | 0.26985 | 1.3486 | 2.5481 | 100.58 | 20.66 | 19.162 | 69.5004 | 111.7487 | 65.969 | 0.004645 | 1.0411 | 0.024933 | 23.9495 | -0.12081 | 0.1418 | -1e-05 | -0.000794 | -0.000794 | -7.3363
2019-03-20 17:42 | 1553118137.073 | -403 | 0.20955 | 0.00045 | 0 | 109.0484 | 110.3475 | 124.6913 | 21.0444 | 0.001945 | 0.001808 | -151.9147 | 0.23839 | 8.4034 | 0.003741 | 0.005185 | 0.004966 | 6e-05 | 0.001276 | 0.29479 | 1.347 | 2.5482 | 100.58 | 20.65 | 19.162 | 68.4245 | 111.7338 | 65.1099 | 0.004642 | 1.0437 | 0.024954 | 23.9086 | -0.10592 | 0.14182 | -4e-06 | -0.000312 | -0.000312 | -18.6813
2019-03-20 17:42 | 1553118138.072 | -402 | 0.20951 | 0.000449 | 0 | 108.9851 | 110.4245 | 124.5345 | 21.0437 | 0.00192 | 0.001848 | -151.9136 | 0.23873 | 8.4029 | 0.003778 | 0.005073 | 0.004842 | 5.6e-05 | 0.001271 | 0.27312 | 1.3564 | 2.5483 | 100.58 | 20.65 | 19.161 | 67.8076 | 112.6221 | 66.0904 | 0.004642 | 1.0407 | 0.024902 | 23.9282 | -0.11736 | 0.14202 | -3e-06 | -0.000249 | -0.000249 | -23.4173
2019-03-20 17:42 | 1553118139.0735 | -401 | 0.20955 | 0.000452 | 0 | 109.0268 | 110.4951 | 124.7004 | 21.0455 | 0.001939 | 0.00182 | -151.923 | 0.23877 | 8.4028 | 0.003699 | 0.004915 | 0.004787 | 5.3e-05 | 0.001276 | 0.26648 | 1.3154 | 2.5482 | 100.58 | 20.65 | 19.161 | 68.2678 | 110.712 | 65.4075 | 0.004642 | 1.0428 | 0.024944 | 23.92 | -0.11966 | 0.14204 | -7e-06 | -0.00055 | -0.00055 | -10.6216
2019-03-20 17:42 | 1553118140.0731 | -400 | 0.2095 | 0.000448 | 0 | 109.0093 | 110.3681 | 124.581 | 21.0444 | 0.00195 | 0.00185 | -151.8995 | 0.23853 | 8.4009 | 0.003693 | 0.004843 | 0.004595 | 5.6e-05 | 0.001279 | 0.26864 | 1.3283 | 2.5481 | 100.58 | 20.66 | 19.163 | 68.5406 | 110.557 | 66.1262 | 0.004645 | 1.0406 | 0.024905 | 23.9339 | -0.1108 | 0.1419 | -2e-06 | -0.000155 | -0.000155 | -37.5591
2019-03-20 17:42 | 1553118141.0732 | -399 | 0.20955 | 0.000451 | 0 | 109.0158 | 110.5755 | 124.8741 | 21.0454 | 0.001936 | 0.00186 | -151.9008 | 0.23949 | 8.4027 | 0.003687 | 0.005056 | 0.004814 | 5.8e-05 | 0.001283 | 0.26128 | 1.3247 | 2.5482 | 100.58 | 20.66 | 19.164 | 68.1999 | 110.4208 | 66.383 | 0.004646 | 1.0398 | 0.024925 | 23.9711 | -0.12706 | 0.14247 | -6e-06 | -0.000466 | -0.000466 | -12.5748
2019-03-20 17:42 | 1553118142.0723 | -398 | 0.20952 | 0.000449 | 1e-06 | 108.9886 | 110.5075 | 125.1295 | 21.0454 | 0.001957 | 0.001845 | -151.8954 | 0.23891 | 8.4012 | 0.003823 | 0.004985 | 0.004758 | 5.7e-05 | 0.00129 | 0.27865 | 1.3445 | 2.5479 | 100.58 | 20.66 | 19.167 | 68.7055 | 113.7387 | 66.0218 | 0.004646 | 1.0409 | 0.024964 | 23.9829 | -0.12379 | 0.14213 | -5e-06 | -0.000383 | -0.000383 | -15.243
2019-03-20 17:42 | 1553118143.0724 | -397 | 0.20953 | 0.000447 | 0 | 108.9944 | 110.2645 | 125.1293 | 21.045 | 0.001923 | 0.001825 | -151.8988 | 0.23802 | 8.4034 | 0.003687 | 0.005367 | 0.005139 | 5.4e-05 | 0.001289 | 0.26217 | 1.362 | 2.5481 | 100.58 | 20.66 | 19.168 | 67.8906 | 110.4109 | 65.5367 | 0.004647 | 1.0424 | 0.024982 | 23.9657 | -0.10363 | 0.1416 | -2e-06 | -0.000134 | -0.000134 | -43.3668
2019-03-20 17:42 | 1553118144.0727 | -396 | 0.20951 | 0.000449 | 1e-06 | 108.8922 | 110.438 | 125.4274 | 21.0443 | 0.002019 | 0.001837 | -151.8885 | 0.23906 | 8.4036 | 0.003747 | 0.005279 | 0.005071 | 5.6e-05 | 0.001289 | 0.29662 | 1.3204 | 2.5481 | 100.58 | 20.66 | 19.17 | 70.1877 | 111.8814 | 65.8122 | 0.004647 | 1.0415 | 0.025001 | 24.004 | -0.12607 | 0.14222 | -5e-06 | -0.000402 | -0.000402 | -14.5264
2019-03-20 17:42 | 1553118145.0715 | -395 | 0.20953 | 0.000449 | 0 | 108.8977 | 110.3449 | 125.5403 | 21.0451 | 0.001929 | 0.001812 | -151.8839 | 0.23852 | 8.403 | 0.003672 | 0.004832 | 0.004617 | 5.6e-05 | 0.00129 | 0.28224 | 1.3422 | 2.5482 | 100.58 | 20.66 | 19.172 | 68.0227 | 110.0505 | 65.2097 | 0.004648 | 1.0434 | 0.025035 | 23.9934 | -0.11808 | 0.14189 | -4e-06 | -0.00031 | -0.00031 | -18.8033
2019-03-20 17:42 | 1553118146.0725 | -394 | 0.20952 | 0.000452 | 1e-06 | 108.9317 | 110.4331 | 125.6379 | 21.045 | 0.001971 | 0.001896 | -151.8921 | 0.23852 | 8.4025 | 0.003664 | 0.005184 | 0.004991 | 5.6e-05 | 0.001297 | 0.26145 | 1.3134 | 2.5481 | 100.59 | 20.66 | 19.174 | 69.0387 | 109.8501 | 67.233 | 0.004648 | 1.0372 | 0.02497 | 24.0744 | -0.12244 | 0.14189 | -8e-06 | -0.000641 | -0.000641 | -9.0997
2019-03-20 17:42 | 1553118147.073 | -393 | 0.20952 | 0.000449 | 0 | 108.9233 | 110.3524 | 125.7931 | 21.0465 | 0.001959 | 0.001814 | -151.9178 | 0.23883 | 8.4042 | 0.003643 | 0.004814 | 0.004533 | 5.6e-05 | 0.001296 | 0.29374 | 1.3363 | 2.5481 | 100.59 | 20.66 | 19.178 | 68.7493 | 109.3587 | 65.2712 | 0.004649 | 1.0432 | 0.025058 | 24.0197 | -0.11659 | 0.14208 | -4e-06 | -0.000302 | -0.000302 | -19.3271
2019-03-20 17:42 | 1553118148.0716 | -392 | 0.20952 | 0.000449 | 0 | 108.9256 | 110.2331 | 126.4902 | 21.0452 | 0.002022 | 0.001795 | -151.9315 | 0.23772 | 8.4008 | 0.003673 | 0.005354 | 0.005158 | 5.6e-05 | 0.001298 | 0.29407 | 1.3597 | 2.5482 | 100.59 | 20.66 | 19.179 | 70.274 | 110.0744 | 64.8087 | 0.004649 | 1.0446 | 0.025144 | 24.0697 | -0.10673 | 0.14142 | -4e-06 | -0.000305 | -0.000305 | -19.0851
2019-03-20 17:42 | 1553118149.0716 | -391 | 0.20952 | 0.000451 | 1e-06 | 108.9032 | 110.392 | 126.8406 | 21.0461 | 0.001961 | 0.001864 | -151.9251 | 0.23841 | 8.4006 | 0.003689 | 0.004526 | 0.004273 | 5.2e-05 | 0.001289 | 0.28064 | 1.3326 | 2.5482 | 100.59 | 20.66 | 19.18 | 68.8083 | 110.4774 | 66.4602 | 0.004649 | 1.0396 | 0.025118 | 24.1619 | -0.12145 | 0.14183 | -6e-06 | -0.000516 | -0.000516 | -11.2905
2019-03-20 17:42 | 1553118150.0731 | -390 | 0.20952 | 0.000448 | 1e-06 | 108.9061 | 110.2731 | 127.1725 | 21.0466 | 0.001987 | 0.001815 | -151.9306 | 0.23772 | 8.4054 | 0.0037 | 0.004666 | 0.004479 | 5.6e-05 | 0.001288 | 0.27295 | 1.3391 | 2.5482 | 100.59 | 20.66 | 19.178 | 69.4186 | 110.7398 | 65.2977 | 0.004649 | 1.0431 | 0.025194 | 24.152 | -0.11157 | 0.14142 | -4e-06 | -0.000338 | -0.000338 | -17.207
2019-03-20 17:42 | 1553118151.0728 | -389 | 0.20951 | 0.00045 | 1e-06 | 108.9349 | 110.5273 | 127.1521 | 21.0457 | 0.001968 | 0.001876 | -151.9319 | 0.23691 | 8.4046 | 0.003632 | 0.005494 | 0.005281 | 5.6e-05 | 0.001289 | 0.27831 | 1.3202 | 2.5481 | 100.59 | 20.66 | 19.176 | 68.9584 | 109.0899 | 66.7563 | 0.004648 | 1.0387 | 0.025138 | 24.2021 | -0.1298 | 0.14094 | -6e-06 | -0.000489 | -0.000489 | -11.8455
2019-03-20 17:43 | 1553118152.0722 | -388 | 0.20952 | 0.000453 | 0 | 108.7765 | 110.0489 | 127.0165 | 21.0472 | 0.001989 | 0.001891 | -151.9306 | 0.2387 | 8.4051 | 0.003669 | 0.005622 | 0.005437 | 5.5e-05 | 0.001291 | 0.2891 | 1.3186 | 2.5481 | 100.59 | 20.66 | 19.171 | 69.4752 | 109.9879 | 67.1204 | 0.004647 | 1.0375 | 0.025111 | 24.2021 | -0.10402 | 0.142 | -9e-06 | -0.000734 | -0.000734 | -7.9535
2019-03-20 17:43 | 1553118153.0733 | -387 | 0.20951 | 0.00045 | 0 | 108.7858 | 110.1958 | 126.8031 | 21.0458 | 0.002019 | 0.00184 | -151.9094 | 0.23894 | 8.4014 | 0.003634 | 0.004778 | 0.004564 | 5.7e-05 | 0.001289 | 0.30469 | 1.3156 | 2.5481 | 100.59 | 20.66 | 19.167 | 70.1992 | 109.1216 | 65.9019 | 0.004646 | 1.0413 | 0.025135 | 24.1384 | -0.11518 | 0.14215 | -5e-06 | -0.000414 | -0.000414 | -14.0995
2019-03-20 17:43 | 1553118154.0729 | -386 | 0.20952 | 0.000449 | 0 | 108.857 | 110.2044 | 126.6482 | 21.0462 | 0.001916 | 0.001849 | -151.9248 | 0.2388 | 8.4027 | 0.003615 | 0.0049 | 0.004637 | 5.6e-05 | 0.001292 | 0.26056 | 1.3516 | 2.5483 | 100.59 | 20.66 | 19.162 | 67.71 | 108.675 | 66.1075 | 0.004645 | 1.0406 | 0.025112 | 24.131 | -0.11003 | 0.14206 | -4e-06 | -0.000347 | -0.000347 | -16.8093
2019-03-20 17:43 | 1553118155.0723 | -385 | 0.20951 | 0.000454 | 0 | 108.8127 | 110.1741 | 126.4167 | 21.0474 | 0.001941 | 0.001866 | -151.9335 | 0.23869 | 8.4032 | 0.003665 | 0.005655 | 0.00547 | 6e-05 | 0.001295 | 0.28296 | 1.3508 | 2.5479 | 100.59 | 20.66 | 19.152 | 68.3252 | 109.8792 | 66.5216 | 0.004642 | 1.0394 | 0.025073 | 24.1237 | -0.11121 | 0.142 | -9e-06 | -0.000713 | -0.000713 | -8.18
2019-03-20 17:43 | 1553118156.0734 | -384 | 0.20955 | 0.000453 | 0 | 108.8256 | 110.2103 | 126.267 | 21.047 | 0.001955 | 0.001873 | -151.9011 | 0.23902 | 8.4032 | 0.003607 | 0.005022 | 0.004819 | 6e-05 | 0.001293 | 0.27959 | 1.3373 | 2.5481 | 100.59 | 20.66 | 19.149 | 68.6531 | 108.478 | 66.6888 | 0.004641 | 1.0389 | 0.025052 | 24.1153 | -0.11309 | 0.14219 | -9e-06 | -0.000661 | -0.000661 | -8.8462
2019-03-20 17:43 | 1553118157.0731 | -383 | 0.20951 | 0.000448 | 1e-06 | 108.8111 | 110.1472 | 126.1947 | 21.0474 | 0.001963 | 0.001814 | -151.8842 | 0.23818 | 8.4048 | 0.003728 | 0.005255 | 0.005051 | 5.9e-05 | 0.001289 | 0.27472 | 1.3462 | 2.548 | 100.59 | 20.66 | 19.146 | 68.8523 | 111.4289 | 65.2674 | 0.004641 | 1.0432 | 0.025098 | 24.0579 | -0.10916 | 0.14169 | -3e-06 | -0.000299 | -0.000299 | -19.446
2019-03-20 17:43 | 1553118158.0727 | -382 | 0.2095 | 0.000451 | 1e-06 | 108.7677 | 110.0164 | 125.8527 | 21.048 | 0.001959 | 0.001831 | -151.8821 | 0.23866 | 8.4024 | 0.003759 | 0.005274 | 0.00509 | 5.8e-05 | 0.001292 | 0.28854 | 1.3711 | 2.5481 | 100.59 | 20.66 | 19.142 | 68.7657 | 112.1714 | 65.6753 | 0.00464 | 1.042 | 0.025049 | 24.0398 | -0.1021 | 0.14198 | -6e-06 | -0.000526 | -0.000526 | -11.0869
2019-03-20 17:43 | 1553118159.0726 | -381 | 0.20952 | 0.000452 | 1e-06 | 108.7845 | 110.0754 | 125.7295 | 21.048 | 0.001989 | 0.001841 | -151.8846 | 0.23932 | 8.4012 | 0.003673 | 0.005464 | 0.005234 | 6.1e-05 | 0.001289 | 0.28622 | 1.3374 | 2.5481 | 100.59 | 20.66 | 19.141 | 69.469 | 110.0772 | 65.9243 | 0.00464 | 1.0412 | 0.025027 | 24.0368 | -0.10551 | 0.14237 | -8e-06 | -0.000694 | -0.000694 | -8.4274
2019-03-20 17:43 | 1553118160.0722 | -380 | 0.20953 | 0.000447 | 0 | 108.7748 | 110.176 | 125.6431 | 21.0471 | 0.00197 | 0.001856 | -151.8958 | 0.23833 | 8.4037 | 0.003732 | 0.004868 | 0.004748 | 5.8e-05 | 0.001282 | 0.29159 | 1.3634 | 2.5481 | 100.59 | 20.66 | 19.141 | 69.03 | 111.5138 | 66.277 | 0.00464 | 1.0401 | 0.025006 | 24.0411 | -0.11448 | 0.14178 | -2e-06 | -0.000164 | -0.000164 | -35.639
2019-03-20 17:43 | 1553118161.0714 | -379 | 0.20952 | 0.000455 | 0 | 108.6694 | 110.081 | 125.7325 | 21.0484 | 0.001902 | 0.001823 | -151.8902 | 0.23891 | 8.4007 | 0.003713 | 0.005326 | 0.005045 | 5.8e-05 | 0.001284 | 0.28456 | 1.3751 | 2.5481 | 100.59 | 20.66 | 19.143 | 67.376 | 111.0455 | 65.4791 | 0.00464 | 1.0426 | 0.025044 | 24.0213 | -0.11544 | 0.14213 | -1.2e-05 | -0.000885 | -0.000885 | -6.6039
2019-03-20 17:43 | 1553118162.0717 | -378 | 0.2095 | 0.000455 | 1e-06 | 108.6474 | 110.0806 | 125.964 | 21.0478 | 0.001975 | 0.001843 | -151.8987 | 0.23909 | 8.4032 | 0.003685 | 0.005719 | 0.005549 | 5.8e-05 | 0.00128 | 0.27544 | 1.3727 | 2.5482 | 100.59 | 20.66 | 19.144 | 69.1327 | 110.3641 | 65.9608 | 0.00464 | 1.0411 | 0.025049 | 24.0605 | -0.11722 | 0.14224 | -1.2e-05 | -0.000999 | -0.000999 | -5.8538
2019-03-20 17:43 | 1553118163.0714 | -377 | 0.20952 | 0.000447 | 0 | 108.7089 | 109.9853 | 126.0539 | 21.0477 | 0.001956 | 0.001813 | -151.9044 | 0.23889 | 8.4025 | 0.003731 | 0.005221 | 0.00496 | 6.3e-05 | 0.001282 | 0.28456 | 1.3059 | 2.5483 | 100.59 | 20.66 | 19.145 | 68.6858 | 111.5048 | 65.2545 | 0.004641 | 1.0433 | 0.025084 | 24.044 | -0.10442 | 0.14211 | -1e-06 | -0.0001 | -0.0001 | -58.4237
2019-03-20 17:43 | 1553118164.0725 | -376 | 0.20949 | 0.000452 | 0 | 108.6335 | 110.075 | 126.0205 | 21.0482 | 0.001936 | 0.001782 | -151.8963 | 0.23833 | 8.4026 | 0.003718 | 0.005371 | 0.005156 | 5.6e-05 | 0.001279 | 0.27616 | 1.3145 | 2.5481 | 100.59 | 20.66 | 19.147 | 68.2144 | 111.1802 | 64.4955 | 0.004641 | 1.0456 | 0.025109 | 24.0139 | -0.11791 | 0.14178 | -8e-06 | -0.000635 | -0.000635 | -9.181
2019-03-20 17:43 | 1553118165.0725 | -375 | 0.20949 | 0.000452 | 0 | 108.6293 | 110.235 | 126.1118 | 21.049 | 0.001973 | 0.001839 | -151.9026 | 0.23824 | 8.4014 | 0.003723 | 0.00513 | 0.004937 | 5.8e-05 | 0.001276 | 0.25105 | 1.3021 | 2.5481 | 100.59 | 20.66 | 19.151 | 69.0863 | 111.2962 | 65.8745 | 0.004642 | 1.0414 | 0.025067 | 24.0716 | -0.13125 | 0.14173 | -7e-06 | -0.000555 | -0.000555 | -10.4909
2019-03-20 17:43 | 1553118166.0725 | -374 | 0.20953 | 0.000451 | 1e-06 | 108.6759 | 109.9876 | 126.2531 | 21.0489 | 0.001939 | 0.001786 | -151.9259 | 0.23848 | 8.4013 | 0.003706 | 0.005486 | 0.00531 | 5.9e-05 | 0.001273 | 0.28888 | 1.3221 | 2.5481 | 100.59 | 20.66 | 19.152 | 68.2783 | 110.8809 | 64.5866 | 0.004642 | 1.0453 | 0.025129 | 24.0393 | -0.10731 | 0.14187 | -7e-06 | -0.00054 | -0.00054 | -10.8015
2019-03-20 17:43 | 1553118167.0716 | -373 | 0.2095 | 0.000451 | 1e-06 | 108.605 | 110.0321 | 126.0954 | 21.0487 | 0.001954 | 0.001808 | -151.9307 | 0.23952 | 8.4036 | 0.003677 | 0.004778 | 0.00453 | 5.9e-05 | 0.001271 | 0.28854 | 1.3158 | 2.5481 | 100.59 | 20.66 | 19.153 | 68.6298 | 110.1794 | 65.1296 | 0.004642 | 1.0436 | 0.025093 | 24.0436 | -0.11676 | 0.14249 | -7e-06 | -0.000561 | -0.000561 | -10.4345
2019-03-20 17:43 | 1553118168.0717 | -372 | 0.20952 | 0.000451 | 1e-06 | 108.6464 | 109.9326 | 125.7828 | 21.0479 | 0.001948 | 0.001817 | -151.9343 | 0.23926 | 8.403 | 0.003744 | 0.005306 | 0.005132 | 5.6e-05 | 0.001278 | 0.29374 | 1.3162 | 2.5481 | 100.59 | 20.66 | 19.153 | 68.487 | 111.8005 | 65.3459 | 0.004642 | 1.043 | 0.025054 | 24.0214 | -0.10527 | 0.14234 | -7e-06 | -0.000592 | -0.000592 | -9.883
2019-03-20 17:43 | 1553118169.0728 | -371 | 0.20954 | 0.000454 | 1e-06 | 108.5575 | 109.9207 | 125.6821 | 21.0484 | 0.001993 | 0.001851 | -151.9368 | 0.23887 | 8.4017 | 0.003728 | 0.004744 | 0.00452 | 5.8e-05 | 0.001282 | 0.28191 | 1.3206 | 2.5482 | 100.59 | 20.66 | 19.152 | 69.5754 | 111.4163 | 66.1659 | 0.004642 | 1.0405 | 0.025014 | 24.0409 | -0.11162 | 0.1421 | -1.1e-05 | -0.000871 | -0.000871 | -6.7049
2019-03-20 17:43 | 1553118170.0718 | -370 | 0.20951 | 0.000448 | 1e-06 | 108.5268 | 110.0377 | 125.6314 | 21.0489 | 0.001951 | 0.001841 | -151.9341 | 0.23896 | 8.4007 | 0.00385 | 0.005216 | 0.004955 | 5.7e-05 | 0.001276 | 0.26913 | 1.3325 | 2.5482 | 100.59 | 20.66 | 19.15 | 68.5745 | 114.4021 | 65.9282 | 0.004642 | 1.0412 | 0.025017 | 24.0276 | -0.12366 | 0.14216 | -3e-06 | -0.000302 | -0.000302 | -19.3316
2019-03-20 17:43 | 1553118171.0729 | -369 | 0.20954 | 0.000449 | 1e-06 | 108.5763 | 109.8261 | 125.5748 | 21.0488 | 0.00196 | 0.001895 | -151.9379 | 0.23963 | 8.4039 | 0.00383 | 0.005018 | 0.004844 | 6.2e-05 | 0.001279 | 0.29535 | 1.3582 | 2.5482 | 100.59 | 20.66 | 19.149 | 68.7741 | 113.9049 | 67.2078 | 0.004641 | 1.0373 | 0.024965 | 24.0675 | -0.10238 | 0.14256 | -5e-06 | -0.000412 | -0.000412 | -14.2271
2019-03-20 17:43 | 1553118172.0731 | -368 | 0.20953 | 0.000443 | 1e-06 | 108.5963 | 109.985 | 125.453 | 21.0485 | 0.001906 | 0.001875 | -151.9234 | 0.23887 | 8.4014 | 0.003833 | 0.005152 | 0.004914 | 5.8e-05 | 0.001282 | 0.26233 | 1.3849 | 2.5481 | 100.59 | 20.66 | 19.147 | 67.4728 | 113.9695 | 66.7271 | 0.004641 | 1.0387 | 0.02497 | 24.0388 | -0.11365 | 0.14211 | 3e-06 | 0.000156 | 0.000156 | 37.5389
2019-03-20 17:43 | 1553118173.0719 | -367 | 0.20949 | 0.000448 | 1e-06 | 108.5012 | 109.8971 | 125.5654 | 21.0484 | 0.00207 | 0.001837 | -151.9364 | 0.23949 | 8.4025 | 0.003831 | 0.004959 | 0.004724 | 5.6e-05 | 0.001291 | 0.27815 | 1.3245 | 2.5481 | 100.58 | 20.66 | 19.142 | 71.4165 | 113.9243 | 65.8335 | 0.00464 | 1.0415 | 0.025014 | 24.0179 | -0.11434 | 0.14247 | -3e-06 | -0.000263 | -0.000263 | -22.2709
2019-03-20 17:43 | 1553118174.0729 | -366 | 0.20952 | 0.000449 | 1e-06 | 108.4886 | 109.8893 | 125.4825 | 21.0492 | 0.001938 | 0.001825 | -151.932 | 0.23904 | 8.4013 | 0.003853 | 0.005281 | 0.005028 | 5.7e-05 | 0.001285 | 0.29142 | 1.3242 | 2.5481 | 100.58 | 20.66 | 19.141 | 68.2536 | 114.457 | 65.543 | 0.00464 | 1.0424 | 0.025017 | 23.9997 | -0.11474 | 0.1422 | -5e-06 | -0.000386 | -0.000386 | -15.1253
2019-03-20 17:43 | 1553118175.0734 | -365 | 0.20952 | 0.000445 | 1e-06 | 108.5294 | 109.7798 | 125.5469 | 21.0483 | 0.001983 | 0.001868 | -151.9203 | 0.23878 | 8.4031 | 0.003758 | 0.005001 | 0.004753 | 6e-05 | 0.001283 | 0.27815 | 1.3324 | 2.5482 | 100.58 | 20.66 | 19.139 | 69.3227 | 112.1591 | 66.5653 | 0.00464 | 1.0392 | 0.024985 | 24.0421 | -0.10246 | 0.14205 | 0 | -8.8e-05 | -8.8e-05 | -66.1272
2019-03-20 17:43 | 1553118176.0731 | -364 | 0.20952 | 0.000451 | 1e-06 | 108.5313 | 109.9009 | 125.4701 | 21.0477 | 0.001977 | 0.001829 | -151.9157 | 0.23844 | 8.4031 | 0.003764 | 0.004895 | 0.00461 | 6e-05 | 0.001284 | 0.24818 | 1.3627 | 2.5481 | 100.58 | 20.66 | 19.136 | 69.1779 | 112.284 | 65.6353 | 0.004639 | 1.0421 | 0.025012 | 24.0018 | -0.11217 | 0.14185 | -6e-06 | -0.000532 | -0.000532 | -10.9623
2019-03-20 17:43 | 1553118177.0725 | -363 | 0.20953 | 0.000449 | 1e-06 | 108.4747 | 109.6097 | 125.648 | 21.0486 | 0.001965 | 0.001853 | -151.9189 | 0.23861 | 8.4023 | 0.003739 | 0.005597 | 0.005322 | 5.8e-05 | 0.001283 | 0.28744 | 1.3056 | 2.5482 | 100.58 | 20.66 | 19.133 | 68.8949 | 111.6763 | 66.2128 | 0.004638 | 1.0403 | 0.025008 | 24.0393 | -0.093104 | 0.14195 | -5e-06 | -0.000416 | -0.000416 | -14.0318
2019-03-20 17:43 | 1553118178.0715 | -362 | 0.20952 | 0.000447 | 1e-06 | 108.445 | 109.7858 | 126.033 | 21.0491 | 0.001928 | 0.001841 | -151.9297 | 0.23831 | 8.401 | 0.003708 | 0.004801 | 0.004608 | 5.9e-05 | 0.001282 | 0.27887 | 1.3175 | 2.5482 | 100.58 | 20.66 | 19.128 | 68.0186 | 110.9341 | 65.9246 | 0.004637 | 1.0412 | 0.025057 | 24.0658 | -0.10991 | 0.14177 | -2e-06 | -0.000191 | -0.000191 | -30.4276
2019-03-20 17:43 | 1553118179.0712 | -361 | 0.20953 | 0.000447 | 0 | 108.4482 | 109.837 | 126.1541 | 21.0492 | 0.001916 | 0.00181 | -151.9183 | 0.23792 | 8.4031 | 0.003714 | 0.005293 | 0.005064 | 5.3e-05 | 0.001277 | 0.28694 | 1.3148 | 2.5481 | 100.58 | 20.66 | 19.124 | 67.7286 | 111.081 | 65.1764 | 0.004636 | 1.0435 | 0.025097 | 24.0508 | -0.11381 | 0.14154 | -1e-06 | -0.000125 | -0.000125 | -46.412
2019-03-20 17:43 | 1553118180.0713 | -360 | 0.20953 | 0.000448 | 1e-06 | 108.3764 | 109.6034 | 126.2562 | 21.05 | 0.001986 | 0.001745 | -151.9262 | 0.23952 | 8.4037 | 0.003637 | 0.005391 | 0.005176 | 5.8e-05 | 0.001274 | 0.31636 | 1.3195 | 2.5481 | 100.58 | 20.66 | 19.122 | 69.4171 | 109.2164 | 63.6195 | 0.004635 | 1.0483 | 0.025165 | 24.0051 | -0.1007 | 0.14249 | -4e-06 | -0.000319 | -0.000319 | -18.3398
2019-03-20 17:43 | 1553118181.072 | -359 | 0.20953 | 0.000445 | 0 | 108.3919 | 109.6194 | 126.2055 | 21.0488 | 0.00197 | 0.001843 | -151.916 | 0.23936 | 8.4037 | 0.003695 | 0.005292 | 0.005131 | 5.6e-05 | 0.001271 | 0.29518 | 1.3277 | 2.548 | 100.58 | 20.66 | 19.116 | 69.016 | 110.6087 | 65.9657 | 0.004634 | 1.0411 | 0.025073 | 24.0838 | -0.10073 | 0.14239 | 1e-06 | 7e-05 | 7e-05 | 83.9957
2019-03-20 17:43 | 1553118182.0724 | -358 | 0.2095 | 0.000446 | 0 | 108.3757 | 109.6695 | 126.082 | 21.0486 | 0.001963 | 0.001855 | -151.9179 | 0.23961 | 8.4037 | 0.003668 | 0.004957 | 0.00469 | 5.6e-05 | 0.00127 | 0.30143 | 1.3341 | 2.5481 | 100.58 | 20.66 | 19.113 | 68.8612 | 109.972 | 66.245 | 0.004633 | 1.0402 | 0.02505 | 24.0819 | -0.10615 | 0.14255 | 0 | -4.2e-05 | -4.2e-05 | -139.1989
2019-03-20 17:43 | 1553118183.0715 | -357 | 0.20951 | 0.000445 | 0 | 108.3961 | 109.8515 | 126.1259 | 21.049 | 0.001986 | 0.00184 | -151.9188 | 0.23848 | 8.4017 | 0.003692 | 0.005005 | 0.00474 | 5.6e-05 | 0.001262 | 0.27632 | 1.3587 | 2.5482 | 100.59 | 20.66 | 19.109 | 69.4084 | 110.5466 | 65.8939 | 0.004632 | 1.0413 | 0.025068 | 24.0736 | -0.1193 | 0.14187 | 1e-06 | 3.6e-05 | 3.6e-05 | 164.1462
2019-03-20 17:43 | 1553118184.0729 | -356 | 0.20951 | 0.000445 | 0 | 108.3625 | 109.5836 | 126.2452 | 21.0491 | 0.001936 | 0.001829 | -151.9374 | 0.23903 | 8.401 | 0.0037 | 0.00527 | 0.005041 | 5.8e-05 | 0.00126 | 0.27151 | 1.3605 | 2.5482 | 100.58 | 20.66 | 19.108 | 68.1958 | 110.7299 | 65.6234 | 0.004632 | 1.0421 | 0.02509 | 24.0754 | -0.10023 | 0.1422 | 1e-06 | 7.3e-05 | 7.3e-05 | 79.8843
2019-03-20 17:43 | 1553118185.0719 | -355 | 0.20951 | 0.00045 | 1e-06 | 108.3964 | 109.6019 | 126.2928 | 21.0486 | 0.002006 | 0.001872 | -151.9343 | 0.23908 | 8.4034 | 0.003734 | 0.004564 | 0.004393 | 5.7e-05 | 0.001267 | 0.26272 | 1.3858 | 2.5481 | 100.58 | 20.66 | 19.108 | 69.8868 | 111.5697 | 66.6663 | 0.004632 | 1.0389 | 0.025056 | 24.117 | -0.098919 | 0.14223 | -6e-06 | -0.000523 | -0.000523 | -11.1801
2019-03-20 17:43 | 1553118186.0726 | -354 | 0.2095 | 0.000446 | 0 | 108.4294 | 109.6 | 126.3872 | 21.0491 | 0.001976 | 0.001821 | -151.9308 | 0.23883 | 8.4035 | 0.003689 | 0.004901 | 0.004716 | 6e-05 | 0.00127 | 0.27273 | 1.3442 | 2.5479 | 100.59 | 20.65 | 19.109 | 69.1691 | 110.4663 | 65.4375 | 0.004629 | 1.0427 | 0.025111 | 24.0823 | -0.096045 | 0.14208 | 0 | 2.9e-05 | 2.9e-05 | 203.258
2019-03-20 17:43 | 1553118187.0731 | -353 | 0.2095 | 0.000448 | 1e-06 | 108.3586 | 109.532 | 126.2475 | 21.049 | 0.001971 | 0.001804 | -151.9219 | 0.23887 | 8.4037 | 0.003645 | 0.004843 | 0.004591 | 5.9e-05 | 0.001272 | 0.27815 | 1.3571 | 2.5482 | 100.59 | 20.65 | 19.111 | 69.0395 | 109.3991 | 65.0315 | 0.004629 | 1.044 | 0.025112 | 24.0546 | -0.096336 | 0.14211 | -4e-06 | -0.000332 | -0.000332 | -17.5913
2019-03-20 17:43 | 1553118188.0731 | -352 | 0.20944 | 0.000452 | 1e-06 | 108.4728 | 109.7413 | 125.9694 | 21.0483 | 0.001932 | 0.001835 | -151.9127 | 0.23865 | 8.4024 | 0.003637 | 0.004336 | 0.004112 | 5.8e-05 | 0.00127 | 0.26609 | 1.3429 | 2.548 | 100.59 | 20.65 | 19.113 | 68.1061 | 109.1996 | 65.767 | 0.00463 | 1.0417 | 0.025057 | 24.0542 | -0.10399 | 0.14197 | -8e-06 | -0.000716 | -0.000716 | -8.1521
2019-03-20 17:43 | 1553118189.0718 | -351 | 0.20925 | 0.000449 | 1e-06 | 108.3677 | 109.8018 | 125.6799 | 21.0497 | 0.00192 | 0.001831 | -151.9215 | 0.23881 | 8.4002 | 0.003613 | 0.005139 | 0.004883 | 5.8e-05 | 0.00127 | 0.26576 | 1.3031 | 2.548 | 100.58 | 20.65 | 19.121 | 67.8297 | 108.6328 | 65.6887 | 0.004632 | 1.0419 | 0.025031 | 24.0237 | -0.1176 | 0.14207 | -4e-06 | -0.000334 | -0.000334 | -17.4777
2019-03-20 17:43 | 1553118190.0729 | -350 | 0.20919 | 0.000449 | 1e-06 | 108.3272 | 109.5267 | 125.2577 | 21.0489 | 0.001952 | 0.001785 | -151.9093 | 0.23885 | 8.4002 | 0.003605 | 0.00452 | 0.00426 | 5.6e-05 | 0.001264 | 0.27455 | 1.3057 | 2.5481 | 100.58 | 20.65 | 19.126 | 68.5769 | 108.4217 | 64.5765 | 0.004633 | 1.0454 | 0.02503 | 23.944 | -0.098499 | 0.14209 | -5e-06 | -0.000425 | -0.000425 | -13.734
2019-03-20 17:43 | 1553118191.073 | -349 | 0.20918 | 0.000448 | 0 | 108.3618 | 109.8336 | 125.0063 | 21.049 | 0.002004 | 0.001798 | -151.9254 | 0.23797 | 8.4046 | 0.003601 | 0.004992 | 0.004796 | 6.1e-05 | 0.001264 | 0.28783 | 1.3021 | 2.5481 | 100.58 | 20.65 | 19.134 | 69.8435 | 108.3406 | 64.8826 | 0.004635 | 1.0444 | 0.024994 | 23.9308 | -0.12067 | 0.14157 | -3e-06 | -0.000275 | -0.000275 | -21.1329
2019-03-20 17:43 | 1553118192.0719 | -348 | 0.20919 | 0.000452 | 0 | 108.431 | 109.662 | 125.2596 | 21.0504 | 0.001951 | 0.001825 | -151.9314 | 0.23849 | 8.4008 | 0.003626 | 0.005168 | 0.004923 | 5.9e-05 | 0.001262 | 0.30574 | 1.3129 | 2.5481 | 100.58 | 20.65 | 19.138 | 68.5706 | 108.9337 | 65.5278 | 0.004636 | 1.0424 | 0.024995 | 23.9778 | -0.10097 | 0.14188 | -7e-06 | -0.000543 | -0.000543 | -10.7431
2019-03-20 17:43 | 1553118193.0714 | -347 | 0.20919 | 0.000452 | 0 | 108.3958 | 109.7439 | 125.4187 | 21.049 | 0.001993 | 0.001912 | -151.9165 | 0.23898 | 8.4013 | 0.003703 | 0.004736 | 0.004528 | 5.8e-05 | 0.001262 | 0.29767 | 1.3018 | 2.548 | 100.58 | 20.65 | 19.143 | 69.577 | 110.8132 | 67.6147 | 0.004638 | 1.036 | 0.024934 | 24.0669 | -0.11056 | 0.14217 | -7e-06 | -0.000521 | -0.000521 | -11.2095
2019-03-20 17:43 | 1553118194.0711 | -346 | 0.20934 | 0.000447 | 0 | 108.4896 | 109.7734 | 125.5497 | 21.0493 | 0.001933 | 0.001809 | -151.8894 | 0.23803 | 8.4037 | 0.003604 | 0.005265 | 0.005086 | 5.8e-05 | 0.001257 | 0.29535 | 1.2958 | 2.548 | 100.58 | 20.65 | 19.145 | 68.1311 | 108.4165 | 65.1445 | 0.004638 | 1.0436 | 0.025038 | 23.992 | -0.10523 | 0.1416 | -1e-06 | -6.1e-05 | -6.1e-05 | -95.1328
2019-03-20 17:43 | 1553118195.071 | -345 | 0.20953 | 0.000449 | 1e-06 | 108.3961 | 109.794 | 125.3457 | 21.0495 | 0.00199 | 0.001864 | -151.8817 | 0.23868 | 8.402 | 0.003599 | 0.004674 | 0.004422 | 5.9e-05 | 0.00126 | 0.28926 | 1.2954 | 2.5481 | 100.58 | 20.65 | 19.145 | 69.4978 | 108.2969 | 66.4791 | 0.004638 | 1.0395 | 0.024969 | 24.0197 | -0.11461 | 0.14199 | -4e-06 | -0.00034 | -0.00034 | -17.1868
2019-03-20 17:43 | 1553118196.0715 | -344 | 0.2095 | 0.000443 | 0 | 108.4146 | 109.711 | 125.2063 | 21.0491 | 0.00196 | 0.001829 | -151.8903 | 0.23852 | 8.402 | 0.003616 | 0.005065 | 0.004856 | 5.9e-05 | 0.001265 | 0.29574 | 1.2846 | 2.5481 | 100.58 | 20.65 | 19.146 | 68.7738 | 108.7049 | 65.6235 | 0.004638 | 1.0421 | 0.024986 | 23.9761 | -0.10633 | 0.14189 | 4e-06 | 0.000341 | 0.000341 | 17.1266
2019-03-20 17:43 | 1553118197.0713 | -343 | 0.20953 | 0.000452 | 0 | 108.4265 | 109.5757 | 125.161 | 21.0488 | 0.001975 | 0.001889 | -151.882 | 0.23943 | 8.4012 | 0.0036 | 0.004737 | 0.004485 | 5.9e-05 | 0.001266 | 0.27815 | 1.3092 | 2.5479 | 100.58 | 20.65 | 19.148 | 69.1362 | 108.316 | 67.0754 | 0.004639 | 1.0377 | 0.024928 | 24.0231 | -0.094302 | 0.14244 | -7e-06 | -0.00051 | -0.00051 | -11.4721
2019-03-20 17:43 | 1553118198.0724 | -342 | 0.20951 | 0.000447 | 1e-06 | 108.4973 | 109.8485 | 125.291 | 21.05 | 0.001945 | 0.001874 | -151.8546 | 0.23863 | 8.4015 | 0.003689 | 0.004768 | 0.004623 | 5.6e-05 | 0.001271 | 0.30359 | 1.3191 | 2.5481 | 100.58 | 20.65 | 19.148 | 68.4267 | 110.4735 | 66.7167 | 0.004639 | 1.0388 | 0.024954 | 24.0229 | -0.1107 | 0.14196 | -2e-06 | -0.000223 | -0.000223 | -26.156
2019-03-20 17:43 | 1553118199.0717 | -341 | 0.20952 | 0.000444 | 0 | 108.399 | 109.7338 | 125.3664 | 21.0491 | 0.002016 | 0.001835 | -151.8588 | 0.23852 | 8.4037 | 0.003567 | 0.004671 | 0.004438 | 5.9e-05 | 0.001275 | 0.27096 | 1.2993 | 2.5481 | 100.58 | 20.65 | 19.146 | 70.1274 | 107.5016 | 65.7661 | 0.004638 | 1.0417 | 0.024997 | 23.9965 | -0.10947 | 0.14189 | 2e-06 | 9.6e-05 | 9.6e-05 | 60.4942
2019-03-20 17:43 | 1553118200.0724 | -340 | 0.20954 | 0.000448 | 1e-06 | 108.3719 | 109.828 | 125.3323 | 21.0485 | 0.001948 | 0.001848 | -151.8565 | 0.23837 | 8.4023 | 0.003581 | 0.004405 | 0.004154 | 6e-05 | 0.001277 | 0.26272 | 1.3607 | 2.5481 | 100.58 | 20.65 | 19.145 | 68.4987 | 107.847 | 66.0972 | 0.004638 | 1.0407 | 0.024981 | 24.005 | -0.11938 | 0.1418 | -4e-06 | -0.000345 | -0.000345 | -16.881
2019-03-20 17:43 | 1553118201.0724 | -339 | 0.20954 | 0.000444 | 1e-06 | 108.3929 | 109.7607 | 125.3434 | 21.0491 | 0.002 | 0.001875 | -151.8649 | 0.23927 | 8.4045 | 0.003646 | 0.005511 | 0.005367 | 5.8e-05 | 0.001281 | 0.30182 | 1.3292 | 2.5482 | 100.58 | 20.65 | 19.145 | 69.7376 | 109.4181 | 66.7378 | 0.004638 | 1.0387 | 0.024959 | 24.0287 | -0.11217 | 0.14234 | 2e-06 | 8.9e-05 | 8.9e-05 | 65.452
2019-03-20 17:43 | 1553118202.0729 | -338 | 0.20954 | 0.000448 | 1e-06 | 108.389 | 109.7118 | 125.2469 | 21.0495 | 0.002012 | 0.001851 | -151.8645 | 0.23936 | 8.4046 | 0.003648 | 0.004726 | 0.004551 | 5.9e-05 | 0.001285 | 0.28838 | 1.3224 | 2.5481 | 100.58 | 20.65 | 19.145 | 70.0242 | 109.4754 | 66.1708 | 0.004638 | 1.0404 | 0.02497 | 23.9994 | -0.1085 | 0.1424 | -4e-06 | -0.000316 | -0.000316 | -18.515
2019-03-20 17:43 | 1553118203.0725 | -337 | 0.20958 | 0.000444 | 0 | 108.4566 | 109.8089 | 124.6882 | 21.0497 | 0.002035 | 0.001835 | -151.8593 | 0.23848 | 8.4061 | 0.003591 | 0.005097 | 0.004834 | 6.3e-05 | 0.001283 | 0.25227 | 1.3192 | 2.5479 | 100.58 | 20.65 | 19.147 | 70.5803 | 108.0829 | 65.7761 | 0.004639 | 1.0417 | 0.024929 | 23.9319 | -0.11084 | 0.14187 | 2e-06 | 0.000134 | 0.000134 | 43.6164
2019-03-20 17:43 | 1553118204.0726 | -336 | 0.20952 | 0.000446 | 1e-06 | 108.5265 | 109.8086 | 124.2897 | 21.0493 | 0.001978 | 0.001784 | -151.8522 | 0.23809 | 8.4043 | 0.003561 | 0.005027 | 0.004823 | 6e-05 | 0.001285 | 0.27184 | 1.3226 | 2.5481 | 100.58 | 20.64 | 19.153 | 69.2208 | 107.3523 | 64.5558 | 0.004637 | 1.0454 | 0.024934 | 23.8505 | -0.10505 | 0.14164 | 0 | -5e-05 | -5e-05 | -116.3977
2019-03-20 17:43 | 1553118205.0727 | -335 | 0.20954 | 0.000448 | 0 | 108.5012 | 109.7091 | 124.2127 | 21.0495 | 0.001932 | 0.00183 | -151.8724 | 0.23831 | 8.4031 | 0.003572 | 0.00548 | 0.005211 | 6e-05 | 0.001289 | 0.27759 | 1.2948 | 2.5482 | 100.58 | 20.64 | 19.157 | 68.0993 | 107.636 | 65.651 | 0.004638 | 1.042 | 0.024886 | 23.8818 | -0.099026 | 0.14177 | -3e-06 | -0.000228 | -0.000228 | -25.5034
2019-03-20 17:43 | 1553118206.0716 | -334 | 0.20954 | 0.00045 | 0 | 108.4721 | 109.8616 | 124.1045 | 21.0503 | 0.001885 | 0.001825 | -151.8651 | 0.23832 | 8.4026 | 0.003571 | 0.004547 | 0.004432 | 6.3e-05 | 0.001301 | 0.30829 | 1.3118 | 2.5481 | 100.58 | 20.64 | 19.161 | 66.9871 | 107.6084 | 65.5402 | 0.004639 | 1.0424 | 0.024879 | 23.8675 | -0.11385 | 0.14178 | -5e-06 | -0.00042 | -0.00042 | -13.8654
2019-03-20 17:43 | 1553118207.0728 | -333 | 0.20954 | 0.000447 | 0 | 108.4582 | 109.7203 | 124.5134 | 21.0507 | 0.001879 | 0.001796 | -151.8697 | 0.23924 | 8.4037 | 0.00352 | 0.004724 | 0.004435 | 5.8e-05 | 0.0013 | 0.29192 | 1.3032 | 2.548 | 100.58 | 20.64 | 19.171 | 66.8461 | 106.3672 | 64.8492 | 0.004642 | 1.0445 | 0.024945 | 23.8824 | -0.10349 | 0.14232 | -2e-06 | -0.000178 | -0.000178 | -32.8915
2019-03-20 17:43 | 1553118208.0708 | -332 | 0.20952 | 0.000449 | 0 | 108.5116 | 109.9136 | 124.5021 | 21.051 | 0.001966 | 0.001825 | -151.8802 | 0.23897 | 8.4001 | 0.003524 | 0.00548 | 0.005315 | 5.9e-05 | 0.001295 | 0.27168 | 1.2808 | 2.548 | 100.58 | 20.64 | 19.174 | 68.9354 | 106.4594 | 65.5293 | 0.004642 | 1.0424 | 0.024919 | 23.9053 | -0.11483 | 0.14216 | -4e-06 | -0.000348 | -0.000348 | -16.8099
2019-03-20 17:43 | 1553118209.0719 | -331 | 0.20952 | 0.000459 | 1e-06 | 108.5365 | 109.8358 | 124.5188 | 21.0498 | 0.001956 | 0.001843 | -151.8808 | 0.23809 | 8.4036 | 0.003537 | 0.005344 | 0.005176 | 6e-05 | 0.001291 | 0.27096 | 1.2873 | 2.5481 | 100.58 | 20.64 | 19.183 | 68.6963 | 106.7827 | 65.9565 | 0.004644 | 1.0411 | 0.024905 | 23.922 | -0.10644 | 0.14164 | -1.7e-05 | -0.001308 | -0.001308 | -4.4518
2019-03-20 17:43 | 1553118210.071 | -330 | 0.20954 | 0.000447 | 1e-06 | 108.5109 | 109.8418 | 124.6333 | 21.0503 | 0.001929 | 0.001847 | -151.8642 | 0.23984 | 8.402 | 0.003467 | 0.004972 | 0.00472 | 6.1e-05 | 0.00129 | 0.26864 | 1.3132 | 2.5481 | 100.58 | 20.64 | 19.187 | 68.0479 | 105.0912 | 66.0756 | 0.004645 | 1.0407 | 0.024912 | 23.9372 | -0.10904 | 0.14268 | -2e-06 | -0.000209 | -0.000209 | -28.014
2019-03-20 17:44 | 1553118211.0719 | -329 | 0.20955 | 0.000452 | 1e-06 | 108.4893 | 109.7723 | 124.8774 | 21.0504 | 0.001921 | 0.001847 | -151.8586 | 0.23794 | 8.403 | 0.003464 | 0.004698 | 0.004509 | 5.8e-05 | 0.001286 | 0.27151 | 1.3533 | 2.548 | 100.58 | 20.65 | 19.194 | 67.8475 | 105.0127 | 66.0694 | 0.00465 | 1.0408 | 0.024937 | 23.9604 | -0.10516 | 0.14155 | -8e-06 | -0.000659 | -0.000659 | -8.8308
2019-03-20 17:44 | 1553118212.072 | -328 | 0.20956 | 0.000449 | 0 | 108.5433 | 109.9218 | 124.9251 | 21.051 | 0.001972 | 0.001775 | -151.8463 | 0.23814 | 8.4048 | 0.003488 | 0.004988 | 0.004785 | 5.8e-05 | 0.001284 | 0.262 | 1.3317 | 2.5481 | 100.58 | 20.65 | 19.196 | 69.0751 | 105.6057 | 64.3319 | 0.00465 | 1.0461 | 0.025006 | 23.9035 | -0.11288 | 0.14167 | -4e-06 | -0.000317 | -0.000317 | -18.3865
2019-03-20 17:44 | 1553118213.0715 | -327 | 0.20963 | 0.000452 | 0 | 108.5552 | 109.9842 | 125.3149 | 21.051 | 0.001991 | 0.001824 | -151.8662 | 0.23906 | 8.4032 | 0.003461 | 0.00526 | 0.005101 | 6e-05 | 0.001283 | 0.29142 | 1.3691 | 2.5482 | 100.58 | 20.65 | 19.197 | 69.5344 | 104.9439 | 65.505 | 0.004651 | 1.0425 | 0.025001 | 23.9823 | -0.11697 | 0.14222 | -7e-06 | -0.000543 | -0.000543 | -10.769
2019-03-20 17:44 | 1553118214.0721 | -326 | 0.20939 | 0.000446 | 0 | 108.5659 | 109.9106 | 125.7372 | 21.0509 | 0.001926 | 0.001845 | -151.8941 | 0.23989 | 8.403 | 0.003556 | 0.005192 | 0.004915 | 6.1e-05 | 0.001282 | 0.29336 | 1.3434 | 2.548 | 100.58 | 20.65 | 19.199 | 67.9572 | 107.2412 | 66.024 | 0.004651 | 1.0409 | 0.025024 | 24.0411 | -0.1101 | 0.14271 | 0 | -2.7e-05 | -2.7e-05 | -216.6539
2019-03-20 17:44 | 1553118215.0722 | -325 | 0.20953 | 0.000448 | 0 | 108.5743 | 109.9117 | 125.8809 | 21.0507 | 0.002001 | 0.001834 | -151.8874 | 0.23861 | 8.4024 | 0.003461 | 0.005528 | 0.005412 | 6.1e-05 | 0.001283 | 0.26792 | 1.3318 | 2.5481 | 100.58 | 20.65 | 19.198 | 69.7656 | 104.9374 | 65.7492 | 0.004651 | 1.0417 | 0.025049 | 24.0451 | -0.1095 | 0.14195 | -2e-06 | -0.000135 | -0.000135 | -43.1191
2019-03-20 17:44 | 1553118216.0717 | -324 | 0.20957 | 0.00045 | 0 | 108.5682 | 109.9831 | 125.956 | 21.0507 | 0.001986 | 0.001846 | -151.9002 | 0.23882 | 8.404 | 0.003565 | 0.004495 | 0.004205 | 5.9e-05 | 0.001283 | 0.27223 | 1.3536 | 2.548 | 100.58 | 20.65 | 19.197 | 69.4021 | 107.4607 | 66.0392 | 0.004651 | 1.0409 | 0.025046 | 24.0626 | -0.11582 | 0.14208 | -5e-06 | -0.000378 | -0.000378 | -15.4548
2019-03-20 17:44 | 1553118217.0713 | -323 | 0.20951 | 0.000445 | 0 | 108.6306 | 109.9517 | 125.6858 | 21.0507 | 0.001974 | 0.001806 | -151.9073 | 0.23841 | 8.4021 | 0.003644 | 0.005875 | 0.005694 | 5.9e-05 | 0.00129 | 0.31614 | 1.3309 | 2.5481 | 100.58 | 20.65 | 19.192 | 69.1246 | 109.3741 | 65.0725 | 0.004649 | 1.0438 | 0.025054 | 24.0025 | -0.10812 | 0.14183 | 1e-06 | 7.7e-05 | 7.7e-05 | 75.8303
2019-03-20 17:44 | 1553118218.0713 | -322 | 0.20952 | 0.000447 | 0 | 108.6455 | 109.9917 | 125.5307 | 21.0504 | 0.002001 | 0.00181 | -151.9097 | 0.23864 | 8.4014 | 0.003645 | 0.004919 | 0.004804 | 6e-05 | 0.001286 | 0.27256 | 1.3356 | 2.548 | 100.58 | 20.65 | 19.188 | 69.7778 | 109.398 | 65.1838 | 0.004649 | 1.0435 | 0.025035 | 23.9916 | -0.11015 | 0.14197 | -1e-06 | -9.5e-05 | -9.5e-05 | -61.6983
2019-03-20 17:44 | 1553118219.0717 | -321 | 0.20953 | 0.000444 | 0 | 108.6539 | 110.1943 | 125.5408 | 21.0511 | 0.001974 | 0.001843 | -151.9015 | 0.23804 | 8.401 | 0.003667 | 0.004704 | 0.004405 | 5.9e-05 | 0.001286 | 0.2672 | 1.3612 | 2.548 | 100.58 | 20.65 | 19.179 | 69.1141 | 109.9278 | 65.9655 | 0.004646 | 1.0411 | 0.025007 | 24.0202 | -0.12592 | 0.14161 | 2e-06 | 0.000175 | 0.000175 | 33.3299
2019-03-20 17:44 | 1553118220.0725 | -320 | 0.20953 | 0.000446 | 0 | 108.6131 | 109.9629 | 125.8029 | 21.0508 | 0.001956 | 0.001892 | -151.9112 | 0.23925 | 8.4042 | 0.003639 | 0.005851 | 0.005572 | 5.8e-05 | 0.001286 | 0.27959 | 1.3556 | 2.548 | 100.58 | 20.65 | 19.174 | 68.6895 | 109.2708 | 67.1419 | 0.004645 | 1.0375 | 0.02499 | 24.087 | -0.11047 | 0.14233 | 1e-06 | 0.000136 | 0.000136 | 42.9853
2019-03-20 17:44 | 1553118221.0725 | -319 | 0.20952 | 0.000449 | 0 | 108.7296 | 110.0418 | 125.8344 | 21.051 | 0.002001 | 0.001843 | -151.8987 | 0.23858 | 8.4003 | 0.003675 | 0.004951 | 0.004737 | 5.9e-05 | 0.00129 | 0.28672 | 1.3348 | 2.5481 | 100.58 | 20.65 | 19.168 | 69.7658 | 110.1384 | 65.9685 | 0.004644 | 1.0411 | 0.025036 | 24.0484 | -0.1073 | 0.14193 | -4e-06 | -0.000314 | -0.000314 | -18.5527
2019-03-20 17:44 | 1553118222.072 | -318 | 0.20957 | 0.000447 | -1e-06 | 108.6445 | 109.8504 | 125.8065 | 21.0517 | 0.001941 | 0.001805 | -151.9116 | 0.23911 | 8.4024 | 0.003649 | 0.004827 | 0.004628 | 5.8e-05 | 0.001292 | 0.28655 | 1.3178 | 2.5481 | 100.58 | 20.65 | 19.159 | 68.3315 | 109.513 | 65.0647 | 0.004641 | 1.0438 | 0.025067 | 24.0137 | -0.098734 | 0.14225 | 0 | 4.2e-05 | 4.2e-05 | 138.7745
2019-03-20 17:44 | 1553118223.0715 | -317 | 0.20953 | 0.00045 | 0 | 108.6202 | 110.0108 | 125.6731 | 21.052 | 0.001982 | 0.001925 | -151.9068 | 0.23926 | 8.4019 | 0.003523 | 0.005013 | 0.004828 | 5.9e-05 | 0.001294 | 0.28279 | 1.3278 | 2.5481 | 100.58 | 20.65 | 19.155 | 69.303 | 106.452 | 67.9397 | 0.004641 | 1.0351 | 0.024948 | 24.1028 | -0.11378 | 0.14234 | -5e-06 | -0.000365 | -0.000365 | -16.0059
2019-03-20 17:44 | 1553118224.0716 | -316 | 0.20953 | 0.000441 | 0 | 108.6202 | 110.1539 | 125.2631 | 21.0511 | 0.001958 | 0.001848 | -151.9115 | 0.23869 | 8.4026 | 0.003555 | 0.004942 | 0.004665 | 5.8e-05 | 0.001297 | 0.29286 | 1.3208 | 2.5483 | 100.58 | 20.65 | 19.152 | 68.7264 | 107.2186 | 66.0865 | 0.00464 | 1.0407 | 0.024975 | 23.9979 | -0.12541 | 0.14199 | 5e-06 | 0.000378 | 0.000378 | 15.4248
2019-03-20 17:44 | 1553118225.0721 | -315 | 0.20956 | 0.000446 | 0 | 108.7244 | 109.973 | 124.9098 | 21.0513 | 0.001975 | 0.001859 | -151.9029 | 0.23881 | 8.4035 | 0.003578 | 0.005422 | 0.005216 | 5.8e-05 | 0.001299 | 0.29535 | 1.3113 | 2.548 | 100.58 | 20.65 | 19.151 | 69.1518 | 107.7851 | 66.3638 | 0.00464 | 1.0399 | 0.024929 | 23.9739 | -0.10214 | 0.14207 | 1e-06 | 5.1e-05 | 5.1e-05 | 113.6396
2019-03-20 17:44 | 1553118226.0722 | -314 | 0.20954 | 0.00045 | 0 | 108.6891 | 110.2713 | 125.0453 | 21.0507 | 0.001973 | 0.001828 | -151.8984 | 0.23837 | 8.4048 | 0.003647 | 0.004515 | 0.004265 | 6.1e-05 | 0.0013 | 0.286 | 1.3456 | 2.5479 | 100.58 | 20.65 | 19.147 | 69.0974 | 109.464 | 65.6118 | 0.004639 | 1.0422 | 0.024971 | 23.9603 | -0.12926 | 0.14181 | -5e-06 | -0.000356 | -0.000356 | -16.378
2019-03-20 17:44 | 1553118227.0707 | -313 | 0.20954 | 0.000447 | 0 | 108.6998 | 109.9719 | 125.068 | 21.0515 | 0.001955 | 0.001812 | -151.8991 | 0.23852 | 8.402 | 0.003609 | 0.004944 | 0.004742 | 6e-05 | 0.0013 | 0.27417 | 1.3513 | 2.5481 | 100.58 | 20.65 | 19.145 | 68.6566 | 108.5376 | 65.23 | 0.004638 | 1.0433 | 0.024987 | 23.949 | -0.10407 | 0.1419 | -1e-06 | -6.9e-05 | -6.9e-05 | -84.9175
2019-03-20 17:44 | 1553118228.0719 | -312 | 0.20953 | 0.00045 | 0 | 108.6193 | 110.0362 | 125.2671 | 21.0517 | 0.001992 | 0.001869 | -151.8866 | 0.23878 | 8.4018 | 0.003582 | 0.005228 | 0.00499 | 5.6e-05 | 0.001296 | 0.29695 | 1.3243 | 2.5481 | 100.58 | 20.65 | 19.145 | 69.5521 | 107.8807 | 66.5912 | 0.004638 | 1.0392 | 0.024957 | 24.0162 | -0.11592 | 0.14205 | -4e-06 | -0.0003 | -0.0003 | -19.4838
2019-03-20 17:44 | 1553118229.0714 | -311 | 0.20952 | 0.000443 | 0 | 108.6623 | 110.0791 | 125.5175 | 21.0513 | 0.00194 | 0.001824 | -151.8882 | 0.23752 | 8.4049 | 0.003618 | 0.005079 | 0.004819 | 5.9e-05 | 0.001297 | 0.28816 | 1.3259 | 2.5481 | 100.58 | 20.65 | 19.144 | 68.2991 | 108.7519 | 65.5029 | 0.004638 | 1.0425 | 0.025022 | 24.0016 | -0.11588 | 0.1413 | 4e-06 | 0.000322 | 0.000322 | 18.0115
2019-03-20 17:44 | 1553118230.072 | -310 | 0.20952 | 0.000446 | 1e-06 | 108.7503 | 110.173 | 125.4534 | 21.0519 | 0.001935 | 0.001833 | -151.8952 | 0.23727 | 8.4005 | 0.00363 | 0.004901 | 0.004767 | 6e-05 | 0.001296 | 0.29446 | 1.3135 | 2.5482 | 100.58 | 20.65 | 19.144 | 68.1913 | 109.0338 | 65.7197 | 0.004638 | 1.0418 | 0.025007 | 24.0032 | -0.11626 | 0.14115 | -1e-06 | -0.00014 | -0.00014 | -41.5083
2019-03-20 17:44 | 1553118231.071 | -309 | 0.20952 | 0.000444 | 1e-06 | 108.6982 | 110.0089 | 125.1849 | 21.0516 | 0.002 | 0.001866 | -151.9056 | 0.23835 | 8.4057 | 0.003736 | 0.005373 | 0.005121 | 6.1e-05 | 0.001304 | 0.27959 | 1.3219 | 2.548 | 100.58 | 20.65 | 19.148 | 69.7432 | 111.6146 | 66.524 | 0.004639 | 1.0394 | 0.024951 | 24.0059 | -0.10721 | 0.14179 | 2e-06 | 0.000134 | 0.000134 | 43.5433
2019-03-20 17:44 | 1553118232.0714 | -308 | 0.20952 | 0.000448 | 1e-06 | 108.7425 | 110.1819 | 125.2499 | 21.0516 | 0.00197 | 0.001858 | -151.9046 | 0.23742 | 8.402 | 0.003716 | 0.005229 | 0.005029 | 6.1e-05 | 0.001299 | 0.26897 | 1.312 | 2.5479 | 100.58 | 20.65 | 19.149 | 69.0132 | 111.1261 | 66.3367 | 0.004639 | 1.0399 | 0.024964 | 24.0055 | -0.11762 | 0.14124 | -3e-06 | -0.000268 | -0.000268 | -21.6524
2019-03-20 17:44 | 1553118233.0721 | -307 | 0.20955 | 0.000445 | 0 | 108.7037 | 109.9573 | 125.6508 | 21.0518 | 0.001985 | 0.001885 | -151.8928 | 0.2378 | 8.4034 | 0.003659 | 0.005025 | 0.004816 | 6.4e-05 | 0.001302 | 0.28854 | 1.3094 | 2.548 | 100.58 | 20.65 | 19.149 | 69.3799 | 109.7448 | 66.9681 | 0.004639 | 1.038 | 0.024981 | 24.0663 | -0.10256 | 0.14147 | 1e-06 | 4e-05 | 4e-05 | 146.6061
2019-03-20 17:44 | 1553118234.0722 | -306 | 0.20949 | 0.000445 | 0 | 108.7765 | 110.0978 | 126.0306 | 21.0513 | 0.001945 | 0.001843 | -151.8848 | 0.23833 | 8.4024 | 0.003646 | 0.00524 | 0.005019 | 6.2e-05 | 0.001302 | 0.26217 | 1.3407 | 2.548 | 100.58 | 20.65 | 19.147 | 68.4107 | 109.4227 | 65.9758 | 0.004639 | 1.041 | 0.025055 | 24.0674 | -0.108 | 0.14178 | 1e-06 | 8.9e-05 | 8.9e-05 | 65.7464
2019-03-20 17:44 | 1553118235.0713 | -305 | 0.20954 | 0.000446 | 0 | 108.7629 | 110.0433 | 126.1578 | 21.052 | 0.00199 | 0.001798 | -151.8963 | 0.23793 | 8.4044 | 0.003675 | 0.004739 | 0.004525 | 6.1e-05 | 0.001298 | 0.28727 | 1.3407 | 2.5479 | 100.58 | 20.65 | 19.145 | 69.5149 | 110.1472 | 64.8838 | 0.004638 | 1.0444 | 0.025108 | 24.0408 | -0.10468 | 0.14155 | 0 | -3.5e-05 | -3.5e-05 | -167.6328
2019-03-20 17:44 | 1553118236.0722 | -304 | 0.2095 | 0.000448 | 0 | 108.7279 | 110.0059 | 126.1292 | 21.0513 | 0.00198 | 0.001873 | -151.8881 | 0.23874 | 8.3998 | 0.003666 | 0.00507 | 0.004907 | 5.9e-05 | 0.001302 | 0.2855 | 1.3128 | 2.5481 | 100.58 | 20.65 | 19.14 | 69.2585 | 109.9153 | 66.6874 | 0.004637 | 1.0389 | 0.025039 | 24.1021 | -0.10452 | 0.14202 | -1e-06 | -6e-05 | -6e-05 | -97.073
2019-03-20 17:44 | 1553118237.0724 | -303 | 0.20956 | 0.000447 | 0 | 108.7486 | 110.2664 | 126.0048 | 21.0512 | 0.001953 | 0.001823 | -151.8958 | 0.23836 | 8.4037 | 0.003703 | 0.005132 | 0.00489 | 6.3e-05 | 0.001298 | 0.26106 | 1.3342 | 2.548 | 100.58 | 20.65 | 19.137 | 68.6058 | 110.8272 | 65.4833 | 0.004636 | 1.0426 | 0.025071 | 24.0475 | -0.12397 | 0.1418 | -1e-06 | -1.8e-05 | -1.8e-05 | -329.0035
2019-03-20 17:44 | 1553118238.0724 | -302 | 0.20952 | 0.000444 | -1e-06 | 108.6872 | 110.0205 | 125.637 | 21.052 | 0.001992 | 0.001866 | -151.9106 | 0.23813 | 8.4003 | 0.003685 | 0.005758 | 0.00554 | 6.2e-05 | 0.001297 | 0.27776 | 1.337 | 2.5481 | 100.58 | 20.65 | 19.135 | 69.5572 | 110.3714 | 66.5299 | 0.004636 | 1.0393 | 0.024996 | 24.0494 | -0.10906 | 0.14167 | 4e-06 | 0.000392 | 0.000392 | 14.8351
2019-03-20 17:44 | 1553118239.0715 | -301 | 0.2095 | 0.000444 | 0 | 108.7208 | 110.1786 | 125.4093 | 21.052 | 0.001954 | 0.001809 | -151.9228 | 0.23839 | 8.4021 | 0.003718 | 0.004924 | 0.004745 | 6.1e-05 | 0.001302 | 0.26698 | 1.3141 | 2.548 | 100.58 | 20.65 | 19.129 | 68.647 | 111.1776 | 65.1603 | 0.004634 | 1.0436 | 0.025024 | 23.9792 | -0.11913 | 0.14182 | 3e-06 | 0.00029 | 0.00029 | 20.11
2019-03-20 17:44 | 1553118240.0716 | -300 | 0.2095 | 0.000441 | 0 | 108.7668 | 110.0545 | 125.6865 | 21.0519 | 0.001902 | 0.001788 | -151.9203 | 0.23835 | 8.4017 | 0.00363 | 0.005094 | 0.004907 | 6.2e-05 | 0.001306 | 0.29535 | 1.31 | 2.548 | 100.58 | 20.65 | 19.126 | 67.3974 | 109.0394 | 64.6511 | 0.004633 | 1.0451 | 0.02507 | 23.9876 | -0.10528 | 0.1418 | 6e-06 | 0.000444 | 0.000444 | 13.1261
2019-03-20 17:44 | 1553118241.0716 | -299 | 0.20951 | 0.000445 | 0 | 108.7292 | 110.0706 | 126.0123 | 21.0517 | 0.001978 | 0.001842 | -151.9146 | 0.23775 | 8.4007 | 0.003617 | 0.004387 | 0.00421 | 5.8e-05 | 0.001298 | 0.28888 | 1.2985 | 2.548 | 100.58 | 20.65 | 19.118 | 69.2246 | 108.7295 | 65.9365 | 0.004632 | 1.0412 | 0.025055 | 24.0643 | -0.10967 | 0.14144 | 1e-06 | 0.000107 | 0.000107 | 54.5711
2019-03-20 17:44 | 1553118242.071 | -298 | 0.20945 | 0.000447 | 0 | 108.6853 | 110.1745 | 125.8687 | 21.0516 | 0.001891 | 0.001824 | -151.9185 | 0.23874 | 8.4024 | 0.003707 | 0.005169 | 0.00486 | 6e-05 | 0.001291 | 0.30917 | 1.3259 | 2.5481 | 100.58 | 20.65 | 19.112 | 67.1278 | 110.9137 | 65.5188 | 0.00463 | 1.0425 | 0.025056 | 24.0358 | -0.12173 | 0.14203 | -1e-06 | -9.3e-05 | -9.3e-05 | -62.9955
2019-03-20 17:44 | 1553118243.0711 | -297 | 0.20946 | 0.000447 | 0 | 108.6901 | 110.1064 | 125.6443 | 21.0521 | 0.001962 | 0.001833 | -151.9366 | 0.23898 | 8.4024 | 0.003617 | 0.004891 | 0.004709 | 6.4e-05 | 0.001292 | 0.28335 | 1.3234 | 2.548 | 100.58 | 20.65 | 19.107 | 68.8295 | 108.7155 | 65.7308 | 0.004629 | 1.0418 | 0.025026 | 24.0218 | -0.1158 | 0.14217 | 0 | -4.1e-05 | -4.1e-05 | -141.7931
2019-03-20 17:44 | 1553118244.0717 | -296 | 0.20947 | 0.000444 | 0 | 108.7121 | 110.1711 | 125.5938 | 21.0517 | 0.002002 | 0.001855 | -151.9373 | 0.23796 | 8.4028 | 0.003725 | 0.004922 | 0.004807 | 6.4e-05 | 0.001295 | 0.27024 | 1.3248 | 2.5482 | 100.58 | 20.65 | 19.1 | 69.7846 | 111.3612 | 66.269 | 0.004627 | 1.0401 | 0.025001 | 24.0361 | -0.11924 | 0.14156 | 3e-06 | 0.000208 | 0.000208 | 28.0319
2019-03-20 17:44 | 1553118245.0708 | -295 | 0.20948 | 0.000444 | 0 | 108.7208 | 110.1505 | 125.5346 | 21.0515 | 0.001981 | 0.001862 | -151.9456 | 0.23911 | 8.401 | 0.0038 | 0.005557 | 0.00532 | 6.3e-05 | 0.001289 | 0.25962 | 1.3011 | 2.5481 | 100.58 | 20.65 | 19.098 | 69.2868 | 113.1679 | 66.431 | 0.004627 | 1.0396 | 0.024989 | 24.0361 | -0.11686 | 0.14225 | 3e-06 | 0.000263 | 0.000263 | 22.2715
2019-03-20 17:44 | 1553118246.0722 | -294 | 0.20949 | 0.000443 | 0 | 108.6998 | 109.8669 | 125.4307 | 21.0529 | 0.001986 | 0.001828 | -151.9537 | 0.23824 | 8.4034 | 0.003793 | 0.004472 | 0.004223 | 6.3e-05 | 0.001285 | 0.28456 | 1.3069 | 2.5481 | 100.58 | 20.65 | 19.095 | 69.4067 | 113.0144 | 65.5983 | 0.004626 | 1.0422 | 0.02501 | 23.9967 | -0.095521 | 0.14173 | 3e-06 | 0.000222 | 0.000222 | 26.1874
2019-03-20 17:44 | 1553118247.0709 | -293 | 0.20949 | 0.000442 | 0 | 108.6798 | 110.0268 | 125.4316 | 21.0526 | 0.001993 | 0.001855 | -151.9587 | 0.23832 | 8.4021 | 0.003749 | 0.005355 | 0.005129 | 6.4e-05 | 0.001287 | 0.30685 | 1.3414 | 2.5481 | 100.58 | 20.65 | 19.089 | 69.585 | 111.9329 | 66.2511 | 0.004625 | 1.0402 | 0.024986 | 24.0199 | -0.11019 | 0.14178 | 6e-06 | 0.000467 | 0.000467 | 12.4757
2019-03-20 17:44 | 1553118248.0709 | -292 | 0.20951 | 0.000442 | 0 | 108.6202 | 110.1289 | 125.5323 | 21.0534 | 0.002005 | 0.001865 | -151.97 | 0.23829 | 8.4039 | 0.003745 | 0.004972 | 0.004715 | 6.2e-05 | 0.001294 | 0.27345 | 1.3669 | 2.5481 | 100.58 | 20.65 | 19.084 | 69.8579 | 111.8271 | 66.5118 | 0.004623 | 1.0394 | 0.024986 | 24.0388 | -0.12337 | 0.14176 | 5e-06 | 0.000404 | 0.000404 | 14.4355
2019-03-20 17:44 | 1553118249.0714 | -291 | 0.2095 | 0.000441 | 0 | 108.7706 | 110.167 | 125.6241 | 21.0526 | 0.001904 | 0.001862 | -151.9748 | 0.2381 | 8.4006 | 0.003759 | 0.004874 | 0.004716 | 6.3e-05 | 0.001295 | 0.27815 | 1.3336 | 2.5481 | 100.58 | 20.65 | 19.076 | 67.4315 | 112.1769 | 66.4266 | 0.004621 | 1.0397 | 0.024998 | 24.0445 | -0.1141 | 0.14164 | 6e-06 | 0.000487 | 0.000487 | 11.9508
2019-03-20 17:44 | 1553118250.0711 | -290 | 0.2095 | 0.000447 | 0 | 108.7208 | 110.2126 | 125.7051 | 21.0532 | 0.001943 | 0.001906 | -151.9776 | 0.23822 | 8.3993 | 0.003709 | 0.005009 | 0.004696 | 6.2e-05 | 0.001298 | 0.28567 | 1.3589 | 2.5481 | 100.58 | 20.65 | 19.072 | 68.3823 | 110.954 | 67.4773 | 0.00462 | 1.0365 | 0.024968 | 24.0895 | -0.12189 | 0.14172 | -2e-06 | -0.000129 | -0.000129 | -45.1476
2019-03-20 17:44 | 1553118251.0721 | -289 | 0.2095 | 0.000442 | 0 | 108.6335 | 110.1266 | 125.7595 | 21.053 | 0.001993 | 0.001829 | -151.9738 | 0.23844 | 8.4039 | 0.003833 | 0.004752 | 0.00455 | 6.1e-05 | 0.0013 | 0.27975 | 1.3416 | 2.5481 | 100.58 | 20.65 | 19.067 | 69.576 | 113.9778 | 65.6335 | 0.004619 | 1.0421 | 0.025041 | 24.0294 | -0.1221 | 0.14185 | 6e-06 | 0.000418 | 0.000418 | 13.9537
2019-03-20 17:44 | 1553118252.0725 | -288 | 0.20952 | 0.000449 | 0 | 108.6613 | 110.0246 | 125.6377 | 21.0531 | 0.001983 | 0.001846 | -151.9646 | 0.23835 | 8.4056 | 0.003771 | 0.00523 | 0.005021 | 6e-05 | 0.001293 | 0.28119 | 1.3407 | 2.5481 | 100.58 | 20.65 | 19.059 | 69.3384 | 112.4798 | 66.0433 | 0.004617 | 1.0408 | 0.025014 | 24.0323 | -0.11152 | 0.1418 | -3e-06 | -0.000239 | -0.000239 | -24.3527
2019-03-20 17:44 | 1553118253.0722 | -287 | 0.20952 | 0.000445 | 0 | 108.7457 | 109.9207 | 125.5227 | 21.0541 | 0.00204 | 0.001873 | -151.9607 | 0.23816 | 8.4016 | 0.003789 | 0.005414 | 0.005165 | 6.1e-05 | 0.001295 | 0.27831 | 1.3449 | 2.5481 | 100.58 | 20.65 | 19.055 | 70.7006 | 112.9121 | 66.6863 | 0.004616 | 1.0389 | 0.024979 | 24.044 | -0.096123 | 0.14168 | 2e-06 | 0.000179 | 0.000179 | 32.4589
2019-03-20 17:44 | 1553118254.0722 | -286 | 0.20953 | 0.000443 | 0 | 108.6781 | 110.004 | 125.2365 | 21.0535 | 0.001983 | 0.00184 | -151.9555 | 0.23822 | 8.4021 | 0.003743 | 0.005656 | 0.00547 | 6.4e-05 | 0.001293 | 0.27815 | 1.379 | 2.5481 | 100.58 | 20.65 | 19.049 | 69.3472 | 111.7797 | 65.9046 | 0.004615 | 1.0413 | 0.024979 | 23.989 | -0.10846 | 0.14172 | 4e-06 | 0.000287 | 0.000287 | 20.3077
2019-03-20 17:44 | 1553118255.0707 | -285 | 0.20951 | 0.000444 | 0 | 108.6694 | 109.9543 | 125.3621 | 21.0544 | 0.001949 | 0.001803 | -151.9537 | 0.23896 | 8.4038 | 0.003756 | 0.004786 | 0.004619 | 6.3e-05 | 0.001295 | 0.27345 | 1.371 | 2.5481 | 100.58 | 20.64 | 19.045 | 68.5138 | 112.1035 | 65.0191 | 0.004611 | 1.044 | 0.025024 | 23.9696 | -0.10514 | 0.14216 | 3e-06 | 0.00027 | 0.00027 | 21.633
2019-03-20 17:44 | 1553118256.0713 | -284 | 0.20952 | 0.000443 | 0 | 108.5792 | 109.9244 | 125.2626 | 21.054 | 0.001973 | 0.001814 | -151.943 | 0.23815 | 8.4039 | 0.00375 | 0.00514 | 0.004931 | 6.2e-05 | 0.001294 | 0.28296 | 1.3371 | 2.5481 | 100.58 | 20.64 | 19.041 | 69.0864 | 111.9634 | 65.2746 | 0.00461 | 1.0432 | 0.025005 | 23.9692 | -0.11014 | 0.14168 | 5e-06 | 0.0004 | 0.0004 | 14.5403
2019-03-20 17:44 | 1553118257.0704 | -283 | 0.20952 | 0.000445 | 0 | 108.6519 | 109.834 | 125.3213 | 21.0534 | 0.001946 | 0.001848 | -151.9355 | 0.23768 | 8.4033 | 0.003691 | 0.005761 | 0.005621 | 6.5e-05 | 0.001296 | 0.274 | 1.3144 | 2.5481 | 100.58 | 20.64 | 19.035 | 68.444 | 110.5193 | 66.1029 | 0.004609 | 1.0407 | 0.02498 | 24.0041 | -0.096782 | 0.14139 | 2e-06 | 0.000175 | 0.000175 | 33.1716
2019-03-20 17:44 | 1553118258.0715 | -282 | 0.20955 | 0.000445 | 0 | 108.6603 | 110.1464 | 125.3161 | 21.0538 | 0.001955 | 0.00181 | -151.9238 | 0.23806 | 8.4006 | 0.003674 | 0.005206 | 0.004951 | 6.2e-05 | 0.001293 | 0.26626 | 1.3229 | 2.548 | 100.58 | 20.64 | 19.032 | 68.6598 | 110.1072 | 65.1777 | 0.004608 | 1.0435 | 0.025014 | 23.9709 | -0.1215 | 0.14162 | 2e-06 | 0.000147 | 0.000147 | 39.7209
2019-03-20 17:44 | 1553118259.0717 | -281 | 0.20953 | 0.000448 | 0 | 108.6054 | 109.9259 | 125.2842 | 21.0536 | 0.001998 | 0.001896 | -151.9311 | 0.23872 | 8.4029 | 0.003666 | 0.004739 | 0.0044 | 6.4e-05 | 0.001284 | 0.25442 | 1.3367 | 2.5481 | 100.58 | 20.64 | 19.031 | 69.6882 | 109.9084 | 67.2469 | 0.004608 | 1.0372 | 0.024934 | 24.041 | -0.1081 | 0.14202 | -1e-06 | -8.4e-05 | -8.4e-05 | -69.6002
2019-03-20 17:44 | 1553118260.0706 | -280 | 0.2095 | 0.000445 | 0 | 108.5514 | 109.8112 | 125.0352 | 21.0542 | 0.001985 | 0.001814 | -151.9324 | 0.23881 | 8.4025 | 0.003662 | 0.005663 | 0.005448 | 6.3e-05 | 0.001292 | 0.28407 | 1.3436 | 2.5481 | 100.58 | 20.64 | 19.031 | 69.3885 | 109.8158 | 65.274 | 0.004608 | 1.0432 | 0.024982 | 23.9474 | -0.10321 | 0.14207 | 1e-06 | 5.1e-05 | 5.1e-05 | 113.7794
2019-03-20 17:44 | 1553118261.071 | -279 | 0.20953 | 0.000447 | 0 | 108.5853 | 109.9323 | 125.2727 | 21.0533 | 0.001983 | 0.001837 | -151.9433 | 0.23927 | 8.4026 | 0.003687 | 0.005171 | 0.004982 | 6.3e-05 | 0.001287 | 0.27704 | 1.3393 | 2.5483 | 100.58 | 20.64 | 19.032 | 69.3376 | 110.4183 | 65.8163 | 0.004608 | 1.0415 | 0.024986 | 23.9893 | -0.11027 | 0.14234 | -1e-06 | -0.0001 | -0.0001 | -58.4588
2019-03-20 17:44 | 1553118262.0716 | -278 | 0.20949 | 0.000444 | 0 | 108.608 | 109.8493 | 125.4384 | 21.054 | 0.001997 | 0.001843 | -151.9409 | 0.23822 | 8.3999 | 0.003742 | 0.005097 | 0.004804 | 6.1e-05 | 0.001282 | 0.27865 | 1.3415 | 2.548 | 100.58 | 20.64 | 19.034 | 69.686 | 111.7558 | 65.9731 | 0.004608 | 1.0411 | 0.024996 | 24.0107 | -0.10165 | 0.14172 | 2e-06 | 0.000154 | 0.000154 | 37.9075
2019-03-20 17:44 | 1553118263.0709 | -277 | 0.20951 | 0.000446 | 0 | 108.6031 | 109.9584 | 125.402 | 21.0547 | 0.001984 | 0.001837 | -151.9497 | 0.23887 | 8.4021 | 0.003729 | 0.004505 | 0.004313 | 6e-05 | 0.001278 | 0.28694 | 1.3208 | 2.5481 | 100.58 | 20.64 | 19.037 | 69.3719 | 111.4393 | 65.8264 | 0.004609 | 1.0415 | 0.024998 | 24.002 | -0.11093 | 0.14211 | 1e-06 | 5.8e-05 | 5.8e-05 | 100.3424
2019-03-20 17:44 | 1553118264.0709 | -276 | 0.20951 | 0.000449 | 0 | 108.6144 | 109.8766 | 125.6184 | 21.0538 | 0.002013 | 0.001862 | -151.9587 | 0.23904 | 8.3998 | 0.00372 | 0.004973 | 0.004694 | 6.3e-05 | 0.001273 | 0.28407 | 1.3293 | 2.5481 | 100.58 | 20.63 | 19.041 | 70.0503 | 111.2356 | 66.4389 | 0.004607 | 1.0396 | 0.024997 | 24.0444 | -0.10334 | 0.14221 | -3e-06 | -0.000233 | -0.000233 | -25.0532
2019-03-20 17:44 | 1553118265.0709 | -275 | 0.20954 | 0.000445 | 0 | 108.6277 | 109.8661 | 125.5572 | 21.0542 | 0.001951 | 0.001875 | -151.944 | 0.23878 | 8.4033 | 0.003742 | 0.005187 | 0.004937 | 6.4e-05 | 0.001277 | 0.25857 | 1.287 | 2.5481 | 100.58 | 20.63 | 19.05 | 68.5782 | 111.7656 | 66.7449 | 0.004609 | 1.0387 | 0.02498 | 24.0494 | -0.1014 | 0.14205 | 2e-06 | 0.000132 | 0.000132 | 44.3726
2019-03-20 17:44 | 1553118266.071 | -274 | 0.20953 | 0.000447 | 0 | 108.5831 | 109.9151 | 125.6743 | 21.0549 | 0.001983 | 0.001802 | -151.9412 | 0.2386 | 8.4037 | 0.003769 | 0.004493 | 0.004253 | 6.5e-05 | 0.001277 | 0.27759 | 1.34 | 2.5481 | 100.58 | 20.63 | 19.054 | 69.3502 | 112.4128 | 64.9882 | 0.00461 | 1.0441 | 0.025056 | 23.9984 | -0.10906 | 0.14194 | -1e-06 | -3.1e-05 | -3.1e-05 | -187.1671
2019-03-20 17:44 | 1553118267.071 | -273 | 0.20952 | 0.000443 | 0 | 108.5808 | 109.9674 | 125.8959 | 21.054 | 0.001989 | 0.001787 | -151.9418 | 0.23883 | 8.4024 | 0.003749 | 0.005373 | 0.005155 | 6.2e-05 | 0.001275 | 0.28031 | 1.3666 | 2.5482 | 100.58 | 20.63 | 19.066 | 69.4734 | 111.9328 | 64.6282 | 0.004613 | 1.0452 | 0.025092 | 24.0067 | -0.1135 | 0.14208 | 5e-06 | 0.000392 | 0.000392 | 14.9096
2019-03-20 17:44 | 1553118268.0721 | -272 | 0.20955 | 0.000447 | 0 | 108.5679 | 110.1206 | 126.0945 | 21.054 | 0.002021 | 0.001802 | -151.9291 | 0.23967 | 8.4004 | 0.003659 | 0.005296 | 0.00504 | 6.6e-05 | 0.001281 | 0.27433 | 1.3801 | 2.548 | 100.58 | 20.63 | 19.071 | 70.2485 | 109.751 | 64.9881 | 0.004614 | 1.0441 | 0.025098 | 24.0384 | -0.12702 | 0.14258 | -1e-06 | -8e-05 | -8e-05 | -73.0262
2019-03-20 17:44 | 1553118269.0713 | -271 | 0.20951 | 0.000449 | 0 | 108.6558 | 109.853 | 126.1161 | 21.0544 | 0.00198 | 0.001859 | -151.9295 | 0.23925 | 8.4004 | 0.003759 | 0.0043 | 0.004094 | 6.4e-05 | 0.001284 | 0.27759 | 1.3587 | 2.5482 | 100.58 | 20.63 | 19.077 | 69.2615 | 112.177 | 66.3629 | 0.004616 | 1.0399 | 0.025049 | 24.0893 | -0.098015 | 0.14233 | -3e-06 | -0.00019 | -0.00019 | -30.8129
2019-03-20 17:44 | 1553118270.1302 | -270 | 0.20956 | 0.000447 | 0 | 108.6377 | 110.0298 | 125.8825 | 21.0538 | 0.001988 | 0.001842 | -151.9342 | 0.23814 | 8.4037 | 0.003739 | 0.00472 | 0.004507 | 6.4e-05 | 0.001288 | 0.28871 | 1.3513 | 2.5481 | 100.58 | 20.63 | 19.09 | 69.4506 | 111.7035 | 65.94 | 0.004619 | 1.0412 | 0.025042 | 24.052 | -0.11389 | 0.14167 | -1e-06 | -0.000119 | -0.000119 | -48.7353
2019-03-20 17:44 | 1553118271.077 | -269 | 0.20955 | 0.000446 | 0 | 108.6442 | 110.0388 | 125.6065 | 21.0543 | 0.001947 | 0.00181 | -151.9319 | 0.23774 | 8.403 | 0.003702 | 0.00547 | 0.005203 | 6.5e-05 | 0.001288 | 0.2688 | 1.3444 | 2.548 | 100.58 | 20.64 | 19.097 | 68.4824 | 110.7853 | 65.1727 | 0.004624 | 1.0435 | 0.025043 | 23.9984 | -0.11409 | 0.14143 | 1e-06 | 0.000109 | 0.000109 | 53.527
2019-03-20 17:45 | 1553118272.0769 | -268 | 0.20951 | 0.000445 | 0 | 108.5918 | 109.7768 | 125.6541 | 21.0546 | 0.001965 | 0.001822 | -151.9322 | 0.23889 | 8.4036 | 0.003709 | 0.005589 | 0.005398 | 6.4e-05 | 0.001298 | 0.27599 | 1.3104 | 2.548 | 100.58 | 20.63 | 19.11 | 68.9171 | 110.9758 | 65.4571 | 0.004624 | 1.0426 | 0.025037 | 24.0131 | -0.097078 | 0.14212 | 1e-06 | 0.000117 | 0.000117 | 50.0267
2019-03-20 17:45 | 1553118273.077 | -267 | 0.20951 | 0.000447 | -1e-06 | 108.6144 | 109.9984 | 125.5262 | 21.0547 | 0.001953 | 0.001829 | -151.9239 | 0.2385 | 8.4038 | 0.003669 | 0.005087 | 0.004832 | 6.4e-05 | 0.001297 | 0.26626 | 1.3162 | 2.5481 | 100.58 | 20.63 | 19.117 | 68.6161 | 109.9826 | 65.6248 | 0.004626 | 1.0421 | 0.025018 | 24.0068 | -0.11325 | 0.14189 | 0 | 7.1e-05 | 7.1e-05 | 82.2926
2019-03-20 17:45 | 1553118274.077 | -266 | 0.20952 | 0.000447 | 0 | 108.5902 | 110.0395 | 125.1628 | 21.0543 | 0.001968 | 0.001876 | -151.9195 | 0.23856 | 8.405 | 0.003617 | 0.005314 | 0.005154 | 6.3e-05 | 0.0013 | 0.27544 | 1.3499 | 2.548 | 100.58 | 20.63 | 19.124 | 68.9761 | 108.7289 | 66.7606 | 0.004627 | 1.0386 | 0.02494 | 24.0122 | -0.11859 | 0.14192 | -1e-06 | -2.4e-05 | -2.4e-05 | -247.6283
2019-03-20 17:45 | 1553118275.0775 | -265 | 0.20951 | 0.000445 | 0 | 108.6384 | 109.9573 | 125.3642 | 21.0547 | 0.001967 | 0.001816 | -151.9191 | 0.23902 | 8.402 | 0.003632 | 0.004802 | 0.004573 | 6.2e-05 | 0.001298 | 0.28567 | 1.3367 | 2.548 | 100.58 | 20.63 | 19.136 | 68.9609 | 109.0808 | 65.3253 | 0.00463 | 1.043 | 0.025013 | 23.9807 | -0.10794 | 0.14219 | 2e-06 | 0.000175 | 0.000175 | 33.3565
2019-03-20 17:45 | 1553118276.0776 | -264 | 0.20949 | 0.000444 | 0 | 108.6549 | 109.8276 | 125.5999 | 21.0552 | 0.002015 | 0.001834 | -151.9417 | 0.23818 | 8.4045 | 0.003604 | 0.005236 | 0.005051 | 6.4e-05 | 0.001285 | 0.27367 | 1.317 | 2.5481 | 100.58 | 20.63 | 19.141 | 70.117 | 108.4145 | 65.7526 | 0.004631 | 1.0417 | 0.025021 | 24.0184 | -0.096025 | 0.1417 | 3e-06 | 0.000199 | 0.000199 | 29.2659
2019-03-20 17:45 | 1553118277.0773 | -263 | 0.2095 | 0.000447 | -1e-06 | 108.6827 | 110.1412 | 125.4403 | 21.0555 | 0.001949 | 0.001857 | -151.9361 | 0.23836 | 8.4011 | 0.00367 | 0.004676 | 0.004388 | 6.6e-05 | 0.001279 | 0.29142 | 1.354 | 2.5481 | 100.58 | 20.63 | 19.145 | 68.5225 | 110.0245 | 66.3094 | 0.004632 | 1.04 | 0.024984 | 24.0228 | -0.11924 | 0.1418 | -1e-06 | -1.4e-05 | -1.4e-05 | -428.1624
2019-03-20 17:45 | 1553118278.0777 | -262 | 0.20951 | 0.000444 | 0 | 108.6374 | 110.0769 | 125.4105 | 21.0549 | 0.001928 | 0.001841 | -151.9359 | 0.23775 | 8.4032 | 0.003639 | 0.005117 | 0.004927 | 6.6e-05 | 0.00128 | 0.26145 | 1.3124 | 2.548 | 100.58 | 20.64 | 19.151 | 68.0281 | 109.2637 | 65.9247 | 0.004637 | 1.0412 | 0.024995 | 24.0063 | -0.11774 | 0.14144 | 2e-06 | 0.000168 | 0.000168 | 34.6483
2019-03-20 17:45 | 1553118279.0775 | -261 | 0.2095 | 0.000443 | 0 | 108.6794 | 109.936 | 125.5149 | 21.0549 | 0.001958 | 0.001888 | -151.9199 | 0.23844 | 8.403 | 0.003683 | 0.004943 | 0.00484 | 6.6e-05 | 0.001282 | 0.26344 | 1.3552 | 2.5481 | 100.58 | 20.64 | 19.153 | 68.731 | 110.3244 | 67.0601 | 0.004637 | 1.0377 | 0.024964 | 24.0565 | -0.10282 | 0.14185 | 4e-06 | 0.000326 | 0.000326 | 17.8865
2019-03-20 17:45 | 1553118280.0774 | -260 | 0.2095 | 0.000445 | 0 | 108.6067 | 109.9188 | 125.729 | 21.0552 | 0.002049 | 0.001892 | -151.9232 | 0.23858 | 8.4026 | 0.003674 | 0.005086 | 0.004811 | 6.5e-05 | 0.001293 | 0.27168 | 1.367 | 2.548 | 100.58 | 20.63 | 19.153 | 70.9209 | 110.1078 | 67.1462 | 0.004634 | 1.0375 | 0.024982 | 24.0801 | -0.10742 | 0.14193 | 1e-06 | 0.000121 | 0.000121 | 48.3744
2019-03-20 17:45 | 1553118281.0789 | -259 | 0.2095 | 0.000446 | 0 | 108.692 | 109.871 | 125.5884 | 21.0546 | 0.00195 | 0.001829 | -151.9412 | 0.23799 | 8.4011 | 0.003639 | 0.00537 | 0.005108 | 6.4e-05 | 0.001295 | 0.2756 | 1.305 | 2.548 | 100.58 | 20.63 | 19.151 | 68.5379 | 109.2627 | 65.6403 | 0.004634 | 1.0421 | 0.025024 | 24.0133 | -0.096494 | 0.14158 | 0 | -5.5e-05 | -5.5e-05 | -105.4855
2019-03-20 17:45 | 1553118282.0776 | -258 | 0.20951 | 0.000444 | 0 | 108.6141 | 110.0762 | 125.4908 | 21.0556 | 0.001953 | 0.001861 | -151.9616 | 0.23896 | 8.4036 | 0.003663 | 0.005515 | 0.005287 | 6.6e-05 | 0.0013 | 0.29822 | 1.3175 | 2.548 | 100.58 | 20.63 | 19.148 | 68.6255 | 109.8591 | 66.417 | 0.004633 | 1.0397 | 0.024985 | 24.0314 | -0.1196 | 0.14216 | 3e-06 | 0.000239 | 0.000239 | 24.4176
2019-03-20 17:45 | 1553118283.0786 | -257 | 0.20952 | 0.000448 | 0 | 108.5928 | 109.8938 | 125.7919 | 21.0555 | 0.001997 | 0.001885 | -151.9678 | 0.23912 | 8.4047 | 0.003706 | 0.005365 | 0.005208 | 6.6e-05 | 0.001293 | 0.29607 | 1.3149 | 2.548 | 100.58 | 20.63 | 19.139 | 69.6736 | 110.8969 | 66.9866 | 0.004631 | 1.038 | 0.024994 | 24.0804 | -0.10652 | 0.14225 | -1e-06 | -7.2e-05 | -7.2e-05 | -81.5536
2019-03-20 17:45 | 1553118284.0787 | -256 | 0.2095 | 0.000443 | 0 | 108.6872 | 110.1214 | 126.0595 | 21.0551 | 0.00199 | 0.001831 | -151.9558 | 0.23832 | 8.4021 | 0.003713 | 0.005647 | 0.005465 | 6.7e-05 | 0.0013 | 0.28727 | 1.2891 | 2.548 | 100.58 | 20.63 | 19.133 | 69.5016 | 111.0694 | 65.6815 | 0.004629 | 1.0419 | 0.025069 | 24.0597 | -0.11726 | 0.14178 | 4e-06 | 0.000337 | 0.000337 | 17.3014
2019-03-20 17:45 | 1553118285.0767 | -255 | 0.2095 | 0.000442 | 0 | 108.6542 | 110.0717 | 126.014 | 21.0543 | 0.001937 | 0.001809 | -151.9604 | 0.23798 | 8.4039 | 0.00382 | 0.005138 | 0.004992 | 6.6e-05 | 0.001301 | 0.26344 | 1.3227 | 2.548 | 100.58 | 20.63 | 19.116 | 68.2288 | 113.6565 | 65.1556 | 0.004625 | 1.0436 | 0.025084 | 24.0367 | -0.11593 | 0.14158 | 5e-06 | 0.000408 | 0.000408 | 14.2704
2019-03-20 17:45 | 1553118286.0772 | -254 | 0.20955 | 0.000444 | 0 | 108.6251 | 109.8149 | 125.9377 | 21.0544 | 0.002012 | 0.001863 | -151.9496 | 0.23839 | 8.4031 | 0.003792 | 0.004864 | 0.004732 | 6.7e-05 | 0.001301 | 0.29551 | 1.3144 | 2.5481 | 100.58 | 20.64 | 19.107 | 70.039 | 112.981 | 66.4651 | 0.004626 | 1.0395 | 0.025028 | 24.0759 | -0.09744 | 0.14182 | 3e-06 | 0.00019 | 0.00019 | 30.6892
2019-03-20 17:45 | 1553118287.0779 | -253 | 0.20949 | 0.000445 | 0 | 108.6464 | 110.0332 | 126.0046 | 21.054 | 0.002008 | 0.001854 | -151.952 | 0.23858 | 8.4055 | 0.003753 | 0.005432 | 0.005086 | 6.8e-05 | 0.001306 | 0.28385 | 1.3103 | 2.548 | 100.58 | 20.64 | 19.098 | 69.9272 | 112.025 | 66.2346 | 0.004624 | 1.0403 | 0.025043 | 24.0741 | -0.11344 | 0.14193 | 2e-06 | 0.0002 | 0.0002 | 29.1425
2019-03-20 17:45 | 1553118288.0778 | -252 | 0.20952 | 0.00044 | 0 | 108.6523 | 109.8859 | 126.0036 | 21.0548 | 0.001979 | 0.001833 | -151.9475 | 0.2384 | 8.4054 | 0.003741 | 0.005604 | 0.005433 | 7e-05 | 0.001308 | 0.29302 | 1.3257 | 2.5481 | 100.58 | 20.64 | 19.079 | 69.2317 | 111.7423 | 65.7407 | 0.004619 | 1.0418 | 0.025061 | 24.0565 | -0.10099 | 0.14182 | 9e-06 | 0.000663 | 0.000663 | 8.7962
2019-03-20 17:45 | 1553118289.0773 | -251 | 0.20951 | 0.000448 | -1e-06 | 108.7839 | 109.8631 | 125.8691 | 21.0544 | 0.001963 | 0.001868 | -151.9556 | 0.23806 | 8.4028 | 0.003682 | 0.004677 | 0.004467 | 6.6e-05 | 0.001301 | 0.27743 | 1.3465 | 2.548 | 100.58 | 20.64 | 19.069 | 68.8567 | 110.3119 | 66.5719 | 0.004617 | 1.0392 | 0.025017 | 24.0731 | -0.0883 | 0.14162 | -1e-06 | -7e-05 | -7e-05 | -83.2503
2019-03-20 17:45 | 1553118290.0771 | -250 | 0.20954 | 0.000444 | 0 | 108.6228 | 110.0167 | 125.7961 | 21.0549 | 0.001982 | 0.001827 | -151.9515 | 0.23719 | 8.4028 | 0.003739 | 0.004942 | 0.004757 | 6.9e-05 | 0.001301 | 0.28368 | 1.3252 | 2.548 | 100.58 | 20.64 | 19.06 | 69.3236 | 111.6833 | 65.5885 | 0.004615 | 1.0422 | 0.025046 | 24.0313 | -0.11405 | 0.14111 | 2e-06 | 0.000156 | 0.000156 | 37.2925
2019-03-20 17:45 | 1553118291.0771 | -249 | 0.20953 | 0.000447 | 0 | 108.6972 | 110.1315 | 125.9433 | 21.0543 | 0.00202 | 0.001841 | -151.9343 | 0.23772 | 8.4034 | 0.003682 | 0.005372 | 0.005175 | 7.2e-05 | 0.0013 | 0.286 | 1.3292 | 2.5481 | 100.58 | 20.64 | 19.044 | 70.2293 | 110.3058 | 65.9171 | 0.004611 | 1.0412 | 0.025049 | 24.057 | -0.11725 | 0.14142 | -1e-06 | -0.000107 | -0.000107 | -54.3214
2019-03-20 17:45 | 1553118292.0784 | -248 | 0.20949 | 0.000445 | -1e-06 | 108.6481 | 109.9767 | 125.8701 | 21.0551 | 0.001921 | 0.001925 | -151.934 | 0.23813 | 8.4011 | 0.003818 | 0.00508 | 0.004818 | 6.8e-05 | 0.001294 | 0.28672 | 1.3361 | 2.548 | 100.58 | 20.64 | 19.039 | 67.8552 | 113.6123 | 67.9372 | 0.00461 | 1.0351 | 0.024967 | 24.1215 | -0.10872 | 0.14166 | 3e-06 | 0.000242 | 0.000242 | 24.0439
2019-03-20 17:45 | 1553118293.0775 | -247 | 0.20954 | 0.000445 | 0 | 108.7179 | 110.1606 | 125.7607 | 21.0557 | 0.001969 | 0.001834 | -151.9266 | 0.23755 | 8.4 | 0.003752 | 0.004897 | 0.004667 | 7.2e-05 | 0.001292 | 0.28711 | 1.3389 | 2.5481 | 100.58 | 20.64 | 19.03 | 68.9977 | 112.0172 | 65.7632 | 0.004607 | 1.0417 | 0.025036 | 24.0341 | -0.11792 | 0.14132 | 2e-06 | 0.000136 | 0.000136 | 42.8422
2019-03-20 17:45 | 1553118294.0782 | -246 | 0.20949 | 0.000449 | 0 | 108.6833 | 110.0463 | 125.7595 | 21.0548 | 0.002008 | 0.001814 | -151.9254 | 0.23836 | 8.4007 | 0.003699 | 0.005266 | 0.004989 | 6.7e-05 | 0.001288 | 0.26864 | 1.3482 | 2.5481 | 100.58 | 20.64 | 19.028 | 69.9366 | 110.7299 | 65.2754 | 0.004607 | 1.0432 | 0.025054 | 24.0167 | -0.11147 | 0.1418 | -3e-06 | -0.000191 | -0.000191 | -30.4472
2019-03-20 17:45 | 1553118295.0776 | -245 | 0.20951 | 0.000446 | 1e-06 | 108.6561 | 110.1094 | 125.8229 | 21.0559 | 0.001982 | 0.001861 | -151.9456 | 0.23845 | 8.4051 | 0.003672 | 0.005429 | 0.005226 | 7.1e-05 | 0.001292 | 0.27831 | 1.3209 | 2.5481 | 100.58 | 20.64 | 19.027 | 69.3175 | 110.0572 | 66.3978 | 0.004607 | 1.0398 | 0.025019 | 24.0625 | -0.11884 | 0.14185 | -1e-06 | -0.000127 | -0.000127 | -45.7451
2019-03-20 17:45 | 1553118296.0777 | -244 | 0.20951 | 0.000448 | 0 | 108.7208 | 110.0851 | 125.6365 | 21.0553 | 0.001929 | 0.001826 | -151.9466 | 0.23873 | 8.4027 | 0.00368 | 0.005091 | 0.004868 | 7e-05 | 0.001292 | 0.26985 | 1.3291 | 2.548 | 100.58 | 20.64 | 19.028 | 68.0372 | 110.2633 | 65.5624 | 0.004607 | 1.0423 | 0.025031 | 24.0151 | -0.11154 | 0.14202 | -3e-06 | -0.000261 | -0.000261 | -22.4059
2019-03-20 17:45 | 1553118297.0773 | -243 | 0.20953 | 0.000451 | 0 | 108.6655 | 109.9749 | 125.4455 | 21.0553 | 0.001958 | 0.001755 | -151.9498 | 0.23963 | 8.4015 | 0.003654 | 0.004948 | 0.004703 | 6.9e-05 | 0.00129 | 0.27024 | 1.3318 | 2.5479 | 100.58 | 20.64 | 19.031 | 68.7336 | 109.6198 | 63.8668 | 0.004608 | 1.0476 | 0.025075 | 23.9367 | -0.10713 | 0.14256 | -6e-06 | -0.000508 | -0.000508 | -11.5233
2019-03-20 17:45 | 1553118298.0787 | -242 | 0.20952 | 0.000448 | 0 | 108.6164 | 109.9162 | 125.5546 | 21.0557 | 0.001911 | 0.001807 | -151.9694 | 0.23841 | 8.4028 | 0.00375 | 0.005411 | 0.005261 | 6.6e-05 | 0.001293 | 0.26322 | 1.3628 | 2.5481 | 100.58 | 20.64 | 19.041 | 67.6024 | 111.9714 | 65.1107 | 0.00461 | 1.0437 | 0.02504 | 23.9913 | -0.1064 | 0.14183 | -1e-06 | -9.8e-05 | -9.8e-05 | -59.7138
2019-03-20 17:45 | 1553118299.0786 | -241 | 0.20954 | 0.000449 | 0 | 108.6904 | 109.9779 | 125.433 | 21.0553 | 0.002005 | 0.001837 | -151.9684 | 0.23841 | 8.4028 | 0.003654 | 0.005028 | 0.004809 | 7e-05 | 0.001283 | 0.27256 | 1.3415 | 2.5481 | 100.58 | 20.64 | 19.048 | 69.87 | 109.6252 | 65.8197 | 0.004612 | 1.0415 | 0.025002 | 24.0048 | -0.10533 | 0.14183 | -3e-06 | -0.000232 | -0.000232 | -25.0966
2019-03-20 17:45 | 1553118300.0775 | -240 | 0.20949 | 0.000445 | 0 | 108.7005 | 110.17 | 125.4555 | 21.0558 | 0.001942 | 0.001812 | -151.9674 | 0.23854 | 8.4028 | 0.003706 | 0.004747 | 0.004511 | 6.9e-05 | 0.001275 | 0.27079 | 1.3859 | 2.5481 | 100.58 | 20.64 | 19.056 | 68.3459 | 110.8831 | 65.2357 | 0.004614 | 1.0433 | 0.025025 | 23.9862 | -0.12011 | 0.14191 | 1e-06 | 6.5e-05 | 6.5e-05 | 89.5993
2019-03-20 17:45 | 1553118301.0775 | -239 | 0.2095 | 0.000447 | 0 | 108.715 | 110.0549 | 125.4095 | 21.0547 | 0.001993 | 0.001783 | -151.9634 | 0.23859 | 8.4045 | 0.003745 | 0.004463 | 0.004289 | 6.8e-05 | 0.001275 | 0.29645 | 1.3062 | 2.5481 | 100.58 | 20.64 | 19.074 | 69.5815 | 111.8339 | 64.5321 | 0.004618 | 1.0455 | 0.025047 | 23.9569 | -0.10956 | 0.14194 | -2e-06 | -0.000136 | -0.000136 | -42.8127
2019-03-20 17:45 | 1553118302.0766 | -238 | 0.20952 | 0.000446 | 0 | 108.8017 | 109.9958 | 125.8065 | 21.0562 | 0.001943 | 0.001854 | -151.943 | 0.23878 | 8.4038 | 0.003712 | 0.005408 | 0.005193 | 6.9e-05 | 0.001281 | 0.28102 | 1.3441 | 2.5481 | 100.58 | 20.64 | 19.082 | 68.3671 | 111.0396 | 66.2475 | 0.00462 | 1.0402 | 0.025023 | 24.0556 | -0.097633 | 0.14205 | 1e-06 | 5.6e-05 | 5.6e-05 | 103.8409
2019-03-20 17:45 | 1553118303.0781 | -237 | 0.20946 | 0.000444 | 0 | 108.7347 | 110.1718 | 125.6769 | 21.0556 | 0.001962 | 0.001849 | -151.9489 | 0.23843 | 8.4036 | 0.003709 | 0.005029 | 0.00487 | 6.8e-05 | 0.001283 | 0.27959 | 1.315 | 2.5481 | 100.58 | 20.64 | 19.1 | 68.8409 | 110.9659 | 66.1198 | 0.004624 | 1.0406 | 0.025015 | 24.0387 | -0.11744 | 0.14184 | 2e-06 | 0.000125 | 0.000125 | 46.817
2019-03-20 17:45 | 1553118304.0777 | -236 | 0.20949 | 0.000448 | 0 | 108.8111 | 110.2021 | 125.6673 | 21.0553 | 0.001969 | 0.001853 | -151.9374 | 0.23915 | 8.402 | 0.003725 | 0.004961 | 0.004748 | 6.9e-05 | 0.00128 | 0.262 | 1.3282 | 2.5481 | 100.58 | 20.64 | 19.108 | 69.0129 | 111.3552 | 66.2133 | 0.004626 | 1.0403 | 0.02501 | 24.0411 | -0.11362 | 0.14227 | -2e-06 | -0.000168 | -0.000168 | -34.7439
2019-03-20 17:45 | 1553118305.0771 | -235 | 0.20952 | 0.000442 | 0 | 108.7842 | 110.1001 | 125.8187 | 21.0557 | 0.002032 | 0.00189 | -151.924 | 0.23798 | 8.4056 | 0.003716 | 0.005053 | 0.004804 | 6.4e-05 | 0.001274 | 0.29767 | 1.3138 | 2.5479 | 100.58 | 20.64 | 19.116 | 70.5246 | 111.1379 | 67.1061 | 0.004628 | 1.0376 | 0.024993 | 24.0872 | -0.10755 | 0.14158 | 6e-06 | 0.000424 | 0.000424 | 13.7335
2019-03-20 17:45 | 1553118306.0768 | -234 | 0.20947 | 0.000448 | 0 | 108.7894 | 110.3284 | 125.9226 | 21.0557 | 0.001978 | 0.00189 | -151.9006 | 0.23863 | 8.4021 | 0.003701 | 0.004435 | 0.00426 | 6.6e-05 | 0.00126 | 0.286 | 1.3297 | 2.5482 | 100.58 | 20.64 | 19.131 | 69.2215 | 110.7828 | 67.1054 | 0.004632 | 1.0376 | 0.025003 | 24.0972 | -0.12565 | 0.14196 | -3e-06 | -0.000236 | -0.000236 | -24.6751
2019-03-20 17:45 | 1553118307.0779 | -233 | 0.20949 | 0.000445 | 0 | 108.8525 | 109.9947 | 125.921 | 21.0552 | 0.001993 | 0.00184 | -151.8908 | 0.23834 | 8.403 | 0.003707 | 0.00484 | 0.004635 | 7.1e-05 | 0.001257 | 0.26377 | 1.344 | 2.5482 | 100.58 | 20.65 | 19.139 | 69.5899 | 110.9117 | 65.8961 | 0.004637 | 1.0413 | 0.025047 | 24.0541 | -0.093369 | 0.14179 | 2e-06 | 0.000165 | 0.000165 | 35.3656
2019-03-20 17:45 | 1553118308.0783 | -232 | 0.20954 | 0.000447 | 0 | 108.8664 | 110.2002 | 125.9172 | 21.0556 | 0.002026 | 0.001811 | -151.8806 | 0.23852 | 8.403 | 0.003713 | 0.005731 | 0.005399 | 6.4e-05 | 0.001249 | 0.25536 | 1.3487 | 2.5481 | 100.58 | 20.65 | 19.147 | 70.3647 | 111.0699 | 65.2059 | 0.004639 | 1.0434 | 0.025072 | 24.0293 | -0.10893 | 0.1419 | -1e-06 | -0.000117 | -0.000117 | -49.8328
2019-03-20 17:45 | 1553118309.078 | -231 | 0.20948 | 0.000442 | 0 | 108.9712 | 110.2391 | 126.0236 | 21.0558 | 0.002027 | 0.001859 | -151.8816 | 0.23831 | 8.4024 | 0.003793 | 0.005067 | 0.004967 | 6.3e-05 | 0.001253 | 0.2724 | 1.3445 | 2.548 | 100.58 | 20.65 | 19.161 | 70.3881 | 113.0014 | 66.3674 | 0.004642 | 1.0398 | 0.02504 | 24.0806 | -0.10348 | 0.14177 | 6e-06 | 0.000466 | 0.000466 | 12.5124
2019-03-20 17:45 | 1553118310.078 | -230 | 0.2095 | 0.00045 | 0 | 109.0413 | 110.3359 | 126.0196 | 21.0552 | 0.001977 | 0.001825 | -151.8782 | 0.23824 | 8.4014 | 0.003712 | 0.0058 | 0.005565 | 6.7e-05 | 0.001257 | 0.25625 | 1.3603 | 2.5481 | 100.58 | 20.65 | 19.166 | 69.1946 | 111.0499 | 65.5488 | 0.004643 | 1.0424 | 0.02507 | 24.0512 | -0.10557 | 0.14173 | -4e-06 | -0.000269 | -0.000269 | -21.6327
2019-03-20 17:45 | 1553118311.077 | -229 | 0.20951 | 0.000447 | 0 | 108.9365 | 110.2899 | 125.6386 | 21.0553 | 0.001954 | 0.001809 | -151.8848 | 0.2377 | 8.4016 | 0.003737 | 0.005373 | 0.005223 | 6.5e-05 | 0.001261 | 0.28064 | 1.3258 | 2.548 | 100.58 | 20.65 | 19.175 | 68.6408 | 111.6379 | 65.1478 | 0.004645 | 1.0436 | 0.025047 | 24.0006 | -0.11044 | 0.14141 | -1e-06 | -6.6e-05 | -6.6e-05 | -87.9417
2019-03-20 17:45 | 1553118312.0772 | -228 | 0.20949 | 0.000444 | 0 | 109.041 | 110.3367 | 125.2403 | 21.0551 | 0.001913 | 0.001845 | -151.9031 | 0.23902 | 8.4018 | 0.003724 | 0.005045 | 0.004896 | 6.8e-05 | 0.001267 | 0.27129 | 1.3268 | 2.5479 | 100.58 | 20.65 | 19.18 | 67.6484 | 111.3244 | 66.0125 | 0.004647 | 1.0409 | 0.024975 | 23.9931 | -0.10565 | 0.14219 | 3e-06 | 0.000208 | 0.000208 | 28.1218
2019-03-20 17:45 | 1553118313.0781 | -227 | 0.2095 | 0.000446 | 0 | 109.0355 | 110.3924 | 125.0535 | 21.0562 | 0.001952 | 0.001829 | -151.8919 | 0.23868 | 8.4039 | 0.003703 | 0.00517 | 0.004869 | 6.6e-05 | 0.001266 | 0.29927 | 1.3411 | 2.548 | 100.58 | 20.65 | 19.183 | 68.6059 | 110.8195 | 65.6467 | 0.004647 | 1.0421 | 0.02497 | 23.9623 | -0.11062 | 0.14199 | 0 | -3.7e-05 | -3.7e-05 | -156.8377
2019-03-20 17:45 | 1553118314.0776 | -226 | 0.20951 | 0.000446 | 0 | 109.0433 | 110.2862 | 125.0363 | 21.056 | 0.001948 | 0.001831 | -151.9134 | 0.23922 | 8.403 | 0.003711 | 0.004972 | 0.004855 | 6.7e-05 | 0.001261 | 0.26897 | 1.3235 | 2.5482 | 100.58 | 20.65 | 19.186 | 68.5011 | 111.0049 | 65.6956 | 0.004648 | 1.0419 | 0.024967 | 23.9624 | -0.10138 | 0.14231 | 0 | -1.5e-05 | -1.5e-05 | -396.4631
2019-03-20 17:45 | 1553118315.0773 | -225 | 0.20952 | 0.000449 | 0 | 109.129 | 110.3116 | 124.6019 | 21.0562 | 0.002019 | 0.001805 | -151.9038 | 0.23824 | 8.4041 | 0.003717 | 0.004956 | 0.004709 | 6.7e-05 | 0.001267 | 0.27007 | 1.2968 | 2.5481 | 100.58 | 20.65 | 19.192 | 70.2076 | 111.1707 | 65.0628 | 0.004649 | 1.0439 | 0.024946 | 23.8984 | -0.096412 | 0.14173 | -3e-06 | -0.000205 | -0.000205 | -28.3756
2019-03-20 17:45 | 1553118316.0783 | -224 | 0.20953 | 0.000446 | 0 | 109.1306 | 110.5026 | 124.3225 | 21.0551 | 0.001995 | 0.001802 | -151.9264 | 0.23802 | 8.401 | 0.003739 | 0.005375 | 0.005203 | 6.6e-05 | 0.001264 | 0.26737 | 1.3204 | 2.5481 | 100.58 | 20.65 | 19.194 | 69.6396 | 111.7032 | 64.9915 | 0.00465 | 1.0441 | 0.024921 | 23.8691 | -0.11175 | 0.14159 | 1e-06 | 7.2e-05 | 7.2e-05 | 81.1177
2019-03-20 17:45 | 1553118317.0787 | -223 | 0.20951 | 0.00045 | 0 | 109.0701 | 110.4944 | 124.2859 | 21.0565 | 0.002012 | 0.001837 | -151.9265 | 0.23885 | 8.4013 | 0.003686 | 0.005296 | 0.005124 | 6.7e-05 | 0.001271 | 0.26504 | 1.3895 | 2.5481 | 100.58 | 20.65 | 19.196 | 70.0262 | 110.4176 | 65.8373 | 0.00465 | 1.0415 | 0.024886 | 23.8954 | -0.11604 | 0.14209 | -4e-06 | -0.000318 | -0.000318 | -18.3696
2019-03-20 17:45 | 1553118318.0765 | -222 | 0.20952 | 0.000446 | 0 | 109.1296 | 110.6319 | 124.0162 | 21.0552 | 0.002007 | 0.001809 | -151.9474 | 0.23825 | 8.4011 | 0.003709 | 0.00546 | 0.005281 | 6.4e-05 | 0.001269 | 0.27273 | 1.3151 | 2.548 | 100.58 | 20.65 | 19.201 | 69.9214 | 110.9598 | 65.1593 | 0.004652 | 1.0436 | 0.024884 | 23.8456 | -0.12229 | 0.14173 | 0 | -7e-06 | -7e-06 | -808.2305
2019-03-20 17:45 | 1553118319.0775 | -221 | 0.20952 | 0.000444 | 0 | 109.1869 | 110.5676 | 123.8226 | 21.0564 | 0.00198 | 0.001881 | -151.9338 | 0.2385 | 8.4031 | 0.003664 | 0.004389 | 0.004284 | 6.5e-05 | 0.001269 | 0.26305 | 1.3445 | 2.548 | 100.58 | 20.65 | 19.202 | 69.2666 | 109.8732 | 66.8791 | 0.004652 | 1.0383 | 0.024802 | 23.8874 | -0.1124 | 0.14188 | 2e-06 | 0.000147 | 0.000147 | 39.6506
2019-03-20 17:45 | 1553118320.0769 | -220 | 0.20952 | 0.00045 | 0 | 109.1882 | 110.6925 | 123.6503 | 21.0568 | 0.001976 | 0.001862 | -151.9365 | 0.23881 | 8.4051 | 0.003626 | 0.004704 | 0.004337 | 6.4e-05 | 0.001275 | 0.27024 | 1.2995 | 2.5481 | 100.58 | 20.65 | 19.205 | 69.179 | 108.9449 | 66.4362 | 0.004653 | 1.0396 | 0.024801 | 23.8552 | -0.12239 | 0.14207 | -6e-06 | -0.000464 | -0.000464 | -12.586
2019-03-20 17:45 | 1553118321.0766 | -219 | 0.20951 | 0.000447 | 0 | 109.2959 | 110.5777 | 124.0885 | 21.0556 | 0.001985 | 0.001835 | -151.9245 | 0.23799 | 8.4019 | 0.003575 | 0.00529 | 0.005064 | 6.5e-05 | 0.001275 | 0.24746 | 1.3642 | 2.5481 | 100.58 | 20.66 | 19.205 | 69.386 | 107.7041 | 65.7693 | 0.004656 | 1.0417 | 0.024869 | 23.874 | -0.10429 | 0.14158 | -2e-06 | -0.000156 | -0.000156 | -37.3353
2019-03-20 17:45 | 1553118322.0766 | -218 | 0.20957 | 0.000447 | 0 | 109.1908 | 110.4529 | 123.977 | 21.056 | 0.001975 | 0.001819 | -151.9215 | 0.23973 | 8.4029 | 0.003619 | 0.005526 | 0.005329 | 6.6e-05 | 0.001275 | 0.26305 | 1.3413 | 2.548 | 100.58 | 20.66 | 19.206 | 69.1384 | 108.7777 | 65.4014 | 0.004656 | 1.0428 | 0.024871 | 23.8503 | -0.1028 | 0.14262 | -1e-06 | -4.1e-05 | -4.1e-05 | -143.9781
2019-03-20 17:45 | 1553118323.0781 | -217 | 0.20952 | 0.000454 | 0 | 109.2286 | 110.6704 | 124.101 | 21.0559 | 0.001995 | 0.001842 | -151.9346 | 0.23885 | 8.4022 | 0.003641 | 0.005242 | 0.005038 | 6.6e-05 | 0.001286 | 0.27793 | 1.3138 | 2.5482 | 100.58 | 20.66 | 19.206 | 69.6163 | 109.316 | 65.9586 | 0.004656 | 1.0411 | 0.024863 | 23.8819 | -0.1173 | 0.14209 | -9e-06 | -0.000692 | -0.000692 | -8.4336
2019-03-20 17:45 | 1553118324.0768 | -216 | 0.20951 | 0.000455 | 1e-06 | 109.3146 | 110.7586 | 124.0965 | 21.0555 | 0.002011 | 0.00184 | -151.9511 | 0.23836 | 8.4022 | 0.003626 | 0.004563 | 0.004363 | 6.4e-05 | 0.001283 | 0.26737 | 1.33 | 2.548 | 100.58 | 20.66 | 19.207 | 70.0214 | 108.9404 | 65.9027 | 0.004656 | 1.0413 | 0.024865 | 23.8795 | -0.11738 | 0.1418 | -1.2e-05 | -0.000938 | -0.000938 | -6.2107
2019-03-20 17:45 | 1553118325.0778 | -215 | 0.20952 | 0.000451 | 0 | 109.3389 | 110.7138 | 124.1812 | 21.0567 | 0.002022 | 0.001842 | -151.9498 | 0.23846 | 8.4041 | 0.003643 | 0.005062 | 0.004877 | 6.4e-05 | 0.001281 | 0.27345 | 1.2968 | 2.5479 | 100.58 | 20.66 | 19.207 | 70.2821 | 109.3687 | 65.9601 | 0.004656 | 1.0411 | 0.024871 | 23.8896 | -0.11177 | 0.14186 | -6e-06 | -0.000414 | -0.000414 | -14.0821
2019-03-20 17:45 | 1553118326.0779 | -214 | 0.20953 | 0.000445 | 1e-06 | 109.3402 | 110.6794 | 124.1887 | 21.0559 | 0.001991 | 0.001843 | -151.9632 | 0.23848 | 8.4023 | 0.003668 | 0.005151 | 0.004911 | 6.7e-05 | 0.001276 | 0.25514 | 1.3545 | 2.5481 | 100.58 | 20.66 | 19.206 | 69.5406 | 109.9807 | 65.9664 | 0.004656 | 1.0411 | 0.024872 | 23.8906 | -0.10889 | 0.14187 | 0 | -1e-06 | -1e-06 | -7416.5122
2019-03-20 17:45 | 1553118327.077 | -213 | 0.20953 | 0.000455 | 0 | 109.3955 | 110.8293 | 124.4049 | 21.0556 | 0.002035 | 0.001847 | -151.9517 | 0.23842 | 8.4032 | 0.003581 | 0.004716 | 0.004513 | 6.9e-05 | 0.001277 | 0.27505 | 1.3465 | 2.5481 | 100.58 | 20.66 | 19.205 | 70.5828 | 107.8507 | 66.0637 | 0.004656 | 1.0408 | 0.02489 | 23.9148 | -0.11647 | 0.14184 | -1.1e-05 | -0.000761 | -0.000761 | -7.6596
2019-03-20 17:45 | 1553118328.078 | -212 | 0.20953 | 0.000453 | 0 | 109.3932 | 111.0199 | 124.6373 | 21.056 | 0.001964 | 0.001812 | -151.9349 | 0.23855 | 8.4027 | 0.003581 | 0.004542 | 0.004352 | 6.8e-05 | 0.001278 | 0.27417 | 1.3721 | 2.548 | 100.58 | 20.67 | 19.199 | 68.8725 | 107.8645 | 65.2351 | 0.004657 | 1.0433 | 0.024944 | 23.9079 | -0.13203 | 0.14191 | -8e-06 | -0.000588 | -0.000588 | -9.9218
2019-03-20 17:45 | 1553118329.0771 | -211 | 0.20955 | 0.000446 | 0 | 109.4249 | 110.8214 | 125.0091 | 21.0564 | 0.001933 | 0.001876 | -151.9435 | 0.23839 | 8.4025 | 0.003643 | 0.004809 | 0.004497 | 6.7e-05 | 0.001279 | 0.27599 | 1.3619 | 2.548 | 100.58 | 20.66 | 19.195 | 68.1442 | 109.3614 | 66.7605 | 0.004653 | 1.0386 | 0.024925 | 23.9974 | -0.11343 | 0.14182 | 1e-06 | 5.2e-05 | 5.2e-05 | 111.4013
2019-03-20 17:45 | 1553118330.0785 | -210 | 0.20955 | 0.000453 | 0 | 109.4712 | 110.7519 | 125.0514 | 21.0558 | 0.002005 | 0.001909 | -151.9308 | 0.23824 | 8.4047 | 0.003624 | 0.005279 | 0.005139 | 6.9e-05 | 0.001282 | 0.26034 | 1.322 | 2.548 | 100.58 | 20.67 | 19.19 | 69.8799 | 108.9028 | 67.5726 | 0.004655 | 1.0362 | 0.024899 | 24.0301 | -0.10404 | 0.14173 | -9e-06 | -0.000662 | -0.000662 | -8.8047
2019-03-20 17:46 | 1553118331.0781 | -209 | 0.20954 | 0.000451 | 0 | 109.4566 | 110.9769 | 124.8493 | 21.0562 | 0.002012 | 0.001861 | -151.9313 | 0.23851 | 8.4026 | 0.00362 | 0.0051 | 0.004821 | 6.7e-05 | 0.001279 | 0.28047 | 1.3449 | 2.5479 | 100.58 | 20.67 | 19.179 | 70.0338 | 108.814 | 66.4169 | 0.004652 | 1.0397 | 0.024921 | 23.9699 | -0.12338 | 0.14189 | -6e-06 | -0.000424 | -0.000424 | -13.7576
2019-03-20 17:46 | 1553118332.0766 | -208 | 0.20953 | 0.000451 | 0 | 109.5763 | 111.1249 | 125.0441 | 21.0553 | 0.00199 | 0.001877 | -151.9178 | 0.23782 | 8.4029 | 0.003659 | 0.004521 | 0.004249 | 6.5e-05 | 0.001285 | 0.26985 | 1.311 | 2.548 | 100.58 | 20.67 | 19.174 | 69.5049 | 109.7576 | 66.7945 | 0.004651 | 1.0385 | 0.024927 | 24.002 | -0.12553 | 0.14148 | -6e-06 | -0.000461 | -0.000461 | -12.6165
2019-03-20 17:46 | 1553118333.0776 | -207 | 0.20948 | 0.000448 | 0 | 109.5662 | 111.0341 | 125.2788 | 21.0565 | 0.001992 | 0.00182 | -151.9145 | 0.23854 | 8.4027 | 0.003765 | 0.005272 | 0.005065 | 6.6e-05 | 0.001285 | 0.24834 | 1.3473 | 2.548 | 100.58 | 20.67 | 19.163 | 69.5529 | 112.326 | 65.4168 | 0.004648 | 1.0428 | 0.025001 | 23.9758 | -0.11903 | 0.14191 | -2e-06 | -0.000125 | -0.000125 | -46.6939
2019-03-20 17:46 | 1553118334.0763 | -206 | 0.20952 | 0.000445 | 0 | 109.5908 | 111.1466 | 125.395 | 21.0549 | 0.001957 | 0.001894 | -151.9036 | 0.23717 | 8.4029 | 0.003709 | 0.005024 | 0.004882 | 6.7e-05 | 0.00129 | 0.25514 | 1.3091 | 2.5479 | 100.58 | 20.67 | 19.159 | 68.7068 | 110.9596 | 67.2068 | 0.004647 | 1.0373 | 0.024947 | 24.0502 | -0.12609 | 0.14109 | 1e-06 | 8.8e-05 | 8.8e-05 | 65.6729
2019-03-20 17:46 | 1553118335.0777 | -205 | 0.20955 | 0.000444 | 0 | 109.6681 | 111.1522 | 125.4797 | 21.0566 | 0.001926 | 0.001826 | -151.9213 | 0.23835 | 8.4022 | 0.003724 | 0.004783 | 0.004551 | 6.4e-05 | 0.001295 | 0.26626 | 1.3322 | 2.5479 | 100.58 | 20.67 | 19.155 | 67.9589 | 111.3242 | 65.573 | 0.004646 | 1.0423 | 0.025015 | 24.0005 | -0.12023 | 0.14179 | 2e-06 | 0.000144 | 0.000144 | 40.3368
2019-03-20 17:46 | 1553118336.0768 | -204 | 0.2095 | 0.00045 | 0 | 109.5591 | 111.0397 | 125.7806 | 21.0562 | 0.001987 | 0.001858 | -151.9086 | 0.23722 | 8.4046 | 0.003749 | 0.004933 | 0.004821 | 7e-05 | 0.001295 | 0.29391 | 1.3114 | 2.5481 | 100.58 | 20.67 | 19.148 | 69.4467 | 111.9459 | 66.3435 | 0.004645 | 1.0399 | 0.025017 | 24.0566 | -0.12007 | 0.14112 | -4e-06 | -0.0003 | -0.0003 | -19.3564
2019-03-20 17:46 | 1553118337.0775 | -203 | 0.20952 | 0.000447 | 0 | 109.66 | 111.1406 | 125.4398 | 21.0567 | 0.001999 | 0.001831 | -151.9179 | 0.23777 | 8.4028 | 0.003725 | 0.005408 | 0.005163 | 6.8e-05 | 0.001286 | 0.2724 | 1.3592 | 2.5481 | 100.58 | 20.67 | 19.145 | 69.7137 | 111.3543 | 65.6799 | 0.004644 | 1.042 | 0.025007 | 24.0005 | -0.11996 | 0.14145 | -1e-06 | -0.000111 | -0.000111 | -52.2079
2019-03-20 17:46 | 1553118338.0767 | -202 | 0.20953 | 0.000448 | 1e-06 | 109.6274 | 111.1062 | 125.6973 | 21.0565 | 0.002011 | 0.001836 | -151.9237 | 0.23731 | 8.4017 | 0.003786 | 0.004777 | 0.004567 | 6.8e-05 | 0.00128 | 0.28926 | 1.3658 | 2.5481 | 100.58 | 20.67 | 19.141 | 70.0192 | 112.8512 | 65.8039 | 0.004643 | 1.0416 | 0.025028 | 24.0295 | -0.11985 | 0.14118 | -3e-06 | -0.000253 | -0.000253 | -22.9253
2019-03-20 17:46 | 1553118339.0766 | -201 | 0.20948 | 0.000448 | 0 | 109.7613 | 111.1466 | 125.3605 | 21.0565 | 0.001918 | 0.00185 | -151.9193 | 0.23775 | 8.3988 | 0.00382 | 0.005321 | 0.005169 | 6.7e-05 | 0.001278 | 0.274 | 1.3566 | 2.548 | 100.58 | 20.67 | 19.14 | 67.777 | 113.6675 | 66.1502 | 0.004643 | 1.0405 | 0.024982 | 24.0095 | -0.11219 | 0.14144 | -3e-06 | -0.000226 | -0.000226 | -25.7504
2019-03-20 17:46 | 1553118340.0767 | -200 | 0.2095 | 0.000449 | 0 | 109.7748 | 111.3181 | 125.4649 | 21.0563 | 0.001942 | 0.001863 | -151.9352 | 0.23851 | 8.4041 | 0.003853 | 0.005401 | 0.005178 | 6.6e-05 | 0.001272 | 0.27759 | 1.3534 | 2.5482 | 100.58 | 20.67 | 19.141 | 68.3446 | 114.4664 | 66.4545 | 0.004643 | 1.0396 | 0.024981 | 24.0303 | -0.12487 | 0.14189 | -4e-06 | -0.000294 | -0.000294 | -19.853
2019-03-20 17:46 | 1553118341.0767 | -199 | 0.20951 | 0.00045 | 0 | 109.8023 | 111.2972 | 125.6358 | 21.0565 | 0.001976 | 0.001824 | -151.9252 | 0.23886 | 8.4034 | 0.003798 | 0.005505 | 0.005246 | 6.8e-05 | 0.001278 | 0.25713 | 1.3634 | 2.548 | 100.58 | 20.67 | 19.146 | 69.1809 | 113.1261 | 65.5135 | 0.004644 | 1.0425 | 0.025033 | 24.0133 | -0.12095 | 0.1421 | -4e-06 | -0.000298 | -0.000298 | -19.5862
2019-03-20 17:46 | 1553118342.0781 | -198 | 0.20951 | 0.000449 | 0 | 109.7794 | 111.3361 | 125.572 | 21.0562 | 0.001969 | 0.001857 | -151.9376 | 0.23772 | 8.4017 | 0.003843 | 0.004938 | 0.004682 | 6.7e-05 | 0.001279 | 0.27455 | 1.3495 | 2.548 | 100.58 | 20.67 | 19.15 | 69.002 | 114.2402 | 66.3202 | 0.004645 | 1.04 | 0.024997 | 24.0358 | -0.12595 | 0.14142 | -3e-06 | -0.000263 | -0.000263 | -22.1048
2019-03-20 17:46 | 1553118343.0768 | -197 | 0.20951 | 0.000443 | 0 | 109.8761 | 111.351 | 125.809 | 21.0567 | 0.001959 | 0.001809 | -151.9335 | 0.23879 | 8.4032 | 0.003705 | 0.005543 | 0.005372 | 6.7e-05 | 0.001281 | 0.26322 | 1.3113 | 2.5479 | 100.58 | 20.67 | 19.157 | 68.7686 | 110.865 | 65.157 | 0.004647 | 1.0436 | 0.025064 | 24.0172 | -0.11927 | 0.14205 | 4e-06 | 0.000299 | 0.000299 | 19.5066
2019-03-20 17:46 | 1553118344.0772 | -196 | 0.20949 | 0.000448 | 0 | 109.8877 | 111.4374 | 125.8074 | 21.0567 | 0.001941 | 0.00185 | -151.9162 | 0.23843 | 8.401 | 0.003776 | 0.004807 | 0.004557 | 6.7e-05 | 0.001269 | 0.25465 | 1.3038 | 2.5479 | 100.58 | 20.67 | 19.16 | 68.3373 | 112.5915 | 66.1318 | 0.004647 | 1.0406 | 0.025027 | 24.0516 | -0.12526 | 0.14184 | -1e-06 | -8.6e-05 | -8.6e-05 | -67.5034
2019-03-20 17:46 | 1553118345.0782 | -195 | 0.2095 | 0.00045 | 0 | 109.9055 | 111.3376 | 125.8489 | 21.0566 | 0.001975 | 0.001854 | -151.9106 | 0.23841 | 8.4053 | 0.003707 | 0.005133 | 0.004914 | 6.8e-05 | 0.001265 | 0.26145 | 1.3097 | 2.548 | 100.58 | 20.67 | 19.163 | 69.149 | 110.9123 | 66.2465 | 0.004648 | 1.0402 | 0.025027 | 24.0597 | -0.11579 | 0.14183 | -5e-06 | -0.000369 | -0.000369 | -15.7964
2019-03-20 17:46 | 1553118346.0769 | -194 | 0.2095 | 0.000451 | 0 | 109.79 | 111.2191 | 125.9205 | 21.0576 | 0.001881 | 0.001815 | -151.8931 | 0.23835 | 8.4033 | 0.003815 | 0.005125 | 0.004941 | 6.7e-05 | 0.001266 | 0.27527 | 1.3056 | 2.548 | 100.58 | 20.67 | 19.167 | 66.8965 | 113.5579 | 65.3084 | 0.004649 | 1.0431 | 0.025069 | 24.0332 | -0.11567 | 0.1418 | -6e-06 | -0.000473 | -0.000473 | -12.3211
2019-03-20 17:46 | 1553118347.078 | -193 | 0.20952 | 0.000449 | 0 | 109.9152 | 111.2393 | 126.251 | 21.057 | 0.001952 | 0.001839 | -151.9081 | 0.23885 | 8.4019 | 0.003851 | 0.005212 | 0.005005 | 6.7e-05 | 0.00127 | 0.27815 | 1.3429 | 2.5479 | 100.58 | 20.68 | 19.17 | 68.6021 | 114.4174 | 65.876 | 0.004653 | 1.0414 | 0.025081 | 24.0849 | -0.10711 | 0.14209 | -3e-06 | -0.000282 | -0.000282 | -20.7169
2019-03-20 17:46 | 1553118348.0774 | -192 | 0.20952 | 0.000446 | 0 | 109.8987 | 111.431 | 125.8424 | 21.0575 | 0.001977 | 0.001848 | -151.9145 | 0.23831 | 8.4031 | 0.003875 | 0.005701 | 0.00549 | 6.4e-05 | 0.001273 | 0.27505 | 1.3123 | 2.5481 | 100.58 | 20.67 | 19.172 | 69.1881 | 115.0085 | 66.1014 | 0.00465 | 1.0407 | 0.025032 | 24.0539 | -0.12385 | 0.14177 | 1e-06 | 7.4e-05 | 7.4e-05 | 79.263
2019-03-20 17:46 | 1553118349.0762 | -191 | 0.20951 | 0.000449 | 0 | 110.0077 | 111.7808 | 126.0834 | 21.0576 | 0.001992 | 0.001838 | -151.91 | 0.23813 | 8.4014 | 0.003824 | 0.005097 | 0.004967 | 6.9e-05 | 0.001274 | 0.25409 | 1.3642 | 2.5481 | 100.58 | 20.67 | 19.175 | 69.5533 | 113.7658 | 65.8658 | 0.004651 | 1.0414 | 0.025065 | 24.0686 | -0.14302 | 0.14166 | -3e-06 | -0.000218 | -0.000218 | -26.6857
2019-03-20 17:46 | 1553118350.0761 | -190 | 0.20955 | 0.000446 | 0 | 109.9563 | 111.6392 | 126.1907 | 21.0567 | 0.002053 | 0.00185 | -151.9189 | 0.23835 | 8.4036 | 0.003809 | 0.004755 | 0.00449 | 6.6e-05 | 0.001276 | 0.2756 | 1.3133 | 2.5481 | 100.58 | 20.67 | 19.176 | 71.03 | 113.4029 | 66.1333 | 0.004651 | 1.0406 | 0.025065 | 24.0883 | -0.13586 | 0.14179 | 1e-06 | 0.000125 | 0.000125 | 46.6042
2019-03-20 17:46 | 1553118351.0772 | -189 | 0.20951 | 0.000447 | 0 | 110.0323 | 111.5368 | 126.2386 | 21.0577 | 0.001996 | 0.001802 | -151.9387 | 0.23905 | 8.4042 | 0.00384 | 0.005061 | 0.004855 | 6.6e-05 | 0.001278 | 0.2959 | 1.3113 | 2.5481 | 100.58 | 20.67 | 19.177 | 69.6532 | 114.1659 | 64.9917 | 0.004652 | 1.0441 | 0.025112 | 24.0523 | -0.12147 | 0.14221 | -1e-06 | -8.1e-05 | -8.1e-05 | -72.5931
2019-03-20 17:46 | 1553118352.0761 | -188 | 0.20953 | 0.000448 | 0 | 109.9851 | 111.5566 | 126.3132 | 21.0577 | 0.001962 | 0.001869 | -151.935 | 0.23877 | 8.4017 | 0.003861 | 0.005269 | 0.005014 | 6.7e-05 | 0.001279 | 0.27223 | 1.3332 | 2.548 | 100.58 | 20.68 | 19.177 | 68.8374 | 114.6796 | 66.5908 | 0.004654 | 1.0392 | 0.025061 | 24.1162 | -0.1269 | 0.14204 | -2e-06 | -0.000151 | -0.000151 | -38.6554
2019-03-20 17:46 | 1553118353.0774 | -187 | 0.2095 | 0.000446 | 0 | 110.0125 | 111.4034 | 126.0217 | 21.0575 | 0.001909 | 0.001856 | -151.927 | 0.23818 | 8.4045 | 0.003836 | 0.004966 | 0.004706 | 6.6e-05 | 0.001279 | 0.26488 | 1.3233 | 2.5481 | 100.58 | 20.67 | 19.178 | 67.5673 | 114.0692 | 66.2776 | 0.004652 | 1.0401 | 0.025043 | 24.0773 | -0.11237 | 0.1417 | 0 | -3e-05 | -3e-05 | -196.5557
2019-03-20 17:46 | 1553118354.0764 | -186 | 0.20953 | 0.000451 | 0 | 110.0494 | 111.6362 | 126.1074 | 21.0576 | 0.001917 | 0.001868 | -151.938 | 0.23835 | 8.4031 | 0.003744 | 0.005498 | 0.005346 | 6.7e-05 | 0.001271 | 0.26825 | 1.2779 | 2.548 | 100.58 | 20.68 | 19.178 | 67.7571 | 111.8192 | 66.5774 | 0.004655 | 1.0392 | 0.025041 | 24.0961 | -0.12805 | 0.14179 | -6e-06 | -0.000415 | -0.000415 | -14.0424
2019-03-20 17:46 | 1553118355.0764 | -185 | 0.20952 | 0.000447 | 0 | 110.0016 | 111.7326 | 126.1933 | 21.0565 | 0.001942 | 0.001885 | -151.9126 | 0.23928 | 8.4013 | 0.003845 | 0.00467 | 0.004395 | 6.4e-05 | 0.001272 | 0.27721 | 1.352 | 2.5481 | 100.58 | 20.67 | 19.178 | 68.3441 | 114.2824 | 66.9744 | 0.004652 | 1.038 | 0.025035 | 24.1184 | -0.13966 | 0.14235 | -1e-06 | -0.000108 | -0.000108 | -54.2657
2019-03-20 17:46 | 1553118356.0774 | -184 | 0.20953 | 0.000442 | 0 | 110.0947 | 111.5431 | 126.0935 | 21.0577 | 0.001926 | 0.001801 | -151.9353 | 0.23818 | 8.4002 | 0.003757 | 0.004886 | 0.004656 | 6.3e-05 | 0.001268 | 0.26089 | 1.3386 | 2.548 | 100.58 | 20.68 | 19.181 | 67.9772 | 112.1267 | 64.9652 | 0.004655 | 1.0442 | 0.025099 | 24.0375 | -0.11691 | 0.14169 | 5e-06 | 0.000395 | 0.000395 | 14.7442
2019-03-20 17:46 | 1553118357.0781 | -183 | 0.20949 | 0.000449 | 0 | 110.0727 | 111.7266 | 126.0029 | 21.0576 | 0.002004 | 0.00182 | -151.9374 | 0.23853 | 8.4021 | 0.003735 | 0.004826 | 0.004594 | 6.4e-05 | 0.00127 | 0.26465 | 1.3545 | 2.5481 | 100.58 | 20.68 | 19.183 | 69.848 | 111.6015 | 65.4179 | 0.004656 | 1.0428 | 0.025073 | 24.045 | -0.1334 | 0.14191 | -3e-06 | -0.000259 | -0.000259 | -22.5543
2019-03-20 17:46 | 1553118358.0777 | -182 | 0.20951 | 0.000446 | 0 | 109.9207 | 111.7371 | 126.125 | 21.0568 | 0.001948 | 0.001837 | -151.938 | 0.23782 | 8.402 | 0.00381 | 0.00586 | 0.005699 | 6.5e-05 | 0.001275 | 0.27599 | 1.3282 | 2.5481 | 100.58 | 20.68 | 19.185 | 68.4879 | 113.4315 | 65.8399 | 0.004656 | 1.0415 | 0.02507 | 24.0716 | -0.1466 | 0.14148 | 0 | 0 | 0 | -21156.1724
2019-03-20 17:46 | 1553118359.0777 | -181 | 0.20948 | 0.000448 | 0 | 110.0226 | 111.7233 | 126.3982 | 21.0569 | 0.002071 | 0.001765 | -151.9372 | 0.23835 | 8.4055 | 0.003865 | 0.00545 | 0.00528 | 6.8e-05 | 0.001276 | 0.26737 | 1.3633 | 2.5481 | 100.58 | 20.68 | 19.188 | 71.4488 | 114.7588 | 64.0876 | 0.004657 | 1.0469 | 0.025162 | 24.0353 | -0.1372 | 0.14179 | -2e-06 | -0.000153 | -0.000153 | -38.1068
2019-03-20 17:46 | 1553118360.0772 | -180 | 0.20949 | 0.000447 | 0 | 110.0701 | 111.6784 | 126.3214 | 21.0579 | 0.001935 | 0.001835 | -151.9273 | 0.23845 | 8.4035 | 0.003877 | 0.005232 | 0.004896 | 6.5e-05 | 0.001273 | 0.26394 | 1.3585 | 2.548 | 100.58 | 20.68 | 19.189 | 68.1856 | 115.0468 | 65.7864 | 0.004657 | 1.0416 | 0.025091 | 24.0885 | -0.12975 | 0.14185 | -1e-06 | -0.000142 | -0.000142 | -41.1013
2019-03-20 17:46 | 1553118361.0769 | -179 | 0.20949 | 0.000444 | 0 | 110.0578 | 111.6945 | 125.9722 | 21.0582 | 0.001951 | 0.001864 | -151.9067 | 0.23802 | 8.403 | 0.003816 | 0.004821 | 0.004741 | 6.3e-05 | 0.001256 | 0.29662 | 1.3757 | 2.5481 | 100.58 | 20.68 | 19.191 | 68.5633 | 113.5773 | 66.4883 | 0.004658 | 1.0395 | 0.025031 | 24.08 | -0.13204 | 0.1416 | 3e-06 | 0.000192 | 0.000192 | 30.2339
2019-03-20 17:46 | 1553118362.0778 | -178 | 0.20951 | 0.000445 | 0 | 110.0792 | 111.6093 | 126.0961 | 21.0572 | 0.001981 | 0.001883 | -151.8942 | 0.23835 | 8.4033 | 0.003784 | 0.004908 | 0.004667 | 6.2e-05 | 0.001256 | 0.27759 | 1.3538 | 2.548 | 100.58 | 20.68 | 19.198 | 69.301 | 112.7832 | 66.9337 | 0.00466 | 1.0381 | 0.025026 | 24.1077 | -0.12347 | 0.14179 | 3e-06 | 0.000209 | 0.000209 | 27.9175
2019-03-20 17:46 | 1553118363.078 | -177 | 0.20951 | 0.000448 | 0 | 110.1138 | 111.5237 | 126.0635 | 21.0573 | 0.001992 | 0.00184 | -151.8887 | 0.23905 | 8.4034 | 0.003803 | 0.005789 | 0.005537 | 6.3e-05 | 0.001253 | 0.24043 | 1.356 | 2.5481 | 100.58 | 20.68 | 19.201 | 69.5538 | 113.2574 | 65.8912 | 0.00466 | 1.0413 | 0.025062 | 24.0676 | -0.1138 | 0.14221 | -3e-06 | -0.000221 | -0.000221 | -26.4013
2019-03-20 17:46 | 1553118364.077 | -176 | 0.20947 | 0.000445 | 0 | 110.0999 | 111.505 | 125.6572 | 21.0585 | 0.00207 | 0.001829 | -151.9074 | 0.23766 | 8.4024 | 0.003755 | 0.004697 | 0.004502 | 6.7e-05 | 0.001255 | 0.2625 | 1.3146 | 2.5481 | 100.58 | 20.68 | 19.208 | 71.441 | 112.0942 | 65.6348 | 0.004662 | 1.0421 | 0.025031 | 24.0197 | -0.11343 | 0.14138 | 1e-06 | 9.8e-05 | 9.8e-05 | 59.0019
2019-03-20 17:46 | 1553118365.0771 | -175 | 0.20949 | 0.000442 | 0 | 110.1403 | 111.7038 | 125.472 | 21.0585 | 0.002005 | 0.001819 | -151.9128 | 0.23849 | 8.4023 | 0.003755 | 0.005723 | 0.005522 | 6.3e-05 | 0.001248 | 0.27848 | 1.3135 | 2.548 | 100.58 | 20.68 | 19.212 | 69.8801 | 112.0934 | 65.3937 | 0.004663 | 1.0428 | 0.025021 | 23.9934 | -0.12608 | 0.14188 | 5e-06 | 0.000427 | 0.000427 | 13.6628
2019-03-20 17:46 | 1553118366.0766 | -174 | 0.2095 | 0.000449 | 0 | 110.0734 | 111.6082 | 125.2584 | 21.0581 | 0.001934 | 0.001822 | -151.9223 | 0.23866 | 8.4028 | 0.003794 | 0.004471 | 0.004268 | 6.7e-05 | 0.001235 | 0.26178 | 1.3418 | 2.548 | 100.58 | 20.68 | 19.216 | 68.1678 | 113.0252 | 65.4754 | 0.004664 | 1.0426 | 0.024997 | 23.9759 | -0.12386 | 0.14198 | -3e-06 | -0.000188 | -0.000188 | -31.0103
2019-03-20 17:46 | 1553118367.0766 | -173 | 0.20952 | 0.000447 | 0 | 110.1054 | 111.7816 | 125.4551 | 21.0585 | 0.001906 | 0.00183 | -151.9199 | 0.23802 | 8.4042 | 0.003728 | 0.005425 | 0.005206 | 6.5e-05 | 0.001235 | 0.2557 | 1.3469 | 2.5481 | 100.58 | 20.68 | 19.227 | 67.4856 | 111.4199 | 65.6614 | 0.004667 | 1.042 | 0.02501 | 24.0013 | -0.13514 | 0.14159 | -1e-06 | -8.8e-05 | -8.8e-05 | -65.8251
2019-03-20 17:46 | 1553118368.0773 | -172 | 0.2095 | 0.000445 | 0 | 110.1018 | 111.607 | 125.3208 | 21.0591 | 0.001955 | 0.001868 | -151.9136 | 0.23786 | 8.4013 | 0.003787 | 0.004957 | 0.004772 | 6.6e-05 | 0.001233 | 0.28313 | 1.3721 | 2.5481 | 100.58 | 20.68 | 19.231 | 68.6765 | 112.8569 | 66.5727 | 0.004668 | 1.0392 | 0.024963 | 24.0207 | -0.12145 | 0.1415 | 6.4e-05 | 0.004799 | 0.004799 | 1.2119
2019-03-20 17:46 | 1553118369.0772 | -171 | 0.2095 | 0.000449 | 0 | 110.0607 | 111.6231 | 124.7643 | 21.0584 | 0.002013 | 0.001824 | -151.9136 | 0.23827 | 8.4024 | 0.003771 | 0.005052 | 0.004895 | 6.3e-05 | 0.001235 | 0.26305 | 1.3478 | 2.548 | 100.58 | 20.68 | 19.241 | 70.0701 | 112.4708 | 65.5122 | 0.00467 | 1.0425 | 0.024946 | 23.9298 | -0.12608 | 0.14175 | 3.7e-05 | 0.002816 | 0.002816 | 2.0687
2019-03-20 17:46 | 1553118370.0778 | -170 | 0.2095 | 0.000446 | -1e-06 | 110.0918 | 111.665 | 124.6324 | 21.0582 | 0.001937 | 0.001821 | -151.9142 | 0.23785 | 8.4029 | 0.003766 | 0.00504 | 0.00478 | 6.4e-05 | 0.001237 | 0.27649 | 1.3433 | 2.5481 | 100.58 | 20.68 | 19.249 | 68.2412 | 112.3489 | 65.4504 | 0.004672 | 1.0427 | 0.024935 | 23.915 | -0.12691 | 0.1415 | 2.2e-05 | 0.001706 | 0.001706 | 3.4095
2019-03-20 17:46 | 1553118371.0763 | -169 | 0.20949 | 0.000449 | 0 | 110.1083 | 111.7057 | 125.0307 | 21.0593 | 0.001967 | 0.001815 | -151.9419 | 0.23856 | 8.402 | 0.003631 | 0.005039 | 0.004896 | 6.6e-05 | 0.001246 | 0.28838 | 1.3677 | 2.5481 | 100.58 | 20.68 | 19.254 | 68.9692 | 109.0836 | 65.311 | 0.004673 | 1.0431 | 0.02498 | 23.9483 | -0.12883 | 0.14192 | 7.4e-05 | 0.005561 | 0.005561 | 1.049
2019-03-20 17:46 | 1553118372.0774 | -168 | 0.20948 | 0.000449 | 0 | 110.0973 | 111.6343 | 125.0683 | 21.0582 | 0.001927 | 0.001833 | -151.9543 | 0.23797 | 8.4035 | 0.003707 | 0.005712 | 0.005462 | 6.4e-05 | 0.001246 | 0.26952 | 1.3586 | 2.548 | 100.58 | 20.68 | 19.268 | 68.0065 | 110.9304 | 65.747 | 0.004677 | 1.0417 | 0.024968 | 23.9673 | -0.12401 | 0.14157 | 0.000177 | 0.013347 | 0.013347 | 0.43592
2019-03-20 17:46 | 1553118373.0775 | -167 | 0.20951 | 0.00045 | 0 | 110.0659 | 111.585 | 125.2373 | 21.0596 | 0.001952 | 0.001814 | -151.961 | 0.23868 | 8.4025 | 0.003741 | 0.004801 | 0.004511 | 6.8e-05 | 0.001254 | 0.28313 | 1.357 | 2.5481 | 100.58 | 20.68 | 19.275 | 68.5921 | 111.7455 | 65.2824 | 0.004678 | 1.0432 | 0.025002 | 23.967 | -0.12261 | 0.14199 | -1e-06 | -0.000128 | -0.000128 | -45.5277
2019-03-20 17:46 | 1553118374.0765 | -166 | 0.20952 | 0.000453 | 0 | 110.1442 | 111.5469 | 125.5633 | 21.0592 | 0.00195 | 0.001852 | -151.9369 | 0.23792 | 8.4032 | 0.003762 | 0.005105 | 0.00498 | 6.6e-05 | 0.001255 | 0.24851 | 1.3406 | 2.548 | 100.58 | 20.68 | 19.283 | 68.5525 | 112.263 | 66.1935 | 0.00468 | 1.0404 | 0.025001 | 24.0305 | -0.11319 | 0.14154 | -0.000108 | -0.008156 | -0.008156 | -0.71321
2019-03-20 17:46 | 1553118375.0766 | -165 | 0.20953 | 0.000453 | 0 | 110.0578 | 111.6235 | 125.7464 | 21.0583 | 0.001987 | 0.001809 | -151.9481 | 0.23857 | 8.4032 | 0.003753 | 0.004686 | 0.0045 | 6.7e-05 | 0.001266 | 0.25713 | 1.3316 | 2.5481 | 100.58 | 20.68 | 19.299 | 69.4425 | 112.0348 | 65.1655 | 0.004684 | 1.0435 | 0.025057 | 24.0115 | -0.12635 | 0.14193 | -0.000162 | -0.01222 | -0.01222 | -0.47735
2019-03-20 17:46 | 1553118376.0777 | -164 | 0.20953 | 0.00045 | 0 | 110.1157 | 111.6627 | 125.8673 | 21.059 | 0.001967 | 0.001865 | -151.9293 | 0.23796 | 8.405 | 0.003709 | 0.005412 | 0.005155 | 6.7e-05 | 0.001264 | 0.26864 | 1.334 | 2.5478 | 100.58 | 20.68 | 19.305 | 68.9495 | 110.9763 | 66.5016 | 0.004686 | 1.0394 | 0.02502 | 24.0705 | -0.12479 | 0.14156 | -0.000142 | -0.010697 | -0.010697 | -0.54388
2019-03-20 17:46 | 1553118377.0768 | -163 | 0.20953 | 0.000446 | 0 | 110.0831 | 111.7229 | 125.8238 | 21.0585 | 0.001916 | 0.001879 | -151.9232 | 0.23854 | 8.4006 | 0.003874 | 0.005227 | 0.005071 | 6.7e-05 | 0.001269 | 0.29806 | 1.3819 | 2.5481 | 100.58 | 20.68 | 19.315 | 67.7323 | 114.9809 | 66.8359 | 0.004688 | 1.0384 | 0.025003 | 24.0782 | -0.13226 | 0.14191 | -0.000159 | -0.011963 | -0.011963 | -0.48753
2019-03-20 17:46 | 1553118378.0771 | -162 | 0.20954 | 0.000441 | 0 | 110.1578 | 111.6313 | 125.7461 | 21.0586 | 0.002036 | 0.001793 | -151.918 | 0.23761 | 8.4027 | 0.003826 | 0.005304 | 0.005081 | 6.4e-05 | 0.001271 | 0.29966 | 1.3302 | 2.5481 | 100.58 | 20.68 | 19.32 | 70.6147 | 113.8087 | 64.7756 | 0.004689 | 1.0447 | 0.025071 | 23.9977 | -0.11886 | 0.14135 | -0.000189 | -0.014305 | -0.014305 | -0.40611
2019-03-20 17:46 | 1553118379.0768 | -161 | 0.20948 | 0.000444 | 0 | 110.0879 | 111.5136 | 125.7093 | 21.06 | 0.001983 | 0.001839 | -151.9303 | 0.23738 | 8.4036 | 0.003766 | 0.004994 | 0.004784 | 6.7e-05 | 0.001274 | 0.27649 | 1.3474 | 2.5481 | 100.58 | 20.68 | 19.323 | 69.3347 | 112.3475 | 65.882 | 0.00469 | 1.0413 | 0.025027 | 24.0334 | -0.11509 | 0.14122 | 0.000153 | 0.011531 | 0.011531 | 0.50331
2019-03-20 17:46 | 1553118380.0768 | -160 | 0.2095 | 0.000447 | 0 | 110.0527 | 111.7293 | 125.7055 | 21.0584 | 0.001911 | 0.001817 | -151.9069 | 0.23784 | 8.4012 | 0.003741 | 0.005289 | 0.005111 | 6.8e-05 | 0.001276 | 0.30287 | 1.3741 | 2.5481 | 100.58 | 20.68 | 19.327 | 67.6089 | 111.7483 | 65.3516 | 0.004691 | 1.043 | 0.025046 | 24.0142 | -0.13524 | 0.14149 | 7.1e-05 | 0.005386 | 0.005386 | 1.0796
2019-03-20 17:46 | 1553118381.0763 | -159 | 0.20947 | 0.000453 | 0 | 110.0966 | 111.6242 | 125.6792 | 21.0585 | 0.001946 | 0.001891 | -151.8947 | 0.23818 | 8.4019 | 0.003781 | 0.00504 | 0.004793 | 6.7e-05 | 0.00127 | 0.26416 | 1.3453 | 2.5481 | 100.58 | 20.68 | 19.327 | 68.4496 | 112.7165 | 67.1376 | 0.004691 | 1.0375 | 0.024978 | 24.075 | -0.12325 | 0.14169 | 0.000236 | 0.017761 | 0.017761 | 0.32787
2019-03-20 17:46 | 1553118382.076 | -158 | 0.20949 | 0.000443 | 1e-06 | 110.1063 | 111.6915 | 125.3828 | 21.0586 | 0.00198 | 0.001881 | -151.8875 | 0.23839 | 8.4008 | 0.003755 | 0.005722 | 0.005537 | 6.4e-05 | 0.001277 | 0.26913 | 1.3501 | 2.5481 | 100.58 | 20.68 | 19.328 | 69.2656 | 112.0825 | 66.8867 | 0.004691 | 1.0383 | 0.024957 | 24.0377 | -0.12785 | 0.14182 | 0.00011 | 0.008283 | 0.008283 | 0.70372
2019-03-20 17:46 | 1553118383.077 | -157 | 0.20947 | 0.000446 | 0 | 110.0523 | 111.7868 | 125.2401 | 21.0597 | 0.001955 | 0.001847 | -151.889 | 0.23881 | 8.403 | 0.003825 | 0.005283 | 0.005106 | 6.8e-05 | 0.001277 | 0.27041 | 1.3865 | 2.548 | 100.58 | 20.68 | 19.327 | 68.6606 | 113.7899 | 66.0799 | 0.004691 | 1.0407 | 0.024973 | 23.9955 | -0.13987 | 0.14207 | 0.000249 | 0.018812 | 0.018812 | 0.3104
2019-03-20 17:46 | 1553118384.0776 | -156 | 0.20951 | 0.00044 | 0 | 110.0168 | 111.5319 | 125.0617 | 21.058 | 0.001962 | 0.001797 | -151.9032 | 0.2382 | 8.404 | 0.003761 | 0.004431 | 0.004274 | 6.7e-05 | 0.001279 | 0.28528 | 1.3737 | 2.5481 | 100.58 | 20.68 | 19.326 | 68.8394 | 112.228 | 64.8695 | 0.004691 | 1.0445 | 0.025 | 23.9356 | -0.12234 | 0.1417 | -6e-06 | -0.00044 | -0.00044 | -13.2444
2019-03-20 17:46 | 1553118385.0761 | -155 | 0.20952 | 0.000448 | 0 | 110.0636 | 111.6145 | 125.1483 | 21.058 | 0.001959 | 0.001855 | -151.9054 | 0.23852 | 8.4016 | 0.003803 | 0.005502 | 0.005237 | 6.9e-05 | 0.001283 | 0.25874 | 1.4056 | 2.548 | 100.58 | 20.68 | 19.321 | 68.7738 | 113.2676 | 66.2525 | 0.004689 | 1.0402 | 0.024957 | 23.9928 | -0.12515 | 0.1419 | -6.6e-05 | -0.004951 | -0.004951 | -1.1779
2019-03-20 17:46 | 1553118386.0762 | -154 | 0.2095 | 0.000444 | 0 | 109.9909 | 111.548 | 125.1251 | 21.0588 | 0.001976 | 0.001883 | -151.8877 | 0.23839 | 8.4029 | 0.003821 | 0.005491 | 0.005317 | 6.9e-05 | 0.001285 | 0.26681 | 1.4087 | 2.5482 | 100.58 | 20.68 | 19.318 | 69.1668 | 113.7029 | 66.928 | 0.004689 | 1.0381 | 0.02493 | 24.0145 | -0.12574 | 0.14182 | 6.3e-05 | 0.004732 | 0.004732 | 1.2318
2019-03-20 17:46 | 1553118387.0777 | -153 | 0.20951 | 0.000441 | 0 | 110.0067 | 111.5783 | 125.1483 | 21.0588 | 0.001956 | 0.001813 | -151.8925 | 0.23813 | 8.4048 | 0.003743 | 0.005013 | 0.004857 | 6.8e-05 | 0.001291 | 0.27577 | 1.3718 | 2.5482 | 100.58 | 20.68 | 19.315 | 68.6894 | 111.8007 | 65.2668 | 0.004688 | 1.0432 | 0.024994 | 23.958 | -0.12688 | 0.14167 | -1.7e-05 | -0.001301 | -0.001301 | -4.4755
2019-03-20 17:46 | 1553118388.0763 | -152 | 0.20952 | 0.000446 | 0 | 109.9889 | 111.44 | 125.1917 | 21.0592 | 0.001947 | 0.001896 | -151.8934 | 0.23865 | 8.4006 | 0.003808 | 0.004633 | 0.004386 | 6.8e-05 | 0.001287 | 0.27312 | 1.3694 | 2.5481 | 100.58 | 20.68 | 19.313 | 68.4739 | 113.3864 | 67.2437 | 0.004687 | 1.0372 | 0.024925 | 24.032 | -0.11723 | 0.14197 | -6.5e-05 | -0.004855 | -0.004855 | -1.2019
2019-03-20 17:46 | 1553118389.0772 | -151 | 0.20955 | 0.000443 | 0 | 109.9828 | 111.2262 | 125.417 | 21.0591 | 0.001906 | 0.001845 | -151.8859 | 0.23786 | 8.4004 | 0.003705 | 0.004415 | 0.00414 | 6.9e-05 | 0.001287 | 0.26609 | 1.3336 | 2.5481 | 100.58 | 20.68 | 19.313 | 67.4962 | 110.8736 | 66.0365 | 0.004687 | 1.0409 | 0.024992 | 24.0109 | -0.10055 | 0.14151 | -0.000242 | -0.018234 | -0.018234 | -0.31896
2019-03-20 17:46 | 1553118390.0773 | -150 | 0.20949 | 0.000447 | 0 | 109.9744 | 111.3338 | 125.3985 | 21.0586 | 0.001955 | 0.001829 | -151.9007 | 0.23797 | 8.4032 | 0.003724 | 0.004851 | 0.004587 | 7e-05 | 0.00128 | 0.28031 | 1.299 | 2.5479 | 100.58 | 20.69 | 19.312 | 68.6596 | 111.3387 | 65.631 | 0.00469 | 1.0421 | 0.025005 | 23.9948 | -0.10989 | 0.14157 | 9.2e-05 | 0.006932 | 0.006932 | 0.83935
2019-03-20 17:46 | 1553118391.0774 | -149 | 0.20949 | 0.000449 | 0 | 110.0012 | 111.5446 | 125.1718 | 21.0584 | 0.001977 | 0.001825 | -151.8994 | 0.23816 | 8.4014 | 0.003768 | 0.004883 | 0.00466 | 7e-05 | 0.00128 | 0.28512 | 1.3125 | 2.5479 | 100.58 | 20.68 | 19.313 | 69.2039 | 112.4107 | 65.5449 | 0.004687 | 1.0424 | 0.024986 | 23.97 | -0.12463 | 0.14168 | 0.000102 | 0.00764 | 0.00764 | 0.76219
2019-03-20 17:47 | 1553118392.0776 | -148 | 0.2095 | 0.000449 | 0 | 109.9964 | 111.5274 | 125.4785 | 21.0577 | 0.002003 | 0.001798 | -151.9062 | 0.23794 | 8.4025 | 0.003754 | 0.005227 | 0.005045 | 6.9e-05 | 0.001282 | 0.25321 | 1.3124 | 2.548 | 100.58 | 20.69 | 19.317 | 69.8199 | 112.068 | 64.8942 | 0.004691 | 1.0444 | 0.02504 | 23.9763 | -0.12364 | 0.14155 | 1.7e-05 | 0.001223 | 0.001223 | 4.7561
2019-03-20 17:47 | 1553118393.0766 | -147 | 0.2095 | 0.000451 | 0 | 109.9734 | 111.5431 | 125.5128 | 21.0591 | 0.001968 | 0.001873 | -151.9046 | 0.23832 | 8.4009 | 0.003752 | 0.004391 | 0.004154 | 6.6e-05 | 0.001283 | 0.28008 | 1.3301 | 2.5481 | 100.58 | 20.69 | 19.319 | 68.9909 | 112.0191 | 66.6866 | 0.004692 | 1.0389 | 0.024978 | 24.0431 | -0.12677 | 0.14178 | 4.3e-05 | 0.003203 | 0.003203 | 1.8191
2019-03-20 17:47 | 1553118394.0776 | -146 | 0.20953 | 0.000445 | 0 | 109.9737 | 111.4422 | 125.7264 | 21.0584 | 0.00191 | 0.001804 | -151.8893 | 0.23779 | 8.4024 | 0.003738 | 0.005425 | 0.005239 | 6.5e-05 | 0.001288 | 0.28263 | 1.3528 | 2.5481 | 100.58 | 20.69 | 19.323 | 67.5783 | 111.667 | 65.0393 | 0.004693 | 1.0439 | 0.02506 | 24.0052 | -0.11865 | 0.14146 | -0.000111 | -0.008431 | -0.008431 | -0.68958
2019-03-20 17:47 | 1553118395.0766 | -145 | 0.20951 | 0.000447 | 0 | 109.9521 | 111.539 | 125.5577 | 21.0585 | 0.002031 | 0.00186 | -151.8954 | 0.23806 | 8.404 | 0.0037 | 0.005233 | 0.005007 | 6.9e-05 | 0.001291 | 0.27544 | 1.3127 | 2.5481 | 100.58 | 20.69 | 19.33 | 70.4985 | 110.7406 | 66.3747 | 0.004695 | 1.0398 | 0.024994 | 24.0364 | -0.12818 | 0.14162 | 1.8e-05 | 0.001361 | 0.001361 | 4.2765
2019-03-20 17:47 | 1553118396.0772 | -144 | 0.20954 | 0.000447 | 0 | 109.956 | 111.4288 | 125.964 | 21.0591 | 0.001985 | 0.001883 | -151.9161 | 0.2384 | 8.4028 | 0.00373 | 0.004481 | 0.004336 | 6.4e-05 | 0.001281 | 0.26681 | 1.3159 | 2.5481 | 100.58 | 20.69 | 19.334 | 69.3839 | 111.4916 | 66.9488 | 0.004695 | 1.0381 | 0.025013 | 24.0956 | -0.11901 | 0.14182 | -0.000176 | -0.013311 | -0.013311 | -0.43791
2019-03-20 17:47 | 1553118397.0757 | -143 | 0.20953 | 0.000447 | 0 | 109.8832 | 111.3084 | 125.8447 | 21.0578 | 0.002006 | 0.001885 | -151.9157 | 0.23808 | 8.4026 | 0.003716 | 0.005627 | 0.005351 | 6.7e-05 | 0.001279 | 0.27024 | 1.3369 | 2.5481 | 100.58 | 20.69 | 19.338 | 69.9044 | 111.1336 | 66.9951 | 0.004696 | 1.0379 | 0.024999 | 24.0858 | -0.11527 | 0.14163 | -0.000147 | -0.01104 | -0.01104 | -0.52726
2019-03-20 17:47 | 1553118398.0769 | -142 | 0.20954 | 0.000449 | 0 | 109.8819 | 111.3178 | 125.8654 | 21.0581 | 0.002 | 0.001812 | -151.9181 | 0.23897 | 8.4029 | 0.003735 | 0.005353 | 0.005239 | 6.8e-05 | 0.001281 | 0.28888 | 1.3029 | 2.5483 | 100.58 | 20.69 | 19.345 | 69.7603 | 111.6122 | 65.2418 | 0.004698 | 1.0433 | 0.025066 | 24.0256 | -0.11613 | 0.14216 | -0.000214 | -0.016179 | -0.016179 | -0.36115
2019-03-20 17:47 | 1553118399.0759 | -141 | 0.20952 | 0.000446 | 0 | 109.8683 | 111.3099 | 125.8992 | 21.0582 | 0.001965 | 0.00185 | -151.9216 | 0.23897 | 8.4009 | 0.003683 | 0.005291 | 0.005086 | 6.6e-05 | 0.001281 | 0.24475 | 1.2985 | 2.5481 | 100.58 | 20.69 | 19.348 | 68.9119 | 110.3315 | 66.1442 | 0.004699 | 1.0405 | 0.025036 | 24.0608 | -0.1166 | 0.14216 | -4.3e-05 | -0.003203 | -0.003203 | -1.8239
2019-03-20 17:47 | 1553118400.0772 | -140 | 0.20955 | 0.000443 | 0 | 109.7997 | 111.1836 | 126.2365 | 21.058 | 0.001974 | 0.001855 | -151.9293 | 0.23874 | 8.4028 | 0.003678 | 0.004775 | 0.004469 | 6.9e-05 | 0.001282 | 0.25371 | 1.2975 | 2.548 | 100.58 | 20.69 | 19.351 | 69.1139 | 110.203 | 66.2618 | 0.0047 | 1.0402 | 0.025065 | 24.0972 | -0.11203 | 0.14202 | -0.000222 | -0.016767 | -0.016767 | -0.34814
2019-03-20 17:47 | 1553118401.0765 | -139 | 0.20953 | 0.000442 | 0 | 109.771 | 111.4463 | 126.2259 | 21.0581 | 0.002035 | 0.001837 | -151.932 | 0.23788 | 8.4042 | 0.00369 | 0.00553 | 0.005286 | 7e-05 | 0.001282 | 0.24585 | 1.2743 | 2.5479 | 100.58 | 20.69 | 19.351 | 70.587 | 110.4972 | 65.8405 | 0.0047 | 1.0415 | 0.02508 | 24.0813 | -0.13548 | 0.14152 | -0.000124 | -0.009382 | -0.009382 | -0.61992
2019-03-20 17:47 | 1553118402.077 | -138 | 0.20951 | 0.000446 | 0 | 109.8369 | 111.2699 | 125.9604 | 21.0588 | 0.001939 | 0.001834 | -151.9367 | 0.23843 | 8.4052 | 0.003637 | 0.004745 | 0.004528 | 6.8e-05 | 0.001282 | 0.27649 | 1.307 | 2.5479 | 100.58 | 20.69 | 19.35 | 68.273 | 109.2148 | 65.7693 | 0.004699 | 1.0417 | 0.025056 | 24.0534 | -0.11594 | 0.14184 | 2e-06 | 9.3e-05 | 9.3e-05 | 62.75
2019-03-20 17:47 | 1553118403.077 | -137 | 0.20953 | 0.000444 | 0 | 109.7545 | 111.3129 | 126.1975 | 21.0591 | 0.001929 | 0.001858 | -151.9472 | 0.23821 | 8.4042 | 0.003697 | 0.004857 | 0.004675 | 6.4e-05 | 0.001284 | 0.2568 | 1.32 | 2.5481 | 100.58 | 20.69 | 19.341 | 68.0502 | 110.6821 | 66.3387 | 0.004697 | 1.0399 | 0.025058 | 24.0962 | -0.12611 | 0.14171 | -0.000151 | -0.011431 | -0.011431 | -0.50953
2019-03-20 17:47 | 1553118404.0776 | -136 | 0.20954 | 0.000449 | 0 | 109.7321 | 111.2363 | 125.914 | 21.0583 | 0.001972 | 0.001819 | -151.9487 | 0.23798 | 8.4038 | 0.003588 | 0.005338 | 0.005132 | 6.6e-05 | 0.001286 | 0.2912 | 1.3351 | 2.548 | 100.58 | 20.69 | 19.335 | 69.0752 | 108.0321 | 65.3935 | 0.004696 | 1.0428 | 0.025065 | 24.0356 | -0.12177 | 0.14157 | -0.000186 | -0.014064 | -0.014064 | -0.41373
2019-03-20 17:47 | 1553118405.0772 | -135 | 0.20951 | 0.000444 | 0 | 109.6367 | 111.1701 | 126.102 | 21.059 | 0.002016 | 0.001837 | -151.9461 | 0.23891 | 8.4019 | 0.003711 | 0.005081 | 0.004834 | 6.4e-05 | 0.001286 | 0.25337 | 1.3133 | 2.5481 | 100.58 | 20.69 | 19.329 | 70.1262 | 111.0142 | 65.8421 | 0.004694 | 1.0415 | 0.025067 | 24.0695 | -0.12423 | 0.14213 | 1.9e-05 | 0.001463 | 0.001463 | 3.9935
2019-03-20 17:47 | 1553118406.0772 | -134 | 0.20954 | 0.00044 | 1e-06 | 109.7286 | 111.1376 | 125.8464 | 21.0577 | 0.001981 | 0.00183 | -151.9328 | 0.23726 | 8.4033 | 0.003673 | 0.005352 | 0.005217 | 6.8e-05 | 0.001293 | 0.28456 | 1.3882 | 2.5481 | 100.58 | 20.69 | 19.322 | 69.3054 | 110.104 | 65.6684 | 0.004693 | 1.042 | 0.025048 | 24.0389 | -0.11413 | 0.14115 | -0.000159 | -0.01208 | -0.01208 | -0.48024
2019-03-20 17:47 | 1553118407.0773 | -133 | 0.20953 | 0.000444 | 0 | 109.7043 | 111.2116 | 125.7215 | 21.0587 | 0.001978 | 0.001835 | -151.923 | 0.23816 | 8.4013 | 0.003634 | 0.005024 | 0.004745 | 6.4e-05 | 0.00129 | 0.26394 | 1.3297 | 2.5481 | 100.58 | 20.69 | 19.304 | 69.2282 | 109.1571 | 65.7803 | 0.004688 | 1.0416 | 0.025032 | 24.031 | -0.12206 | 0.14168 | -9.7e-05 | -0.007279 | -0.007279 | -0.8
2019-03-20 17:47 | 1553118408.0761 | -132 | 0.20951 | 0.000445 | 0 | 109.5138 | 111.1529 | 125.7867 | 21.0595 | 0.002055 | 0.00187 | -151.914 | 0.23858 | 8.4037 | 0.003602 | 0.005076 | 0.004956 | 6.5e-05 | 0.00129 | 0.27383 | 1.3386 | 2.5482 | 100.58 | 20.69 | 19.294 | 71.0841 | 108.3738 | 66.6339 | 0.004686 | 1.039 | 0.025007 | 24.0675 | -0.13288 | 0.14193 | 6e-06 | 0.00047 | 0.00047 | 12.4199
2019-03-20 17:47 | 1553118409.0765 | -131 | 0.20952 | 0.000444 | 0 | 109.5659 | 111.0909 | 125.8203 | 21.059 | 0.001913 | 0.001842 | -151.9039 | 0.23828 | 8.4016 | 0.003559 | 0.004881 | 0.004644 | 6.5e-05 | 0.001284 | 0.27616 | 1.353 | 2.5481 | 100.58 | 20.69 | 19.271 | 67.6553 | 107.3341 | 65.9607 | 0.00468 | 1.0411 | 0.025035 | 24.0468 | -0.12363 | 0.14175 | -4.6e-05 | -0.003486 | -0.003486 | -1.6711
2019-03-20 17:47 | 1553118410.0754 | -130 | 0.20948 | 0.000444 | 0 | 109.5649 | 111.1215 | 125.4926 | 21.0599 | 0.001964 | 0.001809 | -151.8821 | 0.23874 | 8.4041 | 0.003554 | 0.005646 | 0.005452 | 6.5e-05 | 0.00128 | 0.26609 | 1.3251 | 2.5481 | 100.58 | 20.69 | 19.258 | 68.8784 | 107.1949 | 65.1569 | 0.004677 | 1.0436 | 0.025032 | 23.987 | -0.12618 | 0.14202 | 0.000205 | 0.015485 | 0.015485 | 0.37697
2019-03-20 17:47 | 1553118411.0754 | -129 | 0.20951 | 0.000448 | 1e-06 | 109.586 | 111.1376 | 125.4487 | 21.0613 | 0.001968 | 0.001857 | -151.8759 | 0.23781 | 8.4022 | 0.003515 | 0.005059 | 0.004847 | 6.6e-05 | 0.001278 | 0.26432 | 1.3058 | 2.548 | 100.58 | 20.69 | 19.245 | 68.9923 | 106.2493 | 66.3102 | 0.004674 | 1.04 | 0.024985 | 24.0236 | -0.12576 | 0.14147 | 1.5e-05 | 0.001105 | 0.001105 | 5.2614
2019-03-20 17:47 | 1553118412.0758 | -128 | 0.2095 | 0.000451 | -1e-06 | 109.5132 | 110.9575 | 125.7255 | 21.0592 | 0.001977 | 0.001904 | -151.8533 | 0.2385 | 8.4038 | 0.003597 | 0.004959 | 0.00472 | 6.7e-05 | 0.001279 | 0.25586 | 1.3338 | 2.548 | 100.58 | 20.69 | 19.216 | 69.189 | 108.2355 | 67.4396 | 0.004667 | 1.0366 | 0.024971 | 24.0901 | -0.11719 | 0.14189 | 5.4e-05 | 0.004068 | 0.004068 | 1.4333
2019-03-20 17:47 | 1553118413.0755 | -127 | 0.20951 | 0.000449 | 0 | 109.4294 | 110.9993 | 125.5412 | 21.0603 | 0.001915 | 0.001843 | -151.844 | 0.23813 | 8.4031 | 0.003587 | 0.005221 | 0.005131 | 6.5e-05 | 0.001283 | 0.27168 | 1.3263 | 2.5482 | 100.58 | 20.69 | 19.202 | 67.7128 | 108.0035 | 65.9845 | 0.004663 | 1.041 | 0.025006 | 24.021 | -0.12741 | 0.14167 | -2.1e-05 | -0.001625 | -0.001625 | -3.5838
2019-03-20 17:47 | 1553118414.0768 | -126 | 0.2095 | 0.000444 | 0 | 109.3861 | 111.0221 | 125.5917 | 21.0593 | 0.001986 | 0.001879 | -151.8586 | 0.23818 | 8.4049 | 0.003616 | 0.004932 | 0.004664 | 6.6e-05 | 0.001289 | 0.25536 | 1.3235 | 2.5482 | 100.58 | 20.69 | 19.176 | 69.4167 | 108.7022 | 66.8482 | 0.004657 | 1.0384 | 0.02498 | 24.0564 | -0.13279 | 0.1417 | 7e-05 | 0.005257 | 0.005257 | 1.1078
2019-03-20 17:47 | 1553118415.0755 | -125 | 0.20949 | 0.000445 | 0 | 109.3964 | 110.8846 | 125.6034 | 21.059 | 0.001937 | 0.00189 | -151.847 | 0.23751 | 8.4037 | 0.003619 | 0.004879 | 0.004686 | 6.5e-05 | 0.001281 | 0.24635 | 1.3586 | 2.5482 | 100.58 | 20.69 | 19.164 | 68.2342 | 108.7791 | 67.1155 | 0.004654 | 1.0376 | 0.024971 | 24.0669 | -0.12085 | 0.14129 | 0.000104 | 0.007852 | 0.007852 | 0.73959
2019-03-20 17:47 | 1553118416.0766 | -124 | 0.2095 | 0.000444 | 0 | 109.3845 | 110.8401 | 125.5656 | 21.0593 | 0.002025 | 0.00183 | -151.8584 | 0.23892 | 8.4015 | 0.003673 | 0.00568 | 0.005474 | 6.2e-05 | 0.001276 | 0.27417 | 1.3213 | 2.5481 | 100.58 | 20.69 | 19.153 | 70.344 | 110.0819 | 65.6536 | 0.004652 | 1.042 | 0.025021 | 24.0116 | -0.11824 | 0.14214 | 4.6e-05 | 0.003471 | 0.003471 | 1.6831
2019-03-20 17:47 | 1553118417.0757 | -123 | 0.2095 | 0.000442 | 0 | 109.3725 | 110.6641 | 125.7215 | 21.0599 | 0.001977 | 0.001911 | -151.8719 | 0.23779 | 8.4017 | 0.003684 | 0.005215 | 0.00503 | 6.2e-05 | 0.001272 | 0.28655 | 1.3317 | 2.5481 | 100.58 | 20.69 | 19.13 | 69.1974 | 110.3655 | 67.6221 | 0.004646 | 1.036 | 0.024964 | 24.0962 | -0.10501 | 0.14146 | 3.4e-05 | 0.002571 | 0.002571 | 2.2617
2019-03-20 17:47 | 1553118418.0768 | -122 | 0.20949 | 0.000448 | 0 | 109.2606 | 110.7205 | 125.7328 | 21.0597 | 0.001956 | 0.001789 | -151.8785 | 0.23767 | 8.4017 | 0.003605 | 0.004527 | 0.004314 | 6.2e-05 | 0.001273 | 0.27975 | 1.3515 | 2.5481 | 100.58 | 20.69 | 19.118 | 68.692 | 108.4442 | 64.6751 | 0.004643 | 1.0451 | 0.025074 | 23.9928 | -0.11872 | 0.14139 | 9.9e-05 | 0.007492 | 0.007492 | 0.77564
2019-03-20 17:47 | 1553118419.0769 | -121 | 0.20949 | 0.00045 | 0 | 109.2642 | 110.6379 | 125.5736 | 21.0595 | 0.001971 | 0.001788 | -151.905 | 0.23862 | 8.4019 | 0.003686 | 0.005752 | 0.005511 | 6.4e-05 | 0.001274 | 0.29518 | 1.3073 | 2.548 | 100.58 | 20.69 | 19.107 | 69.0638 | 110.4188 | 64.6609 | 0.00464 | 1.0451 | 0.025058 | 23.9771 | -0.11176 | 0.14195 | 9.5e-05 | 0.007232 | 0.007232 | 0.80678
2019-03-20 17:47 | 1553118420.0769 | -120 | 0.20951 | 0.000443 | 0 | 109.2587 | 110.608 | 125.7173 | 21.0598 | 0.001927 | 0.001846 | -151.9369 | 0.23814 | 8.4038 | 0.003615 | 0.004783 | 0.00463 | 6.7e-05 | 0.001271 | 0.27096 | 1.3238 | 2.5481 | 100.58 | 20.69 | 19.084 | 67.9901 | 108.6925 | 66.0536 | 0.004635 | 1.0408 | 0.025021 | 24.0403 | -0.10979 | 0.14167 | -1.8e-05 | -0.001345 | -0.001345 | -4.328
2019-03-20 17:47 | 1553118421.0759 | -119 | 0.2095 | 0.000438 | 0 | 109.2124 | 110.5557 | 125.956 | 21.06 | 0.001917 | 0.001874 | -151.9432 | 0.23729 | 8.402 | 0.003624 | 0.004677 | 0.004413 | 6.2e-05 | 0.001268 | 0.28799 | 1.3481 | 2.5482 | 100.58 | 20.69 | 19.074 | 67.7647 | 108.8967 | 66.7288 | 0.004632 | 1.0387 | 0.02502 | 24.087 | -0.10934 | 0.14116 | 6.6e-05 | 0.004956 | 0.004956 | 1.1705
2019-03-20 17:47 | 1553118422.0759 | -118 | 0.20948 | 0.000448 | 0 | 109.1743 | 110.4522 | 125.545 | 21.0591 | 0.001911 | 0.001842 | -151.958 | 0.23867 | 8.4013 | 0.003673 | 0.005057 | 0.004755 | 6.6e-05 | 0.001273 | 0.2672 | 1.3981 | 2.548 | 100.58 | 20.69 | 19.056 | 67.6057 | 110.0856 | 65.9492 | 0.004628 | 1.0411 | 0.025008 | 24.0201 | -0.10409 | 0.14199 | 0.0002 | 0.015119 | 0.015119 | 0.38598
2019-03-20 17:47 | 1553118423.076 | -117 | 0.20952 | 0.000446 | 0 | 109.073 | 110.6454 | 125.4405 | 21.0602 | 0.001981 | 0.001819 | -151.9548 | 0.23879 | 8.4037 | 0.003628 | 0.004468 | 0.00422 | 6.1e-05 | 0.001276 | 0.27113 | 1.3856 | 2.5483 | 100.58 | 20.69 | 19.049 | 69.3039 | 108.9907 | 65.3936 | 0.004626 | 1.0428 | 0.025018 | 23.9904 | -0.12802 | 0.14205 | -5.9e-05 | -0.004464 | -0.004464 | -1.3079
2019-03-20 17:47 | 1553118424.0761 | -116 | 0.20949 | 0.00045 | 0 | 109.0478 | 110.6061 | 125.8175 | 21.0595 | 0.001934 | 0.001837 | -151.946 | 0.23776 | 8.4013 | 0.00367 | 0.004774 | 0.004581 | 6.2e-05 | 0.001275 | 0.28495 | 1.3451 | 2.5482 | 100.58 | 20.69 | 19.041 | 68.1636 | 110.0207 | 65.8356 | 0.004624 | 1.0415 | 0.025039 | 24.0421 | -0.12692 | 0.14144 | 0.000121 | 0.00914 | 0.00914 | 0.63607
2019-03-20 17:47 | 1553118425.0752 | -115 | 0.20949 | 0.000443 | 0 | 109.0274 | 110.4331 | 125.6593 | 21.0598 | 0.001989 | 0.001835 | -151.9522 | 0.23846 | 8.4014 | 0.003616 | 0.005267 | 0.00508 | 6.3e-05 | 0.001278 | 0.29031 | 1.3257 | 2.5482 | 100.58 | 20.69 | 19.027 | 69.4867 | 108.7168 | 65.7813 | 0.004621 | 1.0416 | 0.025026 | 24.0251 | -0.11458 | 0.14186 | 9.5e-05 | 0.007199 | 0.007199 | 0.80991
2019-03-20 17:47 | 1553118426.0761 | -114 | 0.20951 | 0.000446 | 0 | 108.9734 | 110.3086 | 125.6748 | 21.0593 | 0.001958 | 0.001866 | -151.9564 | 0.23831 | 8.4039 | 0.003677 | 0.004351 | 0.004153 | 6.5e-05 | 0.001273 | 0.27024 | 1.3655 | 2.5481 | 100.58 | 20.69 | 19.021 | 68.7375 | 110.1812 | 66.537 | 0.004619 | 1.0393 | 0.024999 | 24.0533 | -0.10893 | 0.14177 | 4e-06 | 0.000305 | 0.000305 | 19.117
2019-03-20 17:47 | 1553118427.0752 | -113 | 0.20949 | 0.000447 | 0 | 108.8751 | 110.2436 | 125.7173 | 21.0591 | 0.001963 | 0.001838 | -151.9429 | 0.23835 | 8.4048 | 0.003642 | 0.005357 | 0.005133 | 6.5e-05 | 0.001264 | 0.26913 | 1.3026 | 2.5479 | 100.58 | 20.69 | 19.011 | 68.8574 | 109.3514 | 65.8617 | 0.004617 | 1.0414 | 0.025028 | 24.0334 | -0.11173 | 0.14179 | 0.000146 | 0.011016 | 0.011016 | 0.52903
2019-03-20 17:47 | 1553118428.0763 | -112 | 0.20951 | 0.000448 | 0 | 108.8014 | 110.1105 | 125.736 | 21.0589 | 0.001963 | 0.001861 | -151.9675 | 0.23867 | 8.4008 | 0.003632 | 0.005116 | 0.004843 | 6.5e-05 | 0.001266 | 0.28385 | 1.3722 | 2.5481 | 100.58 | 20.69 | 19.007 | 68.8643 | 109.1061 | 66.407 | 0.004616 | 1.0397 | 0.02501 | 24.0546 | -0.10699 | 0.14199 | 1.2e-05 | 0.000902 | 0.000902 | 6.4702
2019-03-20 17:47 | 1553118429.0763 | -111 | 0.20949 | 0.00045 | 0 | 108.7972 | 110.1378 | 125.7553 | 21.0593 | 0.001933 | 0.00184 | -151.9557 | 0.23852 | 8.4023 | 0.003607 | 0.004732 | 0.004547 | 6.7e-05 | 0.001257 | 0.25802 | 1.3283 | 2.5481 | 100.58 | 20.69 | 19.006 | 68.1494 | 108.4887 | 65.915 | 0.004616 | 1.0412 | 0.02503 | 24.039 | -0.10955 | 0.14189 | 9.7e-05 | 0.007317 | 0.007317 | 0.79705
2019-03-20 17:47 | 1553118430.0758 | -110 | 0.20951 | 0.000448 | 0 | 108.7011 | 110.1247 | 125.7893 | 21.0596 | 0.001916 | 0.001859 | -151.9492 | 0.23843 | 8.4031 | 0.003585 | 0.005113 | 0.004845 | 6.6e-05 | 0.001249 | 0.27992 | 1.3397 | 2.5481 | 100.58 | 20.69 | 19.006 | 67.7386 | 107.9681 | 66.3523 | 0.004616 | 1.0399 | 0.025017 | 24.0577 | -0.11639 | 0.14184 | -1e-05 | -0.000794 | -0.000794 | -7.3421
2019-03-20 17:47 | 1553118431.0765 | -109 | 0.20952 | 0.000447 | 0 | 108.6633 | 109.9039 | 125.348 | 21.0592 | 0.002014 | 0.001871 | -151.9467 | 0.23897 | 8.404 | 0.003685 | 0.005997 | 0.005787 | 6.9e-05 | 0.001253 | 0.28152 | 1.3221 | 2.548 | 100.58 | 20.69 | 19.008 | 70.0819 | 110.3775 | 66.6437 | 0.004616 | 1.039 | 0.024963 | 24.0258 | -0.10154 | 0.14216 | -3.4e-05 | -0.002552 | -0.002552 | -2.2895
2019-03-20 17:47 | 1553118432.0765 | -108 | 0.20953 | 0.00045 | 0 | 108.6261 | 109.8403 | 125.7738 | 21.0594 | 0.002034 | 0.001784 | -151.9457 | 0.23843 | 8.4033 | 0.003669 | 0.004896 | 0.004617 | 6.4e-05 | 0.001258 | 0.2844 | 1.3233 | 2.5481 | 100.58 | 20.69 | 19.01 | 70.5757 | 109.9974 | 64.5672 | 0.004617 | 1.0454 | 0.025082 | 23.9929 | -0.09943 | 0.14184 | -0.000146 | -0.01108 | -0.01108 | -0.52617
2019-03-20 17:47 | 1553118433.0767 | -107 | 0.20954 | 0.000445 | 0 | 108.5142 | 109.7061 | 125.8133 | 21.0603 | 0.001981 | 0.001823 | -151.9481 | 0.23825 | 8.4012 | 0.003583 | 0.004561 | 0.00447 | 6.7e-05 | 0.001263 | 0.28279 | 1.3059 | 2.548 | 100.58 | 20.69 | 19.019 | 69.2915 | 107.9003 | 65.5078 | 0.004619 | 1.0425 | 0.025051 | 24.0301 | -0.097716 | 0.14173 | -0.000155 | -0.011743 | -0.011743 | -0.49607
2019-03-20 17:47 | 1553118434.0756 | -106 | 0.20952 | 0.000448 | 0 | 108.5168 | 109.7379 | 125.5114 | 21.0608 | 0.002013 | 0.001865 | -151.9438 | 0.23887 | 8.4017 | 0.003601 | 0.005319 | 0.005004 | 6.4e-05 | 0.001264 | 0.28888 | 1.2894 | 2.548 | 100.58 | 20.68 | 19.024 | 70.0686 | 108.3379 | 66.5158 | 0.004617 | 1.0394 | 0.024984 | 24.0369 | -0.10009 | 0.14211 | -3.9e-05 | -0.002944 | -0.002944 | -1.9839
2019-03-20 17:47 | 1553118435.0771 | -105 | 0.20952 | 0.000448 | 0 | 108.5051 | 109.7936 | 125.7853 | 21.06 | 0.001953 | 0.001831 | -151.9385 | 0.23797 | 8.4039 | 0.003622 | 0.004906 | 0.004698 | 6.7e-05 | 0.001259 | 0.274 | 1.3117 | 2.5481 | 100.58 | 20.69 | 19.04 | 68.6144 | 108.844 | 65.7002 | 0.004624 | 1.0419 | 0.025041 | 24.0342 | -0.10559 | 0.14157 | -7.6e-05 | -0.00576 | -0.00576 | -1.0101
2019-03-20 17:47 | 1553118436.0761 | -104 | 0.20953 | 0.000445 | 0 | 108.3612 | 109.5006 | 125.8677 | 21.0608 | 0.001989 | 0.001822 | -151.9365 | 0.2368 | 8.4028 | 0.003644 | 0.004756 | 0.004569 | 6.8e-05 | 0.001262 | 0.29192 | 1.3162 | 2.548 | 100.58 | 20.69 | 19.05 | 69.48 | 109.3966 | 65.4701 | 0.004627 | 1.0426 | 0.025058 | 24.0339 | -0.093559 | 0.14087 | -0.000112 | -0.008439 | -0.008439 | -0.68611
2019-03-20 17:47 | 1553118437.0758 | -103 | 0.20952 | 0.000445 | 0 | 108.312 | 109.5096 | 125.6844 | 21.0604 | 0.002 | 0.001856 | -151.9369 | 0.23783 | 8.405 | 0.003743 | 0.00536 | 0.00521 | 6.6e-05 | 0.001259 | 0.28528 | 1.3093 | 2.548 | 100.58 | 20.69 | 19.063 | 69.7586 | 111.796 | 66.2993 | 0.00463 | 1.0401 | 0.025009 | 24.0458 | -0.09835 | 0.14149 | -3.7e-05 | -0.002827 | -0.002827 | -2.0572
2019-03-20 17:47 | 1553118438.0768 | -102 | 0.20951 | 0.000444 | 0 | 108.2684 | 109.5275 | 125.6586 | 21.0601 | 0.002005 | 0.001817 | -151.9136 | 0.23845 | 8.4032 | 0.003589 | 0.004894 | 0.004699 | 6.9e-05 | 0.001259 | 0.28854 | 1.3282 | 2.5481 | 100.58 | 20.68 | 19.09 | 69.8809 | 108.0451 | 65.3428 | 0.004633 | 1.043 | 0.025042 | 24.0094 | -0.10342 | 0.14185 | 5e-06 | 0.000375 | 0.000375 | 15.5276
2019-03-20 17:47 | 1553118439.0763 | -101 | 0.20953 | 0.00045 | 0 | 108.1616 | 109.3537 | 125.7856 | 21.0602 | 0.001997 | 0.001842 | -151.9068 | 0.23781 | 8.4028 | 0.003589 | 0.005287 | 0.005033 | 6.7e-05 | 0.001261 | 0.26161 | 1.3265 | 2.5481 | 100.58 | 20.68 | 19.105 | 69.6844 | 108.0435 | 65.9534 | 0.004637 | 1.0411 | 0.025032 | 24.0432 | -0.098039 | 0.14147 | -0.000144 | -0.010825 | -0.010825 | -0.53713
2019-03-20 17:47 | 1553118440.0769 | -100 | 0.20952 | 0.000446 | -1e-06 | 108.04 | 109.3376 | 125.7245 | 21.0605 | 0.001946 | 0.001844 | -151.8956 | 0.23902 | 8.4037 | 0.003652 | 0.004938 | 0.004852 | 6.8e-05 | 0.001263 | 0.27743 | 1.3715 | 2.548 | 100.58 | 20.68 | 19.135 | 68.4634 | 109.5863 | 66.0091 | 0.004644 | 1.0409 | 0.025024 | 24.0394 | -0.10679 | 0.14219 | -3.9e-05 | -0.00292 | -0.00292 | -2.0016
2019-03-20 17:47 | 1553118441.077 | -99 | 0.20953 | 0.000445 | 0 | 107.9618 | 109.178 | 125.8879 | 21.0598 | 0.001952 | 0.001887 | -151.9062 | 0.23863 | 8.402 | 0.003657 | 0.004277 | 0.004156 | 6.8e-05 | 0.001267 | 0.29479 | 1.3533 | 2.5481 | 100.58 | 20.68 | 19.149 | 68.5969 | 109.7062 | 67.0423 | 0.004648 | 1.0378 | 0.025002 | 24.0916 | -0.1002 | 0.14196 | -0.000117 | -0.008793 | -0.008793 | -0.66356
2019-03-20 17:47 | 1553118442.0764 | -98 | 0.20952 | 0.000443 | 0 | 107.8787 | 109.1949 | 125.6353 | 21.0606 | 0.001974 | 0.001802 | -151.9114 | 0.23802 | 8.403 | 0.003701 | 0.0053 | 0.005059 | 6.9e-05 | 0.001268 | 0.27975 | 1.3313 | 2.548 | 100.58 | 20.68 | 19.164 | 69.1186 | 110.7859 | 64.9806 | 0.004651 | 1.0441 | 0.025053 | 23.9944 | -0.10847 | 0.1416 | -7.4e-05 | -0.005628 | -0.005628 | -1.0341
2019-03-20 17:47 | 1553118443.0752 | -97 | 0.20953 | 0.000444 | 0 | 107.8764 | 109.0098 | 125.8581 | 21.0598 | 0.001967 | 0.001871 | -151.9366 | 0.23892 | 8.4025 | 0.003757 | 0.005123 | 0.004875 | 7.2e-05 | 0.001266 | 0.28982 | 1.3373 | 2.5481 | 100.58 | 20.68 | 19.196 | 68.9541 | 112.1436 | 66.6522 | 0.004659 | 1.039 | 0.025013 | 24.0749 | -0.09349 | 0.14214 | -0.000112 | -0.008448 | -0.008448 | -0.69151
2019-03-20 17:47 | 1553118444.0757 | -96 | 0.20949 | 0.000441 | 0 | 107.738 | 109.045 | 125.7588 | 21.0598 | 0.001988 | 0.001885 | -151.9492 | 0.23842 | 8.402 | 0.003742 | 0.005171 | 0.004978 | 6.6e-05 | 0.001265 | 0.2589 | 1.328 | 2.548 | 100.58 | 20.68 | 19.211 | 69.4665 | 111.7735 | 66.9877 | 0.004663 | 1.0379 | 0.024991 | 24.0773 | -0.10786 | 0.14184 | 0.000135 | 0.010188 | 0.010188 | 0.57217
2019-03-20 17:47 | 1553118445.0762 | -95 | 0.20956 | 0.000442 | 0 | 107.619 | 108.7011 | 125.4149 | 21.0607 | 0.001977 | 0.001792 | -151.9504 | 0.23875 | 8.4031 | 0.00363 | 0.004665 | 0.004432 | 7e-05 | 0.001264 | 0.27815 | 1.3317 | 2.548 | 100.58 | 20.68 | 19.244 | 69.2054 | 109.0471 | 64.7526 | 0.004671 | 1.0448 | 0.025039 | 23.9652 | -0.089494 | 0.14203 | -0.000266 | -0.020078 | -0.020078 | -0.29074
2019-03-20 17:47 | 1553118446.0759 | -94 | 0.20953 | 0.000442 | 1e-06 | 107.6707 | 108.9003 | 125.5067 | 21.061 | 0.001975 | 0.001835 | -151.9541 | 0.23879 | 8.4018 | 0.003618 | 0.004934 | 0.004668 | 6.9e-05 | 0.001267 | 0.25514 | 1.343 | 2.548 | 100.58 | 20.68 | 19.259 | 69.1525 | 108.766 | 65.7848 | 0.004674 | 1.0416 | 0.02501 | 24.0106 | -0.10157 | 0.14205 | -0.000125 | -0.009439 | -0.009439 | -0.61853
2019-03-20 17:47 | 1553118447.0753 | -93 | 0.20954 | 0.000444 | 0 | 107.5966 | 108.6387 | 125.636 | 21.0599 | 0.001941 | 0.001874 | -151.9531 | 0.23747 | 8.4026 | 0.003647 | 0.004843 | 0.004639 | 6.9e-05 | 0.001264 | 0.25929 | 1.2907 | 2.548 | 100.58 | 20.68 | 19.273 | 68.3349 | 109.454 | 66.7225 | 0.004678 | 1.0388 | 0.024989 | 24.0562 | -0.086211 | 0.14127 | -0.000179 | -0.013522 | -0.013522 | -0.4294
2019-03-20 17:47 | 1553118448.0754 | -92 | 0.20954 | 0.000448 | 0 | 107.5035 | 108.6398 | 125.53 | 21.0608 | 0.001965 | 0.001865 | -151.95 | 0.23725 | 8.4021 | 0.003651 | 0.004685 | 0.004426 | 6.5e-05 | 0.001262 | 0.28512 | 1.3047 | 2.5481 | 100.58 | 20.68 | 19.299 | 68.9029 | 109.5519 | 66.5111 | 0.004684 | 1.0394 | 0.024986 | 24.0385 | -0.09405 | 0.14114 | -0.000173 | -0.013048 | -0.013048 | -0.44458
2019-03-20 17:47 | 1553118449.0758 | -91 | 0.20954 | 0.000444 | 0 | 107.3877 | 108.5681 | 125.4501 | 21.0613 | 0.001963 | 0.001862 | -151.9542 | 0.2382 | 8.4022 | 0.003583 | 0.004008 | 0.00379 | 7e-05 | 0.001266 | 0.26217 | 1.2992 | 2.548 | 100.58 | 20.68 | 19.31 | 68.8697 | 107.9035 | 66.4368 | 0.004687 | 1.0396 | 0.024981 | 24.0283 | -0.097778 | 0.1417 | -0.000192 | -0.014481 | -0.014481 | -0.40219
2019-03-20 17:47 | 1553118450.0769 | -90 | 0.20953 | 0.000447 | 0 | 107.3049 | 108.3214 | 125.4992 | 21.0612 | 0.001926 | 0.001886 | -151.956 | 0.23818 | 8.4029 | 0.003608 | 0.005321 | 0.005065 | 6.9e-05 | 0.001266 | 0.28495 | 1.2948 | 2.548 | 100.58 | 20.68 | 19.321 | 67.9695 | 108.5092 | 67.021 | 0.004689 | 1.0378 | 0.024964 | 24.0536 | -0.084331 | 0.14169 | -0.000135 | -0.010157 | -0.010157 | -0.57337
2019-03-20 17:48 | 1553118451.0755 | -89 | 0.20957 | 0.000444 | 0 | 107.2516 | 108.2975 | 125.4912 | 21.0601 | 0.002 | 0.001823 | -151.9636 | 0.2382 | 8.4018 | 0.003565 | 0.004638 | 0.004367 | 6.9e-05 | 0.001266 | 0.28351 | 1.286 | 2.5481 | 100.58 | 20.69 | 19.34 | 69.7569 | 107.4685 | 65.4958 | 0.004697 | 1.0425 | 0.025019 | 23.9989 | -0.086805 | 0.1417 | -0.00033 | -0.02495 | -0.02495 | -0.23342
2019-03-20 17:48 | 1553118452.0756 | -88 | 0.20946 | 0.000447 | 0 | 107.1591 | 108.2765 | 125.5185 | 21.0603 | 0.001956 | 0.001896 | -151.957 | 0.23786 | 8.4043 | 0.003603 | 0.004305 | 0.004104 | 7.1e-05 | 0.001272 | 0.27959 | 1.3161 | 2.5479 | 100.58 | 20.69 | 19.347 | 68.6887 | 108.3974 | 67.2416 | 0.004699 | 1.0372 | 0.024958 | 24.0633 | -0.092794 | 0.1415 | 0.00032 | 0.024087 | 0.024087 | 0.24144
2019-03-20 17:48 | 1553118453.0761 | -87 | 0.20954 | 0.000446 | 1e-06 | 107.019 | 108.1996 | 125.3478 | 21.06 | 0.001983 | 0.001882 | -151.9459 | 0.23721 | 8.4017 | 0.003536 | 0.004731 | 0.004497 | 6.9e-05 | 0.001277 | 0.286 | 1.3468 | 2.5481 | 100.58 | 20.69 | 19.361 | 69.3472 | 106.7587 | 66.92 | 0.004702 | 1.0382 | 0.024953 | 24.0355 | -0.098128 | 0.14112 | -0.000154 | -0.011611 | -0.011611 | -0.49955
2019-03-20 17:48 | 1553118454.0758 | -86 | 0.20954 | 0.000448 | 1e-06 | 107.018 | 107.9462 | 125.6929 | 21.0611 | 0.001942 | 0.001805 | -151.9344 | 0.23778 | 8.4017 | 0.003573 | 0.00554 | 0.005296 | 6.8e-05 | 0.001277 | 0.27815 | 1.33 | 2.5481 | 100.58 | 20.69 | 19.366 | 68.3478 | 107.6709 | 65.0551 | 0.004703 | 1.0439 | 0.025056 | 24.0025 | -0.077236 | 0.14146 | -0.000153 | -0.011577 | -0.011577 | -0.50219
2019-03-20 17:48 | 1553118455.0767 | -85 | 0.20953 | 0.000446 | 0 | 106.9757 | 108.074 | 125.8173 | 21.0621 | 0.001979 | 0.001823 | -151.9437 | 0.23781 | 8.4025 | 0.003577 | 0.005295 | 0.005058 | 6.6e-05 | 0.001278 | 0.25371 | 1.3159 | 2.5481 | 100.58 | 20.69 | 19.37 | 69.2613 | 107.7554 | 65.4925 | 0.004704 | 1.0425 | 0.025052 | 24.0299 | -0.091363 | 0.14147 | -0.000144 | -0.010858 | -0.010858 | -0.53548
2019-03-20 17:48 | 1553118456.0752 | -84 | 0.20947 | 0.000447 | 0 | 106.9065 | 107.9682 | 125.6283 | 21.0618 | 0.001967 | 0.001868 | -151.9346 | 0.23834 | 8.403 | 0.003617 | 0.00481 | 0.00465 | 7e-05 | 0.001284 | 0.26537 | 1.3301 | 2.5481 | 100.58 | 20.69 | 19.371 | 68.9724 | 108.7457 | 66.5729 | 0.004704 | 1.0392 | 0.024993 | 24.0501 | -0.088394 | 0.14179 | 0.000232 | 0.017452 | 0.017452 | 0.33391
2019-03-20 17:48 | 1553118457.0759 | -83 | 0.20948 | 0.000447 | 0 | 106.814 | 107.914 | 125.4945 | 21.0605 | 0.001946 | 0.001843 | -151.9348 | 0.23865 | 8.4001 | 0.003608 | 0.005035 | 0.004705 | 6.6e-05 | 0.00129 | 0.28495 | 1.3103 | 2.5481 | 100.58 | 20.69 | 19.368 | 68.4463 | 108.504 | 65.974 | 0.004704 | 1.0411 | 0.025002 | 24.0161 | -0.091645 | 0.14197 | 0.000166 | 0.012547 | 0.012547 | 0.46504
2019-03-20 17:48 | 1553118458.0753 | -82 | 0.20949 | 0.000446 | 0 | 106.7415 | 107.8019 | 125.9116 | 21.0616 | 0.001968 | 0.001887 | -151.9365 | 0.23749 | 8.4035 | 0.003574 | 0.005338 | 0.005224 | 6.9e-05 | 0.001293 | 0.28943 | 1.3536 | 2.5482 | 100.58 | 20.7 | 19.357 | 68.9731 | 107.6904 | 67.0269 | 0.004704 | 1.0378 | 0.025005 | 24.0933 | -0.088415 | 0.14128 | 0.00011 | 0.008296 | 0.008296 | 0.69993
2019-03-20 17:48 | 1553118459.0749 | -81 | 0.20953 | 0.000447 | 0 | 106.7603 | 107.7096 | 126.0884 | 21.0614 | 0.001951 | 0.001843 | -151.9245 | 0.23806 | 8.4026 | 0.003688 | 0.005349 | 0.005175 | 6.6e-05 | 0.001288 | 0.2636 | 1.3406 | 2.548 | 100.58 | 20.69 | 19.349 | 68.5814 | 110.4672 | 65.9703 | 0.004699 | 1.0411 | 0.025061 | 24.0727 | -0.07918 | 0.14162 | -0.000113 | -0.008517 | -0.008517 | -0.68341
2019-03-20 17:48 | 1553118460.075 | -80 | 0.20953 | 0.000445 | 0 | 106.7454 | 107.732 | 126.2292 | 21.0617 | 0.001963 | 0.001847 | -151.9116 | 0.23728 | 8.4024 | 0.003653 | 0.004799 | 0.004605 | 6.5e-05 | 0.001287 | 0.29015 | 1.2988 | 2.5481 | 100.58 | 20.69 | 19.34 | 68.8585 | 109.6174 | 66.0742 | 0.004697 | 1.0407 | 0.025071 | 24.0899 | -0.082289 | 0.14116 | -8.6e-05 | -0.006498 | -0.006498 | -0.89281
2019-03-20 17:48 | 1553118461.0765 | -79 | 0.20955 | 0.000448 | 0 | 106.6778 | 107.757 | 126.4989 | 21.062 | 0.001929 | 0.001808 | -151.8813 | 0.2374 | 8.4022 | 0.003689 | 0.005046 | 0.004837 | 6.8e-05 | 0.001281 | 0.26864 | 1.3579 | 2.5482 | 100.58 | 20.69 | 19.318 | 68.0381 | 110.4741 | 65.1395 | 0.004692 | 1.0436 | 0.025133 | 24.0824 | -0.090034 | 0.14123 | -0.00022 | -0.016709 | -0.016709 | -0.3474
2019-03-20 17:48 | 1553118462.0761 | -78 | 0.20957 | 0.000446 | 0 | 106.6309 | 107.6236 | 126.3137 | 21.0614 | 0.001988 | 0.001828 | -151.8803 | 0.23733 | 8.4031 | 0.00363 | 0.005305 | 0.005111 | 6.7e-05 | 0.001277 | 0.29878 | 1.344 | 2.5481 | 100.58 | 20.69 | 19.305 | 69.4549 | 109.059 | 65.6292 | 0.004688 | 1.0421 | 0.025096 | 24.0821 | -0.082884 | 0.14119 | -0.000372 | -0.028172 | -0.028172 | -0.20598
2019-03-20 17:48 | 1553118463.0762 | -77 | 0.20954 | 0.00045 | 0 | 106.5284 | 107.5275 | 126.2604 | 21.062 | 0.001965 | 0.001846 | -151.8558 | 0.23812 | 8.4038 | 0.003643 | 0.004466 | 0.004275 | 6.7e-05 | 0.00128 | 0.2688 | 1.3298 | 2.548 | 100.58 | 20.69 | 19.278 | 68.9134 | 109.3622 | 66.041 | 0.004682 | 1.0408 | 0.025076 | 24.0917 | -0.083502 | 0.14166 | -0.000201 | -0.015247 | -0.015247 | -0.38186
2019-03-20 17:48 | 1553118464.0766 | -76 | 0.20953 | 0.000449 | 0 | 106.5109 | 107.5814 | 126.0752 | 21.0614 | 0.001962 | 0.001863 | -151.857 | 0.23745 | 8.4033 | 0.003626 | 0.004338 | 0.004084 | 6.6e-05 | 0.001291 | 0.28407 | 1.3071 | 2.548 | 100.58 | 20.69 | 19.264 | 68.8472 | 108.9431 | 66.4547 | 0.004678 | 1.0396 | 0.025042 | 24.0887 | -0.089444 | 0.14126 | -0.000122 | -0.009225 | -0.009225 | -0.62936
2019-03-20 17:48 | 1553118465.0763 | -75 | 0.20948 | 0.000446 | 0 | 106.5941 | 107.4831 | 125.978 | 21.0616 | 0.00195 | 0.00181 | -151.8627 | 0.23775 | 8.3998 | 0.003597 | 0.005248 | 0.00501 | 6.5e-05 | 0.001292 | 0.26626 | 1.325 | 2.5482 | 100.58 | 20.7 | 19.249 | 68.5479 | 108.2459 | 65.1836 | 0.004678 | 1.0435 | 0.025079 | 24.0343 | -0.07429 | 0.14144 | 0.000178 | 0.013449 | 0.013449 | 0.43223
2019-03-20 17:48 | 1553118466.0763 | -74 | 0.20954 | 0.000448 | -1e-06 | 106.3838 | 107.4382 | 125.9454 | 21.062 | 0.002004 | 0.001818 | -151.8709 | 0.23783 | 8.4022 | 0.00358 | 0.005235 | 0.004996 | 6.5e-05 | 0.001298 | 0.29623 | 1.3061 | 2.5481 | 100.58 | 20.69 | 19.218 | 69.8421 | 107.8325 | 65.3714 | 0.004667 | 1.0429 | 0.025069 | 24.0378 | -0.088215 | 0.14148 | -0.000191 | -0.014435 | -0.014435 | -0.40284
2019-03-20 17:48 | 1553118467.0767 | -73 | 0.20954 | 0.000449 | 0 | 106.334 | 107.3425 | 126.1154 | 21.0624 | 0.00194 | 0.001815 | -151.8742 | 0.2379 | 8.4019 | 0.003546 | 0.00497 | 0.0048 | 6.6e-05 | 0.001294 | 0.29535 | 1.3865 | 2.5481 | 100.58 | 20.69 | 19.203 | 68.3179 | 107.0108 | 65.3015 | 0.004664 | 1.0431 | 0.025089 | 24.0516 | -0.084434 | 0.14153 | -0.00017 | -0.012879 | -0.012879 | -0.45166
2019-03-20 17:48 | 1553118468.0758 | -72 | 0.2095 | 0.000448 | 0 | 106.2997 | 107.3089 | 125.8182 | 21.0623 | 0.001956 | 0.001893 | -151.8808 | 0.23786 | 8.4057 | 0.003584 | 0.005035 | 0.004953 | 6.5e-05 | 0.001297 | 0.28816 | 1.3594 | 2.548 | 100.58 | 20.69 | 19.176 | 68.6882 | 107.9423 | 67.1788 | 0.004657 | 1.0374 | 0.02499 | 24.0898 | -0.084515 | 0.14151 | 3.6e-05 | 0.00266 | 0.00266 | 2.1863
2019-03-20 17:48 | 1553118469.0765 | -71 | 0.20958 | 0.000451 | 1e-06 | 106.2677 | 107.4356 | 125.6175 | 21.0611 | 0.001907 | 0.001847 | -151.8894 | 0.23713 | 8.4014 | 0.003646 | 0.005212 | 0.004866 | 6.4e-05 | 0.001293 | 0.27367 | 1.3685 | 2.5481 | 100.58 | 20.69 | 19.165 | 67.528 | 109.4393 | 66.0743 | 0.004654 | 1.0407 | 0.025011 | 24.0314 | -0.097764 | 0.14107 | -0.000444 | -0.033523 | -0.033523 | -0.17295
2019-03-20 17:48 | 1553118470.076 | -70 | 0.20952 | 0.000447 | 0 | 106.2968 | 107.2667 | 125.2973 | 21.0615 | 0.001984 | 0.001844 | -151.8983 | 0.23766 | 8.4053 | 0.003558 | 0.005134 | 0.005035 | 6.5e-05 | 0.001298 | 0.27815 | 1.3871 | 2.5481 | 100.58 | 20.69 | 19.156 | 69.373 | 107.2917 | 65.9944 | 0.004652 | 1.041 | 0.024982 | 23.998 | -0.081239 | 0.14139 | -4.5e-05 | -0.003385 | -0.003385 | -1.7167
2019-03-20 17:48 | 1553118471.076 | -69 | 0.20956 | 0.000449 | 0 | 106.2758 | 107.2891 | 125.2067 | 21.0613 | 0.001989 | 0.001854 | -151.9163 | 0.23818 | 8.4016 | 0.003473 | 0.005492 | 0.005186 | 6.6e-05 | 0.0013 | 0.27041 | 1.3418 | 2.5481 | 100.58 | 20.69 | 19.14 | 69.4961 | 105.2358 | 66.2494 | 0.004648 | 1.0402 | 0.024963 | 23.9983 | -0.084878 | 0.14169 | -0.000282 | -0.021209 | -0.021209 | -0.27458
2019-03-20 17:48 | 1553118472.0757 | -68 | 0.20949 | 0.000449 | 0 | 106.2813 | 107.2476 | 125.3497 | 21.0617 | 0.001953 | 0.001834 | -151.9212 | 0.2382 | 8.4038 | 0.003644 | 0.005009 | 0.004767 | 6.8e-05 | 0.001305 | 0.25282 | 1.3151 | 2.5482 | 100.58 | 20.69 | 19.132 | 68.6292 | 109.3855 | 65.7605 | 0.004646 | 1.0417 | 0.024995 | 23.9947 | -0.080956 | 0.1417 | 0.00013 | 0.009789 | 0.009789 | 0.59497
2019-03-20 17:48 | 1553118473.0751 | -67 | 0.20953 | 0.000447 | 0 | 106.2118 | 107.3863 | 125.3276 | 21.0611 | 0.001983 | 0.001836 | -151.9481 | 0.23745 | 8.4022 | 0.003449 | 0.004912 | 0.004628 | 6.2e-05 | 0.001304 | 0.25752 | 1.3217 | 2.548 | 100.58 | 20.69 | 19.125 | 69.3382 | 104.67 | 65.8138 | 0.004645 | 1.0415 | 0.024991 | 23.9945 | -0.098366 | 0.14126 | -9.1e-05 | -0.006879 | -0.006879 | -0.84393
2019-03-20 17:48 | 1553118474.0752 | -66 | 0.20955 | 0.000446 | 0 | 106.0782 | 106.9138 | 125.8062 | 21.0614 | 0.001987 | 0.001817 | -151.9503 | 0.23813 | 8.4025 | 0.003455 | 0.005111 | 0.0049 | 6.2e-05 | 0.001303 | 0.27704 | 1.3324 | 2.5481 | 100.58 | 20.69 | 19.112 | 69.4438 | 104.7922 | 65.352 | 0.004642 | 1.043 | 0.025056 | 24.0239 | -0.070186 | 0.14166 | -0.000233 | -0.017629 | -0.017629 | -0.33027
2019-03-20 17:48 | 1553118475.0748 | -65 | 0.20953 | 0.000452 | 0 | 106.0245 | 106.8709 | 125.8968 | 21.0626 | 0.001982 | 0.001856 | -151.9364 | 0.23751 | 8.4047 | 0.003537 | 0.004519 | 0.004294 | 6.3e-05 | 0.001304 | 0.29479 | 1.3506 | 2.5481 | 100.58 | 20.69 | 19.108 | 69.3257 | 106.8039 | 66.2913 | 0.004641 | 1.0401 | 0.02503 | 24.0658 | -0.071117 | 0.1413 | -0.000108 | -0.008173 | -0.008173 | -0.71051
2019-03-20 17:48 | 1553118476.0763 | -64 | 0.20953 | 0.000449 | 0 | 106.1296 | 107.0847 | 126.013 | 21.0616 | 0.001937 | 0.001849 | -151.9159 | 0.2376 | 8.4009 | 0.003537 | 0.005139 | 0.005003 | 6.5e-05 | 0.001294 | 0.2995 | 1.3402 | 2.5481 | 100.58 | 20.69 | 19.099 | 68.2476 | 106.7841 | 66.1294 | 0.004638 | 1.0406 | 0.025048 | 24.0712 | -0.080129 | 0.14135 | -8.9e-05 | -0.006735 | -0.006735 | -0.86252
2019-03-20 17:48 | 1553118477.0759 | -63 | 0.20953 | 0.000447 | 0 | 106.0459 | 107.0563 | 125.9668 | 21.0621 | 0.001955 | 0.00189 | -151.9384 | 0.23879 | 8.4001 | 0.003503 | 0.004774 | 0.004542 | 6.6e-05 | 0.001297 | 0.26681 | 1.2896 | 2.5479 | 100.58 | 20.69 | 19.095 | 68.6647 | 105.966 | 67.1084 | 0.004637 | 1.0376 | 0.025007 | 24.1015 | -0.084819 | 0.14205 | -0.000132 | -0.009939 | -0.009939 | -0.58745
2019-03-20 17:48 | 1553118478.076 | -62 | 0.2095 | 0.000446 | 1e-06 | 106.0889 | 107.0316 | 125.9804 | 21.0623 | 0.001984 | 0.001845 | -151.9286 | 0.23883 | 8.4017 | 0.003467 | 0.005233 | 0.005019 | 6.8e-05 | 0.001294 | 0.27743 | 1.3183 | 2.548 | 100.58 | 20.69 | 19.092 | 69.3737 | 105.0997 | 66.0295 | 0.004637 | 1.0409 | 0.025048 | 24.0645 | -0.07913 | 0.14208 | 5.7e-05 | 0.004246 | 0.004246 | 1.3753
2019-03-20 17:48 | 1553118479.0772 | -61 | 0.20953 | 0.000447 | 0 | 106.063 | 107.0256 | 125.9003 | 21.0616 | 0.002006 | 0.001833 | -151.9223 | 0.23802 | 8.4008 | 0.003454 | 0.004502 | 0.00431 | 6.5e-05 | 0.001297 | 0.29264 | 1.3157 | 2.5482 | 100.58 | 20.69 | 19.086 | 69.8987 | 104.7837 | 65.7323 | 0.004635 | 1.0418 | 0.025051 | 24.0464 | -0.080812 | 0.14159 | -0.000118 | -0.008907 | -0.008907 | -0.65337
2019-03-20 17:48 | 1553118480.0753 | -60 | 0.20953 | 0.000455 | 0 | 106.0653 | 107.1474 | 125.9804 | 21.0629 | 0.001919 | 0.001785 | -151.9193 | 0.23733 | 8.4032 | 0.003452 | 0.005063 | 0.004804 | 6.2e-05 | 0.001298 | 0.27704 | 1.2857 | 2.548 | 100.58 | 20.69 | 19.085 | 67.8051 | 104.7425 | 64.58 | 0.004635 | 1.0453 | 0.025102 | 24.0131 | -0.090798 | 0.14119 | -0.000121 | -0.009154 | -0.009154 | -0.63391
2019-03-20 17:48 | 1553118481.0751 | -59 | 0.20952 | 0.000447 | 0 | 106.0598 | 107.1306 | 125.9079 | 21.0625 | 0.001961 | 0.001824 | -151.9134 | 0.23762 | 8.4049 | 0.003483 | 0.004914 | 0.004692 | 6.5e-05 | 0.001297 | 0.26626 | 1.345 | 2.5482 | 100.58 | 20.69 | 19.084 | 68.8149 | 105.4816 | 65.5263 | 0.004635 | 1.0424 | 0.02506 | 24.0398 | -0.089858 | 0.14136 | -6.1e-05 | -0.004642 | -0.004642 | -1.2517
2019-03-20 17:48 | 1553118482.0755 | -58 | 0.20951 | 0.00045 | 0 | 106.0109 | 106.9912 | 126.0478 | 21.0623 | 0.001963 | 0.001854 | -151.9163 | 0.23843 | 8.4021 | 0.003501 | 0.005289 | 0.005011 | 6.8e-05 | 0.001303 | 0.28136 | 1.3042 | 2.548 | 100.58 | 20.69 | 19.086 | 68.8591 | 105.9141 | 66.2521 | 0.004635 | 1.0402 | 0.025047 | 24.0789 | -0.08233 | 0.14184 | 8e-06 | 0.000602 | 0.000602 | 9.6823
2019-03-20 17:48 | 1553118483.0763 | -57 | 0.20951 | 0.000447 | 0 | 106.1021 | 107.2121 | 126.3984 | 21.0634 | 0.001975 | 0.001844 | -151.9072 | 0.23845 | 8.4013 | 0.003493 | 0.004766 | 0.00456 | 6.9e-05 | 0.001304 | 0.27959 | 1.3434 | 2.5481 | 100.58 | 20.7 | 19.088 | 69.1532 | 105.724 | 65.9938 | 0.004639 | 1.041 | 0.025091 | 24.1032 | -0.093085 | 0.14185 | 8e-06 | 0.000632 | 0.000632 | 9.2239
2019-03-20 17:48 | 1553118484.0754 | -56 | 0.20952 | 0.000444 | 0 | 106.0103 | 106.9266 | 126.0834 | 21.0635 | 0.002032 | 0.00186 | -151.9138 | 0.23834 | 8.4029 | 0.003455 | 0.004906 | 0.004708 | 6.8e-05 | 0.00131 | 0.28296 | 1.3099 | 2.548 | 100.58 | 20.7 | 19.093 | 70.518 | 104.7991 | 66.3896 | 0.00464 | 1.0398 | 0.025045 | 24.0872 | -0.076977 | 0.14179 | -6.4e-05 | -0.004852 | -0.004852 | -1.2012
2019-03-20 17:48 | 1553118485.0754 | -55 | 0.20951 | 0.00044 | 0 | 106.0184 | 106.9497 | 126.0172 | 21.0632 | 0.002007 | 0.001839 | -151.9121 | 0.23789 | 8.4042 | 0.003465 | 0.00531 | 0.005118 | 6.7e-05 | 0.001312 | 0.30469 | 1.3146 | 2.5481 | 100.58 | 20.7 | 19.097 | 69.9159 | 105.0418 | 65.8758 | 0.004641 | 1.0414 | 0.025058 | 24.0626 | -0.078234 | 0.14152 | 2.6e-05 | 0.001946 | 0.001946 | 2.9888
2019-03-20 17:48 | 1553118486.0754 | -54 | 0.20953 | 0.000441 | 0 | 105.9281 | 106.912 | 126.1461 | 21.0622 | 0.001975 | 0.00188 | -151.9359 | 0.23945 | 8.4005 | 0.003484 | 0.005067 | 0.004783 | 6.7e-05 | 0.001311 | 0.29911 | 1.3296 | 2.5481 | 100.58 | 20.69 | 19.1 | 69.1467 | 105.5191 | 66.8651 | 0.004639 | 1.0383 | 0.025034 | 24.11 | -0.082692 | 0.14245 | -7.8e-05 | -0.005922 | -0.005922 | -0.98871
2019-03-20 17:48 | 1553118487.0765 | -53 | 0.20952 | 0.000443 | 0 | 105.8935 | 106.8477 | 126.1862 | 21.0618 | 0.001988 | 0.001869 | -151.9122 | 0.23869 | 8.4031 | 0.003436 | 0.005677 | 0.005523 | 6.7e-05 | 0.001315 | 0.27113 | 1.3227 | 2.5479 | 100.58 | 20.7 | 19.108 | 69.4552 | 104.3554 | 66.6118 | 0.004643 | 1.0391 | 0.025047 | 24.1049 | -0.080234 | 0.142 | -5.8e-05 | -0.004358 | -0.004358 | -1.3392
2019-03-20 17:48 | 1553118488.0745 | -52 | 0.20951 | 0.000444 | 0 | 105.7988 | 106.5386 | 126.2219 | 21.0625 | 0.002005 | 0.001887 | -151.9237 | 0.23863 | 8.4013 | 0.003456 | 0.004767 | 0.004558 | 6.4e-05 | 0.001304 | 0.2855 | 1.3338 | 2.5482 | 100.58 | 20.7 | 19.112 | 69.8645 | 104.8357 | 67.0316 | 0.004644 | 1.0378 | 0.025035 | 24.1232 | -0.06233 | 0.14196 | -9e-06 | -0.000718 | -0.000718 | -8.1313
2019-03-20 17:48 | 1553118489.075 | -51 | 0.20953 | 0.000441 | 0 | 105.6885 | 106.5263 | 126.4287 | 21.0634 | 0.002002 | 0.001838 | -151.9396 | 0.23816 | 8.4008 | 0.003488 | 0.004728 | 0.004547 | 6.6e-05 | 0.001299 | 0.2943 | 1.3198 | 2.5481 | 100.58 | 20.7 | 19.124 | 69.8036 | 105.6051 | 65.8706 | 0.004647 | 1.0414 | 0.025099 | 24.1017 | -0.070623 | 0.14168 | -0.000106 | -0.008033 | -0.008033 | -0.72489
2019-03-20 17:48 | 1553118490.075 | -50 | 0.20952 | 0.000442 | 0 | 105.6817 | 106.5831 | 126.4834 | 21.0622 | 0.00195 | 0.001832 | -151.9435 | 0.23785 | 8.4026 | 0.003442 | 0.005152 | 0.004995 | 6.7e-05 | 0.001293 | 0.26952 | 1.3399 | 2.548 | 100.58 | 20.7 | 19.131 | 68.5609 | 104.498 | 65.7127 | 0.004649 | 1.0419 | 0.02511 | 24.1013 | -0.075966 | 0.1415 | -5.7e-05 | -0.004315 | -0.004315 | -1.3477
2019-03-20 17:48 | 1553118491.0746 | -49 | 0.20952 | 0.000445 | 0 | 105.4912 | 106.4448 | 126.6374 | 21.062 | 0.001962 | 0.001814 | -151.9491 | 0.2386 | 8.4036 | 0.00346 | 0.005038 | 0.004779 | 7e-05 | 0.001288 | 0.26504 | 1.3322 | 2.5481 | 100.58 | 20.7 | 19.137 | 68.8517 | 104.9204 | 65.2855 | 0.004651 | 1.0432 | 0.025141 | 24.1007 | -0.080491 | 0.14194 | -3.9e-05 | -0.002982 | -0.002982 | -1.9561
2019-03-20 17:48 | 1553118492.0761 | -48 | 0.20952 | 0.000441 | 0 | 105.5284 | 106.3155 | 126.4519 | 21.0622 | 0.001959 | 0.001806 | -151.9321 | 0.23827 | 8.4047 | 0.003487 | 0.004928 | 0.004738 | 6.7e-05 | 0.001277 | 0.28567 | 1.3066 | 2.548 | 100.58 | 20.7 | 19.148 | 68.7794 | 105.5889 | 65.0889 | 0.004653 | 1.0438 | 0.02513 | 24.0761 | -0.066465 | 0.14175 | -2.9e-05 | -0.002228 | -0.002228 | -2.6149
2019-03-20 17:48 | 1553118493.0752 | -47 | 0.20953 | 0.000437 | 0 | 105.4304 | 106.3286 | 126.3142 | 21.0622 | 0.001928 | 0.001884 | -151.9339 | 0.23836 | 8.4023 | 0.003441 | 0.005137 | 0.004876 | 6.8e-05 | 0.001275 | 0.27345 | 1.305 | 2.548 | 100.58 | 20.7 | 19.152 | 68.0212 | 104.477 | 66.9635 | 0.004654 | 1.038 | 0.025047 | 24.1296 | -0.075875 | 0.1418 | -9.4e-05 | -0.007102 | -0.007102 | -0.82057
2019-03-20 17:48 | 1553118494.0758 | -46 | 0.20953 | 0.000443 | 1e-06 | 105.3379 | 106.1368 | 126.3226 | 21.0613 | 0.001983 | 0.001851 | -151.936 | 0.23913 | 8.4017 | 0.003448 | 0.00432 | 0.004054 | 6.7e-05 | 0.001279 | 0.29087 | 1.3809 | 2.5484 | 100.58 | 20.7 | 19.154 | 69.3564 | 104.6317 | 66.1729 | 0.004655 | 1.0404 | 0.025077 | 24.1023 | -0.067582 | 0.14226 | -8.7e-05 | -0.00659 | -0.00659 | -0.88721
2019-03-20 17:48 | 1553118495.0749 | -45 | 0.20951 | 0.000444 | 0 | 105.2687 | 106.1312 | 126.4571 | 21.0625 | 0.002015 | 0.001817 | -151.9537 | 0.23766 | 8.4033 | 0.003442 | 0.005021 | 0.004773 | 6.8e-05 | 0.001279 | 0.25984 | 1.3727 | 2.5481 | 100.58 | 20.7 | 19.156 | 70.1102 | 104.4899 | 65.3555 | 0.004655 | 1.043 | 0.025121 | 24.0861 | -0.072988 | 0.14138 | 4.1e-05 | 0.003109 | 0.003109 | 1.8692
2019-03-20 17:48 | 1553118496.076 | -44 | 0.20952 | 0.000445 | 0 | 105.1594 | 106.0213 | 126.0285 | 21.0616 | 0.00198 | 0.001858 | -151.9455 | 0.23783 | 8.4033 | 0.003516 | 0.004983 | 0.004861 | 6.6e-05 | 0.001279 | 0.27688 | 1.3235 | 2.5481 | 100.58 | 20.7 | 19.158 | 69.2709 | 106.2944 | 66.3331 | 0.004656 | 1.0399 | 0.025042 | 24.0799 | -0.073016 | 0.14149 | -1.2e-05 | -0.000873 | -0.000873 | -6.6609
2019-03-20 17:48 | 1553118497.0762 | -43 | 0.20952 | 0.000442 | 0 | 105.1717 | 106.0553 | 126.1646 | 21.0624 | 0.002002 | 0.00182 | -151.9407 | 0.23861 | 8.4006 | 0.00344 | 0.00489 | 0.004651 | 7.1e-05 | 0.001282 | 0.2808 | 1.3139 | 2.5481 | 100.58 | 20.7 | 19.156 | 69.8068 | 104.4331 | 65.4257 | 0.004655 | 1.0427 | 0.025089 | 24.0607 | -0.07484 | 0.14195 | -3.5e-05 | -0.002677 | -0.002677 | -2.1794
2019-03-20 17:48 | 1553118498.0752 | -42 | 0.20953 | 0.000439 | 0 | 105.1335 | 105.8584 | 126.1144 | 21.063 | 0.00199 | 0.00186 | -151.9448 | 0.23904 | 8.4011 | 0.003418 | 0.004979 | 0.004833 | 6.9e-05 | 0.001287 | 0.27455 | 1.3216 | 2.5479 | 100.58 | 20.7 | 19.154 | 69.5069 | 103.9107 | 66.39 | 0.004655 | 1.0398 | 0.025048 | 24.0901 | -0.061458 | 0.14221 | -8.8e-05 | -0.006689 | -0.006689 | -0.87372
2019-03-20 17:48 | 1553118499.076 | -41 | 0.20951 | 0.000444 | -1e-06 | 105.0898 | 105.7765 | 126.1623 | 21.0628 | 0.001976 | 0.0018 | -151.9235 | 0.23821 | 8.4016 | 0.003442 | 0.004982 | 0.004691 | 6.6e-05 | 0.001294 | 0.28064 | 1.3262 | 2.548 | 100.58 | 20.69 | 19.149 | 69.1894 | 104.4933 | 64.9382 | 0.004651 | 1.0442 | 0.025107 | 24.0431 | -0.058255 | 0.14171 | 1.7e-05 | 0.001316 | 0.001316 | 4.4258
2019-03-20 17:48 | 1553118500.0754 | -40 | 0.20952 | 0.000445 | 0 | 105.0743 | 105.6872 | 126.1137 | 21.0618 | 0.001953 | 0.001852 | -151.927 | 0.23811 | 8.4023 | 0.003386 | 0.004036 | 0.003872 | 6.6e-05 | 0.001291 | 0.27024 | 1.3597 | 2.5481 | 100.58 | 20.69 | 19.148 | 68.6221 | 103.1269 | 66.1953 | 0.00465 | 1.0404 | 0.025055 | 24.0832 | -0.05202 | 0.14165 | -5.4e-05 | -0.004081 | -0.004081 | -1.4266
2019-03-20 17:48 | 1553118501.0762 | -39 | 0.20955 | 0.000437 | 0 | 104.9223 | 105.7377 | 126.1919 | 21.0624 | 0.001973 | 0.001841 | -151.9116 | 0.23834 | 8.4027 | 0.003393 | 0.004864 | 0.004643 | 6.9e-05 | 0.001296 | 0.29806 | 1.3358 | 2.5481 | 100.58 | 20.7 | 19.149 | 69.097 | 103.3011 | 65.9399 | 0.004653 | 1.0412 | 0.025073 | 24.0816 | -0.069238 | 0.14179 | -0.000219 | -0.016554 | -0.016554 | -0.35203
2019-03-20 17:48 | 1553118502.0752 | -38 | 0.20953 | 0.000443 | 0 | 104.8376 | 105.7724 | 126.5609 | 21.0631 | 0.001962 | 0.001824 | -151.9272 | 0.23871 | 8.4025 | 0.003343 | 0.004822 | 0.004587 | 6.7e-05 | 0.001296 | 0.29535 | 1.3216 | 2.5482 | 100.58 | 20.7 | 19.149 | 68.8339 | 102.0833 | 65.5262 | 0.004653 | 1.0424 | 0.025125 | 24.102 | -0.079404 | 0.14201 | -0.000121 | -0.009144 | -0.009144 | -0.63826
2019-03-20 17:48 | 1553118503.0757 | -37 | 0.20953 | 0.000443 | 0 | 104.7577 | 105.4809 | 126.1313 | 21.0629 | 0.001923 | 0.001864 | -151.9293 | 0.23837 | 8.4027 | 0.003461 | 0.004796 | 0.004608 | 7.1e-05 | 0.001299 | 0.29574 | 1.3438 | 2.5479 | 100.58 | 20.7 | 19.151 | 67.8986 | 104.9588 | 66.4922 | 0.004654 | 1.0395 | 0.025046 | 24.0954 | -0.061536 | 0.1418 | -0.000119 | -0.009012 | -0.009012 | -0.64673
2019-03-20 17:48 | 1553118504.0743 | -36 | 0.20954 | 0.000441 | 0 | 104.8457 | 105.6681 | 126.0663 | 21.0633 | 0.001979 | 0.001813 | -151.9145 | 0.23811 | 8.4033 | 0.003454 | 0.005048 | 0.004849 | 6.7e-05 | 0.0013 | 0.28495 | 1.297 | 2.548 | 100.58 | 20.69 | 19.153 | 69.2437 | 104.7752 | 65.2617 | 0.004652 | 1.0432 | 0.025085 | 24.0455 | -0.069892 | 0.14165 | -0.000171 | -0.012957 | -0.012957 | -0.44933
2019-03-20 17:48 | 1553118505.0744 | -35 | 0.20952 | 0.000441 | 0 | 104.8204 | 105.5642 | 126.0919 | 21.0629 | 0.001997 | 0.001819 | -151.8949 | 0.23762 | 8.4037 | 0.0035 | 0.004565 | 0.004335 | 6.7e-05 | 0.0013 | 0.27312 | 1.3251 | 2.5481 | 100.58 | 20.69 | 19.155 | 69.6885 | 105.9045 | 65.3942 | 0.004652 | 1.0428 | 0.025083 | 24.0526 | -0.063245 | 0.14136 | -1e-05 | -0.000717 | -0.000717 | -8.0978
2019-03-20 17:48 | 1553118506.0745 | -34 | 0.20954 | 0.000442 | -1e-06 | 104.7642 | 105.5601 | 126.328 | 21.0632 | 0.001965 | 0.001828 | -151.8972 | 0.23786 | 8.4029 | 0.003478 | 0.004218 | 0.003998 | 7.1e-05 | 0.0013 | 0.27992 | 1.3898 | 2.5482 | 100.58 | 20.69 | 19.161 | 68.9042 | 105.3628 | 65.6105 | 0.004653 | 1.0422 | 0.025098 | 24.0828 | -0.0677 | 0.1415 | -0.00014 | -0.010499 | -0.010499 | -0.55395
2019-03-20 17:48 | 1553118507.0756 | -33 | 0.20952 | 0.000442 | -1e-06 | 104.8295 | 105.528 | 126.0602 | 21.0634 | 0.001893 | 0.001827 | -151.898 | 0.23752 | 8.403 | 0.003486 | 0.004987 | 0.004795 | 6.9e-05 | 0.001298 | 0.30127 | 1.329 | 2.548 | 100.58 | 20.69 | 19.164 | 67.1831 | 105.5548 | 65.5905 | 0.004654 | 1.0422 | 0.025072 | 24.0566 | -0.0594 | 0.1413 | -3.2e-05 | -0.002397 | -0.002397 | -2.4225
2019-03-20 17:48 | 1553118508.0756 | -32 | 0.20952 | 0.000438 | 0 | 104.6908 | 105.4469 | 126.0935 | 21.0625 | 0.001973 | 0.001822 | -151.9015 | 0.23835 | 8.4021 | 0.0035 | 0.00479 | 0.004591 | 7.2e-05 | 0.001297 | 0.29767 | 1.3422 | 2.5482 | 100.58 | 20.69 | 19.166 | 69.1019 | 105.9048 | 65.4648 | 0.004655 | 1.0426 | 0.02508 | 24.0553 | -0.064369 | 0.14179 | -2.8e-05 | -0.002109 | -0.002109 | -2.7635
2019-03-20 17:48 | 1553118509.0749 | -31 | 0.20952 | 0.000434 | 0 | 104.7283 | 105.5706 | 126.1452 | 21.0628 | 0.001993 | 0.001822 | -151.9112 | 0.23769 | 8.4025 | 0.003512 | 0.005385 | 0.005232 | 7e-05 | 0.001296 | 0.29048 | 1.349 | 2.5481 | 100.58 | 20.69 | 19.177 | 69.5893 | 106.1806 | 65.4813 | 0.004657 | 1.0426 | 0.025085 | 24.0608 | -0.071652 | 0.1414 | 5e-06 | 0.000395 | 0.000395 | 14.6969
2019-03-20 17:48 | 1553118510.0747 | -30 | 0.20951 | 0.000447 | 0 | 104.7422 | 105.4831 | 126.0992 | 21.0636 | 0.001905 | 0.001826 | -151.9155 | 0.23798 | 8.4014 | 0.003554 | 0.004816 | 0.004629 | 6.5e-05 | 0.001288 | 0.2724 | 1.334 | 2.5481 | 100.58 | 20.69 | 19.182 | 67.479 | 107.2151 | 65.5726 | 0.004659 | 1.0423 | 0.025077 | 24.0597 | -0.063051 | 0.14158 | -7e-06 | -0.000544 | -0.000544 | -10.6893
2019-03-20 17:49 | 1553118511.0757 | -29 | 0.20949 | 0.00044 | 0 | 104.7047 | 105.528 | 125.8752 | 21.0629 | 0.001998 | 0.001812 | -151.9378 | 0.23865 | 8.4011 | 0.003481 | 0.005335 | 0.005088 | 6.4e-05 | 0.001288 | 0.27184 | 1.4101 | 2.5481 | 100.58 | 20.69 | 19.187 | 69.7041 | 105.4316 | 65.2332 | 0.00466 | 1.0433 | 0.025067 | 24.0262 | -0.070058 | 0.14197 | 0.000134 | 0.010145 | 0.010145 | 0.57517
2019-03-20 17:49 | 1553118512.0752 | -28 | 0.20955 | 0.000439 | 0 | 104.6442 | 105.4588 | 125.3743 | 21.0628 | 0.001948 | 0.001857 | -151.938 | 0.239 | 8.4022 | 0.003429 | 0.004714 | 0.004516 | 6.5e-05 | 0.001289 | 0.28998 | 1.3448 | 2.5482 | 100.58 | 20.69 | 19.193 | 68.5046 | 104.1724 | 66.329 | 0.004661 | 1.04 | 0.024977 | 24.0172 | -0.069363 | 0.14218 | -0.00021 | -0.015851 | -0.015851 | -0.36867
2019-03-20 17:49 | 1553118513.0748 | -27 | 0.20955 | 0.000438 | 0 | 104.5245 | 105.3037 | 125.7705 | 21.0618 | 0.001927 | 0.001851 | -151.9239 | 0.23928 | 8.4026 | 0.003506 | 0.004632 | 0.004424 | 6.2e-05 | 0.001285 | 0.27295 | 1.3761 | 2.5482 | 100.58 | 20.69 | 19.193 | 68.0045 | 106.0499 | 66.1734 | 0.004661 | 1.0404 | 0.025022 | 24.0496 | -0.066433 | 0.14235 | -0.000204 | -0.015404 | -0.015404 | -0.37982
2019-03-20 17:49 | 1553118514.0759 | -26 | 0.20953 | 0.000445 | 0 | 104.4178 | 105.2391 | 125.7083 | 21.0624 | 0.001926 | 0.001839 | -151.9278 | 0.23879 | 8.4028 | 0.003538 | 0.004814 | 0.004587 | 6.6e-05 | 0.001281 | 0.29502 | 1.375 | 2.5479 | 100.58 | 20.69 | 19.189 | 67.97 | 106.8171 | 65.8822 | 0.00466 | 1.0413 | 0.025027 | 24.0333 | -0.070076 | 0.14205 | -0.000108 | -0.008096 | -0.008096 | -0.72116
2019-03-20 17:49 | 1553118515.0754 | -25 | 0.20953 | 0.000441 | 0 | 104.3554 | 105.078 | 125.998 | 21.0629 | 0.001935 | 0.001843 | -151.9203 | 0.2379 | 8.4034 | 0.003471 | 0.004547 | 0.004351 | 6.5e-05 | 0.001283 | 0.30431 | 1.3442 | 2.548 | 100.58 | 20.69 | 19.186 | 68.1912 | 105.1816 | 65.9749 | 0.00466 | 1.041 | 0.025052 | 24.0643 | -0.061723 | 0.14153 | -0.000118 | -0.008893 | -0.008893 | -0.65405
2019-03-20 17:49 | 1553118516.075 | -24 | 0.20953 | 0.000442 | 0 | 104.3085 | 104.9019 | 126.09 | 21.0631 | 0.001991 | 0.001826 | -151.9258 | 0.23859 | 8.4015 | 0.003539 | 0.004877 | 0.004687 | 6.2e-05 | 0.001284 | 0.27057 | 1.3236 | 2.5482 | 100.58 | 20.69 | 19.182 | 69.5407 | 106.8326 | 65.5808 | 0.004659 | 1.0423 | 0.025076 | 24.0591 | -0.050746 | 0.14194 | -0.00012 | -0.009059 | -0.009059 | -0.64394
2019-03-20 17:49 | 1553118517.0765 | -23 | 0.20952 | 0.000441 | 0 | 104.2069 | 104.6807 | 126.0273 | 21.0625 | 0.001953 | 0.001844 | -151.9328 | 0.23887 | 8.4035 | 0.003538 | 0.004637 | 0.004401 | 6.4e-05 | 0.00128 | 0.26681 | 1.281 | 2.548 | 100.58 | 20.69 | 19.173 | 68.6321 | 106.8196 | 66.0124 | 0.004656 | 1.0409 | 0.025054 | 24.0684 | -0.040573 | 0.14211 | -2e-06 | -0.000157 | -0.000157 | -37.2296
2019-03-20 17:49 | 1553118518.0755 | -22 | 0.20949 | 0.000438 | 0 | 104.2755 | 104.9984 | 126.0593 | 21.0627 | 0.001967 | 0.00187 | -151.9297 | 0.23816 | 8.4019 | 0.003566 | 0.005064 | 0.004853 | 6.5e-05 | 0.001279 | 0.28694 | 1.3208 | 2.548 | 100.58 | 20.69 | 19.167 | 68.96 | 107.4963 | 66.6346 | 0.004655 | 1.039 | 0.025034 | 24.0935 | -0.061794 | 0.14168 | 0.000119 | 0.008981 | 0.008981 | 0.64834
2019-03-20 17:49 | 1553118519.0753 | -21 | 0.20953 | 0.000437 | 0 | 104.193 | 104.6534 | 125.8546 | 21.0632 | 0.00194 | 0.001892 | -151.9252 | 0.23837 | 8.4038 | 0.003553 | 0.004582 | 0.004424 | 6e-05 | 0.001283 | 0.27201 | 1.3514 | 2.5482 | 100.58 | 20.69 | 19.159 | 68.3093 | 107.1919 | 67.1516 | 0.004653 | 1.0374 | 0.024994 | 24.0923 | -0.039435 | 0.1418 | -0.000116 | -0.008777 | -0.008777 | -0.664
2019-03-20 17:49 | 1553118520.076 | -20 | 0.20951 | 0.000441 | 0 | 104.0856 | 104.8418 | 126.2841 | 21.0637 | 0.001969 | 0.001873 | -151.9201 | 0.23798 | 8.404 | 0.00359 | 0.004761 | 0.004623 | 6.1e-05 | 0.001285 | 0.30021 | 1.3404 | 2.548 | 100.58 | 20.69 | 19.154 | 69.0204 | 108.0855 | 66.6911 | 0.004652 | 1.0389 | 0.025054 | 24.117 | -0.064742 | 0.14158 | 3.1e-05 | 0.002335 | 0.002335 | 2.4915
2019-03-20 17:49 | 1553118521.0756 | -19 | 0.20953 | 0.000441 | 1e-06 | 104.1409 | 104.862 | 126.248 | 21.0634 | 0.001963 | 0.001826 | -151.8924 | 0.23866 | 8.403 | 0.003594 | 0.004397 | 0.004097 | 6.5e-05 | 0.001288 | 0.27865 | 1.3528 | 2.5482 | 100.58 | 20.69 | 19.149 | 68.8679 | 108.1782 | 65.5623 | 0.004651 | 1.0423 | 0.025092 | 24.0735 | -0.061713 | 0.14198 | -8.7e-05 | -0.006619 | -0.006619 | -0.88164
2019-03-20 17:49 | 1553118522.0743 | -18 | 0.20952 | 0.000441 | 0 | 104.1293 | 104.7461 | 126.6705 | 21.0633 | 0.00197 | 0.001881 | -151.8893 | 0.23872 | 8.4019 | 0.003601 | 0.005413 | 0.005292 | 6.1e-05 | 0.001284 | 0.29214 | 1.3294 | 2.5482 | 100.58 | 20.69 | 19.139 | 69.0425 | 108.3445 | 66.9045 | 0.004648 | 1.0382 | 0.025084 | 24.1615 | -0.052825 | 0.14202 | -1.2e-05 | -0.000902 | -0.000902 | -6.4729
2019-03-20 17:49 | 1553118523.0749 | -17 | 0.20954 | 0.000441 | 0 | 104.1429 | 104.6781 | 126.505 | 21.0636 | 0.001969 | 0.001856 | -151.8989 | 0.23896 | 8.4017 | 0.003565 | 0.004699 | 0.004446 | 6.2e-05 | 0.001286 | 0.27848 | 1.3462 | 2.5481 | 100.58 | 20.69 | 19.133 | 69.007 | 107.4765 | 66.2922 | 0.004647 | 1.0401 | 0.025091 | 24.1239 | -0.045847 | 0.14216 | -0.000142 | -0.010712 | -0.010712 | -0.54544
2019-03-20 17:49 | 1553118524.0755 | -16 | 0.2095 | 0.000439 | 0 | 104.0439 | 104.6373 | 126.8226 | 21.0635 | 0.001995 | 0.001808 | -151.9093 | 0.23879 | 8.4026 | 0.003569 | 0.005395 | 0.005143 | 6.3e-05 | 0.00128 | 0.27096 | 1.3468 | 2.5481 | 100.58 | 20.69 | 19.127 | 69.6324 | 107.567 | 65.1386 | 0.004645 | 1.0436 | 0.025165 | 24.1131 | -0.050869 | 0.14205 | 6.9e-05 | 0.005212 | 0.005212 | 1.1201
2019-03-20 17:49 | 1553118525.0755 | -15 | 0.20951 | 0.000439 | 0 | 104.0501 | 104.8175 | 126.7827 | 21.0635 | 0.001961 | 0.001841 | -151.9288 | 0.23823 | 8.4024 | 0.003619 | 0.004559 | 0.00439 | 6.1e-05 | 0.001275 | 0.27649 | 1.3369 | 2.5481 | 100.58 | 20.69 | 19.118 | 68.8079 | 108.7945 | 65.9317 | 0.004643 | 1.0412 | 0.025132 | 24.1376 | -0.065727 | 0.14172 | 5e-06 | 0.000371 | 0.000371 | 15.7085
2019-03-20 17:49 | 1553118526.0745 | -14 | 0.20953 | 0.000442 | 0 | 103.955 | 104.5297 | 126.9148 | 21.064 | 0.0019 | 0.00178 | -151.9164 | 0.23811 | 8.4031 | 0.003603 | 0.004951 | 0.004695 | 6.4e-05 | 0.001275 | 0.30198 | 1.355 | 2.5482 | 100.58 | 20.69 | 19.114 | 67.349 | 108.3985 | 64.4631 | 0.004642 | 1.0457 | 0.025199 | 24.0978 | -0.049313 | 0.14165 | -9.4e-05 | -0.007147 | -0.007147 | -0.81454
2019-03-20 17:49 | 1553118527.0745 | -13 | 0.20953 | 0.000441 | 0 | 103.9304 | 104.5835 | 127.2106 | 21.0644 | 0.0019 | 0.001807 | -151.9457 | 0.23808 | 8.4021 | 0.003716 | 0.005348 | 0.005178 | 6.5e-05 | 0.001261 | 0.28152 | 1.3606 | 2.5481 | 100.58 | 20.69 | 19.111 | 67.3487 | 111.1362 | 65.1271 | 0.004641 | 1.0437 | 0.025204 | 24.1496 | -0.056032 | 0.14163 | -6.9e-05 | -0.005261 | -0.005261 | -1.1064
2019-03-20 17:49 | 1553118528.076 | -12 | 0.20952 | 0.000444 | 0 | 103.8754 | 104.3604 | 127.4338 | 21.0638 | 0.001952 | 0.001816 | -151.9692 | 0.23863 | 8.4026 | 0.003747 | 0.004679 | 0.004525 | 6.1e-05 | 0.001258 | 0.29966 | 1.3591 | 2.5479 | 100.58 | 20.69 | 19.111 | 68.5969 | 111.9006 | 65.3217 | 0.004641 | 1.0431 | 0.025219 | 24.1777 | -0.041662 | 0.14196 | -1.1e-05 | -0.000876 | -0.000876 | -6.6638
2019-03-20 17:49 | 1553118529.0751 | -11 | 0.20954 | 0.000443 | 0 | 103.8078 | 104.4262 | 127.6418 | 21.0644 | 0.001914 | 0.001773 | -129.9287 | 0.23816 | 8.4031 | 0.003693 | 0.004746 | 0.004515 | 6.1e-05 | 0.001256 | 0.27168 | 1.323 | 2.5481 | 100.58 | 20.69 | 19.11 | 67.6849 | 110.5783 | 64.287 | 0.004641 | 1.0463 | 0.025278 | 24.1604 | -0.05312 | 0.14168 | -0.000157 | -0.011961 | -0.011961 | -0.48682
2019-03-20 17:49 | 1553118530.0747 | -10 | 0.20955 | 0.000441 | 0 | 103.9508 | 104.5581 | 127.5418 | 21.0646 | 0.001925 | 0.001788 | -97.1418 | 0.23803 | 8.4039 | 0.00369 | 0.005058 | 0.00483 | 6.2e-05 | 0.00126 | 0.28296 | 1.3002 | 2.548 | 100.58 | 20.69 | 19.111 | 67.944 | 110.5226 | 64.6707 | 0.004641 | 1.0451 | 0.025254 | 24.1647 | -0.052105 | 0.14161 | -0.000214 | -0.016304 | -0.016304 | -0.35696
2019-03-20 17:49 | 1553118531.0747 | -9 | 0.20953 | 0.000444 | 1e-06 | 103.8964 | 104.3495 | 127.7171 | 21.065 | 0.001924 | 0.001828 | -64.3745 | 0.239 | 8.4011 | 0.003791 | 0.005346 | 0.005173 | 6.3e-05 | 0.001256 | 0.28313 | 1.3112 | 2.548 | 100.58 | 20.69 | 19.111 | 67.9413 | 112.9662 | 65.6264 | 0.004641 | 1.0421 | 0.025235 | 24.2154 | -0.038923 | 0.14218 | -0.000106 | -0.008082 | -0.008082 | -0.72308
2019-03-20 17:49 | 1553118532.0748 | -8 | 0.20955 | 0.000442 | 0 | 103.9563 | 104.4755 | 127.7101 | 21.0642 | 0.001891 | 0.001752 | -31.5691 | 0.23887 | 8.4019 | 0.003737 | 0.004906 | 0.004703 | 6e-05 | 0.001258 | 0.28008 | 1.3481 | 2.548 | 100.58 | 20.69 | 19.111 | 67.137 | 111.6443 | 63.7816 | 0.004641 | 1.0478 | 0.025304 | 24.1488 | -0.044563 | 0.1421 | -0.000191 | -0.0146 | -0.0146 | -0.40002
2019-03-20 17:49 | 1553118533.0758 | -7 | 0.20951 | 0.000443 | 1e-06 | 104.0145 | 104.539 | 127.9561 | 21.0641 | 0.001944 | 0.001762 | 1.2359 | 0.23794 | 8.4018 | 0.003714 | 0.005111 | 0.004851 | 6.1e-05 | 0.001264 | 0.29551 | 1.3353 | 2.5482 | 100.58 | 20.69 | 19.111 | 68.4137 | 111.0979 | 64.0423 | 0.004641 | 1.047 | 0.025318 | 24.1813 | -0.044993 | 0.14155 | 1.3e-05 | 0.00098 | 0.00098 | 5.9395
2019-03-20 17:49 | 1553118534.0756 | -6 | 0.20953 | 0.000444 | 0 | 104.0811 | 104.8104 | 127.9375 | 21.064 | 0.001908 | 0.001782 | 34.0403 | 0.25503 | 8.4027 | 0.003773 | 0.005419 | 0.005175 | 6.2e-05 | 0.001259 | 0.27273 | 1.33 | 2.5481 | 100.58 | 20.69 | 19.111 | 67.551 | 112.5213 | 64.5098 | 0.004641 | 1.0456 | 0.025299 | 24.1964 | -0.062453 | 0.15171 | -8.2e-05 | -0.006237 | -0.006237 | -0.99983
2019-03-20 17:49 | 1553118535.0752 | -5 | 0.20951 | 0.000442 | 0 | 104.0588 | 104.5977 | 127.635 | 21.0641 | 0.001893 | 0.001771 | 175.9389 | 0.26419 | 8.3999 | 0.003731 | 0.00531 | 0.005207 | 6.1e-05 | 0.001263 | 0.29319 | 1.3397 | 2.5481 | 100.58 | 20.69 | 19.11 | 67.1844 | 111.5132 | 64.2462 | 0.004641 | 1.0464 | 0.025279 | 24.1583 | -0.046203 | 0.15717 | 4.9e-05 | 0.003764 | 0.003764 | 1.7164
2019-03-20 17:49 | 1553118536.074 | -4 | 0.20954 | 0.000445 | 0 | 104.0779 | 104.6194 | 127.3772 | 21.0644 | 0.001894 | 0.001825 | 175.9532 | 0.26478 | 8.4004 | 0.003723 | 0.006148 | 0.006045 | 6.1e-05 | 0.001264 | 0.2672 | 1.3249 | 2.548 | 100.58 | 20.69 | 19.111 | 67.2104 | 111.3223 | 65.5563 | 0.004641 | 1.0423 | 0.025204 | 24.1807 | -0.046417 | 0.15752 | -0.000139 | -0.010554 | -0.010554 | -0.6134
2019-03-20 17:49 | 1553118537.0741 | -3 | 0.20953 | 0.000444 | 0 | 104.0882 | 104.7607 | 127.0174 | 21.0645 | 0.001831 | 0.001846 | 176.1357 | 0.26394 | 8.4017 | 0.00377 | 0.005459 | 0.005373 | 5.9e-05 | 0.001266 | 0.29927 | 1.3486 | 2.548 | 100.58 | 20.69 | 19.109 | 65.6837 | 112.4596 | 66.0442 | 0.004641 | 1.0408 | 0.025151 | 24.1639 | -0.057598 | 0.15702 | -6.7e-05 | -0.005133 | -0.005133 | -1.2571
2019-03-20 17:49 | 1553118538.0751 | -2 | 0.20952 | 0.000448 | 0 | 104.0488 | 104.5431 | 126.9334 | 21.0649 | 0.001931 | 0.0018 | 176.1208 | 0.26408 | 8.4022 | 0.003782 | 0.005164 | 0.004962 | 6.2e-05 | 0.001261 | 0.27544 | 1.3228 | 2.5481 | 100.58 | 20.69 | 19.108 | 68.1027 | 112.752 | 64.9371 | 0.004641 | 1.0442 | 0.025183 | 24.1165 | -0.042398 | 0.1571 | -6.6e-05 | -0.005026 | -0.005026 | -1.2848
2019-03-20 17:49 | 1553118539.0752 | -1 | 0.20953 | 0.000445 | 0 | 104.0886 | 104.8257 | 127.0033 | 21.0637 | 0.001914 | 0.001833 | 176.0608 | 0.26407 | 8.4032 | 0.00372 | 0.005492 | 0.00538 | 5.8e-05 | 0.001258 | 0.28766 | 1.3108 | 2.5481 | 100.58 | 20.69 | 19.106 | 67.6887 | 111.2428 | 65.7288 | 0.00464 | 1.0418 | 0.025161 | 24.1513 | -0.063121 | 0.1571 | -0.000102 | -0.007708 | -0.007708 | -0.83763
2019-03-20 17:49 | 1553118540.0742 | 0 | 0.20949 | 0.000448 | 0 | 104.0138 | 104.5861 | 127.07 | 21.0607 | 0.001893 | 0.001793 | 176.1884 | 0.26377 | 8.4029 | 0.003644 | 0.004999 | 0.004868 | 6e-05 | 0.001254 | 0.27721 | 1.3479 | 2.7588 | 100.58 | 20.69 | 19.101 | 67.1812 | 109.3877 | 64.7783 | 0.004639 | 1.0447 | 0.025203 | 24.1238 | -0.049077 | 0.15692 | 0.000142 | 0.010806 | 0.010806 | 0.59681
2019-03-20 17:49 | 1553118541.0742 | 1 | 0.20952 | 0.000446 | 0 | 103.9996 | 100.1837 | 126.4942 | 21.0488 | 0.001881 | 0.001782 | 176.323 | 0.26382 | 8.4037 | 0.003701 | 0.006138 | 0.005921 | 6.1e-05 | 0.001251 | 0.31542 | 1.3181 | 2.5481 | 100.58 | 20.69 | 19.098 | 66.8818 | 110.7571 | 64.4898 | 0.004638 | 1.0456 | 0.025157 | 24.0588 | 0.33436 | 0.15695 | -3e-05 | -0.002255 | -0.002255 | -2.8607
2019-03-20 17:49 | 1553118542.0754 | 2 | 0.20952 | 0.000443 | 0 | 102.1232 | 78.5969 | 125.5025 | 21.0495 | 0.001845 | 0.001778 | 176.4485 | 0.26302 | 8.4065 | 0.003552 | 0.004416 | 0.004143 | 6e-05 | 0.001245 | 0.29087 | 1.3153 | 2.5484 | 100.58 | 20.69 | 19.094 | 66.0172 | 107.1469 | 64.4108 | 0.004637 | 1.0459 | 0.025061 | 23.9615 | 2.3421 | 0.15647 | -4.7e-05 | -0.003545 | -0.003545 | -1.814
2019-03-20 17:49 | 1553118543.0755 | 3 | 0.20951 | 0.000445 | 0 | 101.0714 | 70.1079 | 123.996 | 21.0429 | 0.001893 | 0.001812 | 176.6216 | 0.26435 | 8.4029 | 0.003693 | 0.006341 | 0.006725 | 6e-05 | 0.001242 | 0.27776 | 1.3426 | 2.5481 | 100.58 | 20.69 | 19.085 | 67.1677 | 110.5579 | 65.2147 | 0.004635 | 1.0434 | 0.02488 | 23.8456 | 3.2718 | 0.15726 | 2.2e-05 | 0.001685 | 0.001685 | 3.8357
2019-03-20 17:49 | 1553118544.0754 | 4 | 0.20951 | 0.000444 | 1e-06 | 102.2645 | 78.7752 | 122.16 | 21.0488 | 0.001902 | 0.001865 | 176.4439 | 0.26422 | 8.4045 | 0.00376 | 0.005043 | 0.004782 | 5.9e-05 | 0.001245 | 1.5849 | 1.2962 | 2.5479 | 100.58 | 20.69 | 19.078 | 67.3834 | 112.2007 | 66.4855 | 0.004633 | 1.0395 | 0.024649 | 23.7128 | 2.3342 | 0.15718 | 4e-05 | 0.00292 | 0.00292 | 2.2126
2019-03-20 17:49 | 1553118545.0755 | 5 | 0.2095 | 0.000443 | 3e-06 | 103.1177 | 87.9699 | 120.5465 | 21.0579 | 0.001966 | 0.001859 | 176.48 | 0.26431 | 8.4035 | 0.003727 | 0.017146 | 0.019538 | 6.6e-05 | 0.001221 | 0.2995 | 1.3406 | 2.5479 | 100.58 | 20.68 | 19.063 | 68.9445 | 111.4108 | 66.3608 | 0.004627 | 1.0399 | 0.02449 | 23.5513 | 1.4211 | 0.15724 | 6.2e-05 | 0.00441 | 0.00441 | 1.4656
2019-03-20 17:49 | 1553118546.0756 | 6 | 0.20935 | 0.00045 | 8e-06 | 103.7797 | 94.792 | 118.645 | 21.0533 | 0.002 | 0.002007 | 176.2724 | 0.26448 | 8.4007 | 0.003613 | 0.005211 | 0.00499 | 5.8e-05 | 0.001233 | 0.30431 | 1.3162 | 2.548 | 100.58 | 20.68 | 19.057 | 69.7573 | 108.6215 | 69.9171 | 0.004625 | 1.0291 | 0.02417 | 23.4869 | 0.81025 | 0.15734 | 0.000957 | 0.069329 | 0.069329 | 0.093273
2019-03-20 17:49 | 1553118547.0747 | 7 | 0.20895 | 0.000457 | 1.4e-05 | 104.0876 | 97.8803 | 116.7365 | 21.0553 | 0.002041 | 0.002002 | 176.2183 | 0.26344 | 8.4023 | 0.003758 | 0.005571 | 0.005409 | 5.9e-05 | 0.001224 | 0.28136 | 1.329 | 1.8127 | 100.58 | 20.68 | 19.05 | 70.7215 | 112.1494 | 69.8044 | 0.004624 | 1.0294 | 0.023979 | 23.2934 | 0.54998 | 0.15672 | 0.003391 | 0.24444 | 0.24444 | 0.026351
2019-03-20 17:49 | 1553118548.0761 | 8 | 0.20771 | 0.000505 | 2.4e-05 | 104.0067 | 99.8271 | 114.4874 | 21.0624 | 0.002098 | 0.002104 | 176.126 | 0.26322 | 8.4032 | 0.00373 | 0.005779 | 0.005602 | 5.6e-05 | 0.00121 | 0.26521 | 1.3164 | 2.5481 | 100.58 | 20.68 | 19.033 | 72.106 | 111.4918 | 72.2494 | 0.00462 | 1.0221 | 0.023662 | 23.15 | 0.36686 | 0.15659 | 0.010781 | 0.76752 | 0.76752 | 0.008385
2019-03-20 17:49 | 1553118549.0757 | 9 | 0.20636 | 0.000636 | 3.6e-05 | 104.0763 | 97.4841 | 112.2462 | 21.0636 | 0.002161 | 0.002165 | 176.0352 | 0.26372 | 8.4011 | 0.00377 | 0.005435 | 0.005314 | 5.8e-05 | 0.001209 | 0.24707 | 1.3165 | 2.5481 | 100.58 | 20.68 | 19.026 | 73.6302 | 112.4625 | 73.716 | 0.004618 | 1.0178 | 0.02338 | 22.9709 | 0.58528 | 0.15689 | 0.01869 | 1.314 | 1.314 | 0.004907
2019-03-20 17:49 | 1553118550.0738 | 10 | 0.20562 | 0.001171 | 4.7e-05 | 104.1642 | 92.2792 | 110.2321 | 21.0643 | 0.002226 | 0.002228 | 175.95 | 0.26393 | 8.4017 | 0.003811 | 0.005725 | 0.005582 | 5.6e-05 | 0.001204 | 0.2677 | 1.3382 | 2.548 | 100.58 | 20.68 | 19.01 | 75.1985 | 113.4608 | 75.2453 | 0.004614 | 1.0133 | 0.023119 | 22.814 | 1.0836 | 0.15701 | 0.02243 | 1.5584 | 1.5584 | 0.004141
2019-03-20 17:49 | 1553118551.0738 | 11 | 0.20417 | 0.001884 | 6.1e-05 | 104.1193 | 88.1526 | 108.87 | 21.0637 | 0.0023 | 0.00239 | 175.8379 | 0.26391 | 8.4054 | 0.00381 | 0.005368 | 0.005284 | 5.6e-05 | 0.001205 | 0.27903 | 1.3334 | 2.5482 | 100.58 | 20.68 | 19.002 | 76.9713 | 113.4398 | 79.1323 | 0.004612 | 1.0022 | 0.022848 | 22.7987 | 1.489 | 0.157 | 0.030214 | 2.0732 | 2.0732 | 0.003112
2019-03-20 17:49 | 1553118552.0749 | 12 | 0.20327 | 0.002708 | 7.2e-05 | 104.2247 | 85.2856 | 107.667 | 21.0629 | 0.002379 | 0.002398 | 175.756 | 0.26401 | 8.4053 | 0.003853 | 0.005798 | 0.00573 | 5.8e-05 | 0.001203 | 0.25857 | 1.3642 | 2.5481 | 100.58 | 20.68 | 18.996 | 78.8648 | 114.4848 | 79.3239 | 0.004611 | 1.0016 | 0.022715 | 22.6786 | 1.7938 | 0.15706 | 0.034564 | 2.3567 | 2.3567 | 0.002739
2019-03-20 17:49 | 1553118553.0749 | 13 | 0.20208 | 0.003617 | 8.3e-05 | 104.215 | 82.5527 | 107.1901 | 21.0635 | 0.002464 | 0.002431 | 175.6339 | 0.2639 | 8.4056 | 0.003889 | 0.005447 | 0.005206 | 6.3e-05 | 0.001198 | 0.23186 | 1.354 | 2.5481 | 100.58 | 20.68 | 18.985 | 80.9035 | 115.345 | 80.1089 | 0.004608 | 0.9994 | 0.02264 | 22.6535 | 2.0842 | 0.15699 | 0.040589 | 2.7566 | 2.7566 | 0.002341
2019-03-20 17:49 | 1553118554.075 | 14 | 0.20093 | 0.004561 | 9.3e-05 | 104.1568 | 79.895 | 107.1922 | 21.0638 | 0.002511 | 0.002445 | 175.5273 | 0.26426 | 8.4051 | 0.003948 | 0.005357 | 0.005192 | 5.8e-05 | 0.001202 | 0.26969 | 1.3261 | 2.5481 | 100.58 | 20.68 | 18.982 | 82.0368 | 116.7969 | 80.4587 | 0.004607 | 0.99841 | 0.022629 | 22.6649 | 2.372 | 0.15721 | 0.046242 | 3.1373 | 3.1373 | 0.00206
2019-03-20 17:49 | 1553118555.075 | 15 | 0.20028 | 0.00541 | 0.0001 | 104.1371 | 77.6424 | 107.1779 | 21.064 | 0.002566 | 0.002559 | 175.4307 | 0.26415 | 8.4056 | 0.003966 | 0.005665 | 0.005675 | 5.9e-05 | 0.001201 | 0.26985 | 1.3273 | 2.548 | 100.58 | 20.68 | 18.978 | 83.3544 | 117.2348 | 83.1969 | 0.004606 | 0.99074 | 0.02254 | 22.751 | 2.6261 | 0.15714 | 0.049096 | 3.3168 | 3.3168 | 0.001947
2019-03-20 17:49 | 1553118556.0743 | 16 | 0.20008 | 0.005972 | 0.000106 | 104.0795 | 75.9956 | 107.7663 | 21.063 | 0.002654 | 0.002533 | 175.3279 | 0.26376 | 8.4028 | 0.004053 | 0.005188 | 0.005012 | 6e-05 | 0.001199 | 0.25841 | 1.2904 | 2.5481 | 100.58 | 20.68 | 18.979 | 85.4821 | 119.3703 | 82.5524 | 0.004606 | 0.99253 | 0.022622 | 22.7927 | 2.8129 | 0.15691 | 0.049623 | 3.3641 | 3.3641 | 0.001917
2019-03-20 17:49 | 1553118557.0755 | 17 | 0.20003 | 0.006354 | 0.000111 | 104.1206 | 76.1847 | 108.543 | 21.0648 | 0.002606 | 0.002584 | 175.2011 | 0.26291 | 8.4022 | 0.003999 | 0.005092 | 0.00494 | 6.2e-05 | 0.001195 | 0.23811 | 1.2993 | 2.5482 | 100.58 | 20.68 | 18.98 | 84.3212 | 118.0421 | 83.789 | 0.004607 | 0.98909 | 0.022664 | 22.9144 | 2.7942 | 0.1564 | 0.049445 | 3.3581 | 3.3581 | 0.001914
2019-03-20 17:49 | 1553118558.0747 | 18 | 0.20007 | 0.006623 | 0.000112 | 104.0242 | 75.1371 | 109.1448 | 21.0647 | 0.002708 | 0.002611 | 175.094 | 0.26407 | 8.4011 | 0.004105 | 0.004973 | 0.004823 | 6.4e-05 | 0.001193 | 0.27649 | 1.3078 | 2.548 | 100.58 | 20.68 | 18.987 | 86.7718 | 120.6428 | 84.4427 | 0.004608 | 0.98728 | 0.022706 | 22.9989 | 2.9097 | 0.15709 | 0.048859 | 3.3245 | 3.3245 | 0.001942
2019-03-20 17:49 | 1553118559.0753 | 19 | 0.20015 | 0.006821 | 0.000112 | 104.01 | 74.1904 | 109.9882 | 21.0648 | 0.002678 | 0.002549 | 174.99 | 0.26456 | 8.4022 | 0.004039 | 0.005742 | 0.005614 | 6.3e-05 | 0.001193 | 0.25802 | 1.299 | 2.5479 | 100.58 | 20.68 | 18.993 | 86.0402 | 119.0121 | 82.9496 | 0.00461 | 0.99142 | 0.022842 | 23.0393 | 3.0219 | 0.15739 | 0.048188 | 3.2985 | 3.2985 | 0.001961
2019-03-20 17:49 | 1553118560.0743 | 20 | 0.20067 | 0.00696 | 0.00011 | 103.9663 | 73.9056 | 110.7525 | 21.0648 | 0.002807 | 0.00266 | 174.876 | 0.26366 | 8.4047 | 0.004067 | 0.005862 | 0.005803 | 6.2e-05 | 0.001195 | 0.28816 | 1.3329 | 2.548 | 100.58 | 20.68 | 18.999 | 89.1616 | 119.7155 | 85.6188 | 0.004611 | 0.98405 | 0.022836 | 23.2057 | 3.0526 | 0.15685 | 0.044872 | 3.0714 | 3.0714 | 0.002099
2019-03-20 17:49 | 1553118561.0744 | 21 | 0.20108 | 0.00705 | 0.000107 | 103.9194 | 74.7936 | 111.7721 | 21.0651 | 0.002763 | 0.002619 | 174.7946 | 0.26383 | 8.4014 | 0.004021 | 0.005766 | 0.005652 | 6.1e-05 | 0.001191 | 0.2636 | 1.3441 | 2.5481 | 100.58 | 20.68 | 19.012 | 88.0925 | 118.5804 | 84.6389 | 0.004614 | 0.98674 | 0.022972 | 23.2805 | 2.9417 | 0.15695 | 0.042332 | 2.9154 | 2.9154 | 0.002213
2019-03-20 17:49 | 1553118562.0745 | 22 | 0.20177 | 0.007072 | 0.000103 | 103.8124 | 76.1847 | 112.6664 | 21.0651 | 0.002793 | 0.002628 | 174.6907 | 0.26353 | 8.4031 | 0.004052 | 0.005448 | 0.005338 | 6.4e-05 | 0.001195 | 0.24674 | 1.3971 | 2.548 | 100.58 | 20.68 | 19.019 | 88.808 | 119.3367 | 84.8402 | 0.004616 | 0.98619 | 0.023057 | 23.38 | 2.7677 | 0.15677 | 0.038223 | 2.6429 | 2.6429 | 0.002438
2019-03-20 17:49 | 1553118563.0745 | 23 | 0.20243 | 0.007031 | 9.8e-05 | 103.747 | 79.088 | 113.215 | 21.0642 | 0.00278 | 0.002684 | 174.6101 | 0.26366 | 8.4043 | 0.004137 | 0.005814 | 0.005667 | 5.9e-05 | 0.001198 | 0.27041 | 1.4115 | 2.5481 | 100.58 | 20.68 | 19.032 | 88.5071 | 121.4175 | 86.1895 | 0.004619 | 0.98248 | 0.02307 | 23.4809 | 2.4275 | 0.15685 | 0.034308 | 2.3742 | 2.3742 | 0.002715
2019-03-20 17:49 | 1553118564.0756 | 24 | 0.20285 | 0.006919 | 9.5e-05 | 103.7558 | 81.7764 | 113.6034 | 21.0647 | 0.0028 | 0.002675 | 174.5535 | 0.26368 | 8.4045 | 0.004193 | 0.005166 | 0.005061 | 6.3e-05 | 0.001193 | 0.25011 | 1.4107 | 2.548 | 100.58 | 20.68 | 19.038 | 88.999 | 122.7952 | 85.9865 | 0.004621 | 0.98304 | 0.023116 | 23.5146 | 2.1293 | 0.15686 | 0.031943 | 2.2153 | 2.2153 | 0.00291
2019-03-20 17:49 | 1553118565.0746 | 25 | 0.20314 | 0.006764 | 9e-05 | 103.7292 | 83.8388 | 114.0311 | 21.0642 | 0.002852 | 0.002651 | 174.4605 | 0.26375 | 8.4029 | 0.004123 | 0.005712 | 0.005545 | 6.4e-05 | 0.001189 | 0.24314 | 1.3683 | 2.5481 | 100.58 | 20.68 | 19.045 | 90.2291 | 121.0794 | 85.3997 | 0.004622 | 0.98465 | 0.023178 | 23.5395 | 1.9042 | 0.1569 | 0.030418 | 2.1156 | 2.1156 | 0.003048
2019-03-20 17:49 | 1553118566.0736 | 26 | 0.20336 | 0.006564 | 8.6e-05 | 103.6325 | 83.8653 | 114.3644 | 21.0642 | 0.002845 | 0.00272 | 174.3937 | 0.26416 | 8.4033 | 0.004115 | 0.005592 | 0.005526 | 6.4e-05 | 0.001182 | 0.2625 | 1.3658 | 2.5481 | 100.58 | 20.68 | 19.059 | 90.071 | 120.8827 | 87.0663 | 0.004626 | 0.98009 | 0.023158 | 23.6286 | 1.893 | 0.15715 | 0.029323 | 2.038 | 2.038 | 0.003169
2019-03-20 17:49 | 1553118567.0741 | 27 | 0.20358 | 0.006347 | 8.3e-05 | 103.5045 | 83.7181 | 114.7043 | 21.0643 | 0.002834 | 0.002724 | 174.3137 | 0.26452 | 8.4007 | 0.004148 | 0.005436 | 0.005335 | 6.1e-05 | 0.001181 | 0.28296 | 1.3545 | 2.5481 | 100.58 | 20.68 | 19.066 | 89.8055 | 121.6931 | 87.1569 | 0.004628 | 0.97985 | 0.02319 | 23.6667 | 1.8977 | 0.15736 | 0.028253 | 1.9665 | 1.9665 | 0.003289
2019-03-20 17:49 | 1553118568.0748 | 28 | 0.20378 | 0.006116 | 7.9e-05 | 103.5669 | 84.5407 | 115.0244 | 21.0652 | 0.002909 | 0.002696 | 174.2366 | 0.26398 | 8.4025 | 0.004056 | 0.005688 | 0.00551 | 6.3e-05 | 0.001184 | 0.27992 | 1.3258 | 2.5481 | 100.58 | 20.68 | 19.078 | 91.6096 | 119.4247 | 86.4927 | 0.00463 | 0.98166 | 0.023243 | 23.6778 | 1.8156 | 0.15704 | 0.027336 | 1.9073 | 1.9073 | 0.003384
2019-03-20 17:49 | 1553118569.0738 | 29 | 0.20403 | 0.005887 | 7.6e-05 | 103.5772 | 85.9613 | 114.8266 | 21.0653 | 0.002877 | 0.002756 | 174.1592 | 0.26407 | 8.4036 | 0.004091 | 0.006274 | 0.00614 | 6.4e-05 | 0.001191 | 0.26681 | 1.3463 | 2.5481 | 100.58 | 20.68 | 19.083 | 90.8351 | 120.2917 | 87.9382 | 0.004632 | 0.97773 | 0.023177 | 23.705 | 1.6674 | 0.15709 | 0.026153 | 1.8197 | 1.8197 | 0.003548
2019-03-20 17:49 | 1553118570.0752 | 30 | 0.2042 | 0.005668 | 7.3e-05 | 103.5556 | 86.6132 | 114.6958 | 21.0648 | 0.002888 | 0.002667 | 174.0972 | 0.26363 | 8.4014 | 0.004097 | 0.005175 | 0.005056 | 6.4e-05 | 0.001202 | 0.27975 | 1.3322 | 2.5481 | 100.58 | 20.68 | 19.088 | 91.0965 | 120.4443 | 85.7874 | 0.004633 | 0.98359 | 0.023233 | 23.6208 | 1.598 | 0.15684 | 0.025394 | 1.7714 | 1.7714 | 0.003639
2019-03-20 17:50 | 1553118571.1309 | 31 | 0.2043 | 0.005471 | 7.1e-05 | 103.3246 | 87.8017 | 114.3975 | 21.0639 | 0.002896 | 0.00274 | 174.0245 | 0.26456 | 8.4035 | 0.004057 | 0.005382 | 0.00527 | 6.2e-05 | 0.001199 | 0.26465 | 1.3251 | 2.5482 | 100.58 | 20.68 | 19.098 | 91.3059 | 119.4513 | 87.5377 | 0.004635 | 0.97881 | 0.023146 | 23.6475 | 1.4561 | 0.15739 | 0.025063 | 1.7419 | 1.7419 | 0.003714
2019-03-20 17:50 | 1553118572.0796 | 32 | 0.20431 | 0.005288 | 6.9e-05 | 103.3298 | 88.5271 | 114.9517 | 21.0651 | 0.002882 | 0.002651 | 173.9578 | 0.26406 | 8.4036 | 0.004034 | 0.005046 | 0.004882 | 6e-05 | 0.001201 | 0.28385 | 1.3117 | 2.5481 | 100.58 | 20.68 | 19.103 | 90.9717 | 118.8887 | 85.3912 | 0.004637 | 0.98467 | 0.023272 | 23.6341 | 1.383 | 0.15709 | 0.025188 | 1.7601 | 1.7601 | 0.003668
2019-03-20 17:50 | 1553118573.0797 | 33 | 0.20444 | 0.005128 | 6.6e-05 | 103.2871 | 87.8432 | 115.561 | 21.0643 | 0.002923 | 0.002634 | 173.8895 | 0.26397 | 8.4024 | 0.00403 | 0.005237 | 0.005101 | 6e-05 | 0.001204 | 0.27328 | 1.3304 | 2.5481 | 100.58 | 20.68 | 19.108 | 91.956 | 118.8008 | 85.0018 | 0.004638 | 0.98574 | 0.023346 | 23.6838 | 1.4487 | 0.15704 | 0.024659 | 1.7289 | 1.7289 | 0.003733
2019-03-20 17:50 | 1553118574.0808 | 34 | 0.20447 | 0.004997 | 6.4e-05 | 103.2173 | 87.8379 | 115.9962 | 21.0639 | 0.002906 | 0.002742 | 173.8016 | 0.26406 | 8.4017 | 0.004 | 0.005759 | 0.005633 | 5.9e-05 | 0.001219 | 0.26626 | 1.3397 | 2.548 | 100.58 | 20.68 | 19.117 | 91.5365 | 118.061 | 87.5875 | 0.00464 | 0.97868 | 0.023306 | 23.8138 | 1.4431 | 0.15709 | 0.024625 | 1.7236 | 1.7236 | 0.003746
2019-03-20 17:50 | 1553118575.0798 | 35 | 0.20457 | 0.004869 | 6.1e-05 | 103.2005 | 87.7677 | 116.2842 | 21.0642 | 0.002925 | 0.002687 | 173.7478 | 0.26339 | 8.402 | 0.004031 | 0.005891 | 0.005725 | 6.2e-05 | 0.001221 | 0.28766 | 1.3471 | 2.5482 | 100.58 | 20.68 | 19.121 | 91.9942 | 118.8282 | 86.2723 | 0.004641 | 0.98226 | 0.023378 | 23.7998 | 1.4488 | 0.15669 | 0.024176 | 1.6976 | 1.6976 | 0.003794
2019-03-20 17:50 | 1553118576.0809 | 36 | 0.2047 | 0.004766 | 5.8e-05 | 103.1587 | 87.8383 | 116.6745 | 21.0643 | 0.003001 | 0.002733 | 173.6736 | 0.26348 | 8.403 | 0.003978 | 0.005943 | 0.005808 | 6.2e-05 | 0.001221 | 0.26521 | 1.3361 | 2.548 | 100.58 | 20.68 | 19.128 | 93.8223 | 117.5218 | 87.3639 | 0.004643 | 0.97928 | 0.023381 | 23.8759 | 1.438 | 0.15675 | 0.023498 | 1.6504 | 1.6504 | 0.003903
2019-03-20 17:50 | 1553118577.0813 | 37 | 0.20497 | 0.004658 | 5.4e-05 | 103.0252 | 88.1104 | 116.9257 | 21.0645 | 0.002928 | 0.002719 | 173.629 | 0.26349 | 8.4029 | 0.004063 | 0.005553 | 0.00544 | 6.3e-05 | 0.001214 | 0.25371 | 1.3255 | 2.548 | 100.58 | 20.68 | 19.131 | 92.0632 | 119.5961 | 87.0377 | 0.004643 | 0.98017 | 0.023417 | 23.8908 | 1.3988 | 0.15675 | 0.022058 | 1.5519 | 1.5519 | 0.004151
2019-03-20 17:50 | 1553118578.081 | 38 | 0.20524 | 0.004551 | 5.1e-05 | 102.9815 | 88.7764 | 117.5454 | 21.0654 | 0.002935 | 0.002729 | 173.5965 | 0.26403 | 8.4002 | 0.004026 | 0.006129 | 0.006032 | 6e-05 | 0.001211 | 0.24818 | 1.3453 | 2.548 | 100.58 | 20.68 | 19.136 | 92.2316 | 118.6987 | 87.2766 | 0.004645 | 0.97952 | 0.023471 | 23.962 | 1.3276 | 0.15707 | 0.020578 | 1.4513 | 1.4513 | 0.004448
2019-03-20 17:50 | 1553118579.0801 | 39 | 0.20557 | 0.004435 | 4.8e-05 | 102.9317 | 89.9691 | 118.3047 | 21.0644 | 0.002926 | 0.002716 | 173.5575 | 0.2634 | 8.4002 | 0.004038 | 0.005766 | 0.005731 | 5.9e-05 | 0.001214 | 0.2521 | 1.3549 | 2.548 | 100.58 | 20.68 | 19.138 | 92.0342 | 118.9824 | 86.9674 | 0.004645 | 0.98036 | 0.023557 | 24.0289 | 1.2039 | 0.1567 | 0.018698 | 1.3237 | 1.3237 | 0.004865
2019-03-20 17:50 | 1553118580.0815 | 40 | 0.20573 | 0.004304 | 4.5e-05 | 102.845 | 91.571 | 118.5648 | 21.0647 | 0.002951 | 0.002686 | 173.4963 | 0.26433 | 8.4036 | 0.003996 | 0.005426 | 0.005346 | 6.1e-05 | 0.001212 | 0.26952 | 1.346 | 2.548 | 100.58 | 20.68 | 19.14 | 92.6188 | 117.9697 | 86.2408 | 0.004645 | 0.98234 | 0.023607 | 24.031 | 1.0385 | 0.15725 | 0.017947 | 1.2734 | 1.2734 | 0.005075
2019-03-20 17:50 | 1553118581.0802 | 41 | 0.20583 | 0.00417 | 4.2e-05 | 102.8424 | 92.2643 | 118.7917 | 21.0637 | 0.002949 | 0.00264 | 173.4409 | 0.26347 | 8.4039 | 0.003911 | 0.006063 | 0.005887 | 5.9e-05 | 0.001214 | 0.28495 | 1.3828 | 2.548 | 100.58 | 20.68 | 19.14 | 92.5881 | 115.9021 | 85.1363 | 0.004645 | 0.98537 | 0.023666 | 24.017 | 0.97085 | 0.15674 | 0.017524 | 1.2466 | 1.2466 | 0.005168
2019-03-20 17:50 | 1553118582.0792 | 42 | 0.20604 | 0.00404 | 3.9e-05 | 102.7674 | 92.4796 | 118.6192 | 21.0638 | 0.002937 | 0.002667 | 173.4002 | 0.26262 | 8.4055 | 0.00396 | 0.005358 | 0.005196 | 5.9e-05 | 0.001213 | 0.26969 | 1.33 | 2.5481 | 100.58 | 20.68 | 19.139 | 92.2879 | 117.0985 | 85.7903 | 0.004645 | 0.98358 | 0.023627 | 24.0215 | 0.94347 | 0.15623 | 0.016415 | 1.166 | 1.166 | 0.005507
2019-03-20 17:50 | 1553118583.0805 | 43 | 0.20617 | 0.003901 | 3.8e-05 | 102.7396 | 93.0469 | 118.4929 | 21.0645 | 0.002923 | 0.002648 | 173.3788 | 0.2627 | 8.4017 | 0.004013 | 0.005558 | 0.005491 | 5.9e-05 | 0.001219 | 0.28816 | 1.338 | 2.5481 | 100.58 | 20.68 | 19.136 | 91.9545 | 118.3744 | 85.3235 | 0.004645 | 0.98486 | 0.02363 | 23.9931 | 0.88635 | 0.15628 | 0.015787 | 1.1216 | 1.1216 | 0.005727
2019-03-20 17:50 | 1553118584.0793 | 44 | 0.20625 | 0.003774 | 3.5e-05 | 102.5387 | 93.0133 | 118.6941 | 21.0649 | 0.002903 | 0.002625 | 173.3287 | 0.2628 | 8.4034 | 0.003953 | 0.005608 | 0.00542 | 5.8e-05 | 0.001214 | 0.30342 | 1.3606 | 2.5482 | 100.58 | 20.68 | 19.132 | 91.4664 | 116.9094 | 84.7815 | 0.004644 | 0.98635 | 0.023668 | 23.9953 | 0.87208 | 0.15634 | 0.015483 | 1.1019 | 1.1019 | 0.005831
2019-03-20 17:50 | 1553118585.0793 | 45 | 0.20636 | 0.003657 | 3.4e-05 | 102.442 | 93.4371 | 119.1851 | 21.065 | 0.002912 | 0.002701 | 173.2708 | 0.26338 | 8.4031 | 0.004011 | 0.00605 | 0.005943 | 5.9e-05 | 0.001216 | 0.27815 | 1.3329 | 2.5479 | 100.58 | 20.68 | 19.121 | 91.6814 | 118.3361 | 86.6087 | 0.004641 | 0.98134 | 0.023656 | 24.1062 | 0.82298 | 0.15669 | 0.014929 | 1.0621 | 1.0621 | 0.006064
2019-03-20 17:50 | 1553118586.0808 | 46 | 0.20647 | 0.003541 | 3.2e-05 | 102.379 | 93.9686 | 119.2051 | 21.0647 | 0.002966 | 0.002599 | 173.2541 | 0.26381 | 8.4019 | 0.003972 | 0.006112 | 0.006045 | 5.9e-05 | 0.001215 | 0.29214 | 1.3246 | 2.5481 | 100.58 | 20.68 | 19.116 | 92.9973 | 117.3733 | 84.1424 | 0.00464 | 0.98811 | 0.02374 | 24.0254 | 0.76673 | 0.15694 | 0.014447 | 1.0315 | 1.0315 | 0.006254
2019-03-20 17:50 | 1553118587.0804 | 47 | 0.20654 | 0.003437 | 3.1e-05 | 102.2406 | 93.9551 | 119.1741 | 21.0646 | 0.002918 | 0.002587 | 173.2076 | 0.26342 | 8.4036 | 0.004022 | 0.005293 | 0.005165 | 5.9e-05 | 0.001216 | 0.27417 | 1.3372 | 2.5481 | 100.58 | 20.68 | 19.108 | 91.8187 | 118.6093 | 83.8739 | 0.004638 | 0.98886 | 0.023746 | 24.0132 | 0.75591 | 0.15671 | 0.014155 | 1.0109 | 1.0109 | 0.006371
2019-03-20 17:50 | 1553118588.0794 | 48 | 0.20659 | 0.003341 | 3e-05 | 102.2008 | 94.1895 | 118.8647 | 21.0656 | 0.002921 | 0.002636 | 173.1328 | 0.2632 | 8.4043 | 0.004097 | 0.005626 | 0.005553 | 5.7e-05 | 0.001212 | 0.2573 | 1.3423 | 2.548 | 100.58 | 20.68 | 19.094 | 91.9078 | 120.4419 | 85.0342 | 0.004634 | 0.98565 | 0.023676 | 24.021 | 0.73015 | 0.15658 | 0.013961 | 0.99422 | 0.99422 | 0.006473
2019-03-20 17:50 | 1553118589.08 | 49 | 0.2066 | 0.003259 | 2.9e-05 | 102.151 | 94.4175 | 118.7572 | 21.0645 | 0.002833 | 0.002631 | 173.0787 | 0.26405 | 8.4042 | 0.004052 | 0.004918 | 0.004759 | 5.7e-05 | 0.001215 | 0.24994 | 1.373 | 2.548 | 100.58 | 20.68 | 19.087 | 89.791 | 119.3375 | 84.9264 | 0.004633 | 0.98595 | 0.023669 | 24.0065 | 0.70417 | 0.15709 | 0.014007 | 0.99726 | 0.99726 | 0.006474
2019-03-20 17:50 | 1553118590.0807 | 50 | 0.20659 | 0.003178 | 2.8e-05 | 102.0921 | 94.2915 | 119.0513 | 21.0646 | 0.002924 | 0.002643 | 173.04 | 0.26362 | 8.4035 | 0.004051 | 0.005235 | 0.005183 | 5.8e-05 | 0.001209 | 0.26576 | 1.3696 | 2.5482 | 100.58 | 20.68 | 19.075 | 91.9814 | 119.3083 | 85.22 | 0.00463 | 0.98514 | 0.023689 | 24.0461 | 0.71095 | 0.15683 | 0.014151 | 1.0084 | 1.0084 | 0.006392
2019-03-20 17:50 | 1553118591.0807 | 51 | 0.20667 | 0.00311 | 2.6e-05 | 102.0064 | 94.4937 | 119.1955 | 21.0655 | 0.002917 | 0.002555 | 173.0075 | 0.26333 | 8.4029 | 0.003997 | 0.005351 | 0.005248 | 6e-05 | 0.001209 | 0.23883 | 1.3341 | 2.5479 | 100.58 | 20.68 | 19.069 | 91.8039 | 118.0023 | 83.0914 | 0.004628 | 0.99103 | 0.023774 | 23.9891 | 0.68428 | 0.15666 | 0.013771 | 0.98494 | 0.98494 | 0.006537
2019-03-20 17:50 | 1553118592.0796 | 52 | 0.20667 | 0.003055 | 2.6e-05 | 101.8567 | 94.1375 | 118.9838 | 21.065 | 0.002907 | 0.002601 | 172.9623 | 0.26362 | 8.4038 | 0.004021 | 0.005973 | 0.005821 | 5.6e-05 | 0.001209 | 0.27079 | 1.3445 | 2.5479 | 100.58 | 20.67 | 19.065 | 91.5554 | 118.5794 | 84.1972 | 0.004624 | 0.98796 | 0.023716 | 24.0049 | 0.70492 | 0.15683 | 0.013828 | 0.98661 | 0.98661 | 0.006533
2019-03-20 17:50 | 1553118593.0798 | 53 | 0.20669 | 0.002998 | 2.5e-05 | 101.8609 | 94.2627 | 119.1682 | 21.0652 | 0.002858 | 0.00259 | 172.8861 | 0.26405 | 8.4033 | 0.004121 | 0.006249 | 0.006184 | 5.6e-05 | 0.001205 | 0.26217 | 1.348 | 2.548 | 100.58 | 20.67 | 19.055 | 90.3801 | 121.0337 | 83.9409 | 0.004622 | 0.98867 | 0.023743 | 24.0149 | 0.6934 | 0.15708 | 0.013799 | 0.98571 | 0.98571 | 0.00655
2019-03-20 17:50 | 1553118594.0812 | 54 | 0.20673 | 0.002958 | 2.5e-05 | 101.7787 | 94.5259 | 119.1802 | 21.0647 | 0.002871 | 0.002657 | 172.8451 | 0.26323 | 8.4024 | 0.004115 | 0.005803 | 0.005695 | 6e-05 | 0.001207 | 0.26034 | 1.3743 | 2.5481 | 100.58 | 20.67 | 19.052 | 90.7063 | 120.8769 | 85.5494 | 0.004621 | 0.98424 | 0.023691 | 24.0701 | 0.66125 | 0.1566 | 0.013606 | 0.96988 | 0.96988 | 0.006636
2019-03-20 17:50 | 1553118595.081 | 55 | 0.20678 | 0.002906 | 2.4e-05 | 101.7564 | 94.9501 | 119.1513 | 21.0648 | 0.002868 | 0.002606 | 172.8201 | 0.2624 | 8.4059 | 0.004153 | 0.005358 | 0.005238 | 5.9e-05 | 0.001203 | 0.25442 | 1.3436 | 2.5481 | 100.58 | 20.67 | 19.051 | 90.6165 | 121.8018 | 84.3298 | 0.004621 | 0.9876 | 0.023728 | 24.0263 | 0.61924 | 0.1561 | 0.013368 | 0.95442 | 0.95442 | 0.006722
2019-03-20 17:50 | 1553118596.0809 | 56 | 0.20678 | 0.002874 | 2.4e-05 | 101.6959 | 95.085 | 119.1492 | 21.0652 | 0.002887 | 0.002644 | 172.7469 | 0.26406 | 8.4018 | 0.004052 | 0.005307 | 0.005219 | 6e-05 | 0.0012 | 0.27975 | 1.31 | 2.5481 | 100.58 | 20.67 | 19.05 | 91.0843 | 119.336 | 85.2306 | 0.004621 | 0.98511 | 0.023698 | 24.0563 | 0.60122 | 0.15709 | 0.013387 | 0.95457 | 0.95457 | 0.006764
2019-03-20 17:50 | 1553118597.079 | 57 | 0.2068 | 0.002839 | 2.4e-05 | 101.5973 | 95.2356 | 119.2539 | 21.0659 | 0.002966 | 0.002586 | 172.6971 | 0.26368 | 8.4022 | 0.004009 | 0.005916 | 0.005787 | 6e-05 | 0.001191 | 0.27721 | 1.3813 | 2.548 | 100.58 | 20.67 | 19.052 | 92.9859 | 118.2981 | 83.8362 | 0.004621 | 0.98896 | 0.023755 | 24.02 | 0.57838 | 0.15687 | 0.013361 | 0.95502 | 0.95502 | 0.006751
2019-03-20 17:50 | 1553118598.0796 | 58 | 0.20681 | 0.002808 | 2.3e-05 | 101.4605 | 95.2435 | 119.525 | 21.0649 | 0.002904 | 0.002581 | 172.6682 | 0.26379 | 8.4025 | 0.004036 | 0.005891 | 0.00581 | 5.8e-05 | 0.001193 | 0.26609 | 1.3427 | 2.5482 | 100.58 | 20.67 | 19.056 | 91.4999 | 118.9543 | 83.7112 | 0.004622 | 0.98931 | 0.023786 | 24.0431 | 0.56559 | 0.15693 | 0.013336 | 0.95451 | 0.95451 | 0.006757
2019-03-20 17:50 | 1553118599.0792 | 59 | 0.20684 | 0.002779 | 2.4e-05 | 101.3347 | 95.3646 | 119.6349 | 21.0647 | 0.002801 | 0.002618 | 172.6266 | 0.26416 | 8.4045 | 0.003972 | 0.005457 | 0.005312 | 6.1e-05 | 0.001195 | 0.24818 | 1.3391 | 2.5481 | 100.58 | 20.67 | 19.059 | 89.016 | 117.3774 | 84.6183 | 0.004623 | 0.9868 | 0.023767 | 24.0847 | 0.54313 | 0.15715 | 0.013185 | 0.94294 | 0.94294 | 0.00685
2019-03-20 17:50 | 1553118600.0803 | 60 | 0.20687 | 0.00275 | 2.4e-05 | 101.2642 | 95.6307 | 119.3971 | 21.0654 | 0.002855 | 0.002576 | 172.5742 | 0.26407 | 8.3999 | 0.003972 | 0.005662 | 0.005586 | 5.8e-05 | 0.001194 | 0.26161 | 1.3671 | 2.5481 | 100.58 | 20.67 | 19.06 | 90.3184 | 117.3947 | 83.6015 | 0.004623 | 0.98961 | 0.023777 | 24.0265 | 0.51198 | 0.15709 | 0.013038 | 0.93281 | 0.93281 | 0.006922
2019-03-20 17:50 | 1553118601.0802 | 61 | 0.20689 | 0.002728 | 2.3e-05 | 101.1539 | 95.7308 | 119.3795 | 21.0645 | 0.002813 | 0.002584 | 172.5264 | 0.2629 | 8.4024 | 0.004007 | 0.005472 | 0.005325 | 5.8e-05 | 0.001198 | 0.28031 | 1.4028 | 2.5481 | 100.58 | 20.67 | 19.063 | 89.3034 | 118.2487 | 83.8012 | 0.004624 | 0.98906 | 0.023769 | 24.0315 | 0.49287 | 0.1564 | 0.012905 | 0.92298 | 0.92298 | 0.006964
2019-03-20 17:50 | 1553118602.0793 | 62 | 0.20692 | 0.002707 | 2.3e-05 | 101.0051 | 95.6464 | 119.8501 | 21.0647 | 0.002829 | 0.00258 | 172.4848 | 0.2631 | 8.4034 | 0.003914 | 0.005408 | 0.005158 | 5.8e-05 | 0.001193 | 0.26913 | 1.383 | 2.5482 | 100.58 | 20.67 | 19.063 | 89.6888 | 115.9641 | 83.7016 | 0.004624 | 0.98933 | 0.023819 | 24.0754 | 0.4876 | 0.15652 | 0.012756 | 0.91427 | 0.91427 | 0.007036
2019-03-20 17:50 | 1553118603.0805 | 63 | 0.20699 | 0.002677 | 2.3e-05 | 100.9126 | 95.3982 | 119.6745 | 21.0646 | 0.00286 | 0.002578 | 172.4535 | 0.26376 | 8.4009 | 0.003967 | 0.005407 | 0.005374 | 6e-05 | 0.001198 | 0.27959 | 1.3843 | 2.5482 | 100.58 | 20.67 | 19.06 | 90.4388 | 117.2584 | 83.6514 | 0.004623 | 0.98947 | 0.023803 | 24.0561 | 0.50264 | 0.15691 | 0.012376 | 0.88647 | 0.88647 | 0.007275
2019-03-20 17:50 | 1553118604.0805 | 64 | 0.2069 | 0.002654 | 2.3e-05 | 100.8311 | 95.7776 | 119.4581 | 21.0643 | 0.002899 | 0.002612 | 172.408 | 0.26349 | 8.4036 | 0.003919 | 0.005415 | 0.005265 | 5.9e-05 | 0.001199 | 0.29231 | 1.3657 | 2.548 | 100.58 | 20.67 | 19.059 | 91.3736 | 116.0871 | 84.4751 | 0.004623 | 0.98719 | 0.023754 | 24.0621 | 0.45992 | 0.15675 | 0.012945 | 0.92531 | 0.92531 | 0.006962
2019-03-20 17:50 | 1553118605.0793 | 65 | 0.20695 | 0.002634 | 2.3e-05 | 100.7386 | 95.7006 | 119.1201 | 21.0644 | 0.002903 | 0.002614 | 172.3407 | 0.26257 | 8.4056 | 0.003905 | 0.005051 | 0.005045 | 6e-05 | 0.001203 | 0.26449 | 1.3086 | 2.548 | 100.58 | 20.67 | 19.058 | 91.476 | 115.7438 | 84.5175 | 0.004623 | 0.98708 | 0.023719 | 24.0294 | 0.4589 | 0.1562 | 0.012706 | 0.90687 | 0.90687 | 0.007079
2019-03-20 17:50 | 1553118606.0805 | 66 | 0.20705 | 0.002607 | 2.3e-05 | 100.782 | 96.583 | 119.0621 | 21.0636 | 0.002843 | 0.002582 | 172.3205 | 0.26372 | 8.4022 | 0.0039 | 0.006009 | 0.005808 | 6e-05 | 0.001204 | 0.27793 | 1.3133 | 2.5482 | 100.58 | 20.67 | 19.056 | 90.0322 | 115.6362 | 83.7317 | 0.004622 | 0.98925 | 0.023739 | 23.9972 | 0.38065 | 0.15688 | 0.012147 | 0.86768 | 0.86768 | 0.007431
2019-03-20 17:50 | 1553118607.0795 | 67 | 0.20706 | 0.002585 | 2.4e-05 | 100.7968 | 96.7063 | 118.9589 | 21.0646 | 0.002844 | 0.00256 | 172.2725 | 0.26381 | 8.4036 | 0.003963 | 0.005366 | 0.005301 | 6.1e-05 | 0.001209 | 0.274 | 1.3069 | 2.5481 | 100.58 | 20.67 | 19.055 | 90.0582 | 117.1546 | 83.2124 | 0.004622 | 0.99069 | 0.023746 | 23.9693 | 0.37056 | 0.15694 | 0.0121 | 0.86456 | 0.86456 | 0.007461
2019-03-20 17:50 | 1553118608.0805 | 68 | 0.20705 | 0.002562 | 2.4e-05 | 100.825 | 96.4559 | 119.0436 | 21.0645 | 0.002807 | 0.00257 | 172.2261 | 0.26327 | 8.4038 | 0.003955 | 0.005698 | 0.005584 | 5.8e-05 | 0.00121 | 0.26161 | 1.3337 | 2.548 | 100.58 | 20.67 | 19.053 | 89.1634 | 116.9651 | 83.4611 | 0.004622 | 0.99 | 0.023746 | 23.9862 | 0.39624 | 0.15662 | 0.012188 | 0.87088 | 0.87088 | 0.007391
2019-03-20 17:50 | 1553118609.0809 | 69 | 0.20706 | 0.002549 | 2.4e-05 | 100.8573 | 96.4735 | 118.9054 | 21.0646 | 0.002858 | 0.002634 | 172.1816 | 0.26299 | 8.4025 | 0.003931 | 0.005912 | 0.005752 | 6.1e-05 | 0.001204 | 0.28871 | 1.332 | 2.5481 | 100.58 | 20.66 | 19.052 | 90.3865 | 116.3747 | 84.9888 | 0.004618 | 0.98578 | 0.023682 | 24.0236 | 0.39748 | 0.15646 | 0.012121 | 0.86372 | 0.86372 | 0.007445
2019-03-20 17:50 | 1553118610.0811 | 70 | 0.20706 | 0.002532 | 2.4e-05 | 100.7975 | 96.9033 | 119.1269 | 21.065 | 0.002871 | 0.00256 | 172.124 | 0.26266 | 8.4047 | 0.003873 | 0.005008 | 0.004949 | 5.8e-05 | 0.001205 | 0.28854 | 1.3543 | 2.5482 | 100.58 | 20.66 | 19.052 | 90.694 | 114.9763 | 83.2072 | 0.004618 | 0.99071 | 0.023763 | 23.9861 | 0.35241 | 0.15625 | 0.012158 | 0.86933 | 0.86933 | 0.007387
2019-03-20 17:50 | 1553118611.0797 | 71 | 0.20708 | 0.002515 | 2.4e-05 | 100.7829 | 97.0416 | 119.126 | 21.0663 | 0.002877 | 0.002551 | 172.0719 | 0.2637 | 8.3998 | 0.003976 | 0.005819 | 0.005661 | 5.7e-05 | 0.001206 | 0.28191 | 1.3478 | 2.5482 | 100.58 | 20.66 | 19.052 | 90.8541 | 117.4817 | 82.9917 | 0.004618 | 0.99131 | 0.02377 | 23.9787 | 0.33837 | 0.15688 | 0.012075 | 0.86359 | 0.86359 | 0.007466
2019-03-20 17:50 | 1553118612.0808 | 72 | 0.20708 | 0.002501 | 2.5e-05 | 100.7752 | 96.9784 | 119.2511 | 21.0649 | 0.002882 | 0.0026 | 172.0161 | 0.26309 | 8.3999 | 0.003943 | 0.005126 | 0.004973 | 5.6e-05 | 0.001201 | 0.27295 | 1.3709 | 2.5482 | 100.58 | 20.66 | 19.054 | 90.9677 | 116.6823 | 84.1637 | 0.004619 | 0.98805 | 0.023744 | 24.0308 | 0.3435 | 0.15651 | 0.012071 | 0.86233 | 0.86233 | 0.00746
2019-03-20 17:50 | 1553118613.0802 | 73 | 0.20704 | 0.002492 | 2.5e-05 | 100.7079 | 96.8846 | 119.3795 | 21.0649 | 0.00288 | 0.002585 | 171.9723 | 0.26277 | 8.4045 | 0.003938 | 0.005649 | 0.005485 | 5.7e-05 | 0.001196 | 0.2724 | 1.3675 | 2.5482 | 100.58 | 20.66 | 19.056 | 90.9245 | 116.5539 | 83.8133 | 0.004619 | 0.98902 | 0.023768 | 24.0319 | 0.34619 | 0.15632 | 0.01233 | 0.88177 | 0.88177 | 0.007286
2019-03-20 17:50 | 1553118614.0789 | 74 | 0.20707 | 0.00248 | 2.4e-05 | 100.7189 | 96.9089 | 119.2844 | 21.0657 | 0.002822 | 0.002499 | 171.9319 | 0.26372 | 8.4034 | 0.003836 | 0.004744 | 0.004703 | 6e-05 | 0.001198 | 0.27831 | 1.3546 | 2.5481 | 100.58 | 20.66 | 19.064 | 89.513 | 114.0621 | 81.7432 | 0.004621 | 0.99479 | 0.023828 | 23.9526 | 0.34492 | 0.15688 | 0.012179 | 0.87315 | 0.87315 | 0.007385
2019-03-20 17:50 | 1553118615.0799 | 75 | 0.20709 | 0.002477 | 2.4e-05 | 100.6866 | 96.7598 | 119.3607 | 21.0653 | 0.002872 | 0.002573 | 171.8979 | 0.26379 | 8.4024 | 0.003937 | 0.004915 | 0.004797 | 5.9e-05 | 0.001203 | 0.28711 | 1.3695 | 2.5483 | 100.58 | 20.66 | 19.068 | 90.7202 | 116.5409 | 83.5291 | 0.004622 | 0.98981 | 0.023776 | 24.0204 | 0.35582 | 0.15693 | 0.012036 | 0.86101 | 0.86101 | 0.007491
2019-03-20 17:50 | 1553118616.0791 | 76 | 0.2071 | 0.002468 | 2.5e-05 | 100.5523 | 96.5363 | 119.3227 | 21.0646 | 0.002831 | 0.00259 | 171.8547 | 0.26331 | 8.4046 | 0.003915 | 0.005839 | 0.005577 | 5.8e-05 | 0.001204 | 0.26217 | 1.3733 | 2.548 | 100.58 | 20.66 | 19.078 | 89.734 | 115.9862 | 83.9386 | 0.004625 | 0.98868 | 0.023758 | 24.0304 | 0.36458 | 0.15664 | 0.012019 | 0.85914 | 0.85914 | 0.007493
2019-03-20 17:50 | 1553118617.0791 | 77 | 0.20706 | 0.002459 | 2.6e-05 | 100.5491 | 96.6633 | 119.7135 | 21.0647 | 0.002765 | 0.00257 | 171.8206 | 0.26399 | 8.4002 | 0.003905 | 0.005185 | 0.00507 | 6.1e-05 | 0.0012 | 0.26737 | 1.3482 | 2.5482 | 100.58 | 20.66 | 19.084 | 88.1554 | 115.7363 | 83.4472 | 0.004626 | 0.99004 | 0.023814 | 24.0531 | 0.35252 | 0.15705 | 0.01222 | 0.87549 | 0.87549 | 0.007373
2019-03-20 17:50 | 1553118618.0801 | 78 | 0.20707 | 0.002452 | 2.5e-05 | 100.401 | 97.1619 | 119.7241 | 21.0656 | 0.00284 | 0.002595 | 171.7821 | 0.2636 | 8.4013 | 0.003846 | 0.005651 | 0.005555 | 5.9e-05 | 0.0012 | 0.27295 | 1.3636 | 2.5483 | 100.58 | 20.65 | 19.089 | 89.9624 | 114.3199 | 84.0449 | 0.004625 | 0.98838 | 0.023795 | 24.0744 | 0.29332 | 0.15682 | 0.012199 | 0.8733 | 0.8733 | 0.00738
2019-03-20 17:50 | 1553118619.0792 | 79 | 0.20706 | 0.002443 | 2.5e-05 | 100.3237 | 97.2841 | 119.9395 | 21.0651 | 0.002829 | 0.002548 | 171.7429 | 0.26345 | 8.4029 | 0.003833 | 0.005524 | 0.005343 | 6e-05 | 0.001198 | 0.27345 | 1.3065 | 2.548 | 100.58 | 20.65 | 19.098 | 89.6837 | 114.0014 | 82.9297 | 0.004627 | 0.99148 | 0.023853 | 24.0584 | 0.27519 | 0.15673 | 0.012298 | 0.88261 | 0.88261 | 0.007298
2019-03-20 17:50 | 1553118620.0792 | 80 | 0.20706 | 0.002434 | 2.5e-05 | 100.2428 | 96.8865 | 119.5518 | 21.0651 | 0.0028 | 0.002604 | 171.7187 | 0.26322 | 8.403 | 0.003827 | 0.005516 | 0.005411 | 5.9e-05 | 0.001198 | 0.26416 | 1.367 | 2.5482 | 100.58 | 20.65 | 19.101 | 88.9844 | 113.8371 | 84.2762 | 0.004627 | 0.98774 | 0.02377 | 24.0648 | 0.30461 | 0.15659 | 0.0123 | 0.87963 | 0.87963 | 0.007317
2019-03-20 17:50 | 1553118621.0792 | 81 | 0.20706 | 0.00243 | 2.5e-05 | 100.2538 | 96.9141 | 119.7666 | 21.0641 | 0.002839 | 0.002556 | 171.7096 | 0.26331 | 8.4023 | 0.003812 | 0.00506 | 0.004943 | 6.2e-05 | 0.001198 | 0.27489 | 1.3395 | 2.5482 | 100.58 | 20.65 | 19.11 | 89.9181 | 113.4891 | 83.1245 | 0.00463 | 0.99094 | 0.02383 | 24.0476 | 0.30304 | 0.15665 | 0.012283 | 0.88063 | 0.88063 | 0.007311
2019-03-20 17:50 | 1553118622.0797 | 82 | 0.20706 | 0.002426 | 2.5e-05 | 100.1254 | 96.7732 | 119.5354 | 21.0646 | 0.00281 | 0.002525 | 171.6668 | 0.26372 | 8.4002 | 0.003837 | 0.005556 | 0.005334 | 6e-05 | 0.001198 | 0.28672 | 1.3505 | 2.5481 | 100.58 | 20.65 | 19.114 | 89.2394 | 114.0888 | 82.369 | 0.004631 | 0.99304 | 0.023832 | 23.9989 | 0.30459 | 0.15688 | 0.012283 | 0.88079 | 0.88079 | 0.007321
2019-03-20 17:50 | 1553118623.0794 | 83 | 0.2071 | 0.002431 | 2.4e-05 | 100.0203 | 96.7246 | 119.3959 | 21.0654 | 0.002813 | 0.00259 | 171.6134 | 0.26383 | 8.4043 | 0.003909 | 0.004833 | 0.004732 | 5.9e-05 | 0.001204 | 0.29767 | 1.3251 | 2.5481 | 100.58 | 20.65 | 19.117 | 89.3105 | 115.8431 | 83.937 | 0.004631 | 0.98868 | 0.023766 | 24.0377 | 0.29969 | 0.15696 | 0.012022 | 0.85966 | 0.85966 | 0.007504
2019-03-20 17:50 | 1553118624.0804 | 84 | 0.20709 | 0.002427 | 2.3e-05 | 99.9915 | 96.6607 | 119.4938 | 21.0645 | 0.002802 | 0.00251 | 171.5736 | 0.26388 | 8.4013 | 0.003867 | 0.005529 | 0.005384 | 6.1e-05 | 0.001202 | 0.27223 | 1.3691 | 2.5481 | 100.58 | 20.65 | 19.122 | 89.0259 | 114.8162 | 82.0172 | 0.004633 | 0.99403 | 0.02384 | 23.9828 | 0.30303 | 0.15698 | 0.01213 | 0.87011 | 0.87011 | 0.007415
2019-03-20 17:50 | 1553118625.0805 | 85 | 0.20712 | 0.002421 | 2.4e-05 | 99.824 | 97.0334 | 119.6708 | 21.0642 | 0.00289 | 0.002518 | 171.5258 | 0.26416 | 8.4031 | 0.003886 | 0.005311 | 0.005192 | 6.3e-05 | 0.001205 | 0.2724 | 1.4068 | 2.5482 | 100.58 | 20.65 | 19.124 | 91.1522 | 115.2793 | 82.2064 | 0.004633 | 0.9935 | 0.023851 | 24.007 | 0.25361 | 0.15715 | 0.011954 | 0.85785 | 0.85785 | 0.007529
2019-03-20 17:50 | 1553118626.0796 | 86 | 0.20707 | 0.002417 | 2.5e-05 | 99.7752 | 97.0345 | 119.6328 | 21.0648 | 0.00283 | 0.002575 | 171.464 | 0.26413 | 8.4033 | 0.003872 | 0.005453 | 0.005347 | 5.8e-05 | 0.001206 | 0.25227 | 1.3972 | 2.5481 | 100.58 | 20.65 | 19.127 | 89.7179 | 114.9335 | 83.5815 | 0.004634 | 0.98967 | 0.023801 | 24.0496 | 0.24913 | 0.15713 | 0.012226 | 0.87549 | 0.87549 | 0.007377
2019-03-20 17:50 | 1553118627.0796 | 87 | 0.20701 | 0.002415 | 2.5e-05 | 99.6073 | 96.4843 | 119.1711 | 21.0647 | 0.002779 | 0.00253 | 171.423 | 0.2637 | 8.4024 | 0.003836 | 0.005825 | 0.0057 | 6e-05 | 0.001213 | 0.27417 | 1.3697 | 2.5482 | 100.58 | 20.66 | 19.132 | 88.4926 | 114.0604 | 82.5001 | 0.004638 | 0.99268 | 0.023791 | 23.9667 | 0.28493 | 0.15687 | 0.012629 | 0.90396 | 0.90396 | 0.007132
2019-03-20 17:50 | 1553118628.0792 | 88 | 0.20699 | 0.002421 | 2.6e-05 | 99.6752 | 96.3258 | 119.0131 | 21.0645 | 0.002804 | 0.002542 | 171.3742 | 0.26394 | 8.4041 | 0.003863 | 0.005755 | 0.005718 | 5.9e-05 | 0.001214 | 0.28191 | 1.3503 | 2.5481 | 100.58 | 20.66 | 19.135 | 89.0729 | 114.7211 | 82.7911 | 0.004639 | 0.99186 | 0.023766 | 23.9606 | 0.30573 | 0.15702 | 0.012743 | 0.91114 | 0.91114 | 0.007083
2019-03-20 17:50 | 1553118629.0796 | 89 | 0.20698 | 0.002425 | 2.6e-05 | 99.5776 | 95.9076 | 118.8643 | 21.063 | 0.002786 | 0.002544 | 171.3279 | 0.26278 | 8.4055 | 0.003863 | 0.005468 | 0.005353 | 6.1e-05 | 0.001215 | 0.27273 | 1.3246 | 2.5482 | 100.58 | 20.66 | 19.143 | 88.6434 | 114.7182 | 82.8207 | 0.00464 | 0.99178 | 0.02375 | 23.9466 | 0.33588 | 0.15633 | 0.012755 | 0.91134 | 0.91134 | 0.00705
2019-03-20 17:50 | 1553118630.08 | 90 | 0.20696 | 0.002435 | 2.6e-05 | 99.5313 | 95.9256 | 119.2581 | 21.0636 | 0.002815 | 0.002599 | 171.2838 | 0.26348 | 8.4032 | 0.003823 | 0.005754 | 0.005657 | 6e-05 | 0.001219 | 0.27704 | 1.348 | 2.5481 | 100.58 | 20.66 | 19.147 | 89.3456 | 113.7417 | 84.1482 | 0.004641 | 0.9881 | 0.023745 | 24.0309 | 0.33005 | 0.15674 | 0.012901 | 0.92161 | 0.92161 | 0.00699
2019-03-20 17:50 | 1553118631.0788 | 91 | 0.20696 | 0.002443 | 2.6e-05 | 99.5155 | 95.9155 | 119.3032 | 21.0643 | 0.002779 | 0.002616 | 171.2089 | 0.26368 | 8.4021 | 0.003801 | 0.005094 | 0.004998 | 6e-05 | 0.001223 | 0.29264 | 1.3469 | 2.5481 | 100.58 | 20.66 | 19.158 | 88.4829 | 113.2246 | 84.5564 | 0.004644 | 0.98697 | 0.023736 | 24.0492 | 0.32957 | 0.15687 | 0.012872 | 0.91918 | 0.91918 | 0.007014
2019-03-20 17:51 | 1553118632.0788 | 92 | 0.20694 | 0.00245 | 2.7e-05 | 99.3939 | 95.8938 | 119.2689 | 21.0636 | 0.002835 | 0.002546 | 171.1615 | 0.26478 | 8.4042 | 0.003824 | 0.004661 | 0.00445 | 6e-05 | 0.001221 | 0.26842 | 1.3275 | 2.5482 | 100.58 | 20.66 | 19.164 | 89.8316 | 113.7756 | 82.865 | 0.004646 | 0.99166 | 0.023789 | 23.9888 | 0.32066 | 0.15752 | 0.012961 | 0.92751 | 0.92751 | 0.00698
2019-03-20 17:51 | 1553118633.0793 | 93 | 0.20691 | 0.002463 | 2.8e-05 | 99.3356 | 95.7432 | 119.4992 | 21.0644 | 0.002884 | 0.002586 | 171.1235 | 0.26422 | 8.4036 | 0.003653 | 0.005666 | 0.005438 | 5.9e-05 | 0.001222 | 0.2912 | 1.3156 | 2.5481 | 100.58 | 20.66 | 19.169 | 91.0194 | 109.6028 | 83.8355 | 0.004647 | 0.98896 | 0.023779 | 24.0447 | 0.32948 | 0.15718 | 0.013135 | 0.93961 | 0.93961 | 0.006875
2019-03-20 17:51 | 1553118634.0804 | 94 | 0.20695 | 0.002466 | 2.7e-05 | 99.281 | 95.2928 | 119.2044 | 21.0643 | 0.002849 | 0.00253 | 171.0742 | 0.26476 | 8.4016 | 0.003677 | 0.004546 | 0.004414 | 6e-05 | 0.001227 | 0.26161 | 1.2984 | 2.5481 | 100.58 | 20.66 | 19.179 | 90.1697 | 110.2059 | 82.4998 | 0.004649 | 0.99268 | 0.023795 | 23.97 | 0.36673 | 0.1575 | 0.012892 | 0.92276 | 0.92276 | 0.007015
2019-03-20 17:51 | 1553118635.08 | 95 | 0.2069 | 0.002479 | 2.7e-05 | 99.2496 | 95.5268 | 119.0427 | 21.0648 | 0.002864 | 0.002572 | 171.0244 | 0.26351 | 8.4031 | 0.003796 | 0.004894 | 0.004758 | 6.1e-05 | 0.001226 | 0.24851 | 1.3432 | 2.5482 | 100.58 | 20.66 | 19.184 | 90.5187 | 113.0901 | 83.5026 | 0.00465 | 0.98989 | 0.023745 | 23.9875 | 0.34197 | 0.15676 | 0.013199 | 0.94284 | 0.94284 | 0.006833
2019-03-20 17:51 | 1553118636.0791 | 96 | 0.20688 | 0.002484 | 2.7e-05 | 99.1536 | 95.2053 | 119.2168 | 21.0643 | 0.002846 | 0.002585 | 170.9707 | 0.26325 | 8.4051 | 0.003877 | 0.005519 | 0.005411 | 5.9e-05 | 0.001225 | 0.2855 | 1.3197 | 2.5481 | 100.58 | 20.66 | 19.188 | 90.0846 | 115.0704 | 83.8058 | 0.004651 | 0.98904 | 0.023752 | 24.0153 | 0.36345 | 0.15661 | 0.013317 | 0.95156 | 0.95156 | 0.006764
2019-03-20 17:51 | 1553118637.0795 | 97 | 0.20687 | 0.002492 | 2.8e-05 | 99.074 | 95.1496 | 119.7081 | 21.0642 | 0.002817 | 0.002592 | 170.9236 | 0.26415 | 8.4047 | 0.003775 | 0.005067 | 0.004921 | 6e-05 | 0.001225 | 0.29214 | 1.3319 | 2.548 | 100.58 | 20.66 | 19.198 | 89.4047 | 112.5712 | 83.9851 | 0.004654 | 0.98855 | 0.023795 | 24.0707 | 0.3615 | 0.15714 | 0.013365 | 0.95669 | 0.95669 | 0.006751
2019-03-20 17:51 | 1553118638.0802 | 98 | 0.2069 | 0.002504 | 2.7e-05 | 98.9463 | 95.2132 | 119.6199 | 21.0642 | 0.002853 | 0.002567 | 170.8642 | 0.26428 | 8.4016 | 0.00377 | 0.004957 | 0.004814 | 6e-05 | 0.001219 | 0.28335 | 1.3315 | 2.5481 | 100.58 | 20.66 | 19.201 | 90.2563 | 112.4497 | 83.3813 | 0.004655 | 0.99022 | 0.023806 | 24.0415 | 0.34399 | 0.15722 | 0.013168 | 0.94303 | 0.94303 | 0.006852
2019-03-20 17:51 | 1553118639.0802 | 99 | 0.20686 | 0.002515 | 2.8e-05 | 98.9527 | 94.9736 | 119.3797 | 21.0638 | 0.002856 | 0.002553 | 170.8163 | 0.26412 | 8.4024 | 0.003878 | 0.004962 | 0.004803 | 6.2e-05 | 0.001223 | 0.28622 | 1.3874 | 2.5481 | 100.58 | 20.66 | 19.207 | 90.3287 | 115.0898 | 83.047 | 0.004656 | 0.99115 | 0.023794 | 24.0061 | 0.36711 | 0.15712 | 0.013401 | 0.95917 | 0.95917 | 0.006733
2019-03-20 17:51 | 1553118640.0793 | 100 | 0.2069 | 0.002519 | 2.7e-05 | 98.911 | 94.7796 | 119.1316 | 21.0645 | 0.002797 | 0.002559 | 170.7658 | 0.26389 | 8.4005 | 0.003767 | 0.005914 | 0.005758 | 5.4e-05 | 0.001214 | 0.27649 | 1.3862 | 2.5482 | 100.58 | 20.67 | 19.21 | 88.9222 | 112.3741 | 83.1818 | 0.00466 | 0.99078 | 0.023764 | 23.9857 | 0.38163 | 0.15699 | 0.013118 | 0.93784 | 0.93784 | 0.00688
2019-03-20 17:51 | 1553118641.0793 | 101 | 0.20691 | 0.002528 | 2.7e-05 | 98.8441 | 94.7875 | 118.926 | 21.0643 | 0.002844 | 0.002589 | 170.7033 | 0.26344 | 8.4048 | 0.003823 | 0.004748 | 0.004629 | 5.6e-05 | 0.001221 | 0.27577 | 1.4301 | 2.5483 | 100.58 | 20.67 | 19.213 | 90.0454 | 113.7438 | 83.9188 | 0.00466 | 0.98873 | 0.023719 | 23.9897 | 0.37483 | 0.15672 | 0.013086 | 0.93373 | 0.93373 | 0.006898
2019-03-20 17:51 | 1553118642.0794 | 102 | 0.20689 | 0.002532 | 2.7e-05 | 98.8227 | 94.8271 | 118.7544 | 21.0649 | 0.002814 | 0.002582 | 170.6747 | 0.26331 | 8.4044 | 0.003908 | 0.006047 | 0.005959 | 5.7e-05 | 0.001225 | 0.27328 | 1.3871 | 2.5482 | 100.58 | 20.67 | 19.22 | 89.3351 | 115.8197 | 83.7383 | 0.004662 | 0.98923 | 0.023708 | 23.9664 | 0.36916 | 0.15665 | 0.0132 | 0.94148 | 0.94148 | 0.006838
2019-03-20 17:51 | 1553118643.0794 | 103 | 0.20689 | 0.002544 | 2.6e-05 | 98.7768 | 95.1665 | 118.2988 | 21.0649 | 0.002804 | 0.00255 | 170.6343 | 0.26375 | 8.4038 | 0.003902 | 0.005123 | 0.005042 | 5.7e-05 | 0.00122 | 0.27688 | 1.3463 | 2.5482 | 100.58 | 20.67 | 19.222 | 89.0841 | 115.6755 | 82.9708 | 0.004663 | 0.99136 | 0.023688 | 23.8946 | 0.33305 | 0.1569 | 0.013139 | 0.93634 | 0.93634 | 0.006887
2019-03-20 17:51 | 1553118644.0795 | 104 | 0.2069 | 0.002545 | 2.6e-05 | 98.7648 | 95.3175 | 118.2117 | 21.0649 | 0.002866 | 0.002555 | 170.5798 | 0.26282 | 8.4018 | 0.003836 | 0.005353 | 0.005224 | 5.7e-05 | 0.00122 | 0.26178 | 1.3355 | 2.5481 | 100.58 | 20.67 | 19.227 | 90.5815 | 114.067 | 83.0918 | 0.004664 | 0.99103 | 0.023676 | 23.8899 | 0.31779 | 0.15635 | 0.013135 | 0.93551 | 0.93551 | 0.006869
2019-03-20 17:51 | 1553118645.0786 | 105 | 0.20689 | 0.002542 | 2.6e-05 | 98.6798 | 95.3205 | 118.2453 | 21.065 | 0.002868 | 0.002595 | 170.5484 | 0.26282 | 8.4009 | 0.003803 | 0.005408 | 0.00525 | 5.5e-05 | 0.001226 | 0.28136 | 1.322 | 2.5481 | 100.58 | 20.67 | 19.231 | 90.6269 | 113.2555 | 84.0592 | 0.004665 | 0.98834 | 0.023647 | 23.9257 | 0.3098 | 0.15635 | 0.0132 | 0.93906 | 0.93906 | 0.006843
2019-03-20 17:51 | 1553118646.0786 | 106 | 0.20686 | 0.002541 | 2.7e-05 | 98.6296 | 95.396 | 118.0836 | 21.0658 | 0.00282 | 0.002562 | 170.5059 | 0.26321 | 8.4024 | 0.00375 | 0.005467 | 0.005326 | 5.8e-05 | 0.001227 | 0.30431 | 1.327 | 2.5481 | 100.58 | 20.67 | 19.235 | 89.4749 | 111.9641 | 83.2682 | 0.004666 | 0.99054 | 0.023657 | 23.8828 | 0.29817 | 0.15659 | 0.013352 | 0.95022 | 0.95022 | 0.006773
2019-03-20 17:51 | 1553118647.0787 | 107 | 0.20681 | 0.002545 | 2.8e-05 | 98.6355 | 95.4658 | 117.7827 | 21.0654 | 0.002882 | 0.002565 | 170.4582 | 0.26327 | 8.4024 | 0.00377 | 0.005424 | 0.005334 | 5.7e-05 | 0.001226 | 0.27848 | 1.3341 | 2.5481 | 100.58 | 20.67 | 19.245 | 90.9515 | 112.4573 | 83.3294 | 0.004668 | 0.99037 | 0.023625 | 23.8544 | 0.29215 | 0.15662 | 0.013626 | 0.96832 | 0.96832 | 0.006648
2019-03-20 17:51 | 1553118648.0787 | 108 | 0.20676 | 0.00255 | 2.9e-05 | 98.6225 | 95.1859 | 117.7667 | 21.065 | 0.002869 | 0.002554 | 170.4164 | 0.26355 | 8.4015 | 0.003775 | 0.006229 | 0.006158 | 5.9e-05 | 0.001229 | 0.27776 | 1.3256 | 2.5481 | 100.58 | 20.67 | 19.25 | 90.6574 | 112.5878 | 83.0589 | 0.004669 | 0.99112 | 0.023632 | 23.8438 | 0.31724 | 0.15678 | 0.013949 | 0.99154 | 0.99154 | 0.006499
2019-03-20 17:51 | 1553118649.0791 | 109 | 0.2067 | 0.002559 | 2.9e-05 | 98.5711 | 94.6567 | 118.0073 | 21.0658 | 0.002842 | 0.002593 | 170.373 | 0.26433 | 8.4008 | 0.003692 | 0.005125 | 0.004989 | 5.7e-05 | 0.001229 | 0.28174 | 1.3012 | 2.5481 | 100.58 | 20.67 | 19.256 | 90.0109 | 110.5614 | 83.9991 | 0.004671 | 0.98851 | 0.023625 | 23.8996 | 0.36245 | 0.15725 | 0.014308 | 1.0168 | 1.0168 | 0.006356
2019-03-20 17:51 | 1553118650.0798 | 110 | 0.20672 | 0.002578 | 3e-05 | 98.5996 | 94.1764 | 118.1655 | 21.0654 | 0.002875 | 0.002579 | 170.3199 | 0.26444 | 8.3989 | 0.003701 | 0.005041 | 0.00491 | 5.9e-05 | 0.001225 | 0.28727 | 1.3086 | 2.5482 | 100.58 | 20.68 | 19.267 | 90.8017 | 110.7719 | 83.6648 | 0.004676 | 0.98944 | 0.023652 | 23.9044 | 0.41053 | 0.15731 | 0.014176 | 1.0085 | 1.0085 | 0.006411
2019-03-20 17:51 | 1553118651.0789 | 111 | 0.2067 | 0.002592 | 3e-05 | 98.5908 | 94.2967 | 118.1049 | 21.0664 | 0.002861 | 0.002553 | 170.2678 | 0.26376 | 8.4011 | 0.003618 | 0.005227 | 0.005116 | 5.6e-05 | 0.001226 | 0.28998 | 1.3213 | 2.548 | 100.58 | 20.68 | 19.273 | 90.4554 | 108.7587 | 83.0456 | 0.004678 | 0.99116 | 0.023666 | 23.8775 | 0.39831 | 0.15691 | 0.014242 | 1.0138 | 1.0138 | 0.006361
2019-03-20 17:51 | 1553118652.0799 | 112 | 0.20669 | 0.002609 | 2.9e-05 | 98.5213 | 94.4096 | 117.8989 | 21.066 | 0.002827 | 0.002612 | 170.2055 | 0.26398 | 8.4027 | 0.003631 | 0.00508 | 0.004974 | 5.7e-05 | 0.00123 | 0.28816 | 1.3578 | 2.5482 | 100.58 | 20.68 | 19.283 | 89.6314 | 109.0768 | 84.462 | 0.00468 | 0.98723 | 0.023599 | 23.9041 | 0.3813 | 0.15704 | 0.014314 | 1.0161 | 1.0161 | 0.006352
2019-03-20 17:51 | 1553118653.0799 | 113 | 0.20673 | 0.002622 | 2.9e-05 | 98.6138 | 94.8301 | 117.8428 | 21.0656 | 0.002795 | 0.002574 | 170.15 | 0.26338 | 8.4049 | 0.003686 | 0.00506 | 0.004819 | 6e-05 | 0.001226 | 0.28241 | 1.3324 | 2.5482 | 100.58 | 20.67 | 19.29 | 88.8646 | 110.4139 | 83.552 | 0.004679 | 0.98975 | 0.023623 | 23.868 | 0.34995 | 0.15668 | 0.014044 | 0.99794 | 0.99794 | 0.006453
2019-03-20 17:51 | 1553118654.079 | 114 | 0.2067 | 0.002633 | 2.9e-05 | 98.5659 | 94.373 | 118.1174 | 21.0656 | 0.002871 | 0.002555 | 170.0922 | 0.26428 | 8.4041 | 0.003779 | 0.005299 | 0.005208 | 5.8e-05 | 0.001226 | 0.27057 | 1.3314 | 2.5481 | 100.58 | 20.67 | 19.296 | 90.6954 | 112.6824 | 83.1026 | 0.00468 | 0.991 | 0.023666 | 23.8807 | 0.38882 | 0.15722 | 0.014188 | 1.01 | 1.01 | 0.006398
2019-03-20 17:51 | 1553118655.0804 | 115 | 0.20671 | 0.002644 | 2.8e-05 | 98.6584 | 94.7124 | 118.1829 | 21.0656 | 0.002875 | 0.002597 | 170.0224 | 0.26308 | 8.4015 | 0.003634 | 0.005003 | 0.004949 | 5.6e-05 | 0.001229 | 0.27151 | 1.3285 | 2.5479 | 100.58 | 20.67 | 19.31 | 90.7946 | 109.1403 | 84.1018 | 0.004684 | 0.98823 | 0.023639 | 23.9208 | 0.36511 | 0.15651 | 0.014108 | 1.0032 | 1.0032 | 0.006412
2019-03-20 17:51 | 1553118656.0801 | 116 | 0.20667 | 0.002644 | 2.9e-05 | 98.7202 | 94.8439 | 118.2096 | 21.0657 | 0.002844 | 0.002645 | 169.9792 | 0.26355 | 8.4005 | 0.003684 | 0.005073 | 0.004895 | 5.7e-05 | 0.001229 | 0.28727 | 1.309 | 2.5482 | 100.58 | 20.68 | 19.317 | 90.0422 | 110.3571 | 85.2552 | 0.004688 | 0.98504 | 0.023604 | 23.9621 | 0.35829 | 0.15679 | 0.014369 | 1.0202 | 1.0202 | 0.006316
2019-03-20 17:51 | 1553118657.0805 | 117 | 0.20667 | 0.002647 | 2.8e-05 | 98.5989 | 95.0428 | 118.0566 | 21.0664 | 0.002832 | 0.002629 | 169.9364 | 0.26374 | 8.403 | 0.003585 | 0.005325 | 0.005149 | 5.9e-05 | 0.001236 | 0.27184 | 1.349 | 2.5481 | 100.58 | 20.68 | 19.331 | 89.756 | 107.973 | 84.8845 | 0.004692 | 0.98606 | 0.023601 | 23.9342 | 0.32856 | 0.1569 | 0.014366 | 1.0199 | 1.0199 | 0.006323
2019-03-20 17:51 | 1553118658.0803 | 118 | 0.20665 | 0.002655 | 2.9e-05 | 98.6051 | 94.7247 | 117.9312 | 21.0652 | 0.002865 | 0.00258 | 169.8981 | 0.26375 | 8.4042 | 0.003676 | 0.005263 | 0.005185 | 5.5e-05 | 0.001233 | 0.26753 | 1.3625 | 2.5481 | 100.58 | 20.68 | 19.338 | 90.5599 | 110.1799 | 83.6865 | 0.004694 | 0.98938 | 0.023628 | 23.8814 | 0.3591 | 0.1569 | 0.014505 | 1.0309 | 1.0309 | 0.006255
2019-03-20 17:51 | 1553118659.0793 | 119 | 0.20664 | 0.002661 | 2.9e-05 | 98.6073 | 94.7568 | 118.0106 | 21.0661 | 0.002875 | 0.002621 | 169.8622 | 0.26381 | 8.4023 | 0.003667 | 0.004696 | 0.004527 | 5.6e-05 | 0.001229 | 0.262 | 1.3115 | 2.5481 | 100.58 | 20.68 | 19.343 | 90.785 | 109.9425 | 84.6843 | 0.004695 | 0.98662 | 0.023603 | 23.9228 | 0.35628 | 0.15694 | 0.014544 | 1.0326 | 1.0326 | 0.006247
2019-03-20 17:51 | 1553118660.0795 | 120 | 0.20663 | 0.002675 | 2.9e-05 | 98.6448 | 94.8084 | 117.7993 | 21.0661 | 0.002813 | 0.002575 | 169.8491 | 0.26402 | 8.4007 | 0.003653 | 0.005442 | 0.005315 | 5.6e-05 | 0.001233 | 0.28998 | 1.3095 | 2.548 | 100.58 | 20.68 | 19.351 | 89.3112 | 109.6147 | 83.5772 | 0.004697 | 0.98968 | 0.023618 | 23.8644 | 0.35481 | 0.15706 | 0.014563 | 1.0346 | 1.0346 | 0.00624
2019-03-20 17:51 | 1553118661.0784 | 121 | 0.20665 | 0.002677 | 3e-05 | 98.8062 | 95.0244 | 117.6153 | 21.0658 | 0.002829 | 0.002642 | 169.8151 | 0.26368 | 8.4022 | 0.00368 | 0.005079 | 0.00495 | 5.7e-05 | 0.00123 | 0.26217 | 1.2715 | 2.5481 | 100.58 | 20.68 | 19.355 | 89.6875 | 110.2634 | 85.1983 | 0.004698 | 0.9852 | 0.023546 | 23.8999 | 0.34907 | 0.15686 | 0.014445 | 1.023 | 1.023 | 0.006302
2019-03-20 17:51 | 1553118662.0788 | 122 | 0.20666 | 0.002682 | 3e-05 | 98.823 | 94.7864 | 117.6604 | 21.0661 | 0.002882 | 0.00262 | 169.7798 | 0.26371 | 8.4018 | 0.003705 | 0.005041 | 0.004904 | 5.7e-05 | 0.001235 | 0.27577 | 1.2999 | 2.5482 | 100.58 | 20.68 | 19.362 | 90.9688 | 110.8727 | 84.6687 | 0.004699 | 0.98666 | 0.023568 | 23.8868 | 0.37303 | 0.15688 | 0.014409 | 1.0214 | 1.0214 | 0.006313
2019-03-20 17:51 | 1553118663.0786 | 123 | 0.20666 | 0.002699 | 3.1e-05 | 98.7994 | 94.9411 | 117.8862 | 21.0649 | 0.002889 | 0.002572 | 169.7383 | 0.26375 | 8.4042 | 0.003629 | 0.005472 | 0.005345 | 5.6e-05 | 0.001228 | 0.28926 | 1.328 | 2.5481 | 100.58 | 20.68 | 19.364 | 91.1425 | 109.0308 | 83.4943 | 0.0047 | 0.98991 | 0.02363 | 23.8704 | 0.35631 | 0.1569 | 0.014365 | 1.0209 | 1.0209 | 0.006317
2019-03-20 17:51 | 1553118664.0796 | 124 | 0.20668 | 0.002704 | 3.1e-05 | 98.8234 | 94.983 | 118.0446 | 21.0657 | 0.002881 | 0.002617 | 169.7079 | 0.26304 | 8.4036 | 0.003572 | 0.005394 | 0.005355 | 5.7e-05 | 0.001225 | 0.26073 | 1.3253 | 2.5482 | 100.58 | 20.69 | 19.365 | 90.9441 | 107.6417 | 84.5796 | 0.004703 | 0.98691 | 0.02361 | 23.9228 | 0.35453 | 0.15648 | 0.014223 | 1.0099 | 1.0099 | 0.006368
2019-03-20 17:51 | 1553118665.0789 | 125 | 0.2067 | 0.002703 | 3.2e-05 | 98.803 | 95.4868 | 117.6503 | 21.0656 | 0.002827 | 0.002634 | 169.6667 | 0.26424 | 8.4002 | 0.0037 | 0.005091 | 0.004941 | 5.7e-05 | 0.001216 | 0.2808 | 1.3226 | 2.5481 | 100.58 | 20.69 | 19.36 | 89.6396 | 110.7563 | 84.9867 | 0.004702 | 0.98578 | 0.023557 | 23.8964 | 0.30537 | 0.15719 | 0.014116 | 1 | 1 | 0.006461
2019-03-20 17:51 | 1553118666.0797 | 126 | 0.2067 | 0.002709 | 3.3e-05 | 98.7519 | 95.2035 | 117.409 | 21.0662 | 0.002842 | 0.002643 | 169.6175 | 0.26385 | 8.4026 | 0.003643 | 0.005262 | 0.005163 | 5.7e-05 | 0.001211 | 0.2792 | 1.339 | 2.5482 | 100.58 | 20.69 | 19.355 | 90.0051 | 109.3797 | 85.2059 | 0.004701 | 0.98518 | 0.023525 | 23.8792 | 0.32732 | 0.15697 | 0.014131 | 0.99972 | 0.99972 | 0.006453
2019-03-20 17:51 | 1553118667.0787 | 127 | 0.20668 | 0.002715 | 3.3e-05 | 98.7713 | 95.1328 | 117.4498 | 21.0661 | 0.002975 | 0.002622 | 169.5605 | 0.2636 | 8.4029 | 0.003642 | 0.005077 | 0.004885 | 6e-05 | 0.001213 | 0.27688 | 1.326 | 2.548 | 100.58 | 20.69 | 19.349 | 93.1966 | 109.3435 | 84.7142 | 0.004699 | 0.98654 | 0.023546 | 23.8669 | 0.33571 | 0.15682 | 0.014204 | 1.0057 | 1.0057 | 0.006408
2019-03-20 17:51 | 1553118668.0787 | 128 | 0.20671 | 0.002713 | 3.4e-05 | 98.7053 | 95.0211 | 117.6219 | 21.0657 | 0.00286 | 0.002599 | 169.5152 | 0.26337 | 8.4022 | 0.003629 | 0.004668 | 0.004553 | 6.1e-05 | 0.001213 | 0.28119 | 1.3311 | 2.5481 | 100.58 | 20.7 | 19.332 | 90.4391 | 109.0277 | 84.1457 | 0.004698 | 0.9881 | 0.023582 | 23.8654 | 0.34025 | 0.15668 | 0.014064 | 0.99725 | 0.99725 | 0.006457
2019-03-20 17:51 | 1553118669.0802 | 129 | 0.20669 | 0.00271 | 3.4e-05 | 98.6523 | 95.4146 | 117.4848 | 21.0658 | 0.002838 | 0.002595 | 169.4585 | 0.26415 | 8.4012 | 0.003605 | 0.005798 | 0.00567 | 6.2e-05 | 0.001213 | 0.26593 | 1.3491 | 2.5481 | 100.58 | 20.7 | 19.323 | 89.9142 | 108.4596 | 84.06 | 0.004696 | 0.98834 | 0.023571 | 23.8487 | 0.29847 | 0.15714 | 0.014173 | 1.0045 | 1.0045 | 0.00643
2019-03-20 17:51 | 1553118670.0789 | 130 | 0.20671 | 0.002706 | 3.6e-05 | 98.6878 | 95.3657 | 117.5785 | 21.0656 | 0.002833 | 0.002664 | 169.3949 | 0.26454 | 8.4041 | 0.003599 | 0.005577 | 0.005497 | 6.1e-05 | 0.001204 | 0.27815 | 1.33 | 2.5481 | 100.58 | 20.7 | 19.305 | 89.7935 | 108.2906 | 85.7126 | 0.004691 | 0.98379 | 0.023526 | 23.9133 | 0.30629 | 0.15738 | 0.014061 | 0.99456 | 0.99456 | 0.006504
2019-03-20 17:51 | 1553118671.0799 | 131 | 0.20672 | 0.002711 | 3.7e-05 | 98.7108 | 95.387 | 117.7752 | 21.0658 | 0.002843 | 0.002591 | 169.3445 | 0.26388 | 8.403 | 0.003536 | 0.004637 | 0.004454 | 5.9e-05 | 0.001198 | 0.27328 | 1.3375 | 2.5481 | 100.58 | 20.7 | 19.295 | 90.0305 | 106.7762 | 83.9623 | 0.004689 | 0.98861 | 0.023603 | 23.8748 | 0.30637 | 0.15699 | 0.013948 | 0.98981 | 0.98981 | 0.006519
2019-03-20 17:51 | 1553118672.08 | 132 | 0.20671 | 0.00271 | 3.7e-05 | 98.6933 | 95.1624 | 117.9338 | 21.0666 | 0.002987 | 0.002642 | 169.2717 | 0.26366 | 8.4037 | 0.003613 | 0.004676 | 0.004612 | 6.2e-05 | 0.001202 | 0.27184 | 1.2885 | 2.5479 | 100.58 | 20.7 | 19.284 | 93.4888 | 108.6542 | 85.1764 | 0.004686 | 0.98526 | 0.023579 | 23.9315 | 0.32588 | 0.15685 | 0.014042 | 0.99544 | 0.99544 | 0.006476
2019-03-20 17:51 | 1553118673.08 | 133 | 0.20675 | 0.002705 | 3.6e-05 | 98.7435 | 95.4894 | 118.1054 | 21.0666 | 0.002886 | 0.002676 | 169.2158 | 0.26406 | 8.4025 | 0.003634 | 0.00543 | 0.00527 | 6.2e-05 | 0.001194 | 0.27632 | 1.3043 | 2.5479 | 100.58 | 20.7 | 19.273 | 91.055 | 109.1529 | 86.0046 | 0.004684 | 0.98299 | 0.023569 | 23.9766 | 0.29973 | 0.15709 | 0.013785 | 0.97682 | 0.97682 | 0.00661
2019-03-20 17:51 | 1553118674.0796 | 134 | 0.20677 | 0.002709 | 3.6e-05 | 98.7823 | 95.5948 | 117.946 | 21.0652 | 0.002853 | 0.002652 | 169.1669 | 0.26355 | 8.4041 | 0.003613 | 0.005661 | 0.005452 | 6.3e-05 | 0.00119 | 0.24475 | 1.3016 | 2.5482 | 100.58 | 20.7 | 19.247 | 90.2684 | 108.6511 | 85.4303 | 0.004677 | 0.98456 | 0.023572 | 23.9412 | 0.29338 | 0.15679 | 0.01368 | 0.96949 | 0.96949 | 0.006647
2019-03-20 17:51 | 1553118675.0782 | 135 | 0.20678 | 0.002696 | 3.7e-05 | 98.769 | 95.4543 | 118.0538 | 21.0662 | 0.002891 | 0.002642 | 169.1078 | 0.26391 | 8.4032 | 0.003645 | 0.00566 | 0.005558 | 5.7e-05 | 0.001185 | 0.26626 | 1.3045 | 2.5482 | 100.58 | 20.71 | 19.234 | 91.1901 | 109.4104 | 85.1902 | 0.004677 | 0.98522 | 0.02359 | 23.9441 | 0.30534 | 0.157 | 0.013624 | 0.96624 | 0.96624 | 0.006678
2019-03-20 17:51 | 1553118676.0792 | 136 | 0.20675 | 0.002699 | 3.8e-05 | 98.7629 | 95.8714 | 118.4948 | 21.0663 | 0.002943 | 0.002609 | 169.0675 | 0.26445 | 8.4033 | 0.003696 | 0.005402 | 0.005268 | 6.1e-05 | 0.001186 | 0.26089 | 1.3067 | 2.548 | 100.58 | 20.71 | 19.221 | 92.431 | 110.6614 | 84.3935 | 0.004674 | 0.98742 | 0.023661 | 23.9621 | 0.26578 | 0.15732 | 0.013847 | 0.98495 | 0.98495 | 0.006565
2019-03-20 17:51 | 1553118677.0797 | 137 | 0.20675 | 0.002694 | 3.9e-05 | 98.7651 | 96.0212 | 118.5272 | 21.0666 | 0.002878 | 0.002672 | 169.028 | 0.26359 | 8.402 | 0.003701 | 0.005454 | 0.005237 | 6e-05 | 0.001182 | 0.2724 | 1.3168 | 2.548 | 100.58 | 20.71 | 19.196 | 90.8582 | 110.7886 | 85.9187 | 0.004668 | 0.98323 | 0.023614 | 24.0165 | 0.25202 | 0.15681 | 0.013794 | 0.97914 | 0.97914 | 0.006582
2019-03-20 17:51 | 1553118678.0787 | 138 | 0.20674 | 0.002686 | 4.1e-05 | 98.8981 | 96.2302 | 118.3648 | 21.0661 | 0.002929 | 0.002651 | 168.9909 | 0.26322 | 8.4017 | 0.003686 | 0.005328 | 0.005194 | 5.9e-05 | 0.001176 | 0.27704 | 1.347 | 2.5481 | 100.58 | 20.71 | 19.184 | 92.0961 | 110.4144 | 85.4001 | 0.004665 | 0.98465 | 0.023614 | 23.9827 | 0.2446 | 0.15659 | 0.013896 | 0.98632 | 0.98632 | 0.006525
2019-03-20 17:51 | 1553118679.0793 | 139 | 0.20673 | 0.002695 | 4.2e-05 | 98.8554 | 96.013 | 118.311 | 21.0672 | 0.00293 | 0.002671 | 168.9302 | 0.26457 | 8.4043 | 0.003709 | 0.005376 | 0.00529 | 6.1e-05 | 0.001174 | 0.27975 | 1.3245 | 2.548 | 100.58 | 20.71 | 19.173 | 92.1222 | 110.9731 | 85.8825 | 0.004662 | 0.98332 | 0.023593 | 23.9933 | 0.26095 | 0.15739 | 0.013941 | 0.98859 | 0.98859 | 0.006543
2019-03-20 17:51 | 1553118680.0794 | 140 | 0.20674 | 0.002689 | 4.2e-05 | 98.781 | 96.1666 | 118.3127 | 21.0667 | 0.002935 | 0.002617 | 168.8639 | 0.26361 | 8.4028 | 0.003683 | 0.005269 | 0.005209 | 6e-05 | 0.001172 | 0.26842 | 1.3444 | 2.548 | 100.58 | 20.71 | 19.151 | 92.2304 | 110.3535 | 84.5928 | 0.004657 | 0.98687 | 0.023636 | 23.9504 | 0.23992 | 0.15682 | 0.013879 | 0.98602 | 0.98602 | 0.006537
2019-03-20 17:51 | 1553118681.0795 | 141 | 0.20674 | 0.002694 | 4.1e-05 | 98.8004 | 96.0702 | 118.343 | 21.0669 | 0.002878 | 0.002623 | 168.8276 | 0.26441 | 8.4024 | 0.003699 | 0.00471 | 0.004531 | 6.1e-05 | 0.001177 | 0.2792 | 1.3369 | 2.548 | 100.58 | 20.71 | 19.14 | 90.8742 | 110.7393 | 84.7243 | 0.004654 | 0.98651 | 0.023635 | 23.9578 | 0.25065 | 0.15729 | 0.013878 | 0.98589 | 0.98589 | 0.006557
2019-03-20 17:51 | 1553118682.0799 | 142 | 0.20676 | 0.002699 | 4.2e-05 | 98.8221 | 96.0209 | 118.4373 | 21.0664 | 0.002898 | 0.002648 | 168.7956 | 0.26458 | 8.4014 | 0.003625 | 0.005399 | 0.005257 | 5.9e-05 | 0.001176 | 0.25769 | 1.3424 | 2.5481 | 100.58 | 20.72 | 19.119 | 91.3387 | 108.9431 | 85.3309 | 0.004652 | 0.98484 | 0.023624 | 23.9877 | 0.2572 | 0.1574 | 0.013775 | 0.97814 | 0.97814 | 0.006614
2019-03-20 17:51 | 1553118683.0789 | 143 | 0.20681 | 0.002701 | 4.2e-05 | 98.7632 | 96.0575 | 118.3347 | 21.0666 | 0.002887 | 0.002687 | 168.725 | 0.2629 | 8.4036 | 0.003604 | 0.005529 | 0.005348 | 6.1e-05 | 0.001179 | 0.29175 | 1.3757 | 2.5482 | 100.58 | 20.72 | 19.109 | 91.077 | 108.4169 | 86.2693 | 0.004649 | 0.98227 | 0.023583 | 24.0087 | 0.24846 | 0.1564 | 0.013443 | 0.95284 | 0.95284 | 0.006746
2019-03-20 17:51 | 1553118684.0786 | 144 | 0.2068 | 0.002699 | 4.1e-05 | 98.673 | 96.1939 | 118.3214 | 21.0653 | 0.002961 | 0.002632 | 168.6564 | 0.26301 | 8.4041 | 0.003669 | 0.004897 | 0.004843 | 5.8e-05 | 0.001178 | 0.29031 | 1.3226 | 2.5481 | 100.58 | 20.72 | 19.1 | 92.8713 | 109.9923 | 84.9411 | 0.004647 | 0.98591 | 0.023625 | 23.9629 | 0.22759 | 0.15646 | 0.013525 | 0.96042 | 0.96042 | 0.006696
2019-03-20 17:51 | 1553118685.0787 | 145 | 0.2068 | 0.002701 | 4.1e-05 | 98.641 | 96.272 | 118.316 | 21.0666 | 0.002867 | 0.002645 | 168.6242 | 0.26333 | 8.4025 | 0.003588 | 0.005705 | 0.0055 | 6.1e-05 | 0.001181 | 0.28888 | 1.3397 | 2.5481 | 100.58 | 20.72 | 19.086 | 90.5925 | 108.03 | 85.2647 | 0.004644 | 0.98502 | 0.023614 | 23.9732 | 0.21743 | 0.15665 | 0.013482 | 0.95691 | 0.95691 | 0.006728
2019-03-20 17:51 | 1553118686.0788 | 146 | 0.20682 | 0.002697 | 4.1e-05 | 98.5802 | 96.1909 | 118.0458 | 21.0659 | 0.002923 | 0.00265 | 168.5651 | 0.26366 | 8.4015 | 0.003482 | 0.005289 | 0.005189 | 6.2e-05 | 0.001184 | 0.27992 | 1.299 | 2.548 | 100.58 | 20.72 | 19.08 | 91.9631 | 105.4723 | 85.385 | 0.004642 | 0.98469 | 0.023583 | 23.9498 | 0.21945 | 0.15685 | 0.013418 | 0.95113 | 0.95113 | 0.006778
2019-03-20 17:51 | 1553118687.0788 | 147 | 0.20677 | 0.002695 | 4.1e-05 | 98.4343 | 95.9323 | 117.8728 | 21.066 | 0.002887 | 0.002612 | 168.502 | 0.26344 | 8.404 | 0.003531 | 0.00548 | 0.005435 | 6e-05 | 0.001185 | 0.27776 | 1.2964 | 2.548 | 100.58 | 20.72 | 19.069 | 91.0945 | 106.6561 | 84.4589 | 0.00464 | 0.98724 | 0.023596 | 23.9013 | 0.23029 | 0.15672 | 0.013697 | 0.97145 | 0.97145 | 0.006631
2019-03-20 17:51 | 1553118688.0798 | 148 | 0.20682 | 0.002694 | 4.1e-05 | 98.5132 | 95.9839 | 118.2059 | 21.066 | 0.002864 | 0.002604 | 168.4263 | 0.26461 | 8.4011 | 0.003513 | 0.005065 | 0.004889 | 6e-05 | 0.001184 | 0.26969 | 1.2897 | 2.548 | 100.58 | 20.72 | 19.065 | 90.5199 | 106.224 | 84.2622 | 0.004639 | 0.98778 | 0.023636 | 23.9285 | 0.23265 | 0.15741 | 0.013399 | 0.95193 | 0.95193 | 0.006796
2019-03-20 17:51 | 1553118689.0798 | 149 | 0.2068 | 0.002696 | 4.1e-05 | 98.3867 | 95.9783 | 118.2178 | 21.0656 | 0.002882 | 0.002706 | 168.3645 | 0.26323 | 8.4036 | 0.003559 | 0.00443 | 0.004333 | 6.1e-05 | 0.001181 | 0.27007 | 1.3095 | 2.548 | 100.58 | 20.72 | 19.062 | 90.9541 | 107.336 | 86.7376 | 0.004638 | 0.98099 | 0.023556 | 24.0124 | 0.22169 | 0.1566 | 0.013525 | 0.95758 | 0.95758 | 0.006721
2019-03-20 17:51 | 1553118690.0779 | 150 | 0.20681 | 0.002688 | 4.2e-05 | 98.3521 | 95.97 | 118.2493 | 21.0663 | 0.00288 | 0.002649 | 168.3098 | 0.26384 | 8.4046 | 0.003566 | 0.005585 | 0.005497 | 6.3e-05 | 0.001186 | 0.28672 | 1.3089 | 2.548 | 100.58 | 20.72 | 19.058 | 90.9127 | 107.5107 | 85.366 | 0.004637 | 0.98474 | 0.023604 | 23.9698 | 0.2193 | 0.15696 | 0.013439 | 0.95341 | 0.95341 | 0.006766
2019-03-20 17:52 | 1553118691.0784 | 151 | 0.20681 | 0.002686 | 4.2e-05 | 98.3192 | 95.9095 | 118.3333 | 21.0663 | 0.002913 | 0.002681 | 168.2553 | 0.26398 | 8.404 | 0.003563 | 0.00507 | 0.004923 | 6.3e-05 | 0.001186 | 0.2907 | 1.3358 | 2.5481 | 100.58 | 20.72 | 19.057 | 91.7109 | 107.4279 | 86.1136 | 0.004637 | 0.98269 | 0.023588 | 24.0033 | 0.22195 | 0.15704 | 0.013491 | 0.95647 | 0.95647 | 0.006748
2019-03-20 17:52 | 1553118692.0795 | 152 | 0.20681 | 0.002682 | 4.3e-05 | 98.3596 | 95.6101 | 118.3918 | 21.0662 | 0.002959 | 0.002637 | 168.1913 | 0.26362 | 8.4033 | 0.003521 | 0.005622 | 0.005453 | 6.5e-05 | 0.001185 | 0.26936 | 1.3252 | 2.5482 | 100.58 | 20.72 | 19.055 | 92.808 | 106.402 | 85.0786 | 0.004636 | 0.98553 | 0.023628 | 23.9746 | 0.25359 | 0.15683 | 0.013497 | 0.95846 | 0.95846 | 0.006725
2019-03-20 17:52 | 1553118693.0791 | 153 | 0.2068 | 0.002681 | 4.3e-05 | 98.2926 | 95.9181 | 118.2124 | 21.0654 | 0.002901 | 0.002662 | 168.162 | 0.26406 | 8.4018 | 0.003537 | 0.004616 | 0.004528 | 6.6e-05 | 0.001179 | 0.25282 | 1.3137 | 2.548 | 100.58 | 20.72 | 19.054 | 91.4144 | 106.8075 | 85.6617 | 0.004636 | 0.98393 | 0.023591 | 23.976 | 0.21874 | 0.15709 | 0.013501 | 0.95718 | 0.95718 | 0.006745
2019-03-20 17:52 | 1553118694.0781 | 154 | 0.20681 | 0.002677 | 4.4e-05 | 98.2755 | 96.0814 | 118.3458 | 21.0648 | 0.002974 | 0.002637 | 168.1404 | 0.26385 | 8.401 | 0.003491 | 0.004836 | 0.004632 | 6.7e-05 | 0.001174 | 0.27489 | 1.3353 | 2.5479 | 100.58 | 20.72 | 19.054 | 93.1758 | 105.6767 | 85.0766 | 0.004636 | 0.98554 | 0.023623 | 23.9699 | 0.20196 | 0.15697 | 0.013479 | 0.95692 | 0.95692 | 0.006742
2019-03-20 17:52 | 1553118695.0782 | 155 | 0.2068 | 0.002679 | 4.6e-05 | 98.3499 | 96.2896 | 118.4049 | 21.0656 | 0.002934 | 0.002626 | 168.1143 | 0.26393 | 8.4012 | 0.003514 | 0.00507 | 0.004964 | 6.4e-05 | 0.001168 | 0.27433 | 1.3143 | 2.5481 | 100.58 | 20.72 | 19.054 | 92.2214 | 106.237 | 84.7971 | 0.004636 | 0.98631 | 0.023638 | 23.9666 | 0.18937 | 0.15701 | 0.013549 | 0.96244 | 0.96244 | 0.006705
2019-03-20 17:52 | 1553118696.0792 | 156 | 0.20683 | 0.002673 | 4.7e-05 | 98.367 | 96.2758 | 118.316 | 21.0657 | 0.002905 | 0.002676 | 168.0647 | 0.26372 | 8.4038 | 0.003536 | 0.005569 | 0.005508 | 6.7e-05 | 0.001168 | 0.28152 | 1.3336 | 2.548 | 100.58 | 20.73 | 19.055 | 91.5257 | 106.7779 | 85.9958 | 0.004639 | 0.98301 | 0.02359 | 23.9976 | 0.19221 | 0.15688 | 0.01338 | 0.94838 | 0.94838 | 0.006799
2019-03-20 17:52 | 1553118697.0782 | 157 | 0.20683 | 0.002669 | 4.8e-05 | 98.4272 | 96.4425 | 118.5303 | 21.0663 | 0.002882 | 0.002637 | 168.0214 | 0.26333 | 8.4018 | 0.003477 | 0.004791 | 0.004589 | 6.7e-05 | 0.001164 | 0.28456 | 1.356 | 2.5481 | 100.58 | 20.73 | 19.06 | 90.9729 | 105.3445 | 85.0701 | 0.00464 | 0.98555 | 0.023642 | 23.9884 | 0.1822 | 0.15665 | 0.013396 | 0.95152 | 0.95152 | 0.006766
2019-03-20 17:52 | 1553118698.0793 | 158 | 0.20685 | 0.002659 | 4.9e-05 | 98.4505 | 96.4316 | 118.4488 | 21.0667 | 0.002969 | 0.002581 | 167.9792 | 0.26321 | 8.4041 | 0.003501 | 0.005372 | 0.005247 | 6.7e-05 | 0.001162 | 0.27489 | 1.322 | 2.5482 | 100.58 | 20.73 | 19.061 | 93.0612 | 105.9201 | 83.7098 | 0.004641 | 0.98931 | 0.023679 | 23.9346 | 0.18532 | 0.15658 | 0.013253 | 0.9428 | 0.9428 | 0.006826
2019-03-20 17:52 | 1553118699.0784 | 159 | 0.20685 | 0.002661 | 4.9e-05 | 98.4346 | 96.4768 | 118.2958 | 21.0662 | 0.002862 | 0.002689 | 167.9324 | 0.26392 | 8.4024 | 0.003456 | 0.005027 | 0.004855 | 6.5e-05 | 0.001158 | 0.28711 | 1.3207 | 2.5481 | 100.58 | 20.73 | 19.063 | 90.4873 | 104.8401 | 86.3118 | 0.004641 | 0.98215 | 0.023578 | 24.0061 | 0.17969 | 0.15701 | 0.013275 | 0.9403 | 0.9403 | 0.006863
2019-03-20 17:52 | 1553118700.0795 | 160 | 0.20685 | 0.00266 | 5e-05 | 98.5019 | 96.5023 | 118.1268 | 21.066 | 0.002931 | 0.002648 | 167.8703 | 0.26359 | 8.4049 | 0.003576 | 0.005147 | 0.005038 | 7.2e-05 | 0.001153 | 0.27041 | 1.3337 | 2.5481 | 100.58 | 20.73 | 19.067 | 92.1414 | 107.7357 | 85.3257 | 0.004642 | 0.98485 | 0.023593 | 23.9561 | 0.18345 | 0.15681 | 0.013228 | 0.93753 | 0.93753 | 0.006874
2019-03-20 17:52 | 1553118701.0795 | 161 | 0.20685 | 0.002655 | 5e-05 | 98.4741 | 96.457 | 118.1071 | 21.0662 | 0.002986 | 0.0027 | 167.7882 | 0.26438 | 8.403 | 0.003567 | 0.005099 | 0.004987 | 7.1e-05 | 0.001153 | 0.27345 | 1.3813 | 2.548 | 100.58 | 20.73 | 19.069 | 93.4769 | 107.5366 | 86.5889 | 0.004643 | 0.98139 | 0.02355 | 23.9962 | 0.18511 | 0.15728 | 0.013233 | 0.93616 | 0.93616 | 0.006905
2019-03-20 17:52 | 1553118702.0801 | 162 | 0.20686 | 0.002658 | 5e-05 | 98.4899 | 96.7736 | 118.1263 | 21.0662 | 0.002935 | 0.002675 | 167.7322 | 0.26272 | 8.4034 | 0.003513 | 0.00492 | 0.00477 | 7.5e-05 | 0.001151 | 0.26394 | 1.3592 | 2.5481 | 100.58 | 20.73 | 19.07 | 92.2408 | 106.2264 | 85.9855 | 0.004643 | 0.98304 | 0.023571 | 23.978 | 0.15724 | 0.15629 | 0.013181 | 0.93329 | 0.93329 | 0.006883
2019-03-20 17:52 | 1553118703.0786 | 163 | 0.20687 | 0.002657 | 4.9e-05 | 98.477 | 96.8562 | 118.1439 | 21.0658 | 0.002912 | 0.002636 | 167.6763 | 0.26321 | 8.4025 | 0.00357 | 0.005187 | 0.005031 | 6.7e-05 | 0.001154 | 0.24674 | 1.3194 | 2.5481 | 100.58 | 20.73 | 19.074 | 91.6892 | 107.5866 | 85.0337 | 0.004644 | 0.98565 | 0.023604 | 23.948 | 0.14844 | 0.15659 | 0.013135 | 0.93142 | 0.93142 | 0.00691
2019-03-20 17:52 | 1553118704.0786 | 164 | 0.20683 | 0.002646 | 4.9e-05 | 98.4986 | 96.959 | 118.0038 | 21.0668 | 0.002926 | 0.00264 | 167.6128 | 0.263 | 8.4036 | 0.00351 | 0.004688 | 0.004574 | 7.1e-05 | 0.001152 | 0.27184 | 1.3041 | 2.5481 | 100.58 | 20.73 | 19.077 | 92.0344 | 106.1499 | 85.13 | 0.004645 | 0.98539 | 0.023587 | 23.937 | 0.14092 | 0.15646 | 0.013372 | 0.94757 | 0.94757 | 0.006786
2019-03-20 17:52 | 1553118705.0787 | 165 | 0.20692 | 0.002642 | 4.9e-05 | 98.5656 | 96.9306 | 117.8625 | 21.0665 | 0.002909 | 0.00266 | 167.5325 | 0.26292 | 8.4047 | 0.003509 | 0.004696 | 0.00458 | 7.4e-05 | 0.001154 | 0.29734 | 1.3103 | 2.548 | 100.58 | 20.73 | 19.082 | 91.6157 | 106.1292 | 85.629 | 0.004646 | 0.98402 | 0.023557 | 23.9394 | 0.14962 | 0.15641 | 0.01286 | 0.91005 | 0.91005 | 0.007064
2019-03-20 17:52 | 1553118706.0777 | 166 | 0.20688 | 0.002639 | 4.9e-05 | 98.5342 | 97.0427 | 117.8636 | 21.066 | 0.002902 | 0.002609 | 167.4511 | 0.26368 | 8.4027 | 0.003443 | 0.005499 | 0.005336 | 7.2e-05 | 0.001152 | 0.27743 | 1.3311 | 2.5482 | 100.58 | 20.73 | 19.085 | 91.4518 | 104.5182 | 84.3988 | 0.004646 | 0.9874 | 0.023597 | 23.8984 | 0.13643 | 0.15686 | 0.013112 | 0.92951 | 0.92951 | 0.006936
2019-03-20 17:52 | 1553118707.0788 | 167 | 0.20689 | 0.002627 | 5.1e-05 | 98.5239 | 97.2445 | 117.8895 | 21.066 | 0.002979 | 0.002683 | 167.3803 | 0.26257 | 8.4042 | 0.003493 | 0.004366 | 0.004238 | 7.2e-05 | 0.001156 | 0.26394 | 1.3123 | 2.5482 | 100.58 | 20.73 | 19.087 | 93.307 | 105.7312 | 86.1687 | 0.004647 | 0.98254 | 0.023542 | 23.9601 | 0.11691 | 0.1562 | 0.013026 | 0.92109 | 0.92109 | 0.00697
2019-03-20 17:52 | 1553118708.0779 | 168 | 0.20692 | 0.002625 | 5.2e-05 | 98.5333 | 96.8674 | 117.9458 | 21.0677 | 0.00286 | 0.002684 | 167.3183 | 0.26362 | 8.4035 | 0.003468 | 0.005389 | 0.00522 | 7.2e-05 | 0.001153 | 0.26217 | 1.3309 | 2.5481 | 100.58 | 20.73 | 19.088 | 90.4348 | 105.1169 | 86.2064 | 0.004647 | 0.98244 | 0.023546 | 23.9671 | 0.15251 | 0.15683 | 0.012841 | 0.90814 | 0.90814 | 0.007098
2019-03-20 17:52 | 1553118709.0789 | 169 | 0.20686 | 0.002623 | 5.3e-05 | 98.4899 | 96.9795 | 118.0526 | 21.0671 | 0.002955 | 0.002647 | 167.2508 | 0.26229 | 8.4038 | 0.003465 | 0.00515 | 0.005056 | 7.4e-05 | 0.001159 | 0.26969 | 1.301 | 2.5481 | 100.58 | 20.73 | 19.088 | 92.7305 | 105.048 | 85.309 | 0.004647 | 0.9849 | 0.023586 | 23.948 | 0.13823 | 0.15604 | 0.013243 | 0.9381 | 0.9381 | 0.006836
2019-03-20 17:52 | 1553118710.078 | 170 | 0.20685 | 0.002619 | 5.4e-05 | 98.4686 | 97.0875 | 118.0611 | 21.0665 | 0.002969 | 0.002678 | 167.1862 | 0.26381 | 8.402 | 0.003457 | 0.005469 | 0.005312 | 7e-05 | 0.00116 | 0.27776 | 1.3117 | 2.5481 | 100.58 | 20.73 | 19.089 | 93.0593 | 104.8593 | 86.0435 | 0.004647 | 0.98288 | 0.023563 | 23.9734 | 0.12634 | 0.15694 | 0.013295 | 0.94082 | 0.94082 | 0.006856
2019-03-20 17:52 | 1553118711.0791 | 171 | 0.20685 | 0.002617 | 5.5e-05 | 98.4805 | 96.8902 | 118.0338 | 21.0662 | 0.002942 | 0.002646 | 167.1542 | 0.26404 | 8.4026 | 0.003415 | 0.005132 | 0.004972 | 6.6e-05 | 0.001164 | 0.28479 | 1.3246 | 2.5481 | 100.58 | 20.73 | 19.089 | 92.406 | 103.8346 | 85.2807 | 0.004647 | 0.98498 | 0.023585 | 23.9451 | 0.14562 | 0.15708 | 0.013311 | 0.94282 | 0.94282 | 0.006847
2019-03-20 17:52 | 1553118712.0781 | 172 | 0.20687 | 0.002609 | 5.4e-05 | 98.4744 | 97.0076 | 118.0904 | 21.0664 | 0.002903 | 0.002724 | 167.1009 | 0.26348 | 8.4031 | 0.003403 | 0.004983 | 0.004847 | 6.8e-05 | 0.001166 | 0.28639 | 1.3332 | 2.5481 | 100.58 | 20.73 | 19.089 | 91.4749 | 103.5519 | 87.1575 | 0.004647 | 0.97984 | 0.023529 | 24.0135 | 0.13424 | 0.15675 | 0.013217 | 0.934 | 0.934 | 0.006897
2019-03-20 17:52 | 1553118713.0782 | 173 | 0.20686 | 0.002615 | 5.3e-05 | 98.4637 | 96.7512 | 118.1681 | 21.067 | 0.002914 | 0.002691 | 167.0187 | 0.26335 | 8.4029 | 0.003509 | 0.005147 | 0.004988 | 7e-05 | 0.001163 | 0.25393 | 1.3022 | 2.5481 | 100.58 | 20.73 | 19.091 | 91.7357 | 106.1258 | 86.3782 | 0.004648 | 0.98197 | 0.023563 | 23.9954 | 0.15694 | 0.15667 | 0.013279 | 0.93974 | 0.93974 | 0.006852
2019-03-20 17:52 | 1553118714.0796 | 174 | 0.20685 | 0.002616 | 5.2e-05 | 98.4666 | 96.833 | 118.147 | 21.0665 | 0.00293 | 0.002688 | 166.9433 | 0.26306 | 8.4036 | 0.00348 | 0.004606 | 0.00455 | 7.2e-05 | 0.00116 | 0.24298 | 1.312 | 2.5481 | 100.58 | 20.73 | 19.091 | 92.1177 | 105.4193 | 86.3051 | 0.004648 | 0.98217 | 0.023563 | 23.9908 | 0.14964 | 0.15649 | 0.013283 | 0.94013 | 0.94013 | 0.006842
2019-03-20 17:52 | 1553118715.0796 | 175 | 0.20685 | 0.002618 | 5.1e-05 | 98.4168 | 96.9164 | 118.1552 | 21.0656 | 0.002907 | 0.002648 | 166.8745 | 0.26321 | 8.4013 | 0.003477 | 0.005438 | 0.005311 | 7e-05 | 0.001163 | 0.27688 | 1.3344 | 2.548 | 100.58 | 20.73 | 19.091 | 91.5757 | 105.3343 | 85.3395 | 0.004648 | 0.98481 | 0.023596 | 23.9594 | 0.13742 | 0.15659 | 0.013317 | 0.94387 | 0.94387 | 0.006818
2019-03-20 17:52 | 1553118716.0784 | 176 | 0.20684 | 0.002619 | 5e-05 | 98.4356 | 96.602 | 118.1153 | 21.0655 | 0.002954 | 0.002737 | 166.8371 | 0.26277 | 8.403 | 0.003553 | 0.005198 | 0.005073 | 7.2e-05 | 0.001158 | 0.26808 | 1.3296 | 2.5481 | 100.58 | 20.73 | 19.091 | 92.7018 | 107.1878 | 87.4729 | 0.004648 | 0.97899 | 0.023522 | 24.0265 | 0.16818 | 0.15632 | 0.013363 | 0.94427 | 0.94427 | 0.006804
2019-03-20 17:52 | 1553118717.0794 | 177 | 0.20682 | 0.002618 | 5.1e-05 | 98.4298 | 96.426 | 118.0697 | 21.0658 | 0.002899 | 0.002671 | 166.8079 | 0.26306 | 8.4014 | 0.003577 | 0.005517 | 0.005287 | 6.8e-05 | 0.001151 | 0.2855 | 1.3391 | 2.5482 | 100.58 | 20.73 | 19.09 | 91.3796 | 107.7734 | 85.8861 | 0.004648 | 0.98332 | 0.023569 | 23.969 | 0.18396 | 0.15649 | 0.013473 | 0.95386 | 0.95386 | 0.006743
2019-03-20 17:52 | 1553118718.08 | 178 | 0.20681 | 0.002621 | 5.1e-05 | 98.4359 | 96.3729 | 117.5792 | 21.0653 | 0.002965 | 0.002686 | 166.7654 | 0.26254 | 8.4044 | 0.003541 | 0.005178 | 0.005049 | 6.9e-05 | 0.001149 | 0.27184 | 1.3039 | 2.5481 | 100.58 | 20.73 | 19.092 | 92.9574 | 106.8908 | 86.2423 | 0.004648 | 0.98234 | 0.023508 | 23.931 | 0.18945 | 0.15618 | 0.013528 | 0.95535 | 0.95535 | 0.006719
2019-03-20 17:52 | 1553118719.0784 | 179 | 0.20678 | 0.002623 | 5.1e-05 | 98.4007 | 96.4787 | 117.443 | 21.0657 | 0.002948 | 0.002642 | 166.7263 | 0.26372 | 8.4032 | 0.003501 | 0.005137 | 0.004997 | 6.6e-05 | 0.00115 | 0.29015 | 1.3052 | 2.5481 | 100.58 | 20.73 | 19.095 | 92.5626 | 105.9152 | 85.1919 | 0.004649 | 0.98522 | 0.023529 | 23.8822 | 0.17643 | 0.15689 | 0.013724 | 0.97002 | 0.97002 | 0.006647
2019-03-20 17:52 | 1553118720.0786 | 180 | 0.20675 | 0.00262 | 5.1e-05 | 98.3754 | 96.4189 | 117.3778 | 21.0652 | 0.002915 | 0.002662 | 166.6891 | 0.26322 | 8.4021 | 0.003514 | 0.005642 | 0.005563 | 6.6e-05 | 0.00115 | 0.29518 | 1.3168 | 2.5481 | 100.58 | 20.73 | 19.099 | 91.7482 | 106.242 | 85.6661 | 0.00465 | 0.98392 | 0.023507 | 23.8913 | 0.17969 | 0.15659 | 0.013939 | 0.98433 | 0.98433 | 0.006538
2019-03-20 17:52 | 1553118721.0795 | 181 | 0.20672 | 0.002632 | 5.1e-05 | 98.3502 | 96.1405 | 117.2496 | 21.0662 | 0.002968 | 0.002656 | 166.6286 | 0.26268 | 8.3998 | 0.003515 | 0.005251 | 0.005227 | 6.8e-05 | 0.001156 | 0.28744 | 1.3389 | 2.5481 | 100.58 | 20.73 | 19.11 | 93.027 | 106.2752 | 85.5273 | 0.004653 | 0.9843 | 0.023499 | 23.8737 | 0.20326 | 0.15627 | 0.014075 | 0.99356 | 0.99356 | 0.006464
2019-03-20 17:52 | 1553118722.079 | 182 | 0.20666 | 0.002641 | 5.1e-05 | 98.345 | 95.7424 | 117.2637 | 21.0654 | 0.002998 | 0.002698 | 166.5814 | 0.26312 | 8.4036 | 0.003572 | 0.005144 | 0.004939 | 6.4e-05 | 0.001151 | 0.26969 | 1.3444 | 2.5482 | 100.58 | 20.73 | 19.118 | 93.7694 | 107.6407 | 86.5225 | 0.004655 | 0.98157 | 0.023468 | 23.9082 | 0.2399 | 0.15653 | 0.014416 | 1.0164 | 1.0164 | 0.00633
2019-03-20 17:52 | 1553118723.0787 | 183 | 0.2066 | 0.002662 | 5.1e-05 | 98.3421 | 95.6912 | 117.417 | 21.065 | 0.002911 | 0.002732 | 166.5348 | 0.26287 | 8.4023 | 0.003581 | 0.004951 | 0.004873 | 6.4e-05 | 0.00115 | 0.2912 | 1.3523 | 2.5482 | 100.58 | 20.73 | 19.136 | 91.654 | 107.8748 | 87.3434 | 0.004659 | 0.97934 | 0.023456 | 23.9511 | 0.24442 | 0.15638 | 0.014731 | 1.0381 | 1.0381 | 0.006191
2019-03-20 17:52 | 1553118724.0788 | 184 | 0.20662 | 0.002672 | 5e-05 | 98.3832 | 95.3974 | 117.3759 | 21.0661 | 0.00297 | 0.002666 | 166.4692 | 0.26378 | 8.4031 | 0.003574 | 0.00473 | 0.00462 | 6.4e-05 | 0.001155 | 0.28279 | 1.4011 | 2.548 | 100.58 | 20.73 | 19.144 | 93.0833 | 107.7018 | 85.773 | 0.004661 | 0.98362 | 0.023503 | 23.8947 | 0.27565 | 0.15692 | 0.014605 | 1.0313 | 1.0313 | 0.006254
2019-03-20 17:52 | 1553118725.0778 | 185 | 0.20661 | 0.002692 | 5e-05 | 98.3809 | 95.4501 | 117.1975 | 21.0668 | 0.002971 | 0.002671 | 166.4302 | 0.26303 | 8.4028 | 0.003571 | 0.004646 | 0.004516 | 6.7e-05 | 0.001155 | 0.26576 | 1.3995 | 2.5482 | 100.58 | 20.73 | 19.153 | 93.1018 | 107.6316 | 85.8947 | 0.004663 | 0.98329 | 0.023482 | 23.8806 | 0.27051 | 0.15648 | 0.014677 | 1.0355 | 1.0355 | 0.006211
2019-03-20 17:52 | 1553118726.0788 | 186 | 0.20654 | 0.002704 | 4.8e-05 | 98.4999 | 95.3421 | 117.1184 | 21.0664 | 0.002925 | 0.002715 | 166.3558 | 0.26374 | 8.4038 | 0.003566 | 0.004876 | 0.004719 | 6.4e-05 | 0.001151 | 0.27527 | 1.3426 | 2.5479 | 100.58 | 20.73 | 19.17 | 91.9947 | 107.5101 | 86.9481 | 0.004667 | 0.98041 | 0.023439 | 23.9075 | 0.29145 | 0.1569 | 0.015054 | 1.0603 | 1.0603 | 0.006082
2019-03-20 17:52 | 1553118727.078 | 187 | 0.20656 | 0.002718 | 4.6e-05 | 98.4715 | 95.2502 | 117.3229 | 21.0661 | 0.002974 | 0.002689 | 166.2617 | 0.26326 | 8.4007 | 0.003502 | 0.005211 | 0.005033 | 6.3e-05 | 0.001154 | 0.26665 | 1.324 | 2.5482 | 100.58 | 20.73 | 19.178 | 93.1725 | 105.958 | 86.3061 | 0.004669 | 0.98216 | 0.023481 | 23.907 | 0.2975 | 0.15661 | 0.01495 | 1.0549 | 1.0549 | 0.006102
2019-03-20 17:52 | 1553118728.078 | 188 | 0.20655 | 0.002729 | 4.5e-05 | 98.4378 | 95.2711 | 117.5933 | 21.0658 | 0.003012 | 0.002719 | 166.1868 | 0.26344 | 8.4013 | 0.003479 | 0.005118 | 0.005005 | 6.4e-05 | 0.001153 | 0.27113 | 1.2915 | 2.5482 | 100.58 | 20.73 | 19.192 | 94.1 | 105.3833 | 87.0499 | 0.004673 | 0.98014 | 0.023483 | 23.9593 | 0.29247 | 0.15672 | 0.014964 | 1.0561 | 1.0561 | 0.006099
2019-03-20 17:52 | 1553118729.0792 | 189 | 0.20648 | 0.002733 | 4.4e-05 | 98.4362 | 94.746 | 117.6794 | 21.0664 | 0.002975 | 0.00272 | 166.1322 | 0.26409 | 8.4013 | 0.003463 | 0.005472 | 0.005291 | 6.3e-05 | 0.001157 | 0.28783 | 1.2931 | 2.548 | 100.58 | 20.73 | 19.198 | 93.2087 | 105.0102 | 87.0661 | 0.004674 | 0.98009 | 0.023491 | 23.9686 | 0.34176 | 0.1571 | 0.01538 | 1.0859 | 1.0859 | 0.005946
2019-03-20 17:52 | 1553118730.0781 | 190 | 0.20652 | 0.002752 | 4.3e-05 | 98.5064 | 94.6578 | 117.804 | 21.0658 | 0.002936 | 0.002722 | 166.087 | 0.26334 | 8.4033 | 0.003517 | 0.00517 | 0.00502 | 6.5e-05 | 0.001159 | 0.28102 | 1.3376 | 2.5481 | 100.58 | 20.73 | 19.203 | 92.2569 | 106.3229 | 87.1117 | 0.004675 | 0.97997 | 0.023502 | 23.9828 | 0.35647 | 0.15666 | 0.015121 | 1.0682 | 1.0682 | 0.006028
2019-03-20 17:52 | 1553118731.0791 | 191 | 0.20653 | 0.002755 | 4.2e-05 | 98.325 | 94.5128 | 117.6212 | 21.0665 | 0.002999 | 0.002723 | 166.0251 | 0.26403 | 8.4058 | 0.003599 | 0.004528 | 0.004436 | 6.4e-05 | 0.001157 | 0.26272 | 1.2956 | 2.548 | 100.58 | 20.73 | 19.208 | 93.7837 | 108.2985 | 87.1465 | 0.004676 | 0.97988 | 0.023483 | 23.9654 | 0.35369 | 0.15707 | 0.015085 | 1.0648 | 1.0648 | 0.006063
2019-03-20 17:52 | 1553118732.0784 | 192 | 0.20654 | 0.002764 | 4.1e-05 | 98.3922 | 94.4264 | 117.5003 | 21.0668 | 0.002961 | 0.002723 | 165.9523 | 0.26335 | 8.4035 | 0.003666 | 0.004873 | 0.004793 | 6.3e-05 | 0.001155 | 0.26808 | 1.3397 | 2.548 | 100.58 | 20.74 | 19.208 | 92.8672 | 109.9393 | 87.1422 | 0.004679 | 0.97989 | 0.023471 | 23.9529 | 0.36799 | 0.15667 | 0.014995 | 1.0579 | 1.0579 | 0.006087
2019-03-20 17:52 | 1553118733.0786 | 193 | 0.20653 | 0.002774 | 4e-05 | 98.4029 | 94.4597 | 117.4494 | 21.0679 | 0.002926 | 0.002743 | 165.9125 | 0.26354 | 8.4041 | 0.003691 | 0.005342 | 0.005207 | 6.3e-05 | 0.001144 | 0.27743 | 1.2958 | 2.548 | 100.58 | 20.74 | 19.207 | 92.033 | 110.5424 | 87.6275 | 0.004679 | 0.97857 | 0.02345 | 23.9638 | 0.36581 | 0.15678 | 0.015056 | 1.0614 | 1.0614 | 0.006071
2019-03-20 17:52 | 1553118734.0782 | 194 | 0.20655 | 0.002782 | 3.9e-05 | 98.3932 | 94.4933 | 117.6515 | 21.0668 | 0.002962 | 0.002764 | 165.8877 | 0.26439 | 8.4016 | 0.003722 | 0.004769 | 0.004691 | 6.3e-05 | 0.001147 | 0.27599 | 1.3261 | 2.5482 | 100.58 | 20.74 | 19.203 | 92.8821 | 111.2986 | 88.1237 | 0.004678 | 0.97722 | 0.023454 | 24.0009 | 0.36174 | 0.15728 | 0.014946 | 1.0539 | 1.0539 | 0.006134
2019-03-20 17:52 | 1553118735.0797 | 195 | 0.20652 | 0.002795 | 3.8e-05 | 98.2671 | 94.0994 | 117.9594 | 21.0664 | 0.002976 | 0.002758 | 165.8245 | 0.26359 | 8.4025 | 0.003598 | 0.00539 | 0.00521 | 6.2e-05 | 0.001141 | 0.27671 | 1.3454 | 2.5482 | 100.58 | 20.74 | 19.201 | 93.2251 | 108.2812 | 87.9751 | 0.004678 | 0.97763 | 0.02349 | 24.0274 | 0.38763 | 0.15681 | 0.015116 | 1.0675 | 1.0675 | 0.006037
2019-03-20 17:52 | 1553118736.0793 | 196 | 0.20652 | 0.002798 | 3.7e-05 | 98.2981 | 93.9619 | 118.1329 | 21.0681 | 0.00295 | 0.002682 | 165.7917 | 0.2641 | 8.4034 | 0.003582 | 0.005012 | 0.004913 | 6.5e-05 | 0.001136 | 0.29159 | 1.3271 | 2.548 | 100.58 | 20.74 | 19.198 | 92.6103 | 107.889 | 86.1606 | 0.004677 | 0.98256 | 0.023566 | 23.9846 | 0.40354 | 0.15711 | 0.015079 | 1.0684 | 1.0684 | 0.006044
2019-03-20 17:52 | 1553118737.0787 | 197 | 0.20654 | 0.002804 | 3.7e-05 | 98.2364 | 93.6389 | 117.8688 | 21.0668 | 0.002983 | 0.002716 | 165.7455 | 0.26314 | 8.4031 | 0.003506 | 0.004797 | 0.004586 | 6.5e-05 | 0.001134 | 0.26985 | 1.2815 | 2.5482 | 100.58 | 20.74 | 19.196 | 93.4092 | 106.0414 | 86.9603 | 0.004676 | 0.98038 | 0.023514 | 23.9844 | 0.42871 | 0.15654 | 0.014949 | 1.0569 | 1.0569 | 0.006087
2019-03-20 17:52 | 1553118738.0789 | 198 | 0.20654 | 0.002801 | 3.6e-05 | 98.2412 | 93.8053 | 117.7592 | 21.0665 | 0.003008 | 0.002727 | 165.6836 | 0.26395 | 8.4015 | 0.003531 | 0.005062 | 0.004994 | 6.2e-05 | 0.001132 | 0.27024 | 1.3275 | 2.5482 | 100.58 | 20.74 | 19.194 | 93.9991 | 106.662 | 87.2334 | 0.004676 | 0.97964 | 0.023494 | 23.9823 | 0.41328 | 0.15702 | 0.014989 | 1.0589 | 1.0589 | 0.006095
2019-03-20 17:52 | 1553118739.0776 | 199 | 0.20656 | 0.002808 | 3.6e-05 | 98.182 | 93.587 | 117.5428 | 21.0675 | 0.002967 | 0.002707 | 165.6175 | 0.26394 | 8.4017 | 0.003548 | 0.004854 | 0.004695 | 6.2e-05 | 0.001135 | 0.28152 | 1.3576 | 2.5482 | 100.58 | 20.74 | 19.193 | 93.007 | 107.0568 | 86.7471 | 0.004676 | 0.98096 | 0.023488 | 23.9441 | 0.42873 | 0.15702 | 0.014847 | 1.0486 | 1.0486 | 0.006154
2019-03-20 17:52 | 1553118740.0776 | 200 | 0.20655 | 0.0028 | 3.5e-05 | 98.1335 | 93.6094 | 117.4721 | 21.068 | 0.003007 | 0.00273 | 165.5443 | 0.26356 | 8.4023 | 0.003439 | 0.004905 | 0.004815 | 6.4e-05 | 0.001141 | 0.28512 | 1.3578 | 2.5482 | 100.58 | 20.74 | 19.194 | 93.9895 | 104.4322 | 87.2989 | 0.004676 | 0.97946 | 0.023463 | 23.9552 | 0.42216 | 0.15679 | 0.014877 | 1.0496 | 1.0496 | 0.00614
2019-03-20 17:52 | 1553118741.0777 | 201 | 0.2066 | 0.002806 | 3.5e-05 | 98.1102 | 93.6647 | 117.7533 | 21.0663 | 0.003008 | 0.002746 | 165.4949 | 0.26319 | 8.403 | 0.003525 | 0.005299 | 0.005151 | 6.5e-05 | 0.001141 | 0.27129 | 1.351 | 2.5481 | 100.58 | 20.74 | 19.198 | 93.9972 | 106.4996 | 87.6921 | 0.004677 | 0.97839 | 0.023478 | 23.997 | 0.41476 | 0.15657 | 0.014585 | 1.0297 | 1.0297 | 0.006249
2019-03-20 17:52 | 1553118742.0778 | 202 | 0.20659 | 0.002801 | 3.4e-05 | 98.0704 | 93.6004 | 117.8477 | 21.0675 | 0.003046 | 0.002752 | 165.4711 | 0.26406 | 8.4049 | 0.003507 | 0.004886 | 0.004801 | 6.1e-05 | 0.00114 | 0.24674 | 1.3331 | 2.5481 | 100.58 | 20.74 | 19.202 | 94.9127 | 106.0671 | 87.837 | 0.004678 | 0.978 | 0.023483 | 24.0114 | 0.41727 | 0.15709 | 0.01467 | 1.0359 | 1.0359 | 0.006232
2019-03-20 17:52 | 1553118743.0778 | 203 | 0.20667 | 0.002801 | 3.3e-05 | 98.0792 | 93.6584 | 117.8425 | 21.0674 | 0.002986 | 0.002743 | 165.4121 | 0.26369 | 8.4041 | 0.00348 | 0.005318 | 0.005208 | 6.1e-05 | 0.001145 | 0.26128 | 1.3289 | 2.5481 | 100.58 | 20.74 | 19.206 | 93.4718 | 105.4093 | 87.6285 | 0.004679 | 0.97857 | 0.023489 | 24.0039 | 0.41253 | 0.15687 | 0.014173 | 1.0012 | 1.0012 | 0.00644
2019-03-20 17:52 | 1553118744.0787 | 204 | 0.20671 | 0.002792 | 3.3e-05 | 98.0543 | 93.9185 | 117.7 | 21.0672 | 0.003025 | 0.002718 | 165.3448 | 0.26378 | 8.4024 | 0.003415 | 0.004885 | 0.004758 | 6.5e-05 | 0.001145 | 0.27417 | 1.3573 | 2.5482 | 100.58 | 20.74 | 19.215 | 94.4168 | 103.8339 | 87.0076 | 0.004681 | 0.98025 | 0.023495 | 23.9688 | 0.38545 | 0.15692 | 0.013985 | 0.98812 | 0.98812 | 0.006527
2019-03-20 17:52 | 1553118745.0779 | 205 | 0.20677 | 0.002781 | 3.2e-05 | 98.1015 | 93.9735 | 117.7834 | 21.0669 | 0.003023 | 0.002723 | 165.2966 | 0.26374 | 8.4025 | 0.003532 | 0.004766 | 0.004594 | 6.3e-05 | 0.001149 | 0.27417 | 1.3367 | 2.5482 | 100.58 | 20.74 | 19.221 | 94.3555 | 106.6687 | 87.1337 | 0.004682 | 0.97991 | 0.0235 | 23.9815 | 0.38453 | 0.1569 | 0.013632 | 0.96344 | 0.96344 | 0.006693
2019-03-20 17:52 | 1553118746.0789 | 206 | 0.20675 | 0.00277 | 3.2e-05 | 98.0455 | 94.179 | 117.6454 | 21.0671 | 0.003002 | 0.002731 | 165.2326 | 0.26414 | 8.3998 | 0.00354 | 0.004717 | 0.004616 | 6.3e-05 | 0.001148 | 0.25929 | 1.3119 | 2.5481 | 100.58 | 20.74 | 19.226 | 93.8484 | 106.8814 | 87.3189 | 0.004684 | 0.97941 | 0.02348 | 23.9736 | 0.35988 | 0.15714 | 0.013778 | 0.97294 | 0.97294 | 0.006638
2019-03-20 17:52 | 1553118747.078 | 207 | 0.20683 | 0.00275 | 3.1e-05 | 98.1154 | 94.0318 | 117.3874 | 21.0676 | 0.003042 | 0.002719 | 165.1883 | 0.26297 | 8.4043 | 0.00355 | 0.004922 | 0.004718 | 6.4e-05 | 0.001147 | 0.28064 | 1.3616 | 2.5482 | 100.58 | 20.74 | 19.236 | 94.8218 | 107.1088 | 87.0484 | 0.004686 | 0.98014 | 0.023463 | 23.9383 | 0.38025 | 0.15644 | 0.013308 | 0.9391 | 0.9391 | 0.006847
2019-03-20 17:52 | 1553118748.0781 | 208 | 0.20672 | 0.002735 | 3e-05 | 98.1801 | 93.6625 | 117.2238 | 21.067 | 0.003026 | 0.002731 | 165.161 | 0.26275 | 8.4029 | 0.00351 | 0.004979 | 0.004896 | 6.4e-05 | 0.001147 | 0.26681 | 1.3299 | 2.5481 | 100.58 | 20.74 | 19.239 | 94.448 | 106.1438 | 87.3282 | 0.004687 | 0.97938 | 0.023437 | 23.9309 | 0.42134 | 0.15631 | 0.013996 | 0.98664 | 0.98664 | 0.006511
2019-03-20 17:52 | 1553118749.0781 | 209 | 0.20695 | 0.002717 | 3e-05 | 98.1484 | 94.1555 | 117.6104 | 21.0665 | 0.002973 | 0.002784 | 165.1276 | 0.26337 | 8.4027 | 0.003485 | 0.004423 | 0.004272 | 6.7e-05 | 0.001151 | 0.27848 | 1.3306 | 2.548 | 100.58 | 20.74 | 19.247 | 93.1492 | 105.5454 | 88.6029 | 0.004689 | 0.97593 | 0.023435 | 24.0127 | 0.3715 | 0.15668 | 0.012626 | 0.88993 | 0.88993 | 0.007236
2019-03-20 17:52 | 1553118750.0791 | 210 | 0.20694 | 0.002691 | 2.9e-05 | 98.1672 | 94.0239 | 117.712 | 21.0674 | 0.002976 | 0.002724 | 165.0933 | 0.26339 | 8.4035 | 0.003452 | 0.004703 | 0.004584 | 6.6e-05 | 0.001155 | 0.24762 | 1.312 | 2.5481 | 100.58 | 20.74 | 19.252 | 93.2379 | 104.7477 | 87.163 | 0.00469 | 0.97983 | 0.023492 | 23.9752 | 0.38571 | 0.15669 | 0.012705 | 0.89774 | 0.89774 | 0.007174
2019-03-20 17:52 | 1553118751.0792 | 211 | 0.207 | 0.002668 | 2.9e-05 | 98.1306 | 94.0856 | 117.8137 | 21.0667 | 0.002986 | 0.002736 | 165.0313 | 0.26399 | 8.4015 | 0.003525 | 0.004706 | 0.004473 | 6e-05 | 0.001149 | 0.24331 | 1.3424 | 2.5481 | 100.58 | 20.74 | 19.255 | 93.4607 | 106.4969 | 87.4566 | 0.004691 | 0.97903 | 0.023492 | 23.9953 | 0.37652 | 0.15705 | 0.012391 | 0.87554 | 0.87554 | 0.007372
2019-03-20 17:53 | 1553118752.0792 | 212 | 0.20698 | 0.002643 | 2.9e-05 | 98.1477 | 93.9581 | 117.578 | 21.0665 | 0.003032 | 0.002734 | 164.9558 | 0.26387 | 8.4009 | 0.003505 | 0.004736 | 0.004641 | 6.2e-05 | 0.001152 | 0.27383 | 1.3132 | 2.5481 | 100.58 | 20.74 | 19.259 | 94.5758 | 106.0229 | 87.4011 | 0.004692 | 0.97918 | 0.02347 | 23.9694 | 0.3902 | 0.15697 | 0.01255 | 0.88595 | 0.88595 | 0.007282
2019-03-20 17:53 | 1553118753.0783 | 213 | 0.20697 | 0.002617 | 2.9e-05 | 98.1258 | 94.3539 | 117.3907 | 21.0669 | 0.003028 | 0.002713 | 164.901 | 0.26391 | 8.4039 | 0.003474 | 0.004813 | 0.004632 | 6.5e-05 | 0.001154 | 0.27632 | 1.3288 | 2.5481 | 100.58 | 20.74 | 19.26 | 94.4956 | 105.264 | 86.8956 | 0.004692 | 0.98056 | 0.023468 | 23.9335 | 0.3506 | 0.157 | 0.01263 | 0.89157 | 0.89157 | 0.007237
2019-03-20 17:53 | 1553118754.0784 | 214 | 0.20701 | 0.00259 | 2.9e-05 | 98.1756 | 94.6066 | 117.5132 | 21.0678 | 0.002949 | 0.002723 | 164.8489 | 0.26319 | 8.3997 | 0.003524 | 0.004968 | 0.004833 | 6.2e-05 | 0.001149 | 0.27129 | 1.2984 | 2.5481 | 100.58 | 20.74 | 19.26 | 92.5818 | 106.472 | 87.1394 | 0.004692 | 0.97989 | 0.023473 | 23.9541 | 0.33122 | 0.15657 | 0.012434 | 0.87782 | 0.87782 | 0.007331
2019-03-20 17:53 | 1553118755.0774 | 215 | 0.20704 | 0.002564 | 2.9e-05 | 98.2276 | 94.6301 | 117.2351 | 21.0671 | 0.002949 | 0.002675 | 164.7899 | 0.26338 | 8.4047 | 0.003468 | 0.00476 | 0.004707 | 6.5e-05 | 0.00115 | 0.27041 | 1.3281 | 2.5481 | 100.58 | 20.74 | 19.259 | 92.5696 | 105.1368 | 85.9879 | 0.004692 | 0.98304 | 0.023482 | 23.8875 | 0.33374 | 0.15668 | 0.012281 | 0.86736 | 0.86736 | 0.007424
2019-03-20 17:53 | 1553118756.0788 | 216 | 0.20703 | 0.002546 | 3e-05 | 98.2092 | 94.8682 | 117.2729 | 21.0675 | 0.003033 | 0.002729 | 164.7581 | 0.26271 | 8.4028 | 0.003524 | 0.005224 | 0.005008 | 6.6e-05 | 0.001149 | 0.2453 | 1.3084 | 2.5482 | 100.58 | 20.74 | 19.259 | 94.612 | 106.4882 | 87.2897 | 0.004692 | 0.97949 | 0.023444 | 23.9346 | 0.30958 | 0.15629 | 0.012374 | 0.87251 | 0.87251 | 0.007362
2019-03-20 17:53 | 1553118757.0775 | 217 | 0.20708 | 0.002526 | 2.9e-05 | 98.2147 | 95.0641 | 117.6123 | 21.0678 | 0.00304 | 0.002734 | 164.7049 | 0.26429 | 8.4 | 0.003502 | 0.004921 | 0.00481 | 6.7e-05 | 0.001154 | 0.27273 | 1.3794 | 2.5482 | 100.58 | 20.74 | 19.254 | 94.7711 | 105.9439 | 87.4118 | 0.004691 | 0.97915 | 0.023474 | 23.9733 | 0.29163 | 0.15723 | 0.01206 | 0.85144 | 0.85144 | 0.007589
2019-03-20 17:53 | 1553118758.0786 | 218 | 0.20706 | 0.002509 | 2.9e-05 | 98.2713 | 95.1874 | 117.7867 | 21.0676 | 0.003018 | 0.002663 | 164.6427 | 0.26342 | 8.4037 | 0.003457 | 0.005291 | 0.005217 | 6.3e-05 | 0.001152 | 0.26488 | 1.3476 | 2.5482 | 100.58 | 20.74 | 19.251 | 94.2505 | 104.8613 | 85.7042 | 0.00469 | 0.98381 | 0.023547 | 23.9342 | 0.28519 | 0.15671 | 0.012222 | 0.86556 | 0.86556 | 0.007441
2019-03-20 17:53 | 1553118759.0786 | 219 | 0.20714 | 0.002495 | 2.9e-05 | 98.2845 | 95.48 | 117.8221 | 21.0671 | 0.002962 | 0.002672 | 164.5891 | 0.26388 | 8.4018 | 0.003486 | 0.004382 | 0.004225 | 6.3e-05 | 0.001147 | 0.28335 | 1.3068 | 2.5481 | 100.58 | 20.74 | 19.248 | 92.8832 | 105.5545 | 85.8981 | 0.004689 | 0.98328 | 0.023544 | 23.9442 | 0.25894 | 0.15699 | 0.011775 | 0.8338 | 0.8338 | 0.007738
2019-03-20 17:53 | 1553118760.0787 | 220 | 0.20712 | 0.002477 | 2.8e-05 | 98.3356 | 95.5387 | 118.0697 | 21.0673 | 0.003038 | 0.002699 | 164.5213 | 0.26292 | 8.4038 | 0.003512 | 0.00502 | 0.004875 | 6.8e-05 | 0.001154 | 0.27383 | 1.3327 | 2.5482 | 100.58 | 20.74 | 19.242 | 94.7378 | 106.204 | 86.5486 | 0.004688 | 0.9815 | 0.023547 | 23.9911 | 0.2581 | 0.15641 | 0.011913 | 0.84375 | 0.84375 | 0.007619
2019-03-20 17:53 | 1553118761.0777 | 221 | 0.20714 | 0.002465 | 2.8e-05 | 98.4831 | 96.0194 | 117.9125 | 21.067 | 0.002966 | 0.00276 | 164.4813 | 0.26371 | 8.4026 | 0.003528 | 0.005262 | 0.005229 | 6.4e-05 | 0.001153 | 0.28385 | 1.3894 | 2.5482 | 100.58 | 20.74 | 19.239 | 92.9817 | 106.58 | 88.0318 | 0.004687 | 0.97747 | 0.023483 | 24.0245 | 0.22661 | 0.15688 | 0.011784 | 0.83234 | 0.83234 | 0.007747
2019-03-20 17:53 | 1553118762.078 | 222 | 0.20716 | 0.002446 | 2.9e-05 | 98.5585 | 96.0848 | 117.7939 | 21.0668 | 0.002975 | 0.002716 | 164.4428 | 0.26377 | 8.4031 | 0.003496 | 0.005215 | 0.005017 | 6.6e-05 | 0.001157 | 0.25912 | 1.3606 | 2.5481 | 100.58 | 20.75 | 19.234 | 93.2091 | 105.8069 | 86.9791 | 0.004689 | 0.98033 | 0.023506 | 23.9774 | 0.22736 | 0.15692 | 0.011676 | 0.82543 | 0.82543 | 0.007813
2019-03-20 17:53 | 1553118763.0783 | 223 | 0.20716 | 0.002434 | 2.9e-05 | 98.6393 | 96.105 | 117.9033 | 21.0674 | 0.003009 | 0.002737 | 164.4037 | 0.26316 | 8.4017 | 0.003543 | 0.004946 | 0.00483 | 6.9e-05 | 0.001151 | 0.26897 | 1.3877 | 2.5482 | 100.58 | 20.75 | 19.233 | 94.0296 | 106.9543 | 87.473 | 0.004688 | 0.97899 | 0.023501 | 24.005 | 0.23282 | 0.15655 | 0.011715 | 0.82805 | 0.82805 | 0.00777
2019-03-20 17:53 | 1553118764.0793 | 224 | 0.20714 | 0.002419 | 2.9e-05 | 98.7063 | 96.2918 | 117.6043 | 21.0675 | 0.003012 | 0.002728 | 164.3685 | 0.26298 | 8.4039 | 0.003481 | 0.004916 | 0.004787 | 6.8e-05 | 0.001152 | 0.25984 | 1.4142 | 2.5481 | 100.58 | 20.74 | 19.233 | 94.0908 | 105.4417 | 87.2483 | 0.004685 | 0.9796 | 0.023478 | 23.967 | 0.22151 | 0.15645 | 0.01182 | 0.83469 | 0.83469 | 0.007703
2019-03-20 17:53 | 1553118765.0795 | 225 | 0.20714 | 0.002407 | 3e-05 | 98.7723 | 96.457 | 117.8423 | 21.0669 | 0.002997 | 0.002735 | 164.3261 | 0.2633 | 8.4041 | 0.003503 | 0.005205 | 0.005075 | 7e-05 | 0.001149 | 0.26377 | 1.4065 | 2.548 | 100.58 | 20.74 | 19.234 | 93.7373 | 105.9627 | 87.4179 | 0.004686 | 0.97914 | 0.023496 | 23.9969 | 0.21216 | 0.15664 | 0.011892 | 0.84038 | 0.84038 | 0.007661
2019-03-20 17:53 | 1553118766.0791 | 226 | 0.20715 | 0.002394 | 3e-05 | 98.9414 | 96.6988 | 118.2873 | 21.0672 | 0.002944 | 0.002698 | 164.2673 | 0.26262 | 8.4014 | 0.003394 | 0.004818 | 0.004673 | 6.8e-05 | 0.001143 | 0.27759 | 1.3793 | 2.5481 | 100.58 | 20.75 | 19.24 | 92.4578 | 103.324 | 86.5329 | 0.00469 | 0.98155 | 0.02357 | 24.0127 | 0.20507 | 0.15624 | 0.011843 | 0.83948 | 0.83948 | 0.007649
2019-03-20 17:53 | 1553118767.078 | 227 | 0.20716 | 0.002387 | 3e-05 | 99.0041 | 96.7792 | 118.4444 | 21.0673 | 0.003029 | 0.002748 | 164.2438 | 0.26266 | 8.4031 | 0.003487 | 0.004405 | 0.004249 | 6.6e-05 | 0.001144 | 0.2672 | 1.3884 | 2.5483 | 100.58 | 20.75 | 19.245 | 94.5202 | 105.5816 | 87.7362 | 0.004691 | 0.97827 | 0.023546 | 24.0688 | 0.20331 | 0.15626 | 0.011785 | 0.83457 | 0.83457 | 0.007695
2019-03-20 17:53 | 1553118768.0775 | 228 | 0.20717 | 0.002385 | 3e-05 | 99.0915 | 96.8427 | 118.681 | 21.067 | 0.002995 | 0.002715 | 164.1742 | 0.26349 | 8.4023 | 0.003448 | 0.005588 | 0.005476 | 6.7e-05 | 0.001141 | 0.24729 | 1.3514 | 2.5481 | 100.58 | 20.75 | 19.251 | 93.6952 | 104.6278 | 86.9528 | 0.004693 | 0.9804 | 0.023595 | 24.0666 | 0.20532 | 0.15675 | 0.011702 | 0.83039 | 0.83039 | 0.007758
2019-03-20 17:53 | 1553118769.078 | 229 | 0.20718 | 0.002374 | 3e-05 | 99.1801 | 96.9597 | 118.5429 | 21.0668 | 0.003 | 0.002697 | 164.1294 | 0.26345 | 8.4047 | 0.003464 | 0.004854 | 0.004727 | 6.5e-05 | 0.001143 | 0.2724 | 1.3924 | 2.5482 | 100.58 | 20.75 | 19.269 | 93.8091 | 105.0281 | 86.5092 | 0.004697 | 0.98161 | 0.023596 | 24.0378 | 0.20252 | 0.15673 | 0.011685 | 0.82915 | 0.82915 | 0.007769
2019-03-20 17:53 | 1553118770.0786 | 230 | 0.2072 | 0.002368 | 3e-05 | 99.2829 | 97.1189 | 118.3439 | 21.0667 | 0.002973 | 0.002697 | 164.0903 | 0.26366 | 8.4033 | 0.003457 | 0.005014 | 0.004925 | 6.4e-05 | 0.001143 | 0.27903 | 1.3939 | 2.5481 | 100.58 | 20.75 | 19.28 | 93.1699 | 104.8611 | 86.5018 | 0.0047 | 0.98163 | 0.023576 | 24.0174 | 0.19711 | 0.15685 | 0.011573 | 0.82059 | 0.82059 | 0.007856
2019-03-20 17:53 | 1553118771.0773 | 231 | 0.2072 | 0.002357 | 3e-05 | 99.3864 | 96.9496 | 118.2106 | 21.0678 | 0.003002 | 0.002709 | 164.043 | 0.26266 | 8.4039 | 0.003534 | 0.00537 | 0.005174 | 6.5e-05 | 0.00115 | 0.27599 | 1.3481 | 2.548 | 100.58 | 20.75 | 19.305 | 93.854 | 106.7341 | 86.8114 | 0.004706 | 0.98079 | 0.023553 | 24.0142 | 0.22204 | 0.15626 | 0.011545 | 0.81774 | 0.81774 | 0.007854
2019-03-20 17:53 | 1553118772.0773 | 232 | 0.20724 | 0.002345 | 3.1e-05 | 99.5297 | 96.9881 | 118.1824 | 21.067 | 0.002943 | 0.002721 | 163.9909 | 0.26355 | 8.4016 | 0.003392 | 0.00493 | 0.004851 | 6.8e-05 | 0.001152 | 0.27057 | 1.2994 | 2.548 | 100.58 | 20.75 | 19.318 | 92.4426 | 103.271 | 87.0767 | 0.004709 | 0.98006 | 0.023541 | 24.0201 | 0.23137 | 0.15679 | 0.011323 | 0.80158 | 0.80158 | 0.008039
2019-03-20 17:53 | 1553118773.0787 | 233 | 0.2073 | 0.002336 | 3.1e-05 | 99.5552 | 97.1735 | 117.9036 | 21.067 | 0.003047 | 0.002743 | 163.9485 | 0.26349 | 8.4054 | 0.003396 | 0.005186 | 0.005058 | 6.6e-05 | 0.001151 | 0.26825 | 1.3267 | 2.5481 | 100.58 | 20.75 | 19.331 | 94.952 | 103.3725 | 87.625 | 0.004712 | 0.97857 | 0.023496 | 24.01 | 0.21659 | 0.15675 | 0.010981 | 0.77582 | 0.77582 | 0.008304
2019-03-20 17:53 | 1553118774.0788 | 234 | 0.20728 | 0.002325 | 3.2e-05 | 99.6565 | 97.1926 | 118.1362 | 21.0672 | 0.00292 | 0.002743 | 163.8893 | 0.26379 | 8.405 | 0.0033 | 0.004211 | 0.004055 | 6.4e-05 | 0.001153 | 0.28119 | 1.3416 | 2.5481 | 100.58 | 20.75 | 19.355 | 91.8721 | 101.0515 | 87.6119 | 0.004718 | 0.97861 | 0.023519 | 24.0333 | 0.22392 | 0.15693 | 0.011127 | 0.78689 | 0.78689 | 0.008197
2019-03-20 17:53 | 1553118775.0775 | 235 | 0.20727 | 0.00231 | 3.3e-05 | 99.7626 | 97.2307 | 118.2136 | 21.0673 | 0.002994 | 0.002773 | 163.8595 | 0.26439 | 8.4024 | 0.003412 | 0.005161 | 0.00504 | 6.6e-05 | 0.001151 | 0.25818 | 1.3295 | 2.5481 | 100.58 | 20.75 | 19.367 | 93.6767 | 103.7569 | 88.3337 | 0.004721 | 0.97666 | 0.023503 | 24.0652 | 0.22993 | 0.15728 | 0.011193 | 0.79091 | 0.79091 | 0.008173
2019-03-20 17:53 | 1553118776.0776 | 236 | 0.20727 | 0.002293 | 3.4e-05 | 99.9996 | 97.6616 | 117.9874 | 21.0673 | 0.002964 | 0.002675 | 163.8361 | 0.26348 | 8.4005 | 0.003329 | 0.005196 | 0.005126 | 6.1e-05 | 0.001151 | 0.27057 | 1.3455 | 2.5482 | 100.58 | 20.75 | 19.388 | 92.953 | 101.7662 | 85.9779 | 0.004726 | 0.98306 | 0.023558 | 23.9637 | 0.21161 | 0.15675 | 0.011193 | 0.79267 | 0.79267 | 0.008127
2019-03-20 17:53 | 1553118777.0786 | 237 | 0.2073 | 0.002285 | 3.5e-05 | 99.9563 | 97.4923 | 117.9099 | 21.0666 | 0.002967 | 0.002748 | 163.7682 | 0.26247 | 8.4053 | 0.003336 | 0.004899 | 0.004731 | 6.6e-05 | 0.001147 | 0.26737 | 1.3167 | 2.5481 | 100.58 | 20.75 | 19.396 | 93.0227 | 101.9352 | 87.7292 | 0.004728 | 0.97829 | 0.023493 | 24.0142 | 0.22325 | 0.15614 | 0.011052 | 0.7805 | 0.7805 | 0.008222
2019-03-20 17:53 | 1553118778.0786 | 238 | 0.2073 | 0.002273 | 3.7e-05 | 100.1109 | 97.6934 | 117.7667 | 21.0667 | 0.002992 | 0.002739 | 163.748 | 0.26335 | 8.4021 | 0.003304 | 0.004688 | 0.004556 | 6.4e-05 | 0.001142 | 0.27649 | 1.319 | 2.5481 | 100.58 | 20.75 | 19.403 | 93.6088 | 101.1507 | 87.5147 | 0.00473 | 0.97887 | 0.023486 | 23.9924 | 0.21864 | 0.15667 | 0.011052 | 0.78013 | 0.78013 | 0.008254
2019-03-20 17:53 | 1553118779.0781 | 239 | 0.20728 | 0.002264 | 3.8e-05 | 100.1422 | 97.8765 | 117.778 | 21.0663 | 0.002948 | 0.002697 | 163.6947 | 0.26321 | 8.4005 | 0.003331 | 0.004405 | 0.004283 | 6.6e-05 | 0.001147 | 0.29087 | 1.355 | 2.5482 | 100.58 | 20.75 | 19.414 | 92.5513 | 101.8068 | 86.5062 | 0.004732 | 0.98162 | 0.02352 | 23.96 | 0.20469 | 0.15658 | 0.011194 | 0.79123 | 0.79123 | 0.008134
2019-03-20 17:53 | 1553118780.0778 | 240 | 0.20731 | 0.002261 | 4e-05 | 100.1668 | 97.9446 | 117.6083 | 21.0672 | 0.00297 | 0.002726 | 163.6572 | 0.26333 | 8.404 | 0.003327 | 0.004572 | 0.004448 | 6.5e-05 | 0.001145 | 0.27041 | 1.3597 | 2.5481 | 100.58 | 20.76 | 19.417 | 93.0865 | 101.7143 | 87.2168 | 0.004736 | 0.97968 | 0.023479 | 23.9664 | 0.20067 | 0.15665 | 0.010987 | 0.77519 | 0.77519 | 0.008306
2019-03-20 17:53 | 1553118781.0779 | 241 | 0.20738 | 0.002251 | 4.1e-05 | 100.1422 | 98.1 | 117.5888 | 21.0673 | 0.002973 | 0.002667 | 163.6074 | 0.2629 | 8.4019 | 0.003326 | 0.005245 | 0.005169 | 6.6e-05 | 0.00115 | 0.24275 | 1.2865 | 2.548 | 100.58 | 20.76 | 19.417 | 93.1657 | 101.6931 | 85.7852 | 0.004736 | 0.98359 | 0.023524 | 23.9168 | 0.18429 | 0.1564 | 0.010612 | 0.75007 | 0.75007 | 0.00857
2019-03-20 17:53 | 1553118782.0789 | 242 | 0.20733 | 0.002239 | 4.2e-05 | 100.2215 | 98.1778 | 117.7353 | 21.0676 | 0.002987 | 0.002687 | 163.5681 | 0.26348 | 8.4032 | 0.003338 | 0.00485 | 0.004746 | 6.6e-05 | 0.001148 | 0.25553 | 1.3137 | 2.548 | 100.58 | 20.76 | 19.415 | 93.4968 | 101.981 | 86.2679 | 0.004736 | 0.98227 | 0.023523 | 23.9477 | 0.18428 | 0.15674 | 0.010943 | 0.77339 | 0.77339 | 0.00833
2019-03-20 17:53 | 1553118783.0788 | 243 | 0.20737 | 0.002232 | 4.3e-05 | 100.2726 | 97.8328 | 117.9477 | 21.0663 | 0.002963 | 0.002717 | 163.5222 | 0.26396 | 8.4038 | 0.003408 | 0.005111 | 0.004981 | 6.4e-05 | 0.001149 | 0.29374 | 1.3437 | 2.5481 | 100.58 | 20.76 | 19.413 | 92.9127 | 103.6774 | 86.9956 | 0.004735 | 0.98028 | 0.023521 | 23.9936 | 0.22032 | 0.15703 | 0.010706 | 0.75643 | 0.75643 | 0.008532
2019-03-20 17:53 | 1553118784.077 | 244 | 0.20737 | 0.002226 | 4.4e-05 | 100.2603 | 97.89 | 118.0587 | 21.0672 | 0.002875 | 0.002709 | 163.4805 | 0.26422 | 8.4005 | 0.003374 | 0.00466 | 0.004486 | 6.4e-05 | 0.001147 | 0.27113 | 1.3532 | 2.548 | 100.58 | 20.76 | 19.408 | 90.7947 | 102.8437 | 86.7951 | 0.004734 | 0.98083 | 0.023538 | 23.9982 | 0.214 | 0.15719 | 0.010677 | 0.75487 | 0.75487 | 0.008558
2019-03-20 17:53 | 1553118785.0771 | 245 | 0.20732 | 0.002222 | 4.6e-05 | 100.3369 | 98.2854 | 117.8888 | 21.0675 | 0.003021 | 0.002644 | 163.4257 | 0.26336 | 8.4018 | 0.003335 | 0.00519 | 0.00501 | 6.6e-05 | 0.001139 | 0.28816 | 1.3587 | 2.548 | 100.58 | 20.76 | 19.407 | 94.3232 | 101.9014 | 85.228 | 0.004734 | 0.98512 | 0.023573 | 23.9286 | 0.18478 | 0.15667 | 0.011005 | 0.77918 | 0.77918 | 0.008264
2019-03-20 17:53 | 1553118786.077 | 246 | 0.20734 | 0.002207 | 4.8e-05 | 100.3599 | 98.2563 | 117.8306 | 21.0669 | 0.00295 | 0.002723 | 163.3812 | 0.26326 | 8.4019 | 0.003387 | 0.004814 | 0.004651 | 6.5e-05 | 0.001136 | 0.28479 | 1.3816 | 2.5482 | 100.58 | 20.76 | 19.405 | 92.6129 | 103.1584 | 87.1259 | 0.004733 | 0.97993 | 0.023505 | 23.986 | 0.18948 | 0.15661 | 0.010903 | 0.76952 | 0.76952 | 0.008365
2019-03-20 17:53 | 1553118787.077 | 247 | 0.20737 | 0.002203 | 5.1e-05 | 100.4605 | 98.4648 | 117.5264 | 21.0665 | 0.00297 | 0.002687 | 163.3465 | 0.26291 | 8.4015 | 0.003408 | 0.004567 | 0.004407 | 6.3e-05 | 0.001137 | 0.28279 | 1.4014 | 2.5481 | 100.58 | 20.76 | 19.401 | 93.0827 | 103.658 | 86.273 | 0.004732 | 0.98226 | 0.023502 | 23.9266 | 0.17947 | 0.1564 | 0.010714 | 0.75594 | 0.75594 | 0.008504
2019-03-20 17:53 | 1553118788.0781 | 248 | 0.20733 | 0.002197 | 5.3e-05 | 100.3023 | 98.2933 | 117.4808 | 21.0667 | 0.002987 | 0.002626 | 163.3002 | 0.26256 | 8.4029 | 0.003386 | 0.005504 | 0.005333 | 6.6e-05 | 0.001138 | 0.28423 | 1.3486 | 2.5481 | 100.58 | 20.76 | 19.4 | 93.5053 | 103.1347 | 84.7912 | 0.004732 | 0.98632 | 0.023546 | 23.8726 | 0.18098 | 0.1562 | 0.010932 | 0.77272 | 0.77272 | 0.008308
2019-03-20 17:53 | 1553118789.0772 | 249 | 0.20735 | 0.002198 | 5.4e-05 | 100.3463 | 98.214 | 117.2435 | 21.067 | 0.002957 | 0.002686 | 163.2629 | 0.26439 | 8.4049 | 0.003416 | 0.004981 | 0.004857 | 6.6e-05 | 0.001136 | 0.28567 | 1.3639 | 2.5481 | 100.58 | 20.76 | 19.397 | 92.7656 | 103.855 | 86.2555 | 0.004731 | 0.9823 | 0.023474 | 23.8973 | 0.19211 | 0.15729 | 0.010848 | 0.76437 | 0.76437 | 0.008457
2019-03-20 17:53 | 1553118790.0782 | 250 | 0.20735 | 0.002194 | 5.5e-05 | 100.314 | 98.0204 | 117.128 | 21.0669 | 0.002935 | 0.00275 | 163.2268 | 0.26372 | 8.404 | 0.003433 | 0.005037 | 0.00496 | 6.3e-05 | 0.001132 | 0.26233 | 1.3075 | 2.5479 | 100.58 | 20.76 | 19.395 | 92.2526 | 104.2756 | 87.7968 | 0.004731 | 0.97811 | 0.023413 | 23.9366 | 0.20688 | 0.15688 | 0.010841 | 0.7618 | 0.7618 | 0.008464
2019-03-20 17:53 | 1553118791.0774 | 251 | 0.20736 | 0.002186 | 5.7e-05 | 100.3353 | 98.1157 | 117.183 | 21.0665 | 0.002919 | 0.002645 | 163.1951 | 0.2635 | 8.4005 | 0.003343 | 0.004886 | 0.004725 | 6.4e-05 | 0.00113 | 0.32084 | 1.3094 | 2.5481 | 100.58 | 20.76 | 19.392 | 91.8446 | 102.0985 | 85.2569 | 0.00473 | 0.98504 | 0.023501 | 23.8579 | 0.20009 | 0.15676 | 0.010754 | 0.75842 | 0.75842 | 0.008495
2019-03-20 17:53 | 1553118792.0774 | 252 | 0.20736 | 0.002181 | 5.9e-05 | 100.3871 | 98.3718 | 117.1799 | 21.067 | 0.002948 | 0.002685 | 163.154 | 0.26305 | 8.4043 | 0.003362 | 0.004715 | 0.004537 | 6.4e-05 | 0.001124 | 0.28943 | 1.3166 | 2.5482 | 100.58 | 20.76 | 19.386 | 92.5506 | 102.5505 | 86.2297 | 0.004729 | 0.98237 | 0.023469 | 23.8899 | 0.18139 | 0.15649 | 0.010789 | 0.75966 | 0.75966 | 0.008466
2019-03-20 17:53 | 1553118793.0774 | 253 | 0.20735 | 0.002178 | 6e-05 | 100.3496 | 98.3198 | 116.9851 | 21.0669 | 0.002971 | 0.002729 | 163.0909 | 0.26273 | 8.4024 | 0.003409 | 0.004596 | 0.00447 | 6.1e-05 | 0.001117 | 0.29286 | 1.3166 | 2.5481 | 100.58 | 20.76 | 19.383 | 93.1183 | 103.6977 | 87.2704 | 0.004728 | 0.97954 | 0.023415 | 23.9046 | 0.18277 | 0.1563 | 0.010837 | 0.76126 | 0.76126 | 0.008439
2019-03-20 17:53 | 1553118794.0789 | 254 | 0.20729 | 0.002175 | 6.3e-05 | 100.5109 | 98.3904 | 116.6896 | 21.066 | 0.002985 | 0.002699 | 163.0461 | 0.2631 | 8.4011 | 0.003409 | 0.005439 | 0.00536 | 6.1e-05 | 0.001124 | 0.24873 | 1.3115 | 2.548 | 100.58 | 20.76 | 19.38 | 93.454 | 103.7038 | 86.5505 | 0.004727 | 0.9815 | 0.023409 | 23.8505 | 0.19072 | 0.15652 | 0.011195 | 0.78612 | 0.78612 | 0.008183
2019-03-20 17:53 | 1553118795.0785 | 255 | 0.20713 | 0.00217 | 6.6e-05 | 100.4356 | 98.3314 | 116.3558 | 21.0653 | 0.002972 | 0.002706 | 162.9845 | 0.26312 | 8.4012 | 0.003449 | 0.004531 | 0.004386 | 6e-05 | 0.001122 | 0.27887 | 1.3156 | 2.548 | 100.58 | 20.76 | 19.378 | 93.1341 | 104.6584 | 86.7332 | 0.004727 | 0.981 | 0.02337 | 23.8224 | 0.18938 | 0.15653 | 0.012207 | 0.85584 | 0.85584 | 0.007517
2019-03-20 17:53 | 1553118796.0789 | 256 | 0.20701 | 0.002177 | 7e-05 | 100.4812 | 98.2955 | 115.9448 | 21.0664 | 0.002927 | 0.002741 | 162.8893 | 0.26245 | 8.4021 | 0.003439 | 0.004544 | 0.004422 | 6e-05 | 0.001123 | 0.30038 | 1.3207 | 2.548 | 100.58 | 20.76 | 19.377 | 92.0441 | 104.4219 | 87.5597 | 0.004726 | 0.97875 | 0.023302 | 23.8076 | 0.19671 | 0.15613 | 0.01288 | 0.90031 | 0.90031 | 0.007128
2019-03-20 17:53 | 1553118797.0777 | 257 | 0.20673 | 0.00219 | 7.6e-05 | 100.454 | 98.7085 | 115.2552 | 21.0661 | 0.002963 | 0.002764 | 162.779 | 0.26265 | 8.4014 | 0.00345 | 0.004813 | 0.00462 | 5.8e-05 | 0.001119 | 0.27455 | 1.3385 | 2.5482 | 100.58 | 20.76 | 19.373 | 92.9094 | 104.6836 | 88.1231 | 0.004725 | 0.97723 | 0.023214 | 23.7552 | 0.15679 | 0.15625 | 0.014557 | 1.0137 | 1.0137 | 0.006335
2019-03-20 17:53 | 1553118798.0777 | 258 | 0.20637 | 0.002237 | 8.3e-05 | 100.4408 | 99.0912 | 114.2522 | 21.065 | 0.002997 | 0.002798 | 162.6699 | 0.26349 | 8.4038 | 0.003397 | 0.004467 | 0.004397 | 6.2e-05 | 0.001121 | 0.2891 | 1.3433 | 2.5481 | 100.58 | 20.76 | 19.372 | 93.7393 | 103.3964 | 88.9453 | 0.004725 | 0.97501 | 0.023087 | 23.6785 | 0.12099 | 0.15675 | 0.01665 | 1.153 | 1.153 | 0.005587
2019-03-20 17:53 | 1553118799.0777 | 259 | 0.20596 | 0.002312 | 9.2e-05 | 100.4074 | 99.2553 | 113.3748 | 21.0662 | 0.002975 | 0.002816 | 162.5482 | 0.26329 | 8.4052 | 0.003412 | 0.005122 | 0.005021 | 5.7e-05 | 0.001112 | 0.26681 | 1.3376 | 2.5482 | 100.58 | 20.76 | 19.369 | 93.2148 | 103.7677 | 89.3729 | 0.004724 | 0.97386 | 0.022984 | 23.6014 | 0.10323 | 0.15663 | 0.019015 | 1.3108 | 1.3108 | 0.004911
2019-03-20 17:53 | 1553118800.0778 | 260 | 0.20519 | 0.002427 | 0.000103 | 100.4314 | 99.1043 | 112.5488 | 21.0671 | 0.003111 | 0.002875 | 162.4058 | 0.26338 | 8.3997 | 0.003368 | 0.004933 | 0.004779 | 5.8e-05 | 0.001109 | 0.26344 | 1.3422 | 2.5481 | 100.58 | 20.76 | 19.364 | 96.4855 | 102.7092 | 90.7939 | 0.004723 | 0.97005 | 0.022856 | 23.5613 | 0.11898 | 0.15669 | 0.023484 | 1.6098 | 1.6098 | 0.004
2019-03-20 17:53 | 1553118801.0768 | 261 | 0.20452 | 0.002629 | 0.000114 | 100.3263 | 98.6192 | 111.5578 | 21.0664 | 0.003077 | 0.002881 | 162.2521 | 0.26311 | 8.4043 | 0.003353 | 0.004276 | 0.004171 | 6.1e-05 | 0.001111 | 0.25498 | 1.3628 | 2.5481 | 100.58 | 20.76 | 19.363 | 95.6609 | 102.3465 | 90.9339 | 0.004723 | 0.96968 | 0.022751 | 23.4618 | 0.1535 | 0.15652 | 0.027237 | 1.858 | 1.858 | 0.003462
2019-03-20 17:53 | 1553118802.0782 | 262 | 0.20356 | 0.003103 | 0.000127 | 100.3428 | 97.7745 | 110.4368 | 21.0664 | 0.00314 | 0.002942 | 162.0909 | 0.26243 | 8.3995 | 0.003393 | 0.004893 | 0.004669 | 5.7e-05 | 0.001116 | 0.286 | 1.3573 | 2.5482 | 100.58 | 20.76 | 19.361 | 97.1837 | 103.3137 | 92.4201 | 0.004722 | 0.96574 | 0.02259 | 23.3913 | 0.23192 | 0.15612 | 0.032393 | 2.1934 | 2.1934 | 0.002925
2019-03-20 17:53 | 1553118803.0769 | 263 | 0.20285 | 0.003683 | 0.00014 | 100.3557 | 96.6974 | 109.2833 | 21.0665 | 0.003154 | 0.003016 | 161.9344 | 0.26273 | 8.3999 | 0.003399 | 0.005326 | 0.005217 | 5.9e-05 | 0.001118 | 0.29856 | 1.3014 | 2.548 | 100.58 | 20.76 | 19.362 | 97.5214 | 103.441 | 94.1899 | 0.004723 | 0.96109 | 0.022417 | 23.325 | 0.33215 | 0.1563 | 0.035857 | 2.4086 | 2.4086 | 0.002667
2019-03-20 17:53 | 1553118804.0774 | 264 | 0.20209 | 0.004319 | 0.000152 | 100.2668 | 95.8359 | 108.4092 | 21.066 | 0.003222 | 0.003037 | 161.7507 | 0.2629 | 8.4052 | 0.003493 | 0.005022 | 0.004895 | 6.1e-05 | 0.001123 | 0.28031 | 1.3044 | 2.548 | 100.58 | 20.76 | 19.361 | 99.1705 | 105.7328 | 94.7123 | 0.004722 | 0.95972 | 0.022312 | 23.2481 | 0.40427 | 0.1564 | 0.039599 | 2.6465 | 2.6465 | 0.002429
2019-03-20 17:53 | 1553118805.0781 | 265 | 0.20088 | 0.005017 | 0.000163 | 100.2949 | 95.3018 | 107.5396 | 21.0661 | 0.003247 | 0.003064 | 161.556 | 0.26337 | 8.4036 | 0.003551 | 0.004566 | 0.004528 | 5.9e-05 | 0.001122 | 0.27079 | 1.3262 | 2.548 | 100.58 | 20.76 | 19.361 | 99.7785 | 107.1484 | 95.3645 | 0.004722 | 0.95802 | 0.022202 | 23.1752 | 0.45677 | 0.15668 | 0.046003 | 3.0582 | 3.0582 | 0.002106
2019-03-20 17:53 | 1553118806.0772 | 266 | 0.20001 | 0.005745 | 0.000174 | 100.3107 | 94.6331 | 106.7589 | 21.0655 | 0.003316 | 0.003095 | 161.3695 | 0.26258 | 8.4019 | 0.003528 | 0.004947 | 0.004785 | 6.2e-05 | 0.001116 | 0.25536 | 1.3259 | 2.5482 | 100.58 | 20.76 | 19.361 | 101.437 | 106.5823 | 96.1008 | 0.004722 | 0.95611 | 0.0221 | 23.1139 | 0.52116 | 0.15621 | 0.050253 | 3.324 | 3.324 | 0.001931
2019-03-20 17:53 | 1553118807.0771 | 267 | 0.19924 | 0.006425 | 0.000184 | 100.3447 | 94.0396 | 106.5105 | 21.0661 | 0.003366 | 0.00313 | 161.2129 | 0.26251 | 8.4027 | 0.003541 | 0.005565 | 0.005379 | 6.2e-05 | 0.001115 | 0.25442 | 1.3648 | 2.548 | 100.58 | 20.76 | 19.359 | 102.6627 | 106.9003 | 96.9474 | 0.004722 | 0.95393 | 0.022049 | 23.1135 | 0.58046 | 0.15616 | 0.053998 | 3.5622 | 3.5622 | 0.001802
2019-03-20 17:53 | 1553118808.0771 | 268 | 0.19862 | 0.007051 | 0.000194 | 100.106 | 93.5048 | 106.4943 | 21.0655 | 0.003379 | 0.00317 | 161.0751 | 0.26266 | 8.4047 | 0.003568 | 0.005866 | 0.005723 | 6.4e-05 | 0.001117 | 0.2756 | 1.3786 | 2.5482 | 100.58 | 20.76 | 19.357 | 102.9701 | 107.5555 | 97.9021 | 0.004721 | 0.95147 | 0.022018 | 23.1415 | 0.61017 | 0.15626 | 0.056952 | 3.7508 | 3.7508 | 0.001712
2019-03-20 17:53 | 1553118809.0776 | 269 | 0.19817 | 0.007568 | 0.000203 | 100.1115 | 92.7655 | 106.5774 | 21.0658 | 0.003384 | 0.003186 | 160.9226 | 0.26313 | 8.4011 | 0.003653 | 0.004871 | 0.004753 | 6.3e-05 | 0.001115 | 0.26272 | 1.3943 | 2.5481 | 100.58 | 20.76 | 19.356 | 103.0961 | 109.6056 | 98.3061 | 0.004721 | 0.95044 | 0.022015 | 23.1632 | 0.68166 | 0.15654 | 0.059009 | 3.8847 | 3.8847 | 0.001656
2019-03-20 17:53 | 1553118810.0784 | 270 | 0.19785 | 0.007982 | 0.000211 | 99.9521 | 91.6913 | 106.5965 | 21.0664 | 0.003494 | 0.003261 | 160.7642 | 0.26338 | 8.4006 | 0.003607 | 0.005711 | 0.005679 | 6.1e-05 | 0.001115 | 0.28136 | 1.3814 | 2.5481 | 100.58 | 20.75 | 19.352 | 105.7458 | 108.5073 | 100.1122 | 0.004717 | 0.94584 | 0.021964 | 23.2215 | 0.77158 | 0.15669 | 0.060396 | 3.9661 | 3.9661 | 0.001624
2019-03-20 17:54 | 1553118811.0784 | 271 | 0.19759 | 0.008319 | 0.000218 | 99.8848 | 90.9476 | 106.7833 | 21.0662 | 0.003498 | 0.003278 | 160.6175 | 0.26331 | 8.4026 | 0.003695 | 0.004894 | 0.004714 | 6.1e-05 | 0.001113 | 0.2688 | 1.3353 | 2.548 | 100.58 | 20.76 | 19.349 | 105.8558 | 110.6265 | 100.5249 | 0.004719 | 0.94479 | 0.021971 | 23.2547 | 0.83841 | 0.15664 | 0.061516 | 4.0402 | 4.0402 | 0.001593
2019-03-20 17:54 | 1553118812.0774 | 272 | 0.19761 | 0.008564 | 0.000223 | 99.7529 | 90.3922 | 107.0746 | 21.0668 | 0.003466 | 0.003258 | 160.4485 | 0.26314 | 8.4045 | 0.003739 | 0.004982 | 0.00483 | 6.6e-05 | 0.001112 | 0.27887 | 1.3186 | 2.5481 | 100.58 | 20.76 | 19.342 | 105.0855 | 111.6966 | 100.0413 | 0.004718 | 0.94602 | 0.022015 | 23.2713 | 0.88138 | 0.15654 | 0.06112 | 4.0221 | 4.0221 | 0.0016
2019-03-20 17:54 | 1553118813.0775 | 273 | 0.19757 | 0.008807 | 0.000227 | 99.7309 | 90.0599 | 107.586 | 21.0663 | 0.003554 | 0.00328 | 160.3075 | 0.26356 | 8.4055 | 0.003799 | 0.006498 | 0.006411 | 6.2e-05 | 0.00112 | 0.26936 | 1.2974 | 2.548 | 100.58 | 20.76 | 19.339 | 107.2075 | 113.1754 | 100.5728 | 0.004717 | 0.94467 | 0.022052 | 23.3434 | 0.91234 | 0.15679 | 0.061058 | 4.0246 | 4.0246 | 0.001601
2019-03-20 17:54 | 1553118814.0775 | 274 | 0.19763 | 0.009026 | 0.000227 | 99.5873 | 90.7416 | 107.9025 | 21.0669 | 0.003541 | 0.003296 | 160.1745 | 0.26323 | 8.3996 | 0.003798 | 0.004816 | 0.004746 | 6.3e-05 | 0.001119 | 0.2792 | 1.3329 | 2.5482 | 100.58 | 20.76 | 19.337 | 106.9058 | 113.1529 | 100.9572 | 0.004717 | 0.9437 | 0.022073 | 23.3897 | 0.832 | 0.1566 | 0.060427 | 3.9869 | 3.9869 | 0.001614
2019-03-20 17:54 | 1553118815.077 | 275 | 0.19768 | 0.009205 | 0.000225 | 99.4896 | 91.2245 | 108.3477 | 21.066 | 0.003626 | 0.003314 | 160.0386 | 0.26291 | 8.4013 | 0.003879 | 0.005549 | 0.005412 | 6.2e-05 | 0.001125 | 0.27041 | 1.3294 | 2.548 | 100.58 | 20.76 | 19.333 | 108.9658 | 115.1211 | 101.3951 | 0.004716 | 0.9426 | 0.022105 | 23.4516 | 0.77575 | 0.1564 | 0.059945 | 3.9612 | 3.9612 | 0.001623
2019-03-20 17:54 | 1553118816.0767 | 276 | 0.19806 | 0.009334 | 0.000222 | 99.4055 | 91.0623 | 108.5512 | 21.0667 | 0.003596 | 0.003347 | 159.9292 | 0.26312 | 8.4027 | 0.003885 | 0.005628 | 0.005519 | 6.6e-05 | 0.001124 | 0.30989 | 1.3388 | 2.5481 | 100.58 | 20.76 | 19.33 | 108.2225 | 115.2588 | 102.1974 | 0.004715 | 0.94058 | 0.022103 | 23.4988 | 0.78411 | 0.15653 | 0.057502 | 3.7998 | 3.7998 | 0.001693
2019-03-20 17:54 | 1553118817.0777 | 277 | 0.19848 | 0.009399 | 0.000219 | 99.2829 | 91.3688 | 108.882 | 21.0669 | 0.003616 | 0.003335 | 159.8182 | 0.26313 | 8.405 | 0.003949 | 0.005191 | 0.005057 | 6.6e-05 | 0.001119 | 0.26377 | 1.3115 | 2.5481 | 100.58 | 20.76 | 19.325 | 108.7111 | 116.8204 | 101.8921 | 0.004714 | 0.94135 | 0.022145 | 23.525 | 0.74302 | 0.15654 | 0.05493 | 3.6375 | 3.6375 | 0.001769
2019-03-20 17:54 | 1553118818.0777 | 278 | 0.19877 | 0.009413 | 0.000216 | 99.2056 | 92.4437 | 109.1803 | 21.0664 | 0.003652 | 0.003302 | 159.6668 | 0.26276 | 8.4035 | 0.0039 | 0.005801 | 0.005669 | 6.8e-05 | 0.001122 | 0.3079 | 1.339 | 2.5482 | 100.58 | 20.76 | 19.316 | 109.5809 | 115.6287 | 101.1018 | 0.004711 | 0.94334 | 0.022199 | 23.5324 | 0.63144 | 0.15632 | 0.053189 | 3.5312 | 3.5312 | 0.001819
2019-03-20 17:54 | 1553118819.0782 | 279 | 0.19935 | 0.009365 | 0.000214 | 99.0863 | 92.9583 | 109.5824 | 21.0668 | 0.003603 | 0.003408 | 159.5407 | 0.26259 | 8.4009 | 0.004065 | 0.006357 | 0.006296 | 6.5e-05 | 0.001121 | 0.2792 | 1.3148 | 2.548 | 100.58 | 20.75 | 19.311 | 108.41 | 119.6651 | 103.6685 | 0.004707 | 0.93691 | 0.022164 | 23.6563 | 0.57102 | 0.15621 | 0.049803 | 3.3017 | 3.3017 | 0.001945
2019-03-20 17:54 | 1553118820.0778 | 280 | 0.19961 | 0.009267 | 0.000212 | 98.9789 | 92.5984 | 110.0666 | 21.0667 | 0.003646 | 0.0033 | 159.4287 | 0.26256 | 8.4006 | 0.004115 | 0.00586 | 0.005722 | 6.7e-05 | 0.001129 | 0.27865 | 1.3078 | 2.5482 | 100.58 | 20.76 | 19.302 | 109.4544 | 120.8884 | 101.059 | 0.004708 | 0.94344 | 0.02229 | 23.6263 | 0.59602 | 0.1562 | 0.048362 | 3.2247 | 3.2247 | 0.001991
2019-03-20 17:54 | 1553118821.0769 | 281 | 0.19983 | 0.009126 | 0.000212 | 98.7564 | 92.6792 | 110.3485 | 21.067 | 0.00365 | 0.003353 | 159.316 | 0.26312 | 8.4048 | 0.004082 | 0.005968 | 0.005847 | 6.4e-05 | 0.001125 | 0.28208 | 1.3277 | 2.548 | 100.58 | 20.75 | 19.297 | 109.5319 | 120.0611 | 102.3352 | 0.004704 | 0.94024 | 0.022281 | 23.6969 | 0.56809 | 0.15653 | 0.047226 | 3.1477 | 3.1477 | 0.002044
2019-03-20 17:54 | 1553118822.0779 | 282 | 0.20011 | 0.008971 | 0.000212 | 98.6726 | 92.9684 | 110.3511 | 21.0665 | 0.003595 | 0.003335 | 159.2293 | 0.26266 | 8.4032 | 0.004191 | 0.004957 | 0.004868 | 6.4e-05 | 0.001131 | 0.29662 | 1.3448 | 2.548 | 100.58 | 20.75 | 19.291 | 108.2111 | 122.7318 | 101.8935 | 0.004702 | 0.94134 | 0.022294 | 23.6832 | 0.53263 | 0.15625 | 0.045753 | 3.0514 | 3.0514 | 0.002105
2019-03-20 17:54 | 1553118823.0785 | 283 | 0.20026 | 0.008799 | 0.000213 | 98.4631 | 92.8092 | 110.4358 | 21.0664 | 0.003652 | 0.003324 | 159.1149 | 0.26266 | 8.4022 | 0.004167 | 0.005626 | 0.005545 | 6.6e-05 | 0.00113 | 0.27577 | 1.3374 | 2.5481 | 100.58 | 20.75 | 19.282 | 109.5848 | 122.148 | 101.6341 | 0.0047 | 0.942 | 0.02231 | 23.6841 | 0.52894 | 0.15626 | 0.045064 | 3.0076 | 3.0076 | 0.002135
2019-03-20 17:54 | 1553118824.0782 | 284 | 0.20038 | 0.008626 | 0.000215 | 98.3052 | 93.1385 | 110.5893 | 21.0669 | 0.003644 | 0.003332 | 158.9922 | 0.26262 | 8.4065 | 0.004152 | 0.005439 | 0.005313 | 6.6e-05 | 0.001127 | 0.2808 | 1.3271 | 2.5481 | 100.58 | 20.75 | 19.277 | 109.4079 | 121.776 | 101.82 | 0.004699 | 0.94153 | 0.02232 | 23.7064 | 0.48291 | 0.15624 | 0.044567 | 2.9757 | 2.9757 | 0.002158
2019-03-20 17:54 | 1553118825.0775 | 285 | 0.20048 | 0.00846 | 0.000217 | 98.1536 | 92.715 | 110.8222 | 21.0658 | 0.003671 | 0.003384 | 158.9144 | 0.2625 | 8.4032 | 0.004224 | 0.004313 | 0.00419 | 6.7e-05 | 0.001124 | 0.27721 | 1.3681 | 2.5482 | 100.58 | 20.75 | 19.264 | 110.056 | 123.5409 | 103.0836 | 0.004696 | 0.93837 | 0.022306 | 23.7713 | 0.50986 | 0.15616 | 0.044138 | 2.9452 | 2.9452 | 0.002179
2019-03-20 17:54 | 1553118826.0772 | 286 | 0.20067 | 0.008298 | 0.000218 | 98.0873 | 92.5151 | 110.9621 | 21.0661 | 0.003695 | 0.003395 | 158.7972 | 0.26248 | 8.4036 | 0.004162 | 0.005848 | 0.005757 | 6.8e-05 | 0.001119 | 0.28102 | 1.3618 | 2.548 | 100.58 | 20.75 | 19.256 | 110.6471 | 122.0236 | 103.3653 | 0.004694 | 0.93767 | 0.022312 | 23.7952 | 0.52313 | 0.15615 | 0.043239 | 2.8859 | 2.8859 | 0.002224
2019-03-20 17:54 | 1553118827.0773 | 287 | 0.20071 | 0.00815 | 0.000218 | 97.966 | 92.3488 | 110.814 | 21.0658 | 0.003684 | 0.003351 | 158.6591 | 0.26307 | 8.4048 | 0.004264 | 0.005684 | 0.005496 | 6.7e-05 | 0.001118 | 0.25465 | 1.3211 | 2.548 | 100.58 | 20.75 | 19.247 | 110.3681 | 124.5309 | 102.2799 | 0.004692 | 0.94038 | 0.022329 | 23.745 | 0.52816 | 0.1565 | 0.043176 | 2.8839 | 2.8839 | 0.00223
2019-03-20 17:54 | 1553118828.0773 | 288 | 0.20072 | 0.008021 | 0.000218 | 97.8813 | 92.348 | 110.6745 | 21.0654 | 0.003642 | 0.003353 | 158.5436 | 0.26279 | 8.4021 | 0.004253 | 0.005739 | 0.005659 | 6.5e-05 | 0.001118 | 0.28351 | 1.2784 | 2.548 | 100.58 | 20.75 | 19.23 | 109.3411 | 124.2545 | 102.348 | 0.004688 | 0.94021 | 0.022313 | 23.7323 | 0.52049 | 0.15633 | 0.043233 | 2.8857 | 2.8857 | 0.002227
2019-03-20 17:54 | 1553118829.0787 | 289 | 0.2007 | 0.007898 | 0.000216 | 97.8586 | 92.1993 | 110.4457 | 21.0657 | 0.003709 | 0.003391 | 158.4349 | 0.26303 | 8.403 | 0.004257 | 0.005463 | 0.005312 | 6.8e-05 | 0.001116 | 0.2625 | 1.3001 | 2.5481 | 100.58 | 20.74 | 19.222 | 110.9821 | 124.3455 | 103.269 | 0.004683 | 0.9379 | 0.022263 | 23.7368 | 0.53284 | 0.15648 | 0.04351 | 2.8977 | 2.8977 | 0.002219
2019-03-20 17:54 | 1553118830.0764 | 290 | 0.20071 | 0.0078 | 0.000214 | 97.7043 | 92.4882 | 110.207 | 21.0665 | 0.003692 | 0.00344 | 158.3183 | 0.26263 | 8.4022 | 0.004264 | 0.006172 | 0.006066 | 6.7e-05 | 0.001114 | 0.2688 | 1.3317 | 2.5482 | 100.58 | 20.75 | 19.206 | 110.5646 | 124.5355 | 104.4465 | 0.004682 | 0.93498 | 0.022204 | 23.7482 | 0.49074 | 0.15624 | 0.04357 | 2.8942 | 2.8942 | 0.002219
2019-03-20 17:54 | 1553118831.0775 | 291 | 0.20067 | 0.007717 | 0.000211 | 97.6684 | 92.8885 | 110.0218 | 21.0666 | 0.003709 | 0.003392 | 158.185 | 0.2626 | 8.4013 | 0.004253 | 0.005926 | 0.005928 | 6.8e-05 | 0.001114 | 0.28031 | 1.3044 | 2.548 | 100.58 | 20.74 | 19.199 | 110.9813 | 124.2572 | 103.28 | 0.004677 | 0.93788 | 0.02222 | 23.6915 | 0.44883 | 0.15622 | 0.043944 | 2.9212 | 2.9212 | 0.002198
2019-03-20 17:54 | 1553118832.0766 | 292 | 0.2006 | 0.007647 | 0.00021 | 97.7855 | 93.559 | 109.9157 | 21.0663 | 0.003721 | 0.003438 | 158.0574 | 0.26308 | 8.404 | 0.004254 | 0.005733 | 0.005459 | 6.7e-05 | 0.001114 | 0.26936 | 1.3652 | 2.5482 | 100.58 | 20.74 | 19.191 | 111.265 | 124.2713 | 104.4007 | 0.004675 | 0.93509 | 0.022176 | 23.7153 | 0.39521 | 0.15651 | 0.044432 | 2.9479 | 2.9479 | 0.002182
2019-03-20 17:54 | 1553118833.0766 | 293 | 0.2006 | 0.007599 | 0.000209 | 97.705 | 93.4465 | 109.8147 | 21.0656 | 0.003767 | 0.003437 | 157.9254 | 0.26266 | 8.4022 | 0.004189 | 0.00536 | 0.005244 | 6.8e-05 | 0.001109 | 0.29911 | 1.3235 | 2.548 | 100.58 | 20.74 | 19.179 | 112.3769 | 122.6881 | 104.3598 | 0.004672 | 0.93519 | 0.022167 | 23.7031 | 0.3986 | 0.15626 | 0.044465 | 2.949 | 2.949 | 0.002178
2019-03-20 17:54 | 1553118834.0766 | 294 | 0.20056 | 0.007563 | 0.000209 | 97.8036 | 93.8008 | 109.3432 | 21.0652 | 0.003821 | 0.003479 | 157.7997 | 0.26164 | 8.4026 | 0.004257 | 0.005757 | 0.005597 | 6.7e-05 | 0.001111 | 0.27743 | 1.317 | 2.548 | 100.58 | 20.74 | 19.174 | 113.7124 | 124.3514 | 105.3827 | 0.004671 | 0.93267 | 0.02209 | 23.6842 | 0.37378 | 0.15565 | 0.04477 | 2.9588 | 2.9588 | 0.002162
2019-03-20 17:54 | 1553118835.0767 | 295 | 0.20049 | 0.007547 | 0.00021 | 97.6613 | 93.8673 | 108.8411 | 21.0656 | 0.003796 | 0.003469 | 157.6629 | 0.26385 | 8.4047 | 0.004174 | 0.005243 | 0.005199 | 6.6e-05 | 0.001109 | 0.27472 | 1.3452 | 2.5481 | 100.58 | 20.74 | 19.17 | 113.0829 | 122.3196 | 105.158 | 0.00467 | 0.93322 | 0.022045 | 23.6228 | 0.35441 | 0.15696 | 0.04523 | 2.9831 | 2.9831 | 0.002163
2019-03-20 17:54 | 1553118836.0781 | 296 | 0.20033 | 0.007544 | 0.000209 | 97.6846 | 93.7896 | 108.6362 | 21.065 | 0.003765 | 0.003518 | 157.5285 | 0.26274 | 8.4028 | 0.004167 | 0.005154 | 0.004913 | 7e-05 | 0.001109 | 0.28479 | 1.3387 | 2.5481 | 100.58 | 20.74 | 19.166 | 112.3383 | 122.1513 | 106.3427 | 0.004669 | 0.93031 | 0.02199 | 23.6374 | 0.36396 | 0.1563 | 0.04619 | 3.0388 | 3.0388 | 0.002114
2019-03-20 17:54 | 1553118837.0778 | 297 | 0.20027 | 0.007562 | 0.00021 | 97.8615 | 93.9563 | 108.8052 | 21.0648 | 0.003829 | 0.003495 | 157.4166 | 0.26312 | 8.4029 | 0.004125 | 0.005862 | 0.005736 | 6.9e-05 | 0.001116 | 0.30879 | 1.309 | 2.548 | 100.58 | 20.74 | 19.167 | 113.8883 | 121.126 | 105.7734 | 0.004669 | 0.93171 | 0.022024 | 23.6381 | 0.36426 | 0.15653 | 0.046506 | 3.0642 | 3.0642 | 0.002099
2019-03-20 17:54 | 1553118838.0771 | 298 | 0.20028 | 0.007578 | 0.00021 | 98.0268 | 94.416 | 108.6808 | 21.0654 | 0.003776 | 0.003471 | 157.299 | 0.26241 | 8.4019 | 0.004096 | 0.005196 | 0.005072 | 6.8e-05 | 0.001122 | 0.29175 | 1.2979 | 2.5481 | 100.58 | 20.74 | 19.174 | 112.615 | 120.4148 | 105.202 | 0.004671 | 0.93311 | 0.022028 | 23.6067 | 0.33569 | 0.15611 | 0.046415 | 3.0588 | 3.0588 | 0.002098
2019-03-20 17:54 | 1553118839.0782 | 299 | 0.20027 | 0.007611 | 0.00021 | 98.1726 | 94.6514 | 108.7015 | 21.0658 | 0.003877 | 0.0035 | 157.1673 | 0.26303 | 8.4026 | 0.004022 | 0.005422 | 0.005264 | 6.9e-05 | 0.001122 | 0.28456 | 1.338 | 2.5481 | 100.58 | 20.74 | 19.18 | 115.0736 | 118.6129 | 105.9097 | 0.004672 | 0.93137 | 0.022009 | 23.631 | 0.32671 | 0.15648 | 0.046421 | 3.0567 | 3.0567 | 0.002104
2019-03-20 17:54 | 1553118840.0784 | 300 | 0.20025 | 0.007639 | 0.00021 | 98.3534 | 95.0925 | 108.7524 | 21.0648 | 0.003873 | 0.003506 | 157.03 | 0.26344 | 8.403 | 0.004083 | 0.005296 | 0.005156 | 6.8e-05 | 0.00112 | 0.26322 | 1.3432 | 2.5482 | 100.58 | 20.74 | 19.186 | 114.961 | 120.1042 | 106.0322 | 0.004674 | 0.93107 | 0.022011 | 23.6404 | 0.30159 | 0.15672 | 0.046506 | 3.0625 | 3.0625 | 0.002103
2019-03-20 17:54 | 1553118841.077 | 301 | 0.2002 | 0.007664 | 0.000209 | 98.477 | 95.2423 | 108.6357 | 21.0658 | 0.003899 | 0.003507 | 156.884 | 0.26285 | 8.4022 | 0.004029 | 0.005623 | 0.005538 | 7e-05 | 0.001116 | 0.27223 | 1.3185 | 2.5481 | 100.58 | 20.74 | 19.201 | 115.6091 | 118.7641 | 106.0773 | 0.004678 | 0.93096 | 0.021998 | 23.6291 | 0.29873 | 0.15637 | 0.046825 | 3.0816 | 3.0816 | 0.002086
2019-03-20 17:54 | 1553118842.077 | 302 | 0.20005 | 0.007691 | 0.000207 | 98.6681 | 95.2382 | 108.8385 | 21.0654 | 0.003865 | 0.003524 | 156.759 | 0.26308 | 8.4016 | 0.004057 | 0.005596 | 0.005411 | 7e-05 | 0.001117 | 0.27273 | 1.3032 | 2.5481 | 100.58 | 20.74 | 19.208 | 114.779 | 119.4591 | 106.4784 | 0.004679 | 0.92998 | 0.022007 | 23.6637 | 0.31646 | 0.15651 | 0.047668 | 3.1384 | 3.1384 | 0.00205
2019-03-20 17:54 | 1553118843.0781 | 303 | 0.19991 | 0.007727 | 0.000205 | 98.7435 | 95.5357 | 108.8439 | 21.0652 | 0.003845 | 0.003552 | 156.6328 | 0.26279 | 8.4025 | 0.004078 | 0.005995 | 0.005914 | 7.2e-05 | 0.001115 | 0.25929 | 1.329 | 2.5481 | 100.58 | 20.74 | 19.221 | 114.2861 | 119.9618 | 107.1711 | 0.004682 | 0.92828 | 0.021987 | 23.6858 | 0.29539 | 0.15634 | 0.048453 | 3.1873 | 3.1873 | 0.002016
2019-03-20 17:54 | 1553118844.0781 | 304 | 0.19977 | 0.007763 | 0.000203 | 98.8441 | 96.0594 | 109.0463 | 21.0654 | 0.003908 | 0.003569 | 156.4885 | 0.26284 | 8.4032 | 0.003916 | 0.005683 | 0.005593 | 7.2e-05 | 0.001113 | 0.26648 | 1.3633 | 2.5482 | 100.58 | 20.74 | 19.228 | 115.8121 | 116.0226 | 107.5768 | 0.004684 | 0.92729 | 0.021996 | 23.7205 | 0.25561 | 0.15637 | 0.049238 | 3.2403 | 3.2403 | 0.001983
2019-03-20 17:54 | 1553118845.0781 | 305 | 0.19969 | 0.007817 | 0.000201 | 98.9589 | 96.3714 | 109.1456 | 21.0658 | 0.00392 | 0.003642 | 156.3507 | 0.26295 | 8.4024 | 0.003993 | 0.005564 | 0.005372 | 7.1e-05 | 0.001112 | 0.24906 | 1.3236 | 2.5481 | 100.58 | 20.75 | 19.234 | 116.1132 | 117.8892 | 109.3485 | 0.004689 | 0.923 | 0.021955 | 23.7864 | 0.23698 | 0.15643 | 0.04964 | 3.2607 | 3.2607 | 0.001972
2019-03-20 17:54 | 1553118846.0774 | 306 | 0.19962 | 0.007873 | 0.0002 | 99.1817 | 96.0298 | 109.2779 | 21.065 | 0.00397 | 0.003588 | 156.2061 | 0.26194 | 8.4047 | 0.003953 | 0.005497 | 0.005417 | 7.5e-05 | 0.001112 | 0.28241 | 1.31 | 2.5481 | 100.58 | 20.74 | 19.244 | 117.3284 | 116.9286 | 108.0365 | 0.004688 | 0.92617 | 0.022006 | 23.76 | 0.28886 | 0.15583 | 0.050016 | 3.293 | 3.293 | 0.001945
2019-03-20 17:54 | 1553118847.0763 | 307 | 0.19948 | 0.007939 | 0.0002 | 99.2862 | 96.5905 | 108.8989 | 21.066 | 0.004005 | 0.003599 | 156.0676 | 0.2621 | 8.4029 | 0.003957 | 0.005164 | 0.005086 | 6.8e-05 | 0.001109 | 0.2354 | 1.3195 | 2.5481 | 100.58 | 20.74 | 19.247 | 118.1993 | 117.0135 | 108.3131 | 0.004689 | 0.9255 | 0.02196 | 23.7273 | 0.24621 | 0.15592 | 0.050762 | 3.335 | 3.335 | 0.001922
2019-03-20 17:54 | 1553118848.0774 | 308 | 0.19942 | 0.008002 | 0.000201 | 99.467 | 96.8371 | 109.1106 | 21.0659 | 0.004 | 0.003608 | 155.9249 | 0.26247 | 8.4016 | 0.003955 | 0.005979 | 0.005835 | 6.7e-05 | 0.001111 | 0.28208 | 1.3156 | 2.548 | 100.58 | 20.74 | 19.253 | 118.0559 | 116.9675 | 108.5195 | 0.00469 | 0.925 | 0.021975 | 23.7568 | 0.23967 | 0.15615 | 0.051044 | 3.3558 | 3.3558 | 0.001912
2019-03-20 17:54 | 1553118849.0775 | 309 | 0.19933 | 0.008071 | 0.0002 | 99.5892 | 96.8902 | 109.2643 | 21.0663 | 0.003953 | 0.003631 | 155.783 | 0.26271 | 8.403 | 0.004045 | 0.005378 | 0.005231 | 6.9e-05 | 0.001113 | 0.27295 | 1.3189 | 2.548 | 100.58 | 20.74 | 19.255 | 116.9323 | 119.1683 | 109.089 | 0.004691 | 0.92363 | 0.021974 | 23.7913 | 0.24575 | 0.15629 | 0.051503 | 3.3858 | 3.3858 | 0.001897
2019-03-20 17:54 | 1553118850.0779 | 310 | 0.19927 | 0.008144 | 0.000201 | 99.7422 | 96.6207 | 109.2678 | 21.0663 | 0.004008 | 0.003616 | 155.662 | 0.26225 | 8.4047 | 0.004028 | 0.005036 | 0.004945 | 6.9e-05 | 0.001109 | 0.27273 | 1.3067 | 2.5481 | 100.58 | 20.74 | 19.255 | 118.267 | 118.7551 | 108.7248 | 0.004691 | 0.92451 | 0.021985 | 23.7803 | 0.2844 | 0.15601 | 0.05179 | 3.4063 | 3.4063 | 0.001882
2019-03-20 17:54 | 1553118851.0776 | 311 | 0.19924 | 0.008201 | 0.000203 | 99.8952 | 97.2296 | 109.1324 | 21.0649 | 0.003952 | 0.003695 | 155.5589 | 0.26207 | 8.4007 | 0.004116 | 0.006098 | 0.005927 | 7.1e-05 | 0.001106 | 0.28766 | 1.3119 | 2.5481 | 100.58 | 20.74 | 19.254 | 116.8921 | 120.8937 | 110.6364 | 0.004691 | 0.9199 | 0.021917 | 23.825 | 0.24192 | 0.15591 | 0.05189 | 3.4021 | 3.4021 | 0.001883
2019-03-20 17:54 | 1553118852.0766 | 312 | 0.19918 | 0.008264 | 0.000205 | 99.9744 | 97.357 | 109.0855 | 21.0654 | 0.004002 | 0.00364 | 155.4196 | 0.2626 | 8.4014 | 0.004097 | 0.005893 | 0.005798 | 7.2e-05 | 0.001101 | 0.28296 | 1.3232 | 2.5481 | 100.58 | 20.74 | 19.254 | 118.1139 | 120.4473 | 109.2936 | 0.004691 | 0.92313 | 0.02195 | 23.7782 | 0.23729 | 0.15622 | 0.052149 | 3.4242 | 3.4242 | 0.001875
2019-03-20 17:54 | 1553118853.0776 | 313 | 0.19908 | 0.008322 | 0.000207 | 100.1277 | 97.7592 | 109.092 | 21.0655 | 0.004032 | 0.003671 | 155.2988 | 0.26172 | 8.4017 | 0.004145 | 0.005261 | 0.005057 | 7.2e-05 | 0.001097 | 0.28567 | 1.3088 | 2.5482 | 100.58 | 20.74 | 19.254 | 118.855 | 121.6197 | 110.0526 | 0.004691 | 0.9213 | 0.021929 | 23.8025 | 0.21413 | 0.1557 | 0.052692 | 3.4564 | 3.4564 | 0.001851
2019-03-20 17:54 | 1553118854.0767 | 314 | 0.19904 | 0.008377 | 0.000208 | 100.2283 | 97.9311 | 108.8451 | 21.0661 | 0.004 | 0.00367 | 155.1727 | 0.26252 | 8.4028 | 0.004238 | 0.005576 | 0.005451 | 7.4e-05 | 0.001098 | 0.28368 | 1.3334 | 2.548 | 100.58 | 20.74 | 19.254 | 118.079 | 123.8839 | 110.0334 | 0.004691 | 0.92135 | 0.021905 | 23.7749 | 0.20739 | 0.15617 | 0.052847 | 3.4626 | 3.4626 | 0.001854
2019-03-20 17:54 | 1553118855.0772 | 315 | 0.19911 | 0.008433 | 0.000208 | 100.4611 | 98.0223 | 108.6167 | 21.0661 | 0.004014 | 0.003677 | 155.0428 | 0.2629 | 8.402 | 0.004294 | 0.005798 | 0.005641 | 7.1e-05 | 0.001096 | 0.29767 | 1.3043 | 2.5482 | 100.58 | 20.74 | 19.253 | 118.4134 | 125.2535 | 110.1958 | 0.00469 | 0.92096 | 0.021877 | 23.755 | 0.21982 | 0.1564 | 0.052377 | 3.4276 | 3.4276 | 0.001875
2019-03-20 17:54 | 1553118856.0779 | 316 | 0.19908 | 0.008487 | 0.000207 | 100.5226 | 97.7252 | 108.2376 | 21.0654 | 0.004061 | 0.003702 | 154.927 | 0.26258 | 8.4022 | 0.004292 | 0.006111 | 0.006015 | 7.4e-05 | 0.001099 | 0.26808 | 1.3611 | 2.5481 | 100.58 | 20.74 | 19.255 | 119.5462 | 125.2153 | 110.8138 | 0.004691 | 0.91948 | 0.021822 | 23.7326 | 0.25244 | 0.15621 | 0.052501 | 3.4269 | 3.4269 | 0.001873
2019-03-20 17:54 | 1553118857.0778 | 317 | 0.19909 | 0.008534 | 0.000207 | 100.5743 | 97.838 | 108.1437 | 21.0657 | 0.004059 | 0.003694 | 154.7949 | 0.26295 | 8.4014 | 0.004307 | 0.005571 | 0.005425 | 7.3e-05 | 0.001099 | 0.28888 | 1.3204 | 2.5481 | 100.58 | 20.74 | 19.257 | 119.4991 | 125.5889 | 110.6056 | 0.004691 | 0.91997 | 0.021818 | 23.7159 | 0.24672 | 0.15643 | 0.052404 | 3.4201 | 3.4201 | 0.00188
2019-03-20 17:54 | 1553118858.0769 | 318 | 0.199 | 0.008573 | 0.000205 | 100.7127 | 98.2252 | 107.6666 | 21.0653 | 0.004065 | 0.003712 | 154.6586 | 0.26253 | 8.4049 | 0.004354 | 0.006423 | 0.006217 | 7.2e-05 | 0.001098 | 0.28174 | 1.367 | 2.5481 | 100.58 | 20.74 | 19.263 | 119.6541 | 126.727 | 111.0534 | 0.004693 | 0.9189 | 0.021757 | 23.6773 | 0.2237 | 0.15618 | 0.052861 | 3.4403 | 3.4403 | 0.001866
2019-03-20 17:54 | 1553118859.0779 | 319 | 0.19892 | 0.008598 | 0.000204 | 100.7943 | 97.9494 | 107.0997 | 21.0647 | 0.004113 | 0.003761 | 154.5223 | 0.26184 | 8.4003 | 0.004431 | 0.005946 | 0.005833 | 7.2e-05 | 0.001107 | 0.28871 | 1.3928 | 2.5481 | 100.58 | 20.74 | 19.265 | 120.8221 | 128.6281 | 112.2448 | 0.004693 | 0.91606 | 0.021666 | 23.6515 | 0.25608 | 0.15577 | 0.053314 | 3.4553 | 3.4553 | 0.001853
2019-03-20 17:54 | 1553118860.0779 | 320 | 0.19881 | 0.008632 | 0.000202 | 100.8919 | 98.013 | 106.791 | 21.0668 | 0.004113 | 0.003761 | 154.3774 | 0.26258 | 8.404 | 0.004395 | 0.006248 | 0.006104 | 7.2e-05 | 0.001104 | 0.25841 | 1.4031 | 2.5482 | 100.58 | 20.74 | 19.269 | 120.84 | 127.751 | 112.2493 | 0.004694 | 0.91605 | 0.021635 | 23.6175 | 0.25895 | 0.15621 | 0.053928 | 3.49 | 3.49 | 0.00184
2019-03-20 17:54 | 1553118861.078 | 321 | 0.19876 | 0.008664 | 0.000201 | 101.0624 | 98.5403 | 106.6723 | 21.0655 | 0.004071 | 0.003745 | 154.2451 | 0.26192 | 8.4008 | 0.004444 | 0.004937 | 0.004842 | 7.2e-05 | 0.001105 | 0.30127 | 1.334 | 2.5481 | 100.58 | 20.74 | 19.271 | 119.8073 | 128.9518 | 111.8421 | 0.004695 | 0.91702 | 0.021634 | 23.5919 | 0.22605 | 0.15582 | 0.054195 | 3.5072 | 3.5072 | 0.001826
2019-03-20 17:54 | 1553118862.0761 | 322 | 0.19864 | 0.008716 | 0.0002 | 101.1571 | 98.5086 | 106.4127 | 21.0652 | 0.00418 | 0.003806 | 154.1073 | 0.26238 | 8.4003 | 0.004363 | 0.006132 | 0.006059 | 7e-05 | 0.00111 | 0.28423 | 1.3755 | 2.5481 | 100.58 | 20.74 | 19.273 | 122.4796 | 126.9618 | 113.3409 | 0.004695 | 0.91346 | 0.021566 | 23.609 | 0.23731 | 0.15609 | 0.054816 | 3.5362 | 3.5362 | 0.001814
2019-03-20 17:54 | 1553118863.0761 | 323 | 0.19858 | 0.008762 | 0.000199 | 101.302 | 98.7549 | 106.3606 | 21.0662 | 0.004206 | 0.003754 | 153.9711 | 0.26227 | 8.4036 | 0.00447 | 0.005548 | 0.005287 | 7.2e-05 | 0.001106 | 0.29087 | 1.3118 | 2.548 | 100.58 | 20.74 | 19.276 | 123.1099 | 129.5881 | 112.0718 | 0.004696 | 0.91647 | 0.021596 | 23.5644 | 0.22778 | 0.15603 | 0.055159 | 3.5632 | 3.5632 | 0.0018
2019-03-20 17:54 | 1553118864.0763 | 324 | 0.19852 | 0.008816 | 0.000198 | 101.3266 | 98.9425 | 106.496 | 21.0655 | 0.004193 | 0.003752 | 153.8463 | 0.26243 | 8.4035 | 0.004428 | 0.006 | 0.005859 | 7.4e-05 | 0.001107 | 0.2756 | 1.3203 | 2.5481 | 100.58 | 20.74 | 19.277 | 122.7825 | 128.5514 | 112.0182 | 0.004696 | 0.9166 | 0.021611 | 23.5778 | 0.21297 | 0.15612 | 0.055425 | 3.583 | 3.583 | 0.001791
2019-03-20 17:54 | 1553118865.0763 | 325 | 0.19857 | 0.00888 | 0.000195 | 101.4734 | 99.138 | 106.5573 | 21.0653 | 0.004134 | 0.003737 | 153.7079 | 0.26266 | 8.4027 | 0.004575 | 0.006476 | 0.006398 | 7.6e-05 | 0.001111 | 0.30541 | 1.2898 | 2.548 | 100.58 | 20.74 | 19.279 | 121.3315 | 132.1605 | 111.6477 | 0.004697 | 0.91748 | 0.021628 | 23.5732 | 0.20827 | 0.15626 | 0.05506 | 3.5623 | 3.5623 | 0.001803
2019-03-20 17:54 | 1553118866.0777 | 326 | 0.19854 | 0.008929 | 0.000192 | 101.6254 | 99.348 | 106.3936 | 21.0661 | 0.00417 | 0.003807 | 153.561 | 0.26238 | 8.4006 | 0.004589 | 0.006023 | 0.00582 | 7.1e-05 | 0.001108 | 0.29607 | 1.2982 | 2.5481 | 100.58 | 20.74 | 19.282 | 122.2288 | 132.5098 | 113.371 | 0.004697 | 0.91339 | 0.021563 | 23.6078 | 0.20273 | 0.15609 | 0.055216 | 3.5618 | 3.5618 | 0.001801
2019-03-20 17:54 | 1553118867.0764 | 327 | 0.19853 | 0.008975 | 0.000188 | 101.7212 | 99.7763 | 106.3871 | 21.0656 | 0.004202 | 0.003815 | 153.4137 | 0.26374 | 8.4013 | 0.00458 | 0.005843 | 0.005739 | 7.6e-05 | 0.001108 | 0.27975 | 1.3076 | 2.5482 | 100.58 | 20.74 | 19.284 | 122.9979 | 132.2776 | 113.5585 | 0.004698 | 0.91295 | 0.021557 | 23.6128 | 0.17267 | 0.1569 | 0.055203 | 3.5602 | 3.5602 | 0.001811
2019-03-20 17:54 | 1553118868.0765 | 328 | 0.1985 | 0.009012 | 0.000187 | 101.8906 | 100.0215 | 106.576 | 21.0662 | 0.004205 | 0.003811 | 153.2777 | 0.26246 | 8.4025 | 0.004618 | 0.005926 | 0.005843 | 7.2e-05 | 0.001113 | 0.29789 | 1.3108 | 2.5481 | 100.58 | 20.74 | 19.283 | 123.0819 | 133.2272 | 113.4517 | 0.004698 | 0.9132 | 0.021579 | 23.6305 | 0.16561 | 0.15614 | 0.055312 | 3.571 | 3.571 | 0.001797
2019-03-20 17:54 | 1553118869.0775 | 329 | 0.19852 | 0.009038 | 0.000186 | 102.0326 | 100.3433 | 106.6427 | 21.0653 | 0.004178 | 0.003814 | 153.1373 | 0.26254 | 8.4 | 0.004631 | 0.005541 | 0.005414 | 7.4e-05 | 0.001111 | 0.28031 | 1.3724 | 2.5482 | 100.58 | 20.74 | 19.283 | 122.416 | 133.5383 | 113.5385 | 0.004698 | 0.913 | 0.021584 | 23.6405 | 0.14933 | 0.15618 | 0.055216 | 3.5655 | 3.5655 | 0.0018
2019-03-20 17:54 | 1553118870.0765 | 330 | 0.19848 | 0.009066 | 0.000187 | 102.1749 | 100.3123 | 106.4983 | 21.0645 | 0.00427 | 0.003861 | 152.9991 | 0.26305 | 8.4033 | 0.004598 | 0.006439 | 0.006208 | 7.2e-05 | 0.001108 | 0.27328 | 1.3448 | 2.548 | 100.58 | 20.74 | 19.284 | 124.6633 | 132.7318 | 114.6707 | 0.004698 | 0.91033 | 0.021538 | 23.6591 | 0.16456 | 0.15649 | 0.055401 | 3.5697 | 3.5697 | 0.001802
2019-03-20 17:54 | 1553118871.0765 | 331 | 0.1984 | 0.009083 | 0.000188 | 102.2871 | 100.5862 | 106.2094 | 21.0652 | 0.004263 | 0.003873 | 152.8378 | 0.26286 | 8.4042 | 0.004621 | 0.006245 | 0.006186 | 7.4e-05 | 0.001104 | 0.29374 | 1.3074 | 2.5481 | 100.58 | 20.74 | 19.288 | 124.5031 | 133.3096 | 114.9622 | 0.004699 | 0.90965 | 0.0215 | 23.6359 | 0.14999 | 0.15638 | 0.055876 | 3.594 | 3.594 | 0.001788
2019-03-20 17:55 | 1553118872.1388 | 332 | 0.19826 | 0.009101 | 0.000191 | 102.4492 | 100.4786 | 105.8247 | 21.0656 | 0.004251 | 0.003869 | 152.6871 | 0.26217 | 8.4027 | 0.004702 | 0.006321 | 0.00622 | 7.3e-05 | 0.0011 | 0.27433 | 1.3219 | 2.5481 | 100.58 | 20.73 | 19.292 | 124.2005 | 135.3042 | 114.8828 | 0.004697 | 0.90983 | 0.021464 | 23.5906 | 0.17372 | 0.15596 | 0.056629 | 3.6359 | 3.6359 | 0.001763
2019-03-20 17:55 | 1553118873.0852 | 333 | 0.19811 | 0.009123 | 0.000191 | 102.6241 | 100.4102 | 105.6615 | 21.0655 | 0.004273 | 0.003832 | 152.5398 | 0.26247 | 8.4035 | 0.004737 | 0.006422 | 0.006272 | 7.3e-05 | 0.001099 | 0.28479 | 1.3211 | 2.548 | 100.58 | 20.74 | 19.303 | 124.7446 | 136.1663 | 113.9623 | 0.004702 | 0.912 | 0.021472 | 23.5444 | 0.19507 | 0.15615 | 0.057529 | 3.695 | 3.695 | 0.001737
2019-03-20 17:55 | 1553118874.0853 | 334 | 0.19816 | 0.009154 | 0.000193 | 102.7179 | 100.1927 | 105.6432 | 21.0652 | 0.004305 | 0.003877 | 152.3909 | 0.26226 | 8.402 | 0.004761 | 0.006365 | 0.006252 | 7.3e-05 | 0.0011 | 0.29192 | 1.316 | 2.548 | 100.58 | 20.73 | 19.308 | 125.5391 | 136.7464 | 115.0555 | 0.004701 | 0.90943 | 0.02144 | 23.5756 | 0.22264 | 0.15602 | 0.057187 | 3.6675 | 3.6675 | 0.001748
2019-03-20 17:55 | 1553118875.0849 | 335 | 0.19797 | 0.00919 | 0.000193 | 102.8221 | 100.3878 | 105.5393 | 21.0655 | 0.004328 | 0.003877 | 152.2315 | 0.26242 | 8.4018 | 0.004753 | 0.006677 | 0.006511 | 7e-05 | 0.001101 | 0.30596 | 1.3139 | 2.5481 | 100.58 | 20.73 | 19.313 | 126.0876 | 136.548 | 115.0621 | 0.004702 | 0.90941 | 0.02143 | 23.5642 | 0.21431 | 0.15611 | 0.058291 | 3.7362 | 3.7362 | 0.001717
2019-03-20 17:55 | 1553118876.086 | 336 | 0.19798 | 0.009248 | 0.000195 | 102.8887 | 100.9155 | 105.8688 | 21.0657 | 0.004292 | 0.003911 | 152.0889 | 0.26263 | 8.4035 | 0.00473 | 0.005943 | 0.00585 | 7.3e-05 | 0.001104 | 0.28479 | 1.3449 | 2.5481 | 100.58 | 20.73 | 19.324 | 125.2156 | 135.9974 | 115.8868 | 0.004705 | 0.90749 | 0.02144 | 23.626 | 0.1732 | 0.15624 | 0.058137 | 3.7281 | 3.7281 | 0.001722
2019-03-20 17:55 | 1553118877.0844 | 337 | 0.19798 | 0.009296 | 0.000197 | 102.9676 | 101.0448 | 105.8305 | 21.0659 | 0.004305 | 0.003889 | 151.9468 | 0.26292 | 8.4033 | 0.004797 | 0.005904 | 0.005775 | 7.1e-05 | 0.001109 | 0.29214 | 1.362 | 2.5482 | 100.58 | 20.73 | 19.329 | 125.5253 | 137.6528 | 115.3483 | 0.004706 | 0.90874 | 0.021451 | 23.6054 | 0.16861 | 0.15641 | 0.058078 | 3.7261 | 3.7261 | 0.001725
2019-03-20 17:55 | 1553118878.0844 | 338 | 0.19795 | 0.009357 | 0.000199 | 103.0811 | 101.1513 | 105.6583 | 21.0657 | 0.004343 | 0.003885 | 151.7956 | 0.26208 | 8.4073 | 0.004672 | 0.006831 | 0.006726 | 7.1e-05 | 0.00111 | 0.29679 | 1.3032 | 2.5481 | 100.58 | 20.73 | 19.334 | 126.4532 | 134.5605 | 115.2562 | 0.004707 | 0.90896 | 0.021436 | 23.5834 | 0.16904 | 0.15591 | 0.058198 | 3.7311 | 3.7311 | 0.001717
2019-03-20 17:55 | 1553118879.0852 | 339 | 0.19792 | 0.009407 | 0.000201 | 103.2244 | 101.2934 | 105.4231 | 21.0653 | 0.004345 | 0.003935 | 151.6505 | 0.26298 | 8.4045 | 0.004763 | 0.006025 | 0.005872 | 7.2e-05 | 0.001106 | 0.27129 | 1.297 | 2.5481 | 100.58 | 20.73 | 19.337 | 126.5205 | 136.7936 | 116.4848 | 0.004708 | 0.90609 | 0.021379 | 23.5943 | 0.16891 | 0.15644 | 0.05831 | 3.7281 | 3.7281 | 0.001725
2019-03-20 17:55 | 1553118880.0852 | 340 | 0.19793 | 0.009444 | 0.000201 | 103.3916 | 101.043 | 105.0853 | 21.0664 | 0.004333 | 0.003948 | 151.4768 | 0.26217 | 8.404 | 0.004778 | 0.005975 | 0.005804 | 6.9e-05 | 0.001115 | 0.29856 | 1.314 | 2.5481 | 100.58 | 20.73 | 19.339 | 126.2184 | 137.1625 | 116.808 | 0.004708 | 0.90534 | 0.021336 | 23.5663 | 0.20553 | 0.15597 | 0.058179 | 3.7123 | 3.7123 | 0.001727
2019-03-20 17:55 | 1553118881.0852 | 341 | 0.19789 | 0.009472 | 0.0002 | 103.5931 | 101.336 | 105.09 | 21.0661 | 0.004391 | 0.003942 | 151.3393 | 0.2629 | 8.4046 | 0.004838 | 0.006622 | 0.006569 | 7e-05 | 0.001111 | 0.28407 | 1.3358 | 2.548 | 100.58 | 20.73 | 19.339 | 127.6548 | 138.6573 | 116.6606 | 0.004708 | 0.90568 | 0.02134 | 23.5624 | 0.19704 | 0.1564 | 0.058445 | 3.73 | 3.73 | 0.001723
2019-03-20 17:55 | 1553118882.0856 | 342 | 0.19785 | 0.009498 | 0.000198 | 103.7056 | 101.4735 | 105.14 | 21.0665 | 0.004395 | 0.003911 | 151.2033 | 0.26246 | 8.4011 | 0.004754 | 0.006426 | 0.006211 | 7.2e-05 | 0.001113 | 0.2975 | 1.3442 | 2.548 | 100.58 | 20.73 | 19.338 | 127.7293 | 136.5707 | 115.9049 | 0.004708 | 0.90744 | 0.021366 | 23.5451 | 0.19462 | 0.15614 | 0.058612 | 3.7452 | 3.7452 | 0.001713
2019-03-20 17:55 | 1553118883.0858 | 343 | 0.19777 | 0.009517 | 0.000196 | 103.8502 | 101.7015 | 105.0064 | 21.0663 | 0.004366 | 0.003931 | 151.0477 | 0.26228 | 8.4017 | 0.00479 | 0.006279 | 0.006156 | 7e-05 | 0.001113 | 0.27489 | 1.3475 | 2.5479 | 100.58 | 20.73 | 19.336 | 127.0409 | 137.4579 | 116.3891 | 0.004708 | 0.90632 | 0.021339 | 23.5448 | 0.18701 | 0.15603 | 0.059072 | 3.7699 | 3.7699 | 0.001701
2019-03-20 17:55 | 1553118884.0858 | 344 | 0.19769 | 0.009539 | 0.000194 | 103.953 | 101.9138 | 104.8473 | 21.0668 | 0.004425 | 0.004021 | 150.8961 | 0.26225 | 8.4046 | 0.004765 | 0.006329 | 0.006213 | 7.1e-05 | 0.001116 | 0.27721 | 1.304 | 2.5481 | 100.58 | 20.74 | 19.33 | 128.482 | 136.8504 | 118.5836 | 0.004709 | 0.90124 | 0.021263 | 23.5931 | 0.17721 | 0.15601 | 0.059525 | 3.7853 | 3.7853 | 0.001694
2019-03-20 17:55 | 1553118885.0855 | 345 | 0.19763 | 0.009577 | 0.000193 | 104.0873 | 102.5959 | 104.6255 | 21.0671 | 0.004404 | 0.004034 | 150.7669 | 0.26221 | 8.4028 | 0.004724 | 0.006441 | 0.006259 | 7e-05 | 0.001113 | 0.28711 | 1.3107 | 2.5481 | 100.58 | 20.74 | 19.326 | 127.9618 | 135.8451 | 118.9034 | 0.004708 | 0.9005 | 0.021232 | 23.5778 | 0.12909 | 0.15599 | 0.059826 | 3.7989 | 3.7989 | 0.001688
2019-03-20 17:55 | 1553118886.0855 | 346 | 0.19743 | 0.009623 | 0.000192 | 104.1768 | 102.5623 | 104.2347 | 21.066 | 0.004438 | 0.004 | 150.6091 | 0.26244 | 8.4033 | 0.00483 | 0.006476 | 0.006309 | 7.3e-05 | 0.001115 | 0.32693 | 1.339 | 2.5482 | 100.58 | 20.74 | 19.32 | 128.8087 | 138.4583 | 118.0718 | 0.004707 | 0.90242 | 0.021215 | 23.5087 | 0.13971 | 0.15613 | 0.060948 | 3.8667 | 3.8667 | 0.001659
2019-03-20 17:55 | 1553118887.0849 | 347 | 0.19742 | 0.009675 | 0.000192 | 104.2897 | 102.588 | 104.1826 | 21.0658 | 0.004523 | 0.004092 | 150.4404 | 0.26227 | 8.4018 | 0.004846 | 0.006421 | 0.006396 | 7.2e-05 | 0.001118 | 0.26753 | 1.3026 | 2.548 | 100.58 | 20.74 | 19.316 | 130.8836 | 138.8588 | 120.3194 | 0.004706 | 0.89726 | 0.021149 | 23.5703 | 0.14715 | 0.15602 | 0.060998 | 3.8579 | 3.8579 | 0.001662
2019-03-20 17:55 | 1553118888.085 | 348 | 0.19725 | 0.009729 | 0.000191 | 104.3143 | 102.6654 | 104.0018 | 21.0666 | 0.004488 | 0.0041 | 150.3048 | 0.26262 | 8.4048 | 0.00487 | 0.005599 | 0.00542 | 7.3e-05 | 0.001122 | 0.27184 | 1.2795 | 2.548 | 100.58 | 20.74 | 19.312 | 130.0303 | 139.44 | 120.5152 | 0.004705 | 0.89682 | 0.021125 | 23.5557 | 0.14252 | 0.15624 | 0.061951 | 3.9137 | 3.9137 | 0.001641
2019-03-20 17:55 | 1553118889.0857 | 349 | 0.19722 | 0.009796 | 0.00019 | 104.4524 | 102.6792 | 103.8903 | 21.0667 | 0.004492 | 0.004033 | 150.142 | 0.26214 | 8.4033 | 0.004866 | 0.006662 | 0.006501 | 7.2e-05 | 0.00112 | 0.29374 | 1.308 | 2.548 | 100.58 | 20.74 | 19.308 | 130.1317 | 139.3512 | 118.8759 | 0.004704 | 0.90057 | 0.021158 | 23.494 | 0.15314 | 0.15595 | 0.062044 | 3.9256 | 3.9256 | 0.001633
2019-03-20 17:55 | 1553118890.0857 | 350 | 0.19712 | 0.00986 | 0.000189 | 104.6115 | 102.5529 | 103.8999 | 21.0668 | 0.004522 | 0.004117 | 149.9999 | 0.26282 | 8.4005 | 0.004886 | 0.00638 | 0.006383 | 7.1e-05 | 0.001116 | 0.3183 | 1.3133 | 2.5481 | 100.58 | 20.74 | 19.307 | 130.8646 | 139.8318 | 120.9281 | 0.004703 | 0.89588 | 0.021104 | 23.5565 | 0.17777 | 0.15635 | 0.06257 | 3.9487 | 3.9487 | 0.001627
2019-03-20 17:55 | 1553118891.0848 | 351 | 0.19704 | 0.00992 | 0.000189 | 104.6487 | 102.7147 | 103.6514 | 21.0665 | 0.004513 | 0.004043 | 149.8235 | 0.26211 | 8.4045 | 0.004867 | 0.006079 | 0.005888 | 7.3e-05 | 0.001114 | 0.27649 | 1.2977 | 2.5482 | 100.58 | 20.74 | 19.304 | 130.6406 | 139.3823 | 119.1264 | 0.004703 | 0.89999 | 0.021127 | 23.4744 | 0.16685 | 0.15593 | 0.062966 | 3.9779 | 3.9779 | 0.001611
2019-03-20 17:55 | 1553118892.0852 | 352 | 0.19687 | 0.009981 | 0.000187 | 104.7839 | 102.8441 | 103.6786 | 21.0665 | 0.004574 | 0.004112 | 149.664 | 0.26231 | 8.4033 | 0.004945 | 0.006041 | 0.00588 | 7.3e-05 | 0.001112 | 0.31006 | 1.3014 | 2.5482 | 100.58 | 20.74 | 19.303 | 132.1354 | 141.29 | 120.8009 | 0.004702 | 0.89617 | 0.021085 | 23.5276 | 0.16714 | 0.15604 | 0.063871 | 4.0269 | 4.0269 | 0.001593
2019-03-20 17:55 | 1553118893.0844 | 353 | 0.19675 | 0.010042 | 0.000187 | 104.7496 | 103.1233 | 103.3932 | 21.0673 | 0.004601 | 0.004108 | 149.4885 | 0.26243 | 8.4052 | 0.004934 | 0.006538 | 0.006435 | 7.3e-05 | 0.001109 | 0.26272 | 1.3908 | 2.5481 | 100.58 | 20.73 | 19.301 | 132.8146 | 141.034 | 120.6992 | 0.004699 | 0.8964 | 0.021058 | 23.4921 | 0.13996 | 0.15612 | 0.064553 | 4.0646 | 4.0646 | 0.001579
2019-03-20 17:55 | 1553118894.0843 | 354 | 0.19659 | 0.010109 | 0.000187 | 104.933 | 103.0724 | 103.0686 | 21.0654 | 0.004531 | 0.00414 | 149.3287 | 0.26237 | 8.4032 | 0.004943 | 0.00725 | 0.007102 | 7.2e-05 | 0.001117 | 0.27455 | 1.3281 | 2.5481 | 100.58 | 20.74 | 19.297 | 131.0858 | 141.2512 | 121.495 | 0.004701 | 0.89459 | 0.021004 | 23.4789 | 0.16002 | 0.15608 | 0.065402 | 4.1073 | 4.1073 | 0.001562
2019-03-20 17:55 | 1553118895.085 | 355 | 0.19656 | 0.010182 | 0.000185 | 104.9469 | 102.9162 | 102.8587 | 21.067 | 0.004538 | 0.004105 | 149.1451 | 0.26253 | 8.4036 | 0.004965 | 0.005837 | 0.005716 | 7.1e-05 | 0.00112 | 0.29999 | 1.3344 | 2.548 | 100.58 | 20.73 | 19.295 | 131.2678 | 141.7937 | 120.6331 | 0.004698 | 0.89655 | 0.021006 | 23.4294 | 0.17477 | 0.15618 | 0.065471 | 4.112 | 4.112 | 0.001561
2019-03-20 17:55 | 1553118896.0839 | 356 | 0.19649 | 0.010259 | 0.000184 | 105.0184 | 102.927 | 102.6137 | 21.0662 | 0.004648 | 0.004113 | 148.9693 | 0.26243 | 8.4014 | 0.004965 | 0.006019 | 0.00593 | 7.1e-05 | 0.001118 | 0.31349 | 1.3261 | 2.5481 | 100.58 | 20.73 | 19.29 | 133.9659 | 141.8057 | 120.8394 | 0.004696 | 0.89608 | 0.020975 | 23.4076 | 0.17992 | 0.15612 | 0.065802 | 4.1267 | 4.1267 | 0.001555
2019-03-20 17:55 | 1553118897.0854 | 357 | 0.19641 | 0.010334 | 0.000184 | 105.162 | 102.9648 | 102.2696 | 21.0667 | 0.004601 | 0.004139 | 148.801 | 0.26225 | 8.4027 | 0.005016 | 0.006407 | 0.006316 | 7.1e-05 | 0.001115 | 0.28783 | 1.315 | 2.5482 | 100.58 | 20.73 | 19.287 | 132.8058 | 143.0631 | 121.4662 | 0.004696 | 0.89465 | 0.020923 | 23.3869 | 0.18886 | 0.15601 | 0.066182 | 4.1402 | 4.1402 | 0.001549
2019-03-20 17:55 | 1553118898.0855 | 358 | 0.19642 | 0.010412 | 0.000183 | 105.1882 | 103.1719 | 102.0428 | 21.0674 | 0.004598 | 0.004146 | 148.6157 | 0.26228 | 8.4035 | 0.005026 | 0.005213 | 0.005033 | 7.1e-05 | 0.00112 | 0.29856 | 1.3502 | 2.5481 | 100.58 | 20.73 | 19.284 | 132.7435 | 143.3 | 121.6273 | 0.004695 | 0.89429 | 0.020896 | 23.3657 | 0.17312 | 0.15603 | 0.066085 | 4.1288 | 4.1288 | 0.001553
2019-03-20 17:55 | 1553118899.0853 | 359 | 0.19636 | 0.010479 | 0.000183 | 105.3159 | 103.5564 | 102.1513 | 21.066 | 0.004642 | 0.00423 | 148.4591 | 0.26177 | 8.402 | 0.005059 | 0.006477 | 0.006329 | 6.7e-05 | 0.001112 | 0.29551 | 1.3493 | 2.5481 | 100.58 | 20.74 | 19.28 | 133.8276 | 144.1221 | 123.6979 | 0.004697 | 0.88962 | 0.020852 | 23.4393 | 0.1507 | 0.15572 | 0.066311 | 4.1342 | 4.1342 | 0.001548
2019-03-20 17:55 | 1553118900.0843 | 360 | 0.19619 | 0.010543 | 0.000184 | 105.4084 | 103.6603 | 102.1334 | 21.066 | 0.004656 | 0.00426 | 148.297 | 0.2613 | 8.4052 | 0.005033 | 0.006958 | 0.006819 | 6.7e-05 | 0.001115 | 0.27041 | 1.3133 | 2.548 | 100.58 | 20.74 | 19.279 | 134.1663 | 143.4816 | 124.4379 | 0.004697 | 0.88797 | 0.020831 | 23.4591 | 0.14958 | 0.15544 | 0.067257 | 4.1886 | 4.1886 | 0.001525
2019-03-20 17:55 | 1553118901.0843 | 361 | 0.19595 | 0.010605 | 0.000185 | 105.532 | 103.5037 | 102.1146 | 21.0662 | 0.004778 | 0.004237 | 148.132 | 0.26256 | 8.4035 | 0.005042 | 0.006473 | 0.006451 | 6.9e-05 | 0.001111 | 0.26808 | 1.2934 | 2.548 | 100.58 | 20.74 | 19.277 | 137.1766 | 143.7012 | 123.8537 | 0.004696 | 0.88928 | 0.020844 | 23.4397 | 0.17358 | 0.1562 | 0.068605 | 4.2749 | 4.2749 | 0.001502
2019-03-20 17:55 | 1553118902.0853 | 362 | 0.19583 | 0.01066 | 0.000185 | 105.529 | 103.525 | 102.0381 | 21.0669 | 0.004688 | 0.004219 | 147.9587 | 0.26147 | 8.4015 | 0.005049 | 0.00667 | 0.00641 | 6.8e-05 | 0.001118 | 0.29407 | 1.3062 | 2.5483 | 100.58 | 20.74 | 19.275 | 134.9497 | 143.8781 | 123.4276 | 0.004696 | 0.89023 | 0.020848 | 23.4184 | 0.17149 | 0.15555 | 0.069269 | 4.3167 | 4.3167 | 0.001481
2019-03-20 17:55 | 1553118903.0854 | 363 | 0.19582 | 0.010735 | 0.000184 | 105.5869 | 103.574 | 102.0538 | 21.0665 | 0.004728 | 0.004214 | 147.8011 | 0.2618 | 8.4032 | 0.005173 | 0.006521 | 0.006396 | 6.9e-05 | 0.001114 | 0.2656 | 1.2915 | 2.5481 | 100.58 | 20.74 | 19.274 | 135.9455 | 146.9504 | 123.3025 | 0.004695 | 0.89051 | 0.020853 | 23.4165 | 0.17217 | 0.15574 | 0.069238 | 4.3158 | 4.3158 | 0.001483
2019-03-20 17:55 | 1553118904.0849 | 364 | 0.19573 | 0.010808 | 0.000183 | 105.703 | 103.5725 | 102.0156 | 21.0661 | 0.004717 | 0.004249 | 147.6485 | 0.26217 | 8.4025 | 0.005126 | 0.0062 | 0.006105 | 6.8e-05 | 0.001117 | 0.26128 | 1.2927 | 2.5481 | 100.58 | 20.74 | 19.274 | 135.6648 | 145.7842 | 124.1567 | 0.004695 | 0.8886 | 0.020826 | 23.4373 | 0.18213 | 0.15597 | 0.069652 | 4.3361 | 4.3361 | 0.001478
2019-03-20 17:55 | 1553118905.0856 | 365 | 0.19572 | 0.010893 | 0.000183 | 105.7554 | 103.8977 | 102.0886 | 21.0675 | 0.004693 | 0.004315 | 147.4701 | 0.26202 | 8.4046 | 0.005084 | 0.006653 | 0.006463 | 6.8e-05 | 0.001113 | 0.29446 | 1.3053 | 2.548 | 100.58 | 20.74 | 19.273 | 135.0814 | 144.7531 | 125.7753 | 0.004695 | 0.88499 | 0.020791 | 23.4934 | 0.15852 | 0.15588 | 0.069644 | 4.3284 | 4.3284 | 0.00148
2019-03-20 17:55 | 1553118906.0855 | 366 | 0.1957 | 0.010971 | 0.000181 | 105.8256 | 103.8648 | 102.1477 | 21.0669 | 0.004752 | 0.004288 | 147.3104 | 0.26239 | 8.4028 | 0.005096 | 0.00692 | 0.006794 | 6.6e-05 | 0.001107 | 0.2756 | 1.3621 | 2.548 | 100.58 | 20.74 | 19.272 | 136.5407 | 145.0409 | 125.1182 | 0.004695 | 0.88645 | 0.020815 | 23.4808 | 0.16729 | 0.1561 | 0.069697 | 4.3366 | 4.3366 | 0.001479
2019-03-20 17:55 | 1553118907.0856 | 367 | 0.19571 | 0.011042 | 0.000179 | 105.8107 | 103.8439 | 102.1236 | 21.0668 | 0.004749 | 0.004285 | 147.1654 | 0.26279 | 8.4018 | 0.005141 | 0.007186 | 0.007114 | 6.7e-05 | 0.001112 | 0.27113 | 1.3474 | 2.5481 | 100.58 | 20.74 | 19.27 | 136.4469 | 146.1627 | 125.033 | 0.004694 | 0.88664 | 0.020814 | 23.4755 | 0.16783 | 0.15633 | 0.069552 | 4.3276 | 4.3276 | 0.001485
2019-03-20 17:55 | 1553118908.084 | 368 | 0.19553 | 0.011107 | 0.000178 | 105.8719 | 103.8588 | 102.0081 | 21.0667 | 0.004803 | 0.004291 | 146.9847 | 0.26191 | 8.4022 | 0.005205 | 0.006369 | 0.006262 | 6.4e-05 | 0.001112 | 0.29679 | 1.3263 | 2.5479 | 100.58 | 20.74 | 19.269 | 137.8003 | 147.7513 | 125.1833 | 0.004694 | 0.88631 | 0.020799 | 23.4667 | 0.17171 | 0.15581 | 0.070539 | 4.3855 | 4.3855 | 0.00146
2019-03-20 17:55 | 1553118909.0852 | 369 | 0.19546 | 0.011175 | 0.000177 | 105.9485 | 103.7317 | 101.8447 | 21.0673 | 0.004817 | 0.004343 | 146.7987 | 0.26236 | 8.4015 | 0.005187 | 0.006488 | 0.006379 | 7.1e-05 | 0.001117 | 0.29048 | 1.3382 | 2.5482 | 100.58 | 20.74 | 19.263 | 138.1382 | 147.2939 | 126.4764 | 0.004693 | 0.88344 | 0.020748 | 23.4859 | 0.18913 | 0.15608 | 0.070862 | 4.3948 | 4.3948 | 0.00146
2019-03-20 17:55 | 1553118910.0838 | 370 | 0.1954 | 0.011234 | 0.000179 | 106.0662 | 104.1989 | 101.8975 | 21.0669 | 0.004858 | 0.004359 | 146.6486 | 0.26213 | 8.4031 | 0.00519 | 0.006549 | 0.006443 | 7e-05 | 0.001115 | 0.26465 | 1.3321 | 2.5482 | 100.58 | 20.74 | 19.26 | 139.1431 | 147.3717 | 126.851 | 0.004692 | 0.88261 | 0.020744 | 23.503 | 0.15887 | 0.15594 | 0.071122 | 4.4099 | 4.4099 | 0.001453
2019-03-20 17:55 | 1553118911.084 | 371 | 0.19535 | 0.011301 | 0.00018 | 106.1817 | 104.0909 | 102.1172 | 21.0668 | 0.004829 | 0.00433 | 146.4979 | 0.26266 | 8.4033 | 0.005183 | 0.006931 | 0.006748 | 6.9e-05 | 0.001118 | 0.28119 | 1.3474 | 2.5481 | 100.58 | 20.74 | 19.256 | 138.4423 | 147.2066 | 126.151 | 0.004691 | 0.88416 | 0.020785 | 23.5077 | 0.17788 | 0.15625 | 0.071354 | 4.4328 | 4.4328 | 0.001449
2019-03-20 17:55 | 1553118912.085 | 372 | 0.19528 | 0.01136 | 0.000182 | 106.2092 | 104.1458 | 101.9693 | 21.0666 | 0.004904 | 0.004423 | 146.3169 | 0.26192 | 8.4014 | 0.005222 | 0.006548 | 0.006465 | 6.4e-05 | 0.001116 | 0.30596 | 1.3649 | 2.5482 | 100.58 | 20.74 | 19.246 | 140.2765 | 148.1658 | 128.4186 | 0.004689 | 0.87917 | 0.020711 | 23.5573 | 0.17548 | 0.15582 | 0.071706 | 4.4387 | 4.4387 | 0.001443
2019-03-20 17:55 | 1553118913.0849 | 373 | 0.19515 | 0.011418 | 0.000182 | 106.2648 | 104.32 | 101.8226 | 21.0669 | 0.004897 | 0.004367 | 146.1555 | 0.26227 | 8.4041 | 0.00516 | 0.006456 | 0.006283 | 6.5e-05 | 0.001115 | 0.26272 | 1.3104 | 2.5481 | 100.58 | 20.73 | 19.241 | 140.1181 | 146.6263 | 127.0461 | 0.004684 | 0.88218 | 0.020731 | 23.5001 | 0.16522 | 0.15603 | 0.072418 | 4.4869 | 4.4869 | 0.001429
2019-03-20 17:55 | 1553118914.0858 | 374 | 0.19506 | 0.011484 | 0.000184 | 106.4159 | 104.4729 | 101.6867 | 21.067 | 0.00491 | 0.00442 | 145.9734 | 0.26298 | 8.3999 | 0.005135 | 0.006319 | 0.006233 | 6.7e-05 | 0.001125 | 0.28528 | 1.3004 | 2.5482 | 100.58 | 20.73 | 19.236 | 140.4413 | 145.9926 | 128.3664 | 0.004683 | 0.87928 | 0.020683 | 23.5231 | 0.16482 | 0.15644 | 0.072841 | 4.5025 | 4.5025 | 0.001428
2019-03-20 17:55 | 1553118915.0841 | 375 | 0.19506 | 0.011541 | 0.000185 | 106.4883 | 104.4497 | 101.8534 | 21.0665 | 0.004955 | 0.004348 | 145.8187 | 0.26143 | 8.4059 | 0.005182 | 0.006969 | 0.006758 | 6.8e-05 | 0.001121 | 0.29231 | 1.3507 | 2.5481 | 100.58 | 20.73 | 19.226 | 141.5493 | 147.1716 | 126.5991 | 0.004681 | 0.88317 | 0.020746 | 23.4905 | 0.17289 | 0.15552 | 0.072768 | 4.5115 | 4.5115 | 0.001417
2019-03-20 17:55 | 1553118916.0841 | 376 | 0.19504 | 0.011608 | 0.000186 | 106.541 | 104.6 | 101.9029 | 21.0676 | 0.004919 | 0.004424 | 145.6552 | 0.26161 | 8.4033 | 0.005121 | 0.006706 | 0.006613 | 6.3e-05 | 0.001121 | 0.2625 | 1.3142 | 2.5481 | 100.58 | 20.73 | 19.22 | 140.6586 | 145.6593 | 128.462 | 0.004679 | 0.87907 | 0.020703 | 23.5509 | 0.16446 | 0.15563 | 0.072816 | 4.5051 | 4.5051 | 0.00142
2019-03-20 17:55 | 1553118917.0842 | 377 | 0.19509 | 0.011661 | 0.000188 | 106.5776 | 104.5869 | 101.8992 | 21.0661 | 0.004916 | 0.004417 | 145.4884 | 0.26169 | 8.403 | 0.00522 | 0.00595 | 0.005851 | 6.6e-05 | 0.001129 | 0.27168 | 1.3388 | 2.5481 | 100.58 | 20.73 | 19.211 | 140.5745 | 148.1233 | 128.2786 | 0.004677 | 0.87947 | 0.020707 | 23.5451 | 0.16865 | 0.15568 | 0.072457 | 4.4839 | 4.4839 | 0.001427
2019-03-20 17:55 | 1553118918.0842 | 378 | 0.19507 | 0.011706 | 0.00019 | 106.6791 | 104.8089 | 101.9989 | 21.0671 | 0.004919 | 0.004452 | 145.3322 | 0.26182 | 8.4042 | 0.005237 | 0.006546 | 0.006373 | 6.7e-05 | 0.001128 | 0.28727 | 1.2934 | 2.548 | 100.58 | 20.73 | 19.207 | 140.6509 | 148.5398 | 129.1512 | 0.004676 | 0.87757 | 0.020695 | 23.5822 | 0.1582 | 0.15576 | 0.072509 | 4.4843 | 4.4843 | 0.001428
2019-03-20 17:55 | 1553118919.0859 | 379 | 0.19508 | 0.011742 | 0.000193 | 106.7913 | 104.7924 | 101.9992 | 21.0673 | 0.005009 | 0.004449 | 145.1918 | 0.26183 | 8.4045 | 0.00521 | 0.0071 | 0.006898 | 6.8e-05 | 0.001126 | 0.29662 | 1.2924 | 2.5481 | 100.58 | 20.73 | 19.203 | 142.8797 | 147.8623 | 129.0715 | 0.004675 | 0.87774 | 0.020697 | 23.5799 | 0.16901 | 0.15576 | 0.072405 | 4.4782 | 4.4782 | 0.00143
2019-03-20 17:55 | 1553118920.0843 | 380 | 0.19515 | 0.011766 | 0.000195 | 106.956 | 104.7476 | 101.9273 | 21.0673 | 0.004994 | 0.004411 | 144.9993 | 0.26212 | 8.4044 | 0.0052 | 0.007503 | 0.007376 | 6.7e-05 | 0.001124 | 0.31293 | 1.3425 | 2.5481 | 100.58 | 20.73 | 19.195 | 142.524 | 147.6199 | 128.127 | 0.004673 | 0.8798 | 0.020714 | 23.5439 | 0.18662 | 0.15594 | 0.071988 | 4.4561 | 4.4561 | 0.001438
2019-03-20 17:55 | 1553118921.0853 | 381 | 0.19511 | 0.011778 | 0.000198 | 107.009 | 104.9326 | 101.6881 | 21.0668 | 0.00497 | 0.00451 | 144.82 | 0.26205 | 8.4019 | 0.00516 | 0.007158 | 0.006961 | 6.6e-05 | 0.001123 | 0.2907 | 1.3317 | 2.5481 | 100.58 | 20.73 | 19.191 | 141.9123 | 146.6272 | 130.5797 | 0.004672 | 0.87446 | 0.020627 | 23.588 | 0.17527 | 0.15589 | 0.072202 | 4.4503 | 4.4503 | 0.00144
2019-03-20 17:55 | 1553118922.0858 | 382 | 0.19491 | 0.011794 | 0.0002 | 107.1574 | 105.3568 | 101.7339 | 21.0665 | 0.005028 | 0.004475 | 144.6558 | 0.26208 | 8.4046 | 0.00514 | 0.007689 | 0.007589 | 6.4e-05 | 0.001126 | 0.24889 | 1.3353 | 2.5481 | 100.58 | 20.72 | 19.185 | 143.3546 | 146.1384 | 129.7103 | 0.004668 | 0.87635 | 0.020654 | 23.5679 | 0.15158 | 0.15591 | 0.073393 | 4.5291 | 4.5291 | 0.001415
2019-03-20 17:55 | 1553118923.0855 | 383 | 0.19491 | 0.011802 | 0.000202 | 107.2208 | 105.2906 | 101.817 | 21.067 | 0.004973 | 0.004465 | 144.4712 | 0.26184 | 8.4014 | 0.005065 | 0.0074 | 0.007279 | 6.5e-05 | 0.001132 | 0.29391 | 1.3471 | 2.5481 | 100.58 | 20.72 | 19.182 | 141.9953 | 144.2655 | 129.4532 | 0.004667 | 0.87691 | 0.020669 | 23.57 | 0.16249 | 0.15577 | 0.073334 | 4.5287 | 4.5287 | 0.001414
2019-03-20 17:55 | 1553118924.0836 | 384 | 0.19493 | 0.011829 | 0.000205 | 107.3343 | 105.1946 | 101.6022 | 21.0672 | 0.005037 | 0.00445 | 144.3051 | 0.26156 | 8.4032 | 0.005125 | 0.006793 | 0.006595 | 6.5e-05 | 0.001134 | 0.2891 | 1.3476 | 2.5481 | 100.58 | 20.72 | 19.18 | 143.5658 | 145.7531 | 129.097 | 0.004667 | 0.87768 | 0.020656 | 23.5347 | 0.18011 | 0.1556 | 0.07322 | 4.5188 | 4.5188 | 0.001415
2019-03-20 17:55 | 1553118925.0845 | 385 | 0.19497 | 0.011848 | 0.000207 | 107.3289 | 105.4659 | 101.2725 | 21.0674 | 0.005039 | 0.004576 | 144.1366 | 0.26174 | 8.404 | 0.004991 | 0.006581 | 0.006507 | 6.6e-05 | 0.001135 | 0.29214 | 1.3107 | 2.548 | 100.58 | 20.72 | 19.176 | 143.6351 | 142.439 | 132.1851 | 0.004666 | 0.871 | 0.020544 | 23.5865 | 0.15661 | 0.15571 | 0.072939 | 4.477 | 4.477 | 0.001429
2019-03-20 17:55 | 1553118926.0839 | 386 | 0.19505 | 0.011876 | 0.000209 | 107.3631 | 105.4267 | 101.2502 | 21.068 | 0.005092 | 0.004473 | 143.9538 | 0.26156 | 8.4018 | 0.004911 | 0.00639 | 0.006251 | 6.9e-05 | 0.001135 | 0.28495 | 1.3493 | 2.5482 | 100.58 | 20.72 | 19.173 | 144.9281 | 140.4585 | 129.6577 | 0.004665 | 0.87646 | 0.020606 | 23.5103 | 0.1628 | 0.1556 | 0.07246 | 4.4611 | 4.4611 | 0.001434
2019-03-20 17:55 | 1553118927.0838 | 387 | 0.1951 | 0.011894 | 0.000211 | 107.4669 | 105.3826 | 101.11 | 21.0667 | 0.005072 | 0.004518 | 143.7881 | 0.26163 | 8.4031 | 0.004933 | 0.006718 | 0.006608 | 6.6e-05 | 0.001135 | 0.31509 | 1.3373 | 2.5481 | 100.58 | 20.72 | 19.169 | 144.4483 | 140.9908 | 130.7646 | 0.004664 | 0.87406 | 0.020563 | 23.5263 | 0.17519 | 0.15564 | 0.07209 | 4.4291 | 4.4291 | 0.001444
2019-03-20 17:55 | 1553118928.0838 | 388 | 0.1951 | 0.011902 | 0.000213 | 107.5543 | 105.3086 | 101.0901 | 21.0668 | 0.005107 | 0.004471 | 143.6082 | 0.26133 | 8.4023 | 0.004944 | 0.00647 | 0.006293 | 6.6e-05 | 0.001131 | 0.2912 | 1.341 | 2.5482 | 100.58 | 20.72 | 19.166 | 145.2991 | 141.2824 | 129.6054 | 0.004663 | 0.87657 | 0.020591 | 23.4902 | 0.18874 | 0.15546 | 0.072129 | 4.4374 | 4.4374 | 0.00144
2019-03-20 17:55 | 1553118929.0848 | 389 | 0.19511 | 0.01189 | 0.000214 | 107.5439 | 105.2604 | 101.2685 | 21.0674 | 0.005084 | 0.004498 | 143.4162 | 0.26121 | 8.4039 | 0.004872 | 0.006826 | 0.006657 | 6.7e-05 | 0.001137 | 0.28385 | 1.3417 | 2.5481 | 100.58 | 20.72 | 19.164 | 144.7361 | 139.4883 | 130.2728 | 0.004663 | 0.87513 | 0.020592 | 23.5304 | 0.19197 | 0.1554 | 0.072071 | 4.434 | 4.434 | 0.00144
2019-03-20 17:55 | 1553118930.0848 | 390 | 0.19516 | 0.011878 | 0.000215 | 107.6558 | 105.673 | 101.4523 | 21.0673 | 0.005091 | 0.004549 | 143.2506 | 0.26181 | 8.4035 | 0.004935 | 0.006113 | 0.006018 | 7.1e-05 | 0.00114 | 0.28838 | 1.3063 | 2.548 | 100.58 | 20.72 | 19.161 | 144.9137 | 141.0451 | 131.5363 | 0.004662 | 0.87239 | 0.020578 | 23.5886 | 0.16628 | 0.15575 | 0.071778 | 4.4131 | 4.4131 | 0.001451
2019-03-20 17:55 | 1553118931.0852 | 391 | 0.19515 | 0.011857 | 0.000216 | 107.6969 | 105.7081 | 101.6802 | 21.0669 | 0.00516 | 0.004551 | 143.0954 | 0.26226 | 8.4006 | 0.004956 | 0.006371 | 0.006257 | 7.3e-05 | 0.001142 | 0.27129 | 1.2923 | 2.5481 | 100.58 | 20.72 | 19.159 | 146.6151 | 141.5672 | 131.5807 | 0.004662 | 0.8723 | 0.0206 | 23.6163 | 0.16672 | 0.15602 | 0.071828 | 4.4208 | 4.4208 | 0.00145
2019-03-20 17:56 | 1553118932.0849 | 392 | 0.1952 | 0.011841 | 0.000217 | 107.6827 | 105.3781 | 101.4886 | 21.0667 | 0.005119 | 0.004558 | 142.9405 | 0.26169 | 8.4036 | 0.004971 | 0.006278 | 0.006107 | 7e-05 | 0.001141 | 0.27815 | 1.3007 | 2.5481 | 100.58 | 20.72 | 19.158 | 145.6014 | 141.9443 | 131.7538 | 0.004661 | 0.87192 | 0.020577 | 23.5991 | 0.1935 | 0.15568 | 0.071571 | 4.4 | 4.4 | 0.001454
2019-03-20 17:56 | 1553118933.0854 | 393 | 0.19517 | 0.011826 | 0.000217 | 107.8538 | 105.5302 | 101.4485 | 21.0661 | 0.005135 | 0.004541 | 142.7898 | 0.26236 | 8.4035 | 0.004939 | 0.006321 | 0.006206 | 6.7e-05 | 0.00114 | 0.29518 | 1.3046 | 2.5481 | 100.58 | 20.72 | 19.157 | 146.0036 | 141.1394 | 131.3333 | 0.004661 | 0.87283 | 0.020583 | 23.5822 | 0.1948 | 0.15608 | 0.071767 | 4.4134 | 4.4134 | 0.001454
2019-03-20 17:56 | 1553118934.0841 | 394 | 0.19535 | 0.011806 | 0.000218 | 107.9537 | 105.4099 | 101.5464 | 21.0671 | 0.00513 | 0.004503 | 142.6318 | 0.26231 | 8.4032 | 0.004975 | 0.006441 | 0.006242 | 7.1e-05 | 0.00113 | 0.27256 | 1.3243 | 2.5481 | 100.58 | 20.72 | 19.156 | 145.8799 | 142.0466 | 130.3994 | 0.004661 | 0.87485 | 0.020617 | 23.5663 | 0.21329 | 0.15605 | 0.070721 | 4.3564 | 4.3564 | 0.001472
2019-03-20 17:56 | 1553118935.0851 | 395 | 0.19546 | 0.011771 | 0.000219 | 107.9909 | 105.6214 | 101.8194 | 21.067 | 0.005099 | 0.004589 | 142.4441 | 0.26157 | 8.4025 | 0.004975 | 0.00647 | 0.00639 | 7.3e-05 | 0.001129 | 0.27295 | 1.2953 | 2.548 | 100.58 | 20.72 | 19.156 | 145.1055 | 142.0466 | 132.5129 | 0.004661 | 0.87029 | 0.020591 | 23.6597 | 0.19844 | 0.15561 | 0.070111 | 4.3135 | 4.3135 | 0.001483
2019-03-20 17:56 | 1553118936.0851 | 396 | 0.1955 | 0.011741 | 0.00022 | 108.0782 | 105.7668 | 101.9562 | 21.0671 | 0.005116 | 0.004598 | 142.3035 | 0.26158 | 8.4022 | 0.004982 | 0.006999 | 0.006838 | 6.9e-05 | 0.001117 | 0.28694 | 1.3128 | 2.5482 | 100.58 | 20.72 | 19.156 | 145.5255 | 142.213 | 132.7432 | 0.004661 | 0.8698 | 0.020599 | 23.6823 | 0.19337 | 0.15562 | 0.069901 | 4.3023 | 4.3023 | 0.001487
2019-03-20 17:56 | 1553118937.0843 | 397 | 0.1956 | 0.011691 | 0.000219 | 108.1141 | 105.8999 | 102.1067 | 21.0665 | 0.005191 | 0.004572 | 142.1343 | 0.26139 | 8.4033 | 0.005093 | 0.006788 | 0.006652 | 6.6e-05 | 0.00111 | 0.27671 | 1.3098 | 2.5481 | 100.58 | 20.72 | 19.157 | 147.4032 | 144.965 | 132.1002 | 0.004661 | 0.87118 | 0.02063 | 23.681 | 0.18509 | 0.1555 | 0.06939 | 4.2776 | 4.2776 | 0.001494
2019-03-20 17:56 | 1553118938.0852 | 398 | 0.19565 | 0.011647 | 0.000219 | 108.1901 | 105.8247 | 102.1503 | 21.0659 | 0.005161 | 0.00456 | 141.9765 | 0.26188 | 8.4026 | 0.004986 | 0.006442 | 0.006344 | 6.6e-05 | 0.00111 | 0.2855 | 1.3207 | 2.5481 | 100.58 | 20.72 | 19.158 | 146.6431 | 142.3194 | 131.8049 | 0.004661 | 0.87181 | 0.020642 | 23.6774 | 0.19772 | 0.15579 | 0.069154 | 4.2656 | 4.2656 | 0.001501
2019-03-20 17:56 | 1553118939.0854 | 399 | 0.19566 | 0.011603 | 0.000218 | 108.2545 | 105.9959 | 101.9029 | 21.0668 | 0.00516 | 0.004594 | 141.8263 | 0.26134 | 8.4019 | 0.005035 | 0.006332 | 0.006247 | 6.7e-05 | 0.001103 | 0.29374 | 1.3242 | 2.5481 | 100.58 | 20.72 | 19.16 | 146.6344 | 143.5164 | 132.6366 | 0.004662 | 0.87003 | 0.020596 | 23.673 | 0.18859 | 0.15547 | 0.069129 | 4.2545 | 4.2545 | 0.001502
2019-03-20 17:56 | 1553118940.0844 | 400 | 0.19564 | 0.011562 | 0.000217 | 108.3166 | 106.4014 | 101.6196 | 21.0663 | 0.005162 | 0.004523 | 141.6652 | 0.26223 | 8.4039 | 0.005032 | 0.007125 | 0.006956 | 6.5e-05 | 0.001103 | 0.25946 | 1.3198 | 2.5482 | 100.58 | 20.72 | 19.162 | 146.6853 | 143.4604 | 130.8897 | 0.004662 | 0.87379 | 0.020612 | 23.5891 | 0.15956 | 0.156 | 0.069284 | 4.2673 | 4.2673 | 0.001503
2019-03-20 17:56 | 1553118941.0843 | 401 | 0.19567 | 0.011519 | 0.000216 | 108.3463 | 106.6059 | 101.6778 | 21.0662 | 0.005156 | 0.004576 | 141.4916 | 0.2621 | 8.4018 | 0.005042 | 0.006225 | 0.00614 | 6.7e-05 | 0.001101 | 0.29717 | 1.3737 | 2.5482 | 100.58 | 20.72 | 19.163 | 146.5274 | 143.6901 | 132.1949 | 0.004663 | 0.87098 | 0.020585 | 23.634 | 0.14485 | 0.15593 | 0.06919 | 4.256 | 4.256 | 0.001506
2019-03-20 17:56 | 1553118942.0835 | 402 | 0.19564 | 0.011484 | 0.000217 | 108.4084 | 106.5603 | 101.7785 | 21.0671 | 0.005252 | 0.004567 | 141.3228 | 0.2615 | 8.4055 | 0.004979 | 0.006987 | 0.006823 | 6.6e-05 | 0.001106 | 0.2959 | 1.3592 | 2.5481 | 100.58 | 20.72 | 19.164 | 148.904 | 142.1496 | 131.9755 | 0.004663 | 0.87145 | 0.0206 | 23.6393 | 0.1538 | 0.15557 | 0.069396 | 4.2718 | 4.2718 | 0.001497
2019-03-20 17:56 | 1553118943.084 | 403 | 0.19569 | 0.011456 | 0.000218 | 108.411 | 106.6597 | 101.7381 | 21.0668 | 0.005158 | 0.004586 | 141.1686 | 0.26154 | 8.4038 | 0.005009 | 0.005983 | 0.005881 | 6.4e-05 | 0.001101 | 0.28368 | 1.3001 | 2.5481 | 100.57 | 20.71 | 19.167 | 146.5731 | 142.8711 | 132.4385 | 0.004661 | 0.87045 | 0.020585 | 23.6481 | 0.14567 | 0.15559 | 0.069146 | 4.2532 | 4.2532 | 0.001504
2019-03-20 17:56 | 1553118944.0846 | 404 | 0.19571 | 0.011433 | 0.000221 | 108.4996 | 106.5397 | 101.9339 | 21.0663 | 0.005148 | 0.004554 | 140.9872 | 0.26152 | 8.4039 | 0.005089 | 0.00634 | 0.006165 | 6.5e-05 | 0.001098 | 0.29175 | 1.2824 | 2.548 | 100.57 | 20.71 | 19.169 | 146.3256 | 144.854 | 131.6488 | 0.004662 | 0.87215 | 0.020624 | 23.6478 | 0.16305 | 0.15558 | 0.069042 | 4.2549 | 4.2549 | 0.001503
2019-03-20 17:56 | 1553118945.0847 | 405 | 0.19572 | 0.01141 | 0.000223 | 108.5578 | 106.4844 | 102.1724 | 21.0668 | 0.005212 | 0.004539 | 140.8435 | 0.2607 | 8.4044 | 0.005035 | 0.006555 | 0.006401 | 6.8e-05 | 0.001097 | 0.28047 | 1.3068 | 2.5482 | 100.57 | 20.71 | 19.17 | 147.9117 | 143.5186 | 131.2853 | 0.004662 | 0.87293 | 0.020658 | 23.6648 | 0.17249 | 0.15509 | 0.068974 | 4.2575 | 4.2575 | 0.001497
2019-03-20 17:56 | 1553118946.0847 | 406 | 0.19579 | 0.011396 | 0.000225 | 108.5009 | 106.0804 | 102.2902 | 21.0662 | 0.005184 | 0.004616 | 140.6741 | 0.26168 | 8.4024 | 0.005054 | 0.005567 | 0.005428 | 6.9e-05 | 0.001092 | 0.29535 | 1.3361 | 2.5481 | 100.57 | 20.71 | 19.172 | 147.2211 | 144.009 | 133.1689 | 0.004662 | 0.86889 | 0.020622 | 23.7335 | 0.2018 | 0.15567 | 0.068585 | 4.2261 | 4.2261 | 0.001514
2019-03-20 17:56 | 1553118947.0847 | 407 | 0.19592 | 0.011374 | 0.000223 | 108.5905 | 106.305 | 102.3069 | 21.0669 | 0.005207 | 0.004579 | 140.5118 | 0.2616 | 8.4018 | 0.004989 | 0.006995 | 0.006788 | 6.6e-05 | 0.001093 | 0.31924 | 1.3989 | 2.5482 | 100.57 | 20.71 | 19.173 | 147.7813 | 142.3823 | 132.2757 | 0.004663 | 0.8708 | 0.020646 | 23.7093 | 0.19026 | 0.15563 | 0.067872 | 4.1874 | 4.1874 | 0.001528
2019-03-20 17:56 | 1553118948.0838 | 408 | 0.19602 | 0.011343 | 0.000224 | 108.5931 | 106.7498 | 102.3911 | 21.0669 | 0.005109 | 0.004604 | 140.3589 | 0.26103 | 8.4047 | 0.005051 | 0.006514 | 0.006359 | 6.7e-05 | 0.001091 | 0.27433 | 1.3665 | 2.5482 | 100.57 | 20.71 | 19.173 | 145.3666 | 143.9196 | 132.8908 | 0.004663 | 0.86948 | 0.020639 | 23.7371 | 0.15313 | 0.15529 | 0.067281 | 4.1496 | 4.1496 | 0.001538
2019-03-20 17:56 | 1553118949.0849 | 409 | 0.19617 | 0.011311 | 0.000223 | 108.6646 | 106.9176 | 102.3454 | 21.067 | 0.005184 | 0.004628 | 140.2047 | 0.26167 | 8.4023 | 0.005024 | 0.006641 | 0.006585 | 6.8e-05 | 0.001095 | 0.29087 | 1.3785 | 2.548 | 100.57 | 20.71 | 19.171 | 147.2251 | 143.257 | 133.4798 | 0.004662 | 0.86822 | 0.020619 | 23.749 | 0.14497 | 0.15567 | 0.066436 | 4.0938 | 4.0938 | 0.001563
2019-03-20 17:56 | 1553118950.0839 | 410 | 0.19614 | 0.011265 | 0.000223 | 108.7244 | 106.838 | 102.6015 | 21.0659 | 0.005258 | 0.004606 | 140.0591 | 0.26256 | 8.4029 | 0.004945 | 0.005846 | 0.005664 | 6.8e-05 | 0.001094 | 0.28871 | 1.3611 | 2.5481 | 100.57 | 20.71 | 19.171 | 149.051 | 141.303 | 132.9418 | 0.004662 | 0.86937 | 0.020659 | 23.7629 | 0.15655 | 0.1562 | 0.066693 | 4.1175 | 4.1175 | 0.001559
2019-03-20 17:56 | 1553118951.0849 | 411 | 0.1962 | 0.011223 | 0.000223 | 108.7224 | 107.1011 | 102.2689 | 21.0669 | 0.005202 | 0.004648 | 139.9133 | 0.261 | 8.4023 | 0.004921 | 0.00658 | 0.006475 | 6.6e-05 | 0.001098 | 0.26056 | 1.3117 | 2.548 | 100.57 | 20.71 | 19.17 | 147.6744 | 140.7069 | 133.9603 | 0.004662 | 0.8672 | 0.0206 | 23.7541 | 0.13439 | 0.15527 | 0.066388 | 4.087 | 4.087 | 0.001561
2019-03-20 17:56 | 1553118952.085 | 412 | 0.19619 | 0.011167 | 0.000224 | 108.7043 | 106.9781 | 102.2156 | 21.0674 | 0.005169 | 0.004583 | 139.764 | 0.2616 | 8.4049 | 0.004962 | 0.006282 | 0.006139 | 6.3e-05 | 0.001104 | 0.28567 | 1.3117 | 2.5481 | 100.58 | 20.71 | 19.166 | 146.8386 | 141.7236 | 132.3624 | 0.00466 | 0.87062 | 0.020635 | 23.7013 | 0.14318 | 0.15563 | 0.066485 | 4.0998 | 4.0998 | 0.00156
2019-03-20 17:56 | 1553118953.0842 | 413 | 0.19628 | 0.011118 | 0.000224 | 108.737 | 106.6253 | 102.4935 | 21.0669 | 0.005193 | 0.004587 | 139.5793 | 0.2616 | 8.4015 | 0.004948 | 0.007194 | 0.006995 | 6.8e-05 | 0.0011 | 0.2756 | 1.3168 | 2.548 | 100.58 | 20.71 | 19.163 | 147.4473 | 141.3676 | 132.4754 | 0.00466 | 0.87037 | 0.02066 | 23.7368 | 0.17541 | 0.15562 | 0.06604 | 4.0775 | 4.0775 | 0.001569
2019-03-20 17:56 | 1553118954.0832 | 414 | 0.19637 | 0.011075 | 0.000222 | 108.7959 | 106.8623 | 102.6644 | 21.067 | 0.005209 | 0.004585 | 139.4543 | 0.2613 | 8.4018 | 0.004912 | 0.007125 | 0.006985 | 6.5e-05 | 0.001108 | 0.28926 | 1.2902 | 2.548 | 100.58 | 20.71 | 19.161 | 147.8337 | 140.4953 | 132.4124 | 0.004659 | 0.87051 | 0.020679 | 23.7547 | 0.1604 | 0.15545 | 0.065524 | 4.0495 | 4.0495 | 0.001578
2019-03-20 17:56 | 1553118955.0842 | 415 | 0.19638 | 0.011025 | 0.00022 | 108.8741 | 106.8944 | 102.7017 | 21.0668 | 0.005235 | 0.004546 | 139.2983 | 0.26128 | 8.4027 | 0.005006 | 0.006478 | 0.006315 | 6.7e-05 | 0.001108 | 0.28744 | 1.28 | 2.5481 | 100.58 | 20.71 | 19.157 | 148.4929 | 142.8117 | 131.4504 | 0.004658 | 0.87258 | 0.020707 | 23.7308 | 0.16414 | 0.15544 | 0.065538 | 4.0561 | 4.0561 | 0.001575
2019-03-20 17:56 | 1553118956.0832 | 416 | 0.19641 | 0.010978 | 0.000218 | 108.7706 | 106.9535 | 102.5752 | 21.0674 | 0.005292 | 0.004556 | 139.148 | 0.26145 | 8.4038 | 0.004976 | 0.006438 | 0.006254 | 6.6e-05 | 0.001113 | 0.29551 | 1.3065 | 2.5481 | 100.58 | 20.71 | 19.155 | 149.8903 | 142.0777 | 131.6898 | 0.004658 | 0.87206 | 0.020688 | 23.7232 | 0.15069 | 0.15554 | 0.065418 | 4.0451 | 4.0451 | 0.00158
2019-03-20 17:56 | 1553118957.0843 | 417 | 0.19646 | 0.010937 | 0.000217 | 108.8415 | 107.0406 | 102.582 | 21.0682 | 0.005178 | 0.004531 | 138.9898 | 0.26165 | 8.4037 | 0.004953 | 0.006534 | 0.006449 | 6.4e-05 | 0.00112 | 0.29856 | 1.3195 | 2.5482 | 100.58 | 20.71 | 19.152 | 147.0639 | 141.4963 | 131.0933 | 0.004657 | 0.87335 | 0.020704 | 23.7065 | 0.14924 | 0.15566 | 0.065149 | 4.0317 | 4.0317 | 0.001587
2019-03-20 17:56 | 1553118958.0843 | 418 | 0.1966 | 0.010896 | 0.000218 | 108.9204 | 107.1007 | 102.5421 | 21.0671 | 0.005176 | 0.004602 | 138.8305 | 0.26111 | 8.4005 | 0.004808 | 0.006155 | 0.005949 | 6.7e-05 | 0.001122 | 0.30254 | 1.3239 | 2.548 | 100.58 | 20.71 | 19.152 | 147.029 | 137.9137 | 132.823 | 0.004657 | 0.86963 | 0.020656 | 23.7526 | 0.15069 | 0.15534 | 0.064372 | 3.9744 | 3.9744 | 0.001606
2019-03-20 17:56 | 1553118959.0848 | 419 | 0.19692 | 0.010853 | 0.00022 | 108.9003 | 107.1594 | 102.8895 | 21.0671 | 0.005219 | 0.004547 | 138.6687 | 0.26097 | 8.4003 | 0.0048 | 0.006901 | 0.006716 | 6.7e-05 | 0.001126 | 0.29878 | 1.3363 | 2.5481 | 100.58 | 20.71 | 19.154 | 148.0971 | 137.7086 | 131.4827 | 0.004658 | 0.87251 | 0.020725 | 23.7535 | 0.14415 | 0.15525 | 0.062512 | 3.8729 | 3.8729 | 0.001648
2019-03-20 17:56 | 1553118960.0844 | 420 | 0.19694 | 0.010786 | 0.000222 | 108.9744 | 107.3003 | 102.7806 | 21.0676 | 0.005165 | 0.00454 | 138.5168 | 0.26119 | 8.4025 | 0.004851 | 0.006635 | 0.006586 | 6.7e-05 | 0.001123 | 0.28727 | 1.3341 | 2.5483 | 100.58 | 20.71 | 19.155 | 146.7549 | 138.973 | 131.3021 | 0.004658 | 0.8729 | 0.020719 | 23.7356 | 0.13847 | 0.15538 | 0.062519 | 3.872 | 3.872 | 0.001649
2019-03-20 17:56 | 1553118961.0845 | 421 | 0.19704 | 0.010714 | 0.000224 | 108.8874 | 107.131 | 103.1695 | 21.0687 | 0.005141 | 0.004594 | 138.3696 | 0.26124 | 8.4024 | 0.004832 | 0.006283 | 0.006206 | 6.4e-05 | 0.001125 | 0.28119 | 1.3487 | 2.5482 | 100.58 | 20.71 | 19.157 | 146.1543 | 138.5074 | 132.6448 | 0.004658 | 0.87001 | 0.020724 | 23.8199 | 0.14546 | 0.15541 | 0.06195 | 3.8376 | 3.8376 | 0.001664
2019-03-20 17:56 | 1553118962.0845 | 422 | 0.19719 | 0.010634 | 0.000227 | 108.9689 | 107.3751 | 103.6518 | 21.0666 | 0.005176 | 0.004531 | 138.2268 | 0.26224 | 8.4034 | 0.00487 | 0.006438 | 0.00622 | 6.3e-05 | 0.001115 | 0.29087 | 1.3324 | 2.5482 | 100.58 | 20.71 | 19.163 | 147.0189 | 139.4433 | 131.0878 | 0.00466 | 0.87336 | 0.020812 | 23.8297 | 0.13179 | 0.15601 | 0.061173 | 3.8057 | 3.8057 | 0.001685
2019-03-20 17:56 | 1553118963.0836 | 423 | 0.19734 | 0.010556 | 0.000228 | 109.0533 | 107.6942 | 103.6159 | 21.0676 | 0.005121 | 0.004577 | 138.1136 | 0.26079 | 8.4031 | 0.004871 | 0.006297 | 0.006154 | 6.3e-05 | 0.001116 | 0.2724 | 1.323 | 2.5481 | 100.58 | 20.71 | 19.168 | 145.6474 | 139.4627 | 132.2069 | 0.004661 | 0.87095 | 0.02078 | 23.8585 | 0.11217 | 0.15515 | 0.060364 | 3.7496 | 3.7496 | 0.001701
2019-03-20 17:56 | 1553118964.0846 | 424 | 0.19757 | 0.01047 | 0.00023 | 108.9915 | 107.5477 | 103.638 | 21.0676 | 0.005154 | 0.004568 | 137.9746 | 0.2619 | 8.4029 | 0.004821 | 0.006315 | 0.006218 | 6.8e-05 | 0.001116 | 0.2792 | 1.3604 | 2.5482 | 100.58 | 20.7 | 19.174 | 146.4832 | 138.2331 | 132.0048 | 0.004659 | 0.87138 | 0.020787 | 23.8551 | 0.11928 | 0.1558 | 0.059089 | 3.6719 | 3.6719 | 0.001744
2019-03-20 17:56 | 1553118965.0851 | 425 | 0.19765 | 0.010379 | 0.000231 | 109.0892 | 107.5604 | 103.869 | 21.0674 | 0.005168 | 0.004547 | 137.8629 | 0.26203 | 8.4036 | 0.004814 | 0.006187 | 0.005991 | 6.8e-05 | 0.001113 | 0.29247 | 1.3202 | 2.5481 | 100.58 | 20.7 | 19.188 | 146.8257 | 138.0566 | 131.4702 | 0.004663 | 0.87253 | 0.020824 | 23.8659 | 0.12623 | 0.15588 | 0.058744 | 3.6569 | 3.6569 | 0.001752
2019-03-20 17:56 | 1553118966.084 | 426 | 0.19774 | 0.010291 | 0.000232 | 109.1257 | 107.6408 | 104.0584 | 21.0678 | 0.005152 | 0.004557 | 137.7351 | 0.2613 | 8.4057 | 0.004817 | 0.006618 | 0.00641 | 6.6e-05 | 0.001109 | 0.2912 | 1.3255 | 2.5481 | 100.58 | 20.7 | 19.195 | 146.4318 | 138.1419 | 131.7342 | 0.004665 | 0.87197 | 0.020836 | 23.8955 | 0.12255 | 0.15545 | 0.058314 | 3.6323 | 3.6323 | 0.001759
2019-03-20 17:56 | 1553118967.0838 | 427 | 0.19792 | 0.010193 | 0.000232 | 109.1028 | 107.4397 | 104.4149 | 21.0669 | 0.005062 | 0.004538 | 137.6111 | 0.26106 | 8.4017 | 0.004727 | 0.004991 | 0.004907 | 6.8e-05 | 0.001103 | 0.28888 | 1.3321 | 2.5481 | 100.58 | 20.7 | 19.214 | 144.1946 | 135.9227 | 131.2645 | 0.004669 | 0.87298 | 0.020884 | 23.9225 | 0.13739 | 0.15531 | 0.057342 | 3.5801 | 3.5801 | 0.001783
2019-03-20 17:56 | 1553118968.0839 | 428 | 0.19825 | 0.010093 | 0.000232 | 109.1655 | 107.6894 | 104.6788 | 21.0669 | 0.005107 | 0.004526 | 137.4965 | 0.26105 | 8.3997 | 0.0047 | 0.005567 | 0.005441 | 6.8e-05 | 0.001106 | 0.27959 | 1.2903 | 2.5481 | 100.58 | 20.7 | 19.224 | 145.3121 | 135.2553 | 130.9636 | 0.004672 | 0.87363 | 0.020918 | 23.9438 | 0.12177 | 0.1553 | 0.055533 | 3.4731 | 3.4731 | 0.001838
2019-03-20 17:56 | 1553118969.0853 | 429 | 0.1982 | 0.009983 | 0.000234 | 109.247 | 107.5877 | 105.1114 | 21.0671 | 0.005081 | 0.00449 | 137.3681 | 0.26205 | 8.399 | 0.004723 | 0.005974 | 0.005776 | 6.8e-05 | 0.001105 | 0.2505 | 1.3129 | 2.5481 | 100.58 | 20.7 | 19.234 | 144.6721 | 135.8212 | 130.0652 | 0.004674 | 0.87558 | 0.020984 | 23.9665 | 0.1369 | 0.15589 | 0.055958 | 3.5106 | 3.5106 | 0.001825
2019-03-20 17:56 | 1553118970.0844 | 430 | 0.19863 | 0.009872 | 0.000234 | 109.3156 | 107.7417 | 105.3322 | 21.067 | 0.005128 | 0.004494 | 137.2254 | 0.26094 | 8.4057 | 0.004772 | 0.006053 | 0.005917 | 6.6e-05 | 0.001103 | 0.28279 | 1.3405 | 2.5482 | 100.58 | 20.7 | 19.253 | 145.8285 | 137.0283 | 130.164 | 0.004679 | 0.87536 | 0.021004 | 23.9946 | 0.12972 | 0.15523 | 0.053487 | 3.359 | 3.359 | 0.001899
2019-03-20 17:56 | 1553118971.0831 | 431 | 0.19846 | 0.009748 | 0.000235 | 109.281 | 107.9618 | 105.5543 | 21.0667 | 0.005072 | 0.004508 | 137.0918 | 0.26123 | 8.4038 | 0.004676 | 0.006163 | 0.006032 | 6.7e-05 | 0.001104 | 0.29446 | 1.3313 | 2.5483 | 100.58 | 20.7 | 19.263 | 144.4436 | 134.6495 | 130.5093 | 0.004681 | 0.87461 | 0.021017 | 24.0302 | 0.10863 | 0.15541 | 0.054645 | 3.4337 | 3.4337 | 0.00186
2019-03-20 17:56 | 1553118972.084 | 432 | 0.19865 | 0.009622 | 0.000236 | 109.2716 | 107.9813 | 105.8261 | 21.0672 | 0.005055 | 0.00445 | 136.9662 | 0.26096 | 8.4048 | 0.004629 | 0.00608 | 0.005901 | 6.8e-05 | 0.001106 | 0.28982 | 1.344 | 2.5481 | 100.58 | 20.7 | 19.273 | 144.0148 | 133.5096 | 129.0849 | 0.004684 | 0.87771 | 0.021081 | 24.0186 | 0.10625 | 0.15524 | 0.053658 | 3.382 | 3.382 | 0.001887
2019-03-20 17:56 | 1553118973.0831 | 433 | 0.19869 | 0.009501 | 0.000236 | 109.3667 | 107.769 | 106.1547 | 21.0683 | 0.005118 | 0.004394 | 136.8237 | 0.26211 | 8.4045 | 0.004639 | 0.006406 | 0.006237 | 6.7e-05 | 0.001108 | 0.29623 | 1.3288 | 2.5482 | 100.58 | 20.71 | 19.292 | 145.5893 | 133.7457 | 127.719 | 0.004691 | 0.8807 | 0.02115 | 24.015 | 0.13163 | 0.15593 | 0.053567 | 3.3873 | 3.3873 | 0.001892
2019-03-20 17:56 | 1553118974.0842 | 434 | 0.19903 | 0.009391 | 0.000235 | 109.3605 | 107.8665 | 106.3087 | 21.0673 | 0.00507 | 0.004431 | 136.6847 | 0.26113 | 8.4038 | 0.00455 | 0.006291 | 0.006151 | 6.3e-05 | 0.001109 | 0.28031 | 1.3378 | 2.548 | 100.58 | 20.71 | 19.3 | 144.3838 | 131.542 | 128.6155 | 0.004693 | 0.87874 | 0.021142 | 24.0592 | 0.12304 | 0.15535 | 0.051684 | 3.2672 | 3.2672 | 0.001954
2019-03-20 17:56 | 1553118975.0846 | 435 | 0.19912 | 0.009281 | 0.000235 | 109.4188 | 107.9895 | 106.2221 | 21.0668 | 0.005077 | 0.004458 | 136.5483 | 0.2618 | 8.4063 | 0.004548 | 0.006738 | 0.00651 | 6.6e-05 | 0.001109 | 0.29015 | 1.3381 | 2.548 | 100.58 | 20.71 | 19.313 | 144.5562 | 131.4999 | 129.2849 | 0.004696 | 0.87727 | 0.021116 | 24.0695 | 0.11761 | 0.15575 | 0.051274 | 3.2373 | 3.2373 | 0.001977
2019-03-20 17:56 | 1553118976.0837 | 436 | 0.19926 | 0.009163 | 0.000235 | 109.4006 | 107.8568 | 106.4995 | 21.067 | 0.005075 | 0.004463 | 136.442 | 0.26136 | 8.4032 | 0.004602 | 0.006584 | 0.006435 | 6.8e-05 | 0.001107 | 0.28799 | 1.3496 | 2.5481 | 100.58 | 20.71 | 19.319 | 144.5155 | 132.8222 | 129.4095 | 0.004698 | 0.877 | 0.02114 | 24.1046 | 0.12712 | 0.15548 | 0.050626 | 3.2001 | 3.2001 | 0.001997
2019-03-20 17:56 | 1553118977.0833 | 437 | 0.19918 | 0.00906 | 0.000235 | 109.4469 | 108.0687 | 106.8708 | 21.0674 | 0.005064 | 0.004477 | 136.3265 | 0.2623 | 8.4006 | 0.004569 | 0.005991 | 0.005852 | 6.8e-05 | 0.001109 | 0.27505 | 1.3497 | 2.5482 | 100.58 | 20.71 | 19.323 | 144.2441 | 132.0181 | 129.7488 | 0.004699 | 0.87626 | 0.021168 | 24.1568 | 0.11335 | 0.15604 | 0.05121 | 3.2412 | 3.2412 | 0.001979
2019-03-20 17:56 | 1553118978.0834 | 438 | 0.1992 | 0.008965 | 0.000234 | 109.4592 | 108.0501 | 107.0265 | 21.0669 | 0.004979 | 0.004384 | 136.2114 | 0.26133 | 8.4003 | 0.004528 | 0.005867 | 0.005695 | 6.9e-05 | 0.001109 | 0.29407 | 1.2913 | 2.548 | 100.58 | 20.71 | 19.331 | 142.1302 | 131.0013 | 127.4781 | 0.004701 | 0.88123 | 0.021243 | 24.1061 | 0.1159 | 0.15547 | 0.051183 | 3.2511 | 3.2511 | 0.001965
2019-03-20 17:56 | 1553118979.0835 | 439 | 0.19922 | 0.008879 | 0.000232 | 109.4834 | 108.0205 | 106.653 | 21.0664 | 0.004968 | 0.004374 | 136.1058 | 0.26167 | 8.4026 | 0.004493 | 0.006174 | 0.00602 | 6.7e-05 | 0.001109 | 0.27616 | 1.2879 | 2.5481 | 100.58 | 20.71 | 19.332 | 141.88 | 130.1516 | 127.2217 | 0.004701 | 0.88179 | 0.021213 | 24.0563 | 0.12032 | 0.15567 | 0.051158 | 3.245 | 3.245 | 0.001972
2019-03-20 17:56 | 1553118980.0845 | 440 | 0.19928 | 0.0088 | 0.00023 | 109.4776 | 108.0108 | 107.0938 | 21.0669 | 0.005043 | 0.004398 | 136.0153 | 0.26069 | 8.4026 | 0.004485 | 0.006726 | 0.006614 | 6.9e-05 | 0.001114 | 0.30845 | 1.2655 | 2.548 | 100.58 | 20.71 | 19.334 | 143.7269 | 129.9485 | 127.8277 | 0.004701 | 0.88046 | 0.02124 | 24.1242 | 0.12065 | 0.15509 | 0.050922 | 3.2343 | 3.2343 | 0.001971
2019-03-20 17:56 | 1553118981.0836 | 441 | 0.19934 | 0.00873 | 0.00023 | 109.4566 | 107.9731 | 107.1241 | 21.0667 | 0.005029 | 0.004392 | 135.9179 | 0.2611 | 8.4048 | 0.004688 | 0.006367 | 0.006254 | 6.9e-05 | 0.001113 | 0.27743 | 1.3532 | 2.5481 | 100.58 | 20.71 | 19.335 | 143.3813 | 134.9567 | 127.6696 | 0.004702 | 0.88081 | 0.021248 | 24.1229 | 0.12206 | 0.15533 | 0.050619 | 3.2163 | 3.2163 | 0.001985
2019-03-20 17:56 | 1553118982.0836 | 442 | 0.19952 | 0.008661 | 0.000229 | 109.5083 | 108.0837 | 107.0014 | 21.0665 | 0.004932 | 0.004368 | 135.8027 | 0.26106 | 8.403 | 0.004627 | 0.005984 | 0.005834 | 6.9e-05 | 0.001114 | 0.2896 | 1.3328 | 2.5481 | 100.58 | 20.71 | 19.335 | 140.9849 | 133.453 | 127.0673 | 0.004702 | 0.88213 | 0.021251 | 24.0909 | 0.11713 | 0.1553 | 0.049647 | 3.1552 | 3.1552 | 0.002023
2019-03-20 17:56 | 1553118983.0837 | 443 | 0.19975 | 0.008598 | 0.000227 | 109.5834 | 108.253 | 106.8814 | 21.0676 | 0.004985 | 0.004348 | 135.7038 | 0.26151 | 8.4025 | 0.004636 | 0.006162 | 0.006054 | 6.6e-05 | 0.001116 | 0.28711 | 1.3146 | 2.5483 | 100.58 | 20.71 | 19.334 | 142.2904 | 133.6635 | 126.5778 | 0.004701 | 0.88321 | 0.021253 | 24.0627 | 0.10925 | 0.15557 | 0.0484 | 3.0763 | 3.0763 | 0.002078
2019-03-20 17:56 | 1553118984.0853 | 444 | 0.19986 | 0.008524 | 0.000225 | 109.6031 | 108.0246 | 106.9532 | 21.0674 | 0.004977 | 0.00434 | 135.608 | 0.26087 | 8.4022 | 0.004651 | 0.005776 | 0.005551 | 6.9e-05 | 0.001105 | 0.27129 | 1.2947 | 2.5481 | 100.58 | 20.71 | 19.334 | 142.0831 | 134.0472 | 126.3842 | 0.004701 | 0.88364 | 0.021265 | 24.065 | 0.12975 | 0.15519 | 0.047835 | 3.0423 | 3.0423 | 0.002097
2019-03-20 17:56 | 1553118985.0848 | 445 | 0.19996 | 0.008454 | 0.000223 | 109.5258 | 107.7062 | 107.1206 | 21.0673 | 0.004974 | 0.004307 | 135.4926 | 0.26203 | 8.4025 | 0.004675 | 0.005596 | 0.005474 | 6.6e-05 | 0.001111 | 0.28982 | 1.3138 | 2.5481 | 100.58 | 20.71 | 19.333 | 142.0196 | 134.6237 | 125.5738 | 0.004701 | 0.88544 | 0.021303 | 24.0593 | 0.14985 | 0.15588 | 0.047319 | 3.0151 | 3.0151 | 0.002125
2019-03-20 17:56 | 1553118986.0839 | 446 | 0.20012 | 0.008381 | 0.000221 | 109.684 | 107.9054 | 107.0546 | 21.0681 | 0.004953 | 0.004311 | 135.3742 | 0.26108 | 8.4003 | 0.004687 | 0.005934 | 0.00579 | 6.8e-05 | 0.001111 | 0.27151 | 1.2785 | 2.5481 | 100.58 | 20.72 | 19.333 | 141.4922 | 134.9221 | 125.6923 | 0.004704 | 0.88518 | 0.021293 | 24.0555 | 0.14623 | 0.15532 | 0.04641 | 2.9559 | 2.9559 | 0.00216
2019-03-20 17:56 | 1553118987.0839 | 447 | 0.20015 | 0.008303 | 0.000221 | 109.6319 | 107.803 | 107.0474 | 21.0677 | 0.004937 | 0.004341 | 135.2734 | 0.26195 | 8.4032 | 0.004699 | 0.006326 | 0.006168 | 6.6e-05 | 0.00111 | 0.29048 | 1.3492 | 2.5481 | 100.58 | 20.71 | 19.333 | 141.113 | 135.2208 | 126.413 | 0.004701 | 0.88358 | 0.021273 | 24.0764 | 0.15047 | 0.15583 | 0.046313 | 2.947 | 2.947 | 0.002173
2019-03-20 17:56 | 1553118988.083 | 448 | 0.20028 | 0.008217 | 0.00022 | 109.5345 | 107.6931 | 107.2938 | 21.0672 | 0.004937 | 0.004301 | 135.1715 | 0.26175 | 8.4037 | 0.004677 | 0.006424 | 0.006281 | 6.8e-05 | 0.001103 | 0.29911 | 1.3175 | 2.548 | 100.58 | 20.71 | 19.33 | 141.1036 | 134.6916 | 125.4329 | 0.0047 | 0.88575 | 0.021324 | 24.0745 | 0.15165 | 0.15571 | 0.045654 | 2.912 | 2.912 | 0.002198
2019-03-20 17:56 | 1553118989.084 | 449 | 0.20031 | 0.008128 | 0.00022 | 109.6953 | 107.9794 | 107.8764 | 21.0677 | 0.004891 | 0.004299 | 135.0604 | 0.26102 | 8.404 | 0.004599 | 0.006983 | 0.006789 | 6.4e-05 | 0.001099 | 0.28672 | 1.304 | 2.5482 | 100.58 | 20.72 | 19.327 | 139.9674 | 132.7534 | 125.3799 | 0.004702 | 0.88587 | 0.021383 | 24.1382 | 0.14102 | 0.15528 | 0.045592 | 2.9162 | 2.9162 | 0.002188
2019-03-20 17:56 | 1553118990.0831 | 450 | 0.20048 | 0.008044 | 0.000221 | 109.5743 | 107.7369 | 107.7501 | 21.0667 | 0.00484 | 0.004288 | 134.9463 | 0.26085 | 8.4008 | 0.004678 | 0.005731 | 0.005693 | 6.3e-05 | 0.001097 | 0.27649 | 1.3051 | 2.5481 | 100.58 | 20.71 | 19.32 | 138.7058 | 134.7015 | 125.1077 | 0.004698 | 0.88648 | 0.021378 | 24.1158 | 0.15126 | 0.15518 | 0.044695 | 2.8581 | 2.8581 | 0.002232
2019-03-20 17:56 | 1553118991.0841 | 451 | 0.20057 | 0.007961 | 0.000221 | 109.6041 | 107.6965 | 107.9912 | 21.0679 | 0.004851 | 0.004244 | 134.8324 | 0.2616 | 8.4007 | 0.004609 | 0.005765 | 0.005581 | 6.4e-05 | 0.001099 | 0.28495 | 1.2918 | 2.5481 | 100.58 | 20.71 | 19.316 | 138.9844 | 132.9957 | 124.0342 | 0.004697 | 0.88887 | 0.021431 | 24.1102 | 0.15704 | 0.15563 | 0.044241 | 2.8361 | 2.8361 | 0.002255
2019-03-20 17:57 | 1553118992.0844 | 452 | 0.20089 | 0.007865 | 0.000221 | 109.574 | 107.8 | 108.155 | 21.0676 | 0.004853 | 0.004256 | 134.7469 | 0.26234 | 8.3999 | 0.004638 | 0.006074 | 0.005856 | 6.3e-05 | 0.001096 | 0.28152 | 1.3102 | 2.5482 | 100.58 | 20.71 | 19.311 | 139.0146 | 133.7124 | 124.3418 | 0.004696 | 0.88818 | 0.021439 | 24.1379 | 0.146 | 0.15607 | 0.042439 | 2.7215 | 2.7215 | 0.002357
2019-03-20 17:57 | 1553118993.0846 | 453 | 0.20101 | 0.007769 | 0.00022 | 109.6041 | 107.7978 | 108.5641 | 21.0674 | 0.004865 | 0.004237 | 134.6685 | 0.26122 | 8.402 | 0.004644 | 0.005959 | 0.005814 | 6.5e-05 | 0.001095 | 0.27184 | 1.3119 | 2.5482 | 100.58 | 20.72 | 19.301 | 139.3189 | 133.8688 | 123.8692 | 0.004696 | 0.88924 | 0.021492 | 24.1691 | 0.14864 | 0.1554 | 0.041859 | 2.6911 | 2.6911 | 0.002373
2019-03-20 17:57 | 1553118994.0832 | 454 | 0.20114 | 0.00766 | 0.000219 | 109.5915 | 107.4842 | 108.8249 | 21.0662 | 0.004862 | 0.004199 | 134.5659 | 0.26139 | 8.403 | 0.004631 | 0.005627 | 0.005548 | 6.3e-05 | 0.001096 | 0.29839 | 1.3133 | 2.5482 | 100.58 | 20.72 | 19.296 | 139.2441 | 133.5359 | 122.9365 | 0.004695 | 0.89133 | 0.021543 | 24.1697 | 0.17366 | 0.1555 | 0.041215 | 2.656 | 2.656 | 0.002406
2019-03-20 17:57 | 1553118995.0847 | 455 | 0.2013 | 0.007559 | 0.000217 | 109.5268 | 107.4756 | 108.9266 | 21.0668 | 0.004792 | 0.004211 | 134.485 | 0.26052 | 8.4021 | 0.004605 | 0.006015 | 0.005882 | 6.6e-05 | 0.001102 | 0.28672 | 1.3168 | 2.5481 | 100.58 | 20.72 | 19.29 | 137.5169 | 132.9042 | 123.2383 | 0.004693 | 0.89065 | 0.021545 | 24.1902 | 0.1691 | 0.15499 | 0.04041 | 2.6045 | 2.6045 | 0.002446
2019-03-20 17:57 | 1553118996.0844 | 456 | 0.20149 | 0.007438 | 0.000214 | 109.5271 | 107.8636 | 109.4357 | 21.0666 | 0.004809 | 0.004179 | 134.3945 | 0.26105 | 8.4029 | 0.004633 | 0.006169 | 0.006024 | 6.3e-05 | 0.001103 | 0.26178 | 1.3114 | 2.548 | 100.58 | 20.72 | 19.278 | 137.9355 | 133.6002 | 122.4349 | 0.004691 | 0.89246 | 0.021617 | 24.222 | 0.1369 | 0.1553 | 0.039368 | 2.546 | 2.546 | 0.002507
2019-03-20 17:57 | 1553118997.0835 | 457 | 0.2017 | 0.007318 | 0.000212 | 109.5394 | 107.9245 | 109.6995 | 21.0682 | 0.004768 | 0.0042 | 134.3029 | 0.26105 | 8.4005 | 0.004644 | 0.006418 | 0.006338 | 6.2e-05 | 0.001103 | 0.26864 | 1.3312 | 2.548 | 100.58 | 20.72 | 19.271 | 136.9271 | 133.8721 | 122.9593 | 0.004689 | 0.89128 | 0.021629 | 24.2673 | 0.13285 | 0.1553 | 0.038285 | 2.4774 | 2.4774 | 0.002576
2019-03-20 17:57 | 1553118998.0844 | 458 | 0.20184 | 0.007193 | 0.000209 | 109.5045 | 107.9626 | 109.892 | 21.068 | 0.004702 | 0.004166 | 134.2178 | 0.2607 | 8.4029 | 0.00466 | 0.005874 | 0.005748 | 6.5e-05 | 0.001104 | 0.27201 | 1.3863 | 2.548 | 100.58 | 20.72 | 19.257 | 135.2923 | 134.2541 | 122.1312 | 0.004685 | 0.89315 | 0.021671 | 24.2632 | 0.12684 | 0.15509 | 0.037586 | 2.4369 | 2.4369 | 0.002616
2019-03-20 17:57 | 1553118999.0849 | 459 | 0.20201 | 0.007063 | 0.000206 | 109.5575 | 107.9282 | 109.9467 | 21.0675 | 0.004754 | 0.00417 | 134.1333 | 0.26095 | 8.4033 | 0.004577 | 0.005902 | 0.005781 | 6.1e-05 | 0.001109 | 0.30254 | 1.3455 | 2.5483 | 100.58 | 20.72 | 19.25 | 136.5717 | 132.2059 | 122.2143 | 0.004684 | 0.89296 | 0.021674 | 24.2717 | 0.13402 | 0.15524 | 0.036761 | 2.3839 | 2.3839 | 0.002676
2019-03-20 17:57 | 1553119000.085 | 460 | 0.20223 | 0.006933 | 0.000202 | 109.553 | 108.1771 | 109.9347 | 21.0682 | 0.004711 | 0.004104 | 134.0238 | 0.26136 | 8.4004 | 0.004697 | 0.005784 | 0.005684 | 6.2e-05 | 0.001108 | 0.26288 | 1.316 | 2.5482 | 100.58 | 20.72 | 19.243 | 135.5143 | 135.1724 | 120.6017 | 0.004682 | 0.89662 | 0.021717 | 24.2209 | 0.11304 | 0.15548 | 0.035598 | 2.3133 | 2.3133 | 0.002763
2019-03-20 17:57 | 1553119001.0829 | 461 | 0.20225 | 0.006814 | 0.000197 | 109.551 | 108.0942 | 110.1375 | 21.0666 | 0.00465 | 0.004132 | 133.919 | 0.26077 | 8.4017 | 0.004636 | 0.005766 | 0.005589 | 6.2e-05 | 0.001108 | 0.28191 | 1.2949 | 2.5481 | 100.58 | 20.72 | 19.232 | 134.0031 | 133.6746 | 121.3025 | 0.004679 | 0.89503 | 0.021718 | 24.2648 | 0.11975 | 0.15513 | 0.035629 | 2.3156 | 2.3156 | 0.002753
2019-03-20 17:57 | 1553119002.0837 | 462 | 0.20245 | 0.006678 | 0.000192 | 109.4349 | 107.8355 | 110.3879 | 21.0669 | 0.004607 | 0.00401 | 133.8417 | 0.25999 | 8.4032 | 0.004601 | 0.005794 | 0.005681 | 6.3e-05 | 0.001109 | 0.30414 | 1.3192 | 2.5481 | 100.58 | 20.72 | 19.227 | 132.9484 | 132.8061 | 118.3055 | 0.004678 | 0.90188 | 0.021825 | 24.1999 | 0.13169 | 0.15467 | 0.034619 | 2.2614 | 2.2614 | 0.002811
2019-03-20 17:57 | 1553119003.0827 | 463 | 0.20247 | 0.006545 | 0.000184 | 109.4278 | 107.9129 | 110.3497 | 21.0663 | 0.004579 | 0.00406 | 133.7436 | 0.26078 | 8.4034 | 0.004615 | 0.005955 | 0.005732 | 5.7e-05 | 0.001112 | 0.27184 | 1.3568 | 2.548 | 100.58 | 20.72 | 19.223 | 132.2669 | 133.1451 | 119.5377 | 0.004677 | 0.89905 | 0.021787 | 24.2337 | 0.12469 | 0.15514 | 0.034621 | 2.258 | 2.258 | 0.002824
2019-03-20 17:57 | 1553119004.0842 | 464 | 0.20269 | 0.006417 | 0.000176 | 109.4676 | 108.062 | 110.5602 | 21.0681 | 0.004648 | 0.004059 | 133.6739 | 0.26067 | 8.4023 | 0.004597 | 0.006425 | 0.006312 | 6.1e-05 | 0.00111 | 0.2568 | 1.3461 | 2.5482 | 100.58 | 20.73 | 19.217 | 133.9696 | 132.7001 | 119.5092 | 0.004679 | 0.89911 | 0.021809 | 24.256 | 0.11559 | 0.15507 | 0.033515 | 2.1883 | 2.1883 | 0.002912
2019-03-20 17:57 | 1553119005.0842 | 465 | 0.20279 | 0.006288 | 0.000169 | 109.4867 | 108.1865 | 110.8083 | 21.0669 | 0.004613 | 0.004043 | 133.584 | 0.26064 | 8.4023 | 0.004635 | 0.006029 | 0.005896 | 6.2e-05 | 0.001108 | 0.30237 | 1.2974 | 2.5481 | 100.58 | 20.73 | 19.213 | 133.1129 | 133.6571 | 119.1155 | 0.004678 | 0.90002 | 0.021844 | 24.271 | 0.10686 | 0.15505 | 0.033045 | 2.1616 | 2.1616 | 0.002948
2019-03-20 17:57 | 1553119006.0831 | 466 | 0.20301 | 0.006153 | 0.000162 | 109.3987 | 108.7598 | 110.8184 | 21.0672 | 0.004616 | 0.004049 | 133.4969 | 0.25996 | 8.4029 | 0.004605 | 0.005629 | 0.005527 | 6.1e-05 | 0.001105 | 0.27256 | 1.3014 | 2.5481 | 100.58 | 20.73 | 19.205 | 133.1691 | 132.9039 | 119.2698 | 0.004676 | 0.89966 | 0.021841 | 24.2769 | 0.05239 | 0.15465 | 0.031944 | 2.0895 | 2.0895 | 0.003042
2019-03-20 17:57 | 1553119007.0829 | 467 | 0.20321 | 0.006022 | 0.000155 | 109.367 | 108.8581 | 110.8586 | 21.0677 | 0.004532 | 0.004025 | 133.4321 | 0.26063 | 8.4044 | 0.004531 | 0.005872 | 0.005753 | 6.1e-05 | 0.001107 | 0.3126 | 1.2764 | 2.548 | 100.58 | 20.73 | 19.201 | 131.1216 | 131.0761 | 118.6734 | 0.004675 | 0.90103 | 0.021862 | 24.2628 | 0.041718 | 0.15505 | 0.030912 | 2.0242 | 2.0242 | 0.003148
2019-03-20 17:57 | 1553119008.084 | 468 | 0.20329 | 0.005903 | 0.00015 | 109.3356 | 108.9646 | 110.8834 | 21.0672 | 0.004526 | 0.004035 | 133.3522 | 0.2602 | 8.4017 | 0.004432 | 0.005734 | 0.005536 | 6.2e-05 | 0.001101 | 0.29374 | 1.2958 | 2.5482 | 100.58 | 20.73 | 19.193 | 130.9734 | 128.6568 | 118.912 | 0.004673 | 0.90048 | 0.021857 | 24.2729 | 0.030404 | 0.15479 | 0.030571 | 2.0018 | 2.0018 | 0.003178
2019-03-20 17:57 | 1553119009.0845 | 469 | 0.20337 | 0.005794 | 0.000145 | 109.3644 | 109.144 | 111.1172 | 21.0659 | 0.004493 | 0.003986 | 133.2759 | 0.26007 | 8.4025 | 0.004499 | 0.006059 | 0.005921 | 6.1e-05 | 0.001099 | 0.31725 | 1.3056 | 2.5481 | 100.58 | 20.73 | 19.188 | 130.1383 | 130.309 | 117.716 | 0.004672 | 0.90324 | 0.021914 | 24.2614 | 0.018044 | 0.15472 | 0.03024 | 1.9855 | 1.9855 | 0.003203
2019-03-20 17:57 | 1553119010.0841 | 470 | 0.20341 | 0.005693 | 0.00014 | 109.3179 | 109.4475 | 111.2792 | 21.0672 | 0.004488 | 0.003999 | 133.2541 | 0.2616 | 8.4002 | 0.00437 | 0.005779 | 0.005668 | 6.3e-05 | 0.001099 | 0.25946 | 1.3406 | 2.5481 | 100.58 | 20.73 | 19.184 | 130.0351 | 127.1292 | 118.054 | 0.004671 | 0.90246 | 0.02192 | 24.2896 | -0.010602 | 0.15562 | 0.030117 | 1.9783 | 1.9783 | 0.003233
2019-03-20 17:57 | 1553119011.0835 | 471 | 0.20348 | 0.005604 | 0.000135 | 109.3039 | 109.2726 | 111.3919 | 21.0669 | 0.004517 | 0.00393 | 133.2204 | 0.26044 | 8.4035 | 0.004484 | 0.005952 | 0.005774 | 6.5e-05 | 0.001099 | 0.30309 | 1.3202 | 2.5481 | 100.58 | 20.73 | 19.178 | 130.7417 | 129.9197 | 116.3558 | 0.004669 | 0.90639 | 0.021979 | 24.2491 | 0.002566 | 0.15494 | 0.029829 | 1.9649 | 1.9649 | 0.003241
2019-03-20 17:57 | 1553119012.0833 | 472 | 0.20355 | 0.005528 | 0.000131 | 109.2454 | 109.23 | 111.6944 | 21.0679 | 0.004449 | 0.003931 | 133.1546 | 0.26097 | 8.4039 | 0.004408 | 0.005816 | 0.005719 | 6.5e-05 | 0.001096 | 0.27649 | 1.3227 | 2.548 | 100.58 | 20.73 | 19.176 | 129.0625 | 128.0559 | 116.3819 | 0.004669 | 0.90633 | 0.022008 | 24.2828 | 0.001263 | 0.15525 | 0.029494 | 1.9456 | 1.9456 | 0.00328
2019-03-20 17:57 | 1553119013.0832 | 473 | 0.2036 | 0.005442 | 0.000129 | 109.2212 | 109.5122 | 112.0981 | 21.0684 | 0.004466 | 0.003932 | 133.1142 | 0.26054 | 8.4008 | 0.004385 | 0.005904 | 0.005719 | 6.3e-05 | 0.001098 | 0.29717 | 1.3092 | 2.5479 | 100.58 | 20.73 | 19.172 | 129.4921 | 127.5057 | 116.4053 | 0.004668 | 0.90628 | 0.022047 | 24.3274 | -0.023801 | 0.15499 | 0.029318 | 1.9375 | 1.9375 | 0.003288
2019-03-20 17:57 | 1553119014.0837 | 474 | 0.20369 | 0.005373 | 0.000127 | 109.2554 | 109.6292 | 112.0589 | 21.0678 | 0.004437 | 0.003891 | 133.0582 | 0.26056 | 8.4029 | 0.004353 | 0.006467 | 0.006365 | 6.1e-05 | 0.001091 | 0.29336 | 1.3148 | 2.5482 | 100.58 | 20.73 | 19.162 | 128.7682 | 126.7142 | 115.4013 | 0.004665 | 0.90862 | 0.022072 | 24.2917 | -0.030544 | 0.15501 | 0.02884 | 1.9082 | 1.9082 | 0.003339
2019-03-20 17:57 | 1553119015.0844 | 475 | 0.20382 | 0.005307 | 0.000125 | 109.1015 | 109.3006 | 111.7803 | 21.0677 | 0.004423 | 0.003895 | 133.0109 | 0.26035 | 8.4038 | 0.004375 | 0.005757 | 0.005601 | 6.3e-05 | 0.001095 | 0.29894 | 1.3013 | 2.548 | 100.58 | 20.73 | 19.155 | 128.4397 | 127.2532 | 115.5176 | 0.004664 | 0.90835 | 0.022041 | 24.2652 | -0.016311 | 0.15488 | 0.028146 | 1.8597 | 1.8597 | 0.003423
2019-03-20 17:57 | 1553119016.0838 | 476 | 0.20395 | 0.00523 | 0.000124 | 109.1616 | 108.9048 | 111.9728 | 21.0676 | 0.004358 | 0.003837 | 132.9873 | 0.26139 | 8.4039 | 0.004326 | 0.005981 | 0.005863 | 6.3e-05 | 0.001086 | 0.28512 | 1.3109 | 2.5481 | 100.58 | 20.73 | 19.143 | 126.8326 | 126.0483 | 114.0918 | 0.004661 | 0.91169 | 0.022101 | 24.2415 | 0.021068 | 0.1555 | 0.027447 | 1.8185 | 1.8185 | 0.003515
2019-03-20 17:57 | 1553119017.0826 | 477 | 0.20413 | 0.005166 | 0.000123 | 109.0433 | 108.2018 | 112.2967 | 21.0674 | 0.004427 | 0.003896 | 132.9235 | 0.26196 | 8.4047 | 0.00432 | 0.006476 | 0.006369 | 6.5e-05 | 0.001084 | 0.29175 | 1.3173 | 2.548 | 100.58 | 20.72 | 19.137 | 128.5281 | 125.909 | 115.5222 | 0.004656 | 0.90834 | 0.022092 | 24.3213 | 0.069292 | 0.15584 | 0.026482 | 1.7539 | 1.7539 | 0.003652
2019-03-20 17:57 | 1553119018.0835 | 478 | 0.20428 | 0.005099 | 0.000122 | 109.0439 | 108.3924 | 112.7631 | 21.0677 | 0.004362 | 0.003852 | 132.8833 | 0.26165 | 8.4006 | 0.00435 | 0.006398 | 0.00623 | 6.3e-05 | 0.001086 | 0.2855 | 1.3113 | 2.5482 | 100.58 | 20.72 | 19.131 | 126.937 | 126.6375 | 114.4649 | 0.004655 | 0.91082 | 0.022168 | 24.3386 | 0.053602 | 0.15566 | 0.025621 | 1.7027 | 1.7027 | 0.003757
2019-03-20 17:57 | 1553119019.083 | 479 | 0.20444 | 0.005015 | 0.000121 | 108.9876 | 108.7071 | 112.5664 | 21.0678 | 0.004312 | 0.003837 | 132.8627 | 0.26137 | 8.4034 | 0.004363 | 0.006045 | 0.005952 | 6.3e-05 | 0.00109 | 0.29806 | 1.3118 | 2.548 | 100.58 | 20.73 | 19.119 | 125.715 | 126.953 | 114.0815 | 0.004655 | 0.91172 | 0.02216 | 24.3053 | 0.023053 | 0.15549 | 0.024809 | 1.648 | 1.648 | 0.003878
2019-03-20 17:57 | 1553119020.0826 | 480 | 0.20453 | 0.00492 | 0.000121 | 108.9931 | 108.6436 | 113.014 | 21.0671 | 0.004319 | 0.003818 | 132.8026 | 0.26037 | 8.4026 | 0.004314 | 0.005805 | 0.00563 | 6.6e-05 | 0.001089 | 0.27688 | 1.3279 | 2.5481 | 100.58 | 20.72 | 19.114 | 125.8831 | 125.7566 | 113.6343 | 0.004651 | 0.91277 | 0.022216 | 24.3395 | 0.028734 | 0.1549 | 0.024369 | 1.623 | 1.623 | 0.003923
2019-03-20 17:57 | 1553119021.0831 | 481 | 0.2047 | 0.004831 | 0.000121 | 108.9456 | 109.1074 | 113.421 | 21.0674 | 0.004283 | 0.003755 | 132.7532 | 0.26163 | 8.403 | 0.004334 | 0.006155 | 0.005964 | 6.4e-05 | 0.001088 | 0.2808 | 1.2918 | 2.548 | 100.58 | 20.73 | 19.106 | 124.9851 | 126.2559 | 112.0919 | 0.004652 | 0.91643 | 0.022301 | 24.3346 | -0.013274 | 0.15564 | 0.023489 | 1.5702 | 1.5702 | 0.004074
2019-03-20 17:57 | 1553119022.0827 | 482 | 0.20492 | 0.004724 | 0.00012 | 108.9197 | 109.1874 | 113.5865 | 21.0672 | 0.004287 | 0.003773 | 132.709 | 0.26054 | 8.407 | 0.004355 | 0.006311 | 0.006189 | 6.2e-05 | 0.00109 | 0.28385 | 1.3113 | 2.5482 | 100.58 | 20.72 | 19.102 | 125.0938 | 126.7669 | 112.5372 | 0.004648 | 0.91537 | 0.022304 | 24.3665 | -0.021953 | 0.155 | 0.02229 | 1.4901 | 1.4901 | 0.004275
2019-03-20 17:57 | 1553119023.0828 | 483 | 0.20513 | 0.004613 | 0.00012 | 108.8386 | 109.363 | 113.5424 | 21.0677 | 0.00425 | 0.003795 | 132.6523 | 0.26055 | 8.4049 | 0.00434 | 0.006671 | 0.006492 | 6.2e-05 | 0.001091 | 0.26272 | 1.3289 | 2.548 | 100.58 | 20.73 | 19.099 | 124.1764 | 126.3892 | 113.0653 | 0.00465 | 0.91412 | 0.022285 | 24.3784 | -0.043 | 0.155 | 0.021134 | 1.4114 | 1.4114 | 0.004514
2019-03-20 17:57 | 1553119024.0829 | 484 | 0.20522 | 0.004501 | 0.00012 | 108.8027 | 109.3746 | 113.8818 | 21.0672 | 0.004196 | 0.003724 | 132.5971 | 0.26156 | 8.4009 | 0.004373 | 0.004973 | 0.004853 | 6.1e-05 | 0.001088 | 0.27936 | 1.3206 | 2.5481 | 100.58 | 20.73 | 19.094 | 122.8692 | 127.1923 | 111.3401 | 0.004649 | 0.91822 | 0.022368 | 24.3602 | -0.046898 | 0.1556 | 0.02075 | 1.3909 | 1.3909 | 0.004598
2019-03-20 17:57 | 1553119025.0838 | 485 | 0.20533 | 0.004377 | 0.000121 | 108.7587 | 109.3952 | 114.4224 | 21.0666 | 0.004227 | 0.003706 | 132.5656 | 0.26116 | 8.4033 | 0.004415 | 0.005517 | 0.005452 | 6e-05 | 0.001096 | 0.27649 | 1.3312 | 2.5481 | 100.58 | 20.73 | 19.093 | 123.6163 | 128.2282 | 110.9087 | 0.004648 | 0.91925 | 0.022434 | 24.4043 | -0.052196 | 0.15536 | 0.020236 | 1.3602 | 1.3602 | 0.004694
2019-03-20 17:57 | 1553119026.0841 | 486 | 0.20543 | 0.004259 | 0.000123 | 108.7069 | 109.5144 | 114.4224 | 21.066 | 0.004207 | 0.003734 | 132.512 | 0.26171 | 8.4018 | 0.00431 | 0.006222 | 0.00607 | 6.5e-05 | 0.001103 | 0.30055 | 1.3218 | 2.5482 | 100.58 | 20.73 | 19.092 | 123.1271 | 125.6498 | 111.5757 | 0.004648 | 0.91765 | 0.022414 | 24.4254 | -0.066196 | 0.15569 | 0.019764 | 1.3272 | 1.3272 | 0.004821
2019-03-20 17:57 | 1553119027.083 | 487 | 0.20559 | 0.004135 | 0.000124 | 108.7202 | 109.4378 | 114.3958 | 21.0667 | 0.004185 | 0.003683 | 132.4729 | 0.26061 | 8.4037 | 0.004319 | 0.005869 | 0.005732 | 6.4e-05 | 0.001103 | 0.28351 | 1.3086 | 2.5482 | 100.58 | 20.73 | 19.091 | 122.5922 | 125.8693 | 110.3416 | 0.004648 | 0.92061 | 0.022448 | 24.3834 | -0.058845 | 0.15504 | 0.018961 | 1.2749 | 1.2749 | 0.004998
2019-03-20 17:57 | 1553119028.0841 | 488 | 0.20573 | 0.004019 | 0.000126 | 108.5788 | 109.4038 | 114.463 | 21.0663 | 0.004128 | 0.003699 | 132.4623 | 0.26111 | 8.4034 | 0.004345 | 0.005582 | 0.005483 | 6.3e-05 | 0.001108 | 0.28744 | 1.2983 | 2.5481 | 100.58 | 20.73 | 19.091 | 121.1855 | 126.5088 | 110.7436 | 0.004648 | 0.91964 | 0.022442 | 24.4033 | -0.0677 | 0.15533 | 0.01825 | 1.2265 | 1.2265 | 0.005205
2019-03-20 17:57 | 1553119029.0831 | 489 | 0.20584 | 0.003903 | 0.000128 | 108.597 | 109.5824 | 114.7364 | 21.0666 | 0.004172 | 0.003671 | 132.4165 | 0.26068 | 8.4013 | 0.004307 | 0.005231 | 0.005067 | 6e-05 | 0.001104 | 0.2896 | 1.3311 | 2.5481 | 100.58 | 20.73 | 19.095 | 122.2665 | 125.5763 | 110.0473 | 0.004649 | 0.92131 | 0.02249 | 24.4103 | -0.080802 | 0.15508 | 0.017745 | 1.1949 | 1.1949 | 0.005334
2019-03-20 17:57 | 1553119030.0832 | 490 | 0.20597 | 0.003795 | 0.00013 | 108.475 | 109.4774 | 114.8195 | 21.0674 | 0.004136 | 0.003671 | 132.3801 | 0.26077 | 8.4036 | 0.004311 | 0.005582 | 0.005467 | 6.4e-05 | 0.001107 | 0.30165 | 1.3136 | 2.5481 | 100.58 | 20.72 | 19.1 | 121.3874 | 125.6788 | 110.0639 | 0.004647 | 0.92127 | 0.022497 | 24.4197 | -0.082274 | 0.15513 | 0.017086 | 1.1506 | 1.1506 | 0.005541
2019-03-20 17:57 | 1553119031.0833 | 491 | 0.20606 | 0.003694 | 0.000131 | 108.4446 | 109.1455 | 114.6062 | 21.0661 | 0.004112 | 0.003653 | 132.3608 | 0.26122 | 8.403 | 0.004248 | 0.005773 | 0.005619 | 6.1e-05 | 0.00111 | 0.28208 | 1.3034 | 2.5481 | 100.58 | 20.72 | 19.106 | 120.8163 | 124.1248 | 109.6105 | 0.004649 | 0.92237 | 0.02249 | 24.3825 | -0.057623 | 0.1554 | 0.016673 | 1.1222 | 1.1222 | 0.005692
2019-03-20 17:57 | 1553119032.0824 | 492 | 0.20616 | 0.003597 | 0.000135 | 108.366 | 109.0663 | 114.5207 | 21.0671 | 0.004048 | 0.003639 | 132.3212 | 0.26092 | 8.4012 | 0.004291 | 0.004918 | 0.004806 | 6.5e-05 | 0.001111 | 0.27831 | 1.3093 | 2.5481 | 100.58 | 20.73 | 19.118 | 119.2418 | 125.1892 | 109.2766 | 0.004655 | 0.92317 | 0.022491 | 24.3628 | -0.057612 | 0.15522 | 0.016204 | 1.0903 | 1.0903 | 0.005851
2019-03-20 17:57 | 1553119033.0824 | 493 | 0.20622 | 0.003496 | 0.000136 | 108.367 | 109.0012 | 114.7782 | 21.0678 | 0.004038 | 0.003656 | 132.3176 | 0.26056 | 8.4022 | 0.004198 | 0.005086 | 0.004942 | 6.5e-05 | 0.001111 | 0.28528 | 1.3158 | 2.5481 | 100.58 | 20.73 | 19.123 | 118.9984 | 122.9096 | 109.6982 | 0.004656 | 0.92215 | 0.022504 | 24.4036 | -0.052197 | 0.15501 | 0.015967 | 1.0749 | 1.0749 | 0.005927
2019-03-20 17:57 | 1553119034.0824 | 494 | 0.20633 | 0.003398 | 0.00014 | 108.3699 | 109.0573 | 114.836 | 21.0673 | 0.004008 | 0.003635 | 132.3005 | 0.26106 | 8.4003 | 0.004252 | 0.005819 | 0.005723 | 6.3e-05 | 0.001109 | 0.26106 | 1.3287 | 2.548 | 100.58 | 20.73 | 19.129 | 118.2624 | 124.239 | 109.1782 | 0.004657 | 0.92341 | 0.022525 | 24.3932 | -0.056556 | 0.15531 | 0.015418 | 1.0385 | 1.0385 | 0.006147
2019-03-20 17:57 | 1553119035.0834 | 495 | 0.20648 | 0.003314 | 0.000143 | 108.2952 | 108.8888 | 115.0026 | 21.0666 | 0.003983 | 0.003524 | 132.2725 | 0.26003 | 8.4052 | 0.00432 | 0.006173 | 0.006127 | 6.8e-05 | 0.001103 | 0.27776 | 1.3032 | 2.548 | 100.58 | 20.73 | 19.136 | 117.6652 | 125.9104 | 106.4835 | 0.004659 | 0.92996 | 0.022621 | 24.3247 | -0.048889 | 0.15469 | 0.014631 | 0.98909 | 0.98909 | 0.006428
2019-03-20 17:57 | 1553119036.0826 | 496 | 0.20665 | 0.003224 | 0.000145 | 108.2325 | 108.731 | 115.4228 | 21.0667 | 0.004001 | 0.003577 | 132.2521 | 0.26011 | 8.4043 | 0.004288 | 0.006181 | 0.005945 | 6.6e-05 | 0.001095 | 0.29264 | 1.318 | 2.5482 | 100.58 | 20.73 | 19.144 | 118.0962 | 125.1131 | 107.7623 | 0.004661 | 0.92684 | 0.022624 | 24.4101 | -0.041107 | 0.15474 | 0.013707 | 0.92625 | 0.92625 | 0.006866
2019-03-20 17:57 | 1553119037.0836 | 497 | 0.20673 | 0.003132 | 0.000148 | 108.2651 | 108.8798 | 115.5986 | 21.0663 | 0.004019 | 0.003505 | 132.2248 | 0.26046 | 8.402 | 0.004189 | 0.005935 | 0.005779 | 6.6e-05 | 0.001097 | 0.27887 | 1.3477 | 2.548 | 100.58 | 20.73 | 19.149 | 118.5325 | 122.6976 | 106.0148 | 0.004662 | 0.93111 | 0.022694 | 24.3726 | -0.050637 | 0.15495 | 0.013335 | 0.90356 | 0.90356 | 0.007048
2019-03-20 17:57 | 1553119038.083 | 498 | 0.20689 | 0.003046 | 0.00015 | 108.249 | 108.8304 | 115.3622 | 21.0674 | 0.004012 | 0.003562 | 132.1639 | 0.25949 | 8.4024 | 0.004202 | 0.005567 | 0.005493 | 6.3e-05 | 0.001099 | 0.29479 | 1.3568 | 2.5481 | 100.58 | 20.73 | 19.159 | 118.3727 | 123.0197 | 107.4062 | 0.004664 | 0.92771 | 0.022629 | 24.3923 | -0.047919 | 0.15437 | 0.012486 | 0.843 | 0.843 | 0.007526
2019-03-20 17:57 | 1553119039.0837 | 499 | 0.20695 | 0.002955 | 0.000153 | 108.1798 | 108.8951 | 115.5791 | 21.0673 | 0.003986 | 0.003531 | 132.1527 | 0.26043 | 8.4035 | 0.004228 | 0.006281 | 0.006144 | 6.6e-05 | 0.001095 | 0.2656 | 1.365 | 2.5481 | 100.58 | 20.73 | 19.163 | 117.7347 | 123.634 | 106.6565 | 0.004665 | 0.92954 | 0.022673 | 24.3911 | -0.058952 | 0.15493 | 0.01223 | 0.82698 | 0.82698 | 0.0077
2019-03-20 17:57 | 1553119040.0827 | 500 | 0.20708 | 0.00286 | 0.000156 | 108.2321 | 108.9407 | 115.711 | 21.0669 | 0.003957 | 0.003485 | 132.149 | 0.26145 | 8.4014 | 0.004152 | 0.005341 | 0.005265 | 6.8e-05 | 0.001088 | 0.26432 | 1.3312 | 2.5481 | 100.58 | 20.73 | 19.168 | 117.0232 | 121.7848 | 105.5433 | 0.004667 | 0.93227 | 0.022719 | 24.3693 | -0.058366 | 0.15554 | 0.011581 | 0.78411 | 0.78411 | 0.008153
2019-03-20 17:57 | 1553119041.0828 | 501 | 0.20723 | 0.002761 | 0.000161 | 108.1038 | 108.7699 | 115.795 | 21.0669 | 0.003921 | 0.003474 | 132.1297 | 0.26019 | 8.4039 | 0.004091 | 0.005683 | 0.005522 | 6.5e-05 | 0.001093 | 0.27793 | 1.3226 | 2.5482 | 100.58 | 20.73 | 19.171 | 116.1348 | 120.2901 | 105.2773 | 0.004667 | 0.93293 | 0.022735 | 24.3696 | -0.054948 | 0.15479 | 0.010774 | 0.72915 | 0.72915 | 0.008725
2019-03-20 17:57 | 1553119042.0838 | 502 | 0.20742 | 0.002657 | 0.000164 | 108.1186 | 108.9583 | 116.0546 | 21.0665 | 0.003879 | 0.003452 | 132.1348 | 0.26068 | 8.4008 | 0.004122 | 0.005905 | 0.00574 | 6.5e-05 | 0.001095 | 0.26913 | 1.2995 | 2.5481 | 100.58 | 20.73 | 19.17 | 115.1258 | 121.0555 | 104.7468 | 0.004667 | 0.93424 | 0.022776 | 24.3797 | -0.069194 | 0.15508 | 0.009744 | 0.65962 | 0.65962 | 0.009663
2019-03-20 17:57 | 1553119043.0843 | 503 | 0.20743 | 0.002561 | 0.000169 | 108.0814 | 109.0248 | 116.2586 | 21.0666 | 0.003882 | 0.003465 | 132.1131 | 0.26066 | 8.4008 | 0.004101 | 0.005908 | 0.005837 | 6.5e-05 | 0.001094 | 0.27721 | 1.3191 | 2.5481 | 100.58 | 20.73 | 19.168 | 115.1978 | 120.5334 | 105.0584 | 0.004667 | 0.93347 | 0.022787 | 24.4112 | -0.077731 | 0.15507 | 0.00982 | 0.66489 | 0.66489 | 0.009585
2019-03-20 17:57 | 1553119044.0829 | 504 | 0.20748 | 0.002473 | 0.000175 | 108.0963 | 109.1459 | 116.6257 | 21.0663 | 0.003882 | 0.003437 | 132.0749 | 0.2607 | 8.4036 | 0.004137 | 0.005484 | 0.005332 | 6.1e-05 | 0.001101 | 0.28385 | 1.3168 | 2.5482 | 100.58 | 20.73 | 19.162 | 115.1999 | 121.4269 | 104.3646 | 0.004665 | 0.93518 | 0.022844 | 24.4273 | -0.08643 | 0.15509 | 0.009623 | 0.65267 | 0.65267 | 0.009767
2019-03-20 17:57 | 1553119045.084 | 505 | 0.20755 | 0.002383 | 0.00018 | 108.0947 | 108.8641 | 116.7224 | 21.0668 | 0.003877 | 0.003412 | 132.0668 | 0.26057 | 8.4035 | 0.004138 | 0.005483 | 0.005358 | 6.4e-05 | 0.001101 | 0.24674 | 1.2952 | 2.548 | 100.58 | 20.73 | 19.158 | 115.0793 | 121.4392 | 103.7669 | 0.004664 | 0.93667 | 0.022872 | 24.4181 | -0.063439 | 0.15502 | 0.009267 | 0.62859 | 0.62859 | 0.010136
2019-03-20 17:57 | 1553119046.0841 | 506 | 0.20758 | 0.002297 | 0.000186 | 108.0174 | 108.8428 | 116.6738 | 21.0657 | 0.00376 | 0.003433 | 132.0443 | 0.26007 | 8.4029 | 0.004125 | 0.006041 | 0.00584 | 6.7e-05 | 0.001104 | 0.27616 | 1.3478 | 2.5482 | 100.58 | 20.73 | 19.149 | 112.2141 | 121.1279 | 104.286 | 0.004662 | 0.93538 | 0.022851 | 24.4298 | -0.068087 | 0.15472 | 0.009186 | 0.62219 | 0.62219 | 0.01022
2019-03-20 17:57 | 1553119047.0825 | 507 | 0.20766 | 0.002219 | 0.000189 | 108.009 | 108.7004 | 116.6499 | 21.0669 | 0.003829 | 0.003362 | 132.04 | 0.26093 | 8.4028 | 0.004064 | 0.005721 | 0.005573 | 6.6e-05 | 0.001102 | 0.28241 | 1.3515 | 2.5482 | 100.58 | 20.73 | 19.144 | 113.9006 | 119.6327 | 102.5574 | 0.004661 | 0.93968 | 0.022901 | 24.3713 | -0.057074 | 0.15523 | 0.00881 | 0.59742 | 0.59742 | 0.010679
2019-03-20 17:57 | 1553119048.083 | 508 | 0.20778 | 0.002134 | 0.000192 | 107.9993 | 108.6615 | 116.8349 | 21.0658 | 0.003741 | 0.003341 | 132.0462 | 0.25949 | 8.403 | 0.004095 | 0.005665 | 0.005502 | 6.5e-05 | 0.001106 | 0.28622 | 1.3533 | 2.5482 | 100.58 | 20.73 | 19.139 | 111.7635 | 120.389 | 102.0552 | 0.00466 | 0.94094 | 0.022935 | 24.3743 | -0.054678 | 0.15437 | 0.008175 | 0.55423 | 0.55423 | 0.011448
2019-03-20 17:57 | 1553119049.0823 | 509 | 0.20788 | 0.002061 | 0.000194 | 108.0287 | 108.814 | 116.9287 | 21.0672 | 0.003785 | 0.003373 | 132.0201 | 0.26168 | 8.4024 | 0.004024 | 0.005698 | 0.005583 | 6.6e-05 | 0.0011 | 0.25282 | 1.3225 | 2.5481 | 100.58 | 20.73 | 19.129 | 112.8215 | 118.6633 | 102.8328 | 0.004657 | 0.93899 | 0.02292 | 24.4093 | -0.064784 | 0.15567 | 0.007674 | 0.51921 | 0.51921 | 0.012323
2019-03-20 17:57 | 1553119050.0833 | 510 | 0.20797 | 0.001992 | 0.000194 | 107.9673 | 108.8005 | 117.0116 | 21.067 | 0.003775 | 0.003319 | 132.0054 | 0.26007 | 8.4032 | 0.00405 | 0.004822 | 0.004703 | 6.8e-05 | 0.001101 | 0.28191 | 1.3086 | 2.548 | 100.58 | 20.73 | 19.123 | 112.5732 | 119.294 | 101.5137 | 0.004656 | 0.9423 | 0.022969 | 24.3751 | -0.068768 | 0.15471 | 0.007255 | 0.49124 | 0.49124 | 0.012944
2019-03-20 17:58 | 1553119051.0823 | 511 | 0.20801 | 0.001918 | 0.000194 | 107.9815 | 108.6578 | 117.2839 | 21.0675 | 0.003681 | 0.003291 | 132.0144 | 0.26071 | 8.4028 | 0.004003 | 0.004991 | 0.004901 | 6.6e-05 | 0.001101 | 0.27345 | 1.3077 | 2.5481 | 100.58 | 20.73 | 19.114 | 110.2903 | 118.1531 | 100.8447 | 0.004654 | 0.94398 | 0.023016 | 24.3817 | -0.055844 | 0.1551 | 0.0071 | 0.48155 | 0.48155 | 0.013237
2019-03-20 17:58 | 1553119052.0827 | 512 | 0.20806 | 0.001844 | 0.000195 | 108.0349 | 108.8562 | 117.5735 | 21.068 | 0.003675 | 0.003385 | 131.9856 | 0.26154 | 8.4026 | 0.003979 | 0.005606 | 0.005414 | 6.9e-05 | 0.001097 | 0.29502 | 1.3072 | 2.5482 | 100.58 | 20.72 | 19.111 | 110.159 | 117.5539 | 103.1038 | 0.00465 | 0.93832 | 0.022975 | 24.4854 | -0.067746 | 0.15559 | 0.006885 | 0.46573 | 0.46573 | 0.013731
2019-03-20 17:58 | 1553119053.0824 | 513 | 0.20816 | 0.001776 | 0.000194 | 107.9201 | 108.777 | 117.4212 | 21.0657 | 0.003717 | 0.003317 | 131.9513 | 0.26102 | 8.4047 | 0.003933 | 0.00571 | 0.005599 | 6.9e-05 | 0.001101 | 0.26537 | 1.3159 | 2.5481 | 100.58 | 20.72 | 19.109 | 111.1628 | 116.4235 | 101.4617 | 0.004649 | 0.94243 | 0.02301 | 24.416 | -0.070744 | 0.15528 | 0.006393 | 0.43243 | 0.43243 | 0.014759
2019-03-20 17:58 | 1553119054.0835 | 514 | 0.20812 | 0.001712 | 0.000193 | 107.8583 | 108.6679 | 117.2632 | 21.0668 | 0.003673 | 0.003328 | 131.9626 | 0.26118 | 8.4028 | 0.003933 | 0.006035 | 0.005856 | 6.8e-05 | 0.001102 | 0.27831 | 1.3408 | 2.5483 | 100.58 | 20.73 | 19.108 | 110.1112 | 116.445 | 101.7304 | 0.004652 | 0.94176 | 0.022987 | 24.4084 | -0.066887 | 0.15538 | 0.006673 | 0.45143 | 0.45143 | 0.014146
2019-03-20 17:58 | 1553119055.0835 | 515 | 0.20822 | 0.001648 | 0.00019 | 107.9802 | 108.6854 | 117.2498 | 21.067 | 0.00371 | 0.00324 | 131.9547 | 0.25977 | 8.4015 | 0.003976 | 0.005488 | 0.005399 | 6.7e-05 | 0.001105 | 0.28368 | 1.323 | 2.5482 | 100.58 | 20.73 | 19.108 | 110.9975 | 117.4786 | 99.5942 | 0.004652 | 0.94715 | 0.023051 | 24.3373 | -0.058227 | 0.15454 | 0.006191 | 0.41933 | 0.41933 | 0.015147
2019-03-20 17:58 | 1553119056.0829 | 516 | 0.20827 | 0.001598 | 0.000188 | 107.9275 | 108.7407 | 117.2142 | 21.067 | 0.003652 | 0.003275 | 131.9583 | 0.26079 | 8.4038 | 0.003951 | 0.006071 | 0.005881 | 6.9e-05 | 0.001107 | 0.27743 | 1.3222 | 2.5482 | 100.58 | 20.72 | 19.11 | 109.6029 | 116.8822 | 100.4494 | 0.00465 | 0.94498 | 0.023021 | 24.3615 | -0.067146 | 0.15514 | 0.005958 | 0.40276 | 0.40276 | 0.015832
2019-03-20 17:58 | 1553119057.0839 | 517 | 0.20835 | 0.001534 | 0.000187 | 107.8861 | 108.5176 | 117.1792 | 21.0661 | 0.003583 | 0.003321 | 131.9478 | 0.261 | 8.3991 | 0.003905 | 0.004757 | 0.0047 | 6.8e-05 | 0.001104 | 0.27992 | 1.3422 | 2.5482 | 100.58 | 20.72 | 19.113 | 107.9155 | 115.7504 | 101.5741 | 0.00465 | 0.94215 | 0.022983 | 24.3945 | -0.052205 | 0.15527 | 0.005542 | 0.37337 | 0.37337 | 0.017092
2019-03-20 17:58 | 1553119058.0828 | 518 | 0.2084 | 0.001486 | 0.000183 | 107.8971 | 108.6963 | 117.3714 | 21.066 | 0.003611 | 0.003214 | 131.9407 | 0.26077 | 8.4016 | 0.003961 | 0.005252 | 0.005161 | 6.9e-05 | 0.001111 | 0.29502 | 1.3213 | 2.5481 | 100.58 | 20.72 | 19.116 | 108.5912 | 117.1039 | 98.9722 | 0.004651 | 0.94873 | 0.023082 | 24.3296 | -0.066007 | 0.15513 | 0.005281 | 0.35701 | 0.35701 | 0.017859
2019-03-20 17:58 | 1553119059.0827 | 519 | 0.20845 | 0.001435 | 0.00018 | 107.9741 | 108.6443 | 117.2663 | 21.0674 | 0.003587 | 0.003243 | 131.9259 | 0.26089 | 8.4041 | 0.00399 | 0.005164 | 0.005024 | 7e-05 | 0.001109 | 0.2625 | 1.3403 | 2.5481 | 100.58 | 20.72 | 19.123 | 108.0221 | 117.8212 | 99.6777 | 0.004653 | 0.94694 | 0.02305 | 24.3418 | -0.055352 | 0.15521 | 0.005074 | 0.34236 | 0.34236 | 0.018632
2019-03-20 17:58 | 1553119060.0827 | 520 | 0.20849 | 0.00139 | 0.000176 | 107.9307 | 108.5303 | 117.3414 | 21.0664 | 0.003577 | 0.0032 | 131.9111 | 0.26097 | 8.4027 | 0.003928 | 0.005591 | 0.005427 | 6.8e-05 | 0.001114 | 0.28943 | 1.3546 | 2.5481 | 100.58 | 20.72 | 19.128 | 107.7761 | 116.3187 | 98.6438 | 0.004654 | 0.94957 | 0.02309 | 24.3158 | -0.049553 | 0.15525 | 0.004887 | 0.33008 | 0.33008 | 0.019331
2019-03-20 17:58 | 1553119061.0838 | 521 | 0.20848 | 0.001347 | 0.000173 | 107.9113 | 108.697 | 117.4174 | 21.0665 | 0.003539 | 0.003217 | 131.8849 | 0.25944 | 8.4039 | 0.003911 | 0.005958 | 0.005804 | 6.8e-05 | 0.001115 | 0.28031 | 1.3453 | 2.548 | 100.58 | 20.72 | 19.132 | 106.851 | 115.8996 | 99.0565 | 0.004655 | 0.94852 | 0.023084 | 24.3371 | -0.064889 | 0.15434 | 0.005012 | 0.33889 | 0.33889 | 0.018718
2019-03-20 17:58 | 1553119062.0839 | 522 | 0.20852 | 0.001308 | 0.00017 | 107.931 | 108.8925 | 117.0736 | 21.0657 | 0.00352 | 0.003204 | 131.8865 | 0.2607 | 8.4035 | 0.003981 | 0.005477 | 0.005323 | 6.9e-05 | 0.00111 | 0.28031 | 1.3458 | 2.5482 | 100.58 | 20.72 | 19.144 | 106.3873 | 117.6135 | 98.7249 | 0.004658 | 0.94937 | 0.023061 | 24.2906 | -0.079324 | 0.15509 | 0.004839 | 0.32671 | 0.32671 | 0.01951
2019-03-20 17:58 | 1553119063.0829 | 523 | 0.20855 | 0.00127 | 0.000166 | 107.9055 | 108.876 | 117.2118 | 21.0661 | 0.003553 | 0.003183 | 131.8947 | 0.26097 | 8.4025 | 0.003949 | 0.004976 | 0.00489 | 6.7e-05 | 0.001108 | 0.26792 | 1.3737 | 2.548 | 100.58 | 20.72 | 19.148 | 107.1961 | 116.814 | 98.2312 | 0.004659 | 0.95063 | 0.02309 | 24.2889 | -0.080091 | 0.15525 | 0.004686 | 0.31668 | 0.31668 | 0.020149
2019-03-20 17:58 | 1553119064.082 | 524 | 0.20855 | 0.001236 | 0.000162 | 107.8699 | 108.7385 | 117.2125 | 21.0664 | 0.003538 | 0.003234 | 131.8847 | 0.26113 | 8.4 | 0.004002 | 0.005457 | 0.005312 | 7.2e-05 | 0.001104 | 0.27383 | 1.3281 | 2.548 | 100.58 | 20.72 | 19.157 | 106.8241 | 118.1086 | 99.454 | 0.004661 | 0.94751 | 0.023052 | 24.3289 | -0.071734 | 0.15535 | 0.004748 | 0.32069 | 0.32069 | 0.019909
2019-03-20 17:58 | 1553119065.083 | 525 | 0.20859 | 0.001206 | 0.000159 | 107.9213 | 108.8551 | 117.2632 | 21.0664 | 0.00349 | 0.003221 | 131.8891 | 0.26121 | 8.3985 | 0.003981 | 0.005828 | 0.005657 | 7.1e-05 | 0.001106 | 0.23236 | 1.3462 | 2.5481 | 100.58 | 20.72 | 19.162 | 105.6467 | 117.5937 | 99.1473 | 0.004662 | 0.94829 | 0.023066 | 24.3241 | -0.077058 | 0.15539 | 0.004509 | 0.30449 | 0.30449 | 0.020975
2019-03-20 17:58 | 1553119066.0831 | 526 | 0.2086 | 0.001176 | 0.000157 | 107.9456 | 108.8028 | 117.2236 | 21.0652 | 0.003507 | 0.003146 | 131.8694 | 0.26065 | 8.4021 | 0.004002 | 0.005572 | 0.005528 | 6.8e-05 | 0.001108 | 0.25498 | 1.363 | 2.5481 | 100.58 | 20.72 | 19.167 | 106.0659 | 118.1067 | 97.3322 | 0.004664 | 0.95293 | 0.023119 | 24.2606 | -0.070747 | 0.15506 | 0.004532 | 0.30685 | 0.30685 | 0.020769
2019-03-20 17:58 | 1553119067.0821 | 527 | 0.20858 | 0.001156 | 0.000154 | 107.9453 | 108.7535 | 117.4926 | 21.0667 | 0.003423 | 0.003157 | 131.8524 | 0.26022 | 8.4022 | 0.003997 | 0.005909 | 0.005821 | 7e-05 | 0.001112 | 0.26504 | 1.3803 | 2.5481 | 100.58 | 20.72 | 19.176 | 104.0316 | 117.9963 | 97.5921 | 0.004666 | 0.95227 | 0.023137 | 24.297 | -0.066717 | 0.15481 | 0.004638 | 0.31472 | 0.31472 | 0.020216
2019-03-20 17:58 | 1553119068.0823 | 528 | 0.20866 | 0.001125 | 0.000152 | 107.901 | 108.7508 | 117.2989 | 21.0666 | 0.00348 | 0.003135 | 131.841 | 0.26046 | 8.4031 | 0.004054 | 0.005601 | 0.005481 | 6.9e-05 | 0.001112 | 0.27383 | 1.3562 | 2.5481 | 100.58 | 20.72 | 19.181 | 105.4267 | 119.389 | 97.0631 | 0.004667 | 0.95363 | 0.023135 | 24.2596 | -0.070174 | 0.15495 | 0.004229 | 0.28628 | 0.28628 | 0.022245
2019-03-20 17:58 | 1553119069.0833 | 529 | 0.20869 | 0.001109 | 0.000149 | 107.9058 | 108.5019 | 117.2714 | 21.0664 | 0.003503 | 0.003122 | 131.8534 | 0.2598 | 8.4017 | 0.00404 | 0.005832 | 0.005719 | 6.8e-05 | 0.001115 | 0.262 | 1.3168 | 2.548 | 100.58 | 20.72 | 19.187 | 105.9826 | 119.0342 | 96.7629 | 0.004668 | 0.9544 | 0.023141 | 24.247 | -0.049276 | 0.15455 | 0.004049 | 0.27396 | 0.27396 | 0.023187
2019-03-20 17:58 | 1553119070.0823 | 530 | 0.20873 | 0.001081 | 0.000146 | 107.8819 | 108.5931 | 117.1628 | 21.0662 | 0.003454 | 0.003116 | 131.852 | 0.2595 | 8.4026 | 0.004065 | 0.005333 | 0.005159 | 7.2e-05 | 0.001112 | 0.2589 | 1.343 | 2.5482 | 100.58 | 20.72 | 19.19 | 104.7836 | 119.654 | 96.613 | 0.004669 | 0.95479 | 0.023135 | 24.2308 | -0.058775 | 0.15438 | 0.003869 | 0.2615 | 0.2615 | 0.024263
2019-03-20 17:58 | 1553119071.0828 | 531 | 0.20875 | 0.001059 | 0.000145 | 107.8628 | 108.6253 | 117.0142 | 21.0663 | 0.003459 | 0.003166 | 131.8438 | 0.26007 | 8.3994 | 0.004041 | 0.005332 | 0.00524 | 6.9e-05 | 0.001113 | 0.27759 | 1.3338 | 2.5481 | 100.58 | 20.72 | 19.192 | 104.9064 | 119.0588 | 97.8239 | 0.00467 | 0.95167 | 0.023083 | 24.2551 | -0.063004 | 0.15472 | 0.003772 | 0.25423 | 0.25423 | 0.025012
2019-03-20 17:58 | 1553119072.0824 | 532 | 0.20875 | 0.001045 | 0.000142 | 107.9699 | 108.7217 | 116.9928 | 21.0654 | 0.003429 | 0.003164 | 131.8357 | 0.25977 | 8.4018 | 0.004055 | 0.005526 | 0.005453 | 6.9e-05 | 0.001111 | 0.28296 | 1.2993 | 2.5481 | 100.58 | 20.72 | 19.195 | 104.1827 | 119.4162 | 97.7682 | 0.00467 | 0.95181 | 0.023082 | 24.251 | -0.062067 | 0.15454 | 0.00381 | 0.25702 | 0.25702 | 0.024712
2019-03-20 17:58 | 1553119073.0835 | 533 | 0.20876 | 0.001027 | 0.000142 | 107.8887 | 108.6398 | 116.9161 | 21.066 | 0.003444 | 0.003119 | 131.8212 | 0.25889 | 8.4033 | 0.004021 | 0.005288 | 0.005141 | 7.2e-05 | 0.001115 | 0.27273 | 1.3146 | 2.5482 | 100.58 | 20.72 | 19.195 | 104.5301 | 118.5845 | 96.6865 | 0.00467 | 0.9546 | 0.023109 | 24.2077 | -0.062059 | 0.15401 | 0.003779 | 0.2552 | 0.2552 | 0.024804
2019-03-20 17:58 | 1553119074.0826 | 534 | 0.20879 | 0.001002 | 0.00014 | 107.9084 | 108.7194 | 116.9792 | 21.0664 | 0.003417 | 0.003085 | 131.8296 | 0.26141 | 8.4003 | 0.003991 | 0.005626 | 0.00544 | 7.1e-05 | 0.001114 | 0.28174 | 1.3499 | 2.5482 | 100.58 | 20.72 | 19.196 | 103.8779 | 117.8527 | 95.8675 | 0.004671 | 0.95672 | 0.02314 | 24.1874 | -0.066975 | 0.15551 | 0.003607 | 0.2436 | 0.2436 | 0.026238
2019-03-20 17:58 | 1553119075.0836 | 535 | 0.20881 | 0.00099 | 0.000138 | 107.9375 | 108.7232 | 116.9053 | 21.0658 | 0.003377 | 0.003093 | 131.8096 | 0.26038 | 8.4013 | 0.004064 | 0.005567 | 0.005443 | 7.3e-05 | 0.001113 | 0.27345 | 1.3776 | 2.548 | 100.58 | 20.72 | 19.195 | 102.9242 | 119.6419 | 96.0501 | 0.00467 | 0.95624 | 0.023127 | 24.1857 | -0.06487 | 0.1549 | 0.003504 | 0.23643 | 0.23643 | 0.026927
2019-03-20 17:58 | 1553119076.084 | 536 | 0.20882 | 0.000976 | 0.000136 | 107.9301 | 108.8005 | 117.0379 | 21.0662 | 0.003396 | 0.003114 | 131.7951 | 0.26163 | 8.3994 | 0.004033 | 0.005513 | 0.00541 | 7.2e-05 | 0.001114 | 0.29159 | 1.3311 | 2.5481 | 100.58 | 20.72 | 19.193 | 103.3692 | 118.8668 | 96.5565 | 0.00467 | 0.95493 | 0.023125 | 24.216 | -0.071849 | 0.15564 | 0.003464 | 0.23373 | 0.23373 | 0.027369
2019-03-20 17:58 | 1553119077.0827 | 537 | 0.20888 | 0.000957 | 0.000134 | 107.9576 | 108.9519 | 117.0635 | 21.0662 | 0.003421 | 0.003083 | 131.8133 | 0.26087 | 8.4019 | 0.003995 | 0.005776 | 0.005646 | 7.1e-05 | 0.001117 | 0.27649 | 1.346 | 2.5482 | 100.58 | 20.72 | 19.188 | 103.9839 | 117.9444 | 95.8101 | 0.004669 | 0.95687 | 0.023151 | 24.1942 | -0.082007 | 0.15519 | 0.003104 | 0.20897 | 0.20897 | 0.030523
2019-03-20 17:58 | 1553119078.0828 | 538 | 0.20888 | 0.000944 | 0.000133 | 108.0135 | 108.6652 | 117.3203 | 21.0669 | 0.003329 | 0.003048 | 131.8003 | 0.2603 | 8.4025 | 0.004032 | 0.005652 | 0.005514 | 6.8e-05 | 0.001118 | 0.27096 | 1.3314 | 2.5482 | 100.58 | 20.72 | 19.186 | 101.7501 | 118.8612 | 94.9613 | 0.004668 | 0.95907 | 0.023203 | 24.1929 | -0.053807 | 0.15485 | 0.003163 | 0.21364 | 0.21364 | 0.02979
2019-03-20 17:58 | 1553119079.0828 | 539 | 0.20892 | 0.000924 | 0.000131 | 107.9226 | 108.7897 | 117.3806 | 21.066 | 0.003312 | 0.003015 | 131.7616 | 0.26007 | 8.4028 | 0.003988 | 0.005054 | 0.004882 | 7e-05 | 0.001114 | 0.2656 | 1.3369 | 2.548 | 100.58 | 20.72 | 19.183 | 101.3347 | 117.7771 | 94.1723 | 0.004667 | 0.96113 | 0.023234 | 24.1731 | -0.071575 | 0.15472 | 0.002936 | 0.19814 | 0.19814 | 0.032094
2019-03-20 17:58 | 1553119080.0828 | 540 | 0.20895 | 0.000907 | 0.000129 | 107.9873 | 108.7837 | 117.3743 | 21.0655 | 0.003358 | 0.003075 | 131.7428 | 0.2612 | 8.4027 | 0.004021 | 0.005605 | 0.005471 | 6.8e-05 | 0.001122 | 0.2677 | 1.3554 | 2.5481 | 100.58 | 20.72 | 19.178 | 102.4682 | 118.5682 | 95.606 | 0.004666 | 0.9574 | 0.023188 | 24.2196 | -0.065724 | 0.15539 | 0.002756 | 0.18521 | 0.18521 | 0.034481
2019-03-20 17:58 | 1553119081.0834 | 541 | 0.20898 | 0.00089 | 0.000128 | 107.9482 | 108.7733 | 117.3428 | 21.0659 | 0.003349 | 0.003097 | 131.7197 | 0.26046 | 8.4033 | 0.003991 | 0.005541 | 0.005401 | 6.7e-05 | 0.001125 | 0.29463 | 1.3665 | 2.5481 | 100.58 | 20.72 | 19.176 | 102.2329 | 117.8514 | 96.1419 | 0.004666 | 0.95601 | 0.023168 | 24.234 | -0.068105 | 0.15495 | 0.002638 | 0.17688 | 0.17688 | 0.036002
2019-03-20 17:58 | 1553119082.083 | 542 | 0.20898 | 0.00088 | 0.000126 | 107.9029 | 108.5998 | 117.416 | 21.0659 | 0.003284 | 0.003062 | 131.6994 | 0.26046 | 8.4031 | 0.003993 | 0.005728 | 0.005632 | 6.8e-05 | 0.001124 | 0.27168 | 1.3958 | 2.5482 | 100.58 | 20.72 | 19.173 | 100.6659 | 117.8854 | 95.3143 | 0.004665 | 0.95815 | 0.023201 | 24.2143 | -0.057586 | 0.15495 | 0.002662 | 0.17894 | 0.17894 | 0.035589
2019-03-20 17:58 | 1553119083.082 | 543 | 0.20898 | 0.000854 | 0.000124 | 107.9754 | 108.6234 | 117.4477 | 21.0655 | 0.003322 | 0.002986 | 131.6939 | 0.26014 | 8.4027 | 0.003973 | 0.004906 | 0.004796 | 7.1e-05 | 0.001123 | 0.28263 | 1.3725 | 2.548 | 100.58 | 20.72 | 19.173 | 101.598 | 117.4082 | 93.4753 | 0.004665 | 0.96296 | 0.023262 | 24.1571 | -0.053522 | 0.15476 | 0.002697 | 0.18195 | 0.18195 | 0.034957
2019-03-20 17:58 | 1553119084.0831 | 544 | 0.20896 | 0.000847 | 0.000124 | 108.0177 | 108.7602 | 117.3534 | 21.0659 | 0.003347 | 0.003026 | 131.6971 | 0.25963 | 8.4026 | 0.003962 | 0.005111 | 0.005004 | 6.7e-05 | 0.001126 | 0.29319 | 1.3132 | 2.5482 | 100.58 | 20.72 | 19.172 | 102.2015 | 117.1281 | 94.4358 | 0.004665 | 0.96044 | 0.023223 | 24.179 | -0.06127 | 0.15446 | 0.002782 | 0.18761 | 0.18761 | 0.033836
2019-03-20 17:58 | 1553119085.0823 | 545 | 0.20896 | 0.00083 | 0.000122 | 108.0436 | 108.8162 | 117.5144 | 21.0656 | 0.003279 | 0.003036 | 131.6907 | 0.25979 | 8.4017 | 0.003992 | 0.005509 | 0.005348 | 6.9e-05 | 0.001122 | 0.26432 | 1.3322 | 2.548 | 100.58 | 20.72 | 19.174 | 100.5422 | 117.8598 | 94.6845 | 0.004665 | 0.95979 | 0.023231 | 24.2038 | -0.063737 | 0.15455 | 0.002828 | 0.19105 | 0.19105 | 0.033248
2019-03-20 17:58 | 1553119086.0831 | 546 | 0.20901 | 0.000823 | 0.000122 | 107.9825 | 108.7426 | 117.578 | 21.0665 | 0.003276 | 0.002972 | 131.7108 | 0.26007 | 8.4015 | 0.003967 | 0.005314 | 0.005233 | 6.8e-05 | 0.00112 | 0.26272 | 1.329 | 2.5481 | 100.58 | 20.72 | 19.173 | 100.4811 | 117.2634 | 93.1283 | 0.004665 | 0.96387 | 0.023286 | 24.1591 | -0.062744 | 0.15471 | 0.002551 | 0.17211 | 0.17211 | 0.036946
2019-03-20 17:58 | 1553119087.0821 | 547 | 0.209 | 0.000811 | 0.000121 | 108.0847 | 108.7019 | 117.6273 | 21.0656 | 0.003192 | 0.003026 | 131.6965 | 0.26031 | 8.4007 | 0.003869 | 0.005432 | 0.005271 | 7.3e-05 | 0.001117 | 0.26825 | 1.3358 | 2.5481 | 100.58 | 20.72 | 19.173 | 98.4439 | 114.8772 | 94.4428 | 0.004665 | 0.96042 | 0.023249 | 24.2074 | -0.050932 | 0.15486 | 0.002626 | 0.17716 | 0.17716 | 0.035927
2019-03-20 17:58 | 1553119088.0823 | 548 | 0.209 | 0.000801 | 0.00012 | 108.0365 | 108.8312 | 117.6876 | 21.0662 | 0.003278 | 0.002995 | 131.6818 | 0.26058 | 8.4001 | 0.003925 | 0.005344 | 0.005136 | 6.7e-05 | 0.001123 | 0.26626 | 1.3153 | 2.548 | 100.58 | 20.72 | 19.172 | 100.5334 | 116.2395 | 93.6857 | 0.004665 | 0.96241 | 0.023279 | 24.1887 | -0.065555 | 0.15502 | 0.002643 | 0.17857 | 0.17857 | 0.035679
2019-03-20 17:58 | 1553119089.0836 | 549 | 0.20901 | 0.000797 | 0.000118 | 108.0682 | 108.9269 | 118.0045 | 21.0645 | 0.003252 | 0.002993 | 131.7088 | 0.25909 | 8.4049 | 0.003967 | 0.005585 | 0.005429 | 6.9e-05 | 0.001123 | 0.27328 | 1.3101 | 2.548 | 100.58 | 20.72 | 19.171 | 99.8992 | 117.2488 | 93.6472 | 0.004665 | 0.96251 | 0.023312 | 24.22 | -0.070791 | 0.15413 | 0.002598 | 0.17577 | 0.17577 | 0.036041
2019-03-20 17:58 | 1553119090.0834 | 550 | 0.20902 | 0.000784 | 0.000116 | 108.0856 | 108.9269 | 118.2798 | 21.0663 | 0.003233 | 0.002965 | 131.7105 | 0.26031 | 8.4028 | 0.003904 | 0.005177 | 0.005044 | 6.9e-05 | 0.001123 | 0.29192 | 1.3005 | 2.548 | 100.58 | 20.72 | 19.168 | 99.4414 | 115.722 | 92.9684 | 0.004664 | 0.96429 | 0.023361 | 24.2258 | -0.069346 | 0.15486 | 0.002548 | 0.17275 | 0.17275 | 0.036843
2019-03-20 17:58 | 1553119091.0838 | 551 | 0.209 | 0.000781 | 0.000116 | 108.1516 | 108.9856 | 118.6946 | 21.067 | 0.003238 | 0.003007 | 131.7112 | 0.25993 | 8.4031 | 0.003865 | 0.005335 | 0.005188 | 6.9e-05 | 0.001124 | 0.29695 | 1.3614 | 2.5481 | 100.58 | 20.72 | 19.167 | 99.5588 | 114.7662 | 93.9714 | 0.004664 | 0.96166 | 0.02337 | 24.3014 | -0.068705 | 0.15463 | 0.002671 | 0.18151 | 0.18151 | 0.035014
2019-03-20 17:58 | 1553119092.0825 | 552 | 0.20903 | 0.000766 | 0.000114 | 108.0992 | 108.8428 | 118.8826 | 21.0657 | 0.003256 | 0.002977 | 131.7091 | 0.25922 | 8.4015 | 0.003898 | 0.005013 | 0.004873 | 7e-05 | 0.001125 | 0.29286 | 1.3554 | 2.5481 | 100.58 | 20.72 | 19.166 | 99.9979 | 115.5719 | 93.2572 | 0.004663 | 0.96353 | 0.023411 | 24.297 | -0.061314 | 0.15421 | 0.002514 | 0.17089 | 0.17089 | 0.037088
2019-03-20 17:58 | 1553119093.0826 | 553 | 0.20908 | 0.000764 | 0.000112 | 108.0743 | 109.1317 | 118.8258 | 21.0668 | 0.003191 | 0.002956 | 131.6873 | 0.25956 | 8.4025 | 0.003852 | 0.005056 | 0.004959 | 7e-05 | 0.001127 | 0.28351 | 1.333 | 2.5481 | 100.58 | 20.72 | 19.167 | 98.4272 | 114.4673 | 92.7529 | 0.004664 | 0.96486 | 0.023421 | 24.2745 | -0.087085 | 0.15441 | 0.002214 | 0.14981 | 0.14981 | 0.042362
2019-03-20 17:58 | 1553119094.0826 | 554 | 0.20903 | 0.000759 | 0.00011 | 108.1186 | 109.0098 | 118.8711 | 21.0661 | 0.00319 | 0.002984 | 131.6512 | 0.26003 | 8.4046 | 0.003883 | 0.005294 | 0.005159 | 6.7e-05 | 0.001124 | 0.27096 | 1.3324 | 2.5481 | 100.58 | 20.72 | 19.167 | 98.3958 | 115.2178 | 93.4158 | 0.004664 | 0.96312 | 0.023405 | 24.3011 | -0.073427 | 0.15469 | 0.0025 | 0.17001 | 0.17001 | 0.037397
2019-03-20 17:58 | 1553119095.0829 | 555 | 0.20902 | 0.000754 | 0.000109 | 108.1151 | 109.0495 | 118.7835 | 21.0663 | 0.003201 | 0.002933 | 131.665 | 0.26029 | 8.4021 | 0.003837 | 0.004765 | 0.004601 | 6.4e-05 | 0.001125 | 0.28495 | 1.2821 | 2.5479 | 100.58 | 20.72 | 19.173 | 98.6541 | 114.0927 | 92.1823 | 0.004665 | 0.96637 | 0.023436 | 24.2512 | -0.076972 | 0.15485 | 0.002613 | 0.17833 | 0.17833 | 0.035689
2019-03-20 17:58 | 1553119096.0821 | 556 | 0.20904 | 0.000751 | 0.000107 | 108.1012 | 108.9343 | 118.749 | 21.0661 | 0.003241 | 0.002922 | 131.6446 | 0.26118 | 8.4045 | 0.003798 | 0.005746 | 0.005611 | 6.8e-05 | 0.001125 | 0.28351 | 1.2967 | 2.5481 | 100.58 | 20.72 | 19.177 | 99.6271 | 113.1404 | 91.9312 | 0.004666 | 0.96703 | 0.02344 | 24.2394 | -0.068675 | 0.15538 | 0.002448 | 0.16675 | 0.16675 | 0.038297
2019-03-20 17:58 | 1553119097.0818 | 557 | 0.20904 | 0.000743 | 0.000107 | 108.1652 | 109.0566 | 118.5857 | 21.0658 | 0.003139 | 0.002968 | 131.6343 | 0.26012 | 8.4039 | 0.00382 | 0.005329 | 0.005259 | 6.8e-05 | 0.001131 | 0.28351 | 1.3365 | 2.548 | 100.58 | 20.72 | 19.182 | 97.1604 | 113.6805 | 93.0297 | 0.004667 | 0.96413 | 0.023389 | 24.2591 | -0.073408 | 0.15474 | 0.002477 | 0.16847 | 0.16847 | 0.037751
2019-03-20 17:58 | 1553119098.0828 | 558 | 0.20901 | 0.000741 | 0.000106 | 108.1283 | 109.2326 | 118.6892 | 21.067 | 0.003163 | 0.002916 | 131.6353 | 0.26095 | 8.4002 | 0.003908 | 0.005858 | 0.005768 | 6.7e-05 | 0.001126 | 0.28672 | 1.3129 | 2.548 | 100.58 | 20.72 | 19.196 | 97.753 | 115.8278 | 91.7945 | 0.004671 | 0.9674 | 0.023439 | 24.2287 | -0.090884 | 0.15524 | 0.002635 | 0.18012 | 0.18012 | 0.035424
2019-03-20 17:58 | 1553119099.0828 | 559 | 0.20901 | 0.000737 | 0.000104 | 108.1419 | 109.0274 | 118.2455 | 21.0657 | 0.003178 | 0.002927 | 131.6001 | 0.26099 | 8.4014 | 0.003843 | 0.005899 | 0.005716 | 6.5e-05 | 0.001126 | 0.28351 | 1.3173 | 2.5481 | 100.58 | 20.72 | 19.204 | 98.1086 | 114.2291 | 92.0559 | 0.004673 | 0.9667 | 0.023387 | 24.192 | -0.072942 | 0.15526 | 0.002682 | 0.18307 | 0.18307 | 0.034858
2019-03-20 17:58 | 1553119100.0833 | 560 | 0.20903 | 0.000732 | 0.000102 | 108.2723 | 109.1003 | 117.9897 | 21.0663 | 0.003126 | 0.002884 | 131.575 | 0.26121 | 8.4035 | 0.003751 | 0.005872 | 0.005775 | 6.7e-05 | 0.001128 | 0.27024 | 1.3334 | 2.5482 | 100.58 | 20.72 | 19.224 | 96.8396 | 112.0086 | 91.0174 | 0.004677 | 0.96946 | 0.023394 | 24.1315 | -0.068146 | 0.15539 | 0.002576 | 0.17582 | 0.17582 | 0.036325
2019-03-20 17:58 | 1553119101.0829 | 561 | 0.20906 | 0.000727 | 0.000101 | 108.1891 | 109.1328 | 117.5033 | 21.0663 | 0.003123 | 0.002897 | 131.5648 | 0.25961 | 8.4041 | 0.003799 | 0.005281 | 0.005137 | 6.7e-05 | 0.001126 | 0.28622 | 1.313 | 2.548 | 100.58 | 20.72 | 19.234 | 96.7863 | 113.1642 | 91.327 | 0.00468 | 0.96864 | 0.023336 | 24.0919 | -0.07768 | 0.15444 | 0.002387 | 0.16212 | 0.16212 | 0.039154
2019-03-20 17:58 | 1553119102.083 | 562 | 0.209 | 0.000733 | 0.0001 | 108.3185 | 109.2715 | 117.2372 | 21.0672 | 0.003127 | 0.002954 | 131.5559 | 0.25951 | 8.4032 | 0.00379 | 0.005516 | 0.005411 | 6.8e-05 | 0.00113 | 0.31293 | 1.2854 | 2.5481 | 100.58 | 20.72 | 19.245 | 96.8695 | 112.9581 | 92.6948 | 0.004683 | 0.96501 | 0.023266 | 24.1097 | -0.078349 | 0.15438 | 0.002724 | 0.18529 | 0.18529 | 0.034244
2019-03-20 17:58 | 1553119103.0824 | 563 | 0.20903 | 0.000724 | 0.0001 | 108.2406 | 109.091 | 117.3491 | 21.0655 | 0.003107 | 0.002913 | 131.5775 | 0.26012 | 8.4024 | 0.003782 | 0.004994 | 0.004851 | 6.3e-05 | 0.00113 | 0.28854 | 1.2925 | 2.5481 | 100.58 | 20.72 | 19.266 | 96.4002 | 112.7463 | 91.7217 | 0.004688 | 0.96759 | 0.023308 | 24.0891 | -0.069998 | 0.15474 | 0.002577 | 0.17536 | 0.17536 | 0.036267
2019-03-20 17:58 | 1553119104.0824 | 564 | 0.20903 | 0.000723 | 9.9e-05 | 108.2496 | 109.1556 | 117.3869 | 21.0664 | 0.003096 | 0.002921 | 131.5885 | 0.26079 | 8.4023 | 0.003803 | 0.004595 | 0.004482 | 6.8e-05 | 0.001132 | 0.28279 | 1.3753 | 2.5482 | 100.58 | 20.72 | 19.275 | 96.1142 | 113.2575 | 91.8922 | 0.00469 | 0.96714 | 0.023307 | 24.0986 | -0.07455 | 0.15514 | 0.002593 | 0.17649 | 0.17649 | 0.036129
2019-03-20 17:58 | 1553119105.0832 | 565 | 0.20905 | 0.000724 | 9.8e-05 | 108.2655 | 109.2199 | 117.2001 | 21.0663 | 0.003121 | 0.002915 | 131.5511 | 0.26011 | 8.4032 | 0.003804 | 0.005149 | 0.004978 | 6.7e-05 | 0.001131 | 0.2724 | 1.3618 | 2.548 | 100.58 | 20.72 | 19.285 | 96.7392 | 113.2779 | 91.7546 | 0.004692 | 0.9675 | 0.023293 | 24.0749 | -0.078507 | 0.15474 | 0.00248 | 0.16854 | 0.16854 | 0.037734
2019-03-20 17:58 | 1553119106.0831 | 566 | 0.20907 | 0.000717 | 9.8e-05 | 108.2962 | 109.2819 | 117.4203 | 21.0657 | 0.003154 | 0.002913 | 131.5492 | 0.26072 | 8.403 | 0.003788 | 0.005219 | 0.005132 | 6.4e-05 | 0.001134 | 0.28816 | 1.3143 | 2.548 | 100.58 | 20.72 | 19.3 | 97.5301 | 112.9053 | 91.7104 | 0.004696 | 0.96762 | 0.023316 | 24.0961 | -0.081048 | 0.1551 | 0.002352 | 0.15968 | 0.15968 | 0.039922
2019-03-20 17:58 | 1553119107.0823 | 567 | 0.20908 | 0.000714 | 9.8e-05 | 108.2865 | 109.3488 | 117.5156 | 21.0667 | 0.003115 | 0.00294 | 131.5455 | 0.26049 | 8.4032 | 0.00375 | 0.005704 | 0.005581 | 6.8e-05 | 0.00113 | 0.27848 | 1.3233 | 2.5481 | 100.58 | 20.72 | 19.307 | 96.5954 | 111.9625 | 92.3705 | 0.004698 | 0.96587 | 0.023304 | 24.1276 | -0.087324 | 0.15496 | 0.002316 | 0.15713 | 0.15713 | 0.040534
2019-03-20 17:58 | 1553119108.0824 | 568 | 0.20905 | 0.000715 | 9.8e-05 | 108.4049 | 109.4464 | 117.6066 | 21.066 | 0.003079 | 0.002853 | 131.5573 | 0.26078 | 8.4047 | 0.003781 | 0.005443 | 0.005301 | 6.6e-05 | 0.001129 | 0.29679 | 1.3854 | 2.5483 | 100.58 | 20.72 | 19.317 | 95.7066 | 112.7349 | 90.2591 | 0.0047 | 0.97148 | 0.023381 | 24.0672 | -0.085528 | 0.15514 | 0.002473 | 0.16866 | 0.16866 | 0.037803
2019-03-20 17:58 | 1553119109.0828 | 569 | 0.20903 | 0.00071 | 9.8e-05 | 108.2777 | 109.3316 | 118.088 | 21.0654 | 0.00309 | 0.002848 | 131.5821 | 0.25968 | 8.4038 | 0.003735 | 0.005101 | 0.004895 | 7e-05 | 0.001121 | 0.27129 | 1.3095 | 2.548 | 100.58 | 20.72 | 19.321 | 95.9771 | 111.6013 | 90.1363 | 0.004701 | 0.97181 | 0.023433 | 24.1123 | -0.086639 | 0.15449 | 0.002591 | 0.17738 | 0.17738 | 0.035795
2019-03-20 17:58 | 1553119110.0825 | 570 | 0.20902 | 0.000705 | 9.8e-05 | 108.4032 | 109.307 | 118.4148 | 21.0659 | 0.003122 | 0.002848 | 131.5986 | 0.26147 | 8.4002 | 0.00369 | 0.005515 | 0.005398 | 6.8e-05 | 0.001118 | 0.27688 | 1.3492 | 2.5481 | 100.58 | 20.72 | 19.324 | 96.7562 | 110.5198 | 90.1378 | 0.004702 | 0.97181 | 0.023465 | 24.1457 | -0.07426 | 0.15555 | 0.002635 | 0.18071 | 0.18071 | 0.035378
2019-03-20 17:58 | 1553119111.0828 | 571 | 0.20904 | 0.000696 | 9.8e-05 | 108.3276 | 109.4688 | 118.6798 | 21.0657 | 0.003097 | 0.002863 | 131.6061 | 0.26042 | 8.4 | 0.003696 | 0.004946 | 0.004871 | 6.8e-05 | 0.001114 | 0.27959 | 1.3361 | 2.5478 | 100.58 | 20.73 | 19.332 | 96.1393 | 110.6545 | 90.5157 | 0.004707 | 0.9708 | 0.023479 | 24.1853 | -0.09374 | 0.15492 | 0.002567 | 0.17602 | 0.17602 | 0.036175
2019-03-20 17:59 | 1553119112.0821 | 572 | 0.20905 | 0.000703 | 9.9e-05 | 108.3799 | 109.3888 | 118.7072 | 21.0665 | 0.00309 | 0.002892 | 131.6118 | 0.26054 | 8.4017 | 0.00369 | 0.004604 | 0.004442 | 7.1e-05 | 0.001116 | 0.28335 | 1.3683 | 2.5482 | 100.58 | 20.72 | 19.336 | 95.9757 | 110.5118 | 91.1966 | 0.004705 | 0.96898 | 0.02346 | 24.2107 | -0.082877 | 0.15499 | 0.002479 | 0.16961 | 0.16961 | 0.037559
2019-03-20 17:59 | 1553119113.0831 | 573 | 0.20907 | 0.000695 | 9.8e-05 | 108.4052 | 109.5817 | 118.8734 | 21.0662 | 0.003082 | 0.002879 | 131.6086 | 0.2605 | 8.4035 | 0.003771 | 0.00454 | 0.004413 | 7e-05 | 0.00112 | 0.30093 | 1.3613 | 2.548 | 100.58 | 20.73 | 19.343 | 95.778 | 112.4792 | 90.8882 | 0.004709 | 0.9698 | 0.023486 | 24.2174 | -0.096551 | 0.15497 | 0.002404 | 0.16453 | 0.16453 | 0.038714
2019-03-20 17:59 | 1553119114.0827 | 574 | 0.20908 | 0.000696 | 9.9e-05 | 108.4986 | 109.6348 | 119.0603 | 21.0662 | 0.00307 | 0.002845 | 131.5992 | 0.26024 | 8.4009 | 0.003715 | 0.004711 | 0.004546 | 7.1e-05 | 0.001122 | 0.29302 | 1.3397 | 2.548 | 100.58 | 20.73 | 19.346 | 95.4985 | 111.1246 | 90.0629 | 0.00471 | 0.97201 | 0.023531 | 24.2089 | -0.093174 | 0.15482 | 0.002299 | 0.15739 | 0.15739 | 0.040429
2019-03-20 17:59 | 1553119115.0827 | 575 | 0.20907 | 0.000694 | 9.8e-05 | 108.5058 | 109.4101 | 119.4602 | 21.0669 | 0.003012 | 0.00283 | 131.6136 | 0.26088 | 8.4033 | 0.003767 | 0.005512 | 0.005319 | 7.4e-05 | 0.001124 | 0.29446 | 1.3086 | 2.5481 | 100.58 | 20.73 | 19.351 | 94.1066 | 112.3772 | 89.7221 | 0.004711 | 0.97292 | 0.023582 | 24.2382 | -0.074242 | 0.1552 | 0.002349 | 0.1613 | 0.1613 | 0.039544
2019-03-20 17:59 | 1553119116.0828 | 576 | 0.20907 | 0.000694 | 9.8e-05 | 108.4915 | 109.7162 | 119.5438 | 21.0662 | 0.003109 | 0.002831 | 131.6039 | 0.26048 | 8.4027 | 0.003743 | 0.005404 | 0.005379 | 7.1e-05 | 0.001125 | 0.27295 | 1.2924 | 2.5482 | 100.58 | 20.73 | 19.359 | 96.4508 | 111.809 | 89.7395 | 0.004713 | 0.97287 | 0.023589 | 24.2472 | -0.1004 | 0.15496 | 0.002392 | 0.1644 | 0.1644 | 0.03874
2019-03-20 17:59 | 1553119117.0819 | 577 | 0.20908 | 0.000689 | 9.8e-05 | 108.5054 | 109.649 | 119.7403 | 21.0669 | 0.003038 | 0.002835 | 131.5915 | 0.26068 | 8.4024 | 0.003801 | 0.005473 | 0.00537 | 7.1e-05 | 0.001123 | 0.28385 | 1.31 | 2.5481 | 100.57 | 20.73 | 19.364 | 94.7178 | 113.2144 | 89.8275 | 0.004715 | 0.97264 | 0.023606 | 24.2701 | -0.093772 | 0.15508 | 0.002335 | 0.16049 | 0.16049 | 0.039714
2019-03-20 17:59 | 1553119118.0831 | 578 | 0.20908 | 0.000689 | 9.7e-05 | 108.5054 | 109.5697 | 119.5811 | 21.0663 | 0.003079 | 0.002849 | 131.5842 | 0.25977 | 8.4033 | 0.003781 | 0.004893 | 0.004725 | 7.1e-05 | 0.001124 | 0.26648 | 1.3376 | 2.5481 | 100.58 | 20.73 | 19.373 | 95.7038 | 112.7257 | 90.1591 | 0.004717 | 0.97175 | 0.02358 | 24.265 | -0.087306 | 0.15454 | 0.002342 | 0.16085 | 0.16085 | 0.039487
2019-03-20 17:59 | 1553119119.082 | 579 | 0.20904 | 0.00068 | 9.6e-05 | 108.5792 | 109.5735 | 119.5633 | 21.0655 | 0.00305 | 0.002854 | 131.5733 | 0.26009 | 8.4032 | 0.003801 | 0.004629 | 0.004486 | 7.3e-05 | 0.001123 | 0.29734 | 1.3451 | 2.5481 | 100.58 | 20.73 | 19.379 | 95.0078 | 113.218 | 90.2802 | 0.004718 | 0.97142 | 0.023574 | 24.2673 | -0.081535 | 0.15473 | 0.002553 | 0.17583 | 0.17583 | 0.036167
2019-03-20 17:59 | 1553119120.083 | 580 | 0.20907 | 0.000693 | 9.6e-05 | 108.5898 | 109.8534 | 119.4731 | 21.0666 | 0.00303 | 0.002821 | 131.5698 | 0.26046 | 8.4014 | 0.003823 | 0.004841 | 0.004728 | 7.2e-05 | 0.00112 | 0.28119 | 1.3595 | 2.5482 | 100.58 | 20.73 | 19.384 | 94.532 | 113.7612 | 89.4913 | 0.004719 | 0.97354 | 0.023591 | 24.2318 | -0.10348 | 0.15495 | 0.002385 | 0.16406 | 0.16406 | 0.038817
2019-03-20 17:59 | 1553119121.082 | 581 | 0.20908 | 0.000686 | 9.5e-05 | 108.6257 | 109.6878 | 119.8351 | 21.0667 | 0.003035 | 0.002847 | 131.5751 | 0.26051 | 8.4023 | 0.003756 | 0.005117 | 0.004971 | 7.4e-05 | 0.001121 | 0.26465 | 1.3547 | 2.5479 | 100.58 | 20.73 | 19.392 | 94.6621 | 112.1085 | 90.1204 | 0.004721 | 0.97185 | 0.023606 | 24.2895 | -0.08703 | 0.15498 | 0.002339 | 0.16094 | 0.16094 | 0.039577
2019-03-20 17:59 | 1553119122.0831 | 582 | 0.20908 | 0.000682 | 9.4e-05 | 108.6943 | 109.7506 | 120.0527 | 21.0659 | 0.003005 | 0.002804 | 131.5838 | 0.26143 | 8.4015 | 0.003748 | 0.005422 | 0.005277 | 7.2e-05 | 0.001124 | 0.28031 | 1.3307 | 2.5481 | 100.58 | 20.73 | 19.395 | 93.9302 | 111.9276 | 89.082 | 0.004722 | 0.97464 | 0.023661 | 24.2768 | -0.086505 | 0.15552 | 0.002318 | 0.15986 | 0.15986 | 0.039985
2019-03-20 17:59 | 1553119123.0835 | 583 | 0.20911 | 0.000683 | 9.3e-05 | 108.6115 | 109.6725 | 119.9147 | 21.0666 | 0.003006 | 0.002854 | 131.5628 | 0.2604 | 8.4031 | 0.00372 | 0.005497 | 0.005304 | 7.4e-05 | 0.001125 | 0.28672 | 1.3384 | 2.5481 | 100.58 | 20.73 | 19.398 | 93.9553 | 111.2337 | 90.2791 | 0.004723 | 0.97143 | 0.023608 | 24.3029 | -0.086953 | 0.15491 | 0.002129 | 0.14612 | 0.14612 | 0.043572
2019-03-20 17:59 | 1553119124.0822 | 584 | 0.20911 | 0.000683 | 9.3e-05 | 108.6591 | 109.686 | 119.9701 | 21.0661 | 0.003008 | 0.002803 | 131.568 | 0.26039 | 8.4011 | 0.003732 | 0.005274 | 0.005141 | 7.3e-05 | 0.001127 | 0.2855 | 1.3042 | 2.5481 | 100.58 | 20.73 | 19.4 | 94.0081 | 111.5434 | 89.0574 | 0.004723 | 0.97471 | 0.023654 | 24.2676 | -0.084136 | 0.1549 | 0.002151 | 0.14799 | 0.14799 | 0.043019
2019-03-20 17:59 | 1553119125.0823 | 585 | 0.20911 | 0.000681 | 9.2e-05 | 108.6623 | 109.7551 | 119.8107 | 21.0654 | 0.003044 | 0.002868 | 131.5669 | 0.2603 | 8.4025 | 0.00376 | 0.005602 | 0.005446 | 7.8e-05 | 0.001127 | 0.26073 | 1.3057 | 2.5481 | 100.58 | 20.73 | 19.404 | 94.8653 | 112.2156 | 90.6164 | 0.004724 | 0.97053 | 0.023587 | 24.3036 | -0.089506 | 0.15485 | 0.00219 | 0.15036 | 0.15036 | 0.042329
2019-03-20 17:59 | 1553119126.0813 | 586 | 0.20913 | 0.000677 | 9.3e-05 | 108.6354 | 109.7742 | 119.6731 | 21.0664 | 0.003059 | 0.002807 | 131.5701 | 0.26173 | 8.4018 | 0.003766 | 0.004971 | 0.004888 | 7.3e-05 | 0.001129 | 0.27367 | 1.3034 | 2.5482 | 100.58 | 20.73 | 19.405 | 95.2412 | 112.3689 | 89.1576 | 0.004724 | 0.97444 | 0.023621 | 24.2409 | -0.093269 | 0.1557 | 0.002011 | 0.13777 | 0.13777 | 0.04645
2019-03-20 17:59 | 1553119127.0813 | 587 | 0.20907 | 0.000678 | 9.3e-05 | 108.7473 | 109.9536 | 119.514 | 21.0671 | 0.003047 | 0.002827 | 131.5913 | 0.26113 | 8.4035 | 0.003688 | 0.005111 | 0.005019 | 7.3e-05 | 0.00113 | 0.28047 | 1.282 | 2.5481 | 100.58 | 20.73 | 19.406 | 94.9465 | 110.4594 | 89.631 | 0.004725 | 0.97316 | 0.02359 | 24.2406 | -0.098666 | 0.15535 | 0.002386 | 0.16429 | 0.16429 | 0.038864
2019-03-20 17:59 | 1553119128.0814 | 588 | 0.2091 | 0.000667 | 9.2e-05 | 108.7648 | 109.8915 | 119.3396 | 21.0658 | 0.003013 | 0.002825 | 131.5884 | 0.26108 | 8.4023 | 0.0038 | 0.005404 | 0.005184 | 7.1e-05 | 0.001125 | 0.28423 | 1.2664 | 2.5482 | 100.58 | 20.73 | 19.406 | 94.126 | 113.1803 | 89.5845 | 0.004725 | 0.97329 | 0.023574 | 24.2213 | -0.092182 | 0.15532 | 0.002231 | 0.1532 | 0.1532 | 0.041668
2019-03-20 17:59 | 1553119129.0814 | 589 | 0.20911 | 0.000669 | 9.3e-05 | 108.7522 | 109.9749 | 119.2044 | 21.0661 | 0.003032 | 0.002781 | 131.6061 | 0.26074 | 8.4011 | 0.003757 | 0.005019 | 0.004875 | 7.5e-05 | 0.001123 | 0.29103 | 1.308 | 2.548 | 100.58 | 20.73 | 19.405 | 94.5699 | 112.1381 | 88.5385 | 0.004724 | 0.9761 | 0.023595 | 24.1727 | -0.1 | 0.15511 | 0.002158 | 0.14814 | 0.14814 | 0.043033
2019-03-20 17:59 | 1553119130.0816 | 590 | 0.20912 | 0.000669 | 9.1e-05 | 108.7376 | 109.8687 | 119.1948 | 21.0664 | 0.002994 | 0.002886 | 131.6061 | 0.26024 | 8.4011 | 0.003761 | 0.005534 | 0.005379 | 7.3e-05 | 0.001126 | 0.27776 | 1.2999 | 2.5481 | 100.58 | 20.73 | 19.401 | 93.6703 | 112.2314 | 91.0542 | 0.004723 | 0.96936 | 0.023512 | 24.2556 | -0.09256 | 0.15482 | 0.002135 | 0.14604 | 0.14604 | 0.043571
2019-03-20 17:59 | 1553119131.0826 | 591 | 0.2091 | 0.000668 | 9e-05 | 108.76 | 109.9083 | 118.8643 | 21.0665 | 0.003003 | 0.002849 | 131.5926 | 0.26083 | 8.4038 | 0.00382 | 0.005127 | 0.005005 | 7.5e-05 | 0.001127 | 0.25426 | 1.3242 | 2.5481 | 100.58 | 20.73 | 19.398 | 93.8816 | 113.6661 | 90.1672 | 0.004723 | 0.97173 | 0.023508 | 24.1925 | -0.093949 | 0.15517 | 0.00224 | 0.15353 | 0.15353 | 0.041537
2019-03-20 17:59 | 1553119132.0821 | 592 | 0.20912 | 0.000665 | 8.9e-05 | 108.8586 | 109.9218 | 118.7556 | 21.0665 | 0.002904 | 0.002843 | 131.5801 | 0.26055 | 8.4021 | 0.003954 | 0.004829 | 0.004732 | 7.4e-05 | 0.00113 | 0.27577 | 1.377 | 2.5481 | 100.58 | 20.73 | 19.394 | 91.4913 | 116.9347 | 90.028 | 0.004722 | 0.9721 | 0.023502 | 24.1768 | -0.086935 | 0.155 | 0.002151 | 0.14725 | 0.14725 | 0.043264
2019-03-20 17:59 | 1553119133.0826 | 593 | 0.20914 | 0.000665 | 8.7e-05 | 108.8638 | 110.0971 | 118.8971 | 21.0663 | 0.003059 | 0.002778 | 131.5895 | 0.26169 | 8.3989 | 0.003828 | 0.004975 | 0.004785 | 7.5e-05 | 0.001134 | 0.29645 | 1.2989 | 2.548 | 100.58 | 20.73 | 19.385 | 95.2433 | 113.8754 | 88.4538 | 0.00472 | 0.97633 | 0.023567 | 24.1387 | -0.10076 | 0.15568 | 0.001991 | 0.13637 | 0.13637 | 0.046919
2019-03-20 17:59 | 1553119134.0821 | 594 | 0.20918 | 0.000664 | 8.6e-05 | 108.8246 | 109.97 | 118.9152 | 21.066 | 0.002916 | 0.002852 | 131.5433 | 0.25999 | 8.4034 | 0.003836 | 0.005518 | 0.00535 | 7.3e-05 | 0.001145 | 0.29264 | 1.3182 | 2.5481 | 100.58 | 20.73 | 19.38 | 91.788 | 114.0673 | 90.2472 | 0.004718 | 0.97151 | 0.023511 | 24.2003 | -0.093649 | 0.15467 | 0.001776 | 0.12085 | 0.12085 | 0.052604
2019-03-20 17:59 | 1553119135.0818 | 595 | 0.20915 | 0.000666 | 8.6e-05 | 108.8777 | 109.9685 | 118.8866 | 21.0657 | 0.002974 | 0.002815 | 131.5369 | 0.26102 | 8.4044 | 0.003796 | 0.005543 | 0.005466 | 7.4e-05 | 0.001144 | 0.26465 | 1.3218 | 2.5482 | 100.58 | 20.74 | 19.37 | 93.1885 | 113.0981 | 89.3554 | 0.004719 | 0.9739 | 0.023537 | 24.1677 | -0.089169 | 0.15528 | 0.001969 | 0.1347 | 0.1347 | 0.047381
2019-03-20 17:59 | 1553119136.0828 | 596 | 0.20915 | 0.00066 | 8.5e-05 | 108.9087 | 110.0242 | 118.7894 | 21.0668 | 0.00296 | 0.002834 | 131.5222 | 0.26102 | 8.4026 | 0.003777 | 0.00467 | 0.004577 | 7.6e-05 | 0.001145 | 0.28456 | 1.3398 | 2.5479 | 100.58 | 20.74 | 19.364 | 92.8453 | 112.6205 | 89.801 | 0.004717 | 0.97271 | 0.023513 | 24.1726 | -0.091147 | 0.15528 | 0.001965 | 0.13433 | 0.13433 | 0.047509
2019-03-20 17:59 | 1553119137.0828 | 597 | 0.20918 | 0.000663 | 8.4e-05 | 108.932 | 110.1442 | 118.584 | 21.0657 | 0.002959 | 0.00276 | 131.5193 | 0.26051 | 8.4026 | 0.003756 | 0.005205 | 0.005056 | 7.3e-05 | 0.001144 | 0.24762 | 1.3333 | 2.548 | 100.58 | 20.73 | 19.359 | 92.8254 | 112.1282 | 88.0203 | 0.004713 | 0.9775 | 0.02355 | 24.0924 | -0.098982 | 0.15498 | 0.001748 | 0.11916 | 0.11916 | 0.053452
2019-03-20 17:59 | 1553119138.0819 | 598 | 0.20916 | 0.000658 | 8.4e-05 | 108.9634 | 110.1397 | 118.6662 | 21.0663 | 0.002967 | 0.002746 | 131.5056 | 0.26098 | 8.401 | 0.003811 | 0.005055 | 0.004905 | 7.7e-05 | 0.001144 | 0.27201 | 1.3781 | 2.548 | 100.58 | 20.74 | 19.348 | 93.0169 | 113.4628 | 87.6775 | 0.004713 | 0.97843 | 0.02357 | 24.0893 | -0.096043 | 0.15526 | 0.00192 | 0.13156 | 0.13156 | 0.048505
2019-03-20 17:59 | 1553119139.0834 | 599 | 0.20916 | 0.00066 | 8.5e-05 | 108.9385 | 110.062 | 118.7265 | 21.0665 | 0.002999 | 0.002808 | 131.4846 | 0.26118 | 8.4022 | 0.003674 | 0.005059 | 0.004974 | 7.5e-05 | 0.001144 | 0.29789 | 1.3427 | 2.548 | 100.58 | 20.74 | 19.343 | 93.7878 | 110.1323 | 89.1811 | 0.004712 | 0.97437 | 0.023527 | 24.1456 | -0.091772 | 0.15538 | 0.001872 | 0.12787 | 0.12787 | 0.04994
2019-03-20 17:59 | 1553119140.0821 | 600 | 0.20915 | 0.00065 | 8.4e-05 | 108.9851 | 110.1401 | 118.6683 | 21.0653 | 0.002956 | 0.002783 | 131.4781 | 0.26027 | 8.4032 | 0.003691 | 0.005694 | 0.005557 | 7.3e-05 | 0.00114 | 0.27079 | 1.3248 | 2.5482 | 100.58 | 20.74 | 19.338 | 92.7368 | 110.5405 | 88.577 | 0.004711 | 0.976 | 0.023541 | 24.1195 | -0.094295 | 0.15484 | 0.001951 | 0.13358 | 0.13358 | 0.04764
2019-03-20 17:59 | 1553119141.0831 | 601 | 0.20916 | 0.000656 | 8.4e-05 | 108.9511 | 110.078 | 118.5946 | 21.0651 | 0.002971 | 0.002775 | 131.4717 | 0.26142 | 8.4009 | 0.003678 | 0.005567 | 0.005413 | 7.5e-05 | 0.001137 | 0.26808 | 1.2829 | 2.5482 | 100.58 | 20.74 | 19.33 | 93.1049 | 110.2121 | 88.3962 | 0.004709 | 0.97649 | 0.023539 | 24.106 | -0.092042 | 0.15552 | 0.001881 | 0.12861 | 0.12861 | 0.049701
2019-03-20 17:59 | 1553119142.0818 | 602 | 0.2092 | 0.000657 | 8.4e-05 | 109.0329 | 110.3415 | 118.6493 | 21.0658 | 0.002907 | 0.002744 | 131.474 | 0.26082 | 8.4039 | 0.00366 | 0.005882 | 0.005757 | 7.3e-05 | 0.00114 | 0.26394 | 1.3301 | 2.5481 | 100.58 | 20.74 | 19.326 | 91.5781 | 109.7893 | 87.637 | 0.004708 | 0.97854 | 0.023569 | 24.0863 | -0.10671 | 0.15516 | 0.001662 | 0.1132 | 0.1132 | 0.056336
2019-03-20 17:59 | 1553119143.0822 | 603 | 0.20915 | 0.000655 | 8.4e-05 | 108.9889 | 110.266 | 118.6138 | 21.0662 | 0.002972 | 0.00273 | 131.4583 | 0.26116 | 8.4013 | 0.003707 | 0.006232 | 0.0061 | 7.5e-05 | 0.001139 | 0.28136 | 1.357 | 2.5481 | 100.58 | 20.74 | 19.319 | 93.1294 | 110.9295 | 87.3118 | 0.004706 | 0.97942 | 0.023577 | 24.0718 | -0.1042 | 0.15536 | 0.001957 | 0.13425 | 0.13425 | 0.047564
2019-03-20 17:59 | 1553119144.0826 | 604 | 0.20918 | 0.000661 | 8.3e-05 | 109.0866 | 110.3613 | 118.4918 | 21.0661 | 0.002953 | 0.002782 | 131.4745 | 0.2613 | 8.4006 | 0.003732 | 0.004754 | 0.004686 | 7.2e-05 | 0.001135 | 0.26936 | 1.3244 | 2.5481 | 100.58 | 20.74 | 19.317 | 92.6787 | 111.5465 | 88.5486 | 0.004706 | 0.97608 | 0.023524 | 24.1007 | -0.10391 | 0.15545 | 0.00179 | 0.12212 | 0.12212 | 0.052316
2019-03-20 17:59 | 1553119145.0814 | 605 | 0.20916 | 0.000651 | 8.2e-05 | 109.0746 | 110.235 | 118.5434 | 21.0651 | 0.002899 | 0.002776 | 131.4493 | 0.26055 | 8.4018 | 0.003818 | 0.005472 | 0.005235 | 7.2e-05 | 0.001137 | 0.30879 | 1.3076 | 2.5481 | 100.58 | 20.74 | 19.315 | 91.3649 | 113.6277 | 88.4165 | 0.004705 | 0.97643 | 0.023533 | 24.1015 | -0.094651 | 0.155 | 0.001942 | 0.13297 | 0.13297 | 0.047912
2019-03-20 17:59 | 1553119146.0816 | 606 | 0.20917 | 0.000653 | 8.2e-05 | 109.1497 | 110.3972 | 118.4669 | 21.0654 | 0.003003 | 0.002711 | 131.4525 | 0.26067 | 8.4044 | 0.003856 | 0.004967 | 0.004839 | 7.2e-05 | 0.001132 | 0.28136 | 1.3053 | 2.5481 | 100.58 | 20.74 | 19.314 | 93.8687 | 114.5553 | 86.8355 | 0.004705 | 0.98072 | 0.023578 | 24.041 | -0.10165 | 0.15507 | 0.001869 | 0.12805 | 0.12805 | 0.049772
2019-03-20 17:59 | 1553119147.0824 | 607 | 0.20916 | 0.000653 | 8.3e-05 | 109.0911 | 110.4129 | 118.7077 | 21.066 | 0.002838 | 0.002738 | 131.4236 | 0.26076 | 8.4026 | 0.003906 | 0.005497 | 0.005372 | 7.1e-05 | 0.001132 | 0.29231 | 1.3649 | 2.5482 | 100.58 | 20.74 | 19.313 | 89.8987 | 115.7693 | 87.489 | 0.004705 | 0.97894 | 0.02358 | 24.0872 | -0.10772 | 0.15513 | 0.001909 | 0.13087 | 0.13087 | 0.048717
2019-03-20 17:59 | 1553119148.0815 | 608 | 0.2092 | 0.000654 | 8.2e-05 | 109.1872 | 110.528 | 118.8267 | 21.0662 | 0.002963 | 0.002704 | 131.4294 | 0.26039 | 8.4036 | 0.003898 | 0.005355 | 0.005249 | 7.1e-05 | 0.001132 | 0.2724 | 1.3532 | 2.5482 | 100.58 | 20.74 | 19.313 | 92.9199 | 115.5778 | 86.6825 | 0.004705 | 0.98114 | 0.023618 | 24.0724 | -0.10917 | 0.15491 | 0.001673 | 0.11433 | 0.11433 | 0.055688
2019-03-20 17:59 | 1553119149.0826 | 609 | 0.20915 | 0.000653 | 8.2e-05 | 109.2063 | 110.497 | 118.9499 | 21.0665 | 0.002935 | 0.002785 | 131.4203 | 0.26082 | 8.4028 | 0.00381 | 0.00536 | 0.0052 | 7.3e-05 | 0.001133 | 0.28279 | 1.3344 | 2.5481 | 100.58 | 20.74 | 19.315 | 92.2514 | 113.4315 | 88.6391 | 0.004705 | 0.97583 | 0.023567 | 24.1502 | -0.1051 | 0.15516 | 0.001951 | 0.13383 | 0.13383 | 0.04765
2019-03-20 17:59 | 1553119150.0816 | 610 | 0.20918 | 0.000649 | 8.1e-05 | 109.2309 | 110.636 | 119.2126 | 21.0663 | 0.002935 | 0.002754 | 131.4058 | 0.26032 | 8.4035 | 0.003932 | 0.005736 | 0.005588 | 7.4e-05 | 0.001133 | 0.27632 | 1.3365 | 2.5481 | 100.58 | 20.74 | 19.316 | 92.2534 | 116.4184 | 87.8741 | 0.004706 | 0.9779 | 0.023618 | 24.1513 | -0.11433 | 0.15487 | 0.001821 | 0.1249 | 0.1249 | 0.05096
2019-03-20 17:59 | 1553119151.0816 | 611 | 0.20915 | 0.000648 | 8.1e-05 | 109.1726 | 110.3849 | 119.2835 | 21.0664 | 0.002873 | 0.00275 | 131.389 | 0.26051 | 8.4039 | 0.003886 | 0.005986 | 0.005852 | 7.3e-05 | 0.001137 | 0.27345 | 1.3594 | 2.5481 | 100.58 | 20.74 | 19.32 | 90.7494 | 115.2932 | 87.7955 | 0.004707 | 0.97811 | 0.023627 | 24.1558 | -0.098772 | 0.15498 | 0.001974 | 0.13585 | 0.13585 | 0.046887
2019-03-20 17:59 | 1553119152.0827 | 612 | 0.20916 | 0.000653 | 8e-05 | 109.3518 | 110.676 | 119.4692 | 21.0658 | 0.002962 | 0.002725 | 131.3701 | 0.26046 | 8.4002 | 0.003848 | 0.005023 | 0.004823 | 7.5e-05 | 0.001137 | 0.26537 | 1.3493 | 2.548 | 100.58 | 20.74 | 19.323 | 92.8851 | 114.3532 | 87.1803 | 0.004707 | 0.97978 | 0.023666 | 24.154 | -0.10766 | 0.15495 | 0.001942 | 0.13388 | 0.13388 | 0.047566
2019-03-20 17:59 | 1553119153.0831 | 613 | 0.20914 | 0.000653 | 8.1e-05 | 109.3757 | 110.7056 | 119.5947 | 21.0668 | 0.002852 | 0.002735 | 131.3905 | 0.26111 | 8.4023 | 0.003927 | 0.005032 | 0.00494 | 7.1e-05 | 0.00114 | 0.26001 | 1.3437 | 2.548 | 100.58 | 20.74 | 19.326 | 90.233 | 116.2962 | 87.4351 | 0.004708 | 0.97909 | 0.02367 | 24.1753 | -0.10809 | 0.15533 | 0.002004 | 0.13834 | 0.13834 | 0.046149
2019-03-20 17:59 | 1553119154.083 | 614 | 0.20914 | 0.000654 | 8e-05 | 109.2571 | 110.5116 | 119.7276 | 21.0661 | 0.002936 | 0.002712 | 131.4252 | 0.26145 | 8.401 | 0.003865 | 0.005925 | 0.005739 | 7.8e-05 | 0.001143 | 0.27903 | 1.3153 | 2.548 | 100.58 | 20.74 | 19.334 | 92.2634 | 114.7734 | 86.8612 | 0.00471 | 0.98065 | 0.023702 | 24.1694 | -0.10212 | 0.15553 | 0.002043 | 0.14134 | 0.14134 | 0.045228
2019-03-20 17:59 | 1553119155.0829 | 615 | 0.20918 | 0.000648 | 8e-05 | 109.3046 | 110.679 | 119.7614 | 21.0659 | 0.002946 | 0.002763 | 131.4242 | 0.26123 | 8.4021 | 0.003816 | 0.005548 | 0.005438 | 7.2e-05 | 0.001143 | 0.27113 | 1.3303 | 2.5482 | 100.58 | 20.74 | 19.34 | 92.5151 | 113.5849 | 88.1003 | 0.004711 | 0.97729 | 0.023664 | 24.2144 | -0.11177 | 0.15541 | 0.001777 | 0.12206 | 0.12206 | 0.05233
2019-03-20 17:59 | 1553119156.0828 | 616 | 0.20914 | 0.000653 | 8e-05 | 109.2955 | 110.6738 | 119.9114 | 21.067 | 0.002938 | 0.002703 | 131.4428 | 0.26041 | 8.4023 | 0.003796 | 0.005618 | 0.005553 | 7.3e-05 | 0.001147 | 0.27151 | 1.3407 | 2.5482 | 100.58 | 20.74 | 19.349 | 92.3237 | 113.0835 | 86.6551 | 0.004714 | 0.98121 | 0.023727 | 24.181 | -0.11209 | 0.15492 | 0.002011 | 0.13916 | 0.13916 | 0.045756
2019-03-20 17:59 | 1553119157.083 | 617 | 0.20917 | 0.000651 | 8e-05 | 109.3554 | 110.7284 | 119.894 | 21.0663 | 0.002901 | 0.002731 | 131.4552 | 0.26011 | 8.4031 | 0.003846 | 0.005438 | 0.005301 | 7.2e-05 | 0.001144 | 0.29302 | 1.3572 | 2.5482 | 100.58 | 20.74 | 19.35 | 91.4207 | 114.3124 | 87.3268 | 0.004714 | 0.97938 | 0.023703 | 24.2019 | -0.1116 | 0.15474 | 0.001824 | 0.12565 | 0.12565 | 0.050615
2019-03-20 17:59 | 1553119158.081 | 618 | 0.20912 | 0.000649 | 8e-05 | 109.3832 | 110.7284 | 120.0522 | 21.0658 | 0.002937 | 0.002698 | 131.4487 | 0.26082 | 8.4029 | 0.003847 | 0.004687 | 0.004539 | 7.4e-05 | 0.001139 | 0.27295 | 1.3256 | 2.5481 | 100.58 | 20.74 | 19.35 | 92.2773 | 114.343 | 86.5433 | 0.004714 | 0.98152 | 0.023744 | 24.1915 | -0.10933 | 0.15516 | 0.002141 | 0.14859 | 0.14859 | 0.042919
2019-03-20 17:59 | 1553119159.0815 | 619 | 0.20914 | 0.000647 | 8.1e-05 | 109.3806 | 110.7758 | 119.9283 | 21.0648 | 0.002884 | 0.002717 | 131.4586 | 0.26102 | 8.4047 | 0.003849 | 0.005198 | 0.005043 | 7.3e-05 | 0.001139 | 0.29031 | 1.3205 | 2.5481 | 100.58 | 20.74 | 19.349 | 91.0213 | 114.3727 | 86.9841 | 0.004714 | 0.98032 | 0.023718 | 24.1938 | -0.11337 | 0.15528 | 0.002031 | 0.14052 | 0.14052 | 0.045417
2019-03-20 17:59 | 1553119160.0812 | 620 | 0.20917 | 0.000653 | 8.1e-05 | 109.4065 | 110.8435 | 119.8989 | 21.0658 | 0.002909 | 0.002712 | 131.4719 | 0.26126 | 8.4001 | 0.003745 | 0.00506 | 0.004944 | 7.3e-05 | 0.001139 | 0.28783 | 1.3654 | 2.5482 | 100.58 | 20.75 | 19.345 | 91.6063 | 111.8628 | 86.8679 | 0.004716 | 0.98063 | 0.023719 | 24.1869 | -0.11672 | 0.15543 | 0.001876 | 0.12943 | 0.12943 | 0.049356
2019-03-20 17:59 | 1553119161.0821 | 621 | 0.20918 | 0.000653 | 8e-05 | 109.454 | 110.7953 | 119.8776 | 21.0658 | 0.002861 | 0.002728 | 131.5021 | 0.261 | 8.4021 | 0.003748 | 0.005284 | 0.005179 | 7.1e-05 | 0.001141 | 0.30038 | 1.3494 | 2.5481 | 100.58 | 20.75 | 19.332 | 90.4503 | 111.919 | 87.2505 | 0.004712 | 0.97959 | 0.023704 | 24.1976 | -0.10894 | 0.15527 | 0.001819 | 0.12531 | 0.12531 | 0.050927
2019-03-20 17:59 | 1553119162.0816 | 622 | 0.20916 | 0.000656 | 8e-05 | 109.4275 | 110.7377 | 119.7189 | 21.0666 | 0.002823 | 0.002677 | 131.505 | 0.26152 | 8.4036 | 0.003778 | 0.006139 | 0.006011 | 7.1e-05 | 0.001137 | 0.27577 | 1.3181 | 2.5481 | 100.58 | 20.75 | 19.324 | 89.5483 | 112.647 | 86.0381 | 0.00471 | 0.9829 | 0.023728 | 24.1409 | -0.10646 | 0.15558 | 0.001892 | 0.13071 | 0.13071 | 0.048921
2019-03-20 17:59 | 1553119163.0813 | 623 | 0.20916 | 0.000654 | 7.9e-05 | 109.4621 | 110.7769 | 119.5593 | 21.0665 | 0.002898 | 0.002683 | 131.5003 | 0.26094 | 8.4011 | 0.003767 | 0.005623 | 0.0056 | 6.9e-05 | 0.001132 | 0.26145 | 1.3616 | 2.5481 | 100.58 | 20.75 | 19.314 | 91.3514 | 112.399 | 86.1848 | 0.004708 | 0.9825 | 0.023707 | 24.1297 | -0.1068 | 0.15524 | 0.001888 | 0.13032 | 0.13032 | 0.048956
2019-03-20 17:59 | 1553119164.0824 | 624 | 0.20917 | 0.000651 | 7.9e-05 | 109.4175 | 110.7 | 119.3537 | 21.0658 | 0.002874 | 0.002732 | 131.4889 | 0.26082 | 8.4039 | 0.003897 | 0.005376 | 0.005224 | 7.1e-05 | 0.001131 | 0.25536 | 1.3782 | 2.5482 | 100.58 | 20.75 | 19.289 | 90.7809 | 115.5633 | 87.3466 | 0.004702 | 0.97933 | 0.023649 | 24.1479 | -0.10423 | 0.15516 | 0.001873 | 0.12894 | 0.12894 | 0.049458
2019-03-20 17:59 | 1553119165.0825 | 625 | 0.20914 | 0.000651 | 7.9e-05 | 109.4126 | 110.7575 | 119.3476 | 21.0656 | 0.002919 | 0.002717 | 131.511 | 0.26109 | 8.4037 | 0.003803 | 0.005334 | 0.005221 | 7.3e-05 | 0.001127 | 0.26145 | 1.3471 | 2.548 | 100.58 | 20.75 | 19.275 | 91.8469 | 113.274 | 86.9825 | 0.004699 | 0.98032 | 0.02366 | 24.1351 | -0.10928 | 0.15532 | 0.002014 | 0.13906 | 0.13906 | 0.045905
2019-03-20 17:59 | 1553119166.0824 | 626 | 0.2092 | 0.000656 | 7.8e-05 | 109.4074 | 110.6364 | 119.2868 | 21.0671 | 0.002818 | 0.002702 | 131.4997 | 0.26048 | 8.4042 | 0.003783 | 0.005182 | 0.005031 | 7.1e-05 | 0.001123 | 0.27312 | 1.3209 | 2.5483 | 100.58 | 20.75 | 19.244 | 89.4169 | 112.7788 | 86.6405 | 0.004691 | 0.98125 | 0.023665 | 24.1175 | -0.099914 | 0.15496 | 0.001663 | 0.11412 | 0.11412 | 0.055809
2019-03-20 17:59 | 1553119167.0818 | 627 | 0.2092 | 0.00065 | 7.7e-05 | 109.4188 | 110.5875 | 119.5628 | 21.0665 | 0.002905 | 0.002777 | 131.5004 | 0.261 | 8.4019 | 0.003768 | 0.005409 | 0.005206 | 6.8e-05 | 0.001118 | 0.25769 | 1.3172 | 2.548 | 100.58 | 20.75 | 19.228 | 91.5272 | 112.4164 | 88.4445 | 0.004687 | 0.97636 | 0.023634 | 24.2058 | -0.095029 | 0.15527 | 0.001683 | 0.1154 | 0.1154 | 0.0553
2019-03-20 17:59 | 1553119168.0816 | 628 | 0.20923 | 0.000647 | 7.8e-05 | 109.445 | 110.7852 | 119.8158 | 21.0671 | 0.002886 | 0.002678 | 131.5055 | 0.26111 | 8.4035 | 0.003764 | 0.005131 | 0.005037 | 6.9e-05 | 0.001115 | 0.28136 | 1.3129 | 2.5481 | 100.58 | 20.75 | 19.213 | 91.0712 | 112.3245 | 86.0448 | 0.004683 | 0.98288 | 0.023737 | 24.1509 | -0.10887 | 0.15533 | 0.00153 | 0.10495 | 0.10495 | 0.060831
2019-03-20 17:59 | 1553119169.0816 | 629 | 0.2092 | 0.000646 | 7.8e-05 | 109.5106 | 110.8442 | 119.8292 | 21.0667 | 0.0029 | 0.002687 | 131.4958 | 0.25962 | 8.4026 | 0.003808 | 0.004846 | 0.004718 | 6.9e-05 | 0.001116 | 0.26056 | 1.3367 | 2.5483 | 100.57 | 20.75 | 19.186 | 91.4079 | 113.374 | 86.2583 | 0.004677 | 0.9823 | 0.023732 | 24.1594 | -0.10827 | 0.15445 | 0.001692 | 0.11652 | 0.11652 | 0.054476
2019-03-20 17:59 | 1553119170.0817 | 630 | 0.20917 | 0.000645 | 7.8e-05 | 109.4614 | 110.7003 | 119.771 | 21.0669 | 0.002873 | 0.002647 | 131.4727 | 0.26072 | 8.4014 | 0.003807 | 0.004668 | 0.00457 | 7e-05 | 0.001117 | 0.27345 | 1.3402 | 2.548 | 100.57 | 20.75 | 19.174 | 90.7412 | 113.3591 | 85.3039 | 0.004674 | 0.98491 | 0.023758 | 24.1215 | -0.10067 | 0.1551 | 0.001866 | 0.12907 | 0.12907 | 0.049391
2019-03-20 18:00 | 1553119171.0831 | 631 | 0.2092 | 0.000641 | 7.8e-05 | 109.5135 | 110.7683 | 119.8043 | 21.0666 | 0.002864 | 0.002777 | 131.4625 | 0.26072 | 8.4049 | 0.003859 | 0.00555 | 0.005313 | 6.7e-05 | 0.001112 | 0.26609 | 1.3524 | 2.5482 | 100.57 | 20.75 | 19.163 | 90.519 | 114.6213 | 88.4406 | 0.004672 | 0.97637 | 0.023658 | 24.2302 | -0.10191 | 0.1551 | 0.00167 | 0.11455 | 0.11455 | 0.055651
2019-03-20 18:00 | 1553119172.0818 | 632 | 0.20921 | 0.000643 | 7.8e-05 | 109.4259 | 110.7601 | 119.81 | 21.0671 | 0.002865 | 0.002712 | 131.4596 | 0.26035 | 8.4053 | 0.003906 | 0.005895 | 0.00585 | 6.7e-05 | 0.001116 | 0.25857 | 1.3902 | 2.5482 | 100.57 | 20.75 | 19.142 | 90.5445 | 115.7672 | 86.8732 | 0.004667 | 0.98062 | 0.02371 | 24.1781 | -0.1084 | 0.15488 | 0.001606 | 0.11024 | 0.11024 | 0.057744
2019-03-20 18:00 | 1553119173.136 | 633 | 0.20919 | 0.000643 | 7.7e-05 | 109.5268 | 110.8872 | 119.8116 | 21.0667 | 0.002797 | 0.002679 | 131.4427 | 0.26089 | 8.4035 | 0.003888 | 0.005236 | 0.005158 | 6.9e-05 | 0.001115 | 0.26344 | 1.3854 | 2.5482 | 100.57 | 20.75 | 19.134 | 88.9111 | 115.3442 | 86.082 | 0.004665 | 0.98278 | 0.023736 | 24.1517 | -0.11041 | 0.1552 | 0.001737 | 0.11975 | 0.11975 | 0.053267
2019-03-20 18:00 | 1553119174.089 | 634 | 0.20921 | 0.000649 | 7.8e-05 | 109.4388 | 110.8379 | 120.1454 | 21.0667 | 0.002862 | 0.002717 | 131.4328 | 0.26084 | 8.4018 | 0.003848 | 0.005077 | 0.004872 | 6.9e-05 | 0.001113 | 0.29087 | 1.3671 | 2.5481 | 100.57 | 20.75 | 19.121 | 90.4903 | 114.3587 | 87 | 0.004661 | 0.98027 | 0.023739 | 24.2162 | -0.11362 | 0.15517 | 0.001644 | 0.11307 | 0.11307 | 0.056404
2019-03-20 18:00 | 1553119175.089 | 635 | 0.20918 | 0.000641 | 7.7e-05 | 109.4831 | 110.756 | 120.3257 | 21.066 | 0.002812 | 0.002717 | 131.4121 | 0.2611 | 8.4024 | 0.003858 | 0.005143 | 0.005111 | 6.6e-05 | 0.001111 | 0.26648 | 1.3844 | 2.5481 | 100.57 | 20.75 | 19.116 | 89.2784 | 114.6056 | 86.9908 | 0.00466 | 0.9803 | 0.023757 | 24.2341 | -0.10339 | 0.15532 | 0.001795 | 0.12406 | 0.12406 | 0.051459
2019-03-20 18:00 | 1553119176.0884 | 636 | 0.20921 | 0.000643 | 7.7e-05 | 109.5213 | 110.771 | 120.5402 | 21.0667 | 0.002865 | 0.002676 | 131.3889 | 0.2607 | 8.4022 | 0.003887 | 0.005153 | 0.00499 | 6.4e-05 | 0.00111 | 0.26609 | 1.3888 | 2.5481 | 100.57 | 20.75 | 19.107 | 90.5637 | 115.3125 | 85.9954 | 0.004658 | 0.98301 | 0.023811 | 24.2221 | -0.10148 | 0.15509 | 0.001664 | 0.11492 | 0.11492 | 0.055469
2019-03-20 18:00 | 1553119177.0894 | 637 | 0.20914 | 0.000643 | 7.6e-05 | 109.487 | 111.0606 | 120.2146 | 21.0657 | 0.002877 | 0.002699 | 131.3933 | 0.2611 | 8.4058 | 0.00393 | 0.005637 | 0.005519 | 6.7e-05 | 0.001111 | 0.28208 | 1.3598 | 2.5482 | 100.57 | 20.75 | 19.105 | 90.8525 | 116.3681 | 86.5516 | 0.004658 | 0.9815 | 0.02376 | 24.2081 | -0.12764 | 0.15532 | 0.002035 | 0.14136 | 0.14136 | 0.04516
2019-03-20 18:00 | 1553119178.0892 | 638 | 0.20918 | 0.000645 | 7.6e-05 | 109.5028 | 110.9803 | 119.8987 | 21.0663 | 0.002868 | 0.002653 | 131.3847 | 0.2611 | 8.4036 | 0.003915 | 0.006023 | 0.005766 | 6.7e-05 | 0.001114 | 0.28102 | 1.358 | 2.5481 | 100.57 | 20.75 | 19.102 | 90.6296 | 115.9959 | 85.4612 | 0.004657 | 0.98448 | 0.023765 | 24.1396 | -0.11987 | 0.15533 | 0.001834 | 0.12694 | 0.12694 | 0.05029
2019-03-20 18:00 | 1553119179.0882 | 639 | 0.20922 | 0.000644 | 7.6e-05 | 109.564 | 110.7224 | 120.0264 | 21.0669 | 0.002815 | 0.002693 | 131.3986 | 0.26015 | 8.4006 | 0.003894 | 0.005316 | 0.005223 | 7e-05 | 0.001119 | 0.29645 | 1.3583 | 2.5481 | 100.57 | 20.75 | 19.1 | 89.3409 | 115.4741 | 86.4195 | 0.004656 | 0.98186 | 0.023746 | 24.1847 | -0.094073 | 0.15476 | 0.001605 | 0.11048 | 0.11048 | 0.057573
2019-03-20 18:00 | 1553119180.0892 | 640 | 0.20921 | 0.000644 | 7.5e-05 | 109.5329 | 111.0887 | 120.0834 | 21.0671 | 0.002833 | 0.002704 | 131.4217 | 0.26097 | 8.3994 | 0.003892 | 0.005073 | 0.004972 | 6.9e-05 | 0.001111 | 0.26128 | 1.3513 | 2.5481 | 100.57 | 20.75 | 19.101 | 89.7863 | 115.4253 | 86.6733 | 0.004657 | 0.98116 | 0.023743 | 24.199 | -0.12615 | 0.15525 | 0.00165 | 0.11371 | 0.11371 | 0.056112
2019-03-20 18:00 | 1553119181.0883 | 641 | 0.20923 | 0.000642 | 7.5e-05 | 109.5446 | 111.1866 | 120.2365 | 21.0671 | 0.002827 | 0.002677 | 131.3944 | 0.26125 | 8.4037 | 0.003894 | 0.005348 | 0.005152 | 6.7e-05 | 0.001112 | 0.28799 | 1.342 | 2.548 | 100.57 | 20.75 | 19.102 | 89.6317 | 115.4792 | 86.0285 | 0.004657 | 0.98293 | 0.02378 | 24.1927 | -0.13308 | 0.15542 | 0.001509 | 0.10377 | 0.10377 | 0.061555
2019-03-20 18:00 | 1553119182.0884 | 642 | 0.2092 | 0.00064 | 7.5e-05 | 109.5898 | 111.1144 | 120.1653 | 21.0669 | 0.002836 | 0.002682 | 131.4055 | 0.26121 | 8.402 | 0.003836 | 0.004859 | 0.004683 | 6.7e-05 | 0.001112 | 0.25105 | 1.3422 | 2.5482 | 100.57 | 20.75 | 19.109 | 89.8676 | 114.0578 | 86.1407 | 0.004659 | 0.98262 | 0.023769 | 24.1893 | -0.12358 | 0.1554 | 0.00171 | 0.11813 | 0.11813 | 0.054065
2019-03-20 18:00 | 1553119183.0884 | 643 | 0.20911 | 0.000638 | 7.5e-05 | 109.5779 | 111.1686 | 120.2311 | 21.0673 | 0.002836 | 0.002668 | 131.4086 | 0.26165 | 8.4031 | 0.003827 | 0.005453 | 0.005385 | 6.9e-05 | 0.001109 | 0.27688 | 1.3493 | 2.5479 | 100.57 | 20.75 | 19.115 | 89.8569 | 113.8574 | 85.8204 | 0.00466 | 0.9835 | 0.023786 | 24.1852 | -0.12892 | 0.15566 | 0.002254 | 0.15728 | 0.15728 | 0.040676
2019-03-20 18:00 | 1553119184.0895 | 644 | 0.20918 | 0.000637 | 7.5e-05 | 109.5543 | 110.8371 | 120.2285 | 21.067 | 0.002853 | 0.002642 | 131.3957 | 0.26061 | 8.4029 | 0.003867 | 0.006059 | 0.005965 | 7.1e-05 | 0.001102 | 0.24851 | 1.3082 | 2.548 | 100.57 | 20.75 | 19.131 | 90.2569 | 114.8248 | 85.1879 | 0.004664 | 0.98523 | 0.023807 | 24.1636 | -0.10413 | 0.15504 | 0.001839 | 0.12764 | 0.12764 | 0.049921
2019-03-20 18:00 | 1553119185.0885 | 645 | 0.20921 | 0.000647 | 7.4e-05 | 109.5905 | 110.9186 | 120.1257 | 21.0665 | 0.002823 | 0.002692 | 131.4044 | 0.25985 | 8.4041 | 0.003885 | 0.005202 | 0.00507 | 6.9e-05 | 0.001106 | 0.27743 | 1.299 | 2.5481 | 100.57 | 20.74 | 19.141 | 89.5357 | 115.2506 | 86.3854 | 0.004663 | 0.98195 | 0.023757 | 24.1936 | -0.10774 | 0.15459 | 0.001665 | 0.11491 | 0.11491 | 0.05529
2019-03-20 18:00 | 1553119186.0899 | 646 | 0.20915 | 0.00064 | 7.4e-05 | 109.6105 | 111.0438 | 119.979 | 21.0666 | 0.00283 | 0.002657 | 131.4355 | 0.26126 | 8.4016 | 0.003784 | 0.004901 | 0.004692 | 6.9e-05 | 0.001109 | 0.28313 | 1.3184 | 2.5481 | 100.57 | 20.75 | 19.151 | 89.7047 | 112.7924 | 85.5411 | 0.004669 | 0.98426 | 0.02377 | 24.1504 | -0.1162 | 0.15543 | 0.00203 | 0.14116 | 0.14116 | 0.045255
2019-03-20 18:00 | 1553119187.09 | 647 | 0.20919 | 0.000635 | 7.4e-05 | 109.6128 | 110.9806 | 119.8853 | 21.0667 | 0.002784 | 0.002685 | 131.4157 | 0.26172 | 8.4025 | 0.00384 | 0.00488 | 0.004756 | 6.7e-05 | 0.001108 | 0.29856 | 1.3045 | 2.5481 | 100.57 | 20.75 | 19.175 | 88.608 | 114.1674 | 86.2249 | 0.004675 | 0.98239 | 0.023738 | 24.164 | -0.11093 | 0.1557 | 0.001787 | 0.12359 | 0.12359 | 0.051776
2019-03-20 18:00 | 1553119188.0887 | 648 | 0.2092 | 0.000635 | 7.4e-05 | 109.6287 | 110.9649 | 119.7081 | 21.0677 | 0.002843 | 0.002685 | 131.4267 | 0.26084 | 8.4017 | 0.003751 | 0.004819 | 0.004687 | 6.7e-05 | 0.00111 | 0.26128 | 1.3146 | 2.5481 | 100.57 | 20.74 | 19.189 | 90.023 | 111.9983 | 86.2127 | 0.004675 | 0.98242 | 0.023721 | 24.1457 | -0.10837 | 0.15517 | 0.001727 | 0.11919 | 0.11919 | 0.053507
2019-03-20 18:00 | 1553119189.0891 | 649 | 0.20919 | 0.000633 | 7.4e-05 | 109.6675 | 111.0524 | 119.6832 | 21.0661 | 0.002857 | 0.002666 | 131.4173 | 0.26092 | 8.403 | 0.00379 | 0.005223 | 0.00509 | 6.5e-05 | 0.001112 | 0.26593 | 1.3698 | 2.548 | 100.57 | 20.74 | 19.203 | 90.3574 | 112.9507 | 85.7653 | 0.004679 | 0.98364 | 0.023734 | 24.1281 | -0.11225 | 0.15522 | 0.00174 | 0.12021 | 0.12021 | 0.053072
2019-03-20 18:00 | 1553119190.0881 | 650 | 0.20928 | 0.000636 | 7.5e-05 | 109.7451 | 111.1982 | 119.6058 | 21.0668 | 0.002789 | 0.002646 | 131.3961 | 0.2613 | 8.403 | 0.00376 | 0.00515 | 0.005108 | 6.8e-05 | 0.001112 | 0.29911 | 1.3482 | 2.548 | 100.57 | 20.74 | 19.236 | 88.7338 | 112.2099 | 85.2907 | 0.004687 | 0.98495 | 0.023742 | 24.1044 | -0.11765 | 0.15544 | 0.001195 | 0.081114 | 0.081114 | 0.078763
2019-03-20 18:00 | 1553119191.0889 | 651 | 0.20922 | 0.000633 | 7.5e-05 | 109.6869 | 111.055 | 119.4959 | 21.0663 | 0.002831 | 0.002626 | 131.4246 | 0.26079 | 8.4019 | 0.003764 | 0.005443 | 0.005296 | 6.5e-05 | 0.001114 | 0.28102 | 1.3294 | 2.5481 | 100.57 | 20.74 | 19.253 | 89.7333 | 112.3128 | 84.7961 | 0.004691 | 0.98631 | 0.023747 | 24.0767 | -0.11088 | 0.15514 | 0.001588 | 0.10934 | 0.10934 | 0.058318
2019-03-20 18:00 | 1553119192.088 | 652 | 0.2092 | 0.000631 | 7.5e-05 | 109.6591 | 111.0001 | 119.576 | 21.0666 | 0.00274 | 0.002666 | 131.4121 | 0.2613 | 8.403 | 0.00369 | 0.005055 | 0.004923 | 6.2e-05 | 0.001109 | 0.29878 | 1.3187 | 2.548 | 100.57 | 20.74 | 19.289 | 87.555 | 110.5119 | 85.7595 | 0.0047 | 0.98366 | 0.023723 | 24.1171 | -0.10872 | 0.15544 | 0.001704 | 0.11745 | 0.11745 | 0.054398
2019-03-20 18:00 | 1553119193.089 | 653 | 0.20911 | 0.000632 | 7.5e-05 | 109.7609 | 111.2725 | 119.4386 | 21.0665 | 0.002863 | 0.002635 | 131.3989 | 0.26087 | 8.4032 | 0.003765 | 0.00523 | 0.005133 | 6.5e-05 | 0.00111 | 0.28871 | 1.3265 | 2.548 | 100.57 | 20.74 | 19.305 | 90.4977 | 112.3391 | 85.0094 | 0.004703 | 0.98572 | 0.023734 | 24.0781 | -0.12234 | 0.15519 | 0.002271 | 0.15809 | 0.15809 | 0.040346
2019-03-20 18:00 | 1553119194.0879 | 654 | 0.20923 | 0.000634 | 7.5e-05 | 109.7089 | 111.0987 | 119.4354 | 21.0653 | 0.002822 | 0.002656 | 131.4045 | 0.26078 | 8.4026 | 0.003752 | 0.005575 | 0.005424 | 6.4e-05 | 0.001109 | 0.27201 | 1.3257 | 2.548 | 100.57 | 20.75 | 19.335 | 89.5182 | 112.0227 | 85.5148 | 0.004714 | 0.98433 | 0.023717 | 24.0947 | -0.1126 | 0.15514 | 0.001507 | 0.10336 | 0.10336 | 0.061688
2019-03-20 18:00 | 1553119195.0884 | 655 | 0.2093 | 0.000636 | 7.4e-05 | 109.7205 | 110.8252 | 119.3938 | 21.0662 | 0.002813 | 0.002702 | 131.4141 | 0.26088 | 8.4051 | 0.003692 | 0.005743 | 0.005576 | 6.5e-05 | 0.00111 | 0.27168 | 1.3439 | 2.548 | 100.57 | 20.75 | 19.348 | 89.3057 | 110.558 | 86.6397 | 0.004717 | 0.98125 | 0.023676 | 24.1283 | -0.089602 | 0.1552 | 0.001115 | 0.07525 | 0.07525 | 0.084766
2019-03-20 18:00 | 1553119196.0893 | 656 | 0.20919 | 0.000632 | 7.4e-05 | 109.7493 | 111.1466 | 119.46 | 21.0659 | 0.002822 | 0.002671 | 131.3895 | 0.26118 | 8.4003 | 0.00372 | 0.005074 | 0.004959 | 6.3e-05 | 0.001113 | 0.27776 | 1.3376 | 2.5481 | 100.57 | 20.75 | 19.359 | 89.525 | 111.2488 | 85.8774 | 0.004719 | 0.98334 | 0.023708 | 24.1094 | -0.11316 | 0.15538 | 0.001743 | 0.12026 | 0.12026 | 0.053099
2019-03-20 18:00 | 1553119197.0896 | 657 | 0.20922 | 0.000632 | 7.4e-05 | 109.7564 | 111.1331 | 119.6724 | 21.0662 | 0.002779 | 0.002622 | 131.3833 | 0.26119 | 8.4017 | 0.003635 | 0.005256 | 0.005077 | 6.3e-05 | 0.001111 | 0.26017 | 1.3316 | 2.5481 | 100.57 | 20.75 | 19.374 | 88.4802 | 109.174 | 84.7119 | 0.004723 | 0.98654 | 0.023767 | 24.0916 | -0.1115 | 0.15538 | 0.001579 | 0.10878 | 0.10878 | 0.058706
2019-03-20 18:00 | 1553119198.0882 | 658 | 0.20924 | 0.000641 | 7.4e-05 | 109.7577 | 111.0502 | 119.6964 | 21.0654 | 0.002755 | 0.002649 | 131.3693 | 0.2607 | 8.403 | 0.00365 | 0.005389 | 0.005268 | 6.8e-05 | 0.00111 | 0.27455 | 1.3484 | 2.548 | 100.57 | 20.75 | 19.378 | 87.9031 | 109.5525 | 85.3509 | 0.004724 | 0.98478 | 0.023749 | 24.1155 | -0.10471 | 0.15509 | 0.001488 | 0.1022 | 0.1022 | 0.062372
2019-03-20 18:00 | 1553119199.0892 | 659 | 0.20902 | 0.000637 | 7.4e-05 | 109.706 | 111.083 | 119.2645 | 21.0656 | 0.002753 | 0.002615 | 131.3477 | 0.26025 | 8.404 | 0.003602 | 0.004872 | 0.004736 | 6.4e-05 | 0.001115 | 0.26808 | 1.3296 | 2.5481 | 100.57 | 20.75 | 19.381 | 87.8477 | 108.3766 | 84.5465 | 0.004725 | 0.987 | 0.023732 | 24.045 | -0.11158 | 0.15482 | 0.002776 | 0.19432 | 0.19432 | 0.032746
2019-03-20 18:00 | 1553119200.0882 | 660 | 0.20947 | 0.000632 | 7.3e-05 | 109.7376 | 111.1006 | 119.353 | 21.0649 | 0.002806 | 0.002617 | 131.3434 | 0.26011 | 8.4022 | 0.003631 | 0.005271 | 0.005178 | 6.3e-05 | 0.001113 | 0.29391 | 1.2803 | 2.548 | 100.57 | 20.75 | 19.384 | 89.1335 | 109.0726 | 84.5771 | 0.004726 | 0.98691 | 0.02374 | 24.0549 | -0.11041 | 0.15474 | 7.7e-05 | 0.001152 | 0.001152 | 5.5203
2019-03-20 18:00 | 1553119201.0894 | 661 | 0.20884 | 0.000635 | 7.3e-05 | 109.8589 | 111.157 | 119.4816 | 21.0658 | 0.002856 | 0.002645 | 131.3143 | 0.2615 | 8.4061 | 0.003659 | 0.005324 | 0.005173 | 6.4e-05 | 0.001111 | 0.29839 | 1.3106 | 2.548 | 100.57 | 20.75 | 19.382 | 90.3267 | 109.7665 | 85.2653 | 0.004725 | 0.98502 | 0.02373 | 24.091 | -0.10507 | 0.15557 | 0.003891 | 0.27415 | 0.27415 | 0.023323
2019-03-20 18:00 | 1553119202.0899 | 662 | 0.20934 | 0.000634 | 7.3e-05 | 109.8001 | 111.3189 | 119.8687 | 21.0652 | 0.002832 | 0.002659 | 131.3036 | 0.26014 | 8.4028 | 0.003657 | 0.005604 | 0.005375 | 6.5e-05 | 0.001113 | 0.28423 | 1.3506 | 2.5481 | 100.57 | 20.75 | 19.375 | 89.7534 | 109.7153 | 85.6027 | 0.004723 | 0.98409 | 0.023757 | 24.1414 | -0.12288 | 0.15476 | 0.000849 | 0.056486 | 0.056486 | 0.1126
2019-03-20 18:00 | 1553119203.0885 | 663 | 0.21016 | 0.000628 | 7.3e-05 | 109.8657 | 111.3028 | 119.9853 | 21.0641 | 0.002784 | 0.002631 | 131.3116 | 0.26098 | 8.4023 | 0.003747 | 0.004726 | 0.004598 | 6.2e-05 | 0.001106 | 0.26089 | 1.3297 | 2.5481 | 100.57 | 20.75 | 19.369 | 88.6093 | 111.9086 | 84.9181 | 0.004722 | 0.98597 | 0.023792 | 24.13 | -0.11624 | 0.15526 | -0.004099 | -0.29879 | -0.29879 | -0.021356
2019-03-20 18:00 | 1553119204.0885 | 664 | 0.20919 | 0.000632 | 7.3e-05 | 109.8221 | 111.2389 | 119.7489 | 21.0653 | 0.0028 | 0.00268 | 131.3187 | 0.26103 | 8.4028 | 0.00379 | 0.005508 | 0.00533 | 6.5e-05 | 0.001102 | 0.28982 | 1.321 | 2.548 | 100.57 | 20.75 | 19.362 | 88.9808 | 112.9436 | 86.1119 | 0.00472 | 0.9827 | 0.023729 | 24.1464 | -0.11466 | 0.15529 | 0.001777 | 0.12286 | 0.12286 | 0.051948
2019-03-20 18:00 | 1553119205.0885 | 665 | 0.20991 | 0.000627 | 7.3e-05 | 109.8508 | 111.5338 | 119.9973 | 21.0655 | 0.002871 | 0.002637 | 131.2895 | 0.26096 | 8.4035 | 0.00376 | 0.005121 | 0.005007 | 6.4e-05 | 0.001094 | 0.27472 | 1.336 | 2.5481 | 100.57 | 20.75 | 19.347 | 90.7016 | 112.2157 | 85.0773 | 0.004717 | 0.98554 | 0.023787 | 24.1366 | -0.13599 | 0.15524 | -0.002542 | -0.18696 | -0.18696 | -0.034128
2019-03-20 18:00 | 1553119206.0877 | 666 | 0.20955 | 0.000631 | 7.3e-05 | 109.9469 | 111.52 | 119.7968 | 21.0654 | 0.002763 | 0.002647 | 131.2967 | 0.261 | 8.3994 | 0.003661 | 0.004991 | 0.004925 | 6.6e-05 | 0.001091 | 0.26089 | 1.3371 | 2.5481 | 100.57 | 20.75 | 19.34 | 88.0942 | 109.811 | 85.3025 | 0.004715 | 0.98491 | 0.02376 | 24.124 | -0.12707 | 0.15527 | -0.000365 | -0.030524 | -0.030524 | -0.20907
2019-03-20 18:00 | 1553119207.0881 | 667 | 0.20884 | 0.000629 | 7.3e-05 | 109.933 | 111.3589 | 119.7353 | 21.0644 | 0.002752 | 0.002624 | 131.3207 | 0.2612 | 8.4029 | 0.00364 | 0.00594 | 0.005821 | 6.3e-05 | 0.001095 | 0.30342 | 1.3616 | 2.5481 | 100.57 | 20.75 | 19.33 | 87.8277 | 109.2983 | 84.7633 | 0.004712 | 0.9864 | 0.023772 | 24.0997 | -0.11527 | 0.15539 | 0.003885 | 0.27421 | 0.27421 | 0.02329
2019-03-20 18:00 | 1553119208.0879 | 668 | 0.20918 | 0.000628 | 7.2e-05 | 109.9712 | 111.2382 | 119.8628 | 21.0648 | 0.002765 | 0.002659 | 131.3201 | 0.2615 | 8.4043 | 0.003655 | 0.005653 | 0.005481 | 6.5e-05 | 0.001097 | 0.26593 | 1.3134 | 2.5481 | 100.57 | 20.75 | 19.327 | 88.1437 | 109.6654 | 85.5899 | 0.004712 | 0.98413 | 0.023757 | 24.1403 | -0.10246 | 0.15557 | 0.001858 | 0.12886 | 0.12886 | 0.049618
2019-03-20 18:00 | 1553119209.0878 | 669 | 0.2095 | 0.000632 | 7.3e-05 | 109.9809 | 111.3649 | 120.2118 | 21.0643 | 0.002782 | 0.002671 | 131.3342 | 0.26033 | 8.4016 | 0.003712 | 0.005119 | 0.005006 | 6.7e-05 | 0.001096 | 0.29175 | 1.3353 | 2.5481 | 100.57 | 20.75 | 19.326 | 88.5486 | 111.0486 | 85.8814 | 0.004711 | 0.98333 | 0.023782 | 24.1853 | -0.11186 | 0.15487 | -9.3e-05 | -0.011028 | -0.011028 | -0.57721
2019-03-20 18:00 | 1553119210.0881 | 670 | 0.20916 | 0.000635 | 7.3e-05 | 109.8871 | 111.3914 | 119.7318 | 21.0651 | 0.002761 | 0.002673 | 131.3443 | 0.26184 | 8.404 | 0.003644 | 0.004783 | 0.004744 | 6.3e-05 | 0.001097 | 0.26576 | 1.2887 | 2.5481 | 100.57 | 20.75 | 19.325 | 88.0409 | 109.3889 | 85.943 | 0.004711 | 0.98316 | 0.023732 | 24.139 | -0.12162 | 0.15577 | 0.001988 | 0.13798 | 0.13798 | 0.046399
2019-03-20 18:00 | 1553119211.0883 | 671 | 0.20911 | 0.000634 | 7.3e-05 | 109.9249 | 111.3155 | 119.7048 | 21.0655 | 0.002767 | 0.002662 | 131.3492 | 0.26066 | 8.4038 | 0.003678 | 0.004568 | 0.004399 | 6.5e-05 | 0.001096 | 0.27273 | 1.2771 | 2.548 | 100.57 | 20.75 | 19.326 | 88.1843 | 110.2322 | 85.6781 | 0.004711 | 0.98388 | 0.023739 | 24.1274 | -0.11244 | 0.15507 | 0.002294 | 0.15994 | 0.15994 | 0.039849
2019-03-20 18:00 | 1553119212.088 | 672 | 0.20928 | 0.000637 | 7.2e-05 | 109.9618 | 111.2733 | 119.4257 | 21.0653 | 0.002765 | 0.002619 | 131.3495 | 0.261 | 8.406 | 0.00362 | 0.005411 | 0.005214 | 6.5e-05 | 0.001097 | 0.28241 | 1.2714 | 2.5481 | 100.57 | 20.75 | 19.327 | 88.1522 | 108.805 | 84.6446 | 0.004712 | 0.98673 | 0.023745 | 24.0645 | -0.10605 | 0.15527 | 0.001202 | 0.081781 | 0.081781 | 0.078033
2019-03-20 18:00 | 1553119213.0882 | 673 | 0.20911 | 0.000636 | 7.2e-05 | 109.988 | 111.3589 | 119.6147 | 21.0647 | 0.002804 | 0.002642 | 131.3456 | 0.26057 | 8.4023 | 0.003647 | 0.005257 | 0.00506 | 6.5e-05 | 0.001094 | 0.27417 | 1.2855 | 2.548 | 100.57 | 20.75 | 19.329 | 89.0934 | 109.4767 | 85.1864 | 0.004712 | 0.98523 | 0.023746 | 24.1018 | -0.1108 | 0.15501 | 0.002234 | 0.15571 | 0.15571 | 0.040915
2019-03-20 18:00 | 1553119214.0881 | 674 | 0.2093 | 0.00063 | 7.2e-05 | 110.0235 | 111.4643 | 119.8607 | 21.0642 | 0.002802 | 0.002632 | 131.3317 | 0.26087 | 8.4041 | 0.003688 | 0.004968 | 0.00488 | 6.6e-05 | 0.001092 | 0.28174 | 1.3231 | 2.548 | 100.57 | 20.75 | 19.332 | 89.0313 | 110.459 | 84.944 | 0.004713 | 0.9859 | 0.023778 | 24.1184 | -0.11637 | 0.15519 | 0.00113 | 0.076711 | 0.076711 | 0.083147
2019-03-20 18:00 | 1553119215.0891 | 675 | 0.20915 | 0.000633 | 7.3e-05 | 110.0307 | 111.5809 | 119.4886 | 21.0643 | 0.002742 | 0.002629 | 131.3507 | 0.26241 | 8.4031 | 0.003588 | 0.00449 | 0.00433 | 6.5e-05 | 0.001095 | 0.29048 | 1.3608 | 2.5481 | 100.57 | 20.75 | 19.338 | 87.5872 | 108.0419 | 84.8613 | 0.004714 | 0.98613 | 0.023744 | 24.0781 | -0.12514 | 0.15611 | 0.002032 | 0.1412 | 0.1412 | 0.045438
2019-03-20 18:00 | 1553119216.0896 | 676 | 0.20924 | 0.000629 | 7.3e-05 | 110.0565 | 111.5517 | 119.6433 | 21.0652 | 0.002763 | 0.002615 | 131.3486 | 0.26191 | 8.4022 | 0.003618 | 0.005395 | 0.005204 | 6.8e-05 | 0.001097 | 0.29048 | 1.3481 | 2.5481 | 100.57 | 20.75 | 19.342 | 88.1012 | 108.7677 | 84.5454 | 0.004715 | 0.987 | 0.02377 | 24.0831 | -0.1207 | 0.15581 | 0.001492 | 0.10268 | 0.10268 | 0.062371
2019-03-20 18:00 | 1553119217.0882 | 677 | 0.2092 | 0.000631 | 7.2e-05 | 110.0569 | 111.5215 | 119.8238 | 21.0649 | 0.002763 | 0.002626 | 131.3306 | 0.26141 | 8.3993 | 0.003648 | 0.005449 | 0.005309 | 6.8e-05 | 0.001096 | 0.28136 | 1.3333 | 2.5482 | 100.57 | 20.75 | 19.349 | 88.0891 | 109.4842 | 84.7895 | 0.004717 | 0.98633 | 0.02378 | 24.1095 | -0.11825 | 0.15551 | 0.001728 | 0.11966 | 0.11966 | 0.053415
2019-03-20 18:00 | 1553119218.0883 | 678 | 0.20927 | 0.000635 | 7.2e-05 | 110.051 | 111.6029 | 119.7123 | 21.0654 | 0.002869 | 0.00266 | 131.333 | 0.26106 | 8.4023 | 0.00361 | 0.005373 | 0.00523 | 6.9e-05 | 0.001098 | 0.2896 | 1.3782 | 2.548 | 100.57 | 20.75 | 19.352 | 90.6492 | 108.5656 | 85.6078 | 0.004718 | 0.98408 | 0.023742 | 24.1258 | -0.12525 | 0.15531 | 0.001319 | 0.0902 | 0.0902 | 0.070766
2019-03-20 18:00 | 1553119219.0883 | 679 | 0.20926 | 0.000631 | 7.1e-05 | 110.0103 | 111.3839 | 119.6842 | 21.0653 | 0.002756 | 0.002653 | 131.3167 | 0.26099 | 8.4025 | 0.003658 | 0.005022 | 0.004879 | 6.9e-05 | 0.001099 | 0.30431 | 1.37 | 2.548 | 100.57 | 20.75 | 19.354 | 87.9247 | 109.7242 | 85.4405 | 0.004718 | 0.98454 | 0.023744 | 24.1173 | -0.11099 | 0.15526 | 0.001347 | 0.092195 | 0.092195 | 0.069216
2019-03-20 18:00 | 1553119220.0873 | 680 | 0.20923 | 0.000631 | 7.1e-05 | 110.0297 | 111.6515 | 120.2606 | 21.0659 | 0.002829 | 0.002679 | 131.3113 | 0.26099 | 8.4015 | 0.003613 | 0.004947 | 0.004897 | 6.4e-05 | 0.001096 | 0.25498 | 1.301 | 2.5481 | 100.57 | 20.75 | 19.358 | 89.6766 | 108.6487 | 86.0719 | 0.004719 | 0.98281 | 0.023781 | 24.1966 | -0.13088 | 0.15526 | 0.001538 | 0.10609 | 0.10609 | 0.060148
2019-03-20 18:00 | 1553119221.0878 | 681 | 0.20928 | 0.000628 | 7.2e-05 | 110.0633 | 111.4426 | 120.2334 | 21.0651 | 0.002766 | 0.002554 | 131.3198 | 0.25994 | 8.4074 | 0.003566 | 0.005511 | 0.00534 | 6.6e-05 | 0.001092 | 0.26288 | 1.3126 | 2.548 | 100.57 | 20.75 | 19.361 | 88.1673 | 107.4954 | 83.0645 | 0.00472 | 0.9911 | 0.023878 | 24.0924 | -0.11139 | 0.15464 | 0.001221 | 0.083648 | 0.083648 | 0.075982
2019-03-20 18:00 | 1553119222.0875 | 682 | 0.20924 | 0.000634 | 7.2e-05 | 110.118 | 111.2359 | 120.3799 | 21.0654 | 0.002741 | 0.002619 | 131.2907 | 0.26106 | 8.4018 | 0.003554 | 0.005008 | 0.00485 | 6.8e-05 | 0.00109 | 0.27433 | 1.3516 | 2.5483 | 100.57 | 20.75 | 19.366 | 87.5618 | 107.2081 | 84.6361 | 0.004721 | 0.98675 | 0.02384 | 24.1602 | -0.090348 | 0.15531 | 0.001473 | 0.10161 | 0.10161 | 0.062816
2019-03-20 18:00 | 1553119223.089 | 683 | 0.20917 | 0.000632 | 7.2e-05 | 110.0475 | 111.4845 | 120.6529 | 21.066 | 0.002828 | 0.002625 | 131.2604 | 0.26131 | 8.4007 | 0.003571 | 0.005287 | 0.005185 | 6.7e-05 | 0.001089 | 0.27848 | 1.3141 | 2.5481 | 100.57 | 20.75 | 19.37 | 89.6603 | 107.6255 | 84.7819 | 0.004722 | 0.98635 | 0.023862 | 24.1925 | -0.11604 | 0.15545 | 0.001907 | 0.13296 | 0.13296 | 0.048055
2019-03-20 18:00 | 1553119224.089 | 684 | 0.20932 | 0.000629 | 7.2e-05 | 110.0527 | 111.5151 | 120.3822 | 21.065 | 0.002769 | 0.002556 | 131.255 | 0.26102 | 8.403 | 0.003657 | 0.004647 | 0.004517 | 6.6e-05 | 0.001092 | 0.26056 | 1.3432 | 2.5481 | 100.57 | 20.75 | 19.374 | 88.2517 | 109.7018 | 83.127 | 0.004723 | 0.99093 | 0.023891 | 24.1094 | -0.11808 | 0.15528 | 0.001009 | 0.068376 | 0.068376 | 0.093336
2019-03-20 18:00 | 1553119225.0875 | 685 | 0.20918 | 0.000624 | 7.1e-05 | 110.0766 | 111.644 | 120.5404 | 21.0658 | 0.002746 | 0.002643 | 131.2258 | 0.26092 | 8.402 | 0.00367 | 0.005286 | 0.005059 | 6.5e-05 | 0.001093 | 0.27433 | 1.3081 | 2.5482 | 100.57 | 20.75 | 19.383 | 87.6895 | 110.0377 | 85.2006 | 0.004725 | 0.9852 | 0.023837 | 24.1953 | -0.12647 | 0.15522 | 0.001859 | 0.12944 | 0.12944 | 0.049284
2019-03-20 18:00 | 1553119226.0877 | 686 | 0.20925 | 0.000629 | 7.1e-05 | 110.0387 | 111.5685 | 120.5491 | 21.0641 | 0.002789 | 0.002636 | 131.2262 | 0.26048 | 8.4035 | 0.003672 | 0.005932 | 0.005813 | 6.7e-05 | 0.001092 | 0.31006 | 1.3183 | 2.5481 | 100.57 | 20.75 | 19.389 | 88.7241 | 110.0779 | 85.0352 | 0.004727 | 0.98565 | 0.023844 | 24.1906 | -0.12349 | 0.15496 | 0.001439 | 0.099287 | 0.099287 | 0.064145
2019-03-20 18:00 | 1553119227.0887 | 687 | 0.20926 | 0.000627 | 7.1e-05 | 110.0336 | 111.5327 | 120.5803 | 21.0642 | 0.002739 | 0.002574 | 131.2412 | 0.2611 | 8.4028 | 0.003654 | 0.005486 | 0.0054 | 6.6e-05 | 0.001097 | 0.30685 | 1.3347 | 2.548 | 100.57 | 20.75 | 19.394 | 87.5137 | 109.6473 | 83.559 | 0.004728 | 0.98973 | 0.023896 | 24.1439 | -0.12104 | 0.15532 | 0.001376 | 0.094949 | 0.094949 | 0.067235
2019-03-20 18:00 | 1553119228.0878 | 688 | 0.20926 | 0.000627 | 7.1e-05 | 109.9883 | 111.348 | 120.7543 | 21.0645 | 0.002732 | 0.002623 | 131.2185 | 0.26118 | 8.4033 | 0.003643 | 0.004784 | 0.0046 | 6.4e-05 | 0.001098 | 0.26897 | 1.3401 | 2.5482 | 100.57 | 20.75 | 19.406 | 87.3491 | 109.3788 | 84.7383 | 0.004731 | 0.98647 | 0.023874 | 24.2012 | -0.1099 | 0.15538 | 0.001346 | 0.092704 | 0.092704 | 0.068885
2019-03-20 18:00 | 1553119229.0883 | 689 | 0.20924 | 0.000633 | 7.1e-05 | 110.0349 | 111.5656 | 120.7449 | 21.0649 | 0.002754 | 0.002601 | 131.2298 | 0.26118 | 8.4024 | 0.003652 | 0.005276 | 0.005133 | 6.7e-05 | 0.0011 | 0.26305 | 1.36 | 2.548 | 100.57 | 20.75 | 19.411 | 87.8916 | 109.5938 | 84.1909 | 0.004732 | 0.98798 | 0.023891 | 24.1817 | -0.12357 | 0.15538 | 0.00146 | 0.10096 | 0.10096 | 0.063251
2019-03-20 18:00 | 1553119230.089 | 690 | 0.20926 | 0.000628 | 7.1e-05 | 109.9511 | 111.3899 | 120.4799 | 21.0653 | 0.002781 | 0.002594 | 131.2298 | 0.26101 | 8.4023 | 0.003622 | 0.005497 | 0.005423 | 6.7e-05 | 0.001104 | 0.28456 | 1.3369 | 2.548 | 100.57 | 20.75 | 19.417 | 88.5284 | 108.8711 | 84.0314 | 0.004734 | 0.98842 | 0.02387 | 24.1498 | -0.11629 | 0.15527 | 0.001347 | 0.092787 | 0.092787 | 0.068779
2019-03-20 18:00 | 1553119231.0879 | 691 | 0.2093 | 0.000632 | 7.1e-05 | 110.0572 | 111.4919 | 120.3848 | 21.065 | 0.00279 | 0.002634 | 131.2298 | 0.26134 | 8.401 | 0.003584 | 0.005102 | 0.004936 | 6.6e-05 | 0.001103 | 0.29717 | 1.302 | 2.548 | 100.57 | 20.76 | 19.421 | 88.7484 | 107.9411 | 85.0002 | 0.004738 | 0.98575 | 0.023828 | 24.173 | -0.11585 | 0.15547 | 0.001156 | 0.078876 | 0.078876 | 0.081013
2019-03-20 18:01 | 1553119232.089 | 692 | 0.20925 | 0.000631 | 7.1e-05 | 109.9537 | 111.3862 | 120.3405 | 21.0642 | 0.002703 | 0.002573 | 131.2519 | 0.26147 | 8.4026 | 0.003587 | 0.004757 | 0.004585 | 6.9e-05 | 0.0011 | 0.29214 | 1.2938 | 2.548 | 100.57 | 20.75 | 19.424 | 86.6541 | 108.0207 | 83.5155 | 0.004735 | 0.98985 | 0.023874 | 24.1184 | -0.11578 | 0.15555 | 0.001444 | 0.099738 | 0.099738 | 0.064099
2019-03-20 18:01 | 1553119233.0882 | 693 | 0.20922 | 0.000633 | 7.1e-05 | 109.9805 | 111.4871 | 119.9841 | 21.0648 | 0.002782 | 0.002595 | 131.2908 | 0.26096 | 8.4046 | 0.003593 | 0.005223 | 0.005087 | 6.8e-05 | 0.001102 | 0.26233 | 1.3138 | 2.5481 | 100.57 | 20.76 | 19.426 | 88.556 | 108.1434 | 84.0472 | 0.004739 | 0.98838 | 0.02382 | 24.1006 | -0.12169 | 0.15524 | 0.001629 | 0.11281 | 0.11281 | 0.056561
2019-03-20 18:01 | 1553119234.0895 | 694 | 0.2093 | 0.000622 | 7.1e-05 | 109.9511 | 111.4411 | 120.097 | 21.0647 | 0.00275 | 0.002611 | 131.291 | 0.26139 | 8.4015 | 0.003568 | 0.005317 | 0.005218 | 6.7e-05 | 0.001101 | 0.26288 | 1.3032 | 2.5479 | 100.57 | 20.76 | 19.427 | 87.7907 | 107.5555 | 84.4451 | 0.004739 | 0.98728 | 0.023818 | 24.1253 | -0.1204 | 0.1555 | 0.001113 | 0.075708 | 0.075708 | 0.084417
2019-03-20 18:01 | 1553119235.0893 | 695 | 0.20923 | 0.000621 | 7.1e-05 | 109.932 | 111.385 | 120.12 | 21.0655 | 0.002825 | 0.002659 | 131.3144 | 0.2618 | 8.4032 | 0.003585 | 0.005144 | 0.004982 | 7e-05 | 0.001106 | 0.28296 | 1.2996 | 2.5481 | 100.57 | 20.76 | 19.423 | 89.5818 | 107.9629 | 85.5891 | 0.004738 | 0.98413 | 0.023783 | 24.1662 | -0.11745 | 0.15575 | 0.001571 | 0.10848 | 0.10848 | 0.059007
2019-03-20 18:01 | 1553119236.0883 | 696 | 0.20929 | 0.000629 | 7.1e-05 | 109.8929 | 111.4045 | 119.9823 | 21.0658 | 0.00274 | 0.002599 | 131.3104 | 0.26142 | 8.4045 | 0.003622 | 0.00516 | 0.005083 | 6.6e-05 | 0.001104 | 0.28479 | 1.3022 | 2.5481 | 100.57 | 20.76 | 19.418 | 87.5458 | 108.8677 | 84.1511 | 0.004737 | 0.98809 | 0.023817 | 24.1039 | -0.12219 | 0.15552 | 0.001212 | 0.082833 | 0.082833 | 0.077165
2019-03-20 18:01 | 1553119237.0877 | 697 | 0.20924 | 0.000623 | 7e-05 | 110.0028 | 111.4587 | 119.8043 | 21.0655 | 0.002776 | 0.002597 | 131.286 | 0.26162 | 8.3997 | 0.003537 | 0.005223 | 0.005084 | 6.6e-05 | 0.001104 | 0.31365 | 1.2862 | 2.5481 | 100.57 | 20.76 | 19.411 | 88.4036 | 106.7925 | 84.1095 | 0.004735 | 0.9882 | 0.023801 | 24.0846 | -0.1176 | 0.15564 | 0.001529 | 0.10562 | 0.10562 | 0.060564
2019-03-20 18:01 | 1553119238.0875 | 698 | 0.20927 | 0.000624 | 7e-05 | 109.8576 | 111.4213 | 119.8917 | 21.0653 | 0.002734 | 0.002586 | 131.2855 | 0.26085 | 8.4025 | 0.003543 | 0.005392 | 0.005295 | 6.4e-05 | 0.001104 | 0.30613 | 1.2827 | 2.5479 | 100.57 | 20.76 | 19.396 | 87.3929 | 106.9442 | 83.8377 | 0.004731 | 0.98896 | 0.023818 | 24.0842 | -0.12641 | 0.15518 | 0.001305 | 0.089601 | 0.089601 | 0.07118
2019-03-20 18:01 | 1553119239.0885 | 699 | 0.20926 | 0.00063 | 7e-05 | 109.9026 | 111.3252 | 120.0114 | 21.0659 | 0.002721 | 0.002562 | 131.2718 | 0.26147 | 8.402 | 0.00355 | 0.005037 | 0.004909 | 6.4e-05 | 0.001108 | 0.26913 | 1.3258 | 2.548 | 100.57 | 20.76 | 19.387 | 87.0993 | 107.1079 | 83.2704 | 0.004729 | 0.99053 | 0.023849 | 24.0771 | -0.11504 | 0.15555 | 0.001375 | 0.094713 | 0.094713 | 0.067499
2019-03-20 18:01 | 1553119240.0876 | 700 | 0.20929 | 0.000632 | 7e-05 | 109.8848 | 111.3215 | 120.0827 | 21.0657 | 0.002818 | 0.002596 | 131.2568 | 0.26139 | 8.4045 | 0.003619 | 0.004957 | 0.004847 | 6.2e-05 | 0.001106 | 0.29695 | 1.3151 | 2.548 | 100.57 | 20.76 | 19.379 | 89.4303 | 108.7879 | 84.0715 | 0.004727 | 0.98831 | 0.023829 | 24.1113 | -0.11619 | 0.1555 | 0.00118 | 0.08064 | 0.08064 | 0.079255
2019-03-20 18:01 | 1553119241.0885 | 701 | 0.20924 | 0.000627 | 7e-05 | 109.8288 | 111.2202 | 120.4902 | 21.0656 | 0.00276 | 0.002631 | 131.2417 | 0.26043 | 8.4011 | 0.003635 | 0.005584 | 0.005402 | 6.9e-05 | 0.001108 | 0.31686 | 1.3025 | 2.548 | 100.57 | 20.76 | 19.364 | 88.0314 | 109.1758 | 84.9213 | 0.004724 | 0.98596 | 0.023841 | 24.1809 | -0.1126 | 0.15493 | 0.001514 | 0.10475 | 0.10475 | 0.060788
2019-03-20 18:01 | 1553119242.0886 | 702 | 0.20921 | 0.000626 | 7e-05 | 109.7871 | 111.237 | 120.3074 | 21.0647 | 0.002758 | 0.002595 | 131.2419 | 0.26055 | 8.4034 | 0.0036 | 0.005524 | 0.005365 | 6.8e-05 | 0.001104 | 0.30182 | 1.2763 | 2.548 | 100.57 | 20.76 | 19.359 | 87.978 | 108.3269 | 84.0589 | 0.004722 | 0.98834 | 0.023852 | 24.1334 | -0.11735 | 0.155 | 0.001664 | 0.11552 | 0.11552 | 0.055148
2019-03-20 18:01 | 1553119243.0877 | 703 | 0.20926 | 0.000626 | 6.9e-05 | 109.7777 | 111.089 | 120.274 | 21.0658 | 0.002738 | 0.002586 | 131.2246 | 0.26133 | 8.4011 | 0.003624 | 0.004447 | 0.004356 | 6.7e-05 | 0.001098 | 0.28766 | 1.3345 | 2.5481 | 100.57 | 20.76 | 19.346 | 87.4965 | 108.907 | 83.8375 | 0.004719 | 0.98896 | 0.023856 | 24.1226 | -0.10621 | 0.15546 | 0.001404 | 0.096919 | 0.096919 | 0.065927
2019-03-20 18:01 | 1553119244.0887 | 704 | 0.20918 | 0.000626 | 6.9e-05 | 109.78 | 111.2269 | 120.5019 | 21.065 | 0.002792 | 0.002556 | 131.2313 | 0.26206 | 8.4 | 0.00357 | 0.004866 | 0.004801 | 6.7e-05 | 0.001098 | 0.28064 | 1.3281 | 2.5481 | 100.57 | 20.76 | 19.339 | 88.7891 | 107.607 | 83.1075 | 0.004718 | 0.99098 | 0.023903 | 24.1207 | -0.11712 | 0.1559 | 0.001887 | 0.13194 | 0.13194 | 0.048563
2019-03-20 18:01 | 1553119245.0888 | 705 | 0.20926 | 0.000622 | 6.9e-05 | 109.7972 | 111.3391 | 120.4646 | 21.0661 | 0.002785 | 0.002558 | 131.3462 | 0.2624 | 8.4013 | 0.003611 | 0.004323 | 0.004177 | 6.6e-05 | 0.001095 | 0.26681 | 1.3361 | 2.5481 | 100.57 | 20.76 | 19.326 | 88.6199 | 108.6025 | 83.1552 | 0.004714 | 0.99085 | 0.023898 | 24.1186 | -0.12474 | 0.1561 | 0.00137 | 0.094639 | 0.094639 | 0.067793
2019-03-20 18:01 | 1553119246.0878 | 706 | 0.20929 | 0.000624 | 7e-05 | 109.7587 | 111.2404 | 120.8122 | 21.0659 | 0.002697 | 0.002574 | 105.5458 | 0.26093 | 8.4022 | 0.003582 | 0.004981 | 0.004901 | 6.7e-05 | 0.001097 | 0.28407 | 1.365 | 2.5481 | 100.57 | 20.76 | 19.318 | 86.516 | 107.8852 | 83.5405 | 0.004712 | 0.98978 | 0.023919 | 24.1664 | -0.11994 | 0.15523 | 0.001172 | 0.080371 | 0.080371 | 0.079379
2019-03-20 18:01 | 1553119247.0891 | 707 | 0.20924 | 0.00062 | 7e-05 | 109.66 | 111.2804 | 121.3958 | 21.0663 | 0.002761 | 0.002561 | 77.2113 | 0.26147 | 8.4036 | 0.003648 | 0.005633 | 0.00548 | 6.6e-05 | 0.001095 | 0.29695 | 1.3704 | 2.5481 | 100.57 | 20.76 | 19.31 | 88.0515 | 109.4846 | 83.2319 | 0.00471 | 0.99064 | 0.023988 | 24.2142 | -0.1312 | 0.15555 | 0.00152 | 0.10579 | 0.10579 | 0.06043
2019-03-20 18:01 | 1553119248.0879 | 708 | 0.20928 | 0.00062 | 7e-05 | 109.726 | 111.3447 | 121.7361 | 21.0654 | 0.002741 | 0.002567 | 48.9061 | 0.26073 | 8.4027 | 0.00363 | 0.005349 | 0.005207 | 6.6e-05 | 0.001095 | 0.28241 | 1.3249 | 2.548 | 100.57 | 20.76 | 19.297 | 87.5664 | 109.0485 | 83.3919 | 0.004707 | 0.99019 | 0.024016 | 24.2536 | -0.13098 | 0.15511 | 0.001279 | 0.088454 | 0.088454 | 0.072071
2019-03-20 18:01 | 1553119249.089 | 709 | 0.20929 | 0.000623 | 6.9e-05 | 109.6028 | 110.9107 | 121.8814 | 21.0659 | 0.002744 | 0.002565 | 20.5795 | 0.26119 | 8.4026 | 0.003582 | 0.004048 | 0.003928 | 6.3e-05 | 0.001097 | 0.27505 | 1.3175 | 2.548 | 100.57 | 20.76 | 19.289 | 87.6419 | 107.8939 | 83.3454 | 0.004705 | 0.99032 | 0.024032 | 24.2665 | -0.10611 | 0.15538 | 0.001224 | 0.084568 | 0.084568 | 0.075516
2019-03-20 18:01 | 1553119250.0889 | 710 | 0.20927 | 0.000627 | 6.7e-05 | 109.5947 | 110.8633 | 122.3654 | 21.066 | 0.002663 | 0.002544 | -7.7365 | 0.26151 | 8.3997 | 0.003567 | 0.005159 | 0.005039 | 6.3e-05 | 0.001099 | 0.28584 | 1.3202 | 2.548 | 100.57 | 20.76 | 19.282 | 85.6817 | 107.53 | 82.8245 | 0.004704 | 0.99177 | 0.024097 | 24.2969 | -0.10294 | 0.15557 | 0.001317 | 0.091732 | 0.091732 | 0.069703
2019-03-20 18:01 | 1553119251.087 | 711 | 0.20924 | 0.000622 | 6.2e-05 | 109.5801 | 110.9014 | 122.4002 | 21.0659 | 0.002715 | 0.002517 | -36.0542 | 0.24112 | 8.4012 | 0.003662 | 0.004606 | 0.004403 | 6.2e-05 | 0.001096 | 0.28727 | 1.29 | 2.548 | 100.57 | 20.76 | 19.267 | 86.9376 | 109.8432 | 82.1778 | 0.0047 | 0.99358 | 0.024122 | 24.2782 | -0.1072 | 0.14344 | 0.001506 | 0.10587 | 0.10587 | 0.055686
2019-03-20 18:01 | 1553119252.0882 | 712 | 0.2093 | 0.000624 | 5.7e-05 | 109.5284 | 111.0397 | 122.9172 | 21.0652 | 0.002715 | 0.002516 | -64.397 | 0.23892 | 8.4043 | 0.003582 | 0.005012 | 0.004809 | 6.5e-05 | 0.001094 | 0.26864 | 1.2561 | 2.548 | 100.57 | 20.76 | 19.259 | 86.9416 | 107.8876 | 82.1537 | 0.004698 | 0.99364 | 0.024174 | 24.3286 | -0.12258 | 0.14213 | 0.001124 | 0.078524 | 0.078524 | 0.074393
2019-03-20 18:01 | 1553119253.0872 | 713 | 0.20932 | 0.000618 | 5.2e-05 | 109.6859 | 110.7411 | 123.5134 | 21.0655 | 0.002685 | 0.002507 | -92.7348 | 0.23863 | 8.4032 | 0.003626 | 0.004569 | 0.004411 | 6.1e-05 | 0.001096 | 0.25354 | 1.289 | 2.5481 | 100.57 | 20.76 | 19.241 | 86.2112 | 108.9463 | 81.9404 | 0.004694 | 0.99424 | 0.02424 | 24.3802 | -0.085634 | 0.14196 | 0.001026 | 0.0719 | 0.0719 | 0.081147
2019-03-20 18:01 | 1553119254.0871 | 714 | 0.20938 | 0.000618 | 4.7e-05 | 109.5921 | 110.9724 | 123.4801 | 21.0658 | 0.002666 | 0.002491 | -121.0824 | 0.23859 | 8.4028 | 0.003668 | 0.005484 | 0.005187 | 6.6e-05 | 0.0011 | 0.30342 | 1.29 | 2.548 | 100.57 | 20.76 | 19.232 | 85.7525 | 109.9724 | 81.5669 | 0.004691 | 0.99529 | 0.024249 | 24.3641 | -0.11195 | 0.14194 | 0.000728 | 0.050417 | 0.050417 | 0.11571
2019-03-20 18:01 | 1553119255.0883 | 715 | 0.20936 | 0.00061 | 4.3e-05 | 109.5543 | 111.0528 | 123.9803 | 21.0664 | 0.002651 | 0.002502 | -149.5388 | 0.23925 | 8.4023 | 0.003648 | 0.004909 | 0.004765 | 6.7e-05 | 0.001098 | 0.27688 | 1.3032 | 2.5482 | 100.57 | 20.76 | 19.224 | 85.412 | 109.4998 | 81.8131 | 0.004689 | 0.9946 | 0.02429 | 24.4219 | -0.12152 | 0.14233 | 0.000814 | 0.057062 | 0.057062 | 0.10252
2019-03-20 18:01 | 1553119256.0877 | 716 | 0.20936 | 0.000597 | 3.9e-05 | 109.4291 | 110.8999 | 124.2777 | 21.0651 | 0.002661 | 0.00248 | -152.0858 | 0.23933 | 8.402 | 0.003652 | 0.004856 | 0.004698 | 6.4e-05 | 0.001099 | 0.26161 | 1.3489 | 2.5483 | 100.57 | 20.76 | 19.211 | 85.6441 | 109.5931 | 81.299 | 0.004686 | 0.99604 | 0.024337 | 24.4334 | -0.11942 | 0.14238 | 0.000832 | 0.058673 | 0.058673 | 0.099733
2019-03-20 18:01 | 1553119257.0883 | 717 | 0.20942 | 0.000589 | 3.5e-05 | 109.5074 | 111.0363 | 124.2479 | 21.0665 | 0.002635 | 0.002432 | -152.0793 | 0.23891 | 8.402 | 0.003606 | 0.005617 | 0.005358 | 6.4e-05 | 0.001102 | 0.27184 | 1.387 | 2.548 | 100.57 | 20.76 | 19.205 | 85.0274 | 108.4673 | 80.1406 | 0.004685 | 0.99931 | 0.024374 | 24.3906 | -0.12402 | 0.14213 | 0.000525 | 0.036442 | 0.036442 | 0.16029
2019-03-20 18:01 | 1553119258.0878 | 718 | 0.20944 | 0.000581 | 3.1e-05 | 109.573 | 111.0203 | 124.0653 | 21.0657 | 0.00258 | 0.002465 | -152.1211 | 0.23986 | 8.4007 | 0.003566 | 0.005011 | 0.004799 | 6.3e-05 | 0.00111 | 0.29214 | 1.4086 | 2.5481 | 100.57 | 20.76 | 19.196 | 83.6853 | 107.5007 | 80.9383 | 0.004683 | 0.99706 | 0.024328 | 24.4001 | -0.11736 | 0.14269 | 0.00038 | 0.025964 | 0.025964 | 0.22587
2019-03-20 18:01 | 1553119259.0875 | 719 | 0.20942 | 0.000571 | 2.7e-05 | 109.5184 | 110.8483 | 124.511 | 21.0666 | 0.002614 | 0.00242 | -152.1154 | 0.23973 | 8.4002 | 0.003642 | 0.005172 | 0.004939 | 6.8e-05 | 0.001108 | 0.29103 | 1.3817 | 2.5481 | 100.57 | 20.76 | 19.191 | 84.5165 | 109.3398 | 79.8399 | 0.004681 | 1.0002 | 0.02441 | 24.406 | -0.10797 | 0.14262 | 0.000512 | 0.036066 | 0.036066 | 0.16252
2019-03-20 18:01 | 1553119260.0885 | 720 | 0.20949 | 0.000557 | 2.4e-05 | 109.531 | 110.9683 | 124.5542 | 21.0654 | 0.002684 | 0.002387 | -152.1147 | 0.23876 | 8.4032 | 0.003655 | 0.005102 | 0.004864 | 6.5e-05 | 0.00111 | 0.28727 | 1.381 | 2.5481 | 100.57 | 20.76 | 19.189 | 86.1938 | 109.6534 | 79.0443 | 0.004681 | 1.0024 | 0.024442 | 24.3827 | -0.11661 | 0.14204 | 0.000167 | 0.01085 | 0.01085 | 0.53804
2019-03-20 18:01 | 1553119261.0879 | 721 | 0.20952 | 0.000546 | 2e-05 | 109.5924 | 111.0341 | 124.5066 | 21.0655 | 0.002561 | 0.002402 | -152.1272 | 0.23975 | 8.4025 | 0.003616 | 0.004706 | 0.004459 | 7.1e-05 | 0.001112 | 0.27865 | 1.3703 | 2.548 | 100.57 | 20.76 | 19.185 | 83.246 | 108.7248 | 79.4185 | 0.00468 | 1.0014 | 0.024424 | 24.391 | -0.1169 | 0.14263 | -1e-06 | -0.001328 | -0.001328 | -4.415
2019-03-20 18:01 | 1553119262.0886 | 722 | 0.20951 | 0.000535 | 1.6e-05 | 109.422 | 110.7725 | 124.1737 | 21.0656 | 0.002536 | 0.00241 | -152.1239 | 0.23945 | 8.4045 | 0.003522 | 0.005828 | 0.00557 | 7e-05 | 0.001113 | 0.29175 | 1.3528 | 2.548 | 100.57 | 20.76 | 19.183 | 82.632 | 106.4218 | 79.6058 | 0.004679 | 1.0008 | 0.024385 | 24.3648 | -0.10972 | 0.14245 | 3.2e-05 | 0.001337 | 0.001337 | 4.3784
2019-03-20 18:01 | 1553119263.0881 | 723 | 0.20957 | 0.000522 | 1.4e-05 | 109.4686 | 110.8894 | 124.2183 | 21.065 | 0.002625 | 0.002342 | -152.1158 | 0.23931 | 8.4025 | 0.003569 | 0.004908 | 0.004758 | 6.9e-05 | 0.001116 | 0.29607 | 1.3275 | 2.548 | 100.57 | 20.76 | 19.18 | 84.7653 | 107.5638 | 77.968 | 0.004679 | 1.0055 | 0.024446 | 24.3126 | -0.11535 | 0.14236 | -0.000311 | -0.023775 | -0.023775 | -0.2461
2019-03-20 18:01 | 1553119264.0888 | 724 | 0.20952 | 0.000518 | 1.3e-05 | 109.5368 | 111.0251 | 124.4763 | 21.0658 | 0.002553 | 0.002375 | -152.0973 | 0.239 | 8.4018 | 0.003579 | 0.004669 | 0.004472 | 6.8e-05 | 0.001107 | 0.27472 | 1.3057 | 2.548 | 100.57 | 20.76 | 19.178 | 83.0583 | 107.8047 | 78.764 | 0.004678 | 1.0032 | 0.024444 | 24.3654 | -0.12071 | 0.14218 | 3.7e-05 | 0.00197 | 0.00197 | 2.966
2019-03-20 18:01 | 1553119265.0878 | 725 | 0.20951 | 0.000511 | 1e-05 | 109.488 | 110.8188 | 124.1899 | 21.065 | 0.002597 | 0.002318 | -152.1331 | 0.23837 | 8.4028 | 0.003521 | 0.004735 | 0.004538 | 7.2e-05 | 0.001105 | 0.27312 | 1.3142 | 2.548 | 100.57 | 20.76 | 19.176 | 84.1037 | 106.4086 | 77.3872 | 0.004678 | 1.0072 | 0.024463 | 24.2897 | -0.10807 | 0.14181 | 9.6e-05 | 0.006463 | 0.006463 | 0.9018
2019-03-20 18:01 | 1553119266.0869 | 726 | 0.20954 | 0.000499 | 8e-06 | 109.4362 | 110.9586 | 124.1362 | 21.0664 | 0.002538 | 0.002315 | -152.125 | 0.23808 | 8.4051 | 0.003531 | 0.004945 | 0.004771 | 7e-05 | 0.001102 | 0.29878 | 1.3136 | 2.5481 | 100.57 | 20.76 | 19.169 | 82.6946 | 106.6619 | 77.3201 | 0.004676 | 1.0073 | 0.024461 | 24.2821 | -0.12357 | 0.14163 | -7.2e-05 | -0.005831 | -0.005831 | -0.99834
2019-03-20 18:01 | 1553119267.0873 | 727 | 0.20953 | 0.000491 | 7e-06 | 109.4301 | 110.8218 | 124.3195 | 21.066 | 0.0026 | 0.002331 | -152.1177 | 0.23932 | 8.4015 | 0.003444 | 0.005372 | 0.005047 | 7e-05 | 0.0011 | 0.27599 | 1.2955 | 2.5481 | 100.57 | 20.76 | 19.164 | 84.1655 | 104.5521 | 77.7187 | 0.004675 | 1.0062 | 0.024465 | 24.3138 | -0.11304 | 0.14237 | 2.3e-05 | 0.001297 | 0.001297 | 4.5133
2019-03-20 18:01 | 1553119268.088 | 728 | 0.20956 | 0.000482 | 6e-06 | 109.3502 | 110.7194 | 124.2892 | 21.0654 | 0.002576 | 0.002372 | -152.0787 | 0.23857 | 8.4033 | 0.003515 | 0.005051 | 0.004958 | 7.1e-05 | 0.001099 | 0.27256 | 1.331 | 2.548 | 100.57 | 20.76 | 19.159 | 83.5977 | 106.265 | 78.6957 | 0.004674 | 1.0034 | 0.024428 | 24.3447 | -0.1113 | 0.14193 | -0.000176 | -0.013312 | -0.013312 | -0.43819
2019-03-20 18:01 | 1553119269.0883 | 729 | 0.20956 | 0.00048 | 5e-06 | 109.4026 | 110.6689 | 124.1326 | 21.0648 | 0.002519 | 0.00239 | -152.0865 | 0.23981 | 8.4021 | 0.003512 | 0.004657 | 0.00435 | 7.1e-05 | 0.0011 | 0.28854 | 1.3412 | 2.548 | 100.57 | 20.76 | 19.149 | 82.2267 | 106.178 | 79.1202 | 0.004671 | 1.0022 | 0.024398 | 24.344 | -0.10294 | 0.14266 | -0.000172 | -0.01294 | -0.01294 | -0.45313
2019-03-20 18:01 | 1553119270.0872 | 730 | 0.20956 | 0.000476 | 4e-06 | 109.4223 | 110.7979 | 124.1899 | 21.0652 | 0.002561 | 0.002302 | -152.0768 | 0.24008 | 8.4011 | 0.003493 | 0.004543 | 0.004353 | 7.3e-05 | 0.001104 | 0.28783 | 1.3293 | 2.548 | 100.57 | 20.76 | 19.143 | 83.2324 | 105.7187 | 77.0226 | 0.00467 | 1.0082 | 0.024476 | 24.277 | -0.11174 | 0.14282 | -0.000156 | -0.011729 | -0.011729 | -0.50046
2019-03-20 18:01 | 1553119271.0882 | 731 | 0.20959 | 0.00047 | 4e-06 | 109.3508 | 110.8367 | 124.3577 | 21.0656 | 0.002457 | 0.002356 | -152.0683 | 0.24002 | 8.4035 | 0.003476 | 0.004569 | 0.004473 | 6.9e-05 | 0.001109 | 0.30718 | 1.356 | 2.548 | 100.57 | 20.76 | 19.128 | 80.7394 | 105.329 | 78.3068 | 0.004666 | 1.0045 | 0.024448 | 24.3379 | -0.12072 | 0.14279 | -0.000301 | -0.02243 | -0.02243 | -0.26164
2019-03-20 18:01 | 1553119272.0872 | 732 | 0.20954 | 0.000466 | 2e-06 | 109.3567 | 110.7097 | 124.5312 | 21.0663 | 0.002494 | 0.002358 | -152.0619 | 0.23931 | 8.4031 | 0.003474 | 0.004719 | 0.004506 | 7e-05 | 0.001103 | 0.2907 | 1.3169 | 2.548 | 100.57 | 20.76 | 19.12 | 81.6344 | 105.2744 | 78.3613 | 0.004664 | 1.0044 | 0.024463 | 24.3568 | -0.10999 | 0.14236 | -1.4e-05 | -0.001148 | -0.001148 | -5.0983
2019-03-20 18:01 | 1553119273.0883 | 733 | 0.20956 | 0.000463 | 2e-06 | 109.3945 | 110.6921 | 124.4047 | 21.0649 | 0.002453 | 0.002313 | -152.0743 | 0.2392 | 8.4028 | 0.003491 | 0.00434 | 0.004065 | 6.7e-05 | 0.001104 | 0.29911 | 1.3414 | 2.5481 | 100.57 | 20.76 | 19.111 | 80.6382 | 105.6806 | 77.278 | 0.004662 | 1.0075 | 0.024488 | 24.3069 | -0.10547 | 0.1423 | -0.000154 | -0.011463 | -0.011463 | -0.51021
2019-03-20 18:01 | 1553119274.0885 | 734 | 0.20956 | 0.000455 | 2e-06 | 109.3505 | 110.6536 | 124.3594 | 21.0653 | 0.002484 | 0.002294 | -152.0892 | 0.23947 | 8.4027 | 0.003442 | 0.004992 | 0.004842 | 7.1e-05 | 0.001103 | 0.27903 | 1.3376 | 2.5479 | 100.57 | 20.77 | 19.093 | 81.3933 | 104.4909 | 76.8224 | 0.00466 | 1.0088 | 0.0245 | 24.2866 | -0.10596 | 0.14246 | -0.000129 | -0.009599 | -0.009599 | -0.60999
2019-03-20 18:01 | 1553119275.0883 | 735 | 0.20956 | 0.000458 | 1e-06 | 109.3728 | 110.6727 | 124.523 | 21.0655 | 0.002462 | 0.00228 | -152.0507 | 0.23934 | 8.4031 | 0.003423 | 0.004787 | 0.004517 | 7e-05 | 0.001105 | 0.28622 | 1.3878 | 2.5481 | 100.57 | 20.77 | 19.084 | 80.8554 | 104.0343 | 76.4844 | 0.004658 | 1.0098 | 0.024528 | 24.2909 | -0.10567 | 0.14238 | -0.000117 | -0.008719 | -0.008719 | -0.67115
2019-03-20 18:01 | 1553119276.0873 | 736 | 0.20955 | 0.000453 | 1e-06 | 109.3075 | 110.6682 | 124.6033 | 21.066 | 0.002458 | 0.002294 | -152.0681 | 0.2396 | 8.4 | 0.003414 | 0.004622 | 0.004447 | 7.1e-05 | 0.001104 | 0.2855 | 1.344 | 2.548 | 100.57 | 20.77 | 19.063 | 80.7617 | 103.8075 | 76.8325 | 0.004653 | 1.0088 | 0.024524 | 24.3108 | -0.11066 | 0.14253 | -6.6e-05 | -0.004935 | -0.004935 | -1.1872
2019-03-20 18:01 | 1553119277.0874 | 737 | 0.20953 | 0.000453 | 1e-06 | 109.3014 | 110.6099 | 124.4737 | 21.0654 | 0.002479 | 0.002227 | -152.0734 | 0.23947 | 8.4023 | 0.003493 | 0.004358 | 0.004119 | 7e-05 | 0.001105 | 0.2625 | 1.3021 | 2.5478 | 100.57 | 20.77 | 19.053 | 81.2586 | 105.719 | 75.2215 | 0.004651 | 1.0134 | 0.024567 | 24.2422 | -0.10644 | 0.14246 | 7.6e-05 | 0.005539 | 0.005539 | 1.057
2019-03-20 18:01 | 1553119278.0888 | 738 | 0.20953 | 0.000451 | 1e-06 | 109.2587 | 110.7257 | 124.6382 | 21.067 | 0.002434 | 0.002284 | -152.0946 | 0.23913 | 8.4019 | 0.003388 | 0.005023 | 0.004747 | 7.1e-05 | 0.001104 | 0.26969 | 1.3549 | 2.5481 | 100.57 | 20.77 | 19.045 | 80.1925 | 103.1791 | 76.5902 | 0.004649 | 1.0094 | 0.024535 | 24.3058 | -0.1193 | 0.14226 | 6e-05 | 0.004418 | 0.004418 | 1.3235
2019-03-20 18:01 | 1553119279.0879 | 739 | 0.20953 | 0.000455 | 0 | 109.291 | 110.6779 | 124.7019 | 21.0654 | 0.002452 | 0.00228 | -152.1046 | 0.23827 | 8.4038 | 0.003423 | 0.00549 | 0.005306 | 6.8e-05 | 0.001104 | 0.2589 | 1.3647 | 2.5481 | 100.57 | 20.77 | 19.028 | 80.62 | 104.0264 | 76.4912 | 0.004645 | 1.0097 | 0.024545 | 24.3085 | -0.11279 | 0.14175 | 5.9e-05 | 0.004372 | 0.004372 | 1.3325
2019-03-20 18:01 | 1553119280.0885 | 740 | 0.20953 | 0.000453 | 0 | 109.2981 | 110.6282 | 124.5382 | 21.0666 | 0.002399 | 0.002226 | -152.1177 | 0.23865 | 8.402 | 0.00342 | 0.004604 | 0.004389 | 6.6e-05 | 0.001099 | 0.28313 | 1.3172 | 2.5482 | 100.57 | 20.77 | 19.019 | 79.3439 | 103.9569 | 75.1789 | 0.004642 | 1.0135 | 0.024575 | 24.247 | -0.10819 | 0.14197 | 4e-05 | 0.002918 | 0.002918 | 1.9998
2019-03-20 18:01 | 1553119281.0876 | 741 | 0.20955 | 0.000453 | 0 | 109.2409 | 110.7706 | 124.2465 | 21.0662 | 0.002384 | 0.002295 | -152.1081 | 0.23933 | 8.4036 | 0.00344 | 0.00434 | 0.004157 | 6.8e-05 | 0.001098 | 0.27007 | 1.3642 | 2.5481 | 100.57 | 20.76 | 19.011 | 78.9788 | 104.4423 | 76.8536 | 0.004637 | 1.0087 | 0.024488 | 24.2767 | -0.12438 | 0.14238 | -6e-05 | -0.004448 | -0.004448 | -1.3154
2019-03-20 18:01 | 1553119282.0868 | 742 | 0.20951 | 0.00046 | 0 | 109.2228 | 110.5669 | 124.749 | 21.0663 | 0.002408 | 0.002282 | -152.1047 | 0.23853 | 8.4014 | 0.003507 | 0.004787 | 0.004624 | 7e-05 | 0.001092 | 0.29103 | 1.3524 | 2.548 | 100.57 | 20.77 | 18.997 | 79.5572 | 106.0802 | 76.531 | 0.004637 | 1.0096 | 0.024548 | 24.3145 | -0.1094 | 0.14191 | 0.000173 | 0.012821 | 0.012821 | 0.4549
2019-03-20 18:01 | 1553119283.0871 | 743 | 0.20955 | 0.000462 | 0 | 109.1397 | 110.6274 | 124.7551 | 21.0654 | 0.002387 | 0.002241 | -152.0999 | 0.23819 | 8.4046 | 0.003517 | 0.005754 | 0.005483 | 7.1e-05 | 0.001087 | 0.28727 | 1.3464 | 2.5482 | 100.57 | 20.77 | 18.991 | 79.0474 | 106.3209 | 75.5456 | 0.004635 | 1.0125 | 0.024584 | 24.2808 | -0.12111 | 0.1417 | -3.8e-05 | -0.002776 | -0.002776 | -2.0978
2019-03-20 18:01 | 1553119284.0868 | 744 | 0.2095 | 0.000463 | 0 | 109.1768 | 110.7519 | 124.7263 | 21.0663 | 0.002452 | 0.002234 | -152.094 | 0.23849 | 8.3997 | 0.003378 | 0.004885 | 0.004744 | 6.9e-05 | 0.001089 | 0.29535 | 1.3532 | 2.5481 | 100.57 | 20.77 | 18.984 | 80.6285 | 102.9426 | 75.3922 | 0.004634 | 1.0129 | 0.024586 | 24.2727 | -0.12812 | 0.14188 | 0.000233 | 0.017296 | 0.017296 | 0.33714
2019-03-20 18:01 | 1553119285.0873 | 745 | 0.20955 | 0.000467 | 1e-06 | 109.1164 | 110.4406 | 125.0675 | 21.0666 | 0.002442 | 0.002249 | -152.1076 | 0.239 | 8.4025 | 0.003389 | 0.004991 | 0.004783 | 6.9e-05 | 0.001089 | 0.24818 | 1.3668 | 2.5481 | 100.57 | 20.77 | 18.984 | 80.3809 | 103.2069 | 75.7433 | 0.004634 | 1.0119 | 0.024608 | 24.3181 | -0.1079 | 0.14218 | -0.000102 | -0.007595 | -0.007595 | -0.76943
2019-03-20 18:01 | 1553119286.088 | 746 | 0.20955 | 0.000463 | 0 | 109.0902 | 110.2784 | 124.922 | 21.0652 | 0.002382 | 0.002254 | -152.1093 | 0.23815 | 8.4036 | 0.003449 | 0.004448 | 0.004243 | 7.2e-05 | 0.001088 | 0.29767 | 1.3673 | 2.5481 | 100.57 | 20.77 | 18.985 | 78.9333 | 104.6701 | 75.8572 | 0.004634 | 1.0116 | 0.024589 | 24.3079 | -0.096896 | 0.14168 | -8.8e-05 | -0.006509 | -0.006509 | -0.89455
2019-03-20 18:01 | 1553119287.088 | 747 | 0.20957 | 0.000467 | 0 | 109.0417 | 110.3531 | 124.7798 | 21.0662 | 0.002334 | 0.002154 | -152.1142 | 0.23846 | 8.4045 | 0.003346 | 0.004413 | 0.004224 | 7.2e-05 | 0.001089 | 0.2573 | 1.3824 | 2.5481 | 100.57 | 20.76 | 18.996 | 77.7906 | 102.1697 | 73.4636 | 0.004634 | 1.0186 | 0.02466 | 24.2106 | -0.10694 | 0.14186 | -0.00018 | -0.013398 | -0.013398 | -0.43516
2019-03-20 18:01 | 1553119288.087 | 748 | 0.20955 | 0.000463 | 0 | 109.0662 | 110.5041 | 125.0368 | 21.0647 | 0.002375 | 0.002174 | -152.1166 | 0.23827 | 8.4027 | 0.003432 | 0.004537 | 0.004314 | 7.4e-05 | 0.001087 | 0.28512 | 1.375 | 2.548 | 100.57 | 20.76 | 19.005 | 78.7567 | 104.2513 | 73.9299 | 0.004636 | 1.0172 | 0.024669 | 24.2519 | -0.11715 | 0.14175 | -4.3e-05 | -0.003212 | -0.003212 | -1.8137
2019-03-20 18:01 | 1553119289.0885 | 749 | 0.20954 | 0.000466 | 0 | 109.0132 | 110.3972 | 124.8004 | 21.065 | 0.002409 | 0.002194 | -152.1056 | 0.23848 | 8.4035 | 0.003373 | 0.004163 | 0.003954 | 7.1e-05 | 0.001089 | 0.29159 | 1.3532 | 2.5482 | 100.57 | 20.76 | 19.024 | 79.5904 | 102.8227 | 74.4144 | 0.004641 | 1.0158 | 0.024628 | 24.2458 | -0.11284 | 0.14187 | -1.6e-05 | -0.001234 | -0.001234 | -4.7268
2019-03-20 18:01 | 1553119290.0882 | 750 | 0.20954 | 0.000466 | 0 | 108.9925 | 110.3669 | 124.9805 | 21.0661 | 0.002347 | 0.002259 | -152.0997 | 0.23877 | 8.4036 | 0.003393 | 0.004741 | 0.004541 | 7.2e-05 | 0.001085 | 0.26432 | 1.3576 | 2.548 | 100.57 | 20.76 | 19.034 | 78.0844 | 103.296 | 75.9922 | 0.004643 | 1.0112 | 0.02459 | 24.3183 | -0.11209 | 0.14204 | 5e-06 | 0.000398 | 0.000398 | 14.6645
2019-03-20 18:02 | 1553119291.0882 | 751 | 0.20955 | 0.000462 | 0 | 109.0501 | 110.4809 | 125.3544 | 21.0659 | 0.002385 | 0.002223 | -152.1173 | 0.23951 | 8.4025 | 0.003349 | 0.004891 | 0.004662 | 7.1e-05 | 0.001078 | 0.27776 | 1.3359 | 2.5481 | 100.57 | 20.76 | 19.044 | 79.0003 | 102.2431 | 75.1056 | 0.004646 | 1.0138 | 0.024658 | 24.3237 | -0.1166 | 0.14249 | -5.3e-05 | -0.003973 | -0.003973 | -1.4739
2019-03-20 18:02 | 1553119292.0882 | 752 | 0.20952 | 0.000458 | 0 | 109.0659 | 110.3684 | 124.9755 | 21.0655 | 0.002295 | 0.002194 | -152.0982 | 0.2393 | 8.4017 | 0.003336 | 0.004753 | 0.004538 | 7.1e-05 | 0.001084 | 0.274 | 1.3153 | 2.548 | 100.57 | 20.76 | 19.067 | 76.8372 | 101.9372 | 74.4092 | 0.004651 | 1.0158 | 0.024646 | 24.2626 | -0.10619 | 0.14236 | 0.000118 | 0.008796 | 0.008796 | 0.66521
2019-03-20 18:02 | 1553119293.0883 | 753 | 0.20959 | 0.000457 | 0 | 109.0798 | 110.3826 | 124.8932 | 21.0667 | 0.002334 | 0.002177 | -152.081 | 0.23784 | 8.4025 | 0.003372 | 0.004853 | 0.004644 | 7.2e-05 | 0.001081 | 0.28296 | 1.2918 | 2.548 | 100.57 | 20.76 | 19.078 | 77.7949 | 102.8056 | 74.0183 | 0.004654 | 1.0169 | 0.024651 | 24.241 | -0.1062 | 0.14149 | -0.000307 | -0.02285 | -0.02285 | -0.2545
2019-03-20 18:02 | 1553119294.0888 | 754 | 0.20958 | 0.000454 | 0 | 108.9359 | 110.3404 | 125.4494 | 21.0672 | 0.002323 | 0.00216 | -152.0937 | 0.23855 | 8.4035 | 0.003424 | 0.004914 | 0.004684 | 7.3e-05 | 0.001084 | 0.29031 | 1.3141 | 2.5482 | 100.57 | 20.76 | 19.101 | 77.5147 | 104.0619 | 73.5994 | 0.004659 | 1.0182 | 0.024721 | 24.2803 | -0.11458 | 0.14191 | -0.000206 | -0.015316 | -0.015316 | -0.38081
2019-03-20 18:02 | 1553119295.0885 | 755 | 0.20952 | 0.000456 | 0 | 109.0462 | 110.3826 | 125.2302 | 21.0667 | 0.002372 | 0.00226 | -152.0762 | 0.23882 | 8.4023 | 0.003313 | 0.004562 | 0.004338 | 6.9e-05 | 0.001082 | 0.29231 | 1.3632 | 2.548 | 100.57 | 20.76 | 19.112 | 78.7017 | 101.3627 | 76.0006 | 0.004662 | 1.0112 | 0.024614 | 24.3429 | -0.10896 | 0.14208 | 0.000103 | 0.0077 | 0.0077 | 0.75836
2019-03-20 18:02 | 1553119296.0877 | 756 | 0.20958 | 0.000454 | 0 | 109.019 | 110.4286 | 125.1091 | 21.0656 | 0.002355 | 0.002195 | -152.0682 | 0.23897 | 8.4037 | 0.003342 | 0.004778 | 0.004577 | 7.1e-05 | 0.001076 | 0.26233 | 1.3626 | 2.5482 | 100.57 | 20.76 | 19.123 | 78.2811 | 102.0637 | 74.4389 | 0.004665 | 1.0157 | 0.024658 | 24.2766 | -0.11491 | 0.14216 | -0.000208 | -0.015468 | -0.015468 | -0.37775
2019-03-20 18:02 | 1553119297.0875 | 757 | 0.20955 | 0.00045 | 1e-06 | 108.9915 | 110.3288 | 125.0373 | 21.0669 | 0.002274 | 0.00219 | -152.0615 | 0.23891 | 8.402 | 0.003353 | 0.004307 | 0.004226 | 7.3e-05 | 0.00108 | 0.29015 | 1.3371 | 2.5483 | 100.57 | 20.76 | 19.142 | 76.3384 | 102.3453 | 74.3162 | 0.004669 | 1.0161 | 0.024655 | 24.2654 | -0.10908 | 0.14213 | -7.7e-05 | -0.005746 | -0.005746 | -1.0165
2019-03-20 18:02 | 1553119298.0875 | 758 | 0.20956 | 0.00045 | 0 | 109.0277 | 110.337 | 124.6993 | 21.0671 | 0.002298 | 0.002161 | -152.0543 | 0.23916 | 8.4055 | 0.003318 | 0.005234 | 0.004925 | 7.5e-05 | 0.001078 | 0.27936 | 1.348 | 2.5481 | 100.57 | 20.76 | 19.15 | 76.9169 | 101.4833 | 73.6297 | 0.004671 | 1.0181 | 0.024646 | 24.2086 | -0.10677 | 0.14228 | -9.6e-05 | -0.007112 | -0.007112 | -0.82217
2019-03-20 18:02 | 1553119299.0866 | 759 | 0.20951 | 0.000447 | 0 | 109.0284 | 110.2944 | 124.8514 | 21.0669 | 0.00232 | 0.00219 | -152.0459 | 0.23833 | 8.4047 | 0.003421 | 0.00424 | 0.004053 | 7.4e-05 | 0.001084 | 0.274 | 1.3631 | 2.5482 | 100.57 | 20.76 | 19.157 | 77.4412 | 103.994 | 74.3131 | 0.004673 | 1.0161 | 0.024637 | 24.2472 | -0.10326 | 0.14178 | 0.000179 | 0.013303 | 0.013303 | 0.43806
2019-03-20 18:02 | 1553119300.0866 | 760 | 0.20948 | 0.000447 | 0 | 109.0083 | 110.429 | 124.6427 | 21.0666 | 0.002328 | 0.002231 | -152.0503 | 0.23774 | 8.4034 | 0.003363 | 0.005349 | 0.005095 | 7.2e-05 | 0.001091 | 0.28263 | 1.3573 | 2.5482 | 100.57 | 20.76 | 19.168 | 77.6319 | 102.5745 | 75.3146 | 0.004676 | 1.0131 | 0.024581 | 24.2619 | -0.11582 | 0.14143 | 0.000394 | 0.029225 | 0.029225 | 0.1989
2019-03-20 18:02 | 1553119301.0877 | 761 | 0.20943 | 0.000456 | -1e-06 | 108.9844 | 110.3497 | 124.7265 | 21.0673 | 0.002213 | 0.002174 | -152.0524 | 0.23903 | 8.4005 | 0.003364 | 0.004959 | 0.004737 | 7.2e-05 | 0.001091 | 0.27992 | 1.3499 | 2.5481 | 100.57 | 20.76 | 19.175 | 74.8726 | 102.5988 | 73.9512 | 0.004678 | 1.0171 | 0.024637 | 24.2225 | -0.11136 | 0.1422 | 0.000694 | 0.051633 | 0.051633 | 0.11319
2019-03-20 18:02 | 1553119302.0868 | 762 | 0.2094 | 0.000464 | 0 | 109.0527 | 110.4884 | 124.688 | 21.0667 | 0.002246 | 0.002248 | -152.0634 | 0.23961 | 8.4034 | 0.003252 | 0.00437 | 0.004147 | 7e-05 | 0.001095 | 0.25658 | 1.314 | 2.5479 | 100.57 | 20.76 | 19.188 | 75.66 | 99.8826 | 75.7153 | 0.004681 | 1.012 | 0.024571 | 24.2802 | -0.11699 | 0.14254 | 0.000844 | 0.062559 | 0.062559 | 0.093648
2019-03-20 18:02 | 1553119303.0868 | 763 | 0.20939 | 0.000475 | 0 | 109.0556 | 110.349 | 125.1424 | 21.0663 | 0.002322 | 0.002181 | -152.0742 | 0.23936 | 8.4035 | 0.003344 | 0.004703 | 0.004495 | 7e-05 | 0.001094 | 0.26665 | 1.3341 | 2.548 | 100.57 | 20.76 | 19.195 | 77.5069 | 102.1288 | 74.1155 | 0.004682 | 1.0166 | 0.024673 | 24.2686 | -0.10546 | 0.1424 | 0.000904 | 0.067294 | 0.067294 | 0.086968
2019-03-20 18:02 | 1553119304.088 | 764 | 0.20936 | 0.000489 | 0 | 109.0203 | 110.3064 | 124.9746 | 21.0677 | 0.002221 | 0.002175 | -152.0787 | 0.23862 | 8.4051 | 0.003418 | 0.00488 | 0.004668 | 7.2e-05 | 0.001096 | 0.30453 | 1.3037 | 2.5481 | 100.57 | 20.76 | 19.203 | 75.0716 | 103.9184 | 73.9631 | 0.004684 | 1.0171 | 0.024661 | 24.247 | -0.1049 | 0.14195 | 0.001019 | 0.075862 | 0.075862 | 0.076907
2019-03-20 18:02 | 1553119305.0879 | 765 | 0.20939 | 0.000509 | 0 | 109.063 | 110.4152 | 125.0079 | 21.0673 | 0.002267 | 0.002176 | -152.0826 | 0.23875 | 8.4041 | 0.00338 | 0.005124 | 0.004834 | 7.2e-05 | 0.001098 | 0.29247 | 1.3369 | 2.548 | 100.57 | 20.76 | 19.211 | 76.1735 | 102.9995 | 73.9978 | 0.004686 | 1.017 | 0.024663 | 24.2514 | -0.11021 | 0.14203 | 0.000814 | 0.060585 | 0.060585 | 0.096353
2019-03-20 18:02 | 1553119306.087 | 766 | 0.20942 | 0.000528 | -1e-06 | 109.0756 | 110.5467 | 125.1762 | 21.0667 | 0.002337 | 0.002154 | -152.0869 | 0.23937 | 8.4003 | 0.003455 | 0.005083 | 0.004911 | 7.3e-05 | 0.00109 | 0.24873 | 1.3295 | 2.548 | 100.57 | 20.76 | 19.228 | 77.8466 | 104.7992 | 73.4596 | 0.00469 | 1.0186 | 0.024699 | 24.2489 | -0.11983 | 0.1424 | 0.000641 | 0.047808 | 0.047808 | 0.12242
2019-03-20 18:02 | 1553119307.0871 | 767 | 0.20942 | 0.000536 | 0 | 109.0724 | 110.309 | 125.5088 | 21.0674 | 0.002284 | 0.002112 | -152.0965 | 0.23971 | 8.4031 | 0.003456 | 0.005267 | 0.004978 | 7.1e-05 | 0.001095 | 0.26416 | 1.3394 | 2.548 | 100.57 | 20.76 | 19.238 | 76.5876 | 104.8422 | 72.4468 | 0.004693 | 1.0216 | 0.024768 | 24.2456 | -0.10084 | 0.1426 | 0.000636 | 0.047531 | 0.047531 | 0.12331
2019-03-20 18:02 | 1553119308.0883 | 768 | 0.20943 | 0.000541 | -1e-06 | 108.9459 | 110.3083 | 125.695 | 21.0672 | 0.002231 | 0.0021 | -152.1056 | 0.2388 | 8.4029 | 0.003455 | 0.004745 | 0.00456 | 7.1e-05 | 0.001097 | 0.27417 | 1.3676 | 2.548 | 100.57 | 20.76 | 19.247 | 75.3184 | 104.8112 | 72.158 | 0.004695 | 1.0224 | 0.024797 | 24.2535 | -0.11116 | 0.14206 | 0.000587 | 0.04398 | 0.04398 | 0.13276
2019-03-20 18:02 | 1553119309.0872 | 769 | 0.20945 | 0.000543 | 0 | 109.0012 | 110.3793 | 125.8048 | 21.0676 | 0.002298 | 0.002141 | -152.1239 | 0.23914 | 8.4018 | 0.003457 | 0.004761 | 0.004536 | 7e-05 | 0.001095 | 0.26106 | 1.3733 | 2.5481 | 100.57 | 20.76 | 19.263 | 76.9125 | 104.847 | 73.1586 | 0.004699 | 1.0195 | 0.024772 | 24.2992 | -0.11237 | 0.14226 | 0.000422 | 0.031552 | 0.031552 | 0.18531
2019-03-20 18:02 | 1553119310.0872 | 770 | 0.20947 | 0.00054 | 0 | 109.02 | 110.3273 | 125.8229 | 21.0661 | 0.002313 | 0.002073 | -152.1298 | 0.23927 | 8.4011 | 0.003426 | 0.004266 | 0.004031 | 6.9e-05 | 0.001093 | 0.26001 | 1.3359 | 2.5482 | 100.57 | 20.76 | 19.272 | 77.2838 | 104.1073 | 71.5053 | 0.004701 | 1.0243 | 0.024833 | 24.2428 | -0.10662 | 0.14234 | 0.000323 | 0.024261 | 0.024261 | 0.24114
2019-03-20 18:02 | 1553119311.0873 | 771 | 0.20946 | 0.000534 | -1e-06 | 108.899 | 110.3471 | 125.59 | 21.066 | 0.00226 | 0.002128 | -152.13 | 0.23902 | 8.4011 | 0.003359 | 0.005191 | 0.00499 | 7e-05 | 0.001093 | 0.26322 | 1.3684 | 2.5482 | 100.57 | 20.76 | 19.281 | 76.0132 | 102.4844 | 72.8344 | 0.004703 | 1.0204 | 0.024762 | 24.2671 | -0.11816 | 0.14219 | 0.000371 | 0.027761 | 0.027761 | 0.21051
2019-03-20 18:02 | 1553119312.0872 | 772 | 0.20949 | 0.000527 | 0 | 108.9433 | 110.2029 | 125.7086 | 21.0665 | 0.002297 | 0.002149 | -152.141 | 0.23893 | 8.4018 | 0.003409 | 0.004763 | 0.004523 | 6.7e-05 | 0.00109 | 0.28385 | 1.3388 | 2.5482 | 100.57 | 20.76 | 19.299 | 76.9038 | 103.7004 | 73.343 | 0.004708 | 1.0189 | 0.024756 | 24.2963 | -0.10282 | 0.14214 | 0.000241 | 0.017999 | 0.017999 | 0.32458
2019-03-20 18:02 | 1553119313.0868 | 773 | 0.20954 | 0.000521 | 0 | 108.956 | 110.3075 | 125.5612 | 21.0664 | 0.002224 | 0.002203 | -152.1488 | 0.23909 | 8.4021 | 0.00342 | 0.004759 | 0.004596 | 7.2e-05 | 0.001094 | 0.27527 | 1.3524 | 2.5481 | 100.57 | 20.76 | 19.307 | 75.1336 | 103.9535 | 74.6283 | 0.00471 | 1.0151 | 0.024696 | 24.3271 | -0.11027 | 0.14223 | -9.6e-05 | -0.007179 | -0.007179 | -0.81434
2019-03-20 18:02 | 1553119314.0874 | 774 | 0.20954 | 0.000516 | 0 | 108.931 | 110.1404 | 125.2894 | 21.0664 | 0.002244 | 0.002108 | -152.152 | 0.2389 | 8.4001 | 0.003352 | 0.004145 | 0.003945 | 7e-05 | 0.001093 | 0.27599 | 1.2971 | 2.548 | 100.57 | 20.76 | 19.323 | 75.6315 | 102.3245 | 72.3444 | 0.004714 | 1.0219 | 0.02475 | 24.2208 | -0.098759 | 0.14212 | -5e-05 | -0.003681 | -0.003681 | -1.5868
2019-03-20 18:02 | 1553119315.0864 | 775 | 0.20954 | 0.000501 | 0 | 108.8605 | 110.2671 | 125.5091 | 21.066 | 0.002222 | 0.002135 | -152.1482 | 0.23904 | 8.4028 | 0.003374 | 0.004524 | 0.004276 | 7.2e-05 | 0.001094 | 0.27328 | 1.2997 | 2.5481 | 100.57 | 20.76 | 19.33 | 75.0933 | 102.8348 | 73.006 | 0.004715 | 1.0199 | 0.024748 | 24.2652 | -0.11483 | 0.1422 | -4.6e-05 | -0.00342 | -0.00342 | -1.7088
2019-03-20 18:02 | 1553119316.0865 | 776 | 0.20957 | 0.000489 | 0 | 108.8557 | 110.0478 | 125.6422 | 21.0655 | 0.002233 | 0.002093 | -152.1392 | 0.23865 | 8.4023 | 0.003397 | 0.005369 | 0.005157 | 7.3e-05 | 0.001094 | 0.28766 | 1.2575 | 2.548 | 100.57 | 20.76 | 19.337 | 75.3477 | 103.4095 | 71.9913 | 0.004717 | 1.0229 | 0.024798 | 24.2425 | -0.097418 | 0.14197 | -0.0002 | -0.014952 | -0.014952 | -0.39024
2019-03-20 18:02 | 1553119317.0865 | 777 | 0.20957 | 0.000484 | 0 | 108.8182 | 110.1389 | 125.4689 | 21.0658 | 0.002234 | 0.002091 | -152.1454 | 0.23932 | 8.405 | 0.003424 | 0.005098 | 0.004926 | 6.9e-05 | 0.001096 | 0.26432 | 1.2923 | 2.5482 | 100.57 | 20.76 | 19.35 | 75.3816 | 104.0581 | 71.9385 | 0.00472 | 1.0231 | 0.024782 | 24.2239 | -0.10791 | 0.14237 | -0.000211 | -0.015777 | -0.015777 | -0.37089
2019-03-20 18:02 | 1553119318.0869 | 778 | 0.20957 | 0.000483 | 0 | 108.8227 | 110.1449 | 125.5544 | 21.0668 | 0.002286 | 0.002051 | -152.1539 | 0.23857 | 8.4025 | 0.003441 | 0.004963 | 0.004823 | 7.2e-05 | 0.001096 | 0.29015 | 1.3374 | 2.548 | 100.57 | 20.76 | 19.355 | 76.6398 | 104.4671 | 70.9782 | 0.004721 | 1.0259 | 0.024825 | 24.1984 | -0.10802 | 0.14192 | -0.000202 | -0.015144 | -0.015144 | -0.38517
2019-03-20 18:02 | 1553119319.0877 | 779 | 0.20956 | 0.000472 | -1e-06 | 108.8059 | 109.9685 | 125.3619 | 21.068 | 0.002294 | 0.00212 | -152.169 | 0.23906 | 8.4028 | 0.003388 | 0.004883 | 0.004634 | 7.3e-05 | 0.001098 | 0.28224 | 1.3844 | 2.5481 | 100.57 | 20.76 | 19.361 | 76.8305 | 103.1752 | 72.6515 | 0.004723 | 1.0209 | 0.024746 | 24.2386 | -0.095069 | 0.14222 | -0.000153 | -0.011403 | -0.011403 | -0.51259
2019-03-20 18:02 | 1553119320.0867 | 780 | 0.20955 | 0.000474 | 0 | 108.8114 | 110.1513 | 125.3872 | 21.0675 | 0.002233 | 0.00216 | -152.1838 | 0.23926 | 8.4049 | 0.003428 | 0.005356 | 0.005166 | 7.2e-05 | 0.001103 | 0.28208 | 1.3611 | 2.5481 | 100.57 | 20.76 | 19.363 | 75.3691 | 104.1511 | 73.6121 | 0.004723 | 1.0181 | 0.024715 | 24.2747 | -0.10947 | 0.14233 | -9.4e-05 | -0.006994 | -0.006994 | -0.83642
2019-03-20 18:02 | 1553119321.0867 | 781 | 0.20953 | 0.000467 | 0 | 108.7849 | 109.9214 | 125.652 | 21.0664 | 0.002198 | 0.002053 | -152.1833 | 0.23835 | 8.4011 | 0.003374 | 0.004654 | 0.004402 | 7.1e-05 | 0.001103 | 0.2756 | 1.3007 | 2.5482 | 100.57 | 20.76 | 19.365 | 74.5278 | 102.8579 | 71.0417 | 0.004724 | 1.0257 | 0.024833 | 24.2101 | -0.092966 | 0.14179 | 6.4e-05 | 0.004822 | 0.004822 | 1.2086
2019-03-20 18:02 | 1553119322.0879 | 782 | 0.20956 | 0.000468 | 0 | 108.7169 | 109.9749 | 125.6544 | 21.0673 | 0.002202 | 0.002109 | -152.1864 | 0.2388 | 8.402 | 0.003406 | 0.004731 | 0.004513 | 7.3e-05 | 0.001102 | 0.28174 | 1.3407 | 2.5482 | 100.57 | 20.76 | 19.367 | 74.6075 | 103.6123 | 72.3842 | 0.004724 | 1.0217 | 0.024785 | 24.2575 | -0.1029 | 0.14206 | -0.000136 | -0.010149 | -0.010149 | -0.57528
2019-03-20 18:02 | 1553119323.087 | 783 | 0.20957 | 0.00046 | 0 | 108.7163 | 110.2201 | 125.399 | 21.0679 | 0.002338 | 0.002141 | -152.1919 | 0.2401 | 8.4009 | 0.003417 | 0.004645 | 0.004517 | 7.5e-05 | 0.001098 | 0.2636 | 1.3235 | 2.548 | 100.57 | 20.76 | 19.368 | 77.8756 | 103.8832 | 73.1512 | 0.004725 | 1.0195 | 0.024732 | 24.2597 | -0.12287 | 0.14283 | -0.000189 | -0.014103 | -0.014103 | -0.41626
2019-03-20 18:02 | 1553119324.088 | 784 | 0.20957 | 0.000459 | 0 | 108.6303 | 109.7701 | 125.6518 | 21.067 | 0.002203 | 0.002091 | -152.1838 | 0.23918 | 8.4018 | 0.003388 | 0.00503 | 0.004784 | 7.2e-05 | 0.001099 | 0.28313 | 1.3378 | 2.5481 | 100.57 | 20.76 | 19.368 | 74.6344 | 103.1901 | 71.9411 | 0.004725 | 1.0231 | 0.0248 | 24.2417 | -0.09336 | 0.14229 | -0.000179 | -0.013366 | -0.013366 | -0.43753
2019-03-20 18:02 | 1553119325.087 | 785 | 0.20961 | 0.000461 | 0 | 108.5792 | 109.985 | 125.6241 | 21.0678 | 0.002211 | 0.002061 | -152.1985 | 0.239 | 8.4031 | 0.003343 | 0.004927 | 0.004677 | 7.3e-05 | 0.001096 | 0.28926 | 1.3414 | 2.5481 | 100.57 | 20.76 | 19.367 | 74.8238 | 102.0855 | 71.2239 | 0.004724 | 1.0252 | 0.024824 | 24.2138 | -0.11506 | 0.14218 | -0.000401 | -0.030062 | -0.030062 | -0.19439
2019-03-20 18:02 | 1553119326.088 | 786 | 0.20956 | 0.000455 | 0 | 108.4608 | 109.8941 | 125.3041 | 21.0665 | 0.002264 | 0.002056 | -152.2058 | 0.23933 | 8.4021 | 0.003395 | 0.004101 | 0.003973 | 7.3e-05 | 0.00109 | 0.25426 | 1.3281 | 2.5481 | 100.57 | 20.76 | 19.367 | 76.1109 | 103.3562 | 71.0954 | 0.004724 | 1.0256 | 0.024797 | 24.1784 | -0.11743 | 0.14238 | -0.00013 | -0.009714 | -0.009714 | -0.6024
2019-03-20 18:02 | 1553119327.0875 | 787 | 0.20953 | 0.000454 | 0 | 108.4886 | 109.8702 | 125.3933 | 21.0678 | 0.002233 | 0.002124 | -152.204 | 0.23833 | 8.4033 | 0.003415 | 0.005109 | 0.004889 | 7.3e-05 | 0.001096 | 0.27776 | 1.3523 | 2.5481 | 100.57 | 20.76 | 19.365 | 75.3571 | 103.8425 | 72.7403 | 0.004724 | 1.0207 | 0.024746 | 24.2447 | -0.11319 | 0.14178 | 9.2e-05 | 0.006863 | 0.006863 | 0.84906
2019-03-20 18:02 | 1553119328.0862 | 788 | 0.20957 | 0.000452 | -1e-06 | 108.4278 | 109.6198 | 125.3718 | 21.0666 | 0.002216 | 0.002052 | -152.1822 | 0.2387 | 8.402 | 0.003387 | 0.004045 | 0.003817 | 7.2e-05 | 0.001096 | 0.27312 | 1.3232 | 2.5481 | 100.57 | 20.76 | 19.365 | 74.9504 | 103.1536 | 71.0076 | 0.004724 | 1.0258 | 0.024806 | 24.1819 | -0.097795 | 0.142 | -0.000175 | -0.013059 | -0.013059 | -0.44691
2019-03-20 18:02 | 1553119329.0862 | 789 | 0.20954 | 0.000452 | 0 | 108.4676 | 109.7607 | 125.256 | 21.0676 | 0.002246 | 0.002067 | -152.1561 | 0.23901 | 8.4034 | 0.003431 | 0.005025 | 0.004843 | 6.9e-05 | 0.001103 | 0.28102 | 1.2779 | 2.5481 | 100.57 | 20.77 | 19.365 | 75.6709 | 104.2171 | 71.3625 | 0.004727 | 1.0248 | 0.024782 | 24.1832 | -0.106 | 0.14219 | -4e-06 | -0.0003 | -0.0003 | -19.4604
2019-03-20 18:02 | 1553119330.0872 | 790 | 0.20957 | 0.000453 | 0 | 108.3512 | 109.5959 | 125.5823 | 21.0668 | 0.002224 | 0.002101 | -152.1541 | 0.23874 | 8.4035 | 0.003384 | 0.004653 | 0.004483 | 7.1e-05 | 0.001102 | 0.27367 | 1.2828 | 2.5481 | 100.57 | 20.77 | 19.365 | 75.1378 | 103.0886 | 72.1968 | 0.004727 | 1.0223 | 0.024784 | 24.2439 | -0.10217 | 0.14202 | -0.000139 | -0.01039 | -0.01039 | -0.56179
2019-03-20 18:02 | 1553119331.0863 | 791 | 0.20959 | 0.000449 | 0 | 108.367 | 109.6482 | 125.6626 | 21.0676 | 0.002209 | 0.002095 | -152.1507 | 0.239 | 8.4024 | 0.003392 | 0.005061 | 0.004855 | 7.3e-05 | 0.001106 | 0.25426 | 1.3088 | 2.5482 | 100.57 | 20.77 | 19.366 | 74.7902 | 103.2929 | 72.0462 | 0.004727 | 1.0227 | 0.024798 | 24.2464 | -0.10513 | 0.14218 | -0.000274 | -0.0205 | -0.0205 | -0.28505
2019-03-20 18:02 | 1553119332.0867 | 792 | 0.20955 | 0.000448 | 0 | 108.2962 | 109.5118 | 125.537 | 21.0672 | 0.002175 | 0.002056 | -152.1483 | 0.23835 | 8.4037 | 0.003469 | 0.005475 | 0.005312 | 7.2e-05 | 0.001112 | 0.28584 | 1.3126 | 2.548 | 100.57 | 20.77 | 19.366 | 73.9713 | 105.1455 | 71.1033 | 0.004727 | 1.0255 | 0.024819 | 24.2012 | -0.099842 | 0.1418 | -3.7e-05 | -0.002732 | -0.002732 | -2.1331
2019-03-20 18:02 | 1553119333.0875 | 793 | 0.20956 | 0.000448 | 0 | 108.2774 | 109.7121 | 125.6142 | 21.0677 | 0.002199 | 0.002072 | -152.1314 | 0.23895 | 8.4046 | 0.003443 | 0.004534 | 0.004274 | 6.6e-05 | 0.001113 | 0.27345 | 1.3212 | 2.548 | 100.57 | 20.77 | 19.365 | 74.5361 | 104.5286 | 71.4991 | 0.004727 | 1.0244 | 0.024813 | 24.2225 | -0.11774 | 0.14215 | -0.000104 | -0.007833 | -0.007833 | -0.74589
2019-03-20 18:02 | 1553119334.0865 | 794 | 0.20958 | 0.000443 | -1e-06 | 108.1895 | 109.52 | 125.7485 | 21.0673 | 0.002235 | 0.002076 | -152.127 | 0.23894 | 8.4022 | 0.003361 | 0.00448 | 0.004271 | 7.3e-05 | 0.001109 | 0.30718 | 1.333 | 2.5481 | 100.57 | 20.77 | 19.365 | 75.3987 | 102.5326 | 71.5867 | 0.004727 | 1.0241 | 0.024823 | 24.2385 | -0.10933 | 0.14215 | -0.000204 | -0.015273 | -0.015273 | -0.38252
2019-03-20 18:02 | 1553119335.0875 | 795 | 0.20959 | 0.000449 | 0 | 108.1312 | 109.3406 | 125.7102 | 21.0671 | 0.002259 | 0.002038 | -152.1214 | 0.23821 | 8.4052 | 0.003378 | 0.00544 | 0.005262 | 7.3e-05 | 0.001116 | 0.27975 | 1.318 | 2.548 | 100.57 | 20.77 | 19.361 | 75.9744 | 102.9475 | 70.6767 | 0.004726 | 1.0268 | 0.024852 | 24.2028 | -0.099483 | 0.14171 | -0.000305 | -0.022834 | -0.022834 | -0.25507
2019-03-20 18:02 | 1553119336.0865 | 796 | 0.20956 | 0.000438 | 0 | 108.1212 | 109.4157 | 125.9926 | 21.0671 | 0.002117 | 0.002088 | -152.1387 | 0.23946 | 8.4042 | 0.003465 | 0.004676 | 0.004422 | 6.9e-05 | 0.001116 | 0.29999 | 1.3329 | 2.5481 | 100.57 | 20.77 | 19.36 | 72.5718 | 105.0479 | 71.8633 | 0.004726 | 1.0233 | 0.024837 | 24.2718 | -0.10646 | 0.14245 | -6.7e-05 | -0.004996 | -0.004996 | -1.172
2019-03-20 18:02 | 1553119337.088 | 797 | 0.20954 | 0.000445 | 0 | 108.02 | 109.212 | 125.94 | 21.0678 | 0.00215 | 0.002052 | -152.1225 | 0.23921 | 8.4034 | 0.003403 | 0.004878 | 0.004694 | 7.1e-05 | 0.001115 | 0.27865 | 1.3511 | 2.5482 | 100.57 | 20.78 | 19.354 | 73.3576 | 103.5431 | 70.9987 | 0.004727 | 1.0259 | 0.024863 | 24.2363 | -0.098167 | 0.1423 | 1.3e-05 | 0.00095 | 0.00095 | 6.1569
2019-03-20 18:02 | 1553119338.0866 | 798 | 0.20958 | 0.000441 | -1e-06 | 108.0843 | 109.3313 | 125.9149 | 21.0662 | 0.002171 | 0.002032 | -152.1386 | 0.23904 | 8.4005 | 0.00327 | 0.005151 | 0.00489 | 7.2e-05 | 0.001112 | 0.27113 | 1.3267 | 2.5482 | 100.57 | 20.78 | 19.351 | 73.872 | 100.3246 | 70.5254 | 0.004726 | 1.0273 | 0.024877 | 24.2172 | -0.1026 | 0.1422 | -0.000184 | -0.01377 | -0.01377 | -0.42443
2019-03-20 18:02 | 1553119339.0867 | 799 | 0.20959 | 0.000444 | 0 | 108.0232 | 109.4019 | 126.1975 | 21.067 | 0.002125 | 0.002099 | -152.1518 | 0.23936 | 8.4032 | 0.003315 | 0.00436 | 0.004091 | 7.5e-05 | 0.001109 | 0.26593 | 1.3059 | 2.548 | 100.57 | 20.78 | 19.347 | 72.7541 | 101.407 | 72.1275 | 0.004725 | 1.0225 | 0.024848 | 24.3008 | -0.11344 | 0.14239 | -0.000252 | -0.018877 | -0.018877 | -0.31002
2019-03-20 18:02 | 1553119340.0872 | 800 | 0.20955 | 0.000443 | 0 | 108.0096 | 109.2281 | 126.1691 | 21.067 | 0.002098 | 0.002053 | -152.1353 | 0.23947 | 8.4025 | 0.003309 | 0.004788 | 0.004598 | 6.7e-05 | 0.001108 | 0.27057 | 1.34 | 2.548 | 100.57 | 20.78 | 19.337 | 72.1064 | 101.272 | 71.0407 | 0.004723 | 1.0257 | 0.024884 | 24.2598 | -0.10034 | 0.14246 | -2.7e-05 | -0.002039 | -0.002039 | -2.8722
2019-03-20 18:02 | 1553119341.0868 | 801 | 0.20956 | 0.000445 | 0 | 107.9731 | 109.1926 | 126.1142 | 21.0672 | 0.002178 | 0.002023 | -152.1352 | 0.23881 | 8.4012 | 0.003364 | 0.004336 | 0.00415 | 6.9e-05 | 0.001105 | 0.28102 | 1.3085 | 2.548 | 100.58 | 20.78 | 19.331 | 74.0297 | 102.6144 | 70.3065 | 0.004721 | 1.0279 | 0.024905 | 24.2286 | -0.10046 | 0.14207 | -8.2e-05 | -0.006197 | -0.006197 | -0.94223
2019-03-20 18:02 | 1553119342.0869 | 802 | 0.20952 | 0.000447 | 0 | 107.9213 | 109.0551 | 126.459 | 21.067 | 0.002163 | 0.00205 | -152.1369 | 0.23869 | 8.402 | 0.003312 | 0.004312 | 0.004095 | 6.9e-05 | 0.001106 | 0.28982 | 1.3 | 2.5481 | 100.58 | 20.78 | 19.324 | 73.6725 | 101.3364 | 70.9573 | 0.004719 | 1.026 | 0.024916 | 24.2847 | -0.093473 | 0.14199 | 0.000127 | 0.009567 | 0.009567 | 0.61004
2019-03-20 18:02 | 1553119343.0879 | 803 | 0.20953 | 0.000442 | 0 | 107.8483 | 109.0723 | 126.52 | 21.0671 | 0.002218 | 0.002015 | -152.1705 | 0.23924 | 8.401 | 0.003442 | 0.005023 | 0.004746 | 6.8e-05 | 0.001106 | 0.27577 | 1.3427 | 2.5481 | 100.57 | 20.78 | 19.305 | 75.0079 | 104.4888 | 70.1247 | 0.004715 | 1.0285 | 0.024952 | 24.2612 | -0.10094 | 0.14232 | 9.8e-05 | 0.007365 | 0.007365 | 0.79427
2019-03-20 18:02 | 1553119344.0873 | 804 | 0.20954 | 0.000446 | 0 | 107.8631 | 108.9672 | 126.1327 | 21.067 | 0.00213 | 0.002008 | -152.1789 | 0.23878 | 8.4025 | 0.003477 | 0.00465 | 0.004487 | 7.2e-05 | 0.00111 | 0.25802 | 1.3585 | 2.5482 | 100.57 | 20.78 | 19.296 | 72.8884 | 105.3373 | 69.9515 | 0.004713 | 1.029 | 0.02492 | 24.2179 | -0.091093 | 0.14205 | 1.5e-05 | 0.00111 | 0.00111 | 5.2616
2019-03-20 18:02 | 1553119345.0865 | 805 | 0.20953 | 0.00044 | 0 | 107.878 | 108.9934 | 126.044 | 21.0664 | 0.002127 | 0.002051 | -152.1895 | 0.23834 | 8.4007 | 0.003496 | 0.004956 | 0.004749 | 7e-05 | 0.001108 | 0.27616 | 1.3733 | 2.5481 | 100.57 | 20.78 | 19.274 | 72.8199 | 105.8111 | 70.9769 | 0.004707 | 1.0259 | 0.024874 | 24.2455 | -0.092006 | 0.14179 | 0.000116 | 0.008738 | 0.008738 | 0.66689
2019-03-20 18:02 | 1553119346.0862 | 806 | 0.20955 | 0.000444 | -1e-06 | 107.8534 | 109.0188 | 126.0632 | 21.0661 | 0.002177 | 0.002057 | -152.1486 | 0.23859 | 8.4052 | 0.003397 | 0.004949 | 0.004732 | 6.9e-05 | 0.001108 | 0.28136 | 1.4076 | 2.5481 | 100.57 | 20.78 | 19.264 | 74.0045 | 103.4076 | 71.1179 | 0.004705 | 1.0255 | 0.024871 | 24.2523 | -0.09613 | 0.14194 | -2.3e-05 | -0.001724 | -0.001724 | -3.3833
2019-03-20 18:02 | 1553119347.0871 | 807 | 0.20954 | 0.000442 | 0 | 107.8124 | 108.9811 | 125.8126 | 21.0663 | 0.002176 | 0.002048 | -152.1682 | 0.23869 | 8.402 | 0.003331 | 0.005429 | 0.005163 | 6.9e-05 | 0.001106 | 0.28672 | 1.3583 | 2.5481 | 100.57 | 20.78 | 19.254 | 73.9978 | 101.7975 | 70.9036 | 0.004703 | 1.0261 | 0.024854 | 24.2207 | -0.096439 | 0.142 | 5.8e-05 | 0.004331 | 0.004331 | 1.3474
2019-03-20 18:02 | 1553119348.0862 | 808 | 0.20952 | 0.000444 | 0 | 107.7693 | 109.1354 | 125.7982 | 21.0671 | 0.002197 | 0.002074 | -152.1625 | 0.23811 | 8.4035 | 0.003419 | 0.005097 | 0.00498 | 6.8e-05 | 0.001103 | 0.28694 | 1.3108 | 2.5481 | 100.57 | 20.78 | 19.233 | 74.4817 | 103.9431 | 71.5296 | 0.004697 | 1.0243 | 0.02483 | 24.2413 | -0.11267 | 0.14165 | 0.000181 | 0.013586 | 0.013586 | 0.42851
2019-03-20 18:02 | 1553119349.0862 | 809 | 0.20952 | 0.000441 | 0 | 107.6791 | 108.9538 | 125.79 | 21.0664 | 0.002114 | 0.002013 | -152.1521 | 0.2388 | 8.403 | 0.003444 | 0.004259 | 0.004045 | 6.8e-05 | 0.001102 | 0.28031 | 1.2925 | 2.548 | 100.57 | 20.78 | 19.223 | 72.4965 | 104.548 | 70.081 | 0.004695 | 1.0286 | 0.024881 | 24.1895 | -0.10526 | 0.14206 | 0.000152 | 0.011436 | 0.011436 | 0.51056
2019-03-20 18:02 | 1553119350.0874 | 810 | 0.20954 | 0.000444 | 0 | 107.8146 | 108.9444 | 125.5241 | 21.0663 | 0.002111 | 0.002062 | -152.146 | 0.239 | 8.4019 | 0.00337 | 0.004865 | 0.004671 | 6.3e-05 | 0.001101 | 0.28655 | 1.319 | 2.5483 | 100.57 | 20.78 | 19.202 | 72.4168 | 102.7416 | 71.2397 | 0.00469 | 1.0251 | 0.024813 | 24.2047 | -0.093245 | 0.14218 | 2.6e-05 | 0.001953 | 0.001953 | 2.9917
2019-03-20 18:02 | 1553119351.0874 | 811 | 0.20952 | 0.000448 | 0 | 107.7347 | 108.9956 | 125.4729 | 21.0661 | 0.002125 | 0.002043 | -152.1429 | 0.23908 | 8.4023 | 0.003372 | 0.005286 | 0.004986 | 6.6e-05 | 0.001106 | 0.28495 | 1.2844 | 2.548 | 100.57 | 20.78 | 19.191 | 72.7668 | 102.8051 | 70.7873 | 0.004687 | 1.0265 | 0.024824 | 24.1839 | -0.10408 | 0.14223 | 0.000161 | 0.012069 | 0.012069 | 0.48433
2019-03-20 18:03 | 1553119352.0876 | 812 | 0.20953 | 0.000448 | -1e-06 | 107.6843 | 108.7449 | 125.3426 | 21.0669 | 0.002082 | 0.002046 | -152.1507 | 0.2386 | 8.4026 | 0.003386 | 0.004707 | 0.004493 | 6.8e-05 | 0.001104 | 0.30215 | 1.2992 | 2.548 | 100.57 | 20.78 | 19.181 | 71.7398 | 103.1294 | 70.8675 | 0.004685 | 1.0262 | 0.024809 | 24.1741 | -0.087663 | 0.14194 | 9.1e-05 | 0.006874 | 0.006874 | 0.84864
2019-03-20 18:03 | 1553119353.0875 | 813 | 0.20953 | 0.000448 | 0 | 107.7163 | 108.8491 | 125.1816 | 21.0663 | 0.0021 | 0.002053 | -152.1359 | 0.23851 | 8.4006 | 0.00337 | 0.004911 | 0.004712 | 7.1e-05 | 0.001104 | 0.28224 | 1.3496 | 2.5481 | 100.57 | 20.78 | 19.163 | 72.1556 | 102.7517 | 71.033 | 0.00468 | 1.0257 | 0.024787 | 24.1644 | -0.093577 | 0.14189 | 9.3e-05 | 0.007018 | 0.007018 | 0.83091
2019-03-20 18:03 | 1553119354.0875 | 814 | 0.20951 | 0.000461 | 0 | 107.6461 | 108.7621 | 125.3605 | 21.0669 | 0.002085 | 0.002054 | -152.1255 | 0.23825 | 8.405 | 0.003441 | 0.005119 | 0.004929 | 7.1e-05 | 0.001105 | 0.27577 | 1.3588 | 2.548 | 100.57 | 20.78 | 19.157 | 71.8006 | 104.4713 | 71.068 | 0.004679 | 1.0256 | 0.024803 | 24.1829 | -0.092248 | 0.14173 | 0.000172 | 0.012914 | 0.012914 | 0.45108
2019-03-20 18:03 | 1553119355.0876 | 815 | 0.20953 | 0.000456 | 0 | 107.6384 | 108.9018 | 125.3506 | 21.0664 | 0.002101 | 0.002063 | -152.1125 | 0.23839 | 8.4017 | 0.003487 | 0.004334 | 0.004073 | 6.8e-05 | 0.001109 | 0.26272 | 1.3879 | 2.5479 | 100.57 | 20.79 | 19.145 | 72.1964 | 105.5948 | 71.2701 | 0.004679 | 1.025 | 0.024795 | 24.189 | -0.10438 | 0.14182 | 9.9e-05 | 0.007358 | 0.007358 | 0.79214
2019-03-20 18:03 | 1553119356.0867 | 816 | 0.20954 | 0.000463 | 0 | 107.584 | 108.7426 | 125.0744 | 21.0665 | 0.002177 | 0.002045 | -152.1365 | 0.2388 | 8.4043 | 0.003412 | 0.004273 | 0.004131 | 7.2e-05 | 0.00111 | 0.27455 | 1.3792 | 2.5481 | 100.57 | 20.78 | 19.139 | 74.0011 | 103.7782 | 70.8297 | 0.004674 | 1.0264 | 0.024783 | 24.1469 | -0.09581 | 0.14206 | 1.7e-05 | 0.001258 | 0.001258 | 4.6412
2019-03-20 18:03 | 1553119357.0867 | 817 | 0.2096 | 0.000462 | 0 | 107.5229 | 108.7561 | 125.2617 | 21.0662 | 0.002058 | 0.002022 | -152.1465 | 0.23869 | 8.4009 | 0.003401 | 0.004886 | 0.004649 | 7.4e-05 | 0.001114 | 0.28888 | 1.3819 | 2.5482 | 100.57 | 20.78 | 19.134 | 71.149 | 103.4918 | 70.285 | 0.004673 | 1.028 | 0.024822 | 24.1459 | -0.102 | 0.14199 | -0.000372 | -0.02787 | -0.02787 | -0.2094
2019-03-20 18:03 | 1553119358.0877 | 818 | 0.20958 | 0.00046 | 0 | 107.575 | 108.8435 | 125.2344 | 21.0667 | 0.00212 | 0.002044 | -152.1516 | 0.23949 | 8.3997 | 0.003434 | 0.005031 | 0.004791 | 7.2e-05 | 0.001116 | 0.27577 | 1.3554 | 2.5481 | 100.57 | 20.78 | 19.126 | 72.6447 | 104.2957 | 70.8118 | 0.004671 | 1.0264 | 0.0248 | 24.1617 | -0.10486 | 0.14247 | -0.000233 | -0.017462 | -0.017462 | -0.33534
2019-03-20 18:03 | 1553119359.0861 | 819 | 0.2096 | 0.00046 | -1e-06 | 107.5197 | 108.6731 | 124.8887 | 21.0664 | 0.002054 | 0.002028 | -152.1599 | 0.23904 | 8.4033 | 0.0034 | 0.005006 | 0.004884 | 7.3e-05 | 0.001116 | 0.28208 | 1.3548 | 2.5481 | 100.57 | 20.78 | 19.124 | 71.0449 | 103.4762 | 70.4273 | 0.004671 | 1.0276 | 0.024779 | 24.1149 | -0.095443 | 0.1422 | -0.000366 | -0.027324 | -0.027324 | -0.2139
2019-03-20 18:03 | 1553119360.0872 | 820 | 0.20951 | 0.000461 | -1e-06 | 107.4272 | 108.6372 | 124.9739 | 21.0667 | 0.002068 | 0.002025 | -152.1759 | 0.23883 | 8.4033 | 0.00352 | 0.005309 | 0.005069 | 7.3e-05 | 0.001117 | 0.27417 | 1.3666 | 2.5481 | 100.57 | 20.78 | 19.124 | 71.3836 | 106.3756 | 70.363 | 0.004671 | 1.0277 | 0.02479 | 24.1208 | -0.10019 | 0.14208 | 0.000213 | 0.01599 | 0.01599 | 0.36519
2019-03-20 18:03 | 1553119361.0869 | 821 | 0.20958 | 0.000457 | -1e-06 | 107.4107 | 108.555 | 125.0624 | 21.067 | 0.002125 | 0.002 | -152.1867 | 0.23961 | 8.4038 | 0.003538 | 0.004835 | 0.004557 | 7e-05 | 0.001122 | 0.28368 | 1.3626 | 2.5481 | 100.57 | 20.78 | 19.126 | 72.7697 | 106.823 | 69.7667 | 0.004671 | 1.0295 | 0.024821 | 24.1084 | -0.094788 | 0.14255 | -0.000229 | -0.01711 | -0.01711 | -0.34242
2019-03-20 18:03 | 1553119362.0874 | 822 | 0.20954 | 0.000455 | 0 | 107.3767 | 108.5393 | 124.7864 | 21.0669 | 0.002088 | 0.002036 | -152.152 | 0.23957 | 8.4035 | 0.003481 | 0.005462 | 0.005167 | 7.1e-05 | 0.001121 | 0.27544 | 1.3518 | 2.548 | 100.57 | 20.78 | 19.129 | 71.8756 | 105.4434 | 70.632 | 0.004672 | 1.0269 | 0.024762 | 24.1122 | -0.096323 | 0.14252 | -4e-06 | -0.000317 | -0.000317 | -18.4913
2019-03-20 18:03 | 1553119363.0861 | 823 | 0.20955 | 0.000457 | -1e-06 | 107.3004 | 108.5344 | 125.0685 | 21.0669 | 0.002083 | 0.002032 | -152.1466 | 0.23897 | 8.4047 | 0.00357 | 0.005427 | 0.005133 | 7e-05 | 0.001124 | 0.29407 | 1.3601 | 2.5481 | 100.57 | 20.78 | 19.136 | 71.7617 | 107.5979 | 70.5385 | 0.004674 | 1.0272 | 0.024793 | 24.1361 | -0.10228 | 0.14216 | -6.8e-05 | -0.005034 | -0.005034 | -1.1607
2019-03-20 18:03 | 1553119364.087 | 824 | 0.20959 | 0.000451 | 0 | 107.3114 | 108.2597 | 124.9892 | 21.0658 | 0.002129 | 0.002051 | -152.1658 | 0.23879 | 8.4035 | 0.003524 | 0.004918 | 0.004745 | 7e-05 | 0.001119 | 0.28368 | 1.3451 | 2.5481 | 100.57 | 20.78 | 19.14 | 72.853 | 106.4764 | 70.9896 | 0.004675 | 1.0259 | 0.024769 | 24.1443 | -0.078698 | 0.14205 | -0.000263 | -0.019636 | -0.019636 | -0.29734
2019-03-20 18:03 | 1553119365.0864 | 825 | 0.20956 | 0.000454 | -1e-06 | 107.3518 | 108.5184 | 124.8014 | 21.0669 | 0.002164 | 0.001981 | -152.1708 | 0.23842 | 8.4009 | 0.003473 | 0.005103 | 0.004889 | 6.8e-05 | 0.00112 | 0.28766 | 1.319 | 2.5481 | 100.57 | 20.78 | 19.142 | 73.7021 | 105.2342 | 69.3073 | 0.004675 | 1.0309 | 0.024811 | 24.0671 | -0.096673 | 0.14183 | -0.00012 | -0.008943 | -0.008943 | -0.65185
2019-03-20 18:03 | 1553119366.0872 | 826 | 0.20957 | 0.000453 | 0 | 107.2166 | 108.3543 | 125.1563 | 21.067 | 0.002117 | 0.002008 | -152.1605 | 0.23859 | 8.4015 | 0.003504 | 0.005069 | 0.004817 | 6.9e-05 | 0.001115 | 0.27489 | 1.3194 | 2.5481 | 100.57 | 20.78 | 19.145 | 72.5672 | 106.0031 | 69.9443 | 0.004676 | 1.029 | 0.024823 | 24.1237 | -0.094409 | 0.14194 | -0.000128 | -0.009579 | -0.009579 | -0.60899
2019-03-20 18:03 | 1553119367.0871 | 827 | 0.20956 | 0.000447 | 0 | 107.1743 | 108.1521 | 125.1154 | 21.0665 | 0.002084 | 0.002044 | -152.1429 | 0.23848 | 8.4031 | 0.003428 | 0.005209 | 0.004985 | 6.8e-05 | 0.001113 | 0.31542 | 1.2849 | 2.5482 | 100.57 | 20.78 | 19.146 | 71.7755 | 104.1565 | 70.8071 | 0.004676 | 1.0264 | 0.024788 | 24.1501 | -0.081238 | 0.14187 | -0.000118 | -0.008811 | -0.008811 | -0.66179
2019-03-20 18:03 | 1553119368.0863 | 828 | 0.20957 | 0.000447 | 0 | 107.2127 | 108.3509 | 125.0009 | 21.067 | 0.002059 | 0.001972 | -152.1284 | 0.23908 | 8.4025 | 0.003378 | 0.004868 | 0.004563 | 6.9e-05 | 0.001113 | 0.26969 | 1.336 | 2.5482 | 100.57 | 20.78 | 19.146 | 71.1741 | 102.9443 | 69.0743 | 0.004676 | 1.0316 | 0.02484 | 24.0782 | -0.094455 | 0.14223 | -0.000166 | -0.01247 | -0.01247 | -0.46877
2019-03-20 18:03 | 1553119369.0864 | 829 | 0.20954 | 0.000449 | 0 | 107.2053 | 108.3883 | 125.6267 | 21.0662 | 0.002097 | 0.001937 | -152.1256 | 0.23883 | 8.4014 | 0.003393 | 0.004932 | 0.004788 | 7e-05 | 0.001112 | 0.27936 | 1.3207 | 2.5481 | 100.57 | 20.78 | 19.144 | 72.0962 | 103.3135 | 68.2434 | 0.004676 | 1.0341 | 0.024932 | 24.109 | -0.098161 | 0.14208 | 2.5e-05 | 0.001905 | 0.001905 | 3.0661
2019-03-20 18:03 | 1553119370.0868 | 830 | 0.20957 | 0.000445 | 0 | 107.2098 | 108.2893 | 125.5424 | 21.0664 | 0.002125 | 0.002015 | -152.1226 | 0.23855 | 8.4003 | 0.003315 | 0.004623 | 0.004351 | 6.9e-05 | 0.00111 | 0.28622 | 1.3037 | 2.5479 | 100.57 | 20.78 | 19.142 | 72.7576 | 101.4134 | 70.1192 | 0.004675 | 1.0285 | 0.024855 | 24.1671 | -0.089606 | 0.14191 | -0.000153 | -0.011444 | -0.011444 | -0.50967
2019-03-20 18:03 | 1553119371.0865 | 831 | 0.20956 | 0.000444 | 0 | 107.2182 | 108.4294 | 125.4682 | 21.0662 | 0.002106 | 0.001989 | -152.1264 | 0.23861 | 8.4024 | 0.003378 | 0.005131 | 0.004879 | 7.1e-05 | 0.001108 | 0.29767 | 1.2914 | 2.5479 | 100.57 | 20.78 | 19.133 | 72.3129 | 102.948 | 69.4929 | 0.004673 | 1.0304 | 0.024871 | 24.1379 | -0.10047 | 0.14195 | -7.3e-05 | -0.005443 | -0.005443 | -1.0719
2019-03-20 18:03 | 1553119372.0865 | 832 | 0.20957 | 0.000439 | -1e-06 | 107.3198 | 108.2743 | 125.8682 | 21.0668 | 0.002081 | 0.002008 | -152.1612 | 0.23822 | 8.3994 | 0.003441 | 0.004711 | 0.00454 | 6.8e-05 | 0.001107 | 0.28744 | 1.286 | 2.5481 | 100.57 | 20.78 | 19.127 | 71.7165 | 104.4626 | 69.9403 | 0.004672 | 1.029 | 0.024894 | 24.1921 | -0.079202 | 0.14172 | -0.000127 | -0.009471 | -0.009471 | -0.61503
2019-03-20 18:03 | 1553119373.0866 | 833 | 0.20952 | 0.000445 | 0 | 107.1552 | 108.3861 | 125.475 | 21.0664 | 0.002116 | 0.001974 | -152.1553 | 0.23829 | 8.4051 | 0.003422 | 0.004651 | 0.004418 | 6.7e-05 | 0.001106 | 0.28799 | 1.3477 | 2.5482 | 100.57 | 20.78 | 19.111 | 72.5379 | 104.0163 | 69.1345 | 0.004668 | 1.0314 | 0.024884 | 24.1259 | -0.10216 | 0.14176 | 0.000146 | 0.010966 | 0.010966 | 0.53131
2019-03-20 18:03 | 1553119374.088 | 834 | 0.20956 | 0.000442 | -1e-06 | 107.127 | 108.1794 | 125.4391 | 21.0669 | 0.002055 | 0.001975 | -152.1417 | 0.23878 | 8.4014 | 0.00333 | 0.004853 | 0.004675 | 6.6e-05 | 0.001109 | 0.29302 | 1.372 | 2.5481 | 100.57 | 20.78 | 19.103 | 71.0843 | 101.7714 | 69.1614 | 0.004666 | 1.0314 | 0.02488 | 24.1234 | -0.087435 | 0.14205 | -7.1e-05 | -0.00527 | -0.00527 | -1.1078
2019-03-20 18:03 | 1553119375.0876 | 835 | 0.20954 | 0.000445 | 0 | 107.2195 | 108.3726 | 125.5795 | 21.0671 | 0.00209 | 0.002002 | -152.1471 | 0.23902 | 8.4008 | 0.003427 | 0.004806 | 0.004581 | 6.7e-05 | 0.001102 | 0.26753 | 1.3753 | 2.5481 | 100.57 | 20.78 | 19.096 | 71.9266 | 104.1209 | 69.8128 | 0.004664 | 1.0294 | 0.02487 | 24.1598 | -0.095678 | 0.14219 | 6e-05 | 0.004536 | 0.004536 | 1.2885
2019-03-20 18:03 | 1553119376.0857 | 836 | 0.20955 | 0.000451 | 0 | 107.1565 | 108.3472 | 125.2412 | 21.0672 | 0.002034 | 0.001952 | -152.151 | 0.23871 | 8.4052 | 0.003426 | 0.004681 | 0.004487 | 6.8e-05 | 0.0011 | 0.27201 | 1.3443 | 2.548 | 100.57 | 20.78 | 19.081 | 70.5879 | 104.1186 | 68.6057 | 0.00466 | 1.033 | 0.02488 | 24.0848 | -0.098843 | 0.14201 | -6e-06 | -0.000475 | -0.000475 | -12.2927
2019-03-20 18:03 | 1553119377.0867 | 837 | 0.20954 | 0.00044 | 0 | 107.1351 | 108.262 | 125.3311 | 21.0678 | 0.002056 | 0.001958 | -152.155 | 0.23953 | 8.4023 | 0.003414 | 0.005051 | 0.004841 | 6.9e-05 | 0.001101 | 0.27312 | 1.3329 | 2.548 | 100.57 | 20.78 | 19.075 | 71.1022 | 103.825 | 68.7418 | 0.004659 | 1.0326 | 0.024884 | 24.0982 | -0.093587 | 0.1425 | 6.1e-05 | 0.004612 | 0.004612 | 1.2699
2019-03-20 18:03 | 1553119378.0858 | 838 | 0.20953 | 0.000446 | 0 | 107.0242 | 108.1663 | 125.71 | 21.0677 | 0.002035 | 0.00207 | -152.1704 | 0.2389 | 8.4037 | 0.003395 | 0.004652 | 0.004435 | 6.9e-05 | 0.001102 | 0.3183 | 1.3639 | 2.5481 | 100.57 | 20.78 | 19.07 | 70.5907 | 103.3542 | 71.4498 | 0.004658 | 1.0245 | 0.024824 | 24.23 | -0.094945 | 0.14212 | 6.5e-05 | 0.00489 | 0.00489 | 1.1945
2019-03-20 18:03 | 1553119379.0861 | 839 | 0.20955 | 0.000452 | 0 | 106.9899 | 107.9413 | 125.3889 | 21.0671 | 0.002111 | 0.001977 | -152.1827 | 0.23872 | 8.4047 | 0.003348 | 0.004816 | 0.004626 | 6.8e-05 | 0.0011 | 0.28871 | 1.3718 | 2.5481 | 100.57 | 20.78 | 19.062 | 72.4183 | 102.2254 | 69.2137 | 0.004656 | 1.0312 | 0.024873 | 24.1204 | -0.079185 | 0.14201 | -1.3e-05 | -0.000935 | -0.000935 | -6.2454
2019-03-20 18:03 | 1553119380.0863 | 840 | 0.20952 | 0.000447 | 0 | 107.0025 | 108.0048 | 125.5825 | 21.0668 | 0.002093 | 0.001982 | -152.1721 | 0.23783 | 8.4046 | 0.003348 | 0.004271 | 0.00406 | 6.9e-05 | 0.001094 | 0.28495 | 1.3442 | 2.5482 | 100.57 | 20.78 | 19.061 | 72.001 | 102.208 | 69.3236 | 0.004655 | 1.0309 | 0.024888 | 24.1429 | -0.083395 | 0.14149 | 0.00018 | 0.013532 | 0.013532 | 0.42974
2019-03-20 18:03 | 1553119381.0872 | 841 | 0.20955 | 0.000446 | 0 | 106.9333 | 108.0101 | 125.5917 | 21.0672 | 0.002137 | 0.001994 | -152.181 | 0.23878 | 8.4018 | 0.003369 | 0.004597 | 0.004372 | 6.9e-05 | 0.001093 | 0.2912 | 1.3561 | 2.5482 | 100.57 | 20.78 | 19.057 | 73.06 | 102.7173 | 69.6117 | 0.004654 | 1.03 | 0.024879 | 24.1539 | -0.089615 | 0.14205 | -5e-06 | -0.000387 | -0.000387 | -15.0874
2019-03-20 18:03 | 1553119382.0871 | 842 | 0.20952 | 0.000445 | 0 | 106.8757 | 108.1121 | 125.5344 | 21.0679 | 0.002113 | 0.002004 | -152.1721 | 0.23859 | 8.4028 | 0.003437 | 0.004913 | 0.004741 | 7e-05 | 0.001089 | 0.27632 | 1.3623 | 2.548 | 100.57 | 20.78 | 19.055 | 72.4705 | 104.362 | 69.8541 | 0.004654 | 1.0293 | 0.024864 | 24.157 | -0.10288 | 0.14194 | 0.000158 | 0.011846 | 0.011846 | 0.49243
2019-03-20 18:03 | 1553119383.0871 | 843 | 0.20953 | 0.000445 | 0 | 106.8806 | 108.0258 | 125.8344 | 21.0668 | 0.002101 | 0.00199 | -152.1932 | 0.239 | 8.4035 | 0.003365 | 0.005231 | 0.004958 | 6.8e-05 | 0.001088 | 0.27223 | 1.3533 | 2.5482 | 100.57 | 20.78 | 19.055 | 72.1763 | 102.6241 | 69.5115 | 0.004654 | 1.0303 | 0.024906 | 24.1737 | -0.095326 | 0.14218 | 8.7e-05 | 0.006539 | 0.006539 | 0.89366
2019-03-20 18:03 | 1553119384.0861 | 844 | 0.20955 | 0.000447 | 0 | 106.857 | 108.0074 | 125.9992 | 21.0668 | 0.002047 | 0.001952 | -152.2059 | 0.23822 | 8.4009 | 0.003416 | 0.004098 | 0.003989 | 7.3e-05 | 0.00109 | 0.28998 | 1.342 | 2.548 | 100.57 | 20.78 | 19.054 | 70.897 | 103.8548 | 68.6168 | 0.004654 | 1.033 | 0.024955 | 24.1579 | -0.095785 | 0.14172 | 2e-06 | 0.000196 | 0.000196 | 29.7386
2019-03-20 18:03 | 1553119385.0866 | 845 | 0.20955 | 0.000443 | 0 | 106.8272 | 107.695 | 126.044 | 21.067 | 0.002066 | 0.001991 | -152.2212 | 0.23889 | 8.4024 | 0.003491 | 0.005003 | 0.004797 | 7.1e-05 | 0.00109 | 0.29087 | 1.3486 | 2.5479 | 100.57 | 20.78 | 19.054 | 71.3433 | 105.6788 | 69.5389 | 0.004654 | 1.0302 | 0.024926 | 24.1948 | -0.072363 | 0.14211 | -1.9e-05 | -0.001415 | -0.001415 | -4.1276
2019-03-20 18:03 | 1553119386.0863 | 846 | 0.20955 | 0.000445 | 0 | 106.6756 | 107.5836 | 126.2376 | 21.0662 | 0.002068 | 0.002037 | -152.2219 | 0.23921 | 8.4011 | 0.003391 | 0.004753 | 0.004536 | 7e-05 | 0.001085 | 0.28943 | 1.3415 | 2.5479 | 100.57 | 20.78 | 19.053 | 71.404 | 103.2523 | 70.6478 | 0.004653 | 1.0269 | 0.024905 | 24.2525 | -0.075817 | 0.1423 | -2.4e-05 | -0.001782 | -0.001782 | -3.2815
2019-03-20 18:03 | 1553119387.0876 | 847 | 0.2095 | 0.000446 | 0 | 106.5953 | 107.8097 | 126.3057 | 21.0676 | 0.002053 | 0.001982 | -152.2164 | 0.23885 | 8.4041 | 0.003421 | 0.005215 | 0.004944 | 7e-05 | 0.001087 | 0.28888 | 1.2978 | 2.5481 | 100.57 | 20.78 | 19.052 | 71.0388 | 103.9907 | 69.325 | 0.004653 | 1.0309 | 0.02496 | 24.2124 | -0.10132 | 0.14209 | 0.000298 | 0.022426 | 0.022426 | 0.26041
2019-03-20 18:03 | 1553119388.0864 | 848 | 0.20953 | 0.000445 | 0 | 106.5782 | 107.5855 | 126.0489 | 21.0658 | 0.002103 | 0.001994 | -152.2228 | 0.24002 | 8.4025 | 0.003371 | 0.005206 | 0.005092 | 6.8e-05 | 0.001082 | 0.28064 | 1.3822 | 2.5481 | 100.57 | 20.78 | 19.051 | 72.2402 | 102.7842 | 69.6228 | 0.004653 | 1.03 | 0.024923 | 24.1983 | -0.084138 | 0.14279 | 0.000103 | 0.007735 | 0.007735 | 0.75874
2019-03-20 18:03 | 1553119389.0866 | 849 | 0.20955 | 0.000452 | 0 | 106.4427 | 107.3388 | 126.1083 | 21.0664 | 0.002106 | 0.001959 | -152.2369 | 0.23883 | 8.4023 | 0.003427 | 0.004344 | 0.004135 | 6.8e-05 | 0.001084 | 0.28479 | 1.3754 | 2.5481 | 100.57 | 20.78 | 19.045 | 72.3048 | 104.1287 | 68.7701 | 0.004652 | 1.0325 | 0.02496 | 24.1738 | -0.074986 | 0.14208 | -1.9e-05 | -0.001435 | -0.001435 | -4.0683
2019-03-20 18:03 | 1553119390.0864 | 850 | 0.20955 | 0.000451 | 0 | 106.324 | 107.2962 | 126.1445 | 21.0672 | 0.002033 | 0.001987 | -152.2332 | 0.23963 | 8.4065 | 0.003348 | 0.004428 | 0.004263 | 6.9e-05 | 0.001086 | 0.24746 | 1.3525 | 2.5481 | 100.57 | 20.78 | 19.041 | 70.5552 | 102.2161 | 69.4393 | 0.004651 | 1.0305 | 0.02494 | 24.2009 | -0.081415 | 0.14256 | -6.1e-05 | -0.004615 | -0.004615 | -1.2696
2019-03-20 18:03 | 1553119391.086 | 851 | 0.20953 | 0.00045 | 0 | 106.2435 | 107.0002 | 125.7034 | 21.0672 | 0.002085 | 0.001998 | -152.2157 | 0.23877 | 8.4005 | 0.003384 | 0.004977 | 0.004805 | 6.7e-05 | 0.00109 | 0.28495 | 1.3263 | 2.5481 | 100.57 | 20.79 | 19.034 | 71.7995 | 103.0897 | 69.7138 | 0.004652 | 1.0297 | 0.024886 | 24.1683 | -0.063481 | 0.14204 | 6.3e-05 | 0.004721 | 0.004721 | 1.2366
2019-03-20 18:03 | 1553119392.086 | 852 | 0.20954 | 0.000452 | 0 | 106.1756 | 107.0518 | 125.7184 | 21.0668 | 0.002066 | 0.002019 | -152.2242 | 0.2401 | 8.3987 | 0.003436 | 0.004808 | 0.004597 | 6.7e-05 | 0.00109 | 0.28241 | 1.3083 | 2.5481 | 100.57 | 20.79 | 19.03 | 71.3461 | 104.3564 | 70.2196 | 0.004651 | 1.0282 | 0.024869 | 24.1875 | -0.073512 | 0.14283 | 7e-06 | 0.000508 | 0.000508 | 11.5633
2019-03-20 18:03 | 1553119393.0867 | 853 | 0.20954 | 0.000444 | 0 | 106.1597 | 107.0962 | 125.8036 | 21.0672 | 0.002105 | 0.001997 | -152.2123 | 0.23872 | 8.4013 | 0.003328 | 0.004972 | 0.004792 | 6.8e-05 | 0.001091 | 0.26985 | 1.3165 | 2.5481 | 100.57 | 20.79 | 19.024 | 72.275 | 101.7417 | 69.6982 | 0.004649 | 1.0297 | 0.024896 | 24.1774 | -0.078562 | 0.14201 | 6.1e-05 | 0.004609 | 0.004609 | 1.2665
2019-03-20 18:03 | 1553119394.087 | 854 | 0.20954 | 0.000453 | 0 | 105.9194 | 106.801 | 125.7041 | 21.0676 | 0.001986 | 0.00203 | -152.1992 | 0.23915 | 8.4008 | 0.003377 | 0.004672 | 0.004464 | 6.8e-05 | 0.001089 | 0.28622 | 1.2976 | 2.5481 | 100.57 | 20.79 | 19.018 | 69.4209 | 102.9305 | 70.4743 | 0.004648 | 1.0274 | 0.024858 | 24.1951 | -0.074137 | 0.14227 | 4.8e-05 | 0.003556 | 0.003556 | 1.6444
2019-03-20 18:03 | 1553119395.0868 | 855 | 0.20955 | 0.000448 | 0 | 105.8916 | 106.6126 | 125.859 | 21.0671 | 0.002022 | 0.002022 | -152.1783 | 0.23883 | 8.4052 | 0.003417 | 0.004708 | 0.004467 | 6.3e-05 | 0.001082 | 0.29103 | 1.3242 | 2.548 | 100.57 | 20.79 | 19.003 | 70.2845 | 103.8849 | 70.2828 | 0.004644 | 1.028 | 0.024881 | 24.2033 | -0.060697 | 0.14208 | -2.2e-05 | -0.001638 | -0.001638 | -3.5651
2019-03-20 18:03 | 1553119396.0857 | 856 | 0.20956 | 0.000446 | 0 | 105.8576 | 106.6362 | 126.006 | 21.0681 | 0.00206 | 0.001983 | -152.1707 | 0.23972 | 8.4018 | 0.00338 | 0.004711 | 0.004508 | 6.6e-05 | 0.001082 | 0.29463 | 1.3664 | 2.548 | 100.57 | 20.79 | 18.995 | 71.2085 | 102.9845 | 69.3411 | 0.004642 | 1.0308 | 0.024929 | 24.1842 | -0.065541 | 0.14261 | -8e-05 | -0.006004 | -0.006004 | -0.97617
2019-03-20 18:03 | 1553119397.0868 | 857 | 0.20957 | 0.000449 | 0 | 105.6248 | 106.5573 | 126.1466 | 21.0684 | 0.002013 | 0.001924 | -152.1771 | 0.23907 | 8.4023 | 0.003436 | 0.004669 | 0.004412 | 6.5e-05 | 0.001082 | 0.27848 | 1.3938 | 2.548 | 100.57 | 20.78 | 18.985 | 70.0683 | 104.3524 | 67.9244 | 0.004637 | 1.0351 | 0.024995 | 24.1476 | -0.078621 | 0.14222 | -0.000116 | -0.008699 | -0.008699 | -0.67195
2019-03-20 18:03 | 1553119398.087 | 858 | 0.20954 | 0.000448 | -1e-06 | 105.5384 | 106.3506 | 126.4203 | 21.0678 | 0.00202 | 0.001926 | -152.1791 | 0.23861 | 8.4035 | 0.003474 | 0.005489 | 0.005317 | 6.6e-05 | 0.001081 | 0.27649 | 1.3486 | 2.548 | 100.57 | 20.78 | 18.968 | 70.2306 | 105.2705 | 67.9853 | 0.004633 | 1.0349 | 0.02502 | 24.1759 | -0.06857 | 0.14195 | 5e-05 | 0.003794 | 0.003794 | 1.5377
2019-03-20 18:03 | 1553119399.0863 | 859 | 0.20954 | 0.000452 | -1e-06 | 105.3987 | 106.1024 | 126.5508 | 21.0688 | 0.002047 | 0.001955 | -152.1445 | 0.23811 | 8.404 | 0.003344 | 0.004569 | 0.004422 | 6.5e-05 | 0.001088 | 0.26416 | 1.3363 | 2.5481 | 100.57 | 20.78 | 18.961 | 70.9006 | 102.1155 | 68.6813 | 0.004631 | 1.0328 | 0.025007 | 24.2131 | -0.059522 | 0.14165 | 3.3e-05 | 0.002567 | 0.002567 | 2.268
2019-03-20 18:03 | 1553119400.0874 | 860 | 0.20954 | 0.000445 | 0 | 105.4223 | 106.3719 | 126.1684 | 21.0673 | 0.00202 | 0.001935 | -152.1544 | 0.23798 | 8.4066 | 0.00345 | 0.004546 | 0.004338 | 6.8e-05 | 0.001089 | 0.28744 | 1.3261 | 2.548 | 100.57 | 20.78 | 18.955 | 70.2357 | 104.6856 | 68.1892 | 0.00463 | 1.0343 | 0.024988 | 24.159 | -0.080207 | 0.14158 | 5.4e-05 | 0.004041 | 0.004041 | 1.4399
2019-03-20 18:03 | 1553119401.086 | 861 | 0.20953 | 0.000447 | 0 | 105.2506 | 106.0415 | 126.5907 | 21.0676 | 0.002078 | 0.001931 | -152.153 | 0.23931 | 8.4009 | 0.00338 | 0.00489 | 0.004696 | 6.6e-05 | 0.001093 | 0.28854 | 1.334 | 2.548 | 100.57 | 20.79 | 18.947 | 71.6283 | 102.9925 | 68.096 | 0.00463 | 1.0346 | 0.025033 | 24.1961 | -0.066964 | 0.14236 | 7.4e-05 | 0.005554 | 0.005554 | 1.0535
2019-03-20 18:03 | 1553119402.0861 | 862 | 0.20954 | 0.000453 | 0 | 105.1471 | 105.867 | 126.4285 | 21.0669 | 0.00209 | 0.001938 | -152.1562 | 0.23852 | 8.406 | 0.003444 | 0.004603 | 0.004384 | 6.5e-05 | 0.001095 | 0.27455 | 1.3501 | 2.5479 | 100.57 | 20.79 | 18.945 | 71.9114 | 104.5492 | 68.2802 | 0.00463 | 1.034 | 0.02501 | 24.1871 | -0.061029 | 0.1419 | 4.8e-05 | 0.003653 | 0.003653 | 1.5966
2019-03-20 18:03 | 1553119403.0862 | 863 | 0.20949 | 0.00045 | 0 | 105.1827 | 106.0337 | 126.2665 | 21.0672 | 0.002043 | 0.001954 | -152.1504 | 0.2387 | 8.4028 | 0.003435 | 0.005023 | 0.004832 | 6.8e-05 | 0.001099 | 0.27959 | 1.3388 | 2.5481 | 100.57 | 20.79 | 18.945 | 70.8001 | 104.3364 | 68.6509 | 0.00463 | 1.0329 | 0.02498 | 24.1848 | -0.072076 | 0.142 | 0.000304 | 0.022902 | 0.022902 | 0.25484
2019-03-20 18:03 | 1553119404.0876 | 864 | 0.20954 | 0.000447 | 0 | 105.0252 | 105.6259 | 126.3299 | 21.0678 | 0.002041 | 0.001931 | -152.1455 | 0.23808 | 8.402 | 0.003408 | 0.004808 | 0.004679 | 6.7e-05 | 0.001101 | 0.2907 | 1.3362 | 2.5481 | 100.57 | 20.79 | 18.946 | 70.7452 | 103.6751 | 68.1087 | 0.00463 | 1.0345 | 0.025007 | 24.1716 | -0.051016 | 0.14163 | 1.8e-05 | 0.001373 | 0.001373 | 4.2384
2019-03-20 18:03 | 1553119405.0876 | 865 | 0.20955 | 0.000449 | 0 | 104.9223 | 105.7152 | 126.0217 | 21.0682 | 0.002005 | 0.001999 | -152.148 | 0.23958 | 8.4007 | 0.003553 | 0.004802 | 0.004617 | 6.9e-05 | 0.001103 | 0.28512 | 1.3028 | 2.5482 | 100.57 | 20.79 | 18.951 | 69.8715 | 107.1764 | 69.7453 | 0.004631 | 1.0296 | 0.024916 | 24.2 | -0.067341 | 0.14253 | -2.5e-05 | -0.001842 | -0.001842 | -3.1808
2019-03-20 18:03 | 1553119406.0863 | 866 | 0.20958 | 0.000444 | 0 | 104.9737 | 105.7773 | 125.9794 | 21.067 | 0.002022 | 0.001944 | -152.1496 | 0.2389 | 8.4005 | 0.003574 | 0.004711 | 0.004499 | 6.7e-05 | 0.001104 | 0.27472 | 1.3482 | 2.5481 | 100.57 | 20.79 | 18.964 | 70.2962 | 107.7083 | 68.4175 | 0.004635 | 1.0336 | 0.024961 | 24.149 | -0.068206 | 0.14212 | -0.000179 | -0.013487 | -0.013487 | -0.43311
2019-03-20 18:03 | 1553119407.0858 | 867 | 0.20961 | 0.000451 | -1e-06 | 104.6966 | 105.5896 | 125.7421 | 21.0672 | 0.001969 | 0.001901 | -152.1557 | 0.23908 | 8.401 | 0.003527 | 0.004924 | 0.004598 | 6.7e-05 | 0.001105 | 0.27383 | 1.3506 | 2.5482 | 100.57 | 20.79 | 18.972 | 69.022 | 106.5447 | 67.3771 | 0.004637 | 1.0368 | 0.024975 | 24.0895 | -0.075974 | 0.14223 | -0.000401 | -0.030214 | -0.030214 | -0.19347
2019-03-20 18:03 | 1553119408.0855 | 868 | 0.20957 | 0.000447 | 0 | 104.6196 | 105.2959 | 125.3053 | 21.0679 | 0.002008 | 0.001935 | -152.1457 | 0.23797 | 8.4021 | 0.003533 | 0.004854 | 0.004789 | 6.5e-05 | 0.001105 | 0.26488 | 1.3528 | 2.5481 | 100.57 | 20.79 | 18.992 | 69.9601 | 106.7131 | 68.1898 | 0.004641 | 1.0343 | 0.024902 | 24.0763 | -0.057633 | 0.14157 | -0.000158 | -0.011866 | -0.011866 | -0.49034
2019-03-20 18:03 | 1553119409.0865 | 869 | 0.20959 | 0.000452 | 0 | 104.5491 | 105.1288 | 125.5936 | 21.0677 | 0.001996 | 0.001912 | -152.1626 | 0.23871 | 8.4021 | 0.003529 | 0.005235 | 0.004998 | 6.8e-05 | 0.001094 | 0.26737 | 1.3295 | 2.5481 | 100.57 | 20.79 | 19.003 | 69.6636 | 106.5971 | 67.6513 | 0.004644 | 1.0359 | 0.02495 | 24.0849 | -0.04946 | 0.14201 | -0.000284 | -0.021386 | -0.021386 | -0.27291
2019-03-20 18:03 | 1553119410.0865 | 870 | 0.20959 | 0.000452 | 0 | 104.4217 | 105.1146 | 125.8853 | 21.0671 | 0.002035 | 0.001941 | -152.171 | 0.2383 | 8.4054 | 0.00356 | 0.004751 | 0.004553 | 6.6e-05 | 0.001094 | 0.27057 | 1.3428 | 2.5481 | 100.57 | 20.79 | 19.013 | 70.6049 | 107.3452 | 68.3528 | 0.004647 | 1.0338 | 0.024954 | 24.1377 | -0.05916 | 0.14177 | -0.000288 | -0.021721 | -0.021721 | -0.26824
2019-03-20 18:04 | 1553119411.0866 | 871 | 0.20957 | 0.000452 | 0 | 104.2095 | 104.8347 | 125.4276 | 21.0669 | 0.00205 | 0.001941 | -152.1832 | 0.23908 | 8.4026 | 0.003527 | 0.005164 | 0.004908 | 6.7e-05 | 0.001097 | 0.29999 | 1.3533 | 2.548 | 100.57 | 20.79 | 19.034 | 70.9711 | 106.5641 | 68.3378 | 0.004652 | 1.0338 | 0.024909 | 24.0933 | -0.053498 | 0.14223 | -0.000172 | -0.012931 | -0.012931 | -0.45208
2019-03-20 18:04 | 1553119412.087 | 872 | 0.20956 | 0.000446 | 0 | 104.1807 | 104.7633 | 125.8297 | 21.0668 | 0.002064 | 0.001962 | -152.1754 | 0.23899 | 8.4028 | 0.003573 | 0.004704 | 0.00455 | 6.5e-05 | 0.001096 | 0.28727 | 1.3335 | 2.548 | 100.57 | 20.79 | 19.043 | 71.294 | 107.6712 | 68.8573 | 0.004654 | 1.0323 | 0.02493 | 24.1502 | -0.049877 | 0.14217 | -7.7e-05 | -0.005786 | -0.005786 | -1.0099
2019-03-20 18:04 | 1553119413.0866 | 873 | 0.20956 | 0.000446 | -1e-06 | 103.9524 | 104.4019 | 126.0226 | 21.0669 | 0.002078 | 0.001922 | -152.168 | 0.2388 | 8.4044 | 0.003555 | 0.005162 | 0.004898 | 6.7e-05 | 0.001091 | 0.29214 | 1.3441 | 2.5481 | 100.57 | 20.79 | 19.06 | 71.6439 | 107.226 | 67.8839 | 0.004658 | 1.0352 | 0.024984 | 24.1343 | -0.038591 | 0.14206 | -5.6e-05 | -0.00417 | -0.00417 | -1.4003
2019-03-20 18:04 | 1553119414.0867 | 874 | 0.20955 | 0.000446 | 0 | 103.8466 | 104.4949 | 125.9135 | 21.0666 | 0.002004 | 0.001928 | -152.185 | 0.23891 | 8.4035 | 0.003587 | 0.005026 | 0.004828 | 6.9e-05 | 0.001088 | 0.29336 | 1.3435 | 2.548 | 100.57 | 20.79 | 19.068 | 69.8591 | 108.0073 | 68.0217 | 0.00466 | 1.0348 | 0.024968 | 24.1287 | -0.055665 | 0.14213 | -5.2e-05 | -0.003937 | -0.003937 | -1.4837
2019-03-20 18:04 | 1553119415.0868 | 875 | 0.20955 | 0.000446 | 0 | 103.6613 | 104.2909 | 126.0196 | 21.0675 | 0.001998 | 0.001933 | -152.2044 | 0.23889 | 8.4033 | 0.003446 | 0.004247 | 0.004083 | 6.8e-05 | 0.001088 | 0.27367 | 1.3429 | 2.5479 | 100.57 | 20.79 | 19.076 | 69.7117 | 104.5866 | 68.1605 | 0.004662 | 1.0344 | 0.024974 | 24.1438 | -0.054156 | 0.14211 | -4.2e-05 | -0.003213 | -0.003213 | -1.8178
2019-03-20 18:04 | 1553119416.0858 | 876 | 0.20954 | 0.00044 | 0 | 103.6254 | 104.1066 | 126.1771 | 21.0668 | 0.001994 | 0.00197 | -152.218 | 0.24015 | 8.4001 | 0.003406 | 0.004616 | 0.004401 | 6.9e-05 | 0.001086 | 0.2677 | 1.354 | 2.5481 | 100.57 | 20.79 | 19.088 | 69.6151 | 103.6314 | 69.0474 | 0.004665 | 1.0317 | 0.024957 | 24.1902 | -0.041437 | 0.14286 | 7.3e-05 | 0.0055 | 0.0055 | 1.0675
2019-03-20 18:04 | 1553119417.0863 | 877 | 0.20956 | 0.000448 | 0 | 103.4831 | 103.8147 | 126.1966 | 21.0663 | 0.001984 | 0.00192 | -152.2059 | 0.23954 | 8.403 | 0.00353 | 0.005423 | 0.005216 | 6.6e-05 | 0.001082 | 0.26272 | 1.3601 | 2.5481 | 100.57 | 20.79 | 19.093 | 69.3738 | 106.6273 | 67.8271 | 0.004666 | 1.0354 | 0.025004 | 24.1489 | -0.028614 | 0.1425 | -5.6e-05 | -0.004252 | -0.004252 | -1.3775
2019-03-20 18:04 | 1553119418.087 | 878 | 0.20957 | 0.000446 | 0 | 103.3072 | 103.7908 | 126.16 | 21.0673 | 0.002024 | 0.00191 | -152.1983 | 0.23907 | 8.4041 | 0.003544 | 0.004978 | 0.004753 | 6.6e-05 | 0.001088 | 0.28495 | 1.3364 | 2.5481 | 100.57 | 20.79 | 19.098 | 70.3417 | 106.9629 | 67.605 | 0.004667 | 1.0361 | 0.025008 | 24.1375 | -0.041773 | 0.14222 | -0.000135 | -0.010185 | -0.010185 | -0.57392
2019-03-20 18:04 | 1553119419.086 | 879 | 0.20955 | 0.00045 | 0 | 103.3379 | 103.9111 | 125.6436 | 21.0684 | 0.002058 | 0.001912 | -152.1928 | 0.23873 | 8.4017 | 0.003456 | 0.004531 | 0.004487 | 6.6e-05 | 0.001091 | 0.25752 | 1.354 | 2.5481 | 100.57 | 20.79 | 19.106 | 71.1489 | 104.8463 | 67.6341 | 0.004669 | 1.036 | 0.024956 | 24.0891 | -0.049479 | 0.14202 | -3.3e-05 | -0.002542 | -0.002542 | -2.2961
2019-03-20 18:04 | 1553119420.0875 | 880 | 0.20957 | 0.000446 | 0 | 103.2364 | 103.7826 | 125.6609 | 21.0673 | 0.002009 | 0.001926 | -152.1773 | 0.23895 | 8.4036 | 0.003454 | 0.005253 | 0.004998 | 6.4e-05 | 0.001092 | 0.25194 | 1.3326 | 2.5481 | 100.57 | 20.79 | 19.109 | 69.9758 | 104.7892 | 67.9848 | 0.00467 | 1.0349 | 0.024945 | 24.1032 | -0.047199 | 0.14215 | -0.00011 | -0.008226 | -0.008226 | -0.71027
2019-03-20 18:04 | 1553119421.0871 | 881 | 0.20953 | 0.000446 | 0 | 103.1975 | 103.6357 | 125.8938 | 21.0678 | 0.002059 | 0.001942 | -152.1793 | 0.2387 | 8.4023 | 0.003455 | 0.004808 | 0.004619 | 6.8e-05 | 0.001091 | 0.26178 | 1.3496 | 2.5481 | 100.57 | 20.79 | 19.115 | 71.1856 | 104.8103 | 68.37 | 0.004672 | 1.0337 | 0.024954 | 24.1391 | -0.037894 | 0.142 | 9.6e-05 | 0.007209 | 0.007209 | 0.80956
2019-03-20 18:04 | 1553119422.0861 | 882 | 0.20954 | 0.000445 | 0 | 103.1756 | 103.7519 | 125.6583 | 21.0682 | 0.002027 | 0.001938 | -152.1687 | 0.23869 | 8.4015 | 0.003481 | 0.005088 | 0.004889 | 6.5e-05 | 0.001092 | 0.27632 | 1.3254 | 2.5481 | 100.57 | 20.79 | 19.119 | 70.4072 | 105.4304 | 68.2783 | 0.004672 | 1.034 | 0.024934 | 24.1133 | -0.049827 | 0.142 | 5.2e-05 | 0.003913 | 0.003913 | 1.4915
2019-03-20 18:04 | 1553119423.0852 | 883 | 0.20955 | 0.000445 | 0 | 103.0963 | 103.6936 | 125.4835 | 21.0672 | 0.002026 | 0.001989 | -152.1924 | 0.23924 | 8.4021 | 0.0034 | 0.00478 | 0.004508 | 6.7e-05 | 0.001098 | 0.30558 | 1.3532 | 2.548 | 100.57 | 20.79 | 19.124 | 70.3923 | 103.4889 | 69.4987 | 0.004674 | 1.0303 | 0.024872 | 24.1395 | -0.051671 | 0.14232 | -5e-06 | -0.00036 | -0.00036 | -16.2582
2019-03-20 18:04 | 1553119424.0863 | 884 | 0.20955 | 0.000441 | 0 | 103.0801 | 103.6196 | 125.5262 | 21.0674 | 0.00197 | 0.001936 | -152.1788 | 0.23862 | 8.4031 | 0.003389 | 0.00547 | 0.00524 | 6.6e-05 | 0.001102 | 0.28119 | 1.3322 | 2.5481 | 100.57 | 20.79 | 19.138 | 69.0442 | 103.2038 | 68.2252 | 0.004677 | 1.0342 | 0.024923 | 24.0987 | -0.046688 | 0.14195 | 2.1e-05 | 0.001557 | 0.001557 | 3.7472
2019-03-20 18:04 | 1553119425.0867 | 885 | 0.20954 | 0.00044 | 0 | 103.0523 | 103.515 | 125.4337 | 21.0674 | 0.001958 | 0.001903 | -152.1439 | 0.23924 | 8.4049 | 0.003441 | 0.004594 | 0.004388 | 6.7e-05 | 0.001102 | 0.26698 | 1.3675 | 2.5481 | 100.57 | 20.79 | 19.145 | 68.754 | 104.4804 | 67.4327 | 0.004679 | 1.0366 | 0.024942 | 24.0619 | -0.040064 | 0.14232 | 6.2e-05 | 0.004705 | 0.004705 | 1.2433
2019-03-20 18:04 | 1553119426.0859 | 886 | 0.20955 | 0.00044 | 0 | 102.9621 | 103.6241 | 125.3638 | 21.0669 | 0.002067 | 0.00195 | -152.1457 | 0.23978 | 8.4022 | 0.003349 | 0.00438 | 0.004168 | 6.5e-05 | 0.001104 | 0.29479 | 1.3788 | 2.5481 | 100.57 | 20.79 | 19.16 | 71.3806 | 102.248 | 68.5641 | 0.004683 | 1.0332 | 0.024894 | 24.0951 | -0.057324 | 0.14264 | -1.4e-05 | -0.001022 | -0.001022 | -5.7381
2019-03-20 18:04 | 1553119427.0864 | 887 | 0.20957 | 0.000442 | 1e-06 | 102.9362 | 103.2219 | 125.5569 | 21.0666 | 0.001952 | 0.00197 | -152.1628 | 0.23904 | 8.4048 | 0.0034 | 0.004882 | 0.004658 | 6.7e-05 | 0.001103 | 0.2792 | 1.3615 | 2.548 | 100.57 | 20.79 | 19.168 | 68.5971 | 103.4662 | 69.0363 | 0.004684 | 1.0317 | 0.024896 | 24.1303 | -0.024792 | 0.14221 | -0.000123 | -0.009246 | -0.009246 | -0.63214
2019-03-20 18:04 | 1553119428.0854 | 888 | 0.20956 | 0.000441 | 0 | 102.825 | 103.2212 | 125.6213 | 21.0673 | 0.002014 | 0.001951 | -152.1685 | 0.23954 | 8.4037 | 0.003419 | 0.004981 | 0.004812 | 6.5e-05 | 0.001111 | 0.27417 | 1.4142 | 2.5481 | 100.57 | 20.79 | 19.177 | 70.1038 | 103.9433 | 68.5885 | 0.004687 | 1.0331 | 0.024919 | 24.1207 | -0.034399 | 0.1425 | -7.9e-05 | -0.005934 | -0.005934 | -0.98704
2019-03-20 18:04 | 1553119429.0855 | 889 | 0.20955 | 0.000442 | 0 | 102.7613 | 103.0586 | 125.9595 | 21.0674 | 0.00202 | 0.001926 | -152.1606 | 0.2398 | 8.3999 | 0.003392 | 0.004832 | 0.004588 | 6.8e-05 | 0.001116 | 0.2656 | 1.3496 | 2.5481 | 100.57 | 20.79 | 19.194 | 70.2396 | 103.2934 | 67.9717 | 0.004691 | 1.035 | 0.024975 | 24.1313 | -0.025844 | 0.14266 | 2e-06 | 0.000176 | 0.000176 | 33.2783
2019-03-20 18:04 | 1553119430.0856 | 890 | 0.20956 | 0.000448 | -1e-06 | 102.758 | 103.1068 | 125.9473 | 21.0676 | 0.00195 | 0.001907 | -152.1677 | 0.23894 | 8.4021 | 0.003364 | 0.00428 | 0.004137 | 6.4e-05 | 0.001118 | 0.31205 | 1.3477 | 2.5483 | 100.57 | 20.79 | 19.202 | 68.56 | 102.6085 | 67.518 | 0.004693 | 1.0363 | 0.02499 | 24.1141 | -0.030309 | 0.14215 | -6.5e-05 | -0.004853 | -0.004853 | -1.2037
2019-03-20 18:04 | 1553119431.0856 | 891 | 0.20958 | 0.000444 | 0 | 102.7315 | 103.1289 | 126.0468 | 21.0687 | 0.002063 | 0.001969 | -152.1707 | 0.23954 | 8.4016 | 0.003359 | 0.00432 | 0.004132 | 6.7e-05 | 0.00112 | 0.27096 | 1.305 | 2.5481 | 100.57 | 20.79 | 19.215 | 71.2771 | 102.4909 | 69.0197 | 0.004696 | 1.0318 | 0.024945 | 24.1768 | -0.03453 | 0.1425 | -0.000172 | -0.012978 | -0.012978 | -0.45131
2019-03-20 18:04 | 1553119432.0857 | 892 | 0.20955 | 0.000444 | 0 | 102.7179 | 103.1162 | 125.9358 | 21.0684 | 0.002019 | 0.001924 | -152.1769 | 0.24017 | 8.4018 | 0.003363 | 0.005047 | 0.004832 | 6.7e-05 | 0.001122 | 0.24602 | 1.2979 | 2.5481 | 100.57 | 20.79 | 19.219 | 70.2093 | 102.57 | 67.9282 | 0.004697 | 1.0351 | 0.024974 | 24.1275 | -0.034611 | 0.14288 | 6e-06 | 0.000443 | 0.000443 | 13.2677
2019-03-20 18:04 | 1553119433.0859 | 893 | 0.20959 | 0.000449 | 0 | 102.595 | 102.8635 | 126.0297 | 21.068 | 0.001978 | 0.00196 | -152.1829 | 0.23858 | 8.4046 | 0.003396 | 0.004476 | 0.004325 | 6.2e-05 | 0.001124 | 0.28766 | 1.3488 | 2.5481 | 100.57 | 20.79 | 19.221 | 69.2365 | 103.3772 | 68.7889 | 0.004697 | 1.0325 | 0.024952 | 24.167 | -0.023376 | 0.14193 | -0.000255 | -0.019201 | -0.019201 | -0.3038
2019-03-20 18:04 | 1553119434.0858 | 894 | 0.20971 | 0.000448 | 0 | 102.7005 | 103.0044 | 126.1318 | 21.0675 | 0.002013 | 0.001898 | -152.1875 | 0.23866 | 8.4052 | 0.003374 | 0.005201 | 0.004988 | 6.4e-05 | 0.001123 | 0.29336 | 1.2995 | 2.548 | 100.57 | 20.79 | 19.216 | 70.083 | 102.8471 | 67.3173 | 0.004696 | 1.0369 | 0.025016 | 24.1247 | -0.026433 | 0.14198 | -0.000969 | -0.073176 | -0.073176 | -0.079743
2019-03-20 18:04 | 1553119435.087 | 895 | 0.20962 | 0.000449 | 0 | 102.695 | 103.0724 | 126.4517 | 21.0676 | 0.00201 | 0.001932 | -152.22 | 0.23928 | 8.4019 | 0.003329 | 0.004189 | 0.003993 | 6.4e-05 | 0.001129 | 0.27024 | 1.3138 | 2.5479 | 100.57 | 20.79 | 19.21 | 69.9956 | 101.7547 | 68.1178 | 0.004695 | 1.0345 | 0.025018 | 24.1836 | -0.032817 | 0.14235 | -0.00045 | -0.03396 | -0.03396 | -0.17228
2019-03-20 18:04 | 1553119436.0859 | 896 | 0.20964 | 0.000447 | 0 | 102.6183 | 102.9936 | 126.2116 | 21.0671 | 0.001991 | 0.001904 | -152.2293 | 0.23915 | 8.4015 | 0.003356 | 0.003856 | 0.003645 | 6.2e-05 | 0.001129 | 0.26952 | 1.3057 | 2.5479 | 100.57 | 20.79 | 19.202 | 69.5476 | 102.4155 | 67.4487 | 0.004693 | 1.0365 | 0.025019 | 24.1369 | -0.032649 | 0.14227 | -0.000542 | -0.040916 | -0.040916 | -0.14291
2019-03-20 18:04 | 1553119437.0859 | 897 | 0.20961 | 0.000447 | 0 | 102.5297 | 102.9674 | 126.0877 | 21.0691 | 0.002053 | 0.001929 | -152.2354 | 0.23913 | 8.4038 | 0.003355 | 0.004334 | 0.004139 | 6.6e-05 | 0.001133 | 0.28694 | 1.3338 | 2.5481 | 100.57 | 20.79 | 19.181 | 71.0365 | 102.3978 | 68.0561 | 0.004688 | 1.0347 | 0.024984 | 24.1466 | -0.038104 | 0.14226 | -0.000379 | -0.028561 | -0.028561 | -0.20471
2019-03-20 18:04 | 1553119438.0851 | 898 | 0.20959 | 0.000446 | 0 | 102.5174 | 102.8355 | 126.3905 | 21.0678 | 0.002042 | 0.001926 | -152.2396 | 0.23926 | 8.4033 | 0.003316 | 0.004844 | 0.004632 | 6.4e-05 | 0.001135 | 0.2756 | 1.3566 | 2.5482 | 100.57 | 20.79 | 19.169 | 70.7584 | 101.4408 | 67.983 | 0.004685 | 1.0349 | 0.025017 | 24.173 | -0.027708 | 0.14233 | -0.000258 | -0.019482 | -0.019482 | -0.30028
2019-03-20 18:04 | 1553119439.0861 | 899 | 0.20963 | 0.000445 | 0 | 102.4844 | 102.8332 | 126.2407 | 21.0686 | 0.00201 | 0.001911 | -152.2415 | 0.23939 | 8.4021 | 0.00333 | 0.00509 | 0.004837 | 6.2e-05 | 0.001131 | 0.29927 | 1.3513 | 2.5481 | 100.57 | 20.79 | 19.144 | 70.0129 | 101.7801 | 67.6093 | 0.004679 | 1.0361 | 0.025016 | 24.1454 | -0.030392 | 0.14242 | -0.000485 | -0.036624 | -0.036624 | -0.15982
2019-03-20 18:04 | 1553119440.0854 | 900 | 0.20953 | 0.000445 | 0 | 102.4078 | 102.7338 | 126.0426 | 21.0695 | 0.002014 | 0.001955 | -152.2396 | 0.23955 | 8.4015 | 0.00339 | 0.004763 | 0.00459 | 6.2e-05 | 0.001131 | 0.24275 | 1.3452 | 2.5481 | 100.57 | 20.79 | 19.13 | 70.1069 | 103.2255 | 68.6782 | 0.004675 | 1.0328 | 0.024957 | 24.1643 | -0.028432 | 0.14251 | 7.6e-05 | 0.005691 | 0.005691 | 1.0293
2019-03-20 18:04 | 1553119441.0855 | 901 | 0.2096 | 0.000445 | 0 | 102.3014 | 102.7873 | 126.3191 | 21.0682 | 0.002029 | 0.001938 | -152.2211 | 0.23956 | 8.4019 | 0.003328 | 0.005169 | 0.004906 | 6.4e-05 | 0.001131 | 0.27383 | 1.3368 | 2.548 | 100.57 | 20.79 | 19.118 | 70.4648 | 101.7262 | 68.2783 | 0.004672 | 1.034 | 0.024999 | 24.1766 | -0.042383 | 0.14251 | -0.000297 | -0.0224 | -0.0224 | -0.26149
2019-03-20 18:04 | 1553119442.0865 | 902 | 0.20957 | 0.000446 | 0 | 102.3169 | 102.7603 | 126.417 | 21.0682 | 0.002041 | 0.001911 | -152.2247 | 0.23979 | 8.4033 | 0.003375 | 0.005426 | 0.005236 | 6.2e-05 | 0.001126 | 0.26681 | 1.3635 | 2.5481 | 100.57 | 20.79 | 19.098 | 70.7503 | 102.8754 | 67.6253 | 0.004667 | 1.036 | 0.025033 | 24.1628 | -0.038683 | 0.14265 | -0.000118 | -0.008896 | -0.008896 | -0.65904
2019-03-20 18:04 | 1553119443.0853 | 903 | 0.20956 | 0.000441 | 0 | 102.1173 | 102.3227 | 126.3996 | 21.0683 | 0.002017 | 0.001929 | -152.2108 | 0.23949 | 8.4029 | 0.003359 | 0.00464 | 0.004489 | 6e-05 | 0.001133 | 0.29806 | 1.3036 | 2.5481 | 100.57 | 20.79 | 19.09 | 70.1593 | 102.4772 | 68.063 | 0.004665 | 1.0347 | 0.025015 | 24.1767 | -0.017968 | 0.14247 | -9e-05 | -0.00677 | -0.00677 | -0.86491
2019-03-20 18:04 | 1553119444.0853 | 904 | 0.20955 | 0.000442 | 0 | 102.1096 | 102.3634 | 126.3658 | 21.0682 | 0.001927 | 0.001911 | -152.2109 | 0.23949 | 8.403 | 0.003397 | 0.005125 | 0.004846 | 6.4e-05 | 0.001133 | 0.25249 | 1.3195 | 2.548 | 100.57 | 20.79 | 19.082 | 67.9975 | 103.416 | 67.6191 | 0.004663 | 1.036 | 0.025028 | 24.1577 | -0.022209 | 0.14247 | 1.6e-05 | 0.001216 | 0.001216 | 4.8149
2019-03-20 18:04 | 1553119445.0855 | 905 | 0.20958 | 0.00045 | 0 | 102.0527 | 102.128 | 126.3034 | 21.0689 | 0.001915 | 0.001935 | -152.2063 | 0.23909 | 8.4042 | 0.003364 | 0.00457 | 0.004407 | 6.3e-05 | 0.00113 | 0.26128 | 1.3517 | 2.5479 | 100.57 | 20.79 | 19.08 | 67.7205 | 102.6129 | 68.1963 | 0.004663 | 1.0343 | 0.025001 | 24.1722 | -0.006598 | 0.14224 | -0.000173 | -0.013069 | -0.013069 | -0.44732
2019-03-20 18:04 | 1553119446.0854 | 906 | 0.20957 | 0.000443 | 0 | 101.9941 | 102.2364 | 126.2768 | 21.0686 | 0.002057 | 0.001898 | -152.1929 | 0.23937 | 8.4058 | 0.003365 | 0.004814 | 0.004486 | 6.5e-05 | 0.001124 | 0.26272 | 1.3335 | 2.5481 | 100.58 | 20.79 | 19.081 | 71.1266 | 102.6417 | 67.3122 | 0.004663 | 1.037 | 0.02503 | 24.1384 | -0.021218 | 0.1424 | -0.00012 | -0.009075 | -0.009075 | -0.64494
2019-03-20 18:04 | 1553119447.0855 | 907 | 0.20955 | 0.000448 | 1e-06 | 101.965 | 102.3631 | 126.5106 | 21.0689 | 0.00197 | 0.001929 | -152.1907 | 0.23989 | 8.4004 | 0.003316 | 0.005242 | 0.005043 | 6.3e-05 | 0.001126 | 0.26449 | 1.3505 | 2.548 | 100.58 | 20.79 | 19.091 | 69.0441 | 101.4428 | 68.0628 | 0.004665 | 1.0347 | 0.025026 | 24.1873 | -0.034849 | 0.14271 | 0 | -2.2e-05 | -2.2e-05 | -264.3135
2019-03-20 18:04 | 1553119448.0865 | 908 | 0.20954 | 0.000444 | 0 | 101.9883 | 102.1881 | 126.5421 | 21.0691 | 0.00201 | 0.001934 | -152.2037 | 0.23877 | 8.4011 | 0.003412 | 0.004661 | 0.004506 | 6.5e-05 | 0.001126 | 0.2636 | 1.3426 | 2.5479 | 100.58 | 20.79 | 19.099 | 70.0134 | 103.7744 | 68.1729 | 0.004667 | 1.0343 | 0.025025 | 24.1942 | -0.017509 | 0.14204 | 4.3e-05 | 0.003226 | 0.003226 | 1.8098
2019-03-20 18:04 | 1553119449.0866 | 909 | 0.20957 | 0.000448 | -1e-06 | 101.9391 | 102.2692 | 126.3285 | 21.0694 | 0.001944 | 0.001887 | -152.2056 | 0.23963 | 8.401 | 0.003354 | 0.004491 | 0.004314 | 6.6e-05 | 0.001122 | 0.28031 | 1.3331 | 2.5482 | 100.58 | 20.79 | 19.108 | 68.4198 | 102.3698 | 67.0523 | 0.004669 | 1.0378 | 0.025045 | 24.1341 | -0.028917 | 0.14256 | -0.000148 | -0.011175 | -0.011175 | -0.52433
2019-03-20 18:04 | 1553119450.0867 | 910 | 0.20959 | 0.000449 | 0 | 102.0132 | 102.4027 | 126.6005 | 21.0691 | 0.001984 | 0.001839 | -152.2194 | 0.23971 | 8.403 | 0.00339 | 0.004651 | 0.00444 | 6.2e-05 | 0.00112 | 0.26665 | 1.3149 | 2.548 | 100.57 | 20.79 | 19.127 | 69.378 | 103.2408 | 65.8794 | 0.004674 | 1.0413 | 0.025115 | 24.1184 | -0.034084 | 0.1426 | -0.00027 | -0.020502 | -0.020502 | -0.28586
2019-03-20 18:04 | 1553119451.0867 | 911 | 0.20956 | 0.000449 | 0 | 101.9786 | 102.2887 | 126.4672 | 21.0697 | 0.001974 | 0.001914 | -152.2222 | 0.23928 | 8.4027 | 0.003397 | 0.00559 | 0.005315 | 6.6e-05 | 0.001115 | 0.27295 | 1.3388 | 2.5481 | 100.57 | 20.79 | 19.138 | 69.1259 | 103.4112 | 67.7008 | 0.004677 | 1.0358 | 0.025035 | 24.1703 | -0.027156 | 0.14235 | -7.4e-05 | -0.005621 | -0.005621 | -1.0408
2019-03-20 18:04 | 1553119452.0861 | 912 | 0.20953 | 0.000452 | 0 | 101.8767 | 102.2644 | 126.4616 | 21.0686 | 0.002014 | 0.001957 | -152.2325 | 0.23917 | 8.4051 | 0.003363 | 0.005517 | 0.005352 | 6.4e-05 | 0.001111 | 0.27024 | 1.3523 | 2.548 | 100.57 | 20.79 | 19.159 | 70.0887 | 102.5734 | 68.7381 | 0.004682 | 1.0326 | 0.024996 | 24.2065 | -0.033971 | 0.14228 | 7.9e-05 | 0.005957 | 0.005957 | 0.98169
2019-03-20 18:04 | 1553119453.0858 | 913 | 0.20954 | 0.00045 | 0 | 101.802 | 102.0614 | 126.2339 | 21.0692 | 0.002039 | 0.001937 | -152.2269 | 0.23919 | 8.401 | 0.00339 | 0.004889 | 0.004669 | 6.8e-05 | 0.001106 | 0.25874 | 1.3715 | 2.548 | 100.57 | 20.79 | 19.168 | 70.6968 | 103.2472 | 68.2425 | 0.004684 | 1.0341 | 0.024992 | 24.1672 | -0.022765 | 0.1423 | 6.7e-05 | 0.005066 | 0.005066 | 1.1545
2019-03-20 18:04 | 1553119454.0848 | 914 | 0.20954 | 0.000447 | 0 | 101.7131 | 101.8738 | 126.3602 | 21.0703 | 0.002013 | 0.001898 | -152.2178 | 0.23889 | 8.4027 | 0.003382 | 0.005017 | 0.004805 | 6.4e-05 | 0.001106 | 0.262 | 1.334 | 2.5481 | 100.57 | 20.79 | 19.176 | 70.0802 | 103.0503 | 67.3176 | 0.004686 | 1.0369 | 0.025039 | 24.1465 | -0.014124 | 0.14212 | 3.6e-05 | 0.002751 | 0.002751 | 2.1233
2019-03-20 18:04 | 1553119455.0849 | 915 | 0.20956 | 0.000448 | 0 | 101.6086 | 101.7366 | 126.2616 | 21.0695 | 0.001992 | 0.00193 | -152.2191 | 0.2393 | 8.4031 | 0.003396 | 0.004315 | 0.004098 | 6.8e-05 | 0.001105 | 0.2855 | 1.3091 | 2.548 | 100.58 | 20.79 | 19.191 | 69.5774 | 103.3836 | 68.0854 | 0.00469 | 1.0346 | 0.025001 | 24.1643 | -0.011264 | 0.14236 | -4.6e-05 | -0.003472 | -0.003472 | -1.6853
2019-03-20 18:04 | 1553119456.0854 | 916 | 0.20955 | 0.000445 | 0 | 101.6303 | 101.8824 | 126.4243 | 21.0694 | 0.002025 | 0.001923 | -152.2212 | 0.23905 | 8.4032 | 0.003501 | 0.004765 | 0.004562 | 6.5e-05 | 0.001102 | 0.25465 | 1.2853 | 2.5481 | 100.58 | 20.79 | 19.199 | 70.3677 | 105.9305 | 67.9097 | 0.004692 | 1.0351 | 0.025023 | 24.1736 | -0.022162 | 0.14221 | 0 | -1.6e-05 | -1.6e-05 | -356.1957
2019-03-20 18:04 | 1553119457.0853 | 917 | 0.20956 | 0.000449 | 0 | 101.4501 | 101.5756 | 126.1851 | 21.0692 | 0.001961 | 0.001902 | -152.1941 | 0.23947 | 8.4011 | 0.003471 | 0.004318 | 0.004163 | 6.7e-05 | 0.001106 | 0.27793 | 1.279 | 2.5479 | 100.58 | 20.79 | 19.214 | 68.8197 | 105.1885 | 67.4018 | 0.004695 | 1.0367 | 0.025018 | 24.1328 | -0.011051 | 0.14246 | -6.7e-05 | -0.00499 | -0.00499 | -1.1734
2019-03-20 18:04 | 1553119458.0861 | 918 | 0.20955 | 0.000449 | 0 | 101.4534 | 101.6402 | 125.902 | 21.0688 | 0.001983 | 0.001873 | -152.186 | 0.23922 | 8.4026 | 0.003372 | 0.004415 | 0.004224 | 6.6e-05 | 0.001105 | 0.28119 | 1.3168 | 2.548 | 100.58 | 20.79 | 19.218 | 69.3536 | 102.7895 | 66.703 | 0.004696 | 1.0388 | 0.025016 | 24.0809 | -0.016458 | 0.14231 | 1.4e-05 | 0.001045 | 0.001045 | 5.5969
2019-03-20 18:04 | 1553119459.0851 | 919 | 0.20957 | 0.000449 | 0 | 101.3651 | 101.4758 | 126.182 | 21.0689 | 0.002012 | 0.001896 | -152.1921 | 0.23889 | 8.4043 | 0.003509 | 0.005197 | 0.004959 | 6.4e-05 | 0.001108 | 0.25371 | 1.3427 | 2.5482 | 100.58 | 20.79 | 19.219 | 70.0515 | 106.1308 | 67.2711 | 0.004696 | 1.0371 | 0.025023 | 24.1278 | -0.009762 | 0.14212 | -0.00013 | -0.009786 | -0.009786 | -0.59687
2019-03-20 18:04 | 1553119460.0863 | 920 | 0.20955 | 0.000452 | -1e-06 | 101.4078 | 101.5513 | 126.2412 | 21.0698 | 0.002038 | 0.001891 | -152.1811 | 0.24039 | 8.4025 | 0.003422 | 0.005112 | 0.004914 | 6.4e-05 | 0.001108 | 0.28799 | 1.38 | 2.5482 | 100.58 | 20.79 | 19.221 | 70.6817 | 104.0165 | 67.1479 | 0.004697 | 1.0375 | 0.025033 | 24.1291 | -0.012648 | 0.14301 | -3.3e-05 | -0.00247 | -0.00247 | -2.3797
2019-03-20 18:04 | 1553119461.0852 | 921 | 0.20957 | 0.000452 | 0 | 101.3289 | 101.512 | 126.1508 | 21.069 | 0.002071 | 0.001922 | -152.2011 | 0.23969 | 8.4032 | 0.003364 | 0.00451 | 0.00432 | 6.7e-05 | 0.00111 | 0.2469 | 1.3664 | 2.5481 | 100.58 | 20.79 | 19.219 | 71.4613 | 102.6067 | 67.8937 | 0.004696 | 1.0352 | 0.024997 | 24.1469 | -0.016154 | 0.14259 | -0.000126 | -0.009501 | -0.009501 | -0.61685
2019-03-20 18:04 | 1553119462.0862 | 922 | 0.20957 | 0.000446 | 0 | 101.2221 | 101.4119 | 126.1567 | 21.0697 | 0.002044 | 0.001881 | -152.1934 | 0.23883 | 8.4045 | 0.003514 | 0.004703 | 0.004456 | 6.9e-05 | 0.001111 | 0.26897 | 1.3772 | 2.5482 | 100.58 | 20.79 | 19.215 | 70.8196 | 106.241 | 66.8903 | 0.004695 | 1.0382 | 0.025034 | 24.1119 | -0.01675 | 0.14208 | -0.000154 | -0.011633 | -0.011633 | -0.50198
2019-03-20 18:04 | 1553119463.0864 | 923 | 0.2096 | 0.00045 | 0 | 101.1801 | 101.3535 | 126.5301 | 21.0689 | 0.001936 | 0.001865 | -152.207 | 0.24 | 8.4009 | 0.003415 | 0.004622 | 0.004452 | 6.6e-05 | 0.001119 | 0.27776 | 1.3744 | 2.5481 | 100.57 | 20.79 | 19.213 | 68.2252 | 103.8397 | 66.5221 | 0.004695 | 1.0394 | 0.025085 | 24.1345 | -0.015321 | 0.14278 | -0.000291 | -0.021992 | -0.021992 | -0.26683
2019-03-20 18:04 | 1553119464.0865 | 924 | 0.20963 | 0.000452 | 0 | 101.1804 | 101.4272 | 126.424 | 21.07 | 0.001955 | 0.001949 | -152.2116 | 0.23893 | 8.4018 | 0.003412 | 0.004768 | 0.004544 | 6.6e-05 | 0.001107 | 0.24458 | 1.3499 | 2.5481 | 100.57 | 20.79 | 19.212 | 68.6714 | 103.776 | 68.5303 | 0.004695 | 1.0333 | 0.025 | 24.1956 | -0.021789 | 0.14214 | -0.00048 | -0.036181 | -0.036181 | -0.16146
2019-03-20 18:04 | 1553119465.0865 | 925 | 0.20957 | 0.000449 | 0 | 101.0575 | 101.1207 | 126.6137 | 21.0701 | 0.001999 | 0.001842 | -152.2056 | 0.23841 | 8.4024 | 0.003418 | 0.005078 | 0.00479 | 6.7e-05 | 0.001113 | 0.28385 | 1.4045 | 2.5482 | 100.57 | 20.79 | 19.211 | 69.7378 | 103.9218 | 65.9715 | 0.004695 | 1.0411 | 0.025113 | 24.1229 | -0.005591 | 0.14183 | -0.000112 | -0.008468 | -0.008468 | -0.68834
2019-03-20 18:04 | 1553119466.0855 | 926 | 0.20958 | 0.000448 | 1e-06 | 100.9 | 100.9805 | 126.4379 | 21.0694 | 0.001983 | 0.001952 | -152.2197 | 0.2398 | 8.4033 | 0.003413 | 0.004989 | 0.004736 | 6.7e-05 | 0.001109 | 0.27544 | 1.371 | 2.5479 | 100.57 | 20.79 | 19.211 | 69.351 | 103.7984 | 68.5984 | 0.004695 | 1.0331 | 0.024999 | 24.1993 | -0.007137 | 0.14266 | -0.000175 | -0.013266 | -0.013266 | -0.44198
2019-03-20 18:04 | 1553119467.0856 | 927 | 0.20956 | 0.000443 | 0 | 100.7917 | 100.8699 | 126.4243 | 21.0691 | 0.001992 | 0.001845 | -152.2328 | 0.23813 | 8.4049 | 0.003434 | 0.004735 | 0.004489 | 6.4e-05 | 0.001111 | 0.29048 | 1.3596 | 2.548 | 100.57 | 20.79 | 19.21 | 69.5784 | 104.2942 | 66.0282 | 0.004695 | 1.0409 | 0.025092 | 24.1069 | -0.006942 | 0.14166 | -5.4e-05 | -0.004114 | -0.004114 | -1.4152
2019-03-20 18:04 | 1553119468.0856 | 928 | 0.20956 | 0.000446 | 0 | 100.6157 | 100.6916 | 126.582 | 21.0706 | 0.001969 | 0.001896 | -152.2362 | 0.23975 | 8.4012 | 0.003441 | 0.004969 | 0.004791 | 6.8e-05 | 0.001104 | 0.26897 | 1.3423 | 2.5481 | 100.57 | 20.79 | 19.207 | 69.0075 | 104.4764 | 67.2727 | 0.004694 | 1.0371 | 0.025062 | 24.1661 | -0.006746 | 0.14262 | -7.4e-05 | -0.005646 | -0.005646 | -1.0382
2019-03-20 18:04 | 1553119469.0857 | 929 | 0.20958 | 0.000444 | 0 | 100.5627 | 100.6023 | 126.539 | 21.07 | 0.002021 | 0.001948 | -152.2485 | 0.23853 | 8.4043 | 0.003451 | 0.004451 | 0.004239 | 6.9e-05 | 0.001103 | 0.28064 | 1.3497 | 2.5481 | 100.57 | 20.79 | 19.204 | 70.2693 | 104.7071 | 68.5147 | 0.004693 | 1.0333 | 0.025012 | 24.206 | -0.003524 | 0.14191 | -0.000179 | -0.01355 | -0.01355 | -0.43041
2019-03-20 18:04 | 1553119470.0849 | 930 | 0.20955 | 0.000441 | -1e-06 | 100.4469 | 100.4644 | 126.2215 | 21.0704 | 0.00202 | 0.001862 | -152.2579 | 0.24001 | 8.4028 | 0.003368 | 0.004697 | 0.004438 | 7.1e-05 | 0.001101 | 0.25818 | 1.3421 | 2.5481 | 100.57 | 20.79 | 19.202 | 70.2526 | 102.6969 | 66.4549 | 0.004693 | 1.0396 | 0.025057 | 24.1027 | -0.001558 | 0.14278 | 4e-06 | 0.000344 | 0.000344 | 17.0604
2019-03-20 18:05 | 1553119471.0858 | 931 | 0.20956 | 0.000447 | 0 | 100.345 | 100.3549 | 126.2238 | 21.0703 | 0.002037 | 0.001947 | -152.2511 | 0.23966 | 8.3997 | 0.003429 | 0.004813 | 0.004644 | 6.7e-05 | 0.001107 | 0.27759 | 1.3368 | 2.5481 | 100.57 | 20.79 | 19.203 | 70.6409 | 104.1876 | 68.4886 | 0.004693 | 1.0334 | 0.024982 | 24.1749 | -0.000879 | 0.14257 | -6.7e-05 | -0.005076 | -0.005076 | -1.1543
2019-03-20 18:05 | 1553119472.0858 | 932 | 0.20957 | 0.000447 | 0 | 100.2622 | 100.4995 | 126.0431 | 21.0702 | 0.002011 | 0.001926 | -152.256 | 0.23945 | 8.4018 | 0.003381 | 0.005202 | 0.004901 | 6.8e-05 | 0.001102 | 0.28622 | 1.3079 | 2.548 | 100.57 | 20.79 | 19.204 | 70.0248 | 103.0275 | 67.9869 | 0.004693 | 1.0349 | 0.024983 | 24.1399 | -0.021144 | 0.14245 | -0.000139 | -0.010497 | -0.010497 | -0.55774
2019-03-20 18:05 | 1553119473.0851 | 933 | 0.20957 | 0.000445 | 0 | 100.2089 | 100.1759 | 125.9895 | 21.0703 | 0.001932 | 0.001882 | -152.24 | 0.23945 | 8.4026 | 0.003406 | 0.005166 | 0.004938 | 6.8e-05 | 0.001098 | 0.27577 | 1.2734 | 2.548 | 100.57 | 20.79 | 19.202 | 68.1267 | 103.6264 | 66.9308 | 0.004693 | 1.0381 | 0.025016 | 24.0974 | 0.002947 | 0.14245 | -0.000134 | -0.010155 | -0.010155 | -0.57653
2019-03-20 18:05 | 1553119474.1419 | 934 | 0.20955 | 0.000444 | 0 | 100.1558 | 100.0929 | 126.1003 | 21.0697 | 0.001998 | 0.001884 | -152.2448 | 0.23981 | 8.4026 | 0.003371 | 0.004758 | 0.004561 | 6.8e-05 | 0.001099 | 0.27793 | 1.3226 | 2.5479 | 100.57 | 20.79 | 19.2 | 69.7201 | 102.7754 | 66.9646 | 0.004692 | 1.038 | 0.025026 | 24.1092 | 0.005623 | 0.14266 | 1.9e-05 | 0.001498 | 0.001498 | 3.9133
2019-03-20 18:05 | 1553119475.0925 | 935 | 0.20957 | 0.000446 | -1e-06 | 100.063 | 100.1299 | 125.9344 | 21.0702 | 0.001996 | 0.001888 | -152.2438 | 0.23958 | 8.4043 | 0.00335 | 0.004717 | 0.004559 | 6.2e-05 | 0.001094 | 0.30613 | 1.3246 | 2.548 | 100.58 | 20.8 | 19.193 | 69.6739 | 102.2771 | 67.0704 | 0.004693 | 1.0377 | 0.025005 | 24.097 | -0.005976 | 0.14252 | -0.000105 | -0.007892 | -0.007892 | -0.74221
2019-03-20 18:05 | 1553119476.0925 | 936 | 0.20955 | 0.000444 | 0 | 99.9505 | 99.8017 | 126.1161 | 21.0703 | 0.001991 | 0.001817 | -152.2161 | 0.23964 | 8.4024 | 0.003338 | 0.004284 | 0.00409 | 6.5e-05 | 0.001089 | 0.27671 | 1.3154 | 2.5481 | 100.58 | 20.8 | 19.189 | 69.5553 | 101.9867 | 65.3707 | 0.004692 | 1.0429 | 0.025086 | 24.0541 | 0.013319 | 0.14256 | -3.5e-05 | -0.002661 | -0.002661 | -2.2019
2019-03-20 18:05 | 1553119477.0916 | 937 | 0.20955 | 0.000441 | -1e-06 | 99.8098 | 99.9034 | 126.1541 | 21.0702 | 0.001949 | 0.001889 | -152.2092 | 0.23902 | 8.403 | 0.003368 | 0.004426 | 0.004202 | 6.6e-05 | 0.001089 | 0.27223 | 1.3356 | 2.5481 | 100.58 | 20.8 | 19.183 | 68.5468 | 102.6967 | 67.1038 | 0.004691 | 1.0376 | 0.025026 | 24.1192 | -0.008386 | 0.14219 | -1.5e-05 | -0.001069 | -0.001069 | -5.468
2019-03-20 18:05 | 1553119478.093 | 938 |  | 0.000441 |  | 99.6219 | 99.5412 | 125.8482 | 21.0704 | 0.001966 | 0.001852 | -152.2015 | 0.23883 | 8.4019 | 0.003448 | 0.004701 | 0.004477 | 6.5e-05 | 0.001089 | 0.27007 | 1.3127 | 2.548 | 100.58 | 20.8 | 19.172 | 68.9425 | 104.6525 | 66.2145 | 0.004688 | 1.0403 | 0.025028 | 24.0585 | 0.007244 | 0.14208
2019-03-20 18:05 | 1553119479.093 | 939 |  | 0.000447 |  | 99.5087 | 99.3697 | 125.8792 | 21.0707 | 0.001898 | 0.001833 | -152.1815 | 0.23931 | 8.4041 | 0.00337 | 0.005469 | 0.005329 | 6.6e-05 | 0.001088 | 0.27256 | 1.3605 | 2.548 | 100.58 | 20.8 | 19.166 | 67.3094 | 102.7624 | 65.7388 | 0.004686 | 1.0418 | 0.025049 | 24.0446 | 0.012503 | 0.14236
2019-03-20 18:05 | 1553119480.093 | 940 |  |  |  | 99.4136 | 99.3099 | 126.0048 | 21.0712 | 0.001961 | 0.001896 | -152.1833 | 0.2393 | 8.4032 | 0.00349 | 0.004468 | 0.0043 | 6.7e-05 | 0.001095 | 0.29502 | 1.3062 | 2.548 | 100.58 | 20.8 | 19.16 | 68.8254 | 105.6654 | 67.2538 | 0.004685 | 1.0371 | 0.025006 | 24.1103 | 0.009337 | 0.14236
2019-03-20 18:05 | 1553119481.093 | 941 |  |  |  | 99.2338 | 99.0296 | 125.8426 | 21.0709 | 0.00196 | 0.001835 | -152.1839 | 0.23888 | 8.4074 | 0.003447 | 0.004829 | 0.004586 | 6.3e-05 | 0.001093 | 0.28854 | 1.318 | 2.5481 | 100.58 | 20.8 | 19.158 | 68.8042 | 104.6168 | 65.7889 | 0.004684 | 1.0416 | 0.025043 | 24.0428 | 0.018426 | 0.14211
2019-03-20 18:05 | 1553119482.093 | 942 |  |  |  | 99.1044 | 98.8296 | 125.952 | 21.0702 | 0.001918 | 0.001896 | -152.1912 | 0.2385 | 8.4056 | 0.00354 | 0.004681 | 0.004485 | 6.6e-05 | 0.001091 | 0.28783 | 1.3089 | 2.548 | 100.58 | 20.8 | 19.157 | 67.7919 | 106.8805 | 67.2692 | 0.004684 | 1.0371 | 0.025 | 24.1058 | 0.024836 | 0.14188
2019-03-20 18:05 | 1553119483.093 | 943 |  |  |  | 98.9382 | 98.8315 |  | 21.0712 | 0.002013 | 0.001862 | -152.1954 | 0.23895 | 8.4014 | 0.003396 | 0.004614 | 0.004456 | 6.3e-05 | 0.001091 | 0.29878 | 1.3405 | 2.548 | 100.57 | 20.8 | 19.155 | 70.0818 | 103.3734 | 66.44 | 0.004684 | 1.0396 |  |  | 0.009651 | 0.14215
2019-03-20 18:05 | 1553119484.093 | 944 |  |  |  | 98.8227 | 98.7668 |  | 21.07 | 0.001949 |  | -152.2065 | 0.23861 | 8.4023 | 0.00345 | 0.005213 | 0.004974 | 6.1e-05 | 0.001095 | 0.25409 | 1.3504 | 2.548 | 100.57 | 20.8 | 19.152 | 68.5428 | 104.7008 |  | 0.004683 |  |  |  | 0.005061 | 0.14195
2019-03-20 18:05 | 1553119485.093 | 945 |  |  |  | 98.6694 | 98.3564 |  | 21.0708 | 0.001972 |  |  | 0.23885 | 8.4062 | 0.003415 | 0.004633 | 0.004401 | 6.3e-05 | 0.001091 | 0.30934 | 1.3702 | 2.548 | 100.57 | 20.8 | 19.153 | 69.0989 | 103.8549 |  | 0.004684 |  |  |  | 0.028417 | 0.14209
2019-03-20 18:05 | 1553119486.093 | 946 |  |  |  | 98.5475 | 98.4215 |  | 21.0706 |  |  |  | 0.2396 | 8.4037 | 0.003399 | 0.004869 | 0.004651 | 6.2e-05 | 0.001094 | 0.26913 | 1.3312 | 2.5481 | 100.57 | 20.8 | 19.155 |  | 103.4443 |  | 0.004684 |  |  |  | 0.011445 | 0.14253
2019-03-20 18:05 | 1553119487.093 | 947 |  |  |  | 98.4259 | 98.1819 |  | 21.0702 |  |  |  | 0.23976 | 8.4027 | 0.003404 | 0.005228 | 0.00498 | 6.3e-05 | 0.001096 | 0.28296 | 1.3476 | 2.548 | 100.57 | 20.8 | 19.16 |  | 103.565 |  | 0.004685 |  |  |  | 0.0222 | 0.14264
2019-03-20 18:05 | 1553119488.093 | 948 |  |  |  | 98.216 | 97.8706 |  | 21.0695 |  |  |  | 0.2385 | 8.4024 | 0.003462 | 0.004156 | 0.004104 | 5.9e-05 | 0.001099 | 0.27273 | 1.3297 | 2.548 | 100.57 | 20.8 | 19.162 |  | 104.9803 |  | 0.004686 |  |  |  | 0.031513 | 0.14188
2019-03-20 18:05 | 1553119489.093 | 949 |  |  |  | 98.0245 | 97.4344 |  | 21.0702 |  |  |  | 0.23956 | 8.401 | 0.003441 | 0.00447 | 0.004271 | 6e-05 | 0.001091 | 0.28727 | 1.3218 | 2.5481 | 100.57 | 20.8 | 19.166 |  | 104.462 |  | 0.004687 |  |  |  | 0.054011 | 0.14252
2019-03-20 18:05 | 1553119490.093

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-20 17:49 | 0 | 0 | 0 | 0 | 0 | 48 | 0.010806
2019-03-20 17:49 | 1 | 0 | 0 | 0 | 0 | 48.1346 | -0.002255
2019-03-20 17:49 | 2 | 0 | 0 | 0 | 0 | 48.2601 | -0.003545
2019-03-20 17:49 | 3 | 0 | 0 | 0 | 0 | 48.4332 | 0.001685
2019-03-20 17:49 | 4 | 0 | 0 | 0 | 0 | 48.2555 | 0.00292
2019-03-20 17:49 | 5 | 0 | 0 | 0 | 0 | 48.2915 | 0.00441
2019-03-20 17:49 | 6 | 0 | 0 | 0 | 0 | 48.084 | 0.069329
2019-03-20 17:49 | 7 | 0 | 0 | 0 | 0 | 48.0298 | 0.24444
2019-03-20 17:49 | 8 | 0 | 0 | 0 | 0 | 47.9375 | 0.76752
2019-03-20 17:49 | 9 | 0 | 0 | 0 | 0 | 47.8468 | 1.314
2019-03-20 17:49 | 10 | 0 | 0 | 0 | 0 | 47.7616 | 1.5584
2019-03-20 17:49 | 11 | 0 | 0 | 0 | 0 | 47.6495 | 2.0732
2019-03-20 17:49 | 12 | 0 | 0 | 0 | 0 | 47.5676 | 2.3567
2019-03-20 17:49 | 13 | 0 | 0 | 0 | 0 | 47.4455 | 2.7566
2019-03-20 17:49 | 14 | 0 | 0 | 0 | 0 | 47.3388 | 3.1373
2019-03-20 17:49 | 15 | 0.096808 | 0 | 0 | 0 | 47.2423 | 3.3168
2019-03-20 17:49 | 16 | 0.11701 | 0 | 0 | 0 | 47.1395 | 3.3641
2019-03-20 17:49 | 17 | 0.11924 | 0 | 0 | 0 | 47.0126 | 3.3581
2019-03-20 17:49 | 18 | 0.10304 | 0 | 0 | 0 | 46.9056 | 3.3245
2019-03-20 17:49 | 19 | 0.11145 | 0 | 0 | 0 | 46.8016 | 3.2985
2019-03-20 17:49 | 20 | 0.096692 | 0 | 0 | 0 | 46.6876 | 3.0714
2019-03-20 17:49 | 21 | 0.091863 | 0 | 0 | 0 | 46.6061 | 2.9154
2019-03-20 17:49 | 22 | 0.096127 | 0 | 0 | 0 | 46.5023 | 2.6429
2019-03-20 17:49 | 23 | 0.063642 | 0 | 0 | 0 | 46.4216 | 2.3742
2019-03-20 17:49 | 24 | 0.074942 | 0 | 0 | 0 | 46.3651 | 2.2153
2019-03-20 17:49 | 25 | 0.081845 | 0 | 0 | 0 | 46.2721 | 2.1156
2019-03-20 17:49 | 26 | 0.071489 | 0 | 0 | 0 | 46.2052 | 2.038
2019-03-20 17:49 | 27 | 0.079581 | 0 | 0 | 0 | 46.1253 | 1.9665
2019-03-20 17:49 | 28 | 0.07829 | 0 | 0 | 0 | 46.0482 | 1.9073
2019-03-20 17:49 | 29 | 0.068826 | 0 | 0 | 0 | 45.9708 | 1.8197
2019-03-20 17:49 | 30 | 0.066542 | 0 | 0 | 0 | 45.9088 | 1.7714
2019-03-20 17:49 | 31 | 0.070501 | 0 | 0 | 0 | 45.8361 | 1.7419
2019-03-20 17:50 | 32 | 0.065371 | 0 | 0 | 0 | 45.7694 | 1.7601
2019-03-20 17:50 | 33 | 0.081069 | 0 | 0 | 0 | 45.7011 | 1.7289
2019-03-20 17:50 | 34 | 0.070783 | 0 | 0 | 0 | 45.6132 | 1.7236
2019-03-20 17:50 | 35 | 0.063648 | 0 | 0 | 0 | 45.5594 | 1.6976
2019-03-20 17:50 | 36 | 0.062123 | 0 | 0 | 0 | 45.4851 | 1.6504
2019-03-20 17:50 | 37 | 0.035498 | 0 | 0 | 0 | 45.4406 | 1.5519
2019-03-20 17:50 | 38 | 0.032878 | 0 | 0 | 0 | 45.4081 | 1.4513
2019-03-20 17:50 | 39 | 0.051141 | 0 | 0 | 0 | 45.3691 | 1.3237
2019-03-20 17:50 | 40 | 0.061396 | 0 | 0 | 0 | 45.3079 | 1.2734
2019-03-20 17:50 | 41 | 0.049161 | 0 | 0 | 0 | 45.2525 | 1.2466
2019-03-20 17:50 | 42 | 0.027449 | 0 | 0 | 0 | 45.2118 | 1.166
2019-03-20 17:50 | 43 | 0.033487 | 0 | 0 | 0 | 45.1903 | 1.1216
2019-03-20 17:50 | 44 | 0.059819 | 0 | 0 | 0 | 45.1403 | 1.1019
2019-03-20 17:50 | 45 | 0.035457 | 0 | 0 | 0 | 45.0824 | 1.0621
2019-03-20 17:50 | 46 | 0.025778 | 0 | 0 | 0 | 45.0657 | 1.0315
2019-03-20 17:50 | 47 | 0.064882 | 0 | 0 | 0 | 45.0192 | 1.0109
2019-03-20 17:50 | 48 | 0.068075 | 0 | 0 | 0 | 44.9444 | 0.99422
2019-03-20 17:50 | 49 | 0.045195 | 0 | 0 | 0 | 44.8903 | 0.99726
2019-03-20 17:50 | 50 | 0.033287 | 0 | 0 | 0 | 44.8516 | 1.0084
2019-03-20 17:50 | 51 | 0.035753 | 0 | 0 | 0 | 44.8191 | 0.98494
2019-03-20 17:50 | 52 | 0.064656 | 0 | 0 | 0 | 44.7739 | 0.98661
2019-03-20 17:50 | 53 | 0.062525 | 0 | 0 | 0 | 44.6977 | 0.98571
2019-03-20 17:50 | 54 | 0.026056 | 0 | 0 | 0 | 44.6567 | 0.96988
2019-03-20 17:50 | 55 | 0.049719 | 0 | 0 | 0 | 44.6317 | 0.95442
2019-03-20 17:50 | 56 | 0.067298 | 0 | 0 | 0 | 44.5585 | 0.95457
2019-03-20 17:50 | 57 | 0.036304 | 0 | 0 | 0 | 44.5087 | 0.95502
2019-03-20 17:50 | 58 | 0.032571 | 0 | 0 | 0 | 44.4798 | 0.95451
2019-03-20 17:50 | 59 | 0.048451 | 0 | 0 | 0 | 44.4382 | 0.94294
2019-03-20 17:50 | 60 | 0.051554 | 0 | 0 | 0 | 44.3858 | 0.93281
2019-03-20 17:50 | 61 | 0.04521 | 0 | 0 | 0 | 44.338 | 0.92298
2019-03-20 17:50 | 62 | 0.034702 | 0 | 0 | 0 | 44.2963 | 0.91427
2019-03-20 17:50 | 63 | 0.035729 | 0 | 0 | 0 | 44.2651 | 0.88647
2019-03-20 17:50 | 64 | 0.061539 | 0 | 0 | 0 | 44.2195 | 0.92531
2019-03-20 17:50 | 65 | 0.043192 | 0 | 0 | 0 | 44.1523 | 0.90687
2019-03-20 17:50 | 66 | 0.030334 | 0 | 0 | 0 | 44.1321 | 0.86768
2019-03-20 17:50 | 67 | 0.049715 | 0 | 0 | 0 | 44.084 | 0.86456
2019-03-20 17:50 | 68 | 0.044175 | 0 | 0 | 0 | 44.0377 | 0.87088
2019-03-20 17:50 | 69 | 0.051358 | 0 | 0 | 0 | 43.9932 | 0.86372
2019-03-20 17:50 | 70 | 0.055628 | 0 | 0 | 0 | 43.9356 | 0.86933
2019-03-20 17:50 | 71 | 0.054495 | 0 | 0 | 0 | 43.8835 | 0.86359
2019-03-20 17:50 | 72 | 0.050427 | 0 | 0 | 0 | 43.8277 | 0.86233
2019-03-20 17:50 | 73 | 0.041609 | 0 | 0 | 0 | 43.7839 | 0.88177
2019-03-20 17:50 | 74 | 0.036119 | 0 | 0 | 0 | 43.7435 | 0.87315
2019-03-20 17:50 | 75 | 0.038825 | 0 | 0 | 0 | 43.7095 | 0.86101
2019-03-20 17:50 | 76 | 0.039111 | 0 | 0 | 0 | 43.6663 | 0.85914
2019-03-20 17:50 | 77 | 0.035509 | 0 | 0 | 0 | 43.6321 | 0.87549
2019-03-20 17:50 | 78 | 0.040427 | 0.014222 | 2.192 | 21.6021 | 43.5936 | 0.8733
2019-03-20 17:50 | 79 | 0.033022 | 0.017454 | 2.6802 | 26.7281 | 43.5545 | 0.88261
2019-03-20 17:50 | 80 | 0.012614 | 0.045534 | 6.9664 | 69.7354 | 43.5302 | 0.87963
2019-03-20 17:50 | 81 | 0.023797 | 0.024196 | 3.6958 | 37.0057 | 43.5211 | 0.88063
2019-03-20 17:50 | 82 | 0.052033 | 0.011067 | 1.6876 | 16.9276 | 43.4783 | 0.88079
2019-03-20 17:50 | 83 | 0.046797 | 0.01178 | 1.8751 | 18.3701 | 43.425 | 0.85966
2019-03-20 17:50 | 84 | 0.041522 | 0.012763 | 2.1164 | 20.9552 | 43.3852 | 0.87011
2019-03-20 17:50 | 85 | 0.057214 | 0.00967 | 1.5329 | 14.9937 | 43.3373 | 0.85785
2019-03-20 17:50 | 86 | 0.051877 | 0.011086 | 1.6842 | 16.8761 | 43.2756 | 0.87549
2019-03-20 17:50 | 87 | 0.043337 | 0.013265 | 2.0136 | 20.8588 | 43.2346 | 0.90396
2019-03-20 17:50 | 88 | 0.048412 | 0.012336 | 1.8051 | 18.8206 | 43.1858 | 0.91114
2019-03-20 17:50 | 89 | 0.042484 | 0.014048 | 2.059 | 21.4514 | 43.1395 | 0.91134
2019-03-20 17:50 | 90 | 0.061575 | 0.0096906 | 1.4262 | 14.9674 | 43.0953 | 0.92161
2019-03-20 17:50 | 91 | 0.064458 | 0.0092537 | 1.3663 | 14.2603 | 43.0205 | 0.91918
2019-03-20 17:50 | 92 | 0.039465 | 0.01573 | 2.243 | 23.5024 | 42.9731 | 0.92751
2019-03-20 17:51 | 93 | 0.042825 | 0.015027 | 2.0772 | 21.941 | 42.9351 | 0.93961
2019-03-20 17:51 | 94 | 0.05021 | 0.012367 | 1.775 | 18.3781 | 42.8858 | 0.92276
2019-03-20 17:51 | 95 | 0.052365 | 0.011833 | 1.7073 | 18.0053 | 42.8359 | 0.94284
2019-03-20 17:51 | 96 | 0.049685 | 0.012475 | 1.8036 | 19.1521 | 42.7823 | 0.95156
2019-03-20 17:51 | 97 | 0.053656 | 0.012001 | 1.6785 | 17.83 | 42.7352 | 0.95669
2019-03-20 17:51 | 98 | 0.054448 | 0.01141 | 1.6628 | 17.32 | 42.6758 | 0.94303
2019-03-20 17:51 | 99 | 0.047223 | 0.013636 | 1.9247 | 20.3113 | 42.6279 | 0.95917
2019-03-20 17:51 | 100 | 0.05952 | 0.010419 | 1.5275 | 15.7567 | 42.5774 | 0.93784
2019-03-20 17:51 | 101 | 0.045608 | 0.013571 | 1.9968 | 20.4729 | 42.5149 | 0.93373
2019-03-20 17:51 | 102 | 0.030502 | 0.020284 | 2.9891 | 30.8666 | 42.4862 | 0.94148
2019-03-20 17:51 | 103 | 0.050318 | 0.01183 | 1.819 | 18.6083 | 42.4459 | 0.93634
2019-03-20 17:51 | 104 | 0.043195 | 0.013774 | 2.1187 | 21.6581 | 42.3914 | 0.93551
2019-03-20 17:51 | 105 | 0.034608 | 0.01717 | 2.638 | 27.1342 | 42.36 | 0.93906
2019-03-20 17:51 | 106 | 0.046543 | 0.013264 | 1.9616 | 20.4161 | 42.3175 | 0.95022
2019-03-20 17:51 | 107 | 0.045031 | 0.014198 | 2.0279 | 21.5037 | 42.2698 | 0.96832
2019-03-20 17:51 | 108 | 0.041284 | 0.016044 | 2.217 | 24.0172 | 42.228 | 0.99154
2019-03-20 17:51 | 109 | 0.048499 | 0.013654 | 1.8933 | 20.9646 | 42.1846 | 1.0168
2019-03-20 17:51 | 110 | 0.052549 | 0.013051 | 1.7624 | 19.1911 | 42.1315 | 1.0085
2019-03-20 17:51 | 111 | 0.057671 | 0.011899 | 1.6155 | 17.5787 | 42.0794 | 1.0138
2019-03-20 17:51 | 112 | 0.059611 | 0.011096 | 1.5687 | 17.0448 | 42.0171 | 1.0161
2019-03-20 17:51 | 113 | 0.055082 | 0.012021 | 1.7079 | 18.1175 | 41.9615 | 0.99794
2019-03-20 17:51 | 114 | 0.066152 | 0.010027 | 1.4307 | 15.2673 | 41.9038 | 1.01
2019-03-20 17:51 | 115 | 0.057559 | 0.011114 | 1.6492 | 17.4292 | 41.834 | 1.0032
2019-03-20 17:51 | 116 | 0.041163 | 0.016073 | 2.3027 | 24.7841 | 41.7907 | 1.0202
2019-03-20 17:51 | 117 | 0.040671 | 0.015704 | 2.333 | 25.0762 | 41.748 | 1.0199
2019-03-20 17:51 | 118 | 0.038665 | 0.017128 | 2.4642 | 26.6624 | 41.7097 | 1.0309
2019-03-20 17:51 | 119 | 0.022558 | 0.029327 | 4.2287 | 45.773 | 41.6738 | 1.0326
2019-03-20 17:51 | 120 | 0.02156 | 0.030705 | 4.4507 | 47.9852 | 41.6607 | 1.0346
2019-03-20 17:51 | 121 | 0.035893 | 0.019021 | 2.6672 | 28.5023 | 41.6266 | 1.023
2019-03-20 17:51 | 122 | 0.039407 | 0.017341 | 2.4362 | 25.9187 | 41.5914 | 1.0214
2019-03-20 17:51 | 123 | 0.035566 | 0.019907 | 2.7235 | 28.7034 | 41.5499 | 1.0209
2019-03-20 17:51 | 124 | 0.034219 | 0.020673 | 2.8336 | 29.513 | 41.5195 | 1.0099
2019-03-20 17:51 | 125 | 0.045421 | 0.016041 | 2.1292 | 22.0165 | 41.4783 | 1
2019-03-20 17:51 | 126 | 0.054718 | 0.013713 | 1.769 | 18.2706 | 41.4291 | 0.99972
2019-03-20 17:51 | 127 | 0.050902 | 0.014754 | 1.9075 | 19.7581 | 41.3721 | 1.0057
2019-03-20 17:51 | 128 | 0.049443 | 0.015673 | 1.9653 | 20.1698 | 41.3267 | 0.99725
2019-03-20 17:51 | 129 | 0.06214 | 0.012465 | 1.5613 | 16.1653 | 41.2701 | 1.0045
2019-03-20 17:51 | 130 | 0.055756 | 0.014682 | 1.7342 | 17.8378 | 41.2065 | 0.99456
2019-03-20 17:51 | 131 | 0.061926 | 0.01363 | 1.5694 | 15.9838 | 41.156 | 0.98981
2019-03-20 17:51 | 132 | 0.066751 | 0.012632 | 1.4539 | 14.9126 | 41.0833 | 0.99544
2019-03-20 17:51 | 133 | 0.050148 | 0.016353 | 1.9309 | 19.4788 | 41.0274 | 0.97682
2019-03-20 17:51 | 134 | 0.055009 | 0.01491 | 1.7631 | 17.6243 | 40.9785 | 0.96949
2019-03-20 17:51 | 135 | 0.050655 | 0.016654 | 1.9069 | 19.0748 | 40.9194 | 0.96624
2019-03-20 17:51 | 136 | 0.038502 | 0.022571 | 2.5192 | 25.582 | 40.879 | 0.98495
2019-03-20 17:51 | 137 | 0.036236 | 0.024564 | 2.6665 | 27.0215 | 40.8396 | 0.97914
2019-03-20 17:51 | 138 | 0.04827 | 0.019386 | 1.9957 | 20.4335 | 40.8025 | 0.98632
2019-03-20 17:51 | 139 | 0.067967 | 0.014091 | 1.4208 | 14.5451 | 40.7418 | 0.98859
2019-03-20 17:51 | 140 | 0.052104 | 0.018415 | 1.8527 | 18.9242 | 40.6755 | 0.98602
2019-03-20 17:51 | 141 | 0.028401 | 0.032978 | 3.4051 | 34.7137 | 40.6392 | 0.98589
2019-03-20 17:51 | 142 | 0.051151 | 0.018748 | 1.8932 | 19.1225 | 40.6072 | 0.97814
2019-03-20 17:51 | 143 | 0.075859 | 0.01262 | 1.2753 | 12.5606 | 40.5365 | 0.95284
2019-03-20 17:51 | 144 | 0.047938 | 0.019529 | 2.0202 | 20.0344 | 40.468 | 0.96042
2019-03-20 17:51 | 145 | 0.042294 | 0.022125 | 2.2905 | 22.6253 | 40.4358 | 0.95691
2019-03-20 17:51 | 146 | 0.062323 | 0.014995 | 1.55 | 15.2615 | 40.3767 | 0.95113
2019-03-20 17:51 | 147 | 0.070896 | 0.013189 | 1.3623 | 13.7024 | 40.3136 | 0.97145
2019-03-20 17:51 | 148 | 0.070371 | 0.01331 | 1.3743 | 13.5274 | 40.2379 | 0.95193
2019-03-20 17:51 | 149 | 0.057094 | 0.016349 | 1.6894 | 16.772 | 40.1761 | 0.95758
2019-03-20 17:51 | 150 | 0.053238 | 0.017998 | 1.8101 | 17.9085 | 40.1214 | 0.95341
2019-03-20 17:51 | 151 | 0.062144 | 0.015408 | 1.5485 | 15.3911 | 40.0669 | 0.95647
2019-03-20 17:52 | 152 | 0.048057 | 0.020434 | 2.0028 | 19.944 | 40.0029 | 0.95846
2019-03-20 17:52 | 153 | 0.022203 | 0.044158 | 4.3264 | 43.1096 | 39.9736 | 0.95718
2019-03-20 17:52 | 154 | 0.021255 | 0.047265 | 4.5189 | 45.0213 | 39.952 | 0.95692
2019-03-20 17:52 | 155 | 0.038734 | 0.027132 | 2.4831 | 24.8477 | 39.9259 | 0.96244
2019-03-20 17:52 | 156 | 0.048477 | 0.022105 | 1.9756 | 19.5636 | 39.8763 | 0.94838
2019-03-20 17:52 | 157 | 0.041853 | 0.026206 | 2.2899 | 22.7348 | 39.833 | 0.95152
2019-03-20 17:52 | 158 | 0.04319 | 0.025965 | 2.2141 | 21.829 | 39.7908 | 0.9428
2019-03-20 17:52 | 159 | 0.05317 | 0.021001 | 1.7922 | 17.6848 | 39.7439 | 0.9403
2019-03-20 17:52 | 160 | 0.075506 | 0.0151 | 1.2624 | 12.4166 | 39.6818 | 0.93753
2019-03-20 17:52 | 161 | 0.070743 | 0.016087 | 1.3424 | 13.2333 | 39.5998 | 0.93616
2019-03-20 17:52 | 162 | 0.053141 | 0.021435 | 1.7906 | 17.5626 | 39.5437 | 0.93329
2019-03-20 17:52 | 163 | 0.058275 | 0.019183 | 1.6346 | 15.9833 | 39.4879 | 0.93142
2019-03-20 17:52 | 164 | 0.072439 | 0.015421 | 1.3086 | 13.0809 | 39.4244 | 0.94757
2019-03-20 17:52 | 165 | 0.083131 | 0.01342 | 1.1371 | 10.9471 | 39.3441 | 0.91005
2019-03-20 17:52 | 166 | 0.076921 | 0.014528 | 1.2296 | 12.0839 | 39.2627 | 0.92951
2019-03-20 17:52 | 167 | 0.065025 | 0.017846 | 1.4445 | 14.1653 | 39.1919 | 0.92109
2019-03-20 17:52 | 168 | 0.064277 | 0.018411 | 1.4605 | 14.1285 | 39.1299 | 0.90814
2019-03-20 17:52 | 169 | 0.069277 | 0.01744 | 1.3563 | 13.5414 | 39.0624 | 0.9381
2019-03-20 17:52 | 170 | 0.046259 | 0.026585 | 2.0262 | 20.3382 | 38.9977 | 0.94082
2019-03-20 17:52 | 171 | 0.037482 | 0.033449 | 2.5011 | 25.154 | 38.9658 | 0.94282
2019-03-20 17:52 | 172 | 0.070128 | 0.017511 | 1.3295 | 13.3186 | 38.9125 | 0.934
2019-03-20 17:52 | 173 | 0.081779 | 0.01476 | 1.1444 | 11.4912 | 38.8303 | 0.93974
2019-03-20 17:52 | 174 | 0.074119 | 0.015978 | 1.2631 | 12.684 | 38.7549 | 0.94013
2019-03-20 17:52 | 175 | 0.053224 | 0.021853 | 1.7628 | 17.7339 | 38.6861 | 0.94387
2019-03-20 17:52 | 176 | 0.029608 | 0.038392 | 3.1601 | 31.8921 | 38.6487 | 0.94427
2019-03-20 17:52 | 177 | 0.035423 | 0.032798 | 2.6457 | 26.9281 | 38.6195 | 0.95386
2019-03-20 17:52 | 178 | 0.042072 | 0.027542 | 2.2243 | 22.7074 | 38.577 | 0.95535
2019-03-20 17:52 | 179 | 0.03598 | 0.032235 | 2.6052 | 26.9602 | 38.5379 | 0.97002
2019-03-20 17:52 | 180 | 0.049784 | 0.023275 | 1.879 | 19.7721 | 38.5006 | 0.98433
2019-03-20 17:52 | 181 | 0.055794 | 0.020761 | 1.6836 | 17.8075 | 38.4402 | 0.99356
2019-03-20 17:52 | 182 | 0.044204 | 0.026169 | 2.1295 | 22.9925 | 38.393 | 1.0164
2019-03-20 17:52 | 183 | 0.058266 | 0.019843 | 1.6276 | 17.8162 | 38.3464 | 1.0381
2019-03-20 17:52 | 184 | 0.050964 | 0.022286 | 1.8715 | 20.2367 | 38.2808 | 1.0313
2019-03-20 17:52 | 185 | 0.052875 | 0.021462 | 1.8158 | 19.5837 | 38.2417 | 1.0355
2019-03-20 17:52 | 186 | 0.088773 | 0.012249 | 1.0844 | 11.9439 | 38.1674 | 1.0603
2019-03-20 17:52 | 187 | 0.087808 | 0.011889 | 1.1039 | 12.0139 | 38.0733 | 1.0549
2019-03-20 17:52 | 188 | 0.06394 | 0.015974 | 1.5222 | 16.5168 | 37.9984 | 1.0561
2019-03-20 17:52 | 189 | 0.046854 | 0.021321 | 2.0811 | 23.1758 | 37.9438 | 1.0859
2019-03-20 17:52 | 190 | 0.051877 | 0.018828 | 1.8936 | 20.5906 | 37.8985 | 1.0682
2019-03-20 17:52 | 191 | 0.071432 | 0.013345 | 1.3756 | 14.9066 | 37.8366 | 1.0648
2019-03-20 17:52 | 192 | 0.058452 | 0.015912 | 1.6857 | 18.0989 | 37.7639 | 1.0579
2019-03-20 17:52 | 193 | 0.0264 | 0.034341 | 3.7425 | 40.2041 | 37.7241 | 1.0614
2019-03-20 17:52 | 194 | 0.045265 | 0.019531 | 2.1893 | 23.2822 | 37.6993 | 1.0539
2019-03-20 17:52 | 195 | 0.050063 | 0.017233 | 1.9918 | 21.3237 | 37.6361 | 1.0675
2019-03-20 17:52 | 196 | 0.035681 | 0.023619 | 2.8067 | 29.9435 | 37.6033 | 1.0684
2019-03-20 17:52 | 197 | 0.054792 | 0.015347 | 1.8276 | 19.2896 | 37.5571 | 1.0569
2019-03-20 17:52 | 198 | 0.0647 | 0.012635 | 1.5448 | 16.366 | 37.4952 | 1.0589
2019-03-20 17:52 | 199 | 0.072026 | 0.011347 | 1.3908 | 14.5589 | 37.4291 | 1.0486
2019-03-20 17:52 | 200 | 0.064046 | 0.012393 | 1.558 | 16.3886 | 37.3559 | 1.0496
2019-03-20 17:52 | 201 | 0.031642 | 0.0251 | 3.1622 | 32.5427 | 37.3064 | 1.0297
2019-03-20 17:52 | 202 | 0.038579 | 0.020003 | 2.5896 | 26.8528 | 37.2827 | 1.0359
2019-03-20 17:52 | 203 | 0.0677 | 0.011066 | 1.476 | 14.7882 | 37.2236 | 1.0012
2019-03-20 17:52 | 204 | 0.057086 | 0.013127 | 1.7453 | 17.3093 | 37.1564 | 0.98812
2019-03-20 17:52 | 205 | 0.056145 | 0.012945 | 1.7679 | 17.1597 | 37.1082 | 0.96344
2019-03-20 17:52 | 206 | 0.056882 | 0.012767 | 1.7366 | 17.1047 | 37.0442 | 0.97294
2019-03-20 17:52 | 207 | 0.033643 | 0.020896 | 2.9129 | 27.9139 | 36.9999 | 0.9391
2019-03-20 17:52 | 208 | 0.028855 | 0.023551 | 3.374 | 34.1929 | 36.9726 | 0.98664
2019-03-20 17:52 | 209 | 0.032102 | 0.021167 | 3.0125 | 27.7221 | 36.9392 | 0.88993
2019-03-20 17:52 | 210 | 0.047084 | 0.013985 | 2.0392 | 19.067 | 36.9048 | 0.89774
2019-03-20 17:52 | 211 | 0.072738 | 0.0090524 | 1.3087 | 12.0369 | 36.8429 | 0.87554
2019-03-20 17:52 | 212 | 0.066506 | 0.0098915 | 1.4166 | 13.3215 | 36.7674 | 0.88595
2019-03-20 17:53 | 213 | 0.051143 | 0.012862 | 1.8239 | 17.4331 | 36.7126 | 0.89157
2019-03-20 17:53 | 214 | 0.057626 | 0.011417 | 1.6023 | 15.233 | 36.6605 | 0.87782
2019-03-20 17:53 | 215 | 0.044187 | 0.014895 | 2.0695 | 19.6294 | 36.6015 | 0.86736
2019-03-20 17:53 | 216 | 0.039456 | 0.017229 | 2.2977 | 22.1137 | 36.5697 | 0.87251
2019-03-20 17:53 | 217 | 0.060259 | 0.010919 | 1.4945 | 14.1297 | 36.5165 | 0.85144
2019-03-20 17:53 | 218 | 0.057455 | 0.011487 | 1.5617 | 15.065 | 36.4542 | 0.86556
2019-03-20 17:53 | 219 | 0.062278 | 0.010596 | 1.4326 | 13.3883 | 36.4007 | 0.8338
2019-03-20 17:53 | 220 | 0.055208 | 0.011542 | 1.6046 | 15.283 | 36.3329 | 0.84375
2019-03-20 17:53 | 221 | 0.036889 | 0.017228 | 2.3833 | 22.5634 | 36.2929 | 0.83234
2019-03-20 17:53 | 222 | 0.039005 | 0.016891 | 2.2388 | 21.1621 | 36.2544 | 0.82543
2019-03-20 17:53 | 223 | 0.03659 | 0.018002 | 2.3744 | 22.6303 | 36.2153 | 0.82805
2019-03-20 17:53 | 224 | 0.037144 | 0.017716 | 2.3223 | 22.4716 | 36.1801 | 0.83469
2019-03-20 17:53 | 225 | 0.054106 | 0.012592 | 1.5876 | 15.5322 | 36.1377 | 0.84038
2019-03-20 17:53 | 226 | 0.038639 | 0.017687 | 2.218 | 21.7261 | 36.0789 | 0.83948
2019-03-20 17:53 | 227 | 0.045711 | 0.014936 | 1.8675 | 18.2576 | 36.0554 | 0.83457
2019-03-20 17:53 | 228 | 0.061541 | 0.011117 | 1.3888 | 13.4932 | 35.9858 | 0.83039
2019-03-20 17:53 | 229 | 0.039219 | 0.017445 | 2.1693 | 21.1415 | 35.941 | 0.82915
2019-03-20 17:53 | 230 | 0.042314 | 0.016155 | 2.0039 | 19.3928 | 35.9018 | 0.82059
2019-03-20 17:53 | 231 | 0.051181 | 0.013343 | 1.6474 | 15.9773 | 35.8546 | 0.81774
2019-03-20 17:53 | 232 | 0.04625 | 0.01525 | 1.8128 | 17.3312 | 35.8024 | 0.80158
2019-03-20 17:53 | 233 | 0.052417 | 0.013431 | 1.5904 | 14.8009 | 35.7601 | 0.77582
2019-03-20 17:53 | 234 | 0.046458 | 0.015657 | 1.7877 | 16.9378 | 35.7009 | 0.78689
2019-03-20 17:53 | 235 | 0.020437 | 0.03668 | 4.0348 | 38.6997 | 35.6711 | 0.79091
2019-03-20 17:53 | 236 | 0.049052 | 0.015782 | 1.6726 | 16.1598 | 35.6477 | 0.79267
2019-03-20 17:53 | 237 | 0.044985 | 0.017666 | 1.8124 | 17.3501 | 35.5798 | 0.7805
2019-03-20 17:53 | 238 | 0.034106 | 0.024626 | 2.3773 | 22.8736 | 35.5596 | 0.78013
2019-03-20 17:53 | 239 | 0.047164 | 0.018315 | 1.7148 | 16.776 | 35.5063 | 0.79123
2019-03-20 17:53 | 240 | 0.043224 | 0.021 | 1.8654 | 17.9344 | 35.4688 | 0.77519
2019-03-20 17:53 | 241 | 0.045 | 0.020715 | 1.7872 | 16.6683 | 35.419 | 0.75007
2019-03-20 17:53 | 242 | 0.042096 | 0.022683 | 1.9002 | 18.3719 | 35.3797 | 0.77339
2019-03-20 17:53 | 243 | 0.043302 | 0.022575 | 1.8414 | 17.4686 | 35.3337 | 0.75643
2019-03-20 17:53 | 244 | 0.048763 | 0.020528 | 1.632 | 15.4803 | 35.292 | 0.75487
2019-03-20 17:53 | 245 | 0.051516 | 0.020344 | 1.5443 | 15.1249 | 35.2372 | 0.77918
2019-03-20 17:53 | 246 | 0.037724 | 0.028906 | 2.0886 | 20.3985 | 35.1928 | 0.76952
2019-03-20 17:53 | 247 | 0.04048 | 0.028619 | 1.9426 | 18.6748 | 35.1581 | 0.75594
2019-03-20 17:53 | 248 | 0.042856 | 0.028144 | 1.8333 | 18.0304 | 35.1117 | 0.77272
2019-03-20 17:53 | 249 | 0.036295 | 0.033756 | 2.1591 | 21.0601 | 35.0745 | 0.76437
2019-03-20 17:53 | 250 | 0.033054 | 0.037654 | 2.3603 | 23.0473 | 35.0384 | 0.7618
2019-03-20 17:53 | 251 | 0.034193 | 0.037865 | 2.282 | 22.1808 | 35.0067 | 0.75842
2019-03-20 17:53 | 252 | 0.054417 | 0.024593 | 1.4286 | 13.9599 | 34.9656 | 0.75966
2019-03-20 17:53 | 253 | 0.054379 | 0.02497 | 1.4244 | 13.9991 | 34.9025 | 0.76126
2019-03-20 17:53 | 254 | 0.048863 | 0.029171 | 1.5826 | 16.0882 | 34.8577 | 0.78612
2019-03-20 17:53 | 255 | 0.078526 | 0.018984 | 0.98086 | 10.8988 | 34.7961 | 0.85584
2019-03-20 17:53 | 256 | 0.10562 | 0.014927 | 0.7295 | 8.5244 | 34.7009 | 0.90031
2019-03-20 17:53 | 257 | 0.10995 | 0.015508 | 0.70225 | 9.2191 | 34.5906 | 1.0137
2019-03-20 17:53 | 258 | 0.11359 | 0.016304 | 0.69054 | 10.1507 | 34.4814 | 1.153
2019-03-20 17:53 | 259 | 0.13213 | 0.015468 | 0.61083 | 9.9208 | 34.3598 | 1.3108
2019-03-20 17:53 | 260 | 0.14919 | 0.015251 | 0.56472 | 10.79 | 34.2174 | 1.6098
2019-03-20 17:53 | 261 | 0.15875 | 0.01579 | 0.57224 | 11.7038 | 34.0636 | 1.858
2019-03-20 17:53 | 262 | 0.15721 | 0.017638 | 0.67722 | 13.9523 | 33.9025 | 2.1934
2019-03-20 17:53 | 263 | 0.1688 | 0.017969 | 0.74285 | 14.2688 | 33.7459 | 2.4086
2019-03-20 17:53 | 264 | 0.19215 | 0.017059 | 0.76172 | 13.7734 | 33.5623 | 2.6465
2019-03-20 17:53 | 265 | 0.19402 | 0.018028 | 0.87196 | 15.7621 | 33.3675 | 3.0582
2019-03-20 17:53 | 266 | 0.17243 | 0.021555 | 1.1184 | 19.2777 | 33.1811 | 3.324
2019-03-20 17:53 | 267 | 0.14348 | 0.027329 | 1.4996 | 24.8273 | 33.0245 | 3.5622
2019-03-20 17:53 | 268 | 0.14307 | 0.028857 | 1.6481 | 26.2173 | 32.8867 | 3.7508
2019-03-20 17:53 | 269 | 0.15768 | 0.027393 | 1.6048 | 24.6369 | 32.7342 | 3.8847
2019-03-20 17:53 | 270 | 0.15122 | 0.029621 | 1.7608 | 26.2272 | 32.5758 | 3.9661
2019-03-20 17:53 | 271 | 0.15921 | 0.029077 | 1.7437 | 25.3771 | 32.429 | 4.0402
2019-03-20 17:54 | 272 | 0.15751 | 0.030125 | 1.818 | 25.5357 | 32.26 | 4.0221
2019-03-20 17:54 | 273 | 0.13439 | 0.036002 | 2.195 | 29.948 | 32.1191 | 4.0246
2019-03-20 17:54 | 274 | 0.13599 | 0.035611 | 2.2251 | 29.3173 | 31.9861 | 3.9869
2019-03-20 17:54 | 275 | 0.12279 | 0.03915 | 2.5169 | 32.2612 | 31.8502 | 3.9612
2019-03-20 17:54 | 276 | 0.10462 | 0.045331 | 2.9951 | 36.3199 | 31.7408 | 3.7998
2019-03-20 17:54 | 277 | 0.13347 | 0.03512 | 2.3685 | 27.2536 | 31.6298 | 3.6375
2019-03-20 17:54 | 278 | 0.14327 | 0.032348 | 2.2152 | 24.6471 | 31.4783 | 3.5312
2019-03-20 17:54 | 279 | 0.11684 | 0.039235 | 2.6981 | 28.2576 | 31.3523 | 3.3017
2019-03-20 17:54 | 280 | 0.11335 | 0.040292 | 2.7677 | 28.4474 | 31.2403 | 3.2247
2019-03-20 17:54 | 281 | 0.097474 | 0.046837 | 3.1683 | 32.2921 | 31.1276 | 3.1477
2019-03-20 17:54 | 282 | 0.097667 | 0.046772 | 3.1102 | 31.2434 | 31.0409 | 3.0514
2019-03-20 17:54 | 283 | 0.12459 | 0.036865 | 2.3931 | 24.1407 | 30.9265 | 3.0076
2019-03-20 17:54 | 284 | 0.097665 | 0.04749 | 2.9941 | 30.4687 | 30.8038 | 2.9757
2019-03-20 17:54 | 285 | 0.09202 | 0.05084 | 3.1147 | 32.0057 | 30.726 | 2.9452
2019-03-20 17:54 | 286 | 0.13282 | 0.035395 | 2.1172 | 21.7283 | 30.6088 | 2.8859
2019-03-20 17:54 | 287 | 0.12912 | 0.036437 | 2.1406 | 22.3353 | 30.4707 | 2.8839
2019-03-20 17:54 | 288 | 0.10958 | 0.042904 | 2.4806 | 26.3343 | 30.3552 | 2.8857
2019-03-20 17:54 | 289 | 0.11071 | 0.041981 | 2.4122 | 26.1733 | 30.2465 | 2.8977
2019-03-20 17:54 | 290 | 0.12605 | 0.036435 | 2.0869 | 22.9607 | 30.1298 | 2.8942
2019-03-20 17:54 | 291 | 0.13144 | 0.034475 | 1.9814 | 22.2248 | 29.9966 | 2.9212
2019-03-20 17:54 | 292 | 0.1299 | 0.034651 | 1.9828 | 22.6944 | 29.869 | 2.9479
2019-03-20 17:54 | 293 | 0.12827 | 0.03491 | 1.9946 | 22.9908 | 29.737 | 2.949
2019-03-20 17:54 | 294 | 0.1309 | 0.03409 | 1.9385 | 22.6044 | 29.6113 | 2.9588
2019-03-20 17:54 | 295 | 0.13842 | 0.032325 | 1.8255 | 21.5509 | 29.4745 | 2.9831
2019-03-20 17:54 | 296 | 0.12251 | 0.036258 | 2.0566 | 24.8041 | 29.3401 | 3.0388
2019-03-20 17:54 | 297 | 0.11168 | 0.040027 | 2.265 | 27.4376 | 29.2281 | 3.0642
2019-03-20 17:54 | 298 | 0.12462 | 0.035878 | 2.0345 | 24.5459 | 29.1106 | 3.0588
2019-03-20 17:54 | 299 | 0.13495 | 0.033102 | 1.8853 | 22.6501 | 28.9789 | 3.0567
2019-03-20 17:54 | 300 | 0.14386 | 0.031055 | 1.7752 | 21.2879 | 28.8416 | 3.0625
2019-03-20 17:54 | 301 | 0.13614 | 0.032641 | 1.8809 | 22.636 | 28.6956 | 3.0816
2019-03-20 17:54 | 302 | 0.12237 | 0.035982 | 2.1008 | 25.6476 | 28.5706 | 3.1384
2019-03-20 17:54 | 303 | 0.13591 | 0.032054 | 1.8986 | 23.4517 | 28.4444 | 3.1873
2019-03-20 17:54 | 304 | 0.14197 | 0.030399 | 1.8268 | 22.8237 | 28.3001 | 3.2403
2019-03-20 17:54 | 305 | 0.14114 | 0.03022 | 1.8469 | 23.1027 | 28.1623 | 3.2607
2019-03-20 17:54 | 306 | 0.14177 | 0.030006 | 1.8561 | 23.2282 | 28.0177 | 3.293
2019-03-20 17:54 | 307 | 0.14018 | 0.030283 | 1.889 | 23.7912 | 27.8792 | 3.335
2019-03-20 17:54 | 308 | 0.14445 | 0.029555 | 1.8489 | 23.2318 | 27.7365 | 3.3558
2019-03-20 17:54 | 309 | 0.13288 | 0.031967 | 2.0272 | 25.4807 | 27.5945 | 3.3858
2019-03-20 17:54 | 310 | 0.10729 | 0.039807 | 2.5345 | 31.7476 | 27.4736 | 3.4063
2019-03-20 17:54 | 311 | 0.12125 | 0.035466 | 2.2515 | 28.0589 | 27.3704 | 3.4021
2019-03-20 17:54 | 312 | 0.13257 | 0.032806 | 2.0782 | 25.8292 | 27.2312 | 3.4242
2019-03-20 17:54 | 313 | 0.12163 | 0.036071 | 2.2788 | 28.4174 | 27.1104 | 3.4564
2019-03-20 17:54 | 314 | 0.12964 | 0.033969 | 2.1498 | 26.7096 | 26.9842 | 3.4626
2019-03-20 17:54 | 315 | 0.12179 | 0.036113 | 2.3008 | 28.144 | 26.8544 | 3.4276
2019-03-20 17:54 | 316 | 0.1224 | 0.035671 | 2.2982 | 27.9985 | 26.7386 | 3.4269
2019-03-20 17:54 | 317 | 0.13556 | 0.032201 | 2.0861 | 25.2295 | 26.6065 | 3.4201
2019-03-20 17:54 | 318 | 0.13595 | 0.031709 | 2.0838 | 25.3058 | 26.4702 | 3.4403
2019-03-20 17:54 | 319 | 0.14161 | 0.030166 | 1.9979 | 24.3994 | 26.3339 | 3.4553
2019-03-20 17:54 | 320 | 0.13885 | 0.03042 | 2.0428 | 25.1343 | 26.189 | 3.49
2019-03-20 17:54 | 321 | 0.13419 | 0.031321 | 2.1215 | 26.1368 | 26.0567 | 3.5072
2019-03-20 17:54 | 322 | 0.13842 | 0.030118 | 2.0626 | 25.5474 | 25.9188 | 3.5362
2019-03-20 17:54 | 323 | 0.12919 | 0.032152 | 2.2246 | 27.5813 | 25.7827 | 3.5632
2019-03-20 17:54 | 324 | 0.12998 | 0.031819 | 2.2263 | 27.5663 | 25.6579 | 3.583
2019-03-20 17:54 | 325 | 0.14374 | 0.028359 | 2.0294 | 24.7835 | 25.5194 | 3.5623
2019-03-20 17:54 | 326 | 0.14873 | 0.026904 | 1.9661 | 23.9479 | 25.3726 | 3.5618
2019-03-20 17:54 | 327 | 0.14135 | 0.027711 | 2.0788 | 25.1864 | 25.2253 | 3.5602
2019-03-20 17:54 | 328 | 0.13742 | 0.028382 | 2.1494 | 25.9868 | 25.0893 | 3.571
2019-03-20 17:54 | 329 | 0.13776 | 0.028166 | 2.1507 | 25.882 | 24.9489 | 3.5655
2019-03-20 17:54 | 330 | 0.15049 | 0.025866 | 1.9706 | 23.7201 | 24.8106 | 3.5697
2019-03-20 17:54 | 331 | 0.15817 | 0.0247 | 1.8752 | 22.7228 | 24.6493 | 3.594
2019-03-20 17:54 | 332 | 0.14799 | 0.026774 | 2.0048 | 24.5683 | 24.4987 | 3.6359
2019-03-20 17:55 | 333 | 0.14697 | 0.026971 | 2.0244 | 25.1416 | 24.3513 | 3.695
2019-03-20 17:55 | 334 | 0.15563 | 0.025697 | 1.9153 | 23.5652 | 24.2025 | 3.6675
2019-03-20 17:55 | 335 | 0.1519 | 0.026316 | 1.9691 | 24.5961 | 24.0431 | 3.7362
2019-03-20 17:55 | 336 | 0.14024 | 0.028813 | 2.1473 | 26.5834 | 23.9005 | 3.7281
2019-03-20 17:55 | 337 | 0.14709 | 0.027768 | 2.059 | 25.3324 | 23.7583 | 3.7261
2019-03-20 17:55 | 338 | 0.14647 | 0.028149 | 2.0799 | 25.4738 | 23.6072 | 3.7311
2019-03-20 17:55 | 339 | 0.16192 | 0.02565 | 1.8864 | 23.0242 | 23.4621 | 3.7281
2019-03-20 17:55 | 340 | 0.15811 | 0.026216 | 1.9357 | 23.4796 | 23.2884 | 3.7123
2019-03-20 17:55 | 341 | 0.13211 | 0.031225 | 2.3238 | 28.234 | 23.1509 | 3.73
2019-03-20 17:55 | 342 | 0.14601 | 0.028003 | 2.1109 | 25.6502 | 23.0149 | 3.7452
2019-03-20 17:55 | 343 | 0.15711 | 0.02573 | 1.9633 | 23.9958 | 22.8593 | 3.7699
2019-03-20 17:55 | 344 | 0.1377 | 0.028954 | 2.2372 | 27.4897 | 22.7077 | 3.7853
2019-03-20 17:55 | 345 | 0.1407 | 0.028149 | 2.195 | 27.0001 | 22.5785 | 3.7989
2019-03-20 17:55 | 346 | 0.16844 | 0.023373 | 1.8408 | 22.9563 | 22.4207 | 3.8667
2019-03-20 17:55 | 347 | 0.1508 | 0.026025 | 2.0608 | 25.583 | 22.2519 | 3.8579
2019-03-20 17:55 | 348 | 0.14813 | 0.026326 | 2.1072 | 26.4198 | 22.1164 | 3.9137
2019-03-20 17:55 | 349 | 0.15188 | 0.025581 | 2.0726 | 25.8462 | 21.9536 | 3.9256
2019-03-20 17:55 | 350 | 0.15892 | 0.024258 | 1.9887 | 24.8466 | 21.8115 | 3.9487
2019-03-20 17:55 | 351 | 0.16944 | 0.022777 | 1.8786 | 23.4769 | 21.6351 | 3.9779
2019-03-20 17:55 | 352 | 0.16741 | 0.022763 | 1.9092 | 24.0535 | 21.4756 | 4.0269
2019-03-20 17:55 | 353 | 0.16706 | 0.022783 | 1.9225 | 24.3309 | 21.3001 | 4.0646
2019-03-20 17:55 | 354 | 0.17104 | 0.022195 | 1.8854 | 24.0134 | 21.1402 | 4.1073
2019-03-20 17:55 | 355 | 0.18224 | 0.02061 | 1.7825 | 22.5632 | 20.9567 | 4.112
2019-03-20 17:55 | 356 | 0.17001 | 0.021941 | 1.9224 | 24.2739 | 20.7809 | 4.1267
2019-03-20 17:55 | 357 | 0.17859 | 0.020835 | 1.8388 | 23.1828 | 20.6125 | 4.1402
2019-03-20 17:55 | 358 | 0.17188 | 0.021502 | 1.9225 | 24.0206 | 20.4273 | 4.1288
2019-03-20 17:55 | 359 | 0.1567 | 0.023536 | 2.1178 | 26.3821 | 20.2707 | 4.1342
2019-03-20 17:55 | 360 | 0.16329 | 0.022687 | 2.0427 | 25.6511 | 20.1086 | 4.1886
2019-03-20 17:55 | 361 | 0.17074 | 0.021829 | 1.9663 | 25.0371 | 19.9436 | 4.2749
2019-03-20 17:55 | 362 | 0.16655 | 0.022382 | 2.0266 | 25.918 | 19.7703 | 4.3167
2019-03-20 17:55 | 363 | 0.15166 | 0.024453 | 2.2419 | 28.4582 | 19.6127 | 4.3158
2019-03-20 17:55 | 364 | 0.16664 | 0.022104 | 2.0515 | 26.0203 | 19.46 | 4.3361
2019-03-20 17:55 | 365 | 0.17239 | 0.021331 | 1.9953 | 25.1079 | 19.2817 | 4.3284
2019-03-20 17:55 | 366 | 0.14783 | 0.024632 | 2.3462 | 29.3345 | 19.122 | 4.3366
2019-03-20 17:55 | 367 | 0.16119 | 0.02234 | 2.1656 | 26.8485 | 18.977 | 4.3276
2019-03-20 17:55 | 368 | 0.18931 | 0.018902 | 1.8534 | 23.1661 | 18.7963 | 4.3855
2019-03-20 17:55 | 369 | 0.16845 | 0.021071 | 2.0905 | 26.0897 | 18.6103 | 4.3948
2019-03-20 17:55 | 370 | 0.1449 | 0.024768 | 2.4427 | 30.4345 | 18.4602 | 4.4099
2019-03-20 17:55 | 371 | 0.16755 | 0.021582 | 2.1292 | 26.4565 | 18.3095 | 4.4328
2019-03-20 17:55 | 372 | 0.17197 | 0.021184 | 2.0779 | 25.8099 | 18.1284 | 4.4387
2019-03-20 17:55 | 373 | 0.17239 | 0.021153 | 2.0854 | 26.0271 | 17.9671 | 4.4869
2019-03-20 17:55 | 374 | 0.16943 | 0.021709 | 2.1292 | 26.5735 | 17.7849 | 4.5025
2019-03-20 17:55 | 375 | 0.15654 | 0.023697 | 2.3231 | 28.8206 | 17.6302 | 4.5115
2019-03-20 17:55 | 376 | 0.16672 | 0.022323 | 2.1893 | 27.0215 | 17.4667 | 4.5051
2019-03-20 17:55 | 377 | 0.16309 | 0.023071 | 2.2487 | 27.4939 | 17.3 | 4.4839
2019-03-20 17:55 | 378 | 0.14309 | 0.02656 | 2.5715 | 31.3397 | 17.1438 | 4.4843
2019-03-20 17:55 | 379 | 0.16622 | 0.023227 | 2.2207 | 26.9418 | 17.0034 | 4.4782
2019-03-20 17:55 | 380 | 0.19151 | 0.020385 | 1.9329 | 23.2681 | 16.8109 | 4.4561
2019-03-20 17:55 | 381 | 0.16892 | 0.023369 | 2.1844 | 26.3464 | 16.6316 | 4.4503
2019-03-20 17:55 | 382 | 0.17472 | 0.022851 | 2.1175 | 25.9219 | 16.4674 | 4.5291
2019-03-20 17:55 | 383 | 0.1769 | 0.022812 | 2.0944 | 25.6004 | 16.2827 | 4.5287
2019-03-20 17:55 | 384 | 0.16455 | 0.024872 | 2.2553 | 27.4617 | 16.1166 | 4.5188
2019-03-20 17:55 | 385 | 0.17726 | 0.023187 | 2.0856 | 25.2564 | 15.9481 | 4.477
2019-03-20 17:55 | 386 | 0.17426 | 0.023886 | 2.1329 | 25.5998 | 15.7654 | 4.4611
2019-03-20 17:55 | 387 | 0.17035 | 0.024618 | 2.1806 | 26.0005 | 15.5996 | 4.4291
2019-03-20 17:55 | 388 | 0.18931 | 0.022392 | 1.9662 | 23.4395 | 15.4198 | 4.4374
2019-03-20 17:55 | 389 | 0.18066 | 0.023575 | 2.0583 | 24.5429 | 15.2278 | 4.434
2019-03-20 17:55 | 390 | 0.15827 | 0.027019 | 2.3457 | 27.8841 | 15.0622 | 4.4131
2019-03-20 17:55 | 391 | 0.15454 | 0.027829 | 2.4005 | 28.6065 | 14.9069 | 4.4208
2019-03-20 17:55 | 392 | 0.15211 | 0.028372 | 2.4328 | 28.926 | 14.7521 | 4.4
2019-03-20 17:56 | 393 | 0.15157 | 0.028482 | 2.4392 | 29.118 | 14.6014 | 4.4134
2019-03-20 17:56 | 394 | 0.17743 | 0.024483 | 2.0835 | 24.5527 | 14.4434 | 4.3564
2019-03-20 17:56 | 395 | 0.16421 | 0.026542 | 2.2418 | 26.2683 | 14.2556 | 4.3135
2019-03-20 17:56 | 396 | 0.15191 | 0.028832 | 2.418 | 28.3206 | 14.1151 | 4.3023
2019-03-20 17:56 | 397 | 0.16653 | 0.026222 | 2.1997 | 25.6865 | 13.9458 | 4.2776
2019-03-20 17:56 | 398 | 0.15212 | 0.028723 | 2.4005 | 28.0416 | 13.7881 | 4.2656
2019-03-20 17:56 | 399 | 0.15396 | 0.028186 | 2.3574 | 27.6334 | 13.6379 | 4.2545
2019-03-20 17:56 | 400 | 0.16864 | 0.025634 | 2.1463 | 25.304 | 13.4768 | 4.2673
2019-03-20 17:56 | 401 | 0.17348 | 0.024772 | 2.076 | 24.5328 | 13.3032 | 4.256
2019-03-20 17:56 | 402 | 0.15889 | 0.027193 | 2.2614 | 26.8858 | 13.1344 | 4.2718
2019-03-20 17:56 | 403 | 0.16971 | 0.025558 | 2.1105 | 25.0619 | 12.9801 | 4.2532
2019-03-20 17:56 | 404 | 0.16266 | 0.027083 | 2.2017 | 26.1586 | 12.7988 | 4.2549
2019-03-20 17:56 | 405 | 0.15402 | 0.028908 | 2.3243 | 27.6418 | 12.6551 | 4.2575
2019-03-20 17:56 | 406 | 0.16879 | 0.026569 | 2.1147 | 25.0373 | 12.4857 | 4.2261
2019-03-20 17:56 | 407 | 0.15689 | 0.028363 | 2.2733 | 26.6894 | 12.3233 | 4.1874
2019-03-20 17:56 | 408 | 0.15344 | 0.02912 | 2.3172 | 27.0431 | 12.1705 | 4.1496
2019-03-20 17:56 | 409 | 0.14999 | 0.02963 | 2.3617 | 27.2943 | 12.0163 | 4.0938
2019-03-20 17:56 | 410 | 0.14473 | 0.030766 | 2.4423 | 28.4498 | 11.8707 | 4.1175
2019-03-20 17:56 | 411 | 0.14462 | 0.0307 | 2.428 | 28.2596 | 11.7248 | 4.087
2019-03-20 17:56 | 412 | 0.17227 | 0.025932 | 2.0315 | 23.7984 | 11.5756 | 4.0998
2019-03-20 17:56 | 413 | 0.1552 | 0.02882 | 2.2479 | 26.2728 | 11.3908 | 4.0775
2019-03-20 17:56 | 414 | 0.13595 | 0.032636 | 2.5585 | 29.7859 | 11.2659 | 4.0495
2019-03-20 17:56 | 415 | 0.15511 | 0.028386 | 2.2354 | 26.1495 | 11.1099 | 4.0561
2019-03-20 17:56 | 416 | 0.15371 | 0.028359 | 2.2441 | 26.3166 | 10.9595 | 4.0451
2019-03-20 17:56 | 417 | 0.15915 | 0.027284 | 2.161 | 25.3325 | 10.8014 | 4.0317
2019-03-20 17:56 | 418 | 0.16145 | 0.026957 | 2.1173 | 24.617 | 10.6421 | 3.9744
2019-03-20 17:56 | 419 | 0.15746 | 0.027988 | 2.1696 | 24.5965 | 10.4803 | 3.8729
2019-03-20 17:56 | 420 | 0.1491 | 0.029816 | 2.2764 | 25.9688 | 10.3284 | 3.872
2019-03-20 17:56 | 421 | 0.14711 | 0.0305 | 2.2924 | 26.0872 | 10.1812 | 3.8376
2019-03-20 17:56 | 422 | 0.12551 | 0.036382 | 2.6783 | 30.3233 | 10.0384 | 3.8057
2019-03-20 17:56 | 423 | 0.12589 | 0.036376 | 2.6465 | 29.7858 | 9.9251 | 3.7496
2019-03-20 17:56 | 424 | 0.12611 | 0.036641 | 2.6211 | 29.1161 | 9.7862 | 3.6719
2019-03-20 17:56 | 425 | 0.11782 | 0.039461 | 2.7861 | 31.0381 | 9.6745 | 3.6569
2019-03-20 17:56 | 426 | 0.12801 | 0.036497 | 2.544 | 28.3741 | 9.5467 | 3.6323
2019-03-20 17:56 | 427 | 0.11779 | 0.039755 | 2.7448 | 30.3933 | 9.4227 | 3.5801
2019-03-20 17:56 | 428 | 0.11953 | 0.039242 | 2.6827 | 29.057 | 9.3081 | 3.4731
2019-03-20 17:56 | 429 | 0.13749 | 0.034518 | 2.3141 | 25.534 | 9.1797 | 3.5106
2019-03-20 17:56 | 430 | 0.14006 | 0.033916 | 2.2485 | 23.9816 | 9.037 | 3.359
2019-03-20 17:56 | 431 | 0.12743 | 0.037462 | 2.4419 | 26.9462 | 8.9033 | 3.4337
2019-03-20 17:56 | 432 | 0.13365 | 0.035977 | 2.305 | 25.3042 | 8.7778 | 3.382
2019-03-20 17:56 | 433 | 0.1424 | 0.033879 | 2.1433 | 23.7877 | 8.6353 | 3.3873
2019-03-20 17:56 | 434 | 0.13991 | 0.034322 | 2.1553 | 23.3515 | 8.4962 | 3.2672
2019-03-20 17:56 | 435 | 0.12036 | 0.039849 | 2.4731 | 26.8978 | 8.3599 | 3.2373
2019-03-20 17:56 | 436 | 0.10839 | 0.044299 | 2.7143 | 29.5238 | 8.2536 | 3.2001
2019-03-20 17:56 | 437 | 0.11686 | 0.041144 | 2.4926 | 27.7367 | 8.1381 | 3.2412
2019-03-20 17:56 | 438 | 0.11162 | 0 | 0 | 0 | 8.023 | 3.2511
2019-03-20 17:56 | 439 | 0.096669 | 0 | 0 | 0 | 7.9173 | 3.245
2019-03-20 17:56 | 440 | 0.09118 | 0 | 0 | 0 | 7.8269 | 3.2343
2019-03-20 17:56 | 441 | 0.10825 | 0 | 0 | 0 | 7.7295 | 3.2163
2019-03-20 17:56 | 442 | 0.10882 | 0 | 0 | 0 | 7.6143 | 3.1552
2019-03-20 17:56 | 443 | 0.094319 | 0 | 0 | 0 | 7.5154 | 3.0763
2019-03-20 17:56 | 444 | 0.10508 | 0 | 0 | 0 | 7.4196 | 3.0423
2019-03-20 17:56 | 445 | 0.12003 | 0 | 0 | 0 | 7.3042 | 3.0151
2019-03-20 17:56 | 446 | 0.10976 | 0 | 0 | 0 | 7.1858 | 2.9559
2019-03-20 17:56 | 447 | 0.099151 | 0 | 0 | 0 | 7.085 | 2.947
2019-03-20 17:56 | 448 | 0.10633 | 0 | 0 | 0 | 6.983 | 2.912
2019-03-20 17:56 | 449 | 0.11336 | 0 | 0 | 0 | 6.872 | 2.9162
2019-03-20 17:56 | 450 | 0.11659 | 0 | 0 | 0 | 6.7579 | 2.8581
2019-03-20 17:56 | 451 | 0.10027 | 0 | 0 | 0 | 6.644 | 2.8361
2019-03-20 17:56 | 452 | 0.077608 | 0 | 0 | 0 | 6.5585 | 2.7215
2019-03-20 17:57 | 453 | 0.091686 | 0 | 0 | 0 | 6.4801 | 2.6911
2019-03-20 17:57 | 454 | 0.09295 | 0 | 0 | 0 | 6.3775 | 2.656
2019-03-20 17:57 | 455 | 0.083799 | 0 | 0 | 0 | 6.2966 | 2.6045
2019-03-20 17:57 | 456 | 0.092382 | 0 | 0 | 0 | 6.2061 | 2.546
2019-03-20 17:57 | 457 | 0.088506 | 0 | 0 | 0 | 6.1145 | 2.4774
2019-03-20 17:57 | 458 | 0.082177 | 0 | 0 | 0 | 6.0294 | 2.4369
2019-03-20 17:57 | 459 | 0.097359 | 0 | 0 | 0 | 5.9449 | 2.3839
2019-03-20 17:57 | 460 | 0.11154 | 0 | 0 | 0 | 5.8354 | 2.3133
2019-03-20 17:57 | 461 | 0.088918 | 0 | 0 | 0 | 5.7306 | 2.3156
2019-03-20 17:57 | 462 | 0.087747 | 0 | 0 | 0 | 5.6533 | 2.2614
2019-03-20 17:57 | 463 | 0.083973 | 0 | 0 | 0 | 5.5552 | 2.258
2019-03-20 17:57 | 464 | 0.077656 | 0 | 0 | 0 | 5.4855 | 2.1883
2019-03-20 17:57 | 465 | 0.092035 | 0 | 0 | 0 | 5.3956 | 2.1616
2019-03-20 17:57 | 466 | 0.07446 | 0 | 0 | 0 | 5.3085 | 2.0895
2019-03-20 17:57 | 467 | 0.07078 | 0 | 0 | 0 | 5.2437 | 2.0242
2019-03-20 17:57 | 468 | 0.083957 | 0 | 0 | 0 | 5.1638 | 2.0018
2019-03-20 17:57 | 469 | 0.047779 | 0 | 0 | 0 | 5.0874 | 1.9855
2019-03-20 17:57 | 470 | 0.020501 | 0 | 0 | 0 | 5.0656 | 1.9783
2019-03-20 17:57 | 471 | 0.052833 | 0 | 0 | 0 | 5.032 | 1.9649
2019-03-20 17:57 | 472 | 0.054495 | 0 | 0 | 0 | 4.9662 | 1.9456
2019-03-20 17:57 | 473 | 0.046798 | 0 | 0 | 0 | 4.9258 | 1.9375
2019-03-20 17:57 | 474 | 0.054945 | 0 | 0 | 0 | 4.8698 | 1.9082
2019-03-20 17:57 | 475 | 0.031389 | 0 | 0 | 0 | 4.8224 | 1.8597
2019-03-20 17:57 | 476 | 0.043649 | 0 | 0 | 0 | 4.7989 | 1.8185
2019-03-20 17:57 | 477 | 0.057013 | 0 | 0 | 0 | 4.7351 | 1.7539
2019-03-20 17:57 | 478 | 0.025156 | 0 | 0 | 0 | 4.6948 | 1.7027
2019-03-20 17:57 | 479 | 0.039569 | 0 | 0 | 0 | 4.6743 | 1.648
2019-03-20 17:57 | 480 | 0.058439 | 0 | 0 | 0 | 4.6142 | 1.623
2019-03-20 17:57 | 481 | 0.044898 | 0 | 0 | 0 | 4.5648 | 1.5702
2019-03-20 17:57 | 482 | 0.050205 | 0 | 0 | 0 | 4.5205 | 1.4901
2019-03-20 17:57 | 483 | 0.05891 | 0 | 0 | 0 | 4.4638 | 1.4114
2019-03-20 17:57 | 484 | 0.04136 | 0 | 0 | 0 | 4.4087 | 1.3909
2019-03-20 17:57 | 485 | 0.041791 | 0 | 0 | 0 | 4.3772 | 1.3602
2019-03-20 17:57 | 486 | 0.050547 | 0 | 0 | 0 | 4.3236 | 1.3272
2019-03-20 17:57 | 487 | 0.02072 | 0 | 0 | 0 | 4.2845 | 1.2749
2019-03-20 17:57 | 488 | 0.026633 | 0 | 0 | 0 | 4.2739 | 1.2265
2019-03-20 17:57 | 489 | 0.045498 | 0 | 0 | 0 | 4.2281 | 1.1949
2019-03-20 17:57 | 490 | 0.025376 | 0 | 0 | 0 | 4.1916 | 1.1506
2019-03-20 17:57 | 491 | 0.03098 | 0 | 0 | 0 | 4.1724 | 1.1222
2019-03-20 17:57 | 492 | 0.022119 | 0 | 0 | 0 | 4.1328 | 1.0903
2019-03-20 17:57 | 493 | 0.0064833 | 0 | 0 | 0 | 4.1292 | 1.0749
2019-03-20 17:57 | 494 | 0.024327 | 0 | 0 | 0 | 4.1121 | 1.0385
2019-03-20 17:57 | 495 | 0.02454 | 0 | 0 | 0 | 4.0841 | 0.98909
2019-03-20 17:57 | 496 | 0.02042 | 0 | 0 | 0 | 4.0636 | 0.92625
2019-03-20 17:57 | 497 | 0.048819 | 0 | 0 | 0 | 4.0364 | 0.90356
2019-03-20 17:57 | 498 | 0.039497 | 0 | 0 | 0 | 3.9754 | 0.843
2019-03-20 17:57 | 499 | 0.0019651 | 0 | 0 | 0 | 3.9643 | 0.82698
2019-03-20 17:57 | 500 | 0.012898 | 0 | 0 | 0 | 3.9606 | 0.78411
2019-03-20 17:57 | 501 | 0.0061957 | 0 | 0 | 0 | 3.9413 | 0.72915
2019-03-20 17:57 | 502 | 0.0048878 | 0 | 0 | 0 | 3.9464 | 0.65962
2019-03-20 17:57 | 503 | 0.034687 | 0 | 0 | 0 | 3.9247 | 0.66489
2019-03-20 17:57 | 504 | 0.023349 | 0 | 0 | 0 | 3.8865 | 0.65267
2019-03-20 17:57 | 505 | 0.014262 | 0 | 0 | 0 | 3.8783 | 0.62859
2019-03-20 17:57 | 506 | 0.015476 | 0 | 0 | 0 | 3.8559 | 0.62219
2019-03-20 17:57 | 507 | -0.0051404 | 0 | 0 | 0 | 3.8515 | 0.59742
2019-03-20 17:57 | 508 | 0.009994 | 0 | 0 | 0 | 3.8578 | 0.55423
2019-03-20 17:57 | 509 | 0.025071 | 0 | 0 | 0 | 3.8317 | 0.51921
2019-03-20 17:57 | 510 | -0.0012188 | 0 | 0 | 0 | 3.817 | 0.49124
2019-03-20 17:57 | 511 | 0.0074737 | 0 | 0 | 0 | 3.8259 | 0.48155
2019-03-20 17:58 | 512 | 0.038511 | 0 | 0 | 0 | 3.7972 | 0.46573
2019-03-20 17:58 | 513 | 0.010348 | 0 | 0 | 0 | 3.7628 | 0.43243
2019-03-20 17:58 | 514 | -0.0045761 | 0 | 0 | 0 | 3.7742 | 0.45143
2019-03-20 17:58 | 515 | 0.0026037 | 0 | 0 | 0 | 3.7663 | 0.41933
2019-03-20 17:58 | 516 | 0.0027506 | 0 | 0 | 0 | 3.7699 | 0.40276
2019-03-20 17:58 | 517 | 0.0093072 | 0 | 0 | 0 | 3.7594 | 0.37337
2019-03-20 17:58 | 518 | 0.01067 | 0 | 0 | 0 | 3.7523 | 0.35701
2019-03-20 17:58 | 519 | 0.014537 | 0 | 0 | 0 | 3.7375 | 0.34236
2019-03-20 17:58 | 520 | 0.022823 | 0 | 0 | 0 | 3.7226 | 0.33008
2019-03-20 17:58 | 521 | 0.013794 | 0 | 0 | 0 | 3.6965 | 0.33889
2019-03-20 17:58 | 522 | -0.0087351 | 0 | 0 | 0 | 3.698 | 0.32671
2019-03-20 17:58 | 523 | 0.0014982 | 0 | 0 | 0 | 3.7063 | 0.31668
2019-03-20 17:58 | 524 | 0.0023335 | 0 | 0 | 0 | 3.6963 | 0.32069

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062444 | 0.062682 | 0 | 0.2095 | 2019-03-20 09:37 | Group A | 0.1 | 0.22 | Do Nothing | 0.20951 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025361 | 0.025478 | 0 | 0.0261 | 2019-03-20 09:37 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044638 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.00275 | 0.002762 | 0 | 0.0011 | 2019-03-20 09:37 | Group A | -0.01 | 0.1 | Do Nothing | -3.157e-09 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -20.1974 | 199.3757 | 0 | 100 | 2019-03-20 09:37 | None | 0 | 1000 | Do Nothing | 109.8733 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -12.5291 | 230.4066 | 0 | 100 | 2019-03-20 09:37 | None | 0 | 1000 | Do Nothing | 111.3885 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -15.3019 | 1447 | 0 |  | 2019-03-20 09:10 | None | -10 | 1000 | Do Nothing | 125.4931 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.0588 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0019684 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018413 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.9076 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23822 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4027 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.003709 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0051032 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0048967 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 6.6373e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0012758 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.27137 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3397 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5481 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.58 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.6854 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 19.293 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 68.9913 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 110.9696 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 65.9346 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0046842 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0412 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025015 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.0256 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN