## Cone Summary
Date: | 2019-03-21
Time: | 15:45
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B2S-Test2
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0411 |  | Start Test | 0 | 2019-03-21 15:50
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 8 | 2019-03-21 15:50
Sample Area | 1 | m^2 | Wetting | 152 | 2019-03-21 15:52
Cone Heat Flux | 75 | kW/m^2 | Dripping | 170 | 2019-03-21 15:53
Cone Orientation |  |  | Bottom Ignition | 291 | 2019-03-21 15:55
Sample Grid |  |  | Flame Out | 607 | 2019-03-21 16:00
Sample Thickness | 106 | mm
Sample Initial Mass | 46.6 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 23.8031 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 3.6868 | kW/m2 | Average Heat Release Rate | 1.8367 | kW/m2
Total Heat Release | 1.0243 | MJ/m2 | Average Mass Loss Rate | 0.081196 | g/m2s
Peak Mass Loss Rate | 0.16444 | g/m2s | Effective Heat of Combustion | 22.6358 | MJ/kg
Total Mass Lost | 43.7138 | g | Average Specific Extinction Area | 69.5508 | m2/kg
Time to Ignition | 8 | s | Peak CO Yield | 0.088838 | kg/kg
Time to Peak Heat Release | 431 | s | Peak CO2 Yield | 4.8766 | kg/kg
Time to Flame Out | 607 | s | Average CO Yield | 0.026505 | kg/kg
Average Heat Release Rate | 1.7071 | kW/m2 | Average CO2 Yield | 1.9985 | kg/kg
Average Mass Loss Rate | 0.072314 | g/m2s
Effective Heat of Combustion | 23.4313 | MJ/kg
Dripping | 170 | s | Time to dripping
HRRD | 0.9439 | kW | HRR at dripping time
THRD | 0.20054 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.2379 | kW | AHRR from Ignition to Dripping
Wetting | 152 | s | Time to wetting
HRRW | 0.98975 | kW | HRR at wetting time
THRW | 0.18296 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.2174 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-21 15:45 | 1553197519.8596 | -271 | 0.20946 | 0.000461 | -3e-06 | 100.2749 | 100.402 | 124.348 | 21.2612 | 0.001867 | 0.001768 | -152.098 | 0.24155 | 8.4007 | 0.003531 | 0.00456 | 0.004482 | 5.2e-05 | 0.001182 | 0.29302 | 1.3414 | 2.5478 | 98.92 | 20.89 | 38.56 | 66.7588 | 106.8331 | 64.3752 | 0.00964 | 1.046 | 0.024946 | 23.8497 | -0.011328 | 0.1437 | 4e-06 | 0.000447 | 0.000447 | 13.2034
2019-03-21 15:45 | 1553197520.8597 | -270 | 0.20946 | 0.000461 | -3e-06 | 100.2749 | 100.402 | 124.348 | 21.2612 | 0.001867 | 0.001768 | -152.098 | 0.24155 | 8.4007 | 0.003531 | 0.00456 | 0.004482 | 5.2e-05 | 0.001182 | 0.29302 | 1.3414 | 2.5478 | 98.92 | 20.89 | 38.562 | 66.7588 | 106.8331 | 64.3752 | 0.009641 | 1.046 | 0.024946 | 23.8497 | -0.011328 | 0.1437 | 4e-06 | 0.000447 | 0.000447 | 13.2024
2019-03-21 15:45 | 1553197521.8581 | -269 | 0.20946 | 0.00046 | -3e-06 | 100.2836 | 100.3707 | 124.348 | 21.2612 | 0.001923 | 0.001758 | -152.098 | 0.24135 | 8.4032 | 0.003516 | 0.004475 | 0.004353 | 4.9e-05 | 0.001183 | 0.29911 | 1.3287 | 2.5478 | 98.92 | 20.89 | 38.567 | 68.0997 | 106.4686 | 64.1299 | 0.009642 | 1.0467 | 0.024955 | 23.8411 | -0.007766 | 0.14358 | 5e-06 | 0.000554 | 0.000554 | 10.6467
2019-03-21 15:45 | 1553197522.8578 | -268 | 0.20948 | 0.000465 | -3e-06 | 100.2055 | 100.2961 | 124.348 | 21.2612 | 0.001895 | 0.001755 | -152.098 | 0.24149 | 8.4001 | 0.003418 | 0.004944 | 0.004696 | 5.2e-05 | 0.001179 | 0.27865 | 1.2994 | 2.548 | 98.92 | 20.89 | 38.568 | 67.4165 | 104.1 | 64.0465 | 0.009642 | 1.047 | 0.024959 | 23.8381 | -0.008081 | 0.14366 | 0 | 0.000168 | 0.000168 | 35.0819
2019-03-21 15:45 | 1553197523.8589 | -267 | 0.20944 | 0.000461 | -3e-06 | 100.176 | 100.2205 | 124.348 | 21.2612 | 0.001826 | 0.001688 | -152.098 | 0.24146 | 8.402 | 0.003501 | 0.005224 | 0.005016 | 5.1e-05 | 0.001181 | 0.26864 | 1.3055 | 2.5479 | 98.92 | 20.89 | 38.57 | 65.7717 | 106.1152 | 62.4541 | 0.009643 | 1.052 | 0.025018 | 23.7818 | -0.003969 | 0.14365 | 4e-06 | 0.00053 | 0.00053 | 11.1323
2019-03-21 15:45 | 1553197524.8583 | -266 | 0.20944 | 0.000464 | -3e-06 | 100.2097 | 100.2992 | 124.348 | 21.2612 | 0.00194 | 0.001799 | -152.098 | 0.24189 | 8.4008 | 0.00349 | 0.004976 | 0.00477 | 5e-05 | 0.001179 | 0.2792 | 1.2966 | 2.5479 | 98.92 | 20.89 | 38.572 | 68.5154 | 105.8399 | 65.109 | 0.009643 | 1.0437 | 0.024919 | 23.8757 | -0.007983 | 0.1439 | 0 | 0.000211 | 0.000211 | 28.0697
2019-03-21 15:45 | 1553197525.858 | -265 | 0.20943 | 0.000464 | -3e-06 | 100.1832 | 100.3271 | 124.348 | 21.2612 | 0.001852 | 0.001816 | -152.098 | 0.24166 | 8.3998 | 0.003458 | 0.004691 | 0.00458 | 5.2e-05 | 0.001181 | 0.28208 | 1.3149 | 2.5478 | 98.92 | 20.89 | 38.573 | 66.3872 | 105.0782 | 65.5281 | 0.009643 | 1.0424 | 0.024904 | 23.8904 | -0.01284 | 0.14376 | 0 | 0.000152 | 0.000152 | 38.8003
2019-03-21 15:45 | 1553197526.858 | -264 | 0.20943 | 0.000462 | -3e-06 | 100.1653 | 100.1898 | 124.348 | 21.2612 | 0.0019 | 0.001763 | -152.098 | 0.24151 | 8.401 | 0.003443 | 0.005035 | 0.004763 | 4.9e-05 | 0.001183 | 0.25299 | 1.3122 | 2.5478 | 98.92 | 20.89 | 38.573 | 67.5343 | 104.7047 | 64.2558 | 0.009643 | 1.0464 | 0.024951 | 23.8455 | -0.002191 | 0.14367 | 2e-06 | 0.000358 | 0.000358 | 16.4989
2019-03-21 15:45 | 1553197527.8587 | -263 | 0.20943 | 0.000463 | -3e-06 | 100.1421 | 100.2144 | 124.348 | 21.2612 | 0.00185 | 0.001777 | -152.098 | 0.24106 | 8.4033 | 0.003579 | 0.004557 | 0.004376 | 5.3e-05 | 0.00118 | 0.27992 | 1.2828 | 2.5479 | 98.92 | 20.89 | 38.573 | 66.3514 | 108.0201 | 64.5845 | 0.009643 | 1.0453 | 0.024939 | 23.8571 | -0.006455 | 0.14341 | 2e-06 | 0.000323 | 0.000323 | 18.2227
2019-03-21 15:45 | 1553197528.8595 | -262 | 0.20943 | 0.000463 | -3e-06 | 100.0646 | 100.1367 | 124.348 | 21.2612 | 0.001859 | 0.001782 | -152.098 | 0.24077 | 8.4028 | 0.003609 | 0.005398 | 0.005222 | 5e-05 | 0.001178 | 0.28208 | 1.3021 | 2.5479 | 98.92 | 20.89 | 38.572 | 66.5464 | 108.7346 | 64.707 | 0.009643 | 1.045 | 0.024934 | 23.8615 | -0.006443 | 0.14324 | 1e-06 | 0.000258 | 0.000258 | 22.8279
2019-03-21 15:45 | 1553197529.8582 | -261 | 0.20945 | 0.000463 | -3e-06 | 100.0238 | 100.213 | 124.348 | 21.2612 | 0.001834 | 0.001822 | -152.098 | 0.24155 | 8.4026 | 0.003555 | 0.004521 | 0.004324 | 4.8e-05 | 0.001177 | 0.28854 | 1.3056 | 2.5478 | 98.92 | 20.88 | 38.568 | 65.9473 | 107.4301 | 65.6577 | 0.009636 | 1.042 | 0.024899 | 23.895 | -0.016905 | 0.1437 | 2e-06 | 0.000318 | 0.000318 | 18.5456
2019-03-21 15:46 | 1553197530.8587 | -260 | 0.20943 | 0.000459 | -3e-06 | 100.0113 | 100.1847 | 124.0182 | 21.2612 | 0.00188 | 0.001755 | -152.1063 | 0.24094 | 8.4047 | 0.003533 | 0.005471 | 0.005376 | 5.1e-05 | 0.001177 | 0.24906 | 1.3108 | 2.5479 | 98.92 | 20.88 | 38.567 | 67.0731 | 106.8899 | 64.045 | 0.009636 | 1.047 | 0.024926 | 23.8064 | -0.015501 | 0.14334 | 7e-06 | 0.000679 | 0.000679 | 8.6804
2019-03-21 15:46 | 1553197531.8587 | -259 | 0.20942 | 0.000461 | -3e-06 | 99.9725 | 99.9364 | 123.825 | 21.2612 | 0.001889 | 0.001761 | -152.1107 | 0.24123 | 8.4076 | 0.003541 | 0.004554 | 0.004325 | 5.3e-05 | 0.001178 | 0.26001 | 1.343 | 2.5478 | 98.92 | 20.88 | 38.565 | 67.2823 | 107.0772 | 64.2106 | 0.009635 | 1.0465 | 0.0249 | 23.7937 | 0.00323 | 0.14351 | 3e-06 | 0.000391 | 0.000391 | 15.085
2019-03-21 15:46 | 1553197532.8584 | -258 | 0.20943 | 0.000462 | -3e-06 | 99.9689 | 100.0172 | 123.3643 | 21.2612 | 0.001909 | 0.001818 | -152.1256 | 0.24123 | 8.4049 | 0.003536 | 0.00416 | 0.004061 | 5.2e-05 | 0.001178 | 0.24818 | 1.3346 | 2.5478 | 98.92 | 20.89 | 38.562 | 67.7668 | 106.9568 | 65.5663 | 0.009641 | 1.0423 | 0.024804 | 23.7971 | -0.004312 | 0.14351 | 3e-06 | 0.000452 | 0.000452 | 13.0442
2019-03-21 15:46 | 1553197533.8593 | -257 | 0.20941 | 0.000458 | -3e-06 | 99.966 | 99.9283 | 123.5302 | 21.2612 | 0.001861 | 0.001785 | -152.118 | 0.24023 | 8.4032 | 0.003647 | 0.004847 | 0.004601 | 5e-05 | 0.001179 | 0.27704 | 1.3265 | 2.5479 | 98.92 | 20.89 | 38.561 | 66.6121 | 109.6712 | 64.789 | 0.00964 | 1.0447 | 0.024849 | 23.7858 | 0.003375 | 0.14291 | 7e-06 | 0.000761 | 0.000761 | 7.7225
2019-03-21 15:46 | 1553197534.8589 | -256 | 0.20941 | 0.000455 | -3e-06 | 99.8551 | 100.0573 | 123.5926 | 21.2612 | 0.00193 | 0.001771 | -152.1321 | 0.24135 | 8.4055 | 0.00363 | 0.005338 | 0.005206 | 5.2e-05 | 0.001184 | 0.26985 | 1.3044 | 2.5478 | 98.92 | 20.89 | 38.558 | 68.2652 | 109.2392 | 64.4447 | 0.00964 | 1.0458 | 0.024868 | 23.7796 | -0.018095 | 0.14358 | 1.1e-05 | 0.001019 | 0.001019 | 5.7896
2019-03-21 15:46 | 1553197535.8585 | -255 | 0.20945 | 0.000457 | -3e-06 | 99.8936 | 99.8356 | 123.8285 | 21.2612 | 0.001915 | 0.001744 | -152.135 | 0.24134 | 8.4033 | 0.003627 | 0.005181 | 0.005101 | 5.3e-05 | 0.001183 | 0.2469 | 1.3238 | 2.548 | 98.92 | 20.89 | 38.558 | 67.9048 | 109.1633 | 63.7841 | 0.00964 | 1.0478 | 0.024916 | 23.779 | 0.00519 | 0.14357 | 8e-06 | 0.000809 | 0.000809 | 7.2966
2019-03-21 15:46 | 1553197536.8596 | -254 | 0.20944 | 0.000458 | -3e-06 | 99.8364 | 100.04 | 123.8206 | 21.2612 | 0.001907 | 0.001733 | -152.134 | 0.24133 | 8.4017 | 0.003642 | 0.005079 | 0.004878 | 5.2e-05 | 0.001186 | 0.25946 | 1.297 | 2.548 | 98.92 | 20.89 | 38.558 | 67.7168 | 109.5412 | 63.5312 | 0.00964 | 1.0486 | 0.024925 | 23.7693 | -0.018223 | 0.14357 | 8e-06 | 0.000778 | 0.000778 | 7.5867
2019-03-21 15:46 | 1553197537.8587 | -253 | 0.20946 | 0.000464 | -3e-06 | 99.8551 | 99.8976 | 123.8729 | 21.2612 | 0.001862 | 0.001718 | -152.1437 | 0.24141 | 8.4013 | 0.003627 | 0.00479 | 0.004636 | 4.9e-05 | 0.001186 | 0.27201 | 1.285 | 2.5479 | 98.92 | 20.88 | 38.557 | 66.6371 | 109.1665 | 63.1564 | 0.009633 | 1.0498 | 0.024944 | 23.7611 | -0.003804 | 0.14361 | 0 | 0.000185 | 0.000185 | 31.9527
2019-03-21 15:46 | 1553197538.8591 | -252 | 0.20944 | 0.000458 | -3e-06 | 99.7938 | 99.8809 | 123.6821 | 21.2612 | 0.001841 | 0.001745 | -152.1655 | 0.24094 | 8.4033 | 0.003613 | 0.005056 | 0.004905 | 5e-05 | 0.001185 | 0.24906 | 1.3386 | 2.548 | 98.92 | 20.89 | 38.556 | 66.1369 | 108.8339 | 63.8255 | 0.009639 | 1.0477 | 0.0249 | 23.7664 | -0.007809 | 0.14333 | 7e-06 | 0.000706 | 0.000706 | 8.3456
2019-03-21 15:46 | 1553197539.8582 | -251 | 0.20944 | 0.00046 | -2e-06 | 99.8533 | 99.868 | 123.6419 | 21.2612 | 0.001893 | 0.001733 | -152.1975 | 0.24127 | 8.4029 | 0.003611 | 0.004667 | 0.004479 | 5.3e-05 | 0.001189 | 0.26842 | 1.3297 | 2.5479 | 98.92 | 20.89 | 38.556 | 67.3713 | 108.7749 | 63.528 | 0.009639 | 1.0486 | 0.024907 | 23.7521 | -0.00131 | 0.14353 | 4e-06 | 0.000486 | 0.000486 | 12.1259
2019-03-21 15:46 | 1553197540.8588 | -250 | 0.20946 | 0.00046 | -3e-06 | 99.6659 | 99.6306 | 123.7816 | 21.2612 | 0.001856 | 0.001753 | -152.2055 | 0.24107 | 8.4029 | 0.003634 | 0.004766 | 0.004565 | 5.2e-05 | 0.001187 | 0.28047 | 1.3326 | 2.5479 | 98.92 | 20.89 | 38.555 | 66.481 | 109.3386 | 64.0067 | 0.009639 | 1.0471 | 0.024903 | 23.7824 | 0.003174 | 0.14341 | 5e-06 | 0.000584 | 0.000584 | 10.1009
2019-03-21 15:46 | 1553197541.858 | -249 | 0.20943 | 0.000459 | -3e-06 | 99.7309 | 99.6568 | 123.994 | 21.2612 | 0.001806 | 0.001751 | -152.1981 | 0.24172 | 8.4031 | 0.003645 | 0.004419 | 0.004249 | 5.3e-05 | 0.001186 | 0.2907 | 1.3394 | 2.5479 | 98.92 | 20.89 | 38.555 | 65.2947 | 109.6153 | 63.949 | 0.009639 | 1.0473 | 0.024927 | 23.8007 | 0.006646 | 0.1438 | 6e-06 | 0.000637 | 0.000637 | 9.2835
2019-03-21 15:46 | 1553197542.8589 | -248 | 0.20943 | 0.000459 | -3e-06 | 99.6665 | 99.539 | 124.3186 | 21.2612 | 0.001856 | 0.001726 | -152.185 | 0.24143 | 8.4023 | 0.003728 | 0.004678 | 0.004505 | 5.1e-05 | 0.001183 | 0.25337 | 1.334 | 2.548 | 98.92 | 20.89 | 38.553 | 66.4748 | 111.6309 | 63.3617 | 0.009638 | 1.0491 | 0.024981 | 23.8111 | 0.011456 | 0.14363 | 6e-06 | 0.000664 | 0.000664 | 8.8836
2019-03-21 15:46 | 1553197543.8589 | -247 | 0.20944 | 0.000458 | -3e-06 | 99.6942 | 99.6721 | 124.2386 | 21.2612 | 0.001902 | 0.001753 | -152.1608 | 0.2416 | 8.4028 | 0.003686 | 0.004763 | 0.004592 | 5.2e-05 | 0.00118 | 0.26217 | 1.3082 | 2.5479 | 98.92 | 20.89 | 38.552 | 67.5944 | 110.6074 | 64.0123 | 0.009638 | 1.0471 | 0.024949 | 23.8264 | 0.001984 | 0.14373 | 7e-06 | 0.00073 | 0.00073 | 8.0882
2019-03-21 15:46 | 1553197544.859 | -246 | 0.20944 | 0.00046 | -3e-06 | 99.6126 | 99.6156 | 124.3431 | 21.2612 | 0.001842 | 0.001783 | -152.1484 | 0.24077 | 8.4027 | 0.00359 | 0.00516 | 0.00499 | 5.3e-05 | 0.00118 | 0.2568 | 1.3028 | 2.5479 | 98.92 | 20.88 | 38.551 | 66.158 | 108.2816 | 64.7265 | 0.009632 | 1.0449 | 0.024933 | 23.8617 | -0.000269 | 0.14323 | 5e-06 | 0.000535 | 0.000535 | 10.9961
2019-03-21 15:46 | 1553197545.858 | -245 | 0.20946 | 0.000461 | -3e-06 | 99.6064 | 99.6946 | 124.6273 | 21.2612 | 0.001886 | 0.00179 | -152.1267 | 0.24166 | 8.4019 | 0.00362 | 0.00519 | 0.005049 | 5.4e-05 | 0.001179 | 0.26985 | 1.3157 | 2.5479 | 98.92 | 20.89 | 38.549 | 67.2027 | 109.0074 | 64.8913 | 0.009637 | 1.0444 | 0.024955 | 23.8948 | -0.007923 | 0.14376 | 4e-06 | 0.000459 | 0.000459 | 12.8748
2019-03-21 15:46 | 1553197546.8582 | -244 | 0.20945 | 0.000459 | -3e-06 | 99.5888 | 99.6108 | 124.7329 | 21.2612 | 0.001852 | 0.00177 | -152.1344 | 0.24058 | 8.4026 | 0.003552 | 0.004451 | 0.004222 | 5.4e-05 | 0.001183 | 0.2724 | 1.3025 | 2.5478 | 98.92 | 20.89 | 38.548 | 66.3935 | 107.3638 | 64.4259 | 0.009637 | 1.0458 | 0.024983 | 23.8884 | -0.001981 | 0.14312 | 7e-06 | 0.000665 | 0.000665 | 8.8427
2019-03-21 15:46 | 1553197547.8582 | -243 | 0.20945 | 0.000461 | -3e-06 | 99.5429 | 99.4995 | 124.8885 | 21.2612 | 0.001896 | 0.001725 | -152.1395 | 0.24027 | 8.4007 | 0.00359 | 0.004795 | 0.004585 | 5.3e-05 | 0.00118 | 0.23828 | 1.3641 | 2.5481 | 98.92 | 20.89 | 38.545 | 67.4499 | 108.2799 | 63.3273 | 0.009636 | 1.0492 | 0.025039 | 23.8644 | 0.003903 | 0.14294 | 4e-06 | 0.000501 | 0.000501 | 11.7165
2019-03-21 15:46 | 1553197548.8582 | -242 | 0.20944 | 0.00046 | -3e-06 | 99.5974 | 99.635 | 124.9285 | 21.2612 | 0.001901 | 0.001743 | -152.1411 | 0.24039 | 8.4027 | 0.003574 | 0.005136 | 0.004996 | 5.3e-05 | 0.001185 | 0.27295 | 1.3526 | 2.5481 | 98.92 | 20.89 | 38.543 | 67.56 | 107.8836 | 63.7601 | 0.009636 | 1.0479 | 0.025027 | 23.8836 | -0.003376 | 0.14301 | 5e-06 | 0.000518 | 0.000518 | 11.3497
2019-03-21 15:46 | 1553197549.8593 | -241 | 0.20947 | 0.000461 | -3e-06 | 99.5497 | 99.5138 | 125.1416 | 21.2612 | 0.001882 | 0.001754 | -152.1151 | 0.24124 | 8.4041 | 0.003606 | 0.005487 | 0.005351 | 5.4e-05 | 0.001185 | 0.27007 | 1.3575 | 2.5479 | 98.92 | 20.88 | 38.542 | 67.1044 | 108.6566 | 64.0224 | 0.00963 | 1.0471 | 0.025039 | 23.9132 | 0.003233 | 0.14351 | 4e-06 | 0.000504 | 0.000504 | 11.7007
2019-03-21 15:46 | 1553197550.8593 | -240 | 0.20945 | 0.000457 | -3e-06 | 99.5423 | 99.4334 | 125.1193 | 21.2612 | 0.001922 | 0.001793 | -152.1108 | 0.24046 | 8.4019 | 0.003616 | 0.004358 | 0.004177 | 5.1e-05 | 0.001184 | 0.25874 | 1.3417 | 2.5479 | 98.92 | 20.88 | 38.539 | 68.0772 | 108.9058 | 64.9724 | 0.009629 | 1.0441 | 0.025002 | 23.9448 | 0.009791 | 0.14305 | 9e-06 | 0.000865 | 0.000865 | 6.7973
2019-03-21 15:46 | 1553197551.8595 | -239 | 0.20944 | 0.000456 | -3e-06 | 99.4964 | 99.493 | 125.0749 | 21.2612 | 0.001891 | 0.001703 | -152.123 | 0.2419 | 8.4013 | 0.003649 | 0.005401 | 0.005193 | 5.3e-05 | 0.001185 | 0.25857 | 1.3738 | 2.5481 | 98.92 | 20.89 | 38.538 | 67.3345 | 109.7036 | 62.8025 | 0.009635 | 1.0509 | 0.025078 | 23.8636 | 0.000308 | 0.14391 | 1e-05 | 0.000912 | 0.000912 | 6.4879
2019-03-21 15:46 | 1553197552.8585 | -238 | 0.20948 | 0.000455 | -3e-06 | 99.5173 | 99.5935 | 125.0651 | 21.2612 | 0.001908 | 0.001797 | -152.1324 | 0.24034 | 8.4031 | 0.003635 | 0.005562 | 0.005441 | 5.2e-05 | 0.001188 | 0.27129 | 1.3677 | 2.5481 | 98.92 | 20.89 | 38.537 | 67.7338 | 109.3646 | 65.0621 | 0.009634 | 1.0439 | 0.024993 | 23.9427 | -0.006846 | 0.14298 | 1.2e-05 | 0.001115 | 0.001115 | 5.2691
2019-03-21 15:46 | 1553197553.8586 | -237 | 0.20947 | 0.000465 | -3e-06 | 99.4931 | 99.6558 | 125.4664 | 21.2612 | 0.001831 | 0.001761 | -152.1474 | 0.24134 | 8.4006 | 0.003683 | 0.004299 | 0.004141 | 5.3e-05 | 0.001187 | 0.26034 | 1.3234 | 2.5481 | 98.92 | 20.88 | 38.537 | 65.881 | 110.5333 | 64.194 | 0.009628 | 1.0465 | 0.025065 | 23.9503 | -0.014611 | 0.14357 | -2e-06 | 4.6e-05 | 4.6e-05 | 129.4765
2019-03-21 15:46 | 1553197554.8576 | -236 | 0.20947 | 0.000457 | -3e-06 | 99.5357 | 99.4882 | 125.6359 | 21.2612 | 0.001871 | 0.001701 | -152.1584 | 0.24131 | 8.402 | 0.003617 | 0.005437 | 0.005181 | 5.2e-05 | 0.001187 | 0.26465 | 1.3427 | 2.5479 | 98.92 | 20.89 | 38.538 | 66.8384 | 108.9255 | 62.7572 | 0.009635 | 1.051 | 0.025136 | 23.9154 | 0.004271 | 0.14356 | 8e-06 | 0.000806 | 0.000806 | 7.323
2019-03-21 15:46 | 1553197555.8587 | -235 | 0.20947 | 0.00046 | -3e-06 | 99.503 | 99.5155 | 125.5732 | 21.2612 | 0.00188 | 0.001754 | -152.169 | 0.24067 | 8.4031 | 0.003696 | 0.005546 | 0.005462 | 5.2e-05 | 0.001187 | 0.28512 | 1.3041 | 2.5478 | 98.92 | 20.89 | 38.538 | 67.0529 | 110.8401 | 64.0262 | 0.009635 | 1.0471 | 0.025082 | 23.9546 | -0.001124 | 0.14317 | 6e-06 | 0.000669 | 0.000669 | 8.7937
2019-03-21 15:46 | 1553197556.8586 | -234 | 0.20946 | 0.000465 | -3e-06 | 99.4872 | 99.3762 | 125.454 | 21.2612 | 0.001873 | 0.001725 | -152.1731 | 0.24093 | 8.4017 | 0.003703 | 0.004979 | 0.004841 | 5.3e-05 | 0.00119 | 0.25586 | 1.2952 | 2.5479 | 98.92 | 20.89 | 38.539 | 66.9048 | 111.0226 | 63.329 | 0.009635 | 1.0492 | 0.025096 | 23.9184 | 0.009986 | 0.14333 | -1e-06 | 5.7e-05 | 5.7e-05 | 103.3445
2019-03-21 15:46 | 1553197557.8592 | -233 | 0.20947 | 0.000463 | -3e-06 | 99.4815 | 99.3762 | 125.2444 | 21.2612 | 0.001864 | 0.001746 | -152.1827 | 0.24133 | 8.3999 | 0.00369 | 0.005206 | 0.00493 | 5.3e-05 | 0.001192 | 0.27992 | 1.3112 | 2.548 | 98.92 | 20.88 | 38.541 | 66.6698 | 110.702 | 63.8457 | 0.009629 | 1.0476 | 0.025056 | 23.9168 | 0.009477 | 0.14357 | 1e-06 | 0.000299 | 0.000299 | 19.7677
2019-03-21 15:46 | 1553197558.8578 | -232 | 0.20944 | 0.000469 | -3e-06 | 99.4035 | 99.3298 | 125.4622 | 21.2612 | 0.001896 | 0.001745 | -152.172 | 0.24091 | 8.4019 | 0.003718 | 0.005286 | 0.005171 | 5.3e-05 | 0.00119 | 0.26626 | 1.3008 | 2.548 | 98.92 | 20.89 | 38.542 | 67.4571 | 111.3823 | 63.8123 | 0.009636 | 1.0477 | 0.025079 | 23.9364 | 0.006626 | 0.14332 | -6e-06 | -0.000247 | -0.000247 | -23.8217
2019-03-21 15:46 | 1553197559.8589 | -231 | 0.20947 | 0.000465 | -3e-06 | 99.455 | 99.1827 | 125.1097 | 21.2612 | 0.001885 | 0.001768 | -152.1605 | 0.2406 | 8.4015 | 0.003678 | 0.005014 | 0.004864 | 5.4e-05 | 0.00119 | 0.28224 | 1.3602 | 2.5481 | 98.92 | 20.89 | 38.544 | 67.1832 | 110.4078 | 64.3667 | 0.009636 | 1.046 | 0.025023 | 23.9224 | 0.024522 | 0.14313 | -1.2e-05 | -0.000711 | -0.000711 | -8.2707
2019-03-21 15:46 | 1553197560.8584 | -230 | 0.20946 | 0.000461 | -3e-06 | 99.3528 | 99.3612 | 125.0397 | 21.2612 | 0.001867 | 0.001762 | -152.1569 | 0.24133 | 8.4 | 0.003657 | 0.005359 | 0.005143 | 5.4e-05 | 0.00119 | 0.26952 | 1.3169 | 2.5479 | 98.92 | 20.89 | 38.548 | 66.7509 | 109.8944 | 64.2339 | 0.009637 | 1.0464 | 0.025021 | 23.911 | -0.000754 | 0.14357 | 6.7e-05 | 0.005238 | 0.005238 | 1.1265
2019-03-21 15:46 | 1553197561.858 | -229 | 0.20949 | 0.000462 | -2e-06 | 99.3921 | 99.2556 | 125.2329 | 21.2612 | 0.001894 | 0.001781 | -152.1608 | 0.24125 | 8.4013 | 0.003657 | 0.005164 | 0.005099 | 5.4e-05 | 0.001192 | 0.25642 | 1.3249 | 2.5478 | 98.92 | 20.89 | 38.55 | 67.4102 | 109.907 | 64.6899 | 0.009638 | 1.045 | 0.025023 | 23.9456 | 0.012295 | 0.14352 | -9.5e-05 | -0.006978 | -0.006978 | -0.84533
2019-03-21 15:46 | 1553197562.8581 | -228 | 0.20949 | 0.000468 | -3e-06 | 99.4392 | 99.4504 | 125.009 | 21.2612 | 0.001882 | 0.00179 | -152.1588 | 0.24118 | 8.3997 | 0.003663 | 0.004397 | 0.004258 | 5.1e-05 | 0.001189 | 0.27704 | 1.3065 | 2.5479 | 98.92 | 20.89 | 38.552 | 67.1092 | 110.0529 | 64.8927 | 0.009638 | 1.0444 | 0.024993 | 23.9314 | -0.001008 | 0.14348 | -9.5e-05 | -0.006931 | -0.006931 | -0.85082
2019-03-21 15:46 | 1553197563.8595 | -227 | 0.2095 | 0.000467 | -3e-06 | 99.3871 | 99.2069 | 124.658 | 21.2612 | 0.001892 | 0.001773 | -152.1594 | 0.24131 | 8.4007 | 0.003734 | 0.005 | 0.004739 | 5.3e-05 | 0.00119 | 0.26808 | 1.338 | 2.5478 | 98.92 | 20.89 | 38.555 | 67.3595 | 111.7667 | 64.501 | 0.009639 | 1.0456 | 0.024973 | 23.8839 | 0.016229 | 0.14356 | -0.000175 | -0.012915 | -0.012915 | -0.45684
2019-03-21 15:46 | 1553197564.8582 | -226 | 0.20943 | 0.000464 | -3e-06 | 99.4011 | 99.3053 | 124.3208 | 21.2612 | 0.001907 | 0.001778 | -152.1552 | 0.24017 | 8.4029 | 0.003809 | 0.004827 | 0.004683 | 5.3e-05 | 0.001184 | 0.27223 | 1.2966 | 2.548 | 98.92 | 20.89 | 38.558 | 67.7206 | 113.6087 | 64.6064 | 0.00964 | 1.0453 | 0.024935 | 23.8553 | 0.00862 | 0.14288 | 0.000242 | 0.018318 | 0.018318 | 0.32059
2019-03-21 15:46 | 1553197565.8593 | -225 | 0.20946 | 0.000463 | -3e-06 | 99.3293 | 99.3615 | 123.9785 | 21.2612 | 0.0019 | 0.001777 | -152.1566 | 0.23987 | 8.4019 | 0.003761 | 0.004965 | 0.004789 | 5e-05 | 0.001185 | 0.25658 | 1.2971 | 2.5478 | 98.92 | 20.89 | 38.56 | 67.5521 | 112.4365 | 64.5821 | 0.00964 | 1.0453 | 0.024902 | 23.8216 | -0.002904 | 0.1427 | 4.7e-05 | 0.003695 | 0.003695 | 1.5871
2019-03-21 15:46 | 1553197566.8583 | -224 | 0.20947 | 0.000463 | -3e-06 | 99.3471 | 99.2764 | 123.9424 | 21.2612 | 0.001891 | 0.001786 | -152.1487 | 0.24125 | 8.406 | 0.003863 | 0.00523 | 0.005073 | 5.2e-05 | 0.001189 | 0.27903 | 1.3176 | 2.5479 | 98.92 | 20.88 | 38.561 | 67.3237 | 114.9077 | 64.8124 | 0.009634 | 1.0446 | 0.02489 | 23.8262 | 0.006374 | 0.14352 | -8e-06 | -0.000382 | -0.000382 | -15.4537
2019-03-21 15:46 | 1553197567.8583 | -223 | 0.20947 | 0.000463 | -3e-06 | 99.2947 | 99.1326 | 123.8994 | 21.2612 | 0.00191 | 0.001719 | -152.1129 | 0.24085 | 8.4046 | 0.003778 | 0.005589 | 0.00541 | 4.9e-05 | 0.001187 | 0.2636 | 1.3884 | 2.5479 | 98.92 | 20.88 | 38.561 | 67.7793 | 112.8447 | 63.1985 | 0.009634 | 1.0496 | 0.024945 | 23.7652 | 0.01461 | 0.14328 | -1.2e-05 | -0.000683 | -0.000683 | -8.6266
2019-03-21 15:46 | 1553197568.8588 | -222 | 0.20944 | 0.000465 | -3e-06 | 99.2592 | 99.0332 | 123.5781 | 21.2612 | 0.001899 | 0.001732 | -152.105 | 0.24121 | 8.4023 | 0.003811 | 0.005513 | 0.005345 | 5.1e-05 | 0.001189 | 0.2688 | 1.3353 | 2.548 | 98.92 | 20.88 | 38.563 | 67.5086 | 113.6389 | 63.51 | 0.009635 | 1.0487 | 0.024901 | 23.7453 | 0.020394 | 0.1435 | 0.000185 | 0.013994 | 0.013994 | 0.42145
2019-03-21 15:46 | 1553197569.8585 | -221 | 0.20948 | 0.000455 | -3e-06 | 99.3036 | 99.2791 | 123.6396 | 21.2612 | 0.001887 | 0.001799 | -152.1335 | 0.24016 | 8.3996 | 0.003701 | 0.005191 | 0.004984 | 5.2e-05 | 0.00119 | 0.26576 | 1.3394 | 2.5479 | 98.92 | 20.88 | 38.563 | 67.2316 | 110.9751 | 65.1215 | 0.009635 | 1.0437 | 0.024848 | 23.808 | 0.002211 | 0.14287 | -7.2e-05 | -0.005196 | -0.005196 | -1.1301
2019-03-21 15:46 | 1553197570.8574 | -220 | 0.20951 | 0.000455 | -3e-06 | 99.3796 | 99.1493 | 123.5746 | 21.2612 | 0.001867 | 0.001789 | -152.1401 | 0.24024 | 8.4028 | 0.003704 | 0.004832 | 0.004758 | 5.2e-05 | 0.001188 | 0.27903 | 1.334 | 2.548 | 98.92 | 20.88 | 38.563 | 66.7518 | 111.0447 | 64.885 | 0.009635 | 1.0444 | 0.02485 | 23.7934 | 0.02075 | 0.14292 | -0.000208 | -0.015367 | -0.015367 | -0.38224
2019-03-21 15:46 | 1553197571.8589 | -219 | 0.20945 | 0.000459 | -3e-06 | 99.3319 | 99.163 | 123.5093 | 21.2612 | 0.001889 | 0.001746 | -152.129 | 0.24114 | 8.4032 | 0.003761 | 0.004698 | 0.004493 | 5.2e-05 | 0.00119 | 0.27223 | 1.3243 | 2.5479 | 98.92 | 20.88 | 38.563 | 67.2698 | 112.4207 | 63.8294 | 0.009635 | 1.0477 | 0.024882 | 23.75 | 0.01523 | 0.14346 | 0.000121 | 0.009186 | 0.009186 | 0.64184
2019-03-21 15:46 | 1553197572.8583 | -218 | 0.20953 | 0.00046 | -3e-06 | 99.2795 | 99.3554 | 123.6046 | 21.2612 | 0.001885 | 0.001746 | -152.1379 | 0.24071 | 8.4026 | 0.003699 | 0.005215 | 0.005007 | 5.3e-05 | 0.001192 | 0.28655 | 1.3608 | 2.5478 | 98.92 | 20.88 | 38.564 | 67.1739 | 110.9294 | 63.8481 | 0.009635 | 1.0476 | 0.024891 | 23.7598 | -0.006834 | 0.1432 | -0.000343 | -0.025457 | -0.025457 | -0.2312
2019-03-21 15:46 | 1553197573.8586 | -217 | 0.20951 | 0.000462 | -3e-06 | 99.3251 | 99.2311 | 123.6678 | 21.2612 | 0.001898 | 0.001743 | -152.1473 | 0.24068 | 8.4003 | 0.003719 | 0.005339 | 0.005132 | 5.1e-05 | 0.001189 | 0.28423 | 1.4056 | 2.5481 | 98.92 | 20.88 | 38.564 | 67.4905 | 111.4154 | 63.767 | 0.009635 | 1.0479 | 0.024901 | 23.763 | 0.008471 | 0.14318 | -0.000225 | -0.016612 | -0.016612 | -0.35423
2019-03-21 15:46 | 1553197574.8578 | -216 | 0.20949 | 0.00046 | -3e-06 | 99.3674 | 99.2369 | 123.7626 | 21.2612 | 0.001858 | 0.00172 | -152.1424 | 0.24093 | 8.4051 | 0.003688 | 0.005202 | 0.005027 | 5.5e-05 | 0.001186 | 0.29927 | 1.3076 | 2.5479 | 98.92 | 20.88 | 38.564 | 66.5332 | 110.6658 | 63.2031 | 0.009635 | 1.0496 | 0.024931 | 23.7522 | 0.011758 | 0.14333 | -0.000145 | -0.01063 | -0.01063 | -0.55417
2019-03-21 15:46 | 1553197575.8582 | -215 | 0.20949 | 0.000459 | -3e-06 | 99.2238 | 99.0696 | 124.0177 | 21.2612 | 0.001867 | 0.001767 | -152.143 | 0.24077 | 8.4005 | 0.003689 | 0.00505 | 0.004897 | 5.2e-05 | 0.001183 | 0.27024 | 1.3055 | 2.5479 | 98.92 | 20.88 | 38.562 | 66.7393 | 110.6833 | 64.3501 | 0.009635 | 1.0461 | 0.024914 | 23.8172 | 0.013907 | 0.14324 | -0.000131 | -0.009614 | -0.009614 | -0.61231
2019-03-21 15:46 | 1553197576.8578 | -214 | 0.20945 | 0.000462 | -3e-06 | 99.2917 | 99.0553 | 124.0611 | 21.2612 | 0.001816 | 0.001721 | -152.1449 | 0.24036 | 8.4045 | 0.003653 | 0.005183 | 0.004982 | 5.1e-05 | 0.001185 | 0.27577 | 1.3253 | 2.5481 | 98.92 | 20.88 | 38.56 | 65.5218 | 109.8013 | 63.2375 | 0.009634 | 1.0495 | 0.02496 | 23.782 | 0.02132 | 0.14299 | 9.7e-05 | 0.007496 | 0.007496 | 0.78403
2019-03-21 15:46 | 1553197577.8579 | -213 | 0.20948 | 0.000463 | -3e-06 | 99.2408 | 99.0509 | 124.1287 | 21.2612 | 0.001875 | 0.001728 | -152.1614 | 0.24098 | 8.4006 | 0.003696 | 0.004626 | 0.004438 | 5.5e-05 | 0.001181 | 0.26609 | 1.3368 | 2.548 | 98.92 | 20.88 | 38.56 | 66.9468 | 110.8535 | 63.4139 | 0.009634 | 1.049 | 0.02496 | 23.7948 | 0.017129 | 0.14336 | -4.8e-05 | -0.003406 | -0.003406 | -1.7301
2019-03-21 15:46 | 1553197578.8589 | -212 | 0.20951 | 0.000461 | -3e-06 | 99.2944 | 99.2219 | 124.1816 | 21.2612 | 0.001847 | 0.001779 | -152.1727 | 0.24114 | 8.402 | 0.003697 | 0.005105 | 0.004977 | 5.1e-05 | 0.00118 | 0.27184 | 1.3333 | 2.5479 | 98.92 | 20.88 | 38.56 | 66.2812 | 110.8781 | 64.6218 | 0.009634 | 1.0452 | 0.024921 | 23.8425 | 0.006535 | 0.14345 | -0.000212 | -0.015729 | -0.015729 | -0.37484
2019-03-21 15:46 | 1553197579.8583 | -211 | 0.20947 | 0.00046 | -3e-06 | 99.2873 | 99.0843 | 124.1276 | 21.2612 | 0.001892 | 0.001746 | -152.1598 | 0.24097 | 8.3995 | 0.003609 | 0.004994 | 0.004791 | 4.8e-05 | 0.001179 | 0.26056 | 1.3666 | 2.548 | 98.92 | 20.88 | 38.56 | 67.3619 | 108.7346 | 63.8426 | 0.009634 | 1.0476 | 0.024944 | 23.8098 | 0.018304 | 0.14336 | -9e-06 | -0.000514 | -0.000514 | -11.4578
2019-03-21 15:46 | 1553197580.8591 | -210 | 0.20949 | 0.000462 | -3e-06 | 99.3239 | 99.166 | 124.0492 | 21.2612 | 0.001916 | 0.001725 | -152.1503 | 0.24136 | 8.3994 | 0.003608 | 0.005205 | 0.005044 | 4.9e-05 | 0.001179 | 0.27815 | 1.34 | 2.548 | 98.92 | 20.88 | 38.56 | 67.925 | 108.7197 | 63.3311 | 0.009634 | 1.0492 | 0.024955 | 23.7842 | 0.014229 | 0.14358 | -0.000107 | -0.007803 | -0.007803 | -0.75633
2019-03-21 15:46 | 1553197581.8581 | -209 | 0.20946 | 0.000458 | -3e-06 | 99.256 | 99.26 | 123.6879 | 21.2612 | 0.001846 | 0.001756 | -152.1563 | 0.24023 | 8.4019 | 0.003571 | 0.004762 | 0.004523 | 5.2e-05 | 0.001183 | 0.27759 | 1.3179 | 2.548 | 98.92 | 20.88 | 38.559 | 66.2492 | 107.8246 | 64.0917 | 0.009634 | 1.0469 | 0.024891 | 23.7764 | -0.000366 | 0.14291 | 8.2e-05 | 0.006283 | 0.006283 | 0.93487
2019-03-21 15:46 | 1553197582.9279 | -208 | 0.20948 | 0.000462 | -3e-06 | 99.2896 | 99.1115 | 123.4204 | 21.2612 | 0.001864 | 0.001812 | -152.1542 | 0.2409 | 8.4036 | 0.00363 | 0.004871 | 0.004773 | 5.2e-05 | 0.001181 | 0.27256 | 1.3116 | 2.5481 | 98.92 | 20.88 | 38.559 | 66.6806 | 109.2487 | 65.422 | 0.009634 | 1.0427 | 0.024815 | 23.7974 | 0.016059 | 0.14331 | -5.8e-05 | -0.00412 | -0.00412 | -1.4297
2019-03-21 15:46 | 1553197583.8668 | -207 | 0.20945 | 0.00046 | -3e-06 | 99.2265 | 99.3663 | 123.173 | 21.2612 | 0.001869 | 0.001718 | -152.1466 | 0.24155 | 8.402 | 0.003675 | 0.004776 | 0.004508 | 5.3e-05 | 0.001179 | 0.23651 | 1.2922 | 2.548 | 98.92 | 20.88 | 38.558 | 66.7884 | 110.34 | 63.1547 | 0.009634 | 1.0498 | 0.024873 | 23.6938 | -0.012595 | 0.1437 | 0.000118 | 0.009003 | 0.009003 | 0.65604
2019-03-21 15:46 | 1553197584.8674 | -206 | 0.20947 | 0.00046 | -3e-06 | 99.2584 | 99.1459 | 122.8986 | 21.2612 | 0.001893 | 0.001767 | -152.147 | 0.24148 | 8.4041 | 0.003691 | 0.004558 | 0.00445 | 5.3e-05 | 0.001184 | 0.23612 | 1.3133 | 2.5479 | 98.92 | 20.88 | 38.557 | 67.3703 | 110.74 | 64.3386 | 0.009633 | 1.0461 | 0.024802 | 23.7091 | 0.010137 | 0.14366 | 3.4e-05 | 0.002752 | 0.002752 | 2.1456
2019-03-21 15:46 | 1553197585.8664 | -205 | 0.20947 | 0.000454 | -3e-06 | 99.2607 | 99.1766 | 122.75 | 21.2612 | 0.001839 | 0.001717 | -152.1398 | 0.24103 | 8.4006 | 0.003584 | 0.005208 | 0.004979 | 5.1e-05 | 0.001185 | 0.27201 | 1.2822 | 2.5479 | 98.92 | 20.88 | 38.555 | 66.0712 | 108.1233 | 63.1391 | 0.009633 | 1.0498 | 0.024831 | 23.6526 | 0.007587 | 0.14339 | -1.7e-05 | -0.001085 | -0.001085 | -5.4336
2019-03-21 15:46 | 1553197586.8678 | -204 | 0.20947 | 0.000458 | -2e-06 | 99.2938 | 99.1957 | 122.6934 | 21.2612 | 0.001904 | 0.001777 | -152.1416 | 0.24096 | 8.4029 | 0.003608 | 0.005192 | 0.005045 | 5.2e-05 | 0.001183 | 0.23595 | 1.2896 | 2.5478 | 98.92 | 20.88 | 38.554 | 67.6396 | 108.7229 | 64.5744 | 0.009633 | 1.0454 | 0.024773 | 23.6975 | 0.008846 | 0.14335 | 1.4e-05 | 0.001224 | 0.001224 | 4.8147
2019-03-21 15:46 | 1553197587.8665 | -203 | 0.20947 | 0.000461 | -3e-06 | 99.2962 | 99.2301 | 122.7155 | 21.2612 | 0.001819 | 0.001782 | -152.1551 | 0.24032 | 8.4037 | 0.003608 | 0.005515 | 0.005389 | 5e-05 | 0.001186 | 0.2401 | 1.3257 | 2.5479 | 98.92 | 20.88 | 38.554 | 65.5938 | 108.7222 | 64.7 | 0.009633 | 1.045 | 0.02477 | 23.7041 | 0.00596 | 0.14297 | -2.5e-05 | -0.001659 | -0.001659 | -3.5424
2019-03-21 15:46 | 1553197588.8667 | -202 | 0.20944 | 0.000465 | -2e-06 | 99.2369 | 99.1994 | 122.9004 | 21.2612 | 0.001895 | 0.001766 | -152.1461 | 0.24125 | 8.4027 | 0.003578 | 0.00526 | 0.005065 | 5e-05 | 0.001185 | 0.27201 | 1.2646 | 2.548 | 98.92 | 20.88 | 38.553 | 67.4196 | 107.99 | 64.3205 | 0.009632 | 1.0462 | 0.024803 | 23.7086 | 0.003381 | 0.14352 | 0.000167 | 0.012567 | 0.012567 | 0.46937
2019-03-21 15:46 | 1553197589.8677 | -201 | 0.20947 | 0.000462 | -3e-06 | 99.2524 | 99.1793 | 123.1044 | 21.2612 | 0.001888 | 0.001752 | -152.1569 | 0.24105 | 8.4045 | 0.003558 | 0.004819 | 0.004689 | 5e-05 | 0.001184 | 0.25409 | 1.3406 | 2.5481 | 98.92 | 20.88 | 38.552 | 67.265 | 107.5049 | 63.9848 | 0.009632 | 1.0472 | 0.024836 | 23.7165 | 0.006589 | 0.1434 | -1e-06 | 0.000159 | 0.000159 | 37.1266
2019-03-21 15:47 | 1553197590.8667 | -200 | 0.20947 | 0.00046 | -3e-06 | 99.2658 | 99.2655 | 123.2877 | 21.2612 | 0.001887 | 0.001703 | -152.1646 | 0.24122 | 8.4033 | 0.003538 | 0.004719 | 0.00459 | 4.8e-05 | 0.001187 | 0.2453 | 1.3288 | 2.5479 | 98.92 | 20.88 | 38.552 | 67.2229 | 107.0245 | 62.8032 | 0.009632 | 1.0509 | 0.024898 | 23.6925 | 2.9e-05 | 0.1435 | 6e-06 | 0.000631 | 0.000631 | 9.3481
2019-03-21 15:47 | 1553197591.8672 | -199 | 0.20946 | 0.000457 | -3e-06 | 99.3647 | 99.2934 | 123.5462 | 21.2612 | 0.001852 | 0.001769 | -152.1849 | 0.24184 | 8.4028 | 0.003565 | 0.004881 | 0.004715 | 5e-05 | 0.001188 | 0.28064 | 1.3032 | 2.548 | 98.92 | 20.87 | 38.551 | 66.3889 | 107.6709 | 64.3821 | 0.009626 | 1.046 | 0.024866 | 23.773 | 0.006422 | 0.14387 | 7.4e-05 | 0.005705 | 0.005705 | 1.0364
2019-03-21 15:47 | 1553197592.8677 | -198 | 0.20937 | 0.000458 | -3e-06 | 99.2604 | 99.2542 | 123.7879 | 21.2612 | 0.001825 | 0.001797 | -152.1797 | 0.24103 | 8.4019 | 0.003508 | 0.005183 | 0.005023 | 5.5e-05 | 0.001191 | 0.26305 | 1.3488 | 2.548 | 98.92 | 20.87 | 38.55 | 65.735 | 106.2821 | 65.077 | 0.009626 | 1.0438 | 0.024864 | 23.8207 | 0.000559 | 0.14339 | 0.000606 | 0.04544 | 0.04544 | 0.12969
2019-03-21 15:47 | 1553197593.8683 | -197 | 0.20952 | 0.000466 | -3e-06 | 99.3602 | 99.1742 | 123.8105 | 21.2612 | 0.001871 | 0.001778 | -152.1717 | 0.2408 | 8.4 | 0.003451 | 0.004469 | 0.004269 | 5.1e-05 | 0.001187 | 0.23922 | 1.3879 | 2.548 | 98.92 | 20.88 | 38.549 | 66.8447 | 104.902 | 64.6023 | 0.009631 | 1.0453 | 0.024884 | 23.8061 | 0.016764 | 0.14325 | -0.000314 | -0.023289 | -0.023289 | -0.25281
2019-03-21 15:47 | 1553197594.8669 | -196 | 0.20953 | 0.000459 | -3e-06 | 99.3632 | 99.3343 | 123.9529 | 21.2612 | 0.001873 | 0.001788 | -152.1918 | 0.24068 | 8.4027 | 0.003489 | 0.004745 | 0.004543 | 5e-05 | 0.001193 | 0.26913 | 1.319 | 2.5479 | 98.92 | 20.87 | 38.549 | 66.8867 | 105.8336 | 64.8475 | 0.009626 | 1.0445 | 0.024889 | 23.8285 | 0.002607 | 0.14318 | -0.00038 | -0.028215 | -0.028215 | -0.20856
2019-03-21 15:47 | 1553197595.8675 | -195 | 0.20946 | 0.00046 | -2e-06 | 99.3457 | 99.2849 | 124.083 | 21.2612 | 0.00194 | 0.001801 | -152.1914 | 0.24025 | 8.4042 | 0.003527 | 0.004572 | 0.00446 | 5.2e-05 | 0.001187 | 0.25874 | 1.3277 | 2.5479 | 98.92 | 20.87 | 38.55 | 68.5171 | 106.7443 | 65.1605 | 0.009626 | 1.0436 | 0.024891 | 23.852 | 0.005472 | 0.14292 | 8.6e-05 | 0.00659 | 0.00659 | 0.89138
2019-03-21 15:47 | 1553197596.868 | -194 | 0.20946 | 0.00046 | -3e-06 | 99.3668 | 99.3322 | 124.1719 | 21.2612 | 0.001872 | 0.001797 | -152.1842 | 0.24158 | 8.3996 | 0.00359 | 0.004841 | 0.004655 | 5.3e-05 | 0.001189 | 0.262 | 1.3354 | 2.5479 | 98.92 | 20.87 | 38.551 | 66.8819 | 108.2665 | 65.0708 | 0.009626 | 1.0438 | 0.024903 | 23.8574 | 0.003113 | 0.14372 | 9.2e-05 | 0.007074 | 0.007074 | 0.83494
2019-03-21 15:47 | 1553197597.8671 | -193 | 0.20944 | 0.000466 | -3e-06 | 99.3781 | 99.352 | 124.0027 | 21.2612 | 0.001962 | 0.001741 | -152.1822 | 0.24179 | 8.4017 | 0.003479 | 0.00455 | 0.004467 | 5.1e-05 | 0.001191 | 0.26128 | 1.3009 | 2.548 | 98.92 | 20.87 | 38.552 | 69.0234 | 105.5714 | 63.7155 | 0.009626 | 1.048 | 0.024936 | 23.7933 | 0.002353 | 0.14384 | 0.000191 | 0.014489 | 0.014489 | 0.40801
2019-03-21 15:47 | 1553197598.8665 | -192 | 0.20945 | 0.000463 | -4e-06 | 99.3713 | 99.3292 | 123.9313 | 21.2612 | 0.001901 | 0.001735 | -152.1747 | 0.24003 | 8.4028 | 0.003522 | 0.004433 | 0.004181 | 5.3e-05 | 0.001188 | 0.28296 | 1.3119 | 2.548 | 98.92 | 20.87 | 38.552 | 67.5756 | 106.6278 | 63.5851 | 0.009626 | 1.0484 | 0.024934 | 23.7819 | 0.003791 | 0.1428 | 0.000116 | 0.008922 | 0.008922 | 0.65776
2019-03-21 15:47 | 1553197599.867 | -191 | 0.20944 | 0.000462 | -3e-06 | 99.3308 | 99.3298 | 123.8208 | 21.2612 | 0.001852 | 0.001756 | -152.17 | 0.24058 | 8.4039 | 0.003546 | 0.004743 | 0.004539 | 5e-05 | 0.001188 | 0.28351 | 1.3192 | 2.548 | 98.92 | 20.87 | 38.553 | 66.3841 | 107.2143 | 64.084 | 0.009627 | 1.0469 | 0.024904 | 23.7889 | 8.2e-05 | 0.14312 | 0.000186 | 0.014095 | 0.014095 | 0.41734
2019-03-21 15:47 | 1553197600.8671 | -190 | 0.20946 | 0.000463 | -4e-06 | 99.5342 | 99.6497 | 123.7565 | 21.2612 | 0.001835 | 0.001748 | -152.1593 | 0.23995 | 8.401 | 0.003592 | 0.005145 | 0.004989 | 4.8e-05 | 0.001184 | 0.25409 | 1.3205 | 2.5479 | 98.92 | 20.87 | 38.554 | 65.9769 | 108.3336 | 63.8763 | 0.009627 | 1.0475 | 0.024905 | 23.7754 | -0.010365 | 0.14275 | 9.3e-05 | 0.007161 | 0.007161 | 0.81931
2019-03-21 15:47 | 1553197601.8666 | -189 | 0.20945 | 0.000462 | -3e-06 | 99.3871 | 99.4276 | 123.8182 | 21.2612 | 0.001843 | 0.001749 | -152.1355 | 0.24108 | 8.4016 | 0.003598 | 0.004695 | 0.004599 | 4.9e-05 | 0.001188 | 0.26056 | 1.3254 | 2.548 | 98.92 | 20.87 | 38.556 | 66.1775 | 108.4771 | 63.9013 | 0.009627 | 1.0475 | 0.024911 | 23.7822 | -0.003647 | 0.14342 | 0.000147 | 0.011211 | 0.011211 | 0.52579
2019-03-21 15:47 | 1553197602.8663 | -188 | 0.20942 | 0.000463 | -3e-06 | 99.4043 | 99.3554 | 123.9417 | 21.2612 | 0.001913 | 0.00178 | -152.1305 | 0.24188 | 8.4037 | 0.003529 | 0.004739 | 0.004629 | 5.1e-05 | 0.001189 | 0.26808 | 1.3538 | 2.5479 | 98.92 | 20.87 | 38.557 | 67.8509 | 106.8008 | 64.6649 | 0.009628 | 1.0451 | 0.024895 | 23.821 | 0.004406 | 0.1439 | 0.00032 | 0.024113 | 0.024113 | 0.24527
2019-03-21 15:47 | 1553197603.8673 | -187 | 0.20946 | 0.000462 | -3e-06 | 99.4714 | 99.4831 | 123.9391 | 21.2612 | 0.001887 | 0.001763 | -152.1347 | 0.24138 | 8.4017 | 0.003534 | 0.004975 | 0.004805 | 5e-05 | 0.001184 | 0.25984 | 1.3414 | 2.548 | 98.92 | 20.87 | 38.557 | 67.2323 | 106.9127 | 64.2416 | 0.009628 | 1.0464 | 0.02491 | 23.8058 | -0.001055 | 0.1436 | 5.1e-05 | 0.004013 | 0.004013 | 1.4708
2019-03-21 15:47 | 1553197604.8668 | -186 | 0.20945 | 0.000465 | -3e-06 | 99.4279 | 99.5492 | 124.0954 | 21.2612 | 0.001926 | 0.001757 | -152.1199 | 0.24171 | 8.4025 | 0.003539 | 0.004667 | 0.004501 | 4.7e-05 | 0.001182 | 0.25514 | 1.3603 | 2.5479 | 98.92 | 20.87 | 38.559 | 68.1644 | 107.0425 | 64.1035 | 0.009628 | 1.0468 | 0.024931 | 23.8159 | -0.010906 | 0.14379 | 0.000147 | 0.011168 | 0.011168 | 0.5292
2019-03-21 15:47 | 1553197605.8674 | -185 | 0.20941 | 0.000458 | -4e-06 | 99.4919 | 99.5053 | 124.1621 | 21.2612 | 0.001914 | 0.001824 | -152.1432 | 0.24085 | 8.4007 | 0.00356 | 0.00432 | 0.004116 | 4.9e-05 | 0.001188 | 0.26288 | 1.3322 | 2.548 | 98.92 | 20.87 | 38.559 | 67.8846 | 107.5418 | 65.7169 | 0.009628 | 1.0418 | 0.024878 | 23.8792 | -0.001196 | 0.14328 | 0.000386 | 0.029041 | 0.029041 | 0.20278
2019-03-21 15:47 | 1553197606.8665 | -184 | 0.20949 | 0.000465 | -3e-06 | 99.4982 | 99.5747 | 124.2168 | 21.2612 | 0.001889 | 0.001742 | -152.1313 | 0.24027 | 8.4023 | 0.003479 | 0.005013 | 0.004735 | 4.8e-05 | 0.001186 | 0.2401 | 1.337 | 2.5481 | 98.92 | 20.87 | 38.559 | 67.29 | 105.5714 | 63.7451 | 0.009628 | 1.0479 | 0.024957 | 23.8149 | -0.006877 | 0.14294 | -0.000101 | -0.007364 | -0.007364 | -0.79778
2019-03-21 15:47 | 1553197607.8665 | -183 | 0.20944 | 0.000457 | -4e-06 | 99.4991 | 99.6057 | 124.647 | 21.2612 | 0.001854 | 0.001811 | -152.1255 | 0.24155 | 8.4025 | 0.003566 | 0.005211 | 0.005145 | 4.7e-05 | 0.001183 | 0.24386 | 1.3591 | 2.548 | 98.92 | 20.87 | 38.558 | 66.4272 | 107.6845 | 65.404 | 0.009628 | 1.0428 | 0.024938 | 23.9148 | -0.009581 | 0.1437 | 0.000171 | 0.01307 | 0.01307 | 0.45186
2019-03-21 15:47 | 1553197608.868 | -182 | 0.20943 | 0.000464 | -4e-06 | 99.5158 | 99.494 | 124.9914 | 21.2612 | 0.001869 | 0.00178 | -152.1489 | 0.24044 | 8.4026 | 0.003519 | 0.004322 | 0.004086 | 5e-05 | 0.001187 | 0.26106 | 1.3092 | 2.5479 | 98.92 | 20.87 | 38.557 | 66.8079 | 106.5465 | 64.6601 | 0.009628 | 1.0451 | 0.025 | 23.9215 | 0.001957 | 0.14304 | 0.000217 | 0.016482 | 0.016482 | 0.35668
2019-03-21 15:47 | 1553197609.8667 | -181 | 0.20944 | 0.000457 | -3e-06 | 99.573 | 99.4978 | 125.0517 | 21.2612 | 0.001803 | 0.001765 | -152.1589 | 0.24149 | 8.4036 | 0.003577 | 0.004929 | 0.004719 | 4.9e-05 | 0.001189 | 0.24065 | 1.3697 | 2.548 | 98.92 | 20.87 | 38.556 | 65.2105 | 107.9703 | 64.2986 | 0.009627 | 1.0462 | 0.02502 | 23.9144 | 0.00676 | 0.14366 | 0.000186 | 0.01419 | 0.01419 | 0.41611
2019-03-21 15:47 | 1553197610.8677 | -180 | 0.20942 | 0.000459 | -3e-06 | 99.578 | 99.6084 | 125.1512 | 21.2612 | 0.001882 | 0.001781 | -152.1755 | 0.24105 | 8.4018 | 0.003577 | 0.004729 | 0.00457 | 5e-05 | 0.001188 | 0.26089 | 1.3194 | 2.5478 | 98.91 | 20.87 | 38.556 | 67.1169 | 107.9499 | 64.6767 | 0.009628 | 1.045 | 0.025016 | 23.9373 | -0.00273 | 0.1434 | 0.000328 | 0.02483 | 0.02483 | 0.23736
2019-03-21 15:47 | 1553197611.8672 | -179 | 0.20948 | 0.000463 | -4e-06 | 99.6233 | 99.3091 | 125.1169 | 21.2612 | 0.001919 | 0.001801 | -152.1671 | 0.24087 | 8.4025 | 0.003539 | 0.004822 | 0.004644 | 4.9e-05 | 0.001188 | 0.26825 | 1.3572 | 2.5479 | 98.91 | 20.87 | 38.556 | 67.9937 | 107.0347 | 65.1518 | 0.009628 | 1.0436 | 0.024995 | 23.9509 | 0.02826 | 0.14329 | -6e-05 | -0.004299 | -0.004299 | -1.3701
2019-03-21 15:47 | 1553197612.8678 | -178 | 0.20945 | 0.000465 | -3e-06 | 99.7017 | 99.6483 | 125.1169 | 21.2612 | 0.00188 | 0.00176 | -152.1761 | 0.24127 | 8.4009 | 0.003499 | 0.004696 | 0.004552 | 4.9e-05 | 0.001187 | 0.28528 | 1.3168 | 2.548 | 98.91 | 20.87 | 38.557 | 67.063 | 106.0744 | 64.1769 | 0.009628 | 1.0466 | 0.025031 | 23.9163 | 0.004792 | 0.14353 | 0.000138 | 0.010585 | 0.010585 | 0.55733
2019-03-21 15:47 | 1553197613.8679 | -177 | 0.20944 | 0.000462 | -3e-06 | 99.7282 | 99.5979 | 125.0313 | 21.2612 | 0.001962 | 0.001763 | -152.1647 | 0.24203 | 8.4022 | 0.003548 | 0.005037 | 0.004862 | 5e-05 | 0.001187 | 0.27649 | 1.3329 | 2.5478 | 98.92 | 20.87 | 38.556 | 69.0241 | 107.2575 | 64.2597 | 0.009627 | 1.0463 | 0.025019 | 23.9111 | 0.011696 | 0.14398 | 0.000197 | 0.015003 | 0.015003 | 0.39443
2019-03-21 15:47 | 1553197614.8673 | -176 | 0.20945 | 0.000458 | -3e-06 | 99.6987 | 99.6554 | 124.9151 | 21.2612 | 0.00191 | 0.001711 | -152.1565 | 0.24172 | 8.4035 | 0.003605 | 0.004702 | 0.004561 | 4.9e-05 | 0.001187 | 0.29231 | 1.3363 | 2.548 | 98.92 | 20.87 | 38.556 | 67.7925 | 108.6323 | 62.9929 | 0.009627 | 1.0503 | 0.025055 | 23.8551 | 0.003882 | 0.1438 | 0.000106 | 0.008146 | 0.008146 | 0.72547
2019-03-21 15:47 | 1553197615.866 | -175 | 0.20945 | 0.000465 | -3e-06 | 99.6969 | 99.6401 | 124.5447 | 21.2612 | 0.001821 | 0.001761 | -152.1557 | 0.2412 | 8.3998 | 0.003566 | 0.004673 | 0.004498 | 5e-05 | 0.001186 | 0.28727 | 1.2992 | 2.548 | 98.92 | 20.87 | 38.556 | 65.6404 | 107.6972 | 64.2065 | 0.009627 | 1.0465 | 0.024972 | 23.8626 | 0.005098 | 0.14349 | 9.7e-05 | 0.007475 | 0.007475 | 0.78896
2019-03-21 15:47 | 1553197616.8661 | -174 | 0.20944 | 0.000466 | -3e-06 | 99.73 | 99.6902 | 124.5013 | 21.2612 | 0.001833 | 0.00175 | -152.1736 | 0.24133 | 8.403 | 0.003564 | 0.004357 | 0.004223 | 4.7e-05 | 0.001183 | 0.28031 | 1.3089 | 2.5479 | 98.92 | 20.87 | 38.558 | 65.9356 | 107.6481 | 63.9466 | 0.009628 | 1.0473 | 0.024978 | 23.8493 | 0.003571 | 0.14357 | 0.000164 | 0.012513 | 0.012513 | 0.47157
2019-03-21 15:47 | 1553197617.8663 | -173 | 0.20945 | 0.000458 | -3e-06 | 99.7196 | 99.5689 | 124.2337 | 21.2612 | 0.001846 | 0.001711 | -152.1806 | 0.24145 | 8.4045 | 0.003618 | 0.004706 | 0.004467 | 5.5e-05 | 0.001186 | 0.26416 | 1.3193 | 2.548 | 98.92 | 20.87 | 38.559 | 66.254 | 108.9586 | 63.0002 | 0.009628 | 1.0503 | 0.024986 | 23.7902 | 0.013521 | 0.14364 | 0.000124 | 0.009456 | 0.009456 | 0.62432
2019-03-21 15:47 | 1553197618.8672 | -172 | 0.20943 | 0.000466 | -3e-06 | 99.786 | 99.8346 | 124.156 | 21.2612 | 0.001881 | 0.001786 | -152.1569 | 0.24086 | 8.3989 | 0.003527 | 0.004503 | 0.004331 | 5.2e-05 | 0.001185 | 0.27184 | 1.3093 | 2.5478 | 98.92 | 20.87 | 38.562 | 67.0887 | 106.7394 | 64.7974 | 0.009629 | 1.0447 | 0.024912 | 23.8462 | -0.004354 | 0.14329 | 0.000243 | 0.018373 | 0.018373 | 0.32053
2019-03-21 15:47 | 1553197619.8676 | -171 | 0.20947 | 0.000461 | -3e-06 | 99.7923 | 99.5158 | 123.9722 | 21.2612 | 0.001816 | 0.001766 | -152.1484 | 0.24244 | 8.4016 | 0.003613 | 0.004827 | 0.004653 | 5.3e-05 | 0.001183 | 0.25658 | 1.3666 | 2.548 | 98.92 | 20.87 | 38.564 | 65.5322 | 108.8427 | 64.3104 | 0.009629 | 1.0462 | 0.024911 | 23.8114 | 0.024814 | 0.14423 | 3e-05 | 0.002432 | 0.002432 | 2.4371
2019-03-21 15:47 | 1553197620.8674 | -170 | 0.20944 | 0.000456 | -3e-06 | 99.8435 | 99.8274 | 123.791 | 21.2612 | 0.00189 | 0.001784 | -152.1251 | 0.2408 | 8.4011 | 0.003572 | 0.005493 | 0.005267 | 5.1e-05 | 0.001186 | 0.27815 | 1.3821 | 2.548 | 98.91 | 20.87 | 38.569 | 67.3104 | 107.8467 | 64.7421 | 0.009631 | 1.0448 | 0.024877 | 23.8092 | 0.00144 | 0.14325 | 0.000215 | 0.016298 | 0.016298 | 0.36124
2019-03-21 15:47 | 1553197621.8673 | -169 | 0.20945 | 0.000458 | -3e-06 | 99.8381 | 99.9678 | 123.5586 | 21.2612 | 0.00183 | 0.001764 | -152.1398 | 0.2422 | 8.4034 | 0.003552 | 0.004425 | 0.004279 | 5.1e-05 | 0.001188 | 0.26985 | 1.3416 | 2.548 | 98.91 | 20.87 | 38.572 | 65.8629 | 107.3512 | 64.2823 | 0.009632 | 1.0463 | 0.024871 | 23.7707 | -0.011605 | 0.14409 | 0.000123 | 0.009414 | 0.009414 | 0.62905
2019-03-21 15:47 | 1553197622.8665 | -168 | 0.20946 | 0.00046 | -4e-06 | 99.9397 | 99.9596 | 123.3082 | 21.2612 | 0.001856 | 0.001739 | -152.1513 | 0.24183 | 8.4014 | 0.003575 | 0.004481 | 0.004358 | 5e-05 | 0.001187 | 0.23396 | 1.3832 | 2.5479 | 98.91 | 20.87 | 38.575 | 66.4887 | 107.9176 | 63.6693 | 0.009633 | 1.0482 | 0.024868 | 23.725 | -0.001776 | 0.14387 | 9.7e-05 | 0.00743 | 0.00743 | 0.79581
2019-03-21 15:47 | 1553197623.8679 | -167 | 0.2095 | 0.000458 | -3e-06 | 99.8361 | 100.0877 | 123.4445 | 21.2612 | 0.001882 | 0.001787 | -152.1497 | 0.24153 | 8.4024 | 0.003553 | 0.005167 | 0.004989 | 5.3e-05 | 0.001186 | 0.25874 | 1.331 | 2.5481 | 98.91 | 20.87 | 38.581 | 67.1217 | 107.3821 | 64.8186 | 0.009634 | 1.0446 | 0.024839 | 23.7786 | -0.022512 | 0.14369 | -0.000179 | -0.013164 | -0.013164 | -0.44862
2019-03-21 15:47 | 1553197624.8674 | -166 | 0.20945 | 0.000458 | -3e-06 | 99.9308 | 99.8584 | 123.4499 | 21.2612 | 0.001851 | 0.001723 | -152.1586 | 0.24099 | 8.4011 | 0.003425 | 0.00439 | 0.004122 | 4.9e-05 | 0.001187 | 0.24347 | 1.3447 | 2.5479 | 98.91 | 20.87 | 38.584 | 66.3678 | 104.2808 | 63.2953 | 0.009635 | 1.0493 | 0.024896 | 23.7254 | 0.00648 | 0.14336 | 0.000122 | 0.0093 | 0.0093 | 0.63358
2019-03-21 15:47 | 1553197625.8668 | -165 | 0.20949 | 0.00046 | -3e-06 | 99.9252 | 99.932 | 123.8353 | 21.2612 | 0.001943 | 0.001748 | -152.1314 | 0.24263 | 8.4031 | 0.003485 | 0.005155 | 0.005046 | 5e-05 | 0.001185 | 0.25984 | 1.3549 | 2.5479 | 98.91 | 20.87 | 38.587 | 68.5803 | 105.7186 | 63.885 | 0.009636 | 1.0475 | 0.024913 | 23.7832 | -0.000614 | 0.14434 | -8.6e-05 | -0.006262 | -0.006262 | -0.94734
2019-03-21 15:47 | 1553197626.8665 | -164 | 0.20944 | 0.00046 | -3e-06 | 100.0053 | 100.0699 | 123.949 | 21.2612 | 0.00184 | 0.001789 | -152.1178 | 0.24181 | 8.4004 | 0.003486 | 0.004397 | 0.004181 | 5.1e-05 | 0.001185 | 0.2688 | 1.3574 | 2.5479 | 98.91 | 20.87 | 38.593 | 66.0948 | 105.7565 | 64.8677 | 0.009637 | 1.0445 | 0.024888 | 23.8288 | -0.00578 | 0.14385 | 0.000171 | 0.012948 | 0.012948 | 0.45662
2019-03-21 15:47 | 1553197627.8676 | -163 | 0.20946 | 0.000454 | -3e-06 | 99.9707 | 100.0083 | 123.7952 | 21.2612 | 0.001863 | 0.001779 | -152.1161 | 0.24195 | 8.4025 | 0.003456 | 0.004785 | 0.004626 | 5.1e-05 | 0.001187 | 0.25354 | 1.327 | 2.548 | 98.91 | 20.87 | 38.596 | 66.6511 | 105.0225 | 64.6228 | 0.009638 | 1.0452 | 0.024882 | 23.8054 | -0.00336 | 0.14394 | 0.000104 | 0.007931 | 0.007931 | 0.74592
2019-03-21 15:47 | 1553197628.8666 | -162 | 0.20945 | 0.000457 | -3e-06 | 100.0253 | 100.026 | 123.5987 | 21.2612 | 0.00192 | 0.001794 | -152.128 | 0.24129 | 8.3996 | 0.003475 | 0.00507 | 0.004862 | 4.8e-05 | 0.001188 | 0.24994 | 1.3638 | 2.548 | 98.91 | 20.87 | 38.605 | 68.0163 | 105.4912 | 64.9904 | 0.00964 | 1.0441 | 0.024849 | 23.7994 | -6.7e-05 | 0.14354 | 0.000133 | 0.01015 | 0.01015 | 0.58123
2019-03-21 15:47 | 1553197629.8677 | -161 | 0.20945 | 0.000459 | -3e-06 | 100.0527 | 100.0802 | 123.7398 | 21.2612 | 0.001879 | 0.001785 | -152.1258 | 0.24052 | 8.4013 | 0.003493 | 0.004916 | 0.004674 | 4.9e-05 | 0.001184 | 0.27975 | 1.3552 | 2.5481 | 98.91 | 20.87 | 38.61 | 67.0483 | 105.9233 | 64.7828 | 0.009642 | 1.0447 | 0.02487 | 23.8057 | -0.002457 | 0.14308 | 0.00015 | 0.011354 | 0.011354 | 0.51793
2019-03-21 15:47 | 1553197630.8668 | -160 | 0.20941 | 0.000457 | -3e-06 | 100.148 | 100.153 | 123.3509 | 21.2612 | 0.001886 | 0.00176 | -152.1229 | 0.2421 | 8.4 | 0.003404 | 0.004616 | 0.004462 | 5.5e-05 | 0.001192 | 0.27616 | 1.2858 | 2.548 | 98.91 | 20.87 | 38.614 | 67.202 | 103.7644 | 64.1738 | 0.009643 | 1.0466 | 0.024854 | 23.7468 | -0.000449 | 0.14402 | 0.000381 | 0.028668 | 0.028668 | 0.20648
2019-03-21 15:47 | 1553197631.8668 | -159 | 0.20942 | 0.00046 | -3e-06 | 100.1251 | 100.1585 | 123.3394 | 21.2612 | 0.001888 | 0.001759 | -152.1233 | 0.24175 | 8.4016 | 0.003447 | 0.004513 | 0.004288 | 5.2e-05 | 0.001191 | 0.27704 | 1.311 | 2.5479 | 98.91 | 20.86 | 38.622 | 67.2465 | 104.7982 | 64.1466 | 0.009639 | 1.0467 | 0.024853 | 23.7448 | -0.002986 | 0.14382 | 0.000295 | 0.022241 | 0.022241 | 0.26577
2019-03-21 15:47 | 1553197632.8669 | -158 | 0.20946 | 0.000459 | -3e-06 | 100.1117 | 100.0546 | 123.223 | 21.2612 | 0.001886 | 0.001788 | -152.104 | 0.24188 | 8.4063 | 0.003498 | 0.004501 | 0.00441 | 5e-05 | 0.001186 | 0.25658 | 1.2782 | 2.548 | 98.91 | 20.86 | 38.626 | 67.201 | 106.0492 | 64.8506 | 0.00964 | 1.0445 | 0.024816 | 23.7583 | 0.005098 | 0.1439 | 5.2e-05 | 0.00412 | 0.00412 | 1.4354
2019-03-21 15:47 | 1553197633.8673 | -157 | 0.20947 | 0.000459 | -3e-06 | 100.1927 | 100.3169 | 123.2025 | 21.2612 | 0.001913 | 0.001756 | -152.1286 | 0.24094 | 8.4025 | 0.003452 | 0.004625 | 0.004405 | 5.1e-05 | 0.001189 | 0.26305 | 1.3078 | 2.5479 | 98.91 | 20.86 | 38.633 | 67.8517 | 104.9352 | 64.0816 | 0.009642 | 1.0469 | 0.024842 | 23.7293 | -0.011078 | 0.14334 | 2.1e-05 | 0.001739 | 0.001739 | 3.3868
2019-03-21 15:47 | 1553197634.867 | -156 | 0.20937 | 0.000461 | -3e-06 | 100.2562 | 100.4007 | 123.0798 | 21.2612 | 0.001869 | 0.001753 | -152.1367 | 0.24213 | 8.4019 | 0.003391 | 0.004567 | 0.004408 | 5.1e-05 | 0.001189 | 0.29159 | 1.3238 | 2.5479 | 98.91 | 20.86 | 38.638 | 66.8026 | 103.4452 | 64.0123 | 0.009643 | 1.0471 | 0.024832 | 23.7151 | -0.012882 | 0.14404 | 0.000619 | 0.04634 | 0.04634 | 0.12775
2019-03-21 15:47 | 1553197635.8661 | -155 | 0.20953 | 0.000459 | -3e-06 | 100.2973 | 100.559 | 122.5594 | 21.2612 | 0.001903 | 0.001746 | -152.1376 | 0.24111 | 8.4025 | 0.003428 | 0.004622 | 0.004329 | 4.9e-05 | 0.001189 | 0.26161 | 1.3187 | 2.548 | 98.91 | 20.86 | 38.641 | 67.6139 | 104.3554 | 63.8397 | 0.009644 | 1.0476 | 0.024786 | 23.6588 | -0.023312 | 0.14344 | -0.000367 | -0.02712 | -0.02712 | -0.21738
2019-03-21 15:47 | 1553197636.8663 | -154 | 0.20944 | 0.00046 | -3e-06 | 100.3509 | 100.4276 | 122.1122 | 21.2612 | 0.001906 | 0.001764 | -152.1361 | 0.24165 | 8.4018 | 0.003417 | 0.004588 | 0.004427 | 4.8e-05 | 0.001189 | 0.24065 | 1.3092 | 2.5479 | 98.91 | 20.86 | 38.648 | 67.6793 | 104.0787 | 64.2628 | 0.009645 | 1.0463 | 0.024725 | 23.6304 | -0.006829 | 0.14376 | 0.000162 | 0.012168 | 0.012168 | 0.48559
2019-03-21 15:47 | 1553197637.8661 | -153 | 0.2094 | 0.000456 | -3e-06 | 100.2973 | 100.4245 | 122.132 | 21.2612 | 0.001893 | 0.001748 | -152.1438 | 0.24185 | 8.4007 | 0.003419 | 0.004502 | 0.004362 | 5.2e-05 | 0.001186 | 0.2688 | 1.308 | 2.5479 | 98.91 | 20.86 | 38.652 | 67.372 | 104.1182 | 63.884 | 0.009646 | 1.0475 | 0.024741 | 23.6191 | -0.011338 | 0.14388 | 0.000414 | 0.030945 | 0.030945 | 0.19109
2019-03-21 15:47 | 1553197638.8666 | -152 | 0.2095 | 0.00046 | -3e-06 | 100.3593 | 100.2804 | 122.1183 | 21.2612 | 0.001863 | 0.001801 | -152.1432 | 0.2417 | 8.4006 | 0.003456 | 0.004471 | 0.004265 | 5e-05 | 0.001185 | 0.2625 | 1.3135 | 2.548 | 98.91 | 20.86 | 38.654 | 66.666 | 105.0218 | 65.1653 | 0.009647 | 1.0435 | 0.024693 | 23.6626 | 0.007028 | 0.14378 | -0.000154 | -0.011237 | -0.011237 | -0.52592
2019-03-21 15:47 | 1553197639.8673 | -151 | 0.20947 | 0.000457 | -3e-06 | 100.4114 | 100.4586 | 122.116 | 21.2612 | 0.001864 | 0.001753 | -152.1346 | 0.24174 | 8.3993 | 0.003431 | 0.004446 | 0.004221 | 4.9e-05 | 0.001182 | 0.26001 | 1.3382 | 2.5481 | 98.91 | 20.86 | 38.658 | 66.6681 | 104.4075 | 64.0161 | 0.009648 | 1.0471 | 0.024735 | 23.6222 | -0.0042 | 0.14381 | -9e-06 | -0.000458 | -0.000458 | -12.9071
2019-03-21 15:47 | 1553197640.8663 | -150 | 0.20944 | 0.000459 | -3e-06 | 100.4087 | 100.5485 | 122.5563 | 21.2612 | 0.001878 | 0.001782 | -152.1472 | 0.24124 | 8.4046 | 0.003441 | 0.004657 | 0.00455 | 4.9e-05 | 0.001179 | 0.27959 | 1.2917 | 2.5479 | 98.91 | 20.86 | 38.659 | 67.0224 | 104.6504 | 64.7014 | 0.009648 | 1.045 | 0.024754 | 23.6887 | -0.012441 | 0.14351 | 0.000157 | 0.011858 | 0.011858 | 0.49741
2019-03-21 15:47 | 1553197641.8677 | -149 | 0.2095 | 0.000456 | -3e-06 | 100.4397 | 100.5737 | 122.8056 | 21.2612 | 0.001894 | 0.001753 | -152.1233 | 0.24207 | 8.4023 | 0.003443 | 0.004101 | 0.003889 | 5.1e-05 | 0.001181 | 0.2792 | 1.3534 | 2.548 | 98.91 | 20.86 | 38.664 | 67.39 | 104.6977 | 64.0161 | 0.009649 | 1.0471 | 0.024804 | 23.6888 | -0.011923 | 0.14401 | -0.000149 | -0.010925 | -0.010925 | -0.54178
2019-03-21 15:47 | 1553197642.8674 | -148 | 0.20943 | 0.000457 | -3e-06 | 100.5854 | 100.6098 | 123.0882 | 21.2612 | 0.001899 | 0.001779 | -152.1307 | 0.24133 | 8.4024 | 0.003464 | 0.004698 | 0.004583 | 4.9e-05 | 0.001184 | 0.2724 | 1.3564 | 2.5481 | 98.91 | 20.86 | 38.666 | 67.5086 | 105.2223 | 64.6437 | 0.00965 | 1.0452 | 0.02481 | 23.7381 | -0.002168 | 0.14356 | 0.000277 | 0.020816 | 0.020816 | 0.28346
2019-03-21 15:47 | 1553197643.8675 | -147 | 0.20948 | 0.00046 | -3e-06 | 100.6512 | 100.6258 | 123.3168 | 21.2612 | 0.001883 | 0.001811 | -152.1146 | 0.24205 | 8.4029 | 0.003509 | 0.004516 | 0.004339 | 5.2e-05 | 0.001183 | 0.2792 | 1.3373 | 2.548 | 98.91 | 20.86 | 38.669 | 67.1464 | 106.2971 | 65.4126 | 0.00965 | 1.0428 | 0.024805 | 23.7871 | 0.002263 | 0.144 | -6.4e-05 | -0.004605 | -0.004605 | -1.2853
2019-03-21 15:47 | 1553197644.8665 | -146 | 0.20946 | 0.000456 | -3e-06 | 100.5851 | 100.5509 | 123.6135 | 21.2612 | 0.001907 | 0.001752 | -152.1084 | 0.24229 | 8.4008 | 0.003423 | 0.005233 | 0.005021 | 5.2e-05 | 0.001182 | 0.28622 | 1.2977 | 2.548 | 98.91 | 20.86 | 38.672 | 67.7136 | 104.2321 | 63.9771 | 0.009651 | 1.0472 | 0.024887 | 23.7652 | 0.003046 | 0.14414 | 7.2e-05 | 0.005589 | 0.005589 | 1.06
2019-03-21 15:47 | 1553197645.866 | -145 | 0.20947 | 0.000459 | -3e-06 | 100.642 | 100.7716 | 123.9487 | 21.2612 | 0.001917 | 0.001742 | -152.1166 | 0.24192 | 8.4011 | 0.003448 | 0.004748 | 0.004641 | 5.3e-05 | 0.001182 | 0.25713 | 1.2825 | 2.5479 | 98.91 | 20.86 | 38.673 | 67.9406 | 104.8273 | 63.7521 | 0.009651 | 1.0479 | 0.024929 | 23.7894 | -0.011507 | 0.14392 | 3.2e-05 | 0.002591 | 0.002591 | 2.2831
2019-03-21 15:47 | 1553197646.8667 | -144 | 0.20948 | 0.00046 | -3e-06 | 100.6286 | 100.7048 | 124.2424 | 21.2612 | 0.001882 | 0.001781 | -152.1176 | 0.24162 | 8.4038 | 0.003537 | 0.004373 | 0.004286 | 5.6e-05 | 0.001188 | 0.27936 | 1.3071 | 2.5479 | 98.91 | 20.86 | 38.673 | 67.1068 | 106.9835 | 64.686 | 0.009651 | 1.045 | 0.024924 | 23.8506 | -0.006771 | 0.14374 | -6.8e-05 | -0.004869 | -0.004869 | -1.2133
2019-03-21 15:47 | 1553197647.8667 | -143 | 0.20945 | 0.000454 | -3e-06 | 100.6307 | 100.7048 | 124.4137 | 21.2612 | 0.001857 | 0.001803 | -152.1049 | 0.24196 | 8.4022 | 0.003499 | 0.004564 | 0.004306 | 5e-05 | 0.001182 | 0.26217 | 1.3291 | 2.5479 | 98.91 | 20.86 | 38.675 | 66.5099 | 106.0727 | 65.2098 | 0.009652 | 1.0434 | 0.024922 | 23.8855 | -0.006586 | 0.14394 | 0.000156 | 0.011903 | 0.011903 | 0.49701
2019-03-21 15:47 | 1553197648.8661 | -142 | 0.20943 | 0.000458 | -3e-06 | 100.6474 | 100.7542 | 124.4311 | 21.2612 | 0.001949 | 0.001805 | -152.1045 | 0.24189 | 8.4025 | 0.003491 | 0.005246 | 0.005044 | 5.2e-05 | 0.001182 | 0.25105 | 1.2646 | 2.5479 | 98.91 | 20.86 | 38.676 | 68.7152 | 105.8746 | 65.2519 | 0.009652 | 1.0433 | 0.024922 | 23.8887 | -0.009488 | 0.1439 | 0.000233 | 0.017654 | 0.017654 | 0.33502
2019-03-21 15:47 | 1553197649.8672 | -141 | 0.20947 | 0.000458 | -3e-06 | 100.7129 | 100.7123 | 124.3837 | 21.2612 | 0.001954 | 0.001727 | -152.1113 | 0.24115 | 8.4023 | 0.003477 | 0.004866 | 0.004753 | 5.5e-05 | 0.001184 | 0.27041 | 1.2895 | 2.5481 | 98.91 | 20.86 | 38.679 | 68.8423 | 105.5251 | 63.375 | 0.009653 | 1.0491 | 0.024987 | 23.8178 | 5.5e-05 | 0.14346 | 3.6e-05 | 0.002889 | 0.002889 | 2.0406
2019-03-21 15:48 | 1553197650.8668 | -140 | 0.20944 | 0.00046 | -3e-06 | 100.6337 | 100.8199 | 124.4006 | 21.2612 | 0.001915 | 0.001733 | -152.1106 | 0.24144 | 8.403 | 0.003384 | 0.004909 | 0.004729 | 5.6e-05 | 0.001186 | 0.24635 | 1.3232 | 2.5481 | 98.91 | 20.86 | 38.681 | 67.894 | 103.2865 | 63.5304 | 0.009653 | 1.0486 | 0.024983 | 23.8249 | -0.01654 | 0.14363 | 0.000169 | 0.012857 | 0.012857 | 0.45914
2019-03-21 15:48 | 1553197651.8663 | -139 | 0.20946 | 0.000457 | -3e-06 | 100.7418 | 100.8509 | 124.1363 | 21.2612 | 0.001845 | 0.00178 | -152.1141 | 0.24096 | 8.4021 | 0.003436 | 0.005058 | 0.004837 | 5.5e-05 | 0.001184 | 0.26272 | 1.3434 | 2.548 | 98.91 | 20.86 | 38.682 | 66.2266 | 104.5467 | 64.6562 | 0.009654 | 1.0451 | 0.024915 | 23.8394 | -0.009681 | 0.14335 | 9.1e-05 | 0.006983 | 0.006983 | 0.84371
2019-03-21 15:48 | 1553197652.8663 | -138 | 0.20943 | 0.000458 | -3e-06 | 100.7403 | 100.7242 | 123.8138 | 21.2612 | 0.001824 | 0.001726 | -152.1092 | 0.24152 | 8.4009 | 0.003518 | 0.004906 | 0.00475 | 5.4e-05 | 0.00118 | 0.26145 | 1.3248 | 2.5478 | 98.91 | 20.86 | 38.684 | 65.7256 | 106.5349 | 63.3641 | 0.009654 | 1.0491 | 0.02493 | 23.7628 | 0.001431 | 0.14368 | 0.00024 | 0.018215 | 0.018215 | 0.3242
2019-03-21 15:48 | 1553197653.866 | -137 | 0.20945 | 0.000455 | -3e-06 | 100.7081 | 100.6844 | 123.6703 | 21.2612 | 0.001921 | 0.001837 | -152.108 | 0.24101 | 8.4025 | 0.003519 | 0.005299 | 0.005205 | 5.6e-05 | 0.001185 | 0.26969 | 1.3593 | 2.548 | 98.91 | 20.86 | 38.684 | 68.0476 | 106.5412 | 66.0291 | 0.009654 | 1.0409 | 0.024818 | 23.8429 | 0.002113 | 0.14338 | 0.000131 | 0.009968 | 0.009968 | 0.59115
2019-03-21 15:48 | 1553197654.8661 | -136 | 0.20944 | 0.000458 | -3e-06 | 100.7058 | 100.8441 | 123.3497 | 21.2612 | 0.001887 | 0.001769 | -152.1197 | 0.24129 | 8.4015 | 0.003473 | 0.004981 | 0.0048 | 5.4e-05 | 0.001186 | 0.26648 | 1.3568 | 2.5481 | 98.91 | 20.86 | 38.685 | 67.2246 | 105.4465 | 64.4009 | 0.009654 | 1.0459 | 0.024845 | 23.7547 | -0.012278 | 0.14355 | 0.000183 | 0.013836 | 0.013836 | 0.42641
2019-03-21 15:48 | 1553197655.8666 | -135 | 0.20946 | 0.000458 | -3e-06 | 100.7531 | 100.8369 | 123.1981 | 21.2612 | 0.00194 | 0.001728 | -152.1328 | 0.24101 | 8.4023 | 0.003513 | 0.005262 | 0.005104 | 5.6e-05 | 0.001185 | 0.2625 | 1.349 | 2.548 | 98.91 | 20.86 | 38.685 | 68.4991 | 106.4136 | 63.4101 | 0.009654 | 1.049 | 0.024866 | 23.7053 | -0.007437 | 0.14338 | 4.6e-05 | 0.003599 | 0.003599 | 1.6374
2019-03-21 15:48 | 1553197656.8672 | -134 | 0.20941 | 0.000461 | -3e-06 | 100.746 | 100.9558 | 123.2563 | 21.2612 | 0.001904 | 0.001753 | -152.125 | 0.24159 | 8.4015 | 0.003518 | 0.005138 | 0.004929 | 5.6e-05 | 0.00118 | 0.25321 | 1.3436 | 2.5479 | 98.91 | 20.86 | 38.685 | 67.6348 | 106.5293 | 64.0005 | 0.009654 | 1.0471 | 0.02485 | 23.7316 | -0.01861 | 0.14372 | 0.000372 | 0.02796 | 0.02796 | 0.21126
2019-03-21 15:48 | 1553197657.8673 | -133 | 0.20946 | 0.000462 | -3e-06 | 100.7722 | 101.0307 | 123.207 | 21.2612 | 0.001853 | 0.001738 | -152.1125 | 0.24097 | 8.4014 | 0.003605 | 0.004591 | 0.004366 | 5.7e-05 | 0.001181 | 0.27367 | 1.2867 | 2.5478 | 98.91 | 20.86 | 38.686 | 66.4115 | 108.6503 | 63.6491 | 0.009655 | 1.0482 | 0.024858 | 23.7145 | -0.022918 | 0.14335 | 4.4e-05 | 0.003477 | 0.003477 | 1.6945
2019-03-21 15:48 | 1553197658.8663 | -132 | 0.20942 | 0.000459 | -3e-06 | 100.7832 | 101.0805 | 123.1368 | 21.2612 | 0.001937 | 0.001719 | -152.1233 | 0.24156 | 8.4013 | 0.003537 | 0.005042 | 0.004846 | 5.8e-05 | 0.00118 | 0.2756 | 1.2715 | 2.5478 | 98.91 | 20.86 | 38.686 | 68.4351 | 106.9789 | 63.179 | 0.009655 | 1.0497 | 0.024869 | 23.6912 | -0.026341 | 0.1437 | 0.000286 | 0.021561 | 0.021561 | 0.27393
2019-03-21 15:48 | 1553197659.8677 | -131 | 0.20943 | 0.000462 | -3e-06 | 100.729 | 100.9173 | 123.2079 | 21.2612 | 0.001915 | 0.001775 | -152.114 | 0.24092 | 8.4032 | 0.003453 | 0.00506 | 0.004907 | 5.3e-05 | 0.001179 | 0.27041 | 1.3275 | 2.5481 | 98.91 | 20.86 | 38.685 | 67.9062 | 104.9595 | 64.5393 | 0.009654 | 1.0455 | 0.024826 | 23.7459 | -0.016707 | 0.14333 | 0.000253 | 0.019016 | 0.019016 | 0.30978
2019-03-21 15:48 | 1553197660.8668 | -130 | 0.20946 | 0.000464 | -3e-06 | 100.7353 | 100.8754 | 123.2729 | 21.2612 | 0.001854 | 0.001752 | -152.115 | 0.24094 | 8.4005 | 0.003446 | 0.004824 | 0.004607 | 5.5e-05 | 0.001179 | 0.28263 | 1.3343 | 2.548 | 98.91 | 20.85 | 38.685 | 66.4341 | 104.7919 | 63.988 | 0.009649 | 1.0472 | 0.024853 | 23.7328 | -0.012438 | 0.14333 | 7.3e-05 | 0.005633 | 0.005633 | 1.0457
2019-03-21 15:48 | 1553197661.8654 | -129 | 0.20944 | 0.000458 | -3e-06 | 100.7925 | 100.934 | 123.4725 | 21.2612 | 0.001885 | 0.001769 | -152.1318 | 0.24101 | 8.4058 | 0.003533 | 0.004321 | 0.004125 | 5.3e-05 | 0.001182 | 0.2568 | 1.3823 | 2.5482 | 98.91 | 20.85 | 38.684 | 67.1847 | 106.8843 | 64.3939 | 0.009648 | 1.0459 | 0.024858 | 23.7663 | -0.012553 | 0.14338 | 0.00021 | 0.015872 | 0.015872 | 0.37127
2019-03-21 15:48 | 1553197662.8655 | -128 | 0.20945 | 0.000461 | -3e-06 | 100.8258 | 100.8111 | 123.7457 | 21.2612 | 0.001834 | 0.001798 | -152.1326 | 0.24149 | 8.4021 | 0.003574 | 0.004816 | 0.00467 | 5.4e-05 | 0.001185 | 0.26322 | 1.341 | 2.5479 | 98.91 | 20.86 | 38.684 | 65.9464 | 107.8868 | 65.0833 | 0.009654 | 1.0438 | 0.02486 | 23.8169 | 0.00131 | 0.14366 | 0.000125 | 0.009501 | 0.009501 | 0.62147
2019-03-21 15:48 | 1553197663.8666 | -127 | 0.20944 | 0.000458 | -3e-06 | 100.7856 | 100.7474 | 123.7987 | 21.2612 | 0.001835 | 0.001782 | -152.1178 | 0.24013 | 8.4019 | 0.003519 | 0.004937 | 0.004829 | 5.2e-05 | 0.00118 | 0.26305 | 1.2861 | 2.5479 | 98.91 | 20.86 | 38.684 | 65.9902 | 106.5538 | 64.7171 | 0.009654 | 1.0449 | 0.024879 | 23.8091 | 0.003394 | 0.14285 | 0.000173 | 0.013082 | 0.013082 | 0.4488
2019-03-21 15:48 | 1553197664.8659 | -126 | 0.20946 | 0.000465 | -3e-06 | 100.8178 | 100.9105 | 123.9499 | 21.2612 | 0.001957 | 0.001788 | -152.0892 | 0.24133 | 8.4012 | 0.003581 | 0.004978 | 0.004777 | 5.1e-05 | 0.00118 | 0.27776 | 1.3588 | 2.5481 | 98.91 | 20.85 | 38.684 | 68.9125 | 108.0538 | 64.8586 | 0.009648 | 1.0445 | 0.024889 | 23.8286 | -0.008223 | 0.14356 | 3.9e-05 | 0.00307 | 0.00307 | 1.9218
2019-03-21 15:48 | 1553197665.8656 | -125 | 0.20944 | 0.000461 | -3e-06 | 100.7841 | 100.8979 | 123.9478 | 21.2612 | 0.001899 | 0.001738 | -152.0819 | 0.24064 | 8.4029 | 0.003546 | 0.005021 | 0.00489 | 5e-05 | 0.00118 | 0.25481 | 1.3207 | 2.548 | 98.91 | 20.85 | 38.684 | 67.5131 | 107.2143 | 63.6349 | 0.009648 | 1.0483 | 0.024934 | 23.7852 | -0.010093 | 0.14316 | 0.000181 | 0.013768 | 0.013768 | 0.42736
2019-03-21 15:48 | 1553197666.8657 | -124 | 0.20946 | 0.000464 | -3e-06 | 100.8175 | 100.9517 | 123.9159 | 21.2612 | 0.001879 | 0.001798 | -152.1021 | 0.24208 | 8.3999 | 0.003548 | 0.004587 | 0.004402 | 5.1e-05 | 0.001181 | 0.27472 | 1.3345 | 2.548 | 98.91 | 20.85 | 38.684 | 67.0474 | 107.2498 | 65.0996 | 0.009648 | 1.0437 | 0.024876 | 23.8338 | -0.011901 | 0.14401 | 5.5e-05 | 0.004271 | 0.004271 | 1.386
2019-03-21 15:48 | 1553197667.8662 | -123 | 0.20945 | 0.000461 | -3e-06 | 100.8708 | 100.9671 | 123.848 | 21.2612 | 0.001826 | 0.001782 | -152.1078 | 0.24222 | 8.4 | 0.003542 | 0.004835 | 0.004636 | 5.3e-05 | 0.001179 | 0.26488 | 1.3306 | 2.5479 | 98.91 | 20.85 | 38.683 | 65.7624 | 107.1039 | 64.7149 | 0.009648 | 1.0449 | 0.024884 | 23.8137 | -0.008529 | 0.1441 | 0.000137 | 0.010402 | 0.010402 | 0.56938
2019-03-21 15:48 | 1553197668.8668 | -122 | 0.20947 | 0.00046 | -3e-06 | 100.886 | 100.7436 | 123.6713 | 21.2612 | 0.001893 | 0.001786 | -152.1047 | 0.2414 | 8.4019 | 0.003503 | 0.005082 | 0.004864 | 5.3e-05 | 0.001184 | 0.28368 | 1.3009 | 2.5479 | 98.91 | 20.86 | 38.683 | 67.3751 | 106.1569 | 64.8131 | 0.009654 | 1.0446 | 0.024862 | 23.8002 | 0.012633 | 0.14361 | -1.4e-05 | -0.000812 | -0.000812 | -7.2677
2019-03-21 15:48 | 1553197669.8668 | -121 | 0.20949 | 0.000461 | -3e-06 | 100.8306 | 100.9895 | 123.6011 | 21.2612 | 0.001851 | 0.00178 | -152.1202 | 0.24157 | 8.4019 | 0.003459 | 0.00408 | 0.003913 | 5.1e-05 | 0.001181 | 0.29391 | 1.2899 | 2.548 | 98.91 | 20.85 | 38.683 | 66.3757 | 105.094 | 64.6687 | 0.009648 | 1.0451 | 0.024861 | 23.7883 | -0.014087 | 0.14371 | -0.000114 | -0.008331 | -0.008331 | -0.709
2019-03-21 15:48 | 1553197670.8672 | -120 | 0.20946 | 0.000463 | -3e-06 | 100.9051 | 101.1254 | 123.5774 | 21.2612 | 0.001889 | 0.001798 | -152.113 | 0.24109 | 8.4037 | 0.00344 | 0.00503 | 0.004882 | 5.1e-05 | 0.00118 | 0.26665 | 1.3032 | 2.5478 | 98.91 | 20.86 | 38.682 | 67.2775 | 104.6402 | 65.0816 | 0.009654 | 1.0438 | 0.024843 | 23.8006 | -0.01951 | 0.14342 | 5.6e-05 | 0.004391 | 0.004391 | 1.3425
2019-03-21 15:48 | 1553197671.8659 | -119 | 0.20944 | 0.000461 | -3e-06 | 100.8556 | 100.8962 | 123.6067 | 21.2612 | 0.001952 | 0.001761 | -152.1103 | 0.24144 | 8.4041 | 0.003493 | 0.004712 | 0.004604 | 5.3e-05 | 0.001177 | 0.25984 | 1.3049 | 2.5479 | 98.91 | 20.86 | 38.683 | 68.7924 | 105.9163 | 64.1901 | 0.009654 | 1.0466 | 0.024879 | 23.772 | -0.003598 | 0.14363 | 0.000166 | 0.012567 | 0.012567 | 0.46974
2019-03-21 15:48 | 1553197672.8661 | -118 | 0.20945 | 0.000456 | -3e-06 | 100.8127 | 100.7538 | 123.6232 | 21.2612 | 0.001942 | 0.001753 | -152.1122 | 0.24166 | 8.404 | 0.003512 | 0.005247 | 0.005055 | 4.9e-05 | 0.001177 | 0.23739 | 1.3404 | 2.5479 | 98.91 | 20.86 | 38.684 | 68.556 | 106.3932 | 64.0144 | 0.009654 | 1.0471 | 0.024887 | 23.7674 | 0.005226 | 0.14376 | 0.000143 | 0.010854 | 0.010854 | 0.54439
2019-03-21 15:48 | 1553197673.866 | -117 | 0.20946 | 0.000461 | -3e-06 | 100.8756 | 100.9453 | 123.6257 | 21.2612 | 0.001903 | 0.001769 | -152.1158 | 0.24086 | 8.403 | 0.003398 | 0.004935 | 0.004754 | 5.1e-05 | 0.001177 | 0.2912 | 1.33 | 2.548 | 98.91 | 20.86 | 38.685 | 67.616 | 103.6244 | 64.404 | 0.009654 | 1.0459 | 0.024873 | 23.7814 | -0.006175 | 0.14329 | 9.5e-05 | 0.007218 | 0.007218 | 0.81586
2019-03-21 15:48 | 1553197674.8675 | -116 | 0.20945 | 0.000466 | -3e-06 | 100.8118 | 100.7675 | 123.6499 | 21.2612 | 0.001842 | 0.001758 | -152.1378 | 0.24166 | 8.4012 | 0.003417 | 0.00503 | 0.00481 | 4.8e-05 | 0.001176 | 0.28224 | 1.326 | 2.548 | 98.91 | 20.86 | 38.688 | 66.1532 | 104.0724 | 64.1167 | 0.009655 | 1.0468 | 0.024886 | 23.7736 | 0.003937 | 0.14376 | 0.000113 | 0.008632 | 0.008632 | 0.68451
2019-03-21 15:48 | 1553197675.8662 | -115 | 0.20943 | 0.000463 | -3e-06 | 100.8455 | 100.9783 | 124.0771 | 21.2612 | 0.001873 | 0.001783 | -152.1398 | 0.24055 | 8.4039 | 0.003402 | 0.0045 | 0.004443 | 5.3e-05 | 0.001176 | 0.28766 | 1.3202 | 2.5479 | 98.91 | 20.86 | 38.689 | 66.893 | 103.7053 | 64.7195 | 0.009655 | 1.0449 | 0.024907 | 23.8359 | -0.01177 | 0.1431 | 0.000219 | 0.016554 | 0.016554 | 0.35529
2019-03-21 15:48 | 1553197676.8652 | -114 | 0.20945 | 0.000462 | -3e-06 | 100.794 | 101.0464 | 124.1665 | 21.2612 | 0.00188 | 0.001743 | -152.1316 | 0.24142 | 8.4029 | 0.003453 | 0.004569 | 0.004291 | 5.2e-05 | 0.001174 | 0.23794 | 1.2937 | 2.5479 | 98.91 | 20.86 | 38.69 | 67.0637 | 104.9437 | 63.7779 | 0.009656 | 1.0478 | 0.02495 | 23.8112 | -0.022375 | 0.14362 | 0.000141 | 0.01076 | 0.01076 | 0.54857
2019-03-21 15:48 | 1553197677.8657 | -113 | 0.20945 | 0.000462 | -4e-06 | 100.8011 | 101.0127 | 124.2522 | 21.2612 | 0.001904 | 0.001725 | -152.1515 | 0.24098 | 8.4014 | 0.00344 | 0.004457 | 0.004277 | 5.3e-05 | 0.001181 | 0.27417 | 1.3063 | 2.548 | 98.91 | 20.86 | 38.692 | 67.6302 | 104.648 | 63.3328 | 0.009656 | 1.0492 | 0.024975 | 23.8037 | -0.018755 | 0.14336 | 0.00012 | 0.0092 | 0.0092 | 0.64043
2019-03-21 15:48 | 1553197678.8654 | -112 | 0.20947 | 0.000462 | -3e-06 | 100.9337 | 101.1094 | 124.5984 | 21.2612 | 0.001899 | 0.0018 | -152.1559 | 0.24182 | 8.3991 | 0.003485 | 0.004309 | 0.004121 | 4.7e-05 | 0.00118 | 0.26034 | 1.3231 | 2.5479 | 98.91 | 20.85 | 38.693 | 67.5138 | 105.7186 | 65.1333 | 0.009651 | 1.0436 | 0.024944 | 23.9005 | -0.015559 | 0.14386 | 3.4e-05 | 0.002712 | 0.002712 | 2.1799
2019-03-21 15:48 | 1553197679.8654 | -111 | 0.20945 | 0.00046 | -3e-06 | 100.7266 | 100.901 | 124.6045 | 21.2612 | 0.001854 | 0.001779 | -152.1442 | 0.24164 | 8.4059 | 0.003489 | 0.004619 | 0.004463 | 5.2e-05 | 0.001181 | 0.26936 | 1.3695 | 2.5479 | 98.91 | 20.86 | 38.695 | 66.439 | 105.8319 | 64.6423 | 0.009657 | 1.0452 | 0.024962 | 23.8838 | -0.01547 | 0.14375 | 0.000119 | 0.009154 | 0.009154 | 0.64543
2019-03-21 15:48 | 1553197680.8658 | -110 | 0.20946 | 0.000464 | -3e-06 | 100.858 | 100.9684 | 124.4923 | 21.2612 | 0.001941 | 0.001758 | -152.1362 | 0.24138 | 8.403 | 0.003493 | 0.004352 | 0.004213 | 5.6e-05 | 0.001174 | 0.26648 | 1.3212 | 2.5479 | 98.91 | 20.86 | 38.697 | 68.5226 | 105.9216 | 64.1309 | 0.009657 | 1.0467 | 0.02497 | 23.8549 | -0.009786 | 0.1436 | 8.7e-05 | 0.006683 | 0.006683 | 0.88309
2019-03-21 15:48 | 1553197681.8665 | -109 | 0.20944 | 0.000459 | -3e-06 | 100.8446 | 100.8962 | 124.4928 | 21.2612 | 0.00187 | 0.001779 | -152.127 | 0.24108 | 8.4001 | 0.003509 | 0.004829 | 0.004601 | 5.1e-05 | 0.001175 | 0.29502 | 1.2767 | 2.548 | 98.91 | 20.86 | 38.699 | 66.8117 | 106.3105 | 64.6447 | 0.009658 | 1.0451 | 0.024951 | 23.8732 | -0.004576 | 0.14342 | 0.000217 | 0.016502 | 0.016502 | 0.35721
2019-03-21 15:48 | 1553197682.8666 | -108 | 0.20944 | 0.000462 | -3e-06 | 100.8047 | 100.8983 | 124.4153 | 21.2612 | 0.001849 | 0.001831 | -152.1205 | 0.24127 | 8.4 | 0.003467 | 0.004515 | 0.004346 | 5.3e-05 | 0.001175 | 0.28279 | 1.3301 | 2.548 | 98.91 | 20.86 | 38.701 | 66.3108 | 105.2977 | 65.8911 | 0.009658 | 1.0413 | 0.024897 | 23.9097 | -0.008299 | 0.14353 | 0.00018 | 0.013653 | 0.013653 | 0.43207
2019-03-21 15:48 | 1553197683.8659 | -107 | 0.20946 | 0.00046 | -3e-06 | 100.8455 | 100.7617 | 124.2226 | 21.2612 | 0.001878 | 0.001791 | -152.1451 | 0.24149 | 8.4025 | 0.003514 | 0.004617 | 0.00445 | 5.3e-05 | 0.001176 | 0.27328 | 1.2785 | 2.548 | 98.91 | 20.86 | 38.703 | 67.0154 | 106.4276 | 64.9115 | 0.009659 | 1.0443 | 0.024914 | 23.8567 | 0.007438 | 0.14366 | 7.1e-05 | 0.005513 | 0.005513 | 1.071
2019-03-21 15:48 | 1553197684.8657 | -106 | 0.20946 | 0.000462 | -3e-06 | 100.7633 | 100.9139 | 124.0527 | 21.2612 | 0.00197 | 0.001759 | -152.1351 | 0.24155 | 8.4018 | 0.003491 | 0.004908 | 0.00477 | 5.1e-05 | 0.001176 | 0.29231 | 1.31 | 2.548 | 98.91 | 20.86 | 38.703 | 69.219 | 105.8721 | 64.141 | 0.009659 | 1.0467 | 0.024925 | 23.8131 | -0.013363 | 0.1437 | 5.1e-05 | 0.004006 | 0.004006 | 1.4743
2019-03-21 15:48 | 1553197685.8671 | -105 | 0.20943 | 0.000463 | -3e-06 | 100.7531 | 100.9292 | 123.5957 | 21.2612 | 0.001851 | 0.001763 | -152.1296 | 0.24075 | 8.405 | 0.003528 | 0.004182 | 0.004062 | 5.4e-05 | 0.001176 | 0.2688 | 1.3387 | 2.548 | 98.91 | 20.86 | 38.703 | 66.3702 | 106.7639 | 64.2457 | 0.009659 | 1.0464 | 0.024876 | 23.7729 | -0.015621 | 0.14322 | 0.000234 | 0.017713 | 0.017713 | 0.33233
2019-03-21 15:48 | 1553197686.8668 | -104 | 0.20947 | 0.000463 | -3e-06 | 100.7999 | 100.9531 | 123.3582 | 21.2612 | 0.001946 | 0.001737 | -152.1217 | 0.24202 | 8.4022 | 0.003535 | 0.005116 | 0.004914 | 5.4e-05 | 0.001179 | 0.29535 | 1.3536 | 2.548 | 98.91 | 20.86 | 38.704 | 68.6558 | 106.9488 | 63.6286 | 0.009659 | 1.0483 | 0.024874 | 23.7283 | -0.013582 | 0.14397 | -1.8e-05 | -0.001148 | -0.001148 | -5.1525
2019-03-21 15:48 | 1553197687.8659 | -103 | 0.20945 | 0.000461 | -4e-06 | 100.7975 | 100.7971 | 123.1624 | 21.2612 | 0.001877 | 0.001715 | -152.1055 | 0.24178 | 8.4048 | 0.003574 | 0.004723 | 0.004514 | 5.4e-05 | 0.001176 | 0.27312 | 1.3737 | 2.548 | 98.91 | 20.86 | 38.705 | 66.9904 | 107.8766 | 63.1046 | 0.009659 | 1.0499 | 0.024874 | 23.6911 | 3.8e-05 | 0.14384 | 0.000105 | 0.008092 | 0.008092 | 0.73055
2019-03-21 15:48 | 1553197688.8663 | -102 | 0.20946 | 0.000464 | -3e-06 | 100.7749 | 100.872 | 122.8108 | 21.2612 | 0.001833 | 0.001759 | -152.1172 | 0.24187 | 8.4036 | 0.003539 | 0.005049 | 0.004828 | 5.5e-05 | 0.001175 | 0.28279 | 1.3402 | 2.5479 | 98.91 | 20.86 | 38.707 | 65.9324 | 107.0394 | 64.1449 | 0.00966 | 1.0467 | 0.0248 | 23.6938 | -0.008618 | 0.14389 | 6.8e-05 | 0.005229 | 0.005229 | 1.131
2019-03-21 15:48 | 1553197689.8659 | -101 | 0.20944 | 0.00046 | -3e-06 | 100.7969 | 100.8618 | 122.7798 | 21.2612 | 0.001848 | 0.001779 | -152.1299 | 0.24128 | 8.4031 | 0.003507 | 0.0044 | 0.004301 | 5.5e-05 | 0.001174 | 0.30525 | 1.3075 | 2.5479 | 98.91 | 20.86 | 38.707 | 66.2922 | 106.2537 | 64.6266 | 0.00966 | 1.0452 | 0.024779 | 23.7077 | -0.005755 | 0.14354 | 0.000184 | 0.013907 | 0.013907 | 0.4242
2019-03-21 15:48 | 1553197690.867 | -100 | 0.20945 | 0.000458 | -3e-06 | 100.6667 | 100.9136 | 122.7549 | 21.2612 | 0.001915 | 0.001767 | -152.1514 | 0.24183 | 8.4025 | 0.003547 | 0.005033 | 0.004817 | 5.3e-05 | 0.001174 | 0.28528 | 1.3242 | 2.5481 | 98.91 | 20.86 | 38.707 | 67.9094 | 107.2284 | 64.3542 | 0.00966 | 1.046 | 0.024787 | 23.6957 | -0.021906 | 0.14387 | 0.000118 | 0.009009 | 0.009009 | 0.65634
2019-03-21 15:48 | 1553197691.867 | -99 | 0.20944 | 0.000461 | -3e-06 | 100.7508 | 100.7228 | 122.6134 | 21.2612 | 0.001958 | 0.001783 | -152.1527 | 0.24089 | 8.3993 | 0.003544 | 0.00492 | 0.004769 | 5.2e-05 | 0.001173 | 0.26178 | 1.3466 | 2.5482 | 98.91 | 20.86 | 38.708 | 68.9375 | 107.1505 | 64.7375 | 0.00966 | 1.0449 | 0.024759 | 23.6955 | 0.002477 | 0.1433 | 0.000201 | 0.015151 | 0.015151 | 0.38875
2019-03-21 15:48 | 1553197692.8661 | -98 | 0.20946 | 0.000461 | -4e-06 | 100.7782 | 100.9149 | 122.5493 | 21.2612 | 0.001917 | 0.001768 | -152.1562 | 0.24145 | 8.4008 | 0.003503 | 0.00519 | 0.004989 | 5.3e-05 | 0.001176 | 0.27024 | 1.3157 | 2.548 | 98.91 | 20.86 | 38.709 | 67.9594 | 106.1554 | 64.3675 | 0.00966 | 1.046 | 0.024766 | 23.6764 | -0.012132 | 0.14364 | 9e-05 | 0.006892 | 0.006892 | 0.85662
2019-03-21 15:48 | 1553197693.8665 | -97 | 0.20945 | 0.000463 | -4e-06 | 100.7436 | 100.6615 | 122.7385 | 21.2612 | 0.001885 | 0.001787 | -152.1346 | 0.2416 | 8.4041 | 0.003483 | 0.005132 | 0.004959 | 5.2e-05 | 0.001175 | 0.27776 | 1.3213 | 2.548 | 98.91 | 20.86 | 38.708 | 67.1856 | 105.6849 | 64.8234 | 0.00966 | 1.0446 | 0.024768 | 23.7106 | 0.007288 | 0.14373 | 0.000133 | 0.010089 | 0.010089 | 0.58551
2019-03-21 15:48 | 1553197694.8652 | -96 | 0.20946 | 0.00046 | -3e-06 | 100.7061 | 100.8284 | 123.1152 | 21.2612 | 0.001925 | 0.001799 | -152.1264 | 0.24166 | 8.4016 | 0.003419 | 0.004759 | 0.004617 | 5.4e-05 | 0.001179 | 0.27417 | 1.3254 | 2.548 | 98.91 | 20.85 | 38.708 | 68.1529 | 104.1354 | 65.1215 | 0.009654 | 1.0437 | 0.024795 | 23.7575 | -0.010862 | 0.14376 | 7.2e-05 | 0.0056 | 0.0056 | 1.0551
2019-03-21 15:48 | 1553197695.8663 | -95 | 0.20948 | 0.000457 | -3e-06 | 100.7224 | 100.9085 | 123.3927 | 21.2612 | 0.001875 | 0.001791 | -152.1272 | 0.23972 | 8.401 | 0.003422 | 0.004608 | 0.004494 | 5.3e-05 | 0.001176 | 0.27367 | 1.3377 | 2.548 | 98.91 | 20.86 | 38.708 | 66.9459 | 104.189 | 64.931 | 0.00966 | 1.0443 | 0.02483 | 23.7775 | -0.016504 | 0.14261 | -5e-05 | -0.003543 | -0.003543 | -1.6541
2019-03-21 15:48 | 1553197696.8667 | -94 | 0.20949 | 0.000459 | -3e-06 | 100.6941 | 100.7106 | 123.425 | 21.2612 | 0.001889 | 0.001772 | -152.1242 | 0.24101 | 8.4015 | 0.003495 | 0.004763 | 0.004587 | 5.1e-05 | 0.001179 | 0.27649 | 1.3234 | 2.5479 | 98.91 | 20.86 | 38.708 | 67.2722 | 105.9626 | 64.4579 | 0.00966 | 1.0457 | 0.024851 | 23.764 | -0.001461 | 0.14338 | -0.000123 | -0.008957 | -0.008957 | -0.65789
2019-03-21 15:48 | 1553197697.8664 | -93 | 0.20942 | 0.000463 | -4e-06 | 100.6673 | 100.7596 | 123.4757 | 21.2612 | 0.001894 | 0.001749 | -152.1204 | 0.24037 | 8.4029 | 0.003482 | 0.004939 | 0.00482 | 5.2e-05 | 0.001176 | 0.29015 | 1.2914 | 2.548 | 98.91 | 20.86 | 38.708 | 67.3984 | 105.666 | 63.9037 | 0.00966 | 1.0474 | 0.024876 | 23.7493 | -0.008198 | 0.143 | 0.000315 | 0.023778 | 0.023778 | 0.24717
2019-03-21 15:48 | 1553197698.8664 | -92 | 0.20947 | 0.00046 | -4e-06 | 100.6778 | 100.7733 | 123.7821 | 21.2612 | 0.001885 | 0.001757 | -152.0819 | 0.24124 | 8.4038 | 0.003497 | 0.004223 | 0.004048 | 5.3e-05 | 0.001177 | 0.2896 | 1.3146 | 2.548 | 98.91 | 20.85 | 38.707 | 67.184 | 106.0184 | 64.1059 | 0.009654 | 1.0468 | 0.0249 | 23.7859 | -0.008481 | 0.14351 | -2.2e-05 | -0.001413 | -0.001413 | -4.1732
2019-03-21 15:48 | 1553197699.8659 | -91 | 0.20945 | 0.000457 | -3e-06 | 100.7102 | 100.6976 | 123.7508 | 21.2612 | 0.001906 | 0.001743 | -152.0764 | 0.24111 | 8.4034 | 0.003476 | 0.004898 | 0.004765 | 5.5e-05 | 0.001176 | 0.26626 | 1.3258 | 2.5479 | 98.91 | 20.85 | 38.706 | 67.6771 | 105.4992 | 63.7786 | 0.009654 | 1.0478 | 0.024908 | 23.7714 | 0.001118 | 0.14343 | 0.000136 | 0.01037 | 0.01037 | 0.5685
2019-03-21 15:48 | 1553197700.8665 | -90 | 0.20944 | 0.000458 | -4e-06 | 100.5526 | 100.793 | 123.8227 | 21.2612 | 0.001882 | 0.001779 | -152.0744 | 0.24142 | 8.4018 | 0.003416 | 0.004593 | 0.004435 | 5e-05 | 0.001177 | 0.26322 | 1.3201 | 2.548 | 98.91 | 20.85 | 38.703 | 67.1207 | 104.0632 | 64.6235 | 0.009653 | 1.0452 | 0.024885 | 23.8081 | -0.021358 | 0.14362 | 0.000211 | 0.015984 | 0.015984 | 0.36929
2019-03-21 15:48 | 1553197701.8656 | -89 | 0.20946 | 0.000464 | -4e-06 | 100.5779 | 100.65 | 123.8964 | 21.2612 | 0.001883 | 0.001767 | -152.0874 | 0.24072 | 8.4022 | 0.003475 | 0.004819 | 0.004697 | 5.2e-05 | 0.001177 | 0.27649 | 1.2944 | 2.548 | 98.91 | 20.85 | 38.703 | 67.1238 | 105.4788 | 64.3501 | 0.009653 | 1.0461 | 0.024902 | 23.8055 | -0.006402 | 0.1432 | 8e-05 | 0.006192 | 0.006192 | 0.95059
2019-03-21 15:48 | 1553197702.8666 | -88 | 0.20959 | 0.000458 | -4e-06 | 100.6125 | 100.6149 | 123.8217 | 21.2612 | 0.001886 | 0.00179 | -152.076 | 0.24037 | 8.4023 | 0.003501 | 0.004981 | 0.004846 | 5.3e-05 | 0.001176 | 0.2844 | 1.3228 | 2.5479 | 98.91 | 20.85 | 38.701 | 67.1989 | 106.1152 | 64.9028 | 0.009653 | 1.0443 | 0.024874 | 23.8178 | -0.000211 | 0.143 | -0.000724 | -0.053818 | -0.053818 | -0.1092
2019-03-21 15:48 | 1553197703.866 | -87 | 0.20947 | 0.000457 | -3e-06 | 100.6486 | 100.6581 | 123.9722 | 21.2612 | 0.001923 | 0.00172 | -152.0836 | 0.24114 | 8.4009 | 0.003443 | 0.005091 | 0.004911 | 5.7e-05 | 0.001182 | 0.28998 | 1.4053 | 2.5479 | 98.91 | 20.85 | 38.699 | 68.0935 | 104.7086 | 63.2132 | 0.009652 | 1.0496 | 0.024952 | 23.7727 | -0.000851 | 0.14346 | -1.1e-05 | -0.000635 | -0.000635 | -9.2809
2019-03-21 15:48 | 1553197704.8661 | -86 | 0.20951 | 0.000453 | -4e-06 | 100.6882 | 100.6632 | 123.725 | 21.2612 | 0.00192 | 0.001813 | -152.0878 | 0.24151 | 8.3985 | 0.00343 | 0.004804 | 0.004661 | 5.4e-05 | 0.001177 | 0.28423 | 1.3602 | 2.5479 | 98.91 | 20.85 | 38.698 | 68.0139 | 104.394 | 65.4408 | 0.009652 | 1.0427 | 0.024845 | 23.8275 | 0.002215 | 0.14367 | -0.000233 | -0.017121 | -0.017121 | -0.34491
2019-03-21 15:48 | 1553197705.8668 | -85 | 0.20947 | 0.000455 | -3e-06 | 100.6301 | 100.8758 | 123.5539 | 21.2612 | 0.001819 | 0.001761 | -152.0907 | 0.24149 | 8.4005 | 0.003392 | 0.0047 | 0.004496 | 5.2e-05 | 0.001181 | 0.26842 | 1.3604 | 2.5478 | 98.91 | 20.85 | 38.697 | 65.5858 | 103.4711 | 64.2026 | 0.009652 | 1.0465 | 0.024873 | 23.7674 | -0.021811 | 0.14366 | 2.7e-05 | 0.00222 | 0.00222 | 2.6593
2019-03-21 15:48 | 1553197706.8661 | -84 | 0.20952 | 0.000454 | -4e-06 | 100.5845 | 100.7194 | 123.6837 | 21.2612 | 0.00195 | 0.00179 | -152.0958 | 0.24131 | 8.4039 | 0.003452 | 0.00522 | 0.005059 | 5.5e-05 | 0.001183 | 0.26825 | 1.3583 | 2.5481 | 98.91 | 20.85 | 38.696 | 68.7486 | 104.9209 | 64.8882 | 0.009651 | 1.0444 | 0.024861 | 23.804 | -0.011992 | 0.14356 | -0.000264 | -0.019518 | -0.019518 | -0.30229
2019-03-21 15:48 | 1553197707.8659 | -83 | 0.20951 | 0.000456 | -4e-06 | 100.5943 | 100.728 | 123.9262 | 21.2612 | 0.001892 | 0.001819 | -152.0973 | 0.24117 | 8.4033 | 0.003395 | 0.00416 | 0.004048 | 5.8e-05 | 0.001186 | 0.26377 | 1.3149 | 2.548 | 98.91 | 20.85 | 38.695 | 67.3518 | 103.5435 | 65.5851 | 0.009651 | 1.0422 | 0.02486 | 23.8519 | -0.011874 | 0.14347 | -0.000205 | -0.015046 | -0.015046 | -0.39192
2019-03-21 15:48 | 1553197708.8659 | -82 | 0.20947 | 0.000454 | -3e-06 | 100.5654 | 100.7613 | 123.6994 | 21.2612 | 0.001945 | 0.001724 | -152.1187 | 0.241 | 8.4024 | 0.003474 | 0.005056 | 0.004852 | 5.5e-05 | 0.001188 | 0.2672 | 1.3465 | 2.548 | 98.91 | 20.85 | 38.694 | 68.62 | 105.4558 | 63.3172 | 0.009651 | 1.0493 | 0.02492 | 23.7502 | -0.017407 | 0.14337 | 7e-06 | 0.000701 | 0.000701 | 8.4003
2019-03-21 15:48 | 1553197709.865 | -81 | 0.20949 | 0.000457 | -4e-06 | 100.6086 | 100.5274 | 123.6077 | 21.2612 | 0.001894 | 0.001778 | -152.1178 | 0.24118 | 8.4042 | 0.003502 | 0.004883 | 0.004766 | 5.8e-05 | 0.00119 | 0.27848 | 1.3184 | 2.5479 | 98.91 | 20.85 | 38.693 | 67.3939 | 106.1341 | 64.6023 | 0.009651 | 1.0453 | 0.024864 | 23.7866 | 0.007229 | 0.14348 | -9.6e-05 | -0.006917 | -0.006917 | -0.85258
2019-03-21 15:48 | 1553197710.8654 | -80 | 0.20948 | 0.000457 | -3e-06 | 100.5339 | 100.4994 | 123.7464 | 21.2612 | 0.001865 | 0.001785 | -152.1188 | 0.24108 | 8.4018 | 0.003376 | 0.004184 | 0.004048 | 5.4e-05 | 0.001189 | 0.2557 | 1.3368 | 2.5479 | 98.91 | 20.85 | 38.692 | 66.6917 | 103.0782 | 64.7686 | 0.00965 | 1.0448 | 0.024872 | 23.8058 | 0.003063 | 0.14342 | -3.1e-05 | -0.002108 | -0.002108 | -2.796
2019-03-21 15:49 | 1553197711.8655 | -79 | 0.20949 | 0.000452 | -3e-06 | 100.5282 | 100.7215 | 123.7102 | 21.2612 | 0.001864 | 0.001788 | -152.1268 | 0.24086 | 8.4012 | 0.003394 | 0.004307 | 0.004052 | 5.5e-05 | 0.00119 | 0.26985 | 1.3279 | 2.5479 | 98.91 | 20.85 | 38.691 | 66.6768 | 103.5127 | 64.8451 | 0.00965 | 1.0445 | 0.024865 | 23.8051 | -0.017182 | 0.14329 | -0.000115 | -0.008348 | -0.008348 | -0.70544
2019-03-21 15:49 | 1553197712.8652 | -78 | 0.20948 | 0.000459 | -4e-06 | 100.5371 | 100.5873 | 123.9351 | 21.2612 | 0.001898 | 0.001758 | -152.1385 | 0.24163 | 8.4011 | 0.003444 | 0.005015 | 0.004921 | 5.6e-05 | 0.001193 | 0.24475 | 1.3147 | 2.5479 | 98.91 | 20.85 | 38.691 | 67.4898 | 104.7418 | 64.1223 | 0.00965 | 1.0468 | 0.024914 | 23.8012 | -0.004462 | 0.14374 | -7.1e-05 | -0.0051 | -0.0051 | -1.1583
2019-03-21 15:49 | 1553197713.8662 | -77 | 0.20948 | 0.000453 | -3e-06 | 100.507 | 100.5801 | 124.0461 | 21.2612 | 0.001852 | 0.001769 | -152.1496 | 0.24188 | 8.4029 | 0.003528 | 0.004755 | 0.004544 | 5.4e-05 | 0.001193 | 0.29302 | 1.3136 | 2.5479 | 98.91 | 20.85 | 38.69 | 66.3921 | 106.778 | 64.3947 | 0.00965 | 1.0459 | 0.024915 | 23.8215 | -0.006504 | 0.14389 | -3e-05 | -0.002041 | -0.002041 | -2.8976
2019-03-21 15:49 | 1553197714.8663 | -76 | 0.2095 | 0.000456 | -4e-06 | 100.6095 | 100.5781 | 124.1191 | 21.2612 | 0.001901 | 0.001801 | -152.1545 | 0.24116 | 8.4019 | 0.003494 | 0.005192 | 0.004993 | 5.5e-05 | 0.001194 | 0.25818 | 1.3275 | 2.5478 | 98.91 | 20.85 | 38.689 | 67.5576 | 105.9507 | 65.1684 | 0.00965 | 1.0435 | 0.024894 | 23.8558 | 0.002794 | 0.14347 | -0.000166 | -0.01219 | -0.01219 | -0.48372
2019-03-21 15:49 | 1553197715.8667 | -75 | 0.20946 | 0.000456 | -4e-06 | 100.6143 | 100.6224 | 124.1107 | 21.2612 | 0.001867 | 0.001788 | -152.1511 | 0.2405 | 8.4033 | 0.003431 | 0.005027 | 0.004905 | 5.8e-05 | 0.001191 | 0.28423 | 1.3495 | 2.5481 | 98.91 | 20.85 | 38.689 | 66.747 | 104.4177 | 64.8607 | 0.00965 | 1.0445 | 0.024905 | 23.8441 | -0.000719 | 0.14307 | 5.2e-05 | 0.004137 | 0.004137 | 1.4213
2019-03-21 15:49 | 1553197716.8654 | -74 | 0.20948 | 0.000457 | -4e-06 | 100.5267 | 100.7157 | 124.0644 | 21.2612 | 0.001873 | 0.001741 | -152.1663 | 0.24111 | 8.4025 | 0.003445 | 0.005135 | 0.004932 | 5.4e-05 | 0.001191 | 0.30469 | 1.3339 | 2.548 | 98.91 | 20.84 | 38.689 | 66.8867 | 104.7527 | 63.7232 | 0.009644 | 1.048 | 0.024942 | 23.7995 | -0.016801 | 0.14344 | -6.5e-05 | -0.004622 | -0.004622 | -1.2756
2019-03-21 15:49 | 1553197717.8654 | -73 | 0.20948 | 0.000455 | -4e-06 | 100.5374 | 100.5485 | 123.8668 | 21.2612 | 0.001865 | 0.001791 | -152.1625 | 0.24129 | 8.4033 | 0.003445 | 0.004841 | 0.00469 | 5.7e-05 | 0.001189 | 0.25321 | 1.3297 | 2.548 | 98.91 | 20.85 | 38.689 | 66.6987 | 104.7614 | 64.9264 | 0.00965 | 1.0443 | 0.024878 | 23.823 | -0.000983 | 0.14355 | -4.5e-05 | -0.003153 | -0.003153 | -1.8709
2019-03-21 15:49 | 1553197718.8669 | -72 | 0.20947 | 0.000455 | -4e-06 | 100.5303 | 100.5348 | 124.1205 | 21.2612 | 0.001898 | 0.001774 | -152.1529 | 0.24123 | 8.4023 | 0.00347 | 0.004751 | 0.004574 | 5.7e-05 | 0.001193 | 0.27577 | 1.3531 | 2.5478 | 98.91 | 20.84 | 38.689 | 67.4898 | 105.3607 | 64.5025 | 0.009644 | 1.0456 | 0.024919 | 23.8324 | -0.000407 | 0.14351 | 1.9e-05 | 0.001613 | 0.001613 | 3.6566
2019-03-21 15:49 | 1553197719.8655 | -71 | 0.20947 | 0.000459 | -4e-06 | 100.5121 | 100.6609 | 124.224 | 21.2612 | 0.001962 | 0.001777 | -152.1379 | 0.24198 | 8.4021 | 0.003482 | 0.004512 | 0.004331 | 5.5e-05 | 0.001188 | 0.28335 | 1.3259 | 2.5479 | 98.91 | 20.85 | 38.688 | 69.0234 | 105.6565 | 64.5835 | 0.009649 | 1.0453 | 0.024926 | 23.8452 | -0.013229 | 0.14396 | -1e-06 | 0.000188 | 0.000188 | 31.4246
2019-03-21 15:49 | 1553197720.8666 | -70 | 0.20947 | 0.000456 | -4e-06 | 100.5255 | 100.4892 | 123.9818 | 21.2612 | 0.001868 | 0.001838 | -152.1375 | 0.24182 | 8.4029 | 0.003472 | 0.004299 | 0.004145 | 5.4e-05 | 0.00119 | 0.25299 | 1.3064 | 2.5478 | 98.91 | 20.85 | 38.688 | 66.7797 | 105.4133 | 66.0597 | 0.009649 | 1.0408 | 0.024848 | 23.874 | 0.003231 | 0.14386 | 1.8e-05 | 0.001563 | 0.001563 | 3.7835
2019-03-21 15:49 | 1553197721.866 | -69 | 0.20948 | 0.000452 | -4e-06 | 100.4909 | 100.5151 | 123.979 | 21.2612 | 0.001903 | 0.001794 | -152.1336 | 0.24145 | 8.4022 | 0.003488 | 0.004777 | 0.004622 | 5.5e-05 | 0.001189 | 0.2688 | 1.3254 | 2.5479 | 98.91 | 20.85 | 38.688 | 67.6192 | 105.8006 | 64.9897 | 0.009649 | 1.0441 | 0.024887 | 23.836 | -0.002149 | 0.14364 | -1.7e-05 | -0.001049 | -0.001049 | -5.6302
2019-03-21 15:49 | 1553197722.8661 | -68 | 0.20946 | 0.000454 | -4e-06 | 100.5327 | 100.7314 | 124.1236 | 21.2612 | 0.001881 | 0.001754 | -152.1362 | 0.24112 | 8.401 | 0.003493 | 0.004826 | 0.004674 | 6e-05 | 0.001193 | 0.2688 | 1.3138 | 2.5479 | 98.91 | 20.85 | 38.688 | 67.0794 | 105.9192 | 64.0231 | 0.009649 | 1.0471 | 0.024937 | 23.8158 | -0.017662 | 0.14344 | 6.5e-05 | 0.00508 | 0.00508 | 1.1606
2019-03-21 15:49 | 1553197723.8657 | -67 | 0.20947 | 0.000458 | -4e-06 | 100.516 | 100.5216 | 123.8694 | 21.2612 | 0.001873 | 0.00181 | -152.1294 | 0.24127 | 8.402 | 0.003468 | 0.004955 | 0.00474 | 6e-05 | 0.001192 | 0.25802 | 1.3523 | 2.548 | 98.91 | 20.85 | 38.687 | 66.8937 | 105.3253 | 65.3799 | 0.009649 | 1.0429 | 0.024861 | 23.8392 | -0.000497 | 0.14353 | -1.9e-05 | -0.001202 | -0.001202 | -4.9079
2019-03-21 15:49 | 1553197724.8658 | -66 | 0.20947 | 0.000458 | -4e-06 | 100.5514 | 100.4524 | 123.6105 | 21.2612 | 0.001915 | 0.00172 | -152.1285 | 0.24196 | 8.4 | 0.003518 | 0.005229 | 0.005094 | 6e-05 | 0.001192 | 0.26377 | 1.3305 | 2.5479 | 98.91 | 20.85 | 38.687 | 67.8978 | 106.5286 | 63.2055 | 0.009649 | 1.0496 | 0.024916 | 23.7377 | 0.008811 | 0.14394 | -7e-06 | -0.000314 | -0.000314 | -18.8564
2019-03-21 15:49 | 1553197725.8662 | -65 | 0.20951 | 0.000455 | -4e-06 | 100.488 | 100.5982 | 123.429 | 21.2612 | 0.001879 | 0.001785 | -152.1305 | 0.24123 | 8.4048 | 0.003597 | 0.004751 | 0.0046 | 5.9e-05 | 0.001193 | 0.27704 | 1.3578 | 2.548 | 98.91 | 20.85 | 38.687 | 67.0279 | 108.4509 | 64.789 | 0.009649 | 1.0447 | 0.024839 | 23.776 | -0.009806 | 0.14351 | -0.000214 | -0.015733 | -0.015733 | -0.3749
2019-03-21 15:49 | 1553197726.8659 | -64 | 0.20952 | 0.000457 | -4e-06 | 100.5219 | 100.5311 | 123.5074 | 21.2612 | 0.001855 | 0.001788 | -152.1276 | 0.24064 | 8.403 | 0.003563 | 0.004365 | 0.004176 | 6e-05 | 0.001189 | 0.28313 | 1.3415 | 2.548 | 98.91 | 20.85 | 38.688 | 66.4498 | 107.6088 | 64.8513 | 0.009649 | 1.0445 | 0.024844 | 23.7858 | -0.000816 | 0.14316 | -0.000301 | -0.022211 | -0.022211 | -0.26491
2019-03-21 15:49 | 1553197727.865 | -63 | 0.20952 | 0.000456 | -4e-06 | 100.5368 | 100.4831 | 123.7666 | 21.2612 | 0.001901 | 0.001787 | -152.1392 | 0.24066 | 8.4039 | 0.003487 | 0.005498 | 0.00536 | 6.1e-05 | 0.001185 | 0.27472 | 1.3138 | 2.548 | 98.91 | 20.85 | 38.688 | 67.5639 | 105.7856 | 64.8266 | 0.009649 | 1.0446 | 0.024871 | 23.8098 | 0.004784 | 0.14317 | -0.00028 | -0.020661 | -0.020661 | -0.28481
2019-03-21 15:49 | 1553197728.865 | -62 | 0.20944 | 0.00046 | -4e-06 | 100.4528 | 100.4756 | 123.7095 | 21.2612 | 0.001867 | 0.00176 | -152.1406 | 0.24193 | 8.4021 | 0.003502 | 0.005033 | 0.004839 | 6.3e-05 | 0.001187 | 0.28263 | 1.319 | 2.548 | 98.91 | 20.85 | 38.688 | 66.7557 | 106.138 | 64.1699 | 0.009649 | 1.0466 | 0.02489 | 23.7812 | -0.002028 | 0.14392 | 0.000184 | 0.013961 | 0.013961 | 0.4237
2019-03-21 15:49 | 1553197729.8664 | -61 | 0.20947 | 0.000458 | -4e-06 | 100.4999 | 100.5239 | 123.8483 | 21.2612 | 0.001908 | 0.001766 | -152.1658 | 0.24233 | 8.4016 | 0.003496 | 0.00496 | 0.004762 | 6.1e-05 | 0.001187 | 0.27936 | 1.3226 | 2.5478 | 98.91 | 20.85 | 38.687 | 67.7355 | 106.0036 | 64.3088 | 0.009649 | 1.0462 | 0.024899 | 23.7994 | -0.002141 | 0.14416 | -1.2e-05 | -0.000634 | -0.000634 | -9.3449
2019-03-21 15:49 | 1553197730.8651 | -60 | 0.20947 | 0.000462 | -4e-06 | 100.4954 | 100.8145 | 124.0079 | 21.2612 | 0.001957 | 0.001805 | -152.1699 | 0.24174 | 8.4016 | 0.00361 | 0.004795 | 0.004664 | 6.3e-05 | 0.001184 | 0.25465 | 1.3624 | 2.548 | 98.91 | 20.85 | 38.689 | 68.9234 | 108.7662 | 65.2512 | 0.00965 | 1.0433 | 0.02488 | 23.848 | -0.028352 | 0.14381 | 3.2e-05 | 0.002661 | 0.002661 | 2.221
2019-03-21 15:49 | 1553197731.8661 | -59 | 0.20948 | 0.00046 | -4e-06 | 100.4156 | 100.4756 | 124.1349 | 21.2612 | 0.001902 | 0.001802 | -152.1667 | 0.24228 | 8.4016 | 0.003655 | 0.005144 | 0.004922 | 5.8e-05 | 0.001186 | 0.29518 | 1.3481 | 2.5481 | 98.91 | 20.85 | 38.692 | 67.5802 | 109.8596 | 65.1963 | 0.00965 | 1.0434 | 0.024895 | 23.8583 | -0.005344 | 0.14413 | -4e-05 | -0.002741 | -0.002741 | -2.1611
2019-03-21 15:49 | 1553197732.8666 | -58 | 0.20946 | 0.000463 | -3e-06 | 100.5261 | 100.6081 | 123.9802 | 21.2612 | 0.00194 | 0.001793 | -152.156 | 0.24096 | 8.4004 | 0.00354 | 0.004873 | 0.004642 | 6.4e-05 | 0.001186 | 0.27887 | 1.3647 | 2.5481 | 98.91 | 20.85 | 38.695 | 68.5154 | 107.0646 | 64.9772 | 0.009651 | 1.0441 | 0.024887 | 23.8357 | -0.00729 | 0.14334 | 5.2e-05 | 0.004079 | 0.004079 | 1.4445
2019-03-21 15:49 | 1553197733.8667 | -57 | 0.20947 | 0.000458 | -4e-06 | 100.496 | 100.4282 | 123.8546 | 21.2612 | 0.001924 | 0.001794 | -152.1691 | 0.24117 | 8.4012 | 0.003571 | 0.005263 | 0.005063 | 5.7e-05 | 0.001186 | 0.27848 | 1.3116 | 2.5479 | 98.91 | 20.85 | 38.702 | 68.1231 | 107.8049 | 64.9974 | 0.009653 | 1.0441 | 0.024874 | 23.8243 | 0.006034 | 0.14347 | -1.3e-05 | -0.000762 | -0.000762 | -7.7351
2019-03-21 15:49 | 1553197734.8656 | -56 | 0.20945 | 0.000465 | -4e-06 | 100.3998 | 100.3986 | 123.9727 | 21.2612 | 0.001942 | 0.001831 | -152.1835 | 0.24157 | 8.4042 | 0.003463 | 0.004729 | 0.004572 | 6.1e-05 | 0.001184 | 0.25658 | 1.3042 | 2.5478 | 98.91 | 20.85 | 38.705 | 68.556 | 105.182 | 65.8793 | 0.009654 | 1.0413 | 0.024853 | 23.8667 | 0.000104 | 0.14371 | 9.8e-05 | 0.007545 | 0.007545 | 0.78278
2019-03-21 15:49 | 1553197735.8658 | -55 | 0.20947 | 0.000464 | -4e-06 | 100.3867 | 100.3714 | 124.0203 | 21.2612 | 0.001878 | 0.001763 | -152.1758 | 0.24092 | 8.4024 | 0.003526 | 0.004164 | 0.004025 | 6e-05 | 0.001188 | 0.28584 | 1.2981 | 2.5479 | 98.9 | 20.85 | 38.71 | 67.0154 | 106.7103 | 64.2503 | 0.009656 | 1.0464 | 0.024918 | 23.8139 | 0.001364 | 0.14333 | 3e-05 | 0.002489 | 0.002489 | 2.367
2019-03-21 15:49 | 1553197736.8664 | -54 | 0.20944 | 0.00046 | -3e-06 | 100.375 | 100.4545 | 124.0109 | 21.2612 | 0.001909 | 0.001799 | -152.1778 | 0.24059 | 8.4016 | 0.00356 | 0.004941 | 0.004697 | 5.7e-05 | 0.001191 | 0.27865 | 1.3297 | 2.548 | 98.9 | 20.85 | 38.718 | 67.7627 | 107.5386 | 65.1059 | 0.009658 | 1.0437 | 0.024886 | 23.8432 | -0.007075 | 0.14313 | 0.000179 | 0.013571 | 0.013571 | 0.43345
2019-03-21 15:49 | 1553197737.8665 | -53 | 0.20946 | 0.000464 | -4e-06 | 100.3363 | 100.2508 | 123.6309 | 21.2612 | 0.001941 | 0.001798 | -152.1884 | 0.24175 | 8.4058 | 0.003568 | 0.004674 | 0.004475 | 6e-05 | 0.001191 | 0.29856 | 1.3511 | 2.5479 | 98.9 | 20.85 | 38.721 | 68.5296 | 107.7443 | 65.0847 | 0.009658 | 1.0438 | 0.024848 | 23.8059 | 0.007627 | 0.14382 | 8.1e-05 | 0.006305 | 0.006305 | 0.93744
2019-03-21 15:49 | 1553197738.8655 | -52 | 0.20946 | 0.000461 | -3e-06 | 100.3327 | 100.4354 | 123.4703 | 21.2612 | 0.001917 | 0.00178 | -152.1874 | 0.24225 | 8.4035 | 0.003562 | 0.005098 | 0.004982 | 6.2e-05 | 0.001191 | 0.24475 | 1.3258 | 2.5479 | 98.9 | 20.85 | 38.725 | 67.9531 | 107.6049 | 64.6625 | 0.009659 | 1.0451 | 0.024848 | 23.7755 | -0.009147 | 0.14411 | 9.2e-05 | 0.007091 | 0.007091 | 0.83535
2019-03-21 15:49 | 1553197739.865 | -51 | 0.20949 | 0.000457 | -4e-06 | 100.3986 | 100.3101 | 123.2891 | 21.2612 | 0.00198 | 0.001734 | -152.1857 | 0.24124 | 8.3994 | 0.003528 | 0.0046 | 0.004467 | 6e-05 | 0.001191 | 0.28136 | 1.3278 | 2.548 | 98.9 | 20.85 | 38.728 | 69.4608 | 106.7811 | 63.5451 | 0.00966 | 1.0486 | 0.024871 | 23.7188 | 0.00789 | 0.14351 | -0.000116 | -0.008409 | -0.008409 | -0.70139
2019-03-21 15:49 | 1553197740.866 | -50 | 0.20943 | 0.000458 | -4e-06 | 100.3572 | 100.1936 | 123.5664 | 21.2612 | 0.001931 | 0.001801 | -152.1895 | 0.24052 | 8.404 | 0.003559 | 0.0053 | 0.005105 | 6.1e-05 | 0.001191 | 0.25481 | 1.3514 | 2.5479 | 98.9 | 20.85 | 38.729 | 68.2808 | 107.5207 | 65.1511 | 0.00966 | 1.0436 | 0.024839 | 23.802 | 0.014592 | 0.14308 | 0.000258 | 0.019486 | 0.019486 | 0.3018
2019-03-21 15:49 | 1553197741.8656 | -49 | 0.20946 | 0.000459 | -4e-06 | 100.3155 | 100.3257 | 123.6919 | 21.2612 | 0.001924 | 0.001797 | -152.1834 | 0.24039 | 8.4023 | 0.003532 | 0.004554 | 0.004416 | 5.9e-05 | 0.001197 | 0.25642 | 1.3283 | 2.5481 | 98.9 | 20.85 | 38.731 | 68.1137 | 106.8734 | 65.0551 | 0.009661 | 1.0439 | 0.024855 | 23.8107 | -0.000916 | 0.14301 | 8.7e-05 | 0.006726 | 0.006726 | 0.87387
2019-03-21 15:49 | 1553197742.8661 | -48 | 0.20946 | 0.000455 | -4e-06 | 100.3384 | 100.0941 | 123.8379 | 21.2612 | 0.00192 | 0.001787 | -152.1983 | 0.24131 | 8.4031 | 0.003492 | 0.004264 | 0.004093 | 6.3e-05 | 0.001196 | 0.23147 | 1.3654 | 2.548 | 98.9 | 20.85 | 38.73 | 68.0195 | 105.9012 | 64.8343 | 0.009661 | 1.0446 | 0.024878 | 23.817 | 0.021802 | 0.14356 | 7.4e-05 | 0.005778 | 0.005778 | 1.0211
2019-03-21 15:49 | 1553197743.8651 | -47 | 0.20948 | 0.000458 | -4e-06 | 100.3667 | 100.2532 | 123.983 | 21.2612 | 0.001832 | 0.00184 | -152.1845 | 0.24006 | 8.4022 | 0.003464 | 0.00541 | 0.00528 | 6.1e-05 | 0.001201 | 0.27184 | 1.3157 | 2.5479 | 98.9 | 20.85 | 38.725 | 65.9098 | 105.2089 | 66.0962 | 0.009659 | 1.0407 | 0.024847 | 23.8754 | 0.010123 | 0.14281 | -3e-05 | -0.002029 | -0.002029 | -2.8935
2019-03-21 15:49 | 1553197744.8658 | -46 | 0.20946 | 0.000455 | -4e-06 | 100.331 | 100.4531 | 123.9309 | 21.2612 | 0.001933 | 0.001776 | -152.1784 | 0.2407 | 8.4007 | 0.003469 | 0.004992 | 0.00479 | 6.1e-05 | 0.001202 | 0.28313 | 1.2972 | 2.5481 | 98.9 | 20.86 | 38.722 | 68.3392 | 105.3372 | 64.5713 | 0.009665 | 1.0454 | 0.024897 | 23.8166 | -0.010884 | 0.14319 | 9.7e-05 | 0.007485 | 0.007485 | 0.78632
2019-03-21 15:49 | 1553197745.8655 | -45 | 0.20947 | 0.000458 | -4e-06 | 100.3321 | 100.3468 | 124.0478 | 21.2612 | 0.001901 | 0.001832 | -152.1702 | 0.24176 | 8.4006 | 0.00348 | 0.004525 | 0.004433 | 6e-05 | 0.001202 | 0.28119 | 1.3045 | 2.548 | 98.9 | 20.86 | 38.718 | 67.5646 | 105.6022 | 65.9168 | 0.009664 | 1.0412 | 0.02486 | 23.8753 | -0.00131 | 0.14382 | 3e-06 | 0.000509 | 0.000509 | 11.6103
2019-03-21 15:49 | 1553197746.866 | -44 | 0.20946 | 0.000453 | -4e-06 | 100.3902 | 100.3172 | 123.8921 | 21.2612 | 0.001899 | 0.00171 | -152.1729 | 0.23988 | 8.4023 | 0.003485 | 0.004736 | 0.004497 | 6.2e-05 | 0.001203 | 0.24137 | 1.3363 | 2.5479 | 98.9 | 20.86 | 38.712 | 67.5078 | 105.7368 | 62.964 | 0.009662 | 1.0504 | 0.024953 | 23.7562 | 0.006509 | 0.1427 | 0.000108 | 0.008322 | 0.008322 | 0.70473
2019-03-21 15:49 | 1553197747.866 | -43 | 0.20946 | 0.000457 | -4e-06 | 100.3753 | 100.3281 | 124.1346 | 21.2612 | 0.001967 | 0.001809 | -152.1626 | 0.24085 | 8.4009 | 0.003508 | 0.004889 | 0.004743 | 5.6e-05 | 0.0012 | 0.2349 | 1.3718 | 2.548 | 98.9 | 20.86 | 38.709 | 69.1659 | 106.2813 | 65.3455 | 0.009661 | 1.043 | 0.024889 | 23.8635 | 0.00421 | 0.14328 | 9.1e-05 | 0.007077 | 0.007077 | 0.83207
2019-03-21 15:49 | 1553197748.8651 | -42 | 0.20946 | 0.000455 | -4e-06 | 100.3816 | 100.4596 | 124.2766 | 21.2612 | 0.001898 | 0.001751 | -152.1735 | 0.24079 | 8.4006 | 0.003509 | 0.004674 | 0.004501 | 5.8e-05 | 0.001205 | 0.28136 | 1.3562 | 2.5481 | 98.9 | 20.86 | 38.708 | 67.4874 | 106.3049 | 63.9598 | 0.009661 | 1.0473 | 0.024955 | 23.8282 | -0.006945 | 0.14324 | 6.7e-05 | 0.005251 | 0.005251 | 1.1211
2019-03-21 15:49 | 1553197749.8651 | -41 | 0.20944 | 0.000454 | -4e-06 | 100.3819 | 100.2184 | 124.4196 | 21.2612 | 0.001887 | 0.001795 | -152.1591 | 0.24088 | 8.403 | 0.003495 | 0.005077 | 0.004891 | 5.6e-05 | 0.001203 | 0.25874 | 1.3625 | 2.5479 | 98.9 | 20.86 | 38.707 | 67.2215 | 105.9672 | 65.0186 | 0.009661 | 1.044 | 0.02493 | 23.8793 | 0.014577 | 0.1433 | 0.000212 | 0.016091 | 0.016091 | 0.36602
2019-03-21 15:49 | 1553197750.8665 | -40 | 0.20945 | 0.000452 | -4e-06 | 100.3533 | 100.3458 | 124.178 | 21.2612 | 0.001862 | 0.001807 | -152.1442 | 0.24132 | 8.4012 | 0.003562 | 0.004644 | 0.004462 | 5.6e-05 | 0.001204 | 0.29286 | 1.3544 | 2.5478 | 98.9 | 20.86 | 38.708 | 66.627 | 107.6049 | 65.3135 | 0.009661 | 1.0431 | 0.024895 | 23.8665 | 0.000666 | 0.14356 | 0.000157 | 0.011976 | 0.011976 | 0.49269
2019-03-21 15:49 | 1553197751.8652 | -39 | 0.20948 | 0.000453 | -4e-06 | 100.2726 | 100.3349 | 124.2172 | 21.2612 | 0.001937 | 0.001759 | -152.1296 | 0.24006 | 8.4032 | 0.003453 | 0.005168 | 0.004963 | 5.5e-05 | 0.001205 | 0.26842 | 1.3466 | 2.548 | 98.9 | 20.86 | 38.709 | 68.4445 | 104.9493 | 64.1466 | 0.009661 | 1.0467 | 0.024942 | 23.8291 | -0.005562 | 0.14281 | -6.1e-05 | -0.004327 | -0.004327 | -1.3564
2019-03-21 15:49 | 1553197752.8663 | -38 | 0.2095 | 0.000457 | -4e-06 | 100.3128 | 100.2603 | 124.0708 | 21.2612 | 0.001904 | 0.001749 | -152.1209 | 0.24069 | 8.4014 | 0.0035 | 0.004584 | 0.004448 | 5.6e-05 | 0.001204 | 0.26377 | 1.2969 | 2.5478 | 98.9 | 20.86 | 38.711 | 67.6483 | 106.0948 | 63.9177 | 0.009662 | 1.0474 | 0.024935 | 23.807 | 0.004677 | 0.14319 | -0.000148 | -0.010823 | -0.010823 | -0.54377
2019-03-21 15:49 | 1553197753.8658 | -37 | 0.20946 | 0.00046 | -4e-06 | 100.2943 | 100.3986 | 124.3175 | 21.2612 | 0.001853 | 0.001795 | -152.1381 | 0.24077 | 8.4033 | 0.003472 | 0.00491 | 0.004733 | 5.4e-05 | 0.001203 | 0.25658 | 1.2996 | 2.548 | 98.9 | 20.86 | 38.712 | 66.4248 | 105.4063 | 65.02 | 0.009662 | 1.044 | 0.02492 | 23.8696 | -0.009298 | 0.14323 | 4.4e-05 | 0.003512 | 0.003512 | 1.6761
2019-03-21 15:49 | 1553197754.8644 | -36 | 0.20945 | 0.000457 | -4e-06 | 100.2237 | 100.1421 | 124.479 | 21.2612 | 0.001918 | 0.001738 | -152.1328 | 0.24107 | 8.4034 | 0.003498 | 0.004774 | 0.004636 | 5.6e-05 | 0.001204 | 0.26913 | 1.2637 | 2.5479 | 98.9 | 20.87 | 38.715 | 67.9673 | 106.0334 | 63.6412 | 0.009669 | 1.0483 | 0.024987 | 23.8363 | 0.00728 | 0.14341 | 0.000119 | 0.009139 | 0.009139 | 0.64496
2019-03-21 15:49 | 1553197755.8654 | -35 | 0.20948 | 0.000456 | -4e-06 | 100.1951 | 100.3475 | 124.6625 | 21.2612 | 0.00193 | 0.001818 | -152.1525 | 0.24065 | 8.4028 | 0.003481 | 0.004613 | 0.004439 | 5.7e-05 | 0.001205 | 0.32068 | 1.3174 | 2.548 | 98.9 | 20.87 | 38.715 | 68.2582 | 105.6236 | 65.5726 | 0.009669 | 1.0423 | 0.024934 | 23.9222 | -0.013598 | 0.14316 | -2.7e-05 | -0.001812 | -0.001812 | -3.2471
2019-03-21 15:49 | 1553197756.8649 | -34 | 0.20948 | 0.000457 | -4e-06 | 100.2225 | 100.1299 | 124.9327 | 21.2612 | 0.001873 | 0.001791 | -152.1415 | 0.24192 | 8.4026 | 0.003516 | 0.00472 | 0.004549 | 5.5e-05 | 0.001205 | 0.26272 | 1.3198 | 2.548 | 98.9 | 20.87 | 38.716 | 66.8923 | 106.4844 | 64.9122 | 0.009669 | 1.0443 | 0.024985 | 23.9248 | 0.008269 | 0.14392 | -4.1e-05 | -0.002859 | -0.002859 | -2.0687
2019-03-21 15:49 | 1553197757.8656 | -33 | 0.20946 | 0.000459 | -3e-06 | 100.2058 | 100.0982 | 124.8076 | 21.2612 | 0.001895 | 0.001816 | -152.1301 | 0.2418 | 8.4014 | 0.003528 | 0.004938 | 0.004868 | 5.6e-05 | 0.001202 | 0.26233 | 1.3088 | 2.5481 | 98.9 | 20.87 | 38.715 | 67.4242 | 106.7717 | 65.5141 | 0.009669 | 1.0425 | 0.02495 | 23.9341 | 0.00961 | 0.14385 | 7e-05 | 0.005459 | 0.005459 | 1.0831
2019-03-21 15:49 | 1553197758.8656 | -32 | 0.20946 | 0.000453 | -4e-06 | 100.1999 | 100.2872 | 124.8097 | 21.2612 | 0.001922 | 0.001792 | -152.1324 | 0.24096 | 8.4014 | 0.003549 | 0.00427 | 0.00392 | 5.7e-05 | 0.001207 | 0.24945 | 1.2901 | 2.548 | 98.9 | 20.87 | 38.715 | 68.0615 | 107.2898 | 64.9428 | 0.009669 | 1.0442 | 0.024972 | 23.9141 | -0.007797 | 0.14335 | 5.4e-05 | 0.004268 | 0.004268 | 1.3803
2019-03-21 15:49 | 1553197759.8649 | -31 | 0.20946 | 0.000459 | -4e-06 | 100.227 | 100.0686 | 124.8348 | 21.2612 | 0.001963 | 0.001777 | -152.1344 | 0.2409 | 8.4053 | 0.003563 | 0.004676 | 0.004558 | 5.7e-05 | 0.001208 | 0.26969 | 1.3188 | 2.5481 | 98.9 | 20.87 | 38.712 | 69.0654 | 107.6309 | 64.5915 | 0.009668 | 1.0453 | 0.024987 | 23.904 | 0.014146 | 0.14331 | 5.6e-05 | 0.004417 | 0.004417 | 1.3336
2019-03-21 15:49 | 1553197760.8661 | -30 | 0.20947 | 0.000457 | -4e-06 | 100.2743 | 100.2702 | 124.7252 | 21.2612 | 0.001915 | 0.001779 | -152.1428 | 0.24188 | 8.4014 | 0.003552 | 0.004449 | 0.004292 | 5.4e-05 | 0.00121 | 0.26089 | 1.3505 | 2.548 | 98.9 | 20.87 | 38.709 | 67.9149 | 107.3473 | 64.6399 | 0.009667 | 1.0452 | 0.024974 | 23.8953 | 0.000368 | 0.14389 | -7e-06 | -0.000242 | -0.000242 | -24.4849
2019-03-21 15:49 | 1553197761.8658 | -29 | 0.2095 | 0.00046 | -4e-06 | 100.1561 | 100.0635 | 124.6376 | 21.2612 | 0.001986 | 0.001753 | -152.1627 | 0.24169 | 8.4035 | 0.003614 | 0.004548 | 0.004383 | 5.4e-05 | 0.001211 | 0.24563 | 1.3046 | 2.548 | 98.9 | 20.87 | 38.705 | 69.6168 | 108.8466 | 64.0154 | 0.009666 | 1.0471 | 0.024989 | 23.8648 | 0.008272 | 0.14378 | -0.000165 | -0.012148 | -0.012148 | -0.48645
2019-03-21 15:49 | 1553197762.8662 | -28 | 0.20949 | 0.000448 | -4e-06 | 100.2317 | 100.1909 | 124.6233 | 21.2612 | 0.001887 | 0.00179 | -152.1614 | 0.24134 | 8.4051 | 0.003642 | 0.004533 | 0.0044 | 5.5e-05 | 0.001211 | 0.28854 | 1.299 | 2.548 | 98.9 | 20.87 | 38.697 | 67.2361 | 109.5466 | 64.8944 | 0.009664 | 1.0444 | 0.024955 | 23.8945 | 0.003649 | 0.14357 | -6.9e-05 | -0.004953 | -0.004953 | -1.1913
2019-03-21 15:49 | 1553197763.8653 | -27 | 0.20948 | 0.000458 | -4e-06 | 100.2365 | 100.3489 | 124.16 | 21.2612 | 0.00191 | 0.001785 | -152.1364 | 0.24166 | 8.4035 | 0.003651 | 0.004263 | 0.004088 | 5.8e-05 | 0.001213 | 0.26626 | 1.3472 | 2.548 | 98.9 | 20.87 | 38.693 | 67.7947 | 109.7549 | 64.7821 | 0.009663 | 1.0447 | 0.024913 | 23.8461 | -0.010023 | 0.14376 | -7.8e-05 | -0.005628 | -0.005628 | -1.0498
2019-03-21 15:49 | 1553197764.8659 | -26 | 0.20947 | 0.000457 | -4e-06 | 100.2353 | 100.2208 | 124.209 | 21.2612 | 0.00193 | 0.001746 | -152.1307 | 0.24121 | 8.4021 | 0.003603 | 0.005426 | 0.005198 | 6.2e-05 | 0.001213 | 0.26449 | 1.3893 | 2.548 | 98.9 | 20.87 | 38.686 | 68.2729 | 108.5896 | 63.8506 | 0.009662 | 1.0476 | 0.024952 | 23.8179 | 0.001293 | 0.1435 | 2.1e-05 | 0.001774 | 0.001774 | 3.3253
2019-03-21 15:49 | 1553197765.8654 | -25 | 0.20948 | 0.000454 | -4e-06 | 100.2422 | 100.2651 | 124.128 | 21.2612 | 0.001909 | 0.001755 | -152.141 | 0.24121 | 8.4028 | 0.003672 | 0.005056 | 0.004909 | 5.7e-05 | 0.001215 | 0.26504 | 1.3541 | 2.5477 | 98.9 | 20.87 | 38.683 | 67.762 | 110.2627 | 64.0597 | 0.009661 | 1.047 | 0.024936 | 23.8175 | -0.002047 | 0.1435 | -6e-05 | -0.004268 | -0.004268 | -1.3819
2019-03-21 15:49 | 1553197766.8651 | -24 | 0.20946 | 0.000459 | -4e-06 | 100.2031 | 100.4112 | 123.9546 | 21.2612 | 0.001915 | 0.001811 | -152.1424 | 0.24109 | 8.4008 | 0.003708 | 0.005469 | 0.00529 | 5.3e-05 | 0.001213 | 0.2573 | 1.3808 | 2.548 | 98.9 | 20.87 | 38.679 | 67.9125 | 111.1425 | 65.4001 | 0.00966 | 1.0428 | 0.024869 | 23.8481 | -0.018555 | 0.14342 | 7.4e-05 | 0.00578 | 0.00578 | 1.0199
2019-03-21 15:49 | 1553197767.865 | -23 | 0.20946 | 0.000455 | -3e-06 | 100.2836 | 100.4855 | 123.856 | 21.2612 | 0.00186 | 0.001757 | -152.1555 | 0.24067 | 8.4033 | 0.003732 | 0.005232 | 0.005087 | 5.5e-05 | 0.001218 | 0.26576 | 1.3348 | 2.5481 | 98.9 | 20.87 | 38.671 | 66.584 | 111.7368 | 64.1042 | 0.009658 | 1.0468 | 0.024907 | 23.793 | -0.017988 | 0.14317 | 0.000104 | 0.007992 | 0.007992 | 0.73631
2019-03-21 15:49 | 1553197768.8663 | -22 | 0.20944 | 0.000456 | -4e-06 | 100.187 | 100.1803 | 123.4725 | 21.2612 | 0.001906 | 0.001754 | -152.143 | 0.24129 | 8.4026 | 0.003643 | 0.005055 | 0.004868 | 5.4e-05 | 0.001223 | 0.26753 | 1.3184 | 2.5481 | 98.9 | 20.87 | 38.667 | 67.6918 | 109.5546 | 64.0363 | 0.009657 | 1.047 | 0.024871 | 23.7537 | 0.000603 | 0.14354 | 0.0002 | 0.015163 | 0.015163 | 0.38907
2019-03-21 15:49 | 1553197769.8661 | -21 | 0.20944 | 0.000459 | -4e-06 | 100.1283 | 100.024 | 123.2143 | 21.2612 | 0.001901 | 0.001795 | -152.1474 | 0.24052 | 8.4037 | 0.003689 | 0.004945 | 0.004845 | 5.1e-05 | 0.001218 | 0.26969 | 1.3083 | 2.548 | 98.9 | 20.87 | 38.66 | 67.56 | 110.6855 | 65.0193 | 0.009655 | 1.044 | 0.024809 | 23.7634 | 0.00933 | 0.14308 | 0.000157 | 0.011934 | 0.011934 | 0.49276
2019-03-21 15:50 | 1553197770.8646 | -20 | 0.20947 | 0.000458 | -4e-06 | 100.1626 | 100.0975 | 122.7838 | 21.2612 | 0.001874 | 0.001801 | -152.1496 | 0.2409 | 8.4005 | 0.003646 | 0.004668 | 0.004404 | 5.2e-05 | 0.001217 | 0.26128 | 1.2807 | 2.548 | 98.9 | 20.87 | 38.656 | 66.9226 | 109.6467 | 65.1581 | 0.009654 | 1.0436 | 0.02476 | 23.7267 | 0.005814 | 0.14331 | -1.5e-05 | -0.000891 | -0.000891 | -6.6086
2019-03-21 15:50 | 1553197771.8646 | -19 | 0.20947 | 0.000458 | -3e-06 | 100.2618 | 100.2791 | 122.6371 | 21.2612 | 0.001946 | 0.001797 | -152.1272 | 0.24105 | 8.4017 | 0.003534 | 0.00451 | 0.00432 | 5.3e-05 | 0.001216 | 0.26178 | 1.3085 | 2.5479 | 98.9 | 20.87 | 38.651 | 68.6589 | 106.9183 | 65.0558 | 0.009653 | 1.0439 | 0.024749 | 23.709 | -0.001538 | 0.1434 | -1e-06 | 0.000139 | 0.000139 | 42.3543
2019-03-21 15:50 | 1553197772.8643 | -18 | 0.20946 | 0.000459 | -4e-06 | 100.1954 | 100.1251 | 122.6317 | 21.2612 | 0.001924 | 0.001772 | -152.1265 | 0.24086 | 8.4018 | 0.003589 | 0.004723 | 0.004563 | 5.3e-05 | 0.001216 | 0.24945 | 1.3275 | 2.5479 | 98.9 | 20.87 | 38.643 | 68.1293 | 108.258 | 64.4604 | 0.009651 | 1.0457 | 0.024771 | 23.6876 | 0.006275 | 0.14329 | 7.7e-05 | 0.005967 | 0.005967 | 0.98698
2019-03-21 15:50 | 1553197773.8644 | -17 | 0.20949 | 0.000454 | -4e-06 | 100.117 | 100.2327 | 123.0725 | 21.2612 | 0.001891 | 0.001702 | -152.133 | 0.2413 | 8.4019 | 0.003492 | 0.004597 | 0.004408 | 5e-05 | 0.001215 | 0.24403 | 1.3056 | 2.5479 | 98.9 | 20.87 | 38.641 | 67.3376 | 105.8964 | 62.7844 | 0.00965 | 1.0509 | 0.024877 | 23.6711 | -0.010332 | 0.14355 | -8.3e-05 | -0.005965 | -0.005965 | -0.98909
2019-03-21 15:50 | 1553197774.8649 | -16 | 0.20944 | 0.00046 | -4e-06 | 100.0461 | 100.0481 | 123.1685 | 21.2612 | 0.001876 | 0.001769 | -152.1384 | 0.24092 | 8.404 | 0.003578 | 0.00502 | 0.004815 | 5.2e-05 | 0.001219 | 0.27472 | 1.3277 | 2.5479 | 98.9 | 20.87 | 38.633 | 66.9687 | 107.9861 | 64.3821 | 0.009648 | 1.046 | 0.024828 | 23.7366 | -0.000182 | 0.14333 | 0.000186 | 0.014113 | 0.014113 | 0.41739
2019-03-21 15:50 | 1553197775.8646 | -15 | 0.20947 | 0.000456 | -4e-06 | 100.1126 | 100.0696 | 123.2042 | 21.2612 | 0.001861 | 0.001767 | -152.1227 | 0.24132 | 8.3993 | 0.003574 | 0.005189 | 0.005027 | 5.2e-05 | 0.001215 | 0.28816 | 1.3048 | 2.548 | 98.9 | 20.87 | 38.629 | 66.601 | 107.879 | 64.3424 | 0.009647 | 1.0461 | 0.024833 | 23.7386 | 0.003839 | 0.14356 | 4.3e-05 | 0.003468 | 0.003468 | 1.7016
2019-03-21 15:50 | 1553197776.8645 | -14 | 0.20945 | 0.000454 | -4e-06 | 99.9961 | 99.9184 | 123.118 | 21.2612 | 0.001865 | 0.001746 | -152.128 | 0.24077 | 8.4022 | 0.003617 | 0.004878 | 0.004725 | 4.9e-05 | 0.001213 | 0.27024 | 1.2774 | 2.5479 | 98.9 | 20.87 | 38.626 | 66.7112 | 108.9287 | 63.8388 | 0.009647 | 1.0476 | 0.024842 | 23.7126 | 0.006951 | 0.14323 | 0.000157 | 0.011945 | 0.011945 | 0.49282
2019-03-21 15:50 | 1553197777.8646 | -13 | 0.2095 | 0.000456 | -3e-06 | 99.9964 | 100.0587 | 123.3276 | 21.2612 | 0.001868 | 0.001765 | -152.121 | 0.24154 | 8.4038 | 0.003671 | 0.004957 | 0.004868 | 5.3e-05 | 0.001212 | 0.27007 | 1.3013 | 2.548 | 98.9 | 20.87 | 38.62 | 66.7696 | 110.243 | 64.2994 | 0.009645 | 1.0462 | 0.024847 | 23.749 | -0.005575 | 0.14369 | -0.000174 | -0.012751 | -0.012751 | -0.46315
2019-03-21 15:50 | 1553197778.8651 | -12 | 0.20946 | 0.000456 | -4e-06 | 99.9931 | 100.1326 | 123.9313 | 21.2612 | 0.0019 | 0.001721 | -152.1406 | 0.24136 | 8.4011 | 0.003747 | 0.005005 | 0.004773 | 5.3e-05 | 0.001209 | 0.24906 | 1.3443 | 2.5479 | 98.9 | 20.87 | 38.617 | 67.553 | 112.0935 | 63.2421 | 0.009644 | 1.0495 | 0.024946 | 23.7698 | -0.012473 | 0.14359 | 6.5e-05 | 0.005091 | 0.005091 | 1.1591
2019-03-21 15:50 | 1553197779.8654 | -11 | 0.20949 | 0.000459 | -3e-06 | 99.9889 | 99.9473 | 124.0694 | 21.2612 | 0.001885 | 0.001752 | -151.9803 | 0.24046 | 8.3998 | 0.003706 | 0.005158 | 0.004973 | 5.3e-05 | 0.001204 | 0.22733 | 1.3221 | 2.5479 | 98.9 | 20.87 | 38.608 | 67.1832 | 111.0841 | 63.9803 | 0.009642 | 1.0472 | 0.024933 | 23.8091 | 0.00372 | 0.14305 | -8.8e-05 | -0.006362 | -0.006362 | -0.92412
2019-03-21 15:50 | 1553197780.8647 | -10 | 0.20947 | 0.000458 | -3e-06 | 99.9406 | 99.8363 | 124.9632 | 21.2612 | 0.001875 | 0.001728 | -121.2971 | 0.24053 | 8.4018 | 0.003731 | 0.005002 | 0.004753 | 5.2e-05 | 0.001203 | 0.26897 | 1.3405 | 2.5479 | 98.9 | 20.87 | 38.603 | 66.9459 | 111.71 | 63.4007 | 0.009641 | 1.049 | 0.025044 | 23.8741 | 0.009344 | 0.14309 | -1.3e-05 | -0.000733 | -0.000733 | -8.0258
2019-03-21 15:50 | 1553197781.8648 | -9 | 0.20949 | 0.000457 | -3e-06 | 99.859 | 99.7842 | 125.1801 | 21.2612 | 0.001855 | 0.001741 | -79.6691 | 0.24113 | 8.3998 | 0.003769 | 0.005011 | 0.004864 | 5.4e-05 | 0.001199 | 0.26449 | 1.318 | 2.548 | 98.9 | 20.87 | 38.599 | 66.4693 | 112.6234 | 63.7288 | 0.00964 | 1.048 | 0.025054 | 23.9065 | 0.006704 | 0.14345 | -8.4e-05 | -0.006131 | -0.006131 | -0.96165
2019-03-21 15:50 | 1553197782.8648 | -8 | 0.20947 | 0.00046 | -3e-06 | 99.8575 | 99.8816 | 125.3726 | 21.2612 | 0.001929 | 0.001741 | -38.0448 | 0.24079 | 8.4038 | 0.003868 | 0.00523 | 0.005124 | 5.2e-05 | 0.001195 | 0.26504 | 1.305 | 2.5479 | 98.9 | 20.87 | 38.592 | 68.244 | 115.0413 | 63.7187 | 0.009638 | 1.048 | 0.025073 | 23.9245 | -0.002157 | 0.14325 | 3.7e-05 | 0.00295 | 0.00295 | 1.9955
2019-03-21 15:50 | 1553197783.8659 | -7 | 0.20947 | 0.00046 | -3e-06 | 99.7437 | 99.7651 | 125.3202 | 21.2612 | 0.00185 | 0.001766 | 3.6041 | 0.24133 | 8.4012 | 0.003858 | 0.004824 | 0.004598 | 5.4e-05 | 0.001195 | 0.26017 | 1.33 | 2.5481 | 98.9 | 20.87 | 38.589 | 66.3413 | 114.7938 | 64.3174 | 0.009637 | 1.0462 | 0.025046 | 23.9408 | -0.00192 | 0.14356 | 3.5e-05 | 0.002825 | 0.002825 | 2.0885
2019-03-21 15:50 | 1553197784.865 | -6 | 0.20945 | 0.000462 | -4e-06 | 99.699 | 99.5737 | 124.8461 | 21.2612 | 0.001913 | 0.001754 | 45.2644 | 0.2417 | 8.4019 | 0.003838 | 0.004738 | 0.004654 | 5e-05 | 0.001194 | 0.24978 | 1.3216 | 2.5479 | 98.9 | 20.87 | 38.587 | 67.8572 | 114.3035 | 64.0262 | 0.009637 | 1.0471 | 0.025009 | 23.8851 | 0.011249 | 0.14378 | 9.9e-05 | 0.007653 | 0.007653 | 0.77219
2019-03-21 15:50 | 1553197785.864 | -5 | 0.20948 | 0.000458 | -4e-06 | 99.629 | 99.5407 | 124.8796 | 21.2612 | 0.00183 | 0.001754 | 154.6324 | 0.26514 | 8.4027 | 0.003851 | 0.005201 | 0.005087 | 4.8e-05 | 0.001188 | 0.26432 | 1.2781 | 2.5478 | 98.9 | 20.88 | 38.583 | 65.8552 | 114.6184 | 64.0231 | 0.009642 | 1.0471 | 0.025013 | 23.8882 | 0.007933 | 0.15773 | -6.6e-05 | -0.004722 | -0.004722 | -1.3728
2019-03-21 15:50 | 1553197786.865 | -4 | 0.20945 | 0.000457 | -3e-06 | 99.6239 | 99.4961 | 125.3331 | 21.2612 | 0.001944 | 0.001749 | 264.3325 | 0.27299 | 8.4023 | 0.003897 | 0.005541 | 0.005467 | 5e-05 | 0.001184 | 0.27129 | 1.2901 | 2.548 | 98.9 | 20.87 | 38.575 | 68.596 | 115.7391 | 63.9052 | 0.009634 | 1.0474 | 0.025062 | 23.9274 | 0.011488 | 0.1624 | 0.00012 | 0.009254 | 0.009254 | 0.72129
2019-03-21 15:50 | 1553197787.8651 | -3 | 0.20945 | 0.000455 | -4e-06 | 99.5101 | 99.6265 | 125.4352 | 21.2612 | 0.001879 | 0.001769 | 264.1614 | 0.27275 | 8.4042 | 0.003815 | 0.00554 | 0.005503 | 5.1e-05 | 0.00118 | 0.29407 | 1.2972 | 2.548 | 98.9 | 20.87 | 38.571 | 67.0428 | 113.7361 | 64.3862 | 0.009633 | 1.0459 | 0.025055 | 23.9542 | -0.010455 | 0.16226 | 0.000133 | 0.010261 | 0.010261 | 0.6499
2019-03-21 15:50 | 1553197788.8646 | -2 | 0.20948 | 0.000458 | -4e-06 | 99.4514 | 99.2593 | 125.1061 | 21.2612 | 0.00189 | 0.00171 | 264.3425 | 0.27398 | 8.406 | 0.003814 | 0.005402 | 0.005372 | 5.3e-05 | 0.00118 | 0.27113 | 1.2863 | 2.5478 | 98.9 | 20.87 | 38.568 | 67.3042 | 113.7336 | 62.9766 | 0.009632 | 1.0503 | 0.025074 | 23.8727 | 0.017292 | 0.16299 | -3.7e-05 | -0.002536 | -0.002536 | -2.6412
2019-03-21 15:50 | 1553197789.8642 | -1 | 0.20946 | 0.000454 | -3e-06 | 99.4064 | 99.1143 | 125.036 | 21.2612 | 0.001914 | 0.00183 | 264.423 | 0.27304 | 8.4014 | 0.003866 | 0.00541 | 0.005394 | 5.4e-05 | 0.001173 | 0.26576 | 1.3181 | 2.5479 | 98.9 | 20.88 | 38.56 | 67.886 | 114.9859 | 65.866 | 0.009636 | 1.0414 | 0.02496 | 23.9684 | 0.026329 | 0.16243 | 0.000107 | 0.008268 | 0.008268 | 0.80741
2019-03-21 15:50 | 1553197790.8653 | 0 | 0.20945 | 0.000462 | -4e-06 | 99.068 | 96.6399 | 124.2968 | 21.2545 | 0.001823 | 0.001723 | 264.6584 | 0.27339 | 8.402 | 0.003742 | 0.005669 | 0.007098 | 5.2e-05 | 0.001176 | 0.24945 | 1.3336 | 2.6963 | 98.9 | 20.87 | 38.556 | 65.6753 | 111.9719 | 63.2811 | 0.009629 | 1.0494 | 0.024982 | 23.8062 | 0.22196 | 0.16264 | 0.000133 | 0.010205 | 0.010205 | 0.65504
2019-03-21 15:50 | 1553197791.8659 | 1 | 0.20946 | 0.000462 | -3e-06 | 96.1705 | 92.8992 | 123.8626 | 21.2523 | 0.001889 | 0.001749 | 264.8594 | 0.27352 | 8.404 | 0.003877 | 0.010103 | 0.014534 | 5.4e-05 | 0.001172 | 0.37144 | 1.3378 | 2.5485 | 98.9 | 20.88 | 38.55 | 67.2798 | 115.2438 | 63.9058 | 0.009634 | 1.0474 | 0.024915 | 23.7866 | 0.30955 | 0.16272 | 7.3e-05 | 0.005648 | 0.005648 | 1.1841
2019-03-21 15:50 | 1553197792.8654 | 2 | 0.20948 | 0.00046 | -3e-06 | 96.3364 | 88.115 | 122.615 | 21.2551 | 0.001869 | 0.001702 | 264.7613 | 0.27343 | 8.4032 | 0.003754 | 0.006216 | 0.007047 | 5.3e-05 | 0.001173 | 0.27184 | 1.3619 | 2.548 | 98.9 | 20.88 | 38.547 | 66.7973 | 112.2683 | 62.7652 | 0.009633 | 1.051 | 0.024831 | 23.6264 | 0.79789 | 0.16266 | -4.5e-05 | -0.003137 | -0.003137 | -2.131
2019-03-21 15:50 | 1553197793.8644 | 3 | 0.20946 | 0.000466 | -3e-06 | 97.4907 | 91.621 | 121.301 | 21.2556 | 0.001868 | 0.001779 | 264.5765 | 0.27371 | 8.4062 | 0.003887 | 0.005646 | 0.005805 | 5.5e-05 | 0.001166 | 0.30215 | 1.3263 | 2.5479 | 98.9 | 20.88 | 38.545 | 66.762 | 115.5033 | 64.6351 | 0.009632 | 1.0452 | 0.024629 | 23.5648 | 0.55542 | 0.16283 | 3.7e-05 | 0.002943 | 0.002943 | 2.2744
2019-03-21 15:50 | 1553197794.8645 | 4 | 0.20947 | 0.00046 | -1e-06 | 98.2596 | 95.0163 | 120.099 | 21.2572 | 0.001833 | 0.001821 | 264.8119 | 0.27328 | 8.4019 | 0.003807 | 0.005515 | 0.005634 | 5e-05 | 0.001177 | 0.34954 | 1.3224 | 2.548 | 98.9 | 20.88 | 38.541 | 65.9247 | 113.5455 | 65.6476 | 0.009631 | 1.0421 | 0.02447 | 23.4829 | 0.30022 | 0.16257 | 7e-06 | 0.0006 | 0.0006 | 11.1361
2019-03-21 15:50 | 1553197795.8655 | 5 | 0.20943 | 0.000464 | 2e-06 | 98.4453 | 96.7172 | 118.6811 | 21.259 | 0.001861 | 0.00184 | 264.5662 | 0.27492 | 8.4036 | 0.003809 | 0.006894 | 0.007472 | 5.7e-05 | 0.001162 | 0.50638 | 1.3142 | 2.548 | 98.9 | 20.88 | 38.538 | 66.595 | 113.5916 | 66.1104 | 0.009631 | 1.0406 | 0.024309 | 23.3598 | 0.1584 | 0.16355 | 0.000216 | 0.015626 | 0.015626 | 0.43017
2019-03-21 15:50 | 1553197796.8646 | 6 | 0.20925 | 0.000462 | 7e-06 | 98.8102 | 97.2727 | 116.462 | 21.2605 | 0.0019 | 0.001917 | 264.5571 | 0.27286 | 8.4021 | 0.003684 | 0.009666 | 0.01173 | 0.000132 | 0.001221 | 0.28744 | 1.2965 | 2.548 | 98.9 | 20.88 | 38.535 | 67.5422 | 110.5571 | 67.9579 | 0.00963 | 1.035 | 0.024015 | 23.2033 | 0.14028 | 0.16233 | 0.001293 | 0.092805 | 0.092805 | 0.071888
2019-03-21 15:50 | 1553197797.8656 | 7 | 0.20901 | 0.000466 | 1.3e-05 | 98.9083 | 97.2942 | 114.1542 | 21.2605 | 0.001875 | 0.001947 | 264.4952 | 0.27447 | 8.4025 | 0.003789 | 0.010269 | 0.012813 | 0.000291 | 0.001165 | 0.25514 | 1.2989 | 2.5479 | 98.9 | 20.87 | 38.531 | 66.932 | 113.108 | 68.67 | 0.009623 | 1.0328 | 0.023751 | 22.9962 | 0.14717 | 0.16328 | 0.002758 | 0.19584 | 0.19584 | 0.034268
2019-03-21 15:50 | 1553197798.8657 | 8 | 0.20806 | 0.000482 | 2.2e-05 | 98.9381 | 95.8133 | 111.6513 | 21.2588 | 0.002042 | 0.002013 | 264.4466 | 0.27516 | 8.4028 | 0.003787 | 0.005155 | 0.005647 | 5.4e-05 | 0.001166 | 0.23258 | 1.317 | 1.7688 | 98.9 | 20.87 | 38.53 | 70.9624 | 113.0722 | 70.2482 | 0.009623 | 1.0281 | 0.023436 | 22.7951 | 0.28706 | 0.1637 | 0.008474 | 0.59433 | 0.59433 | 0.01132
2019-03-21 15:50 | 1553197799.8657 | 9 | 0.20716 | 0.000523 | 3.2e-05 | 98.9255 | 93.562 | 109.5576 | 21.2612 | 0.002063 | 0.002105 | 264.3645 | 0.27379 | 8.404 | 0.003759 | 0.005136 | 0.005229 | 5.3e-05 | 0.001173 | 0.23418 | 1.3059 | 2.5479 | 98.9 | 20.87 | 38.526 | 71.4507 | 112.3803 | 72.4671 | 0.009622 | 1.0215 | 0.02314 | 22.6532 | 0.4986 | 0.16288 | 0.013791 | 0.95476 | 0.95476 | 0.007012
2019-03-21 15:50 | 1553197800.8658 | 10 | 0.20623 | 0.000616 | 4.4e-05 | 98.9369 | 88.5652 | 107.984 | 21.2612 | 0.002184 | 0.002201 | 264.3142 | 0.27375 | 8.4035 | 0.003705 | 0.004834 | 0.004765 | 5.4e-05 | 0.00117 | 0.26001 | 1.3054 | 2.5479 | 98.9 | 20.87 | 38.525 | 74.3582 | 111.0802 | 74.7721 | 0.009621 | 1.0147 | 0.022897 | 22.5648 | 0.99054 | 0.16285 | 0.019241 | 1.3173 | 1.3173 | 0.005081
2019-03-21 15:50 | 1553197801.8638 | 11 | 0.20427 | 0.001022 | 5.6e-05 | 98.9241 | 83.8059 | 106.3618 | 21.2612 | 0.002252 | 0.002264 | 264.2164 | 0.27383 | 8.4023 | 0.003809 | 0.00504 | 0.004951 | 5.1e-05 | 0.001167 | 0.25265 | 1.2754 | 2.5479 | 98.9 | 20.87 | 38.524 | 76.0061 | 113.5914 | 76.2915 | 0.009621 | 1.0103 | 0.022675 | 22.4435 | 1.4834 | 0.1629 | 0.030458 | 2.0635 | 2.0635 | 0.003245
2019-03-21 15:50 | 1553197802.8639 | 12 | 0.20238 | 0.00166 | 6.9e-05 | 98.8657 | 80.8162 | 105.6774 | 21.2612 | 0.002273 | 0.002322 | 264.1387 | 0.2734 | 8.4026 | 0.003838 | 0.005369 | 0.005277 | 5.1e-05 | 0.001165 | 0.26106 | 1.3003 | 2.5479 | 98.9 | 20.87 | 38.523 | 76.5002 | 114.3067 | 77.6953 | 0.009621 | 1.0063 | 0.022557 | 22.4161 | 1.8031 | 0.16265 | 0.040925 | 2.7554 | 2.7554 | 0.002426
2019-03-21 15:50 | 1553197803.8649 | 13 | 0.20124 | 0.002347 | 8.1e-05 | 98.8514 | 78.4181 | 104.9488 | 21.2612 | 0.002386 | 0.002335 | 264.0088 | 0.27388 | 8.404 | 0.003885 | 0.005229 | 0.005213 | 5.3e-05 | 0.001165 | 0.24065 | 1.3216 | 2.548 | 98.9 | 20.87 | 38.523 | 79.2326 | 115.4423 | 78.0055 | 0.009621 | 1.0054 | 0.022469 | 22.3486 | 2.0712 | 0.16293 | 0.046872 | 3.1414 | 3.1414 | 0.002132
2019-03-21 15:50 | 1553197804.8639 | 14 | 0.20057 | 0.003073 | 9.3e-05 | 98.7682 | 75.9189 | 104.8164 | 21.2612 | 0.002461 | 0.002372 | 263.8978 | 0.27364 | 8.402 | 0.003905 | 0.005529 | 0.005436 | 5.2e-05 | 0.00116 | 0.26106 | 1.293 | 2.5479 | 98.9 | 20.87 | 38.523 | 81.0331 | 115.9503 | 78.8842 | 0.009621 | 1.0029 | 0.022427 | 22.3624 | 2.3534 | 0.16279 | 0.049936 | 3.3391 | 3.3391 | 0.002004
2019-03-21 15:50 | 1553197805.8641 | 15 | 0.20033 | 0.003686 | 0.000104 | 98.8147 | 73.297 | 104.5164 | 21.2612 | 0.002505 | 0.002438 | 263.7855 | 0.27295 | 8.4009 | 0.003954 | 0.005094 | 0.005049 | 5.5e-05 | 0.001155 | 0.24923 | 1.334 | 2.548 | 98.9 | 20.87 | 38.523 | 82.0755 | 117.1385 | 80.4707 | 0.009621 | 0.99837 | 0.022344 | 22.3806 | 2.672 | 0.16238 | 0.050596 | 3.3701 | 3.3701 | 0.00198
2019-03-21 15:50 | 1553197806.8646 | 16 | 0.20013 | 0.004209 | 0.000111 | 98.8239 | 71.1878 | 104.9338 | 21.2612 | 0.002511 | 0.002481 | 263.7023 | 0.27296 | 8.4025 | 0.003902 | 0.005908 | 0.005849 | 5.4e-05 | 0.001156 | 0.27903 | 1.3279 | 2.548 | 98.9 | 20.87 | 38.523 | 82.2238 | 115.861 | 81.4955 | 0.009621 | 0.99549 | 0.022356 | 22.4577 | 2.934 | 0.16238 | 0.051154 | 3.4085 | 3.4085 | 0.001958
2019-03-21 15:50 | 1553197807.8655 | 17 | 0.20016 | 0.00461 | 0.000117 | 98.7316 | 70.5893 | 105.9964 | 21.2612 | 0.002579 | 0.002487 | 263.5825 | 0.27401 | 8.4034 | 0.004001 | 0.005647 | 0.005635 | 5.3e-05 | 0.001154 | 0.27151 | 1.3109 | 2.5479 | 98.9 | 20.87 | 38.524 | 83.8667 | 118.2746 | 81.645 | 0.009621 | 0.99507 | 0.022465 | 22.5759 | 3.0011 | 0.16301 | 0.05053 | 3.3831 | 3.3831 | 0.00198
2019-03-21 15:50 | 1553197808.8653 | 18 | 0.20012 | 0.004858 | 0.000118 | 98.6648 | 68.8779 | 107.2679 | 21.2612 | 0.002614 | 0.002551 | 263.4838 | 0.27421 | 8.4005 | 0.004084 | 0.004809 | 0.004763 | 5.6e-05 | 0.001154 | 0.2636 | 1.3136 | 2.548 | 98.9 | 20.86 | 38.524 | 84.692 | 120.3091 | 83.1809 | 0.009615 | 0.99078 | 0.02255 | 22.76 | 3.2146 | 0.16313 | 0.050438 | 3.3897 | 3.3897 | 0.001978
2019-03-21 15:50 | 1553197809.8652 | 19 | 0.2002 | 0.005074 | 0.000118 | 98.6908 | 69.6664 | 108.2342 | 21.2612 | 0.002623 | 0.002554 | 263.3841 | 0.27345 | 8.3991 | 0.004161 | 0.005327 | 0.005265 | 5e-05 | 0.001153 | 0.25033 | 1.2722 | 2.5479 | 98.9 | 20.86 | 38.527 | 84.9162 | 122.1967 | 83.2487 | 0.009616 | 0.99059 | 0.022649 | 22.8645 | 3.1152 | 0.16268 | 0.049726 | 3.3568 | 3.3568 | 0.001992
2019-03-21 15:50 | 1553197810.8643 | 20 | 0.20059 | 0.005272 | 0.000116 | 98.694 | 71.1312 | 108.8372 | 21.2612 | 0.00266 | 0.00253 | 263.2545 | 0.27286 | 8.4024 | 0.004113 | 0.006044 | 0.005926 | 5.4e-05 | 0.001155 | 0.26272 | 1.3212 | 2.5479 | 98.9 | 20.86 | 38.528 | 85.8057 | 121.0282 | 82.6871 | 0.009616 | 0.99216 | 0.02273 | 22.91 | 2.9293 | 0.16233 | 0.047158 | 3.1954 | 3.1954 | 0.002088
2019-03-21 15:50 | 1553197811.8653 | 21 | 0.20086 | 0.00544 | 0.000113 | 98.7459 | 72.3151 | 109.3691 | 21.2612 | 0.002693 | 0.002548 | 263.1564 | 0.2737 | 8.3999 | 0.004137 | 0.006092 | 0.005964 | 5.4e-05 | 0.001153 | 0.24259 | 1.3113 | 2.5479 | 98.9 | 20.86 | 38.53 | 86.5917 | 121.61 | 83.1052 | 0.009617 | 0.99099 | 0.022772 | 22.9794 | 2.7864 | 0.16282 | 0.045361 | 3.0798 | 3.0798 | 0.002173
2019-03-21 15:50 | 1553197812.8644 | 22 | 0.20132 | 0.00557 | 0.000109 | 98.7024 | 73.9915 | 110.0125 | 21.2612 | 0.002682 | 0.002572 | 263.0801 | 0.27396 | 8.4016 | 0.004163 | 0.005993 | 0.006013 | 5.3e-05 | 0.001151 | 0.2453 | 1.2909 | 2.5479 | 98.9 | 20.86 | 38.534 | 86.3273 | 122.2444 | 83.7044 | 0.009618 | 0.98933 | 0.02282 | 23.0663 | 2.5775 | 0.16298 | 0.042436 | 2.888 | 2.888 | 0.002319
2019-03-21 15:50 | 1553197813.8645 | 23 | 0.20158 | 0.005654 | 0.000105 | 98.6994 | 76.0479 | 110.7612 | 21.2612 | 0.002688 | 0.002606 | 262.9937 | 0.27341 | 8.4027 | 0.004176 | 0.005481 | 0.005411 | 5.5e-05 | 0.001152 | 0.2557 | 1.2914 | 2.5477 | 98.9 | 20.86 | 38.536 | 86.4824 | 122.5573 | 84.5188 | 0.009618 | 0.98707 | 0.022872 | 23.1711 | 2.332 | 0.16265 | 0.040846 | 2.7865 | 2.7865 | 0.002399
2019-03-21 15:50 | 1553197814.8657 | 24 | 0.20235 | 0.005696 | 0.0001 | 98.6932 | 77.5169 | 111.1903 | 21.2612 | 0.002768 | 0.002597 | 262.935 | 0.27303 | 8.4013 | 0.004188 | 0.005103 | 0.00504 | 5.3e-05 | 0.001147 | 0.25586 | 1.2982 | 2.5479 | 98.9 | 20.86 | 38.538 | 88.4073 | 122.8592 | 84.3012 | 0.009619 | 0.98767 | 0.022923 | 23.2088 | 2.1603 | 0.16242 | 0.036167 | 2.4736 | 2.4736 | 0.002699
2019-03-21 15:50 | 1553197815.8646 | 25 | 0.20286 | 0.005698 | 9.6e-05 | 98.7027 | 79.0444 | 111.4656 | 21.2612 | 0.002736 | 0.002582 | 262.8674 | 0.27365 | 8.4019 | 0.004155 | 0.005078 | 0.004967 | 5.2e-05 | 0.001142 | 0.25442 | 1.3752 | 2.548 | 98.9 | 20.86 | 38.546 | 87.645 | 122.041 | 83.9227 | 0.009621 | 0.98872 | 0.022963 | 23.2252 | 1.9866 | 0.16279 | 0.033124 | 2.2699 | 2.2699 | 0.002948
2019-03-21 15:50 | 1553197816.8636 | 26 | 0.20315 | 0.005658 | 9.2e-05 | 98.8108 | 79.9487 | 111.6351 | 21.2612 | 0.00275 | 0.002646 | 262.7824 | 0.27363 | 8.4033 | 0.004137 | 0.005753 | 0.005661 | 5.6e-05 | 0.001147 | 0.23667 | 1.3395 | 2.5479 | 98.9 | 20.86 | 38.551 | 87.9839 | 121.6068 | 85.4819 | 0.009622 | 0.98442 | 0.022931 | 23.2936 | 1.8947 | 0.16278 | 0.031425 | 2.1508 | 2.1508 | 0.003111
2019-03-21 15:50 | 1553197817.8647 | 27 | 0.20344 | 0.005595 | 8.8e-05 | 98.7095 | 81.0376 | 111.5551 | 21.2612 | 0.002731 | 0.002625 | 262.7257 | 0.27446 | 8.4011 | 0.004107 | 0.005589 | 0.0056 | 5.4e-05 | 0.001147 | 0.27328 | 1.3463 | 2.5479 | 98.89 | 20.86 | 38.557 | 87.5256 | 120.8652 | 84.9597 | 0.009624 | 0.98586 | 0.022939 | 23.2683 | 1.7645 | 0.16327 | 0.029761 | 2.0379 | 2.0379 | 0.003293
2019-03-21 15:50 | 1553197818.8637 | 28 | 0.20355 | 0.005514 | 8.4e-05 | 98.7504 | 81.3904 | 111.6823 | 21.2612 | 0.002797 | 0.00265 | 262.6419 | 0.274 | 8.406 | 0.004112 | 0.00551 | 0.005424 | 5.4e-05 | 0.001147 | 0.24458 | 1.3305 | 2.5479 | 98.89 | 20.86 | 38.561 | 89.1024 | 120.9978 | 85.5755 | 0.009625 | 0.98417 | 0.022933 | 23.3015 | 1.7293 | 0.163 | 0.029234 | 2.0016 | 2.0016 | 0.003347
2019-03-21 15:50 | 1553197819.8647 | 29 | 0.20362 | 0.005427 | 8.1e-05 | 98.7816 | 81.2671 | 111.8659 | 21.2612 | 0.002843 | 0.002672 | 262.5386 | 0.27326 | 8.4003 | 0.004096 | 0.005332 | 0.005305 | 5.6e-05 | 0.00115 | 0.27256 | 1.3323 | 2.548 | 98.89 | 20.86 | 38.565 | 90.208 | 120.6089 | 86.0915 | 0.009626 | 0.98275 | 0.022935 | 23.3374 | 1.7457 | 0.16256 | 0.028894 | 1.9786 | 1.9786 | 0.003377
2019-03-21 15:50 | 1553197820.8648 | 30 | 0.20378 | 0.005331 | 7.8e-05 | 98.8174 | 81.6193 | 112.2891 | 21.2612 | 0.002878 | 0.002631 | 262.4784 | 0.2739 | 8.4034 | 0.004111 | 0.006373 | 0.006318 | 5.5e-05 | 0.00115 | 0.28279 | 1.3047 | 2.548 | 98.89 | 20.86 | 38.574 | 91.0685 | 120.9756 | 85.1057 | 0.009629 | 0.98546 | 0.02301 | 23.3494 | 1.7103 | 0.16294 | 0.028096 | 1.9305 | 1.9305 | 0.003469
2019-03-21 15:50 | 1553197821.8649 | 31 | 0.2039 | 0.005244 | 7.5e-05 | 98.8749 | 81.6424 | 112.4693 | 21.2612 | 0.002864 | 0.002672 | 262.3763 | 0.27368 | 8.3997 | 0.004097 | 0.005603 | 0.005525 | 5.3e-05 | 0.001145 | 0.27328 | 1.3034 | 2.548 | 98.89 | 20.86 | 38.578 | 90.724 | 120.6414 | 86.0884 | 0.00963 | 0.98276 | 0.022997 | 23.4002 | 1.7129 | 0.16281 | 0.027471 | 1.8867 | 1.8867 | 0.003547
2019-03-21 15:50 | 1553197822.864 | 32 | 0.20405 | 0.005155 | 7.3e-05 | 98.9309 | 82.2538 | 112.5238 | 21.2612 | 0.002819 | 0.002681 | 262.304 | 0.27325 | 8.4028 | 0.00411 | 0.005611 | 0.005585 | 5.1e-05 | 0.001145 | 0.27887 | 1.2912 | 2.5478 | 98.89 | 20.86 | 38.587 | 89.6464 | 120.9597 | 86.3248 | 0.009632 | 0.98211 | 0.022995 | 23.4136 | 1.6513 | 0.16256 | 0.026681 | 1.8324 | 1.8324 | 0.003646
2019-03-21 15:50 | 1553197823.864 | 33 | 0.20417 | 0.005072 | 6.9e-05 | 99.0039 | 82.4823 | 112.7958 | 21.2612 | 0.002886 | 0.00271 | 262.2103 | 0.27275 | 8.4024 | 0.004071 | 0.005764 | 0.005683 | 5.1e-05 | 0.001143 | 0.24259 | 1.3106 | 2.5479 | 98.89 | 20.86 | 38.591 | 91.2391 | 120.0052 | 87.0226 | 0.009633 | 0.98021 | 0.023 | 23.4646 | 1.6331 | 0.16226 | 0.026072 | 1.7912 | 1.7912 | 0.003723
2019-03-21 15:50 | 1553197824.8651 | 34 | 0.20431 | 0.004997 | 6.6e-05 | 99.0286 | 83.7602 | 113.2319 | 21.2612 | 0.002943 | 0.002707 | 262.1297 | 0.27313 | 8.4019 | 0.004101 | 0.006453 | 0.006358 | 5.1e-05 | 0.001138 | 0.27704 | 1.3694 | 2.5479 | 98.89 | 20.86 | 38.595 | 92.623 | 120.7336 | 86.9377 | 0.009634 | 0.98044 | 0.023047 | 23.5071 | 1.4978 | 0.16248 | 0.025291 | 1.7413 | 1.7413 | 0.003835
2019-03-21 15:50 | 1553197825.8641 | 35 | 0.2045 | 0.004917 | 6.2e-05 | 99.0534 | 83.9091 | 114.0056 | 21.2612 | 0.002877 | 0.002674 | 262.0585 | 0.27374 | 8.4019 | 0.004074 | 0.005804 | 0.005729 | 4.8e-05 | 0.001139 | 0.28479 | 1.3318 | 2.548 | 98.89 | 20.86 | 38.601 | 91.0389 | 120.0663 | 86.1555 | 0.009635 | 0.98258 | 0.023151 | 23.5617 | 1.4841 | 0.16285 | 0.024288 | 1.68 | 1.68 | 0.003984
2019-03-21 15:50 | 1553197826.8652 | 36 | 0.20476 | 0.004838 | 5.8e-05 | 99.1246 | 84.9111 | 114.3635 | 21.2612 | 0.002876 | 0.002659 | 261.9984 | 0.27382 | 8.3985 | 0.004042 | 0.005661 | 0.005611 | 4.9e-05 | 0.001134 | 0.28047 | 1.3458 | 2.5479 | 98.89 | 20.86 | 38.603 | 91.0186 | 119.2914 | 85.7941 | 0.009636 | 0.98357 | 0.023199 | 23.5868 | 1.3844 | 0.1629 | 0.022822 | 1.5821 | 1.5821 | 0.004232
2019-03-21 15:50 | 1553197827.8642 | 37 | 0.20496 | 0.004763 | 5.5e-05 | 99.1216 | 84.971 | 114.9015 | 21.2612 | 0.002861 | 0.002672 | 261.9302 | 0.27315 | 8.4023 | 0.004056 | 0.005122 | 0.005069 | 4.9e-05 | 0.001135 | 0.27168 | 1.3665 | 2.548 | 98.89 | 20.86 | 38.605 | 90.6544 | 119.6294 | 86.0898 | 0.009636 | 0.98276 | 0.023244 | 23.6519 | 1.3778 | 0.1625 | 0.021731 | 1.5096 | 1.5096 | 0.004424
2019-03-21 15:50 | 1553197828.8653 | 38 | 0.20522 | 0.004676 | 5.1e-05 | 99.2327 | 85.2639 | 115.3522 | 21.2612 | 0.002843 | 0.002656 | 261.8928 | 0.27318 | 8.4003 | 0.004091 | 0.005926 | 0.005907 | 4.9e-05 | 0.001133 | 0.25625 | 1.3144 | 2.5479 | 98.89 | 20.86 | 38.606 | 90.2097 | 120.4766 | 85.7027 | 0.009637 | 0.98382 | 0.023302 | 23.6855 | 1.357 | 0.16252 | 0.020245 | 1.4102 | 1.4102 | 0.004737
2019-03-21 15:50 | 1553197829.8657 | 39 | 0.20544 | 0.004582 | 4.8e-05 | 99.2369 | 86.0125 | 115.8735 | 21.2612 | 0.00288 | 0.002685 | 261.8586 | 0.27371 | 8.3996 | 0.003968 | 0.005278 | 0.005178 | 5e-05 | 0.001135 | 0.27079 | 1.3121 | 2.548 | 98.89 | 20.86 | 38.607 | 91.1131 | 117.4893 | 86.4052 | 0.009637 | 0.98189 | 0.023332 | 23.7622 | 1.2792 | 0.16283 | 0.019072 | 1.3303 | 1.3303 | 0.005031
2019-03-21 15:50 | 1553197830.8644 | 40 | 0.20556 | 0.004479 | 4.4e-05 | 99.3153 | 86.4114 | 116.116 | 21.2612 | 0.002846 | 0.002657 | 261.804 | 0.27377 | 8.403 | 0.004065 | 0.006099 | 0.006054 | 5.1e-05 | 0.001131 | 0.24563 | 1.3363 | 2.548 | 98.89 | 20.86 | 38.609 | 90.2877 | 119.8578 | 85.7472 | 0.009637 | 0.98369 | 0.023378 | 23.7652 | 1.2449 | 0.16286 | 0.018478 | 1.2916 | 1.2916 | 0.005183
2019-03-21 15:51 | 1553197831.8644 | 41 | 0.20585 | 0.004374 | 4.1e-05 | 99.376 | 87.4822 | 116.6911 | 21.2612 | 0.002918 | 0.002611 | 261.7578 | 0.27372 | 8.4008 | 0.003989 | 0.005843 | 0.005787 | 4.7e-05 | 0.001132 | 0.26792 | 1.3224 | 2.5479 | 98.89 | 20.86 | 38.609 | 92.017 | 117.9813 | 84.6304 | 0.009637 | 0.98677 | 0.023472 | 23.7869 | 1.1402 | 0.16284 | 0.016881 | 1.1849 | 1.1849 | 0.005648
2019-03-21 15:51 | 1553197832.8635 | 42 | 0.20598 | 0.004259 | 3.8e-05 | 99.3629 | 88.6204 | 116.8378 | 21.2612 | 0.002844 | 0.002642 | 261.6892 | 0.27331 | 8.4001 | 0.004009 | 0.005734 | 0.005634 | 4.9e-05 | 0.001128 | 0.2521 | 1.3513 | 2.5478 | 98.89 | 20.86 | 38.61 | 90.2487 | 118.4822 | 85.3859 | 0.009638 | 0.98469 | 0.023462 | 23.827 | 1.0234 | 0.16259 | 0.016205 | 1.1371 | 1.1371 | 0.005877
2019-03-21 15:51 | 1553197833.8635 | 43 | 0.20607 | 0.004143 | 3.6e-05 | 99.3623 | 89.198 | 116.6549 | 21.2612 | 0.002865 | 0.002615 | 261.6392 | 0.27336 | 8.4041 | 0.003971 | 0.005267 | 0.005297 | 4.7e-05 | 0.00113 | 0.2505 | 1.3606 | 2.548 | 98.89 | 20.86 | 38.607 | 90.7512 | 117.5504 | 84.7288 | 0.009637 | 0.98649 | 0.023465 | 23.7865 | 0.96524 | 0.16262 | 0.015814 | 1.1099 | 1.1099 | 0.006022
2019-03-21 15:51 | 1553197834.8646 | 44 | 0.2061 | 0.004031 | 3.4e-05 | 99.4091 | 89.8721 | 116.5253 | 21.2612 | 0.002873 | 0.002605 | 261.5834 | 0.27329 | 8.4017 | 0.003993 | 0.005621 | 0.005517 | 4.6e-05 | 0.001129 | 0.27887 | 1.3113 | 2.5479 | 98.89 | 20.86 | 38.606 | 90.9288 | 118.0946 | 84.4869 | 0.009637 | 0.98716 | 0.02346 | 23.7653 | 0.90212 | 0.16258 | 0.015819 | 1.1102 | 1.1102 | 0.006019
2019-03-21 15:51 | 1553197835.8637 | 45 | 0.20622 | 0.00392 | 3.2e-05 | 99.4982 | 89.9106 | 116.2538 | 21.2612 | 0.002891 | 0.002622 | 261.5412 | 0.27325 | 8.4027 | 0.004007 | 0.00516 | 0.005068 | 4.9e-05 | 0.00113 | 0.2724 | 1.3581 | 2.5481 | 98.89 | 20.86 | 38.604 | 91.3822 | 118.4194 | 84.9051 | 0.009636 | 0.98601 | 0.023419 | 23.7514 | 0.9063 | 0.16256 | 0.015236 | 1.0675 | 1.0675 | 0.006258
2019-03-21 15:51 | 1553197836.864 | 46 | 0.20623 | 0.003812 | 3e-05 | 99.4365 | 89.8966 | 116.6645 | 21.2612 | 0.002767 | 0.002633 | 261.4919 | 0.27332 | 8.4036 | 0.004049 | 0.00524 | 0.005131 | 4.6e-05 | 0.001129 | 0.24873 | 1.3077 | 2.5481 | 98.89 | 20.86 | 38.604 | 88.3816 | 119.4532 | 85.1533 | 0.009636 | 0.98533 | 0.023452 | 23.8016 | 0.90214 | 0.1626 | 0.015278 | 1.0721 | 1.0721 | 0.006233
2019-03-21 15:51 | 1553197837.8638 | 47 | 0.20633 | 0.00372 | 2.8e-05 | 99.4329 | 90.3261 | 116.6833 | 21.2612 | 0.002893 | 0.00261 | 261.4376 | 0.27281 | 8.4032 | 0.003997 | 0.005246 | 0.005183 | 4.6e-05 | 0.001128 | 0.25409 | 1.3009 | 2.5479 | 98.89 | 20.86 | 38.603 | 91.4244 | 118.1984 | 84.6047 | 0.009636 | 0.98684 | 0.023472 | 23.7853 | 0.85919 | 0.16229 | 0.01482 | 1.0409 | 1.0409 | 0.006408
2019-03-21 15:51 | 1553197838.8639 | 48 | 0.20642 | 0.003628 | 2.7e-05 | 99.503 | 90.496 | 116.5897 | 21.2612 | 0.002874 | 0.002593 | 261.3494 | 0.27338 | 8.4059 | 0.00402 | 0.005281 | 0.005275 | 5.1e-05 | 0.001128 | 0.22251 | 1.2948 | 2.548 | 98.89 | 20.86 | 38.602 | 90.9724 | 118.7573 | 84.2052 | 0.009636 | 0.98794 | 0.023476 | 23.7625 | 0.84867 | 0.16264 | 0.014397 | 1.0114 | 1.0114 | 0.006609
2019-03-21 15:51 | 1553197839.8653 | 49 | 0.20643 | 0.003551 | 2.6e-05 | 99.5697 | 91.1721 | 116.2251 | 21.2612 | 0.002812 | 0.002648 | 261.312 | 0.27329 | 8.4 | 0.003961 | 0.005968 | 0.005907 | 4.7e-05 | 0.00113 | 0.2568 | 1.3087 | 2.5478 | 98.89 | 20.86 | 38.6 | 89.4751 | 117.3057 | 85.5185 | 0.009635 | 0.98432 | 0.023396 | 23.7688 | 0.7881 | 0.16258 | 0.014417 | 1.0095 | 1.0095 | 0.006619
2019-03-21 15:51 | 1553197840.865 | 50 | 0.20648 | 0.003476 | 2.5e-05 | 99.6442 | 91.513 | 115.974 | 21.2612 | 0.002789 | 0.002581 | 261.2745 | 0.27384 | 8.4037 | 0.004046 | 0.00541 | 0.005372 | 4.6e-05 | 0.001136 | 0.28368 | 1.3497 | 2.548 | 98.89 | 20.85 | 38.6 | 88.9119 | 119.3804 | 83.9032 | 0.009629 | 0.98877 | 0.023424 | 23.6896 | 0.7614 | 0.16291 | 0.014193 | 0.99502 | 0.99502 | 0.006729
2019-03-21 15:51 | 1553197841.864 | 51 | 0.20651 | 0.003406 | 2.4e-05 | 99.6099 | 91.2017 | 115.8144 | 21.2612 | 0.002887 | 0.002586 | 261.2389 | 0.2731 | 8.4042 | 0.003967 | 0.005224 | 0.005183 | 5e-05 | 0.001132 | 0.25265 | 1.3492 | 2.548 | 98.89 | 20.85 | 38.599 | 91.2625 | 117.4656 | 84.0227 | 0.009629 | 0.98844 | 0.023404 | 23.6773 | 0.7888 | 0.16246 | 0.014108 | 0.98826 | 0.98826 | 0.006757
2019-03-21 15:51 | 1553197842.8651 | 52 | 0.20652 | 0.003344 | 2.4e-05 | 99.6549 | 91.5106 | 116.2057 | 21.2612 | 0.002862 | 0.00257 | 261.2014 | 0.27314 | 8.4009 | 0.003957 | 0.00623 | 0.006113 | 4.5e-05 | 0.001129 | 0.27201 | 1.3107 | 2.5479 | 98.89 | 20.85 | 38.597 | 90.6652 | 117.2027 | 83.6551 | 0.009629 | 0.98946 | 0.023455 | 23.705 | 0.7626 | 0.16249 | 0.014168 | 0.99465 | 0.99465 | 0.006714
2019-03-21 15:51 | 1553197843.8645 | 53 | 0.20655 | 0.003288 | 2.4e-05 | 99.7321 | 91.8247 | 116.5082 | 21.2612 | 0.002834 | 0.002598 | 261.1874 | 0.27299 | 8.4018 | 0.003922 | 0.005792 | 0.005738 | 4.6e-05 | 0.001132 | 0.24347 | 1.3078 | 2.5479 | 98.89 | 20.85 | 38.597 | 90.0064 | 116.3579 | 84.3082 | 0.009629 | 0.98765 | 0.023464 | 23.7576 | 0.73888 | 0.1624 | 0.014045 | 0.98646 | 0.98646 | 0.006766
2019-03-21 15:51 | 1553197844.8652 | 54 | 0.20657 | 0.003234 | 2.4e-05 | 99.7851 | 91.7644 | 116.635 | 21.2612 | 0.002844 | 0.002533 | 261.1388 | 0.27232 | 8.4061 | 0.003917 | 0.005364 | 0.00537 | 4.9e-05 | 0.001131 | 0.25929 | 1.3273 | 2.5479 | 98.89 | 20.85 | 38.598 | 90.2417 | 116.2423 | 82.7455 | 0.009629 | 0.99199 | 0.023529 | 23.7185 | 0.74951 | 0.16201 | 0.013993 | 0.98544 | 0.98544 | 0.006757
2019-03-21 15:51 | 1553197845.8652 | 55 | 0.20654 | 0.003187 | 2.4e-05 | 99.7223 | 91.8917 | 116.6601 | 21.2612 | 0.002832 | 0.00262 | 261.091 | 0.27265 | 8.4038 | 0.003941 | 0.005428 | 0.005389 | 4.8e-05 | 0.00113 | 0.24635 | 1.3127 | 2.5479 | 98.89 | 20.85 | 38.6 | 89.9479 | 116.8091 | 84.8397 | 0.009629 | 0.98619 | 0.023462 | 23.7907 | 0.73146 | 0.1622 | 0.014209 | 0.99786 | 0.99786 | 0.006681
2019-03-21 15:51 | 1553197846.8653 | 56 | 0.20659 | 0.00315 | 2.4e-05 | 99.7482 | 91.8093 | 116.2991 | 21.2612 | 0.00279 | 0.002533 | 261.0301 | 0.27296 | 8.4019 | 0.003856 | 0.005429 | 0.005305 | 4.8e-05 | 0.001127 | 0.27616 | 1.3073 | 2.548 | 98.89 | 20.85 | 38.603 | 88.9314 | 114.7463 | 82.7611 | 0.00963 | 0.99195 | 0.023494 | 23.6848 | 0.74181 | 0.16238 | 0.013981 | 0.98321 | 0.98321 | 0.006788
2019-03-21 15:51 | 1553197847.8638 | 57 | 0.20657 | 0.003114 | 2.3e-05 | 99.7526 | 92.0937 | 116.3277 | 21.2612 | 0.002834 | 0.002629 | 260.9858 | 0.27362 | 8.4021 | 0.00387 | 0.005492 | 0.005391 | 5e-05 | 0.001129 | 0.24906 | 1.3032 | 2.5478 | 98.89 | 20.85 | 38.604 | 89.9963 | 115.0778 | 85.0549 | 0.00963 | 0.9856 | 0.023422 | 23.764 | 0.71455 | 0.16278 | 0.014117 | 0.98977 | 0.98977 | 0.006759
2019-03-21 15:51 | 1553197848.8644 | 58 | 0.20666 | 0.003082 | 2.3e-05 | 99.8343 | 92.3464 | 116.5922 | 21.2612 | 0.002783 | 0.002625 | 260.9468 | 0.27256 | 8.4011 | 0.00386 | 0.005343 | 0.005295 | 4.6e-05 | 0.001127 | 0.262 | 1.3607 | 2.548 | 98.89 | 20.85 | 38.607 | 88.7745 | 114.8428 | 84.9785 | 0.009631 | 0.98581 | 0.023451 | 23.7884 | 0.69736 | 0.16215 | 0.013634 | 0.95707 | 0.95707 | 0.006963
2019-03-21 15:51 | 1553197849.8644 | 59 | 0.20668 | 0.003052 | 2.3e-05 | 99.7905 | 92.5283 | 116.8845 | 21.2612 | 0.00286 | 0.002556 | 260.9074 | 0.27339 | 8.4051 | 0.003858 | 0.004906 | 0.004886 | 4.8e-05 | 0.001125 | 0.24834 | 1.329 | 2.5479 | 98.89 | 20.85 | 38.609 | 90.6255 | 114.8048 | 83.3026 | 0.009632 | 0.99044 | 0.023535 | 23.7624 | 0.67583 | 0.16264 | 0.013554 | 0.9549 | 0.9549 | 0.007
2019-03-21 15:51 | 1553197850.864 | 60 | 0.20669 | 0.003017 | 2.2e-05 | 99.859 | 93.2224 | 116.9502 | 21.2612 | 0.002884 | 0.002536 | 260.8448 | 0.27315 | 8.4013 | 0.003847 | 0.005893 | 0.005815 | 5e-05 | 0.001127 | 0.23291 | 1.3103 | 2.5479 | 98.89 | 20.85 | 38.615 | 91.2001 | 114.5162 | 82.8393 | 0.009633 | 0.99173 | 0.023557 | 23.7536 | 0.61513 | 0.1625 | 0.013556 | 0.956 | 0.956 | 0.006986
2019-03-21 15:51 | 1553197851.8645 | 61 | 0.20675 | 0.002987 | 2.1e-05 | 99.9761 | 93.3593 | 116.846 | 21.2612 | 0.002791 | 0.00261 | 260.8143 | 0.27315 | 8.4025 | 0.003833 | 0.005318 | 0.005287 | 4.9e-05 | 0.001126 | 0.25769 | 1.3 | 2.5479 | 98.89 | 20.85 | 38.618 | 88.9565 | 114.1873 | 84.6054 | 0.009634 | 0.98683 | 0.023489 | 23.8019 | 0.61248 | 0.1625 | 0.013233 | 0.93051 | 0.93051 | 0.007178
2019-03-21 15:51 | 1553197852.8646 | 62 | 0.20679 | 0.00296 | 2.1e-05 | 99.9782 | 93.6529 | 116.5289 | 21.2612 | 0.002859 | 0.00259 | 260.7404 | 0.2736 | 8.4036 | 0.003932 | 0.005294 | 0.005256 | 4.9e-05 | 0.001125 | 0.27168 | 1.2834 | 2.5479 | 98.89 | 20.85 | 38.621 | 90.5886 | 116.5995 | 84.1187 | 0.009635 | 0.98818 | 0.023473 | 23.7534 | 0.58458 | 0.16277 | 0.012997 | 0.91335 | 0.91335 | 0.007324
2019-03-21 15:51 | 1553197853.8653 | 63 | 0.20681 | 0.002933 | 1.9e-05 | 99.9818 | 93.7793 | 116.3944 | 21.2612 | 0.002824 | 0.002551 | 260.699 | 0.27302 | 8.4022 | 0.003845 | 0.00564 | 0.005534 | 4.8e-05 | 0.001128 | 0.25752 | 1.3639 | 2.5481 | 98.89 | 20.85 | 38.626 | 89.7627 | 114.4846 | 83.1824 | 0.009636 | 0.99078 | 0.02349 | 23.7085 | 0.57284 | 0.16242 | 0.012926 | 0.9091 | 0.9091 | 0.007343
2019-03-21 15:51 | 1553197854.8646 | 64 | 0.20682 | 0.002902 | 1.9e-05 | 100.0229 | 93.8515 | 116.3847 | 21.2612 | 0.00287 | 0.002537 | 260.6579 | 0.27321 | 8.4015 | 0.003793 | 0.005445 | 0.005381 | 4.6e-05 | 0.001125 | 0.25874 | 1.2784 | 2.5479 | 98.89 | 20.85 | 38.629 | 90.8529 | 113.208 | 82.8487 | 0.009637 | 0.9917 | 0.0235 | 23.6965 | 0.56964 | 0.16254 | 0.012904 | 0.90795 | 0.90795 | 0.007357
2019-03-21 15:51 | 1553197855.8642 | 65 | 0.20682 | 0.002871 | 1.9e-05 | 100.0133 | 94.0422 | 116.4298 | 21.2612 | 0.002808 | 0.002552 | 260.6129 | 0.27381 | 8.4029 | 0.003839 | 0.004637 | 0.004668 | 4.8e-05 | 0.001122 | 0.20957 | 1.2809 | 2.5481 | 98.89 | 20.85 | 38.63 | 89.3658 | 114.3398 | 83.2206 | 0.009637 | 0.99067 | 0.023492 | 23.7134 | 0.55062 | 0.16289 | 0.012914 | 0.90833 | 0.90833 | 0.00737
2019-03-21 15:51 | 1553197856.8648 | 66 | 0.20683 | 0.002844 | 1.9e-05 | 100.117 | 94.6795 | 116.631 | 21.2612 | 0.002842 | 0.002577 | 260.57 | 0.27355 | 8.4025 | 0.003821 | 0.005241 | 0.005134 | 4.6e-05 | 0.001126 | 0.2557 | 1.308 | 2.548 | 98.89 | 20.85 | 38.637 | 90.1847 | 113.9005 | 83.8152 | 0.009639 | 0.98902 | 0.023493 | 23.7537 | 0.49949 | 0.16274 | 0.012907 | 0.90792 | 0.90792 | 0.007367
2019-03-21 15:51 | 1553197857.8639 | 67 | 0.20685 | 0.002819 | 1.9e-05 | 100.0756 | 94.8859 | 116.5188 | 21.2612 | 0.002854 | 0.002551 | 260.5466 | 0.27372 | 8.4033 | 0.003811 | 0.005374 | 0.005276 | 4.7e-05 | 0.001119 | 0.27936 | 1.3163 | 2.5481 | 98.89 | 20.85 | 38.64 | 90.4706 | 113.6608 | 83.1973 | 0.009639 | 0.99073 | 0.023502 | 23.7217 | 0.47631 | 0.16284 | 0.012803 | 0.90096 | 0.90096 | 0.007428
2019-03-21 15:51 | 1553197858.8653 | 68 | 0.20681 | 0.002795 | 1.9e-05 | 100.1447 | 95.3583 | 116.2118 | 21.2612 | 0.002799 | 0.002537 | 260.5049 | 0.27249 | 8.403 | 0.003763 | 0.005377 | 0.005351 | 4.8e-05 | 0.001117 | 0.26465 | 1.3367 | 2.548 | 98.89 | 20.85 | 38.644 | 89.151 | 112.4876 | 82.8494 | 0.00964 | 0.9917 | 0.023482 | 23.6789 | 0.43806 | 0.1621 | 0.013062 | 0.91836 | 0.91836 | 0.007255
2019-03-21 15:51 | 1553197859.8654 | 69 | 0.20684 | 0.002768 | 1.8e-05 | 100.1799 | 95.4015 | 116.2737 | 21.2612 | 0.002767 | 0.002564 | 260.4484 | 0.27292 | 8.405 | 0.003863 | 0.004953 | 0.004934 | 4.6e-05 | 0.001117 | 0.25984 | 1.2981 | 2.5479 | 98.89 | 20.85 | 38.646 | 88.3736 | 114.9164 | 83.5108 | 0.009641 | 0.98986 | 0.023467 | 23.7072 | 0.43715 | 0.16236 | 0.01294 | 0.90925 | 0.90925 | 0.007339
2019-03-21 15:51 | 1553197860.865 | 70 | 0.20687 | 0.002753 | 1.9e-05 | 100.1355 | 95.158 | 116.3265 | 21.2612 | 0.002816 | 0.002468 | 260.3806 | 0.27466 | 8.4027 | 0.003859 | 0.005666 | 0.005601 | 4.7e-05 | 0.001121 | 0.25874 | 1.3298 | 2.5479 | 98.89 | 20.85 | 38.648 | 89.5681 | 114.8207 | 81.1811 | 0.009641 | 0.99637 | 0.023549 | 23.635 | 0.45604 | 0.16339 | 0.012785 | 0.90145 | 0.90145 | 0.00745
2019-03-21 15:51 | 1553197861.8654 | 71 | 0.20687 | 0.002734 | 1.8e-05 | 100.2025 | 95.4877 | 116.638 | 21.2612 | 0.002821 | 0.002529 | 260.3376 | 0.27295 | 8.403 | 0.003931 | 0.005209 | 0.005219 | 4.3e-05 | 0.001118 | 0.23092 | 1.2968 | 2.548 | 98.89 | 20.85 | 38.649 | 89.6815 | 116.5836 | 82.67 | 0.009642 | 0.9922 | 0.023531 | 23.7163 | 0.43109 | 0.16238 | 0.012816 | 0.903 | 0.903 | 0.007391
2019-03-21 15:51 | 1553197862.8646 | 72 | 0.20692 | 0.002716 | 1.9e-05 | 100.2368 | 95.2707 | 116.6798 | 21.2612 | 0.002823 | 0.002552 | 260.3176 | 0.27361 | 8.4024 | 0.003893 | 0.005215 | 0.005129 | 4.8e-05 | 0.001118 | 0.25282 | 1.3354 | 2.548 | 98.89 | 20.85 | 38.65 | 89.7415 | 115.6442 | 83.2042 | 0.009642 | 0.99072 | 0.023518 | 23.7383 | 0.4545 | 0.16277 | 0.01254 | 0.88302 | 0.88302 | 0.007576
2019-03-21 15:51 | 1553197863.8642 | 73 | 0.20694 | 0.002702 | 1.8e-05 | 100.2437 | 95.7135 | 116.5474 | 21.2612 | 0.002798 | 0.002576 | 260.2573 | 0.27287 | 8.3983 | 0.003902 | 0.005612 | 0.005506 | 5e-05 | 0.001119 | 0.28102 | 1.3011 | 2.5478 | 98.89 | 20.85 | 38.65 | 89.1281 | 115.8735 | 83.7847 | 0.009642 | 0.9891 | 0.023485 | 23.7442 | 0.41364 | 0.16233 | 0.012438 | 0.87471 | 0.87471 | 0.007628
2019-03-21 15:51 | 1553197864.8642 | 74 | 0.20697 | 0.002684 | 1.8e-05 | 100.2073 | 95.5016 | 116.7533 | 21.2612 | 0.002893 | 0.002537 | 260.2062 | 0.27353 | 8.4025 | 0.003956 | 0.004822 | 0.00479 | 4.9e-05 | 0.001118 | 0.28119 | 1.3405 | 2.548 | 98.89 | 20.85 | 38.651 | 91.4189 | 117.1964 | 82.8587 | 0.009642 | 0.99168 | 0.023537 | 23.7343 | 0.43021 | 0.16272 | 0.012266 | 0.8645 | 0.8645 | 0.007736
2019-03-21 15:51 | 1553197865.8651 | 75 | 0.20691 | 0.00267 | 1.7e-05 | 100.398 | 95.6079 | 116.8246 | 21.2612 | 0.002806 | 0.002525 | 260.1573 | 0.27286 | 8.3995 | 0.003913 | 0.005751 | 0.00565 | 4.8e-05 | 0.001118 | 0.22876 | 1.2869 | 2.5479 | 98.89 | 20.85 | 38.65 | 89.3119 | 116.1259 | 82.571 | 0.009642 | 0.99248 | 0.023553 | 23.7319 | 0.43727 | 0.16233 | 0.012656 | 0.89261 | 0.89261 | 0.007474
2019-03-21 15:51 | 1553197866.8644 | 76 | 0.20697 | 0.002653 | 1.8e-05 | 100.3158 | 95.2952 | 116.8234 | 21.2612 | 0.002798 | 0.002554 | 260.093 | 0.27359 | 8.4023 | 0.003953 | 0.005794 | 0.005732 | 5e-05 | 0.001115 | 0.25498 | 1.3339 | 2.5481 | 98.89 | 20.85 | 38.65 | 89.1407 | 117.1173 | 83.2689 | 0.009642 | 0.99054 | 0.02353 | 23.7551 | 0.45924 | 0.16276 | 0.012269 | 0.86448 | 0.86448 | 0.007738
2019-03-21 15:51 | 1553197867.8644 | 77 | 0.20696 | 0.002639 | 1.8e-05 | 100.3375 | 95.393 | 116.9664 | 21.2612 | 0.00281 | 0.002497 | 260.0362 | 0.27361 | 8.4055 | 0.00386 | 0.005056 | 0.004956 | 5e-05 | 0.001119 | 0.25177 | 1.333 | 2.5479 | 98.89 | 20.85 | 38.649 | 89.4135 | 114.8475 | 81.8979 | 0.009642 | 0.99436 | 0.02359 | 23.7239 | 0.45201 | 0.16277 | 0.012375 | 0.87412 | 0.87412 | 0.007653
2019-03-21 15:51 | 1553197868.8637 | 78 | 0.207 | 0.002626 | 1.7e-05 | 100.4218 | 95.6638 | 116.8812 | 21.2612 | 0.002845 | 0.002499 | 259.9909 | 0.27347 | 8.4021 | 0.003773 | 0.005438 | 0.005367 | 5.3e-05 | 0.001117 | 0.26842 | 1.3252 | 2.548 | 98.89 | 20.85 | 38.647 | 90.2738 | 112.7349 | 81.9351 | 0.009641 | 0.99426 | 0.02358 | 23.7165 | 0.43417 | 0.16269 | 0.012123 | 0.85604 | 0.85604 | 0.007811
2019-03-21 15:51 | 1553197869.8635 | 79 | 0.207 | 0.002609 | 1.7e-05 | 100.4168 | 96.111 | 116.8359 | 21.2612 | 0.002766 | 0.002559 | 259.9468 | 0.2728 | 8.4006 | 0.00375 | 0.004856 | 0.004877 | 5e-05 | 0.001121 | 0.23667 | 1.3326 | 2.548 | 98.89 | 20.85 | 38.644 | 88.3652 | 112.1687 | 83.3822 | 0.00964 | 0.99022 | 0.023528 | 23.7601 | 0.392 | 0.16229 | 0.012147 | 0.85583 | 0.85583 | 0.007794
2019-03-21 15:51 | 1553197870.8649 | 80 | 0.20702 | 0.002597 | 1.7e-05 | 100.3801 | 96.3082 | 116.7265 | 21.2612 | 0.002816 | 0.00251 | 259.9099 | 0.27368 | 8.4023 | 0.00382 | 0.005226 | 0.0051 | 4.7e-05 | 0.001116 | 0.25249 | 1.3571 | 2.548 | 98.89 | 20.85 | 38.642 | 89.5611 | 113.8649 | 82.2051 | 0.00964 | 0.9935 | 0.023556 | 23.7098 | 0.3704 | 0.16281 | 0.012087 | 0.85261 | 0.85261 | 0.007848
2019-03-21 15:51 | 1553197871.8636 | 81 | 0.20702 | 0.002583 | 1.8e-05 | 100.4007 | 96.6365 | 116.3249 | 21.2612 | 0.002842 | 0.002489 | 259.8538 | 0.27442 | 8.3997 | 0.003805 | 0.005883 | 0.005764 | 4.9e-05 | 0.001118 | 0.25033 | 1.3161 | 2.5479 | 98.89 | 20.85 | 38.638 | 90.1955 | 113.5122 | 81.688 | 0.009639 | 0.99495 | 0.023532 | 23.6517 | 0.34179 | 0.16325 | 0.012094 | 0.85223 | 0.85223 | 0.007873
2019-03-21 15:51 | 1553197872.8636 | 82 | 0.207 | 0.002563 | 1.7e-05 | 100.4981 | 97.0016 | 116.3711 | 21.2612 | 0.002768 | 0.00254 | 259.7963 | 0.27407 | 8.4003 | 0.003782 | 0.005626 | 0.005622 | 4.9e-05 | 0.00112 | 0.28368 | 1.3202 | 2.5479 | 98.89 | 20.85 | 38.637 | 88.4049 | 112.9538 | 82.9164 | 0.009639 | 0.99152 | 0.023496 | 23.6973 | 0.31674 | 0.16305 | 0.012251 | 0.86199 | 0.86199 | 0.007774
2019-03-21 15:51 | 1553197873.864 | 83 | 0.207 | 0.002554 | 1.8e-05 | 100.4349 | 96.803 | 116.7505 | 21.2612 | 0.002764 | 0.002536 | 259.7563 | 0.27373 | 8.401 | 0.003753 | 0.005014 | 0.004971 | 5.1e-05 | 0.001121 | 0.29175 | 1.3522 | 2.5479 | 98.89 | 20.86 | 38.635 | 88.3134 | 112.2262 | 82.8188 | 0.009644 | 0.99179 | 0.023538 | 23.7327 | 0.32944 | 0.16284 | 0.012246 | 0.86309 | 0.86309 | 0.007754
2019-03-21 15:51 | 1553197874.8638 | 84 | 0.207 | 0.002537 | 1.7e-05 | 100.5148 | 96.659 | 116.6854 | 21.2612 | 0.00277 | 0.002519 | 259.708 | 0.27405 | 8.4022 | 0.003799 | 0.005548 | 0.005398 | 5e-05 | 0.001122 | 0.27975 | 1.3233 | 2.548 | 98.89 | 20.86 | 38.632 | 88.4449 | 113.3488 | 82.4202 | 0.009643 | 0.9929 | 0.023544 | 23.7128 | 0.34987 | 0.16303 | 0.012277 | 0.86554 | 0.86554 | 0.007741
2019-03-21 15:51 | 1553197875.8652 | 85 | 0.207 | 0.002529 | 1.7e-05 | 100.5246 | 96.8425 | 116.6596 | 21.2612 | 0.002829 | 0.002533 | 259.6629 | 0.27293 | 8.4005 | 0.003826 | 0.004466 | 0.004471 | 4.5e-05 | 0.001123 | 0.24298 | 1.2765 | 2.5479 | 98.89 | 20.86 | 38.628 | 89.8884 | 114.0094 | 82.7479 | 0.009642 | 0.99199 | 0.023531 | 23.7211 | 0.33378 | 0.16237 | 0.012257 | 0.86369 | 0.86369 | 0.007726
2019-03-21 15:51 | 1553197876.8649 | 86 | 0.207 | 0.002518 | 1.7e-05 | 100.566 | 97.004 | 116.7575 | 21.2612 | 0.002785 | 0.002564 | 259.6413 | 0.27332 | 8.4025 | 0.003803 | 0.004897 | 0.004812 | 4.7e-05 | 0.001123 | 0.29103 | 1.2785 | 2.5479 | 98.89 | 20.86 | 38.626 | 88.8065 | 113.4516 | 83.4993 | 0.009642 | 0.98989 | 0.023516 | 23.7561 | 0.32256 | 0.1626 | 0.012296 | 0.86588 | 0.86588 | 0.007718
2019-03-21 15:51 | 1553197877.8639 | 87 | 0.20698 | 0.002509 | 1.7e-05 | 100.563 | 97.4825 | 116.4305 | 21.2612 | 0.002803 | 0.002536 | 259.6036 | 0.27385 | 8.4018 | 0.003766 | 0.005254 | 0.005198 | 4.8e-05 | 0.001123 | 0.24563 | 1.2932 | 2.548 | 98.89 | 20.86 | 38.625 | 89.2565 | 112.547 | 82.8236 | 0.009641 | 0.99177 | 0.023505 | 23.7003 | 0.27828 | 0.16291 | 0.012402 | 0.87288 | 0.87288 | 0.007671
2019-03-21 15:51 | 1553197878.8634 | 88 | 0.20699 | 0.002497 | 1.7e-05 | 100.6509 | 97.2114 | 116.6307 | 21.2612 | 0.002781 | 0.002599 | 259.5749 | 0.27411 | 8.4004 | 0.003812 | 0.005456 | 0.005416 | 4.7e-05 | 0.001126 | 0.25625 | 1.3128 | 2.548 | 98.89 | 20.86 | 38.622 | 88.7119 | 113.6808 | 84.3378 | 0.009641 | 0.98757 | 0.023476 | 23.7711 | 0.31101 | 0.16307 | 0.0124 | 0.87171 | 0.87171 | 0.007688
2019-03-21 15:51 | 1553197879.8641 | 89 | 0.20696 | 0.002494 | 1.7e-05 | 100.611 | 97.1126 | 116.4789 | 21.2612 | 0.002785 | 0.002536 | 259.5173 | 0.27354 | 8.4026 | 0.003746 | 0.005643 | 0.005567 | 4.9e-05 | 0.001129 | 0.29717 | 1.3205 | 2.548 | 98.89 | 20.86 | 38.62 | 88.8253 | 112.0787 | 82.8251 | 0.00964 | 0.99177 | 0.02351 | 23.7053 | 0.31655 | 0.16273 | 0.012544 | 0.88311 | 0.88311 | 0.007574
2019-03-21 15:51 | 1553197880.8642 | 90 | 0.20697 | 0.002493 | 1.8e-05 | 100.65 | 96.7458 | 116.5777 | 21.2612 | 0.002888 | 0.002544 | 259.4761 | 0.27394 | 8.4011 | 0.003781 | 0.005609 | 0.005503 | 4.8e-05 | 0.001127 | 0.24618 | 1.3011 | 2.5479 | 98.89 | 20.86 | 38.618 | 91.3046 | 112.9141 | 83.0263 | 0.00964 | 0.99121 | 0.023513 | 23.722 | 0.35387 | 0.16296 | 0.012518 | 0.88138 | 0.88138 | 0.007599
2019-03-21 15:51 | 1553197881.8644 | 91 | 0.20693 | 0.002494 | 1.8e-05 | 100.6054 | 96.7369 | 117.0549 | 21.2612 | 0.002804 | 0.002481 | 259.4315 | 0.27307 | 8.4015 | 0.003764 | 0.005364 | 0.005281 | 4.8e-05 | 0.001125 | 0.27959 | 1.2691 | 2.5479 | 98.88 | 20.86 | 38.614 | 89.2844 | 112.5178 | 81.514 | 0.00964 | 0.99544 | 0.023612 | 23.72 | 0.35072 | 0.16245 | 0.012713 | 0.8988 | 0.8988 | 0.007428
2019-03-21 15:51 | 1553197882.8633 | 92 | 0.20692 | 0.002487 | 1.8e-05 | 100.6173 | 96.472 | 117.1504 | 21.2612 | 0.002812 | 0.002547 | 259.4039 | 0.27327 | 8.4031 | 0.003685 | 0.005708 | 0.005648 | 5e-05 | 0.001128 | 0.27168 | 1.2797 | 2.548 | 98.88 | 20.86 | 38.612 | 89.4633 | 110.579 | 83.0944 | 0.009639 | 0.99102 | 0.023569 | 23.7825 | 0.37631 | 0.16257 | 0.012808 | 0.9039 | 0.9039 | 0.007392
2019-03-21 15:51 | 1553197883.9168 | 93 | 0.20693 | 0.002487 | 1.9e-05 | 100.6879 | 96.8647 | 116.8765 | 21.2612 | 0.00278 | 0.002534 | 259.368 | 0.27347 | 8.4035 | 0.003789 | 0.004869 | 0.004851 | 4.7e-05 | 0.001126 | 0.27345 | 1.2862 | 2.548 | 98.88 | 20.86 | 38.608 | 88.6861 | 113.1204 | 82.7712 | 0.009638 | 0.99192 | 0.023552 | 23.7439 | 0.34625 | 0.16269 | 0.012766 | 0.90026 | 0.90026 | 0.007427
2019-03-21 15:51 | 1553197884.8657 | 94 | 0.20693 | 0.002487 | 1.9e-05 | 100.7141 | 96.9818 | 116.7178 | 21.2612 | 0.002789 | 0.002502 | 259.3342 | 0.2717 | 8.4023 | 0.003787 | 0.005325 | 0.005194 | 4.9e-05 | 0.001121 | 0.2589 | 1.3425 | 2.5481 | 98.88 | 20.87 | 38.606 | 88.9182 | 113.0707 | 82.021 | 0.009644 | 0.99401 | 0.023561 | 23.7027 | 0.33777 | 0.16164 | 0.012737 | 0.89848 | 0.89848 | 0.007394
2019-03-21 15:51 | 1553197885.8659 | 95 | 0.20691 | 0.002486 | 1.9e-05 | 100.7171 | 96.9774 | 116.5704 | 21.2612 | 0.002794 | 0.002531 | 259.2864 | 0.27396 | 8.4004 | 0.003772 | 0.00526 | 0.005217 | 5e-05 | 0.001123 | 0.28136 | 1.3594 | 2.548 | 98.88 | 20.87 | 38.603 | 89.0251 | 112.6938 | 82.6982 | 0.009643 | 0.99212 | 0.023524 | 23.7103 | 0.33844 | 0.16298 | 0.012863 | 0.90594 | 0.90594 | 0.007394
2019-03-21 15:51 | 1553197886.8668 | 96 | 0.20687 | 0.002492 | 2e-05 | 100.6786 | 96.8371 | 116.5305 | 21.2612 | 0.002838 | 0.002599 | 259.2367 | 0.27327 | 8.4063 | 0.003775 | 0.00559 | 0.005501 | 4.9e-05 | 0.001117 | 0.26593 | 1.3177 | 2.5479 | 98.88 | 20.86 | 38.6 | 90.091 | 112.7783 | 84.3418 | 0.009636 | 0.98756 | 0.023465 | 23.761 | 0.34798 | 0.16257 | 0.013082 | 0.91904 | 0.91904 | 0.00727
2019-03-21 15:51 | 1553197887.8673 | 97 | 0.20687 | 0.002491 | 1.9e-05 | 100.7388 | 96.9008 | 116.6624 | 21.2612 | 0.002821 | 0.002548 | 259.1767 | 0.27365 | 8.4018 | 0.003775 | 0.005058 | 0.004938 | 4.7e-05 | 0.001119 | 0.28479 | 1.3661 | 2.548 | 98.88 | 20.86 | 38.598 | 89.6931 | 112.7617 | 83.1069 | 0.009636 | 0.99099 | 0.023519 | 23.7333 | 0.34744 | 0.16279 | 0.013084 | 0.92141 | 0.92141 | 0.007262
2019-03-21 15:51 | 1553197888.8659 | 98 | 0.20689 | 0.002497 | 1.9e-05 | 100.7394 | 97.0595 | 116.6455 | 21.2612 | 0.0028 | 0.002525 | 259.1424 | 0.27388 | 8.4017 | 0.003689 | 0.005169 | 0.00512 | 5e-05 | 0.001119 | 0.25337 | 1.3592 | 2.5481 | 98.88 | 20.86 | 38.596 | 89.1869 | 110.6886 | 82.5599 | 0.009635 | 0.99251 | 0.023536 | 23.7134 | 0.33286 | 0.16293 | 0.012951 | 0.91261 | 0.91261 | 0.007338
2019-03-21 15:51 | 1553197889.8664 | 99 | 0.2069 | 0.002494 | 1.9e-05 | 100.7901 | 96.9481 | 116.6347 | 21.2612 | 0.002818 | 0.002494 | 259.106 | 0.27403 | 8.4003 | 0.003694 | 0.00553 | 0.005423 | 4.9e-05 | 0.001116 | 0.25442 | 1.3764 | 2.5479 | 98.88 | 20.86 | 38.594 | 89.6172 | 110.7952 | 81.8198 | 0.009635 | 0.99458 | 0.023559 | 23.6876 | 0.34762 | 0.16302 | 0.012904 | 0.91023 | 0.91023 | 0.007361
2019-03-21 15:52 | 1553197890.8664 | 100 | 0.20684 | 0.002502 | 1.9e-05 | 100.7123 | 96.7676 | 116.8117 | 21.2612 | 0.002753 | 0.002543 | 259.1008 | 0.27381 | 8.4001 | 0.003714 | 0.005025 | 0.004926 | 5.1e-05 | 0.001113 | 0.26665 | 1.2917 | 2.5479 | 98.88 | 20.87 | 38.591 | 88.0364 | 111.3011 | 83.0023 | 0.00964 | 0.99128 | 0.023538 | 23.745 | 0.35739 | 0.16289 | 0.013252 | 0.9339 | 0.9339 | 0.007168
2019-03-21 15:52 | 1553197891.8661 | 101 | 0.20686 | 0.002499 | 1.9e-05 | 100.8014 | 96.9301 | 116.4826 | 21.2612 | 0.002789 | 0.002548 | 259.0638 | 0.27378 | 8.3999 | 0.00367 | 0.005274 | 0.005216 | 4.8e-05 | 0.001113 | 0.23938 | 1.2843 | 2.5479 | 98.88 | 20.87 | 38.591 | 88.9081 | 110.2194 | 83.1146 | 0.00964 | 0.99096 | 0.023501 | 23.7153 | 0.35029 | 0.16287 | 0.013158 | 0.92585 | 0.92585 | 0.00723
2019-03-21 15:52 | 1553197892.8672 | 102 | 0.20686 | 0.002503 | 1.9e-05 | 100.7362 | 96.9757 | 116.5655 | 21.2612 | 0.002826 | 0.002555 | 259.0256 | 0.27337 | 8.4023 | 0.003629 | 0.005322 | 0.005234 | 5e-05 | 0.001114 | 0.30541 | 1.2533 | 2.5479 | 98.88 | 20.87 | 38.59 | 89.8128 | 109.2321 | 83.2776 | 0.00964 | 0.99051 | 0.023504 | 23.7291 | 0.34029 | 0.16263 | 0.01314 | 0.92469 | 0.92469 | 0.007228
2019-03-21 15:52 | 1553197893.8662 | 103 | 0.20686 | 0.002502 | 1.9e-05 | 100.7403 | 96.6934 | 116.6399 | 21.2612 | 0.002768 | 0.002502 | 258.9475 | 0.27379 | 8.4044 | 0.003747 | 0.005228 | 0.005184 | 5e-05 | 0.001113 | 0.25409 | 1.2713 | 2.5479 | 98.88 | 20.87 | 38.59 | 88.3994 | 112.0977 | 82.0157 | 0.00964 | 0.99403 | 0.023553 | 23.6947 | 0.36674 | 0.16287 | 0.013125 | 0.92559 | 0.92559 | 0.007232
2019-03-21 15:52 | 1553197894.8653 | 104 | 0.20685 | 0.00251 | 1.9e-05 | 100.7788 | 97.226 | 116.4721 | 21.2612 | 0.002774 | 0.002552 | 258.9168 | 0.27357 | 8.4031 | 0.003761 | 0.005183 | 0.005166 | 4.9e-05 | 0.001113 | 0.286 | 1.2684 | 2.548 | 98.88 | 20.87 | 38.591 | 88.5611 | 112.434 | 83.2081 | 0.00964 | 0.9907 | 0.023497 | 23.7173 | 0.32101 | 0.16275 | 0.013206 | 0.92906 | 0.92906 | 0.0072
2019-03-21 15:52 | 1553197895.8653 | 105 | 0.20684 | 0.002513 | 1.9e-05 | 100.8145 | 97.1007 | 116.5369 | 21.2612 | 0.002764 | 0.002498 | 258.8509 | 0.27403 | 8.4019 | 0.003813 | 0.005557 | 0.005463 | 5e-05 | 0.001115 | 0.24314 | 1.3211 | 2.5479 | 98.88 | 20.87 | 38.592 | 88.3127 | 113.6964 | 81.9219 | 0.00964 | 0.99429 | 0.023546 | 23.6811 | 0.33572 | 0.16302 | 0.013269 | 0.93542 | 0.93542 | 0.007163
2019-03-21 15:52 | 1553197896.8653 | 106 | 0.20679 | 0.002513 | 1.9e-05 | 100.7874 | 96.7921 | 116.3653 | 21.2612 | 0.002801 | 0.002575 | 258.8265 | 0.27285 | 8.4001 | 0.003781 | 0.006013 | 0.005935 | 4.8e-05 | 0.001115 | 0.25553 | 1.2887 | 2.5481 | 98.88 | 20.87 | 38.594 | 89.2127 | 112.9268 | 83.7635 | 0.009641 | 0.98916 | 0.023468 | 23.7249 | 0.36179 | 0.16232 | 0.01353 | 0.95064 | 0.95064 | 0.007018
2019-03-21 15:52 | 1553197897.8664 | 107 | 0.2068 | 0.002517 | 1.9e-05 | 100.8836 | 96.9079 | 116.3115 | 21.2612 | 0.002745 | 0.002508 | 258.7745 | 0.27339 | 8.4005 | 0.003763 | 0.005534 | 0.005431 | 4.9e-05 | 0.001117 | 0.2589 | 1.3116 | 2.548 | 98.88 | 20.87 | 38.595 | 87.852 | 112.4917 | 82.1582 | 0.009641 | 0.99363 | 0.023515 | 23.666 | 0.35963 | 0.16264 | 0.013515 | 0.95155 | 0.95155 | 0.007025
2019-03-21 15:52 | 1553197898.8655 | 108 | 0.20681 | 0.00252 | 1.9e-05 | 100.7895 | 96.8088 | 116.0165 | 21.2612 | 0.00282 | 0.002575 | 258.7031 | 0.2732 | 8.4005 | 0.003811 | 0.004985 | 0.004911 | 5.3e-05 | 0.001115 | 0.26609 | 1.3367 | 2.5479 | 98.88 | 20.87 | 38.595 | 89.6659 | 113.6586 | 83.7551 | 0.009641 | 0.98919 | 0.023433 | 23.689 | 0.36043 | 0.16253 | 0.013439 | 0.94284 | 0.94284 | 0.007085
2019-03-21 15:52 | 1553197899.8655 | 109 | 0.20675 | 0.002523 | 2e-05 | 100.8547 | 96.9672 | 115.9397 | 21.2612 | 0.002828 | 0.002597 | 258.6585 | 0.2741 | 8.4001 | 0.003906 | 0.005785 | 0.00568 | 5e-05 | 0.001112 | 0.2453 | 1.3323 | 2.5479 | 98.88 | 20.87 | 38.594 | 89.8573 | 115.9669 | 84.3005 | 0.009641 | 0.98768 | 0.023407 | 23.6993 | 0.3516 | 0.16306 | 0.013787 | 0.9662 | 0.9662 | 0.006936
2019-03-21 15:52 | 1553197900.8666 | 110 | 0.20675 | 0.002528 | 2e-05 | 100.7573 | 96.6828 | 115.6461 | 21.2612 | 0.002822 | 0.002537 | 258.5935 | 0.27349 | 8.4043 | 0.00388 | 0.005166 | 0.005107 | 4.8e-05 | 0.001109 | 0.24419 | 1.3678 | 2.5481 | 98.88 | 20.87 | 38.592 | 89.7049 | 115.3309 | 82.8633 | 0.00964 | 0.99166 | 0.023425 | 23.6216 | 0.36922 | 0.1627 | 0.013799 | 0.9677 | 0.9677 | 0.00691
2019-03-21 15:52 | 1553197901.867 | 111 | 0.20672 | 0.002537 | 2.1e-05 | 100.7549 | 96.9376 | 115.7719 | 21.2612 | 0.002805 | 0.002552 | 258.5256 | 0.27406 | 8.4035 | 0.003796 | 0.00523 | 0.005118 | 4.9e-05 | 0.001111 | 0.25066 | 1.3357 | 2.5481 | 98.88 | 20.87 | 38.59 | 89.2955 | 113.2784 | 83.2081 | 0.00964 | 0.9907 | 0.023426 | 23.6459 | 0.34547 | 0.16304 | 0.013924 | 0.97648 | 0.97648 | 0.006862
2019-03-21 15:52 | 1553197902.8667 | 112 | 0.20673 | 0.00254 | 2.2e-05 | 100.9045 | 97.0588 | 115.7815 | 21.2612 | 0.002812 | 0.002523 | 258.4788 | 0.27319 | 8.4039 | 0.003827 | 0.005491 | 0.005501 | 4.9e-05 | 0.001111 | 0.27544 | 1.3142 | 2.5479 | 98.88 | 20.87 | 38.587 | 89.479 | 114.0428 | 82.5046 | 0.009639 | 0.99266 | 0.02345 | 23.6235 | 0.34756 | 0.16252 | 0.013906 | 0.97619 | 0.97619 | 0.006843
2019-03-21 15:52 | 1553197903.8658 | 113 | 0.20668 | 0.002551 | 2.3e-05 | 100.8401 | 96.9413 | 115.7625 | 21.2612 | 0.002813 | 0.002568 | 258.4442 | 0.27403 | 8.4002 | 0.003848 | 0.005421 | 0.005311 | 5.1e-05 | 0.001114 | 0.26377 | 1.3549 | 2.548 | 98.88 | 20.88 | 38.585 | 89.4828 | 114.544 | 83.606 | 0.009644 | 0.9896 | 0.023412 | 23.6582 | 0.35269 | 0.16302 | 0.014159 | 0.99224 | 0.99224 | 0.006753
2019-03-21 15:52 | 1553197904.8658 | 114 | 0.20667 | 0.00256 | 2.4e-05 | 100.9093 | 96.7397 | 115.7672 | 21.2612 | 0.002781 | 0.002532 | 258.3869 | 0.27245 | 8.4035 | 0.003811 | 0.005613 | 0.005559 | 5.1e-05 | 0.001109 | 0.26305 | 1.348 | 2.5479 | 98.88 | 20.88 | 38.58 | 88.715 | 113.6586 | 82.7198 | 0.009643 | 0.99206 | 0.023442 | 23.6292 | 0.37744 | 0.16208 | 0.014226 | 0.99814 | 0.99814 | 0.006674
2019-03-21 15:52 | 1553197905.8664 | 115 | 0.20664 | 0.00257 | 2.4e-05 | 100.816 | 96.5868 | 115.6327 | 21.2612 | 0.002844 | 0.002577 | 258.3397 | 0.27367 | 8.4037 | 0.003797 | 0.005426 | 0.005428 | 4.9e-05 | 0.001105 | 0.25194 | 1.3313 | 2.5479 | 98.88 | 20.88 | 38.577 | 90.2323 | 113.3093 | 83.8104 | 0.009642 | 0.98903 | 0.023392 | 23.6517 | 0.38332 | 0.16281 | 0.014349 | 1.0046 | 1.0046 | 0.00666
2019-03-21 15:52 | 1553197906.8673 | 116 | 0.20663 | 0.002583 | 2.5e-05 | 100.8473 | 96.8837 | 115.5796 | 21.2612 | 0.002768 | 0.002495 | 258.2701 | 0.2736 | 8.4035 | 0.003781 | 0.005079 | 0.005072 | 5.3e-05 | 0.001108 | 0.2557 | 1.2843 | 2.5478 | 98.88 | 20.88 | 38.57 | 88.4136 | 112.9292 | 81.8315 | 0.009641 | 0.99455 | 0.023452 | 23.5806 | 0.35864 | 0.16277 | 0.014436 | 1.0133 | 1.0133 | 0.006602
2019-03-21 15:52 | 1553197907.8669 | 117 | 0.2066 | 0.002592 | 2.5e-05 | 100.8586 | 96.5506 | 115.6379 | 21.2612 | 0.002824 | 0.002626 | 258.2167 | 0.27404 | 8.4036 | 0.003747 | 0.005689 | 0.00559 | 5.1e-05 | 0.001112 | 0.27295 | 1.297 | 2.548 | 98.88 | 20.88 | 38.567 | 89.7665 | 112.0945 | 85.0003 | 0.00964 | 0.98575 | 0.023354 | 23.6916 | 0.39045 | 0.16303 | 0.014603 | 1.0207 | 1.0207 | 0.006564
2019-03-21 15:52 | 1553197908.867 | 118 | 0.20659 | 0.002607 | 2.6e-05 | 100.9152 | 96.8657 | 116.0399 | 21.2612 | 0.002802 | 0.002545 | 258.1851 | 0.27405 | 8.4017 | 0.00382 | 0.005062 | 0.004998 | 5e-05 | 0.001113 | 0.26233 | 1.2688 | 2.5479 | 98.88 | 20.88 | 38.564 | 89.2322 | 113.8805 | 83.0523 | 0.009639 | 0.99114 | 0.023458 | 23.6681 | 0.36633 | 0.16303 | 0.014644 | 1.0281 | 1.0281 | 0.006518
2019-03-21 15:52 | 1553197909.866 | 119 | 0.20657 | 0.002619 | 2.6e-05 | 100.7799 | 96.833 | 116.3639 | 21.2612 | 0.002838 | 0.002612 | 258.1254 | 0.27275 | 8.4041 | 0.003842 | 0.004895 | 0.004801 | 5.1e-05 | 0.001111 | 0.27256 | 1.2707 | 2.548 | 98.88 | 20.88 | 38.558 | 90.0948 | 114.4078 | 84.661 | 0.009638 | 0.98668 | 0.023438 | 23.7546 | 0.35735 | 0.16226 | 0.014751 | 1.0347 | 1.0347 | 0.006445
2019-03-21 15:52 | 1553197910.8655 | 120 | 0.20654 | 0.002633 | 2.7e-05 | 100.8163 | 97.0193 | 116.6183 | 21.2612 | 0.002817 | 0.002607 | 258.0796 | 0.27316 | 8.4019 | 0.003763 | 0.005142 | 0.005156 | 4.6e-05 | 0.001112 | 0.2688 | 1.3116 | 2.5481 | 98.88 | 20.88 | 38.555 | 89.5931 | 112.471 | 84.5422 | 0.009637 | 0.98701 | 0.023468 | 23.7766 | 0.34338 | 0.1625 | 0.014885 | 1.0454 | 1.0454 | 0.006389
2019-03-21 15:52 | 1553197911.8661 | 121 | 0.20652 | 0.002643 | 2.7e-05 | 100.8553 | 97.0166 | 116.261 | 21.2612 | 0.00285 | 0.002552 | 258.0132 | 0.27319 | 8.4018 | 0.003847 | 0.005363 | 0.005315 | 5.1e-05 | 0.001112 | 0.27671 | 1.2815 | 2.5479 | 98.88 | 20.88 | 38.548 | 90.3776 | 114.5314 | 83.2098 | 0.009635 | 0.9907 | 0.023475 | 23.6959 | 0.3471 | 0.16252 | 0.015005 | 1.0541 | 1.0541 | 0.006337
2019-03-21 15:52 | 1553197912.8652 | 122 | 0.20653 | 0.002653 | 2.7e-05 | 100.8315 | 96.6692 | 115.8531 | 21.2612 | 0.002833 | 0.002576 | 257.9518 | 0.27328 | 8.4016 | 0.003834 | 0.00539 | 0.005308 | 5e-05 | 0.001113 | 0.25857 | 1.3045 | 2.5479 | 98.88 | 20.88 | 38.545 | 89.9823 | 114.2182 | 83.7871 | 0.009634 | 0.9891 | 0.023415 | 23.6734 | 0.37707 | 0.16257 | 0.014916 | 1.0452 | 1.0452 | 0.006393
2019-03-21 15:52 | 1553197913.8656 | 123 | 0.2065 | 0.002665 | 2.7e-05 | 100.8595 | 96.6654 | 115.5557 | 21.2612 | 0.002734 | 0.002595 | 257.8969 | 0.27386 | 8.4037 | 0.003847 | 0.005427 | 0.005316 | 5.1e-05 | 0.001112 | 0.25337 | 1.307 | 2.548 | 98.88 | 20.88 | 38.543 | 87.5872 | 114.5384 | 84.2543 | 0.009634 | 0.9878 | 0.02337 | 23.6585 | 0.3799 | 0.16292 | 0.015098 | 1.0559 | 1.0559 | 0.006341
2019-03-21 15:52 | 1553197914.8667 | 124 | 0.20652 | 0.002674 | 2.7e-05 | 100.7829 | 97.0128 | 115.8184 | 21.2612 | 0.002793 | 0.002576 | 257.839 | 0.27329 | 8.4011 | 0.003809 | 0.005133 | 0.005084 | 4.9e-05 | 0.001114 | 0.27527 | 1.2851 | 2.5479 | 98.88 | 20.88 | 38.538 | 89.008 | 113.6009 | 83.7926 | 0.009633 | 0.98908 | 0.023412 | 23.6701 | 0.34102 | 0.16258 | 0.01496 | 1.0481 | 1.0481 | 0.006375
2019-03-21 15:52 | 1553197915.8655 | 125 | 0.20654 | 0.002689 | 2.7e-05 | 100.8014 | 97.1157 | 115.9151 | 21.2612 | 0.002833 | 0.002584 | 257.8054 | 0.27371 | 8.403 | 0.003848 | 0.005619 | 0.005551 | 5.2e-05 | 0.001112 | 0.27312 | 1.3106 | 2.548 | 98.88 | 20.88 | 38.535 | 89.9698 | 114.543 | 83.9727 | 0.009632 | 0.98858 | 0.023415 | 23.6859 | 0.33318 | 0.16283 | 0.014815 | 1.0381 | 1.0381 | 0.006447
2019-03-21 15:52 | 1553197916.8655 | 126 | 0.20653 | 0.002692 | 2.7e-05 | 100.8649 | 97.0384 | 115.9698 | 21.2612 | 0.002839 | 0.002587 | 257.7514 | 0.27449 | 8.4032 | 0.003843 | 0.005623 | 0.005634 | 5e-05 | 0.001113 | 0.25553 | 1.2694 | 2.5479 | 98.88 | 20.87 | 38.532 | 90.1199 | 114.4253 | 84.0547 | 0.009625 | 0.98836 | 0.023418 | 23.6942 | 0.34593 | 0.1633 | 0.014905 | 1.0446 | 1.0446 | 0.006425
2019-03-21 15:52 | 1553197917.8655 | 127 | 0.20649 | 0.002704 | 2.8e-05 | 100.8151 | 97.0329 | 115.6268 | 21.2612 | 0.002865 | 0.002534 | 257.7277 | 0.27401 | 8.403 | 0.003758 | 0.005748 | 0.005704 | 4.9e-05 | 0.001117 | 0.28119 | 1.3197 | 2.5479 | 98.88 | 20.88 | 38.53 | 90.7481 | 112.3683 | 82.7878 | 0.009631 | 0.99187 | 0.023425 | 23.6172 | 0.34202 | 0.16301 | 0.015091 | 1.0578 | 1.0578 | 0.006333
2019-03-21 15:52 | 1553197918.8655 | 128 | 0.20649 | 0.002708 | 2.9e-05 | 100.8166 | 97.0166 | 115.1616 | 21.2612 | 0.002815 | 0.002593 | 257.6648 | 0.2743 | 8.4036 | 0.003719 | 0.005476 | 0.00532 | 4.9e-05 | 0.001114 | 0.27383 | 1.3108 | 2.5481 | 98.88 | 20.88 | 38.528 | 89.5378 | 111.403 | 84.1959 | 0.00963 | 0.98796 | 0.023332 | 23.6162 | 0.34366 | 0.16318 | 0.015082 | 1.053 | 1.053 | 0.006369
2019-03-21 15:52 | 1553197919.867 | 129 | 0.20648 | 0.002716 | 3e-05 | 100.8422 | 97.2042 | 114.7246 | 21.2612 | 0.002834 | 0.002628 | 257.6189 | 0.27383 | 8.4018 | 0.003716 | 0.005412 | 0.005381 | 5.1e-05 | 0.001114 | 0.27433 | 1.2976 | 2.5479 | 98.88 | 20.88 | 38.525 | 90.0042 | 111.332 | 85.0285 | 0.009629 | 0.98567 | 0.023261 | 23.5988 | 0.32865 | 0.1629 | 0.015173 | 1.056 | 1.056 | 0.00634
2019-03-21 15:52 | 1553197920.867 | 130 | 0.20647 | 0.002725 | 3.1e-05 | 100.7615 | 96.9784 | 114.722 | 21.2612 | 0.002833 | 0.002552 | 257.5563 | 0.27387 | 8.4038 | 0.00378 | 0.005203 | 0.0051 | 4.9e-05 | 0.001114 | 0.27367 | 1.2848 | 2.5479 | 98.88 | 20.87 | 38.525 | 89.9775 | 112.8873 | 83.2237 | 0.009623 | 0.99066 | 0.023319 | 23.539 | 0.34229 | 0.16293 | 0.015223 | 1.0621 | 1.0621 | 0.006305
2019-03-21 15:52 | 1553197921.8656 | 131 | 0.20647 | 0.002735 | 3.1e-05 | 100.8223 | 96.9198 | 115.0069 | 21.2612 | 0.002824 | 0.002568 | 257.5486 | 0.27362 | 8.4041 | 0.003743 | 0.005609 | 0.005554 | 5.1e-05 | 0.001112 | 0.24331 | 1.2773 | 2.5479 | 98.88 | 20.87 | 38.523 | 89.7651 | 111.9901 | 83.5959 | 0.009623 | 0.98963 | 0.023336 | 23.5805 | 0.35308 | 0.16278 | 0.01519 | 1.0605 | 1.0605 | 0.006308
2019-03-21 15:52 | 1553197922.8659 | 132 | 0.20651 | 0.002741 | 3.2e-05 | 100.8464 | 96.9226 | 115.1433 | 21.2612 | 0.002764 | 0.002585 | 257.4772 | 0.27357 | 8.405 | 0.003745 | 0.005008 | 0.005025 | 4.8e-05 | 0.001113 | 0.28368 | 1.2907 | 2.548 | 98.88 | 20.87 | 38.522 | 88.3221 | 112.0446 | 83.9992 | 0.009623 | 0.98851 | 0.023337 | 23.6078 | 0.35497 | 0.16274 | 0.014936 | 1.0429 | 1.0429 | 0.006414
2019-03-21 15:52 | 1553197923.8658 | 133 | 0.20651 | 0.002747 | 3.1e-05 | 100.7665 | 97.2454 | 115.1419 | 21.2612 | 0.002841 | 0.002585 | 257.424 | 0.27304 | 8.402 | 0.003743 | 0.004922 | 0.004858 | 4.5e-05 | 0.001109 | 0.2808 | 1.2924 | 2.5479 | 98.88 | 20.87 | 38.522 | 90.1635 | 112.002 | 84.0148 | 0.009623 | 0.98847 | 0.023336 | 23.6082 | 0.31814 | 0.16243 | 0.01493 | 1.0424 | 1.0424 | 0.006405
2019-03-21 15:52 | 1553197924.8672 | 134 | 0.2065 | 0.002755 | 3.1e-05 | 100.8181 | 97.2008 | 114.6952 | 21.2612 | 0.002815 | 0.002591 | 257.354 | 0.2734 | 8.4021 | 0.00378 | 0.005838 | 0.005792 | 4.7e-05 | 0.001113 | 0.26488 | 1.3113 | 2.5479 | 98.88 | 20.87 | 38.52 | 89.5361 | 112.8937 | 84.1483 | 0.009622 | 0.9881 | 0.023286 | 23.5667 | 0.32682 | 0.16264 | 0.014975 | 1.0433 | 1.0433 | 0.006407
2019-03-21 15:52 | 1553197925.8659 | 135 | 0.20648 | 0.00276 | 3.1e-05 | 100.8935 | 97.0707 | 114.8013 | 21.2612 | 0.002889 | 0.002594 | 257.3123 | 0.27321 | 8.403 | 0.003742 | 0.005066 | 0.005052 | 4.8e-05 | 0.001111 | 0.27328 | 1.2948 | 2.548 | 98.88 | 20.87 | 38.52 | 91.3273 | 111.9736 | 84.2153 | 0.009622 | 0.98791 | 0.023295 | 23.5798 | 0.34549 | 0.16253 | 0.015123 | 1.054 | 1.054 | 0.006338
2019-03-21 15:52 | 1553197926.8659 | 136 | 0.20647 | 0.002772 | 3.1e-05 | 100.881 | 97.3401 | 114.8548 | 21.2612 | 0.00286 | 0.002551 | 257.2821 | 0.27282 | 8.4006 | 0.003743 | 0.005523 | 0.005425 | 4.8e-05 | 0.001112 | 0.28655 | 1.3437 | 2.548 | 98.88 | 20.87 | 38.519 | 90.6144 | 111.9838 | 83.1896 | 0.009622 | 0.99076 | 0.023334 | 23.5515 | 0.31959 | 0.1623 | 0.015171 | 1.0592 | 1.0592 | 0.006298
2019-03-21 15:52 | 1553197927.866 | 137 | 0.20646 | 0.002778 | 3.1e-05 | 100.8532 | 97.5074 | 115.075 | 21.2612 | 0.002884 | 0.002618 | 257.2076 | 0.27326 | 8.4016 | 0.003779 | 0.005012 | 0.004959 | 5.1e-05 | 0.001109 | 0.27959 | 1.3297 | 2.5479 | 98.88 | 20.87 | 38.52 | 91.2133 | 112.8796 | 84.7928 | 0.009622 | 0.98632 | 0.023304 | 23.627 | 0.30178 | 0.16256 | 0.015204 | 1.0601 | 1.0601 | 0.006303
2019-03-21 15:52 | 1553197928.8661 | 138 | 0.20645 | 0.002782 | 3.2e-05 | 100.8026 | 97.6198 | 115.297 | 21.2612 | 0.002804 | 0.002625 | 257.1349 | 0.27312 | 8.4006 | 0.003702 | 0.004777 | 0.004775 | 4.7e-05 | 0.001104 | 0.26089 | 1.3328 | 2.548 | 98.88 | 20.87 | 38.52 | 89.2719 | 110.9948 | 84.959 | 0.009622 | 0.98586 | 0.023321 | 23.6553 | 0.28698 | 0.16248 | 0.015225 | 1.0623 | 1.0623 | 0.006286
2019-03-21 15:52 | 1553197929.8651 | 139 | 0.20645 | 0.002783 | 3.2e-05 | 100.8467 | 97.4556 | 115.2001 | 21.2612 | 0.002864 | 0.002639 | 257.0816 | 0.2736 | 8.4015 | 0.003693 | 0.005363 | 0.005284 | 4.9e-05 | 0.001103 | 0.2656 | 1.3429 | 2.548 | 98.88 | 20.87 | 38.518 | 90.7177 | 110.7857 | 85.3008 | 0.009622 | 0.98492 | 0.0233 | 23.6566 | 0.30595 | 0.16277 | 0.015219 | 1.0609 | 1.0609 | 0.006306
2019-03-21 15:52 | 1553197930.8651 | 140 | 0.20649 | 0.002794 | 3.3e-05 | 100.8827 | 97.2567 | 114.8121 | 21.2612 | 0.002876 | 0.002596 | 257.007 | 0.27291 | 8.4005 | 0.003656 | 0.005546 | 0.005405 | 4.9e-05 | 0.001102 | 0.25586 | 1.3412 | 2.5479 | 98.88 | 20.87 | 38.516 | 91.0044 | 109.8761 | 84.2769 | 0.009621 | 0.98774 | 0.023294 | 23.583 | 0.32742 | 0.16235 | 0.014998 | 1.0452 | 1.0452 | 0.006384
2019-03-21 15:52 | 1553197931.8667 | 141 | 0.20648 | 0.002798 | 3.3e-05 | 100.7692 | 97.2649 | 114.7635 | 21.2612 | 0.002845 | 0.002611 | 256.9321 | 0.27341 | 8.4036 | 0.003645 | 0.00496 | 0.004981 | 4.7e-05 | 0.001102 | 0.2625 | 1.335 | 2.5481 | 98.88 | 20.87 | 38.514 | 90.2714 | 109.6081 | 84.636 | 0.009621 | 0.98675 | 0.023277 | 23.5898 | 0.31659 | 0.16265 | 0.015039 | 1.0472 | 1.0472 | 0.006383
2019-03-21 15:52 | 1553197932.8653 | 142 | 0.20647 | 0.002803 | 3.4e-05 | 100.85 | 97.4287 | 114.8555 | 21.2612 | 0.002864 | 0.002638 | 256.8864 | 0.27278 | 8.4006 | 0.003662 | 0.005274 | 0.00519 | 5e-05 | 0.001102 | 0.26576 | 1.3215 | 2.5478 | 98.88 | 20.87 | 38.512 | 90.7271 | 110.0222 | 85.2858 | 0.00962 | 0.98496 | 0.023265 | 23.6207 | 0.30871 | 0.16228 | 0.015095 | 1.0506 | 1.0506 | 0.006348
2019-03-21 15:52 | 1553197933.8653 | 143 | 0.2065 | 0.002812 | 3.6e-05 | 100.8562 | 97.3507 | 114.9982 | 21.2612 | 0.002863 | 0.002593 | 256.8295 | 0.27283 | 8.4009 | 0.003682 | 0.005124 | 0.005021 | 5e-05 | 0.001102 | 0.2568 | 1.2991 | 2.5481 | 98.88 | 20.87 | 38.511 | 90.6973 | 110.5089 | 84.2045 | 0.00962 | 0.98794 | 0.023315 | 23.5997 | 0.31642 | 0.16231 | 0.014912 | 1.04 | 1.04 | 0.006414
2019-03-21 15:52 | 1553197934.8653 | 144 | 0.20645 | 0.002819 | 3.6e-05 | 100.8964 | 97.4086 | 115.2808 | 21.2612 | 0.00286 | 0.00261 | 256.7753 | 0.27282 | 8.4017 | 0.003642 | 0.005078 | 0.004998 | 4.9e-05 | 0.0011 | 0.24889 | 1.2806 | 2.5478 | 98.88 | 20.87 | 38.51 | 90.6161 | 109.5475 | 84.6054 | 0.00962 | 0.98683 | 0.023331 | 23.6419 | 0.31467 | 0.1623 | 0.015194 | 1.0603 | 1.0603 | 0.006291
2019-03-21 15:52 | 1553197935.8664 | 145 | 0.20651 | 0.002821 | 3.7e-05 | 100.7633 | 97.5479 | 115.0029 | 21.2612 | 0.002922 | 0.00262 | 256.7048 | 0.27332 | 8.4019 | 0.00365 | 0.00538 | 0.005267 | 4.9e-05 | 0.0011 | 0.24441 | 1.2752 | 2.5479 | 98.88 | 20.87 | 38.511 | 92.1296 | 109.7366 | 84.8418 | 0.00962 | 0.98618 | 0.023295 | 23.6212 | 0.29008 | 0.1626 | 0.01484 | 1.034 | 1.034 | 0.006463
2019-03-21 15:52 | 1553197936.8663 | 146 | 0.20648 | 0.002826 | 3.8e-05 | 100.8607 | 97.7621 | 114.875 | 21.2612 | 0.002909 | 0.002631 | 256.6299 | 0.27359 | 8.4025 | 0.003687 | 0.005336 | 0.005299 | 5.2e-05 | 0.001098 | 0.25393 | 1.2873 | 2.5479 | 98.88 | 20.87 | 38.511 | 91.8076 | 110.6296 | 85.1151 | 0.00962 | 0.98543 | 0.023273 | 23.6171 | 0.2791 | 0.16276 | 0.015024 | 1.0457 | 1.0457 | 0.006397
2019-03-21 15:52 | 1553197937.8655 | 147 | 0.20654 | 0.00283 | 3.9e-05 | 100.805 | 97.9096 | 115.0032 | 21.2612 | 0.002893 | 0.002665 | 256.5869 | 0.27365 | 8.4015 | 0.003663 | 0.005048 | 0.004999 | 4.4e-05 | 0.001094 | 0.27079 | 1.2967 | 2.548 | 98.87 | 20.86 | 38.512 | 91.4133 | 110.0609 | 85.9243 | 0.009615 | 0.98321 | 0.02326 | 23.6569 | 0.26067 | 0.16279 | 0.014618 | 1.0169 | 1.0169 | 0.00658
2019-03-21 15:52 | 1553197938.8659 | 148 | 0.2065 | 0.002826 | 4e-05 | 100.8947 | 97.9498 | 115.0438 | 21.2612 | 0.002863 | 0.002611 | 256.5286 | 0.27258 | 8.4012 | 0.003631 | 0.005658 | 0.005576 | 5e-05 | 0.001097 | 0.23756 | 1.3182 | 2.548 | 98.87 | 20.86 | 38.513 | 90.7076 | 109.2817 | 84.6328 | 0.009615 | 0.98676 | 0.023306 | 23.6185 | 0.26496 | 0.16216 | 0.01488 | 1.0371 | 1.0371 | 0.006426
2019-03-21 15:52 | 1553197939.8666 | 149 | 0.20657 | 0.002829 | 4.1e-05 | 100.8208 | 97.8388 | 115.1752 | 21.2612 | 0.002943 | 0.002663 | 256.4599 | 0.27354 | 8.4022 | 0.003706 | 0.005357 | 0.005215 | 4.6e-05 | 0.001095 | 0.2688 | 1.2833 | 2.548 | 98.87 | 20.86 | 38.514 | 92.6121 | 111.095 | 85.8738 | 0.009616 | 0.98335 | 0.023279 | 23.673 | 0.26855 | 0.16273 | 0.014457 | 1.0064 | 1.0064 | 0.006646
2019-03-21 15:52 | 1553197940.8666 | 150 | 0.20658 | 0.002829 | 4.1e-05 | 100.8503 | 97.9924 | 115.121 | 21.2612 | 0.002879 | 0.002675 | 256.4043 | 0.27281 | 8.4026 | 0.003642 | 0.005364 | 0.005307 | 5.1e-05 | 0.001093 | 0.27903 | 1.338 | 2.5479 | 98.87 | 20.86 | 38.518 | 91.0765 | 109.5458 | 86.1718 | 0.009617 | 0.98253 | 0.023264 | 23.6772 | 0.25713 | 0.16229 | 0.014391 | 1.0011 | 1.0011 | 0.006663
2019-03-21 15:52 | 1553197941.8657 | 151 | 0.20656 | 0.002825 | 4.2e-05 | 100.8708 | 98.0213 | 114.9027 | 21.2612 | 0.002925 | 0.002634 | 256.3566 | 0.27409 | 8.4016 | 0.003659 | 0.005911 | 0.005869 | 4.5e-05 | 0.001092 | 0.27041 | 1.2851 | 2.548 | 98.87 | 20.86 | 38.521 | 92.1922 | 109.9527 | 85.1829 | 0.009617 | 0.98524 | 0.023274 | 23.6222 | 0.25631 | 0.16306 | 0.01452 | 1.0105 | 1.0105 | 0.006632
2019-03-21 15:52 | 1553197942.8657 | 152 | 0.20661 | 0.002826 | 4.2e-05 | 100.8643 | 98.5741 | 114.5999 | 21.2612 | 0.002923 | 0.002655 | 256.2971 | 0.27304 | 8.4027 | 0.003615 | 0.005696 | 0.005558 | 4.5e-05 | 0.001095 | 0.26128 | 1.3251 | 2.5481 | 98.87 | 20.86 | 38.527 | 92.1366 | 108.8796 | 85.6965 | 0.009619 | 0.98383 | 0.023226 | 23.6079 | 0.20543 | 0.16243 | 0.014252 | 0.98975 | 0.98975 | 0.006745
2019-03-21 15:52 | 1553197943.8648 | 153 | 0.20658 | 0.002826 | 4.2e-05 | 100.7898 | 98.6684 | 114.7201 | 21.2612 | 0.002879 | 0.002637 | 256.2569 | 0.27336 | 8.4025 | 0.003648 | 0.004104 | 0.004107 | 4.7e-05 | 0.001097 | 0.26808 | 1.3258 | 2.548 | 98.87 | 20.86 | 38.531 | 91.0803 | 109.6878 | 85.2555 | 0.00962 | 0.98504 | 0.023253 | 23.6058 | 0.19027 | 0.16262 | 0.01441 | 1.0019 | 1.0019 | 0.006671
2019-03-21 15:52 | 1553197944.866 | 154 | 0.2066 | 0.002822 | 4.3e-05 | 100.7737 | 99.0434 | 114.8335 | 21.2612 | 0.002912 | 0.002647 | 256.2012 | 0.27312 | 8.4028 | 0.003653 | 0.004835 | 0.004789 | 4.4e-05 | 0.00109 | 0.24563 | 1.3099 | 2.5479 | 98.87 | 20.86 | 38.536 | 91.8825 | 109.8076 | 85.4966 | 0.009621 | 0.98438 | 0.023256 | 23.6254 | 0.15491 | 0.16248 | 0.014284 | 0.99324 | 0.99324 | 0.006723
2019-03-21 15:52 | 1553197945.8649 | 155 | 0.20661 | 0.002809 | 4.2e-05 | 100.8193 | 98.9569 | 115.2412 | 21.2612 | 0.0029 | 0.002607 | 256.15 | 0.27384 | 8.4005 | 0.003583 | 0.005448 | 0.005306 | 4.7e-05 | 0.001087 | 0.23706 | 1.3257 | 2.5481 | 98.87 | 20.86 | 38.544 | 91.5878 | 108.0949 | 84.5359 | 0.009623 | 0.98703 | 0.023329 | 23.6356 | 0.16677 | 0.16291 | 0.014254 | 0.99427 | 0.99427 | 0.006734
2019-03-21 15:52 | 1553197946.866 | 156 | 0.2066 | 0.002808 | 4.3e-05 | 100.763 | 99.0243 | 115.2576 | 21.2612 | 0.002926 | 0.002636 | 256.0884 | 0.27302 | 8.4014 | 0.003621 | 0.004692 | 0.004659 | 4.4e-05 | 0.001093 | 0.25083 | 1.3274 | 2.5481 | 98.87 | 20.86 | 38.548 | 92.2048 | 109.0161 | 85.2298 | 0.009624 | 0.98511 | 0.023308 | 23.6602 | 0.15568 | 0.16242 | 0.014328 | 0.9985 | 0.9985 | 0.006685
2019-03-21 15:52 | 1553197947.865 | 157 | 0.2066 | 0.002799 | 4.4e-05 | 100.8124 | 99.0679 | 115.0081 | 21.2612 | 0.002908 | 0.002651 | 256.0323 | 0.27259 | 8.4013 | 0.003619 | 0.00569 | 0.005589 | 4.9e-05 | 0.001092 | 0.26681 | 1.3394 | 2.548 | 98.87 | 20.86 | 38.554 | 91.7693 | 108.9901 | 85.5817 | 0.009626 | 0.98415 | 0.023271 | 23.6461 | 0.15613 | 0.16216 | 0.014316 | 0.99602 | 0.99602 | 0.006692
2019-03-21 15:52 | 1553197948.8651 | 158 | 0.20663 | 0.002794 | 4.4e-05 | 100.7901 | 98.8411 | 114.7584 | 21.2612 | 0.002879 | 0.002602 | 255.9777 | 0.27225 | 8.4019 | 0.003592 | 0.004735 | 0.004646 | 4.9e-05 | 0.001092 | 0.25912 | 1.3629 | 2.5481 | 98.87 | 20.86 | 38.558 | 91.082 | 108.3146 | 84.4174 | 0.009627 | 0.98735 | 0.023284 | 23.5821 | 0.17466 | 0.16196 | 0.01412 | 0.98289 | 0.98289 | 0.006773
2019-03-21 15:52 | 1553197949.8655 | 159 | 0.20664 | 0.002792 | 4.5e-05 | 100.6733 | 98.978 | 114.8698 | 21.2612 | 0.002904 | 0.002641 | 255.9285 | 0.27299 | 8.406 | 0.003659 | 0.00559 | 0.005559 | 4.5e-05 | 0.001092 | 0.29087 | 1.3018 | 2.5481 | 98.87 | 20.86 | 38.562 | 91.6871 | 109.9512 | 85.3616 | 0.009628 | 0.98475 | 0.023264 | 23.6247 | 0.1519 | 0.1624 | 0.01406 | 0.97783 | 0.97783 | 0.006826
2019-03-21 15:52 | 1553197950.8662 | 160 | 0.20667 | 0.002788 | 4.7e-05 | 100.8026 | 99.1803 | 115.0069 | 21.2612 | 0.002867 | 0.00269 | 255.8568 | 0.27342 | 8.4019 | 0.003649 | 0.005402 | 0.005376 | 4.6e-05 | 0.001091 | 0.25442 | 1.3234 | 2.5479 | 98.87 | 20.86 | 38.568 | 90.7982 | 109.7028 | 86.5325 | 0.009629 | 0.98155 | 0.02324 | 23.6773 | 0.14512 | 0.16265 | 0.013872 | 0.96363 | 0.96363 | 0.006937
2019-03-21 15:53 | 1553197951.8653 | 161 | 0.20668 | 0.00278 | 4.9e-05 | 100.8214 | 99.2025 | 115.1032 | 21.2612 | 0.002884 | 0.002691 | 255.795 | 0.2739 | 8.4025 | 0.003658 | 0.00523 | 0.005133 | 4.7e-05 | 0.001088 | 0.25354 | 1.2931 | 2.5479 | 98.87 | 20.86 | 38.571 | 91.2032 | 109.9316 | 86.5433 | 0.00963 | 0.98152 | 0.02325 | 23.6876 | 0.14479 | 0.16295 | 0.013837 | 0.96147 | 0.96147 | 0.006965
2019-03-21 15:53 | 1553197952.8663 | 162 | 0.20666 | 0.002778 | 5.1e-05 | 100.8223 | 99.3465 | 115.148 | 21.2612 | 0.002899 | 0.002654 | 255.7228 | 0.27286 | 8.4009 | 0.003635 | 0.004829 | 0.004805 | 4.6e-05 | 0.00109 | 0.25138 | 1.2935 | 2.5479 | 98.87 | 20.86 | 38.575 | 91.5753 | 109.3716 | 85.6621 | 0.009631 | 0.98393 | 0.023283 | 23.6632 | 0.13189 | 0.16232 | 0.013935 | 0.96958 | 0.96958 | 0.006881
2019-03-21 15:53 | 1553197953.8664 | 163 | 0.20666 | 0.002765 | 5.2e-05 | 100.7862 | 99.6296 | 115.0849 | 21.2612 | 0.002893 | 0.002682 | 255.6338 | 0.27385 | 8.4009 | 0.00359 | 0.005403 | 0.005281 | 4.7e-05 | 0.00109 | 0.26609 | 1.2676 | 2.548 | 98.87 | 20.85 | 38.58 | 91.4189 | 108.2777 | 86.3489 | 0.009626 | 0.98205 | 0.023254 | 23.6793 | 0.10324 | 0.16291 | 0.014003 | 0.97298 | 0.97298 | 0.006882
2019-03-21 15:53 | 1553197954.8658 | 164 | 0.20663 | 0.00276 | 5.4e-05 | 100.7558 | 99.4882 | 114.8654 | 21.2612 | 0.002883 | 0.002614 | 255.5774 | 0.27372 | 8.4021 | 0.003601 | 0.004906 | 0.004857 | 4.7e-05 | 0.001089 | 0.27527 | 1.3314 | 2.548 | 98.87 | 20.85 | 38.583 | 91.1851 | 108.5322 | 84.692 | 0.009627 | 0.9866 | 0.023286 | 23.6021 | 0.11324 | 0.16284 | 0.014141 | 0.98383 | 0.98383 | 0.006803
2019-03-21 15:53 | 1553197955.8655 | 165 | 0.20666 | 0.002758 | 5.6e-05 | 100.7192 | 99.4596 | 114.4577 | 21.2612 | 0.002904 | 0.002604 | 255.5099 | 0.27332 | 8.4029 | 0.003616 | 0.004759 | 0.004745 | 4.6e-05 | 0.001087 | 0.25105 | 1.3703 | 2.548 | 98.87 | 20.85 | 38.588 | 91.6729 | 108.8971 | 84.4635 | 0.009628 | 0.98723 | 0.023252 | 23.5527 | 0.11256 | 0.1626 | 0.013996 | 0.97223 | 0.97223 | 0.006874
2019-03-21 15:53 | 1553197956.8646 | 166 | 0.20666 | 0.002757 | 5.6e-05 | 100.7418 | 99.3585 | 114.2741 | 21.2612 | 0.00291 | 0.002634 | 255.4374 | 0.27257 | 8.4016 | 0.003566 | 0.004573 | 0.004498 | 4.6e-05 | 0.001088 | 0.27455 | 1.3659 | 2.5481 | 98.87 | 20.85 | 38.59 | 91.8191 | 107.6853 | 85.1899 | 0.009629 | 0.98523 | 0.02321 | 23.5577 | 0.12367 | 0.16215 | 0.01399 | 0.97004 | 0.97004 | 0.00687
2019-03-21 15:53 | 1553197957.866 | 167 | 0.20665 | 0.002754 | 5.7e-05 | 100.7966 | 99.4422 | 114.5842 | 21.2612 | 0.002876 | 0.002682 | 255.3702 | 0.27327 | 8.4046 | 0.003571 | 0.004631 | 0.004617 | 4.8e-05 | 0.001088 | 0.25946 | 1.3172 | 2.5479 | 98.87 | 20.85 | 38.592 | 91.0211 | 107.8222 | 86.3475 | 0.009629 | 0.98205 | 0.023204 | 23.6277 | 0.121 | 0.16257 | 0.014036 | 0.97296 | 0.97296 | 0.006867
2019-03-21 15:53 | 1553197958.8647 | 168 | 0.20665 | 0.002746 | 5.7e-05 | 100.6617 | 99.3312 | 115.018 | 21.2612 | 0.002892 | 0.002667 | 255.3176 | 0.27435 | 8.4045 | 0.003596 | 0.004679 | 0.00461 | 4.7e-05 | 0.001086 | 0.24386 | 1.3065 | 2.548 | 98.87 | 20.84 | 38.596 | 91.4039 | 108.414 | 85.9666 | 0.009624 | 0.98309 | 0.02326 | 23.6598 | 0.11901 | 0.16321 | 0.014073 | 0.97785 | 0.97785 | 0.00686
2019-03-21 15:53 | 1553197959.8661 | 169 | 0.20673 | 0.002741 | 5.6e-05 | 100.7162 | 99.2481 | 115.0795 | 21.2612 | 0.002939 | 0.002687 | 255.2527 | 0.27261 | 8.4023 | 0.003587 | 0.004981 | 0.005016 | 4.6e-05 | 0.001086 | 0.28224 | 1.2809 | 2.548 | 98.87 | 20.84 | 38.597 | 92.5362 | 108.1935 | 86.4709 | 0.009625 | 0.98172 | 0.02325 | 23.6828 | 0.13134 | 0.16218 | 0.0136 | 0.94454 | 0.94454 | 0.007057
2019-03-21 15:53 | 1553197960.8647 | 170 | 0.20673 | 0.002744 | 5.5e-05 | 100.6706 | 99.5594 | 115.1151 | 21.2612 | 0.002912 | 0.002654 | 255.2276 | 0.27348 | 8.4011 | 0.003647 | 0.004596 | 0.004436 | 4.4e-05 | 0.001086 | 0.26233 | 1.2943 | 2.5479 | 98.87 | 20.84 | 38.6 | 91.8731 | 109.6688 | 85.6613 | 0.009625 | 0.98393 | 0.02328 | 23.6598 | 0.099279 | 0.16269 | 0.013572 | 0.9439 | 0.9439 | 0.007084
2019-03-21 15:53 | 1553197961.8648 | 171 | 0.20662 | 0.002739 | 5.5e-05 | 100.6989 | 99.6418 | 115.416 | 21.2612 | 0.002944 | 0.002671 | 255.1794 | 0.2726 | 8.4006 | 0.003578 | 0.005045 | 0.004974 | 4.7e-05 | 0.001085 | 0.26305 | 1.2879 | 2.5479 | 98.87 | 20.84 | 38.602 | 92.6473 | 107.9892 | 86.0696 | 0.009626 | 0.98281 | 0.023297 | 23.7041 | 0.094392 | 0.16217 | 0.014257 | 0.99237 | 0.99237 | 0.006717
2019-03-21 15:53 | 1553197962.8658 | 172 | 0.2066 | 0.002731 | 5.5e-05 | 100.6736 | 99.5087 | 115.6515 | 21.2612 | 0.002934 | 0.002656 | 255.1279 | 0.27278 | 8.4003 | 0.003623 | 0.004822 | 0.004815 | 4.9e-05 | 0.00108 | 0.24961 | 1.267 | 2.548 | 98.87 | 20.84 | 38.604 | 92.4027 | 109.065 | 85.7215 | 0.009626 | 0.98377 | 0.023332 | 23.7168 | 0.1041 | 0.16228 | 0.01438 | 1.0024 | 1.0024 | 0.006654
2019-03-21 15:53 | 1553197963.8659 | 173 | 0.20682 | 0.002722 | 5.8e-05 | 100.7412 | 99.3241 | 115.7423 | 21.2612 | 0.002921 | 0.002622 | 255.0765 | 0.27289 | 8.4046 | 0.003611 | 0.004753 | 0.004693 | 4.7e-05 | 0.001084 | 0.25249 | 1.2746 | 2.5479 | 98.87 | 20.84 | 38.606 | 92.0891 | 108.789 | 84.8849 | 0.009627 | 0.98606 | 0.023368 | 23.6985 | 0.12672 | 0.16234 | 0.013047 | 0.9106 | 0.9106 | 0.007327
2019-03-21 15:53 | 1553197964.8659 | 174 | 0.20669 | 0.002717 | 6e-05 | 100.7236 | 99.4661 | 115.863 | 21.2612 | 0.002881 | 0.002716 | 255.0083 | 0.27272 | 8.3991 | 0.003584 | 0.00516 | 0.005122 | 4.7e-05 | 0.001079 | 0.27057 | 1.3076 | 2.548 | 98.87 | 20.84 | 38.606 | 91.121 | 108.1318 | 87.161 | 0.009627 | 0.97984 | 0.023306 | 23.7861 | 0.11238 | 0.16224 | 0.013833 | 0.96293 | 0.96293 | 0.006925
2019-03-21 15:53 | 1553197965.8649 | 175 | 0.20664 | 0.002716 | 6.2e-05 | 100.7213 | 99.3653 | 116.2467 | 21.2612 | 0.002924 | 0.002723 | 254.9447 | 0.2734 | 8.4007 | 0.003608 | 0.00527 | 0.005182 | 5.1e-05 | 0.00108 | 0.28351 | 1.308 | 2.5479 | 98.87 | 20.84 | 38.606 | 92.1578 | 108.7236 | 87.3263 | 0.009627 | 0.97939 | 0.02334 | 23.8309 | 0.12124 | 0.16264 | 0.014118 | 0.98405 | 0.98405 | 0.006793
2019-03-21 15:53 | 1553197966.866 | 176 | 0.20651 | 0.002711 | 6.3e-05 | 100.7311 | 98.898 | 116.3899 | 21.2612 | 0.002954 | 0.002717 | 254.8698 | 0.27322 | 8.4003 | 0.003585 | 0.00574 | 0.005678 | 4.8e-05 | 0.001076 | 0.2724 | 1.3549 | 2.548 | 98.87 | 20.84 | 38.607 | 92.8961 | 108.1658 | 87.1836 | 0.009627 | 0.97977 | 0.023359 | 23.8408 | 0.16427 | 0.16254 | 0.014951 | 1.043 | 1.043 | 0.006405
2019-03-21 15:53 | 1553197967.8655 | 177 | 0.20647 | 0.002716 | 6.4e-05 | 100.6992 | 98.4651 | 116.2939 | 21.2612 | 0.002943 | 0.002757 | 254.793 | 0.27224 | 8.4052 | 0.003687 | 0.004964 | 0.004909 | 4.9e-05 | 0.001077 | 0.21604 | 1.3175 | 2.548 | 98.87 | 20.84 | 38.609 | 92.6317 | 110.6271 | 88.1346 | 0.009628 | 0.9772 | 0.023318 | 23.8624 | 0.20068 | 0.16196 | 0.015149 | 1.055 | 1.055 | 0.006309
2019-03-21 15:53 | 1553197968.8651 | 178 | 0.20651 | 0.002726 | 6.3e-05 | 100.7713 | 98.3193 | 116.1428 | 21.2612 | 0.002983 | 0.002637 | 254.7131 | 0.27238 | 8.4018 | 0.003705 | 0.005253 | 0.005172 | 5e-05 | 0.001077 | 0.27472 | 1.3054 | 2.5479 | 98.87 | 20.84 | 38.61 | 93.587 | 111.0729 | 85.2517 | 0.009628 | 0.98506 | 0.023397 | 23.7516 | 0.22033 | 0.16204 | 0.014935 | 1.0436 | 1.0436 | 0.006381
2019-03-21 15:53 | 1553197969.8661 | 179 | 0.20663 | 0.002743 | 6.2e-05 | 100.7353 | 98.1661 | 116.1118 | 21.2612 | 0.002987 | 0.002682 | 254.6254 | 0.27255 | 8.4019 | 0.003693 | 0.005129 | 0.005052 | 4.7e-05 | 0.001078 | 0.27671 | 1.3179 | 2.5479 | 98.87 | 20.84 | 38.612 | 93.6854 | 110.7881 | 86.3506 | 0.009628 | 0.98204 | 0.023358 | 23.7848 | 0.23109 | 0.16214 | 0.014166 | 0.98819 | 0.98819 | 0.006744
2019-03-21 15:53 | 1553197970.8662 | 180 | 0.20647 | 0.002747 | 6.2e-05 | 100.7531 | 98.3834 | 116.2174 | 21.2612 | 0.002907 | 0.00269 | 254.5421 | 0.27304 | 8.4016 | 0.003641 | 0.005215 | 0.005136 | 5e-05 | 0.001079 | 0.26792 | 1.3517 | 2.5481 | 98.87 | 20.84 | 38.615 | 91.7536 | 109.5057 | 86.5349 | 0.009629 | 0.98154 | 0.023362 | 23.8017 | 0.2129 | 0.16243 | 0.015157 | 1.0577 | 1.0577 | 0.006312
2019-03-21 15:53 | 1553197971.8663 | 181 | 0.20649 | 0.002762 | 6.1e-05 | 100.7034 | 98.4896 | 115.7684 | 21.2612 | 0.002945 | 0.002697 | 254.445 | 0.2721 | 8.4022 | 0.003717 | 0.005082 | 0.005068 | 5e-05 | 0.001078 | 0.30862 | 1.3654 | 2.5479 | 98.87 | 20.84 | 38.617 | 92.6833 | 111.3517 | 86.6939 | 0.00963 | 0.98111 | 0.023312 | 23.7609 | 0.19882 | 0.16187 | 0.014989 | 1.0437 | 1.0437 | 0.006374
2019-03-21 15:53 | 1553197972.8653 | 182 | 0.20654 | 0.002773 | 5.9e-05 | 100.7424 | 98.2437 | 115.5963 | 21.2612 | 0.002975 | 0.002677 | 254.3626 | 0.27325 | 8.4037 | 0.003769 | 0.005352 | 0.005289 | 5e-05 | 0.001077 | 0.27273 | 1.3405 | 2.5479 | 98.87 | 20.84 | 38.617 | 93.3953 | 112.6181 | 86.2132 | 0.00963 | 0.98242 | 0.02331 | 23.7274 | 0.22465 | 0.16255 | 0.014669 | 1.0214 | 1.0214 | 0.006541
2019-03-21 15:53 | 1553197973.8644 | 183 | 0.20646 | 0.002778 | 5.7e-05 | 100.7031 | 97.2812 | 115.3665 | 21.2612 | 0.003023 | 0.002726 | 254.2898 | 0.27259 | 8.4046 | 0.003766 | 0.005767 | 0.005782 | 5.2e-05 | 0.001073 | 0.23435 | 1.3474 | 2.5481 | 98.87 | 20.84 | 38.617 | 94.5511 | 112.5579 | 87.4029 | 0.00963 | 0.97918 | 0.023249 | 23.743 | 0.30922 | 0.16216 | 0.015145 | 1.052 | 1.052 | 0.006336
2019-03-21 15:53 | 1553197974.8654 | 184 | 0.20647 | 0.002788 | 5.5e-05 | 100.7946 | 97.0162 | 115.3196 | 21.2612 | 0.002965 | 0.002698 | 254.2412 | 0.2726 | 8.4009 | 0.003763 | 0.00484 | 0.004741 | 5.1e-05 | 0.001076 | 0.28296 | 1.3056 | 2.548 | 98.87 | 20.84 | 38.618 | 93.1565 | 112.4844 | 86.7151 | 0.00963 | 0.98105 | 0.023266 | 23.7155 | 0.34174 | 0.16217 | 0.01507 | 1.0477 | 1.0477 | 0.006362
2019-03-21 15:53 | 1553197975.8655 | 185 | 0.20644 | 0.002788 | 5.2e-05 | 100.7403 | 97.2914 | 115.0447 | 21.2612 | 0.00298 | 0.002737 | 254.1613 | 0.2717 | 8.4019 | 0.003775 | 0.004936 | 0.004897 | 5.4e-05 | 0.001075 | 0.2557 | 1.2888 | 2.548 | 98.87 | 20.84 | 38.619 | 93.5204 | 112.7839 | 87.6529 | 0.00963 | 0.9785 | 0.023208 | 23.7181 | 0.31159 | 0.16163 | 0.015278 | 1.0596 | 1.0596 | 0.006269
2019-03-21 15:53 | 1553197976.8646 | 186 | 0.20648 | 0.002795 | 5.1e-05 | 100.8047 | 97.2611 | 115.2226 | 21.2612 | 0.002919 | 0.002728 | 254.1159 | 0.27148 | 8.4007 | 0.003729 | 0.005579 | 0.005556 | 5.5e-05 | 0.001073 | 0.26753 | 1.2923 | 2.5479 | 98.87 | 20.83 | 38.62 | 92.0522 | 111.6443 | 87.438 | 0.009624 | 0.97908 | 0.023233 | 23.7293 | 0.32009 | 0.1615 | 0.015053 | 1.0452 | 1.0452 | 0.006351
2019-03-21 15:53 | 1553197977.8656 | 187 | 0.2064 | 0.002798 | 5e-05 | 100.6316 | 97.504 | 115.0715 | 21.2612 | 0.002987 | 0.002731 | 254.0641 | 0.27189 | 8.4022 | 0.003722 | 0.005713 | 0.005638 | 5.3e-05 | 0.001069 | 0.24945 | 1.3024 | 2.5479 | 98.87 | 20.83 | 38.621 | 93.6902 | 111.4889 | 87.5218 | 0.009625 | 0.97886 | 0.023215 | 23.7165 | 0.28241 | 0.16175 | 0.01548 | 1.0742 | 1.0742 | 0.006189
2019-03-21 15:53 | 1553197978.8656 | 188 | 0.20653 | 0.00281 | 4.8e-05 | 100.8011 | 97.4195 | 114.9987 | 21.2612 | 0.002976 | 0.002716 | 254.0072 | 0.27197 | 8.4021 | 0.003662 | 0.005163 | 0.005111 | 5.2e-05 | 0.001072 | 0.28943 | 1.2965 | 2.5479 | 98.87 | 20.83 | 38.62 | 93.4148 | 110.0198 | 87.1625 | 0.009624 | 0.97983 | 0.023219 | 23.6972 | 0.30522 | 0.1618 | 0.01471 | 1.0209 | 1.0209 | 0.006513
2019-03-21 15:53 | 1553197979.8658 | 189 | 0.20645 | 0.002816 | 4.7e-05 | 100.6959 | 97.304 | 114.679 | 21.2612 | 0.002984 | 0.002715 | 253.9567 | 0.27305 | 8.4061 | 0.003713 | 0.005985 | 0.005917 | 5e-05 | 0.001071 | 0.26985 | 1.2987 | 2.5479 | 98.87 | 20.83 | 38.621 | 93.6166 | 111.2529 | 87.1242 | 0.009625 | 0.97994 | 0.023188 | 23.663 | 0.30649 | 0.16244 | 0.015158 | 1.0507 | 1.0507 | 0.006354
2019-03-21 15:53 | 1553197980.8657 | 190 | 0.20642 | 0.002821 | 4.6e-05 | 100.6968 | 97.1276 | 114.5358 | 21.2612 | 0.002984 | 0.002716 | 253.8969 | 0.27298 | 8.4015 | 0.003649 | 0.005233 | 0.005239 | 5e-05 | 0.001072 | 0.26305 | 1.3275 | 2.5481 | 98.87 | 20.83 | 38.618 | 93.6034 | 109.7123 | 87.1586 | 0.009624 | 0.97984 | 0.023173 | 23.6494 | 0.3228 | 0.1624 | 0.015376 | 1.0652 | 1.0652 | 0.006266
2019-03-21 15:53 | 1553197981.8649 | 191 | 0.20643 | 0.002826 | 4.4e-05 | 100.7508 | 97.0016 | 114.6774 | 21.2612 | 0.002957 | 0.002713 | 253.8359 | 0.27265 | 8.4029 | 0.003617 | 0.005258 | 0.005217 | 5.2e-05 | 0.001076 | 0.29927 | 1.2845 | 2.5481 | 98.87 | 20.83 | 38.618 | 92.969 | 108.9255 | 87.096 | 0.009624 | 0.98001 | 0.023189 | 23.6619 | 0.3392 | 0.1622 | 0.01531 | 1.0615 | 1.0615 | 0.006281
2019-03-21 15:53 | 1553197982.8649 | 192 | 0.20641 | 0.002832 | 4.4e-05 | 100.7266 | 96.6 | 114.8863 | 21.2612 | 0.003029 | 0.002739 | 253.794 | 0.27245 | 8.4033 | 0.003619 | 0.004686 | 0.004615 | 5.1e-05 | 0.001073 | 0.27815 | 1.307 | 2.548 | 98.87 | 20.83 | 38.617 | 94.6912 | 108.9892 | 87.709 | 0.009624 | 0.97835 | 0.02319 | 23.7036 | 0.37416 | 0.16208 | 0.015415 | 1.0689 | 1.0689 | 0.006232
2019-03-21 15:53 | 1553197983.8649 | 193 | 0.20642 | 0.002838 | 4.3e-05 | 100.6968 | 96.6535 | 114.9177 | 21.2612 | 0.003014 | 0.002733 | 253.7243 | 0.27206 | 8.4004 | 0.003605 | 0.005816 | 0.005687 | 4.7e-05 | 0.001072 | 0.27256 | 1.3556 | 2.548 | 98.87 | 20.83 | 38.618 | 94.3255 | 108.6394 | 87.5723 | 0.009624 | 0.97872 | 0.023198 | 23.7023 | 0.36656 | 0.16185 | 0.015341 | 1.0641 | 1.0641 | 0.006251
2019-03-21 15:53 | 1553197984.865 | 194 | 0.20636 | 0.002845 | 4.2e-05 | 100.6724 | 96.5806 | 115.0151 | 21.2612 | 0.003012 | 0.002765 | 253.6937 | 0.27268 | 8.404 | 0.003676 | 0.005172 | 0.005192 | 4.8e-05 | 0.001069 | 0.25514 | 1.3438 | 2.548 | 98.87 | 20.83 | 38.619 | 94.2951 | 110.3762 | 88.344 | 0.009624 | 0.97663 | 0.023183 | 23.7377 | 0.37114 | 0.16222 | 0.015713 | 1.0893 | 1.0893 | 0.006121
2019-03-21 15:53 | 1553197985.866 | 195 | 0.20637 | 0.002851 | 4.1e-05 | 100.6858 | 96.5908 | 114.9912 | 21.2612 | 0.002906 | 0.002759 | 253.6358 | 0.2729 | 8.4057 | 0.003621 | 0.004803 | 0.00477 | 4.5e-05 | 0.001067 | 0.25713 | 1.3487 | 2.5479 | 98.87 | 20.83 | 38.619 | 91.7285 | 109.0169 | 88.201 | 0.009624 | 0.97701 | 0.023185 | 23.7306 | 0.37139 | 0.16235 | 0.015615 | 1.0827 | 1.0827 | 0.006163
2019-03-21 15:53 | 1553197986.8651 | 196 | 0.2064 | 0.002857 | 4e-05 | 100.6906 | 96.0156 | 114.6487 | 21.2612 | 0.002984 | 0.002719 | 253.5658 | 0.27424 | 8.4031 | 0.003605 | 0.00576 | 0.005667 | 4.7e-05 | 0.001068 | 0.2672 | 1.3403 | 2.548 | 98.87 | 20.83 | 38.619 | 93.6149 | 108.6299 | 87.2344 | 0.009624 | 0.97964 | 0.023182 | 23.6635 | 0.42524 | 0.16315 | 0.015461 | 1.0719 | 1.0719 | 0.006256
2019-03-21 15:53 | 1553197987.8661 | 197 | 0.20637 | 0.002864 | 3.9e-05 | 100.6807 | 96.1726 | 114.6464 | 21.2612 | 0.003012 | 0.002692 | 253.5077 | 0.27338 | 8.4006 | 0.003647 | 0.005517 | 0.005455 | 4.8e-05 | 0.001065 | 0.27527 | 1.2996 | 2.5481 | 98.87 | 20.83 | 38.619 | 94.284 | 109.6523 | 86.5801 | 0.009624 | 0.98142 | 0.023202 | 23.6418 | 0.40975 | 0.16263 | 0.015629 | 1.0845 | 1.0845 | 0.006163
2019-03-21 15:53 | 1553197988.8652 | 198 | 0.20638 | 0.00288 | 3.8e-05 | 100.7344 | 96.1539 | 114.7133 | 21.2612 | 0.002994 | 0.002755 | 253.459 | 0.27238 | 8.4039 | 0.003628 | 0.005408 | 0.005323 | 4.8e-05 | 0.001064 | 0.26593 | 1.2931 | 2.548 | 98.87 | 20.83 | 38.62 | 93.8544 | 109.1886 | 88.0956 | 0.009624 | 0.9773 | 0.02316 | 23.6984 | 0.41625 | 0.16204 | 0.015541 | 1.0765 | 1.0765 | 0.006186
2019-03-21 15:53 | 1553197989.8642 | 199 | 0.20638 | 0.002881 | 3.7e-05 | 100.6706 | 96.2114 | 114.9236 | 21.2612 | 0.00297 | 0.002731 | 253.408 | 0.27319 | 8.4062 | 0.003694 | 0.005346 | 0.005261 | 4.8e-05 | 0.001067 | 0.2625 | 1.3186 | 2.5479 | 98.87 | 20.83 | 38.621 | 93.2716 | 110.8054 | 87.5107 | 0.009625 | 0.97889 | 0.0232 | 23.7009 | 0.40524 | 0.16252 | 0.015519 | 1.077 | 1.077 | 0.006202
2019-03-21 15:53 | 1553197990.8643 | 200 | 0.20643 | 0.002884 | 3.5e-05 | 100.6426 | 96.1532 | 114.7785 | 21.2612 | 0.003 | 0.002681 | 253.3533 | 0.27322 | 8.4052 | 0.003667 | 0.005506 | 0.005365 | 4.9e-05 | 0.00107 | 0.27815 | 1.3298 | 2.5479 | 98.87 | 20.83 | 38.622 | 93.9907 | 110.1499 | 86.321 | 0.009625 | 0.98213 | 0.023224 | 23.6469 | 0.40817 | 0.16254 | 0.015225 | 1.0577 | 1.0577 | 0.006316
2019-03-21 15:53 | 1553197991.8643 | 201 | 0.20647 | 0.002887 | 3.4e-05 | 100.6316 | 96.5193 | 114.6081 | 21.2612 | 0.002997 | 0.002699 | 253.3012 | 0.27299 | 8.4023 | 0.00367 | 0.00499 | 0.004934 | 4.9e-05 | 0.001066 | 0.26432 | 1.2902 | 2.5479 | 98.87 | 20.83 | 38.624 | 93.9249 | 110.2272 | 86.7363 | 0.009625 | 0.98099 | 0.023194 | 23.6429 | 0.37319 | 0.1624 | 0.015028 | 1.0427 | 1.0427 | 0.006401
2019-03-21 15:53 | 1553197992.8644 | 202 | 0.20647 | 0.002882 | 3.3e-05 | 100.6804 | 96.3051 | 114.2114 | 21.2612 | 0.002975 | 0.002749 | 253.2621 | 0.27272 | 8.4023 | 0.003652 | 0.00499 | 0.004956 | 4.8e-05 | 0.001069 | 0.24923 | 1.3138 | 2.5479 | 98.87 | 20.83 | 38.625 | 93.3851 | 109.7847 | 87.9567 | 0.009626 | 0.97768 | 0.023114 | 23.642 | 0.39741 | 0.16224 | 0.015027 | 1.0391 | 1.0391 | 0.006417
2019-03-21 15:53 | 1553197993.8654 | 203 | 0.20646 | 0.002877 | 3.2e-05 | 100.6795 | 96.5687 | 114.3466 | 21.2612 | 0.002969 | 0.002734 | 253.2153 | 0.27215 | 8.4023 | 0.00369 | 0.004664 | 0.004637 | 5.1e-05 | 0.001066 | 0.27992 | 1.3299 | 2.548 | 98.87 | 20.82 | 38.628 | 93.2402 | 110.7164 | 87.5786 | 0.00962 | 0.9787 | 0.02314 | 23.6436 | 0.37288 | 0.1619 | 0.015079 | 1.0439 | 1.0439 | 0.006374
2019-03-21 15:53 | 1553197994.8656 | 204 | 0.20649 | 0.002866 | 3.2e-05 | 100.6217 | 96.5629 | 114.2462 | 21.2612 | 0.003017 | 0.002779 | 253.1329 | 0.27279 | 8.4048 | 0.003631 | 0.004994 | 0.004887 | 5.2e-05 | 0.001066 | 0.27328 | 1.3248 | 2.548 | 98.87 | 20.82 | 38.629 | 94.3984 | 109.2691 | 88.6712 | 0.009621 | 0.97575 | 0.023095 | 23.669 | 0.36828 | 0.16228 | 0.014899 | 1.0295 | 1.0295 | 0.006479
2019-03-21 15:53 | 1553197995.8659 | 205 | 0.2065 | 0.002859 | 3.1e-05 | 100.6113 | 96.4349 | 114.356 | 21.2612 | 0.002989 | 0.002808 | 253.0673 | 0.27252 | 8.4027 | 0.003636 | 0.00539 | 0.005259 | 4.6e-05 | 0.001065 | 0.2724 | 1.2934 | 2.5479 | 98.87 | 20.82 | 38.631 | 93.7425 | 109.392 | 89.361 | 0.009621 | 0.97389 | 0.023084 | 23.7029 | 0.37922 | 0.16212 | 0.014858 | 1.0262 | 1.0262 | 0.006493
2019-03-21 15:53 | 1553197996.8657 | 206 | 0.20651 | 0.002855 | 3e-05 | 100.5681 | 95.8909 | 114.4236 | 21.2612 | 0.003028 | 0.002743 | 253.0392 | 0.27271 | 8.4034 | 0.003612 | 0.005321 | 0.005261 | 4.7e-05 | 0.001065 | 0.29175 | 1.2749 | 2.5481 | 98.87 | 20.82 | 38.632 | 94.6801 | 108.815 | 87.8183 | 0.009621 | 0.97805 | 0.02314 | 23.6594 | 0.42597 | 0.16224 | 0.014773 | 1.0228 | 1.0228 | 0.006519
2019-03-21 15:53 | 1553197997.8656 | 207 | 0.20651 | 0.002849 | 2.9e-05 | 100.6393 | 95.8531 | 114.349 | 21.2612 | 0.003017 | 0.002745 | 252.9592 | 0.27258 | 8.4021 | 0.003654 | 0.004915 | 0.004883 | 5e-05 | 0.001063 | 0.23396 | 1.2779 | 2.548 | 98.87 | 20.82 | 38.635 | 94.4179 | 109.8368 | 87.8443 | 0.009622 | 0.97798 | 0.023132 | 23.6525 | 0.43583 | 0.16216 | 0.014783 | 1.0232 | 1.0232 | 0.006514
2019-03-21 15:53 | 1553197998.8647 | 208 | 0.20657 | 0.002838 | 3e-05 | 100.5955 | 96.096 | 114.0887 | 21.2612 | 0.003047 | 0.002763 | 252.9095 | 0.27319 | 8.4004 | 0.003611 | 0.00501 | 0.004837 | 5.2e-05 | 0.001062 | 0.27472 | 1.3747 | 2.548 | 98.87 | 20.82 | 38.64 | 95.1224 | 108.7795 | 88.2845 | 0.009623 | 0.97679 | 0.023091 | 23.64 | 0.40931 | 0.16252 | 0.014468 | 0.99964 | 0.99964 | 0.006682
2019-03-21 15:53 | 1553197999.8658 | 209 | 0.20667 | 0.002826 | 3e-05 | 100.6 | 96.237 | 114.1835 | 21.2612 | 0.003 | 0.002693 | 252.8506 | 0.27217 | 8.4017 | 0.00356 | 0.004903 | 0.004844 | 5.2e-05 | 0.001065 | 0.28224 | 1.2989 | 2.5479 | 98.87 | 20.82 | 38.641 | 93.9977 | 107.5435 | 86.5941 | 0.009624 | 0.98138 | 0.023155 | 23.5944 | 0.39659 | 0.16192 | 0.013887 | 0.96217 | 0.96217 | 0.006916
2019-03-21 15:53 | 1553198000.8658 | 210 | 0.20658 | 0.002816 | 2.9e-05 | 100.5744 | 96.3092 | 114.6288 | 21.2612 | 0.002967 | 0.002765 | 252.7927 | 0.27289 | 8.4012 | 0.0035 | 0.00501 | 0.004935 | 5e-05 | 0.001062 | 0.29103 | 1.2776 | 2.5479 | 98.87 | 20.82 | 38.643 | 93.2091 | 106.0885 | 88.3259 | 0.009624 | 0.97668 | 0.023145 | 23.6972 | 0.3876 | 0.16234 | 0.014414 | 0.99826 | 0.99826 | 0.006684
2019-03-21 15:53 | 1553198001.866 | 211 | 0.20674 | 0.002804 | 2.9e-05 | 100.5845 | 96.4124 | 114.8863 | 21.2612 | 0.002997 | 0.002745 | 252.7178 | 0.2721 | 8.4012 | 0.003543 | 0.004653 | 0.004576 | 5e-05 | 0.00106 | 0.25785 | 1.2794 | 2.548 | 98.87 | 20.82 | 38.648 | 93.9367 | 107.1347 | 87.8474 | 0.009625 | 0.97797 | 0.023186 | 23.7081 | 0.37892 | 0.16187 | 0.013466 | 0.93423 | 0.93423 | 0.007121
2019-03-21 15:53 | 1553198002.8659 | 212 | 0.2068 | 0.00279 | 2.9e-05 | 100.502 | 96.3664 | 115.2262 | 21.2612 | 0.003066 | 0.002716 | 252.6468 | 0.27306 | 8.4045 | 0.003561 | 0.005104 | 0.005009 | 5.1e-05 | 0.001057 | 0.30055 | 1.2842 | 2.5479 | 98.87 | 20.82 | 38.65 | 95.5821 | 107.5663 | 87.1531 | 0.009626 | 0.97986 | 0.023243 | 23.7203 | 0.37585 | 0.16244 | 0.013161 | 0.91527 | 0.91527 | 0.007294
2019-03-21 15:53 | 1553198003.865 | 213 | 0.20675 | 0.002773 | 2.8e-05 | 100.5869 | 96.6528 | 115.324 | 21.2612 | 0.00296 | 0.002745 | 252.6082 | 0.27305 | 8.4049 | 0.003597 | 0.004929 | 0.004886 | 5.1e-05 | 0.00106 | 0.2896 | 1.276 | 2.548 | 98.87 | 20.82 | 38.653 | 93.0394 | 108.4375 | 87.8551 | 0.009627 | 0.97795 | 0.02323 | 23.7535 | 0.35685 | 0.16244 | 0.013465 | 0.93601 | 0.93601 | 0.007133
2019-03-21 15:53 | 1553198004.8641 | 214 | 0.20668 | 0.002758 | 2.8e-05 | 100.5771 | 96.7533 | 115.3367 | 21.2612 | 0.002961 | 0.002729 | 252.5491 | 0.27354 | 8.403 | 0.003675 | 0.005439 | 0.005405 | 5.4e-05 | 0.001059 | 0.24513 | 1.2661 | 2.548 | 98.87 | 20.82 | 38.653 | 93.0635 | 110.3392 | 87.4756 | 0.009627 | 0.97898 | 0.023243 | 23.7423 | 0.34669 | 0.16273 | 0.013893 | 0.96634 | 0.96634 | 0.006921
2019-03-21 15:53 | 1553198005.8648 | 215 | 0.2068 | 0.002742 | 2.7e-05 | 100.6229 | 96.3402 | 115.0628 | 21.2612 | 0.003056 | 0.002755 | 252.4996 | 0.27356 | 8.3996 | 0.003599 | 0.005008 | 0.004965 | 5.1e-05 | 0.001057 | 0.27616 | 1.2595 | 2.5479 | 98.87 | 20.82 | 38.654 | 95.3495 | 108.4825 | 88.0917 | 0.009627 | 0.97731 | 0.023196 | 23.7344 | 0.38904 | 0.16274 | 0.013214 | 0.91724 | 0.91724 | 0.007292
2019-03-21 15:53 | 1553198006.8642 | 216 | 0.20669 | 0.002729 | 2.7e-05 | 100.5934 | 96.5258 | 115.182 | 21.2612 | 0.002991 | 0.002766 | 252.4483 | 0.2722 | 8.4035 | 0.003608 | 0.005302 | 0.005216 | 5.4e-05 | 0.001058 | 0.29286 | 1.3349 | 2.5479 | 98.87 | 20.82 | 38.655 | 93.7707 | 108.7096 | 88.3698 | 0.009627 | 0.97656 | 0.023199 | 23.7558 | 0.3692 | 0.16193 | 0.013869 | 0.96292 | 0.96292 | 0.006912
2019-03-21 15:53 | 1553198007.8642 | 217 | 0.20681 | 0.002711 | 2.6e-05 | 100.6057 | 96.5752 | 115.3473 | 21.2612 | 0.003044 | 0.00275 | 252.4221 | 0.27199 | 8.4033 | 0.00365 | 0.004855 | 0.004869 | 5.3e-05 | 0.001055 | 0.26897 | 1.3209 | 2.5481 | 98.87 | 20.82 | 38.655 | 95.0645 | 109.7327 | 87.9651 | 0.009627 | 0.97765 | 0.023229 | 23.7595 | 0.36571 | 0.16181 | 0.013157 | 0.91466 | 0.91466 | 0.007271
2019-03-21 15:53 | 1553198008.8648 | 218 | 0.20684 | 0.002698 | 2.5e-05 | 100.6882 | 97.194 | 115.6379 | 21.2612 | 0.003019 | 0.002762 | 252.3745 | 0.27304 | 8.4005 | 0.003559 | 0.005151 | 0.005113 | 5.3e-05 | 0.001054 | 0.24314 | 1.3402 | 2.548 | 98.87 | 20.82 | 38.652 | 94.4562 | 107.527 | 88.266 | 0.009626 | 0.97684 | 0.023248 | 23.7993 | 0.31592 | 0.16243 | 0.013009 | 0.90515 | 0.90515 | 0.007376
2019-03-21 15:53 | 1553198009.8643 | 219 | 0.2068 | 0.002686 | 2.6e-05 | 100.6018 | 97.0775 | 115.5569 | 21.2612 | 0.003027 | 0.002715 | 252.3411 | 0.27309 | 8.405 | 0.003599 | 0.004641 | 0.004529 | 5.4e-05 | 0.001054 | 0.27295 | 1.3145 | 2.5481 | 98.87 | 20.82 | 38.652 | 94.6543 | 108.5038 | 87.1367 | 0.009626 | 0.9799 | 0.023276 | 23.7538 | 0.31896 | 0.16246 | 0.013274 | 0.92472 | 0.92472 | 0.007221
2019-03-21 15:54 | 1553198010.8644 | 220 | 0.20683 | 0.002666 | 2.6e-05 | 100.636 | 97.3868 | 115.2088 | 21.2612 | 0.003035 | 0.002808 | 252.3149 | 0.27236 | 8.4028 | 0.003594 | 0.005142 | 0.005132 | 5.2e-05 | 0.001053 | 0.22804 | 1.3235 | 2.5479 | 98.87 | 20.82 | 38.65 | 94.8343 | 108.3715 | 89.3704 | 0.009626 | 0.97386 | 0.02317 | 23.7914 | 0.29356 | 0.16203 | 0.013124 | 0.91006 | 0.91006 | 0.007317
2019-03-21 15:54 | 1553198011.8644 | 221 | 0.20681 | 0.002665 | 2.6e-05 | 100.681 | 97.4617 | 114.9804 | 21.2612 | 0.00295 | 0.002731 | 252.2928 | 0.27261 | 8.4016 | 0.003613 | 0.005618 | 0.005493 | 5e-05 | 0.001054 | 0.29336 | 1.3339 | 2.5481 | 98.87 | 20.82 | 38.65 | 92.793 | 108.8441 | 87.5263 | 0.009626 | 0.97884 | 0.023206 | 23.7073 | 0.29068 | 0.16218 | 0.013232 | 0.91895 | 0.91895 | 0.007253
2019-03-21 15:54 | 1553198012.8645 | 222 | 0.20674 | 0.002648 | 2.7e-05 | 100.628 | 97.553 | 115.0217 | 21.2612 | 0.003018 | 0.002733 | 252.2609 | 0.27243 | 8.4022 | 0.003622 | 0.004413 | 0.004401 | 5.2e-05 | 0.001053 | 0.27632 | 1.3063 | 2.5479 | 98.87 | 20.82 | 38.648 | 94.4259 | 109.0548 | 87.56 | 0.009625 | 0.97875 | 0.023209 | 23.7127 | 0.27759 | 0.16207 | 0.013669 | 0.94943 | 0.94943 | 0.007016
2019-03-21 15:54 | 1553198013.8645 | 223 | 0.20685 | 0.002638 | 2.6e-05 | 100.6289 | 97.5479 | 115.0168 | 21.2612 | 0.002962 | 0.002768 | 252.2115 | 0.27199 | 8.4018 | 0.003578 | 0.005309 | 0.005199 | 5.1e-05 | 0.001051 | 0.26161 | 1.3281 | 2.548 | 98.87 | 20.82 | 38.644 | 93.0746 | 107.99 | 88.3994 | 0.009624 | 0.97648 | 0.023181 | 23.7397 | 0.27814 | 0.16181 | 0.013014 | 0.90281 | 0.90281 | 0.007366
2019-03-21 15:54 | 1553198014.8645 | 224 | 0.20691 | 0.002626 | 2.7e-05 | 100.6298 | 97.7349 | 114.8574 | 21.2612 | 0.002985 | 0.00274 | 252.1742 | 0.27285 | 8.4022 | 0.003574 | 0.004935 | 0.004891 | 5.4e-05 | 0.001053 | 0.27704 | 1.333 | 2.5479 | 98.87 | 20.82 | 38.643 | 93.6347 | 107.8948 | 87.7357 | 0.009624 | 0.97828 | 0.023187 | 23.7015 | 0.26109 | 0.16232 | 0.012707 | 0.88172 | 0.88172 | 0.007566
2019-03-21 15:54 | 1553198015.8661 | 225 | 0.20694 | 0.002614 | 2.7e-05 | 100.6506 | 97.8701 | 115.0597 | 21.2612 | 0.002997 | 0.002741 | 252.1308 | 0.27251 | 8.4032 | 0.003577 | 0.00523 | 0.005209 | 5.1e-05 | 0.001049 | 0.26288 | 1.3241 | 2.5479 | 98.87 | 20.82 | 38.642 | 93.9335 | 107.9594 | 87.7543 | 0.009624 | 0.97822 | 0.023206 | 23.723 | 0.25058 | 0.16212 | 0.012501 | 0.86818 | 0.86818 | 0.007675
2019-03-21 15:54 | 1553198016.8658 | 226 | 0.20692 | 0.002608 | 2.6e-05 | 100.5562 | 97.7175 | 115.0393 | 21.2612 | 0.00296 | 0.002748 | 252.0985 | 0.2725 | 8.4016 | 0.003517 | 0.005214 | 0.005176 | 5.1e-05 | 0.001047 | 0.26322 | 1.3608 | 2.5479 | 98.87 | 20.82 | 38.639 | 93.0237 | 106.5096 | 87.9317 | 0.009623 | 0.97774 | 0.023199 | 23.7267 | 0.25614 | 0.16211 | 0.012664 | 0.87923 | 0.87923 | 0.007578
2019-03-21 15:54 | 1553198017.8647 | 227 | 0.207 | 0.002596 | 2.5e-05 | 100.6858 | 97.7485 | 115.1975 | 21.2612 | 0.003017 | 0.002721 | 252.0511 | 0.27282 | 8.4023 | 0.003611 | 0.0052 | 0.005125 | 5.1e-05 | 0.001052 | 0.25083 | 1.3574 | 2.5481 | 98.87 | 20.82 | 38.638 | 94.3976 | 108.7946 | 87.2859 | 0.009623 | 0.9795 | 0.023235 | 23.7218 | 0.26482 | 0.1623 | 0.012207 | 0.8489 | 0.8489 | 0.007858
2019-03-21 15:54 | 1553198018.8638 | 228 | 0.20705 | 0.002577 | 2.5e-05 | 100.625 | 97.7195 | 115.2189 | 21.2612 | 0.002984 | 0.002737 | 252.014 | 0.27129 | 8.4023 | 0.003655 | 0.005239 | 0.005152 | 4.8e-05 | 0.00105 | 0.26842 | 1.3406 | 2.548 | 98.87 | 20.82 | 38.636 | 93.6229 | 109.862 | 87.663 | 0.009622 | 0.97847 | 0.023225 | 23.7364 | 0.26207 | 0.16139 | 0.011898 | 0.82698 | 0.82698 | 0.008021
2019-03-21 15:54 | 1553198019.8649 | 229 | 0.20702 | 0.002565 | 2.5e-05 | 100.6575 | 97.5118 | 115.4203 | 21.2612 | 0.002979 | 0.002729 | 251.9605 | 0.27149 | 8.402 | 0.003603 | 0.005241 | 0.005173 | 5e-05 | 0.001051 | 0.2792 | 1.3122 | 2.5481 | 98.87 | 20.81 | 38.635 | 93.4876 | 108.6015 | 87.47 | 0.009616 | 0.979 | 0.023252 | 23.7507 | 0.28399 | 0.16151 | 0.012131 | 0.84422 | 0.84422 | 0.007863
2019-03-21 15:54 | 1553198020.8639 | 230 | 0.20697 | 0.002553 | 2.5e-05 | 100.5875 | 97.6123 | 115.4348 | 21.2612 | 0.003009 | 0.002733 | 251.8861 | 0.27201 | 8.3983 | 0.00357 | 0.004975 | 0.00494 | 5.4e-05 | 0.00105 | 0.2672 | 1.2904 | 2.5481 | 98.87 | 20.81 | 38.635 | 94.2191 | 107.7986 | 87.5639 | 0.009616 | 0.97874 | 0.02325 | 23.7553 | 0.26856 | 0.16182 | 0.012446 | 0.86606 | 0.86606 | 0.007679
2019-03-21 15:54 | 1553198021.8639 | 231 | 0.20702 | 0.002539 | 2.6e-05 | 100.6185 | 97.962 | 115.5914 | 21.2612 | 0.002994 | 0.002724 | 251.842 | 0.27196 | 8.4004 | 0.003573 | 0.005616 | 0.005541 | 5.1e-05 | 0.001047 | 0.27096 | 1.3012 | 2.5479 | 98.87 | 20.81 | 38.635 | 93.8467 | 107.8656 | 87.3538 | 0.009616 | 0.97931 | 0.023273 | 23.7645 | 0.23932 | 0.16179 | 0.012156 | 0.84669 | 0.84669 | 0.007854
2019-03-21 15:54 | 1553198022.865 | 232 | 0.20708 | 0.002522 | 2.6e-05 | 100.5958 | 98.0969 | 115.4888 | 21.2612 | 0.002998 | 0.002669 | 251.7821 | 0.27254 | 8.3995 | 0.003569 | 0.004936 | 0.004977 | 5.4e-05 | 0.001045 | 0.24994 | 1.3375 | 2.5478 | 98.87 | 20.81 | 38.634 | 93.9618 | 107.7765 | 86.0258 | 0.009616 | 0.98293 | 0.023305 | 23.7102 | 0.225 | 0.16214 | 0.011782 | 0.82176 | 0.82176 | 0.008109
2019-03-21 15:54 | 1553198023.865 | 233 | 0.20708 | 0.002515 | 2.7e-05 | 100.5705 | 97.6722 | 115.5665 | 21.2612 | 0.002971 | 0.002723 | 251.737 | 0.27178 | 8.4054 | 0.00354 | 0.004826 | 0.004769 | 5.6e-05 | 0.001047 | 0.25929 | 1.3331 | 2.5481 | 98.87 | 20.81 | 38.632 | 93.3084 | 107.067 | 87.3304 | 0.009615 | 0.97937 | 0.023271 | 23.7612 | 0.26156 | 0.16168 | 0.011794 | 0.82132 | 0.82132 | 0.008091
2019-03-21 15:54 | 1553198024.8641 | 234 | 0.20699 | 0.002493 | 2.7e-05 | 100.5914 | 97.6368 | 115.7069 | 21.2612 | 0.002924 | 0.002721 | 251.6809 | 0.27272 | 8.4015 | 0.003496 | 0.005357 | 0.005295 | 5.3e-05 | 0.001049 | 0.24746 | 1.2925 | 2.5479 | 98.87 | 20.81 | 38.632 | 92.1734 | 105.9988 | 87.2758 | 0.009615 | 0.97952 | 0.023287 | 23.7738 | 0.26665 | 0.16224 | 0.012397 | 0.86392 | 0.86392 | 0.007718
2019-03-21 15:54 | 1553198025.8644 | 235 | 0.2071 | 0.002483 | 2.8e-05 | 100.6027 | 97.8892 | 115.6862 | 21.2612 | 0.002951 | 0.002698 | 251.6649 | 0.27235 | 8.4058 | 0.003542 | 0.004675 | 0.004586 | 5.4e-05 | 0.001047 | 0.27079 | 1.3037 | 2.5479 | 98.87 | 20.81 | 38.632 | 92.8282 | 107.1008 | 86.712 | 0.009615 | 0.98106 | 0.023303 | 23.7531 | 0.24457 | 0.16202 | 0.011719 | 0.81713 | 0.81713 | 0.008149
2019-03-21 15:54 | 1553198026.8652 | 236 | 0.20703 | 0.002465 | 2.9e-05 | 100.5175 | 98.3309 | 115.5886 | 21.2612 | 0.002911 | 0.00266 | 251.5785 | 0.27272 | 8.4025 | 0.00353 | 0.004223 | 0.004135 | 5.4e-05 | 0.001046 | 0.25299 | 1.3027 | 2.5481 | 98.87 | 20.81 | 38.632 | 91.8637 | 106.8134 | 85.8011 | 0.009615 | 0.98355 | 0.023323 | 23.713 | 0.19672 | 0.16224 | 0.012167 | 0.84905 | 0.84905 | 0.007854
2019-03-21 15:54 | 1553198027.8643 | 237 | 0.20693 | 0.002453 | 3e-05 | 100.5365 | 98.5186 | 115.4564 | 21.2612 | 0.002978 | 0.002701 | 251.5342 | 0.27253 | 8.4019 | 0.003512 | 0.00516 | 0.005134 | 5.6e-05 | 0.001049 | 0.25354 | 1.3163 | 2.5479 | 98.87 | 20.81 | 38.633 | 93.4782 | 106.3798 | 86.8049 | 0.009616 | 0.98081 | 0.023277 | 23.7325 | 0.18136 | 0.16213 | 0.012749 | 0.88789 | 0.88789 | 0.007505
2019-03-21 15:54 | 1553198028.8657 | 238 | 0.20701 | 0.002443 | 3e-05 | 100.5601 | 98.5479 | 115.4921 | 21.2612 | 0.002979 | 0.00272 | 251.4767 | 0.27232 | 8.3993 | 0.00344 | 0.00527 | 0.005211 | 5.3e-05 | 0.001047 | 0.26288 | 1.3214 | 2.548 | 98.87 | 20.81 | 38.635 | 93.4845 | 104.6269 | 87.2546 | 0.009616 | 0.97958 | 0.023266 | 23.751 | 0.1808 | 0.162 | 0.012317 | 0.85733 | 0.85733 | 0.007766
2019-03-21 15:54 | 1553198029.8654 | 239 | 0.20712 | 0.002432 | 3.1e-05 | 100.5371 | 98.5836 | 115.232 | 21.2612 | 0.003011 | 0.002708 | 251.4435 | 0.27241 | 8.4029 | 0.00347 | 0.005043 | 0.00501 | 5e-05 | 0.001046 | 0.23922 | 1.2801 | 2.5479 | 98.86 | 20.81 | 38.639 | 94.2575 | 105.3527 | 86.9719 | 0.009618 | 0.98035 | 0.023249 | 23.715 | 0.17551 | 0.16206 | 0.011649 | 0.81021 | 0.81021 | 0.008221
2019-03-21 15:54 | 1553198030.8654 | 240 | 0.20714 | 0.002421 | 3.2e-05 | 100.5264 | 98.4862 | 115.1773 | 21.2612 | 0.002924 | 0.002671 | 251.4158 | 0.27216 | 8.4011 | 0.003473 | 0.00477 | 0.004771 | 5.4e-05 | 0.00105 | 0.27184 | 1.2856 | 2.5479 | 98.86 | 20.81 | 38.64 | 92.1602 | 105.4284 | 86.0734 | 0.009618 | 0.9828 | 0.023272 | 23.6798 | 0.1834 | 0.16191 | 0.011541 | 0.80341 | 0.80341 | 0.008283
2019-03-21 15:54 | 1553198031.8649 | 241 | 0.20709 | 0.002415 | 3.3e-05 | 100.5112 | 98.1058 | 115.328 | 21.2612 | 0.002946 | 0.002751 | 251.3787 | 0.27172 | 8.403 | 0.003403 | 0.004825 | 0.004732 | 5e-05 | 0.001046 | 0.2943 | 1.2676 | 2.5478 | 98.86 | 20.81 | 38.643 | 92.6857 | 103.7361 | 87.9964 | 0.009619 | 0.97757 | 0.023226 | 23.7586 | 0.21666 | 0.16165 | 0.011844 | 0.82283 | 0.82283 | 0.008074
2019-03-21 15:54 | 1553198032.8655 | 242 | 0.20712 | 0.002406 | 3.3e-05 | 100.4701 | 98.1521 | 115.4562 | 21.2612 | 0.002989 | 0.002693 | 251.3083 | 0.27252 | 8.4001 | 0.003487 | 0.004597 | 0.00454 | 5e-05 | 0.001048 | 0.28031 | 1.3047 | 2.5481 | 98.86 | 20.81 | 38.643 | 93.7223 | 105.7691 | 86.5948 | 0.009619 | 0.98138 | 0.023284 | 23.7256 | 0.20878 | 0.16212 | 0.011687 | 0.81395 | 0.81395 | 0.008186
2019-03-21 15:54 | 1553198033.8646 | 243 | 0.20713 | 0.002397 | 3.4e-05 | 100.4698 | 98.3285 | 115.4766 | 21.2612 | 0.002959 | 0.002689 | 251.2714 | 0.27206 | 8.4034 | 0.003521 | 0.005064 | 0.004913 | 5.1e-05 | 0.001051 | 0.27472 | 1.2995 | 2.5479 | 98.86 | 20.81 | 38.644 | 93.0119 | 106.6096 | 86.5082 | 0.009619 | 0.98161 | 0.023289 | 23.7248 | 0.19269 | 0.16185 | 0.011615 | 0.80901 | 0.80901 | 0.008222
2019-03-21 15:54 | 1553198034.8646 | 244 | 0.2072 | 0.002384 | 3.4e-05 | 100.5005 | 98.4389 | 115.72 | 21.2612 | 0.003001 | 0.002716 | 251.2059 | 0.27241 | 8.4012 | 0.003454 | 0.0062 | 0.006135 | 4.8e-05 | 0.001051 | 0.27776 | 1.3149 | 2.5479 | 98.86 | 20.81 | 38.647 | 94.0189 | 104.9784 | 87.1475 | 0.00962 | 0.97987 | 0.023292 | 23.7709 | 0.18539 | 0.16206 | 0.011231 | 0.78237 | 0.78237 | 0.008513
2019-03-21 15:54 | 1553198035.8646 | 245 | 0.2072 | 0.00238 | 3.4e-05 | 100.5291 | 98.4518 | 115.6613 | 21.2612 | 0.002997 | 0.002673 | 251.1708 | 0.27196 | 8.4021 | 0.003475 | 0.004931 | 0.004874 | 5.1e-05 | 0.001051 | 0.28335 | 1.3295 | 2.548 | 98.86 | 20.81 | 38.648 | 93.9186 | 105.4819 | 86.1235 | 0.00962 | 0.98267 | 0.02332 | 23.7311 | 0.18676 | 0.16179 | 0.011232 | 0.78336 | 0.78336 | 0.008488
2019-03-21 15:54 | 1553198036.8637 | 246 | 0.20715 | 0.002364 | 3.4e-05 | 100.5553 | 98.6531 | 115.6585 | 21.2612 | 0.002982 | 0.002711 | 251.1403 | 0.27213 | 8.4009 | 0.003478 | 0.005047 | 0.005012 | 5.2e-05 | 0.001049 | 0.25177 | 1.3371 | 2.548 | 98.86 | 20.81 | 38.649 | 93.558 | 105.5692 | 87.0352 | 0.009621 | 0.98018 | 0.02329 | 23.7609 | 0.17083 | 0.16189 | 0.011567 | 0.80575 | 0.80575 | 0.008258
2019-03-21 15:54 | 1553198037.8649 | 247 | 0.20722 | 0.002363 | 3.4e-05 | 100.4567 | 98.5894 | 115.8449 | 21.2612 | 0.002955 | 0.002713 | 251.0972 | 0.27102 | 8.404 | 0.003445 | 0.005162 | 0.00509 | 5.5e-05 | 0.001049 | 0.26521 | 1.2879 | 2.548 | 98.86 | 20.81 | 38.651 | 92.9039 | 104.7636 | 87.0797 | 0.009621 | 0.98006 | 0.023307 | 23.7815 | 0.16782 | 0.16123 | 0.011154 | 0.77751 | 0.77751 | 0.008523
2019-03-21 15:54 | 1553198038.8638 | 248 | 0.20724 | 0.002347 | 3.4e-05 | 100.5908 | 98.6858 | 115.793 | 21.2612 | 0.002978 | 0.002719 | 251.0489 | 0.27237 | 8.4014 | 0.003479 | 0.005315 | 0.005289 | 5.4e-05 | 0.001048 | 0.29535 | 1.3073 | 2.5479 | 98.86 | 20.81 | 38.651 | 93.4608 | 105.5857 | 87.232 | 0.009621 | 0.97964 | 0.023297 | 23.7812 | 0.17101 | 0.16203 | 0.01104 | 0.76918 | 0.76918 | 0.008658
2019-03-21 15:54 | 1553198039.864 | 249 | 0.20738 | 0.002335 | 3.5e-05 | 100.4612 | 98.7553 | 115.8517 | 21.2612 | 0.002968 | 0.00269 | 250.9958 | 0.27232 | 8.4023 | 0.003448 | 0.00463 | 0.0046 | 5.2e-05 | 0.001049 | 0.28263 | 1.3175 | 2.5481 | 98.86 | 20.81 | 38.649 | 93.2325 | 104.821 | 86.5371 | 0.009621 | 0.98154 | 0.023325 | 23.7643 | 0.15319 | 0.162 | 0.010199 | 0.71134 | 0.71134 | 0.00936
2019-03-21 15:54 | 1553198040.865 | 250 | 0.20723 | 0.002322 | 3.5e-05 | 100.4781 | 98.5567 | 116.1845 | 21.2612 | 0.002975 | 0.002673 | 250.9571 | 0.27203 | 8.3999 | 0.003399 | 0.004654 | 0.004538 | 5e-05 | 0.001047 | 0.30326 | 1.2979 | 2.548 | 98.86 | 20.81 | 38.65 | 93.3984 | 103.6387 | 86.1117 | 0.009621 | 0.9827 | 0.023373 | 23.7843 | 0.1727 | 0.16183 | 0.011149 | 0.77928 | 0.77928 | 0.008535
2019-03-21 15:54 | 1553198041.8649 | 251 | 0.20731 | 0.002311 | 3.6e-05 | 100.4248 | 98.6248 | 116.2085 | 21.2612 | 0.002986 | 0.002704 | 250.92 | 0.27241 | 8.4039 | 0.003391 | 0.005014 | 0.004973 | 5.4e-05 | 0.001049 | 0.27201 | 1.296 | 2.548 | 98.86 | 20.81 | 38.648 | 93.6543 | 103.4602 | 86.8689 | 0.00962 | 0.98063 | 0.023351 | 23.8118 | 0.16177 | 0.16206 | 0.010657 | 0.74412 | 0.74412 | 0.008951
2019-03-21 15:54 | 1553198042.8651 | 252 | 0.20726 | 0.002299 | 3.7e-05 | 100.476 | 98.2822 | 116.1031 | 21.2612 | 0.00298 | 0.002698 | 250.9055 | 0.27213 | 8.403 | 0.003462 | 0.005262 | 0.005203 | 5.5e-05 | 0.001049 | 0.26521 | 1.2801 | 2.548 | 98.86 | 20.81 | 38.646 | 93.5235 | 105.1655 | 86.7269 | 0.00962 | 0.98102 | 0.023345 | 23.7963 | 0.19746 | 0.16189 | 0.010972 | 0.76585 | 0.76585 | 0.008688
2019-03-21 15:54 | 1553198043.8652 | 253 | 0.20723 | 0.002293 | 3.7e-05 | 100.4528 | 98.2389 | 115.9592 | 21.2612 | 0.002959 | 0.002716 | 250.8627 | 0.27269 | 8.3994 | 0.003429 | 0.004927 | 0.004853 | 5.1e-05 | 0.001054 | 0.28008 | 1.3132 | 2.548 | 98.86 | 20.81 | 38.646 | 93.0143 | 104.3666 | 87.1641 | 0.00962 | 0.97983 | 0.023316 | 23.796 | 0.19933 | 0.16222 | 0.011174 | 0.77904 | 0.77904 | 0.008558
2019-03-21 15:54 | 1553198044.8652 | 254 | 0.2073 | 0.002282 | 3.7e-05 | 100.4269 | 98.2819 | 115.735 | 21.2612 | 0.002987 | 0.00266 | 250.8251 | 0.27211 | 8.403 | 0.00351 | 0.004967 | 0.00491 | 5.4e-05 | 0.001054 | 0.24585 | 1.3296 | 2.548 | 98.86 | 20.81 | 38.644 | 93.6808 | 106.342 | 85.7994 | 0.009619 | 0.98355 | 0.023338 | 23.728 | 0.19312 | 0.16188 | 0.010723 | 0.74825 | 0.74825 | 0.008892
2019-03-21 15:54 | 1553198045.8653 | 255 | 0.2073 | 0.002275 | 3.7e-05 | 100.4486 | 98.4613 | 116.0949 | 21.2612 | 0.002976 | 0.002698 | 250.7487 | 0.27199 | 8.3987 | 0.00352 | 0.005045 | 0.004905 | 5.2e-05 | 0.001052 | 0.26305 | 1.3126 | 2.5479 | 98.86 | 20.81 | 38.644 | 93.4203 | 106.5734 | 86.7308 | 0.009619 | 0.98101 | 0.023344 | 23.7956 | 0.17873 | 0.16181 | 0.010736 | 0.74935 | 0.74935 | 0.008875
2019-03-21 15:54 | 1553198046.8643 | 256 | 0.20738 | 0.00227 | 3.7e-05 | 100.3393 | 98.4951 | 116.4056 | 21.2612 | 0.00296 | 0.002667 | 250.6723 | 0.27249 | 8.4001 | 0.003455 | 0.005669 | 0.005696 | 5.2e-05 | 0.001048 | 0.2724 | 1.315 | 2.5481 | 98.86 | 20.81 | 38.643 | 93.0338 | 105.0004 | 85.9681 | 0.009619 | 0.98309 | 0.0234 | 23.8022 | 0.16593 | 0.1621 | 0.010293 | 0.72006 | 0.72006 | 0.009253
2019-03-21 15:54 | 1553198047.8651 | 257 | 0.20735 | 0.002256 | 3.8e-05 | 100.3301 | 98.4464 | 116.3413 | 21.2612 | 0.003002 | 0.002686 | 250.6173 | 0.27199 | 8.4002 | 0.003378 | 0.005129 | 0.004973 | 5.3e-05 | 0.001053 | 0.31022 | 1.3313 | 2.548 | 98.86 | 20.81 | 38.64 | 94.0385 | 103.1332 | 86.4348 | 0.009618 | 0.98181 | 0.023378 | 23.8111 | 0.16953 | 0.16181 | 0.010496 | 0.73356 | 0.73356 | 0.009066
2019-03-21 15:54 | 1553198048.8643 | 258 | 0.20749 | 0.002243 | 3.8e-05 | 100.3864 | 98.1736 | 116.2247 | 21.2612 | 0.003002 | 0.002683 | 250.5462 | 0.27189 | 8.4038 | 0.003506 | 0.005437 | 0.005423 | 5.3e-05 | 0.001056 | 0.27632 | 1.3293 | 2.5481 | 98.86 | 20.81 | 38.639 | 94.0563 | 106.2428 | 86.3624 | 0.009618 | 0.98201 | 0.023369 | 23.7967 | 0.19937 | 0.16174 | 0.009636 | 0.6731 | 0.6731 | 0.009876
2019-03-21 15:54 | 1553198049.8645 | 259 | 0.20734 | 0.002236 | 3.8e-05 | 100.3494 | 98.3173 | 116.2718 | 21.2612 | 0.002989 | 0.002672 | 250.4912 | 0.27224 | 8.4031 | 0.003466 | 0.004649 | 0.004635 | 5.3e-05 | 0.001059 | 0.29607 | 1.3541 | 2.548 | 98.87 | 20.81 | 38.636 | 93.7449 | 105.2758 | 86.0999 | 0.009616 | 0.98273 | 0.023382 | 23.7929 | 0.18299 | 0.16196 | 0.010549 | 0.7374 | 0.7374 | 0.009027
2019-03-21 15:54 | 1553198050.8645 | 260 | 0.20744 | 0.002228 | 3.8e-05 | 100.387 | 98.7515 | 116.0155 | 21.2612 | 0.002962 | 0.002665 | 250.4141 | 0.27261 | 8.4025 | 0.003415 | 0.005602 | 0.005494 | 5.5e-05 | 0.001065 | 0.28838 | 1.309 | 2.5479 | 98.87 | 20.81 | 38.635 | 93.0729 | 104.0247 | 85.926 | 0.009616 | 0.9832 | 0.023362 | 23.7609 | 0.14692 | 0.16218 | 0.009987 | 0.69742 | 0.69742 | 0.009558
2019-03-21 15:54 | 1553198051.8635 | 261 | 0.20725 | 0.002211 | 3.8e-05 | 100.3971 | 98.6729 | 115.5095 | 21.2612 | 0.002966 | 0.00264 | 250.3629 | 0.27279 | 8.4015 | 0.00347 | 0.005106 | 0.005023 | 5.2e-05 | 0.001065 | 0.28479 | 1.3206 | 2.5481 | 98.87 | 20.81 | 38.633 | 93.1684 | 105.3622 | 85.3282 | 0.009616 | 0.98484 | 0.02333 | 23.6893 | 0.15495 | 0.16228 | 0.011136 | 0.77678 | 0.77678 | 0.008587
2019-03-21 15:54 | 1553198052.8648 | 262 | 0.20743 | 0.002202 | 3.8e-05 | 100.4051 | 98.8932 | 115.3254 | 21.2612 | 0.002948 | 0.002673 | 250.3123 | 0.27224 | 8.4009 | 0.003449 | 0.004609 | 0.004552 | 5.5e-05 | 0.001061 | 0.25138 | 1.3212 | 2.5479 | 98.87 | 20.81 | 38.628 | 92.7474 | 104.8644 | 86.1141 | 0.009614 | 0.98269 | 0.023286 | 23.6963 | 0.13571 | 0.16196 | 0.010041 | 0.69898 | 0.69898 | 0.009523
2019-03-21 15:54 | 1553198053.8637 | 263 | 0.20741 | 0.002199 | 3.9e-05 | 100.4376 | 98.5703 | 115.2783 | 21.2612 | 0.002929 | 0.002696 | 250.2618 | 0.27304 | 8.403 | 0.003416 | 0.005114 | 0.00509 | 5.2e-05 | 0.001064 | 0.27743 | 1.3246 | 2.5481 | 98.87 | 20.81 | 38.626 | 92.2853 | 104.0443 | 86.6761 | 0.009614 | 0.98115 | 0.023263 | 23.71 | 0.16786 | 0.16243 | 0.010158 | 0.7064 | 0.7064 | 0.009451
2019-03-21 15:54 | 1553198054.8647 | 264 | 0.20746 | 0.002186 | 3.9e-05 | 100.345 | 98.6115 | 115.2984 | 21.2612 | 0.002886 | 0.002657 | 250.2376 | 0.27253 | 8.4013 | 0.003402 | 0.004571 | 0.00447 | 4.9e-05 | 0.001061 | 0.27688 | 1.3436 | 2.548 | 98.87 | 20.81 | 38.625 | 91.261 | 103.7211 | 85.727 | 0.009614 | 0.98375 | 0.023296 | 23.6808 | 0.15586 | 0.16213 | 0.009887 | 0.68839 | 0.68839 | 0.00968
2019-03-21 15:54 | 1553198055.8648 | 265 | 0.20743 | 0.002177 | 4e-05 | 100.3959 | 98.6786 | 115.2982 | 21.2612 | 0.002954 | 0.002692 | 250.2182 | 0.27266 | 8.4011 | 0.003383 | 0.004747 | 0.004641 | 5.3e-05 | 0.001065 | 0.24618 | 1.3271 | 2.548 | 98.87 | 20.81 | 38.623 | 92.9007 | 103.2456 | 86.5753 | 0.009613 | 0.98143 | 0.023268 | 23.7087 | 0.15432 | 0.16221 | 0.010101 | 0.70247 | 0.70247 | 0.00949
2019-03-21 15:54 | 1553198056.8638 | 266 | 0.20744 | 0.002164 | 4.3e-05 | 100.4987 | 98.5114 | 115.279 | 21.2612 | 0.002961 | 0.002686 | 250.1968 | 0.27229 | 8.4035 | 0.003441 | 0.004996 | 0.004976 | 5e-05 | 0.001062 | 0.28888 | 1.3578 | 2.548 | 98.87 | 20.81 | 38.62 | 93.0666 | 104.6574 | 86.4372 | 0.009612 | 0.98181 | 0.023271 | 23.7022 | 0.17864 | 0.16198 | 0.01002 | 0.69676 | 0.69676 | 0.009555
2019-03-21 15:54 | 1553198057.8653 | 267 | 0.20739 | 0.002155 | 4.5e-05 | 100.4525 | 98.0506 | 115.1072 | 21.2612 | 0.002918 | 0.002698 | 250.1702 | 0.272 | 8.4007 | 0.003411 | 0.004523 | 0.004448 | 4.9e-05 | 0.001058 | 0.25553 | 1.3317 | 2.5479 | 98.87 | 20.81 | 38.619 | 92.0271 | 103.9277 | 86.7245 | 0.009612 | 0.98102 | 0.023244 | 23.694 | 0.21647 | 0.16181 | 0.010327 | 0.71724 | 0.71724 | 0.009272
2019-03-21 15:54 | 1553198058.8639 | 268 | 0.2074 | 0.002149 | 4.8e-05 | 100.5127 | 97.71 | 115.3154 | 21.2612 | 0.002925 | 0.002636 | 250.1575 | 0.27215 | 8.403 | 0.00338 | 0.004981 | 0.004971 | 5e-05 | 0.001057 | 0.28385 | 1.3171 | 2.5479 | 98.87 | 20.8 | 38.618 | 92.179 | 103.1717 | 85.243 | 0.009606 | 0.98508 | 0.023313 | 23.6665 | 0.25295 | 0.1619 | 0.010289 | 0.71651 | 0.71651 | 0.009287
2019-03-21 15:54 | 1553198059.865 | 269 | 0.20739 | 0.002144 | 5.1e-05 | 100.465 | 97.0867 | 115.524 | 21.2612 | 0.003013 | 0.002648 | 250.1294 | 0.27242 | 8.4016 | 0.003375 | 0.005181 | 0.005077 | 5.4e-05 | 0.001057 | 0.26089 | 1.3376 | 2.548 | 98.87 | 20.8 | 38.617 | 94.3028 | 103.0508 | 85.5163 | 0.009606 | 0.98433 | 0.023326 | 23.697 | 0.30595 | 0.16206 | 0.010394 | 0.72406 | 0.72406 | 0.009199
2019-03-21 15:54 | 1553198060.8644 | 270 | 0.2074 | 0.002134 | 5.5e-05 | 100.4775 | 97.397 | 115.5203 | 21.2612 | 0.002952 | 0.002679 | 250.0996 | 0.27208 | 8.4026 | 0.003403 | 0.005109 | 0.005073 | 5.5e-05 | 0.001057 | 0.3027 | 1.331 | 2.5479 | 98.87 | 20.8 | 38.616 | 92.8499 | 103.733 | 86.257 | 0.009606 | 0.9823 | 0.023301 | 23.721 | 0.27852 | 0.16186 | 0.010301 | 0.71656 | 0.71656 | 0.009284
2019-03-21 15:54 | 1553198061.8641 | 271 | 0.20738 | 0.002125 | 5.9e-05 | 100.5115 | 97.6368 | 115.7871 | 21.2612 | 0.002901 | 0.002648 | 250.0329 | 0.27189 | 8.4013 | 0.003464 | 0.004734 | 0.004667 | 5.1e-05 | 0.001056 | 0.25642 | 1.3116 | 2.548 | 98.87 | 20.8 | 38.617 | 91.608 | 105.2237 | 85.5122 | 0.009606 | 0.98434 | 0.023352 | 23.7238 | 0.25955 | 0.16174 | 0.010466 | 0.72945 | 0.72945 | 0.009113
2019-03-21 15:54 | 1553198062.8652 | 272 | 0.20738 | 0.00212 | 6.2e-05 | 100.4519 | 97.8071 | 115.8099 | 21.2612 | 0.002938 | 0.002655 | 249.972 | 0.27257 | 8.4045 | 0.003445 | 0.005229 | 0.005199 | 5.2e-05 | 0.001054 | 0.27975 | 1.3403 | 2.548 | 98.87 | 20.8 | 38.617 | 92.5058 | 104.7636 | 85.6888 | 0.009606 | 0.98386 | 0.023349 | 23.732 | 0.23866 | 0.16215 | 0.010468 | 0.72929 | 0.72929 | 0.009138
2019-03-21 15:54 | 1553198063.8642 | 273 | 0.20736 | 0.002118 | 6.4e-05 | 100.5273 | 97.8588 | 115.8512 | 21.2612 | 0.00294 | 0.002636 | 249.9164 | 0.27198 | 8.4027 | 0.003454 | 0.004378 | 0.00435 | 5.3e-05 | 0.001055 | 0.28102 | 1.3125 | 2.548 | 98.87 | 20.8 | 38.617 | 92.5567 | 104.9854 | 85.2267 | 0.009606 | 0.98512 | 0.023368 | 23.7209 | 0.24064 | 0.1618 | 0.010538 | 0.73464 | 0.73464 | 0.009052
2019-03-21 15:54 | 1553198064.8642 | 274 | 0.20741 | 0.002111 | 6.6e-05 | 100.586 | 98.1214 | 115.717 | 21.2612 | 0.002916 | 0.002653 | 249.8533 | 0.2713 | 8.4041 | 0.003426 | 0.004795 | 0.004706 | 5e-05 | 0.001055 | 0.30397 | 1.3276 | 2.5481 | 98.87 | 20.8 | 38.619 | 91.9811 | 104.2982 | 85.6481 | 0.009606 | 0.98397 | 0.023341 | 23.7211 | 0.22189 | 0.16139 | 0.010235 | 0.71255 | 0.71255 | 0.009309
2019-03-21 15:54 | 1553198065.8643 | 275 | 0.20738 | 0.002109 | 6.6e-05 | 100.5589 | 98.2406 | 115.452 | 21.2612 | 0.002931 | 0.002655 | 249.8084 | 0.2726 | 8.4002 | 0.003449 | 0.004861 | 0.004876 | 4.9e-05 | 0.001055 | 0.30862 | 1.3087 | 2.548 | 98.87 | 20.8 | 38.621 | 92.3431 | 104.8613 | 85.695 | 0.009607 | 0.98384 | 0.023313 | 23.6955 | 0.20862 | 0.16217 | 0.010475 | 0.7284 | 0.7284 | 0.009151
2019-03-21 15:54 | 1553198066.8637 | 276 | 0.20739 | 0.002104 | 6.7e-05 | 100.6149 | 98.2699 | 115.267 | 21.2612 | 0.002924 | 0.002671 | 249.753 | 0.27257 | 8.4043 | 0.00346 | 0.004746 | 0.004655 | 5e-05 | 0.001052 | 0.27776 | 1.3245 | 2.548 | 98.87 | 20.8 | 38.625 | 92.1664 | 105.1216 | 86.0797 | 0.009608 | 0.98279 | 0.023281 | 23.6892 | 0.21093 | 0.16215 | 0.010404 | 0.72241 | 0.72241 | 0.009225
2019-03-21 15:54 | 1553198067.864 | 277 | 0.20742 | 0.002103 | 6.7e-05 | 100.5553 | 98.2791 | 115.2717 | 21.2612 | 0.002908 | 0.002615 | 249.6863 | 0.27252 | 8.4043 | 0.003384 | 0.005383 | 0.00533 | 5.1e-05 | 0.001053 | 0.29407 | 1.3361 | 2.548 | 98.87 | 20.8 | 38.626 | 91.777 | 103.2889 | 84.7264 | 0.009608 | 0.9865 | 0.023326 | 23.645 | 0.2048 | 0.16212 | 0.010223 | 0.71113 | 0.71113 | 0.00937
2019-03-21 15:54 | 1553198068.8635 | 278 | 0.20736 | 0.002095 | 7e-05 | 100.6 | 98.3789 | 114.9001 | 21.2612 | 0.002978 | 0.002638 | 249.6394 | 0.27208 | 8.3993 | 0.00334 | 0.004543 | 0.004486 | 5.6e-05 | 0.001054 | 0.28982 | 1.3152 | 2.548 | 98.87 | 20.81 | 38.627 | 93.4587 | 102.2039 | 85.2813 | 0.009614 | 0.98497 | 0.02327 | 23.6251 | 0.19969 | 0.16186 | 0.010557 | 0.73254 | 0.73254 | 0.009081
2019-03-21 15:54 | 1553198069.8645 | 279 | 0.20736 | 0.00209 | 7.4e-05 | 100.673 | 98.4007 | 114.8161 | 21.2612 | 0.00291 | 0.002686 | 249.5865 | 0.27265 | 8.4041 | 0.003409 | 0.004598 | 0.00456 | 5.1e-05 | 0.001053 | 0.29391 | 1.2778 | 2.548 | 98.87 | 20.81 | 38.629 | 91.8396 | 103.8955 | 86.4396 | 0.009615 | 0.9818 | 0.023224 | 23.6546 | 0.2042 | 0.1622 | 0.010608 | 0.7344 | 0.7344 | 0.009077
2019-03-21 15:54 | 1553198070.8645 | 280 | 0.20736 | 0.002083 | 8e-05 | 100.6566 | 98.1691 | 115.0046 | 21.2612 | 0.002915 | 0.002654 | 249.5568 | 0.27285 | 8.4031 | 0.00342 | 0.004901 | 0.004798 | 5.3e-05 | 0.001055 | 0.29983 | 1.3084 | 2.548 | 98.87 | 20.81 | 38.629 | 91.9584 | 104.148 | 85.6669 | 0.009615 | 0.98391 | 0.023268 | 23.6486 | 0.22382 | 0.16232 | 0.010581 | 0.73358 | 0.73358 | 0.009094
2019-03-21 15:55 | 1553198071.8636 | 281 | 0.20732 | 0.002074 | 8.7e-05 | 100.7245 | 98.195 | 115.1839 | 21.2612 | 0.002868 | 0.002668 | 249.5078 | 0.27312 | 8.4026 | 0.003404 | 0.005544 | 0.005462 | 5.2e-05 | 0.001051 | 0.29214 | 1.2853 | 2.548 | 98.87 | 20.8 | 38.629 | 90.8208 | 103.7729 | 85.9945 | 0.009609 | 0.98302 | 0.023276 | 23.6778 | 0.2275 | 0.16248 | 0.010827 | 0.75058 | 0.75058 | 0.008897
2019-03-21 15:55 | 1553198072.8637 | 282 | 0.2073 | 0.002077 | 9.4e-05 | 100.6447 | 97.9345 | 115.1632 | 21.2612 | 0.0029 | 0.002665 | 249.496 | 0.27221 | 8.4016 | 0.003356 | 0.004973 | 0.004922 | 5e-05 | 0.001052 | 0.28368 | 1.3013 | 2.5479 | 98.87 | 20.81 | 38.631 | 91.5909 | 102.5975 | 85.9361 | 0.009615 | 0.98318 | 0.023276 | 23.6738 | 0.24417 | 0.16194 | 0.010967 | 0.7599 | 0.7599 | 0.008759
2019-03-21 15:55 | 1553198073.8642 | 283 | 0.20727 | 0.002075 | 9.9e-05 | 100.7877 | 97.5874 | 114.9128 | 21.2612 | 0.002872 | 0.002634 | 249.4572 | 0.27274 | 8.4018 | 0.00341 | 0.00539 | 0.005345 | 5.2e-05 | 0.00105 | 0.27936 | 1.3259 | 2.5481 | 98.87 | 20.81 | 38.632 | 90.9107 | 103.9028 | 85.1774 | 0.009615 | 0.98526 | 0.023275 | 23.623 | 0.28862 | 0.16225 | 0.011091 | 0.76821 | 0.76821 | 0.008681
2019-03-21 15:55 | 1553198074.8638 | 284 | 0.20725 | 0.002073 | 0.000105 | 100.7639 | 97.5196 | 114.8436 | 21.2612 | 0.002895 | 0.002688 | 249.4374 | 0.27206 | 8.4036 | 0.003473 | 0.004796 | 0.004739 | 5.2e-05 | 0.00105 | 0.27815 | 1.3248 | 2.5481 | 98.87 | 20.81 | 38.636 | 91.476 | 105.4402 | 86.4731 | 0.009616 | 0.98171 | 0.023226 | 23.6586 | 0.29272 | 0.16185 | 0.011255 | 0.7777 | 0.7777 | 0.008553
2019-03-21 15:55 | 1553198075.8648 | 285 | 0.20724 | 0.002077 | 0.000109 | 100.7296 | 97.4117 | 114.6884 | 21.2612 | 0.00291 | 0.00268 | 249.3497 | 0.27275 | 8.4058 | 0.003482 | 0.004568 | 0.004469 | 5.3e-05 | 0.00105 | 0.27201 | 1.3255 | 2.548 | 98.87 | 20.81 | 38.638 | 91.8309 | 105.6439 | 86.2943 | 0.009617 | 0.9822 | 0.023216 | 23.6367 | 0.29959 | 0.16226 | 0.011311 | 0.78106 | 0.78106 | 0.008538
2019-03-21 15:55 | 1553198076.8641 | 286 | 0.20725 | 0.002083 | 0.000113 | 100.7406 | 97.0707 | 114.3321 | 21.2612 | 0.002942 | 0.002754 | 249.3059 | 0.27213 | 8.4038 | 0.003488 | 0.004684 | 0.004625 | 5.2e-05 | 0.001048 | 0.29159 | 1.3206 | 2.5479 | 98.87 | 20.81 | 38.64 | 92.5972 | 105.8006 | 88.0746 | 0.009617 | 0.97736 | 0.023123 | 23.6583 | 0.33193 | 0.16189 | 0.011236 | 0.77241 | 0.77241 | 0.008614
2019-03-21 15:55 | 1553198077.8639 | 287 | 0.20725 | 0.002081 | 0.000117 | 100.7448 | 97.5213 | 113.9777 | 21.2612 | 0.002904 | 0.002666 | 249.2432 | 0.27322 | 8.3996 | 0.003513 | 0.005013 | 0.004922 | 4.9e-05 | 0.001044 | 0.2792 | 1.32 | 2.5481 | 98.87 | 20.81 | 38.642 | 91.6948 | 106.4121 | 85.944 | 0.009618 | 0.98316 | 0.023155 | 23.5519 | 0.29087 | 0.16254 | 0.011176 | 0.7691 | 0.7691 | 0.008686
2019-03-21 15:55 | 1553198078.8645 | 288 | 0.20711 | 0.002089 | 0.00012 | 100.698 | 97.8926 | 113.6483 | 21.2612 | 0.002965 | 0.002698 | 249.1849 | 0.27203 | 8.4018 | 0.003488 | 0.005333 | 0.00529 | 5.1e-05 | 0.001044 | 0.28799 | 1.3697 | 2.548 | 98.87 | 20.81 | 38.643 | 93.1558 | 105.8013 | 86.7199 | 0.009618 | 0.98104 | 0.023097 | 23.5432 | 0.25273 | 0.16183 | 0.012006 | 0.82442 | 0.82442 | 0.008068
2019-03-21 15:55 | 1553198079.8651 | 289 | 0.20698 | 0.002091 | 0.000121 | 100.7814 | 98.2628 | 112.9862 | 21.2612 | 0.002934 | 0.002753 | 249.1068 | 0.27325 | 8.4034 | 0.003554 | 0.005259 | 0.005274 | 5.3e-05 | 0.001045 | 0.28744 | 1.3762 | 2.5479 | 98.87 | 20.81 | 38.648 | 92.408 | 107.4105 | 88.0433 | 0.009619 | 0.97744 | 0.022987 | 23.5176 | 0.22638 | 0.16256 | 0.012805 | 0.87542 | 0.87542 | 0.007632
2019-03-21 15:55 | 1553198080.8651 | 290 | 0.20655 | 0.002095 | 0.000121 | 100.7805 | 98.3592 | 112.3691 | 21.2612 | 0.00295 | 0.002731 | 249.0242 | 0.27286 | 8.4024 | 0.003517 | 0.005015 | 0.004999 | 5.4e-05 | 0.001046 | 0.27057 | 1.32 | 2.5478 | 98.87 | 20.81 | 38.65 | 92.785 | 106.497 | 87.51 | 0.00962 | 0.97889 | 0.022941 | 23.436 | 0.21753 | 0.16232 | 0.01539 | 1.0512 | 1.0512 | 0.006346
2019-03-21 15:55 | 1553198081.8632 | 291 | 0.20616 | 0.002109 | 0.00012 | 100.827 | 98.5659 | 111.5727 | 21.2612 | 0.002965 | 0.002735 | 248.942 | 0.2721 | 8.4021 | 0.003548 | 0.004697 | 0.004586 | 5e-05 | 0.001047 | 0.27527 | 1.304 | 2.5478 | 98.87 | 20.81 | 38.651 | 93.1471 | 107.2488 | 87.6123 | 0.00962 | 0.97861 | 0.022857 | 23.3561 | 0.20287 | 0.16187 | 0.017732 | 1.2075 | 1.2075 | 0.00551
2019-03-21 15:55 | 1553198082.8639 | 292 | 0.20575 | 0.00214 | 0.000121 | 100.81 | 98.9552 | 111.0485 | 21.2612 | 0.002948 | 0.002756 | 248.835 | 0.27241 | 8.4011 | 0.003557 | 0.005377 | 0.005293 | 5e-05 | 0.001048 | 0.28694 | 1.2971 | 2.5479 | 98.87 | 20.81 | 38.657 | 92.7327 | 107.4758 | 88.1175 | 0.009622 | 0.97724 | 0.022787 | 23.3175 | 0.16611 | 0.16206 | 0.02013 | 1.3672 | 1.3672 | 0.004872
2019-03-21 15:55 | 1553198083.8632 | 293 | 0.20521 | 0.0022 | 0.000121 | 100.8187 | 99.3217 | 110.4958 | 21.2612 | 0.002961 | 0.00278 | 248.7169 | 0.27243 | 8.4006 | 0.003578 | 0.004781 | 0.004679 | 5.4e-05 | 0.001046 | 0.27793 | 1.2816 | 2.5479 | 98.87 | 20.81 | 38.66 | 93.0697 | 107.9822 | 88.7063 | 0.009622 | 0.97565 | 0.022712 | 23.2783 | 0.13381 | 0.16207 | 0.02328 | 1.5765 | 1.5765 | 0.004225
2019-03-21 15:55 | 1553198084.8643 | 294 | 0.20492 | 0.002293 | 0.000121 | 100.8097 | 99.1514 | 109.9099 | 21.2612 | 0.003027 | 0.002803 | 248.5836 | 0.27139 | 8.4022 | 0.003573 | 0.005622 | 0.005627 | 5.3e-05 | 0.001042 | 0.28064 | 1.3023 | 2.548 | 98.87 | 20.81 | 38.663 | 94.6582 | 107.8734 | 89.2471 | 0.009623 | 0.9742 | 0.022634 | 23.2339 | 0.14837 | 0.16145 | 0.024906 | 1.681 | 1.681 | 0.003947
2019-03-21 15:55 | 1553198085.8644 | 295 | 0.20449 | 0.002425 | 0.000123 | 100.8297 | 98.7556 | 109.1898 | 21.2612 | 0.003087 | 0.002814 | 248.483 | 0.27208 | 8.4019 | 0.0035 | 0.004948 | 0.004891 | 5.3e-05 | 0.001035 | 0.29623 | 1.3083 | 2.548 | 98.87 | 20.81 | 38.667 | 96.1052 | 106.0798 | 89.5173 | 0.009624 | 0.97347 | 0.022552 | 23.1663 | 0.18591 | 0.16186 | 0.027362 | 1.8401 | 1.8401 | 0.003615
2019-03-21 15:55 | 1553198086.8634 | 296 | 0.2038 | 0.002585 | 0.000128 | 100.799 | 98.2359 | 108.5086 | 21.2612 | 0.003081 | 0.00289 | 248.3876 | 0.27192 | 8.4029 | 0.00349 | 0.00525 | 0.005163 | 5.2e-05 | 0.001033 | 0.29767 | 1.3316 | 2.548 | 98.87 | 20.81 | 38.668 | 95.9464 | 105.8578 | 91.3439 | 0.009624 | 0.96859 | 0.022425 | 23.152 | 0.23039 | 0.16176 | 0.031252 | 2.0897 | 2.0897 | 0.003182
2019-03-21 15:55 | 1553198087.8645 | 297 | 0.20315 | 0.00285 | 0.000133 | 100.8938 | 97.7621 | 107.9706 | 21.2612 | 0.003022 | 0.002911 | 248.2453 | 0.27179 | 8.4026 | 0.003552 | 0.005029 | 0.004949 | 5.3e-05 | 0.001033 | 0.28567 | 1.3374 | 2.5481 | 98.87 | 20.81 | 38.673 | 94.54 | 107.3592 | 91.8505 | 0.009626 | 0.96725 | 0.022354 | 23.1106 | 0.28203 | 0.16169 | 0.034794 | 2.3189 | 2.3189 | 0.002866
2019-03-21 15:55 | 1553198088.8645 | 298 | 0.20248 | 0.003175 | 0.00014 | 100.8226 | 97.3013 | 107.696 | 21.2612 | 0.003138 | 0.0029 | 248.1093 | 0.27282 | 8.4002 | 0.003575 | 0.005034 | 0.004912 | 5.4e-05 | 0.001035 | 0.29789 | 1.3252 | 2.548 | 98.87 | 20.8 | 38.675 | 97.3258 | 107.905 | 91.5902 | 0.00962 | 0.96794 | 0.022333 | 23.0729 | 0.31798 | 0.1623 | 0.038373 | 2.5545 | 2.5545 | 0.002611
2019-03-21 15:55 | 1553198089.8641 | 299 | 0.20206 | 0.003508 | 0.000147 | 100.8389 | 96.5592 | 107.2261 | 21.2612 | 0.003131 | 0.002949 | 247.9904 | 0.27152 | 8.4035 | 0.003639 | 0.005337 | 0.005341 | 5.3e-05 | 0.001034 | 0.27312 | 1.3222 | 2.5482 | 98.87 | 20.8 | 38.677 | 97.1722 | 109.4717 | 92.7662 | 0.009621 | 0.96483 | 0.022249 | 23.0596 | 0.38791 | 0.16153 | 0.040504 | 2.6856 | 2.6856 | 0.002472
2019-03-21 15:55 | 1553198090.8637 | 300 | 0.20167 | 0.003851 | 0.000154 | 100.8237 | 95.614 | 107.0399 | 21.2612 | 0.003159 | 0.002976 | 247.8667 | 0.27171 | 8.405 | 0.003683 | 0.005013 | 0.004924 | 5.2e-05 | 0.001033 | 0.28047 | 1.3267 | 2.548 | 98.87 | 20.8 | 38.679 | 97.8297 | 110.535 | 93.4124 | 0.009621 | 0.96313 | 0.02221 | 23.0599 | 0.47455 | 0.16164 | 0.042405 | 2.8062 | 2.8062 | 0.002367
2019-03-21 15:55 | 1553198091.8648 | 301 | 0.20138 | 0.004166 | 0.000159 | 100.8401 | 95.6362 | 107.3291 | 21.2612 | 0.003232 | 0.002983 | 247.752 | 0.27207 | 8.3997 | 0.003592 | 0.005447 | 0.005426 | 5.3e-05 | 0.001036 | 0.30287 | 1.2977 | 2.5479 | 98.87 | 20.8 | 38.678 | 99.5932 | 108.3217 | 93.586 | 0.009621 | 0.96267 | 0.022234 | 23.0965 | 0.47393 | 0.16185 | 0.043716 | 2.8957 | 2.8957 | 0.002297
2019-03-21 15:55 | 1553198092.8648 | 302 | 0.20111 | 0.004454 | 0.000164 | 100.8625 | 95.2476 | 107.5875 | 21.2612 | 0.00322 | 0.003042 | 247.6339 | 0.27228 | 8.4038 | 0.003638 | 0.005324 | 0.005297 | 5.4e-05 | 0.001032 | 0.27959 | 1.3393 | 2.548 | 98.87 | 20.8 | 38.682 | 99.3038 | 109.4503 | 95.0208 | 0.009622 | 0.95892 | 0.022218 | 23.1695 | 0.51233 | 0.16198 | 0.044986 | 2.9772 | 2.9772 | 0.002236
2019-03-21 15:55 | 1553198093.8638 | 303 | 0.20095 | 0.004782 | 0.000166 | 100.8109 | 95.0708 | 107.9875 | 21.2612 | 0.003218 | 0.002974 | 247.5125 | 0.2721 | 8.4013 | 0.003685 | 0.004972 | 0.004961 | 5.4e-05 | 0.001036 | 0.28119 | 1.3081 | 2.548 | 98.87 | 20.8 | 38.684 | 99.2668 | 110.5838 | 93.3716 | 0.009622 | 0.96323 | 0.022309 | 23.1605 | 0.52437 | 0.16187 | 0.045541 | 3.0261 | 3.0261 | 0.002199
2019-03-21 15:55 | 1553198094.8649 | 304 | 0.20092 | 0.004998 | 0.000167 | 100.7734 | 94.8474 | 108.557 | 21.2612 | 0.003263 | 0.003012 | 247.421 | 0.27234 | 8.4006 | 0.003669 | 0.005194 | 0.005127 | 5.2e-05 | 0.001034 | 0.30879 | 1.3006 | 2.5479 | 98.87 | 20.8 | 38.686 | 100.3472 | 110.2028 | 94.299 | 0.009623 | 0.9608 | 0.022339 | 23.2508 | 0.54209 | 0.16201 | 0.045483 | 3.0263 | 3.0263 | 0.0022
2019-03-21 15:55 | 1553198095.8643 | 305 | 0.20093 | 0.005208 | 0.000166 | 100.8246 | 94.6839 | 108.9227 | 21.2612 | 0.00331 | 0.003023 | 247.3318 | 0.27161 | 8.4021 | 0.003676 | 0.005235 | 0.005166 | 5.6e-05 | 0.001032 | 0.28799 | 1.2931 | 2.5479 | 98.87 | 20.8 | 38.688 | 101.4807 | 110.3582 | 94.5525 | 0.009623 | 0.96014 | 0.022369 | 23.298 | 0.56207 | 0.16158 | 0.045148 | 3.0082 | 3.0082 | 0.002208
2019-03-21 15:55 | 1553198096.863 | 306 | 0.20105 | 0.005414 | 0.000164 | 100.81 | 94.593 | 109.5834 | 21.2612 | 0.003306 | 0.003033 | 247.2282 | 0.27206 | 8.4018 | 0.00366 | 0.005703 | 0.005647 | 5.4e-05 | 0.001036 | 0.30055 | 1.2632 | 2.548 | 98.87 | 20.8 | 38.689 | 101.3904 | 109.9756 | 94.7928 | 0.009624 | 0.95951 | 0.02243 | 23.3762 | 0.56937 | 0.16185 | 0.044176 | 2.9516 | 2.9516 | 0.002254
2019-03-21 15:55 | 1553198097.8631 | 307 | 0.20124 | 0.005593 | 0.00016 | 100.8339 | 94.4485 | 110.0667 | 21.2612 | 0.003276 | 0.003073 | 247.1606 | 0.27256 | 8.4036 | 0.003715 | 0.005248 | 0.005176 | 5.4e-05 | 0.001039 | 0.2724 | 1.2974 | 2.548 | 98.87 | 20.8 | 38.692 | 100.6573 | 111.3144 | 95.7528 | 0.009624 | 0.95701 | 0.02245 | 23.4582 | 0.58515 | 0.16214 | 0.042856 | 2.8663 | 2.8663 | 0.002325
2019-03-21 15:55 | 1553198098.8642 | 308 | 0.20161 | 0.005747 | 0.000156 | 100.794 | 94.5991 | 110.5052 | 21.2612 | 0.00332 | 0.00308 | 247.0837 | 0.27117 | 8.4038 | 0.003739 | 0.005428 | 0.005375 | 5.3e-05 | 0.001034 | 0.27383 | 1.2921 | 2.5479 | 98.87 | 20.8 | 38.693 | 101.728 | 111.9081 | 95.9251 | 0.009625 | 0.95657 | 0.022489 | 23.5104 | 0.56736 | 0.16132 | 0.040452 | 2.7106 | 2.7106 | 0.002446
2019-03-21 15:55 | 1553198099.8636 | 309 | 0.20199 | 0.005855 | 0.000151 | 100.7606 | 95.173 | 110.9819 | 21.2612 | 0.003325 | 0.003044 | 246.9807 | 0.27217 | 8.4012 | 0.003814 | 0.004678 | 0.0047 | 5.3e-05 | 0.001037 | 0.28313 | 1.3278 | 2.5479 | 98.87 | 20.8 | 38.694 | 101.8473 | 113.7163 | 95.0551 | 0.009625 | 0.95883 | 0.022564 | 23.5333 | 0.5103 | 0.16192 | 0.038053 | 2.5588 | 2.5588 | 0.002601
2019-03-21 15:55 | 1553198100.8632 | 310 | 0.20239 | 0.005923 | 0.000146 | 100.7043 | 95.6927 | 111.4023 | 21.2612 | 0.003337 | 0.003074 | 246.8995 | 0.27258 | 8.4039 | 0.003764 | 0.005568 | 0.005432 | 5.4e-05 | 0.001037 | 0.2959 | 1.2812 | 2.548 | 98.87 | 20.8 | 38.698 | 102.1458 | 112.5178 | 95.7779 | 0.009626 | 0.95695 | 0.022585 | 23.6009 | 0.45658 | 0.16216 | 0.035616 | 2.3974 | 2.3974 | 0.00278
2019-03-21 15:55 | 1553198101.8633 | 311 | 0.20285 | 0.005947 | 0.000142 | 100.7925 | 96.2543 | 111.4539 | 21.2612 | 0.003313 | 0.003017 | 246.8276 | 0.27225 | 8.4014 | 0.003787 | 0.005543 | 0.005402 | 5.1e-05 | 0.001036 | 0.28783 | 1.2976 | 2.5479 | 98.87 | 20.81 | 38.699 | 101.5506 | 113.0605 | 94.4116 | 0.009632 | 0.96051 | 0.022632 | 23.5626 | 0.41207 | 0.16196 | 0.032824 | 2.2143 | 2.2143 | 0.003006
2019-03-21 15:55 | 1553198102.8634 | 312 | 0.20327 | 0.005936 | 0.000137 | 100.7219 | 96.5105 | 111.5391 | 21.2612 | 0.003311 | 0.003036 | 246.7565 | 0.27165 | 8.4013 | 0.003744 | 0.005383 | 0.005379 | 5.3e-05 | 0.001038 | 0.30934 | 1.2938 | 2.5481 | 98.87 | 20.8 | 38.7 | 101.5058 | 112.0069 | 94.8712 | 0.009626 | 0.95931 | 0.022627 | 23.5864 | 0.38203 | 0.16161 | 0.030347 | 2.0469 | 2.0469 | 0.003245
2019-03-21 15:55 | 1553198103.8644 | 313 | 0.20345 | 0.005884 | 0.000133 | 100.6387 | 96.5289 | 111.6652 | 21.2612 | 0.003257 | 0.002985 | 246.7073 | 0.27216 | 8.404 | 0.003684 | 0.00468 | 0.004583 | 5.1e-05 | 0.001036 | 0.26377 | 1.3133 | 2.5481 | 98.86 | 20.8 | 38.701 | 100.2169 | 110.5688 | 93.6323 | 0.009628 | 0.96255 | 0.022678 | 23.5599 | 0.37294 | 0.16191 | 0.029336 | 1.9834 | 1.9834 | 0.003355
2019-03-21 15:55 | 1553198104.8638 | 314 | 0.20364 | 0.005818 | 0.000129 | 100.6375 | 97.0585 | 111.7776 | 21.2612 | 0.003284 | 0.002987 | 246.6366 | 0.2728 | 8.4057 | 0.003698 | 0.005785 | 0.005735 | 5.1e-05 | 0.001036 | 0.24043 | 1.3251 | 2.5479 | 98.86 | 20.8 | 38.701 | 100.8699 | 110.8971 | 93.6731 | 0.009628 | 0.96244 | 0.022688 | 23.5731 | 0.3239 | 0.16229 | 0.028226 | 1.9094 | 1.9094 | 0.003493
2019-03-21 15:55 | 1553198105.8639 | 315 | 0.20379 | 0.005722 | 0.000125 | 100.5881 | 97.57 | 112.224 | 21.2612 | 0.003293 | 0.003008 | 246.5516 | 0.27142 | 8.4023 | 0.00368 | 0.005744 | 0.005699 | 5.4e-05 | 0.001041 | 0.30541 | 1.3227 | 2.5479 | 98.86 | 20.8 | 38.701 | 101.0795 | 110.4521 | 94.194 | 0.009628 | 0.96108 | 0.022717 | 23.6369 | 0.27248 | 0.16147 | 0.027499 | 1.8627 | 1.8627 | 0.003563
2019-03-21 15:55 | 1553198106.8646 | 316 | 0.20393 | 0.005618 | 0.000122 | 100.6071 | 97.8009 | 112.3616 | 21.2612 | 0.003269 | 0.003055 | 246.4992 | 0.27271 | 8.4023 | 0.0037 | 0.004366 | 0.004314 | 5.2e-05 | 0.001041 | 0.25769 | 1.3162 | 2.5479 | 98.86 | 20.8 | 38.701 | 100.4892 | 110.9365 | 95.3369 | 0.009628 | 0.9581 | 0.022695 | 23.6882 | 0.25303 | 0.16224 | 0.026794 | 1.8134 | 1.8134 | 0.003677
2019-03-21 15:55 | 1553198107.8647 | 317 | 0.20405 | 0.005508 | 0.000118 | 100.6563 | 98.0353 | 112.5111 | 21.2612 | 0.003298 | 0.003067 | 246.4381 | 0.27199 | 8.4026 | 0.003675 | 0.005565 | 0.005469 | 5e-05 | 0.001043 | 0.26913 | 1.3341 | 2.5479 | 98.86 | 20.8 | 38.7 | 101.2043 | 110.3375 | 95.6229 | 0.009627 | 0.95735 | 0.022702 | 23.7131 | 0.236 | 0.16181 | 0.026168 | 1.7717 | 1.7717 | 0.003754
2019-03-21 15:55 | 1553198108.8636 | 318 | 0.20423 | 0.005391 | 0.000117 | 100.7007 | 98.5083 | 112.5759 | 21.2612 | 0.003361 | 0.003061 | 246.373 | 0.27239 | 8.4018 | 0.003712 | 0.004982 | 0.004904 | 4.9e-05 | 0.001045 | 0.28224 | 1.3486 | 2.5481 | 98.86 | 20.8 | 38.699 | 102.7152 | 111.2473 | 95.481 | 0.009627 | 0.95772 | 0.022713 | 23.7154 | 0.19688 | 0.16205 | 0.02524 | 1.7097 | 1.7097 | 0.003895
2019-03-21 15:55 | 1553198109.8641 | 319 | 0.20441 | 0.005273 | 0.000115 | 100.6071 | 98.7519 | 112.7263 | 21.2612 | 0.003376 | 0.003064 | 246.3365 | 0.27101 | 8.4012 | 0.003672 | 0.005558 | 0.00547 | 4.9e-05 | 0.001043 | 0.28479 | 1.3402 | 2.5482 | 98.86 | 20.8 | 38.699 | 103.0932 | 110.277 | 95.5343 | 0.009627 | 0.95758 | 0.022726 | 23.7329 | 0.16648 | 0.16122 | 0.02434 | 1.6497 | 1.6497 | 0.004017
2019-03-21 15:55 | 1553198110.8638 | 320 | 0.20447 | 0.005167 | 0.000114 | 100.4793 | 99.2127 | 112.5069 | 21.2612 | 0.003341 | 0.003059 | 246.2542 | 0.27213 | 8.4023 | 0.003607 | 0.004183 | 0.004051 | 5e-05 | 0.001042 | 0.26593 | 1.3068 | 2.548 | 98.86 | 20.8 | 38.698 | 102.2283 | 108.6882 | 95.4294 | 0.009627 | 0.95785 | 0.022707 | 23.7064 | 0.11347 | 0.16189 | 0.024071 | 1.6301 | 1.6301 | 0.004082
2019-03-21 15:55 | 1553198111.8638 | 321 | 0.2045 | 0.005064 | 0.000114 | 100.4597 | 99.2805 | 112.466 | 21.2612 | 0.003304 | 0.003035 | 246.1874 | 0.27241 | 8.4038 | 0.003682 | 0.005178 | 0.005128 | 5.1e-05 | 0.001042 | 0.26305 | 1.2795 | 2.5479 | 98.86 | 20.8 | 38.697 | 101.3386 | 110.5111 | 94.8406 | 0.009627 | 0.95939 | 0.022721 | 23.6832 | 0.10561 | 0.16206 | 0.024028 | 1.6283 | 1.6283 | 0.004091
2019-03-21 15:55 | 1553198112.8633 | 322 | 0.20451 | 0.004963 | 0.000114 | 100.4853 | 99.3411 | 112.4252 | 21.2612 | 0.003363 | 0.003034 | 246.115 | 0.27163 | 8.4033 | 0.003598 | 0.005095 | 0.005026 | 5.2e-05 | 0.001044 | 0.28783 | 1.3012 | 2.5479 | 98.86 | 20.8 | 38.695 | 102.782 | 108.4754 | 94.8266 | 0.009626 | 0.95942 | 0.022718 | 23.6784 | 0.10243 | 0.16159 | 0.024099 | 1.6328 | 1.6328 | 0.004068
2019-03-21 15:55 | 1553198113.8643 | 323 | 0.20455 | 0.00487 | 0.000114 | 100.4263 | 99.3295 | 112.3691 | 21.2612 | 0.003354 | 0.002997 | 246.0382 | 0.27228 | 8.4025 | 0.003582 | 0.005161 | 0.005089 | 5.1e-05 | 0.001042 | 0.27489 | 1.2962 | 2.5481 | 98.86 | 20.8 | 38.694 | 102.5464 | 108.0766 | 93.9217 | 0.009626 | 0.96179 | 0.02274 | 23.6434 | 0.098223 | 0.16198 | 0.023989 | 1.6269 | 1.6269 | 0.004092
2019-03-21 15:55 | 1553198114.863 | 324 | 0.20457 | 0.004793 | 0.000116 | 100.4814 | 99.4814 | 111.9393 | 21.2612 | 0.003373 | 0.003015 | 245.9683 | 0.2714 | 8.403 | 0.003608 | 0.005406 | 0.005329 | 5.1e-05 | 0.001042 | 0.30663 | 1.3081 | 2.5479 | 98.86 | 20.8 | 38.692 | 103.0106 | 108.7127 | 94.3687 | 0.009625 | 0.96062 | 0.022683 | 23.6125 | 0.089463 | 0.16145 | 0.023971 | 1.6215 | 1.6215 | 0.004092
2019-03-21 15:55 | 1553198115.864 | 325 | 0.20454 | 0.004717 | 0.000116 | 100.4772 | 99.3894 | 111.6027 | 21.2612 | 0.00326 | 0.003013 | 245.8998 | 0.2722 | 8.4023 | 0.00355 | 0.005688 | 0.005658 | 4.9e-05 | 0.00104 | 0.27168 | 1.3255 | 2.548 | 98.87 | 20.8 | 38.69 | 100.2689 | 107.2947 | 94.3132 | 0.009624 | 0.96076 | 0.02265 | 23.5752 | 0.097364 | 0.16193 | 0.024249 | 1.638 | 1.638 | 0.004063
2019-03-21 15:55 | 1553198116.8637 | 326 | 0.20455 | 0.004655 | 0.000118 | 100.3727 | 99.4473 | 111.5497 | 21.2612 | 0.003319 | 0.003045 | 245.8087 | 0.27252 | 8.4013 | 0.003623 | 0.004974 | 0.004836 | 4.8e-05 | 0.001041 | 0.28694 | 1.2856 | 2.548 | 98.87 | 20.8 | 38.69 | 101.7014 | 109.0664 | 95.0787 | 0.009624 | 0.95877 | 0.022621 | 23.5941 | 0.08284 | 0.16212 | 0.024236 | 1.635 | 1.635 | 0.004076
2019-03-21 15:55 | 1553198117.8635 | 327 | 0.20458 | 0.0046 | 0.000118 | 100.4025 | 99.573 | 111.4703 | 21.2612 | 0.003334 | 0.003066 | 245.7321 | 0.27242 | 8.4025 | 0.00361 | 0.005485 | 0.005472 | 5e-05 | 0.001037 | 0.28241 | 1.3058 | 2.548 | 98.87 | 20.8 | 38.689 | 102.0601 | 108.7552 | 95.5923 | 0.009624 | 0.95743 | 0.022597 | 23.6022 | 0.074198 | 0.16206 | 0.024143 | 1.6269 | 1.6269 | 0.004094
2019-03-21 15:55 | 1553198118.8631 | 328 | 0.2046 | 0.004562 | 0.000118 | 100.3229 | 99.5628 | 111.2497 | 21.2612 | 0.003351 | 0.00304 | 245.6661 | 0.27232 | 8.4015 | 0.00365 | 0.00484 | 0.004697 | 5e-05 | 0.001037 | 0.30326 | 1.298 | 2.5481 | 98.87 | 20.8 | 38.688 | 102.4772 | 109.7208 | 94.9682 | 0.009623 | 0.95905 | 0.022594 | 23.5589 | 0.068026 | 0.16201 | 0.02404 | 1.6198 | 1.6198 | 0.004111
2019-03-21 15:55 | 1553198119.8642 | 329 | 0.20459 | 0.004517 | 0.000119 | 100.3029 | 99.5291 | 111.3664 | 21.2612 | 0.003303 | 0.003047 | 245.5786 | 0.27104 | 8.4036 | 0.003634 | 0.005388 | 0.005356 | 5e-05 | 0.001038 | 0.29231 | 1.3199 | 2.5479 | 98.87 | 20.8 | 38.688 | 101.3253 | 109.3551 | 95.1241 | 0.009623 | 0.95865 | 0.022601 | 23.5762 | 0.069276 | 0.16124 | 0.024154 | 1.6279 | 1.6279 | 0.004071
2019-03-21 15:55 | 1553198120.8637 | 330 | 0.20461 | 0.004483 | 0.000119 | 100.2737 | 99.62 | 111.4734 | 21.2612 | 0.003309 | 0.003064 | 245.5255 | 0.27206 | 8.4014 | 0.003597 | 0.005225 | 0.005121 | 5.1e-05 | 0.001044 | 0.27577 | 1.36 | 2.5482 | 98.87 | 20.8 | 38.689 | 101.4579 | 108.4353 | 95.5438 | 0.009624 | 0.95756 | 0.022599 | 23.601 | 0.058503 | 0.16185 | 0.024095 | 1.6238 | 1.6238 | 0.004097
2019-03-21 15:55 | 1553198121.8644 | 331 | 0.20459 | 0.004454 | 0.000119 | 100.2097 | 99.4003 | 111.5002 | 21.2612 | 0.003304 | 0.003068 | 245.4616 | 0.27169 | 8.4009 | 0.003572 | 0.005436 | 0.005345 | 5.1e-05 | 0.001042 | 0.31542 | 1.3236 | 2.5479 | 98.87 | 20.8 | 38.686 | 101.3345 | 107.8309 | 95.6314 | 0.009623 | 0.95733 | 0.022599 | 23.6066 | 0.072533 | 0.16162 | 0.024263 | 1.6351 | 1.6351 | 0.004063
2019-03-21 15:55 | 1553198122.8638 | 332 | 0.20457 | 0.004422 | 0.00012 | 100.1567 | 99.3476 | 111.5337 | 21.2612 | 0.003282 | 0.003053 | 245.3813 | 0.27172 | 8.4028 | 0.003604 | 0.004921 | 0.004929 | 5.2e-05 | 0.001044 | 0.29336 | 1.3322 | 2.5478 | 98.87 | 20.8 | 38.684 | 100.8 | 108.6063 | 95.2797 | 0.009622 | 0.95824 | 0.022613 | 23.5989 | 0.07255 | 0.16165 | 0.024411 | 1.6461 | 1.6461 | 0.004036
2019-03-21 15:55 | 1553198123.8645 | 333 | 0.20449 | 0.004396 | 0.000122 | 100.1704 | 99.4525 | 111.5776 | 21.2612 | 0.003326 | 0.003047 | 245.3304 | 0.27252 | 8.4011 | 0.003593 | 0.004389 | 0.004337 | 5.3e-05 | 0.001045 | 0.30596 | 1.2758 | 2.5479 | 98.87 | 20.8 | 38.68 | 101.8732 | 108.3367 | 95.121 | 0.009621 | 0.95866 | 0.022623 | 23.5985 | 0.064335 | 0.16212 | 0.024943 | 1.6826 | 1.6826 | 0.00396
2019-03-21 15:55 | 1553198124.8644 | 334 | 0.20448 | 0.004383 | 0.000123 | 100.0402 | 99.3775 | 111.621 | 21.2612 | 0.003348 | 0.003071 | 245.2436 | 0.27274 | 8.4006 | 0.00355 | 0.005011 | 0.004895 | 4.9e-05 | 0.001044 | 0.29789 | 1.3397 | 2.5481 | 98.87 | 20.8 | 38.679 | 102.4104 | 107.2969 | 95.7081 | 0.009621 | 0.95713 | 0.022609 | 23.6219 | 0.059442 | 0.16225 | 0.02501 | 1.686 | 1.686 | 0.003955
2019-03-21 15:55 | 1553198125.8639 | 335 | 0.20442 | 0.004375 | 0.000126 | 100.0896 | 99.3636 | 111.7457 | 21.2612 | 0.003307 | 0.00307 | 245.1652 | 0.27261 | 8.4036 | 0.003591 | 0.005373 | 0.005358 | 5e-05 | 0.001047 | 0.25984 | 1.2784 | 2.5481 | 98.87 | 20.8 | 38.678 | 101.4076 | 108.2886 | 95.683 | 0.009621 | 0.9572 | 0.022623 | 23.6342 | 0.06512 | 0.16218 | 0.025386 | 1.7123 | 1.7123 | 0.003893
2019-03-21 15:55 | 1553198126.864 | 336 | 0.20437 | 0.00437 | 0.000128 | 100.0637 | 99.3993 | 111.5213 | 21.2612 | 0.00334 | 0.003082 | 245.0919 | 0.27239 | 8.4025 | 0.003556 | 0.004727 | 0.004696 | 5e-05 | 0.001045 | 0.27041 | 1.2954 | 2.5479 | 98.87 | 20.8 | 38.677 | 102.2094 | 107.4505 | 95.9681 | 0.009621 | 0.95646 | 0.022591 | 23.6196 | 0.059585 | 0.16204 | 0.025664 | 1.7285 | 1.7285 | 0.003853
2019-03-21 15:55 | 1553198127.8627 | 337 | 0.20431 | 0.004373 | 0.00013 | 99.9925 | 99.2587 | 111.4279 | 21.2612 | 0.003331 | 0.003024 | 245.0075 | 0.27136 | 8.4006 | 0.003525 | 0.004751 | 0.004674 | 5.3e-05 | 0.001044 | 0.28871 | 1.2964 | 2.5479 | 98.87 | 20.8 | 38.675 | 102.0044 | 106.7025 | 94.566 | 0.00962 | 0.9601 | 0.022625 | 23.5648 | 0.065884 | 0.16143 | 0.026023 | 1.7552 | 1.7552 | 0.00378
2019-03-21 15:55 | 1553198128.8627 | 338 | 0.20423 | 0.004373 | 0.000133 | 99.9335 | 98.9957 | 111.4218 | 21.2612 | 0.003349 | 0.003073 | 244.9185 | 0.27185 | 8.4012 | 0.003504 | 0.004933 | 0.004927 | 5.3e-05 | 0.001048 | 0.28655 | 1.2894 | 2.548 | 98.87 | 20.8 | 38.67 | 102.4254 | 106.1972 | 95.766 | 0.009619 | 0.95698 | 0.022587 | 23.6026 | 0.084331 | 0.16173 | 0.026462 | 1.7818 | 1.7818 | 0.00373
2019-03-21 15:55 | 1553198129.8642 | 339 | 0.20419 | 0.004384 | 0.000134 | 99.9707 | 99.0591 | 111.2584 | 21.2612 | 0.003326 | 0.003021 | 244.8442 | 0.2715 | 8.4022 | 0.003534 | 0.005306 | 0.005254 | 5.3e-05 | 0.001045 | 0.27223 | 1.2953 | 2.5479 | 98.87 | 20.8 | 38.668 | 101.8693 | 106.9229 | 94.4985 | 0.009618 | 0.96028 | 0.02261 | 23.5447 | 0.081942 | 0.16152 | 0.026682 | 1.7983 | 1.7983 | 0.003691
2019-03-21 15:56 | 1553198130.8632 | 340 | 0.20419 | 0.004401 | 0.000137 | 99.9201 | 98.7764 | 111.4706 | 21.2612 | 0.0034 | 0.003084 | 244.7573 | 0.27164 | 8.4017 | 0.003526 | 0.005221 | 0.005078 | 5.3e-05 | 0.001041 | 0.26681 | 1.2934 | 2.5479 | 98.87 | 20.8 | 38.666 | 103.6574 | 106.7236 | 96.0333 | 0.009618 | 0.95629 | 0.022584 | 23.6163 | 0.10297 | 0.1616 | 0.026669 | 1.7953 | 1.7953 | 0.0037
2019-03-21 15:56 | 1553198131.8628 | 341 | 0.20416 | 0.004412 | 0.000137 | 99.9111 | 98.7451 | 111.2539 | 21.2612 | 0.003361 | 0.003058 | 244.6638 | 0.27099 | 8.4029 | 0.003519 | 0.005307 | 0.005361 | 4.8e-05 | 0.001038 | 0.2907 | 1.3416 | 2.548 | 98.87 | 20.8 | 38.66 | 102.7263 | 106.5608 | 95.3934 | 0.009616 | 0.95795 | 0.022582 | 23.5729 | 0.10501 | 0.16121 | 0.026829 | 1.8058 | 1.8058 | 0.003669
2019-03-21 15:56 | 1553198132.8639 | 342 | 0.20412 | 0.004427 | 0.000139 | 99.7893 | 98.6109 | 111.2011 | 21.2612 | 0.003394 | 0.003058 | 244.5957 | 0.27164 | 8.401 | 0.003604 | 0.005208 | 0.005162 | 4.9e-05 | 0.001039 | 0.27151 | 1.3642 | 2.548 | 98.87 | 20.8 | 38.658 | 103.5326 | 108.6126 | 95.4067 | 0.009616 | 0.95791 | 0.022576 | 23.5677 | 0.10626 | 0.1616 | 0.027065 | 1.8211 | 1.8211 | 0.003647
2019-03-21 15:56 | 1553198133.863 | 343 | 0.20412 | 0.004446 | 0.000139 | 99.7938 | 98.4304 | 111.0284 | 21.2612 | 0.003392 | 0.003089 | 244.5037 | 0.27224 | 8.4011 | 0.003589 | 0.005037 | 0.00498 | 4.8e-05 | 0.001039 | 0.2672 | 1.3768 | 2.548 | 98.87 | 20.8 | 38.657 | 103.4742 | 108.2556 | 96.139 | 0.009616 | 0.95601 | 0.022536 | 23.5728 | 0.12305 | 0.16195 | 0.027047 | 1.8167 | 1.8167 | 0.003664
2019-03-21 15:56 | 1553198134.8644 | 344 | 0.20409 | 0.004459 | 0.000141 | 99.7175 | 98.2536 | 111.0596 | 21.2612 | 0.003319 | 0.003083 | 244.4308 | 0.27174 | 8.4022 | 0.003565 | 0.005498 | 0.005398 | 5.4e-05 | 0.001042 | 0.27184 | 1.3628 | 2.548 | 98.87 | 20.8 | 38.654 | 101.7077 | 107.6639 | 96.0002 | 0.009615 | 0.95637 | 0.022543 | 23.5717 | 0.13228 | 0.16166 | 0.027213 | 1.8284 | 1.8284 | 0.003634
2019-03-21 15:56 | 1553198135.8641 | 345 | 0.20405 | 0.004478 | 0.000143 | 99.7026 | 98.1473 | 110.9542 | 21.2612 | 0.003358 | 0.003035 | 244.3707 | 0.27172 | 8.4008 | 0.003556 | 0.004758 | 0.004714 | 5.4e-05 | 0.00104 | 0.25442 | 1.3502 | 2.5481 | 98.87 | 20.8 | 38.653 | 102.661 | 107.4529 | 94.8493 | 0.009615 | 0.95936 | 0.022568 | 23.5238 | 0.14063 | 0.16165 | 0.027424 | 1.8445 | 1.8445 | 0.003602
2019-03-21 15:56 | 1553198136.8641 | 346 | 0.20402 | 0.004497 | 0.000146 | 99.6776 | 98.4392 | 111.1455 | 21.2612 | 0.00343 | 0.003055 | 244.3036 | 0.27185 | 8.4019 | 0.003509 | 0.005003 | 0.004969 | 5.1e-05 | 0.001044 | 0.27959 | 1.3287 | 2.5479 | 98.87 | 20.8 | 38.653 | 104.3995 | 106.2995 | 95.3205 | 0.009615 | 0.95814 | 0.022573 | 23.5591 | 0.11182 | 0.16173 | 0.027558 | 1.8538 | 1.8538 | 0.003586
2019-03-21 15:56 | 1553198137.8632 | 347 | 0.20395 | 0.004519 | 0.000149 | 99.6487 | 98.6003 | 111.0208 | 21.2612 | 0.003358 | 0.003088 | 244.2382 | 0.27127 | 8.4045 | 0.003556 | 0.004978 | 0.004912 | 5.2e-05 | 0.001043 | 0.28031 | 1.3651 | 2.5481 | 98.86 | 20.8 | 38.653 | 102.6547 | 107.452 | 96.1194 | 0.009616 | 0.95606 | 0.022536 | 23.5714 | 0.094599 | 0.16138 | 0.027951 | 1.877 | 1.877 | 0.003534
2019-03-21 15:56 | 1553198138.8632 | 348 | 0.20378 | 0.004535 | 0.000151 | 99.5915 | 98.3731 | 111.1227 | 21.2612 | 0.003405 | 0.003129 | 244.1608 | 0.27111 | 8.4031 | 0.003547 | 0.005196 | 0.005074 | 5e-05 | 0.001041 | 0.23811 | 1.3095 | 2.5478 | 98.86 | 20.8 | 38.654 | 103.7935 | 107.226 | 97.1212 | 0.009616 | 0.95348 | 0.022516 | 23.6142 | 0.11009 | 0.16129 | 0.028985 | 1.9447 | 1.9447 | 0.003409
2019-03-21 15:56 | 1553198139.8636 | 349 | 0.20371 | 0.004563 | 0.000154 | 99.6531 | 98.1034 | 111.2103 | 21.2612 | 0.003371 | 0.003086 | 244.0674 | 0.2713 | 8.4013 | 0.003579 | 0.004682 | 0.00463 | 5e-05 | 0.001045 | 0.27223 | 1.305 | 2.548 | 98.86 | 20.8 | 38.656 | 102.9574 | 108.0026 | 96.0692 | 0.009616 | 0.95619 | 0.022557 | 23.5899 | 0.14019 | 0.16139 | 0.029331 | 1.9713 | 1.9713 | 0.003365
2019-03-21 15:56 | 1553198140.8647 | 350 | 0.20363 | 0.004588 | 0.000157 | 99.5926 | 97.8769 | 110.9774 | 21.2612 | 0.0034 | 0.003093 | 244.0032 | 0.27237 | 8.4044 | 0.003551 | 0.004817 | 0.004737 | 4.9e-05 | 0.001039 | 0.26969 | 1.3074 | 2.548 | 98.86 | 20.8 | 38.656 | 103.6559 | 107.3355 | 96.2329 | 0.009616 | 0.95577 | 0.022528 | 23.5704 | 0.15544 | 0.16203 | 0.029768 | 1.9981 | 1.9981 | 0.003333
2019-03-21 15:56 | 1553198141.8645 | 351 | 0.20365 | 0.004629 | 0.000159 | 99.5471 | 97.7257 | 111.2983 | 21.2612 | 0.00341 | 0.003069 | 243.9262 | 0.2719 | 8.4032 | 0.003468 | 0.00497 | 0.004894 | 5.1e-05 | 0.001036 | 0.26504 | 1.326 | 2.5479 | 98.86 | 20.8 | 38.658 | 103.9067 | 105.326 | 95.6666 | 0.009617 | 0.95724 | 0.022578 | 23.5864 | 0.16517 | 0.16175 | 0.029627 | 1.9928 | 1.9928 | 0.003336
2019-03-21 15:56 | 1553198142.8644 | 352 | 0.2036 | 0.004661 | 0.000161 | 99.5605 | 97.8537 | 111.2126 | 21.2612 | 0.003399 | 0.003153 | 243.8252 | 0.27143 | 8.401 | 0.003437 | 0.005284 | 0.005195 | 4.9e-05 | 0.001038 | 0.28584 | 1.3768 | 2.5479 | 98.86 | 20.8 | 38.66 | 103.6426 | 104.5513 | 97.6973 | 0.009617 | 0.952 | 0.022507 | 23.6421 | 0.15466 | 0.16147 | 0.029885 | 2.0038 | 2.0038 | 0.003312
2019-03-21 15:56 | 1553198143.8639 | 353 | 0.20348 | 0.004704 | 0.000163 | 99.5167 | 97.8333 | 111.272 | 21.2612 | 0.003358 | 0.003129 | 243.7197 | 0.27147 | 8.4008 | 0.00348 | 0.004684 | 0.004672 | 4.8e-05 | 0.001041 | 0.28224 | 1.3755 | 2.5481 | 98.86 | 20.79 | 38.662 | 102.6564 | 105.6015 | 97.115 | 0.009612 | 0.95349 | 0.022531 | 23.6298 | 0.1526 | 0.1615 | 0.030545 | 2.0502 | 2.0502 | 0.003238
2019-03-21 15:56 | 1553198144.864 | 354 | 0.2034 | 0.004738 | 0.000166 | 99.5161 | 97.8425 | 111.3534 | 21.2612 | 0.003445 | 0.003101 | 243.623 | 0.2723 | 8.4049 | 0.003554 | 0.004611 | 0.004528 | 5e-05 | 0.001039 | 0.27168 | 1.3451 | 2.548 | 98.86 | 20.79 | 38.666 | 104.7684 | 107.4119 | 96.4294 | 0.009613 | 0.95526 | 0.02256 | 23.6166 | 0.1517 | 0.16199 | 0.030974 | 2.0815 | 2.0815 | 0.003199
2019-03-21 15:56 | 1553198145.8637 | 355 | 0.20334 | 0.004781 | 0.000169 | 99.4696 | 97.7257 | 111.199 | 21.2612 | 0.003445 | 0.003185 | 243.5239 | 0.27172 | 8.4041 | 0.003479 | 0.004911 | 0.004822 | 4.8e-05 | 0.001039 | 0.25608 | 1.3207 | 2.5479 | 98.86 | 20.79 | 38.668 | 104.7456 | 105.5809 | 98.4717 | 0.009614 | 0.95001 | 0.022482 | 23.6653 | 0.15821 | 0.16165 | 0.031257 | 2.0931 | 2.0931 | 0.003174
2019-03-21 15:56 | 1553198146.863 | 356 | 0.20331 | 0.004821 | 0.000172 | 99.3281 | 97.5612 | 110.8786 | 21.2612 | 0.003376 | 0.003145 | 243.4214 | 0.27232 | 8.4008 | 0.003426 | 0.00437 | 0.004319 | 4.7e-05 | 0.001042 | 0.2896 | 1.3059 | 2.5481 | 98.86 | 20.79 | 38.671 | 103.0814 | 104.2965 | 97.4944 | 0.009614 | 0.95252 | 0.02248 | 23.6001 | 0.16054 | 0.162 | 0.031427 | 2.1041 | 2.1041 | 0.003164
2019-03-21 15:56 | 1553198147.8631 | 357 | 0.20305 | 0.004864 | 0.000175 | 99.475 | 97.5005 | 110.7136 | 21.2612 | 0.003406 | 0.003181 | 243.3066 | 0.27184 | 8.4022 | 0.003484 | 0.005122 | 0.005042 | 4.9e-05 | 0.001041 | 0.27079 | 1.3213 | 2.548 | 98.86 | 20.79 | 38.672 | 103.8156 | 105.7077 | 98.3769 | 0.009615 | 0.95025 | 0.022436 | 23.6106 | 0.17932 | 0.16172 | 0.032907 | 2.1989 | 2.1989 | 0.003023
2019-03-21 15:56 | 1553198148.8638 | 358 | 0.20299 | 0.004914 | 0.000178 | 99.4326 | 97.4699 | 110.4054 | 21.2612 | 0.003513 | 0.00311 | 243.2001 | 0.27149 | 8.4009 | 0.00353 | 0.005076 | 0.005066 | 4.8e-05 | 0.00104 | 0.27383 | 1.2931 | 2.5481 | 98.86 | 20.79 | 38.675 | 106.4058 | 106.8071 | 96.6636 | 0.009615 | 0.95466 | 0.022457 | 23.5233 | 0.17833 | 0.16151 | 0.033181 | 2.2191 | 2.2191 | 0.002991
2019-03-21 15:56 | 1553198149.8632 | 359 | 0.20289 | 0.004968 | 0.000181 | 99.3564 | 97.4134 | 109.9529 | 21.2612 | 0.003459 | 0.003147 | 243.1062 | 0.27171 | 8.4022 | 0.003502 | 0.004822 | 0.004701 | 4.8e-05 | 0.001048 | 0.29518 | 1.2991 | 2.5479 | 98.86 | 20.79 | 38.674 | 105.0933 | 106.1404 | 97.5367 | 0.009615 | 0.95241 | 0.022384 | 23.5027 | 0.17665 | 0.16164 | 0.033706 | 2.2468 | 2.2468 | 0.002957
2019-03-21 15:56 | 1553198150.8629 | 360 | 0.20269 | 0.005026 | 0.000183 | 99.3504 | 97.3793 | 109.5259 | 21.2612 | 0.003441 | 0.003163 | 243.0106 | 0.27182 | 8.4017 | 0.003472 | 0.005264 | 0.005146 | 4.9e-05 | 0.001045 | 0.26106 | 1.2951 | 2.5479 | 98.86 | 20.79 | 38.676 | 104.6661 | 105.4094 | 97.9229 | 0.009616 | 0.95142 | 0.022329 | 23.4693 | 0.17924 | 0.16171 | 0.034872 | 2.3188 | 2.3188 | 0.002866
2019-03-21 15:56 | 1553198151.8639 | 361 | 0.20255 | 0.005093 | 0.000184 | 99.3248 | 97.4253 | 109.1273 | 21.2612 | 0.003429 | 0.003125 | 242.8955 | 0.27236 | 8.4025 | 0.003474 | 0.004795 | 0.004693 | 5.1e-05 | 0.001048 | 0.286 | 1.2864 | 2.5479 | 98.86 | 20.79 | 38.677 | 104.3768 | 105.4589 | 97.0116 | 0.009616 | 0.95376 | 0.022316 | 23.3977 | 0.17272 | 0.16202 | 0.035631 | 2.3677 | 2.3677 | 0.002812
2019-03-21 15:56 | 1553198152.863 | 362 | 0.20246 | 0.005168 | 0.000186 | 99.3567 | 97.5976 | 108.7907 | 21.2612 | 0.003482 | 0.003162 | 242.815 | 0.27208 | 8.4041 | 0.003394 | 0.00463 | 0.004599 | 4.9e-05 | 0.001052 | 0.26089 | 1.2765 | 2.5479 | 98.86 | 20.79 | 38.678 | 105.6495 | 103.5302 | 97.8978 | 0.009616 | 0.95148 | 0.022255 | 23.3896 | 0.15978 | 0.16186 | 0.036032 | 2.3877 | 2.3877 | 0.002786
2019-03-21 15:56 | 1553198153.8635 | 363 | 0.20235 | 0.005235 | 0.00019 | 99.261 | 97.2311 | 108.4882 | 21.2612 | 0.003426 | 0.003216 | 242.723 | 0.27251 | 8.4043 | 0.003447 | 0.004228 | 0.004164 | 4.7e-05 | 0.001049 | 0.27096 | 1.2618 | 2.548 | 98.86 | 20.79 | 38.679 | 104.2878 | 104.7999 | 99.2238 | 0.009616 | 0.94809 | 0.022184 | 23.3987 | 0.18481 | 0.16212 | 0.036623 | 2.419 | 2.419 | 0.002754
2019-03-21 15:56 | 1553198154.8645 | 364 | 0.20225 | 0.005306 | 0.000193 | 99.3376 | 97.2376 | 108.306 | 21.2612 | 0.003471 | 0.003208 | 242.6124 | 0.27193 | 8.4014 | 0.003422 | 0.004786 | 0.004677 | 4.5e-05 | 0.001051 | 0.2975 | 1.2581 | 2.5479 | 98.86 | 20.79 | 38.679 | 105.3779 | 104.2016 | 99.0167 | 0.009616 | 0.94862 | 0.022172 | 23.3726 | 0.19111 | 0.16177 | 0.037112 | 2.4498 | 2.4498 | 0.002714
2019-03-21 15:56 | 1553198155.8641 | 365 | 0.20218 | 0.005372 | 0.000195 | 99.334 | 97.2294 | 108.0931 | 21.2612 | 0.003499 | 0.003241 | 242.491 | 0.27213 | 8.4033 | 0.003498 | 0.005073 | 0.005038 | 4.9e-05 | 0.001054 | 0.2625 | 1.2892 | 2.5478 | 98.86 | 20.79 | 38.679 | 106.0555 | 106.0485 | 99.8137 | 0.009616 | 0.94659 | 0.022126 | 23.3746 | 0.19154 | 0.16189 | 0.037459 | 2.4676 | 2.4676 | 0.002696
2019-03-21 15:56 | 1553198156.8632 | 366 | 0.20211 | 0.005447 | 0.000197 | 99.295 | 97.1072 | 108.2377 | 21.2612 | 0.003498 | 0.003244 | 242.3812 | 0.27205 | 8.4023 | 0.0034 | 0.005536 | 0.005409 | 4.5e-05 | 0.001054 | 0.26305 | 1.2881 | 2.548 | 98.86 | 20.79 | 38.68 | 106.0397 | 103.67 | 99.8826 | 0.009616 | 0.94642 | 0.022139 | 23.3924 | 0.19928 | 0.16184 | 0.037766 | 2.4891 | 2.4891 | 0.002672
2019-03-21 15:56 | 1553198157.8634 | 367 | 0.20199 | 0.005518 | 0.000198 | 99.2881 | 96.9168 | 108.4361 | 21.2612 | 0.003567 | 0.00324 | 242.2755 | 0.27202 | 8.4018 | 0.003391 | 0.005398 | 0.005444 | 4.7e-05 | 0.001054 | 0.24635 | 1.284 | 2.548 | 98.86 | 20.79 | 38.68 | 107.712 | 103.4539 | 99.7973 | 0.009616 | 0.94664 | 0.022162 | 23.4111 | 0.21622 | 0.16182 | 0.038416 | 2.5346 | 2.5346 | 0.002624
2019-03-21 15:56 | 1553198158.8637 | 368 | 0.20189 | 0.005593 | 0.000199 | 99.2998 | 96.6951 | 108.4023 | 21.2612 | 0.003565 | 0.003235 | 242.1668 | 0.27123 | 8.4027 | 0.003415 | 0.00482 | 0.004669 | 4.6e-05 | 0.001056 | 0.25177 | 1.2979 | 2.5479 | 98.86 | 20.79 | 38.679 | 107.6765 | 104.0419 | 99.6747 | 0.009616 | 0.94695 | 0.022162 | 23.4036 | 0.23775 | 0.16135 | 0.038926 | 2.5682 | 2.5682 | 0.002582
2019-03-21 15:56 | 1553198159.8634 | 369 | 0.20176 | 0.005667 | 0.0002 | 99.2825 | 96.9846 | 108.2591 | 21.2612 | 0.003545 | 0.003242 | 242.0518 | 0.27206 | 8.4021 | 0.003413 | 0.005251 | 0.005174 | 4.5e-05 | 0.001056 | 0.27168 | 1.3093 | 2.5479 | 98.86 | 20.79 | 38.679 | 107.1898 | 103.9743 | 99.8396 | 0.009616 | 0.94653 | 0.022142 | 23.3933 | 0.20946 | 0.16185 | 0.039601 | 2.6103 | 2.6103 | 0.002548
2019-03-21 15:56 | 1553198160.8634 | 370 | 0.20173 | 0.005743 | 0.0002 | 99.2604 | 97.2073 | 108.2321 | 21.2612 | 0.003526 | 0.003253 | 241.9197 | 0.27189 | 8.404 | 0.003351 | 0.005113 | 0.005109 | 4.6e-05 | 0.001054 | 0.26001 | 1.3064 | 2.548 | 98.86 | 20.79 | 38.681 | 106.7198 | 102.4779 | 100.1015 | 0.009617 | 0.94586 | 0.022132 | 23.3986 | 0.18695 | 0.16174 | 0.039695 | 2.6153 | 2.6153 | 0.002542
2019-03-21 15:56 | 1553198161.8635 | 371 | 0.20166 | 0.005815 | 0.0002 | 99.2339 | 97.1797 | 107.8901 | 21.2612 | 0.003579 | 0.003286 | 241.8082 | 0.27263 | 8.4013 | 0.003396 | 0.004909 | 0.004922 | 4.6e-05 | 0.001055 | 0.2625 | 1.3748 | 2.5482 | 98.86 | 20.79 | 38.68 | 108.0145 | 103.5702 | 100.9186 | 0.009616 | 0.9438 | 0.022073 | 23.3872 | 0.1871 | 0.16219 | 0.04003 | 2.6303 | 2.6303 | 0.002534
2019-03-21 15:56 | 1553198162.8636 | 372 | 0.20159 | 0.005889 | 0.000201 | 99.2533 | 97.224 | 107.7359 | 21.2612 | 0.003669 | 0.003326 | 241.6857 | 0.27166 | 8.4026 | 0.0034 | 0.00446 | 0.004334 | 4.7e-05 | 0.001058 | 0.25586 | 1.3613 | 2.5481 | 98.86 | 20.79 | 38.68 | 110.2011 | 103.6583 | 101.878 | 0.009616 | 0.94138 | 0.022029 | 23.4004 | 0.18477 | 0.16161 | 0.040377 | 2.6478 | 2.6478 | 0.002509
2019-03-21 15:56 | 1553198163.8626 | 373 | 0.20152 | 0.005954 | 0.000202 | 99.303 | 97.066 | 107.5094 | 21.2612 | 0.003615 | 0.003262 | 241.5732 | 0.27194 | 8.4018 | 0.00343 | 0.004942 | 0.00491 | 4.5e-05 | 0.001056 | 0.27527 | 1.3479 | 2.5481 | 98.86 | 20.79 | 38.681 | 108.8813 | 104.3908 | 100.3206 | 0.009617 | 0.94531 | 0.022051 | 23.3272 | 0.20381 | 0.16178 | 0.040706 | 2.672 | 2.672 | 0.002488
2019-03-21 15:56 | 1553198164.8626 | 374 | 0.20136 | 0.006021 | 0.000203 | 99.2592 | 97.1218 | 107.1186 | 21.2612 | 0.003629 | 0.003321 | 241.4847 | 0.27254 | 8.4035 | 0.003419 | 0.004813 | 0.004734 | 5.1e-05 | 0.001059 | 0.27151 | 1.3148 | 2.5479 | 98.86 | 20.79 | 38.681 | 109.2321 | 104.115 | 101.7616 | 0.009617 | 0.94168 | 0.021969 | 23.3296 | 0.19471 | 0.16214 | 0.041545 | 2.7168 | 2.7168 | 0.002453
2019-03-21 15:56 | 1553198165.8629 | 375 | 0.20127 | 0.006088 | 0.000206 | 99.2735 | 97.036 | 107.1214 | 21.2612 | 0.003569 | 0.003332 | 241.3815 | 0.2714 | 8.4025 | 0.003468 | 0.005021 | 0.005064 | 4.8e-05 | 0.001058 | 0.2677 | 1.2861 | 2.5479 | 98.86 | 20.79 | 38.681 | 107.7734 | 105.3112 | 102.009 | 0.009617 | 0.94105 | 0.021962 | 23.3376 | 0.20391 | 0.16145 | 0.042007 | 2.746 | 2.746 | 0.002416
2019-03-21 15:56 | 1553198166.8627 | 376 | 0.2011 | 0.006153 | 0.000208 | 99.2265 | 97.0336 | 107.0003 | 21.2612 | 0.003628 | 0.003386 | 241.2689 | 0.27142 | 8.4034 | 0.0034 | 0.004935 | 0.004778 | 4.5e-05 | 0.001054 | 0.23579 | 1.2801 | 2.548 | 98.86 | 20.79 | 38.681 | 109.1886 | 103.6787 | 103.3376 | 0.009617 | 0.93773 | 0.021911 | 23.3657 | 0.19989 | 0.16147 | 0.042956 | 2.8014 | 2.8014 | 0.002369
2019-03-21 15:56 | 1553198167.8638 | 377 | 0.20099 | 0.00624 | 0.000208 | 99.2828 | 97.2815 | 106.8242 | 21.2612 | 0.003673 | 0.003389 | 241.1589 | 0.27102 | 8.401 | 0.003414 | 0.00446 | 0.00445 | 4.5e-05 | 0.001056 | 0.26089 | 1.3135 | 2.548 | 98.86 | 20.79 | 38.682 | 110.2802 | 104.0058 | 103.4098 | 0.009617 | 0.93756 | 0.021891 | 23.3487 | 0.18214 | 0.16123 | 0.043489 | 2.8335 | 2.8335 | 0.002339
2019-03-21 15:56 | 1553198168.8633 | 378 | 0.2009 | 0.006326 | 0.000208 | 99.2604 | 96.895 | 106.8 | 21.2612 | 0.003741 | 0.003355 | 241.0453 | 0.27152 | 8.4033 | 0.003481 | 0.004426 | 0.00432 | 4.7e-05 | 0.001053 | 0.26233 | 1.3112 | 2.548 | 98.86 | 20.79 | 38.683 | 111.9514 | 105.6408 | 102.5794 | 0.009617 | 0.93963 | 0.021912 | 23.3203 | 0.21574 | 0.16153 | 0.043909 | 2.8637 | 2.8637 | 0.002318
2019-03-21 15:56 | 1553198169.864 | 379 | 0.20088 | 0.006407 | 0.000208 | 99.3051 | 96.8466 | 106.8423 | 21.2612 | 0.003695 | 0.003379 | 240.9428 | 0.27214 | 8.4018 | 0.003437 | 0.005469 | 0.005371 | 4.8e-05 | 0.00105 | 0.2521 | 1.2989 | 2.5478 | 98.86 | 20.78 | 38.684 | 110.8387 | 104.5687 | 103.1695 | 0.009612 | 0.93815 | 0.0219 | 23.3432 | 0.22423 | 0.1619 | 0.043981 | 2.8667 | 2.8667 | 0.002321
2019-03-21 15:56 | 1553198170.863 | 380 | 0.2009 | 0.006485 | 0.000206 | 99.2801 | 96.882 | 106.8418 | 21.2612 | 0.003685 | 0.003351 | 240.8371 | 0.27191 | 8.4006 | 0.003412 | 0.005259 | 0.005191 | 4.6e-05 | 0.00105 | 0.26681 | 1.2815 | 2.5479 | 98.86 | 20.78 | 38.684 | 110.5807 | 103.9624 | 102.4796 | 0.009612 | 0.93988 | 0.02192 | 23.3218 | 0.2187 | 0.16176 | 0.043729 | 2.853 | 2.853 | 0.00233
2019-03-21 15:56 | 1553198171.8634 | 381 | 0.2009 | 0.00656 | 0.000204 | 99.2095 | 96.8585 | 107.0862 | 21.2612 | 0.003712 | 0.003397 | 240.7152 | 0.27221 | 8.4037 | 0.003404 | 0.005231 | 0.005179 | 4.5e-05 | 0.001049 | 0.27632 | 1.3007 | 2.5479 | 98.86 | 20.78 | 38.685 | 111.2434 | 103.7565 | 103.5953 | 0.009612 | 0.93709 | 0.021912 | 23.3831 | 0.21451 | 0.16194 | 0.043631 | 2.8458 | 2.8458 | 0.002339
2019-03-21 15:56 | 1553198172.8635 | 382 | 0.2009 | 0.006626 | 0.000204 | 99.2461 | 97.0513 | 107.2322 | 21.2612 | 0.00368 | 0.003444 | 240.5905 | 0.27191 | 8.4019 | 0.003453 | 0.005248 | 0.005153 | 4.5e-05 | 0.001046 | 0.27831 | 1.2976 | 2.548 | 98.86 | 20.78 | 38.685 | 110.4598 | 104.943 | 104.7369 | 0.009612 | 0.93426 | 0.021894 | 23.4345 | 0.20003 | 0.16176 | 0.043556 | 2.8385 | 2.8385 | 0.002342
2019-03-21 15:56 | 1553198173.8636 | 383 | 0.2008 | 0.006684 | 0.000204 | 99.2399 | 97.1525 | 107.2944 | 21.2612 | 0.003747 | 0.003403 | 240.471 | 0.27194 | 8.403 | 0.003483 | 0.004874 | 0.004918 | 4.3e-05 | 0.001044 | 0.26449 | 1.3075 | 2.5479 | 98.86 | 20.78 | 38.687 | 112.0833 | 105.6706 | 103.7509 | 0.009612 | 0.93671 | 0.021929 | 23.4106 | 0.19015 | 0.16178 | 0.044078 | 2.8771 | 2.8771 | 0.002311
2019-03-21 15:56 | 1553198174.8628 | 384 | 0.20073 | 0.006738 | 0.000203 | 99.3245 | 97.4576 | 107.366 | 21.2612 | 0.003822 | 0.003423 | 240.3551 | 0.27159 | 8.4015 | 0.003494 | 0.004895 | 0.004801 | 4.4e-05 | 0.001045 | 0.262 | 1.3459 | 2.548 | 98.86 | 20.78 | 38.689 | 113.9193 | 105.9342 | 104.2251 | 0.009613 | 0.93553 | 0.021922 | 23.4332 | 0.16972 | 0.16157 | 0.044478 | 2.9024 | 2.9024 | 0.002288
2019-03-21 15:56 | 1553198175.8632 | 385 | 0.20073 | 0.006778 | 0.000202 | 99.2807 | 97.3326 | 107.1343 | 21.2612 | 0.003685 | 0.003423 | 240.2265 | 0.27183 | 8.4033 | 0.003519 | 0.005251 | 0.005201 | 4.5e-05 | 0.001039 | 0.24082 | 1.3281 | 2.5481 | 98.86 | 20.78 | 38.69 | 110.5831 | 106.5506 | 104.2219 | 0.009613 | 0.93554 | 0.021899 | 23.4078 | 0.17725 | 0.16171 | 0.044422 | 2.8957 | 2.8957 | 0.002295
2019-03-21 15:56 | 1553198176.8626 | 386 | 0.20073 | 0.006826 | 0.000202 | 99.3737 | 97.179 | 106.558 | 21.2612 | 0.003758 | 0.003416 | 240.1045 | 0.27129 | 8.4015 | 0.003518 | 0.00511 | 0.004977 | 4.9e-05 | 0.001041 | 0.25946 | 1.3033 | 2.5479 | 98.86 | 20.78 | 38.692 | 112.3496 | 106.5293 | 104.0537 | 0.009614 | 0.93595 | 0.021845 | 23.3395 | 0.19975 | 0.16139 | 0.044329 | 2.8824 | 2.8824 | 0.002301
2019-03-21 15:56 | 1553198177.8633 | 387 | 0.20072 | 0.006866 | 0.000203 | 99.3349 | 97.2972 | 106.2649 | 21.2612 | 0.003796 | 0.003414 | 239.9742 | 0.27183 | 8.4016 | 0.003548 | 0.005027 | 0.004958 | 4.6e-05 | 0.001038 | 0.24347 | 1.3366 | 2.5479 | 98.86 | 20.78 | 38.692 | 113.2872 | 107.2435 | 104.0079 | 0.009614 | 0.93607 | 0.021816 | 23.306 | 0.18539 | 0.16171 | 0.044396 | 2.883 | 2.883 | 0.002305
2019-03-21 15:56 | 1553198178.8634 | 388 | 0.20074 | 0.00691 | 0.000202 | 99.2342 | 97.5142 | 105.9419 | 21.2612 | 0.003769 | 0.003502 | 239.871 | 0.27194 | 8.4049 | 0.003558 | 0.005147 | 0.005087 | 5.2e-05 | 0.001035 | 0.26416 | 1.3099 | 2.548 | 98.86 | 20.78 | 38.694 | 112.6195 | 107.5025 | 106.1506 | 0.009614 | 0.93078 | 0.021721 | 23.3366 | 0.1564 | 0.16178 | 0.044194 | 2.8574 | 2.8574 | 0.002327
2019-03-21 15:56 | 1553198179.8639 | 389 | 0.2007 | 0.006944 | 0.000204 | 99.3325 | 97.4481 | 105.9954 | 21.2612 | 0.003724 | 0.003486 | 239.7496 | 0.27094 | 8.3999 | 0.003556 | 0.004639 | 0.004597 | 4.9e-05 | 0.001033 | 0.2656 | 1.3115 | 2.548 | 98.86 | 20.77 | 38.695 | 111.5324 | 107.4498 | 105.7477 | 0.009608 | 0.93177 | 0.021738 | 23.3301 | 0.17132 | 0.16118 | 0.044404 | 2.8732 | 2.8732 | 0.002306
2019-03-21 15:56 | 1553198180.8638 | 390 | 0.20067 | 0.006976 | 0.000205 | 99.3376 | 97.0721 | 105.8593 | 21.2612 | 0.00381 | 0.003475 | 239.6305 | 0.27236 | 8.4019 | 0.003572 | 0.004703 | 0.00465 | 4.9e-05 | 0.001032 | 0.26034 | 1.3304 | 2.548 | 98.86 | 20.77 | 38.696 | 113.6347 | 107.838 | 105.4778 | 0.009609 | 0.93243 | 0.021732 | 23.3068 | 0.20635 | 0.16203 | 0.04454 | 2.8811 | 2.8811 | 0.002311
2019-03-21 15:56 | 1553198181.863 | 391 | 0.20059 | 0.007008 | 0.000205 | 99.3725 | 96.9917 | 105.6213 | 21.2612 | 0.003798 | 0.003453 | 239.5241 | 0.27269 | 8.401 | 0.003523 | 0.005446 | 0.005393 | 5e-05 | 0.001034 | 0.27704 | 1.2916 | 2.548 | 98.86 | 20.77 | 38.696 | 113.3315 | 106.6489 | 104.9422 | 0.009609 | 0.93376 | 0.021723 | 23.2641 | 0.2169 | 0.16223 | 0.044951 | 2.9064 | 2.9064 | 0.002294
2019-03-21 15:56 | 1553198182.8636 | 392 | 0.20057 | 0.007036 | 0.000206 | 99.3844 | 97.2798 | 105.7213 | 21.2612 | 0.003773 | 0.003527 | 239.4189 | 0.27099 | 8.4012 | 0.003563 | 0.004349 | 0.004266 | 4.8e-05 | 0.001033 | 0.30502 | 1.3311 | 2.5479 | 98.86 | 20.77 | 38.697 | 112.7159 | 107.6221 | 106.7348 | 0.009609 | 0.92935 | 0.021682 | 23.3302 | 0.19144 | 0.16121 | 0.045065 | 2.9082 | 2.9082 | 0.002278
2019-03-21 15:56 | 1553198183.8636 | 393 | 0.20054 | 0.007068 | 0.000206 | 99.3793 | 97.3486 | 105.618 | 21.2612 | 0.003826 | 0.003515 | 239.306 | 0.27138 | 8.4005 | 0.003646 | 0.005499 | 0.005485 | 4.7e-05 | 0.001033 | 0.27129 | 1.2986 | 2.548 | 98.86 | 20.77 | 38.699 | 114.0157 | 109.6277 | 106.4546 | 0.009609 | 0.93003 | 0.021679 | 23.3102 | 0.18466 | 0.16144 | 0.0452 | 2.9166 | 2.9166 | 0.002275
2019-03-21 15:56 | 1553198184.9291 | 394 | 0.2005 | 0.007096 | 0.000205 | 99.3906 | 97.3592 | 105.4548 | 21.2612 | 0.003804 | 0.003499 | 239.1754 | 0.27147 | 8.4079 | 0.003694 | 0.004742 | 0.004703 | 5.1e-05 | 0.001036 | 0.2218 | 1.3337 | 2.548 | 98.86 | 20.77 | 38.703 | 113.484 | 110.8102 | 106.0625 | 0.00961 | 0.931 | 0.021674 | 23.2802 | 0.18471 | 0.1615 | 0.045373 | 2.927 | 2.927 | 0.002268
2019-03-21 15:56 | 1553198185.8728 | 395 | 0.20045 | 0.007122 | 0.000203 | 99.3361 | 97.4328 | 105.6419 | 21.2612 | 0.003872 | 0.003559 | 239.0668 | 0.27183 | 8.4001 | 0.003646 | 0.005249 | 0.005169 | 4.9e-05 | 0.001034 | 0.27041 | 1.3198 | 2.5479 | 98.86 | 20.77 | 38.704 | 115.1441 | 109.6404 | 107.5144 | 0.009611 | 0.92744 | 0.021652 | 23.3454 | 0.17305 | 0.16171 | 0.045656 | 2.9424 | 2.9424 | 0.002259
2019-03-21 15:56 | 1553198186.8725 | 396 | 0.20034 | 0.007156 | 0.0002 | 99.4365 | 97.3711 | 105.7891 | 21.2612 | 0.003805 | 0.003552 | 238.9502 | 0.27146 | 8.4011 | 0.003573 | 0.005142 | 0.005097 | 5.1e-05 | 0.001033 | 0.2453 | 1.289 | 2.5478 | 98.86 | 20.77 | 38.706 | 113.4942 | 107.872 | 107.3442 | 0.009611 | 0.92786 | 0.021671 | 23.3564 | 0.18774 | 0.16149 | 0.046275 | 2.9851 | 2.9851 | 0.002223
2019-03-21 15:56 | 1553198187.8729 | 397 | 0.20019 | 0.007192 | 0.000196 | 99.337 | 97.3074 | 106.0337 | 21.2612 | 0.003894 | 0.003543 | 238.8296 | 0.27213 | 8.4053 | 0.003739 | 0.005321 | 0.005277 | 5.3e-05 | 0.00103 | 0.30021 | 1.3043 | 2.5479 | 98.86 | 20.77 | 38.707 | 115.6757 | 111.8867 | 107.1418 | 0.009611 | 0.92835 | 0.021702 | 23.3772 | 0.18464 | 0.16189 | 0.047156 | 3.0464 | 3.0464 | 0.002184
2019-03-21 15:56 | 1553198188.8726 | 398 | 0.20008 | 0.007221 | 0.000193 | 99.4502 | 97.4437 | 106.1949 | 21.2612 | 0.003878 | 0.003525 | 238.7091 | 0.27208 | 8.402 | 0.003813 | 0.005194 | 0.005138 | 5e-05 | 0.00103 | 0.25514 | 1.3097 | 2.5479 | 98.86 | 20.77 | 38.709 | 115.2746 | 113.6949 | 106.6938 | 0.009612 | 0.92945 | 0.021732 | 23.3811 | 0.18232 | 0.16186 | 0.047778 | 3.0909 | 3.0909 | 0.002152
2019-03-21 15:56 | 1553198189.8731 | 399 | 0.19997 | 0.007269 | 0.00019 | 99.4815 | 97.8221 | 106.0346 | 21.2612 | 0.003928 | 0.003535 | 238.589 | 0.27159 | 8.4023 | 0.003702 | 0.005277 | 0.005171 | 4.7e-05 | 0.00103 | 0.26416 | 1.3424 | 2.548 | 98.86 | 20.77 | 38.71 | 116.4933 | 111.0075 | 106.9505 | 0.009612 | 0.92882 | 0.021708 | 23.3714 | 0.15046 | 0.16157 | 0.048344 | 3.1242 | 3.1242 | 0.002125
2019-03-21 15:56 | 1553198190.8716 | 400 | 0.19992 | 0.007326 | 0.000189 | 99.4836 | 97.7744 | 106.0233 | 21.2612 | 0.003932 | 0.003606 | 238.4445 | 0.27233 | 8.4026 | 0.003845 | 0.005342 | 0.00524 | 4.6e-05 | 0.001028 | 0.26128 | 1.3547 | 2.5482 | 98.86 | 20.77 | 38.712 | 116.608 | 114.4807 | 108.6527 | 0.009613 | 0.92468 | 0.021658 | 23.4224 | 0.15501 | 0.16201 | 0.048628 | 3.1354 | 3.1354 | 0.002124
2019-03-21 15:57 | 1553198191.8717 | 401 | 0.1999 | 0.00738 | 0.000189 | 99.4687 | 97.805 | 105.8597 | 21.2612 | 0.003955 | 0.003643 | 238.3199 | 0.27179 | 8.4016 | 0.003736 | 0.005093 | 0.005054 | 4.9e-05 | 0.001027 | 0.29662 | 1.3095 | 2.548 | 98.86 | 20.77 | 38.714 | 117.1695 | 111.8251 | 109.5577 | 0.009613 | 0.92249 | 0.021616 | 23.4321 | 0.15087 | 0.16169 | 0.048644 | 3.1303 | 3.1303 | 0.002123
2019-03-21 15:57 | 1553198192.8719 | 402 | 0.1998 | 0.007444 | 0.00019 | 99.5042 | 97.7744 | 105.6398 | 21.2612 | 0.004003 | 0.003652 | 238.1934 | 0.27072 | 8.4026 | 0.003797 | 0.005491 | 0.005391 | 4.5e-05 | 0.001029 | 0.26089 | 1.296 | 2.5479 | 98.85 | 20.77 | 38.72 | 118.3269 | 113.3196 | 109.7847 | 0.009616 | 0.92195 | 0.021587 | 23.4146 | 0.15686 | 0.16105 | 0.049201 | 3.1617 | 3.1617 | 0.002094
2019-03-21 15:57 | 1553198193.8728 | 403 | 0.19979 | 0.007509 | 0.000191 | 99.5444 | 97.3013 | 105.8459 | 21.2612 | 0.003956 | 0.003665 | 238.0681 | 0.2716 | 8.4035 | 0.003739 | 0.004496 | 0.004466 | 4.9e-05 | 0.001029 | 0.23667 | 1.2908 | 2.5479 | 98.85 | 20.77 | 38.723 | 117.1947 | 111.886 | 110.0859 | 0.009616 | 0.92122 | 0.0216 | 23.4467 | 0.20386 | 0.16157 | 0.049146 | 3.16 | 3.16 | 0.002101
2019-03-21 15:57 | 1553198194.8719 | 404 | 0.19977 | 0.00757 | 0.000193 | 99.4544 | 97.2914 | 106.1297 | 21.2612 | 0.003984 | 0.003572 | 237.9311 | 0.2726 | 8.4016 | 0.003655 | 0.005215 | 0.005088 | 4.9e-05 | 0.001033 | 0.29048 | 1.2883 | 2.5479 | 98.85 | 20.77 | 38.726 | 117.8729 | 109.8557 | 107.8326 | 0.009617 | 0.92667 | 0.021692 | 23.409 | 0.19668 | 0.16217 | 0.049193 | 3.1765 | 3.1765 | 0.002098
2019-03-21 15:57 | 1553198195.8724 | 405 | 0.19975 | 0.007632 | 0.000194 | 99.5015 | 97.3398 | 105.9177 | 21.2612 | 0.004056 | 0.003642 | 237.7972 | 0.27239 | 8.4034 | 0.00369 | 0.005121 | 0.005024 | 4.8e-05 | 0.001033 | 0.28407 | 1.3188 | 2.548 | 98.85 | 20.77 | 38.728 | 119.6316 | 110.6974 | 109.5434 | 0.009618 | 0.92253 | 0.021622 | 23.438 | 0.19646 | 0.16204 | 0.049244 | 3.1694 | 3.1694 | 0.002101
2019-03-21 15:57 | 1553198196.873 | 406 | 0.1998 | 0.007685 | 0.000196 | 99.4845 | 97.4294 | 106.1806 | 21.2612 | 0.003982 | 0.00366 | 237.6782 | 0.272 | 8.405 | 0.003596 | 0.005292 | 0.005258 | 5.1e-05 | 0.001035 | 0.29767 | 1.2899 | 2.5481 | 98.85 | 20.77 | 38.729 | 117.8111 | 108.4329 | 109.9794 | 0.009618 | 0.92148 | 0.021637 | 23.4805 | 0.18671 | 0.16181 | 0.048896 | 3.1491 | 3.1491 | 0.002112
2019-03-21 15:57 | 1553198197.8731 | 407 | 0.19978 | 0.007723 | 0.000196 | 99.4812 | 97.3442 | 105.7553 | 21.2612 | 0.004054 | 0.003628 | 237.5739 | 0.27148 | 8.4034 | 0.003636 | 0.005426 | 0.005313 | 4.9e-05 | 0.001032 | 0.29015 | 1.2942 | 2.5479 | 98.85 | 20.77 | 38.729 | 119.5874 | 109.3842 | 109.1988 | 0.009618 | 0.92336 | 0.021615 | 23.4095 | 0.19423 | 0.1615 | 0.04895 | 3.1495 | 3.1495 | 0.002108
2019-03-21 15:57 | 1553198198.8721 | 408 | 0.19978 | 0.007763 | 0.000196 | 99.5214 | 97.5166 | 105.7041 | 21.2612 | 0.004012 | 0.00364 | 237.445 | 0.27184 | 8.4008 | 0.003589 | 0.004625 | 0.004573 | 4.9e-05 | 0.001033 | 0.29286 | 1.323 | 2.548 | 98.85 | 20.77 | 38.73 | 118.5558 | 108.2469 | 109.4868 | 0.009618 | 0.92266 | 0.021602 | 23.4127 | 0.18203 | 0.16172 | 0.048866 | 3.1421 | 3.1421 | 0.002115
2019-03-21 15:57 | 1553198199.8732 | 409 | 0.19981 | 0.007784 | 0.000196 | 99.5447 | 97.5288 | 105.6346 | 21.2612 | 0.004046 | 0.003696 | 237.3164 | 0.27109 | 8.402 | 0.003642 | 0.005267 | 0.005318 | 4.8e-05 | 0.001036 | 0.29247 | 1.321 | 2.5479 | 98.85 | 20.77 | 38.733 | 119.3819 | 109.5286 | 110.8426 | 0.009619 | 0.91941 | 0.021557 | 23.4464 | 0.18299 | 0.16127 | 0.048702 | 3.125 | 3.125 | 0.002121
2019-03-21 15:57 | 1553198200.8732 | 410 | 0.1998 | 0.007811 | 0.000197 | 99.5751 | 97.7124 | 105.5117 | 21.2612 | 0.003988 | 0.003706 | 237.2064 | 0.27133 | 8.3999 | 0.003612 | 0.005365 | 0.005266 | 4.8e-05 | 0.001035 | 0.25929 | 1.3416 | 2.5479 | 98.85 | 20.77 | 38.735 | 117.9584 | 108.8213 | 111.1006 | 0.009619 | 0.91879 | 0.021537 | 23.4406 | 0.1689 | 0.16141 | 0.048705 | 3.1223 | 3.1223 | 0.002125
2019-03-21 15:57 | 1553198201.8733 | 411 | 0.19981 | 0.007829 | 0.000197 | 99.5235 | 97.6177 | 105.4567 | 21.2612 | 0.00399 | 0.003695 | 237.0782 | 0.27161 | 8.4025 | 0.003576 | 0.004579 | 0.004414 | 4.8e-05 | 0.001035 | 0.29087 | 1.3391 | 2.5479 | 98.85 | 20.77 | 38.735 | 118.0218 | 107.9404 | 110.8338 | 0.009619 | 0.91943 | 0.021539 | 23.4264 | 0.17294 | 0.16158 | 0.048643 | 3.1186 | 3.1186 | 0.002129
2019-03-21 15:57 | 1553198202.8734 | 412 | 0.19969 | 0.007846 | 0.000198 | 99.5596 | 97.7764 | 105.2246 | 21.2612 | 0.00407 | 0.003734 | 236.9523 | 0.27231 | 8.4056 | 0.003548 | 0.005335 | 0.005278 | 4.5e-05 | 0.001033 | 0.26488 | 1.3312 | 2.5479 | 98.85 | 20.77 | 38.735 | 119.9807 | 107.2614 | 111.7801 | 0.009619 | 0.91717 | 0.021489 | 23.4294 | 0.16165 | 0.162 | 0.04931 | 3.1538 | 3.1538 | 0.002111
2019-03-21 15:57 | 1553198203.8728 | 413 | 0.19959 | 0.007853 | 0.000199 | 99.6081 | 97.6545 | 104.8143 | 21.2612 | 0.004069 | 0.0037 | 236.8263 | 0.27122 | 8.4001 | 0.003501 | 0.004677 | 0.004637 | 4.6e-05 | 0.001035 | 0.28639 | 1.3449 | 2.5481 | 98.85 | 20.77 | 38.735 | 119.9443 | 106.1137 | 110.9435 | 0.009619 | 0.91917 | 0.02147 | 23.3582 | 0.17717 | 0.16135 | 0.04992 | 3.19 | 3.19 | 0.002079
2019-03-21 15:57 | 1553198204.8739 | 414 | 0.19948 | 0.007875 | 0.0002 | 99.6427 | 97.426 | 104.3519 | 21.2612 | 0.004073 | 0.003723 | 236.7228 | 0.27204 | 8.4037 | 0.003546 | 0.005339 | 0.005288 | 4.7e-05 | 0.001036 | 0.27527 | 1.3334 | 2.548 | 98.85 | 20.77 | 38.735 | 120.0465 | 107.2007 | 111.5142 | 0.009619 | 0.9178 | 0.021407 | 23.324 | 0.20123 | 0.16183 | 0.050538 | 3.2199 | 3.2199 | 0.002066
2019-03-21 15:57 | 1553198205.8725 | 415 | 0.19947 | 0.007896 | 0.000199 | 99.6618 | 97.7505 | 104.0883 | 21.2612 | 0.00408 | 0.003769 | 236.5915 | 0.2713 | 8.4033 | 0.00351 | 0.005227 | 0.005175 | 4.8e-05 | 0.001032 | 0.25442 | 1.3497 | 2.5481 | 98.85 | 20.77 | 38.735 | 120.2211 | 106.3371 | 112.6266 | 0.009619 | 0.91516 | 0.021349 | 23.3281 | 0.1732 | 0.16139 | 0.050568 | 3.2131 | 3.2131 | 0.002064
2019-03-21 15:57 | 1553198206.8726 | 416 | 0.19928 | 0.007929 | 0.000198 | 99.6507 | 97.7247 | 104.0078 | 21.2612 | 0.004136 | 0.003789 | 236.4603 | 0.2704 | 8.401 | 0.003516 | 0.004693 | 0.004638 | 4.6e-05 | 0.001035 | 0.23595 | 1.2998 | 2.5481 | 98.86 | 20.76 | 38.735 | 121.586 | 106.4781 | 113.1133 | 0.009612 | 0.914 | 0.021327 | 23.3338 | 0.17458 | 0.16086 | 0.05167 | 3.2797 | 3.2797 | 0.002016
2019-03-21 15:57 | 1553198207.873 | 417 | 0.19925 | 0.007969 | 0.000197 | 99.6647 | 97.864 | 104.1021 | 21.2612 | 0.00416 | 0.003721 | 236.3105 | 0.27142 | 8.4022 | 0.003534 | 0.005293 | 0.00525 | 4.6e-05 | 0.001034 | 0.27887 | 1.3189 | 2.548 | 98.86 | 20.77 | 38.734 | 122.1817 | 106.9056 | 111.4614 | 0.009618 | 0.91793 | 0.021383 | 23.2944 | 0.16309 | 0.16147 | 0.051826 | 3.2981 | 3.2981 | 0.002012
2019-03-21 15:57 | 1553198208.8732 | 418 | 0.19922 | 0.008009 | 0.000197 | 99.7202 | 97.8623 | 104.0019 | 21.2612 | 0.004165 | 0.003757 | 236.1841 | 0.27168 | 8.4031 | 0.003552 | 0.004816 | 0.004702 | 4.6e-05 | 0.001035 | 0.27184 | 1.3596 | 2.5481 | 98.86 | 20.76 | 38.735 | 122.2929 | 107.3614 | 112.3313 | 0.009612 | 0.91586 | 0.021348 | 23.3095 | 0.16822 | 0.16162 | 0.051923 | 3.299 | 3.299 | 0.002014
2019-03-21 15:57 | 1553198209.8721 | 419 | 0.19913 | 0.008056 | 0.000198 | 99.7869 | 97.9225 | 103.9489 | 21.2612 | 0.004212 | 0.003834 | 236.0565 | 0.27176 | 8.4017 | 0.003483 | 0.005448 | 0.005369 | 4.8e-05 | 0.001035 | 0.27671 | 1.3849 | 2.548 | 98.86 | 20.76 | 38.736 | 123.4385 | 105.6825 | 114.2063 | 0.009613 | 0.91142 | 0.021291 | 23.3602 | 0.1687 | 0.16167 | 0.052419 | 3.3215 | 3.3215 | 0.002
2019-03-21 15:57 | 1553198210.8722 | 420 | 0.19909 | 0.008107 | 0.000198 | 99.7237 | 97.8817 | 103.7766 | 21.2612 | 0.004152 | 0.003783 | 235.9062 | 0.27148 | 8.4054 | 0.003598 | 0.004957 | 0.004888 | 4.8e-05 | 0.001035 | 0.25498 | 1.3699 | 2.5479 | 98.85 | 20.76 | 38.735 | 121.9871 | 108.4589 | 112.9718 | 0.009613 | 0.91434 | 0.021307 | 23.3036 | 0.16677 | 0.16151 | 0.052575 | 3.3339 | 3.3339 | 0.001991
2019-03-21 15:57 | 1553198211.8728 | 421 | 0.19899 | 0.008151 | 0.000198 | 99.7321 | 98.0646 | 103.6461 | 21.2612 | 0.004198 | 0.003806 | 235.7726 | 0.27107 | 8.4021 | 0.003608 | 0.005224 | 0.005126 | 4.5e-05 | 0.001036 | 0.28208 | 1.3416 | 2.548 | 98.85 | 20.76 | 38.735 | 123.1088 | 108.7149 | 113.5375 | 0.009613 | 0.913 | 0.021278 | 23.306 | 0.15082 | 0.16126 | 0.05314 | 3.3651 | 3.3651 | 0.00197
2019-03-21 15:57 | 1553198212.8728 | 422 | 0.19896 | 0.008199 | 0.000198 | 99.7151 | 97.7591 | 103.5116 | 21.2612 | 0.004152 | 0.003789 | 235.6423 | 0.27235 | 8.4042 | 0.00371 | 0.004867 | 0.0048 | 4.6e-05 | 0.001035 | 0.28224 | 1.3164 | 2.548 | 98.85 | 20.76 | 38.734 | 121.9822 | 111.1882 | 113.1055 | 0.009613 | 0.91402 | 0.021276 | 23.2779 | 0.1772 | 0.16202 | 0.053256 | 3.372 | 3.372 | 0.001975
2019-03-21 15:57 | 1553198213.873 | 423 | 0.19894 | 0.008246 | 0.000197 | 99.7604 | 97.7471 | 103.1762 | 21.2612 | 0.004202 | 0.003794 | 235.4974 | 0.27135 | 8.4014 | 0.003691 | 0.00491 | 0.00483 | 4.9e-05 | 0.001036 | 0.28567 | 1.3275 | 2.548 | 98.85 | 20.76 | 38.734 | 123.1994 | 110.7273 | 113.2421 | 0.009613 | 0.9137 | 0.021238 | 23.2442 | 0.18235 | 0.16143 | 0.053302 | 3.3689 | 3.3689 | 0.001969
2019-03-21 15:57 | 1553198214.8721 | 424 | 0.19889 | 0.008288 | 0.000197 | 99.7503 | 97.7662 | 103.2036 | 21.2612 | 0.004208 | 0.003953 | 235.3568 | 0.27206 | 8.4054 | 0.003734 | 0.00508 | 0.004968 | 4.7e-05 | 0.001033 | 0.27184 | 1.3272 | 2.5481 | 98.85 | 20.76 | 38.734 | 123.3417 | 111.7786 | 117.1085 | 0.009613 | 0.90465 | 0.021136 | 23.3633 | 0.1797 | 0.16185 | 0.053549 | 3.3682 | 3.3682 | 0.001975
2019-03-21 15:57 | 1553198215.8734 | 425 | 0.19873 | 0.008334 | 0.000196 | 99.783 | 97.7829 | 103.1081 | 21.2612 | 0.004192 | 0.003892 | 235.2208 | 0.27175 | 8.4028 | 0.003748 | 0.005535 | 0.005442 | 4.8e-05 | 0.001029 | 0.27649 | 1.3035 | 2.548 | 98.85 | 20.77 | 38.733 | 122.9601 | 112.1196 | 115.6242 | 0.009619 | 0.9081 | 0.021166 | 23.3081 | 0.18111 | 0.16166 | 0.054439 | 3.429 | 3.429 | 0.001938
2019-03-21 15:57 | 1553198216.8731 | 426 | 0.19855 | 0.00838 | 0.000195 | 99.7389 | 97.6681 | 102.8717 | 21.2612 | 0.004301 | 0.003913 | 235.0682 | 0.272 | 8.4052 | 0.003801 | 0.005186 | 0.005123 | 4.8e-05 | 0.00103 | 0.24994 | 1.3433 | 2.5478 | 98.85 | 20.77 | 38.732 | 125.6265 | 113.4143 | 116.145 | 0.009619 | 0.90688 | 0.021128 | 23.2969 | 0.18766 | 0.16181 | 0.055493 | 3.4889 | 3.4889 | 0.001906
2019-03-21 15:57 | 1553198217.8721 | 427 | 0.19834 | 0.008432 | 0.000194 | 99.8557 | 97.6882 | 102.8971 | 21.2612 | 0.004295 | 0.003917 | 234.9445 | 0.27237 | 8.4049 | 0.003684 | 0.005592 | 0.005608 | 4.8e-05 | 0.001029 | 0.28816 | 1.3403 | 2.548 | 98.85 | 20.77 | 38.731 | 125.4865 | 110.5688 | 116.2455 | 0.009618 | 0.90665 | 0.021127 | 23.3028 | 0.19629 | 0.16203 | 0.056661 | 3.5622 | 3.5622 | 0.001869
2019-03-21 15:57 | 1553198218.8726 | 428 | 0.19824 | 0.008493 | 0.000192 | 99.8408 | 97.7465 | 102.7497 | 21.2612 | 0.004304 | 0.00385 | 234.7929 | 0.27157 | 8.4029 | 0.003737 | 0.005218 | 0.005122 | 4.5e-05 | 0.001027 | 0.28407 | 1.3195 | 2.5481 | 98.85 | 20.77 | 38.73 | 125.691 | 111.8536 | 114.6008 | 0.009618 | 0.9105 | 0.021157 | 23.2369 | 0.18963 | 0.16156 | 0.057211 | 3.6017 | 3.6017 | 0.001844
2019-03-21 15:57 | 1553198219.8732 | 429 | 0.19809 | 0.008556 | 0.000192 | 99.8894 | 97.9546 | 102.521 | 21.2612 | 0.00431 | 0.003909 | 234.6262 | 0.27123 | 8.4006 | 0.003724 | 0.004878 | 0.004863 | 4.8e-05 | 0.001027 | 0.26073 | 1.3452 | 2.5481 | 98.85 | 20.77 | 38.729 | 125.8566 | 111.5314 | 116.0391 | 0.009618 | 0.90713 | 0.021094 | 23.254 | 0.17495 | 0.16135 | 0.058044 | 3.6432 | 3.6432 | 0.00182
2019-03-21 15:57 | 1553198220.8733 | 430 | 0.19805 | 0.008623 | 0.000192 | 99.7613 | 97.8442 | 102.2884 | 21.2612 | 0.004258 | 0.003952 | 234.4951 | 0.27194 | 8.406 | 0.003762 | 0.004815 | 0.004675 | 4.7e-05 | 0.001026 | 0.28854 | 1.3183 | 2.5479 | 98.85 | 20.77 | 38.726 | 124.5731 | 112.4616 | 117.0761 | 0.009617 | 0.90472 | 0.021042 | 23.2586 | 0.17356 | 0.16178 | 0.058186 | 3.6431 | 3.6431 | 0.001825
2019-03-21 15:57 | 1553198221.8718 | 431 | 0.19793 | 0.0087 | 0.000195 | 99.8173 | 97.883 | 102.1894 | 21.2612 | 0.004374 | 0.003888 | 234.3497 | 0.27164 | 8.4019 | 0.003803 | 0.005141 | 0.004997 | 4.4e-05 | 0.001026 | 0.27831 | 1.3113 | 2.5479 | 98.85 | 20.77 | 38.726 | 127.4283 | 113.4633 | 115.5333 | 0.009617 | 0.90831 | 0.021074 | 23.2013 | 0.17503 | 0.1616 | 0.058801 | 3.6868 | 3.6868 | 0.001801
2019-03-21 15:57 | 1553198222.8718 | 432 | 0.19793 | 0.00878 | 0.000197 | 99.8608 | 97.8442 | 102.1462 | 21.2612 | 0.004321 | 0.003988 | 234.2022 | 0.27144 | 8.3992 | 0.003736 | 0.005592 | 0.00556 | 4.7e-05 | 0.00103 | 0.28495 | 1.2887 | 2.5479 | 98.85 | 20.76 | 38.726 | 126.128 | 111.8181 | 117.9735 | 0.009611 | 0.90264 | 0.021004 | 23.2691 | 0.18247 | 0.16148 | 0.05869 | 3.6675 | 3.6675 | 0.00181
2019-03-21 15:57 | 1553198223.8728 | 433 | 0.19793 | 0.008863 | 0.000197 | 99.8653 | 97.4232 | 102.2903 | 21.2612 | 0.004342 | 0.003945 | 234.0598 | 0.27071 | 8.4028 | 0.003765 | 0.004867 | 0.004787 | 4.7e-05 | 0.001031 | 0.30934 | 1.3275 | 2.548 | 98.85 | 20.77 | 38.725 | 126.6374 | 112.5192 | 116.9138 | 0.009617 | 0.9051 | 0.021047 | 23.2539 | 0.22144 | 0.16104 | 0.0586 | 3.6694 | 3.6694 | 0.001804
2019-03-21 15:57 | 1553198224.8725 | 434 | 0.19793 | 0.008942 | 0.000198 | 99.8656 | 97.5758 | 102.3516 | 21.2612 | 0.004411 | 0.003971 | 233.9125 | 0.27199 | 8.4008 | 0.00374 | 0.005924 | 0.005833 | 4.9e-05 | 0.001031 | 0.27096 | 1.3108 | 2.548 | 98.85 | 20.77 | 38.723 | 128.3393 | 111.9327 | 117.5521 | 0.009616 | 0.90362 | 0.021036 | 23.2799 | 0.20747 | 0.16181 | 0.058522 | 3.6627 | 3.6627 | 0.001816
2019-03-21 15:57 | 1553198225.8715 | 435 | 0.19791 | 0.009007 | 0.000198 | 99.7982 | 97.4621 | 102.5084 | 21.2612 | 0.004397 | 0.00397 | 233.7543 | 0.27137 | 8.4011 | 0.003766 | 0.005794 | 0.00579 | 4.7e-05 | 0.001031 | 0.28368 | 1.3034 | 2.548 | 98.85 | 20.77 | 38.721 | 127.9838 | 112.5555 | 117.5362 | 0.009616 | 0.90365 | 0.021053 | 23.2973 | 0.21187 | 0.16144 | 0.058564 | 3.6682 | 3.6682 | 0.001809
2019-03-21 15:57 | 1553198226.8716 | 436 | 0.19791 | 0.009063 | 0.000199 | 99.9064 | 97.8864 | 102.3387 | 21.2612 | 0.004381 | 0.003996 | 233.6266 | 0.2717 | 8.4025 | 0.003728 | 0.005586 | 0.005517 | 5e-05 | 0.00103 | 0.24154 | 1.299 | 2.5479 | 98.85 | 20.77 | 38.72 | 127.5916 | 111.6397 | 118.1652 | 0.009616 | 0.9022 | 0.021018 | 23.2967 | 0.1827 | 0.16164 | 0.058493 | 3.6577 | 3.6577 | 0.001816
2019-03-21 15:57 | 1553198227.8717 | 437 | 0.19788 | 0.009107 | 0.000199 | 99.8617 | 97.8101 | 101.9504 | 21.2612 | 0.004351 | 0.004016 | 233.4806 | 0.2717 | 8.4031 | 0.00381 | 0.005248 | 0.005239 | 4.3e-05 | 0.001029 | 0.27345 | 1.2849 | 2.5479 | 98.85 | 20.77 | 38.718 | 126.8598 | 113.623 | 118.6415 | 0.009615 | 0.90111 | 0.020966 | 23.2666 | 0.18567 | 0.16163 | 0.05858 | 3.6539 | 3.6539 | 0.001818
2019-03-21 15:57 | 1553198228.8721 | 438 | 0.19784 | 0.009144 | 0.0002 | 99.8974 | 97.7206 | 102.0938 | 21.2612 | 0.004407 | 0.003951 | 233.3326 | 0.27081 | 8.4036 | 0.003801 | 0.005024 | 0.004957 | 4.8e-05 | 0.001031 | 0.28744 | 1.3392 | 2.548 | 98.85 | 20.77 | 38.717 | 128.226 | 113.4065 | 117.0753 | 0.009615 | 0.90472 | 0.021022 | 23.2364 | 0.19707 | 0.16111 | 0.058771 | 3.6757 | 3.6757 | 0.001801
2019-03-21 15:57 | 1553198229.8732 | 439 | 0.19783 | 0.009184 | 0.0002 | 99.9198 | 97.7475 | 102.3387 | 21.2612 | 0.00442 | 0.004085 | 233.211 | 0.27254 | 8.401 | 0.003773 | 0.005223 | 0.005134 | 4.4e-05 | 0.001032 | 0.29192 | 1.3226 | 2.548 | 98.85 | 20.77 | 38.717 | 128.5599 | 112.7167 | 120.3296 | 0.009615 | 0.89724 | 0.02096 | 23.361 | 0.1966 | 0.16213 | 0.058803 | 3.6669 | 3.6669 | 0.001817
2019-03-21 15:57 | 1553198230.8719 | 440 | 0.19784 | 0.009219 | 0.0002 | 99.9749 | 98.0523 | 102.6957 | 21.2612 | 0.004398 | 0.004007 | 233.0622 | 0.27124 | 8.4007 | 0.003725 | 0.005406 | 0.005321 | 4.9e-05 | 0.00103 | 0.26864 | 1.2987 | 2.5479 | 98.85 | 20.77 | 38.718 | 128.0204 | 111.556 | 118.4258 | 0.009615 | 0.9016 | 0.021048 | 23.3451 | 0.17369 | 0.16136 | 0.058691 | 3.6752 | 3.6752 | 0.001804
2019-03-21 15:57 | 1553198231.8719 | 441 | 0.19794 | 0.00925 | 0.0002 | 99.9436 | 98.0295 | 102.8285 | 21.2612 | 0.004516 | 0.004008 | 232.9154 | 0.27176 | 8.4003 | 0.003746 | 0.005118 | 0.00502 | 4.9e-05 | 0.001031 | 0.24298 | 1.2861 | 2.548 | 98.85 | 20.77 | 38.718 | 130.9155 | 112.058 | 118.46 | 0.009615 | 0.90152 | 0.021061 | 23.3612 | 0.17297 | 0.16167 | 0.058077 | 3.639 | 3.639 | 0.001826
2019-03-21 15:57 | 1553198232.8729 | 442 | 0.1981 | 0.009281 | 0.000201 | 100.0285 | 97.6259 | 102.84 | 21.2612 | 0.004468 | 0.004003 | 232.7613 | 0.27198 | 8.4048 | 0.003834 | 0.004601 | 0.004446 | 4.5e-05 | 0.00103 | 0.26626 | 1.2751 | 2.5478 | 98.85 | 20.77 | 38.718 | 129.7306 | 114.2189 | 118.3332 | 0.009615 | 0.90182 | 0.021065 | 23.3587 | 0.21747 | 0.1618 | 0.057098 | 3.5786 | 3.5786 | 0.001858
2019-03-21 15:57 | 1553198233.872 | 443 | 0.19816 | 0.0093 | 0.000201 | 99.9418 | 97.5867 | 103.1217 | 21.2612 | 0.004455 | 0.004018 | 232.6182 | 0.27155 | 8.4029 | 0.00375 | 0.005377 | 0.005419 | 4.9e-05 | 0.001031 | 0.26449 | 1.3141 | 2.5479 | 98.85 | 20.77 | 38.717 | 129.4011 | 112.1529 | 118.6906 | 0.009615 | 0.90099 | 0.021084 | 23.4014 | 0.2133 | 0.16154 | 0.05669 | 3.5564 | 3.5564 | 0.001867
2019-03-21 15:57 | 1553198234.8734 | 444 | 0.19825 | 0.009304 | 0.000201 | 99.9448 | 97.5946 | 102.8086 | 21.2612 | 0.004389 | 0.00406 | 232.4772 | 0.2728 | 8.4042 | 0.003777 | 0.005534 | 0.005405 | 4.9e-05 | 0.001033 | 0.24187 | 1.3039 | 2.5479 | 98.85 | 20.77 | 38.717 | 127.7837 | 112.8218 | 119.715 | 0.009615 | 0.89864 | 0.021025 | 23.3963 | 0.21285 | 0.16229 | 0.056157 | 3.5131 | 3.5131 | 0.001899
2019-03-21 15:57 | 1553198235.8725 | 445 | 0.19826 | 0.009297 | 0.000202 | 99.9597 | 97.8258 | 102.8926 | 21.2612 | 0.004492 | 0.004005 | 232.3644 | 0.27127 | 8.3995 | 0.003864 | 0.005085 | 0.004979 | 4.4e-05 | 0.001035 | 0.28672 | 1.3032 | 2.5479 | 98.85 | 20.77 | 38.716 | 130.3138 | 114.9479 | 118.3901 | 0.009615 | 0.90168 | 0.021069 | 23.3664 | 0.19301 | 0.16138 | 0.056106 | 3.5172 | 3.5172 | 0.001886
2019-03-21 15:57 | 1553198236.8711 | 446 | 0.1983 | 0.009288 | 0.000202 | 100.0187 | 97.8503 | 103.2752 | 21.2612 | 0.004484 | 0.004071 | 232.1982 | 0.27111 | 8.4019 | 0.003854 | 0.005292 | 0.005185 | 5e-05 | 0.001036 | 0.2672 | 1.3214 | 2.5481 | 98.85 | 20.77 | 38.717 | 130.12 | 114.7051 | 120.0013 | 0.009615 | 0.89799 | 0.021065 | 23.4579 | 0.19605 | 0.16129 | 0.055889 | 3.5029 | 3.5029 | 0.001892
2019-03-21 15:57 | 1553198237.8712 | 447 | 0.19828 | 0.009272 | 0.000202 | 100.109 | 98.1712 | 103.2583 | 21.2612 | 0.004467 | 0.004049 | 232.0457 | 0.27144 | 8.4042 | 0.003882 | 0.005373 | 0.00537 | 4.6e-05 | 0.001033 | 0.27688 | 1.3267 | 2.5479 | 98.85 | 20.77 | 38.719 | 129.7028 | 115.3806 | 119.4557 | 0.009615 | 0.89924 | 0.021078 | 23.4397 | 0.17484 | 0.16148 | 0.056038 | 3.5144 | 3.5144 | 0.001888
2019-03-21 15:57 | 1553198238.8713 | 448 | 0.19826 | 0.009255 | 0.0002 | 100.1394 | 98.0782 | 103.1212 | 21.2612 | 0.004543 | 0.004026 | 231.8922 | 0.27126 | 8.4028 | 0.00389 | 0.006069 | 0.005959 | 4.8e-05 | 0.001035 | 0.27328 | 1.2821 | 2.548 | 98.85 | 20.77 | 38.72 | 131.575 | 115.5713 | 118.8936 | 0.009616 | 0.90053 | 0.021079 | 23.4074 | 0.18603 | 0.16137 | 0.056134 | 3.5207 | 3.5207 | 0.001884
2019-03-21 15:57 | 1553198239.8726 | 449 | 0.19824 | 0.009244 | 0.000199 | 100.0494 | 97.7321 | 102.887 | 21.2612 | 0.00458 | 0.004098 | 231.7594 | 0.27209 | 8.4036 | 0.003879 | 0.005051 | 0.004943 | 4.7e-05 | 0.001037 | 0.28031 | 1.3067 | 2.5481 | 98.85 | 20.77 | 38.721 | 132.4856 | 115.307 | 120.6591 | 0.009616 | 0.89649 | 0.021008 | 23.4333 | 0.2096 | 0.16187 | 0.056305 | 3.5195 | 3.5195 | 0.00189
2019-03-21 15:57 | 1553198240.8713 | 450 | 0.19816 | 0.009228 | 0.000198 | 100.1185 | 98.0646 | 102.8065 | 21.2612 | 0.004456 | 0.004072 | 231.6233 | 0.27104 | 8.4024 | 0.003857 | 0.005124 | 0.005033 | 4.8e-05 | 0.001036 | 0.25409 | 1.2933 | 2.548 | 98.85 | 20.77 | 38.722 | 129.4395 | 114.7614 | 120.0189 | 0.009616 | 0.89795 | 0.021017 | 23.4051 | 0.18541 | 0.16124 | 0.056791 | 3.5513 | 3.5513 | 0.001866
2019-03-21 15:57 | 1553198241.8718 | 451 | 0.19802 | 0.009218 | 0.0002 | 100.1444 | 98.2713 | 102.7236 | 21.2612 | 0.004566 | 0.004077 | 231.4782 | 0.27138 | 8.4026 | 0.003841 | 0.005921 | 0.005796 | 5.1e-05 | 0.001035 | 0.26737 | 1.2965 | 2.548 | 98.85 | 20.77 | 38.722 | 132.1325 | 114.381 | 120.1409 | 0.009616 | 0.89767 | 0.021005 | 23.3993 | 0.16889 | 0.16144 | 0.057655 | 3.6031 | 3.6031 | 0.001842
2019-03-21 15:57 | 1553198242.8715 | 452 | 0.19799 | 0.009217 | 0.0002 | 100.137 | 98.0434 | 102.7886 | 21.2612 | 0.004576 | 0.004069 | 231.3449 | 0.27113 | 8.4002 | 0.003812 | 0.005623 | 0.005514 | 5e-05 | 0.001035 | 0.27865 | 1.2956 | 2.548 | 98.85 | 20.77 | 38.722 | 132.3873 | 113.664 | 119.956 | 0.009616 | 0.89809 | 0.021016 | 23.4012 | 0.18898 | 0.1613 | 0.057798 | 3.614 | 3.614 | 0.001834
2019-03-21 15:57 | 1553198243.8716 | 453 | 0.19795 | 0.009227 | 0.000203 | 100.1447 | 97.9362 | 102.6102 | 21.2612 | 0.004558 | 0.004113 | 231.2049 | 0.2714 | 8.4011 | 0.003866 | 0.005261 | 0.005233 | 5.2e-05 | 0.001034 | 0.25033 | 1.3283 | 2.5481 | 98.85 | 20.77 | 38.723 | 131.9351 | 114.9947 | 121.0321 | 0.009616 | 0.89564 | 0.02097 | 23.4129 | 0.19947 | 0.16145 | 0.058027 | 3.62 | 3.62 | 0.001833
2019-03-21 15:57 | 1553198244.8726 | 454 | 0.19789 | 0.00925 | 0.000205 | 100.1876 | 97.913 | 102.9008 | 21.2612 | 0.004557 | 0.004073 | 231.0696 | 0.27135 | 8.403 | 0.003872 | 0.00522 | 0.005154 | 4.9e-05 | 0.001034 | 0.2625 | 1.2954 | 2.548 | 98.85 | 20.77 | 38.722 | 131.9297 | 115.1402 | 120.0409 | 0.009616 | 0.8979 | 0.021026 | 23.4165 | 0.20542 | 0.16143 | 0.058378 | 3.6515 | 3.6515 | 0.001817
2019-03-21 15:57 | 1553198245.872 | 455 | 0.19794 | 0.009274 | 0.000208 | 100.2055 | 98.0874 | 102.8743 | 21.2612 | 0.004593 | 0.004129 | 230.9211 | 0.27059 | 8.4048 | 0.003942 | 0.005036 | 0.00496 | 4.9e-05 | 0.001033 | 0.27417 | 1.2953 | 2.548 | 98.85 | 20.77 | 38.723 | 132.7988 | 116.8432 | 121.424 | 0.009616 | 0.89475 | 0.020986 | 23.4547 | 0.1911 | 0.16097 | 0.058039 | 3.6234 | 3.6234 | 0.001826
2019-03-21 15:57 | 1553198246.8727 | 456 | 0.19802 | 0.0093 | 0.000212 | 100.1832 | 98.1606 | 102.5152 | 21.2612 | 0.00456 | 0.004117 | 230.8108 | 0.26993 | 8.4029 | 0.003959 | 0.005878 | 0.005748 | 5e-05 | 0.001031 | 0.26698 | 1.3379 | 2.548 | 98.85 | 20.77 | 38.723 | 131.9848 | 117.2574 | 121.113 | 0.009616 | 0.89546 | 0.020958 | 23.4044 | 0.18243 | 0.16058 | 0.057549 | 3.5878 | 3.5878 | 0.00184
2019-03-21 15:57 | 1553198247.8728 | 457 | 0.19814 | 0.009325 | 0.000214 | 100.2684 | 98.3071 | 102.8663 | 21.2612 | 0.004563 | 0.004097 | 230.6802 | 0.27036 | 8.4018 | 0.003985 | 0.00542 | 0.005338 | 4.9e-05 | 0.00103 | 0.24027 | 1.3034 | 2.548 | 98.85 | 20.77 | 38.724 | 132.0702 | 117.8895 | 120.6351 | 0.009617 | 0.89654 | 0.021006 | 23.4303 | 0.1767 | 0.16084 | 0.056784 | 3.5484 | 3.5484 | 0.001863
2019-03-21 15:57 | 1553198248.8718 | 458 | 0.19818 | 0.009336 | 0.000216 | 100.2922 | 97.9522 | 102.8551 | 21.2612 | 0.004562 | 0.004097 | 230.56 | 0.27144 | 8.4 | 0.004029 | 0.005499 | 0.00542 | 4.6e-05 | 0.00103 | 0.27312 | 1.3167 | 2.5479 | 98.85 | 20.77 | 38.725 | 132.0471 | 118.9672 | 120.6392 | 0.009617 | 0.89653 | 0.021005 | 23.4291 | 0.21117 | 0.16148 | 0.056518 | 3.5314 | 3.5314 | 0.001879
2019-03-21 15:57 | 1553198249.8721 | 459 | 0.19821 | 0.009342 | 0.000218 | 100.2413 | 98.0308 | 103.032 | 21.2612 | 0.004591 | 0.004131 | 230.4182 | 0.27165 | 8.4048 | 0.003988 | 0.005584 | 0.005479 | 5e-05 | 0.001031 | 0.26432 | 1.3536 | 2.5479 | 98.85 | 20.77 | 38.727 | 132.7547 | 117.9679 | 121.471 | 0.009617 | 0.89464 | 0.021001 | 23.474 | 0.19944 | 0.1616 | 0.056335 | 3.5192 | 3.5192 | 0.001887
2019-03-21 15:58 | 1553198250.8719 | 460 | 0.19822 | 0.009335 | 0.000219 | 100.2875 | 97.9603 | 103.0496 | 21.2612 | 0.004585 | 0.004076 | 230.2825 | 0.27167 | 8.4033 | 0.003972 | 0.005584 | 0.005496 | 4.8e-05 | 0.001029 | 0.28567 | 1.3361 | 2.5481 | 98.85 | 20.77 | 38.727 | 132.6109 | 117.5838 | 120.1091 | 0.009617 | 0.89774 | 0.021039 | 23.4355 | 0.21 | 0.16162 | 0.056294 | 3.5231 | 3.5231 | 0.001885
2019-03-21 15:58 | 1553198251.871 | 461 | 0.19823 | 0.009321 | 0.000218 | 100.3455 | 98.3292 | 103.0003 | 21.2612 | 0.004607 | 0.004097 | 230.1881 | 0.2714 | 8.3999 | 0.003947 | 0.005463 | 0.005374 | 4.4e-05 | 0.001025 | 0.28008 | 1.3586 | 2.5481 | 98.85 | 20.77 | 38.728 | 133.1375 | 116.9557 | 120.6248 | 0.009618 | 0.89657 | 0.02102 | 23.4452 | 0.18156 | 0.16145 | 0.056235 | 3.5163 | 3.5163 | 0.001887
2019-03-21 15:58 | 1553198252.8711 | 462 | 0.19825 | 0.009302 | 0.000217 | 100.3601 | 98.212 | 103.0637 | 21.2612 | 0.004537 | 0.004083 | 230.0751 | 0.27101 | 8.4018 | 0.003956 | 0.005539 | 0.005456 | 4.5e-05 | 0.001027 | 0.29031 | 1.3525 | 2.5479 | 98.85 | 20.77 | 38.729 | 131.4273 | 117.1791 | 120.2893 | 0.009618 | 0.89733 | 0.021036 | 23.4425 | 0.19353 | 0.16122 | 0.056133 | 3.5125 | 3.5125 | 0.001886
2019-03-21 15:58 | 1553198253.871 | 463 | 0.19847 | 0.009296 | 0.000216 | 100.4269 | 98.0135 | 102.921 | 21.2612 | 0.004592 | 0.004122 | 229.954 | 0.27148 | 8.4018 | 0.003986 | 0.005698 | 0.005624 | 4.7e-05 | 0.001025 | 0.25642 | 1.3716 | 2.548 | 98.85 | 20.77 | 38.729 | 132.786 | 117.9156 | 121.2368 | 0.009618 | 0.89518 | 0.020996 | 23.4544 | 0.21758 | 0.1615 | 0.054863 | 3.4268 | 3.4268 | 0.001937
2019-03-21 15:58 | 1553198254.8715 | 464 | 0.19873 | 0.009272 | 0.000215 | 100.3849 | 98.3299 | 102.6898 | 21.2612 | 0.004564 | 0.004156 | 229.801 | 0.27201 | 8.4031 | 0.003914 | 0.005133 | 0.005087 | 4.8e-05 | 0.001026 | 0.26161 | 1.3261 | 2.5479 | 98.85 | 20.77 | 38.73 | 132.099 | 116.172 | 122.0792 | 0.009618 | 0.89327 | 0.02095 | 23.4531 | 0.18501 | 0.16182 | 0.053346 | 3.325 | 3.325 | 0.002
2019-03-21 15:58 | 1553198255.8726 | 465 | 0.19881 | 0.009248 | 0.000217 | 100.3905 | 98.4191 | 102.5077 | 21.2612 | 0.004621 | 0.004145 | 229.671 | 0.2712 | 8.4036 | 0.003934 | 0.005137 | 0.005051 | 4.7e-05 | 0.001026 | 0.26145 | 1.2942 | 2.5479 | 98.85 | 20.77 | 38.732 | 133.4836 | 116.6602 | 121.8131 | 0.009619 | 0.89387 | 0.020938 | 23.4244 | 0.1774 | 0.16134 | 0.052876 | 3.2938 | 3.2938 | 0.002013
2019-03-21 15:58 | 1553198256.8715 | 466 | 0.19885 | 0.009214 | 0.000218 | 100.4963 | 98.148 | 102.311 | 21.2612 | 0.004638 | 0.004078 | 229.533 | 0.27081 | 8.4004 | 0.003858 | 0.005135 | 0.005057 | 4.9e-05 | 0.001024 | 0.26808 | 1.3185 | 2.5479 | 98.85 | 20.77 | 38.733 | 133.9121 | 114.7914 | 120.1766 | 0.009619 | 0.89759 | 0.020962 | 23.3533 | 0.21149 | 0.16111 | 0.052696 | 3.2863 | 3.2863 | 0.002015
2019-03-21 15:58 | 1553198257.8713 | 467 | 0.19889 | 0.009168 | 0.000218 | 100.4019 | 98.4034 | 102.3281 | 21.2612 | 0.004635 | 0.0041 | 229.4195 | 0.27059 | 8.4027 | 0.003984 | 0.005383 | 0.005286 | 4.8e-05 | 0.001024 | 0.28871 | 1.3005 | 2.548 | 98.85 | 20.77 | 38.733 | 133.8298 | 117.869 | 120.7004 | 0.009619 | 0.8964 | 0.020949 | 23.3708 | 0.17983 | 0.16097 | 0.052523 | 3.2736 | 3.2736 | 0.002021
2019-03-21 15:58 | 1553198258.8714 | 468 | 0.19887 | 0.009112 | 0.000218 | 100.5404 | 98.2226 | 102.1183 | 21.2612 | 0.004617 | 0.004105 | 229.2901 | 0.27135 | 8.4031 | 0.003907 | 0.005383 | 0.005297 | 5e-05 | 0.001022 | 0.25946 | 1.2914 | 2.5479 | 98.85 | 20.77 | 38.735 | 133.4038 | 115.9803 | 120.8193 | 0.009619 | 0.89612 | 0.020925 | 23.3504 | 0.20862 | 0.16143 | 0.052701 | 3.2808 | 3.2808 | 0.002022
2019-03-21 15:58 | 1553198259.8719 | 469 | 0.1989 | 0.009062 | 0.000216 | 100.5261 | 98.1099 | 102.1485 | 21.2612 | 0.004559 | 0.004144 | 229.1488 | 0.27085 | 8.4016 | 0.003938 | 0.00565 | 0.00556 | 4.9e-05 | 0.001019 | 0.26305 | 1.3589 | 2.548 | 98.85 | 20.77 | 38.736 | 131.9663 | 116.7388 | 121.7869 | 0.00962 | 0.89393 | 0.020902 | 23.3825 | 0.21762 | 0.16113 | 0.052545 | 3.2677 | 3.2677 | 0.002027
2019-03-21 15:58 | 1553198260.8716 | 470 | 0.19881 | 0.009009 | 0.000214 | 100.5493 | 98.3166 | 102.5321 | 21.2612 | 0.004578 | 0.004132 | 229.0157 | 0.27122 | 8.4018 | 0.003946 | 0.004946 | 0.004834 | 5e-05 | 0.001024 | 0.2912 | 1.3563 | 2.5481 | 98.85 | 20.77 | 38.735 | 132.4464 | 116.9509 | 121.4756 | 0.009619 | 0.89463 | 0.02095 | 23.4171 | 0.20086 | 0.16135 | 0.05317 | 3.3141 | 3.3141 | 0.002001
2019-03-21 15:58 | 1553198261.8719 | 471 | 0.19871 | 0.008968 | 0.00021 | 100.5249 | 98.5196 | 103.1426 | 21.2612 | 0.004658 | 0.004141 | 228.8663 | 0.27059 | 8.406 | 0.003907 | 0.005369 | 0.005302 | 5e-05 | 0.001018 | 0.24043 | 1.3707 | 2.548 | 98.85 | 20.77 | 38.735 | 134.4127 | 115.9898 | 121.7194 | 0.009619 | 0.89408 | 0.021006 | 23.494 | 0.18023 | 0.16098 | 0.053817 | 3.3635 | 3.3635 | 0.001967
2019-03-21 15:58 | 1553198262.8729 | 472 | 0.19858 | 0.008933 | 0.000206 | 100.5565 | 98.4637 | 103.462 | 21.2612 | 0.004614 | 0.004151 | 228.7288 | 0.27183 | 8.4022 | 0.003888 | 0.005514 | 0.005433 | 4.9e-05 | 0.001019 | 0.24043 | 1.3861 | 2.548 | 98.85 | 20.77 | 38.736 | 133.3109 | 115.5374 | 121.9507 | 0.00962 | 0.89356 | 0.021032 | 23.5372 | 0.18812 | 0.16171 | 0.05466 | 3.4205 | 3.4205 | 0.001943
2019-03-21 15:58 | 1553198263.8719 | 473 | 0.19851 | 0.008919 | 0.000204 | 100.6116 | 98.7137 | 103.9158 | 21.2612 | 0.004615 | 0.004154 | 228.596 | 0.27093 | 8.4012 | 0.003879 | 0.005397 | 0.005325 | 5.1e-05 | 0.001025 | 0.24027 | 1.3435 | 2.5479 | 98.85 | 20.77 | 38.736 | 133.3351 | 115.3048 | 122.0285 | 0.00962 | 0.89338 | 0.021076 | 23.5911 | 0.17034 | 0.16118 | 0.055071 | 3.4535 | 3.4535 | 0.001918
2019-03-21 15:58 | 1553198264.8717 | 474 | 0.19855 | 0.008919 | 0.000203 | 100.5368 | 98.6354 | 104.5547 | 21.2612 | 0.004663 | 0.004123 | 228.4796 | 0.27086 | 8.4007 | 0.003803 | 0.005728 | 0.005622 | 4.7e-05 | 0.001027 | 0.26808 | 1.3184 | 2.5479 | 98.85 | 20.77 | 38.738 | 134.5406 | 113.4467 | 121.2655 | 0.00962 | 0.89511 | 0.021161 | 23.6407 | 0.17079 | 0.16113 | 0.054836 | 3.4528 | 3.4528 | 0.001918
2019-03-21 15:58 | 1553198265.8718 | 475 | 0.19875 | 0.008923 | 0.000203 | 100.5553 | 98.8643 | 104.8453 | 21.2612 | 0.004605 | 0.004117 | 228.3653 | 0.27014 | 8.4026 | 0.003874 | 0.004929 | 0.004863 | 5e-05 | 0.001026 | 0.23828 | 1.325 | 2.5479 | 98.85 | 20.77 | 38.739 | 133.1121 | 115.1797 | 121.1289 | 0.00962 | 0.89542 | 0.021194 | 23.6694 | 0.1517 | 0.1607 | 0.053646 | 3.3832 | 3.3832 | 0.001952
2019-03-21 15:58 | 1553198266.8727 | 476 | 0.19892 | 0.008935 | 0.000205 | 100.5809 | 98.7246 | 104.8901 | 21.2612 | 0.004583 | 0.004139 | 228.2472 | 0.27071 | 8.4023 | 0.003817 | 0.00517 | 0.005157 | 4.5e-05 | 0.001025 | 0.26808 | 1.2959 | 2.548 | 98.85 | 20.77 | 38.739 | 132.5599 | 113.7945 | 121.6624 | 0.00962 | 0.89421 | 0.021184 | 23.6905 | 0.16662 | 0.16104 | 0.052625 | 3.3174 | 3.3174 | 0.001995
2019-03-21 15:58 | 1553198267.8708 | 477 | 0.19917 | 0.008925 | 0.000207 | 100.6137 | 98.4195 | 104.8887 | 21.2612 | 0.004597 | 0.004121 | 228.1194 | 0.27145 | 8.4017 | 0.003804 | 0.005171 | 0.005059 | 4.7e-05 | 0.001026 | 0.25658 | 1.3026 | 2.548 | 98.85 | 20.77 | 38.738 | 132.8922 | 113.4745 | 121.2075 | 0.00962 | 0.89524 | 0.021196 | 23.6767 | 0.19723 | 0.16149 | 0.051101 | 3.2232 | 3.2232 | 0.002059
2019-03-21 15:58 | 1553198268.8708 | 478 | 0.19926 | 0.008907 | 0.000208 | 100.6229 | 98.6902 | 105.3117 | 21.2612 | 0.00459 | 0.004098 | 228.0031 | 0.27136 | 8.4019 | 0.003846 | 0.004884 | 0.004807 | 4.8e-05 | 0.001029 | 0.27577 | 1.3215 | 2.5479 | 98.85 | 20.77 | 38.737 | 132.722 | 114.498 | 120.6559 | 0.00962 | 0.8965 | 0.021254 | 23.7078 | 0.17347 | 0.16143 | 0.050609 | 3.2008 | 3.2008 | 0.002073
2019-03-21 15:58 | 1553198269.8723 | 479 | 0.19926 | 0.008874 | 0.000209 | 100.6527 | 98.2917 | 105.195 | 21.2612 | 0.004657 | 0.004133 | 227.8888 | 0.27048 | 8.4006 | 0.003878 | 0.005091 | 0.005039 | 4.8e-05 | 0.001023 | 0.28136 | 1.2926 | 2.548 | 98.85 | 20.77 | 38.736 | 134.388 | 115.2904 | 121.5226 | 0.00962 | 0.89453 | 0.021219 | 23.7207 | 0.21231 | 0.16091 | 0.050617 | 3.1961 | 3.1961 | 0.002069
2019-03-21 15:58 | 1553198270.8719 | 480 | 0.19928 | 0.00882 | 0.00021 | 100.594 | 98.555 | 104.8943 | 21.2612 | 0.004658 | 0.004107 | 227.7856 | 0.27046 | 8.4026 | 0.003857 | 0.005126 | 0.005087 | 4.9e-05 | 0.001027 | 0.29374 | 1.2934 | 2.5481 | 98.85 | 20.77 | 38.735 | 134.408 | 114.7773 | 120.8716 | 0.009619 | 0.89601 | 0.021206 | 23.6672 | 0.18317 | 0.16089 | 0.050563 | 3.1907 | 3.1907 | 0.002073
2019-03-21 15:58 | 1553198271.872 | 481 | 0.19934 | 0.008772 | 0.000209 | 100.6366 | 98.4961 | 104.6641 | 21.2612 | 0.004607 | 0.004118 | 227.7015 | 0.27161 | 8.4003 | 0.003894 | 0.00547 | 0.005306 | 4.8e-05 | 0.001022 | 0.26305 | 1.3538 | 2.5481 | 98.85 | 20.76 | 38.731 | 133.1599 | 115.683 | 121.1541 | 0.009612 | 0.89536 | 0.021175 | 23.6497 | 0.1923 | 0.16158 | 0.050283 | 3.1685 | 3.1685 | 0.002096
2019-03-21 15:58 | 1553198272.8711 | 482 | 0.19938 | 0.008715 | 0.000208 | 100.6325 | 98.4069 | 104.3998 | 21.2612 | 0.004626 | 0.004149 | 227.553 | 0.27123 | 8.402 | 0.003924 | 0.005548 | 0.005455 | 4.7e-05 | 0.001025 | 0.25962 | 1.3626 | 2.5481 | 98.85 | 20.77 | 38.73 | 133.6253 | 116.4055 | 121.9125 | 0.009618 | 0.89364 | 0.021128 | 23.6425 | 0.20004 | 0.16135 | 0.050138 | 3.1524 | 3.1524 | 0.002104
2019-03-21 15:58 | 1553198273.8715 | 483 | 0.19945 | 0.008667 | 0.000209 | 100.6646 | 98.4433 | 104.1336 | 21.2612 | 0.004643 | 0.004121 | 227.464 | 0.27112 | 8.403 | 0.004014 | 0.004943 | 0.00491 | 5e-05 | 0.001029 | 0.27831 | 1.3014 | 2.5479 | 98.85 | 20.77 | 38.729 | 134.0289 | 118.6034 | 121.2178 | 0.009618 | 0.89522 | 0.02112 | 23.5916 | 0.19959 | 0.16129 | 0.049781 | 3.1287 | 3.1287 | 0.002119
2019-03-21 15:58 | 1553198274.8712 | 484 | 0.19945 | 0.00862 | 0.000208 | 100.5955 | 98.647 | 103.7313 | 21.2612 | 0.004634 | 0.004136 | 227.3308 | 0.27079 | 8.4017 | 0.003924 | 0.005839 | 0.005735 | 4.5e-05 | 0.001027 | 0.28783 | 1.3051 | 2.548 | 98.85 | 20.77 | 38.728 | 133.8116 | 116.3992 | 121.5956 | 0.009618 | 0.89436 | 0.021069 | 23.5573 | 0.17496 | 0.16109 | 0.049842 | 3.125 | 3.125 | 0.002119
2019-03-21 15:58 | 1553198275.8722 | 485 | 0.19965 | 0.008579 | 0.000207 | 100.7305 | 98.666 | 103.6425 | 21.2612 | 0.0046 | 0.004109 | 227.1929 | 0.27088 | 8.4005 | 0.003965 | 0.005508 | 0.005432 | 4.9e-05 | 0.001029 | 0.28008 | 1.3246 | 2.548 | 98.85 | 20.77 | 38.725 | 132.9665 | 117.3992 | 120.9225 | 0.009617 | 0.89589 | 0.021078 | 23.5271 | 0.18522 | 0.16115 | 0.048698 | 3.0547 | 3.0547 | 0.002168
2019-03-21 15:58 | 1553198276.8725 | 486 | 0.19973 | 0.008533 | 0.000206 | 100.6584 | 98.8343 | 103.9744 | 21.2612 | 0.004631 | 0.004094 | 227.0663 | 0.27064 | 8.404 | 0.003928 | 0.00579 | 0.005664 | 5e-05 | 0.001028 | 0.2959 | 1.3196 | 2.5481 | 98.85 | 20.76 | 38.723 | 133.7354 | 116.5116 | 120.5512 | 0.00961 | 0.89673 | 0.021121 | 23.5536 | 0.16358 | 0.16101 | 0.048229 | 3.0317 | 3.0317 | 0.002183
2019-03-21 15:58 | 1553198277.8717 | 487 | 0.19975 | 0.008482 | 0.000206 | 100.7078 | 98.7348 | 104.1028 | 21.2612 | 0.004608 | 0.004058 | 226.9486 | 0.27031 | 8.4033 | 0.003886 | 0.005897 | 0.005939 | 4.8e-05 | 0.001026 | 0.24746 | 1.3021 | 2.5479 | 98.85 | 20.76 | 38.722 | 133.1776 | 115.4818 | 119.6825 | 0.00961 | 0.89872 | 0.021158 | 23.5422 | 0.17697 | 0.16081 | 0.048178 | 3.0337 | 3.0337 | 0.002179
2019-03-21 15:58 | 1553198278.8723 | 488 | 0.19986 | 0.008434 | 0.000205 | 100.7189 | 98.6255 | 104.1068 | 21.2612 | 0.004578 | 0.004077 | 226.8135 | 0.2703 | 8.4041 | 0.003908 | 0.005352 | 0.0053 | 4.7e-05 | 0.00103 | 0.26753 | 1.3412 | 2.5479 | 98.85 | 20.76 | 38.722 | 132.4415 | 116.0088 | 120.1394 | 0.00961 | 0.89767 | 0.021146 | 23.5563 | 0.18786 | 0.1608 | 0.047611 | 2.9964 | 2.9964 | 0.002206
2019-03-21 15:58 | 1553198279.8725 | 489 | 0.19989 | 0.008381 | 0.000206 | 100.7293 | 98.9021 | 104.0575 | 21.2612 | 0.004545 | 0.004078 | 226.7046 | 0.2706 | 8.4029 | 0.003986 | 0.005271 | 0.005238 | 4.9e-05 | 0.001026 | 0.2286 | 1.3068 | 2.5478 | 98.85 | 20.76 | 38.721 | 131.6148 | 117.9227 | 120.1719 | 0.00961 | 0.8976 | 0.02114 | 23.5517 | 0.16374 | 0.16098 | 0.047473 | 2.9869 | 2.9869 | 0.002215
2019-03-21 15:58 | 1553198280.8724 | 490 | 0.20006 | 0.008334 | 0.000206 | 100.7695 | 98.947 | 104.1315 | 21.2612 | 0.004593 | 0.004089 | 226.5878 | 0.27122 | 8.4032 | 0.003865 | 0.00523 | 0.005103 | 4.8e-05 | 0.001025 | 0.26001 | 1.3473 | 2.548 | 98.85 | 20.76 | 38.72 | 132.7924 | 114.9701 | 120.4257 | 0.00961 | 0.89702 | 0.021141 | 23.5677 | 0.16325 | 0.16135 | 0.046537 | 2.9281 | 2.9281 | 0.002265
2019-03-21 15:58 | 1553198281.8716 | 491 | 0.20016 | 0.008278 | 0.000206 | 100.7138 | 98.9038 | 103.9423 | 21.2612 | 0.004661 | 0.004059 | 226.4695 | 0.27066 | 8.403 | 0.003892 | 0.005521 | 0.005515 | 4.9e-05 | 0.001027 | 0.29463 | 1.3168 | 2.5481 | 98.85 | 20.76 | 38.72 | 134.4679 | 115.6323 | 119.7079 | 0.00961 | 0.89866 | 0.021141 | 23.5248 | 0.16221 | 0.16101 | 0.046004 | 2.8947 | 2.8947 | 0.002286
2019-03-21 15:58 | 1553198282.8716 | 492 | 0.20021 | 0.008225 | 0.000207 | 100.7075 | 98.6156 | 104.0857 | 21.2612 | 0.004607 | 0.004048 | 226.3724 | 0.27167 | 8.3988 | 0.003891 | 0.005135 | 0.004951 | 4.8e-05 | 0.001024 | 0.29479 | 1.3071 | 2.5481 | 98.85 | 20.76 | 38.72 | 133.1392 | 115.5999 | 119.4383 | 0.00961 | 0.89928 | 0.021163 | 23.5329 | 0.18776 | 0.16162 | 0.045741 | 2.881 | 2.881 | 0.002306
2019-03-21 15:58 | 1553198283.8707 | 493 | 0.20027 | 0.008167 | 0.000207 | 100.7075 | 98.5754 | 104.1841 | 21.2612 | 0.00459 | 0.00409 | 226.2398 | 0.27083 | 8.4015 | 0.003909 | 0.005403 | 0.005317 | 5.1e-05 | 0.001025 | 0.26825 | 1.3184 | 2.5481 | 98.85 | 20.76 | 38.721 | 132.7205 | 116.0437 | 120.4685 | 0.00961 | 0.89692 | 0.021145 | 23.5749 | 0.1914 | 0.16112 | 0.045499 | 2.8634 | 2.8634 | 0.002313
2019-03-21 15:58 | 1553198284.8706 | 494 | 0.2003 | 0.008107 | 0.000207 | 100.7138 | 98.5884 | 104.139 | 21.2612 | 0.004586 | 0.004007 | 226.1424 | 0.27098 | 8.4016 | 0.003873 | 0.005078 | 0.005035 | 4.5e-05 | 0.001028 | 0.27168 | 1.3219 | 2.5479 | 98.85 | 20.76 | 38.721 | 132.6351 | 115.1593 | 118.4378 | 0.00961 | 0.90157 | 0.021195 | 23.5089 | 0.19078 | 0.1612 | 0.045369 | 2.862 | 2.862 | 0.002315
2019-03-21 15:58 | 1553198285.8718 | 495 | 0.20034 | 0.008045 | 0.000205 | 100.6831 | 98.5128 | 104.3678 | 21.2612 | 0.004564 | 0.004078 | 226.0748 | 0.27047 | 8.4035 | 0.003994 | 0.005155 | 0.005106 | 4.7e-05 | 0.001029 | 0.27433 | 1.3064 | 2.548 | 98.85 | 20.76 | 38.72 | 132.0788 | 118.12 | 120.1751 | 0.00961 | 0.89759 | 0.021171 | 23.5869 | 0.19492 | 0.1609 | 0.045195 | 2.848 | 2.848 | 0.002322
2019-03-21 15:58 | 1553198286.8708 | 496 | 0.20042 | 0.007987 | 0.000201 | 100.7826 | 98.7246 | 104.3582 | 21.2612 | 0.004568 | 0.004096 | 225.9525 | 0.27036 | 8.401 | 0.003964 | 0.005806 | 0.0057 | 4.7e-05 | 0.001029 | 0.27273 | 1.344 | 2.548 | 98.85 | 20.76 | 38.719 | 132.1987 | 117.3802 | 120.6162 | 0.009609 | 0.89659 | 0.021159 | 23.599 | 0.18454 | 0.16084 | 0.044826 | 2.8232 | 2.8232 | 0.002341
2019-03-21 15:58 | 1553198287.8723 | 497 | 0.20051 | 0.007924 | 0.000199 | 100.7516 | 98.4992 | 104.5237 | 21.2612 | 0.004601 | 0.00405 | 225.8624 | 0.27055 | 8.4022 | 0.003928 | 0.004525 | 0.004397 | 4.7e-05 | 0.001027 | 0.27887 | 1.3019 | 2.548 | 98.85 | 20.76 | 38.719 | 133.0096 | 116.5011 | 119.4794 | 0.009609 | 0.89918 | 0.021206 | 23.5836 | 0.20224 | 0.16095 | 0.044368 | 2.8008 | 2.8008 | 0.002362
2019-03-21 15:58 | 1553198288.8709 | 498 | 0.20053 | 0.007865 | 0.000197 | 100.7388 | 98.4317 | 104.7227 | 21.2612 | 0.004608 | 0.004019 | 225.7867 | 0.27091 | 8.4016 | 0.003936 | 0.005829 | 0.005774 | 4.7e-05 | 0.001028 | 0.30038 | 1.3165 | 2.548 | 98.85 | 20.76 | 38.718 | 133.184 | 116.6959 | 118.7175 | 0.009609 | 0.90093 | 0.021247 | 23.5831 | 0.20723 | 0.16117 | 0.044307 | 2.8025 | 2.8025 | 0.002364
2019-03-21 15:58 | 1553198289.872 | 499 | 0.20054 | 0.007808 | 0.000195 | 100.7656 | 98.7508 | 105.2485 | 21.2612 | 0.0046 | 0.003994 | 225.6973 | 0.27051 | 8.4033 | 0.003944 | 0.005588 | 0.005513 | 4.7e-05 | 0.001023 | 0.25752 | 1.3569 | 2.5481 | 98.85 | 20.76 | 38.717 | 132.985 | 116.8859 | 118.1161 | 0.009609 | 0.90231 | 0.021316 | 23.6241 | 0.18066 | 0.16093 | 0.044314 | 2.8122 | 2.8122 | 0.002352
2019-03-21 15:58 | 1553198290.8724 | 500 | 0.2006 | 0.007753 | 0.000194 | 100.8652 | 98.8036 | 105.6027 | 21.2612 | 0.004567 | 0.004023 | 225.5933 | 0.27107 | 8.4025 | 0.003972 | 0.005314 | 0.00522 | 4.6e-05 | 0.001019 | 0.26792 | 1.3551 | 2.5481 | 98.85 | 20.76 | 38.716 | 132.1763 | 117.5789 | 118.8294 | 0.009609 | 0.90067 | 0.021333 | 23.6854 | 0.18471 | 0.16126 | 0.044 | 2.7945 | 2.7945 | 0.002372
2019-03-21 15:58 | 1553198291.872 | 501 | 0.20067 | 0.007699 | 0.000194 | 100.7868 | 98.914 | 106.1804 | 21.2612 | 0.004482 | 0.003989 | 225.4691 | 0.27086 | 8.4023 | 0.00388 | 0.005536 | 0.005503 | 4.7e-05 | 0.001023 | 0.28263 | 1.3167 | 2.5481 | 98.85 | 20.76 | 38.714 | 130.0664 | 115.326 | 117.9972 | 0.009608 | 0.90259 | 0.021414 | 23.7249 | 0.16777 | 0.16113 | 0.043675 | 2.7844 | 2.7844 | 0.002378
2019-03-21 15:58 | 1553198292.8712 | 502 | 0.20073 | 0.007652 | 0.000195 | 100.8482 | 98.7468 | 106.3254 | 21.2612 | 0.00455 | 0.003966 | 225.3729 | 0.27138 | 8.4016 | 0.003913 | 0.005603 | 0.00555 | 4.6e-05 | 0.001019 | 0.26792 | 1.3729 | 2.5482 | 98.85 | 20.76 | 38.713 | 131.7396 | 116.1276 | 117.442 | 0.009608 | 0.90387 | 0.021444 | 23.7242 | 0.18835 | 0.16144 | 0.043344 | 2.7671 | 2.7671 | 0.002398
2019-03-21 15:58 | 1553198293.8715 | 503 | 0.2009 | 0.007606 | 0.000197 | 100.83 | 98.8649 | 106.8026 | 21.2612 | 0.00451 | 0.004014 | 225.2952 | 0.27127 | 8.4028 | 0.003955 | 0.005627 | 0.005511 | 4.4e-05 | 0.001023 | 0.2505 | 1.3077 | 2.5479 | 98.85 | 20.76 | 38.713 | 130.7711 | 117.1647 | 118.5924 | 0.009608 | 0.90122 | 0.02146 | 23.8124 | 0.17604 | 0.16138 | 0.042386 | 2.708 | 2.708 | 0.002449
2019-03-21 15:58 | 1553198294.8722 | 504 | 0.20104 | 0.007557 | 0.000197 | 100.8491 | 98.4814 | 107.3526 | 21.2612 | 0.004447 | 0.004006 | 225.2035 | 0.27057 | 8.4048 | 0.003959 | 0.005186 | 0.005163 | 4.5e-05 | 0.001021 | 0.24137 | 1.2854 | 2.5477 | 98.85 | 20.76 | 38.711 | 129.2106 | 117.255 | 118.407 | 0.009607 | 0.90165 | 0.02152 | 23.8679 | 0.21249 | 0.16096 | 0.041606 | 2.6656 | 2.6656 | 0.002482
2019-03-21 15:58 | 1553198295.8713 | 505 | 0.20122 | 0.007507 | 0.000196 | 100.802 | 98.5417 | 107.4136 | 21.2612 | 0.00446 | 0.003997 | 225.1055 | 0.27115 | 8.4026 | 0.003928 | 0.006169 | 0.006039 | 4.4e-05 | 0.001021 | 0.25409 | 1.3066 | 2.548 | 98.85 | 20.76 | 38.71 | 129.5287 | 116.4965 | 118.1952 | 0.009607 | 0.90213 | 0.021532 | 23.8683 | 0.20284 | 0.16131 | 0.040609 | 2.6032 | 2.6032 | 0.002547
2019-03-21 15:58 | 1553198296.8717 | 506 | 0.20144 | 0.007444 | 0.000196 | 100.858 | 98.6698 | 107.6413 | 21.2612 | 0.004457 | 0.003986 | 225.0103 | 0.27029 | 8.4026 | 0.003886 | 0.005114 | 0.005029 | 4.4e-05 | 0.001025 | 0.2672 | 1.3276 | 2.548 | 98.85 | 20.76 | 38.708 | 129.4594 | 115.4843 | 117.9181 | 0.009607 | 0.90277 | 0.021563 | 23.8851 | 0.1962 | 0.1608 | 0.03938 | 2.528 | 2.528 | 0.002614
2019-03-21 15:58 | 1553198297.8714 | 507 | 0.2015 | 0.007378 | 0.000196 | 100.9319 | 98.3486 | 107.922 | 21.2612 | 0.004465 | 0.00397 | 224.9064 | 0.27025 | 8.4019 | 0.00382 | 0.005956 | 0.005895 | 4.3e-05 | 0.001024 | 0.28655 | 1.3282 | 2.548 | 98.85 | 20.76 | 38.708 | 129.6627 | 113.8791 | 117.5362 | 0.009607 | 0.90365 | 0.021601 | 23.9045 | 0.23191 | 0.16077 | 0.039121 | 2.5159 | 2.5159 | 0.002626
2019-03-21 15:58 | 1553198298.8705 | 508 | 0.20157 | 0.007309 | 0.000195 | 100.9423 | 98.5308 | 107.9896 | 21.2612 | 0.00453 | 0.003941 | 224.7768 | 0.27035 | 8.3996 | 0.00379 | 0.005276 | 0.005277 | 4.6e-05 | 0.001022 | 0.28064 | 1.336 | 2.548 | 98.85 | 20.75 | 38.708 | 131.2445 | 113.1474 | 116.821 | 0.009601 | 0.90531 | 0.021628 | 23.8901 | 0.21628 | 0.16083 | 0.038756 | 2.4956 | 2.4956 | 0.002649
2019-03-21 15:58 | 1553198299.872 | 509 | 0.20162 | 0.007224 | 0.000194 | 100.9018 | 98.6112 | 108.242 | 21.2612 | 0.004459 | 0.003959 | 224.6758 | 0.27068 | 8.4034 | 0.003814 | 0.005244 | 0.005175 | 4.4e-05 | 0.001023 | 0.27328 | 1.3236 | 2.548 | 98.85 | 20.76 | 38.705 | 129.5088 | 113.7154 | 117.2599 | 0.009606 | 0.90429 | 0.021641 | 23.9315 | 0.20539 | 0.16103 | 0.038535 | 2.4829 | 2.4829 | 0.002666
2019-03-21 15:58 | 1553198300.8716 | 510 | 0.20166 | 0.007155 | 0.000194 | 100.9671 | 98.5778 | 108.4422 | 21.2612 | 0.004401 | 0.003957 | 224.5719 | 0.27079 | 8.4026 | 0.003751 | 0.00575 | 0.005642 | 4.6e-05 | 0.001021 | 0.27815 | 1.3285 | 2.548 | 98.85 | 20.75 | 38.705 | 128.0897 | 112.1899 | 117.2179 | 0.0096 | 0.90439 | 0.021662 | 23.9523 | 0.2142 | 0.16109 | 0.038423 | 2.4781 | 2.4781 | 0.002672
2019-03-21 15:58 | 1553198301.871 | 511 | 0.20176 | 0.007076 | 0.000193 | 100.9867 | 98.4695 | 108.4814 | 21.2612 | 0.004484 | 0.003899 | 224.4581 | 0.27034 | 8.4005 | 0.003835 | 0.005241 | 0.005175 | 4.6e-05 | 0.001023 | 0.28191 | 1.3091 | 2.548 | 98.85 | 20.75 | 38.705 | 130.1191 | 114.245 | 115.7937 | 0.0096 | 0.9077 | 0.021706 | 23.9129 | 0.22578 | 0.16082 | 0.037895 | 2.449 | 2.449 | 0.002699
2019-03-21 15:58 | 1553198302.8721 | 512 | 0.20191 | 0.00701 | 0.000192 | 101.0174 | 98.4453 | 108.6255 | 21.2612 | 0.004439 | 0.003907 | 224.3527 | 0.26989 | 8.4032 | 0.003844 | 0.005702 | 0.005658 | 4.4e-05 | 0.00102 | 0.25752 | 1.3168 | 2.548 | 98.85 | 20.75 | 38.704 | 129.0106 | 114.457 | 115.993 | 0.0096 | 0.90724 | 0.021715 | 23.9349 | 0.23069 | 0.16056 | 0.037064 | 2.3963 | 2.3963 | 0.002754
2019-03-21 15:58 | 1553198303.8708 | 513 | 0.20199 | 0.006936 | 0.000192 | 101.0591 | 98.6872 | 108.7213 | 21.2612 | 0.004389 | 0.003882 | 224.2404 | 0.27035 | 8.4006 | 0.003779 | 0.0052 | 0.005118 | 4.1e-05 | 0.001021 | 0.28313 | 1.3039 | 2.5479 | 98.85 | 20.75 | 38.703 | 127.7783 | 112.8691 | 115.3838 | 0.009599 | 0.90866 | 0.021741 | 23.9267 | 0.21244 | 0.16083 | 0.036716 | 2.3767 | 2.3767 | 0.002781
2019-03-21 15:58 | 1553198304.8718 | 514 | 0.20201 | 0.006866 | 0.000192 | 101.0642 | 98.8094 | 108.7583 | 21.2612 | 0.004336 | 0.003896 | 224.1422 | 0.27086 | 8.4012 | 0.00387 | 0.005297 | 0.005238 | 4.6e-05 | 0.00102 | 0.26609 | 1.3099 | 2.5479 | 98.85 | 20.75 | 38.702 | 126.4766 | 115.0871 | 115.725 | 0.009599 | 0.90786 | 0.021735 | 23.9413 | 0.20181 | 0.16113 | 0.036652 | 2.372 | 2.372 | 0.002792
2019-03-21 15:58 | 1553198305.8718 | 515 | 0.20224 | 0.006807 | 0.000192 | 100.9566 | 98.8861 | 108.8813 | 21.2612 | 0.004386 | 0.00388 | 224.0767 | 0.27059 | 8.4026 | 0.003861 | 0.00565 | 0.005592 | 4.7e-05 | 0.001022 | 0.29142 | 1.3047 | 2.548 | 98.85 | 20.75 | 38.702 | 127.7159 | 114.8704 | 115.3394 | 0.009599 | 0.90876 | 0.021759 | 23.9429 | 0.18536 | 0.16097 | 0.035381 | 2.2921 | 2.2921 | 0.002886
2019-03-21 15:58 | 1553198306.8719 | 516 | 0.2022 | 0.006729 | 0.000192 | 101.0123 | 98.7267 | 109.0546 | 21.2612 | 0.00441 | 0.0039 | 223.9907 | 0.27119 | 8.4005 | 0.003804 | 0.006121 | 0.005981 | 4.7e-05 | 0.001021 | 0.28191 | 1.3363 | 2.5481 | 98.85 | 20.75 | 38.702 | 128.2985 | 113.4681 | 115.811 | 0.009599 | 0.90766 | 0.021763 | 23.9765 | 0.20472 | 0.16133 | 0.035684 | 2.3121 | 2.3121 | 0.002868
2019-03-21 15:58 | 1553198307.8719 | 517 | 0.20227 | 0.00666 | 0.000191 | 101.0392 | 98.4184 | 109.0579 | 21.2612 | 0.004351 | 0.003832 | 223.918 | 0.27038 | 8.402 | 0.003798 | 0.004801 | 0.004732 | 4.2e-05 | 0.001022 | 0.28407 | 1.2638 | 2.548 | 98.85 | 20.75 | 38.701 | 126.8509 | 113.333 | 114.1597 | 0.009599 | 0.91153 | 0.021809 | 23.9259 | 0.23506 | 0.16085 | 0.035362 | 2.2962 | 2.2962 | 0.002879
2019-03-21 15:58 | 1553198308.8714 | 518 | 0.20234 | 0.006587 | 0.00019 | 101.102 | 98.2267 | 108.9215 | 21.2612 | 0.004398 | 0.00389 | 223.8528 | 0.27036 | 8.4025 | 0.003823 | 0.004851 | 0.004802 | 4.5e-05 | 0.001022 | 0.26432 | 1.2795 | 2.5479 | 98.85 | 20.75 | 38.7 | 128.0155 | 113.9526 | 115.5752 | 0.009599 | 0.90821 | 0.021756 | 23.9546 | 0.25807 | 0.16084 | 0.035052 | 2.2705 | 2.2705 | 0.002911
2019-03-21 15:58 | 1553198309.871 | 519 | 0.20238 | 0.006521 | 0.000188 | 101.0198 | 98.3898 | 108.7058 | 21.2612 | 0.004339 | 0.003859 | 223.7596 | 0.27035 | 8.4041 | 0.00379 | 0.005191 | 0.00515 | 4.9e-05 | 0.001025 | 0.25929 | 1.3214 | 2.5479 | 98.85 | 20.75 | 38.697 | 126.5508 | 113.1433 | 114.8104 | 0.009598 | 0.91 | 0.021756 | 23.9073 | 0.23595 | 0.16083 | 0.0349 | 2.2608 | 2.2608 | 0.002924
2019-03-21 15:58 | 1553198310.8721 | 520 | 0.20247 | 0.00646 | 0.000187 | 101.0469 | 98.3169 | 108.2753 | 21.2612 | 0.004362 | 0.003881 | 223.689 | 0.26991 | 8.3997 | 0.003712 | 0.005082 | 0.005046 | 4.5e-05 | 0.001024 | 0.27649 | 1.3114 | 2.5479 | 98.85 | 20.75 | 38.696 | 127.117 | 111.2395 | 115.3506 | 0.009598 | 0.90874 | 0.021698 | 23.8766 | 0.24498 | 0.16057 | 0.034386 | 2.2216 | 2.2216 | 0.002971
2019-03-21 15:59 | 1553198311.8715 | 521 | 0.2026 | 0.006392 | 0.000186 | 101.1232 | 98.0959 | 108.1368 | 21.2612 | 0.004309 | 0.003862 | 223.6544 | 0.27001 | 8.4034 | 0.00391 | 0.004988 | 0.004904 | 4.6e-05 | 0.001027 | 0.27168 | 1.3298 | 2.5481 | 98.85 | 20.75 | 38.695 | 125.8326 | 116.0747 | 114.8911 | 0.009597 | 0.90981 | 0.021697 | 23.8472 | 0.27186 | 0.16063 | 0.033708 | 2.1777 | 2.1777 | 0.003032
2019-03-21 15:59 | 1553198312.8726 | 522 | 0.20273 | 0.006328 | 0.000183 | 101.0794 | 98.2383 | 108.1957 | 21.2612 | 0.004284 | 0.003814 | 223.5977 | 0.27005 | 8.4003 | 0.003854 | 0.005483 | 0.005463 | 4.2e-05 | 0.001027 | 0.27201 | 1.3066 | 2.5481 | 98.85 | 20.75 | 38.694 | 125.2088 | 114.691 | 113.7139 | 0.009597 | 0.91258 | 0.021735 | 23.8175 | 0.25501 | 0.16065 | 0.033019 | 2.137 | 2.137 | 0.00309
2019-03-21 15:59 | 1553198313.8722 | 523 | 0.20278 | 0.006261 | 0.00018 | 101.0839 | 98.635 | 107.9352 | 21.2612 | 0.004337 | 0.003825 | 223.5262 | 0.27049 | 8.4028 | 0.00388 | 0.005298 | 0.005268 | 4.5e-05 | 0.001025 | 0.26593 | 1.3517 | 2.5481 | 98.85 | 20.75 | 38.693 | 126.5068 | 115.3419 | 113.9984 | 0.009597 | 0.91191 | 0.021701 | 23.7975 | 0.21935 | 0.16092 | 0.032812 | 2.1205 | 2.1205 | 0.003119
2019-03-21 15:59 | 1553198314.8713 | 524 | 0.20288 | 0.006194 | 0.000178 | 101.1032 | 98.6347 | 107.8986 | 21.2612 | 0.004326 | 0.003871 | 223.4593 | 0.27105 | 8.4022 | 0.003816 | 0.005694 | 0.005477 | 4.6e-05 | 0.001029 | 0.30215 | 1.3583 | 2.5479 | 98.85 | 20.75 | 38.694 | 126.2401 | 113.7668 | 115.1227 | 0.009597 | 0.90927 | 0.021666 | 23.828 | 0.2211 | 0.16125 | 0.032317 | 2.0852 | 2.0852 | 0.003178
2019-03-21 15:59 | 1553198315.8707 | 525 | 0.20303 | 0.006124 | 0.000177 | 101.1712 | 98.7897 | 107.8237 | 21.2612 | 0.004269 | 0.00383 | 223.3707 | 0.27147 | 8.4009 | 0.003768 | 0.004803 | 0.004775 | 4.4e-05 | 0.001027 | 0.2907 | 1.367 | 2.548 | 98.85 | 20.75 | 38.693 | 124.831 | 112.5935 | 114.1027 | 0.009597 | 0.91167 | 0.021687 | 23.7884 | 0.21306 | 0.1615 | 0.031448 | 2.0311 | 2.0311 | 0.003268
2019-03-21 15:59 | 1553198316.8704 | 526 | 0.20318 | 0.006051 | 0.000175 | 101.1437 | 98.8196 | 107.7817 | 21.2612 | 0.004295 | 0.00379 | 223.3054 | 0.27109 | 8.4029 | 0.003821 | 0.004764 | 0.004613 | 4.2e-05 | 0.00103 | 0.27632 | 1.281 | 2.5481 | 98.85 | 20.75 | 38.693 | 125.4847 | 113.9051 | 113.1306 | 0.009597 | 0.91396 | 0.02171 | 23.7539 | 0.20793 | 0.16127 | 0.030659 | 1.9823 | 1.9823 | 0.003344
2019-03-21 15:59 | 1553198317.8715 | 527 | 0.20333 | 0.005988 | 0.000174 | 101.1288 | 98.7212 | 107.5833 | 21.2612 | 0.004297 | 0.003788 | 223.211 | 0.27049 | 8.405 | 0.00373 | 0.005405 | 0.005364 | 4.6e-05 | 0.001029 | 0.24331 | 1.2874 | 2.5479 | 98.85 | 20.75 | 38.693 | 125.5365 | 111.6721 | 113.0824 | 0.009597 | 0.91407 | 0.021692 | 23.7306 | 0.21552 | 0.16091 | 0.029852 | 1.9283 | 1.9283 | 0.00343
2019-03-21 15:59 | 1553198318.8716 | 528 | 0.20324 | 0.005911 | 0.000173 | 101.1425 | 98.6476 | 107.6507 | 21.2612 | 0.004222 | 0.003765 | 223.1395 | 0.27047 | 8.4057 | 0.003816 | 0.005394 | 0.005288 | 4.7e-05 | 0.001028 | 0.30525 | 1.309 | 2.5481 | 98.85 | 20.75 | 38.693 | 123.6984 | 113.778 | 112.5304 | 0.009597 | 0.91538 | 0.021714 | 23.721 | 0.2234 | 0.1609 | 0.030471 | 1.9705 | 1.9705 | 0.003356
2019-03-21 15:59 | 1553198319.8709 | 529 | 0.2034 | 0.005844 | 0.00017 | 101.2093 | 98.6868 | 107.6962 | 21.2612 | 0.004212 | 0.00371 | 223.09 | 0.26999 | 8.4031 | 0.003765 | 0.005072 | 0.005041 | 4.5e-05 | 0.001028 | 0.28567 | 1.3143 | 2.548 | 98.85 | 20.75 | 38.694 | 123.4576 | 112.5382 | 111.1836 | 0.009597 | 0.91859 | 0.021756 | 23.6846 | 0.22575 | 0.16062 | 0.029623 | 1.9195 | 1.9195 | 0.003439
2019-03-21 15:59 | 1553198320.8716 | 530 | 0.2035 | 0.005777 | 0.000168 | 101.1714 | 98.7052 | 108.0037 | 21.2612 | 0.004169 | 0.003735 | 223.0084 | 0.2706 | 8.4037 | 0.003736 | 0.005585 | 0.005469 | 4.4e-05 | 0.001028 | 0.26416 | 1.3006 | 2.548 | 98.85 | 20.75 | 38.695 | 122.3962 | 111.8276 | 111.7896 | 0.009597 | 0.91714 | 0.02177 | 23.7371 | 0.22074 | 0.16098 | 0.029096 | 1.8867 | 1.8867 | 0.003507
2019-03-21 15:59 | 1553198321.8716 | 531 | 0.20358 | 0.005708 | 0.000165 | 101.1622 | 98.7846 | 108.1783 | 21.2612 | 0.004182 | 0.003705 | 222.9293 | 0.27053 | 8.404 | 0.003723 | 0.005427 | 0.005345 | 4.9e-05 | 0.001025 | 0.29336 | 1.3336 | 2.548 | 98.85 | 20.75 | 38.696 | 122.714 | 111.5173 | 111.0778 | 0.009598 | 0.91884 | 0.021808 | 23.7343 | 0.21274 | 0.16094 | 0.028674 | 1.8626 | 1.8626 | 0.003551
2019-03-21 15:59 | 1553198322.8718 | 532 | 0.2037 | 0.005643 | 0.000163 | 101.2444 | 98.9119 | 108.133 | 21.2612 | 0.004217 | 0.003736 | 222.8468 | 0.26981 | 8.4011 | 0.003709 | 0.004999 | 0.0049 | 4.3e-05 | 0.001026 | 0.29319 | 1.3551 | 2.548 | 98.85 | 20.75 | 38.699 | 123.5627 | 111.1622 | 111.8332 | 0.009598 | 0.91704 | 0.021782 | 23.7526 | 0.20848 | 0.16051 | 0.028085 | 1.8222 | 1.8222 | 0.00362
2019-03-21 15:59 | 1553198323.871 | 533 | 0.2038 | 0.005572 | 0.000161 | 101.1467 | 99.1177 | 108.368 | 21.2612 | 0.004215 | 0.003673 | 222.7632 | 0.2707 | 8.4028 | 0.003719 | 0.005796 | 0.005739 | 4.4e-05 | 0.001027 | 0.29231 | 1.362 | 2.548 | 98.85 | 20.74 | 38.7 | 123.5259 | 111.4195 | 110.284 | 0.009593 | 0.92075 | 0.02185 | 23.7305 | 0.18126 | 0.16104 | 0.027579 | 1.795 | 1.795 | 0.003687
2019-03-21 15:59 | 1553198324.8718 | 534 | 0.20385 | 0.0055 | 0.000158 | 101.2105 | 99.2209 | 108.6417 | 21.2612 | 0.00414 | 0.003719 | 222.6991 | 0.2712 | 8.402 | 0.003732 | 0.005489 | 0.005369 | 4.6e-05 | 0.001027 | 0.28854 | 1.3536 | 2.548 | 98.85 | 20.74 | 38.703 | 121.6822 | 111.7256 | 111.4052 | 0.009594 | 0.91806 | 0.021845 | 23.7952 | 0.17758 | 0.16134 | 0.027361 | 1.7805 | 1.7805 | 0.003724
2019-03-21 15:59 | 1553198325.8709 | 535 | 0.20391 | 0.005426 | 0.000155 | 101.2275 | 99.0087 | 108.8323 | 21.2612 | 0.004155 | 0.003653 | 222.6029 | 0.2706 | 8.4008 | 0.003743 | 0.005599 | 0.005505 | 4.4e-05 | 0.001032 | 0.27815 | 1.3143 | 2.5479 | 98.85 | 20.74 | 38.705 | 122.0554 | 111.9831 | 109.7991 | 0.009594 | 0.92191 | 0.02191 | 23.7663 | 0.19823 | 0.16098 | 0.02709 | 1.7683 | 1.7683 | 0.003742
2019-03-21 15:59 | 1553198326.8723 | 536 | 0.20404 | 0.005353 | 0.000152 | 101.178 | 98.9651 | 108.9222 | 21.2612 | 0.004126 | 0.003673 | 222.5349 | 0.27066 | 8.4017 | 0.003659 | 0.005307 | 0.005208 | 4.5e-05 | 0.001033 | 0.27184 | 1.3806 | 2.5481 | 98.85 | 20.74 | 38.709 | 121.3313 | 109.9471 | 110.2918 | 0.009595 | 0.92073 | 0.021905 | 23.7914 | 0.1978 | 0.16102 | 0.026388 | 1.7221 | 1.7221 | 0.003843
2019-03-21 15:59 | 1553198327.871 | 537 | 0.20408 | 0.005283 | 0.000149 | 101.1944 | 99.1027 | 109.0475 | 21.2612 | 0.004118 | 0.003627 | 222.47 | 0.27081 | 8.4009 | 0.003635 | 0.005544 | 0.00551 | 4.3e-05 | 0.001028 | 0.29015 | 1.3565 | 2.548 | 98.85 | 20.74 | 38.711 | 121.136 | 109.3755 | 109.1777 | 0.009596 | 0.92341 | 0.02195 | 23.7704 | 0.18682 | 0.16111 | 0.026217 | 1.7145 | 1.7145 | 0.003862
2019-03-21 15:59 | 1553198328.8711 | 538 | 0.20425 | 0.005214 | 0.000147 | 101.2167 | 99.0175 | 109.186 | 21.2612 | 0.004106 | 0.003687 | 222.4052 | 0.27067 | 8.4025 | 0.003701 | 0.005191 | 0.005083 | 4.2e-05 | 0.001028 | 0.27959 | 1.3205 | 2.5479 | 98.85 | 20.74 | 38.714 | 120.84 | 110.961 | 110.6342 | 0.009596 | 0.91991 | 0.021922 | 23.8308 | 0.19649 | 0.16102 | 0.025303 | 1.6527 | 1.6527 | 0.004004
2019-03-21 15:59 | 1553198329.8717 | 539 | 0.20438 | 0.005142 | 0.000144 | 101.1747 | 99.2127 | 109.4191 | 21.2612 | 0.004104 | 0.00364 | 222.3393 | 0.27074 | 8.4039 | 0.003611 | 0.005258 | 0.005132 | 4.4e-05 | 0.00103 | 0.27383 | 1.3313 | 2.5479 | 98.85 | 20.74 | 38.719 | 120.8043 | 108.7764 | 109.4868 | 0.009598 | 0.92266 | 0.021978 | 23.8205 | 0.17516 | 0.16106 | 0.024648 | 1.6141 | 1.6141 | 0.004101
2019-03-21 15:59 | 1553198330.8706 | 540 | 0.20451 | 0.005067 | 0.000143 | 101.2984 | 99.5328 | 109.4158 | 21.2612 | 0.004115 | 0.003605 | 222.2671 | 0.27068 | 8.3972 | 0.003711 | 0.005045 | 0.004983 | 4.2e-05 | 0.001029 | 0.28711 | 1.3495 | 2.548 | 98.85 | 20.74 | 38.722 | 121.0717 | 111.2181 | 108.6471 | 0.009598 | 0.92469 | 0.022002 | 23.794 | 0.15727 | 0.16103 | 0.023949 | 1.57 | 1.57 | 0.004215
2019-03-21 15:59 | 1553198331.8702 | 541 | 0.20459 | 0.004989 | 0.000143 | 101.1834 | 99.367 | 109.3714 | 21.2612 | 0.004032 | 0.003595 | 222.1949 | 0.27052 | 8.402 | 0.0036 | 0.00502 | 0.004974 | 4.4e-05 | 0.001029 | 0.25265 | 1.3099 | 2.5481 | 98.85 | 20.74 | 38.726 | 119.0498 | 108.5126 | 108.3863 | 0.009599 | 0.92533 | 0.022005 | 23.7811 | 0.16203 | 0.16093 | 0.023566 | 1.5451 | 1.5451 | 0.004281
2019-03-21 15:59 | 1553198332.8713 | 542 | 0.20461 | 0.004911 | 0.000142 | 101.2689 | 99.4232 | 109.7505 | 21.2612 | 0.00403 | 0.003606 | 222.1514 | 0.27064 | 8.4043 | 0.003641 | 0.005505 | 0.005418 | 4.5e-05 | 0.001027 | 0.28102 | 1.3057 | 2.548 | 98.85 | 20.74 | 38.728 | 119.0006 | 109.5057 | 108.6654 | 0.0096 | 0.92465 | 0.022035 | 23.8309 | 0.16453 | 0.161 | 0.023534 | 1.5451 | 1.5451 | 0.004283
2019-03-21 15:59 | 1553198333.8707 | 543 | 0.20463 | 0.004837 | 0.000141 | 101.1944 | 99.5751 | 109.8003 | 21.2612 | 0.004022 | 0.003612 | 222.1166 | 0.27145 | 8.4022 | 0.003619 | 0.005047 | 0.00499 | 4.2e-05 | 0.001023 | 0.25155 | 1.2885 | 2.5479 | 98.85 | 20.74 | 38.731 | 118.8025 | 108.9884 | 108.8009 | 0.0096 | 0.92432 | 0.022036 | 23.8406 | 0.14429 | 0.16149 | 0.023509 | 1.5436 | 1.5436 | 0.0043
2019-03-21 15:59 | 1553198334.8704 | 544 | 0.20471 | 0.004768 | 0.00014 | 101.175 | 99.4695 | 109.6691 | 21.2612 | 0.004088 | 0.003594 | 222.0659 | 0.27124 | 8.405 | 0.003665 | 0.005122 | 0.005035 | 4.2e-05 | 0.001022 | 0.25033 | 1.2974 | 2.5479 | 98.85 | 20.74 | 38.737 | 120.4218 | 110.1017 | 108.3809 | 0.009602 | 0.92534 | 0.022035 | 23.8132 | 0.15207 | 0.16136 | 0.023129 | 1.5185 | 1.5185 | 0.004368
2019-03-21 15:59 | 1553198335.8715 | 545 | 0.20483 | 0.004692 | 0.000139 | 101.2239 | 99.8888 | 109.9076 | 21.2612 | 0.003993 | 0.003558 | 222.0259 | 0.27081 | 8.4002 | 0.00358 | 0.004749 | 0.004668 | 4.3e-05 | 0.001025 | 0.27312 | 1.3179 | 2.548 | 98.85 | 20.74 | 38.738 | 118.0868 | 108.0436 | 107.4947 | 0.009602 | 0.92749 | 0.022085 | 23.8114 | 0.11876 | 0.16111 | 0.022478 | 1.479 | 1.479 | 0.004477
2019-03-21 15:59 | 1553198336.8709 | 546 | 0.20494 | 0.004623 | 0.000138 | 101.2978 | 100.0093 | 109.9653 | 21.2612 | 0.003938 | 0.00356 | 221.9731 | 0.27036 | 8.4005 | 0.003578 | 0.005016 | 0.004891 | 4.4e-05 | 0.001025 | 0.26322 | 1.2922 | 2.548 | 98.85 | 20.74 | 38.742 | 116.7483 | 107.972 | 107.5396 | 0.009603 | 0.92738 | 0.022089 | 23.8191 | 0.1145 | 0.16084 | 0.021933 | 1.4434 | 1.4434 | 0.00458
2019-03-21 15:59 | 1553198337.8709 | 547 | 0.20502 | 0.004554 | 0.000138 | 101.3067 | 100.5607 | 109.9749 | 21.2612 | 0.003983 | 0.003592 | 221.9294 | 0.27081 | 8.4033 | 0.003582 | 0.004765 | 0.004725 | 4.3e-05 | 0.001028 | 0.27616 | 1.3376 | 2.548 | 98.85 | 20.73 | 38.745 | 117.8475 | 108.0839 | 108.3139 | 0.009598 | 0.9255 | 0.022068 | 23.8443 | 0.066108 | 0.16111 | 0.021536 | 1.4159 | 1.4159 | 0.004677
2019-03-21 15:59 | 1553198338.8706 | 548 | 0.20508 | 0.004484 | 0.000138 | 101.2206 | 100.3686 | 110.2502 | 21.2612 | 0.003991 | 0.003546 | 221.8775 | 0.26954 | 8.4019 | 0.003571 | 0.004756 | 0.004687 | 4.3e-05 | 0.001031 | 0.26001 | 1.3182 | 2.548 | 98.85 | 20.73 | 38.746 | 118.0328 | 107.8159 | 107.2189 | 0.009598 | 0.92817 | 0.022127 | 23.8399 | 0.075604 | 0.16035 | 0.02125 | 1.4008 | 1.4008 | 0.004705
2019-03-21 15:59 | 1553198339.8718 | 549 | 0.20516 | 0.004418 | 0.000138 | 101.2513 | 100.1789 | 110.0442 | 21.2612 | 0.003923 | 0.003516 | 221.8285 | 0.26984 | 8.4026 | 0.003535 | 0.004995 | 0.004917 | 4.4e-05 | 0.001026 | 0.29264 | 1.3173 | 2.5478 | 98.85 | 20.73 | 38.749 | 116.3826 | 106.9474 | 106.4805 | 0.009599 | 0.92997 | 0.022128 | 23.7944 | 0.095238 | 0.16053 | 0.020831 | 1.3731 | 1.3731 | 0.004805
2019-03-21 15:59 | 1553198340.8711 | 550 | 0.20526 | 0.004353 | 0.00014 | 101.2844 | 99.9739 | 110.3514 | 21.2612 | 0.003972 | 0.003562 | 221.8241 | 0.27057 | 8.4055 | 0.003562 | 0.004817 | 0.004781 | 4.4e-05 | 0.001026 | 0.26034 | 1.3512 | 2.5478 | 98.85 | 20.73 | 38.751 | 117.5743 | 107.5884 | 107.5835 | 0.0096 | 0.92728 | 0.022127 | 23.8622 | 0.11648 | 0.16096 | 0.020309 | 1.3384 | 1.3384 | 0.004943
2019-03-21 15:59 | 1553198341.8707 | 551 | 0.20539 | 0.004293 | 0.000141 | 101.2239 | 99.7368 | 110.7103 | 21.2612 | 0.003924 | 0.003521 | 221.7759 | 0.27004 | 8.4023 | 0.003506 | 0.004769 | 0.004627 | 4.2e-05 | 0.001027 | 0.29695 | 1.3538 | 2.548 | 98.85 | 20.73 | 38.752 | 116.3992 | 106.2452 | 106.6072 | 0.0096 | 0.92966 | 0.022191 | 23.8703 | 0.13237 | 0.16065 | 0.019597 | 1.295 | 1.295 | 0.005099
2019-03-21 15:59 | 1553198342.8717 | 552 | 0.20545 | 0.004231 | 0.000143 | 101.2379 | 99.5284 | 110.7706 | 21.2612 | 0.003965 | 0.003505 | 221.7181 | 0.27146 | 8.3994 | 0.003559 | 0.00502 | 0.004974 | 4.4e-05 | 0.001027 | 0.2677 | 1.3359 | 2.548 | 98.85 | 20.73 | 38.756 | 117.4158 | 107.5333 | 106.2066 | 0.009601 | 0.93064 | 0.022209 | 23.8642 | 0.15232 | 0.16149 | 0.019309 | 1.2768 | 1.2768 | 0.005198
2019-03-21 15:59 | 1553198343.8722 | 553 | 0.20558 | 0.004163 | 0.000145 | 101.2695 | 99.8894 | 110.9997 | 21.2612 | 0.003934 | 0.0035 | 221.6753 | 0.27078 | 8.4033 | 0.003565 | 0.005075 | 0.005091 | 4.5e-05 | 0.001026 | 0.27328 | 1.3173 | 2.548 | 98.85 | 20.73 | 38.76 | 116.6531 | 107.6593 | 106.1018 | 0.009602 | 0.9309 | 0.022235 | 23.8856 | 0.12273 | 0.16109 | 0.018644 | 1.234 | 1.234 | 0.005365
2019-03-21 15:59 | 1553198344.8709 | 554 | 0.20563 | 0.004105 | 0.000149 | 101.2364 | 100.1674 | 111.2199 | 21.2612 | 0.003861 | 0.00349 | 221.6432 | 0.27131 | 8.4016 | 0.003569 | 0.004805 | 0.004666 | 4.6e-05 | 0.00103 | 0.27903 | 1.3114 | 2.5481 | 98.85 | 20.73 | 38.764 | 114.8611 | 107.7617 | 105.8578 | 0.009603 | 0.9315 | 0.022264 | 23.9016 | 0.094956 | 0.1614 | 0.018405 | 1.2194 | 1.2194 | 0.00544
2019-03-21 15:59 | 1553198345.8713 | 555 | 0.20562 | 0.004034 | 0.000153 | 101.257 | 100.3874 | 111.2976 | 21.2612 | 0.003893 | 0.00346 | 221.5903 | 0.27049 | 8.403 | 0.003553 | 0.005156 | 0.005027 | 4.1e-05 | 0.001024 | 0.26322 | 1.296 | 2.548 | 98.85 | 20.73 | 38.766 | 115.6464 | 107.3867 | 105.1216 | 0.009603 | 0.93331 | 0.022294 | 23.8867 | 0.077147 | 0.16091 | 0.018538 | 1.2297 | 1.2297 | 0.005378
2019-03-21 15:59 | 1553198346.8701 | 556 | 0.20569 | 0.003966 | 0.000157 | 101.2611 | 100.2263 | 111.1042 | 21.2612 | 0.003905 | 0.003449 | 221.5535 | 0.2709 | 8.3998 | 0.003548 | 0.005249 | 0.005206 | 4.5e-05 | 0.00103 | 0.27184 | 1.2715 | 2.5479 | 98.85 | 20.73 | 38.767 | 115.9471 | 107.2466 | 104.8511 | 0.009603 | 0.93398 | 0.022282 | 23.8574 | 0.091881 | 0.16116 | 0.018184 | 1.2052 | 1.2052 | 0.005496
2019-03-21 15:59 | 1553198347.871 | 557 | 0.20569 | 0.003896 | 0.000161 | 101.2212 | 100.1227 | 111.4457 | 21.2612 | 0.003885 | 0.003493 | 221.5082 | 0.27081 | 8.4006 | 0.003495 | 0.00501 | 0.004964 | 4.4e-05 | 0.001032 | 0.28008 | 1.2626 | 2.548 | 98.85 | 20.72 | 38.772 | 115.4423 | 105.9808 | 105.9163 | 0.009599 | 0.93136 | 0.022285 | 23.9277 | 0.097598 | 0.1611 | 0.018247 | 1.2094 | 1.2094 | 0.005475
2019-03-21 15:59 | 1553198348.8712 | 558 | 0.20579 | 0.003834 | 0.000165 | 101.265 | 99.5223 | 111.5804 | 21.2612 | 0.003855 | 0.003478 | 221.4798 | 0.27069 | 8.4031 | 0.003483 | 0.004651 | 0.004547 | 4.3e-05 | 0.001032 | 0.27865 | 1.2908 | 2.5479 | 98.85 | 20.72 | 38.773 | 114.7171 | 105.6832 | 105.5496 | 0.009599 | 0.93226 | 0.022309 | 23.9306 | 0.15527 | 0.16103 | 0.017759 | 1.1779 | 1.1779 | 0.005619
2019-03-21 15:59 | 1553198349.8712 | 559 | 0.20587 | 0.003775 | 0.00017 | 101.2474 | 99.5294 | 111.909 | 21.2612 | 0.003872 | 0.003421 | 221.4536 | 0.27074 | 8.403 | 0.003504 | 0.00524 | 0.005161 | 4.3e-05 | 0.001034 | 0.28119 | 1.3241 | 2.5479 | 98.85 | 20.72 | 38.777 | 115.1276 | 106.1821 | 104.1771 | 0.0096 | 0.93565 | 0.022383 | 23.9223 | 0.15308 | 0.16106 | 0.01732 | 1.1521 | 1.1521 | 0.005745
2019-03-21 15:59 | 1553198350.8712 | 560 | 0.20598 | 0.003713 | 0.000175 | 101.2882 | 99.5396 | 112.0271 | 21.2612 | 0.003838 | 0.003418 | 221.4372 | 0.27102 | 8.4018 | 0.00348 | 0.004629 | 0.004529 | 4.2e-05 | 0.001038 | 0.26626 | 1.3063 | 2.5479 | 98.85 | 20.72 | 38.779 | 114.3052 | 105.6061 | 104.093 | 0.009601 | 0.93586 | 0.022397 | 23.9322 | 0.15576 | 0.16123 | 0.016724 | 1.1127 | 1.1127 | 0.005955
2019-03-21 15:59 | 1553198351.8703 | 561 | 0.20606 | 0.003657 | 0.00018 | 101.3198 | 100.0611 | 112.2367 | 21.2612 | 0.003807 | 0.003429 | 221.4127 | 0.2719 | 8.4001 | 0.003474 | 0.004965 | 0.004886 | 4e-05 | 0.001035 | 0.29048 | 1.3095 | 2.5479 | 98.85 | 20.72 | 38.781 | 113.5614 | 105.4543 | 104.3666 | 0.009601 | 0.93518 | 0.02241 | 23.9633 | 0.11182 | 0.16175 | 0.016317 | 1.0857 | 1.0857 | 0.006123
2019-03-21 15:59 | 1553198352.8716 | 562 | 0.20614 | 0.003604 | 0.000186 | 101.3183 | 100.1401 | 112.3043 | 21.2612 | 0.003789 | 0.003375 | 221.4129 | 0.27158 | 8.4013 | 0.003543 | 0.00546 | 0.005373 | 4.3e-05 | 0.00103 | 0.286 | 1.3365 | 2.548 | 98.85 | 20.72 | 38.785 | 113.1204 | 107.1372 | 103.068 | 0.009602 | 0.93841 | 0.022455 | 23.9293 | 0.10463 | 0.16156 | 0.015902 | 1.0597 | 1.0597 | 0.006266
2019-03-21 15:59 | 1553198353.8704 | 563 | 0.20627 | 0.003539 | 0.000192 | 101.2728 | 100.2137 | 112.4463 | 21.2612 | 0.003809 | 0.003432 | 221.3786 | 0.27137 | 8.4024 | 0.003477 | 0.005372 | 0.005275 | 4.8e-05 | 0.001031 | 0.28926 | 1.3439 | 2.548 | 98.85 | 20.72 | 38.788 | 113.597 | 105.5433 | 104.4514 | 0.009603 | 0.93497 | 0.022428 | 23.9884 | 0.094032 | 0.16144 | 0.015187 | 1.0101 | 1.0101 | 0.006569
2019-03-21 15:59 | 1553198354.8704 | 564 | 0.20637 | 0.003475 | 0.000195 | 101.2129 | 100.4633 | 112.5644 | 21.2612 | 0.003791 | 0.003379 | 221.3375 | 0.27037 | 8.4024 | 0.003486 | 0.00418 | 0.004189 | 4.5e-05 | 0.001033 | 0.2891 | 1.3353 | 2.548 | 98.85 | 20.72 | 38.791 | 113.163 | 105.7502 | 103.1521 | 0.009603 | 0.9382 | 0.022479 | 23.9596 | 0.066486 | 0.16084 | 0.01465 | 0.97606 | 0.97606 | 0.006773
2019-03-21 15:59 | 1553198355.8717 | 565 | 0.20654 | 0.003414 | 0.000199 | 101.3174 | 100.4933 | 112.5569 | 21.2612 | 0.003816 | 0.003381 | 221.3201 | 0.27048 | 8.4012 | 0.003455 | 0.005514 | 0.005434 | 4.6e-05 | 0.001036 | 0.29231 | 1.2966 | 2.5479 | 98.85 | 20.72 | 38.793 | 113.7811 | 104.9902 | 103.2032 | 0.009604 | 0.93807 | 0.022477 | 23.9605 | 0.073055 | 0.16091 | 0.013684 | 0.91076 | 0.91076 | 0.007261
2019-03-21 15:59 | 1553198356.8706 | 566 | 0.20668 | 0.003343 | 0.000203 | 101.3028 | 100.5113 | 113.2369 | 21.2612 | 0.003763 | 0.003366 | 221.3085 | 0.27102 | 8.4041 | 0.003476 | 0.005248 | 0.005211 | 4.6e-05 | 0.001033 | 0.2469 | 1.2755 | 2.5478 | 98.85 | 20.72 | 38.796 | 112.4774 | 105.5062 | 102.8457 | 0.009605 | 0.93896 | 0.022555 | 24.0213 | 0.07016 | 0.16123 | 0.012936 | 0.86321 | 0.86321 | 0.007677
2019-03-21 15:59 | 1553198357.8715 | 567 | 0.20675 | 0.003271 | 0.000206 | 101.3037 | 100.6091 | 113.1146 | 21.2612 | 0.003756 | 0.003344 | 221.2883 | 0.27033 | 8.4032 | 0.00347 | 0.005284 | 0.005133 | 4.7e-05 | 0.001034 | 0.29927 | 1.2768 | 2.5479 | 98.85 | 20.72 | 38.8 | 112.3021 | 105.3677 | 102.3179 | 0.009606 | 0.94028 | 0.022559 | 23.9915 | 0.061544 | 0.16082 | 0.012588 | 0.83968 | 0.83968 | 0.007872
2019-03-21 15:59 | 1553198358.8706 | 568 | 0.2068 | 0.003191 | 0.000209 | 101.3085 | 100.6905 | 113.1913 | 21.2612 | 0.003734 | 0.003309 | 221.2878 | 0.27122 | 8.3988 | 0.003443 | 0.005418 | 0.005399 | 4.9e-05 | 0.001035 | 0.28799 | 1.2854 | 2.548 | 98.85 | 20.72 | 38.803 | 111.7874 | 104.7071 | 101.4611 | 0.009606 | 0.94243 | 0.022592 | 23.9722 | 0.054731 | 0.16135 | 0.012419 | 0.82933 | 0.82933 | 0.007996
2019-03-21 15:59 | 1553198359.8718 | 569 | 0.20702 | 0.003102 | 0.00021 | 101.2999 | 100.8451 | 113.3127 | 21.2612 | 0.003716 | 0.003303 | 221.24 | 0.27059 | 8.4017 | 0.003468 | 0.004715 | 0.004655 | 5.1e-05 | 0.001034 | 0.28655 | 1.3293 | 2.5479 | 98.85 | 20.72 | 38.805 | 111.3342 | 105.3112 | 101.3253 | 0.009607 | 0.94277 | 0.022608 | 23.9807 | 0.040243 | 0.16098 | 0.011169 | 0.74523 | 0.74523 | 0.008878
2019-03-21 15:59 | 1553198360.8698 | 570 | 0.2072 | 0.003017 | 0.00021 | 101.3282 | 101.0818 | 113.3904 | 21.2612 | 0.003649 | 0.003298 | 221.2065 | 0.27032 | 8.4021 | 0.003458 | 0.005135 | 0.005038 | 5e-05 | 0.00103 | 0.26488 | 1.3394 | 2.548 | 98.85 | 20.72 | 38.81 | 109.7186 | 105.0768 | 101.1886 | 0.009608 | 0.94312 | 0.02262 | 23.9846 | 0.021771 | 0.16082 | 0.010207 | 0.68039 | 0.68039 | 0.009714
2019-03-21 15:59 | 1553198361.8697 | 571 | 0.20746 | 0.002922 | 0.000209 | 101.2853 | 100.9453 | 113.4326 | 21.2612 | 0.003687 | 0.003291 | 221.1925 | 0.27133 | 8.3997 | 0.003424 | 0.004298 | 0.004257 | 5.2e-05 | 0.001029 | 0.2891 | 1.3363 | 2.5478 | 98.85 | 20.72 | 38.812 | 110.6422 | 104.251 | 101.0333 | 0.009609 | 0.94351 | 0.022629 | 23.984 | 0.030077 | 0.16142 | 0.008744 | 0.58157 | 0.58157 | 0.011407
2019-03-21 15:59 | 1553198362.8713 | 572 | 0.2076 | 0.002823 | 0.000207 | 101.2325 | 101.0478 | 113.754 | 21.2612 | 0.003662 | 0.003249 | 221.1675 | 0.27076 | 8.4014 | 0.003428 | 0.004729 | 0.004581 | 5.1e-05 | 0.001031 | 0.286 | 1.2991 | 2.5479 | 98.85 | 20.72 | 38.817 | 110.023 | 104.3571 | 100.0036 | 0.00961 | 0.94611 | 0.022692 | 23.9849 | 0.016339 | 0.16107 | 0.008046 | 0.5358 | 0.5358 | 0.012355
2019-03-21 15:59 | 1553198363.871 | 573 | 0.2077 | 0.002724 | 0.000204 | 101.3267 | 101.0485 | 113.661 | 21.2612 | 0.003615 | 0.003175 | 221.1446 | 0.27077 | 8.4035 | 0.003482 | 0.005052 | 0.005043 | 5.2e-05 | 0.001031 | 0.24403 | 1.319 | 2.5479 | 98.85 | 20.71 | 38.82 | 108.8704 | 105.6643 | 98.2327 | 0.009605 | 0.95062 | 0.022737 | 23.9182 | 0.024594 | 0.16108 | 0.007587 | 0.5058 | 0.5058 | 0.013089
2019-03-21 15:59 | 1553198364.8701 | 574 | 0.2078 | 0.002621 | 0.000201 | 101.3484 | 101.1387 | 113.9579 | 21.2612 | 0.003618 | 0.003218 | 221.1567 | 0.27162 | 8.4016 | 0.00351 | 0.004101 | 0.004082 | 5.3e-05 | 0.001028 | 0.29192 | 1.3531 | 2.548 | 98.85 | 20.71 | 38.823 | 108.9522 | 106.3396 | 99.2637 | 0.009605 | 0.94799 | 0.022735 | 23.9826 | 0.018528 | 0.16159 | 0.007101 | 0.47283 | 0.47283 | 0.014046
2019-03-21 15:59 | 1553198365.871 | 575 | 0.20783 | 0.002525 | 0.000197 | 101.3067 | 100.9548 | 114.1103 | 21.2612 | 0.003574 | 0.003217 | 221.1372 | 0.27096 | 8.4015 | 0.003518 | 0.004998 | 0.004931 | 5.4e-05 | 0.001029 | 0.26913 | 1.3469 | 2.548 | 98.85 | 20.71 | 38.827 | 107.8924 | 106.5254 | 99.2403 | 0.009606 | 0.94805 | 0.022751 | 23.9979 | 0.031126 | 0.1612 | 0.007065 | 0.47092 | 0.47092 | 0.014069
2019-03-21 15:59 | 1553198366.8701 | 576 | 0.20791 | 0.002425 | 0.000192 | 101.3416 | 100.9139 | 114.0159 | 21.2612 | 0.003604 | 0.003234 | 221.1294 | 0.27109 | 8.4015 | 0.003456 | 0.005294 | 0.005218 | 5.1e-05 | 0.001028 | 0.27096 | 1.3339 | 2.5479 | 98.85 | 20.71 | 38.829 | 108.6095 | 105.031 | 99.6561 | 0.009607 | 0.947 | 0.022729 | 24.0013 | 0.037825 | 0.16127 | 0.006684 | 0.44477 | 0.44477 | 0.014902
2019-03-21 15:59 | 1553198367.8705 | 577 | 0.20798 | 0.002338 | 0.000187 | 101.2966 | 101.1482 | 114.3753 | 21.2612 | 0.003581 | 0.003201 | 221.1079 | 0.27068 | 8.4042 | 0.003565 | 0.005119 | 0.005 | 5.1e-05 | 0.001025 | 0.27759 | 1.3462 | 2.548 | 98.85 | 20.71 | 38.833 | 108.057 | 107.6702 | 98.8434 | 0.009608 | 0.94906 | 0.02279 | 24.0129 | 0.013108 | 0.16103 | 0.006416 | 0.42793 | 0.42793 | 0.015466
2019-03-21 15:59 | 1553198368.8712 | 578 | 0.20811 | 0.002248 | 0.000182 | 101.2838 | 100.8216 | 114.5394 | 21.2612 | 0.003548 | 0.003236 | 221.0997 | 0.2707 | 8.4007 | 0.003495 | 0.005678 | 0.005586 | 5.3e-05 | 0.001025 | 0.25802 | 1.3728 | 2.548 | 98.85 | 20.71 | 38.836 | 107.2575 | 105.9745 | 99.703 | 0.009609 | 0.94688 | 0.02278 | 24.0579 | 0.040907 | 0.16104 | 0.005741 | 0.38194 | 0.38194 | 0.017329
2019-03-21 15:59 | 1553198369.8712 | 579 | 0.20813 | 0.002164 | 0.000176 | 101.3413 | 101.185 | 114.8344 | 21.2612 | 0.003545 | 0.003155 | 221.1107 | 0.27012 | 8.4003 | 0.003474 | 0.004502 | 0.004391 | 5.3e-05 | 0.001024 | 0.2688 | 1.3264 | 2.548 | 98.85 | 20.71 | 38.839 | 107.1748 | 105.4715 | 97.7434 | 0.009609 | 0.95188 | 0.022869 | 24.0255 | 0.013801 | 0.16069 | 0.005728 | 0.38284 | 0.38284 | 0.017251
2019-03-21 15:59 | 1553198370.8707 | 580 | 0.20823 | 0.002085 | 0.000172 | 101.2701 | 101.1029 | 114.9999 | 21.2612 | 0.003567 | 0.003125 | 221.1158 | 0.27107 | 8.4012 | 0.003532 | 0.005913 | 0.005804 | 5.7e-05 | 0.001026 | 0.27903 | 1.3008 | 2.548 | 98.85 | 20.71 | 38.845 | 107.7175 | 106.8787 | 97.021 | 0.009611 | 0.95374 | 0.022908 | 24.0193 | 0.014773 | 0.16126 | 0.005226 | 0.34933 | 0.34933 | 0.018973
2019-03-21 16:00 | 1553198371.8707 | 581 | 0.20825 | 0.002011 | 0.000166 | 101.2513 | 100.9269 | 115.1384 | 21.2612 | 0.003491 | 0.003121 | 221.1046 | 0.27206 | 8.4035 | 0.003574 | 0.004636 | 0.004702 | 5.5e-05 | 0.001024 | 0.27527 | 1.2744 | 2.5479 | 98.85 | 20.71 | 38.847 | 105.8683 | 107.8868 | 96.9184 | 0.009611 | 0.954 | 0.022925 | 24.0305 | 0.028707 | 0.16185 | 0.005187 | 0.3472 | 0.3472 | 0.019159
2019-03-21 16:00 | 1553198372.871 | 582 | 0.2083 | 0.001934 | 0.000161 | 101.3034 | 101.0437 | 115.0097 | 21.2612 | 0.003481 | 0.003144 | 221.0905 | 0.2704 | 8.4036 | 0.00356 | 0.004503 | 0.004385 | 5.5e-05 | 0.001022 | 0.27168 | 1.2807 | 2.5481 | 98.85 | 20.71 | 38.852 | 105.6282 | 107.5442 | 97.4661 | 0.009613 | 0.95259 | 0.022895 | 24.0348 | 0.022964 | 0.16086 | 0.005014 | 0.33515 | 0.33515 | 0.019727
2019-03-21 16:00 | 1553198373.8715 | 583 | 0.20835 | 0.001866 | 0.000157 | 101.2245 | 101.1724 | 115.2639 | 21.2612 | 0.003442 | 0.003138 | 221.0826 | 0.27094 | 8.4043 | 0.003551 | 0.004622 | 0.004612 | 5.6e-05 | 0.001018 | 0.25299 | 1.3034 | 2.548 | 98.85 | 20.71 | 38.853 | 104.6906 | 107.3277 | 97.3265 | 0.009613 | 0.95295 | 0.022925 | 24.0568 | 0.004601 | 0.16118 | 0.004801 | 0.32119 | 0.32119 | 0.020625
2019-03-21 16:00 | 1553198374.8705 | 584 | 0.20838 | 0.001804 | 0.000152 | 101.2406 | 101.1425 | 115.2022 | 21.2612 | 0.00349 | 0.003099 | 221.0491 | 0.27166 | 8.4041 | 0.003522 | 0.004903 | 0.004832 | 5.9e-05 | 0.001016 | 0.274 | 1.2866 | 2.548 | 98.85 | 20.71 | 38.855 | 105.8525 | 106.6293 | 96.384 | 0.009613 | 0.95538 | 0.022948 | 24.0198 | 0.008673 | 0.16161 | 0.004662 | 0.31226 | 0.31226 | 0.021272
2019-03-21 16:00 | 1553198375.8715 | 585 | 0.20845 | 0.001743 | 0.000148 | 101.2415 | 101.1132 | 115.3477 | 21.2612 | 0.003514 | 0.003083 | 221.0566 | 0.2703 | 8.4025 | 0.003598 | 0.004904 | 0.004804 | 6e-05 | 0.001014 | 0.27489 | 1.3314 | 2.5481 | 98.85 | 20.71 | 38.859 | 106.4213 | 108.4754 | 96.0082 | 0.009614 | 0.95635 | 0.022974 | 24.0227 | 0.011342 | 0.1608 | 0.004338 | 0.29052 | 0.29052 | 0.022748
2019-03-21 16:00 | 1553198376.8706 | 586 | 0.20848 | 0.001688 | 0.000145 | 101.2751 | 101.323 | 115.6083 | 21.2612 | 0.00347 | 0.003081 | 221.0652 | 0.27054 | 8.4034 | 0.003629 | 0.005207 | 0.005218 | 6.1e-05 | 0.001015 | 0.26521 | 1.3586 | 2.5479 | 98.85 | 20.71 | 38.861 | 105.36 | 109.221 | 95.9415 | 0.009615 | 0.95652 | 0.023002 | 24.0477 | -0.004223 | 0.16094 | 0.004239 | 0.28418 | 0.28418 | 0.023277
2019-03-21 16:00 | 1553198377.8696 | 587 | 0.20851 | 0.001637 | 0.000142 | 101.2444 | 101.2603 | 115.7503 | 21.2612 | 0.003374 | 0.00314 | 221.0794 | 0.27118 | 8.4 | 0.003647 | 0.004756 | 0.004627 | 6e-05 | 0.001015 | 0.27168 | 1.3295 | 2.5479 | 98.85 | 20.71 | 38.864 | 103.0271 | 109.6547 | 97.3799 | 0.009616 | 0.95281 | 0.022972 | 24.1093 | -0.0014 | 0.16132 | 0.00411 | 0.27512 | 0.27512 | 0.0241
2019-03-21 16:00 | 1553198378.8711 | 588 | 0.20852 | 0.001584 | 0.00014 | 101.2144 | 101.2181 | 115.8151 | 21.2612 | 0.003401 | 0.003082 | 221.0607 | 0.2706 | 8.4039 | 0.00369 | 0.005121 | 0.005113 | 5.9e-05 | 0.001016 | 0.26753 | 1.3028 | 2.5478 | 98.85 | 20.72 | 38.867 | 103.692 | 110.7076 | 95.9845 | 0.009622 | 0.95641 | 0.023021 | 24.0706 | -0.000328 | 0.16098 | 0.004173 | 0.28015 | 0.28015 | 0.023617
2019-03-21 16:00 | 1553198379.8708 | 589 | 0.20857 | 0.001543 | 0.000137 | 101.1664 | 101.0856 | 115.5545 | 21.2612 | 0.003421 | 0.00307 | 221.0307 | 0.27181 | 8.402 | 0.003661 | 0.005037 | 0.004926 | 6.3e-05 | 0.001015 | 0.28031 | 1.3165 | 2.5479 | 98.85 | 20.71 | 38.869 | 104.1778 | 109.9969 | 95.6994 | 0.009617 | 0.95715 | 0.023004 | 24.0342 | 0.007147 | 0.1617 | 0.003875 | 0.25961 | 0.25961 | 0.0256
2019-03-21 16:00 | 1553198380.8698 | 590 | 0.20857 | 0.001501 | 0.000135 | 101.2561 | 101.121 | 115.5893 | 21.2612 | 0.003365 | 0.003058 | 221.0131 | 0.27083 | 8.4037 | 0.00371 | 0.005753 | 0.00564 | 6.2e-05 | 0.001015 | 0.26626 | 1.2866 | 2.5479 | 98.85 | 20.71 | 38.869 | 102.8128 | 111.1843 | 95.3879 | 0.009617 | 0.95796 | 0.023018 | 24.0276 | 0.011939 | 0.16112 | 0.00393 | 0.26368 | 0.26368 | 0.025114
2019-03-21 16:00 | 1553198381.8703 | 591 | 0.20858 | 0.001457 | 0.000133 | 101.2164 | 100.9156 | 115.5721 | 21.2612 | 0.003402 | 0.003062 | 221.0066 | 0.27091 | 8.4028 | 0.00366 | 0.004213 | 0.004132 | 6.1e-05 | 0.001016 | 0.27113 | 1.3064 | 2.548 | 98.85 | 20.71 | 38.869 | 103.7032 | 109.9787 | 95.4873 | 0.009617 | 0.9577 | 0.023013 | 24.0291 | 0.026623 | 0.16117 | 0.003974 | 0.26673 | 0.26673 | 0.024834
2019-03-21 16:00 | 1553198382.8702 | 592 | 0.20863 | 0.001418 | 0.000131 | 101.1747 | 100.9231 | 115.7679 | 21.2612 | 0.003353 | 0.003029 | 220.9869 | 0.27066 | 8.4019 | 0.003671 | 0.005407 | 0.005345 | 6e-05 | 0.001012 | 0.29231 | 1.2776 | 2.548 | 98.85 | 20.71 | 38.869 | 102.5314 | 110.2366 | 94.7006 | 0.009617 | 0.95975 | 0.023057 | 24.0238 | 0.022272 | 0.16102 | 0.003718 | 0.24962 | 0.24962 | 0.026512
2019-03-21 16:00 | 1553198383.87 | 593 | 0.20861 | 0.001388 | 0.00013 | 101.2292 | 101.0777 | 115.7832 | 21.2612 | 0.003364 | 0.003025 | 220.9874 | 0.27032 | 8.4031 | 0.003646 | 0.004817 | 0.00474 | 6.1e-05 | 0.001006 | 0.28351 | 1.284 | 2.5479 | 98.85 | 20.71 | 38.87 | 102.8048 | 109.6404 | 94.5957 | 0.009617 | 0.96003 | 0.023062 | 24.0219 | 0.013397 | 0.16081 | 0.003865 | 0.25993 | 0.25993 | 0.025427
2019-03-21 16:00 | 1553198384.8701 | 594 | 0.20866 | 0.001352 | 0.000128 | 101.2865 | 101.0689 | 115.6139 | 21.2612 | 0.003362 | 0.003041 | 220.9919 | 0.27057 | 8.399 | 0.00365 | 0.005099 | 0.005125 | 5.8e-05 | 0.001006 | 0.2568 | 1.3023 | 2.548 | 98.85 | 20.71 | 38.87 | 102.7372 | 109.7429 | 94.9894 | 0.009617 | 0.959 | 0.023032 | 24.0172 | 0.019234 | 0.16096 | 0.003625 | 0.24319 | 0.24319 | 0.027203
2019-03-21 16:00 | 1553198385.87 | 595 | 0.20874 | 0.001324 | 0.000126 | 101.2135 | 101.0961 | 115.6484 | 21.2612 | 0.003331 | 0.003005 | 220.9685 | 0.27055 | 8.4025 | 0.003671 | 0.005181 | 0.005058 | 5.8e-05 | 0.001009 | 0.3027 | 1.3041 | 2.5479 | 98.85 | 20.71 | 38.869 | 101.9911 | 110.2486 | 94.119 | 0.009617 | 0.96127 | 0.023063 | 23.9924 | 0.010374 | 0.16095 | 0.00318 | 0.2128 | 0.2128 | 0.031085
2019-03-21 16:00 | 1553198386.8714 | 596 | 0.20869 | 0.001288 | 0.000123 | 101.1935 | 101.0331 | 115.5977 | 21.2612 | 0.003298 | 0.003041 | 220.9452 | 0.26928 | 8.4002 | 0.003584 | 0.005145 | 0.005104 | 6.1e-05 | 0.001006 | 0.2625 | 1.2699 | 2.5479 | 98.85 | 20.71 | 38.869 | 101.2012 | 108.124 | 94.9863 | 0.009617 | 0.95901 | 0.023031 | 24.0154 | 0.014187 | 0.16019 | 0.003486 | 0.23378 | 0.23378 | 0.028163
2019-03-21 16:00 | 1553198387.8711 | 597 | 0.20872 | 0.001262 | 0.000122 | 101.2051 | 100.9994 | 115.609 | 21.2612 | 0.003334 | 0.003032 | 220.9313 | 0.26996 | 8.3995 | 0.003595 | 0.005287 | 0.005192 | 5.7e-05 | 0.001011 | 0.27057 | 1.2953 | 2.5479 | 98.85 | 20.71 | 38.868 | 102.079 | 108.3989 | 94.7803 | 0.009617 | 0.95954 | 0.023039 | 24.0099 | 0.018199 | 0.1606 | 0.003373 | 0.22615 | 0.22615 | 0.029187
2019-03-21 16:00 | 1553198388.8702 | 598 | 0.20873 | 0.001229 | 0.00012 | 101.1649 | 101.1816 | 115.7179 | 21.2612 | 0.003321 | 0.00298 | 220.928 | 0.27085 | 8.4028 | 0.003564 | 0.005159 | 0.005125 | 6.2e-05 | 0.001003 | 0.27599 | 1.2973 | 2.548 | 98.85 | 20.71 | 38.865 | 101.7595 | 107.6435 | 93.5141 | 0.009616 | 0.96286 | 0.023089 | 23.9798 | -0.001479 | 0.16113 | 0.003373 | 0.22681 | 0.22681 | 0.029198
2019-03-21 16:00 | 1553198389.8712 | 599 | 0.20877 | 0.001208 | 0.000118 | 101.1968 | 100.9436 | 116.0496 | 21.2612 | 0.003325 | 0.002992 | 220.9373 | 0.27113 | 8.4026 | 0.003579 | 0.005214 | 0.005157 | 5.6e-05 | 0.001008 | 0.2724 | 1.2898 | 2.548 | 98.85 | 20.71 | 38.863 | 101.8543 | 108.0121 | 93.7973 | 0.009615 | 0.96212 | 0.023113 | 24.0234 | 0.02241 | 0.16129 | 0.003104 | 0.20845 | 0.20845 | 0.031802
2019-03-21 16:00 | 1553198390.8713 | 600 | 0.20876 | 0.001176 | 0.000116 | 101.3011 | 101.5008 | 116.3512 | 21.2612 | 0.003285 | 0.002952 | 220.9341 | 0.27035 | 8.4013 | 0.003653 | 0.004445 | 0.004276 | 5.8e-05 | 0.001008 | 0.3011 | 1.2781 | 2.5478 | 98.85 | 20.71 | 38.863 | 100.8833 | 109.8146 | 92.8414 | 0.009615 | 0.96463 | 0.023174 | 24.0233 | -0.017615 | 0.16083 | 0.003214 | 0.21674 | 0.21674 | 0.030498
2019-03-21 16:00 | 1553198391.8717 | 601 | 0.20881 | 0.001155 | 0.000115 | 101.1643 | 101.0127 | 116.5573 | 21.2612 | 0.003282 | 0.002948 | 220.9356 | 0.2712 | 8.402 | 0.003632 | 0.005274 | 0.00529 | 6e-05 | 0.001009 | 0.27367 | 1.3121 | 2.5479 | 98.85 | 20.71 | 38.862 | 100.8141 | 109.2856 | 92.7529 | 0.009615 | 0.96486 | 0.023197 | 24.0416 | 0.013415 | 0.16134 | 0.00295 | 0.19868 | 0.19868 | 0.033376
2019-03-21 16:00 | 1553198392.8708 | 602 | 0.20883 | 0.001133 | 0.000114 | 101.2078 | 101.0042 | 116.411 | 21.2612 | 0.003278 | 0.003011 | 220.9497 | 0.27143 | 8.4031 | 0.003657 | 0.004967 | 0.004785 | 5.7e-05 | 0.001007 | 0.27367 | 1.3473 | 2.548 | 98.85 | 20.71 | 38.861 | 100.7137 | 109.8961 | 94.2551 | 0.009615 | 0.96092 | 0.023135 | 24.0758 | 0.018014 | 0.16147 | 0.002884 | 0.19364 | 0.19364 | 0.034272
2019-03-21 16:00 | 1553198393.8709 | 603 | 0.20882 | 0.001108 | 0.000112 | 101.1652 | 101.0086 | 116.5613 | 21.2612 | 0.00325 | 0.002996 | 220.9455 | 0.2703 | 8.4066 | 0.003664 | 0.005051 | 0.005074 | 5.8e-05 | 0.001004 | 0.26504 | 1.3808 | 2.548 | 98.85 | 20.71 | 38.858 | 100.0435 | 110.0852 | 93.914 | 0.009614 | 0.96181 | 0.023161 | 24.0802 | 0.013856 | 0.1608 | 0.00296 | 0.19923 | 0.19923 | 0.033172
2019-03-21 16:00 | 1553198394.8705 | 604 | 0.20882 | 0.001084 | 0.000111 | 101.1631 | 101.1741 | 116.6514 | 21.2612 | 0.003321 | 0.002948 | 220.9207 | 0.2707 | 8.4025 | 0.003652 | 0.005442 | 0.005389 | 5.6e-05 | 0.001 | 0.29031 | 1.3636 | 2.548 | 98.85 | 20.71 | 38.857 | 101.7452 | 109.7935 | 92.7349 | 0.009614 | 0.96491 | 0.023207 | 24.0507 | -0.000975 | 0.16104 | 0.002989 | 0.2017 | 0.2017 | 0.032815
2019-03-21 16:00 | 1553198395.8696 | 605 | 0.20888 | 0.001065 | 0.00011 | 101.1583 | 101.1401 | 116.4085 | 21.2612 | 0.003316 | 0.002904 | 220.9442 | 0.27048 | 8.4024 | 0.003688 | 0.005151 | 0.004993 | 5.6e-05 | 0.001004 | 0.26322 | 1.3517 | 2.5481 | 98.85 | 20.71 | 38.855 | 101.6213 | 110.6556 | 91.6854 | 0.009613 | 0.96768 | 0.023216 | 23.9912 | 0.001616 | 0.16091 | 0.002657 | 0.17872 | 0.17872 | 0.037005
2019-03-21 16:00 | 1553198396.8711 | 606 | 0.20886 | 0.001043 | 0.000109 | 101.0645 | 101.1997 | 116.3319 | 21.2612 | 0.003238 | 0.00294 | 220.935 | 0.27137 | 8.4024 | 0.00373 | 0.005132 | 0.005064 | 5.8e-05 | 0.001003 | 0.24801 | 1.3125 | 2.5479 | 98.85 | 20.72 | 38.851 | 99.7431 | 111.6855 | 92.5425 | 0.009618 | 0.96542 | 0.023181 | 24.0115 | -0.011957 | 0.16144 | 0.00284 | 0.19126 | 0.19126 | 0.034692
2019-03-21 16:00 | 1553198397.8698 | 607 | 0.20884 | 0.001031 | 0.000107 | 101.1378 | 101.1571 | 116.124 | 21.2612 | 0.003195 | 0.002964 | 220.9446 | 0.27014 | 8.4031 | 0.003775 | 0.005263 | 0.005208 | 6e-05 | 0.000999 | 0.27455 | 1.3407 | 2.5479 | 98.85 | 20.72 | 38.849 | 98.7015 | 112.7624 | 93.1331 | 0.009618 | 0.96386 | 0.023142 | 24.0094 | -0.001708 | 0.16071 | 0.002969 | 0.19994 | 0.19994 | 0.033036
2019-03-21 16:00 | 1553198398.8707 | 608 | 0.20884 | 0.001013 | 0.000106 | 101.1095 | 100.8785 | 115.7404 | 21.2612 | 0.00324 | 0.002884 | 220.9246 | 0.27101 | 8.4036 | 0.003737 | 0.005674 | 0.005592 | 5.4e-05 | 0.001001 | 0.27129 | 1.3129 | 2.548 | 98.85 | 20.71 | 38.845 | 99.7893 | 111.8448 | 91.2054 | 0.009611 | 0.96896 | 0.023164 | 23.9065 | 0.020457 | 0.16122 | 0.002949 | 0.19887 | 0.19887 | 0.033319
2019-03-21 16:00 | 1553198399.8698 | 609 | 0.2089 | 0.000998 | 0.000105 | 101.0803 | 100.9064 | 115.6536 | 21.2612 | 0.003244 | 0.002946 | 220.9332 | 0.27105 | 8.4026 | 0.003681 | 0.004807 | 0.004748 | 5.8e-05 | 0.000996 | 0.28174 | 1.2874 | 2.5479 | 98.85 | 20.71 | 38.843 | 99.8991 | 110.4963 | 92.7059 | 0.00961 | 0.96498 | 0.023108 | 23.9467 | 0.015397 | 0.16125 | 0.002617 | 0.17543 | 0.17543 | 0.037778
2019-03-21 16:00 | 1553198400.8699 | 610 | 0.20889 | 0.000979 | 0.000104 | 101.0934 | 101.0205 | 115.2374 | 21.2612 | 0.003166 | 0.002886 | 220.9131 | 0.27056 | 8.4013 | 0.003718 | 0.00522 | 0.005109 | 6e-05 | 0.000998 | 0.27903 | 1.2763 | 2.5479 | 98.85 | 20.71 | 38.842 | 98.0092 | 111.3903 | 91.2492 | 0.00961 | 0.96884 | 0.023113 | 23.856 | 0.006449 | 0.16096 | 0.002733 | 0.18356 | 0.18356 | 0.036039
2019-03-21 16:00 | 1553198401.8699 | 611 | 0.20893 | 0.000968 | 0.000103 | 101.0582 | 100.964 | 115.0226 | 21.2612 | 0.003194 | 0.00292 | 220.9063 | 0.2717 | 8.4014 | 0.003739 | 0.005405 | 0.005337 | 5.9e-05 | 0.000997 | 0.27345 | 1.2914 | 2.548 | 98.85 | 20.71 | 38.837 | 98.6732 | 111.8916 | 92.0734 | 0.009609 | 0.96666 | 0.023065 | 23.8606 | 0.008345 | 0.16163 | 0.002504 | 0.16732 | 0.16732 | 0.039702
2019-03-21 16:00 | 1553198402.8711 | 612 | 0.20888 | 0.000957 | 0.000103 | 101.0633 | 100.9088 | 115.0757 | 21.2612 | 0.003165 | 0.002829 | 220.9077 | 0.27124 | 8.4017 | 0.003789 | 0.005391 | 0.005344 | 5.9e-05 | 0.000995 | 0.28102 | 1.2917 | 2.5479 | 98.85 | 20.71 | 38.834 | 97.9889 | 113.1211 | 89.8831 | 0.009608 | 0.97249 | 0.02314 | 23.7945 | 0.013683 | 0.16136 | 0.002787 | 0.18757 | 0.18757 | 0.035357
2019-03-21 16:00 | 1553198403.8715 | 613 | 0.20892 | 0.000942 | 0.000103 | 100.9992 | 100.9003 | 115.1452 | 21.2612 | 0.003207 | 0.002862 | 220.9328 | 0.27107 | 8.4047 | 0.003825 | 0.005816 | 0.005689 | 5.7e-05 | 0.000995 | 0.2724 | 1.3084 | 2.548 | 98.85 | 20.71 | 38.83 | 98.9877 | 113.9967 | 90.6732 | 0.009607 | 0.97038 | 0.023122 | 23.8275 | 0.008767 | 0.16126 | 0.002554 | 0.17127 | 0.17127 | 0.038697
2019-03-21 16:00 | 1553198404.8701 | 614 | 0.20893 | 0.000937 | 0.000102 | 101.0082 | 100.7913 | 115.2961 | 21.2612 | 0.003167 | 0.002878 | 220.9541 | 0.27072 | 8.4034 | 0.003792 | 0.005744 | 0.005716 | 5.8e-05 | 0.000997 | 0.27273 | 1.2969 | 2.5479 | 98.85 | 20.71 | 38.829 | 98.0203 | 113.1932 | 91.0615 | 0.009607 | 0.96934 | 0.023124 | 23.8559 | 0.019225 | 0.16105 | 0.002527 | 0.16947 | 0.16947 | 0.039059
2019-03-21 16:00 | 1553198405.8712 | 615 | 0.20896 | 0.000917 | 0.000101 | 101.0466 | 100.9418 | 115.6059 | 21.2612 | 0.003117 | 0.002881 | 220.9343 | 0.26977 | 8.4026 | 0.00381 | 0.005477 | 0.005482 | 5.9e-05 | 0.000997 | 0.2401 | 1.3151 | 2.548 | 98.85 | 20.71 | 38.828 | 96.8165 | 113.6143 | 91.121 | 0.009607 | 0.96918 | 0.023154 | 23.8899 | 0.009278 | 0.16048 | 0.00239 | 0.16023 | 0.16023 | 0.041166
2019-03-21 16:00 | 1553198406.8716 | 616 | 0.20893 | 0.000908 | 0.0001 | 100.979 | 100.7702 | 115.9357 | 21.2612 | 0.003164 | 0.002875 | 220.9417 | 0.27071 | 8.4016 | 0.003797 | 0.004954 | 0.004844 | 6e-05 | 0.000999 | 0.24546 | 1.3454 | 2.5479 | 98.85 | 20.71 | 38.825 | 97.9676 | 113.3125 | 90.9849 | 0.009606 | 0.96955 | 0.023191 | 23.9194 | 0.018513 | 0.16104 | 0.002569 | 0.17297 | 0.17297 | 0.038267
2019-03-21 16:00 | 1553198407.8703 | 617 | 0.20897 | 0.000902 | 0.0001 | 101.0362 | 101.0914 | 116.2995 | 21.2612 | 0.003088 | 0.002872 | 220.9454 | 0.27096 | 8.4052 | 0.003787 | 0.00491 | 0.004806 | 6e-05 | 0.000998 | 0.27455 | 1.3088 | 2.5479 | 98.85 | 20.71 | 38.825 | 96.1194 | 113.0675 | 90.9124 | 0.009606 | 0.96974 | 0.02323 | 23.9546 | -0.004884 | 0.1612 | 0.002326 | 0.15636 | 0.15636 | 0.042371
2019-03-21 16:00 | 1553198408.87 | 618 | 0.20896 | 0.000884 | 9.9e-05 | 101.0338 | 100.9418 | 116.7591 | 21.2612 | 0.003094 | 0.002834 | 220.948 | 0.2703 | 8.4012 | 0.003794 | 0.005223 | 0.005196 | 6e-05 | 0.001001 | 0.25321 | 1.3314 | 2.548 | 98.85 | 20.71 | 38.824 | 96.2619 | 113.2373 | 90.0002 | 0.009606 | 0.97218 | 0.023305 | 23.9718 | 0.008144 | 0.1608 | 0.00242 | 0.16348 | 0.16348 | 0.040425
2019-03-21 16:00 | 1553198409.8704 | 619 | 0.20901 | 0.000875 | 9.9e-05 | 101.108 | 101.0757 | 116.7824 | 21.2612 | 0.003097 | 0.002849 | 220.9602 | 0.27014 | 8.4033 | 0.003823 | 0.004671 | 0.004643 | 5.9e-05 | 0.000999 | 0.26001 | 1.3362 | 2.548 | 98.85 | 20.72 | 38.823 | 96.3425 | 113.9368 | 90.3692 | 0.009611 | 0.97119 | 0.023295 | 23.9863 | 0.002857 | 0.16071 | 0.002123 | 0.14264 | 0.14264 | 0.046306
2019-03-21 16:00 | 1553198410.8694 | 620 | 0.209 | 0.000867 | 9.8e-05 | 101.0284 | 100.7695 | 117.1722 | 21.2612 | 0.003053 | 0.002877 | 220.9526 | 0.27059 | 8.3999 | 0.003748 | 0.005548 | 0.005428 | 5.9e-05 | 0.001002 | 0.27455 | 1.3361 | 2.548 | 98.85 | 20.72 | 38.823 | 95.287 | 112.1111 | 91.0343 | 0.009611 | 0.96941 | 0.023313 | 24.0483 | 0.022954 | 0.16097 | 0.002201 | 0.14825 | 0.14825 | 0.044625
2019-03-21 16:00 | 1553198411.8708 | 621 | 0.209 | 0.000858 | 9.7e-05 | 101.0085 | 100.9139 | 117.3051 | 21.2612 | 0.003072 | 0.00285 | 220.9682 | 0.27038 | 8.4038 | 0.0038 | 0.004897 | 0.004919 | 5.9e-05 | 0.001003 | 0.28241 | 1.2649 | 2.5479 | 98.85 | 20.72 | 38.822 | 95.7419 | 113.3805 | 90.3824 | 0.009611 | 0.97115 | 0.023347 | 24.0404 | 0.008377 | 0.16085 | 0.00219 | 0.14777 | 0.14777 | 0.044737
2019-03-21 16:00 | 1553198412.8696 | 622 | 0.20901 | 0.000849 | 9.7e-05 | 101.0269 | 101.1166 | 117.245 | 21.2612 | 0.003065 | 0.002824 | 220.936 | 0.26996 | 8.402 | 0.00386 | 0.005205 | 0.005068 | 5.9e-05 | 0.001 | 0.2891 | 1.2967 | 2.5479 | 98.85 | 20.72 | 38.822 | 95.5609 | 114.8372 | 89.7603 | 0.009611 | 0.97282 | 0.023361 | 24.0136 | -0.007931 | 0.1606 | 0.002142 | 0.14451 | 0.14451 | 0.045676
2019-03-21 16:00 | 1553198413.8696 | 623 | 0.20891 | 0.000841 | 9.5e-05 | 100.951 | 100.872 | 117.1704 | 21.2612 | 0.00306 | 0.002838 | 220.8954 | 0.27018 | 8.3999 | 0.003813 | 0.005383 | 0.00534 | 5.5e-05 | 0.000997 | 0.27417 | 1.3007 | 2.548 | 98.85 | 20.72 | 38.822 | 95.4419 | 113.6917 | 90.0878 | 0.009611 | 0.97194 | 0.023343 | 24.0168 | 0.006997 | 0.16073 | 0.002782 | 0.18928 | 0.18928 | 0.034901
2019-03-21 16:00 | 1553198414.8698 | 624 | 0.20905 | 0.000833 | 9.5e-05 | 100.937 | 100.9395 | 117.2107 | 21.2612 | 0.00304 | 0.002828 | 220.8893 | 0.26997 | 8.3987 | 0.003795 | 0.005294 | 0.005275 | 5.6e-05 | 0.000997 | 0.26609 | 1.2761 | 2.5478 | 98.85 | 20.72 | 38.823 | 94.9714 | 113.2626 | 89.8494 | 0.009611 | 0.97258 | 0.023355 | 24.0131 | -0.000221 | 0.1606 | 0.001979 | 0.13317 | 0.13317 | 0.049565
2019-03-21 16:00 | 1553198415.87 | 625 | 0.20901 | 0.000821 | 9.3e-05 | 100.9688 | 100.8461 | 116.9321 | 21.2612 | 0.003098 | 0.002834 | 220.8851 | 0.27081 | 8.4036 | 0.003807 | 0.005552 | 0.005469 | 5.7e-05 | 0.000997 | 0.27113 | 1.3074 | 2.548 | 98.85 | 20.72 | 38.822 | 96.3693 | 113.5551 | 90.0042 | 0.009611 | 0.97216 | 0.023322 | 23.9896 | 0.010877 | 0.16111 | 0.002203 | 0.14874 | 0.14874 | 0.044518
2019-03-21 16:00 | 1553198416.8707 | 626 | 0.209 | 0.00082 | 9.2e-05 | 100.9867 | 101.1799 | 116.849 | 21.2612 | 0.00307 | 0.002823 | 220.8718 | 0.26999 | 8.3993 | 0.003756 | 0.005767 | 0.005651 | 6e-05 | 0.000995 | 0.28528 | 1.3613 | 2.5479 | 98.85 | 20.72 | 38.822 | 95.6924 | 112.3109 | 89.7307 | 0.009611 | 0.9729 | 0.023322 | 23.9721 | -0.017095 | 0.16062 | 0.002271 | 0.15357 | 0.15357 | 0.042987
2019-03-21 16:00 | 1553198417.8698 | 627 | 0.20904 | 0.000808 | 9.2e-05 | 101.0201 | 100.9613 | 116.8096 | 21.2612 | 0.003031 | 0.002831 | 220.873 | 0.27001 | 8.4043 | 0.003832 | 0.005392 | 0.005285 | 6.1e-05 | 0.000997 | 0.26576 | 1.3168 | 2.5479 | 98.85 | 20.72 | 38.821 | 94.7499 | 114.1612 | 89.9228 | 0.009611 | 0.97238 | 0.023312 | 23.9744 | 0.005211 | 0.16063 | 0.002045 | 0.13764 | 0.13764 | 0.047962
2019-03-21 16:00 | 1553198418.8712 | 628 | 0.20901 | 0.000797 | 9.2e-05 | 101.0275 | 101.1506 | 116.9413 | 21.2612 | 0.003095 | 0.002806 | 220.8703 | 0.27081 | 8.3994 | 0.003786 | 0.005665 | 0.005573 | 6.4e-05 | 0.000998 | 0.26217 | 1.3691 | 2.5479 | 98.85 | 20.72 | 38.821 | 96.28 | 113.0344 | 89.3213 | 0.009611 | 0.974 | 0.023345 | 23.968 | -0.010891 | 0.1611 | 0.002261 | 0.153 | 0.153 | 0.043277
2019-03-21 16:00 | 1553198419.8699 | 629 | 0.20904 | 0.000798 | 9.1e-05 | 100.9396 | 100.9143 | 117.1072 | 21.2612 | 0.003008 | 0.002798 | 220.8459 | 0.27088 | 8.4061 | 0.003746 | 0.004752 | 0.004735 | 5.6e-05 | 0.000997 | 0.25769 | 1.3252 | 2.548 | 98.85 | 20.72 | 38.82 | 94.1996 | 112.0668 | 89.1392 | 0.009611 | 0.97449 | 0.023367 | 23.979 | 0.00225 | 0.16115 | 0.002063 | 0.13938 | 0.13938 | 0.047517
2019-03-21 16:00 | 1553198420.8714 | 630 | 0.20904 | 0.00079 | 9.1e-05 | 100.9334 | 100.9129 | 116.8173 | 21.2612 | 0.003012 | 0.002788 | 220.8552 | 0.27142 | 8.4032 | 0.003762 | 0.004666 | 0.00459 | 5.6e-05 | 0.001003 | 0.29087 | 1.319 | 2.5479 | 98.85 | 20.72 | 38.82 | 94.2857 | 112.4601 | 88.8915 | 0.009611 | 0.97515 | 0.023346 | 23.9411 | 0.001816 | 0.16147 | 0.002087 | 0.14087 | 0.14087 | 0.047108
2019-03-21 16:00 | 1553198421.8705 | 631 | 0.20905 | 0.000781 | 9e-05 | 100.9909 | 100.8424 | 116.7286 | 21.2612 | 0.002999 | 0.002793 | 220.8364 | 0.27011 | 8.4016 | 0.003784 | 0.005029 | 0.004911 | 5.7e-05 | 0.000998 | 0.25752 | 1.3284 | 2.5479 | 98.85 | 20.72 | 38.82 | 93.9789 | 112.9916 | 89.0056 | 0.009611 | 0.97484 | 0.023334 | 23.9358 | 0.013162 | 0.16069 | 0.002007 | 0.13529 | 0.13529 | 0.048817
2019-03-21 16:00 | 1553198422.8692 | 632 | 0.20906 | 0.000778 | 9.1e-05 | 101.0284 | 100.9555 | 116.7542 | 21.2612 | 0.003023 | 0.002817 | 220.8391 | 0.27011 | 8.4052 | 0.003805 | 0.005682 | 0.005628 | 6.1e-05 | 0.001001 | 0.26145 | 1.3186 | 2.548 | 98.85 | 20.72 | 38.82 | 94.5636 | 113.4988 | 89.59 | 0.009611 | 0.97327 | 0.023317 | 23.9577 | 0.006462 | 0.16069 | 0.001966 | 0.13225 | 0.13225 | 0.049938
2019-03-21 16:00 | 1553198423.8691 | 633 | 0.20905 | 0.00077 | 9e-05 | 100.9855 | 100.856 | 116.78 | 21.2612 | 0.003052 | 0.002803 | 220.8622 | 0.27082 | 8.4027 | 0.003803 | 0.004394 | 0.00438 | 6e-05 | 0.000998 | 0.25481 | 1.2867 | 2.5479 | 98.85 | 20.72 | 38.819 | 95.2455 | 113.4474 | 89.2517 | 0.00961 | 0.97418 | 0.023331 | 23.9492 | 0.011479 | 0.16111 | 0.002028 | 0.13675 | 0.13675 | 0.04842
2019-03-21 16:00 | 1553198424.8705 | 634 | 0.20905 | 0.000773 | 9e-05 | 101.0237 | 101.0461 | 116.8983 | 21.2612 | 0.002998 | 0.002813 | 220.8629 | 0.27007 | 8.4033 | 0.00373 | 0.005571 | 0.00547 | 5.8e-05 | 0.000998 | 0.26897 | 1.3114 | 2.5478 | 98.85 | 20.72 | 38.819 | 93.9413 | 111.6872 | 89.4869 | 0.00961 | 0.97355 | 0.023335 | 23.9691 | -0.001983 | 0.16067 | 0.002002 | 0.13499 | 0.13499 | 0.048917
2019-03-21 16:00 | 1553198425.8693 | 635 | 0.20905 | 0.000771 | 9e-05 | 100.9435 | 100.8543 | 117.2948 | 21.2612 | 0.002896 | 0.002802 | 220.8569 | 0.27064 | 8.402 | 0.003679 | 0.005218 | 0.005265 | 5.8e-05 | 0.000997 | 0.25299 | 1.3032 | 2.548 | 98.85 | 20.72 | 38.819 | 91.4994 | 110.4387 | 89.2305 | 0.00961 | 0.97424 | 0.023383 | 24.0012 | 0.007908 | 0.161 | 0.002038 | 0.13778 | 0.13778 | 0.048026
2019-03-21 16:00 | 1553198426.8693 | 636 | 0.20905 | 0.00076 | 8.9e-05 | 100.9537 | 100.9429 | 116.9993 | 21.2612 | 0.002981 | 0.00278 | 220.8369 | 0.2704 | 8.4043 | 0.003709 | 0.004938 | 0.004869 | 5.8e-05 | 0.001001 | 0.2891 | 1.2951 | 2.5481 | 98.85 | 20.72 | 38.818 | 93.5361 | 111.1709 | 88.6868 | 0.00961 | 0.9757 | 0.023371 | 23.953 | 0.000955 | 0.16086 | 0.002038 | 0.13775 | 0.13775 | 0.047996
2019-03-21 16:00 | 1553198427.8703 | 637 | 0.20906 | 0.000756 | 8.9e-05 | 101.0448 | 100.9092 | 116.719 | 21.2612 | 0.003014 | 0.002792 | 220.8485 | 0.2711 | 8.4039 | 0.003615 | 0.005089 | 0.005049 | 6.1e-05 | 0.001003 | 0.27815 | 1.2836 | 2.548 | 98.85 | 20.73 | 38.818 | 94.3272 | 108.882 | 88.9916 | 0.009616 | 0.97488 | 0.023333 | 23.9343 | 0.012018 | 0.16128 | 0.001992 | 0.13426 | 0.13426 | 0.049369
2019-03-21 16:00 | 1553198428.8704 | 638 | 0.20906 | 0.000753 | 8.8e-05 | 100.993 | 101.0587 | 116.3883 | 21.2612 | 0.00303 | 0.00283 | 220.8495 | 0.27087 | 8.4009 | 0.00364 | 0.005179 | 0.005079 | 5.9e-05 | 0.001 | 0.2557 | 1.2997 | 2.548 | 98.85 | 20.73 | 38.817 | 94.7296 | 109.4994 | 89.911 | 0.009616 | 0.97241 | 0.023271 | 23.9307 | -0.005816 | 0.16114 | 0.001989 | 0.13374 | 0.13374 | 0.049518
2019-03-21 16:00 | 1553198429.8709 | 639 | 0.20905 | 0.000749 | 8.8e-05 | 101.0266 | 101.1888 | 116.4099 | 21.2612 | 0.002981 | 0.00276 | 220.8552 | 0.27004 | 8.4028 | 0.003629 | 0.005301 | 0.005221 | 5.7e-05 | 0.001001 | 0.2505 | 1.2916 | 2.5479 | 98.85 | 20.73 | 38.815 | 93.5493 | 109.2202 | 88.2135 | 0.009615 | 0.97698 | 0.023327 | 23.8769 | -0.014343 | 0.16065 | 0.002073 | 0.13998 | 0.13998 | 0.047166
2019-03-21 16:00 | 1553198430.8695 | 640 | 0.20907 | 0.000742 | 8.8e-05 | 101.0302 | 101.0927 | 116.5768 | 21.2612 | 0.003002 | 0.002753 | 220.8289 | 0.2712 | 8.4031 | 0.003718 | 0.004451 | 0.004371 | 5.6e-05 | 0.001002 | 0.2677 | 1.3088 | 2.5479 | 98.85 | 20.73 | 38.813 | 94.0486 | 111.3957 | 88.0472 | 0.009615 | 0.97743 | 0.023349 | 23.8885 | -0.005532 | 0.16134 | 0.001957 | 0.13201 | 0.13201 | 0.050232
2019-03-21 16:01 | 1553198431.8706 | 641 | 0.20905 | 0.000741 | 8.7e-05 | 101.085 | 100.9742 | 116.4042 | 21.2612 | 0.002971 | 0.002747 | 220.8397 | 0.27066 | 8.4027 | 0.003713 | 0.005383 | 0.005371 | 5.6e-05 | 0.000997 | 0.29695 | 1.2836 | 2.5479 | 98.85 | 20.73 | 38.813 | 93.2889 | 111.2536 | 87.9035 | 0.009615 | 0.97782 | 0.023337 | 23.8661 | 0.009813 | 0.16101 | 0.002067 | 0.13966 | 0.13966 | 0.047386
2019-03-21 16:01 | 1553198432.87 | 642 | 0.20905 | 0.000737 | 8.7e-05 | 101.102 | 101.1159 | 116.3376 | 21.2612 | 0.002963 | 0.002768 | 220.8456 | 0.27079 | 8.4018 | 0.003728 | 0.005518 | 0.005418 | 6.2e-05 | 0.001003 | 0.25083 | 1.3128 | 2.548 | 98.85 | 20.73 | 38.812 | 93.1081 | 111.6397 | 88.4174 | 0.009615 | 0.97643 | 0.023313 | 23.8762 | -0.001226 | 0.16109 | 0.002083 | 0.14067 | 0.14067 | 0.047067
2019-03-21 16:01 | 1553198433.8711 | 643 | 0.20902 | 0.000735 | 8.6e-05 | 101.1303 | 100.9827 | 116.1554 | 21.2612 | 0.002991 | 0.002785 | 220.8478 | 0.27108 | 8.4031 | 0.003822 | 0.005458 | 0.005389 | 6.1e-05 | 0.000999 | 0.24403 | 1.372 | 2.5478 | 98.85 | 20.73 | 38.81 | 93.7801 | 113.9098 | 88.8197 | 0.009614 | 0.97535 | 0.023282 | 23.8708 | 0.013065 | 0.16127 | 0.002278 | 0.15418 | 0.15418 | 0.042989
2019-03-21 16:01 | 1553198434.8707 | 644 | 0.20907 | 0.000735 | 8.5e-05 | 101.0964 | 101.119 | 116.1366 | 21.2612 | 0.002966 | 0.00277 | 220.8359 | 0.27044 | 8.4014 | 0.003756 | 0.004932 | 0.004862 | 5.9e-05 | 0.001 | 0.26648 | 1.3261 | 2.5479 | 98.85 | 20.73 | 38.808 | 93.1778 | 112.3067 | 88.4502 | 0.009614 | 0.97634 | 0.023292 | 23.8567 | -0.001998 | 0.16089 | 0.001974 | 0.13302 | 0.13302 | 0.049709
2019-03-21 16:01 | 1553198435.8708 | 645 | 0.20904 | 0.000726 | 8.4e-05 | 101.0514 | 101.0532 | 115.9231 | 21.2612 | 0.003003 | 0.002749 | 220.835 | 0.27049 | 8.4012 | 0.003791 | 0.005032 | 0.004876 | 5.6e-05 | 0.001001 | 0.27975 | 1.3105 | 2.5479 | 98.85 | 20.73 | 38.807 | 94.072 | 113.171 | 87.9466 | 0.009613 | 0.9777 | 0.023287 | 23.8181 | -0.000163 | 0.16091 | 0.002176 | 0.14717 | 0.14717 | 0.044938
2019-03-21 16:01 | 1553198436.8712 | 646 | 0.20905 | 0.000726 | 8.4e-05 | 101.0603 | 101.2092 | 116.2972 | 21.2612 | 0.002987 | 0.002715 | 220.8434 | 0.26996 | 8.4049 | 0.003831 | 0.005131 | 0.005123 | 6.1e-05 | 0.000999 | 0.23684 | 1.3071 | 2.5478 | 98.85 | 20.73 | 38.805 | 93.6885 | 114.1256 | 87.1211 | 0.009613 | 0.97994 | 0.023351 | 23.8293 | -0.013169 | 0.1606 | 0.002084 | 0.14118 | 0.14118 | 0.046752
2019-03-21 16:01 | 1553198437.8709 | 647 | 0.20907 | 0.000725 | 8.3e-05 | 101.0386 | 101.1745 | 116.4195 | 21.2612 | 0.002898 | 0.002762 | 220.8303 | 0.27066 | 8.4007 | 0.00373 | 0.005944 | 0.005801 | 5.5e-05 | 0.000996 | 0.286 | 1.3054 | 2.5479 | 98.85 | 20.73 | 38.805 | 91.5362 | 111.6879 | 88.2713 | 0.009613 | 0.97683 | 0.023326 | 23.8798 | -0.012023 | 0.16102 | 0.002004 | 0.13541 | 0.13541 | 0.048873
2019-03-21 16:01 | 1553198438.8699 | 648 | 0.20905 | 0.000723 | 8.2e-05 | 101.0418 | 100.9803 | 116.5554 | 21.2612 | 0.002957 | 0.002726 | 220.8157 | 0.27072 | 8.4016 | 0.003828 | 0.0054 | 0.005423 | 5.4e-05 | 0.000992 | 0.26288 | 1.3463 | 2.5479 | 98.85 | 20.73 | 38.804 | 92.9697 | 114.0679 | 87.4053 | 0.009613 | 0.97917 | 0.023368 | 23.8651 | 0.005447 | 0.16105 | 0.00207 | 0.14038 | 0.14038 | 0.04715
2019-03-21 16:01 | 1553198439.8704 | 649 | 0.20909 | 0.00072 | 8.2e-05 | 101.0138 | 101.0304 | 116.673 | 21.2612 | 0.002945 | 0.002747 | 220.8204 | 0.27031 | 8.4038 | 0.003775 | 0.005452 | 0.005392 | 5.2e-05 | 0.000993 | 0.28711 | 1.3146 | 2.5479 | 98.85 | 20.73 | 38.804 | 92.6693 | 112.7688 | 87.891 | 0.009613 | 0.97785 | 0.023364 | 23.8932 | -0.001467 | 0.16081 | 0.001843 | 0.12442 | 0.12442 | 0.053119
2019-03-21 16:01 | 1553198440.8691 | 650 | 0.20905 | 0.000715 | 8.1e-05 | 101.0332 | 101.0055 | 116.8326 | 21.2612 | 0.002915 | 0.00269 | 220.8179 | 0.27042 | 8.4036 | 0.003819 | 0.005771 | 0.005722 | 5.2e-05 | 0.000996 | 0.29391 | 1.3142 | 2.5479 | 98.85 | 20.72 | 38.802 | 91.9575 | 113.8442 | 86.527 | 0.009606 | 0.98156 | 0.023424 | 23.8644 | 0.002449 | 0.16087 | 0.002085 | 0.14186 | 0.14186 | 0.04661
2019-03-21 16:01 | 1553198441.8695 | 651 | 0.20907 | 0.00072 | 8e-05 | 100.9557 | 100.918 | 117.0406 | 21.2612 | 0.002966 | 0.002714 | 220.788 | 0.27081 | 8.4006 | 0.003809 | 0.004771 | 0.004689 | 5.3e-05 | 0.000991 | 0.28263 | 1.309 | 2.548 | 98.85 | 20.73 | 38.804 | 93.1708 | 113.5984 | 87.1078 | 0.009613 | 0.97998 | 0.023426 | 23.9049 | 0.003343 | 0.1611 | 0.001964 | 0.13339 | 0.13339 | 0.04964
2019-03-21 16:01 | 1553198442.8691 | 652 | 0.20909 | 0.000715 | 8.1e-05 | 101.0135 | 101.1339 | 117.173 | 21.2612 | 0.002915 | 0.002738 | 220.7944 | 0.2719 | 8.402 | 0.003856 | 0.005609 | 0.005524 | 5e-05 | 0.000991 | 0.26504 | 1.2753 | 2.548 | 98.85 | 20.73 | 38.803 | 91.9536 | 114.7558 | 87.6982 | 0.009612 | 0.97838 | 0.02342 | 23.938 | -0.010655 | 0.16175 | 0.001866 | 0.12643 | 0.12643 | 0.052582
2019-03-21 16:01 | 1553198443.8702 | 653 | 0.20908 | 0.000713 | 8e-05 | 100.9641 | 100.9391 | 117.0929 | 21.2612 | 0.002898 | 0.002781 | 220.7823 | 0.27023 | 8.4013 | 0.003771 | 0.004997 | 0.004908 | 5.2e-05 | 0.000993 | 0.24873 | 1.3406 | 2.548 | 98.85 | 20.73 | 38.804 | 91.5408 | 112.6733 | 88.7229 | 0.009613 | 0.97561 | 0.023379 | 23.9637 | 0.002211 | 0.16076 | 0.001964 | 0.13315 | 0.13315 | 0.049621
2019-03-21 16:01 | 1553198444.8696 | 654 | 0.20907 | 0.000711 | 7.9e-05 | 100.9 | 101.1448 | 116.8615 | 21.2612 | 0.002917 | 0.002687 | 220.7748 | 0.27107 | 8.4024 | 0.00376 | 0.005524 | 0.005463 | 5e-05 | 0.000991 | 0.28783 | 1.3413 | 2.5481 | 98.85 | 20.72 | 38.802 | 92.0038 | 112.4126 | 86.4536 | 0.009606 | 0.98176 | 0.02343 | 23.8649 | -0.021676 | 0.16126 | 0.00203 | 0.13813 | 0.13813 | 0.047984
2019-03-21 16:01 | 1553198445.8707 | 655 | 0.20911 | 0.000708 | 7.8e-05 | 100.9131 | 100.9878 | 116.9162 | 21.2612 | 0.002988 | 0.002731 | 220.7837 | 0.27049 | 8.4033 | 0.003811 | 0.004795 | 0.004801 | 4.8e-05 | 0.000993 | 0.26034 | 1.3347 | 2.548 | 98.85 | 20.73 | 38.801 | 93.7066 | 113.6452 | 87.5162 | 0.009612 | 0.97887 | 0.023401 | 23.9057 | -0.006618 | 0.16092 | 0.001747 | 0.11812 | 0.11812 | 0.055989
2019-03-21 16:01 | 1553198446.8703 | 656 | 0.20908 | 0.000708 | 7.8e-05 | 101.0231 | 100.8925 | 116.8152 | 21.2612 | 0.002903 | 0.002724 | 220.7852 | 0.27005 | 8.4023 | 0.003728 | 0.005355 | 0.005215 | 4.9e-05 | 0.00099 | 0.27455 | 1.375 | 2.5479 | 98.85 | 20.73 | 38.8 | 91.6488 | 111.6412 | 87.342 | 0.009612 | 0.97934 | 0.023396 | 23.8896 | 0.011572 | 0.16065 | 0.001938 | 0.1315 | 0.1315 | 0.050211
2019-03-21 16:01 | 1553198447.8704 | 657 | 0.20908 | 0.000707 | 7.8e-05 | 100.911 | 100.9183 | 116.9476 | 21.2612 | 0.002898 | 0.00271 | 220.7787 | 0.27092 | 8.4025 | 0.003694 | 0.005086 | 0.005048 | 5.1e-05 | 0.000991 | 0.26626 | 1.3184 | 2.548 | 98.85 | 20.73 | 38.799 | 91.5471 | 110.7905 | 87.0024 | 0.009611 | 0.98027 | 0.02342 | 23.8919 | -0.000647 | 0.16117 | 0.001966 | 0.13362 | 0.13362 | 0.049573
2019-03-21 16:01 | 1553198448.8695 | 658 | 0.20905 | 0.000703 | 7.7e-05 | 100.9075 | 100.869 | 116.8049 | 21.2612 | 0.002924 | 0.002738 | 220.7816 | 0.27055 | 8.4041 | 0.003782 | 0.004873 | 0.004829 | 5.3e-05 | 0.000991 | 0.27577 | 1.2993 | 2.548 | 98.85 | 20.73 | 38.798 | 92.1585 | 112.9538 | 87.6951 | 0.009611 | 0.97838 | 0.023384 | 23.9002 | 0.003414 | 0.16095 | 0.00211 | 0.14356 | 0.14356 | 0.04608
2019-03-21 16:01 | 1553198449.8695 | 659 | 0.20908 | 0.000703 | 7.7e-05 | 101.0061 | 100.8189 | 116.5089 | 21.2612 | 0.002913 | 0.002738 | 220.7762 | 0.27044 | 8.4018 | 0.003717 | 0.005287 | 0.005161 | 5.4e-05 | 0.000994 | 0.28584 | 1.3472 | 2.5478 | 98.85 | 20.73 | 38.798 | 91.8902 | 111.3704 | 87.6871 | 0.009611 | 0.97841 | 0.023354 | 23.8697 | 0.016592 | 0.16088 | 0.001928 | 0.13061 | 0.13061 | 0.050625
2019-03-21 16:01 | 1553198450.8695 | 660 | 0.20906 | 0.000698 | 7.7e-05 | 100.9069 | 100.9078 | 116.3045 | 21.2612 | 0.002895 | 0.002784 | 220.7845 | 0.26993 | 8.4047 | 0.003673 | 0.00496 | 0.004942 | 5.4e-05 | 0.000995 | 0.28102 | 1.3235 | 2.548 | 98.85 | 20.73 | 38.797 | 91.4586 | 110.302 | 88.786 | 0.009611 | 0.97544 | 0.023298 | 23.885 | -8.1e-05 | 0.16058 | 0.002077 | 0.14077 | 0.14077 | 0.046886
2019-03-21 16:01 | 1553198451.871 | 661 | 0.20908 | 0.000702 | 7.6e-05 | 100.9015 | 100.9827 | 116.3197 | 21.2612 | 0.002881 | 0.00269 | 220.7914 | 0.26984 | 8.4028 | 0.00365 | 0.004788 | 0.004754 | 5e-05 | 0.000992 | 0.26106 | 1.3229 | 2.548 | 98.85 | 20.73 | 38.797 | 91.1389 | 109.7383 | 86.5255 | 0.009611 | 0.98157 | 0.023373 | 23.8119 | -0.007196 | 0.16053 | 0.001956 | 0.13275 | 0.13275 | 0.049702
2019-03-21 16:01 | 1553198452.8702 | 662 | 0.20909 | 0.000696 | 7.6e-05 | 100.934 | 101.1023 | 116.2629 | 21.2612 | 0.002869 | 0.002714 | 220.7791 | 0.27082 | 8.4012 | 0.003678 | 0.005065 | 0.005006 | 5.4e-05 | 0.000993 | 0.27129 | 1.3119 | 2.548 | 98.85 | 20.73 | 38.798 | 90.8334 | 110.4243 | 87.1203 | 0.009611 | 0.97995 | 0.023348 | 23.8257 | -0.0149 | 0.16111 | 0.001875 | 0.12693 | 0.12693 | 0.052166
2019-03-21 16:01 | 1553198453.8707 | 663 | 0.2091 | 0.000701 | 7.5e-05 | 100.982 | 101.0205 | 116.3491 | 21.2612 | 0.00294 | 0.002687 | 220.77 | 0.27061 | 8.4009 | 0.003668 | 0.005138 | 0.005118 | 5.6e-05 | 0.000991 | 0.25249 | 1.3074 | 2.5481 | 98.85 | 20.73 | 38.8 | 92.5456 | 110.1705 | 86.4584 | 0.009612 | 0.98175 | 0.023378 | 23.8126 | -0.003415 | 0.16099 | 0.001864 | 0.12637 | 0.12637 | 0.052361
2019-03-21 16:01 | 1553198454.8691 | 664 | 0.20911 | 0.000695 | 7.5e-05 | 100.9414 | 100.8046 | 116.3643 | 21.2612 | 0.00284 | 0.002641 | 220.7887 | 0.27094 | 8.4029 | 0.003631 | 0.004829 | 0.004755 | 5.7e-05 | 0.000995 | 0.27256 | 1.2948 | 2.548 | 98.85 | 20.73 | 38.803 | 90.1471 | 109.2802 | 85.3477 | 0.009612 | 0.98479 | 0.023416 | 23.7774 | 0.012134 | 0.16118 | 0.001758 | 0.11919 | 0.11919 | 0.05558
2019-03-21 16:01 | 1553198455.8692 | 665 | 0.20914 | 0.000697 | 7.4e-05 | 100.9522 | 101.0696 | 116.6275 | 21.2612 | 0.002903 | 0.002727 | 220.7793 | 0.27094 | 8.4035 | 0.003663 | 0.005621 | 0.00558 | 5.4e-05 | 0.00099 | 0.25066 | 1.2696 | 2.5479 | 98.85 | 20.73 | 38.806 | 91.6644 | 110.0553 | 87.4241 | 0.009613 | 0.97912 | 0.023375 | 23.8731 | -0.010396 | 0.16118 | 0.001595 | 0.10754 | 0.10754 | 0.061603
2019-03-21 16:01 | 1553198456.87 | 666 | 0.20913 | 0.000691 | 7.3e-05 | 100.9939 | 101.1605 | 116.646 | 21.2612 | 0.00286 | 0.002665 | 220.7773 | 0.27051 | 8.4034 | 0.00359 | 0.005688 | 0.005626 | 5.7e-05 | 0.000992 | 0.25752 | 1.2525 | 2.548 | 98.85 | 20.73 | 38.81 | 90.6154 | 108.2658 | 85.9329 | 0.009614 | 0.98319 | 0.023425 | 23.8256 | -0.014746 | 0.16093 | 0.001677 | 0.11358 | 0.11358 | 0.058234
2019-03-21 16:01 | 1553198457.8699 | 667 | 0.20913 | 0.000691 | 7.3e-05 | 100.9233 | 100.8884 | 116.3113 | 21.2612 | 0.002897 | 0.002655 | 220.7731 | 0.27057 | 8.403 | 0.003593 | 0.005288 | 0.005243 | 5.5e-05 | 0.000995 | 0.29789 | 1.2801 | 2.5479 | 98.85 | 20.73 | 38.818 | 91.5268 | 108.343 | 85.6902 | 0.009616 | 0.98385 | 0.023399 | 23.7833 | 0.003092 | 0.16096 | 0.001699 | 0.11507 | 0.11507 | 0.057489
2019-03-21 16:01 | 1553198458.8703 | 668 | 0.20912 | 0.000692 | 7.3e-05 | 101.0237 | 100.966 | 116.5925 | 21.2612 | 0.00285 | 0.0027 | 220.7812 | 0.27058 | 8.4025 | 0.003615 | 0.005184 | 0.005114 | 5.9e-05 | 0.000994 | 0.25752 | 1.2602 | 2.5479 | 98.85 | 20.73 | 38.822 | 90.3713 | 108.8781 | 86.7652 | 0.009617 | 0.98091 | 0.023393 | 23.8477 | 0.005105 | 0.16097 | 0.001726 | 0.11689 | 0.11689 | 0.056596
2019-03-21 16:01 | 1553198459.8691 | 669 | 0.20913 | 0.000691 | 7.3e-05 | 101.0058 | 101.1006 | 116.781 | 21.2612 | 0.002865 | 0.002666 | 220.7757 | 0.26969 | 8.4009 | 0.003614 | 0.005185 | 0.005262 | 6e-05 | 0.000996 | 0.29031 | 1.3377 | 2.5479 | 98.84 | 20.72 | 38.825 | 90.7505 | 108.8672 | 85.9556 | 0.009613 | 0.98312 | 0.023438 | 23.8401 | -0.008388 | 0.16044 | 0.00169 | 0.11461 | 0.11461 | 0.057532
2019-03-21 16:01 | 1553198460.8691 | 670 | 0.20915 | 0.000689 | 7.2e-05 | 101.049 | 101.1162 | 116.7418 | 21.2612 | 0.002867 | 0.002725 | 220.7769 | 0.2703 | 8.4014 | 0.003607 | 0.004854 | 0.004666 | 5.8e-05 | 0.000992 | 0.28263 | 1.3501 | 2.5481 | 98.84 | 20.73 | 38.833 | 90.7951 | 108.6811 | 87.3819 | 0.009621 | 0.97923 | 0.023387 | 23.8834 | -0.005949 | 0.1608 | 0.001528 | 0.10302 | 0.10302 | 0.06415
2019-03-21 16:01 | 1553198461.8708 | 671 | 0.20912 | 0.000686 | 7.2e-05 | 101.0868 | 101.1643 | 116.6129 | 21.2612 | 0.002868 | 0.002716 | 220.771 | 0.27027 | 8.4007 | 0.003598 | 0.005098 | 0.005028 | 5.7e-05 | 0.000995 | 0.30862 | 1.3245 | 2.5479 | 98.84 | 20.73 | 38.836 | 90.8184 | 108.481 | 87.1548 | 0.009622 | 0.97985 | 0.023382 | 23.8627 | -0.006848 | 0.16078 | 0.001717 | 0.11627 | 0.11627 | 0.056832
2019-03-21 16:01 | 1553198462.8706 | 672 | 0.2091 | 0.000688 | 7.2e-05 | 101.0323 | 101.1823 | 116.7497 | 21.2612 | 0.002871 | 0.002697 | 220.7747 | 0.27014 | 8.4036 | 0.003554 | 0.005147 | 0.005093 | 5.8e-05 | 0.00099 | 0.26504 | 1.3218 | 2.5481 | 98.84 | 20.73 | 38.842 | 90.8919 | 107.3891 | 86.7089 | 0.009623 | 0.98107 | 0.02341 | 23.8619 | -0.013269 | 0.1607 | 0.001847 | 0.12557 | 0.12557 | 0.052599
2019-03-21 16:01 | 1553198463.8702 | 673 | 0.2091 | 0.000688 | 7.3e-05 | 101.0561 | 101.1169 | 116.8573 | 21.2612 | 0.002892 | 0.002698 | 220.7629 | 0.27093 | 8.4004 | 0.003548 | 0.004978 | 0.004881 | 6e-05 | 0.000992 | 0.28031 | 1.2781 | 2.5479 | 98.84 | 20.73 | 38.844 | 91.3945 | 107.2614 | 86.7182 | 0.009624 | 0.98104 | 0.023421 | 23.8732 | -0.005377 | 0.16118 | 0.00187 | 0.12718 | 0.12718 | 0.052087
2019-03-21 16:01 | 1553198464.8707 | 674 | 0.20911 | 0.000689 | 7.3e-05 | 101.1068 | 101.044 | 116.8239 | 21.2612 | 0.002881 | 0.00265 | 220.7397 | 0.27108 | 8.4002 | 0.003587 | 0.005128 | 0.005101 | 5.8e-05 | 0.000992 | 0.28174 | 1.2988 | 2.5479 | 98.84 | 20.73 | 38.846 | 91.1304 | 108.2131 | 85.5709 | 0.009624 | 0.98418 | 0.023455 | 23.8317 | 0.005555 | 0.16127 | 0.001811 | 0.12316 | 0.12316 | 0.053818
2019-03-21 16:01 | 1553198465.8708 | 675 | 0.20908 | 0.000694 | 7.3e-05 | 101.1548 | 101.3424 | 116.5714 | 21.2612 | 0.002847 | 0.002702 | 220.7368 | 0.27048 | 8.4036 | 0.003556 | 0.004925 | 0.004765 | 5.9e-05 | 0.000993 | 0.26449 | 1.2937 | 2.5479 | 98.84 | 20.73 | 38.846 | 90.3215 | 107.4529 | 86.8299 | 0.009624 | 0.98074 | 0.023388 | 23.8477 | -0.016574 | 0.16091 | 0.001985 | 0.13509 | 0.13509 | 0.048956
2019-03-21 16:01 | 1553198466.8704 | 676 | 0.20911 | 0.000692 | 7.2e-05 | 101.1822 | 101.2906 | 116.4075 | 21.2612 | 0.002852 | 0.002696 | 220.7423 | 0.27027 | 8.4022 | 0.003523 | 0.005184 | 0.005139 | 5.9e-05 | 0.00099 | 0.26952 | 1.3213 | 2.548 | 98.84 | 20.73 | 38.848 | 90.4284 | 106.6433 | 86.6754 | 0.009625 | 0.98116 | 0.023377 | 23.8258 | -0.009581 | 0.16079 | 0.001765 | 0.1196 | 0.1196 | 0.055253
2019-03-21 16:01 | 1553198467.8705 | 677 | 0.2091 | 0.000694 | 7.2e-05 | 101.1661 | 101.3053 | 116.8441 | 21.2612 | 0.002863 | 0.002681 | 220.7195 | 0.27013 | 8.4026 | 0.003515 | 0.005633 | 0.005549 | 5.8e-05 | 0.00099 | 0.29391 | 1.3108 | 2.5479 | 98.84 | 20.73 | 38.847 | 90.6997 | 106.4444 | 86.3123 | 0.009624 | 0.98215 | 0.023432 | 23.8584 | -0.012296 | 0.1607 | 0.001844 | 0.12548 | 0.12548 | 0.052636
2019-03-21 16:01 | 1553198468.8706 | 678 | 0.20909 | 0.000693 | 7.2e-05 | 101.1706 | 101.3669 | 116.8136 | 21.2612 | 0.002814 | 0.00269 | 220.7168 | 0.27028 | 8.4034 | 0.003519 | 0.005631 | 0.005542 | 6e-05 | 0.000993 | 0.28711 | 1.3214 | 2.548 | 98.84 | 20.73 | 38.848 | 89.5211 | 106.5402 | 86.5224 | 0.009625 | 0.98157 | 0.023423 | 23.8623 | -0.017342 | 0.16079 | 0.001895 | 0.12898 | 0.12898 | 0.051238
2019-03-21 16:01 | 1553198469.87 | 679 | 0.20911 | 0.000694 | 7.2e-05 | 101.2027 | 101.2041 | 116.3608 | 21.2612 | 0.002869 | 0.002696 | 220.7018 | 0.27029 | 8.4008 | 0.003502 | 0.004536 | 0.004464 | 5.9e-05 | 0.000994 | 0.3027 | 1.2702 | 2.548 | 98.84 | 20.73 | 38.848 | 90.8302 | 106.1428 | 86.6785 | 0.009625 | 0.98115 | 0.023372 | 23.8211 | -0.000121 | 0.1608 | 0.001814 | 0.12303 | 0.12303 | 0.053716
2019-03-21 16:01 | 1553198470.8696 | 680 | 0.20913 | 0.000698 | 7.2e-05 | 101.2158 | 101.1435 | 116.4075 | 21.2612 | 0.002881 | 0.002664 | 220.7099 | 0.2699 | 8.4035 | 0.003552 | 0.004671 | 0.004605 | 6.1e-05 | 0.000991 | 0.24818 | 1.2816 | 2.5479 | 98.84 | 20.73 | 38.847 | 91.1203 | 107.3646 | 85.9058 | 0.009624 | 0.98326 | 0.023402 | 23.8003 | 0.006397 | 0.16056 | 0.001631 | 0.11033 | 0.11033 | 0.05981
2019-03-21 16:01 | 1553198471.8688 | 681 | 0.20911 | 0.000699 | 7.2e-05 | 101.1619 | 101.1888 | 116.6805 | 21.2612 | 0.002883 | 0.002636 | 220.7167 | 0.27038 | 8.4034 | 0.003494 | 0.004512 | 0.004467 | 5.7e-05 | 0.00099 | 0.26465 | 1.3095 | 2.548 | 98.84 | 20.73 | 38.844 | 91.1883 | 105.943 | 85.2267 | 0.009624 | 0.98512 | 0.023452 | 23.8057 | -0.002375 | 0.16085 | 0.001792 | 0.12191 | 0.12191 | 0.054229
2019-03-21 16:01 | 1553198472.8691 | 682 | 0.20911 | 0.000696 | 7.2e-05 | 101.2233 | 101.4473 | 116.8572 | 21.2612 | 0.002819 | 0.002693 | 220.7208 | 0.27036 | 8.4024 | 0.003526 | 0.005329 | 0.005245 | 5.7e-05 | 0.000992 | 0.28528 | 1.3331 | 2.5481 | 98.84 | 20.73 | 38.839 | 89.6307 | 106.7151 | 86.5972 | 0.009622 | 0.98137 | 0.023425 | 23.8692 | -0.019771 | 0.16084 | 0.001757 | 0.1193 | 0.1193 | 0.05541
2019-03-21 16:01 | 1553198473.8692 | 683 | 0.20911 | 0.000695 | 7.2e-05 | 101.1586 | 101.2835 | 117.1244 | 21.2612 | 0.002838 | 0.002656 | 220.7206 | 0.27029 | 8.4025 | 0.003534 | 0.005432 | 0.005383 | 6e-05 | 0.000989 | 0.26897 | 1.2822 | 2.5479 | 98.84 | 20.73 | 38.837 | 90.0823 | 106.9229 | 85.709 | 0.009622 | 0.9838 | 0.02348 | 23.8669 | -0.01103 | 0.16079 | 0.001774 | 0.12082 | 0.12082 | 0.054697
2019-03-21 16:01 | 1553198474.8698 | 684 | 0.20912 | 0.000691 | 7.1e-05 | 101.3058 | 101.323 | 117.1833 | 21.2612 | 0.002819 | 0.002694 | 220.7267 | 0.27026 | 8.4028 | 0.003564 | 0.005072 | 0.005075 | 5.6e-05 | 0.000991 | 0.25962 | 1.3324 | 2.5479 | 98.84 | 20.73 | 38.835 | 89.6338 | 107.6467 | 86.6292 | 0.009621 | 0.98128 | 0.023456 | 23.9035 | -0.001514 | 0.16078 | 0.00175 | 0.119 | 0.119 | 0.055531
2019-03-21 16:01 | 1553198475.8693 | 685 | 0.20911 | 0.000695 | 7.1e-05 | 101.2921 | 101.2167 | 117.2347 | 21.2612 | 0.002875 | 0.002697 | 220.7292 | 0.27018 | 8.4024 | 0.003575 | 0.005194 | 0.005076 | 5.5e-05 | 0.000985 | 0.24945 | 1.3225 | 2.5479 | 98.84 | 20.73 | 38.83 | 90.9859 | 107.912 | 86.7019 | 0.00962 | 0.98108 | 0.023459 | 23.9112 | 0.006662 | 0.16073 | 0.001798 | 0.1224 | 0.1224 | 0.053971
2019-03-21 16:01 | 1553198476.869 | 686 | 0.20912 | 0.000691 | 7.1e-05 | 101.2266 | 101.3761 | 117.3499 | 21.2612 | 0.002847 | 0.002659 | 220.734 | 0.2696 | 8.4044 | 0.003487 | 0.004964 | 0.004974 | 5.7e-05 | 0.000988 | 0.27489 | 1.3307 | 2.548 | 98.85 | 20.73 | 38.829 | 90.3058 | 105.7839 | 85.7816 | 0.009619 | 0.9836 | 0.023501 | 23.8923 | -0.013203 | 0.16038 | 0.001719 | 0.11705 | 0.11705 | 0.056316
2019-03-21 16:01 | 1553198477.87 | 687 | 0.20907 | 0.000691 | 7.1e-05 | 101.3752 | 101.5151 | 117.4274 | 21.2612 | 0.002845 | 0.002692 | 220.755 | 0.26935 | 8.4 | 0.003404 | 0.005033 | 0.004931 | 5.3e-05 | 0.00099 | 0.26432 | 1.3392 | 2.548 | 98.85 | 20.73 | 38.824 | 90.2627 | 103.7627 | 86.5816 | 0.009618 | 0.98141 | 0.023482 | 23.9268 | -0.012328 | 0.16024 | 0.002046 | 0.14003 | 0.14003 | 0.047029
2019-03-21 16:01 | 1553198478.8705 | 688 | 0.20914 | 0.000693 | 7e-05 | 101.2543 | 101.4371 | 117.4483 | 21.2612 | 0.002821 | 0.002631 | 220.7537 | 0.27012 | 8.4008 | 0.003488 | 0.005308 | 0.005185 | 5.6e-05 | 0.00099 | 0.29878 | 1.3461 | 2.5482 | 98.85 | 20.73 | 38.821 | 89.6822 | 105.8021 | 85.1026 | 0.009617 | 0.98547 | 0.023533 | 23.8797 | -0.016132 | 0.1607 | 0.001616 | 0.11005 | 0.11005 | 0.060017
2019-03-21 16:01 | 1553198479.8705 | 689 | 0.20917 | 0.000689 | 7e-05 | 101.3392 | 101.2038 | 117.6572 | 21.2612 | 0.002876 | 0.002642 | 220.7481 | 0.27072 | 8.4024 | 0.003498 | 0.005496 | 0.005487 | 5.5e-05 | 0.000987 | 0.25177 | 1.3126 | 2.5479 | 98.85 | 20.73 | 38.818 | 91.0061 | 106.0444 | 85.3647 | 0.009616 | 0.98474 | 0.023545 | 23.9097 | 0.011962 | 0.16105 | 0.001432 | 0.097086 | 0.097086 | 0.068177
2019-03-21 16:01 | 1553198480.8702 | 690 | 0.20913 | 0.000686 | 6.9e-05 | 101.2418 | 101.4681 | 117.7335 | 21.2612 | 0.002809 | 0.00264 | 220.7466 | 0.27124 | 8.4031 | 0.003567 | 0.004447 | 0.00437 | 5.3e-05 | 0.00099 | 0.27489 | 1.342 | 2.548 | 98.85 | 20.73 | 38.811 | 89.3986 | 107.7253 | 85.3376 | 0.009614 | 0.98482 | 0.023553 | 23.9165 | -0.01997 | 0.16136 | 0.001653 | 0.11276 | 0.11276 | 0.058817
2019-03-21 16:01 | 1553198481.8692 | 691 | 0.20913 | 0.000687 | 6.8e-05 | 101.22 | 101.2763 | 117.7976 | 21.2612 | 0.002774 | 0.002634 | 220.7319 | 0.27124 | 8.406 | 0.003602 | 0.004756 | 0.004623 | 5.5e-05 | 0.000988 | 0.27345 | 1.31 | 2.5479 | 98.85 | 20.73 | 38.809 | 88.5518 | 108.5779 | 85.1846 | 0.009614 | 0.98524 | 0.023565 | 23.9179 | -0.004973 | 0.16136 | 0.001711 | 0.11698 | 0.11698 | 0.056695
2019-03-21 16:01 | 1553198482.8697 | 692 | 0.20912 | 0.000684 | 6.9e-05 | 101.324 | 101.3369 | 117.7354 | 21.2612 | 0.002811 | 0.002606 | 220.7458 | 0.27014 | 8.401 | 0.003551 | 0.004832 | 0.004774 | 5.1e-05 | 0.000991 | 0.26465 | 1.2943 | 2.548 | 98.85 | 20.73 | 38.807 | 89.4487 | 107.3197 | 84.5039 | 0.009613 | 0.98711 | 0.023581 | 23.8889 | -0.001142 | 0.1607 | 0.001763 | 0.12067 | 0.12067 | 0.054736
2019-03-21 16:01 | 1553198483.8704 | 693 | 0.20911 | 0.00068 | 6.9e-05 | 101.2397 | 101.3543 | 117.5469 | 21.2612 | 0.002787 | 0.002672 | 220.7211 | 0.27063 | 8.4033 | 0.003512 | 0.004878 | 0.004855 | 5.1e-05 | 0.000988 | 0.24889 | 1.279 | 2.548 | 98.85 | 20.73 | 38.805 | 88.8556 | 106.375 | 86.1078 | 0.009613 | 0.98271 | 0.02351 | 23.9232 | -0.010121 | 0.16099 | 0.001815 | 0.12399 | 0.12399 | 0.053368
2019-03-21 16:01 | 1553198484.8704 | 694 | 0.20911 | 0.000687 | 6.9e-05 | 101.2182 | 101.2072 | 117.5612 | 21.2612 | 0.0028 | 0.002624 | 220.713 | 0.27113 | 8.4036 | 0.003509 | 0.005478 | 0.005477 | 5.5e-05 | 0.000989 | 0.274 | 1.3146 | 2.5479 | 98.85 | 20.73 | 38.804 | 89.1869 | 106.3105 | 84.9433 | 0.009613 | 0.9859 | 0.023549 | 23.8859 | 0.000978 | 0.1613 | 0.001774 | 0.12127 | 0.12127 | 0.054663
2019-03-21 16:01 | 1553198485.9332 | 695 | 0.20913 | 0.000679 | 6.9e-05 | 101.3207 | 101.4875 | 117.575 | 21.2612 | 0.002831 | 0.002648 | 220.7035 | 0.2702 | 8.4008 | 0.003492 | 0.004515 | 0.004433 | 5.4e-05 | 0.000991 | 0.28622 | 1.2967 | 2.5481 | 98.84 | 20.74 | 38.802 | 89.9221 | 105.8911 | 85.51 | 0.009619 | 0.98435 | 0.023532 | 23.9062 | -0.014709 | 0.16074 | 0.001698 | 0.11586 | 0.11586 | 0.057021
2019-03-21 16:01 | 1553198486.8767 | 696 | 0.20912 | 0.000681 | 6.9e-05 | 101.3323 | 101.5655 | 117.4173 | 21.2612 | 0.002787 | 0.00257 | 220.6978 | 0.27035 | 8.4001 | 0.003525 | 0.005638 | 0.005517 | 5.3e-05 | 0.000986 | 0.25946 | 1.3308 | 2.5479 | 98.84 | 20.74 | 38.802 | 88.8604 | 106.6899 | 83.656 | 0.009619 | 0.98946 | 0.023577 | 23.8283 | -0.020554 | 0.16083 | 0.001754 | 0.11998 | 0.11998 | 0.055092
2019-03-21 16:01 | 1553198487.8768 | 697 | 0.2091 | 0.000676 | 6.9e-05 | 101.2555 | 101.3148 | 117.5647 | 21.2612 | 0.002804 | 0.002581 | 220.6866 | 0.27016 | 8.4018 | 0.003494 | 0.005403 | 0.005345 | 4.7e-05 | 0.000992 | 0.25962 | 1.2928 | 2.548 | 98.84 | 20.74 | 38.801 | 89.275 | 105.9391 | 83.9008 | 0.009619 | 0.98878 | 0.023584 | 23.8514 | -0.005239 | 0.16072 | 0.00187 | 0.1283 | 0.1283 | 0.051485
2019-03-21 16:01 | 1553198488.8768 | 698 | 0.20913 | 0.000678 | 6.8e-05 | 101.2355 | 101.2596 | 117.4607 | 21.2612 | 0.00276 | 0.002603 | 220.6794 | 0.26965 | 8.4022 | 0.003483 | 0.004714 | 0.004683 | 5.2e-05 | 0.000992 | 0.28031 | 1.3003 | 2.548 | 98.84 | 20.74 | 38.8 | 88.2268 | 105.6776 | 84.4486 | 0.009619 | 0.98727 | 0.023555 | 23.8592 | -0.00213 | 0.16041 | 0.001692 | 0.11561 | 0.11561 | 0.057029
2019-03-21 16:01 | 1553198489.8769 | 699 | 0.20911 | 0.000672 | 6.8e-05 | 101.2507 | 101.5198 | 117.3269 | 21.2612 | 0.002752 | 0.002594 | 220.6581 | 0.27044 | 8.402 | 0.003576 | 0.004819 | 0.004751 | 5e-05 | 0.000994 | 0.2625 | 1.2891 | 2.5478 | 98.85 | 20.74 | 38.799 | 88.0145 | 107.9327 | 84.224 | 0.009617 | 0.98789 | 0.023549 | 23.8381 | -0.023743 | 0.16089 | 0.001808 | 0.1238 | 0.1238 | 0.053411
2019-03-21 16:02 | 1553198490.8768 | 700 | 0.20913 | 0.000675 | 6.8e-05 | 101.2382 | 101.4394 | 117.3833 | 21.2612 | 0.002814 | 0.002663 | 220.6371 | 0.27006 | 8.4 | 0.003617 | 0.005452 | 0.005413 | 5.1e-05 | 0.000993 | 0.26272 | 1.3018 | 2.5481 | 98.85 | 20.74 | 38.797 | 89.5252 | 108.927 | 85.8839 | 0.009617 | 0.98332 | 0.023501 | 23.8991 | -0.017764 | 0.16066 | 0.001712 | 0.11671 | 0.11671 | 0.056575
2019-03-21 16:02 | 1553198491.8759 | 701 | 0.20917 | 0.000677 | 6.8e-05 | 101.2894 | 101.4735 | 117.0835 | 21.2612 | 0.002797 | 0.002614 | 220.6377 | 0.27088 | 8.4044 | 0.003659 | 0.004403 | 0.0043 | 5.2e-05 | 0.000996 | 0.25912 | 1.3337 | 2.548 | 98.85 | 20.74 | 38.796 | 89.1149 | 109.9559 | 84.692 | 0.009617 | 0.9866 | 0.02351 | 23.8289 | -0.01624 | 0.16115 | 0.00146 | 0.099031 | 0.099031 | 0.066879
2019-03-21 16:02 | 1553198492.8764 | 702 | 0.20913 | 0.000673 | 6.8e-05 | 101.2424 | 101.573 | 117.0875 | 21.2612 | 0.002832 | 0.002625 | 220.6055 | 0.27053 | 8.402 | 0.003616 | 0.005022 | 0.004887 | 4.8e-05 | 0.000995 | 0.28512 | 1.3233 | 2.548 | 98.85 | 20.74 | 38.795 | 89.9573 | 108.8964 | 84.9739 | 0.009616 | 0.98582 | 0.023501 | 23.8387 | -0.02916 | 0.16094 | 0.001706 | 0.11633 | 0.11633 | 0.056861
2019-03-21 16:02 | 1553198493.8762 | 703 | 0.20914 | 0.00067 | 6.7e-05 | 101.2885 | 101.559 | 117.0204 | 21.2612 | 0.002784 | 0.002639 | 220.6307 | 0.26969 | 8.4003 | 0.003554 | 0.005198 | 0.005179 | 4.9e-05 | 0.000993 | 0.26808 | 1.3264 | 2.5481 | 98.85 | 20.75 | 38.793 | 88.7891 | 107.4095 | 85.3094 | 0.009622 | 0.9849 | 0.023483 | 23.8431 | -0.023851 | 0.16044 | 0.001657 | 0.11284 | 0.11284 | 0.058439
2019-03-21 16:02 | 1553198494.876 | 704 | 0.20913 | 0.000671 | 6.8e-05 | 101.2406 | 101.3662 | 117.2129 | 21.2612 | 0.002798 | 0.002613 | 220.6326 | 0.26968 | 8.4053 | 0.003521 | 0.004994 | 0.004902 | 5.3e-05 | 0.00099 | 0.27759 | 1.2868 | 2.548 | 98.85 | 20.75 | 38.793 | 89.133 | 106.604 | 84.6742 | 0.009622 | 0.98664 | 0.023523 | 23.8415 | -0.011094 | 0.16043 | 0.001709 | 0.11665 | 0.11665 | 0.056526
2019-03-21 16:02 | 1553198495.8761 | 705 | 0.20915 | 0.000665 | 6.7e-05 | 101.2859 | 101.2899 | 117.1042 | 21.2612 | 0.002765 | 0.002639 | 220.6451 | 0.27054 | 8.4016 | 0.003529 | 0.005635 | 0.005622 | 5.3e-05 | 0.000994 | 0.27632 | 1.284 | 2.5478 | 98.85 | 20.75 | 38.793 | 88.3401 | 106.7993 | 85.3087 | 0.009622 | 0.9849 | 0.023491 | 23.8516 | -0.000359 | 0.16094 | 0.001605 | 0.10923 | 0.10923 | 0.060558
2019-03-21 16:02 | 1553198496.8772 | 706 | 0.20913 | 0.000665 | 6.9e-05 | 101.2385 | 101.3683 | 117.391 | 21.2612 | 0.002789 | 0.002635 | 220.6479 | 0.26961 | 8.4017 | 0.003519 | 0.004591 | 0.004507 | 5e-05 | 0.000994 | 0.25962 | 1.305 | 2.5481 | 98.85 | 20.75 | 38.794 | 88.9102 | 106.5412 | 85.1961 | 0.009622 | 0.98521 | 0.023524 | 23.877 | -0.011459 | 0.16039 | 0.001679 | 0.11452 | 0.11452 | 0.057562
2019-03-21 16:02 | 1553198497.8763 | 707 | 0.20913 | 0.000667 | 6.9e-05 | 101.3222 | 101.5546 | 117.088 | 21.2612 | 0.002768 | 0.002626 | 220.6539 | 0.26997 | 8.4014 | 0.003456 | 0.005282 | 0.00527 | 4.8e-05 | 0.000993 | 0.27688 | 1.3461 | 2.548 | 98.85 | 20.75 | 38.794 | 88.4174 | 105.0232 | 85.0035 | 0.009622 | 0.98574 | 0.0235 | 23.8398 | -0.020489 | 0.16061 | 0.001682 | 0.11457 | 0.11457 | 0.057616
2019-03-21 16:02 | 1553198498.8763 | 708 | 0.20911 | 0.000667 | 6.9e-05 | 101.2635 | 101.4173 | 116.8394 | 21.2612 | 0.002814 | 0.002663 | 220.6573 | 0.27055 | 8.4001 | 0.003485 | 0.005017 | 0.004934 | 4.8e-05 | 0.000994 | 0.30038 | 1.3454 | 2.5479 | 98.85 | 20.75 | 38.794 | 89.504 | 105.7383 | 85.8908 | 0.009622 | 0.9833 | 0.023446 | 23.8439 | -0.013574 | 0.16095 | 0.001809 | 0.12326 | 0.12326 | 0.053668
2019-03-21 16:02 | 1553198499.8764 | 709 | 0.20913 | 0.000665 | 6.9e-05 | 101.3231 | 101.3843 | 117.1328 | 21.2612 | 0.002737 | 0.00261 | 220.6652 | 0.27041 | 8.4024 | 0.003497 | 0.004746 | 0.004703 | 4.8e-05 | 0.000993 | 0.28495 | 1.3174 | 2.5479 | 98.85 | 20.75 | 38.795 | 87.6686 | 106.0271 | 84.5984 | 0.009622 | 0.98685 | 0.023518 | 23.8308 | -0.005398 | 0.16087 | 0.001718 | 0.11716 | 0.11716 | 0.056434
2019-03-21 16:02 | 1553198500.8775 | 710 | 0.20917 | 0.000661 | 6.9e-05 | 101.3475 | 101.3693 | 117.0302 | 21.2612 | 0.002755 | 0.002622 | 220.6728 | 0.26921 | 8.4014 | 0.003465 | 0.005109 | 0.00508 | 4.6e-05 | 0.000988 | 0.28926 | 1.308 | 2.5479 | 98.85 | 20.75 | 38.795 | 88.0879 | 105.2404 | 84.9012 | 0.009622 | 0.98602 | 0.023497 | 23.8305 | -0.001919 | 0.16015 | 0.001468 | 0.099457 | 0.099457 | 0.066183
2019-03-21 16:02 | 1553198501.8758 | 711 | 0.20914 | 0.000665 | 6.8e-05 | 101.2748 | 101.3362 | 117.4394 | 21.2612 | 0.002785 | 0.002592 | 220.6645 | 0.26964 | 8.4039 | 0.003428 | 0.00413 | 0.004043 | 5.1e-05 | 0.000989 | 0.26178 | 1.3511 | 2.5479 | 98.85 | 20.75 | 38.795 | 88.8096 | 104.3477 | 84.1757 | 0.009622 | 0.98802 | 0.023562 | 23.8479 | -0.005422 | 0.16041 | 0.00166 | 0.11331 | 0.11331 | 0.058181
2019-03-21 16:02 | 1553198502.8756 | 712 | 0.20912 | 0.000668 | 6.9e-05 | 101.2969 | 101.3196 | 117.4067 | 21.2612 | 0.002792 | 0.002599 | 220.6517 | 0.27009 | 8.4018 | 0.003437 | 0.005033 | 0.00498 | 4.3e-05 | 0.000988 | 0.29607 | 1.3207 | 2.548 | 98.85 | 20.75 | 38.795 | 88.9892 | 104.5694 | 84.3378 | 0.009622 | 0.98757 | 0.023554 | 23.85 | -0.002002 | 0.16067 | 0.001777 | 0.12154 | 0.12154 | 0.054334
2019-03-21 16:02 | 1553198503.8767 | 713 | 0.20914 | 0.000662 | 6.9e-05 | 101.2436 | 101.5181 | 117.3321 | 21.2612 | 0.002738 | 0.002581 | 220.6545 | 0.27075 | 8.4017 | 0.003446 | 0.00461 | 0.004586 | 4.8e-05 | 0.000988 | 0.29142 | 1.3358 | 2.548 | 98.85 | 20.75 | 38.796 | 87.6763 | 104.7723 | 83.9039 | 0.009623 | 0.98877 | 0.02356 | 23.8279 | -0.024225 | 0.16107 | 0.001643 | 0.11214 | 0.11214 | 0.059032
2019-03-21 16:02 | 1553198504.8769 | 714 | 0.20912 | 0.000667 | 6.8e-05 | 101.3055 | 101.4074 | 117.2265 | 21.2612 | 0.002714 | 0.002605 | 220.6466 | 0.27029 | 8.4017 | 0.00342 | 0.00498 | 0.004816 | 4.9e-05 | 0.000986 | 0.25553 | 1.3396 | 2.5479 | 98.84 | 20.75 | 38.797 | 87.0984 | 104.1582 | 84.497 | 0.009624 | 0.98713 | 0.02353 | 23.837 | -0.008993 | 0.16079 | 0.001777 | 0.12144 | 0.12144 | 0.054417
2019-03-21 16:02 | 1553198505.8757 | 715 | 0.20916 | 0.000664 | 6.9e-05 | 101.3192 | 101.5137 | 117.6586 | 21.2612 | 0.002767 | 0.002595 | 220.6154 | 0.27068 | 8.4008 | 0.003496 | 0.004275 | 0.004303 | 4.9e-05 | 0.000985 | 0.30182 | 1.2994 | 2.548 | 98.84 | 20.74 | 38.797 | 88.3869 | 105.9886 | 84.256 | 0.009618 | 0.9878 | 0.023582 | 23.8728 | -0.017151 | 0.16103 | 0.001554 | 0.10594 | 0.10594 | 0.062474
2019-03-21 16:02 | 1553198506.8767 | 716 | 0.20914 | 0.000662 | 6.8e-05 | 101.2641 | 101.3308 | 117.6053 | 21.2612 | 0.002764 | 0.002599 | 220.5891 | 0.27068 | 8.4035 | 0.003458 | 0.004703 | 0.004562 | 5.3e-05 | 0.000985 | 0.262 | 1.3071 | 2.5477 | 98.84 | 20.75 | 38.796 | 88.319 | 105.0712 | 84.3534 | 0.009624 | 0.98753 | 0.023573 | 23.8707 | -0.005889 | 0.16103 | 0.00167 | 0.11413 | 0.11413 | 0.057987
2019-03-21 16:02 | 1553198507.8768 | 717 | 0.20914 | 0.000661 | 6.8e-05 | 101.3452 | 101.4991 | 117.5706 | 21.2612 | 0.002818 | 0.002571 | 220.5365 | 0.27032 | 8.401 | 0.003413 | 0.004567 | 0.004443 | 4.7e-05 | 0.000987 | 0.27472 | 1.2649 | 2.548 | 98.85 | 20.75 | 38.794 | 89.6071 | 103.9806 | 83.6668 | 0.009622 | 0.98943 | 0.023592 | 23.8442 | -0.013573 | 0.16081 | 0.001678 | 0.11478 | 0.11478 | 0.057581
2019-03-21 16:02 | 1553198508.8759 | 718 | 0.20915 | 0.00066 | 6.8e-05 | 101.2415 | 101.3465 | 117.8384 | 21.2612 | 0.00278 | 0.002584 | 220.5014 | 0.27058 | 8.4024 | 0.003426 | 0.004943 | 0.004879 | 5.4e-05 | 0.000986 | 0.26106 | 1.292 | 2.5479 | 98.85 | 20.75 | 38.792 | 88.6924 | 104.3045 | 83.9727 | 0.009622 | 0.98858 | 0.023609 | 23.8816 | -0.009272 | 0.16097 | 0.001611 | 0.1101 | 0.1101 | 0.060086
2019-03-21 16:02 | 1553198509.8769 | 719 | 0.20913 | 0.000663 | 6.7e-05 | 101.2742 | 101.4207 | 117.9684 | 21.2612 | 0.002771 | 0.002581 | 220.5272 | 0.27081 | 8.3996 | 0.00345 | 0.005515 | 0.005361 | 5.1e-05 | 0.000984 | 0.26936 | 1.2907 | 2.548 | 98.85 | 20.75 | 38.792 | 88.487 | 104.8816 | 83.9095 | 0.009622 | 0.98876 | 0.023624 | 23.8927 | -0.012927 | 0.1611 | 0.001694 | 0.11617 | 0.11617 | 0.056999
2019-03-21 16:02 | 1553198510.877 | 720 | 0.20897 | 0.000661 | 6.7e-05 | 101.2424 | 101.3805 | 118.2593 | 21.2612 | 0.002751 | 0.00261 | 220.5098 | 0.26993 | 8.4023 | 0.003496 | 0.004965 | 0.004993 | 5.1e-05 | 0.000984 | 0.24746 | 1.3184 | 2.5478 | 98.85 | 20.75 | 38.79 | 88.0043 | 106.0019 | 84.6134 | 0.009621 | 0.98681 | 0.02363 | 23.9457 | -0.012198 | 0.16058 | 0.002693 | 0.18706 | 0.18706 | 0.035282
2019-03-21 16:02 | 1553198511.8767 | 721 | 0.20902 | 0.000662 | 6.6e-05 | 101.2695 | 101.3213 | 118.2309 | 21.2612 | 0.002775 | 0.002563 | 220.5181 | 0.27067 | 8.402 | 0.003485 | 0.004901 | 0.004849 | 5.5e-05 | 0.000982 | 0.27688 | 1.3221 | 2.5478 | 98.85 | 20.75 | 38.788 | 88.5657 | 105.7351 | 83.4788 | 0.009621 | 0.98995 | 0.023665 | 23.9048 | -0.004573 | 0.16102 | 0.00241 | 0.16724 | 0.16724 | 0.039572
2019-03-21 16:02 | 1553198512.8761 | 722 | 0.20901 | 0.000661 | 6.6e-05 | 101.3019 | 101.3771 | 118.0764 | 21.2612 | 0.002733 | 0.002601 | 220.5277 | 0.27068 | 8.4041 | 0.003524 | 0.004733 | 0.004605 | 5.1e-05 | 0.000984 | 0.26576 | 1.3317 | 2.548 | 98.85 | 20.75 | 38.788 | 87.574 | 106.6836 | 84.3871 | 0.009621 | 0.98744 | 0.023619 | 23.9196 | -0.006635 | 0.16103 | 0.002463 | 0.17062 | 0.17062 | 0.038788
2019-03-21 16:02 | 1553198513.8771 | 723 | 0.2092 | 0.000663 | 6.6e-05 | 101.2587 | 101.4725 | 117.9889 | 21.2612 | 0.002729 | 0.002579 | 220.5041 | 0.26983 | 8.4026 | 0.003517 | 0.004844 | 0.004853 | 5.1e-05 | 0.000984 | 0.28456 | 1.3151 | 2.5478 | 98.85 | 20.75 | 38.786 | 87.4748 | 106.5048 | 83.858 | 0.00962 | 0.9889 | 0.023628 | 23.893 | -0.01886 | 0.16052 | 0.001265 | 0.085809 | 0.085809 | 0.076886
2019-03-21 16:02 | 1553198514.8772 | 724 | 0.20912 | 0.000658 | 6.6e-05 | 101.3273 | 101.34 | 118.1959 | 21.2612 | 0.002757 | 0.002542 | 220.5052 | 0.26895 | 8.4 | 0.003469 | 0.005491 | 0.00539 | 5.1e-05 | 0.000981 | 0.26416 | 1.3291 | 2.5479 | 98.85 | 20.75 | 38.786 | 88.144 | 105.3299 | 82.9821 | 0.00962 | 0.99133 | 0.023678 | 23.8846 | -0.001123 | 0.16 | 0.001765 | 0.12151 | 0.12151 | 0.054117
2019-03-21 16:02 | 1553198515.8763 | 725 | 0.20914 | 0.000659 | 6.7e-05 | 101.2587 | 101.4585 | 118.0327 | 21.2612 | 0.002784 | 0.002603 | 220.5513 | 0.27038 | 8.4013 | 0.003545 | 0.004758 | 0.004622 | 5.2e-05 | 0.000983 | 0.28456 | 1.3576 | 2.5481 | 98.85 | 20.76 | 38.787 | 88.7877 | 107.1859 | 84.4361 | 0.009626 | 0.9873 | 0.023613 | 23.9168 | -0.017629 | 0.16085 | 0.001675 | 0.11468 | 0.11468 | 0.057646
2019-03-21 16:02 | 1553198516.8763 | 726 | 0.20913 | 0.000661 | 6.7e-05 | 101.2468 | 101.3921 | 118.0656 | 21.2612 | 0.002725 | 0.002568 | 220.5507 | 0.27074 | 8.4044 | 0.003486 | 0.004748 | 0.004739 | 5.3e-05 | 0.000981 | 0.25122 | 1.3532 | 2.5478 | 98.85 | 20.76 | 38.786 | 87.3803 | 105.7453 | 83.6029 | 0.009626 | 0.98961 | 0.023644 | 23.8922 | -0.012825 | 0.16106 | 0.001701 | 0.1167 | 0.1167 | 0.056724
2019-03-21 16:02 | 1553198517.8767 | 727 | 0.20914 | 0.000658 | 6.8e-05 | 101.212 | 101.4609 | 118.0318 | 21.2612 | 0.002736 | 0.002536 | 220.592 | 0.26971 | 8.4046 | 0.003487 | 0.004845 | 0.0048 | 5.4e-05 | 0.000979 | 0.28744 | 1.3503 | 2.5481 | 98.85 | 20.76 | 38.787 | 87.6373 | 105.7793 | 82.8188 | 0.009626 | 0.99179 | 0.023667 | 23.8625 | -0.021973 | 0.16045 | 0.001661 | 0.11394 | 0.11394 | 0.057879
2019-03-21 16:02 | 1553198518.8768 | 728 | 0.20915 | 0.000655 | 6.7e-05 | 101.2689 | 101.3615 | 118.0384 | 21.2612 | 0.002729 | 0.002589 | 220.6097 | 0.27073 | 8.4008 | 0.003481 | 0.004692 | 0.004732 | 5.4e-05 | 0.000977 | 0.26305 | 1.3144 | 2.548 | 98.85 | 20.76 | 38.789 | 87.4662 | 105.6267 | 84.0951 | 0.009627 | 0.98824 | 0.023625 | 23.906 | -0.008173 | 0.16106 | 0.001615 | 0.11055 | 0.11055 | 0.059878
2019-03-21 16:02 | 1553198519.8756 | 729 | 0.20915 | 0.000657 | 6.8e-05 | 101.293 | 101.5008 | 117.9968 | 21.2612 | 0.002754 | 0.002517 | 220.6011 | 0.2702 | 8.4024 | 0.003544 | 0.004972 | 0.004839 | 5.6e-05 | 0.000979 | 0.27223 | 1.2882 | 2.5479 | 98.85 | 20.76 | 38.79 | 88.0809 | 107.1653 | 82.3674 | 0.009627 | 0.99305 | 0.023678 | 23.8439 | -0.018326 | 0.16074 | 0.001589 | 0.1089 | 0.1089 | 0.060667
2019-03-21 16:02 | 1553198520.8754 | 730 | 0.20914 | 0.000655 | 6.9e-05 | 101.2921 | 101.435 | 117.8299 | 21.2612 | 0.00275 | 0.002554 | 220.5997 | 0.2704 | 8.4016 | 0.003478 | 0.004807 | 0.004755 | 5.3e-05 | 0.00098 | 0.28479 | 1.2732 | 2.548 | 98.85 | 20.76 | 38.791 | 87.9832 | 105.5574 | 83.2658 | 0.009627 | 0.99054 | 0.023631 | 23.8571 | -0.01261 | 0.16086 | 0.001679 | 0.115 | 0.115 | 0.057489
2019-03-21 16:02 | 1553198521.8765 | 731 | 0.20912 | 0.00066 | 6.8e-05 | 101.3019 | 101.4411 | 117.7584 | 21.2612 | 0.002737 | 0.002578 | 220.585 | 0.27022 | 8.4077 | 0.003469 | 0.004901 | 0.004785 | 5.5e-05 | 0.000979 | 0.26106 | 1.2471 | 2.5479 | 98.85 | 20.76 | 38.791 | 87.6662 | 105.3277 | 83.8393 | 0.009627 | 0.98895 | 0.023605 | 23.869 | -0.012282 | 0.16076 | 0.001779 | 0.12198 | 0.12198 | 0.054165
2019-03-21 16:02 | 1553198522.8758 | 732 | 0.20916 | 0.000659 | 6.8e-05 | 101.2212 | 101.481 | 118.1062 | 21.2612 | 0.002726 | 0.002597 | 220.5637 | 0.26991 | 8.4013 | 0.003511 | 0.004881 | 0.004807 | 5.5e-05 | 0.000977 | 0.27865 | 1.2362 | 2.548 | 98.85 | 20.76 | 38.79 | 87.3966 | 106.3546 | 84.2911 | 0.009627 | 0.9877 | 0.023625 | 23.9194 | -0.022927 | 0.16057 | 0.001522 | 0.1039 | 0.1039 | 0.063513
2019-03-21 16:02 | 1553198523.877 | 733 | 0.20915 | 0.000655 | 6.7e-05 | 101.296 | 101.3059 | 118.4114 | 21.2612 | 0.002696 | 0.002585 | 220.5759 | 0.27027 | 8.4037 | 0.003495 | 0.005251 | 0.005174 | 5.3e-05 | 0.000975 | 0.25321 | 1.2907 | 2.5479 | 98.85 | 20.76 | 38.792 | 86.6723 | 105.9815 | 84.0054 | 0.009627 | 0.98849 | 0.023665 | 23.9407 | -0.000878 | 0.16078 | 0.001583 | 0.10847 | 0.10847 | 0.060922
2019-03-21 16:02 | 1553198524.8771 | 734 | 0.20916 | 0.000656 | 6.6e-05 | 101.3112 | 101.3471 | 117.9835 | 21.2612 | 0.002718 | 0.002591 | 220.5598 | 0.27053 | 8.4011 | 0.003562 | 0.004962 | 0.004859 | 5.1e-05 | 0.00098 | 0.2636 | 1.2947 | 2.5479 | 98.85 | 20.76 | 38.793 | 87.2046 | 107.6064 | 84.1437 | 0.009628 | 0.98811 | 0.023618 | 23.902 | -0.003175 | 0.16094 | 0.001567 | 0.10713 | 0.10713 | 0.061742
2019-03-21 16:02 | 1553198525.8758 | 735 | 0.20915 | 0.000658 | 6.7e-05 | 101.3011 | 101.5001 | 117.9827 | 21.2612 | 0.002687 | 0.002562 | 220.5372 | 0.26897 | 8.4014 | 0.003587 | 0.004655 | 0.004587 | 5.1e-05 | 0.000982 | 0.27041 | 1.3523 | 2.5479 | 98.85 | 20.76 | 38.794 | 86.4661 | 108.1973 | 83.4423 | 0.009628 | 0.99005 | 0.023641 | 23.8785 | -0.017555 | 0.16001 | 0.001591 | 0.1089 | 0.1089 | 0.060391
2019-03-21 16:02 | 1553198526.8768 | 736 | 0.20916 | 0.000656 | 6.6e-05 | 101.2564 | 101.2664 | 118.0905 | 21.2612 | 0.002715 | 0.002553 | 220.9869 | 0.26961 | 8.4008 | 0.003568 | 0.004194 | 0.004128 | 5.1e-05 | 0.000977 | 0.26698 | 1.3498 | 2.5478 | 98.85 | 20.76 | 38.794 | 87.1398 | 107.7341 | 83.2331 | 0.009628 | 0.99064 | 0.023659 | 23.8824 | -0.000889 | 0.16039 | 0.001558 | 0.10673 | 0.10673 | 0.061766
2019-03-21 16:02 | 1553198527.8769 | 737 | 0.20914 | 0.000661 | 6.6e-05 | 101.2802 | 101.4115 | 118.3935 | 21.2612 | 0.002695 | 0.002544 | 193.4648 | 0.27036 | 8.4036 | 0.003597 | 0.005451 | 0.005331 | 5.2e-05 | 0.000982 | 0.25442 | 1.3654 | 2.5478 | 98.85 | 20.76 | 38.793 | 86.6542 | 108.4375 | 83.0179 | 0.009628 | 0.99123 | 0.023696 | 23.9058 | -0.01159 | 0.16084 | 0.00168 | 0.11557 | 0.11557 | 0.0572
2019-03-21 16:02 | 1553198528.876 | 738 | 0.20919 | 0.000659 | 6.6e-05 | 101.2906 | 101.3257 | 118.7017 | 21.2612 | 0.002717 | 0.002578 | 156.187 | 0.2704 | 8.4012 | 0.003485 | 0.004879 | 0.004895 | 5.3e-05 | 0.000981 | 0.28047 | 1.3403 | 2.5478 | 98.85 | 20.77 | 38.79 | 87.1875 | 105.7257 | 83.8361 | 0.009633 | 0.98896 | 0.0237 | 23.9643 | -0.003096 | 0.16086 | 0.001341 | 0.091499 | 0.091499 | 0.072257
2019-03-21 16:02 | 1553198529.8769 | 739 | 0.20918 | 0.000657 | 6.5e-05 | 101.3028 | 101.467 | 119.0723 | 21.2612 | 0.002741 | 0.002525 | 118.9076 | 0.27014 | 8.4007 | 0.003521 | 0.00524 | 0.005155 | 5e-05 | 0.00098 | 0.24674 | 1.3303 | 2.5478 | 98.85 | 20.77 | 38.789 | 87.7622 | 106.5892 | 82.5606 | 0.009633 | 0.99251 | 0.023779 | 23.9588 | -0.014486 | 0.16071 | 0.001443 | 0.099109 | 0.099109 | 0.066645
2019-03-21 16:02 | 1553198530.8764 | 740 | 0.20915 | 0.000656 | 6.3e-05 | 101.2766 | 101.3397 | 119.2334 | 21.2612 | 0.002699 | 0.002517 | 81.6345 | 0.26944 | 8.4034 | 0.003567 | 0.005167 | 0.005106 | 5.1e-05 | 0.000982 | 0.24115 | 1.3365 | 2.5479 | 98.85 | 20.77 | 38.788 | 86.7527 | 107.7127 | 82.376 | 0.009632 | 0.99302 | 0.023802 | 23.9688 | -0.005565 | 0.16029 | 0.001593 | 0.11 | 0.11 | 0.05989
2019-03-21 16:02 | 1553198531.8767 | 741 | 0.20919 | 0.000656 | 6.1e-05 | 101.2677 | 101.3574 | 119.2772 | 21.2612 | 0.002659 | 0.002557 | 44.3594 | 0.26952 | 8.4024 | 0.003453 | 0.005226 | 0.005192 | 5.1e-05 | 0.000983 | 0.2505 | 1.301 | 2.5478 | 98.85 | 20.77 | 38.784 | 85.7823 | 104.9563 | 83.3228 | 0.009631 | 0.99038 | 0.023774 | 24.0051 | -0.007918 | 0.16034 | 0.001368 | 0.09397 | 0.09397 | 0.070127
2019-03-21 16:02 | 1553198532.8752 | 742 | 0.20923 | 0.000652 | 5.6e-05 | 101.2233 | 101.3253 | 119.1597 | 21.2612 | 0.002653 | 0.002478 | 7.1098 | 0.24907 | 8.4044 | 0.003518 | 0.004799 | 0.004637 | 5e-05 | 0.000986 | 0.27848 | 1.3323 | 2.5479 | 98.85 | 20.77 | 38.781 | 85.6363 | 106.5191 | 81.4376 | 0.009631 | 0.99565 | 0.023826 | 23.9297 | -0.009014 | 0.14817 | 0.001159 | 0.079524 | 0.079524 | 0.076578
2019-03-21 16:02 | 1553198533.8752 | 743 | 0.20924 | 0.000652 | 5.2e-05 | 101.324 | 101.5076 | 119.1754 | 21.2612 | 0.00269 | 0.002511 | -30.1721 | 0.24071 | 8.4016 | 0.003516 | 0.004579 | 0.004452 | 5e-05 | 0.000984 | 0.25785 | 1.3307 | 2.548 | 98.85 | 20.77 | 38.777 | 86.527 | 106.4742 | 82.2301 | 0.00963 | 0.99343 | 0.023801 | 23.958 | -0.01619 | 0.1432 | 0.001063 | 0.072858 | 0.072858 | 0.080782
2019-03-21 16:02 | 1553198534.8762 | 744 | 0.2093 | 0.00065 | 4.8e-05 | 101.2876 | 101.34 | 119.6483 | 21.2612 | 0.002673 | 0.002492 | -67.4555 | 0.24087 | 8.4005 | 0.003515 | 0.0046 | 0.004413 | 5.1e-05 | 0.000983 | 0.27831 | 1.3479 | 2.5479 | 98.85 | 20.77 | 38.774 | 86.1288 | 106.4553 | 81.7777 | 0.009629 | 0.9947 | 0.023863 | 23.9902 | -0.004622 | 0.14329 | 0.000718 | 0.048541 | 0.048541 | 0.12133
2019-03-21 16:02 | 1553198535.8763 | 745 | 0.20928 | 0.000651 | 4.5e-05 | 101.2847 | 101.4381 | 119.9138 | 21.2612 | 0.00271 | 0.002485 | -104.7527 | 0.24123 | 8.4028 | 0.003469 | 0.004904 | 0.004709 | 5e-05 | 0.000987 | 0.274 | 1.3207 | 2.548 | 98.85 | 20.77 | 38.774 | 87.0226 | 105.3505 | 81.6109 | 0.009629 | 0.99516 | 0.023895 | 24.0112 | -0.013538 | 0.14351 | 0.000821 | 0.056231 | 0.056231 | 0.10489
2019-03-21 16:02 | 1553198536.8764 | 746 | 0.2093 | 0.000641 | 4.2e-05 | 101.2835 | 101.3778 | 120.3671 | 21.2612 | 0.00266 | 0.002488 | -142.4784 | 0.24114 | 8.4032 | 0.003437 | 0.004665 | 0.004536 | 5.1e-05 | 0.000989 | 0.30093 | 1.3192 | 2.5479 | 98.85 | 20.77 | 38.769 | 85.8112 | 104.5709 | 81.6763 | 0.009628 | 0.99498 | 0.023938 | 24.0587 | -0.008326 | 0.14345 | 0.000744 | 0.050941 | 0.050941 | 0.11574
2019-03-21 16:02 | 1553198537.8754 | 747 | 0.20932 | 0.000638 | 3.9e-05 | 101.2972 | 101.4854 | 120.6321 | 21.2612 | 0.002687 | 0.002516 | -152.2529 | 0.24126 | 8.4028 | 0.003506 | 0.004761 | 0.00456 | 5.2e-05 | 0.000986 | 0.23994 | 1.3229 | 2.5479 | 98.85 | 20.78 | 38.767 | 86.4615 | 106.2326 | 82.3392 | 0.009633 | 0.99313 | 0.023942 | 24.1077 | -0.016607 | 0.14353 | 0.00066 | 0.04503 | 0.04503 | 0.131
2019-03-21 16:02 | 1553198538.8755 | 748 | 0.20936 | 0.000626 | 3.6e-05 | 101.3228 | 101.4023 | 120.6983 | 21.2612 | 0.002637 | 0.002448 | -152.2565 | 0.24038 | 8.4041 | 0.00344 | 0.004384 | 0.004265 | 5.3e-05 | 0.000987 | 0.26952 | 1.3213 | 2.548 | 98.85 | 20.78 | 38.763 | 85.257 | 104.6325 | 80.7069 | 0.009632 | 0.99771 | 0.024004 | 24.0589 | -0.007017 | 0.143 | 0.000399 | 0.026587 | 0.026587 | 0.22107
2019-03-21 16:02 | 1553198539.8755 | 749 | 0.20934 | 0.000623 | 3.3e-05 | 101.3535 | 101.4922 | 120.4431 | 21.2612 | 0.002667 | 0.002447 | -152.2602 | 0.2405 | 8.4044 | 0.00358 | 0.004734 | 0.004537 | 5.1e-05 | 0.00099 | 0.24347 | 1.3687 | 2.5479 | 98.85 | 20.78 | 38.763 | 85.9673 | 108.0422 | 80.6836 | 0.009632 | 0.99777 | 0.023979 | 24.0327 | -0.012236 | 0.14307 | 0.000537 | 0.036635 | 0.036635 | 0.16051
2019-03-21 16:02 | 1553198540.8766 | 750 | 0.20935 | 0.000611 | 2.9e-05 | 101.2826 | 101.5038 | 120.6152 | 21.2612 | 0.002572 | 0.002435 | -152.2654 | 0.24096 | 8.4016 | 0.003456 | 0.004988 | 0.00486 | 5e-05 | 0.000994 | 0.27024 | 1.357 | 2.548 | 98.85 | 20.78 | 38.763 | 83.688 | 105.0225 | 80.3897 | 0.009632 | 0.9986 | 0.024006 | 24.0398 | -0.019517 | 0.14335 | 0.000479 | 0.032747 | 0.032747 | 0.17991
2019-03-21 16:02 | 1553198541.8758 | 751 | 0.20937 | 0.000607 | 2.5e-05 | 101.2915 | 101.3744 | 120.7246 | 21.2612 | 0.002599 | 0.00243 | -152.2623 | 0.24231 | 8.4024 | 0.00348 | 0.005145 | 0.004982 | 5.3e-05 | 0.000993 | 0.26465 | 1.3409 | 2.548 | 98.85 | 20.78 | 38.763 | 84.3488 | 105.6029 | 80.2907 | 0.009632 | 0.99888 | 0.02402 | 24.0474 | -0.007315 | 0.14415 | 0.000413 | 0.02823 | 0.02823 | 0.20987
2019-03-21 16:02 | 1553198542.8756 | 752 | 0.20941 | 0.000599 | 2.2e-05 | 101.3046 | 101.4425 | 120.9318 | 21.2612 | 0.002576 | 0.002398 | -152.2613 | 0.24136 | 8.4013 | 0.003522 | 0.00468 | 0.004517 | 5.4e-05 | 0.000995 | 0.26465 | 1.3166 | 2.548 | 98.85 | 20.78 | 38.763 | 83.7839 | 106.6363 | 79.5032 | 0.009632 | 1.0011 | 0.024068 | 24.0412 | -0.012166 | 0.14358 | 0.000129 | 0.008022 | 0.008022 | 0.73563
2019-03-21 16:02 | 1553198543.8758 | 753 | 0.20939 | 0.000588 | 1.8e-05 | 101.2316 | 101.3996 | 121.2414 | 21.2612 | 0.002544 | 0.002364 | -152.2628 | 0.24127 | 8.4033 | 0.003569 | 0.004933 | 0.004834 | 5.3e-05 | 0.000989 | 0.25371 | 1.3748 | 2.5478 | 98.85 | 20.78 | 38.765 | 83.0186 | 107.7663 | 78.6838 | 0.009633 | 1.0034 | 0.024127 | 24.0439 | -0.014829 | 0.14353 | 0.00027 | 0.018488 | 0.018488 | 0.31908
2019-03-21 16:02 | 1553198544.8757 | 754 | 0.20944 | 0.000587 | 1.5e-05 | 101.3088 | 101.4176 | 121.5728 | 21.2612 | 0.002594 | 0.002389 | -152.2458 | 0.24166 | 8.4041 | 0.00352 | 0.005443 | 0.005234 | 4.9e-05 | 0.000988 | 0.26073 | 1.3634 | 2.5479 | 98.85 | 20.78 | 38.771 | 84.2254 | 106.5812 | 79.2989 | 0.009634 | 1.0017 | 0.024139 | 24.0978 | -0.009605 | 0.14376 | 1.8e-05 | 0.000382 | 0.000382 | 15.4616
2019-03-21 16:02 | 1553198545.8768 | 755 | 0.20944 | 0.000576 | 1.2e-05 | 101.2388 | 101.4241 | 121.781 | 21.2612 | 0.002534 | 0.002351 | -152.2529 | 0.24145 | 8.4 | 0.003525 | 0.004753 | 0.004561 | 5.2e-05 | 0.000988 | 0.26626 | 1.3353 | 2.548 | 98.85 | 20.78 | 38.775 | 82.7674 | 106.6977 | 78.3829 | 0.009635 | 1.0043 | 0.024191 | 24.0871 | -0.016359 | 0.14364 | -9e-06 | -0.001345 | -0.001345 | -4.3899
2019-03-21 16:02 | 1553198546.8759 | 756 | 0.20945 | 0.00057 | 1e-05 | 101.2975 | 101.3281 | 121.5818 | 21.2612 | 0.002536 | 0.002346 | -152.2643 | 0.24128 | 8.4022 | 0.003546 | 0.005169 | 0.005001 | 5.2e-05 | 0.000991 | 0.29336 | 1.3208 | 2.5478 | 98.85 | 20.78 | 38.782 | 82.8236 | 107.2189 | 78.2729 | 0.009637 | 1.0046 | 0.024175 | 24.0636 | -0.002701 | 0.14354 | -9e-06 | -0.001264 | -0.001264 | -4.668
2019-03-21 16:02 | 1553198547.8754 | 757 | 0.20944 | 0.000561 | 8e-06 | 101.3088 | 101.4374 | 121.6505 | 21.2612 | 0.002511 | 0.002365 | -152.2663 | 0.24177 | 8.403 | 0.003505 | 0.005648 | 0.005515 | 5.3e-05 | 0.00099 | 0.26808 | 1.3141 | 2.5478 | 98.85 | 20.78 | 38.786 | 82.2378 | 106.2224 | 78.7119 | 0.009638 | 1.0034 | 0.024166 | 24.0854 | -0.011347 | 0.14383 | 7.9e-05 | 0.00529 | 0.00529 | 1.1174
2019-03-21 16:02 | 1553198548.876 | 758 | 0.20942 | 0.000552 | 5e-06 | 101.3225 | 101.4612 | 121.6109 | 21.2612 | 0.002491 | 0.002349 | -152.2555 | 0.24113 | 8.3997 | 0.003429 | 0.004377 | 0.004292 | 5.4e-05 | 0.000991 | 0.26952 | 1.3386 | 2.548 | 98.85 | 20.78 | 38.79 | 81.7544 | 104.3697 | 78.3423 | 0.009639 | 1.0044 | 0.024175 | 24.0689 | -0.012239 | 0.14345 | 0.000157 | 0.011038 | 0.011038 | 0.53413
2019-03-21 16:02 | 1553198549.876 | 759 | 0.20943 | 0.000547 | 4e-06 | 101.273 | 101.3826 | 121.7651 | 21.2612 | 0.002538 | 0.00232 | -152.2871 | 0.24144 | 8.4001 | 0.003508 | 0.004888 | 0.004606 | 5.3e-05 | 0.000995 | 0.24082 | 1.3341 | 2.5479 | 98.85 | 20.79 | 38.796 | 82.8643 | 106.2964 | 77.6251 | 0.009646 | 1.0065 | 0.024215 | 24.0595 | -0.009668 | 0.14363 | 0.000106 | 0.007436 | 0.007436 | 0.79391
2019-03-21 16:03 | 1553198550.8765 | 760 | 0.20944 | 0.000539 | 3e-06 | 101.3034 | 101.3553 | 121.5653 | 21.2612 | 0.002548 | 0.002366 | -152.2882 | 0.24124 | 8.4017 | 0.003484 | 0.0052 | 0.005091 | 5.4e-05 | 0.000994 | 0.26936 | 1.3031 | 2.548 | 98.85 | 20.79 | 38.799 | 83.1201 | 105.706 | 78.7297 | 0.009647 | 1.0033 | 0.024157 | 24.0776 | -0.00458 | 0.14351 | 5.3e-05 | 0.003654 | 0.003654 | 1.614
2019-03-21 16:03 | 1553198551.8761 | 761 | 0.20944 | 0.000533 | 2e-06 | 101.2382 | 101.3253 | 121.7911 | 21.2612 | 0.002481 | 0.002317 | -152.2892 | 0.24171 | 8.4033 | 0.003428 | 0.005198 | 0.005114 | 5.3e-05 | 0.000994 | 0.26272 | 1.3473 | 2.5479 | 98.85 | 20.79 | 38.801 | 81.4969 | 104.3462 | 77.5706 | 0.009648 | 1.0066 | 0.02422 | 24.0602 | -0.007697 | 0.14379 | 9.2e-05 | 0.006571 | 0.006571 | 0.89942
2019-03-21 16:03 | 1553198552.8752 | 762 | 0.20946 | 0.000527 | 1e-06 | 101.257 | 101.3856 | 122.1294 | 21.2612 | 0.002497 | 0.002297 | -152.2835 | 0.24147 | 8.4044 | 0.003416 | 0.004494 | 0.004257 | 5.4e-05 | 0.000997 | 0.30143 | 1.3308 | 2.5479 | 98.85 | 20.79 | 38.8 | 81.8955 | 104.06 | 77.0802 | 0.009647 | 1.008 | 0.02427 | 24.0768 | -0.011359 | 0.14365 | -1e-06 | -0.000103 | -0.000103 | -57.1929
2019-03-21 16:03 | 1553198553.8766 | 763 | 0.20945 | 0.000524 | 0 | 101.1866 | 101.3689 | 122.3392 | 21.2612 | 0.002461 | 0.002333 | -152.2803 | 0.24144 | 8.4033 | 0.0035 | 0.004712 | 0.004557 | 5.1e-05 | 0.000994 | 0.28423 | 1.3317 | 2.548 | 98.85 | 20.79 | 38.8 | 81.0134 | 106.0846 | 77.9454 | 0.009647 | 1.0056 | 0.024261 | 24.1272 | -0.0161 | 0.14363 | 5e-05 | 0.003617 | 0.003617 | 1.6319
2019-03-21 16:03 | 1553198554.876 | 764 | 0.20944 | 0.000514 | 0 | 101.2611 | 101.5157 | 122.365 | 21.2612 | 0.002456 | 0.002245 | -152.2804 | 0.24131 | 8.4034 | 0.00346 | 0.00477 | 0.004585 | 5.4e-05 | 0.001 | 0.26128 | 1.3492 | 2.548 | 98.85 | 20.79 | 38.797 | 80.9035 | 105.1207 | 75.8432 | 0.009647 | 1.0116 | 0.024337 | 24.0574 | -0.022462 | 0.14355 | 9.5e-05 | 0.006996 | 0.006996 | 0.84332
2019-03-21 16:03 | 1553198555.8754 | 765 | 0.20946 | 0.000514 | -1e-06 | 101.1673 | 101.3982 | 122.5838 | 21.2612 | 0.002474 | 0.002301 | -152.2657 | 0.2416 | 8.403 | 0.00346 | 0.005301 | 0.005083 | 5.4e-05 | 0.000995 | 0.26089 | 1.3255 | 2.548 | 98.85 | 20.79 | 38.794 | 81.3371 | 105.1098 | 77.1698 | 0.009646 | 1.0078 | 0.024312 | 24.1246 | -0.020397 | 0.14373 | 1.6e-05 | 0.001224 | 0.001224 | 4.8244
2019-03-21 16:03 | 1553198556.8764 | 766 | 0.20945 | 0.000516 | -2e-06 | 101.2167 | 101.3373 | 122.6103 | 21.2612 | 0.002492 | 0.002275 | -152.2507 | 0.24177 | 8.4034 | 0.003497 | 0.004952 | 0.004757 | 5.5e-05 | 0.000993 | 0.2688 | 1.3187 | 2.548 | 98.85 | 20.79 | 38.783 | 81.7582 | 106.0233 | 76.544 | 0.009643 | 1.0096 | 0.024337 | 24.1057 | -0.010645 | 0.14383 | 7.3e-05 | 0.005446 | 0.005446 | 1.0854
2019-03-21 16:03 | 1553198557.8755 | 767 | 0.20947 | 0.000509 | -2e-06 | 101.2385 | 101.4136 | 122.1613 | 21.2612 | 0.002429 | 0.002253 | -152.2512 | 0.24152 | 8.4036 | 0.003479 | 0.004226 | 0.004144 | 5.3e-05 | 0.000993 | 0.31653 | 1.3603 | 2.5479 | 98.85 | 20.79 | 38.777 | 80.2469 | 105.592 | 76.0202 | 0.009642 | 1.0111 | 0.02431 | 24.0434 | -0.015455 | 0.14368 | -7e-05 | -0.004974 | -0.004974 | -1.1873
2019-03-21 16:03 | 1553198558.8769 | 768 | 0.20946 | 0.000507 | -3e-06 | 101.2555 | 101.5474 | 122.1284 | 21.2612 | 0.002421 | 0.002288 | -152.2594 | 0.24139 | 8.4033 | 0.003499 | 0.004576 | 0.004368 | 5.6e-05 | 0.000992 | 0.286 | 1.3464 | 2.5478 | 98.85 | 20.79 | 38.771 | 80.0552 | 106.0672 | 76.8652 | 0.00964 | 1.0087 | 0.024278 | 24.0693 | -0.025752 | 0.1436 | 3.1e-05 | 0.002405 | 0.002405 | 2.4542
2019-03-21 16:03 | 1553198559.8756 | 769 | 0.20944 | 0.000501 | -3e-06 | 101.2036 | 101.3703 | 122.0662 | 21.2612 | 0.002423 | 0.002218 | -152.2315 | 0.24059 | 8.4034 | 0.003524 | 0.004614 | 0.004373 | 5.4e-05 | 0.000994 | 0.29479 | 1.2987 | 2.5479 | 98.85 | 20.79 | 38.76 | 80.1152 | 106.6819 | 75.1753 | 0.009637 | 1.0136 | 0.02433 | 24.005 | -0.01472 | 0.14313 | 0.000108 | 0.008062 | 0.008062 | 0.72964
2019-03-21 16:03 | 1553198560.8757 | 770 | 0.20945 | 0.000503 | -3e-06 | 101.2093 | 101.4779 | 122.511 | 21.2612 | 0.002398 | 0.002283 | -152.2197 | 0.24131 | 8.4029 | 0.003497 | 0.004443 | 0.004329 | 5.1e-05 | 0.000993 | 0.28208 | 1.3033 | 2.5479 | 98.85 | 20.79 | 38.756 | 79.5001 | 106.0169 | 76.7412 | 0.009636 | 1.009 | 0.02432 | 24.1027 | -0.023711 | 0.14356 | 5.1e-05 | 0.003877 | 0.003877 | 1.522
2019-03-21 16:03 | 1553198561.876 | 771 | 0.20945 | 0.000498 | -3e-06 | 101.1974 | 101.4875 | 122.631 | 21.2612 | 0.002404 | 0.0022 | -152.2131 | 0.24111 | 8.4022 | 0.003495 | 0.004945 | 0.00478 | 4.8e-05 | 0.000993 | 0.23883 | 1.3031 | 2.5479 | 98.85 | 20.79 | 38.749 | 79.6606 | 105.9609 | 74.7613 | 0.009635 | 1.0148 | 0.024401 | 24.0462 | -0.025604 | 0.14344 | 5.8e-05 | 0.00445 | 0.00445 | 1.3248
2019-03-21 16:03 | 1553198562.8767 | 772 | 0.20946 | 0.000496 | -3e-06 | 101.2087 | 101.1918 | 122.3772 | 21.2612 | 0.002432 | 0.002236 | -152.2356 | 0.24103 | 8.4028 | 0.003464 | 0.004377 | 0.004117 | 4.8e-05 | 0.000993 | 0.25393 | 1.3529 | 2.548 | 98.85 | 20.79 | 38.746 | 80.3282 | 105.2223 | 75.6195 | 0.009634 | 1.0123 | 0.024346 | 24.0509 | 0.001489 | 0.14339 | 2.8e-05 | 0.002234 | 0.002234 | 2.6379
2019-03-21 16:03 | 1553198563.8748 | 773 | 0.20947 | 0.000495 | -4e-06 | 101.1664 | 101.2589 | 122.0467 | 21.2612 | 0.002412 | 0.002252 | -152.2423 | 0.24178 | 8.4034 | 0.003424 | 0.004991 | 0.00484 | 5.1e-05 | 0.000994 | 0.24298 | 1.2983 | 2.5479 | 98.85 | 20.79 | 38.743 | 79.8478 | 104.2432 | 75.9967 | 0.009633 | 1.0112 | 0.0243 | 24.0314 | -0.008179 | 0.14384 | -1.3e-05 | -0.000714 | -0.000714 | -8.2753
2019-03-21 16:03 | 1553198564.8753 | 774 | 0.20943 | 0.000493 | -3e-06 | 101.1846 | 101.4044 | 121.7982 | 21.2612 | 0.002379 | 0.002241 | -152.2668 | 0.24079 | 8.4036 | 0.003463 | 0.004727 | 0.004602 | 4.9e-05 | 0.000992 | 0.27472 | 1.2738 | 2.5478 | 98.85 | 20.79 | 38.738 | 79.0432 | 105.1995 | 75.7473 | 0.009632 | 1.0119 | 0.024284 | 23.9983 | -0.019409 | 0.14325 | 0.000189 | 0.013973 | 0.013973 | 0.42135
2019-03-21 16:03 | 1553198565.8749 | 775 | 0.20945 | 0.000491 | -4e-06 | 101.203 | 101.1643 | 121.3719 | 21.2612 | 0.002429 | 0.002266 | -152.2727 | 0.24157 | 8.3991 | 0.003452 | 0.004508 | 0.004312 | 5e-05 | 0.000995 | 0.27273 | 1.322 | 2.5479 | 98.85 | 20.79 | 38.736 | 80.2525 | 104.9342 | 76.3381 | 0.009631 | 1.0102 | 0.024221 | 23.9765 | 0.003428 | 0.14371 | 8.3e-05 | 0.00628 | 0.00628 | 0.94052
2019-03-21 16:03 | 1553198566.875 | 776 | 0.20947 | 0.000488 | -3e-06 | 101.1214 | 101.2136 | 121.5757 | 21.2612 | 0.002328 | 0.002163 | -152.2872 | 0.24132 | 8.404 | 0.003415 | 0.004489 | 0.004268 | 4.9e-05 | 0.000991 | 0.25929 | 1.2967 | 2.548 | 98.85 | 20.79 | 38.734 | 77.8238 | 104.037 | 73.8539 | 0.009631 | 1.0174 | 0.024328 | 23.9112 | -0.008155 | 0.14356 | -7e-06 | -0.000317 | -0.000317 | -18.629
2019-03-21 16:03 | 1553198567.8753 | 777 | 0.20944 | 0.000483 | -4e-06 | 101.1467 | 101.3761 | 122.055 | 21.2612 | 0.002272 | 0.002231 | -152.2804 | 0.24076 | 8.4 | 0.003445 | 0.004496 | 0.004333 | 5.1e-05 | 0.000992 | 0.25946 | 1.2839 | 2.5479 | 98.85 | 20.8 | 38.73 | 76.4902 | 104.7597 | 75.5097 | 0.009636 | 1.0126 | 0.024317 | 24.0154 | -0.020261 | 0.14323 | 0.000133 | 0.009915 | 0.009915 | 0.59371
2019-03-21 16:03 | 1553198568.875 | 778 | 0.20946 | 0.000486 | -4e-06 | 101.1914 | 101.512 | 122.3373 | 21.2612 | 0.002326 | 0.002246 | -152.2988 | 0.24112 | 8.4035 | 0.003461 | 0.004566 | 0.004414 | 5.1e-05 | 0.000993 | 0.28423 | 1.287 | 2.548 | 98.85 | 20.8 | 38.728 | 77.7693 | 105.1364 | 75.8689 | 0.009635 | 1.0115 | 0.024333 | 24.0555 | -0.028292 | 0.14344 | 4.1e-05 | 0.003222 | 0.003222 | 1.83
2019-03-21 16:03 | 1553198569.8751 | 779 | 0.20944 | 0.000481 | -4e-06 | 101.1184 | 101.1891 | 122.3648 | 21.2612 | 0.002339 | 0.002174 | -152.3073 | 0.24087 | 8.4028 | 0.003506 | 0.004477 | 0.004251 | 5.3e-05 | 0.000996 | 0.29214 | 1.3071 | 2.5481 | 98.85 | 20.8 | 38.723 | 78.085 | 106.2319 | 74.1299 | 0.009634 | 1.0166 | 0.024397 | 23.9982 | -0.006252 | 0.14329 | 0.000171 | 0.012729 | 0.012729 | 0.46268
2019-03-21 16:03 | 1553198570.8755 | 780 | 0.20946 | 0.000484 | -4e-06 | 101.1533 | 101.3822 | 121.9085 | 21.2612 | 0.002367 | 0.002225 | -152.3147 | 0.24119 | 8.4028 | 0.003442 | 0.005238 | 0.005088 | 5.5e-05 | 0.000996 | 0.29607 | 1.3417 | 2.548 | 98.85 | 20.8 | 38.718 | 78.7696 | 104.6749 | 75.3475 | 0.009633 | 1.0131 | 0.024309 | 23.9954 | -0.020222 | 0.14348 | 5.8e-05 | 0.004435 | 0.004435 | 1.3296
2019-03-21 16:03 | 1553198571.8752 | 781 | 0.20944 | 0.000485 | -4e-06 | 101.1467 | 101.2794 | 121.7442 | 21.2612 | 0.002361 | 0.00221 | -152.3379 | 0.24081 | 8.4019 | 0.003462 | 0.005314 | 0.005224 | 5.5e-05 | 0.000992 | 0.26913 | 1.3163 | 2.548 | 98.85 | 20.8 | 38.713 | 78.6191 | 105.1774 | 74.9843 | 0.009632 | 1.0141 | 0.024305 | 23.9667 | -0.011723 | 0.14326 | 0.00019 | 0.014063 | 0.014063 | 0.41869
2019-03-21 16:03 | 1553198572.8763 | 782 | 0.20947 | 0.000482 | -4e-06 | 101.1753 | 101.3741 | 121.8179 | 21.2612 | 0.002378 | 0.002229 | -152.3485 | 0.24123 | 8.4036 | 0.003453 | 0.004511 | 0.004317 | 5.7e-05 | 0.001002 | 0.28224 | 1.3365 | 2.548 | 98.85 | 20.8 | 38.701 | 79.0392 | 104.9468 | 75.4403 | 0.009629 | 1.0128 | 0.024296 | 23.9897 | -0.017552 | 0.14351 | -2.1e-05 | -0.00129 | -0.00129 | -4.5725
2019-03-21 16:03 | 1553198573.8756 | 783 | 0.20945 | 0.000477 | -4e-06 | 101.1396 | 101.294 | 122.0024 | 21.2612 | 0.002316 | 0.002162 | -152.3269 | 0.24129 | 8.4022 | 0.003487 | 0.00433 | 0.004151 | 5.5e-05 | 0.001004 | 0.25553 | 1.3391 | 2.548 | 98.85 | 20.8 | 38.695 | 77.5394 | 105.7659 | 73.8383 | 0.009627 | 1.0175 | 0.024371 | 23.9526 | -0.013649 | 0.14354 | 8.1e-05 | 0.006186 | 0.006186 | 0.95368
2019-03-21 16:03 | 1553198574.8764 | 784 | 0.20949 | 0.000482 | -4e-06 | 101.0737 | 101.2933 | 122.1305 | 21.2612 | 0.002287 | 0.002163 | -152.3168 | 0.24141 | 8.4028 | 0.003451 | 0.005232 | 0.00504 | 5.5e-05 | 0.001004 | 0.26017 | 1.3497 | 2.5479 | 98.85 | 20.8 | 38.683 | 76.8479 | 104.9083 | 73.8712 | 0.009624 | 1.0174 | 0.024382 | 23.9663 | -0.019413 | 0.14361 | -0.000136 | -0.009694 | -0.009694 | -0.60886
2019-03-21 16:03 | 1553198575.8755 | 785 | 0.20947 | 0.00048 | -4e-06 | 101.1065 | 101.1782 | 122.0909 | 21.2612 | 0.002297 | 0.002132 | -152.3184 | 0.24135 | 8.403 | 0.003512 | 0.004682 | 0.004418 | 5.2e-05 | 0.001006 | 0.28584 | 1.3327 | 2.5481 | 98.85 | 20.8 | 38.675 | 77.092 | 106.3798 | 73.1149 | 0.009622 | 1.0196 | 0.024405 | 23.9363 | -0.006342 | 0.14358 | -3e-06 | 3.7e-05 | 3.7e-05 | 160.7087
2019-03-21 16:03 | 1553198576.8769 | 786 | 0.20949 | 0.000481 | -4e-06 | 101.0076 | 101.1012 | 121.9008 | 21.2612 | 0.002281 | 0.002167 | -152.3271 | 0.2417 | 8.4001 | 0.003461 | 0.004951 | 0.004793 | 5.2e-05 | 0.001004 | 0.27743 | 1.3211 | 2.5481 | 98.85 | 20.8 | 38.668 | 76.6905 | 105.1529 | 73.949 | 0.00962 | 1.0171 | 0.024357 | 23.9465 | -0.00829 | 0.14378 | -0.000162 | -0.011601 | -0.011601 | -0.50938
2019-03-21 16:03 | 1553198577.876 | 787 | 0.20947 | 0.000482 | -5e-06 | 101.0719 | 101.1227 | 121.7308 | 21.2612 | 0.002321 | 0.002137 | -152.3383 | 0.24114 | 8.4006 | 0.003534 | 0.004762 | 0.004627 | 5.1e-05 | 0.001007 | 0.30127 | 1.274 | 2.5478 | 98.85 | 20.8 | 38.654 | 77.6672 | 106.9204 | 73.2429 | 0.009617 | 1.0192 | 0.024365 | 23.9054 | -0.004491 | 0.14345 | -4e-06 | -2.9e-05 | -2.9e-05 | -201.0959
2019-03-21 16:03 | 1553198578.8757 | 788 | 0.20947 | 0.000476 | -4e-06 | 101.0639 | 101.1956 | 121.8585 | 21.2612 | 0.002302 | 0.002166 | -152.3143 | 0.2406 | 8.4033 | 0.003515 | 0.004914 | 0.004767 | 5.2e-05 | 0.001007 | 0.30237 | 1.2822 | 2.5479 | 98.85 | 20.81 | 38.646 | 77.2018 | 106.4577 | 73.9272 | 0.009621 | 1.0172 | 0.024353 | 23.9415 | -0.011648 | 0.14313 | 8e-06 | 0.000813 | 0.000813 | 7.2374
2019-03-21 16:03 | 1553198579.8751 | 789 | 0.20945 | 0.000472 | -4e-06 | 101.029 | 101.2044 | 121.9615 | 21.2612 | 0.00225 | 0.00215 | -152.3088 | 0.24146 | 8.4014 | 0.003508 | 0.004476 | 0.004238 | 5.3e-05 | 0.001006 | 0.27544 | 1.3197 | 2.5479 | 98.85 | 20.81 | 38.639 | 75.9491 | 106.2869 | 73.5538 | 0.009619 | 1.0183 | 0.024377 | 23.9388 | -0.015516 | 0.14365 | 0.000125 | 0.009434 | 0.009434 | 0.62579
2019-03-21 16:03 | 1553198580.8748 | 790 | 0.20955 | 0.000479 | -4e-06 | 100.9912 | 101.1394 | 122.0768 | 21.2612 | 0.002259 | 0.002173 | -152.3221 | 0.24208 | 8.4024 | 0.003526 | 0.004689 | 0.004512 | 5.2e-05 | 0.00101 | 0.28174 | 1.3087 | 2.548 | 98.85 | 20.81 | 38.626 | 76.1689 | 106.7236 | 74.0949 | 0.009616 | 1.0167 | 0.024369 | 23.9688 | -0.013116 | 0.14401 | -0.000523 | -0.038069 | -0.038069 | -0.15548
2019-03-21 16:03 | 1553198581.8758 | 791 | 0.20954 | 0.000476 | -4e-06 | 100.9876 | 101.1118 | 121.9132 | 21.2612 | 0.002303 | 0.002143 | -152.2991 | 0.24114 | 8.4047 | 0.003491 | 0.004826 | 0.004599 | 5.4e-05 | 0.001013 | 0.27865 | 1.347 | 2.5478 | 98.85 | 20.81 | 38.621 | 77.2376 | 105.8816 | 73.3839 | 0.009615 | 1.0188 | 0.024378 | 23.9282 | -0.010992 | 0.14346 | -0.00044 | -0.031937 | -0.031937 | -0.18461
2019-03-21 16:03 | 1553198582.8748 | 792 | 0.20936 | 0.000476 | -4e-06 | 100.9441 | 101.0852 | 121.908 | 21.2612 | 0.002329 | 0.002163 | -152.2865 | 0.24202 | 8.4043 | 0.003473 | 0.004766 | 0.004629 | 5.2e-05 | 0.001013 | 0.29264 | 1.3453 | 2.548 | 98.85 | 20.81 | 38.612 | 77.8519 | 105.4245 | 73.8563 | 0.009612 | 1.0174 | 0.024361 | 23.944 | -0.012496 | 0.14398 | 0.000656 | 0.04818 | 0.04818 | 0.12282
2019-03-21 16:03 | 1553198583.8753 | 793 | 0.20943 | 0.000477 | -4e-06 | 100.886 | 101.06 | 121.8271 | 21.2612 | 0.002311 | 0.002169 | -152.3005 | 0.24065 | 8.4035 | 0.003475 | 0.005427 | 0.005275 | 5.7e-05 | 0.001016 | 0.26272 | 1.3723 | 2.548 | 98.85 | 20.81 | 38.607 | 77.4185 | 105.4921 | 74.0053 | 0.009611 | 1.017 | 0.024347 | 23.9412 | -0.015415 | 0.14316 | 0.000202 | 0.015036 | 0.015036 | 0.39132
2019-03-21 16:03 | 1553198584.8759 | 794 | 0.20939 | 0.000477 | -3e-06 | 100.9069 | 101.058 | 121.6656 | 21.2612 | 0.002315 | 0.002172 | -152.2865 | 0.24107 | 8.402 | 0.003447 | 0.004869 | 0.004704 | 5.5e-05 | 0.001017 | 0.28191 | 1.342 | 2.5479 | 98.85 | 20.81 | 38.603 | 77.5043 | 104.7999 | 74.0848 | 0.00961 | 1.0167 | 0.024328 | 23.928 | -0.013385 | 0.14341 | 0.000489 | 0.035909 | 0.035909 | 0.16414
2019-03-21 16:03 | 1553198585.8752 | 795 | 0.20953 | 0.000478 | -4e-06 | 100.8762 | 100.9878 | 121.6308 | 21.2612 | 0.002266 | 0.002172 | -152.2846 | 0.2409 | 8.4006 | 0.003398 | 0.005148 | 0.004935 | 5.4e-05 | 0.001018 | 0.28313 | 1.3484 | 2.5482 | 98.85 | 20.81 | 38.593 | 76.339 | 103.6125 | 74.0807 | 0.009608 | 1.0167 | 0.024325 | 23.9245 | -0.009894 | 0.14331 | -0.000361 | -0.026119 | -0.026119 | -0.2255
2019-03-21 16:03 | 1553198586.875 | 796 | 0.20941 | 0.00048 | -4e-06 | 100.9262 | 101.2436 | 121.6623 | 21.2612 | 0.002274 | 0.002217 | -152.2741 | 0.24164 | 8.4054 | 0.003405 | 0.00448 | 0.004254 | 5.6e-05 | 0.001018 | 0.26969 | 1.3582 | 2.5479 | 98.85 | 20.81 | 38.587 | 76.5228 | 103.7848 | 75.1548 | 0.009606 | 1.0136 | 0.024291 | 23.9645 | -0.028082 | 0.14375 | 0.000337 | 0.024789 | 0.024789 | 0.23835
2019-03-21 16:03 | 1553198587.8765 | 797 | 0.20949 | 0.000475 | -4e-06 | 100.8097 | 100.9599 | 121.7827 | 21.2612 | 0.002293 | 0.002165 | -152.2574 | 0.242 | 8.4034 | 0.003434 | 0.004673 | 0.004567 | 5.2e-05 | 0.001019 | 0.2792 | 1.3472 | 2.548 | 98.85 | 20.82 | 38.58 | 76.9798 | 104.4868 | 73.9164 | 0.00961 | 1.0172 | 0.024346 | 23.9337 | -0.013313 | 0.14397 | -0.000104 | -0.007365 | -0.007365 | -0.80337
2019-03-21 16:03 | 1553198588.8751 | 798 | 0.20948 | 0.000478 | -5e-06 | 100.9444 | 101.0876 | 121.6778 | 21.2612 | 0.002295 | 0.00215 | -152.2597 | 0.24044 | 8.4017 | 0.003427 | 0.004677 | 0.004476 | 5.5e-05 | 0.00102 | 0.27273 | 1.3404 | 2.5479 | 98.85 | 20.82 | 38.577 | 77.0413 | 104.3154 | 73.5476 | 0.00961 | 1.0183 | 0.024348 | 23.9107 | -0.01268 | 0.14304 | -5.7e-05 | -0.003911 | -0.003911 | -1.5033
2019-03-21 16:03 | 1553198589.8755 | 799 | 0.2095 | 0.000473 | -4e-06 | 100.8276 | 100.9919 | 121.5369 | 21.2612 | 0.002214 | 0.002059 | -152.2692 | 0.24158 | 8.4027 | 0.003446 | 0.005044 | 0.004945 | 5.7e-05 | 0.001021 | 0.28335 | 1.3337 | 2.5478 | 98.85 | 20.82 | 38.573 | 75.0794 | 104.7927 | 71.3577 | 0.009609 | 1.0248 | 0.024412 | 23.8213 | -0.014563 | 0.14372 | -0.000171 | -0.012279 | -0.012279 | -0.48104
2019-03-21 16:03 | 1553198590.8767 | 800 | 0.20948 | 0.000475 | -4e-06 | 100.911 | 101.0641 | 121.6937 | 21.2612 | 0.002268 | 0.00213 | -152.2519 | 0.24153 | 8.4013 | 0.003463 | 0.004809 | 0.004599 | 5.6e-05 | 0.001014 | 0.29894 | 1.3147 | 2.5477 | 98.85 | 20.82 | 38.568 | 76.3943 | 105.1869 | 73.0714 | 0.009607 | 1.0197 | 0.024367 | 23.8958 | -0.013558 | 0.14369 | -8.9e-05 | -0.006272 | -0.006272 | -0.94163
2019-03-21 16:03 | 1553198591.8763 | 801 | 0.20948 | 0.000475 | -4e-06 | 100.8363 | 100.7705 | 121.9993 | 21.2612 | 0.002231 | 0.002087 | -152.2649 | 0.24189 | 8.4039 | 0.003405 | 0.005095 | 0.004959 | 5.4e-05 | 0.001015 | 0.29695 | 1.3693 | 2.548 | 98.85 | 20.82 | 38.565 | 75.4962 | 103.795 | 72.0368 | 0.009607 | 1.0228 | 0.024434 | 23.8901 | 0.005832 | 0.1439 | -6.4e-05 | -0.004478 | -0.004478 | -1.3208
2019-03-21 16:03 | 1553198592.8758 | 802 | 0.20948 | 0.000477 | -5e-06 | 100.8115 | 100.8032 | 122.2115 | 21.2612 | 0.002261 | 0.002129 | -152.2617 | 0.24152 | 8.4033 | 0.003489 | 0.004761 | 0.004638 | 5.1e-05 | 0.001016 | 0.26897 | 1.3204 | 2.548 | 98.85 | 20.82 | 38.563 | 76.2071 | 105.821 | 73.0519 | 0.009606 | 1.0198 | 0.024419 | 23.946 | 0.000737 | 0.14368 | -8.7e-05 | -0.006086 | -0.006086 | -0.97029
2019-03-21 16:03 | 1553198593.8754 | 803 | 0.20949 | 0.000474 | -4e-06 | 100.8681 | 101.1166 | 122.4197 | 21.2612 | 0.002196 | 0.00213 | -152.2622 | 0.24067 | 8.405 | 0.003434 | 0.005038 | 0.004828 | 5e-05 | 0.001017 | 0.31564 | 1.2867 | 2.5479 | 98.85 | 20.82 | 38.558 | 74.6638 | 104.4899 | 73.0627 | 0.009605 | 1.0197 | 0.02444 | 23.9667 | -0.022002 | 0.14317 | -0.000123 | -0.008821 | -0.008821 | -0.66706
2019-03-21 16:03 | 1553198594.8759 | 804 | 0.2095 | 0.000468 | -4e-06 | 100.7621 | 100.9735 | 122.6223 | 21.2612 | 0.002262 | 0.002083 | -152.2554 | 0.24013 | 8.4046 | 0.003313 | 0.004687 | 0.004615 | 5.2e-05 | 0.001015 | 0.26626 | 1.2997 | 2.5479 | 98.85 | 20.81 | 38.555 | 76.2353 | 101.5496 | 71.9394 | 0.009598 | 1.0231 | 0.0245 | 23.9476 | -0.018751 | 0.14285 | -0.000207 | -0.014977 | -0.014977 | -0.39201
2019-03-21 16:03 | 1553198595.8748 | 805 | 0.20947 | 0.000476 | -5e-06 | 100.788 | 100.809 | 123.0176 | 21.2612 | 0.002217 | 0.002118 | -152.2585 | 0.24084 | 8.4011 | 0.003415 | 0.004735 | 0.004456 | 5.2e-05 | 0.001016 | 0.26969 | 1.3207 | 2.548 | 98.85 | 20.81 | 38.555 | 75.158 | 104.0268 | 72.7766 | 0.009598 | 1.0206 | 0.024509 | 24.0152 | -0.001866 | 0.14328 | 9e-06 | 0.000986 | 0.000986 | 5.9703
2019-03-21 16:03 | 1553198596.8756 | 806 | 0.20943 | 0.000474 | -4e-06 | 100.7749 | 100.9773 | 123.1314 | 21.2612 | 0.00218 | 0.002084 | -152.2609 | 0.24216 | 8.4018 | 0.003479 | 0.004777 | 0.004656 | 5.4e-05 | 0.001017 | 0.28351 | 1.3397 | 2.5481 | 98.85 | 20.81 | 38.553 | 74.271 | 105.5833 | 71.9627 | 0.009598 | 1.023 | 0.02455 | 23.9981 | -0.017945 | 0.14406 | 0.000218 | 0.016345 | 0.016345 | 0.36226
2019-03-21 16:03 | 1553198597.8746 | 807 | 0.20945 | 0.000473 | -4e-06 | 100.7999 | 100.9057 | 122.9469 | 21.2612 | 0.002204 | 0.002068 | -152.2627 | 0.24083 | 8.4009 | 0.003445 | 0.004991 | 0.004791 | 5.2e-05 | 0.001011 | 0.27759 | 1.3407 | 2.5481 | 98.85 | 20.81 | 38.552 | 74.8401 | 104.7573 | 71.5879 | 0.009597 | 1.0241 | 0.024545 | 23.9671 | -0.009386 | 0.14327 | 8.5e-05 | 0.00654 | 0.00654 | 0.90044
2019-03-21 16:03 | 1553198598.8761 | 808 | 0.20946 | 0.000474 | -4e-06 | 100.796 | 100.9449 | 122.5906 | 21.2612 | 0.002207 | 0.00207 | -152.2828 | 0.24107 | 8.4027 | 0.003436 | 0.005164 | 0.00502 | 5.1e-05 | 0.001015 | 0.28479 | 1.3301 | 2.5479 | 98.85 | 20.82 | 38.553 | 74.9249 | 104.5379 | 71.6261 | 0.009604 | 1.024 | 0.024508 | 23.9336 | -0.013201 | 0.14341 | 7.9e-05 | 0.006057 | 0.006057 | 0.97318
2019-03-21 16:03 | 1553198599.8747 | 809 | 0.20945 | 0.000474 | -4e-06 | 100.7898 | 100.8192 | 122.8042 | 21.2612 | 0.002244 | 0.002136 | -152.2967 | 0.24126 | 8.403 | 0.003466 | 0.005049 | 0.004895 | 5.2e-05 | 0.001017 | 0.25586 | 1.3191 | 2.548 | 98.85 | 20.82 | 38.553 | 75.8011 | 105.2615 | 73.2131 | 0.009604 | 1.0193 | 0.024473 | 24.0095 | -0.002614 | 0.14353 | 0.000116 | 0.008774 | 0.008774 | 0.6723
2019-03-21 16:03 | 1553198600.8757 | 810 | 0.20947 | 0.000476 | -4e-06 | 100.7227 | 100.808 | 122.8096 | 21.2612 | 0.002186 | 0.002066 | -152.2985 | 0.24098 | 8.4005 | 0.003475 | 0.004561 | 0.004314 | 4.9e-05 | 0.001021 | 0.26681 | 1.3206 | 2.548 | 98.85 | 20.82 | 38.553 | 74.4137 | 105.4904 | 71.5425 | 0.009604 | 1.0242 | 0.024533 | 23.9521 | -0.007568 | 0.14336 | -4.3e-05 | -0.002924 | -0.002924 | -2.0148
2019-03-21 16:03 | 1553198601.8752 | 811 | 0.2095 | 0.000472 | -4e-06 | 100.7096 | 100.9255 | 122.6713 | 21.2612 | 0.002159 | 0.00209 | -152.2967 | 0.24113 | 8.4035 | 0.003457 | 0.004776 | 0.004648 | 5.4e-05 | 0.001023 | 0.23994 | 1.3041 | 2.548 | 98.85 | 20.82 | 38.553 | 73.7792 | 105.0414 | 72.1117 | 0.009604 | 1.0225 | 0.024498 | 23.9584 | -0.019151 | 0.14345 | -0.000196 | -0.014169 | -0.014169 | -0.41611
2019-03-21 16:03 | 1553198602.8759 | 812 | 0.20948 | 0.000474 | -4e-06 | 100.6813 | 100.8213 | 122.2686 | 21.2612 | 0.002195 | 0.002036 | -152.2921 | 0.2413 | 8.4013 | 0.003545 | 0.004133 | 0.003992 | 5.3e-05 | 0.00102 | 0.27201 | 1.3402 | 2.548 | 98.85 | 20.82 | 38.554 | 74.6429 | 107.1835 | 70.8067 | 0.009604 | 1.0264 | 0.024505 | 23.8737 | -0.012424 | 0.14355 | -7.8e-05 | -0.005485 | -0.005485 | -1.0756
2019-03-21 16:03 | 1553198603.875 | 813 | 0.2095 | 0.000472 | -4e-06 | 100.6366 | 100.666 | 122.3838 | 21.2612 | 0.002155 | 0.002106 | -152.2902 | 0.24207 | 8.4044 | 0.003515 | 0.004593 | 0.004405 | 5.2e-05 | 0.001019 | 0.26128 | 1.3528 | 2.5479 | 98.85 | 20.82 | 38.555 | 73.6684 | 106.4609 | 72.4938 | 0.009604 | 1.0214 | 0.024456 | 23.9435 | -0.002607 | 0.14401 | -0.00018 | -0.012973 | -0.012973 | -0.45626
2019-03-21 16:03 | 1553198604.8753 | 814 | 0.20947 | 0.000477 | -4e-06 | 100.6703 | 100.8036 | 122.4082 | 21.2612 | 0.002194 | 0.002097 | -152.3141 | 0.24141 | 8.4024 | 0.003517 | 0.005089 | 0.004921 | 5.2e-05 | 0.001021 | 0.28047 | 1.323 | 2.548 | 98.85 | 20.82 | 38.556 | 74.6155 | 106.5152 | 72.2761 | 0.009604 | 1.0221 | 0.024466 | 23.9383 | -0.011833 | 0.14361 | -3.5e-05 | -0.002339 | -0.002339 | -2.5241
2019-03-21 16:03 | 1553198605.875 | 815 | 0.20949 | 0.00047 | -4e-06 | 100.6313 | 100.7102 | 122.1852 | 21.2612 | 0.002191 | 0.002014 | -152.316 | 0.24129 | 8.4028 | 0.00346 | 0.00464 | 0.004468 | 4.9e-05 | 0.001017 | 0.26985 | 1.2813 | 2.5479 | 98.85 | 20.82 | 38.556 | 74.536 | 105.1287 | 70.2874 | 0.009604 | 1.028 | 0.024515 | 23.8476 | -0.007017 | 0.14355 | -0.000116 | -0.008293 | -0.008293 | -0.71141
2019-03-21 16:03 | 1553198606.8751 | 816 | 0.20949 | 0.000472 | -4e-06 | 100.6733 | 100.8812 | 121.8714 | 21.2612 | 0.002232 | 0.002046 | -152.3109 | 0.24136 | 8.4007 | 0.003458 | 0.004464 | 0.004295 | 5e-05 | 0.001021 | 0.30989 | 1.28 | 2.5479 | 98.85 | 20.82 | 38.557 | 75.5167 | 105.0775 | 71.0419 | 0.009605 | 1.0257 | 0.024456 | 23.8431 | -0.018456 | 0.14358 | -9.7e-05 | -0.006859 | -0.006859 | -0.86041
2019-03-21 16:03 | 1553198607.8761 | 817 | 0.20951 | 0.000471 | -4e-06 | 100.6063 | 100.7409 | 121.7848 | 21.2612 | 0.002199 | 0.002058 | -152.3324 | 0.24136 | 8.4023 | 0.003428 | 0.004595 | 0.004442 | 4.6e-05 | 0.00102 | 0.27295 | 1.3478 | 2.5479 | 98.85 | 20.82 | 38.555 | 74.7255 | 104.3501 | 71.3486 | 0.009604 | 1.0248 | 0.024437 | 23.8452 | -0.011963 | 0.14358 | -0.000265 | -0.019185 | -0.019185 | -0.3076
2019-03-21 16:03 | 1553198608.8752 | 818 | 0.20951 | 0.000474 | -4e-06 | 100.6173 | 100.7246 | 122.1289 | 21.2612 | 0.002161 | 0.00209 | -152.3231 | 0.2406 | 8.4002 | 0.00344 | 0.004939 | 0.004807 | 4.8e-05 | 0.001017 | 0.2896 | 1.291 | 2.5479 | 98.85 | 20.82 | 38.556 | 73.8063 | 104.6283 | 72.1204 | 0.009604 | 1.0225 | 0.024444 | 23.9056 | -0.009532 | 0.14314 | -0.000247 | -0.017879 | -0.017879 | -0.32904
2019-03-21 16:03 | 1553198609.8753 | 819 | 0.20949 | 0.000473 | -5e-06 | 100.6113 | 100.5277 | 122.3026 | 21.2612 | 0.002168 | 0.002113 | -152.2998 | 0.24193 | 8.4036 | 0.003448 | 0.004782 | 0.004629 | 5e-05 | 0.001018 | 0.26233 | 1.3198 | 2.548 | 98.85 | 20.82 | 38.554 | 73.9834 | 104.8203 | 72.673 | 0.009604 | 1.0209 | 0.024442 | 23.9418 | 0.007437 | 0.14392 | -0.000134 | -0.009551 | -0.009551 | -0.61933
2019-03-21 16:04 | 1553198610.8754 | 820 | 0.2095 | 0.000475 | -4e-06 | 100.5678 | 100.7375 | 122.3441 | 21.2612 | 0.002215 | 0.002101 | -152.3073 | 0.24137 | 8.4027 | 0.003398 | 0.005112 | 0.004922 | 4.7e-05 | 0.001017 | 0.26288 | 1.3466 | 2.5481 | 98.85 | 20.82 | 38.552 | 75.1066 | 103.6118 | 72.3854 | 0.009603 | 1.0217 | 0.024456 | 23.9359 | -0.015079 | 0.14359 | -0.000183 | -0.013182 | -0.013182 | -0.44772
2019-03-21 16:04 | 1553198611.8754 | 821 | 0.2095 | 0.000471 | -4e-06 | 100.5866 | 100.7099 | 122.1615 | 21.2612 | 0.002195 | 0.002064 | -152.3153 | 0.24123 | 8.4005 | 0.003422 | 0.004229 | 0.004105 | 4.9e-05 | 0.001019 | 0.2724 | 1.3229 | 2.5479 | 98.85 | 20.82 | 38.55 | 74.6335 | 104.2054 | 71.474 | 0.009603 | 1.0244 | 0.02447 | 23.8864 | -0.01096 | 0.14351 | -0.000177 | -0.012715 | -0.012715 | -0.46388
2019-03-21 16:04 | 1553198612.8754 | 822 | 0.20947 | 0.000471 | -4e-06 | 100.6125 | 100.6684 | 122.297 | 21.2612 | 0.002233 | 0.002038 | -152.3255 | 0.24157 | 8.4014 | 0.003366 | 0.004486 | 0.004245 | 4.9e-05 | 0.00102 | 0.30287 | 1.3113 | 2.5479 | 98.85 | 20.82 | 38.549 | 75.557 | 102.8324 | 70.8603 | 0.009603 | 1.0263 | 0.024506 | 23.8784 | -0.004964 | 0.14371 | -3.6e-05 | -0.002364 | -0.002364 | -2.4989
2019-03-21 16:04 | 1553198613.8755 | 823 | 0.20947 | 0.000473 | -4e-06 | 100.5532 | 100.7838 | 121.7717 | 21.2612 | 0.002207 | 0.002091 | -152.3383 | 0.24129 | 8.4018 | 0.003363 | 0.004959 | 0.004806 | 4.9e-05 | 0.001023 | 0.28512 | 1.3126 | 2.548 | 98.85 | 20.82 | 38.547 | 74.9187 | 102.7796 | 72.1305 | 0.009602 | 1.0225 | 0.024408 | 23.871 | -0.020489 | 0.14354 | 1e-05 | 0.001028 | 0.001028 | 5.7371
2019-03-21 16:04 | 1553198614.8745 | 824 | 0.20943 | 0.000475 | -4e-06 | 100.5216 | 100.7222 | 121.8428 | 21.2612 | 0.002163 | 0.002016 | -152.3376 | 0.24134 | 8.4058 | 0.003518 | 0.0047 | 0.004476 | 5.1e-05 | 0.001024 | 0.28584 | 1.3192 | 2.5479 | 98.85 | 20.82 | 38.545 | 73.8688 | 106.5237 | 70.3263 | 0.009602 | 1.0279 | 0.024479 | 23.8155 | -0.017826 | 0.14357 | 0.000236 | 0.017601 | 0.017601 | 0.33526
2019-03-21 16:04 | 1553198615.8745 | 825 | 0.20943 | 0.000473 | -4e-06 | 100.5282 | 100.6084 | 121.8226 | 21.2612 | 0.002142 | 0.002049 | -152.3254 | 0.2418 | 8.4022 | 0.003496 | 0.004603 | 0.004506 | 4.9e-05 | 0.001023 | 0.2589 | 1.3021 | 2.548 | 98.85 | 20.82 | 38.542 | 73.3496 | 105.9854 | 71.1277 | 0.009601 | 1.0255 | 0.024448 | 23.8413 | -0.007135 | 0.14385 | 0.000239 | 0.017775 | 0.017775 | 0.3326
2019-03-21 16:04 | 1553198616.8756 | 826 | 0.20945 | 0.000471 | -4e-06 | 100.6131 | 100.7399 | 122.381 | 21.2612 | 0.002144 | 0.002061 | -152.337 | 0.24116 | 8.4011 | 0.003467 | 0.004476 | 0.004307 | 4.9e-05 | 0.00102 | 0.26897 | 1.3281 | 2.548 | 98.85 | 20.82 | 38.54 | 73.4003 | 105.2882 | 71.4038 | 0.0096 | 1.0246 | 0.024495 | 23.9054 | -0.011264 | 0.14346 | 0.000122 | 0.009209 | 0.009209 | 0.64028
2019-03-21 16:04 | 1553198617.8761 | 827 | 0.20944 | 0.000475 | -4e-06 | 100.4763 | 100.5056 | 122.8352 | 21.2612 | 0.002132 | 0.002041 | -152.323 | 0.24227 | 8.4035 | 0.003463 | 0.004629 | 0.004439 | 5e-05 | 0.001023 | 0.27815 | 1.3042 | 2.548 | 98.85 | 20.82 | 38.537 | 73.1149 | 105.1907 | 70.9422 | 0.0096 | 1.026 | 0.024556 | 23.9337 | -0.0026 | 0.14413 | 0.000152 | 0.011454 | 0.011454 | 0.51718
2019-03-21 16:04 | 1553198618.8747 | 828 | 0.20945 | 0.000472 | -4e-06 | 100.5845 | 100.6571 | 122.9018 | 21.2612 | 0.002146 | 0.002046 | -152.3278 | 0.24109 | 8.4017 | 0.003474 | 0.005243 | 0.005095 | 5e-05 | 0.001023 | 0.30326 | 1.3465 | 2.5481 | 98.85 | 20.82 | 38.535 | 73.4455 | 105.4606 | 71.0412 | 0.009599 | 1.0257 | 0.02456 | 23.9437 | -0.006455 | 0.14342 | 0.000123 | 0.009285 | 0.009285 | 0.63489
2019-03-21 16:04 | 1553198619.8748 | 829 | 0.20944 | 0.000473 | -4e-06 | 100.5178 | 100.6353 | 122.5671 | 21.2612 | 0.002137 | 0.00208 | -152.3402 | 0.24143 | 8.4025 | 0.003419 | 0.004597 | 0.00443 | 5.1e-05 | 0.001027 | 0.26609 | 1.3221 | 2.5479 | 98.85 | 20.82 | 38.531 | 73.2343 | 104.1182 | 71.8818 | 0.009598 | 1.0232 | 0.024496 | 23.9402 | -0.010455 | 0.14363 | 0.00019 | 0.014173 | 0.014173 | 0.41651
2019-03-21 16:04 | 1553198620.8758 | 830 | 0.20948 | 0.000476 | -4e-06 | 100.4903 | 100.5866 | 122.4308 | 21.2612 | 0.002177 | 0.002075 | -152.3336 | 0.24138 | 8.4019 | 0.003399 | 0.004165 | 0.003954 | 5.3e-05 | 0.001025 | 0.29767 | 1.3497 | 2.5479 | 98.85 | 20.82 | 38.529 | 74.1914 | 103.6387 | 71.7508 | 0.009598 | 1.0236 | 0.024487 | 23.9223 | -0.008565 | 0.1436 | -7e-05 | -0.004885 | -0.004885 | -1.2081
2019-03-21 16:04 | 1553198621.8749 | 831 | 0.20945 | 0.000478 | -5e-06 | 100.5374 | 100.7099 | 122.431 | 21.2612 | 0.002169 | 0.00203 | -152.3259 | 0.24165 | 8.4027 | 0.003387 | 0.00505 | 0.004893 | 5e-05 | 0.001027 | 0.31454 | 1.314 | 2.5479 | 98.85 | 20.82 | 38.527 | 74.0105 | 103.3572 | 70.6757 | 0.009597 | 1.0268 | 0.024526 | 23.885 | -0.015332 | 0.14376 | 0.000113 | 0.008607 | 0.008607 | 0.68649
2019-03-21 16:04 | 1553198622.8753 | 832 | 0.20946 | 0.000471 | -4e-06 | 100.4763 | 100.6598 | 122.6272 | 21.2612 | 0.002216 | 0.002041 | -152.3376 | 0.24168 | 8.4025 | 0.003422 | 0.004648 | 0.004486 | 4.8e-05 | 0.001026 | 0.27759 | 1.2892 | 2.5481 | 98.85 | 20.82 | 38.525 | 75.1291 | 104.1984 | 70.9352 | 0.009597 | 1.026 | 0.024536 | 23.9132 | -0.01632 | 0.14377 | 7.6e-05 | 0.005845 | 0.005845 | 1.011
2019-03-21 16:04 | 1553198623.876 | 833 | 0.20948 | 0.000473 | -4e-06 | 100.5103 | 100.6207 | 123.2821 | 21.2612 | 0.002132 | 0.002097 | -152.3428 | 0.24199 | 8.4001 | 0.003444 | 0.004716 | 0.004514 | 5.1e-05 | 0.001025 | 0.28064 | 1.2893 | 2.548 | 98.85 | 20.82 | 38.523 | 73.1266 | 104.7219 | 72.2809 | 0.009596 | 1.022 | 0.024553 | 24.0238 | -0.009816 | 0.14396 | -7.1e-05 | -0.005018 | -0.005018 | -1.1791
2019-03-21 16:04 | 1553198624.875 | 834 | 0.20948 | 0.000475 | -4e-06 | 100.4737 | 100.6738 | 123.4089 | 21.2612 | 0.002109 | 0.00201 | -152.3437 | 0.24183 | 8.4025 | 0.003397 | 0.004891 | 0.004733 | 5.1e-05 | 0.001025 | 0.27113 | 1.3125 | 2.5479 | 98.85 | 20.82 | 38.52 | 72.5654 | 103.5876 | 70.1814 | 0.009595 | 1.0283 | 0.024641 | 23.963 | -0.0178 | 0.14386 | -8.6e-05 | -0.006074 | -0.006074 | -0.97351
2019-03-21 16:04 | 1553198625.8745 | 835 | 0.20943 | 0.000469 | -4e-06 | 100.4516 | 100.6646 | 123.6851 | 21.2612 | 0.00216 | 0.002069 | -152.3355 | 0.24191 | 8.4041 | 0.003467 | 0.005218 | 0.004991 | 5.2e-05 | 0.001024 | 0.26792 | 1.2923 | 2.548 | 98.85 | 20.82 | 38.516 | 73.7806 | 105.2994 | 71.6057 | 0.009594 | 1.024 | 0.024618 | 24.0395 | -0.018945 | 0.14391 | 0.000252 | 0.018885 | 0.018885 | 0.31319
2019-03-21 16:04 | 1553198626.8746 | 836 | 0.20948 | 0.000464 | -4e-06 | 100.4406 | 100.5921 | 123.3661 | 21.2612 | 0.002146 | 0.002025 | -152.3364 | 0.24248 | 8.4035 | 0.003501 | 0.005324 | 0.005136 | 5e-05 | 0.001027 | 0.25409 | 1.2951 | 2.548 | 98.85 | 20.82 | 38.512 | 73.4628 | 106.1035 | 70.5385 | 0.009593 | 1.0272 | 0.024624 | 23.9713 | -0.013479 | 0.14425 | -6e-05 | -0.004176 | -0.004176 | -1.4197
2019-03-21 16:04 | 1553198627.8752 | 837 | 0.20946 | 0.000467 | -4e-06 | 100.5035 | 100.5222 | 123.0727 | 21.2612 | 0.002198 | 0.002049 | -152.3456 | 0.24155 | 8.4016 | 0.00347 | 0.004914 | 0.004785 | 5e-05 | 0.001022 | 0.25912 | 1.3193 | 2.5482 | 98.85 | 20.82 | 38.508 | 74.7152 | 105.3694 | 71.1277 | 0.009592 | 1.0255 | 0.024574 | 23.9633 | -0.001672 | 0.1437 | 5e-05 | 0.003917 | 0.003917 | 1.508
2019-03-21 16:04 | 1553198628.8742 | 838 | 0.20949 | 0.00047 | -4e-06 | 100.4835 | 100.3863 | 122.6836 | 21.2612 | 0.002082 | 0.002027 | -152.3457 | 0.24133 | 8.4044 | 0.003475 | 0.004837 | 0.004672 | 5.2e-05 | 0.001026 | 0.25514 | 1.2907 | 2.5479 | 98.85 | 20.82 | 38.507 | 71.9183 | 105.4873 | 70.5882 | 0.009592 | 1.0271 | 0.024554 | 23.9066 | 0.008652 | 0.14356 | -9.8e-05 | -0.006977 | -0.006977 | -0.84567
2019-03-21 16:04 | 1553198629.8757 | 839 | 0.20947 | 0.000467 | -5e-06 | 100.4719 | 100.6479 | 122.7082 | 21.2612 | 0.002146 | 0.00204 | -152.3434 | 0.24154 | 8.3999 | 0.003415 | 0.004544 | 0.004349 | 5.1e-05 | 0.001026 | 0.25282 | 1.2996 | 2.548 | 98.85 | 20.82 | 38.503 | 73.4524 | 104.0268 | 70.9134 | 0.009591 | 1.0261 | 0.024545 | 23.9203 | -0.015659 | 0.14369 | 4e-06 | 0.000559 | 0.000559 | 10.5737
2019-03-21 16:04 | 1553198630.8744 | 840 | 0.20944 | 0.000465 | -4e-06 | 100.4784 | 100.7521 | 122.8511 | 21.2612 | 0.002138 | 0.001977 | -152.3588 | 0.24093 | 8.4014 | 0.003422 | 0.00443 | 0.004242 | 4.6e-05 | 0.001025 | 0.28456 | 1.2874 | 2.5481 | 98.85 | 20.82 | 38.499 | 73.2708 | 104.2093 | 69.3834 | 0.00959 | 1.0307 | 0.024614 | 23.881 | -0.024332 | 0.14333 | 0.000187 | 0.014075 | 0.014075 | 0.41854
2019-03-21 16:04 | 1553198631.8744 | 841 | 0.20945 | 0.000468 | -4e-06 | 100.4847 | 100.6397 | 123.0037 | 21.2612 | 0.002114 | 0.002013 | -152.3404 | 0.24053 | 8.4017 | 0.003454 | 0.005294 | 0.00513 | 5.2e-05 | 0.001024 | 0.26537 | 1.2846 | 2.5479 | 98.85 | 20.82 | 38.497 | 72.6901 | 104.9665 | 70.2578 | 0.00959 | 1.0281 | 0.024598 | 23.9263 | -0.013792 | 0.14309 | 0.000131 | 0.009934 | 0.009934 | 0.59198
2019-03-21 16:04 | 1553198632.8745 | 842 | 0.20949 | 0.000471 | -4e-06 | 100.4367 | 100.5457 | 122.9603 | 21.2612 | 0.002106 | 0.002067 | -152.3446 | 0.24137 | 8.4011 | 0.003417 | 0.004596 | 0.004501 | 4.8e-05 | 0.001024 | 0.28008 | 1.3083 | 2.5479 | 98.85 | 20.82 | 38.493 | 72.5046 | 104.0678 | 71.5644 | 0.009589 | 1.0242 | 0.024547 | 23.9675 | -0.009704 | 0.14359 | -0.000106 | -0.007561 | -0.007561 | -0.78056
2019-03-21 16:04 | 1553198633.8745 | 843 | 0.20949 | 0.000466 | -4e-06 | 100.451 | 100.6128 | 122.4561 | 21.2612 | 0.002084 | 0.001982 | -152.3438 | 0.24147 | 8.4013 | 0.003473 | 0.004224 | 0.004033 | 5.2e-05 | 0.001028 | 0.2756 | 1.3027 | 2.5479 | 98.85 | 20.82 | 38.491 | 71.9565 | 105.4385 | 69.5036 | 0.009588 | 1.0303 | 0.02457 | 23.8467 | -0.014397 | 0.14365 | -0.000137 | -0.009852 | -0.009852 | -0.59927
2019-03-21 16:04 | 1553198634.8747 | 844 | 0.20947 | 0.000468 | -4e-06 | 100.4629 | 100.5962 | 122.596 | 21.2612 | 0.002105 | 0.00204 | -152.3574 | 0.2424 | 8.4023 | 0.00341 | 0.004869 | 0.004649 | 5.2e-05 | 0.001024 | 0.27544 | 1.3229 | 2.5479 | 98.85 | 20.82 | 38.49 | 72.4719 | 103.913 | 70.9064 | 0.009588 | 1.0261 | 0.024534 | 23.9092 | -0.011853 | 0.1442 | 1e-06 | 0.000325 | 0.000325 | 18.2489
2019-03-21 16:04 | 1553198635.8746 | 845 | 0.20949 | 0.00047 | -4e-06 | 100.4623 | 100.4586 | 122.5291 | 21.2612 | 0.002145 | 0.00207 | -152.3874 | 0.24159 | 8.4027 | 0.003393 | 0.004912 | 0.004756 | 5.1e-05 | 0.001028 | 0.28584 | 1.3296 | 2.5479 | 98.85 | 20.82 | 38.487 | 73.4222 | 103.4972 | 71.6292 | 0.009587 | 1.024 | 0.024501 | 23.9277 | 0.000337 | 0.14372 | -9.7e-05 | -0.006878 | -0.006878 | -0.85886
2019-03-21 16:04 | 1553198636.8753 | 846 | 0.20946 | 0.000472 | -5e-06 | 100.4293 | 100.4337 | 122.4552 | 21.2612 | 0.002143 | 0.00201 | -152.3803 | 0.24155 | 8.4014 | 0.00343 | 0.004025 | 0.003846 | 4.9e-05 | 0.001031 | 0.28174 | 1.3142 | 2.5479 | 98.85 | 20.82 | 38.485 | 73.3739 | 104.3831 | 70.1967 | 0.009587 | 1.0282 | 0.024545 | 23.8708 | -0.000394 | 0.1437 | 7.3e-05 | 0.005642 | 0.005642 | 1.0468
2019-03-21 16:04 | 1553198637.8758 | 847 | 0.20949 | 0.000473 | -4e-06 | 100.4737 | 100.3506 | 122.2458 | 21.2612 | 0.002118 | 0.002027 | -152.3803 | 0.24081 | 8.4013 | 0.003387 | 0.004995 | 0.004761 | 5.4e-05 | 0.001031 | 0.30127 | 1.3193 | 2.548 | 98.85 | 20.82 | 38.485 | 72.7775 | 103.3463 | 70.607 | 0.009587 | 1.027 | 0.024509 | 23.8646 | 0.010964 | 0.14326 | -0.000104 | -0.007408 | -0.007408 | -0.79478
2019-03-21 16:04 | 1553198638.8758 | 848 | 0.20945 | 0.000469 | -4e-06 | 100.4546 | 100.5713 | 122.354 | 21.2612 | 0.002097 | 0.00206 | -152.3892 | 0.24194 | 8.4079 | 0.003432 | 0.004724 | 0.004657 | 5.3e-05 | 0.001025 | 0.25752 | 1.3087 | 2.5478 | 98.85 | 20.82 | 38.484 | 72.2722 | 104.442 | 71.3827 | 0.009586 | 1.0247 | 0.024493 | 23.9021 | -0.010384 | 0.14393 | 0.000107 | 0.008119 | 0.008119 | 0.72862
2019-03-21 16:04 | 1553198639.8758 | 849 | 0.20949 | 0.000471 | -5e-06 | 100.5738 | 100.8209 | 122.5969 | 21.2612 | 0.002137 | 0.002031 | -152.3818 | 0.24075 | 8.402 | 0.003378 | 0.004656 | 0.004397 | 4.9e-05 | 0.001028 | 0.27079 | 1.2722 | 2.5481 | 98.85 | 20.82 | 38.483 | 73.2444 | 103.1223 | 70.7045 | 0.009586 | 1.0267 | 0.024541 | 23.9022 | -0.021955 | 0.14322 | -0.000127 | -0.009071 | -0.009071 | -0.64897
2019-03-21 16:04 | 1553198640.8758 | 850 | 0.20948 | 0.000467 | -4e-06 | 100.4889 | 100.4695 | 122.5106 | 21.2612 | 0.00214 | 0.002042 | -152.3663 | 0.24181 | 8.4046 | 0.003488 | 0.004578 | 0.004449 | 5.6e-05 | 0.001031 | 0.28567 | 1.2985 | 2.5479 | 98.85 | 20.82 | 38.481 | 73.3169 | 105.8045 | 70.9617 | 0.009586 | 1.026 | 0.024523 | 23.9027 | 0.001728 | 0.14385 | -3.5e-05 | -0.002326 | -0.002326 | -2.5418
2019-03-21 16:04 | 1553198641.8753 | 851 | 0.20948 | 0.000471 | -4e-06 | 100.4585 | 100.4804 | 122.9138 | 21.2612 | 0.002123 | 0.002002 | -152.3902 | 0.24126 | 8.4026 | 0.003476 | 0.004752 | 0.004531 | 5.2e-05 | 0.001026 | 0.26808 | 1.3016 | 2.548 | 98.85 | 20.82 | 38.481 | 72.8998 | 105.5118 | 70.0028 | 0.009586 | 1.0288 | 0.024598 | 23.9087 | -0.001948 | 0.14353 | -4.6e-05 | -0.003095 | -0.003095 | -1.9057
2019-03-21 16:04 | 1553198642.875 | 852 | 0.20949 | 0.000468 | -4e-06 | 100.5473 | 100.5795 | 122.9767 | 21.2612 | 0.002105 | 0.002016 | -152.3662 | 0.24111 | 8.3994 | 0.003408 | 0.004816 | 0.004693 | 5.2e-05 | 0.001025 | 0.26842 | 1.324 | 2.548 | 98.85 | 20.82 | 38.48 | 72.4796 | 103.872 | 70.3349 | 0.009585 | 1.0278 | 0.024592 | 23.9264 | -0.002865 | 0.14344 | -0.000151 | -0.010863 | -0.010863 | -0.54268
2019-03-21 16:04 | 1553198643.8753 | 853 | 0.20947 | 0.000467 | -4e-06 | 100.544 | 100.5474 | 123.099 | 21.2612 | 0.00213 | 0.002012 | -152.3566 | 0.2422 | 8.4015 | 0.003425 | 0.003938 | 0.003763 | 4.8e-05 | 0.001028 | 0.27312 | 1.3066 | 2.5479 | 98.85 | 20.82 | 38.479 | 73.0776 | 104.2692 | 70.246 | 0.009585 | 1.0281 | 0.024608 | 23.9352 | -0.000308 | 0.14409 | -7e-06 | -0.000224 | -0.000224 | -26.4311
2019-03-21 16:04 | 1553198644.8742 | 854 | 0.20942 | 0.000465 | -4e-06 | 100.4871 | 100.4885 | 122.754 | 21.2612 | 0.002111 | 0.002013 | -152.3403 | 0.24107 | 8.4035 | 0.003459 | 0.004861 | 0.00469 | 5.5e-05 | 0.001029 | 0.29048 | 1.3034 | 2.5479 | 98.85 | 20.82 | 38.48 | 72.6036 | 105.0964 | 70.2624 | 0.009585 | 1.0281 | 0.024573 | 23.9022 | -0.000129 | 0.14341 | 0.000317 | 0.023644 | 0.023644 | 0.24929
2019-03-21 16:04 | 1553198645.8746 | 855 | 0.20943 | 0.000468 | -5e-06 | 100.5049 | 100.5917 | 122.5974 | 21.2612 | 0.002113 | 0.001993 | -152.3455 | 0.24217 | 8.4051 | 0.003452 | 0.004534 | 0.004334 | 5.2e-05 | 0.001027 | 0.26432 | 1.3183 | 2.548 | 98.85 | 20.82 | 38.478 | 72.6706 | 104.917 | 69.7726 | 0.009585 | 1.0295 | 0.024574 | 23.8699 | -0.007719 | 0.14407 | 0.00025 | 0.018693 | 0.018693 | 0.31675
2019-03-21 16:04 | 1553198646.8746 | 856 | 0.20948 | 0.000469 | -4e-06 | 100.5508 | 100.6534 | 122.6279 | 21.2612 | 0.00203 | 0.002025 | -152.3263 | 0.24134 | 8.4009 | 0.003362 | 0.005011 | 0.004859 | 5.2e-05 | 0.001032 | 0.27505 | 1.3438 | 2.5479 | 98.85 | 20.82 | 38.477 | 70.6623 | 102.7365 | 70.5493 | 0.009585 | 1.0272 | 0.02455 | 23.8999 | -0.009117 | 0.14357 | -4.5e-05 | -0.003037 | -0.003037 | -1.9429
2019-03-21 16:04 | 1553198647.8743 | 857 | 0.20943 | 0.000466 | -5e-06 | 100.5112 | 100.5832 | 122.8117 | 21.2612 | 0.002051 | 0.001968 | -152.3214 | 0.24039 | 8.402 | 0.003388 | 0.004665 | 0.004557 | 4.8e-05 | 0.001033 | 0.2995 | 1.3584 | 2.548 | 98.85 | 20.82 | 38.476 | 71.1847 | 103.3848 | 69.1707 | 0.009584 | 1.0313 | 0.024618 | 23.8697 | -0.006405 | 0.14301 | 0.000218 | 0.016402 | 0.016402 | 0.35837
2019-03-21 16:04 | 1553198648.8754 | 858 | 0.20944 | 0.000464 | -4e-06 | 100.5666 | 100.479 | 122.6845 | 21.2612 | 0.002104 | 0.001977 | -152.296 | 0.24143 | 8.403 | 0.003384 | 0.004714 | 0.004539 | 5e-05 | 0.001033 | 0.27776 | 1.3554 | 2.548 | 98.85 | 20.82 | 38.476 | 72.4532 | 103.2809 | 69.3921 | 0.009584 | 1.0307 | 0.024597 | 23.8651 | 0.007798 | 0.14363 | 0.000161 | 0.012139 | 0.012139 | 0.48629
2019-03-21 16:04 | 1553198649.8754 | 859 | 0.20943 | 0.000468 | -4e-06 | 100.5773 | 100.715 | 122.7404 | 21.2612 | 0.002081 | 0.001984 | -152.3067 | 0.24174 | 8.3998 | 0.003403 | 0.005275 | 0.005091 | 5.1e-05 | 0.001034 | 0.27489 | 1.3179 | 2.5479 | 98.85 | 20.82 | 38.474 | 71.8902 | 103.7361 | 69.5567 | 0.009584 | 1.0302 | 0.024597 | 23.8763 | -0.012235 | 0.14381 | 0.000227 | 0.017022 | 0.017022 | 0.34723
2019-03-21 16:04 | 1553198650.8755 | 860 | 0.20944 | 0.000472 | -4e-06 | 100.5496 | 100.6169 | 122.846 | 21.2612 | 0.00208 | 0.001955 | -152.316 | 0.24223 | 8.401 | 0.003359 | 0.004534 | 0.004344 | 5.2e-05 | 0.001034 | 0.27489 | 1.3376 | 2.548 | 98.85 | 20.82 | 38.473 | 71.8613 | 102.6784 | 68.8642 | 0.009584 | 1.0323 | 0.024632 | 23.8624 | -0.005984 | 0.1441 | 0.000183 | 0.013817 | 0.013817 | 0.42863
2019-03-21 16:04 | 1553198651.8746 | 861 | 0.20943 | 0.000473 | -4e-06 | 100.5827 | 100.7559 | 122.5446 | 21.2612 | 0.002093 | 0.002007 | -152.313 | 0.2414 | 8.4023 | 0.003364 | 0.004989 | 0.004849 | 5.2e-05 | 0.001035 | 0.27975 | 1.3642 | 2.5481 | 98.85 | 20.82 | 38.474 | 72.1843 | 102.8024 | 70.1227 | 0.009584 | 1.0285 | 0.024557 | 23.8769 | -0.015387 | 0.14361 | 0.000212 | 0.015843 | 0.015843 | 0.37256
2019-03-21 16:04 | 1553198652.8745 | 862 | 0.20945 | 0.000467 | -4e-06 | 100.5785 | 100.8615 | 122.5171 | 21.2612 | 0.002102 | 0.002022 | -152.3223 | 0.24072 | 8.4041 | 0.003357 | 0.00412 | 0.003968 | 5.2e-05 | 0.001037 | 0.28926 | 1.304 | 2.5479 | 98.85 | 20.82 | 38.472 | 72.3948 | 102.6217 | 70.4798 | 0.009583 | 1.0274 | 0.024541 | 23.8866 | -0.025126 | 0.1432 | 9.8e-05 | 0.007462 | 0.007462 | 0.78881
2019-03-21 16:04 | 1553198653.8745 | 863 | 0.20947 | 0.000468 | -4e-06 | 100.5836 | 100.7058 | 122.6779 | 21.2612 | 0.00208 | 0.001982 | -152.3324 | 0.2409 | 8.4017 | 0.003302 | 0.004932 | 0.004762 | 5e-05 | 0.001038 | 0.26034 | 1.3215 | 2.5479 | 98.85 | 20.82 | 38.472 | 71.8613 | 101.2977 | 69.5036 | 0.009583 | 1.0303 | 0.024592 | 23.8683 | -0.010861 | 0.14331 | -3.1e-05 | -0.002015 | -0.002015 | -2.9237
2019-03-21 16:04 | 1553198654.875 | 864 | 0.20948 | 0.000467 | -4e-06 | 100.5759 | 100.7307 | 122.8054 | 21.2612 | 0.002089 | 0.00203 | -152.3332 | 0.24057 | 8.4022 | 0.003327 | 0.004525 | 0.004352 | 5e-05 | 0.001037 | 0.28296 | 1.319 | 2.5481 | 98.85 | 20.82 | 38.473 | 72.0899 | 101.9071 | 70.657 | 0.009584 | 1.0269 | 0.024564 | 23.9209 | -0.013758 | 0.14311 | -5e-05 | -0.003466 | -0.003466 | -1.6969
2019-03-21 16:04 | 1553198655.8747 | 865 | 0.20945 | 0.000468 | -4e-06 | 100.552 | 100.6476 | 122.7713 | 21.2612 | 0.002071 | 0.002013 | -152.3063 | 0.24198 | 8.4033 | 0.003327 | 0.004644 | 0.00446 | 5e-05 | 0.001036 | 0.26936 | 1.3442 | 2.5481 | 98.85 | 20.82 | 38.472 | 71.6619 | 101.8875 | 70.2648 | 0.009583 | 1.028 | 0.024574 | 23.9039 | -0.008496 | 0.14396 | 0.000107 | 0.008189 | 0.008189 | 0.7225
2019-03-21 16:04 | 1553198656.8738 | 866 | 0.20944 | 0.000473 | -4e-06 | 100.5607 | 100.698 | 122.7345 | 21.2612 | 0.00207 | 0.001963 | -152.3202 | 0.24075 | 8.4015 | 0.003347 | 0.004466 | 0.0043 | 4.9e-05 | 0.001035 | 0.27936 | 1.3321 | 2.548 | 98.85 | 20.82 | 38.472 | 71.6362 | 102.3774 | 69.0676 | 0.009583 | 1.0316 | 0.024614 | 23.8586 | -0.012206 | 0.14322 | 0.000199 | 0.014942 | 0.014942 | 0.39395
2019-03-21 16:04 | 1553198657.8748 | 867 | 0.20945 | 0.000468 | -5e-06 | 100.5917 | 100.6609 | 123.0063 | 21.2612 | 0.002102 | 0.002025 | -152.3163 | 0.24122 | 8.4042 | 0.003362 | 0.004233 | 0.004082 | 5e-05 | 0.001035 | 0.27455 | 1.2828 | 2.5479 | 98.85 | 20.81 | 38.472 | 72.401 | 102.7569 | 70.5375 | 0.009578 | 1.0272 | 0.024588 | 23.9363 | -0.006152 | 0.1435 | 0.000131 | 0.009958 | 0.009958 | 0.59228
2019-03-21 16:04 | 1553198658.8748 | 868 | 0.20946 | 0.00047 | -4e-06 | 100.5607 | 100.6758 | 123.283 | 21.2612 | 0.002102 | 0.00201 | -152.3388 | 0.24085 | 8.4031 | 0.003329 | 0.004789 | 0.00456 | 5.3e-05 | 0.001036 | 0.26521 | 1.3461 | 2.548 | 98.85 | 20.81 | 38.475 | 72.4049 | 101.9376 | 70.1773 | 0.009578 | 1.0283 | 0.024629 | 23.9506 | -0.010239 | 0.14328 | 8e-05 | 0.006175 | 0.006175 | 0.95365
2019-03-21 16:04 | 1553198659.8739 | 869 | 0.20945 | 0.000464 | -4e-06 | 100.5717 | 100.6203 | 123.4666 | 21.2612 | 0.002092 | 0.001965 | -152.3464 | 0.24211 | 8.4009 | 0.003399 | 0.005208 | 0.00508 | 5e-05 | 0.001036 | 0.26537 | 1.3113 | 2.548 | 98.85 | 20.81 | 38.476 | 72.1482 | 103.6387 | 69.1099 | 0.009579 | 1.0315 | 0.024685 | 23.9312 | -0.004325 | 0.14403 | 0.000137 | 0.010381 | 0.010381 | 0.57026
2019-03-21 16:04 | 1553198660.874 | 870 | 0.20944 | 0.000471 | -5e-06 | 100.6247 | 100.7501 | 123.7053 | 21.2612 | 0.00209 | 0.001975 | -152.333 | 0.24164 | 8.4036 | 0.003417 | 0.004551 | 0.00439 | 5.1e-05 | 0.001035 | 0.27936 | 1.3039 | 2.548 | 98.85 | 20.81 | 38.477 | 72.1016 | 104.0724 | 69.3476 | 0.009579 | 1.0308 | 0.024701 | 23.9626 | -0.011138 | 0.14375 | 0.000175 | 0.013287 | 0.013287 | 0.44466
2019-03-21 16:04 | 1553198661.874 | 871 | 0.20946 | 0.000472 | -4e-06 | 100.5827 | 100.5866 | 123.7694 | 21.2612 | 0.002029 | 0.001958 | -152.3228 | 0.24096 | 8.4038 | 0.00339 | 0.005139 | 0.00493 | 5.6e-05 | 0.001039 | 0.30143 | 1.3344 | 2.5482 | 98.85 | 20.81 | 38.481 | 70.6334 | 103.4265 | 68.9397 | 0.00958 | 1.032 | 0.024722 | 23.9545 | -0.000347 | 0.14335 | 4.3e-05 | 0.003422 | 0.003422 | 1.7218
2019-03-21 16:04 | 1553198662.8751 | 872 | 0.20946 | 0.000466 | -4e-06 | 100.5988 | 100.7515 | 123.6417 | 21.2612 | 0.002053 | 0.001972 | -152.3328 | 0.24141 | 8.4019 | 0.003398 | 0.004674 | 0.004525 | 5.4e-05 | 0.00104 | 0.27096 | 1.3192 | 2.5479 | 98.85 | 20.81 | 38.482 | 71.233 | 103.6072 | 69.2861 | 0.00958 | 1.031 | 0.024696 | 23.9543 | -0.013563 | 0.14361 | 6.5e-05 | 0.005048 | 0.005048 | 1.1693
2019-03-21 16:04 | 1553198663.8741 | 873 | 0.20945 | 0.000472 | -4e-06 | 100.5282 | 100.7429 | 123.5908 | 21.2612 | 0.002086 | 0.001935 | -152.3269 | 0.24145 | 8.4034 | 0.003432 | 0.004551 | 0.004384 | 5.3e-05 | 0.001038 | 0.27831 | 1.3324 | 2.5479 | 98.85 | 20.81 | 38.485 | 72.0103 | 104.4437 | 68.3774 | 0.009581 | 1.0337 | 0.024724 | 23.9176 | -0.019088 | 0.14364 | 0.000134 | 0.010199 | 0.010199 | 0.57886
2019-03-21 16:04 | 1553198664.8745 | 874 | 0.20945 | 0.00047 | -5e-06 | 100.5306 | 100.5774 | 123.3603 | 21.2612 | 0.002043 | 0.00201 | -152.3234 | 0.24236 | 8.406 | 0.003422 | 0.004953 | 0.00474 | 5.5e-05 | 0.001039 | 0.28152 | 1.3281 | 2.548 | 98.85 | 20.81 | 38.487 | 70.9711 | 104.1936 | 70.1991 | 0.009581 | 1.0282 | 0.024636 | 23.9589 | -0.004167 | 0.14418 | 0.000115 | 0.008849 | 0.008849 | 0.66966
2019-03-21 16:04 | 1553198665.8749 | 875 | 0.20944 | 0.00047 | -4e-06 | 100.5839 | 100.6919 | 123.3079 | 21.2612 | 0.002051 | 0.001988 | -152.3503 | 0.24122 | 8.4032 | 0.003379 | 0.004977 | 0.004774 | 5.1e-05 | 0.001035 | 0.25818 | 1.3355 | 2.5481 | 98.85 | 20.81 | 38.487 | 71.1669 | 103.1552 | 69.669 | 0.009581 | 1.0298 | 0.024649 | 23.9353 | -0.009595 | 0.1435 | 0.000183 | 0.013774 | 0.013774 | 0.42818
2019-03-21 16:04 | 1553198666.8753 | 876 | 0.20946 | 0.000472 | -4e-06 | 100.5473 | 100.7269 | 123.4135 | 21.2612 | 0.002075 | 0.001906 | -152.3584 | 0.24196 | 8.3991 | 0.003448 | 0.004271 | 0.00417 | 5.5e-05 | 0.001038 | 0.30127 | 1.3058 | 2.548 | 98.85 | 20.81 | 38.488 | 71.7508 | 104.8336 | 67.698 | 0.009582 | 1.0358 | 0.024731 | 23.8766 | -0.01597 | 0.14394 | 7.4e-05 | 0.00573 | 0.00573 | 1.0324
2019-03-21 16:04 | 1553198667.8757 | 877 | 0.20946 | 0.000475 | -4e-06 | 100.6396 | 100.5509 | 123.1633 | 21.2612 | 0.002027 | 0.001991 | -152.3713 | 0.24126 | 8.4051 | 0.0035 | 0.004694 | 0.004528 | 5.4e-05 | 0.00104 | 0.25752 | 1.3659 | 2.5479 | 98.85 | 20.81 | 38.489 | 70.5945 | 106.0963 | 69.7235 | 0.009582 | 1.0297 | 0.024633 | 23.9232 | 0.007893 | 0.14353 | 7.2e-05 | 0.005609 | 0.005609 | 1.0516
2019-03-21 16:04 | 1553198668.8744 | 878 | 0.20944 | 0.000468 | -4e-06 | 100.583 | 100.4497 | 123.0584 | 21.2612 | 0.001983 | 0.002002 | -152.3691 | 0.24085 | 8.402 | 0.003483 | 0.005053 | 0.004891 | 5e-05 | 0.001042 | 0.28744 | 1.3721 | 2.548 | 98.85 | 20.81 | 38.491 | 69.5435 | 105.6818 | 69.9949 | 0.009582 | 1.0289 | 0.024613 | 23.9225 | 0.011863 | 0.14328 | 0.000147 | 0.01108 | 0.01108 | 0.53149
2019-03-21 16:04 | 1553198669.8745 | 879 | 0.20946 | 0.00047 | -5e-06 | 100.6006 | 100.6125 | 122.7617 | 21.2612 | 0.00211 | 0.001993 | -152.3771 | 0.24128 | 8.4001 | 0.003489 | 0.00549 | 0.005384 | 5.3e-05 | 0.001045 | 0.28296 | 1.3581 | 2.548 | 98.85 | 20.81 | 38.494 | 72.6028 | 105.8343 | 69.7694 | 0.009583 | 1.0295 | 0.024591 | 23.8857 | -0.001059 | 0.14354 | 6.7e-05 | 0.005247 | 0.005247 | 1.1242
2019-03-21 16:04 | 1553198670.8746 | 880 | 0.20944 | 0.000473 | -4e-06 | 100.5762 | 100.6145 | 122.558 | 21.2612 | 0.002123 | 0.001958 | -152.401 | 0.24194 | 8.4044 | 0.003521 | 0.004665 | 0.004536 | 5.2e-05 | 0.001047 | 0.2724 | 1.3229 | 2.5481 | 98.85 | 20.81 | 38.496 | 72.9015 | 106.5977 | 68.9484 | 0.009584 | 1.032 | 0.0246 | 23.8373 | -0.003413 | 0.14393 | 0.000168 | 0.012695 | 0.012695 | 0.46598
2019-03-21 16:05 | 1553198671.8759 | 881 | 0.20946 | 0.000473 | -4e-06 | 100.6453 | 100.5485 | 122.5448 | 21.2612 | 0.002097 | 0.001956 | -152.4221 | 0.2414 | 8.4021 | 0.003498 | 0.004982 | 0.004866 | 5.5e-05 | 0.00105 | 0.26272 | 1.3106 | 2.5479 | 98.85 | 20.82 | 38.5 | 72.2854 | 106.0373 | 68.8868 | 0.00959 | 1.0322 | 0.024601 | 23.8339 | 0.008608 | 0.14361 | 3.1e-05 | 0.002584 | 0.002584 | 2.2844
2019-03-21 16:05 | 1553198672.8737 | 882 | 0.20945 | 0.000472 | -5e-06 | 100.6596 | 100.5706 | 122.4707 | 21.2612 | 0.002018 | 0.001968 | -152.4231 | 0.24133 | 8.4002 | 0.003455 | 0.005354 | 0.005142 | 5.5e-05 | 0.00105 | 0.2656 | 1.3351 | 2.548 | 98.85 | 20.81 | 38.502 | 70.3746 | 104.9997 | 69.177 | 0.009585 | 1.0313 | 0.024583 | 23.8368 | 0.00791 | 0.14356 | 0.000133 | 0.010067 | 0.010067 | 0.58611
2019-03-21 16:05 | 1553198673.8741 | 883 | 0.20947 | 0.00047 | -4e-06 | 100.6268 | 100.7314 | 122.339 | 21.2612 | 0.002014 | 0.001963 | -152.4142 | 0.24051 | 8.4015 | 0.003433 | 0.004644 | 0.00447 | 5.2e-05 | 0.001048 | 0.26576 | 1.2904 | 2.548 | 98.85 | 20.81 | 38.504 | 70.2811 | 104.4647 | 69.0474 | 0.009586 | 1.0317 | 0.024575 | 23.8195 | -0.009289 | 0.14308 | 2.6e-05 | 0.002169 | 0.002169 | 2.7108
2019-03-21 16:05 | 1553198674.8737 | 884 | 0.20945 | 0.000474 | -4e-06 | 100.5476 | 100.6197 | 122.3498 | 21.2612 | 0.002066 | 0.001988 | -152.4221 | 0.24152 | 8.4021 | 0.00349 | 0.004824 | 0.004747 | 5e-05 | 0.001047 | 0.29806 | 1.2838 | 2.5481 | 98.85 | 20.81 | 38.507 | 71.5456 | 105.8508 | 69.6488 | 0.009586 | 1.0299 | 0.024554 | 23.8414 | -0.006411 | 0.14368 | 0.000101 | 0.007692 | 0.007692 | 0.76769
2019-03-21 16:05 | 1553198675.8751 | 885 | 0.20946 | 0.000472 | -4e-06 | 100.5568 | 100.7198 | 122.5369 | 21.2612 | 0.002042 | 0.001938 | -152.4028 | 0.24114 | 8.4013 | 0.003435 | 0.004794 | 0.004485 | 5.1e-05 | 0.001048 | 0.28655 | 1.2651 | 2.5481 | 98.85 | 20.81 | 38.51 | 70.9626 | 104.5207 | 68.4469 | 0.009587 | 1.0335 | 0.024616 | 23.8178 | -0.014487 | 0.14346 | 5e-05 | 0.00394 | 0.00394 | 1.4966
2019-03-21 16:05 | 1553198676.8751 | 886 | 0.20948 | 0.000468 | -4e-06 | 100.4987 | 100.4599 | 122.8955 | 21.2612 | 0.002044 | 0.00193 | -152.4077 | 0.24211 | 8.4001 | 0.003505 | 0.00442 | 0.004224 | 4.8e-05 | 0.001047 | 0.26521 | 1.3414 | 2.5481 | 98.85 | 20.81 | 38.517 | 71.0061 | 106.2168 | 68.262 | 0.009589 | 1.0341 | 0.024659 | 23.8462 | 0.003452 | 0.14403 | -7.4e-05 | -0.005206 | -0.005206 | -1.1371
2019-03-21 16:05 | 1553198677.8749 | 887 | 0.20945 | 0.000471 | -5e-06 | 100.5827 | 100.4565 | 123.1361 | 21.2612 | 0.002034 | 0.001952 | -152.3895 | 0.24152 | 8.4031 | 0.003474 | 0.005271 | 0.005067 | 4.6e-05 | 0.001049 | 0.23435 | 1.3226 | 2.5481 | 98.85 | 20.82 | 38.52 | 70.7622 | 105.463 | 68.7823 | 0.009595 | 1.0325 | 0.024664 | 23.8877 | 0.01123 | 0.14368 | 0.000101 | 0.007742 | 0.007742 | 0.76275
2019-03-21 16:05 | 1553198678.874 | 888 | 0.20947 | 0.000471 | -4e-06 | 100.5523 | 100.5985 | 123.0514 | 21.2612 | 0.002036 | 0.00194 | -152.376 | 0.24166 | 8.4033 | 0.003447 | 0.004903 | 0.004721 | 4.9e-05 | 0.001052 | 0.28998 | 1.3292 | 2.548 | 98.85 | 20.82 | 38.523 | 70.8057 | 104.7968 | 68.5053 | 0.009596 | 1.0333 | 0.024666 | 23.8698 | -0.00411 | 0.14376 | -3.4e-05 | -0.002264 | -0.002264 | -2.61
2019-03-21 16:05 | 1553198679.8746 | 889 | 0.20945 | 0.000469 | -4e-06 | 100.569 | 100.6758 | 123.1016 | 21.2612 | 0.001999 | 0.001962 | -152.3561 | 0.24162 | 8.4005 | 0.003456 | 0.004673 | 0.004519 | 5e-05 | 0.001052 | 0.24491 | 1.3071 | 2.548 | 98.85 | 20.82 | 38.528 | 69.9247 | 105.013 | 69.0262 | 0.009597 | 1.0318 | 0.024652 | 23.8929 | -0.009497 | 0.14374 | 0.00014 | 0.010621 | 0.010621 | 0.55622
2019-03-21 16:05 | 1553198680.8745 | 890 | 0.20947 | 0.000465 | -4e-06 | 100.5479 | 100.6047 | 122.8507 | 21.2612 | 0.00206 | 0.002003 | -152.3466 | 0.24186 | 8.4011 | 0.003582 | 0.004974 | 0.0048 | 5e-05 | 0.001054 | 0.26073 | 1.3513 | 2.5481 | 98.85 | 20.82 | 38.529 | 71.3913 | 108.0776 | 70.0285 | 0.009598 | 1.0288 | 0.024591 | 23.9034 | -0.005053 | 0.14388 | 1.9e-05 | 0.001665 | 0.001665 | 3.5527
2019-03-21 16:05 | 1553198681.8751 | 891 | 0.20943 | 0.000473 | -4e-06 | 100.5282 | 100.6391 | 122.9763 | 21.2612 | 0.002058 | 0.001943 | -152.3431 | 0.24127 | 8.4021 | 0.003506 | 0.00541 | 0.005207 | 5.1e-05 | 0.001058 | 0.28224 | 1.3452 | 2.548 | 98.85 | 20.82 | 38.533 | 71.3507 | 106.234 | 68.5748 | 0.009599 | 1.0331 | 0.024656 | 23.8649 | -0.00986 | 0.14353 | 0.000231 | 0.017342 | 0.017342 | 0.34016
2019-03-21 16:05 | 1553198682.8741 | 892 | 0.20945 | 0.000471 | -4e-06 | 100.5619 | 100.7232 | 123.2861 | 21.2612 | 0.002026 | 0.001915 | -152.3303 | 0.24206 | 8.402 | 0.003565 | 0.005411 | 0.00531 | 5e-05 | 0.001058 | 0.25642 | 1.3583 | 2.5479 | 98.85 | 20.82 | 38.534 | 70.5827 | 107.6561 | 67.9048 | 0.009599 | 1.0352 | 0.024711 | 23.8715 | -0.014338 | 0.144 | 0.000104 | 0.00795 | 0.00795 | 0.74443
2019-03-21 16:05 | 1553198683.8755 | 893 | 0.20946 | 0.000467 | -4e-06 | 100.5318 | 100.7876 | 123.556 | 21.2612 | 0.002067 | 0.001962 | -152.3547 | 0.24108 | 8.4004 | 0.00352 | 0.004283 | 0.004085 | 5e-05 | 0.001061 | 0.26145 | 1.362 | 2.5479 | 98.85 | 20.81 | 38.536 | 71.5629 | 106.5742 | 69.0342 | 0.009593 | 1.0317 | 0.024697 | 23.9372 | -0.02273 | 0.14342 | 8.6e-05 | 0.006629 | 0.006629 | 0.88924
2019-03-21 16:05 | 1553198684.8742 | 894 | 0.20948 | 0.000469 | -4e-06 | 100.5452 | 100.7164 | 123.6307 | 21.2612 | 0.002071 | 0.001955 | -152.3464 | 0.24144 | 8.4007 | 0.003514 | 0.00485 | 0.004661 | 4.9e-05 | 0.001062 | 0.25929 | 1.2961 | 2.5479 | 98.85 | 20.81 | 38.538 | 71.6658 | 106.4206 | 68.8618 | 0.009594 | 1.0323 | 0.024711 | 23.9384 | -0.015218 | 0.14363 | -8.8e-05 | -0.006319 | -0.006319 | -0.93425
2019-03-21 16:05 | 1553198685.8753 | 895 | 0.20947 | 0.000468 | -4e-06 | 100.5681 | 100.6067 | 123.4274 | 21.2612 | 0.002038 | 0.001967 | -152.3758 | 0.24153 | 8.4026 | 0.003415 | 0.004266 | 0.004166 | 5.2e-05 | 0.001064 | 0.25769 | 1.3055 | 2.548 | 98.85 | 20.82 | 38.541 | 70.8595 | 104.0293 | 69.1481 | 0.009601 | 1.0314 | 0.02468 | 23.9287 | -0.003432 | 0.14369 | -1.2e-05 | -0.000604 | -0.000604 | -9.7724
2019-03-21 16:05 | 1553198686.8755 | 896 | 0.20948 | 0.00047 | -4e-06 | 100.5234 | 100.5812 | 122.9392 | 21.2612 | 0.002074 | 0.001979 | -152.3688 | 0.24112 | 8.4016 | 0.003412 | 0.004853 | 0.004623 | 4.9e-05 | 0.001068 | 0.29518 | 1.2794 | 2.5478 | 98.85 | 20.82 | 38.544 | 71.7376 | 103.9672 | 69.4365 | 0.009601 | 1.0305 | 0.024621 | 23.8914 | -0.005137 | 0.14344 | -6.9e-05 | -0.004861 | -0.004861 | -1.2129
2019-03-21 16:05 | 1553198687.8737 | 897 | 0.20949 | 0.000471 | -5e-06 | 100.5321 | 100.8363 | 122.6753 | 21.2612 | 0.002023 | 0.00194 | -152.3967 | 0.24141 | 8.4035 | 0.003416 | 0.00477 | 0.0046 | 5.2e-05 | 0.001067 | 0.25083 | 1.2705 | 2.548 | 98.85 | 20.82 | 38.545 | 70.4972 | 104.0457 | 68.5085 | 0.009602 | 1.0333 | 0.024628 | 23.8334 | -0.027025 | 0.14361 | -0.000114 | -0.008123 | -0.008123 | -0.72665
2019-03-21 16:05 | 1553198688.8745 | 898 | 0.2095 | 0.000468 | -4e-06 | 100.5294 | 100.6861 | 123.0516 | 21.2612 | 0.002065 | 0.001969 | -152.3983 | 0.24184 | 8.4047 | 0.003454 | 0.004637 | 0.004464 | 4.7e-05 | 0.001062 | 0.26001 | 1.2815 | 2.5479 | 98.85 | 20.82 | 38.548 | 71.5115 | 104.9658 | 69.2065 | 0.009602 | 1.0312 | 0.02464 | 23.8943 | -0.01393 | 0.14387 | -0.000183 | -0.013232 | -0.013232 | -0.44688
2019-03-21 16:05 | 1553198689.8745 | 899 | 0.20949 | 0.000469 | -4e-06 | 100.5285 | 100.5042 | 123.2129 | 21.2612 | 0.002003 | 0.001969 | -152.4044 | 0.24143 | 8.4 | 0.0034 | 0.004999 | 0.004776 | 5e-05 | 0.001065 | 0.29927 | 1.2817 | 2.548 | 98.85 | 20.81 | 38.549 | 70.0199 | 103.6709 | 69.2019 | 0.009597 | 1.0312 | 0.024657 | 23.9098 | 0.002162 | 0.14363 | -0.000152 | -0.010988 | -0.010988 | -0.53722
2019-03-21 16:05 | 1553198690.8746 | 900 | 0.20948 | 0.00047 | -4e-06 | 100.4871 | 100.5924 | 123.5018 | 21.2612 | 0.002098 | 0.001973 | -152.4113 | 0.24078 | 8.4006 | 0.003399 | 0.004509 | 0.004268 | 5e-05 | 0.001064 | 0.274 | 1.3225 | 2.548 | 98.85 | 20.81 | 38.55 | 72.302 | 103.652 | 69.3024 | 0.009597 | 1.0309 | 0.024682 | 23.9413 | -0.009371 | 0.14324 | -9.1e-05 | -0.006461 | -0.006461 | -0.91119
2019-03-21 16:05 | 1553198691.875 | 901 | 0.20953 | 0.00047 | -4e-06 | 100.4677 | 100.4609 | 123.3992 | 21.2612 | 0.002021 | 0.001928 | -152.411 | 0.24244 | 8.4046 | 0.003368 | 0.004839 | 0.004685 | 5e-05 | 0.001067 | 0.26952 | 1.3651 | 2.548 | 98.85 | 20.81 | 38.555 | 70.4527 | 102.8857 | 68.2075 | 0.009598 | 1.0342 | 0.024711 | 23.8931 | 0.000602 | 0.14423 | -0.000373 | -0.027448 | -0.027448 | -0.21596
2019-03-21 16:05 | 1553198692.8747 | 902 | 0.20949 | 0.00047 | -4e-06 | 100.5857 | 100.6074 | 123.2349 | 21.2612 | 0.002013 | 0.001968 | -152.42 | 0.24102 | 8.4034 | 0.003386 | 0.004895 | 0.004739 | 5.1e-05 | 0.001068 | 0.25752 | 1.3577 | 2.5479 | 98.85 | 20.81 | 38.557 | 70.27 | 103.3369 | 69.1753 | 0.009599 | 1.0313 | 0.02466 | 23.911 | -0.001929 | 0.14338 | -0.00013 | -0.00935 | -0.00935 | -0.6303
2019-03-21 16:05 | 1553198693.8737 | 903 | 0.20947 | 0.00047 | -4e-06 | 100.4835 | 100.59 | 122.8864 | 21.2612 | 0.002028 | 0.001957 | -152.4248 | 0.24159 | 8.4023 | 0.003405 | 0.004826 | 0.004672 | 4.8e-05 | 0.001064 | 0.22357 | 1.347 | 2.5479 | 98.85 | 20.81 | 38.563 | 70.6288 | 103.7942 | 68.9039 | 0.0096 | 1.0321 | 0.024635 | 23.8677 | -0.009477 | 0.14372 | -2.3e-05 | -0.001424 | -0.001424 | -4.1467
2019-03-21 16:05 | 1553198694.8748 | 904 | 0.20947 | 0.000466 | -4e-06 | 100.5496 | 100.5587 | 122.6962 | 21.2612 | 0.002039 | 0.001974 | -152.4483 | 0.24066 | 8.4019 | 0.003402 | 0.005029 | 0.0048 | 5e-05 | 0.001063 | 0.26698 | 1.3237 | 2.5481 | 98.85 | 20.81 | 38.567 | 70.8821 | 103.7228 | 69.3236 | 0.009601 | 1.0309 | 0.0246 | 23.8638 | -0.000805 | 0.14317 | 2.3e-05 | 0.001945 | 0.001945 | 3.0259
2019-03-21 16:05 | 1553198695.8739 | 905 | 0.2095 | 0.000462 | -4e-06 | 100.5401 | 100.6231 | 122.6908 | 21.2612 | 0.002015 | 0.002008 | -152.4236 | 0.24098 | 8.4016 | 0.003405 | 0.004509 | 0.004388 | 4.8e-05 | 0.001065 | 0.30055 | 1.3635 | 2.5479 | 98.85 | 20.81 | 38.571 | 70.3075 | 103.7778 | 70.1463 | 0.009602 | 1.0284 | 0.02457 | 23.892 | -0.007377 | 0.14336 | -0.000181 | -0.013125 | -0.013125 | -0.44892
2019-03-21 16:05 | 1553198696.8739 | 906 | 0.20948 | 0.000469 | -4e-06 | 100.4686 | 100.6411 | 122.4453 | 21.2612 | 0.002022 | 0.001936 | -152.4145 | 0.24174 | 8.402 | 0.003471 | 0.004362 | 0.004112 | 4.8e-05 | 0.001058 | 0.274 | 1.3227 | 2.5481 | 98.85 | 20.81 | 38.578 | 70.4854 | 105.3859 | 68.4212 | 0.009604 | 1.0336 | 0.024608 | 23.808 | -0.015345 | 0.14381 | -4e-05 | -0.002742 | -0.002742 | -2.1555
2019-03-21 16:05 | 1553198697.8749 | 907 | 0.20946 | 0.000471 | -4e-06 | 100.454 | 100.5727 | 122.6563 | 21.2612 | 0.002043 | 0.002007 | -152.4072 | 0.2417 | 8.4013 | 0.003478 | 0.004753 | 0.004564 | 4.9e-05 | 0.001061 | 0.27328 | 1.3081 | 2.548 | 98.85 | 20.81 | 38.584 | 70.9835 | 105.5455 | 70.1196 | 0.009605 | 1.0285 | 0.024568 | 23.8877 | -0.010559 | 0.14379 | 8.1e-05 | 0.006223 | 0.006223 | 0.94962
2019-03-21 16:05 | 1553198698.874 | 908 | 0.20949 | 0.000464 | -4e-06 | 100.4814 | 100.6442 | 122.6368 | 21.2612 | 0.00204 | 0.001949 | -152.4324 | 0.24228 | 8.4019 | 0.003472 | 0.005512 | 0.005379 | 4.8e-05 | 0.001062 | 0.27273 | 1.3268 | 2.5477 | 98.85 | 20.81 | 38.593 | 70.897 | 105.4203 | 68.7222 | 0.009608 | 1.0327 | 0.024616 | 23.8371 | -0.014477 | 0.14413 | -0.000127 | -0.009149 | -0.009149 | -0.64752
2019-03-21 16:05 | 1553198699.8753 | 909 | 0.20947 | 0.000466 | -4e-06 | 100.5416 | 100.5641 | 122.4364 | 21.2612 | 0.002064 | 0.001926 | -152.4316 | 0.24175 | 8.3977 | 0.003451 | 0.004544 | 0.004323 | 5.1e-05 | 0.001065 | 0.27992 | 1.3558 | 2.548 | 98.85 | 20.81 | 38.598 | 71.4927 | 104.903 | 68.1808 | 0.009609 | 1.0343 | 0.024616 | 23.7988 | -0.002005 | 0.14382 | -2.8e-05 | -0.001814 | -0.001814 | -3.258
2019-03-21 16:05 | 1553198700.8745 | 910 | 0.20949 | 0.000469 | -4e-06 | 100.4802 | 100.6959 | 122.673 | 21.2612 | 0.002012 | 0.001958 | -152.4197 | 0.24109 | 8.4011 | 0.003476 | 0.005233 | 0.00509 | 4.9e-05 | 0.001061 | 0.28871 | 1.3673 | 2.548 | 98.85 | 20.81 | 38.604 | 70.2318 | 105.5188 | 68.9452 | 0.00961 | 1.032 | 0.024612 | 23.8484 | -0.019183 | 0.14342 | -0.000108 | -0.007725 | -0.007725 | -0.76311
2019-03-21 16:05 | 1553198701.8741 | 911 | 0.20949 | 0.000464 | -4e-06 | 100.4671 | 100.3795 | 122.7237 | 21.2612 | 0.002074 | 0.001903 | -152.4152 | 0.24051 | 8.4053 | 0.003519 | 0.004684 | 0.004554 | 4.9e-05 | 0.001064 | 0.27616 | 1.3172 | 2.5479 | 98.85 | 20.81 | 38.615 | 71.7289 | 106.5497 | 67.6177 | 0.009613 | 1.036 | 0.024665 | 23.807 | 0.0078 | 0.14308 | -0.0001 | -0.00718 | -0.00718 | -0.81899
2019-03-21 16:05 | 1553198702.8751 | 912 | 0.20949 | 0.000463 | -4e-06 | 100.5157 | 100.4361 | 122.5735 | 21.2612 | 0.002045 | 0.001938 | -152.4164 | 0.24123 | 8.4006 | 0.003497 | 0.005499 | 0.005317 | 4.7e-05 | 0.001062 | 0.2896 | 1.365 | 2.548 | 98.85 | 20.82 | 38.621 | 71.0381 | 106.0184 | 68.4601 | 0.009621 | 1.0335 | 0.024619 | 23.8218 | 0.007086 | 0.14351 | -9.6e-05 | -0.006816 | -0.006816 | -0.86537
2019-03-21 16:05 | 1553198703.8739 | 913 | 0.2095 | 0.000464 | -4e-06 | 100.4236 | 100.7603 | 122.8082 | 21.2612 | 0.002051 | 0.001968 | -152.3894 | 0.24146 | 8.4036 | 0.003436 | 0.005258 | 0.005087 | 5e-05 | 0.001064 | 0.28208 | 1.3165 | 2.5479 | 98.85 | 20.82 | 38.628 | 71.1638 | 104.5341 | 69.1822 | 0.009622 | 1.0313 | 0.024617 | 23.8698 | -0.02994 | 0.14365 | -0.000171 | -0.012398 | -0.012398 | -0.47619
2019-03-21 16:05 | 1553198704.8733 | 914 | 0.20948 | 0.000466 | -4e-06 | 100.4886 | 100.6888 | 122.9648 | 21.2612 | 0.002065 | 0.00192 | -152.3725 | 0.24133 | 8.3997 | 0.003384 | 0.004797 | 0.004583 | 5.1e-05 | 0.001063 | 0.27577 | 1.2973 | 2.5479 | 98.85 | 20.82 | 38.643 | 71.5153 | 103.288 | 68.0219 | 0.009626 | 1.0348 | 0.024674 | 23.8445 | -0.017805 | 0.14356 | -6e-05 | -0.00419 | -0.00419 | -1.4082
2019-03-21 16:05 | 1553198705.8734 | 915 | 0.20949 | 0.000462 | -4e-06 | 100.4838 | 100.7825 | 123.1455 | 21.2612 | 0.002019 | 0.001945 | -152.3756 | 0.24159 | 8.4012 | 0.003386 | 0.004818 | 0.004643 | 5.1e-05 | 0.001064 | 0.27831 | 1.3128 | 2.548 | 98.85 | 20.82 | 38.651 | 70.4027 | 103.318 | 68.6262 | 0.009628 | 1.033 | 0.024671 | 23.8831 | -0.026546 | 0.14372 | -8.3e-05 | -0.005861 | -0.005861 | -1.0078
2019-03-21 16:05 | 1553198706.8738 | 916 | 0.20948 | 0.000466 | -5e-06 | 100.4623 | 100.4299 | 123.3957 | 21.2612 | 0.002011 | 0.001934 | -152.3707 | 0.24191 | 8.4033 | 0.003403 | 0.003935 | 0.003758 | 5.4e-05 | 0.001066 | 0.27367 | 1.3356 | 2.548 | 98.85 | 20.82 | 38.669 | 70.2217 | 103.7448 | 68.3517 | 0.009633 | 1.0338 | 0.024706 | 23.8978 | 0.002885 | 0.14391 | -6.2e-05 | -0.004314 | -0.004314 | -1.3709
2019-03-21 16:05 | 1553198707.8735 | 917 | 0.20949 | 0.000465 | -4e-06 | 100.4752 | 100.5165 | 123.1924 | 21.2612 | 0.002006 | 0.001925 | -152.3776 | 0.24124 | 8.4003 | 0.003445 | 0.004967 | 0.004805 | 5.1e-05 | 0.001065 | 0.24491 | 1.3017 | 2.548 | 98.85 | 20.81 | 38.678 | 70.0977 | 104.7551 | 68.1512 | 0.009629 | 1.0344 | 0.024692 | 23.8711 | -0.003676 | 0.14351 | -0.00011 | -0.00794 | -0.00794 | -0.74292
2019-03-21 16:05 | 1553198708.8736 | 918 | 0.20948 | 0.000467 | -4e-06 | 100.4248 | 100.4497 | 122.98 | 21.2612 | 0.002004 | 0.001967 | -152.3479 | 0.24144 | 8.4027 | 0.003465 | 0.005253 | 0.005055 | 4.8e-05 | 0.001064 | 0.27848 | 1.3356 | 2.548 | 98.85 | 20.81 | 38.689 | 70.0456 | 105.2538 | 69.1637 | 0.009632 | 1.0313 | 0.024635 | 23.8859 | -0.002217 | 0.14363 | -4.3e-05 | -0.002953 | -0.002953 | -1.9994
2019-03-21 16:05 | 1553198709.8739 | 919 | 0.20945 | 0.000467 | -4e-06 | 100.451 | 100.5318 | 122.976 | 21.2612 | 0.002056 | 0.001927 | -152.3516 | 0.24181 | 8.3998 | 0.003445 | 0.004638 | 0.004473 | 5.2e-05 | 0.001069 | 0.27079 | 1.3075 | 2.548 | 98.85 | 20.81 | 38.712 | 71.2969 | 104.7604 | 68.1988 | 0.009637 | 1.0343 | 0.024669 | 23.8518 | -0.007188 | 0.14386 | 0.000132 | 0.00998 | 0.00998 | 0.59241
2019-03-21 16:05 | 1553198710.8736 | 920 | 0.20951 | 0.000464 | -4e-06 | 100.395 | 100.4548 | 122.7789 | 21.2612 | 0.002022 | 0.001944 | -152.3497 | 0.24241 | 8.4039 | 0.003426 | 0.004508 | 0.00429 | 4.9e-05 | 0.00107 | 0.27489 | 1.3038 | 2.548 | 98.85 | 20.81 | 38.724 | 70.4868 | 104.3077 | 68.6044 | 0.00964 | 1.033 | 0.024635 | 23.8468 | -0.005326 | 0.14421 | -0.000249 | -0.018112 | -0.018112 | -0.32725
2019-03-21 16:05 | 1553198711.8737 | 921 | 0.20946 | 0.000468 | -4e-06 | 100.3783 | 100.4838 | 122.4127 | 21.2612 | 0.002008 | 0.001955 | -152.3369 | 0.24227 | 8.4013 | 0.003507 | 0.005051 | 0.004909 | 5.3e-05 | 0.001069 | 0.26488 | 1.3276 | 2.5479 | 98.85 | 20.81 | 38.748 | 70.1283 | 106.2561 | 68.857 | 0.009646 | 1.0323 | 0.024589 | 23.82 | -0.00939 | 0.14413 | 3e-05 | 0.002483 | 0.002483 | 2.3855
2019-03-21 16:05 | 1553198712.8747 | 922 | 0.20948 | 0.000467 | -4e-06 | 100.4078 | 100.4569 | 122.4547 | 21.2612 | 0.002027 | 0.001925 | -152.3386 | 0.24126 | 8.4031 | 0.003467 | 0.004916 | 0.004701 | 5.3e-05 | 0.001071 | 0.26106 | 1.3477 | 2.5479 | 98.85 | 20.81 | 38.761 | 70.6046 | 105.2931 | 68.1411 | 0.00965 | 1.0344 | 0.024619 | 23.7992 | -0.004367 | 0.14353 | -5e-05 | -0.003447 | -0.003447 | -1.7115
2019-03-21 16:05 | 1553198713.8742 | 923 | 0.20951 | 0.000465 | -4e-06 | 100.3572 | 100.2583 | 122.4427 | 21.2612 | 0.001979 | 0.001939 | -152.3507 | 0.2408 | 8.4054 | 0.003552 | 0.004806 | 0.004685 | 5.5e-05 | 0.001073 | 0.29518 | 1.2952 | 2.548 | 98.85 | 20.81 | 38.772 | 69.4358 | 107.3561 | 68.4702 | 0.009652 | 1.0334 | 0.024606 | 23.8095 | 0.008817 | 0.14325 | -0.000249 | -0.018107 | -0.018107 | -0.32516
2019-03-21 16:05 | 1553198714.8748 | 924 | 0.20949 | 0.000464 | -4e-06 | 100.3747 | 100.3996 | 122.4669 | 21.2612 | 0.002018 | 0.001907 | -152.3535 | 0.24166 | 8.4027 | 0.003495 | 0.00532 | 0.005134 | 5.4e-05 | 0.00107 | 0.25465 | 1.2878 | 2.5478 | 98.85 | 20.81 | 38.79 | 70.3707 | 105.9641 | 67.7004 | 0.009657 | 1.0358 | 0.024636 | 23.785 | -0.002217 | 0.14377 | -0.000116 | -0.008311 | -0.008311 | -0.71094
2019-03-21 16:05 | 1553198715.8739 | 925 | 0.20949 | 0.000465 | -4e-06 | 100.3435 | 100.34 | 122.2519 | 21.2612 | 0.002009 | 0.001918 | -152.3572 | 0.24211 | 8.401 | 0.003496 | 0.005196 | 0.004936 | 4.8e-05 | 0.001069 | 0.30165 | 1.3026 | 2.5479 | 98.85 | 20.81 | 38.797 | 70.1571 | 105.9949 | 67.9774 | 0.009658 | 1.0349 | 0.024604 | 23.7737 | 0.000306 | 0.14403 | -0.000119 | -0.00853 | -0.00853 | -0.69396
2019-03-21 16:05 | 1553198716.8743 | 926 | 0.20949 | 0.000467 | -4e-06 | 100.3828 | 100.3523 | 122.0993 | 21.2612 | 0.002002 | 0.001916 | -152.3702 | 0.24047 | 8.4016 | 0.003496 | 0.005055 | 0.004854 | 5.7e-05 | 0.001071 | 0.26952 | 1.3401 | 2.548 | 98.85 | 20.81 | 38.805 | 69.9941 | 105.9995 | 67.936 | 0.00966 | 1.0351 | 0.02459 | 23.7575 | 0.002718 | 0.14305 | -0.000107 | -0.007656 | -0.007656 | -0.76795
2019-03-21 16:05 | 1553198717.8754 | 927 | 0.20949 | 0.000468 | -4e-06 | 100.3804 | 100.493 | 122.0733 | 21.2612 | 0.002029 | 0.001911 | -152.375 | 0.24125 | 8.398 | 0.003507 | 0.0045 | 0.004341 | 5.1e-05 | 0.001068 | 0.25642 | 1.3172 | 2.5481 | 98.85 | 20.81 | 38.82 | 70.6406 | 106.2663 | 67.8057 | 0.009664 | 1.0355 | 0.024593 | 23.7504 | -0.010023 | 0.14352 | -0.000126 | -0.009017 | -0.009017 | -0.65413
2019-03-21 16:05 | 1553198718.8741 | 928 | 0.20948 | 0.000467 | -3e-06 | 100.3747 | 100.1939 | 122.1967 | 21.2612 | 0.001971 | 0.001943 | -152.404 | 0.24086 | 8.4005 | 0.003487 | 0.004823 | 0.004604 | 5.2e-05 | 0.001072 | 0.30309 | 1.3651 | 2.548 | 98.85 | 20.81 | 38.825 | 69.2572 | 105.7768 | 68.5818 | 0.009665 | 1.0331 | 0.024577 | 23.7894 | 0.016129 | 0.14329 | -5.6e-05 | -0.003924 | -0.003924 | -1.5007
2019-03-21 16:05 | 1553198719.8731 | 929 | 0.20946 | 0.000467 | -4e-06 | 100.272 | 100.3727 | 122.4347 | 21.2612 | 0.001993 | 0.001943 | -152.4086 | 0.24142 | 8.4032 | 0.003554 | 0.005045 | 0.004927 | 5.2e-05 | 0.00107 | 0.29247 | 1.3141 | 2.5479 | 98.85 | 20.81 | 38.83 | 69.7694 | 107.3923 | 68.5717 | 0.009667 | 1.0331 | 0.024601 | 23.8122 | -0.008984 | 0.14362 | 7.4e-05 | 0.00569 | 0.00569 | 1.0374
2019-03-21 16:05 | 1553198720.8746 | 930 | 0.20946 | 0.000465 | -4e-06 | 100.3035 | 100.1149 | 122.3446 | 21.2612 | 0.002033 | 0.001906 | -152.4144 | 0.24187 | 8.4009 | 0.003564 | 0.004848 | 0.004628 | 5.1e-05 | 0.001069 | 0.27489 | 1.3466 | 2.5481 | 98.85 | 20.81 | 38.832 | 70.7466 | 107.6428 | 67.6872 | 0.009667 | 1.0358 | 0.024624 | 23.7726 | 0.016838 | 0.14389 | 3.5e-05 | 0.002804 | 0.002804 | 2.1089
2019-03-21 16:05 | 1553198721.8733 | 931 | 0.20947 | 0.000466 | -4e-06 | 100.2386 | 100.2181 | 122.6366 | 21.2612 | 0.001993 | 0.001938 | -152.4269 | 0.24212 | 8.4045 | 0.003572 | 0.005123 | 0.004971 | 5.1e-05 | 0.001071 | 0.25514 | 1.3182 | 2.548 | 98.85 | 20.81 | 38.834 | 69.786 | 107.8317 | 68.4647 | 0.009668 | 1.0335 | 0.024625 | 23.8281 | 0.001829 | 0.14404 | -4e-06 | -6.6e-05 | -6.6e-05 | -90.1095
2019-03-21 16:05 | 1553198722.8733 | 932 | 0.20948 | 0.000465 | -4e-06 | 100.2964 | 100.2849 | 122.7105 | 21.2612 | 0.002001 | 0.00195 | -152.4201 | 0.24142 | 8.4012 | 0.003517 | 0.00475 | 0.004611 | 4.9e-05 | 0.001068 | 0.29048 | 1.3028 | 2.5479 | 98.85 | 20.81 | 38.841 | 69.9792 | 106.5048 | 68.7503 | 0.009669 | 1.0326 | 0.024623 | 23.8453 | 0.001029 | 0.14362 | -6.7e-05 | -0.004668 | -0.004668 | -1.2646
2019-03-21 16:05 | 1553198723.8737 | 933 | 0.20944 | 0.000464 | -4e-06 | 100.2353 | 100.3145 | 122.8249 | 21.2612 | 0.00207 | 0.001925 | -152.4307 | 0.24133 | 8.401 | 0.003556 | 0.005441 | 0.005174 | 5.2e-05 | 0.001066 | 0.28152 | 1.32 | 2.548 | 98.85 | 20.81 | 38.848 | 71.6331 | 107.46 | 68.1505 | 0.009671 | 1.0344 | 0.024656 | 23.8354 | -0.007063 | 0.14356 | 0.000155 | 0.011762 | 0.011762 | 0.50166
2019-03-21 16:05 | 1553198724.8734 | 934 | 0.20948 | 0.000463 | -4e-06 | 100.2168 | 100.2675 | 122.9035 | 21.2612 | 0.002011 | 0.001948 | -152.4448 | 0.24146 | 8.403 | 0.003533 | 0.004855 | 0.00475 | 5.4e-05 | 0.001068 | 0.26913 | 1.3254 | 2.5478 | 98.85 | 20.82 | 38.861 | 70.2092 | 106.8937 | 68.7003 | 0.00968 | 1.0327 | 0.024644 | 23.8622 | -0.004519 | 0.14364 | -8.2e-05 | -0.005789 | -0.005789 | -1.0199
2019-03-21 16:05 | 1553198725.8734 | 935 | 0.20946 | 0.000465 | -4e-06 | 100.2711 | 100.2392 | 122.8479 | 21.2612 | 0.002038 | 0.001894 | -152.435 | 0.24107 | 8.4032 | 0.003543 | 0.004124 | 0.003988 | 4.9e-05 | 0.001066 | 0.31614 | 1.3064 | 2.548 | 98.85 | 20.82 | 38.869 | 70.8634 | 107.127 | 67.3922 | 0.009682 | 1.0367 | 0.024685 | 23.8112 | 0.002841 | 0.14341 | 4.6e-05 | 0.003681 | 0.003681 | 1.6013
2019-03-21 16:05 | 1553198726.8748 | 936 | 0.20947 | 0.000468 | -4e-06 | 100.3354 | 100.3962 | 123.061 | 21.2612 | 0.001984 | 0.001906 | -152.4456 | 0.24192 | 8.4036 | 0.003497 | 0.005188 | 0.00494 | 5.4e-05 | 0.001064 | 0.28174 | 1.3315 | 2.5479 | 98.85 | 20.82 | 38.88 | 69.5567 | 106.0194 | 67.6848 | 0.009685 | 1.0358 | 0.024696 | 23.842 | -0.005419 | 0.14392 | 2.9e-05 | 0.002387 | 0.002387 | 2.4782
2019-03-21 16:05 | 1553198727.874 | 937 | 0.20951 | 0.000466 | -4e-06 | 100.2213 | 100.1616 | 123.3326 | 21.2612 | 0.002009 | 0.001914 | -152.4165 | 0.24095 | 8.4027 | 0.003632 | 0.005099 | 0.004969 | 5.4e-05 | 0.001063 | 0.27975 | 1.3035 | 2.548 | 98.85 | 20.82 | 38.904 | 70.1525 | 109.3038 | 67.8721 | 0.009691 | 1.0353 | 0.024717 | 23.8749 | 0.005334 | 0.14334 | -0.000267 | -0.0195 | -0.0195 | -0.30212
2019-03-21 16:05 | 1553198728.8736 | 938 | 0.20948 | 0.000464 | -4e-06 | 100.3211 | 100.2678 | 123.2354 | 21.2612 | 0.002003 | 0.001909 | -152.3993 | 0.24103 | 8.4035 | 0.003587 | 0.004828 | 0.004661 | 5.5e-05 | 0.001068 | 0.28943 | 1.3484 | 2.548 | 98.85 | 20.82 | 38.914 | 70.03 | 108.1956 | 67.7536 | 0.009694 | 1.0356 | 0.024711 | 23.8613 | 0.004753 | 0.14339 | -5.8e-05 | -0.00403 | -0.00403 | -1.4625
2019-03-21 16:05 | 1553198729.874 | 939 | 0.20948 | 0.000462 | -4e-06 | 100.2556 | 100.2978 | 122.9457 | 21.2612 | 0.001989 | 0.001922 | -152.4041 | 0.24111 | 8.4012 | 0.003617 | 0.004848 | 0.00456 | 5.6e-05 | 0.001066 | 0.28655 | 1.3281 | 2.5479 | 98.85 | 20.82 | 38.929 | 69.6752 | 108.9428 | 68.0663 | 0.009697 | 1.0347 | 0.024671 | 23.8442 | -0.003766 | 0.14343 | -4.8e-05 | -0.00328 | -0.00328 | -1.7973
2019-03-21 16:06 | 1553198730.8751 | 940 | 0.20948 | 0.000462 | -4e-06 | 100.2553 | 100.2845 | 122.5772 | 21.2612 | 0.002021 | 0.001875 | -152.4097 | 0.242 | 8.4028 | 0.003675 | 0.00525 | 0.00514 | 5.2e-05 | 0.001069 | 0.29048 | 1.3346 | 2.548 | 98.85 | 20.82 | 38.935 | 70.4409 | 110.3344 | 66.9469 | 0.009699 | 1.0381 | 0.024674 | 23.7694 | -0.002608 | 0.14397 | -7.5e-05 | -0.005294 | -0.005294 | -1.1177
2019-03-21 16:06 | 1553198731.8748 | 941 | 0.20948 | 0.000466 | -4e-06 | 100.2612 | 100.3049 | 122.2341 | 21.2612 | 0.00204 | 0.00194 | -152.4077 | 0.24159 | 8.4021 | 0.003653 | 0.005164 | 0.004934 | 4.8e-05 | 0.001068 | 0.28982 | 1.3325 | 2.548 | 98.85 | 20.82 | 38.942 | 70.9033 | 109.8125 | 68.5147 | 0.009701 | 1.0333 | 0.024583 | 23.7907 | -0.003899 | 0.14372 | -4.8e-05 | -0.003274 | -0.003274 | -1.8043
2019-03-21 16:06 | 1553198732.8739 | 942 | 0.20947 | 0.000469 | -4e-06 | 100.2517 | 100.3703 | 122.1965 | 21.2612 | 0.001982 | 0.001937 | -152.3916 | 0.24146 | 8.4019 | 0.003598 | 0.005214 | 0.005097 | 5.2e-05 | 0.001067 | 0.30735 | 1.3289 | 2.548 | 98.85 | 20.82 | 38.949 | 69.5115 | 108.4621 | 68.4282 | 0.009702 | 1.0336 | 0.024583 | 23.784 | -0.010579 | 0.14364 | 2.7e-05 | 0.002229 | 0.002229 | 2.6483
2019-03-21 16:06 | 1553198733.8753 | 943 | 0.20947 | 0.000469 | -4e-06 | 100.2386 | 100.2276 | 122.0613 | 21.2612 | 0.00203 | 0.001951 | -152.3899 | 0.24194 | 8.4029 | 0.003604 | 0.005198 | 0.005027 | 5.3e-05 | 0.001063 | 0.30646 | 1.3349 | 2.548 | 98.85 | 20.82 | 38.949 | 70.6716 | 108.6071 | 68.7767 | 0.009702 | 1.0325 | 0.024556 | 23.783 | 0.000977 | 0.14393 | 1e-05 | 0.000973 | 0.000973 | 6.0791
2019-03-21 16:06 | 1553198734.8742 | 944 | 0.20946 | 0.000465 | -4e-06 | 100.3006 | 100.4017 | 122.254 | 21.2612 | 0.00204 | 0.001919 | -152.3775 | 0.2421 | 8.4031 | 0.00367 | 0.005448 | 0.005253 | 5.6e-05 | 0.001066 | 0.2808 | 1.3066 | 2.5478 | 98.85 | 20.82 | 38.947 | 70.911 | 110.2169 | 68.007 | 0.009702 | 1.0348 | 0.024603 | 23.775 | -0.009013 | 0.14402 | 6.8e-05 | 0.005281 | 0.005281 | 1.1209
2019-03-21 16:06 | 1553198735.874 | 945 | 0.20948 | 0.000464 | -4e-06 | 100.2839 | 100.3925 | 122.7542 | 21.2612 | 0.002036 | 0.001958 | -152.3924 | 0.2413 | 8.405 | 0.003613 | 0.004782 | 0.004589 | 5.3e-05 | 0.001063 | 0.26449 | 1.3222 | 2.548 | 98.85 | 20.82 | 38.932 | 70.8074 | 108.8245 | 68.939 | 0.009698 | 1.032 | 0.02462 | 23.8561 | -0.009682 | 0.14355 | -6.6e-05 | -0.00463 | -0.00463 | -1.2742
2019-03-21 16:06 | 1553198736.873 | 946 | 0.20946 | 0.000466 | -4e-06 | 100.3155 | 100.3537 | 122.7542 | 21.2612 | 0.002022 | 0.001987 | -152.3826 | 0.24166 | 8.4024 | 0.003643 | 0.005072 | 0.004908 | 5.2e-05 | 0.001064 | 0.26842 | 1.3714 | 2.5481 | 98.85 | 20.82 | 38.924 | 70.4666 | 109.5592 | 69.6408 | 0.009696 | 1.0299 | 0.024595 | 23.8805 | -0.003405 | 0.14377 | 4.4e-05 | 0.00352 | 0.00352 | 1.6784
2019-03-21 16:06 | 1553198737.8745 | 947 | 0.20944 | 0.000465 | -4e-06 | 100.274 | 100.4698 | 122.5589 | 21.2612 | 0.00201 | 0.001995 | -152.3915 | 0.24182 | 8.4055 | 0.00362 | 0.005321 | 0.005127 | 5.4e-05 | 0.001064 | 0.27455 | 1.3223 | 2.548 | 98.85 | 20.82 | 38.901 | 70.1905 | 109.0091 | 69.817 | 0.00969 | 1.0294 | 0.024569 | 23.8677 | -0.017444 | 0.14386 | 0.000203 | 0.015208 | 0.015208 | 0.38878
2019-03-21 16:06 | 1553198738.8742 | 948 | 0.20948 | 0.000469 | -4e-06 | 100.3497 | 100.3172 | 122.5688 | 21.2612 | 0.002009 | 0.001907 | -152.3889 | 0.24155 | 8.4022 | 0.003627 | 0.00518 | 0.005069 | 5.5e-05 | 0.001066 | 0.27079 | 1.3169 | 2.5481 | 98.85 | 20.83 | 38.886 | 70.1609 | 109.1721 | 67.706 | 0.009693 | 1.0358 | 0.024646 | 23.7951 | 0.002898 | 0.1437 | -5e-05 | -0.003447 | -0.003447 | -1.7136
2019-03-21 16:06 | 1553198739.8741 | 949 | 0.20949 | 0.000457 | -4e-06 | 100.35 | 100.3049 | 122.7195 | 21.2612 | 0.00204 | 0.00194 | -152.3843 | 0.24045 | 8.4003 | 0.003648 | 0.005055 | 0.004904 | 5.6e-05 | 0.001066 | 0.25393 | 1.3057 | 2.5479 | 98.85 | 20.83 | 38.871 | 70.9119 | 109.6846 | 68.5147 | 0.009689 | 1.0333 | 0.024632 | 23.8379 | 0.004018 | 0.14305 | -0.0001 | -0.007129 | -0.007129 | -0.82464
2019-03-21 16:06 | 1553198740.8742 | 950 | 0.20949 | 0.000464 | -4e-06 | 100.3068 | 100.5801 | 123.2103 | 21.2612 | 0.001993 | 0.001923 | -152.3905 | 0.24233 | 8.4049 | 0.003683 | 0.004958 | 0.004803 | 5.3e-05 | 0.001061 | 0.27383 | 1.2943 | 2.5481 | 98.85 | 20.83 | 38.843 | 69.7875 | 110.5435 | 68.0983 | 0.009682 | 1.0346 | 0.024696 | 23.871 | -0.02434 | 0.14416 | -0.000138 | -0.010014 | -0.010014 | -0.5917
2019-03-21 16:06 | 1553198741.8733 | 951 | 0.20948 | 0.000464 | -4e-06 | 100.339 | 100.3683 | 123.5985 | 21.2612 | 0.001989 | 0.001938 | -152.391 | 0.24132 | 8.4017 | 0.003634 | 0.005524 | 0.005325 | 5.5e-05 | 0.001066 | 0.26936 | 1.2765 | 2.5481 | 98.85 | 20.83 | 38.83 | 69.694 | 109.3502 | 68.4671 | 0.009679 | 1.0335 | 0.024722 | 23.9215 | -0.002612 | 0.14356 | -6.2e-05 | -0.004372 | -0.004372 | -1.3495
2019-03-21 16:06 | 1553198742.8733 | 952 | 0.20947 | 0.000468 | -4e-06 | 100.3003 | 100.51 | 123.4858 | 21.2612 | 0.002054 | 0.001925 | -152.4031 | 0.24128 | 8.4021 | 0.003605 | 0.004823 | 0.004654 | 5.5e-05 | 0.001065 | 0.27759 | 1.3306 | 2.5479 | 98.85 | 20.83 | 38.805 | 71.2556 | 108.6362 | 68.1474 | 0.009672 | 1.0344 | 0.024722 | 23.8994 | -0.018683 | 0.14354 | 1.3e-05 | 0.001203 | 0.001203 | 4.9052
2019-03-21 16:06 | 1553198743.8735 | 953 | 0.20949 | 0.000467 | -4e-06 | 100.2502 | 100.4722 | 123.5553 | 21.2612 | 0.00201 | 0.001874 | -152.3983 | 0.24118 | 8.4032 | 0.003553 | 0.005339 | 0.005236 | 5.3e-05 | 0.001066 | 0.27417 | 1.2879 | 2.548 | 98.85 | 20.83 | 38.792 | 70.1867 | 107.3842 | 66.9156 | 0.009669 | 1.0382 | 0.024774 | 23.8629 | -0.019783 | 0.14348 | -0.000153 | -0.011164 | -0.011164 | -0.52819
2019-03-21 16:06 | 1553198744.8744 | 954 | 0.20949 | 0.000465 | -4e-06 | 100.3441 | 100.3036 | 123.3959 | 21.2612 | 0.002069 | 0.001928 | -152.4116 | 0.24144 | 8.4029 | 0.003525 | 0.005323 | 0.00509 | 5.7e-05 | 0.001073 | 0.26698 | 1.3015 | 2.548 | 98.85 | 20.83 | 38.78 | 71.598 | 106.7033 | 68.2058 | 0.009666 | 1.0342 | 0.024711 | 23.8927 | 0.003608 | 0.14363 | -0.000116 | -0.00838 | -0.00838 | -0.70446
2019-03-21 16:06 | 1553198745.8745 | 955 | 0.20945 | 0.000468 | -4e-06 | 100.3712 | 100.2494 | 123.0469 | 21.2612 | 0.002032 | 0.001934 | -152.4127 | 0.24146 | 8.4029 | 0.003519 | 0.004655 | 0.00449 | 5.4e-05 | 0.00107 | 0.27887 | 1.2997 | 2.5479 | 98.85 | 20.83 | 38.758 | 70.713 | 106.5514 | 68.3705 | 0.009661 | 1.0337 | 0.02467 | 23.8646 | 0.010855 | 0.14365 | 0.000118 | 0.008991 | 0.008991 | 0.65663
2019-03-21 16:06 | 1553198746.8742 | 956 | 0.20949 | 0.000467 | -4e-06 | 100.3759 | 100.3254 | 122.9763 | 21.2612 | 0.002006 | 0.001956 | -152.4113 | 0.24093 | 8.4006 | 0.003517 | 0.004772 | 0.004583 | 5.4e-05 | 0.001074 | 0.29048 | 1.3075 | 2.5479 | 98.85 | 20.83 | 38.748 | 70.0814 | 106.5065 | 68.8844 | 0.009658 | 1.0322 | 0.024644 | 23.8757 | 0.004506 | 0.14333 | -0.000112 | -0.008053 | -0.008053 | -0.7315
2019-03-21 16:06 | 1553198747.8735 | 957 | 0.2095 | 0.00046 | -3e-06 | 100.3742 | 100.3329 | 122.8915 | 21.2612 | 0.002042 | 0.001943 | -152.4005 | 0.24157 | 8.4029 | 0.003674 | 0.004446 | 0.004261 | 5.4e-05 | 0.001071 | 0.26825 | 1.3165 | 2.548 | 98.85 | 20.83 | 38.739 | 70.9554 | 110.3195 | 68.5803 | 0.009656 | 1.0331 | 0.024647 | 23.8569 | 0.003679 | 0.14371 | -0.000175 | -0.012709 | -0.012709 | -0.46473
2019-03-21 16:06 | 1553198748.8736 | 958 | 0.20949 | 0.000466 | -4e-06 | 100.3902 | 100.3325 | 122.9624 | 21.2612 | 0.002023 | 0.001953 | -152.4095 | 0.24178 | 8.4033 | 0.003604 | 0.00454 | 0.004374 | 5.8e-05 | 0.001075 | 0.28711 | 1.3017 | 2.548 | 98.85 | 20.83 | 38.722 | 70.5041 | 108.6063 | 68.808 | 0.009652 | 1.0324 | 0.024646 | 23.8717 | 0.005143 | 0.14384 | -0.000115 | -0.00826 | -0.00826 | -0.71571
2019-03-21 16:06 | 1553198749.8727 | 959 | 0.20949 | 0.000467 | -4e-06 | 100.4287 | 100.3097 | 123.1331 | 21.2612 | 0.002088 | 0.001925 | -152.3935 | 0.24115 | 8.4025 | 0.003538 | 0.004543 | 0.004367 | 5.6e-05 | 0.001069 | 0.26145 | 1.3597 | 2.5479 | 98.84 | 20.83 | 38.714 | 72.0541 | 107.0245 | 68.1551 | 0.009651 | 1.0344 | 0.024686 | 23.8655 | 0.010601 | 0.14346 | -0.000109 | -0.007869 | -0.007869 | -0.74926
2019-03-21 16:06 | 1553198750.8727 | 960 | 0.20948 | 0.000464 | -4e-06 | 100.3393 | 100.4391 | 123.1953 | 21.2612 | 0.002022 | 0.001896 | -152.382 | 0.24141 | 8.4056 | 0.003537 | 0.005088 | 0.004941 | 5.6e-05 | 0.001071 | 0.24403 | 1.3744 | 2.548 | 98.84 | 20.83 | 38.701 | 70.4784 | 106.9954 | 67.4398 | 0.009647 | 1.0366 | 0.024719 | 23.8465 | -0.008896 | 0.14361 | -5.1e-05 | -0.003567 | -0.003567 | -1.6547
2019-03-21 16:06 | 1553198751.8738 | 961 | 0.2095 | 0.000466 | -4e-06 | 100.3917 | 100.48 | 123.1091 | 21.2612 | 0.002036 | 0.001914 | -152.3755 | 0.24214 | 8.4039 | 0.003526 | 0.004532 | 0.004356 | 5.7e-05 | 0.001072 | 0.26842 | 1.3017 | 2.548 | 98.84 | 20.83 | 38.696 | 70.8012 | 106.7103 | 67.8743 | 0.009646 | 1.0353 | 0.024694 | 23.8533 | -0.007862 | 0.14405 | -0.000212 | -0.015481 | -0.015481 | -0.38244
2019-03-21 16:06 | 1553198752.8728 | 962 | 0.20947 | 0.000464 | -4e-06 | 100.426 | 100.4977 | 123.2868 | 21.2612 | 0.001983 | 0.001915 | -152.3933 | 0.2406 | 8.4047 | 0.003585 | 0.005186 | 0.005085 | 5.5e-05 | 0.00107 | 0.27505 | 1.314 | 2.5479 | 98.84 | 20.83 | 38.689 | 69.5394 | 108.1641 | 67.8954 | 0.009644 | 1.0352 | 0.024711 | 23.8713 | -0.006386 | 0.14313 | -2.9e-05 | -0.001953 | -0.001953 | -3.0127
2019-03-21 16:06 | 1553198753.8739 | 963 | 0.20947 | 0.000465 | -4e-06 | 100.4177 | 100.5761 | 123.1575 | 21.2612 | 0.001973 | 0.001886 | -152.3806 | 0.24105 | 8.4015 | 0.0035 | 0.004769 | 0.004598 | 5.4e-05 | 0.00107 | 0.31454 | 1.3675 | 2.5481 | 98.84 | 20.83 | 38.687 | 69.3048 | 106.0917 | 67.2058 | 0.009644 | 1.0373 | 0.024723 | 23.8346 | -0.014099 | 0.1434 | -2.8e-05 | -0.001805 | -0.001805 | -3.2654
2019-03-21 16:06 | 1553198754.8739 | 964 | 0.2095 | 0.000464 | -4e-06 | 100.3828 | 100.5481 | 123.1523 | 21.2612 | 0.002029 | 0.00194 | -152.3732 | 0.24167 | 8.4022 | 0.003573 | 0.00496 | 0.004793 | 5e-05 | 0.00107 | 0.29645 | 1.3539 | 2.548 | 98.84 | 20.83 | 38.685 | 70.6351 | 107.8744 | 68.5101 | 0.009643 | 1.0333 | 0.024675 | 23.8797 | -0.01472 | 0.14377 | -0.000183 | -0.013266 | -0.013266 | -0.44543
2019-03-21 16:06 | 1553198755.8744 | 965 | 0.20938 | 0.000466 | -4e-06 | 100.3175 | 100.5096 | 123.4668 | 21.2612 | 0.002011 | 0.001928 | -152.3659 | 0.24097 | 8.401 | 0.003497 | 0.004636 | 0.004411 | 5.4e-05 | 0.001073 | 0.25393 | 1.3455 | 2.5481 | 98.84 | 20.83 | 38.683 | 70.2124 | 106.0296 | 68.2221 | 0.009643 | 1.0342 | 0.024717 | 23.9001 | -0.017112 | 0.14336 | 0.000535 | 0.039975 | 0.039975 | 0.14739
2019-03-21 16:06 | 1553198756.8733 | 966 | 0.2095 | 0.000462 | -4e-06 | 100.2666 | 100.3601 | 123.9837 | 21.2612 | 0.001996 | 0.001921 | -152.3683 | 0.24132 | 8.4041 | 0.003516 | 0.004953 | 0.00479 | 5e-05 | 0.001071 | 0.262 | 1.3563 | 2.548 | 98.84 | 20.84 | 38.682 | 69.8593 | 106.4766 | 68.0435 | 0.009649 | 1.0347 | 0.024776 | 23.9438 | -0.00834 | 0.14356 | -0.000169 | -0.012309 | -0.012309 | -0.47934
2019-03-21 16:06 | 1553198757.8735 | 967 | 0.20947 | 0.000463 | -4e-06 | 100.406 | 100.4766 | 123.9642 | 21.2612 | 0.002031 | 0.001945 | -152.3793 | 0.24068 | 8.4021 | 0.003532 | 0.00433 | 0.004171 | 4.7e-05 | 0.001074 | 0.27743 | 1.3263 | 2.548 | 98.84 | 20.83 | 38.682 | 70.6966 | 106.8622 | 68.6248 | 0.009643 | 1.033 | 0.024753 | 23.9624 | -0.006285 | 0.14318 | -2.8e-05 | -0.001828 | -0.001828 | -3.2191
2019-03-21 16:06 | 1553198758.8742 | 968 | 0.20938 | 0.000468 | -3e-06 | 100.3259 | 100.4337 | 124.0438 | 21.2612 | 0.002005 | 0.001857 | -152.3871 | 0.24111 | 8.4016 | 0.003508 | 0.004286 | 0.004095 | 4.8e-05 | 0.001075 | 0.26753 | 1.3035 | 2.5479 | 98.85 | 20.84 | 38.681 | 70.0619 | 106.2813 | 66.5051 | 0.009647 | 1.0394 | 0.024838 | 23.8956 | -0.009606 | 0.14343 | 0.000525 | 0.03938 | 0.03938 | 0.1497
2019-03-21 16:06 | 1553198759.8733 | 969 | 0.20949 | 0.000466 | -3e-06 | 100.361 | 100.5274 | 123.6849 | 21.2612 | 0.002019 | 0.00188 | -152.4092 | 0.24072 | 8.4017 | 0.003565 | 0.005549 | 0.005341 | 4.9e-05 | 0.001071 | 0.26537 | 1.3383 | 2.5481 | 98.85 | 20.84 | 38.68 | 70.3996 | 107.6646 | 67.0717 | 0.009647 | 1.0377 | 0.024781 | 23.8809 | -0.01481 | 0.1432 | -0.000146 | -0.010619 | -0.010619 | -0.55425
2019-03-21 16:06 | 1553198760.8744 | 970 | 0.20944 | 0.000465 | -4e-06 | 100.3771 | 100.3611 | 123.3265 | 21.2612 | 0.00197 | 0.001848 | -152.4328 | 0.24104 | 8.4034 | 0.003571 | 0.005239 | 0.005025 | 5e-05 | 0.00107 | 0.25299 | 1.303 | 2.5479 | 98.85 | 20.84 | 38.678 | 69.2315 | 107.82 | 66.3048 | 0.009647 | 1.04 | 0.024773 | 23.8194 | 0.001425 | 0.14339 | 0.000213 | 0.016151 | 0.016151 | 0.36491
2019-03-21 16:06 | 1553198761.8737 | 971 | 0.20952 | 0.000467 | -4e-06 | 100.4099 | 100.4551 | 123.5675 | 21.2612 | 0.001922 | 0.00189 | -152.4437 | 0.2404 | 8.4003 | 0.003547 | 0.00479 | 0.004699 | 4.6e-05 | 0.001073 | 0.2688 | 1.2874 | 2.5478 | 98.85 | 20.84 | 38.68 | 68.0702 | 107.2221 | 67.3032 | 0.009647 | 1.037 | 0.024761 | 23.8777 | -0.004029 | 0.14302 | -0.000281 | -0.020623 | -0.020623 | -0.28501
2019-03-21 16:06 | 1553198762.8735 | 972 | 0.20947 | 0.000464 | -4e-06 | 100.3849 | 100.5427 | 123.7957 | 21.2612 | 0.001973 | 0.001934 | -152.4344 | 0.24015 | 8.4033 | 0.0035 | 0.005035 | 0.00484 | 5.1e-05 | 0.001078 | 0.27328 | 1.3133 | 2.548 | 98.85 | 20.84 | 38.682 | 69.2938 | 106.0948 | 68.3563 | 0.009648 | 1.0338 | 0.024745 | 23.9367 | -0.014049 | 0.14286 | -1.2e-05 | -0.000644 | -0.000644 | -9.1133
2019-03-21 16:06 | 1553198763.8735 | 973 | 0.20943 | 0.000468 | -4e-06 | 100.3518 | 100.4623 | 123.8527 | 21.2612 | 0.002039 | 0.001866 | -152.4316 | 0.24149 | 8.4019 | 0.003514 | 0.004658 | 0.004533 | 4.9e-05 | 0.001077 | 0.25946 | 1.3167 | 2.548 | 98.85 | 20.84 | 38.686 | 70.8769 | 106.4191 | 66.7345 | 0.009649 | 1.0387 | 0.02481 | 23.8852 | -0.009843 | 0.14366 | 0.000267 | 0.020123 | 0.020123 | 0.29342
2019-03-21 16:06 | 1553198764.8739 | 974 | 0.20947 | 0.000463 | -4e-06 | 100.3065 | 100.2246 | 123.8363 | 21.2612 | 0.002047 | 0.001896 | -152.4336 | 0.24129 | 8.4036 | 0.003574 | 0.004466 | 0.004267 | 4.7e-05 | 0.001079 | 0.26753 | 1.3339 | 2.548 | 98.85 | 20.84 | 38.688 | 71.0739 | 107.8814 | 67.4376 | 0.009649 | 1.0366 | 0.024783 | 23.9084 | 0.00731 | 0.14354 | 2e-05 | 0.001772 | 0.001772 | 3.3293
2019-03-21 16:06 | 1553198765.8736 | 975 | 0.20949 | 0.000462 | -5e-06 | 100.3268 | 100.3332 | 123.4851 | 21.2612 | 0.001921 | 0.001941 | -152.4314 | 0.2412 | 8.4009 | 0.003551 | 0.004912 | 0.00465 | 4.9e-05 | 0.001076 | 0.2855 | 1.2805 | 2.5479 | 98.85 | 20.84 | 38.693 | 68.049 | 107.3316 | 68.538 | 0.00965 | 1.0332 | 0.024708 | 23.913 | -0.000574 | 0.14349 | -9.2e-05 | -0.006511 | -0.006511 | -0.90573
2019-03-21 16:06 | 1553198766.8737 | 976 | 0.20947 | 0.000463 | -4e-06 | 100.3914 | 100.3516 | 123.1014 | 21.2612 | 0.00198 | 0.001906 | -152.4286 | 0.2414 | 8.3979 | 0.003535 | 0.005232 | 0.00505 | 4.5e-05 | 0.00108 | 0.28102 | 1.3054 | 2.548 | 98.85 | 20.84 | 38.696 | 69.4716 | 106.9449 | 67.6862 | 0.009651 | 1.0358 | 0.0247 | 23.846 | 0.003549 | 0.14361 | -8e-06 | -0.000338 | -0.000338 | -17.4493
2019-03-21 16:06 | 1553198767.8731 | 977 | 0.20946 | 0.000462 | -4e-06 | 100.3026 | 100.4068 | 123.1842 | 21.2612 | 0.002011 | 0.001925 | -152.4317 | 0.24159 | 8.4035 | 0.00352 | 0.00478 | 0.00455 | 4.9e-05 | 0.001079 | 0.26576 | 1.3263 | 2.548 | 98.85 | 20.84 | 38.703 | 70.2109 | 106.5725 | 68.1397 | 0.009653 | 1.0344 | 0.024692 | 23.8699 | -0.009282 | 0.14372 | 4.7e-05 | 0.003767 | 0.003767 | 1.5679
2019-03-21 16:06 | 1553198768.8738 | 978 | 0.20945 | 0.000465 | -4e-06 | 100.3685 | 100.5093 | 123.2145 | 21.2612 | 0.002002 | 0.001915 | -152.4159 | 0.24173 | 8.4017 | 0.00353 | 0.004871 | 0.004751 | 4.8e-05 | 0.001075 | 0.27831 | 1.3212 | 2.548 | 98.85 | 20.84 | 38.727 | 70.0066 | 106.8285 | 67.911 | 0.009659 | 1.0351 | 0.024703 | 23.8648 | -0.01254 | 0.1438 | 0.000121 | 0.009199 | 0.009199 | 0.64252
2019-03-21 16:06 | 1553198769.8728 | 979 | 0.20949 | 0.000461 | -5e-06 | 100.3062 | 100.247 | 123.3337 | 21.2612 | 0.001985 | 0.001874 | -152.409 | 0.24212 | 8.4004 | 0.003519 | 0.004472 | 0.004219 | 4.3e-05 | 0.001075 | 0.26969 | 1.3236 | 2.5479 | 98.85 | 20.84 | 38.741 | 69.5901 | 106.5599 | 66.9139 | 0.009662 | 1.0382 | 0.024752 | 23.8414 | 0.005278 | 0.14404 | -0.000123 | -0.008887 | -0.008887 | -0.66611
2019-03-21 16:06 | 1553198770.8733 | 980 | 0.20947 | 0.000462 | -4e-06 | 100.277 | 100.2692 | 123.4272 | 21.2612 | 0.001975 | 0.001904 | -152.3652 | 0.24164 | 8.4014 | 0.003458 | 0.004866 | 0.004669 | 4.7e-05 | 0.001078 | 0.274 | 1.3773 | 2.548 | 98.85 | 20.84 | 38.758 | 69.3462 | 105.0618 | 67.6511 | 0.009667 | 1.0359 | 0.024734 | 23.8763 | 0.000699 | 0.14375 | -2e-05 | -0.001229 | -0.001229 | -4.8079
2019-03-21 16:06 | 1553198771.8743 | 981 | 0.20948 | 0.000462 | -4e-06 | 100.2863 | 100.3012 | 123.2887 | 21.2612 | 0.001942 | 0.001903 | -152.3707 | 0.24088 | 8.4021 | 0.003513 | 0.004443 | 0.004269 | 5e-05 | 0.001076 | 0.29463 | 1.3494 | 2.5482 | 98.84 | 20.84 | 38.792 | 68.556 | 106.4048 | 67.6146 | 0.009676 | 1.036 | 0.024722 | 23.8616 | -0.001333 | 0.1433 | -4.1e-05 | -0.00274 | -0.00274 | -2.1496
2019-03-21 16:06 | 1553198772.8729 | 982 | 0.20945 | 0.000464 | -4e-06 | 100.3137 | 100.3806 | 123.0441 | 21.2612 | 0.001943 | 0.001934 | -152.3644 | 0.24151 | 8.4029 | 0.003503 | 0.004913 | 0.004743 | 4.7e-05 | 0.001073 | 0.28816 | 1.3449 | 2.548 | 98.84 | 20.84 | 38.809 | 68.5762 | 106.1554 | 68.3712 | 0.00968 | 1.0337 | 0.02467 | 23.8644 | -0.005962 | 0.14368 | 0.000128 | 0.009709 | 0.009709 | 0.60819
2019-03-21 16:06 | 1553198773.873 | 983 | 0.20943 | 0.000463 | -4e-06 | 100.2997 | 100.4572 | 123.0176 | 21.2612 | 0.002006 | 0.001891 | -152.376 | 0.24147 | 8.4017 | 0.003505 | 0.004809 | 0.004649 | 4.7e-05 | 0.001076 | 0.27256 | 1.3195 | 2.548 | 98.84 | 20.84 | 38.835 | 70.0884 | 106.2224 | 67.322 | 0.009687 | 1.0369 | 0.024705 | 23.8251 | -0.014037 | 0.14365 | 0.000251 | 0.018857 | 0.018857 | 0.31311
2019-03-21 16:06 | 1553198774.8745 | 984 | 0.20948 | 0.000464 | -4e-06 | 100.3059 | 100.6227 | 122.9368 | 21.2612 | 0.00204 | 0.001887 | -152.369 | 0.24205 | 8.4023 | 0.003451 | 0.004916 | 0.004758 | 4.7e-05 | 0.001074 | 0.28495 | 1.3167 | 2.5479 | 98.84 | 20.84 | 38.843 | 70.9026 | 104.8974 | 67.2361 | 0.009689 | 1.0372 | 0.0247 | 23.8143 | -0.028205 | 0.144 | -3.4e-05 | -0.002292 | -0.002292 | -2.582
2019-03-21 16:06 | 1553198775.8742 | 985 | 0.20943 | 0.000466 | -4e-06 | 100.2448 | 100.5624 | 123.1281 | 21.2612 | 0.001962 | 0.001858 | -152.3743 | 0.24069 | 8.404 | 0.003458 | 0.004175 | 0.003976 | 5.3e-05 | 0.001079 | 0.27704 | 1.2798 | 2.548 | 98.84 | 20.84 | 38.85 | 69.0397 | 105.0681 | 66.5388 | 0.009691 | 1.0393 | 0.024745 | 23.8084 | -0.028293 | 0.14319 | 0.000271 | 0.020387 | 0.020387 | 0.28867
2019-03-21 16:06 | 1553198776.8741 | 986 | 0.20945 | 0.000464 | -4e-06 | 100.277 | 100.3724 | 123.0464 | 21.2612 | 0.001948 | 0.001897 | -152.3916 | 0.24239 | 8.4006 | 0.003486 | 0.004734 | 0.004582 | 5e-05 | 0.001073 | 0.274 | 1.3082 | 2.5478 | 98.84 | 20.84 | 38.856 | 68.6996 | 105.7431 | 67.4756 | 0.009692 | 1.0365 | 0.024702 | 23.8333 | -0.008502 | 0.14419 | 0.000117 | 0.008962 | 0.008962 | 0.6613
2019-03-21 16:06 | 1553198777.8736 | 987 | 0.20943 | 0.000459 | -4e-06 | 100.2779 | 100.1275 | 123.0831 | 21.2612 | 0.001954 | 0.001862 | -152.3967 | 0.24049 | 8.402 | 0.003477 | 0.004451 | 0.004216 | 5e-05 | 0.001074 | 0.25658 | 1.3095 | 2.5479 | 98.84 | 20.84 | 38.863 | 68.8486 | 105.5212 | 66.6299 | 0.009694 | 1.039 | 0.024737 | 23.8073 | 0.013427 | 0.14307 | 0.000252 | 0.019002 | 0.019002 | 0.30944
2019-03-21 16:06 | 1553198778.8743 | 988 | 0.20944 | 0.000464 | -4e-06 | 100.2392 | 100.3121 | 123.264 | 21.2612 | 0.001997 | 0.001933 | -152.4017 | 0.2416 | 8.4001 | 0.003505 | 0.004887 | 0.004713 | 4.9e-05 | 0.001072 | 0.31509 | 1.3249 | 2.5479 | 98.84 | 20.84 | 38.865 | 69.8733 | 106.2105 | 68.333 | 0.009694 | 1.0339 | 0.024693 | 23.8844 | -0.006504 | 0.14373 | 0.000151 | 0.011409 | 0.011409 | 0.51779
2019-03-21 16:06 | 1553198779.8733 | 989 | 0.20947 | 0.000466 | -4e-06 | 100.2165 | 100.1956 | 123.4286 | 21.2612 | 0.001952 | 0.001909 | -152.3997 | 0.24152 | 8.4023 | 0.00348 | 0.004961 | 0.004807 | 5e-05 | 0.001071 | 0.24961 | 1.3482 | 2.548 | 98.84 | 20.84 | 38.863 | 68.7931 | 105.6163 | 67.7637 | 0.009694 | 1.0356 | 0.02473 | 23.8804 | 0.001867 | 0.14368 | -1e-06 | 0.000206 | 0.000206 | 28.6456
2019-03-21 16:06 | 1553198780.8735 | 990 | 0.20948 | 0.000465 | -4e-06 | 100.2183 | 100.4058 | 123.5851 | 21.2612 | 0.002035 | 0.001868 | -152.4034 | 0.24158 | 8.4008 | 0.003471 | 0.005107 | 0.00484 | 4.8e-05 | 0.001071 | 0.28711 | 1.3391 | 2.548 | 98.84 | 20.84 | 38.847 | 70.7942 | 105.3842 | 66.7797 | 0.00969 | 1.0386 | 0.024782 | 23.861 | -0.016713 | 0.14372 | -8.2e-05 | -0.005844 | -0.005844 | -1.0107
2019-03-21 16:06 | 1553198781.8735 | 991 | 0.20949 | 0.000466 | -5e-06 | 100.3339 | 100.2978 | 123.5257 | 21.2612 | 0.001998 | 0.00195 | -152.3873 | 0.24041 | 8.4043 | 0.003378 | 0.004373 | 0.004274 | 5.3e-05 | 0.001069 | 0.23667 | 1.3416 | 2.5478 | 98.84 | 20.84 | 38.837 | 69.8896 | 103.131 | 68.751 | 0.009687 | 1.0326 | 0.024704 | 23.9244 | 0.003222 | 0.14302 | -0.000135 | -0.009732 | -0.009732 | -0.604
2019-03-21 16:06 | 1553198782.8735 | 992 | 0.20947 | 0.000463 | -5e-06 | 100.2785 | 100.2723 | 123.364 | 21.2612 | 0.001965 | 0.001931 | -152.3959 | 0.24092 | 8.4025 | 0.003448 | 0.005684 | 0.005494 | 5.3e-05 | 0.001073 | 0.30237 | 1.3406 | 2.548 | 98.84 | 20.84 | 38.813 | 69.0974 | 104.8336 | 68.2964 | 0.009681 | 1.034 | 0.024704 | 23.8928 | 0.000558 | 0.14333 | 3e-05 | 0.002511 | 0.002511 | 2.3455
2019-03-21 16:06 | 1553198783.8725 | 993 | 0.20946 | 0.000456 | -5e-06 | 100.319 | 100.3663 | 123.3502 | 21.2612 | 0.001991 | 0.001912 | -152.4061 | 0.24119 | 8.4004 | 0.003491 | 0.004376 | 0.004222 | 4.9e-05 | 0.00107 | 0.25426 | 1.3355 | 2.5478 | 98.84 | 20.84 | 38.8 | 69.7336 | 105.8799 | 67.8401 | 0.009678 | 1.0354 | 0.02472 | 23.8755 | -0.004209 | 0.14348 | 8.6e-05 | 0.006676 | 0.006676 | 0.88338
2019-03-21 16:06 | 1553198784.873 | 994 | 0.20947 | 0.000461 | -4e-06 | 100.2612 | 100.2413 | 123.3614 | 21.2612 | 0.002015 | 0.001891 | -152.3962 | 0.24023 | 8.4037 | 0.00351 | 0.005015 | 0.004879 | 5.2e-05 | 0.001073 | 0.27793 | 1.3593 | 2.548 | 98.84 | 20.84 | 38.786 | 70.3013 | 106.3255 | 67.3275 | 0.009675 | 1.0369 | 0.024739 | 23.8586 | 0.001782 | 0.14291 | -2.1e-05 | -0.001249 | -0.001249 | -4.7017
2019-03-21 16:06 | 1553198785.8726 | 995 | 0.20948 | 0.000461 | -4e-06 | 100.2332 | 100.1071 | 123.1206 | 21.2612 | 0.001989 | 0.001902 | -152.4119 | 0.24084 | 8.4027 | 0.00344 | 0.005309 | 0.005104 | 4.8e-05 | 0.001071 | 0.25514 | 1.3732 | 2.5478 | 98.84 | 20.84 | 38.762 | 69.6807 | 104.6363 | 67.592 | 0.009669 | 1.0361 | 0.024706 | 23.8446 | 0.011265 | 0.14327 | -7.9e-05 | -0.005591 | -0.005591 | -1.0532
2019-03-21 16:06 | 1553198786.8941 | 996 | 0.20946 | 0.000466 | -4e-06 | 100.342 | 100.4041 | 123.0953 | 21.2612 | 0.00197 | 0.001839 | -152.3977 | 0.24085 | 8.4021 | 0.003492 | 0.005104 | 0.004955 | 5.2e-05 | 0.001073 | 0.29031 | 1.3886 | 2.548 | 98.84 | 20.84 | 38.751 | 69.2322 | 105.8848 | 66.069 | 0.009666 | 1.0408 | 0.024758 | 23.7888 | -0.005533 | 0.14328 | 4.2e-05 | 0.003327 | 0.003327 | 1.77
2019-03-21 16:06 | 1553198787.8816 | 997 | 0.2095 | 0.000467 | -4e-06 | 100.3077 | 100.3802 | 123.0748 | 21.2612 | 0.002011 | 0.001871 | -152.383 | 0.24183 | 8.3997 | 0.003382 | 0.00406 | 0.003909 | 5e-05 | 0.001074 | 0.26394 | 1.3916 | 2.548 | 98.84 | 20.84 | 38.734 | 70.2133 | 103.2339 | 66.8454 | 0.009662 | 1.0384 | 0.024728 | 23.814 | -0.006463 | 0.14386 | -0.000199 | -0.014546 | -0.014546 | -0.4065
2019-03-21 16:06 | 1553198788.8826 | 998 | 0.20949 | 0.000465 | -4e-06 | 100.2836 | 100.2198 | 122.8152 | 21.2612 | 0.00202 | 0.001872 | -152.4021 | 0.24113 | 8.4023 | 0.003466 | 0.004281 | 0.004086 | 5.1e-05 | 0.001072 | 0.29374 | 1.3745 | 2.5482 | 98.84 | 20.84 | 38.725 | 70.4229 | 105.2749 | 66.8697 | 0.009659 | 1.0383 | 0.024701 | 23.7897 | 0.00569 | 0.14345 | -0.000106 | -0.007594 | -0.007594 | -0.77633
2019-03-21 16:06 | 1553198789.8818 | 999 | 0.2095 | 0.000465 | -4e-06 | 100.2818 | 100.3451 | 122.6711 | 21.2612 | 0.00197 | 0.001915 | -152.4182 | 0.24203 | 8.4034 | 0.003447 | 0.00538 | 0.005165 | 5.1e-05 | 0.001076 | 0.27759 | 1.3257 | 2.5479 | 98.84 | 20.84 | 38.718 | 69.237 | 104.8147 | 67.911 | 0.009658 | 1.0351 | 0.024649 | 23.8121 | -0.005649 | 0.14398 | -0.000177 | -0.012863 | -0.012863 | -0.46005
2019-03-21 16:06 | 1553198790.8822 | 1000 | 0.20945 | 0.000467 | -4e-06 | 100.302 | 100.134 | 122.3819 | 21.2612 | 0.00197 | 0.001834 | -152.4707 | 0.24071 | 8.4001 | 0.003453 | 0.004745 | 0.004591 | 5e-05 | 0.001076 | 0.28998 | 1.3366 | 2.548 | 98.84 | 20.84 | 38.705 | 69.2361 | 104.9524 | 65.9473 | 0.009654 | 1.0411 | 0.024691 | 23.7155 | 0.015001 | 0.1432 | 0.000102 | 0.007842 | 0.007842 | 0.75051
2019-03-21 16:07 | 1553198791.8832 | 1001 | 0.20947 | 0.000467 | -4e-06 | 100.3166 | 100.3468 | 122.3134 | 21.2612 | 0.001955 | 0.001886 | -152.4774 | 0.2417 | 8.4029 | 0.003423 | 0.004849 | 0.004652 | 5.2e-05 | 0.001079 | 0.27688 | 1.3466 | 2.5479 | 98.84 | 20.84 | 38.702 | 68.8594 | 104.2345 | 67.2066 | 0.009654 | 1.0373 | 0.024638 | 23.7528 | -0.002692 | 0.14379 | 2.7e-05 | 0.002251 | 0.002251 | 2.625
2019-03-21 16:07 | 1553198792.8829 | 1002 | 0.20946 | 0.000468 | -4e-06 | 100.2678 | 100.3642 | 122.5606 | 21.2612 | 0.001922 | 0.001903 | -152.4666 | 0.24073 | 8.4049 | 0.00338 | 0.005258 | 0.005121 | 5.5e-05 | 0.001078 | 0.24818 | 1.3251 | 2.5478 | 98.84 | 20.85 | 38.7 | 68.0663 | 103.1937 | 67.6069 | 0.009659 | 1.0361 | 0.024649 | 23.7908 | -0.008598 | 0.14321 | 5.7e-05 | 0.004456 | 0.004456 | 1.3208
2019-03-21 16:07 | 1553198793.883 | 1003 | 0.20946 | 0.000465 | -5e-06 | 100.2889 | 100.3039 | 122.8965 | 21.2612 | 0.001989 | 0.001917 | -152.46 | 0.24125 | 8.4029 | 0.003435 | 0.005393 | 0.005254 | 4.7e-05 | 0.001083 | 0.26322 | 1.2992 | 2.5479 | 98.84 | 20.85 | 38.695 | 69.6829 | 104.5113 | 67.9548 | 0.009658 | 1.035 | 0.02467 | 23.8355 | -0.001337 | 0.14352 | 7.6e-05 | 0.005948 | 0.005948 | 0.99162
2019-03-21 16:07 | 1553198794.8834 | 1004 | 0.20948 | 0.000462 | -4e-06 | 100.2508 | 100.4024 | 123.0955 | 21.2612 | 0.002007 | 0.001906 | -152.4732 | 0.24125 | 8.3997 | 0.003426 | 0.004658 | 0.004416 | 4.7e-05 | 0.001083 | 0.27815 | 1.3555 | 2.5482 | 98.84 | 20.85 | 38.693 | 70.1119 | 104.2864 | 67.6886 | 0.009657 | 1.0358 | 0.0247 | 23.8455 | -0.013512 | 0.14352 | -4.4e-05 | -0.003032 | -0.003032 | -1.9456
2019-03-21 16:07 | 1553198795.8831 | 1005 | 0.20946 | 0.000467 | -4e-06 | 100.302 | 100.2634 | 123.5124 | 21.2612 | 0.002003 | 0.001906 | -152.4673 | 0.24217 | 8.3993 | 0.003423 | 0.005341 | 0.005117 | 4.8e-05 | 0.001083 | 0.27599 | 1.3212 | 2.5481 | 98.84 | 20.85 | 38.692 | 70.0307 | 104.2338 | 67.6894 | 0.009657 | 1.0358 | 0.024741 | 23.8859 | 0.003447 | 0.14407 | 5.7e-05 | 0.004482 | 0.004482 | 1.321
2019-03-21 16:07 | 1553198796.8821 | 1006 | 0.20947 | 0.000462 | -4e-06 | 100.2696 | 100.2685 | 123.5764 | 21.2612 | 0.001971 | 0.001923 | -152.4674 | 0.24131 | 8.4006 | 0.003403 | 0.004878 | 0.004686 | 4.8e-05 | 0.001088 | 0.28943 | 1.2998 | 2.5479 | 98.84 | 20.85 | 38.694 | 69.2604 | 103.7424 | 68.1036 | 0.009658 | 1.0346 | 0.024733 | 23.9066 | 9.5e-05 | 0.14355 | -1.8e-05 | -0.001072 | -0.001072 | -5.5015
2019-03-21 16:07 | 1553198797.8816 | 1007 | 0.20945 | 0.000465 | -4e-06 | 100.3003 | 100.5314 | 123.7945 | 21.2612 | 0.001971 | 0.001878 | -152.4905 | 0.24097 | 8.4018 | 0.003445 | 0.005057 | 0.004871 | 5e-05 | 0.001089 | 0.27865 | 1.3551 | 2.5479 | 98.84 | 20.85 | 38.697 | 69.2503 | 104.7566 | 67.0101 | 0.009658 | 1.0379 | 0.024794 | 23.8893 | -0.020592 | 0.14336 | 0.000126 | 0.009578 | 0.009578 | 0.61513
2019-03-21 16:07 | 1553198798.8821 | 1008 | 0.20946 | 0.000461 | -4e-06 | 100.3128 | 100.4569 | 123.8433 | 21.2612 | 0.001998 | 0.001877 | -152.4728 | 0.24043 | 8.4023 | 0.003462 | 0.003902 | 0.003715 | 4.9e-05 | 0.001086 | 0.27273 | 1.353 | 2.5479 | 98.84 | 20.85 | 38.711 | 69.9014 | 105.1726 | 66.9858 | 0.009662 | 1.038 | 0.0248 | 23.8932 | -0.012837 | 0.14303 | 5.3e-05 | 0.004182 | 0.004182 | 1.4057
2019-03-21 16:07 | 1553198799.8823 | 1009 | 0.20946 | 0.000463 | -4e-06 | 100.3006 | 100.1966 | 123.7163 | 21.2612 | 0.001959 | 0.001839 | -152.4469 | 0.24136 | 8.4008 | 0.003427 | 0.004959 | 0.004734 | 4.7e-05 | 0.001084 | 0.25912 | 1.3558 | 2.548 | 98.84 | 20.85 | 38.718 | 68.956 | 104.3154 | 66.0854 | 0.009664 | 1.0407 | 0.02482 | 23.8493 | 0.009272 | 0.14358 | 7e-05 | 0.005471 | 0.005471 | 1.0786
2019-03-21 16:07 | 1553198800.8817 | 1010 | 0.20945 | 0.000463 | -4e-06 | 100.3205 | 100.4024 | 124.0335 | 21.2612 | 0.001949 | 0.001855 | -152.4293 | 0.24278 | 8.4012 | 0.003473 | 0.0047 | 0.004501 | 4.2e-05 | 0.001087 | 0.27383 | 1.3091 | 2.548 | 98.84 | 20.85 | 38.737 | 68.7207 | 105.4259 | 66.4637 | 0.009668 | 1.0395 | 0.024838 | 23.8931 | -0.007294 | 0.14443 | 0.000106 | 0.008139 | 0.008139 | 0.72933
2019-03-21 16:07 | 1553198801.8824 | 1011 | 0.20944 | 0.000464 | -4e-06 | 100.3286 | 100.3526 | 124.1144 | 21.2612 | 0.001955 | 0.001924 | -152.4378 | 0.24118 | 8.4038 | 0.003526 | 0.004273 | 0.00409 | 4.6e-05 | 0.001085 | 0.27793 | 1.3233 | 2.548 | 98.84 | 20.85 | 38.746 | 68.8758 | 106.7103 | 68.1137 | 0.009671 | 1.0345 | 0.024786 | 23.9589 | -0.002145 | 0.14348 | 0.00017 | 0.012867 | 0.012867 | 0.45831
2019-03-21 16:07 | 1553198802.8824 | 1012 | 0.20946 | 0.000465 | -4e-06 | 100.2511 | 100.1602 | 124.4165 | 21.2612 | 0.002015 | 0.001881 | -152.4679 | 0.2414 | 8.4025 | 0.003471 | 0.005308 | 0.005189 | 4.7e-05 | 0.001087 | 0.2688 | 1.3167 | 2.548 | 98.84 | 20.85 | 38.753 | 70.3051 | 105.3976 | 67.0786 | 0.009672 | 1.0377 | 0.024854 | 23.9517 | 0.008114 | 0.14361 | 8e-05 | 0.006177 | 0.006177 | 0.95559
2019-03-21 16:07 | 1553198803.8815 | 1013 | 0.20944 | 0.000461 | -4e-06 | 100.2919 | 100.2573 | 124.4071 | 21.2612 | 0.001971 | 0.00187 | -152.4723 | 0.24143 | 8.4006 | 0.003458 | 0.004834 | 0.00464 | 5e-05 | 0.001089 | 0.26521 | 1.3502 | 2.5482 | 98.84 | 20.85 | 38.76 | 69.2399 | 105.0751 | 66.8127 | 0.009674 | 1.0385 | 0.024863 | 23.9414 | 0.003091 | 0.14363 | 0.000187 | 0.014219 | 0.014219 | 0.41517
2019-03-21 16:07 | 1553198804.8829 | 1014 | 0.20947 | 0.000463 | -4e-06 | 100.3175 | 100.3935 | 124.1856 | 21.2612 | 0.001993 | 0.001872 | -152.4651 | 0.24112 | 8.4017 | 0.00356 | 0.005195 | 0.005082 | 5e-05 | 0.001094 | 0.27903 | 1.3101 | 2.548 | 98.84 | 20.85 | 38.765 | 69.7781 | 107.5505 | 66.8718 | 0.009675 | 1.0383 | 0.024838 | 23.9222 | -0.00677 | 0.14344 | 0 | 0.000274 | 0.000274 | 21.5405
2019-03-21 16:07 | 1553198805.8819 | 1015 | 0.20944 | 0.000463 | -4e-06 | 100.3581 | 100.3169 | 123.913 | 21.2612 | 0.001953 | 0.001866 | -152.4577 | 0.24191 | 8.3996 | 0.003564 | 0.004341 | 0.004125 | 4.8e-05 | 0.001087 | 0.24801 | 1.2941 | 2.5481 | 98.84 | 20.86 | 38.771 | 68.8219 | 107.6442 | 66.7314 | 0.009683 | 1.0387 | 0.024816 | 23.891 | 0.003672 | 0.14391 | 0.000168 | 0.012743 | 0.012743 | 0.46416
2019-03-21 16:07 | 1553198806.8828 | 1016 | 0.20943 | 0.000467 | -4e-06 | 100.3211 | 100.3043 | 123.7525 | 21.2612 | 0.001979 | 0.001836 | -152.4544 | 0.24089 | 8.403 | 0.003549 | 0.004669 | 0.004456 | 5.1e-05 | 0.001091 | 0.28711 | 1.2952 | 2.548 | 98.84 | 20.86 | 38.785 | 69.4445 | 107.2804 | 65.9957 | 0.009686 | 1.041 | 0.024827 | 23.8496 | 0.001503 | 0.1433 | 0.000226 | 0.017044 | 0.017044 | 0.34557
2019-03-21 16:07 | 1553198807.8827 | 1017 | 0.20946 | 0.000464 | -4e-06 | 100.2991 | 100.4024 | 123.5673 | 21.2612 | 0.001938 | 0.00188 | -152.442 | 0.2413 | 8.4011 | 0.003545 | 0.004821 | 0.004666 | 5e-05 | 0.001096 | 0.25249 | 1.3147 | 2.5478 | 98.84 | 20.86 | 38.791 | 68.4587 | 107.1772 | 67.0553 | 0.009688 | 1.0377 | 0.02477 | 23.869 | -0.009206 | 0.14355 | 7.6e-05 | 0.005874 | 0.005874 | 1.0044
2019-03-21 16:07 | 1553198808.8817 | 1018 | 0.20944 | 0.000467 | -4e-06 | 100.2434 | 100.337 | 123.4295 | 21.2612 | 0.001976 | 0.001863 | -152.4635 | 0.24194 | 8.4047 | 0.003591 | 0.004686 | 0.004607 | 5e-05 | 0.001093 | 0.26344 | 1.3308 | 2.5479 | 98.84 | 20.86 | 38.798 | 69.3594 | 108.2967 | 66.6573 | 0.00969 | 1.039 | 0.024771 | 23.8417 | -0.008349 | 0.14393 | 0.000177 | 0.013394 | 0.013394 | 0.44166
2019-03-21 16:07 | 1553198809.8818 | 1019 | 0.20946 | 0.000462 | -4e-06 | 100.3441 | 100.4623 | 123.6832 | 21.2612 | 0.001934 | 0.001898 | -152.4581 | 0.2409 | 8.4019 | 0.003559 | 0.004827 | 0.004555 | 5e-05 | 0.001092 | 0.24746 | 1.2959 | 2.5479 | 98.84 | 20.86 | 38.799 | 68.3534 | 107.5323 | 67.5037 | 0.00969 | 1.0364 | 0.024765 | 23.8959 | -0.010534 | 0.14331 | 4.4e-05 | 0.003534 | 0.003534 | 1.6667
2019-03-21 16:07 | 1553198810.8822 | 1020 | 0.20949 | 0.000467 | -4e-06 | 100.3166 | 100.399 | 123.5936 | 21.2612 | 0.001991 | 0.001917 | -152.4399 | 0.24145 | 8.4037 | 0.003568 | 0.00451 | 0.004371 | 5e-05 | 0.001092 | 0.25371 | 1.3264 | 2.5478 | 98.84 | 20.86 | 38.801 | 69.7399 | 107.7443 | 67.9469 | 0.00969 | 1.035 | 0.02474 | 23.9028 | -0.007335 | 0.14364 | -0.000125 | -0.009066 | -0.009066 | -0.65116
2019-03-21 16:07 | 1553198811.8809 | 1021 | 0.20949 | 0.000466 | -4e-06 | 100.3637 | 100.3247 | 123.4274 | 21.2612 | 0.00198 | 0.001916 | -152.4392 | 0.24174 | 8.4007 | 0.00361 | 0.004986 | 0.004868 | 5.3e-05 | 0.001093 | 0.28263 | 1.3822 | 2.5479 | 98.84 | 20.86 | 38.802 | 69.4661 | 108.7615 | 67.9243 | 0.009691 | 1.0351 | 0.024724 | 23.8859 | 0.003478 | 0.14381 | -0.000101 | -0.007221 | -0.007221 | -0.81846
2019-03-21 16:07 | 1553198812.8812 | 1022 | 0.20945 | 0.000468 | -4e-06 | 100.3229 | 100.2934 | 122.942 | 21.2612 | 0.00194 | 0.001891 | -152.4334 | 0.2415 | 8.4041 | 0.003549 | 0.004761 | 0.004532 | 5.2e-05 | 0.001094 | 0.26465 | 1.386 | 2.548 | 98.84 | 20.86 | 38.803 | 68.4991 | 107.2757 | 67.3181 | 0.009691 | 1.0369 | 0.024698 | 23.8177 | 0.002634 | 0.14367 | 9.9e-05 | 0.007595 | 0.007595 | 0.77751
2019-03-21 16:07 | 1553198813.881 | 1023 | 0.20948 | 0.000464 | -4e-06 | 100.3268 | 100.4075 | 122.5108 | 21.2612 | 0.001989 | 0.001841 | -152.4307 | 0.24168 | 8.3989 | 0.003504 | 0.004543 | 0.004487 | 4.9e-05 | 0.001092 | 0.28335 | 1.3561 | 2.548 | 98.84 | 20.86 | 38.805 | 69.6815 | 106.1923 | 66.1292 | 0.009691 | 1.0406 | 0.024697 | 23.7343 | -0.007191 | 0.14377 | -5.3e-05 | -0.003662 | -0.003662 | -1.6135
2019-03-21 16:07 | 1553198814.8811 | 1024 | 0.2095 | 0.000463 | -4e-06 | 100.3381 | 100.4139 | 122.3805 | 21.2612 | 0.001982 | 0.001894 | -152.4369 | 0.24225 | 8.4038 | 0.003601 | 0.00499 | 0.004766 | 5.1e-05 | 0.001098 | 0.28119 | 1.3309 | 2.5481 | 98.84 | 20.86 | 38.808 | 69.5115 | 108.5424 | 67.4071 | 0.009692 | 1.0367 | 0.024638 | 23.7663 | -0.006758 | 0.14411 | -0.000179 | -0.013016 | -0.013016 | -0.45507
2019-03-21 16:07 | 1553198815.881 | 1025 | 0.2095 | 0.000464 | -4e-06 | 100.3488 | 100.2675 | 122.3474 | 21.2612 | 0.001936 | 0.001851 | -152.4447 | 0.24055 | 8.4012 | 0.00362 | 0.004604 | 0.004429 | 5e-05 | 0.001095 | 0.28351 | 1.3293 | 2.5481 | 98.84 | 20.86 | 38.813 | 68.4212 | 108.9933 | 66.3577 | 0.009693 | 1.0399 | 0.024673 | 23.7265 | 0.007253 | 0.1431 | -0.000178 | -0.012932 | -0.012932 | -0.45481
2019-03-21 16:07 | 1553198816.8815 | 1026 | 0.20948 | 0.000468 | -5e-06 | 100.3676 | 100.3819 | 122.4268 | 21.2612 | 0.001996 | 0.001848 | -152.4482 | 0.2412 | 8.4005 | 0.003591 | 0.004872 | 0.004661 | 5.2e-05 | 0.001094 | 0.28152 | 1.3322 | 2.548 | 98.84 | 20.86 | 38.824 | 69.8552 | 108.3044 | 66.2937 | 0.009696 | 1.0401 | 0.024683 | 23.7319 | -0.001277 | 0.14349 | -5.6e-05 | -0.003841 | -0.003841 | -1.5355
2019-03-21 16:07 | 1553198817.8812 | 1027 | 0.20948 | 0.000468 | -4e-06 | 100.4001 | 100.5502 | 122.4796 | 21.2612 | 0.001949 | 0.001831 | -152.4448 | 0.24136 | 8.4025 | 0.003615 | 0.004665 | 0.004591 | 5e-05 | 0.001094 | 0.27489 | 1.3129 | 2.5481 | 98.84 | 20.86 | 38.831 | 68.7332 | 108.875 | 65.8761 | 0.009698 | 1.0414 | 0.024703 | 23.7225 | -0.013362 | 0.14359 | -4.8e-05 | -0.003282 | -0.003282 | -1.7982
2019-03-21 16:07 | 1553198818.8817 | 1028 | 0.20948 | 0.000466 | -4e-06 | 100.3312 | 100.2688 | 122.6054 | 21.2612 | 0.002015 | 0.001867 | -152.4206 | 0.2416 | 8.4025 | 0.003562 | 0.005118 | 0.004895 | 4.7e-05 | 0.001091 | 0.2589 | 1.3445 | 2.548 | 98.84 | 20.86 | 38.844 | 70.3044 | 107.5884 | 66.7446 | 0.009701 | 1.0387 | 0.024685 | 23.765 | 0.005565 | 0.14373 | -4.2e-05 | -0.002862 | -0.002862 | -2.0637
2019-03-21 16:07 | 1553198819.8813 | 1029 | 0.20945 | 0.000468 | -4e-06 | 100.3893 | 100.4623 | 122.4094 | 21.2612 | 0.001963 | 0.001824 | -152.4289 | 0.24064 | 8.4016 | 0.003564 | 0.005043 | 0.004874 | 4.8e-05 | 0.001095 | 0.28672 | 1.3123 | 2.5481 | 98.84 | 20.86 | 38.851 | 69.0637 | 107.6365 | 65.7224 | 0.009703 | 1.0418 | 0.024702 | 23.7103 | -0.006498 | 0.14316 | 0.000111 | 0.008497 | 0.008497 | 0.69246
2019-03-21 16:07 | 1553198820.8827 | 1030 | 0.20947 | 0.000469 | -4e-06 | 100.2937 | 100.4596 | 122.2665 | 21.2612 | 0.00193 | 0.001895 | -152.4427 | 0.24106 | 8.4032 | 0.00361 | 0.004414 | 0.004206 | 4.8e-05 | 0.001093 | 0.27007 | 1.3083 | 2.548 | 98.84 | 20.86 | 38.859 | 68.2769 | 108.7693 | 67.4196 | 0.009705 | 1.0366 | 0.024626 | 23.7557 | -0.014781 | 0.14341 | 2.3e-05 | 0.001961 | 0.001961 | 3.0058
2019-03-21 16:07 | 1553198821.8815 | 1031 | 0.20947 | 0.000463 | -4e-06 | 100.3747 | 100.3812 | 122.4876 | 21.2612 | 0.001972 | 0.001897 | -152.4526 | 0.2414 | 8.4034 | 0.003598 | 0.004949 | 0.004796 | 5.1e-05 | 0.001096 | 0.23612 | 1.3153 | 2.5479 | 98.84 | 20.86 | 38.867 | 69.2666 | 108.4786 | 67.4665 | 0.009707 | 1.0365 | 0.024646 | 23.7788 | -0.000579 | 0.14361 | -3e-06 | -1.6e-05 | -1.6e-05 | -368.1772
2019-03-21 16:07 | 1553198822.8819 | 1032 | 0.20949 | 0.000464 | -4e-06 | 100.325 | 100.3176 | 122.4007 | 21.2612 | 0.001996 | 0.001835 | -152.4313 | 0.24141 | 8.4007 | 0.003571 | 0.005124 | 0.004954 | 5.1e-05 | 0.001093 | 0.28982 | 1.3431 | 2.548 | 98.84 | 20.86 | 38.869 | 69.8569 | 107.8183 | 65.97 | 0.009707 | 1.0411 | 0.024692 | 23.7181 | 0.000663 | 0.14361 | -0.000135 | -0.009803 | -0.009803 | -0.60211
2019-03-21 16:07 | 1553198823.8819 | 1033 | 0.20949 | 0.000462 | -4e-06 | 100.3533 | 100.3029 | 122.347 | 21.2612 | 0.001981 | 0.00188 | -152.4389 | 0.24124 | 8.4032 | 0.003532 | 0.005006 | 0.004839 | 4.7e-05 | 0.001092 | 0.27865 | 1.3465 | 2.5478 | 98.84 | 20.86 | 38.867 | 69.4981 | 106.8695 | 67.0661 | 0.009707 | 1.0377 | 0.024647 | 23.7512 | 0.004493 | 0.14351 | -0.000149 | -0.010783 | -0.010783 | -0.54702
2019-03-21 16:07 | 1553198824.8816 | 1034 | 0.2095 | 0.000468 | -3e-06 | 100.3241 | 100.3622 | 122.3566 | 21.2612 | 0.001994 | 0.001862 | -152.4432 | 0.24167 | 8.4029 | 0.00356 | 0.005588 | 0.005358 | 5.2e-05 | 0.00109 | 0.27616 | 1.3515 | 2.5479 | 98.84 | 20.86 | 38.865 | 69.7968 | 107.5403 | 66.633 | 0.009706 | 1.039 | 0.024663 | 23.737 | -0.003394 | 0.14377 | -0.000159 | -0.011592 | -0.011592 | -0.50972
2019-03-21 16:07 | 1553198825.8826 | 1035 | 0.20947 | 0.000464 | -4e-06 | 100.2714 | 100.3955 | 122.4502 | 21.2612 | 0.001981 | 0.001889 | -152.4286 | 0.24177 | 8.4023 | 0.003505 | 0.004821 | 0.004675 | 5.4e-05 | 0.00109 | 0.28241 | 1.3313 | 2.5481 | 98.84 | 20.86 | 38.862 | 69.5005 | 106.2192 | 67.2847 | 0.009706 | 1.037 | 0.024649 | 23.7688 | -0.011071 | 0.14383 | -1.1e-05 | -0.000584 | -0.000584 | -10.1145
2019-03-21 16:07 | 1553198826.8816 | 1036 | 0.20947 | 0.000465 | -4e-06 | 100.3268 | 100.4054 | 122.822 | 21.2612 | 0.001978 | 0.001869 | -152.4242 | 0.24103 | 8.4036 | 0.003611 | 0.004566 | 0.004422 | 5e-05 | 0.001094 | 0.31741 | 1.3373 | 2.548 | 98.84 | 20.87 | 38.852 | 69.4125 | 108.7732 | 66.8016 | 0.009709 | 1.0385 | 0.024704 | 23.788 | -0.007009 | 0.14339 | -1.7e-05 | -0.001058 | -0.001058 | -5.5702
2019-03-21 16:07 | 1553198827.8809 | 1037 | 0.20949 | 0.000465 | -4e-06 | 100.2734 | 100.3945 | 123.0443 | 21.2612 | 0.001931 | 0.00185 | -152.4402 | 0.24041 | 8.4009 | 0.003577 | 0.00536 | 0.005163 | 5.1e-05 | 0.001087 | 0.25138 | 1.3624 | 2.5481 | 98.84 | 20.86 | 38.844 | 68.2902 | 107.9492 | 66.3367 | 0.009701 | 1.0399 | 0.024744 | 23.7932 | -0.010794 | 0.14302 | -0.000132 | -0.00959 | -0.00959 | -0.61294
2019-03-21 16:07 | 1553198828.8812 | 1038 | 0.2095 | 0.00047 | -3e-06 | 100.305 | 100.2981 | 123.0814 | 21.2612 | 0.001962 | 0.001878 | -152.4583 | 0.2422 | 8.4039 | 0.003618 | 0.004616 | 0.004449 | 5.3e-05 | 0.001086 | 0.25155 | 1.3362 | 2.548 | 98.84 | 20.87 | 38.834 | 69.0279 | 108.9442 | 67.0171 | 0.009705 | 1.0379 | 0.024722 | 23.8207 | 0.000615 | 0.14408 | -0.000184 | -0.013468 | -0.013468 | -0.43971
2019-03-21 16:07 | 1553198829.8818 | 1039 | 0.2095 | 0.000466 | -4e-06 | 100.288 | 100.2276 | 122.8648 | 21.2612 | 0.001901 | 0.001904 | -152.4553 | 0.24146 | 8.4025 | 0.003537 | 0.005228 | 0.004983 | 5.3e-05 | 0.001086 | 0.27793 | 1.3262 | 2.5479 | 98.84 | 20.87 | 38.806 | 67.5583 | 106.9843 | 67.6309 | 0.009698 | 1.036 | 0.024678 | 23.8211 | 0.00539 | 0.14365 | -0.000198 | -0.014415 | -0.014415 | -0.40957
2019-03-21 16:07 | 1553198830.8809 | 1040 | 0.20948 | 0.000467 | -4e-06 | 100.291 | 100.3462 | 122.8415 | 21.2612 | 0.001955 | 0.001879 | -152.4724 | 0.24137 | 8.4015 | 0.003519 | 0.004982 | 0.004872 | 5.5e-05 | 0.001085 | 0.26897 | 1.3474 | 2.548 | 98.84 | 20.87 | 38.791 | 68.858 | 106.5545 | 67.0435 | 0.009694 | 1.0378 | 0.024697 | 23.7984 | -0.004916 | 0.14359 | -5.6e-05 | -0.003927 | -0.003927 | -1.503
2019-03-21 16:07 | 1553198831.881 | 1041 | 0.20951 | 0.000466 | -4e-06 | 100.3071 | 100.4497 | 122.6737 | 21.2612 | 0.002013 | 0.001893 | -152.4554 | 0.24179 | 8.4013 | 0.003565 | 0.00508 | 0.004865 | 5.5e-05 | 0.001088 | 0.25409 | 1.2903 | 2.5479 | 98.84 | 20.87 | 38.761 | 70.2662 | 107.6734 | 67.3696 | 0.009686 | 1.0368 | 0.024669 | 23.7935 | -0.012705 | 0.14384 | -0.000246 | -0.018009 | -0.018009 | -0.32827
2019-03-21 16:07 | 1553198832.8811 | 1042 | 0.20949 | 0.000464 | -4e-06 | 100.3354 | 100.2794 | 123.0279 | 21.2612 | 0.002001 | 0.001903 | -152.4593 | 0.24085 | 8.4 | 0.00353 | 0.004269 | 0.00402 | 5.6e-05 | 0.001093 | 0.27367 | 1.3055 | 2.548 | 98.84 | 20.87 | 38.747 | 69.9615 | 106.8134 | 67.6045 | 0.009683 | 1.0361 | 0.024696 | 23.836 | 0.004995 | 0.14328 | -9.3e-05 | -0.006664 | -0.006664 | -0.8837
2019-03-21 16:07 | 1553198833.882 | 1043 | 0.2095 | 0.000465 | -3e-06 | 100.3497 | 100.4034 | 123.0305 | 21.2612 | 0.001966 | 0.001838 | -152.4268 | 0.24172 | 8.4005 | 0.003563 | 0.005406 | 0.005271 | 5.5e-05 | 0.001095 | 0.31094 | 1.3145 | 2.548 | 98.84 | 20.87 | 38.733 | 69.1387 | 107.6136 | 66.0424 | 0.009679 | 1.0408 | 0.024753 | 23.7816 | -0.004782 | 0.1438 | -0.000196 | -0.014352 | -0.014352 | -0.41178
2019-03-21 16:07 | 1553198834.8811 | 1044 | 0.20946 | 0.000463 | -3e-06 | 100.2627 | 100.2947 | 123.0831 | 21.2612 | 0.001972 | 0.00189 | -152.4159 | 0.24113 | 8.4015 | 0.003564 | 0.004743 | 0.004532 | 5.8e-05 | 0.001095 | 0.28512 | 1.3256 | 2.5481 | 98.84 | 20.87 | 38.706 | 69.276 | 107.6442 | 67.2979 | 0.009673 | 1.037 | 0.024712 | 23.8306 | -0.002855 | 0.14345 | 5.3e-05 | 0.004177 | 0.004177 | 1.4116
2019-03-21 16:07 | 1553198835.8821 | 1045 | 0.20947 | 0.00047 | -4e-06 | 100.2681 | 100.4943 | 123.0535 | 21.2612 | 0.001973 | 0.001889 | -152.4275 | 0.24114 | 8.403 | 0.003532 | 0.005318 | 0.005138 | 5.6e-05 | 0.001095 | 0.27671 | 1.3059 | 2.5478 | 98.84 | 20.87 | 38.694 | 69.2938 | 106.8615 | 67.283 | 0.00967 | 1.037 | 0.02471 | 23.8273 | -0.020159 | 0.14346 | -3.7e-05 | -0.002494 | -0.002494 | -2.3641
2019-03-21 16:07 | 1553198836.8826 | 1046 | 0.20947 | 0.000467 | -4e-06 | 100.2216 | 100.1592 | 122.7124 | 21.2612 | 0.002008 | 0.00187 | -152.4493 | 0.24173 | 8.4025 | 0.003521 | 0.004775 | 0.004597 | 5.4e-05 | 0.001096 | 0.28888 | 1.3207 | 2.5479 | 98.84 | 20.87 | 38.671 | 70.1391 | 106.6009 | 66.8314 | 0.009664 | 1.0384 | 0.024692 | 23.7784 | 0.005573 | 0.1438 | -3.5e-05 | -0.002361 | -0.002361 | -2.5029
2019-03-21 16:07 | 1553198837.8816 | 1047 | 0.20951 | 0.000466 | -4e-06 | 100.213 | 100.3332 | 122.531 | 21.2612 | 0.001976 | 0.001834 | -152.4475 | 0.24072 | 8.4033 | 0.003589 | 0.004883 | 0.004696 | 5.6e-05 | 0.001092 | 0.26217 | 1.3471 | 2.548 | 98.84 | 20.87 | 38.662 | 69.3632 | 108.2423 | 65.9473 | 0.009662 | 1.0411 | 0.024706 | 23.7299 | -0.010727 | 0.1432 | -0.000243 | -0.017803 | -0.017803 | -0.33059
2019-03-21 16:07 | 1553198838.8813 | 1048 | 0.2095 | 0.000467 | -3e-06 | 100.1844 | 100.1714 | 122.4632 | 21.2612 | 0.001953 | 0.001899 | -152.4443 | 0.24261 | 8.4026 | 0.00362 | 0.005161 | 0.004968 | 5.5e-05 | 0.001093 | 0.26792 | 1.3323 | 2.548 | 98.84 | 20.87 | 38.652 | 68.8087 | 109.0059 | 67.5288 | 0.009659 | 1.0363 | 0.024642 | 23.7786 | 0.001154 | 0.14433 | -0.00017 | -0.01236 | -0.01236 | -0.47993
2019-03-21 16:07 | 1553198839.8814 | 1049 | 0.20948 | 0.000465 | -4e-06 | 100.1814 | 100.1776 | 122.7542 | 21.2612 | 0.001928 | 0.001887 | -152.4674 | 0.24112 | 8.4004 | 0.003586 | 0.00437 | 0.00425 | 5.6e-05 | 0.001093 | 0.27616 | 1.3166 | 2.548 | 98.84 | 20.87 | 38.636 | 68.2089 | 108.1728 | 67.2198 | 0.009655 | 1.0372 | 0.024682 | 23.7961 | 0.00034 | 0.14344 | -7.4e-05 | -0.005261 | -0.005261 | -1.1205
2019-03-21 16:07 | 1553198840.8825 | 1050 | 0.20951 | 0.000466 | -4e-06 | 100.1391 | 100.3281 | 122.8432 | 21.2612 | 0.001941 | 0.001866 | -152.445 | 0.2414 | 8.4014 | 0.003594 | 0.00512 | 0.004957 | 5.5e-05 | 0.001092 | 0.2907 | 1.3066 | 2.5481 | 98.84 | 20.87 | 38.63 | 68.5265 | 108.3778 | 66.7174 | 0.009654 | 1.0388 | 0.024709 | 23.7871 | -0.016869 | 0.14361 | -0.000231 | -0.016891 | -0.016891 | -0.34944
2019-03-21 16:07 | 1553198841.8814 | 1051 | 0.20952 | 0.000462 | -4e-06 | 100.0965 | 99.9637 | 122.8298 | 21.2612 | 0.001967 | 0.001822 | -152.454 | 0.24142 | 8.4035 | 0.003523 | 0.004497 | 0.004338 | 5.3e-05 | 0.001093 | 0.26952 | 1.3071 | 2.548 | 98.84 | 20.87 | 38.625 | 69.1503 | 106.645 | 65.6772 | 0.009652 | 1.042 | 0.024746 | 23.7494 | 0.011874 | 0.14362 | -0.000269 | -0.019758 | -0.019758 | -0.29875
2019-03-21 16:07 | 1553198842.8826 | 1052 | 0.20948 | 0.000465 | -3e-06 | 100.1632 | 100.1176 | 122.6 | 21.2612 | 0.002017 | 0.001855 | -152.4555 | 0.24142 | 8.4015 | 0.003541 | 0.004554 | 0.004418 | 5.5e-05 | 0.00109 | 0.29247 | 1.3012 | 2.5479 | 98.84 | 20.87 | 38.612 | 70.3575 | 107.093 | 66.4553 | 0.009649 | 1.0396 | 0.024694 | 23.7544 | 0.004071 | 0.14362 | -4.5e-05 | -0.00313 | -0.00313 | -1.8857
2019-03-21 16:07 | 1553198843.8806 | 1053 | 0.20961 | 0.000469 | -3e-06 | 100.0944 | 100.0556 | 122.5843 | 21.2612 | 0.001964 | 0.001881 | -152.4701 | 0.24085 | 8.3978 | 0.003556 | 0.004719 | 0.004459 | 5.3e-05 | 0.001086 | 0.26106 | 1.2946 | 2.548 | 98.84 | 20.87 | 38.607 | 69.081 | 107.4418 | 67.0842 | 0.009648 | 1.0377 | 0.02467 | 23.7748 | 0.003463 | 0.14328 | -0.000879 | -0.06491 | -0.06491 | -0.090724
2019-03-21 16:07 | 1553198844.8806 | 1054 | 0.2095 | 0.000468 | -4e-06 | 100.067 | 100.152 | 122.6636 | 21.2612 | 0.001988 | 0.001832 | -152.477 | 0.24181 | 8.4023 | 0.003654 | 0.004781 | 0.004644 | 5.5e-05 | 0.00109 | 0.30198 | 1.3038 | 2.548 | 98.84 | 20.88 | 38.603 | 69.662 | 109.8241 | 65.9122 | 0.009653 | 1.0412 | 0.024721 | 23.7415 | -0.007599 | 0.14385 | -0.000191 | -0.013966 | -0.013966 | -0.42333
2019-03-21 16:07 | 1553198845.8811 | 1055 | 0.20945 | 0.000465 | -4e-06 | 100.0786 | 100.368 | 122.5188 | 21.2612 | 0.002031 | 0.001804 | -152.4789 | 0.2425 | 8.4051 | 0.003658 | 0.00477 | 0.004591 | 5.4e-05 | 0.001093 | 0.26465 | 1.3133 | 2.548 | 98.84 | 20.87 | 38.602 | 70.6834 | 109.9337 | 65.2387 | 0.009647 | 1.0433 | 0.024731 | 23.7039 | -0.025825 | 0.14427 | 0.000104 | 0.007979 | 0.007979 | 0.74308
2019-03-21 16:07 | 1553198846.881 | 1056 | 0.20951 | 0.000467 | -4e-06 | 100.0914 | 100.0441 | 122.7059 | 21.2612 | 0.001969 | 0.001851 | -152.4627 | 0.24181 | 8.4045 | 0.003678 | 0.004623 | 0.004477 | 5.2e-05 | 0.00109 | 0.24818 | 1.3116 | 2.548 | 98.84 | 20.87 | 38.604 | 69.2111 | 110.4127 | 66.3719 | 0.009647 | 1.0398 | 0.024708 | 23.7617 | 0.004232 | 0.14385 | -0.000231 | -0.016905 | -0.016905 | -0.34974
2019-03-21 16:07 | 1553198847.8819 | 1057 | 0.20951 | 0.000463 | -4e-06 | 100.0568 | 100.1551 | 122.5932 | 21.2612 | 0.001932 | 0.001877 | -152.4638 | 0.24136 | 8.403 | 0.003718 | 0.004785 | 0.004633 | 5.3e-05 | 0.001095 | 0.25299 | 1.3184 | 2.5479 | 98.84 | 20.87 | 38.608 | 68.3243 | 111.3777 | 66.9827 | 0.009648 | 1.038 | 0.024675 | 23.7722 | -0.008779 | 0.14358 | -0.000233 | -0.016981 | -0.016981 | -0.34753
2019-03-21 16:07 | 1553198848.8818 | 1058 | 0.20947 | 0.000461 | -4e-06 | 100.1126 | 100.1275 | 122.5617 | 21.2612 | 0.001943 | 0.001876 | -152.4487 | 0.24164 | 8.4018 | 0.00367 | 0.005169 | 0.005003 | 5.4e-05 | 0.001092 | 0.27041 | 1.3196 | 2.5479 | 98.84 | 20.88 | 38.613 | 68.5741 | 110.2296 | 66.9678 | 0.009655 | 1.038 | 0.024672 | 23.7686 | -0.001335 | 0.14375 | -1.6e-05 | -0.000951 | -0.000951 | -6.2132
2019-03-21 16:07 | 1553198849.8822 | 1059 | 0.2095 | 0.000465 | -4e-06 | 100.1021 | 100.2723 | 122.5542 | 21.2612 | 0.001987 | 0.001789 | -152.4421 | 0.24137 | 8.4042 | 0.003633 | 0.005629 | 0.005422 | 5.2e-05 | 0.001095 | 0.28622 | 1.3277 | 2.548 | 98.84 | 20.87 | 38.622 | 69.6269 | 109.3298 | 64.8694 | 0.009652 | 1.0445 | 0.024748 | 23.6944 | -0.015188 | 0.14359 | -0.000203 | -0.014868 | -0.014868 | -0.39693
2019-03-21 16:08 | 1553198850.8814 | 1060 | 0.20946 | 0.000468 | -4e-06 | 100.1468 | 100.0624 | 122.657 | 21.2612 | 0.002 | 0.001862 | -152.4521 | 0.24139 | 8.4035 | 0.003599 | 0.004963 | 0.004801 | 5.2e-05 | 0.001094 | 0.26626 | 1.3215 | 2.5479 | 98.84 | 20.88 | 38.625 | 69.9458 | 108.4895 | 66.6309 | 0.009658 | 1.039 | 0.024694 | 23.766 | 0.007539 | 0.1436 | 6.2e-05 | 0.004852 | 0.004852 | 1.2165
2019-03-21 16:08 | 1553198851.8821 | 1061 | 0.20947 | 0.000466 | -4e-06 | 100.1355 | 100.1272 | 122.9115 | 21.2612 | 0.001925 | 0.001877 | -152.4407 | 0.24051 | 8.4017 | 0.003627 | 0.004548 | 0.00436 | 5.4e-05 | 0.001092 | 0.27527 | 1.3672 | 2.5481 | 98.84 | 20.88 | 38.629 | 68.1474 | 109.1753 | 66.9873 | 0.009659 | 1.0379 | 0.024706 | 23.8032 | 0.000745 | 0.14308 | 1.7e-05 | 0.001503 | 0.001503 | 3.9116
2019-03-21 16:08 | 1553198852.8811 | 1062 | 0.20952 | 0.000467 | -5e-06 | 99.9627 | 99.8853 | 123.1492 | 21.2612 | 0.001955 | 0.001859 | -152.423 | 0.2416 | 8.4019 | 0.003558 | 0.005176 | 0.005007 | 5.1e-05 | 0.001095 | 0.26842 | 1.3331 | 2.5479 | 98.84 | 20.88 | 38.635 | 68.8642 | 107.4898 | 66.5621 | 0.009661 | 1.0392 | 0.024746 | 23.8113 | 0.006924 | 0.14373 | -0.000332 | -0.02437 | -0.02437 | -0.24239
2019-03-21 16:08 | 1553198853.8821 | 1063 | 0.20947 | 0.000466 | -4e-06 | 100.103 | 100.0529 | 123.3333 | 21.2612 | 0.00194 | 0.001842 | -152.4424 | 0.24058 | 8.4044 | 0.003637 | 0.005701 | 0.005539 | 5.2e-05 | 0.001092 | 0.25984 | 1.3416 | 2.548 | 98.84 | 20.88 | 38.639 | 68.5077 | 109.4078 | 66.1532 | 0.009662 | 1.0405 | 0.024779 | 23.8147 | 0.004478 | 0.14312 | -7e-06 | -0.000247 | -0.000247 | -23.7672
2019-03-21 16:08 | 1553198854.8812 | 1064 | 0.20952 | 0.000463 | -4e-06 | 100.019 | 100.01 | 123.3023 | 21.2612 | 0.001946 | 0.001882 | -152.4258 | 0.24153 | 8.4019 | 0.003617 | 0.004533 | 0.004463 | 4.9e-05 | 0.001094 | 0.27975 | 1.3055 | 2.5479 | 98.84 | 20.87 | 38.646 | 68.6582 | 108.9294 | 67.1185 | 0.009658 | 1.0375 | 0.024741 | 23.8456 | 0.000805 | 0.14369 | -0.000312 | -0.022916 | -0.022916 | -0.25771
2019-03-21 16:08 | 1553198855.8822 | 1065 | 0.20949 | 0.000465 | -4e-06 | 100.0393 | 100.2109 | 123.1887 | 21.2612 | 0.001958 | 0.001853 | -152.4226 | 0.24233 | 8.4029 | 0.003666 | 0.004907 | 0.004671 | 5.3e-05 | 0.001097 | 0.28871 | 1.3405 | 2.5479 | 98.84 | 20.87 | 38.647 | 68.9296 | 110.1333 | 66.4077 | 0.009658 | 1.0397 | 0.024755 | 23.8097 | -0.015337 | 0.14416 | -0.000149 | -0.010782 | -0.010782 | -0.54956
2019-03-21 16:08 | 1553198856.8823 | 1066 | 0.20948 | 0.000468 | -4e-06 | 100.0857 | 100.0713 | 123.0042 | 21.2612 | 0.001949 | 0.001834 | -152.4533 | 0.24211 | 8.4016 | 0.003656 | 0.004843 | 0.004745 | 5.2e-05 | 0.001096 | 0.28335 | 1.3407 | 2.548 | 98.84 | 20.87 | 38.648 | 68.7176 | 109.8771 | 65.9582 | 0.009658 | 1.0411 | 0.024753 | 23.7761 | 0.00129 | 0.14403 | -8.8e-05 | -0.006333 | -0.006333 | -0.9347
2019-03-21 16:08 | 1553198857.8813 | 1067 | 0.20948 | 0.000467 | -3e-06 | 100.0116 | 100.0073 | 123.1429 | 21.2612 | 0.001912 | 0.001864 | -152.4311 | 0.24138 | 8.4011 | 0.003539 | 0.004815 | 0.004642 | 5.4e-05 | 0.001098 | 0.2896 | 1.3556 | 2.5481 | 98.84 | 20.87 | 38.646 | 67.8221 | 107.0457 | 66.6698 | 0.009658 | 1.0389 | 0.024741 | 23.8144 | 0.000383 | 0.1436 | -6.6e-05 | -0.004672 | -0.004672 | -1.2632
2019-03-21 16:08 | 1553198858.8818 | 1068 | 0.20949 | 0.000463 | -4e-06 | 100.0238 | 99.9712 | 123.08 | 21.2612 | 0.001953 | 0.001843 | -152.4466 | 0.24136 | 8.4021 | 0.003557 | 0.005242 | 0.00505 | 5.3e-05 | 0.001096 | 0.30326 | 1.3615 | 2.548 | 98.84 | 20.88 | 38.642 | 68.8111 | 107.4835 | 66.1782 | 0.009662 | 1.0404 | 0.024753 | 23.7911 | 0.004705 | 0.14359 | -0.0001 | -0.007227 | -0.007227 | -0.81663
2019-03-21 16:08 | 1553198859.8815 | 1069 | 0.20948 | 0.000465 | -3e-06 | 100.0545 | 100.1234 | 123.1598 | 21.2612 | 0.001973 | 0.001794 | -152.44 | 0.24191 | 8.3997 | 0.003566 | 0.004715 | 0.004552 | 5.3e-05 | 0.001096 | 0.29391 | 1.3375 | 2.5479 | 98.84 | 20.87 | 38.635 | 69.2993 | 107.6821 | 65.0022 | 0.009655 | 1.044 | 0.024804 | 23.7576 | -0.006163 | 0.14391 | -6.5e-05 | -0.004641 | -0.004641 | -1.2744
2019-03-21 16:08 | 1553198860.8805 | 1070 | 0.20947 | 0.000467 | -4e-06 | 100.098 | 99.9027 | 123.376 | 21.2612 | 0.001875 | 0.001837 | -152.4369 | 0.24168 | 8.4044 | 0.003537 | 0.004648 | 0.004535 | 5.5e-05 | 0.001092 | 0.2959 | 1.3051 | 2.5478 | 98.84 | 20.87 | 38.63 | 66.942 | 106.9804 | 66.019 | 0.009654 | 1.0409 | 0.024788 | 23.8141 | 0.017464 | 0.14378 | 1.3e-05 | 0.001242 | 0.001242 | 4.7586
2019-03-21 16:08 | 1553198861.8805 | 1071 | 0.20949 | 0.000463 | -4e-06 | 100.0458 | 99.962 | 122.9373 | 21.2612 | 0.001854 | 0.00185 | -152.4542 | 0.24122 | 8.4009 | 0.003567 | 0.004441 | 0.004153 | 5.1e-05 | 0.001086 | 0.26576 | 1.3381 | 2.548 | 98.84 | 20.87 | 38.626 | 66.4421 | 107.7277 | 66.3305 | 0.009653 | 1.04 | 0.024733 | 23.7827 | 0.007499 | 0.1435 | -9.4e-05 | -0.006709 | -0.006709 | -0.87908
2019-03-21 16:08 | 1553198862.8821 | 1072 | 0.20949 | 0.000465 | -4e-06 | 100.0381 | 100.0035 | 122.6347 | 21.2612 | 0.001901 | 0.001887 | -152.4778 | 0.24051 | 8.402 | 0.003582 | 0.004659 | 0.004513 | 5.1e-05 | 0.001084 | 0.26593 | 1.3417 | 2.5479 | 98.84 | 20.87 | 38.619 | 67.5646 | 108.0752 | 67.2246 | 0.009651 | 1.0372 | 0.02467 | 23.7846 | 0.003089 | 0.14308 | -0.000107 | -0.007677 | -0.007677 | -0.76601
2019-03-21 16:08 | 1553198863.8807 | 1073 | 0.20948 | 0.000467 | -5e-06 | 100.0062 | 99.9051 | 122.8645 | 21.2612 | 0.001956 | 0.00183 | -152.4731 | 0.24087 | 8.4062 | 0.003538 | 0.004922 | 0.004813 | 5.5e-05 | 0.001085 | 0.27616 | 1.3764 | 2.548 | 98.84 | 20.87 | 38.614 | 68.8806 | 107.0088 | 65.8605 | 0.00965 | 1.0414 | 0.024743 | 23.7592 | 0.009046 | 0.14329 | -7.9e-05 | -0.00557 | -0.00557 | -1.0573
2019-03-21 16:08 | 1553198864.8808 | 1074 | 0.20947 | 0.000465 | -4e-06 | 100.005 | 99.9971 | 123.0213 | 21.2612 | 0.002 | 0.001909 | -152.4902 | 0.24163 | 8.402 | 0.003574 | 0.004715 | 0.00451 | 5e-05 | 0.00109 | 0.26488 | 1.3655 | 2.548 | 98.84 | 20.87 | 38.61 | 69.9396 | 107.8846 | 67.7519 | 0.009649 | 1.0356 | 0.02469 | 23.8405 | 0.00071 | 0.14375 | -1.9e-05 | -0.001163 | -0.001163 | -5.0779
2019-03-21 16:08 | 1553198865.8818 | 1075 | 0.2095 | 0.000467 | -4e-06 | 100.0717 | 99.9364 | 123.6126 | 21.2612 | 0.00194 | 0.001865 | -152.4908 | 0.24164 | 8.3999 | 0.003489 | 0.004988 | 0.004821 | 5.1e-05 | 0.001087 | 0.27273 | 1.3325 | 2.5479 | 98.84 | 20.87 | 38.606 | 68.5164 | 105.8336 | 66.691 | 0.009648 | 1.0389 | 0.024788 | 23.8606 | 0.012102 | 0.14375 | -0.000194 | -0.014185 | -0.014185 | -0.41651
2019-03-21 16:08 | 1553198866.8809 | 1076 | 0.20948 | 0.000465 | -4e-06 | 100.0428 | 100.0621 | 123.9321 | 21.2612 | 0.002002 | 0.001887 | -152.4465 | 0.24133 | 8.4026 | 0.003521 | 0.005333 | 0.005174 | 5e-05 | 0.00109 | 0.27201 | 1.3228 | 2.5479 | 98.84 | 20.88 | 38.598 | 69.9987 | 106.6064 | 67.2299 | 0.009651 | 1.0372 | 0.0248 | 23.9103 | -0.001723 | 0.14357 | -7.5e-05 | -0.005377 | -0.005377 | -1.0974
2019-03-21 16:08 | 1553198867.8819 | 1077 | 0.20949 | 0.00047 | -4e-06 | 99.9669 | 99.8922 | 123.8123 | 21.2612 | 0.001922 | 0.001853 | -152.4654 | 0.24146 | 8.4026 | 0.003587 | 0.005238 | 0.00501 | 5.3e-05 | 0.00109 | 0.28512 | 1.3101 | 2.5481 | 98.84 | 20.88 | 38.594 | 68.074 | 108.21 | 66.424 | 0.00965 | 1.0397 | 0.024817 | 23.8704 | 0.006687 | 0.14365 | -0.000151 | -0.011014 | -0.011014 | -0.53604
2019-03-21 16:08 | 1553198868.8823 | 1078 | 0.20949 | 0.000465 | -4e-06 | 100.0452 | 99.8516 | 123.8471 | 21.2612 | 0.001963 | 0.001875 | -152.4736 | 0.24088 | 8.401 | 0.003544 | 0.004419 | 0.004272 | 5.3e-05 | 0.00109 | 0.24043 | 1.3296 | 2.5478 | 98.84 | 20.88 | 38.591 | 69.0685 | 107.1466 | 66.9483 | 0.00965 | 1.0381 | 0.024802 | 23.8922 | 0.017325 | 0.1433 | -0.000136 | -0.009892 | -0.009892 | -0.59541
2019-03-21 16:08 | 1553198869.881 | 1079 | 0.20948 | 0.000467 | -4e-06 | 99.974 | 100.1571 | 123.6621 | 21.2612 | 0.00195 | 0.001883 | -152.4737 | 0.24148 | 8.4017 | 0.003508 | 0.005126 | 0.004978 | 5.2e-05 | 0.001091 | 0.27936 | 1.3567 | 2.5479 | 98.84 | 20.88 | 38.586 | 68.7479 | 106.2964 | 67.1294 | 0.009648 | 1.0375 | 0.024777 | 23.8807 | -0.016369 | 0.14366 | -7.3e-05 | -0.005169 | -0.005169 | -1.1422
2019-03-21 16:08 | 1553198870.8811 | 1080 | 0.20948 | 0.000465 | -4e-06 | 99.9669 | 100.009 | 123.4159 | 21.2612 | 0.001954 | 0.001864 | -152.4784 | 0.24151 | 8.4052 | 0.003508 | 0.004396 | 0.004252 | 5.1e-05 | 0.001091 | 0.27759 | 1.3357 | 2.5479 | 98.84 | 20.88 | 38.584 | 68.8407 | 106.2789 | 66.6792 | 0.009648 | 1.0389 | 0.024768 | 23.8411 | -0.003767 | 0.14367 | -6.1e-05 | -0.004258 | -0.004258 | -1.3867
2019-03-21 16:08 | 1553198871.8814 | 1081 | 0.20947 | 0.000468 | -4e-06 | 100.092 | 100.0877 | 123.6288 | 21.2612 | 0.001947 | 0.0019 | -152.474 | 0.24086 | 8.4007 | 0.003528 | 0.005133 | 0.004972 | 5.3e-05 | 0.001092 | 0.28766 | 1.3211 | 2.548 | 98.84 | 20.88 | 38.577 | 68.6621 | 106.7646 | 67.5413 | 0.009646 | 1.0363 | 0.024758 | 23.8919 | 0.000388 | 0.14329 | -2.3e-05 | -0.001468 | -0.001468 | -4.0116
2019-03-21 16:08 | 1553198872.8825 | 1082 | 0.20948 | 0.000468 | -4e-06 | 100.075 | 99.9596 | 123.7055 | 21.2612 | 0.00194 | 0.001871 | -152.4657 | 0.24077 | 8.4009 | 0.003487 | 0.005128 | 0.004959 | 5.3e-05 | 0.001093 | 0.27273 | 1.3364 | 2.5479 | 98.84 | 20.88 | 38.573 | 68.4984 | 105.7754 | 66.8437 | 0.009645 | 1.0384 | 0.024791 | 23.8749 | 0.010322 | 0.14323 | -8.3e-05 | -0.005946 | -0.005946 | -0.9901
2019-03-21 16:08 | 1553198873.8802 | 1083 | 0.20945 | 0.000468 | -4e-06 | 100.0488 | 100.118 | 123.7783 | 21.2612 | 0.001955 | 0.00186 | -152.4595 | 0.24189 | 8.4012 | 0.003462 | 0.005065 | 0.004838 | 5.2e-05 | 0.001089 | 0.27201 | 1.2906 | 2.5481 | 98.84 | 20.89 | 38.569 | 68.8642 | 105.1592 | 66.5746 | 0.00965 | 1.0392 | 0.024809 | 23.8725 | -0.006182 | 0.1439 | 0.000123 | 0.009383 | 0.009383 | 0.6303
2019-03-21 16:08 | 1553198874.8812 | 1084 | 0.20946 | 0.000468 | -4e-06 | 100.1522 | 100.1857 | 123.8802 | 21.2612 | 0.002007 | 0.001884 | -152.4477 | 0.24144 | 8.4027 | 0.003469 | 0.004661 | 0.0045 | 5.4e-05 | 0.001091 | 0.2959 | 1.3107 | 2.5481 | 98.84 | 20.89 | 38.559 | 70.1119 | 105.3505 | 67.1474 | 0.009648 | 1.0375 | 0.024798 | 23.9024 | -0.002997 | 0.14363 | 6.5e-05 | 0.005092 | 0.005092 | 1.1592
2019-03-21 16:08 | 1553198875.8807 | 1085 | 0.20949 | 0.000468 | -4e-06 | 100.1263 | 100.1333 | 123.507 | 21.2612 | 0.001916 | 0.001845 | -152.4454 | 0.24175 | 8.4017 | 0.003482 | 0.004547 | 0.004342 | 5.2e-05 | 0.001091 | 0.29822 | 1.3254 | 2.548 | 98.84 | 20.89 | 38.556 | 67.9281 | 105.6495 | 66.2242 | 0.009647 | 1.0403 | 0.024794 | 23.834 | -0.000628 | 0.14382 | -0.000152 | -0.011029 | -0.011029 | -0.53594
2019-03-21 16:08 | 1553198876.8813 | 1086 | 0.20949 | 0.000463 | -3e-06 | 100.1814 | 100.2191 | 123.3023 | 21.2612 | 0.001982 | 0.001861 | -152.4811 | 0.2418 | 8.4014 | 0.003442 | 0.004028 | 0.003846 | 5.4e-05 | 0.001086 | 0.27959 | 1.3248 | 2.5479 | 98.84 | 20.89 | 38.554 | 69.5248 | 104.6889 | 66.609 | 0.009646 | 1.0391 | 0.02476 | 23.8277 | -0.003369 | 0.14385 | -9.6e-05 | -0.006939 | -0.006939 | -0.85204
2019-03-21 16:08 | 1553198877.8808 | 1087 | 0.20947 | 0.000466 | -4e-06 | 100.1379 | 100.181 | 123.5344 | 21.2612 | 0.001928 | 0.001881 | -152.486 | 0.24122 | 8.4029 | 0.003402 | 0.004993 | 0.004824 | 5.5e-05 | 0.001087 | 0.27345 | 1.2893 | 2.548 | 98.84 | 20.89 | 38.549 | 68.2176 | 103.7032 | 67.0755 | 0.009645 | 1.0377 | 0.024766 | 23.8665 | -0.003848 | 0.1435 | -2.6e-05 | -0.001681 | -0.001681 | -3.5079
2019-03-21 16:08 | 1553198878.8808 | 1088 | 0.20946 | 0.000467 | -4e-06 | 100.1978 | 100.1708 | 123.7109 | 21.2612 | 0.001945 | 0.001856 | -152.4731 | 0.24153 | 8.4047 | 0.003434 | 0.004841 | 0.00469 | 5.2e-05 | 0.00108 | 0.27992 | 1.3373 | 2.548 | 98.84 | 20.89 | 38.548 | 68.6311 | 104.5025 | 66.495 | 0.009645 | 1.0395 | 0.024805 | 23.8632 | 0.002411 | 0.14369 | 8.1e-05 | 0.006273 | 0.006273 | 0.94143
2019-03-21 16:08 | 1553198879.8819 | 1089 | 0.20947 | 0.000471 | -4e-06 | 100.1802 | 100.2828 | 124.0729 | 21.2612 | 0.001944 | 0.001829 | -152.47 | 0.24159 | 8.403 | 0.003484 | 0.004988 | 0.004839 | 5.2e-05 | 0.001084 | 0.24978 | 1.365 | 2.5481 | 98.84 | 20.89 | 38.546 | 68.6113 | 105.6934 | 65.8247 | 0.009644 | 1.0415 | 0.024865 | 23.8744 | -0.009158 | 0.14372 | 2.1e-05 | 0.001866 | 0.001866 | 3.1652
2019-03-21 16:08 | 1553198880.8806 | 1090 | 0.2095 | 0.000466 | -4e-06 | 100.2204 | 100.1159 | 123.9877 | 21.2612 | 0.002044 | 0.001857 | -152.4545 | 0.24171 | 8.4036 | 0.003436 | 0.004542 | 0.004382 | 5e-05 | 0.001085 | 0.23258 | 1.3331 | 2.5479 | 98.84 | 20.89 | 38.545 | 71.0124 | 104.5467 | 66.5193 | 0.009644 | 1.0394 | 0.024832 | 23.8907 | 0.009331 | 0.14379 | -0.0002 | -0.014644 | -0.014644 | -0.40356
2019-03-21 16:08 | 1553198881.8819 | 1091 | 0.20949 | 0.000465 | -4e-06 | 100.1951 | 100.3806 | 124.0313 | 21.2612 | 0.001908 | 0.001855 | -152.4443 | 0.24196 | 8.3982 | 0.003397 | 0.004553 | 0.004328 | 5.2e-05 | 0.001085 | 0.26593 | 1.2958 | 2.548 | 98.84 | 20.89 | 38.543 | 67.7411 | 103.582 | 66.456 | 0.009644 | 1.0396 | 0.024838 | 23.8927 | -0.016542 | 0.14394 | -0.000105 | -0.007576 | -0.007576 | -0.78092
2019-03-21 16:08 | 1553198882.8821 | 1092 | 0.20948 | 0.000465 | -3e-06 | 100.3077 | 100.2913 | 124.113 | 21.2612 | 0.001969 | 0.001775 | -152.4524 | 0.2413 | 8.4017 | 0.003399 | 0.004986 | 0.004798 | 4.9e-05 | 0.001088 | 0.29966 | 1.3186 | 2.5481 | 98.84 | 20.89 | 38.539 | 69.1926 | 103.6387 | 64.5337 | 0.009643 | 1.0455 | 0.024917 | 23.8328 | 0.001461 | 0.14355 | -8.1e-05 | -0.005846 | -0.005846 | -1.0092
2019-03-21 16:08 | 1553198883.8807 | 1093 | 0.20948 | 0.000464 | -4e-06 | 100.2648 | 100.3652 | 124.1438 | 21.2612 | 0.002003 | 0.00184 | -152.448 | 0.24198 | 8.4015 | 0.003447 | 0.004649 | 0.004456 | 5.3e-05 | 0.001084 | 0.26089 | 1.3708 | 2.548 | 98.84 | 20.89 | 38.537 | 70.0276 | 104.8014 | 66.0941 | 0.009642 | 1.0407 | 0.024863 | 23.8908 | -0.008955 | 0.14395 | -4.2e-05 | -0.002912 | -0.002912 | -2.0318
2019-03-21 16:08 | 1553198884.8817 | 1094 | 0.2095 | 0.000466 | -5e-06 | 100.2931 | 100.1309 | 124.033 | 21.2612 | 0.001961 | 0.001844 | -152.4533 | 0.24121 | 8.4014 | 0.003412 | 0.005064 | 0.004854 | 5.3e-05 | 0.001086 | 0.26913 | 1.3483 | 2.548 | 98.84 | 20.89 | 38.535 | 69.0036 | 103.9631 | 66.1932 | 0.009642 | 1.0404 | 0.024848 | 23.8836 | 0.014478 | 0.1435 | -0.000204 | -0.014915 | -0.014915 | -0.3954
2019-03-21 16:08 | 1553198885.8813 | 1095 | 0.20949 | 0.000465 | -4e-06 | 100.3137 | 100.4763 | 124.1919 | 21.2612 | 0.001977 | 0.001842 | -152.4509 | 0.24055 | 8.4036 | 0.003412 | 0.004828 | 0.004676 | 5.5e-05 | 0.00109 | 0.28838 | 1.3038 | 2.548 | 98.84 | 20.89 | 38.532 | 69.3897 | 103.9491 | 66.1414 | 0.009641 | 1.0405 | 0.024866 | 23.8971 | -0.014486 | 0.1431 | -9.1e-05 | -0.006553 | -0.006553 | -0.89756
2019-03-21 16:08 | 1553198886.8809 | 1096 | 0.2095 | 0.000463 | -4e-06 | 100.2192 | 100.3853 | 124.1313 | 21.2612 | 0.001994 | 0.00184 | -152.4445 | 0.24073 | 8.4001 | 0.003415 | 0.004708 | 0.004523 | 5.3e-05 | 0.001087 | 0.25354 | 1.3043 | 2.5479 | 98.84 | 20.89 | 38.529 | 69.8038 | 104.0363 | 66.0994 | 0.00964 | 1.0407 | 0.024861 | 23.8897 | -0.014813 | 0.14321 | -0.000211 | -0.015467 | -0.015467 | -0.38056
2019-03-21 16:08 | 1553198887.8819 | 1097 | 0.20948 | 0.000465 | -4e-06 | 100.3843 | 100.3874 | 124.0626 | 21.2612 | 0.001955 | 0.001836 | -152.4234 | 0.24118 | 8.404 | 0.003409 | 0.004753 | 0.004555 | 5.3e-05 | 0.001088 | 0.29103 | 1.2951 | 2.5479 | 98.84 | 20.89 | 38.522 | 68.8594 | 103.8744 | 66.0144 | 0.009638 | 1.0409 | 0.024857 | 23.8801 | -0.000275 | 0.14348 | -4.1e-05 | -0.0028 | -0.0028 | -2.1061
2019-03-21 16:08 | 1553198888.8813 | 1098 | 0.20947 | 0.000464 | -4e-06 | 100.3342 | 100.372 | 124.0236 | 21.2612 | 0.00193 | 0.001844 | -152.424 | 0.2405 | 8.402 | 0.00342 | 0.004334 | 0.004215 | 5.6e-05 | 0.001086 | 0.26017 | 1.2977 | 2.548 | 98.84 | 20.89 | 38.52 | 68.2697 | 104.1558 | 66.2054 | 0.009638 | 1.0403 | 0.024847 | 23.8831 | -0.003371 | 0.14307 | -3e-06 | 1.6e-05 | 1.6e-05 | 356.4858
2019-03-21 16:08 | 1553198889.8801 | 1099 | 0.20951 | 0.000468 | -4e-06 | 100.3849 | 100.511 | 123.676 | 21.2612 | 0.001936 | 0.001867 | -152.4238 | 0.24101 | 8.4027 | 0.003458 | 0.004677 | 0.00445 | 5e-05 | 0.001087 | 0.27743 | 1.3042 | 2.548 | 98.84 | 20.89 | 38.509 | 68.3976 | 105.0838 | 66.755 | 0.009635 | 1.0387 | 0.024792 | 23.8689 | -0.011231 | 0.14338 | -0.000252 | -0.01851 | -0.01851 | -0.31836
2019-03-21 16:08 | 1553198890.8812 | 1100 | 0.20951 | 0.000463 | -4e-06 | 100.482 | 100.4466 | 123.6818 | 21.2612 | 0.00204 | 0.001872 | -152.4386 | 0.24017 | 8.4035 | 0.003513 | 0.004555 | 0.00437 | 5.4e-05 | 0.001086 | 0.26737 | 1.3168 | 2.5479 | 98.84 | 20.89 | 38.502 | 70.9088 | 106.4128 | 66.8805 | 0.009633 | 1.0383 | 0.024788 | 23.8739 | 0.003149 | 0.14288 | -0.000225 | -0.016519 | -0.016519 | -0.35549
2019-03-21 16:08 | 1553198891.8802 | 1101 | 0.20951 | 0.000463 | -4e-06 | 100.4391 | 100.6323 | 123.6358 | 21.2612 | 0.001968 | 0.001835 | -152.4552 | 0.24126 | 8.4003 | 0.0035 | 0.004661 | 0.004445 | 5.3e-05 | 0.001091 | 0.27704 | 1.3466 | 2.5481 | 98.84 | 20.88 | 38.498 | 69.1854 | 106.0892 | 65.9909 | 0.009626 | 1.041 | 0.024816 | 23.8382 | -0.017184 | 0.14353 | -0.000259 | -0.019024 | -0.019024 | -0.31009
2019-03-21 16:08 | 1553198892.8806 | 1102 | 0.20951 | 0.000466 | -4e-06 | 100.4677 | 100.667 | 123.8898 | 21.2612 | 0.001976 | 0.001841 | -152.4517 | 0.24126 | 8.4014 | 0.003539 | 0.004158 | 0.00404 | 5.1e-05 | 0.001092 | 0.28279 | 1.3554 | 2.5481 | 98.84 | 20.88 | 38.489 | 69.3632 | 107.033 | 66.1268 | 0.009624 | 1.0406 | 0.024836 | 23.8675 | -0.017725 | 0.14353 | -0.00026 | -0.019131 | -0.019131 | -0.30835
2019-03-21 16:08 | 1553198893.8813 | 1103 | 0.2095 | 0.000466 | -4e-06 | 100.4936 | 100.5733 | 123.6368 | 21.2612 | 0.001968 | 0.001864 | -152.4543 | 0.24108 | 8.4016 | 0.003458 | 0.005261 | 0.005073 | 4.9e-05 | 0.001087 | 0.25299 | 1.3125 | 2.548 | 98.84 | 20.88 | 38.486 | 69.1784 | 105.0736 | 66.6754 | 0.009623 | 1.0389 | 0.024791 | 23.8623 | -0.007091 | 0.14342 | -0.000201 | -0.01474 | -0.01474 | -0.3999
2019-03-21 16:08 | 1553198894.8813 | 1104 | 0.2095 | 0.00046 | -4e-06 | 100.4722 | 100.6162 | 123.5497 | 21.2612 | 0.001909 | 0.001853 | -152.4633 | 0.24168 | 8.4032 | 0.003483 | 0.005033 | 0.00481 | 5.3e-05 | 0.001094 | 0.26217 | 1.3895 | 2.5481 | 98.84 | 20.88 | 38.481 | 67.7596 | 105.6692 | 66.406 | 0.009622 | 1.0397 | 0.024792 | 23.8445 | -0.012817 | 0.14378 | -0.000204 | -0.014963 | -0.014963 | -0.39492
2019-03-21 16:08 | 1553198895.8814 | 1105 | 0.20947 | 0.00047 | -4e-06 | 100.5157 | 100.4116 | 123.4898 | 21.2612 | 0.001946 | 0.001849 | -152.4565 | 0.24168 | 8.4049 | 0.003551 | 0.004796 | 0.004629 | 5.3e-05 | 0.001095 | 0.26288 | 1.3358 | 2.548 | 98.84 | 20.88 | 38.473 | 68.6488 | 107.3364 | 66.3139 | 0.00962 | 1.04 | 0.024789 | 23.8355 | 0.00927 | 0.14377 | -1.7e-05 | -0.001058 | -0.001058 | -5.5861
2019-03-21 16:08 | 1553198896.8814 | 1106 | 0.20946 | 0.000462 | -4e-06 | 100.5386 | 100.7106 | 123.5354 | 21.2612 | 0.001916 | 0.00185 | -152.4504 | 0.24111 | 8.4031 | 0.003535 | 0.005087 | 0.004866 | 5.1e-05 | 0.00109 | 0.25265 | 1.3268 | 2.5481 | 98.84 | 20.88 | 38.47 | 67.9235 | 106.934 | 66.3399 | 0.009619 | 1.0399 | 0.024793 | 23.8408 | -0.015287 | 0.14344 | 9.2e-05 | 0.007111 | 0.007111 | 0.82904
2019-03-21 16:08 | 1553198897.8804 | 1107 | 0.20948 | 0.000464 | -4e-06 | 100.5532 | 100.7426 | 123.5584 | 21.2612 | 0.001932 | 0.001813 | -152.478 | 0.24189 | 8.4003 | 0.003487 | 0.004997 | 0.004863 | 5.1e-05 | 0.001094 | 0.27671 | 1.2799 | 2.5479 | 98.84 | 20.88 | 38.464 | 68.3152 | 105.7691 | 65.4557 | 0.009618 | 1.0426 | 0.024827 | 23.8119 | -0.016831 | 0.1439 | -4.1e-05 | -0.002814 | -0.002814 | -2.1015
2019-03-21 16:08 | 1553198898.8819 | 1108 | 0.2095 | 0.000464 | -4e-06 | 100.6486 | 100.822 | 123.4943 | 21.2612 | 0.001938 | 0.001795 | -152.488 | 0.24118 | 8.401 | 0.003493 | 0.005252 | 0.005121 | 5.3e-05 | 0.00109 | 0.24585 | 1.3273 | 2.5481 | 98.84 | 20.88 | 38.462 | 68.4587 | 105.9177 | 65.027 | 0.009617 | 1.044 | 0.024837 | 23.7907 | -0.015397 | 0.14348 | -0.000189 | -0.013823 | -0.013823 | -0.4266
2019-03-21 16:08 | 1553198899.8805 | 1109 | 0.2095 | 0.000463 | -4e-06 | 100.6232 | 100.6408 | 123.5757 | 21.2612 | 0.001976 | 0.001845 | -152.4789 | 0.24034 | 8.3986 | 0.003545 | 0.004704 | 0.004507 | 5.2e-05 | 0.001093 | 0.25227 | 1.3089 | 2.5479 | 98.84 | 20.88 | 38.46 | 69.3688 | 107.1772 | 66.2126 | 0.009617 | 1.0403 | 0.024801 | 23.8402 | -0.001559 | 0.14298 | -0.000199 | -0.014517 | -0.014517 | -0.40478
2019-03-21 16:08 | 1553198900.8807 | 1110 | 0.20954 | 0.000464 | -4e-06 | 100.6408 | 100.665 | 123.5598 | 21.2612 | 0.001915 | 0.001845 | -152.4626 | 0.24103 | 8.4003 | 0.003486 | 0.004884 | 0.004652 | 5.2e-05 | 0.001094 | 0.26792 | 1.3342 | 2.5479 | 98.84 | 20.88 | 38.456 | 67.9017 | 105.7414 | 66.2102 | 0.009616 | 1.0403 | 0.0248 | 23.8386 | -0.002145 | 0.14339 | -0.000433 | -0.031965 | -0.031965 | -0.18437
2019-03-21 16:08 | 1553198901.8817 | 1111 | 0.20953 | 0.000462 | -4e-06 | 100.6229 | 100.7317 | 123.6196 | 21.2612 | 0.001918 | 0.001824 | -152.4625 | 0.24076 | 8.4059 | 0.003503 | 0.005034 | 0.004823 | 5.3e-05 | 0.001092 | 0.2401 | 1.2967 | 2.5479 | 98.84 | 20.88 | 38.454 | 67.975 | 106.1712 | 65.7085 | 0.009615 | 1.0419 | 0.024824 | 23.8267 | -0.009664 | 0.14323 | -0.000366 | -0.027067 | -0.027067 | -0.21749
2019-03-21 16:08 | 1553198902.8822 | 1112 | 0.20955 | 0.000468 | -4e-06 | 100.62 | 100.59 | 123.7403 | 21.2612 | 0.001966 | 0.001854 | -152.4716 | 0.24093 | 8.4012 | 0.003505 | 0.004708 | 0.004591 | 5.3e-05 | 0.001095 | 0.29839 | 1.301 | 2.5478 | 98.84 | 20.88 | 38.45 | 69.1245 | 106.2231 | 66.4358 | 0.009614 | 1.0396 | 0.02481 | 23.8639 | 0.002661 | 0.14333 | -0.000482 | -0.035697 | -0.035697 | -0.16502
2019-03-21 16:08 | 1553198903.8819 | 1113 | 0.2095 | 0.000464 | -3e-06 | 100.6617 | 100.762 | 123.829 | 21.2612 | 0.001937 | 0.001843 | -152.485 | 0.24057 | 8.4033 | 0.003556 | 0.004697 | 0.004555 | 5.4e-05 | 0.001092 | 0.25337 | 1.3207 | 2.548 | 98.84 | 20.88 | 38.448 | 68.4265 | 107.4614 | 66.1626 | 0.009614 | 1.0405 | 0.024829 | 23.8629 | -0.008913 | 0.14311 | -0.000156 | -0.011404 | -0.011404 | -0.51577
2019-03-21 16:08 | 1553198904.8809 | 1114 | 0.20947 | 0.000466 | -4e-06 | 100.6778 | 100.8955 | 123.7485 | 21.2612 | 0.001986 | 0.001848 | -152.4776 | 0.24153 | 8.4019 | 0.003477 | 0.004878 | 0.004696 | 5e-05 | 0.001092 | 0.28208 | 1.3279 | 2.548 | 98.84 | 20.88 | 38.446 | 69.5995 | 105.5275 | 66.285 | 0.009613 | 1.0401 | 0.024816 | 23.8594 | -0.019327 | 0.14369 | 1.1e-05 | 0.001023 | 0.001023 | 5.7721
2019-03-21 16:08 | 1553198905.8809 | 1115 | 0.20956 | 0.000471 | -4e-06 | 100.726 | 100.7722 | 123.6447 | 21.2612 | 0.001946 | 0.001842 | -152.4965 | 0.24176 | 8.4046 | 0.003536 | 0.004692 | 0.004431 | 4.9e-05 | 0.00109 | 0.25586 | 1.3021 | 2.5479 | 98.84 | 20.88 | 38.443 | 68.6387 | 106.9527 | 66.154 | 0.009613 | 1.0405 | 0.02481 | 23.8448 | -0.004103 | 0.14382 | -0.000574 | -0.042469 | -0.042469 | -0.13919
2019-03-21 16:08 | 1553198906.88 | 1116 | 0.2095 | 0.000466 | -4e-06 | 100.6834 | 100.6384 | 123.6311 | 21.2612 | 0.00193 | 0.001865 | -152.5051 | 0.24177 | 8.4032 | 0.003454 | 0.005084 | 0.004993 | 5.2e-05 | 0.001094 | 0.27599 | 1.3314 | 2.5479 | 98.84 | 20.88 | 38.442 | 68.2589 | 104.9721 | 66.7018 | 0.009612 | 1.0388 | 0.024789 | 23.8627 | 0.004002 | 0.14383 | -0.000206 | -0.015044 | -0.015044 | -0.39294
2019-03-21 16:08 | 1553198907.8813 | 1117 | 0.20947 | 0.000464 | -4e-06 | 100.681 | 100.9449 | 123.5652 | 21.2612 | 0.001966 | 0.001867 | -152.4721 | 0.2415 | 8.4023 | 0.003521 | 0.00514 | 0.004946 | 5e-05 | 0.001091 | 0.26432 | 1.345 | 2.548 | 98.84 | 20.88 | 38.439 | 69.1301 | 106.5953 | 66.7487 | 0.009612 | 1.0387 | 0.024781 | 23.858 | -0.023413 | 0.14367 | 2e-06 | 0.000392 | 0.000392 | 15.0444
2019-03-21 16:08 | 1553198908.88 | 1118 | 0.20947 | 0.000467 | -4e-06 | 100.7281 | 100.6156 | 123.4894 | 21.2612 | 0.001977 | 0.001843 | -152.4553 | 0.2416 | 8.4019 | 0.003452 | 0.005243 | 0.005064 | 5.1e-05 | 0.001095 | 0.27383 | 1.3655 | 2.548 | 98.84 | 20.88 | 38.438 | 69.4022 | 104.9342 | 66.1713 | 0.009611 | 1.0404 | 0.024794 | 23.8304 | 0.009999 | 0.14373 | 2.3e-05 | 0.001942 | 0.001942 | 3.0425
2019-03-21 16:08 | 1553198909.8805 | 1119 | 0.20949 | 0.000463 | -4e-06 | 100.7055 | 100.6251 | 123.745 | 21.2612 | 0.00194 | 0.001797 | -152.4838 | 0.24192 | 8.4045 | 0.003502 | 0.004754 | 0.004573 | 5e-05 | 0.001094 | 0.26808 | 1.3359 | 2.548 | 98.84 | 20.88 | 38.437 | 68.5077 | 106.1404 | 65.0551 | 0.009611 | 1.0439 | 0.024861 | 23.8158 | 0.007141 | 0.14392 | -0.00011 | -0.007948 | -0.007948 | -0.7442
2019-03-21 16:08 | 1553198910.8801 | 1120 | 0.20951 | 0.00046 | -4e-06 | 100.7302 | 100.888 | 123.6091 | 21.2612 | 0.001912 | 0.001866 | -152.4907 | 0.2418 | 8.4032 | 0.00339 | 0.005576 | 0.005412 | 5.2e-05 | 0.001096 | 0.2625 | 1.2782 | 2.5481 | 98.84 | 20.88 | 38.435 | 67.8401 | 103.4202 | 66.7275 | 0.009611 | 1.0387 | 0.024786 | 23.8615 | -0.014005 | 0.14385 | -0.000237 | -0.017421 | -0.017421 | -0.33937
2019-03-21 16:09 | 1553198911.8812 | 1121 | 0.20948 | 0.000465 | -4e-06 | 100.695 | 100.7542 | 123.3429 | 21.2612 | 0.001922 | 0.001858 | -152.5045 | 0.2415 | 8.4031 | 0.003387 | 0.004313 | 0.004239 | 4.8e-05 | 0.001097 | 0.26521 | 1.2904 | 2.5478 | 98.84 | 20.88 | 38.433 | 68.0671 | 103.3432 | 66.5356 | 0.00961 | 1.0393 | 0.024766 | 23.8291 | -0.005252 | 0.14367 | -8.5e-05 | -0.006056 | -0.006056 | -0.97507
2019-03-21 16:09 | 1553198912.8807 | 1122 | 0.2095 | 0.000465 | -4e-06 | 100.7794 | 100.8965 | 122.8932 | 21.2612 | 0.001963 | 0.001842 | -152.4758 | 0.2414 | 8.4028 | 0.003458 | 0.004332 | 0.004111 | 5.1e-05 | 0.001097 | 0.26233 | 1.293 | 2.5479 | 98.84 | 20.88 | 38.432 | 69.0613 | 105.0625 | 66.1549 | 0.00961 | 1.0405 | 0.024735 | 23.7722 | -0.010395 | 0.14361 | -0.000171 | -0.012445 | -0.012445 | -0.47425
2019-03-21 16:09 | 1553198913.8817 | 1123 | 0.20949 | 0.000463 | -4e-06 | 100.715 | 100.8192 | 122.5336 | 21.2612 | 0.001956 | 0.001848 | -152.4664 | 0.24107 | 8.4016 | 0.00348 | 0.004575 | 0.004355 | 5.1e-05 | 0.001099 | 0.26001 | 1.3155 | 2.548 | 98.84 | 20.88 | 38.43 | 68.8945 | 105.5945 | 66.2985 | 0.009609 | 1.0401 | 0.024693 | 23.7425 | -0.009253 | 0.14341 | -0.000136 | -0.009866 | -0.009866 | -0.59741
2019-03-21 16:09 | 1553198914.8814 | 1124 | 0.20946 | 0.000466 | -4e-06 | 100.7636 | 100.7419 | 122.8519 | 21.2612 | 0.001961 | 0.0018 | -152.4834 | 0.24159 | 8.4014 | 0.003426 | 0.004507 | 0.004364 | 4.9e-05 | 0.001097 | 0.25818 | 1.3725 | 2.5479 | 98.84 | 20.88 | 38.428 | 69.0029 | 104.2951 | 65.1355 | 0.009609 | 1.0436 | 0.024768 | 23.7325 | 0.001921 | 0.14372 | 8.6e-05 | 0.006664 | 0.006664 | 0.88635
2019-03-21 16:09 | 1553198915.8818 | 1125 | 0.20946 | 0.000465 | -4e-06 | 100.7505 | 100.8209 | 123.3072 | 21.2612 | 0.001916 | 0.00184 | -152.4687 | 0.24118 | 8.4021 | 0.003414 | 0.004827 | 0.004632 | 5.4e-05 | 0.0011 | 0.25371 | 1.3168 | 2.548 | 98.84 | 20.88 | 38.424 | 67.9267 | 103.9985 | 66.1001 | 0.009608 | 1.0407 | 0.024779 | 23.8103 | -0.006255 | 0.14348 | 3.8e-05 | 0.003071 | 0.003071 | 1.9203
2019-03-21 16:09 | 1553198916.8805 | 1126 | 0.20951 | 0.000462 | -4e-06 | 100.7868 | 100.8186 | 123.5081 | 21.2612 | 0.001908 | 0.001835 | -152.4629 | 0.24058 | 8.4041 | 0.003373 | 0.004775 | 0.004579 | 5e-05 | 0.001097 | 0.28456 | 1.3648 | 2.548 | 98.84 | 20.88 | 38.422 | 67.7245 | 103.006 | 65.969 | 0.009607 | 1.0411 | 0.024804 | 23.8251 | -0.002817 | 0.14312 | -0.000245 | -0.018041 | -0.018041 | -0.32605
2019-03-21 16:09 | 1553198917.8815 | 1127 | 0.20949 | 0.000461 | -4e-06 | 100.7251 | 100.824 | 123.4706 | 21.2612 | 0.001983 | 0.00181 | -152.4959 | 0.2417 | 8.4041 | 0.003373 | 0.004123 | 0.00398 | 4.9e-05 | 0.001097 | 0.26985 | 1.3258 | 2.5479 | 98.84 | 20.88 | 38.42 | 69.5293 | 103.024 | 65.3681 | 0.009607 | 1.0429 | 0.024822 | 23.8004 | -0.008776 | 0.14378 | -0.000115 | -0.008336 | -0.008336 | -0.70891
2019-03-21 16:09 | 1553198918.882 | 1128 | 0.20945 | 0.000465 | -4e-06 | 100.8121 | 100.9538 | 123.4903 | 21.2612 | 0.001915 | 0.001859 | -152.5115 | 0.24108 | 8.4049 | 0.003349 | 0.004822 | 0.004623 | 4.9e-05 | 0.001095 | 0.27743 | 1.3286 | 2.548 | 98.84 | 20.88 | 38.415 | 67.8971 | 102.4222 | 66.5582 | 0.009606 | 1.0393 | 0.02478 | 23.8441 | -0.012558 | 0.14342 | 0.000103 | 0.007912 | 0.007912 | 0.74502
2019-03-21 16:09 | 1553198919.882 | 1129 | 0.20946 | 0.000464 | -5e-06 | 100.8506 | 100.8669 | 123.0748 | 21.2612 | 0.001919 | 0.001882 | -152.4919 | 0.24138 | 8.4008 | 0.003353 | 0.00425 | 0.00405 | 5.3e-05 | 0.001096 | 0.25227 | 1.301 | 2.548 | 98.84 | 20.88 | 38.412 | 68.007 | 102.5188 | 67.102 | 0.009605 | 1.0376 | 0.024719 | 23.823 | -0.00145 | 0.1436 | 6e-05 | 0.004769 | 0.004769 | 1.2375
2019-03-21 16:09 | 1553198920.8807 | 1130 | 0.20946 | 0.000462 | -4e-06 | 100.8506 | 100.9105 | 123.1192 | 21.2612 | 0.001987 | 0.001835 | -152.5047 | 0.24052 | 8.4019 | 0.003344 | 0.003909 | 0.00377 | 5.1e-05 | 0.001099 | 0.274 | 1.3139 | 2.548 | 98.84 | 20.88 | 38.405 | 69.6353 | 102.3201 | 65.9856 | 0.009603 | 1.041 | 0.024764 | 23.7882 | -0.005315 | 0.14308 | 7.7e-05 | 0.005986 | 0.005986 | 0.98247
2019-03-21 16:09 | 1553198921.8798 | 1131 | 0.20948 | 0.000465 | -4e-06 | 100.7654 | 101.0522 | 123.3366 | 21.2612 | 0.001902 | 0.001844 | -152.4794 | 0.24127 | 8.4001 | 0.003365 | 0.004318 | 0.004173 | 5.3e-05 | 0.001097 | 0.25625 | 1.3228 | 2.548 | 98.84 | 20.88 | 38.403 | 67.5809 | 102.8205 | 66.204 | 0.009603 | 1.0403 | 0.024778 | 23.8168 | -0.025426 | 0.14353 | -4.1e-05 | -0.002799 | -0.002799 | -2.1074
2019-03-21 16:09 | 1553198922.8802 | 1132 | 0.20946 | 0.000464 | -4e-06 | 100.8798 | 100.8155 | 123.2288 | 21.2612 | 0.001913 | 0.001821 | -152.5108 | 0.24077 | 8.4016 | 0.003411 | 0.004431 | 0.004232 | 4.8e-05 | 0.001095 | 0.24746 | 1.3512 | 2.5479 | 98.84 | 20.87 | 38.4 | 67.8634 | 103.9403 | 65.6529 | 0.009596 | 1.042 | 0.024787 | 23.7871 | 0.005701 | 0.14324 | 4.8e-05 | 0.003789 | 0.003789 | 1.5537
2019-03-21 16:09 | 1553198923.881 | 1133 | 0.20947 | 0.000466 | -3e-06 | 100.8997 | 101.2068 | 123.4818 | 21.2612 | 0.001944 | 0.001868 | -152.5267 | 0.2417 | 8.4038 | 0.003445 | 0.004355 | 0.004064 | 5.3e-05 | 0.001096 | 0.25299 | 1.3059 | 2.548 | 98.84 | 20.87 | 38.391 | 68.5936 | 104.752 | 66.7776 | 0.009594 | 1.0386 | 0.024771 | 23.851 | -0.027183 | 0.14379 | -1.5e-05 | -0.000879 | -0.000879 | -6.7257
2019-03-21 16:09 | 1553198924.8809 | 1134 | 0.20946 | 0.000463 | -4e-06 | 100.895 | 101.1459 | 123.3673 | 21.2612 | 0.001951 | 0.001833 | -152.5127 | 0.24126 | 8.4039 | 0.003472 | 0.004621 | 0.004522 | 5.3e-05 | 0.001096 | 0.25498 | 1.3348 | 2.5478 | 98.84 | 20.87 | 38.389 | 68.7683 | 105.4007 | 65.9293 | 0.009593 | 1.0412 | 0.024791 | 23.8101 | -0.022215 | 0.14353 | 7.4e-05 | 0.005719 | 0.005719 | 1.0315
2019-03-21 16:09 | 1553198925.8811 | 1135 | 0.20946 | 0.000463 | -4e-06 | 100.9155 | 101.0495 | 123.0382 | 21.2612 | 0.00196 | 0.001806 | -152.5348 | 0.24131 | 8.3993 | 0.003498 | 0.004702 | 0.004498 | 4.9e-05 | 0.001094 | 0.27295 | 1.3655 | 2.5481 | 98.84 | 20.88 | 38.386 | 68.9967 | 106.0317 | 65.2738 | 0.009598 | 1.0432 | 0.024782 | 23.7554 | -0.011865 | 0.14356 | 9.7e-05 | 0.007447 | 0.007447 | 0.79227
2019-03-21 16:09 | 1553198926.8804 | 1136 | 0.20948 | 0.000464 | -4e-06 | 100.8866 | 100.9865 | 122.7737 | 21.2612 | 0.001952 | 0.001825 | -152.5274 | 0.24167 | 8.4034 | 0.003492 | 0.004976 | 0.004798 | 5.3e-05 | 0.00109 | 0.27383 | 1.2934 | 2.5478 | 98.84 | 20.88 | 38.379 | 68.7955 | 105.8894 | 65.746 | 0.009597 | 1.0418 | 0.024738 | 23.7464 | -0.008848 | 0.14377 | -7.5e-05 | -0.005368 | -0.005368 | -1.1008
2019-03-21 16:09 | 1553198927.881 | 1137 | 0.20947 | 0.000462 | -4e-06 | 100.9766 | 100.9027 | 122.7988 | 21.2612 | 0.001946 | 0.001827 | -152.5095 | 0.24199 | 8.4039 | 0.003499 | 0.004691 | 0.004595 | 5.3e-05 | 0.001095 | 0.27975 | 1.3278 | 2.548 | 98.84 | 20.88 | 38.376 | 68.638 | 106.072 | 65.7811 | 0.009596 | 1.0416 | 0.024739 | 23.75 | 0.00655 | 0.14396 | -5e-06 | -0.000145 | -0.000145 | -40.6715
2019-03-21 16:09 | 1553198928.8811 | 1138 | 0.20947 | 0.000464 | -3e-06 | 100.9671 | 100.9309 | 122.723 | 21.2612 | 0.002 | 0.00185 | -152.4831 | 0.241 | 8.4004 | 0.003481 | 0.005159 | 0.005012 | 5.5e-05 | 0.001095 | 0.25818 | 1.3276 | 2.548 | 98.84 | 20.88 | 38.371 | 69.9357 | 105.6289 | 66.3305 | 0.009595 | 1.04 | 0.024711 | 23.7619 | 0.003199 | 0.14337 | -2.8e-05 | -0.001848 | -0.001848 | -3.1882
2019-03-21 16:09 | 1553198929.8815 | 1139 | 0.20947 | 0.000468 | -4e-06 | 100.9665 | 101.0113 | 123.1164 | 21.2612 | 0.001985 | 0.001837 | -152.4874 | 0.24098 | 8.4037 | 0.003496 | 0.004613 | 0.004441 | 5.5e-05 | 0.001098 | 0.27505 | 1.3019 | 2.548 | 98.84 | 20.88 | 38.369 | 69.5964 | 106.0012 | 66.0371 | 0.009594 | 1.0409 | 0.024762 | 23.7897 | -0.003974 | 0.14336 | -1.2e-05 | -0.000652 | -0.000652 | -9.0314
2019-03-21 16:09 | 1553198930.8802 | 1140 | 0.20948 | 0.000468 | -4e-06 | 100.9924 | 101.0791 | 123.3992 | 21.2612 | 0.001956 | 0.001812 | -152.4876 | 0.24124 | 8.4026 | 0.003462 | 0.005003 | 0.004842 | 5.5e-05 | 0.001094 | 0.26698 | 1.2911 | 2.5479 | 98.84 | 20.88 | 38.363 | 68.8798 | 105.1648 | 65.4157 | 0.009593 | 1.0428 | 0.024813 | 23.7952 | -0.007677 | 0.14351 | -4.8e-05 | -0.003307 | -0.003307 | -1.7836
2019-03-21 16:09 | 1553198931.8803 | 1141 | 0.20946 | 0.000463 | -4e-06 | 100.968 | 101.0137 | 123.7044 | 21.2612 | 0.001929 | 0.001827 | -152.4753 | 0.2415 | 8.4028 | 0.003518 | 0.005013 | 0.004795 | 5.5e-05 | 0.001094 | 0.24618 | 1.3039 | 2.548 | 98.84 | 20.88 | 38.361 | 68.2464 | 106.523 | 65.7802 | 0.009592 | 1.0416 | 0.02483 | 23.8374 | -0.004053 | 0.14367 | 6.2e-05 | 0.00485 | 0.00485 | 1.2175
2019-03-21 16:09 | 1553198932.8814 | 1142 | 0.20947 | 0.000468 | -4e-06 | 101.0499 | 101.1561 | 123.9757 | 21.2612 | 0.001888 | 0.001826 | -152.4765 | 0.24051 | 8.4016 | 0.003516 | 0.005081 | 0.004931 | 5.4e-05 | 0.001096 | 0.26488 | 1.2948 | 2.5479 | 98.84 | 20.88 | 38.36 | 67.2635 | 106.4813 | 65.7624 | 0.009592 | 1.0417 | 0.024858 | 23.8629 | -0.009395 | 0.14308 | -1.8e-05 | -0.001124 | -0.001124 | -5.2323
2019-03-21 16:09 | 1553198933.8813 | 1143 | 0.20945 | 0.000467 | -4e-06 | 100.9712 | 101.2385 | 123.7933 | 21.2612 | 0.001975 | 0.001827 | -152.4629 | 0.24118 | 8.4033 | 0.003461 | 0.004876 | 0.004725 | 5.4e-05 | 0.001097 | 0.27903 | 1.352 | 2.548 | 98.84 | 20.88 | 38.355 | 69.3469 | 105.1372 | 65.7944 | 0.009591 | 1.0416 | 0.024839 | 23.8465 | -0.023645 | 0.14348 | 0.000104 | 0.007989 | 0.007989 | 0.73811
2019-03-21 16:09 | 1553198934.8815 | 1144 | 0.20945 | 0.000464 | -4e-06 | 100.9972 | 101.088 | 123.817 | 21.2612 | 0.001955 | 0.001911 | -152.4421 | 0.24153 | 8.4023 | 0.003462 | 0.003976 | 0.003804 | 5.4e-05 | 0.001101 | 0.28368 | 1.3681 | 2.5481 | 98.84 | 20.88 | 38.354 | 68.8712 | 105.1655 | 67.8057 | 0.00959 | 1.0355 | 0.024768 | 23.9194 | -0.008039 | 0.14369 | 0.000115 | 0.008821 | 0.008821 | 0.66946
2019-03-21 16:09 | 1553198935.8809 | 1145 | 0.20946 | 0.000464 | -4e-06 | 101.0341 | 101.0348 | 123.8842 | 21.2612 | 0.001951 | 0.001872 | -152.4244 | 0.24192 | 8.4001 | 0.003468 | 0.004963 | 0.004709 | 5.2e-05 | 0.001096 | 0.27959 | 1.3465 | 2.548 | 98.84 | 20.88 | 38.353 | 68.7791 | 105.3127 | 66.8665 | 0.00959 | 1.0383 | 0.024808 | 23.8929 | -6.5e-05 | 0.14392 | 8.8e-05 | 0.006799 | 0.006799 | 0.87004
2019-03-21 16:09 | 1553198936.8805 | 1146 | 0.20949 | 0.000464 | -3e-06 | 101.1104 | 101.2542 | 123.8515 | 21.2612 | 0.001981 | 0.001856 | -152.4303 | 0.24092 | 8.4041 | 0.003462 | 0.004787 | 0.00464 | 5.2e-05 | 0.001098 | 0.2557 | 1.3453 | 2.5481 | 98.84 | 20.88 | 38.347 | 69.4834 | 105.1813 | 66.488 | 0.009589 | 1.0395 | 0.024819 | 23.8765 | -0.012712 | 0.14333 | -8.8e-05 | -0.006361 | -0.006361 | -0.92601
2019-03-21 16:09 | 1553198937.8806 | 1147 | 0.20948 | 0.000464 | -4e-06 | 101.032 | 101.293 | 123.6976 | 21.2612 | 0.001915 | 0.001859 | -152.4221 | 0.24151 | 8.4021 | 0.003421 | 0.004595 | 0.004394 | 5.4e-05 | 0.001094 | 0.25857 | 1.3136 | 2.548 | 98.84 | 20.88 | 38.345 | 67.8985 | 104.1827 | 66.5614 | 0.009588 | 1.0393 | 0.024801 | 23.8642 | -0.023075 | 0.14368 | -5.8e-05 | -0.004047 | -0.004047 | -1.4592
2019-03-21 16:09 | 1553198938.88 | 1148 | 0.20947 | 0.000463 | -4e-06 | 101.0237 | 101.2388 | 123.9417 | 21.2612 | 0.001924 | 0.001851 | -152.4247 | 0.24216 | 8.4 | 0.003458 | 0.005138 | 0.004994 | 4.8e-05 | 0.001088 | 0.27024 | 1.3242 | 2.5481 | 98.84 | 20.88 | 38.34 | 68.1272 | 105.0751 | 66.367 | 0.009587 | 1.0398 | 0.024832 | 23.8809 | -0.019031 | 0.14406 | -7e-06 | -0.000302 | -0.000302 | -19.6129
2019-03-21 16:09 | 1553198939.8797 | 1149 | 0.20946 | 0.000464 | -4e-06 | 101.0764 | 101.2508 | 123.6504 | 21.2612 | 0.001974 | 0.001837 | -152.4258 | 0.24115 | 8.4026 | 0.003511 | 0.004904 | 0.004736 | 5e-05 | 0.001093 | 0.26161 | 1.338 | 2.548 | 98.84 | 20.88 | 38.337 | 69.3164 | 106.3655 | 66.0402 | 0.009586 | 1.0408 | 0.024815 | 23.8413 | -0.015416 | 0.14346 | 4.1e-05 | 0.00327 | 0.00327 | 1.8032
2019-03-21 16:09 | 1553198940.8798 | 1150 | 0.20944 | 0.000462 | -4e-06 | 101.1023 | 101.2916 | 123.5875 | 21.2612 | 0.001978 | 0.001825 | -152.4145 | 0.24057 | 8.4033 | 0.003485 | 0.004918 | 0.004826 | 5.1e-05 | 0.001089 | 0.26985 | 1.329 | 2.548 | 98.84 | 20.88 | 38.334 | 69.4178 | 105.7274 | 65.7506 | 0.009585 | 1.0417 | 0.024819 | 23.8251 | -0.016732 | 0.14312 | 0.000203 | 0.015401 | 0.015401 | 0.38192
2019-03-21 16:09 | 1553198941.8797 | 1151 | 0.20945 | 0.000459 | -4e-06 | 101.0967 | 101.3448 | 123.1762 | 21.2612 | 0.001945 | 0.001807 | -152.434 | 0.24132 | 8.4048 | 0.003489 | 0.004356 | 0.004106 | 5.3e-05 | 0.001089 | 0.24585 | 1.3726 | 2.5481 | 98.84 | 20.88 | 38.329 | 68.6169 | 105.8343 | 65.2979 | 0.009584 | 1.0431 | 0.024795 | 23.7695 | -0.021924 | 0.14356 | 0.000124 | 0.009482 | 0.009482 | 0.62225
2019-03-21 16:09 | 1553198942.8798 | 1152 | 0.20944 | 0.000464 | -4e-06 | 101.1047 | 101.3533 | 123.2349 | 21.2612 | 0.001957 | 0.001838 | -152.4316 | 0.24174 | 8.4031 | 0.003512 | 0.004854 | 0.004709 | 5e-05 | 0.00109 | 0.26504 | 1.347 | 2.548 | 98.84 | 20.88 | 38.327 | 68.9024 | 106.3845 | 66.0409 | 0.009584 | 1.0408 | 0.024773 | 23.8013 | -0.021963 | 0.14381 | 0.000174 | 0.013165 | 0.013165 | 0.44898
2019-03-21 16:09 | 1553198943.8803 | 1153 | 0.20945 | 0.000462 | -4e-06 | 101.1265 | 101.4922 | 123.1924 | 21.2612 | 0.001976 | 0.001874 | -152.4394 | 0.24113 | 8.4015 | 0.003487 | 0.004655 | 0.004417 | 5.3e-05 | 0.001091 | 0.26488 | 1.3019 | 2.548 | 98.84 | 20.89 | 38.32 | 69.3632 | 105.7768 | 66.9086 | 0.009588 | 1.0382 | 0.024738 | 23.8276 | -0.032294 | 0.14345 | 0.000101 | 0.007751 | 0.007751 | 0.7606
2019-03-21 16:09 | 1553198944.88 | 1154 | 0.20948 | 0.000457 | -4e-06 | 101.1986 | 101.2562 | 123.0117 | 21.2612 | 0.001969 | 0.001806 | -152.4663 | 0.24036 | 8.4038 | 0.003511 | 0.005805 | 0.005648 | 5.2e-05 | 0.001092 | 0.26394 | 1.3464 | 2.5479 | 98.84 | 20.88 | 38.316 | 69.2019 | 106.3592 | 65.2892 | 0.009581 | 1.0432 | 0.024778 | 23.7533 | -0.005094 | 0.14299 | -4.8e-05 | -0.003366 | -0.003366 | -1.7459
2019-03-21 16:09 | 1553198945.881 | 1155 | 0.20949 | 0.000459 | -4e-06 | 101.1634 | 101.3271 | 123.1488 | 21.2612 | 0.001951 | 0.001867 | -152.4775 | 0.24101 | 8.4021 | 0.003617 | 0.005028 | 0.004787 | 5.5e-05 | 0.001089 | 0.26217 | 1.3676 | 2.548 | 98.84 | 20.89 | 38.313 | 68.7597 | 108.9318 | 66.7619 | 0.009586 | 1.0386 | 0.024739 | 23.8182 | -0.014458 | 0.14338 | -8.9e-05 | -0.006418 | -0.006418 | -0.91817
2019-03-21 16:09 | 1553198946.8801 | 1156 | 0.20946 | 0.000462 | -4e-06 | 101.2051 | 101.3131 | 123.4884 | 21.2612 | 0.001968 | 0.001812 | -152.463 | 0.24098 | 8.4025 | 0.003549 | 0.00519 | 0.005 | 5.2e-05 | 0.001094 | 0.27616 | 1.3504 | 2.5478 | 98.84 | 20.89 | 38.306 | 69.1721 | 107.2765 | 65.4213 | 0.009584 | 1.0428 | 0.024822 | 23.804 | -0.009538 | 0.14336 | 6.8e-05 | 0.005258 | 0.005258 | 1.1206
2019-03-21 16:09 | 1553198947.8805 | 1157 | 0.20949 | 0.000459 | -4e-06 | 101.2671 | 101.4704 | 123.7304 | 21.2612 | 0.001982 | 0.001839 | -152.4841 | 0.24155 | 8.4018 | 0.003603 | 0.004777 | 0.004621 | 5.6e-05 | 0.001092 | 0.28279 | 1.3567 | 2.5479 | 98.84 | 20.89 | 38.304 | 69.5123 | 108.6007 | 66.069 | 0.009584 | 1.0408 | 0.024822 | 23.8501 | -0.017943 | 0.1437 | -0.000131 | -0.009528 | -0.009528 | -0.61982
2019-03-21 16:09 | 1553198948.8802 | 1158 | 0.20946 | 0.000464 | -4e-06 | 101.212 | 101.2814 | 123.5896 | 21.2612 | 0.001942 | 0.001781 | -152.5115 | 0.24146 | 8.4031 | 0.003583 | 0.005084 | 0.004894 | 5.5e-05 | 0.001099 | 0.29535 | 1.3438 | 2.5479 | 98.84 | 20.89 | 38.301 | 68.5536 | 108.1036 | 64.6781 | 0.009583 | 1.045 | 0.024859 | 23.7876 | -0.006135 | 0.14365 | 5.8e-05 | 0.004577 | 0.004577 | 1.2898
2019-03-21 16:09 | 1553198949.8813 | 1159 | 0.20949 | 0.000459 | -4e-06 | 101.1911 | 101.466 | 123.6112 | 21.2612 | 0.001988 | 0.001834 | -152.5023 | 0.24071 | 8.4022 | 0.003582 | 0.004932 | 0.004676 | 4.9e-05 | 0.001096 | 0.25177 | 1.3174 | 2.5478 | 98.84 | 20.89 | 38.298 | 69.661 | 108.0815 | 65.9683 | 0.009582 | 1.0411 | 0.024814 | 23.835 | -0.024266 | 0.1432 | -0.000139 | -0.010136 | -0.010136 | -0.58067
2019-03-21 16:09 | 1553198950.8806 | 1160 | 0.20951 | 0.000459 | -4e-06 | 101.2799 | 101.354 | 123.4307 | 21.2612 | 0.001918 | 0.001839 | -152.4948 | 0.24157 | 8.4009 | 0.003477 | 0.004927 | 0.004754 | 5.4e-05 | 0.001098 | 0.25984 | 1.2998 | 2.5479 | 98.84 | 20.89 | 38.296 | 67.9836 | 105.5377 | 66.0767 | 0.009582 | 1.0407 | 0.024792 | 23.8214 | -0.006538 | 0.14371 | -0.000212 | -0.015549 | -0.015549 | -0.37985
2019-03-21 16:09 | 1553198951.8798 | 1161 | 0.20942 | 0.000459 | -4e-06 | 101.2164 | 101.3884 | 123.6344 | 21.2612 | 0.001896 | 0.001839 | -152.4949 | 0.24148 | 8.405 | 0.003556 | 0.004957 | 0.004764 | 5e-05 | 0.001096 | 0.24618 | 1.31 | 2.5481 | 98.84 | 20.89 | 38.291 | 67.454 | 107.441 | 66.0791 | 0.009581 | 1.0407 | 0.024812 | 23.8412 | -0.01518 | 0.14366 | 0.000323 | 0.024283 | 0.024283 | 0.24314
2019-03-21 16:09 | 1553198952.8795 | 1162 | 0.20947 | 0.000462 | -3e-06 | 101.206 | 101.389 | 123.276 | 21.2612 | 0.001977 | 0.001866 | -152.48 | 0.24192 | 8.4037 | 0.003636 | 0.005149 | 0.00502 | 5.2e-05 | 0.001095 | 0.23308 | 1.3865 | 2.5479 | 98.84 | 20.89 | 38.287 | 69.3865 | 109.3898 | 66.7283 | 0.00958 | 1.0387 | 0.024753 | 23.8294 | -0.016161 | 0.14392 | -2.3e-05 | -0.001463 | -0.001463 | -4.0444
2019-03-21 16:09 | 1553198953.8808 | 1163 | 0.20946 | 0.000464 | -4e-06 | 101.2197 | 101.417 | 123.6706 | 21.2612 | 0.001987 | 0.001837 | -152.4687 | 0.24211 | 8.3992 | 0.003583 | 0.00522 | 0.00503 | 5.1e-05 | 0.001099 | 0.29518 | 1.3917 | 2.5479 | 98.84 | 20.89 | 38.286 | 69.6384 | 108.0956 | 66.0176 | 0.009579 | 1.0409 | 0.024818 | 23.8425 | -0.017414 | 0.14403 | 8.8e-05 | 0.006774 | 0.006774 | 0.87387
2019-03-21 16:09 | 1553198954.8809 | 1164 | 0.20946 | 0.000459 | -4e-06 | 101.2888 | 101.3686 | 123.8957 | 21.2612 | 0.001992 | 0.001834 | -152.4619 | 0.24112 | 8.4013 | 0.003524 | 0.004689 | 0.004514 | 5.4e-05 | 0.001102 | 0.28727 | 1.3109 | 2.5481 | 98.84 | 20.89 | 38.285 | 69.7617 | 106.6647 | 65.9644 | 0.009579 | 1.0411 | 0.024843 | 23.8623 | -0.007041 | 0.14344 | 8.1e-05 | 0.006246 | 0.006246 | 0.94381
2019-03-21 16:09 | 1553198955.8796 | 1165 | 0.20949 | 0.000457 | -4e-06 | 101.234 | 101.4493 | 123.7994 | 21.2612 | 0.001943 | 0.001786 | -152.4414 | 0.24162 | 8.4013 | 0.003454 | 0.004729 | 0.004534 | 5e-05 | 0.0011 | 0.24458 | 1.2912 | 2.5481 | 98.84 | 20.89 | 38.284 | 68.5686 | 104.9864 | 64.7905 | 0.009579 | 1.0447 | 0.024876 | 23.8117 | -0.019003 | 0.14374 | -9.1e-05 | -0.006586 | -0.006586 | -0.89699
2019-03-21 16:09 | 1553198956.8796 | 1166 | 0.20948 | 0.000458 | -3e-06 | 101.223 | 101.4187 | 123.6947 | 21.2612 | 0.001932 | 0.001797 | -152.4362 | 0.24134 | 8.4044 | 0.003536 | 0.005018 | 0.004897 | 5.3e-05 | 0.001108 | 0.29319 | 1.2944 | 2.548 | 98.84 | 20.89 | 38.284 | 68.3135 | 106.9663 | 65.0645 | 0.009579 | 1.0439 | 0.024855 | 23.8113 | -0.017274 | 0.14357 | -2.6e-05 | -0.00176 | -0.00176 | -3.3525
2019-03-21 16:09 | 1553198957.8801 | 1167 | 0.20946 | 0.000461 | -3e-06 | 101.2337 | 101.3369 | 123.6332 | 21.2612 | 0.00192 | 0.001826 | -152.4156 | 0.24113 | 8.4019 | 0.003484 | 0.004419 | 0.004246 | 4.9e-05 | 0.001103 | 0.27256 | 1.3343 | 2.548 | 98.84 | 20.89 | 38.283 | 68.0226 | 105.6927 | 65.7756 | 0.009579 | 1.0417 | 0.024823 | 23.8304 | -0.009115 | 0.14345 | 4.2e-05 | 0.003337 | 0.003337 | 1.7669
2019-03-21 16:09 | 1553198958.8808 | 1168 | 0.20947 | 0.000457 | -4e-06 | 101.2796 | 101.4664 | 123.4025 | 21.2612 | 0.001918 | 0.001831 | -152.4215 | 0.24145 | 8.4017 | 0.003532 | 0.005129 | 0.004873 | 5e-05 | 0.001107 | 0.28313 | 1.312 | 2.5481 | 98.84 | 20.89 | 38.282 | 67.9781 | 106.8702 | 65.874 | 0.009578 | 1.0414 | 0.024796 | 23.8116 | -0.016478 | 0.14364 | 4.2e-05 | 0.00336 | 0.00336 | 1.757
2019-03-21 16:09 | 1553198959.8808 | 1169 | 0.20945 | 0.000462 | -4e-06 | 101.217 | 101.4258 | 123.4079 | 21.2612 | 0.001916 | 0.00184 | -152.4254 | 0.24103 | 8.3984 | 0.00357 | 0.004618 | 0.004581 | 4.3e-05 | 0.001106 | 0.26073 | 1.3182 | 2.5479 | 98.84 | 20.89 | 38.282 | 67.9274 | 107.7846 | 66.108 | 0.009578 | 1.0406 | 0.024788 | 23.8203 | -0.018432 | 0.14339 | 9.9e-05 | 0.00756 | 0.00756 | 0.77952
2019-03-21 16:09 | 1553198960.8812 | 1170 | 0.20947 | 0.000455 | -4e-06 | 101.2695 | 101.3049 | 123.7879 | 21.2612 | 0.001951 | 0.001785 | -152.4385 | 0.24053 | 8.4023 | 0.003567 | 0.005145 | 0.004922 | 4.6e-05 | 0.001105 | 0.26034 | 1.2945 | 2.548 | 98.84 | 20.89 | 38.282 | 68.7659 | 107.7105 | 64.7898 | 0.009578 | 1.0447 | 0.024875 | 23.8106 | -0.00313 | 0.14309 | 8e-06 | 0.000877 | 0.000877 | 6.7095
2019-03-21 16:09 | 1553198961.8803 | 1171 | 0.2095 | 0.000461 | -4e-06 | 101.2251 | 101.3471 | 123.9109 | 21.2612 | 0.001948 | 0.001851 | -152.442 | 0.24111 | 8.4031 | 0.003547 | 0.004579 | 0.004387 | 4.6e-05 | 0.001107 | 0.28241 | 1.3048 | 2.5481 | 98.84 | 20.89 | 38.282 | 68.695 | 107.2386 | 66.3709 | 0.009578 | 1.0398 | 0.024829 | 23.8781 | -0.01078 | 0.14343 | -0.00016 | -0.011721 | -0.011721 | -0.50296
2019-03-21 16:09 | 1553198962.8808 | 1172 | 0.20948 | 0.00046 | -4e-06 | 101.307 | 101.5334 | 124.2994 | 21.2612 | 0.001948 | 0.001877 | -152.4697 | 0.24144 | 8.4009 | 0.003563 | 0.005271 | 0.00511 | 4.9e-05 | 0.00111 | 0.26073 | 1.3319 | 2.5478 | 98.84 | 20.89 | 38.283 | 68.6887 | 107.6081 | 66.9889 | 0.009579 | 1.0379 | 0.024846 | 23.9372 | -0.01997 | 0.14363 | -5.6e-05 | -0.003922 | -0.003922 | -1.5052
2019-03-21 16:09 | 1553198963.8799 | 1173 | 0.20949 | 0.000459 | -4e-06 | 101.2707 | 101.4323 | 124.022 | 21.2612 | 0.001985 | 0.001799 | -152.4753 | 0.24205 | 8.4014 | 0.003525 | 0.004536 | 0.004284 | 4.6e-05 | 0.001107 | 0.27223 | 1.329 | 2.5478 | 98.84 | 20.89 | 38.286 | 69.5863 | 106.6945 | 65.1145 | 0.009579 | 1.0437 | 0.024886 | 23.8445 | -0.014264 | 0.14399 | -0.000137 | -0.009953 | -0.009953 | -0.59459
2019-03-21 16:09 | 1553198964.8814 | 1174 | 0.20949 | 0.00046 | -4e-06 | 101.2868 | 101.3359 | 123.7602 | 21.2612 | 0.001966 | 0.001891 | -152.4714 | 0.24185 | 8.4002 | 0.003547 | 0.005542 | 0.005349 | 4.6e-05 | 0.001108 | 0.28926 | 1.3518 | 2.5481 | 98.84 | 20.89 | 38.287 | 69.1332 | 107.2403 | 67.3275 | 0.00958 | 1.0369 | 0.024779 | 23.8971 | -0.00434 | 0.14388 | -9.4e-05 | -0.006785 | -0.006785 | -0.87148
2019-03-21 16:09 | 1553198965.8801 | 1175 | 0.20949 | 0.000456 | -4e-06 | 101.2424 | 101.5321 | 123.6271 | 21.2612 | 0.001971 | 0.001847 | -152.492 | 0.24127 | 8.4017 | 0.003491 | 0.004887 | 0.004754 | 4.5e-05 | 0.001109 | 0.26161 | 1.2867 | 2.548 | 98.84 | 20.89 | 38.287 | 69.2447 | 105.8784 | 66.2687 | 0.00958 | 1.0401 | 0.024805 | 23.8471 | -0.02556 | 0.14353 | -0.000111 | -0.007998 | -0.007998 | -0.73761
2019-03-21 16:09 | 1553198966.8812 | 1176 | 0.20947 | 0.000465 | -4e-06 | 101.2248 | 101.4854 | 123.337 | 21.2612 | 0.001982 | 0.00179 | -152.5131 | 0.24153 | 8.4023 | 0.003553 | 0.005116 | 0.004854 | 4.5e-05 | 0.001108 | 0.27489 | 1.3145 | 2.5479 | 98.84 | 20.89 | 38.287 | 69.5185 | 107.3656 | 64.892 | 0.00958 | 1.0444 | 0.024826 | 23.7708 | -0.023002 | 0.14369 | -1.5e-05 | -0.000866 | -0.000866 | -6.8231
2019-03-21 16:09 | 1553198967.8806 | 1177 | 0.20949 | 0.000461 | -3e-06 | 101.2757 | 101.2228 | 123.4473 | 21.2612 | 0.001945 | 0.00187 | -152.5226 | 0.24148 | 8.4023 | 0.00359 | 0.004688 | 0.004535 | 4.5e-05 | 0.001098 | 0.2417 | 1.2928 | 2.5479 | 98.84 | 20.89 | 38.287 | 68.6224 | 108.2784 | 66.8314 | 0.00958 | 1.0384 | 0.024766 | 23.8495 | 0.004673 | 0.14366 | -0.0001 | -0.007224 | -0.007224 | -0.81727
2019-03-21 16:09 | 1553198968.8807 | 1178 | 0.20948 | 0.000461 | -4e-06 | 101.2507 | 101.5488 | 123.6933 | 21.2612 | 0.001931 | 0.00182 | -152.4826 | 0.24173 | 8.4021 | 0.003561 | 0.004985 | 0.004785 | 4.7e-05 | 0.001098 | 0.28622 | 1.3099 | 2.548 | 98.84 | 20.89 | 38.289 | 68.2892 | 107.5631 | 65.6132 | 0.00958 | 1.0422 | 0.024835 | 23.8305 | -0.026293 | 0.1438 | -7.9e-05 | -0.005662 | -0.005662 | -1.0439
2019-03-21 16:09 | 1553198969.8803 | 1179 | 0.20947 | 0.000457 | -3e-06 | 101.2656 | 101.5379 | 123.7194 | 21.2612 | 0.0019 | 0.001854 | -152.4881 | 0.24157 | 8.4011 | 0.003561 | 0.004377 | 0.004208 | 4.6e-05 | 0.001098 | 0.26432 | 1.3082 | 2.548 | 98.84 | 20.89 | 38.289 | 67.5506 | 107.5607 | 66.4442 | 0.00958 | 1.0396 | 0.024807 | 23.8622 | -0.024017 | 0.14371 | 2.6e-05 | 0.002162 | 0.002162 | 2.7323
2019-03-21 16:10 | 1553198970.8798 | 1180 | 0.20946 | 0.00046 | -4e-06 | 101.2453 | 101.4153 | 123.3389 | 21.2612 | 0.001962 | 0.001891 | -152.4735 | 0.24151 | 8.4023 | 0.003569 | 0.004959 | 0.004779 | 4.6e-05 | 0.0011 | 0.25122 | 1.321 | 2.548 | 98.84 | 20.89 | 38.289 | 69.0428 | 107.76 | 67.3261 | 0.00958 | 1.0369 | 0.024737 | 23.8564 | -0.015 | 0.14367 | 7.6e-05 | 0.005897 | 0.005897 | 1.0013
2019-03-21 16:10 | 1553198971.8809 | 1181 | 0.20948 | 0.000462 | -4e-06 | 101.2912 | 101.4125 | 123.2582 | 21.2612 | 0.001932 | 0.001811 | -152.4816 | 0.24145 | 8.4016 | 0.003559 | 0.004797 | 0.004558 | 4.3e-05 | 0.001102 | 0.28567 | 1.3157 | 2.5479 | 98.84 | 20.89 | 38.29 | 68.3027 | 107.526 | 65.4025 | 0.00958 | 1.0428 | 0.024799 | 23.7811 | -0.010707 | 0.14364 | -4e-05 | -0.002757 | -0.002757 | -2.1413
2019-03-21 16:10 | 1553198972.8795 | 1182 | 0.20946 | 0.000464 | -4e-06 | 101.2963 | 101.4575 | 123.3866 | 21.2612 | 0.001913 | 0.001885 | -152.473 | 0.24068 | 8.403 | 0.003522 | 0.004979 | 0.004802 | 4.4e-05 | 0.001101 | 0.26034 | 1.3198 | 2.5479 | 98.84 | 20.89 | 38.29 | 67.8634 | 106.6152 | 67.1801 | 0.00958 | 1.0374 | 0.024747 | 23.8559 | -0.014224 | 0.14318 | 7.1e-05 | 0.00551 | 0.00551 | 1.068
2019-03-21 16:10 | 1553198973.8809 | 1183 | 0.20948 | 0.00046 | -4e-06 | 101.2292 | 101.3349 | 123.3711 | 21.2612 | 0.00192 | 0.001829 | -152.463 | 0.24144 | 8.4023 | 0.003498 | 0.004972 | 0.004779 | 4.5e-05 | 0.001102 | 0.26913 | 1.3443 | 2.5481 | 98.84 | 20.88 | 38.289 | 68.0178 | 106.0444 | 65.8427 | 0.009574 | 1.0415 | 0.024794 | 23.8075 | -0.009329 | 0.14363 | -4.7e-05 | -0.003305 | -0.003305 | -1.7861
2019-03-21 16:10 | 1553198974.88 | 1184 | 0.20946 | 0.000461 | -4e-06 | 101.2623 | 101.2443 | 123.5213 | 21.2612 | 0.001951 | 0.001808 | -152.454 | 0.24091 | 8.4037 | 0.003487 | 0.004958 | 0.004801 | 4.7e-05 | 0.0011 | 0.24226 | 1.3305 | 2.5481 | 98.84 | 20.88 | 38.287 | 68.758 | 105.7776 | 65.3416 | 0.009574 | 1.043 | 0.024828 | 23.8043 | 0.001593 | 0.14332 | 8.1e-05 | 0.006263 | 0.006263 | 0.94046
2019-03-21 16:10 | 1553198975.8795 | 1185 | 0.20945 | 0.00046 | -4e-06 | 101.2677 | 101.5093 | 123.5619 | 21.2612 | 0.001881 | 0.001792 | -152.4482 | 0.24135 | 8.4005 | 0.003448 | 0.0053 | 0.005062 | 4.4e-05 | 0.001106 | 0.26985 | 1.3605 | 2.5481 | 98.84 | 20.88 | 38.286 | 67.0779 | 104.821 | 64.9466 | 0.009573 | 1.0442 | 0.024846 | 23.7944 | -0.021313 | 0.14358 | 0.000136 | 0.010363 | 0.010363 | 0.56943
2019-03-21 16:10 | 1553198976.88 | 1186 | 0.20948 | 0.000457 | -4e-06 | 101.2757 | 101.464 | 123.8832 | 21.2612 | 0.001985 | 0.001844 | -152.4516 | 0.24125 | 8.4037 | 0.003453 | 0.005548 | 0.005425 | 4.6e-05 | 0.001104 | 0.2521 | 1.3152 | 2.5479 | 98.84 | 20.88 | 38.285 | 69.5786 | 104.9602 | 66.204 | 0.009573 | 1.0403 | 0.024833 | 23.8695 | -0.01661 | 0.14352 | -4.1e-05 | -0.00284 | -0.00284 | -2.0768
2019-03-21 16:10 | 1553198977.881 | 1187 | 0.20947 | 0.000457 | -4e-06 | 101.2522 | 101.4228 | 123.7668 | 21.2612 | 0.001904 | 0.001831 | -152.4693 | 0.24144 | 8.4034 | 0.003456 | 0.00473 | 0.004606 | 4.5e-05 | 0.001104 | 0.27527 | 1.3279 | 2.548 | 98.84 | 20.88 | 38.286 | 67.6285 | 105.0281 | 65.8911 | 0.009573 | 1.0413 | 0.024832 | 23.8473 | -0.015055 | 0.14363 | 1.3e-05 | 0.001163 | 0.001163 | 5.0768
2019-03-21 16:10 | 1553198978.8808 | 1188 | 0.20946 | 0.000459 | -4e-06 | 101.2054 | 101.3036 | 123.8048 | 21.2612 | 0.001957 | 0.001824 | -152.4729 | 0.24196 | 8.4028 | 0.00346 | 0.005171 | 0.005013 | 4.4e-05 | 0.001102 | 0.24994 | 1.3156 | 2.5479 | 98.84 | 20.88 | 38.285 | 68.9157 | 105.1325 | 65.7162 | 0.009573 | 1.0418 | 0.024843 | 23.8448 | -0.008669 | 0.14394 | 5.9e-05 | 0.004684 | 0.004684 | 1.2632
2019-03-21 16:10 | 1553198979.8798 | 1189 | 0.20947 | 0.000458 | -4e-06 | 101.2751 | 101.3557 | 123.7436 | 21.2612 | 0.001907 | 0.001823 | -152.4817 | 0.24124 | 8.4029 | 0.003328 | 0.004914 | 0.004711 | 4.5e-05 | 0.001108 | 0.26272 | 1.3033 | 2.5479 | 98.84 | 20.88 | 38.282 | 67.7192 | 101.9291 | 65.7006 | 0.009572 | 1.0419 | 0.024837 | 23.8384 | -0.007109 | 0.14351 | 1.9e-05 | 0.00164 | 0.00164 | 3.596
2019-03-21 16:10 | 1553198980.8803 | 1190 | 0.2095 | 0.000461 | -4e-06 | 101.3332 | 101.4987 | 124.156 | 21.2612 | 0.001942 | 0.001819 | -152.48 | 0.2412 | 8.4004 | 0.003369 | 0.004166 | 0.004055 | 4.6e-05 | 0.001107 | 0.27793 | 1.3179 | 2.548 | 98.84 | 20.88 | 38.281 | 68.5435 | 102.9126 | 65.5945 | 0.009572 | 1.0422 | 0.024882 | 23.8743 | -0.014594 | 0.14349 | -0.00018 | -0.01316 | -0.01316 | -0.44814
2019-03-21 16:10 | 1553198981.8796 | 1191 | 0.2095 | 0.000461 | -4e-06 | 101.2793 | 101.3747 | 124.3337 | 21.2612 | 0.001916 | 0.001842 | -152.4727 | 0.24079 | 8.4025 | 0.003351 | 0.004275 | 0.004146 | 4.7e-05 | 0.001109 | 0.28119 | 1.3312 | 2.548 | 98.84 | 20.88 | 38.279 | 67.9274 | 102.4694 | 66.1525 | 0.009572 | 1.0405 | 0.02488 | 23.9111 | -0.008424 | 0.14324 | -0.000174 | -0.012754 | -0.012754 | -0.46161
2019-03-21 16:10 | 1553198982.8801 | 1192 | 0.20952 | 0.000457 | -3e-06 | 101.2337 | 101.4486 | 124.2309 | 21.2612 | 0.001949 | 0.001815 | -152.4686 | 0.24166 | 8.4033 | 0.003409 | 0.00475 | 0.004573 | 4.3e-05 | 0.001107 | 0.30055 | 1.3197 | 2.5479 | 98.84 | 20.88 | 38.278 | 68.727 | 103.8885 | 65.5062 | 0.009571 | 1.0425 | 0.024893 | 23.8784 | -0.01897 | 0.14377 | -0.000314 | -0.023237 | -0.023237 | -0.25428
2019-03-21 16:10 | 1553198983.8804 | 1193 | 0.20953 | 0.000458 | -4e-06 | 101.1938 | 101.2644 | 124.4114 | 21.2612 | 0.00194 | 0.001832 | -152.4568 | 0.24185 | 8.404 | 0.003395 | 0.004751 | 0.004642 | 4.8e-05 | 0.001107 | 0.2912 | 1.3114 | 2.5482 | 98.84 | 20.88 | 38.278 | 68.496 | 103.5544 | 65.9091 | 0.009571 | 1.0413 | 0.024896 | 23.91 | -0.006237 | 0.14388 | -0.000362 | -0.026778 | -0.026778 | -0.22083
2019-03-21 16:10 | 1553198984.8791 | 1194 | 0.2095 | 0.000463 | -4e-06 | 101.2528 | 101.558 | 124.5313 | 21.2612 | 0.001914 | 0.001813 | -152.4681 | 0.2417 | 8.4034 | 0.003403 | 0.005031 | 0.004794 | 4.3e-05 | 0.00111 | 0.25321 | 1.288 | 2.5479 | 98.84 | 20.88 | 38.275 | 67.8743 | 103.7281 | 65.454 | 0.009571 | 1.0427 | 0.024925 | 23.9054 | -0.026919 | 0.14378 | -0.000155 | -0.01137 | -0.01137 | -0.51975
2019-03-21 16:10 | 1553198985.8802 | 1195 | 0.2095 | 0.000464 | -4e-06 | 101.212 | 101.3376 | 124.5017 | 21.2612 | 0.001965 | 0.001749 | -152.4694 | 0.24137 | 8.4021 | 0.003457 | 0.004536 | 0.00433 | 4.5e-05 | 0.001105 | 0.25625 | 1.2687 | 2.548 | 98.84 | 20.88 | 38.274 | 69.1058 | 105.039 | 63.917 | 0.00957 | 1.0474 | 0.024979 | 23.8483 | -0.011097 | 0.14359 | -0.000169 | -0.012389 | -0.012389 | -0.47636
2019-03-21 16:10 | 1553198986.8796 | 1196 | 0.20951 | 0.000459 | -4e-06 | 101.2546 | 101.6166 | 124.6235 | 21.2612 | 0.0019 | 0.001793 | -152.4558 | 0.24128 | 8.4017 | 0.003371 | 0.004791 | 0.004749 | 4.9e-05 | 0.001109 | 0.28335 | 1.3027 | 2.5479 | 98.84 | 20.88 | 38.273 | 67.5326 | 102.9729 | 64.9639 | 0.00957 | 1.0442 | 0.024952 | 23.897 | -0.031919 | 0.14354 | -0.000214 | -0.015782 | -0.015782 | -0.3738
2019-03-21 16:10 | 1553198987.8803 | 1197 | 0.20949 | 0.000458 | -3e-06 | 101.2388 | 101.2773 | 124.7883 | 21.2612 | 0.001912 | 0.001877 | -152.4502 | 0.24052 | 8.4019 | 0.003414 | 0.004308 | 0.00407 | 4.5e-05 | 0.00111 | 0.3011 | 1.3057 | 2.5478 | 98.84 | 20.88 | 38.272 | 67.8346 | 104.0104 | 66.9983 | 0.00957 | 1.0379 | 0.024894 | 23.9846 | -0.003405 | 0.14308 | -0.000138 | -0.010138 | -0.010138 | -0.58009
2019-03-21 16:10 | 1553198988.8802 | 1198 | 0.20947 | 0.000455 | -4e-06 | 101.1998 | 101.4367 | 124.8709 | 21.2612 | 0.001929 | 0.001892 | -152.4695 | 0.24169 | 8.401 | 0.003421 | 0.004779 | 0.004506 | 4.7e-05 | 0.001112 | 0.2636 | 1.309 | 2.5479 | 98.84 | 20.88 | 38.272 | 68.251 | 104.1802 | 67.3508 | 0.00957 | 1.0368 | 0.024889 | 24.005 | -0.02092 | 0.14378 | 1.6e-05 | 0.001477 | 0.001477 | 4.0011
2019-03-21 16:10 | 1553198989.8803 | 1199 | 0.20947 | 0.000465 | -4e-06 | 101.2462 | 101.2763 | 125.2038 | 21.2612 | 0.001894 | 0.001791 | -152.458 | 0.24084 | 8.4035 | 0.003474 | 0.004599 | 0.004474 | 4.6e-05 | 0.001108 | 0.26178 | 1.3149 | 2.5479 | 98.84 | 20.88 | 38.271 | 67.404 | 105.4511 | 64.931 | 0.00957 | 1.0443 | 0.025012 | 23.9514 | -0.002656 | 0.14327 | -7e-06 | -0.000279 | -0.000279 | -21.1376
2019-03-21 16:10 | 1553198990.8804 | 1200 | 0.20947 | 0.000459 | -4e-06 | 101.2144 | 101.4643 | 124.97 | 21.2612 | 0.001936 | 0.001844 | -152.476 | 0.24163 | 8.4033 | 0.003548 | 0.004768 | 0.004662 | 4.5e-05 | 0.001111 | 0.26897 | 1.3113 | 2.5481 | 98.84 | 20.88 | 38.271 | 68.4142 | 107.2498 | 66.1922 | 0.00957 | 1.0404 | 0.024942 | 23.9736 | -0.022062 | 0.14375 | 1.1e-05 | 0.001027 | 0.001027 | 5.751
2019-03-21 16:10 | 1553198991.8805 | 1201 | 0.20945 | 0.000463 | -4e-06 | 101.2403 | 101.276 | 125.0709 | 21.2612 | 0.001921 | 0.001819 | -152.4663 | 0.24149 | 8.4022 | 0.003551 | 0.004774 | 0.004541 | 4.8e-05 | 0.001113 | 0.26952 | 1.3663 | 2.548 | 98.84 | 20.88 | 38.271 | 68.0483 | 107.3301 | 65.607 | 0.00957 | 1.0422 | 0.024973 | 23.9626 | -0.003153 | 0.14366 | 0.000113 | 0.008704 | 0.008704 | 0.67839
2019-03-21 16:10 | 1553198992.8796 | 1202 | 0.20949 | 0.000457 | -3e-06 | 101.2546 | 101.4875 | 125.0808 | 21.2612 | 0.001987 | 0.001822 | -152.4572 | 0.24076 | 8.4048 | 0.003553 | 0.005157 | 0.004999 | 4.9e-05 | 0.001109 | 0.28711 | 1.3288 | 2.548 | 98.84 | 20.88 | 38.271 | 69.6331 | 107.3811 | 65.6678 | 0.00957 | 1.042 | 0.024972 | 23.9657 | -0.02055 | 0.14323 | -0.000106 | -0.007746 | -0.007746 | -0.76001
2019-03-21 16:10 | 1553198993.8806 | 1203 | 0.20945 | 0.000462 | -4e-06 | 101.1279 | 101.6254 | 125.1364 | 21.2612 | 0.001902 | 0.001805 | -152.4703 | 0.2408 | 8.4037 | 0.00348 | 0.004989 | 0.004833 | 4.7e-05 | 0.001107 | 0.2557 | 1.3664 | 2.5479 | 98.84 | 20.88 | 38.271 | 67.5865 | 105.6163 | 65.2613 | 0.00957 | 1.0432 | 0.024993 | 23.9566 | -0.043891 | 0.14325 | 0.000103 | 0.007946 | 0.007946 | 0.741
2019-03-21 16:10 | 1553198994.8807 | 1204 | 0.20947 | 0.000463 | -4e-06 | 101.1586 | 101.6315 | 125.0841 | 21.2612 | 0.001917 | 0.001799 | -152.4746 | 0.24097 | 8.4021 | 0.003497 | 0.004658 | 0.004505 | 4.7e-05 | 0.001113 | 0.2891 | 1.3354 | 2.548 | 98.84 | 20.88 | 38.269 | 67.9437 | 106.0106 | 65.1184 | 0.009569 | 1.0437 | 0.024993 | 23.9466 | -0.041717 | 0.14335 | 2.1e-05 | 0.001823 | 0.001823 | 3.2313
2019-03-21 16:10 | 1553198995.8797 | 1205 | 0.20946 | 0.000456 | -4e-06 | 101.17 | 101.4977 | 125.0195 | 21.2612 | 0.001927 | 0.001877 | -152.4787 | 0.24112 | 8.4018 | 0.00342 | 0.004845 | 0.004709 | 4.4e-05 | 0.001107 | 0.2943 | 1.3706 | 2.5479 | 98.84 | 20.88 | 38.269 | 68.1988 | 104.1613 | 66.9921 | 0.009569 | 1.0379 | 0.024917 | 24.0066 | -0.028928 | 0.14344 | 6.3e-05 | 0.004974 | 0.004974 | 1.1852
2019-03-21 16:10 | 1553198996.8798 | 1206 | 0.20946 | 0.00046 | -4e-06 | 101.2433 | 101.5358 | 125.0585 | 21.2612 | 0.001979 | 0.001847 | -152.4906 | 0.24028 | 8.4053 | 0.003516 | 0.004781 | 0.004669 | 4.7e-05 | 0.001114 | 0.2896 | 1.3595 | 2.5479 | 98.84 | 20.88 | 38.266 | 69.4498 | 106.482 | 66.2812 | 0.009568 | 1.0401 | 0.024947 | 23.9852 | -0.025811 | 0.14294 | 5.3e-05 | 0.004182 | 0.004182 | 1.4049
2019-03-21 16:10 | 1553198997.8812 | 1207 | 0.20946 | 0.000464 | -4e-06 | 101.1378 | 101.3042 | 125.1047 | 21.2612 | 0.001891 | 0.001851 | -152.4913 | 0.24052 | 8.4035 | 0.003428 | 0.005135 | 0.004939 | 4.4e-05 | 0.001113 | 0.26698 | 1.3617 | 2.5481 | 98.84 | 20.88 | 38.264 | 67.3362 | 104.3547 | 66.3639 | 0.009568 | 1.0399 | 0.024949 | 23.9926 | -0.014709 | 0.14308 | 3.8e-05 | 0.003082 | 0.003082 | 1.9079
2019-03-21 16:10 | 1553198998.8799 | 1208 | 0.20948 | 0.000459 | -4e-06 | 101.1843 | 101.4377 | 125.096 | 21.2612 | 0.001941 | 0.001827 | -152.4887 | 0.24099 | 8.4018 | 0.003547 | 0.004732 | 0.004648 | 4.3e-05 | 0.001115 | 0.27472 | 1.2946 | 2.5481 | 98.84 | 20.88 | 38.261 | 68.5359 | 107.2333 | 65.7989 | 0.009567 | 1.0416 | 0.024969 | 23.9718 | -0.02238 | 0.14336 | -3.2e-05 | -0.002165 | -0.002165 | -2.7216
2019-03-21 16:10 | 1553198999.8789 | 1209 | 0.20946 | 0.000459 | -4e-06 | 101.1396 | 101.4698 | 124.8836 | 21.2612 | 0.001972 | 0.001831 | -152.4997 | 0.2406 | 8.4042 | 0.003479 | 0.005033 | 0.004826 | 4.6e-05 | 0.001116 | 0.27721 | 1.3257 | 2.5479 | 98.84 | 20.88 | 38.256 | 69.268 | 105.5787 | 65.8786 | 0.009566 | 1.0413 | 0.024945 | 23.9542 | -0.029154 | 0.14313 | 4.9e-05 | 0.003915 | 0.003915 | 1.5028
2019-03-21 16:10 | 1553199000.8804 | 1210 | 0.20947 | 0.000461 | -3e-06 | 101.1679 | 101.4446 | 124.6707 | 21.2612 | 0.001996 | 0.001813 | -152.5067 | 0.24111 | 8.4012 | 0.003442 | 0.004713 | 0.004563 | 4.3e-05 | 0.001115 | 0.25409 | 1.3188 | 2.548 | 98.84 | 20.88 | 38.252 | 69.8545 | 104.6749 | 65.4564 | 0.009565 | 1.0426 | 0.024939 | 23.9189 | -0.024429 | 0.14344 | 2e-05 | 0.001701 | 0.001701 | 3.4658
2019-03-21 16:10 | 1553199001.88 | 1211 | 0.20944 | 0.000462 | -3e-06 | 101.1062 | 101.1455 | 124.2886 | 21.2612 | 0.001922 | 0.001858 | -152.4981 | 0.24138 | 8.4016 | 0.003543 | 0.004755 | 0.004566 | 4.3e-05 | 0.001117 | 0.29336 | 1.2948 | 2.548 | 98.84 | 20.88 | 38.249 | 68.0685 | 107.1403 | 66.5356 | 0.009564 | 1.0393 | 0.024861 | 23.9202 | -0.003478 | 0.1436 | 0.000168 | 0.012738 | 0.012738 | 0.46331
2019-03-21 16:10 | 1553199002.8791 | 1212 | 0.20949 | 0.00046 | -4e-06 | 101.1732 | 101.2773 | 124.014 | 21.2612 | 0.001978 | 0.00181 | -152.5115 | 0.24016 | 8.4019 | 0.003555 | 0.004703 | 0.00451 | 4.6e-05 | 0.00112 | 0.3027 | 1.2944 | 2.548 | 98.84 | 20.87 | 38.244 | 69.4108 | 107.434 | 65.3736 | 0.009557 | 1.0429 | 0.024876 | 23.8529 | -0.009198 | 0.14287 | -0.000118 | -0.008601 | -0.008601 | -0.68275
2019-03-21 16:10 | 1553199003.8807 | 1213 | 0.20948 | 0.000462 | -4e-06 | 101.1363 | 101.2896 | 123.8422 | 21.2612 | 0.001953 | 0.001811 | -152.5266 | 0.24107 | 8.4045 | 0.003548 | 0.004559 | 0.004458 | 4.6e-05 | 0.00112 | 0.27616 | 1.3077 | 2.5479 | 98.84 | 20.87 | 38.241 | 68.8097 | 107.2435 | 65.4032 | 0.009556 | 1.0428 | 0.024858 | 23.8374 | -0.013547 | 0.14341 | -5.2e-05 | -0.003659 | -0.003659 | -1.611
2019-03-21 16:10 | 1553199004.8806 | 1214 | 0.2095 | 0.000466 | -4e-06 | 101.0302 | 101.3267 | 123.7048 | 21.2612 | 0.001939 | 0.001856 | -152.5329 | 0.24065 | 8.4039 | 0.003588 | 0.005213 | 0.004998 | 4.6e-05 | 0.001118 | 0.29999 | 1.3067 | 2.548 | 98.84 | 20.87 | 38.237 | 68.482 | 108.2216 | 66.481 | 0.009555 | 1.0395 | 0.024805 | 23.8621 | -0.026211 | 0.14316 | -0.000159 | -0.011621 | -0.011621 | -0.50631
2019-03-21 16:10 | 1553199005.8798 | 1215 | 0.20947 | 0.00046 | -4e-06 | 101.1726 | 101.4916 | 123.6548 | 21.2612 | 0.001914 | 0.001845 | -152.5407 | 0.24111 | 8.4032 | 0.003591 | 0.005257 | 0.005101 | 4.5e-05 | 0.001113 | 0.25752 | 1.3144 | 2.5479 | 98.84 | 20.87 | 38.233 | 67.8798 | 108.2918 | 66.229 | 0.009554 | 1.0403 | 0.024809 | 23.8484 | -0.028151 | 0.14344 | -1.2e-05 | -0.000661 | -0.000661 | -8.9184
2019-03-21 16:10 | 1553199006.8793 | 1216 | 0.20949 | 0.000462 | -4e-06 | 101.1271 | 101.3277 | 123.3565 | 21.2612 | 0.001967 | 0.001843 | -152.5553 | 0.24047 | 8.4064 | 0.003684 | 0.004959 | 0.004849 | 4.7e-05 | 0.001114 | 0.26985 | 1.333 | 2.548 | 98.84 | 20.87 | 38.23 | 69.1457 | 110.5498 | 66.1782 | 0.009554 | 1.0404 | 0.024781 | 23.8178 | -0.017732 | 0.14305 | -8.8e-05 | -0.006285 | -0.006285 | -0.93548
2019-03-21 16:10 | 1553199007.8804 | 1217 | 0.20949 | 0.000463 | -4e-06 | 101.0967 | 101.1173 | 122.9967 | 21.2612 | 0.001931 | 0.001801 | -152.5295 | 0.24028 | 8.4035 | 0.003609 | 0.004937 | 0.004771 | 4.8e-05 | 0.001116 | 0.27312 | 1.2998 | 2.548 | 98.84 | 20.87 | 38.225 | 68.2784 | 108.7419 | 65.1511 | 0.009552 | 1.0436 | 0.024782 | 23.7471 | -0.001821 | 0.14294 | -0.000122 | -0.008844 | -0.008844 | -0.66429
2019-03-21 16:10 | 1553199008.8804 | 1218 | 0.20949 | 0.000461 | -4e-06 | 101.0368 | 101.1394 | 122.8746 | 21.2612 | 0.001934 | 0.001821 | -152.5367 | 0.24083 | 8.4035 | 0.003571 | 0.00461 | 0.004454 | 5.1e-05 | 0.001112 | 0.28279 | 1.2937 | 2.5479 | 98.84 | 20.87 | 38.224 | 68.3594 | 107.812 | 65.632 | 0.009552 | 1.0421 | 0.024752 | 23.7521 | -0.009079 | 0.14327 | -0.000139 | -0.010034 | -0.010034 | -0.58683
2019-03-21 16:10 | 1553199009.8798 | 1219 | 0.20948 | 0.000462 | -4e-06 | 101.111 | 101.2405 | 122.8927 | 21.2612 | 0.001964 | 0.00187 | -152.514 | 0.24169 | 8.4006 | 0.003555 | 0.004671 | 0.004522 | 4.9e-05 | 0.001111 | 0.26842 | 1.3143 | 2.5479 | 98.83 | 20.87 | 38.222 | 69.0779 | 107.4347 | 66.829 | 0.009553 | 1.0384 | 0.02471 | 23.7958 | -0.011455 | 0.14378 | -6.6e-05 | -0.004674 | -0.004674 | -1.2644
2019-03-21 16:10 | 1553199010.8795 | 1220 | 0.20947 | 0.00046 | -4e-06 | 101.1455 | 101.0096 | 122.991 | 21.2612 | 0.001892 | 0.001821 | -152.5015 | 0.24129 | 8.3997 | 0.003576 | 0.004483 | 0.004282 | 4.6e-05 | 0.001117 | 0.29662 | 1.3086 | 2.5479 | 98.83 | 20.87 | 38.222 | 67.3588 | 107.9271 | 65.6539 | 0.009553 | 1.042 | 0.024763 | 23.7641 | 0.012026 | 0.14355 | -6e-06 | -0.000164 | -0.000164 | -35.9795
2019-03-21 16:10 | 1553199011.8795 | 1221 | 0.20948 | 0.000463 | -3e-06 | 101.1172 | 101.2702 | 122.9525 | 21.2612 | 0.00197 | 0.001875 | -152.5155 | 0.24094 | 8.3999 | 0.003568 | 0.004391 | 0.004193 | 4.8e-05 | 0.001114 | 0.27007 | 1.3253 | 2.5481 | 98.83 | 20.87 | 38.222 | 69.2236 | 107.7365 | 66.9319 | 0.009553 | 1.0381 | 0.024713 | 23.8052 | -0.01352 | 0.14333 | -4.5e-05 | -0.00313 | -0.00313 | -1.8821
2019-03-21 16:10 | 1553199012.88 | 1222 | 0.20949 | 0.000462 | -4e-06 | 101.1399 | 101.325 | 122.7382 | 21.2612 | 0.001955 | 0.001842 | -152.5137 | 0.24163 | 8.4038 | 0.003541 | 0.004815 | 0.004674 | 4.7e-05 | 0.001115 | 0.28672 | 1.3169 | 2.548 | 98.83 | 20.87 | 38.222 | 68.8649 | 107.0976 | 66.1571 | 0.009553 | 1.0405 | 0.024719 | 23.7573 | -0.016358 | 0.14375 | -0.000129 | -0.009363 | -0.009363 | -0.63102
2019-03-21 16:10 | 1553199013.8799 | 1223 | 0.20946 | 0.000457 | -3e-06 | 101.1583 | 101.3574 | 122.6131 | 21.2612 | 0.001956 | 0.001845 | -152.5105 | 0.24187 | 8.4014 | 0.003578 | 0.004869 | 0.004653 | 5.1e-05 | 0.001116 | 0.26842 | 1.3292 | 2.5478 | 98.83 | 20.87 | 38.221 | 68.8806 | 107.9885 | 66.2273 | 0.009552 | 1.0403 | 0.024704 | 23.7477 | -0.017581 | 0.14389 | 7e-05 | 0.005392 | 0.005392 | 1.0969
2019-03-21 16:10 | 1553199014.8787 | 1224 | 0.20946 | 0.00046 | -4e-06 | 101.1032 | 101.186 | 122.4697 | 21.2612 | 0.00195 | 0.001838 | -152.5045 | 0.24174 | 8.4034 | 0.003633 | 0.005053 | 0.004841 | 4.6e-05 | 0.001117 | 0.28727 | 1.2993 | 2.5481 | 98.83 | 20.87 | 38.22 | 68.7541 | 109.3101 | 66.0496 | 0.009552 | 1.0408 | 0.024696 | 23.7276 | -0.007325 | 0.14381 | 7.5e-05 | 0.005797 | 0.005797 | 1.0196
2019-03-21 16:10 | 1553199015.8798 | 1225 | 0.20948 | 0.000464 | -4e-06 | 101.172 | 101.2845 | 122.5448 | 21.2612 | 0.001949 | 0.001843 | -152.4936 | 0.24118 | 8.4026 | 0.003653 | 0.004851 | 0.004691 | 4.8e-05 | 0.001113 | 0.2808 | 1.3328 | 2.548 | 98.83 | 20.87 | 38.218 | 68.727 | 109.8013 | 66.1727 | 0.009552 | 1.0404 | 0.024699 | 23.7392 | -0.009935 | 0.14348 | -3.7e-05 | -0.002473 | -0.002473 | -2.3841
2019-03-21 16:10 | 1553199016.8802 | 1226 | 0.20944 | 0.000461 | -4e-06 | 101.1476 | 101.201 | 122.6115 | 21.2612 | 0.001949 | 0.001865 | -152.4695 | 0.24097 | 8.4024 | 0.003601 | 0.005023 | 0.004855 | 4.8e-05 | 0.001116 | 0.26034 | 1.31 | 2.5479 | 98.83 | 20.87 | 38.216 | 68.7128 | 108.5424 | 66.6994 | 0.009551 | 1.0388 | 0.024687 | 23.764 | -0.004723 | 0.14335 | 0.000173 | 0.01305 | 0.01305 | 0.45146
2019-03-21 16:10 | 1553199017.8803 | 1227 | 0.20948 | 0.00046 | -3e-06 | 101.0779 | 100.98 | 122.4254 | 21.2612 | 0.001957 | 0.001815 | -152.4869 | 0.24144 | 8.4018 | 0.003654 | 0.005083 | 0.004903 | 4.7e-05 | 0.001116 | 0.28224 | 1.3049 | 2.5481 | 98.83 | 20.87 | 38.216 | 68.9181 | 109.829 | 65.5023 | 0.009551 | 1.0425 | 0.024712 | 23.7041 | 0.008668 | 0.14363 | -7.1e-05 | -0.005053 | -0.005053 | -1.1684
2019-03-21 16:10 | 1553199018.88 | 1228 | 0.20945 | 0.000464 | -4e-06 | 101.1133 | 101.3819 | 122.3822 | 21.2612 | 0.001947 | 0.001788 | -152.4878 | 0.2415 | 8.4032 | 0.003621 | 0.005146 | 0.005008 | 5e-05 | 0.001116 | 0.2724 | 1.3464 | 2.5481 | 98.83 | 20.87 | 38.216 | 68.6661 | 109.031 | 64.8468 | 0.009551 | 1.0445 | 0.024731 | 23.677 | -0.023724 | 0.14367 | 0.000123 | 0.009344 | 0.009344 | 0.63194
2019-03-21 16:10 | 1553199019.8793 | 1229 | 0.20947 | 0.000463 | -4e-06 | 101.0174 | 101.326 | 122.1568 | 21.2612 | 0.001922 | 0.001788 | -152.5122 | 0.24156 | 8.4025 | 0.003645 | 0.00529 | 0.005092 | 4.6e-05 | 0.001119 | 0.29822 | 1.3213 | 2.5479 | 98.83 | 20.87 | 38.215 | 68.0827 | 109.6098 | 64.8436 | 0.009551 | 1.0445 | 0.024709 | 23.6551 | -0.027285 | 0.1437 | 2.6e-05 | 0.002207 | 0.002207 | 2.676
2019-03-21 16:10 | 1553199020.8794 | 1230 | 0.20948 | 0.000462 | -4e-06 | 101.1062 | 101.2903 | 122.3226 | 21.2612 | 0.00194 | 0.001844 | -152.5116 | 0.24046 | 8.4029 | 0.003693 | 0.004982 | 0.004868 | 4.7e-05 | 0.001117 | 0.26073 | 1.3141 | 2.5478 | 98.83 | 20.87 | 38.213 | 68.5029 | 110.7803 | 66.2054 | 0.00955 | 1.0403 | 0.024676 | 23.7188 | -0.016269 | 0.14305 | -3.8e-05 | -0.002575 | -0.002575 | -2.283
2019-03-21 16:10 | 1553199021.8796 | 1231 | 0.20946 | 0.000461 | -4e-06 | 101.1002 | 101.2293 | 122.735 | 21.2612 | 0.001919 | 0.001878 | -152.5246 | 0.2414 | 8.4033 | 0.003692 | 0.005325 | 0.005077 | 4.7e-05 | 0.00112 | 0.25912 | 1.3227 | 2.5479 | 98.83 | 20.87 | 38.212 | 68.0101 | 110.7438 | 67.0233 | 0.00955 | 1.0378 | 0.024687 | 23.7873 | -0.011411 | 0.14361 | 4e-05 | 0.003161 | 0.003161 | 1.867
2019-03-21 16:10 | 1553199022.8802 | 1232 | 0.20946 | 0.000461 | -3e-06 | 101.1399 | 101.2405 | 123.4344 | 21.2612 | 0.00193 | 0.001799 | -152.5166 | 0.24064 | 8.4012 | 0.003709 | 0.004929 | 0.004741 | 4.8e-05 | 0.001117 | 0.24729 | 1.3187 | 2.548 | 98.83 | 20.87 | 38.209 | 68.2738 | 111.1678 | 65.1177 | 0.009549 | 1.0437 | 0.024827 | 23.7881 | -0.008899 | 0.14316 | 5.1e-05 | 0.003969 | 0.003969 | 1.4826
2019-03-21 16:10 | 1553199023.8796 | 1233 | 0.2095 | 0.000463 | -4e-06 | 101.1047 | 101.2518 | 123.8652 | 21.2612 | 0.001996 | 0.001854 | -152.5178 | 0.24069 | 8.3988 | 0.003674 | 0.004815 | 0.004646 | 4.9e-05 | 0.001112 | 0.25929 | 1.3032 | 2.548 | 98.83 | 20.87 | 38.208 | 69.8552 | 110.3069 | 66.4303 | 0.009549 | 1.0397 | 0.024822 | 23.8758 | -0.013002 | 0.14319 | -0.000156 | -0.011405 | -0.011405 | -0.51598
2019-03-21 16:10 | 1553199024.8802 | 1234 | 0.20948 | 0.000463 | -4e-06 | 101.1542 | 101.3451 | 124.0036 | 21.2612 | 0.001905 | 0.00184 | -152.5004 | 0.24099 | 8.4013 | 0.003652 | 0.005595 | 0.005352 | 5.2e-05 | 0.001117 | 0.28064 | 1.3057 | 2.5479 | 98.83 | 20.86 | 38.207 | 67.6559 | 109.7697 | 66.1049 | 0.009543 | 1.0406 | 0.024848 | 23.8776 | -0.016868 | 0.14336 | -5.3e-05 | -0.003704 | -0.003704 | -1.5906
2019-03-21 16:10 | 1553199025.8803 | 1235 | 0.2095 | 0.000463 | -4e-06 | 101.0698 | 101.2113 | 124.091 | 21.2612 | 0.001937 | 0.001839 | -152.4994 | 0.2406 | 8.4035 | 0.003559 | 0.005257 | 0.0051 | 4.7e-05 | 0.001117 | 0.28783 | 1.2805 | 2.5478 | 98.83 | 20.86 | 38.204 | 68.432 | 107.5105 | 66.076 | 0.009542 | 1.0407 | 0.024858 | 23.885 | -0.012505 | 0.14313 | -0.000182 | -0.01338 | -0.01338 | -0.43967
2019-03-21 16:10 | 1553199026.8807 | 1236 | 0.20949 | 0.000461 | -4e-06 | 101.1467 | 101.311 | 124.2395 | 21.2612 | 0.001921 | 0.001782 | -152.5033 | 0.24063 | 8.4031 | 0.003527 | 0.004983 | 0.004848 | 4.8e-05 | 0.00112 | 0.29927 | 1.3143 | 2.548 | 98.83 | 20.86 | 38.202 | 68.0382 | 106.7481 | 64.7014 | 0.009542 | 1.045 | 0.024923 | 23.8509 | -0.01452 | 0.14315 | -0.000106 | -0.007711 | -0.007711 | -0.76296
2019-03-21 16:10 | 1553199027.88 | 1237 | 0.20948 | 0.000463 | -4e-06 | 101.1107 | 101.2095 | 124.5017 | 21.2612 | 0.001961 | 0.001862 | -152.5222 | 0.24164 | 8.4049 | 0.003626 | 0.004872 | 0.004754 | 4.9e-05 | 0.001115 | 0.23955 | 1.3179 | 2.548 | 98.83 | 20.86 | 38.199 | 69.0171 | 109.1468 | 66.6191 | 0.009541 | 1.0391 | 0.024879 | 23.9437 | -0.008743 | 0.14375 | -3.8e-05 | -0.002569 | -0.002569 | -2.2996
2019-03-21 16:10 | 1553199028.8794 | 1238 | 0.20949 | 0.000458 | -4e-06 | 101.119 | 101.2521 | 124.3527 | 21.2612 | 0.001857 | 0.001817 | -152.5083 | 0.24097 | 8.4035 | 0.003554 | 0.004355 | 0.00412 | 5e-05 | 0.001118 | 0.29607 | 1.321 | 2.548 | 98.83 | 20.86 | 38.198 | 66.5113 | 107.4032 | 65.5555 | 0.009541 | 1.0423 | 0.024903 | 23.8919 | -0.011767 | 0.14335 | -0.000128 | -0.009272 | -0.009272 | -0.63546
2019-03-21 16:10 | 1553199029.8798 | 1239 | 0.20949 | 0.000462 | -4e-06 | 101.0979 | 101.3601 | 124.228 | 21.2612 | 0.001917 | 0.001812 | -152.5004 | 0.24087 | 8.4027 | 0.003538 | 0.004168 | 0.00399 | 4.7e-05 | 0.001115 | 0.25426 | 1.3198 | 2.548 | 98.83 | 20.86 | 38.197 | 67.9413 | 107.0175 | 65.4196 | 0.00954 | 1.0428 | 0.024896 | 23.8751 | -0.023171 | 0.14329 | -0.000106 | -0.007646 | -0.007646 | -0.77031
2019-03-21 16:10 | 1553199030.879 | 1240 | 0.20948 | 0.00046 | -3e-06 | 101.0535 | 101.2586 | 123.8781 | 21.2612 | 0.001961 | 0.001833 | -152.4787 | 0.24129 | 8.402 | 0.003629 | 0.004561 | 0.004339 | 4.9e-05 | 0.001119 | 0.26449 | 1.3229 | 2.5479 | 98.83 | 20.86 | 38.196 | 69.0077 | 109.2124 | 65.9356 | 0.00954 | 1.0412 | 0.024842 | 23.8596 | -0.018139 | 0.14355 | -3.4e-05 | -0.002333 | -0.002333 | -2.5289
2019-03-21 16:11 | 1553199031.8786 | 1241 | 0.20948 | 0.000462 | -4e-06 | 101.0839 | 101.2957 | 123.1816 | 21.2612 | 0.001896 | 0.001785 | -152.4502 | 0.24171 | 8.4029 | 0.003606 | 0.00532 | 0.005174 | 4.7e-05 | 0.001117 | 0.2792 | 1.3135 | 2.548 | 98.83 | 20.86 | 38.196 | 67.4391 | 108.6636 | 64.7703 | 0.00954 | 1.0448 | 0.024815 | 23.7515 | -0.018728 | 0.14379 | -7.3e-05 | -0.005206 | -0.005206 | -1.1352
2019-03-21 16:11 | 1553199032.8797 | 1242 | 0.20949 | 0.000462 | -3e-06 | 101.0931 | 101.2814 | 122.9105 | 21.2612 | 0.001932 | 0.001767 | -152.4419 | 0.24051 | 8.4015 | 0.003581 | 0.004388 | 0.004184 | 4.7e-05 | 0.001117 | 0.27616 | 1.3128 | 2.5479 | 98.83 | 20.86 | 38.194 | 68.308 | 108.0618 | 64.3557 | 0.00954 | 1.046 | 0.024803 | 23.7108 | -0.016647 | 0.14308 | -9.4e-05 | -0.006783 | -0.006783 | -0.86688
2019-03-21 16:11 | 1553199033.8791 | 1243 | 0.20948 | 0.000463 | -4e-06 | 101.1178 | 101.1731 | 122.9415 | 21.2612 | 0.001905 | 0.001822 | -152.4439 | 0.24181 | 8.4038 | 0.003644 | 0.005146 | 0.005063 | 4.7e-05 | 0.001119 | 0.28744 | 1.3189 | 2.5479 | 98.83 | 20.86 | 38.194 | 67.6699 | 109.5901 | 65.6654 | 0.00954 | 1.042 | 0.024758 | 23.7598 | -0.004889 | 0.14386 | -7.6e-05 | -0.005412 | -0.005412 | -1.0925
2019-03-21 16:11 | 1553199034.8802 | 1244 | 0.2095 | 0.000462 | -3e-06 | 101.1223 | 101.3822 | 122.8756 | 21.2612 | 0.001919 | 0.001788 | -152.4623 | 0.24185 | 8.4023 | 0.003591 | 0.004863 | 0.004635 | 4.4e-05 | 0.00112 | 0.26913 | 1.3369 | 2.5478 | 98.83 | 20.87 | 38.193 | 67.9993 | 108.3027 | 64.8388 | 0.009545 | 1.0445 | 0.024781 | 23.7244 | -0.022963 | 0.14387 | -0.000191 | -0.013987 | -0.013987 | -0.42278
2019-03-21 16:11 | 1553199035.8798 | 1245 | 0.20949 | 0.000461 | -4e-06 | 101.1178 | 101.3192 | 123.0288 | 21.2612 | 0.001903 | 0.001784 | -152.4734 | 0.24185 | 8.4032 | 0.00368 | 0.005117 | 0.004934 | 4.6e-05 | 0.001122 | 0.26449 | 1.3041 | 2.5479 | 98.83 | 20.87 | 38.19 | 67.6264 | 110.4703 | 64.7428 | 0.009545 | 1.0448 | 0.0248 | 23.7358 | -0.017798 | 0.14388 | -0.000114 | -0.008268 | -0.008268 | -0.71525
2019-03-21 16:11 | 1553199036.879 | 1246 | 0.20951 | 0.000464 | -4e-06 | 101.1038 | 101.3114 | 122.9518 | 21.2612 | 0.001913 | 0.0018 | -152.4611 | 0.24115 | 8.4019 | 0.003678 | 0.004992 | 0.004843 | 4.5e-05 | 0.001118 | 0.27256 | 1.3515 | 2.5478 | 98.83 | 20.87 | 38.189 | 67.8596 | 110.418 | 65.1278 | 0.009544 | 1.0437 | 0.024778 | 23.7419 | -0.018345 | 0.14346 | -0.000216 | -0.015852 | -0.015852 | -0.37195
2019-03-21 16:11 | 1553199037.8804 | 1247 | 0.2095 | 0.000463 | -4e-06 | 101.0931 | 101.3032 | 123.1744 | 21.2612 | 0.00192 | 0.001864 | -152.4474 | 0.24115 | 8.4036 | 0.003642 | 0.0053 | 0.005124 | 4.6e-05 | 0.001121 | 0.28622 | 1.3734 | 2.548 | 98.83 | 20.87 | 38.188 | 68.0264 | 109.549 | 66.6744 | 0.009544 | 1.0389 | 0.024744 | 23.8176 | -0.018572 | 0.14346 | -0.000188 | -0.01375 | -0.01375 | -0.42882
2019-03-21 16:11 | 1553199038.88 | 1248 | 0.20946 | 0.000456 | -3e-06 | 101.1211 | 101.2279 | 123.4863 | 21.2612 | 0.001954 | 0.001811 | -152.4398 | 0.24187 | 8.3984 | 0.003677 | 0.005136 | 0.004842 | 4.5e-05 | 0.001122 | 0.27383 | 1.3448 | 2.5479 | 98.83 | 20.87 | 38.187 | 68.8306 | 110.3881 | 65.4119 | 0.009544 | 1.0428 | 0.024822 | 23.8034 | -0.009446 | 0.14389 | 6.4e-05 | 0.004984 | 0.004984 | 1.1867
2019-03-21 16:11 | 1553199039.8791 | 1249 | 0.2095 | 0.000462 | -4e-06 | 101.1896 | 101.1929 | 123.8645 | 21.2612 | 0.001934 | 0.001853 | -152.4463 | 0.24141 | 8.4024 | 0.00364 | 0.005032 | 0.005043 | 4.3e-05 | 0.001121 | 0.26089 | 1.303 | 2.5478 | 98.83 | 20.87 | 38.186 | 68.3729 | 109.4805 | 66.4077 | 0.009544 | 1.0397 | 0.024823 | 23.8749 | -0.000286 | 0.14361 | -0.000151 | -0.011011 | -0.011011 | -0.53607
2019-03-21 16:11 | 1553199040.8797 | 1250 | 0.20948 | 0.000463 | -3e-06 | 101.1291 | 101.2167 | 124.2656 | 21.2612 | 0.001938 | 0.001758 | -152.4658 | 0.24127 | 8.4036 | 0.003595 | 0.004716 | 0.004431 | 4.8e-05 | 0.001119 | 0.27616 | 1.322 | 2.5478 | 98.83 | 20.87 | 38.184 | 68.4556 | 108.3957 | 64.1208 | 0.009543 | 1.0468 | 0.024948 | 23.8329 | -0.007741 | 0.14353 | -7.6e-05 | -0.00549 | -0.00549 | -1.0745
2019-03-21 16:11 | 1553199041.8801 | 1251 | 0.20948 | 0.000462 | -4e-06 | 101.0967 | 101.2497 | 124.6085 | 21.2612 | 0.00192 | 0.001767 | -152.4731 | 0.24222 | 8.4036 | 0.003525 | 0.004718 | 0.004531 | 4.5e-05 | 0.001116 | 0.24818 | 1.3186 | 2.5479 | 98.83 | 20.87 | 38.182 | 68.0296 | 106.6994 | 64.3487 | 0.009543 | 1.0461 | 0.024974 | 23.8738 | -0.013533 | 0.1441 | -7.8e-05 | -0.005606 | -0.005606 | -1.0563
2019-03-21 16:11 | 1553199042.8792 | 1252 | 0.20948 | 0.000463 | -4e-06 | 101.0725 | 101.0965 | 124.532 | 21.2612 | 0.001882 | 0.00188 | -152.4963 | 0.24174 | 8.4038 | 0.003517 | 0.004948 | 0.004807 | 4.4e-05 | 0.001118 | 0.26305 | 1.3183 | 2.548 | 98.83 | 20.87 | 38.18 | 67.1154 | 106.5048 | 67.0623 | 0.009542 | 1.0377 | 0.024866 | 23.9622 | -0.002119 | 0.14381 | -7.7e-05 | -0.005556 | -0.005556 | -1.0639
2019-03-21 16:11 | 1553199043.8806 | 1253 | 0.20949 | 0.000457 | -4e-06 | 101.0761 | 101.4973 | 124.2475 | 21.2612 | 0.001943 | 0.001792 | -152.4908 | 0.24181 | 8.4015 | 0.003557 | 0.004712 | 0.004529 | 4.4e-05 | 0.001116 | 0.2808 | 1.2965 | 2.548 | 98.83 | 20.87 | 38.175 | 68.5873 | 107.4663 | 64.9435 | 0.009541 | 1.0442 | 0.024915 | 23.8602 | -0.037199 | 0.14386 | -9.7e-05 | -0.007021 | -0.007021 | -0.84214
2019-03-21 16:11 | 1553199044.8794 | 1254 | 0.20946 | 0.000463 | -3e-06 | 101.1205 | 101.261 | 124.2647 | 21.2612 | 0.001929 | 0.001835 | -152.487 | 0.24177 | 8.4045 | 0.00359 | 0.005044 | 0.004918 | 4.5e-05 | 0.001107 | 0.24873 | 1.3282 | 2.548 | 98.83 | 20.87 | 38.172 | 68.2394 | 108.2753 | 65.9745 | 0.00954 | 1.041 | 0.024879 | 23.8982 | -0.012417 | 0.14383 | 3.8e-05 | 0.003066 | 0.003066 | 1.9281
2019-03-21 16:11 | 1553199045.8794 | 1255 | 0.20952 | 0.000467 | -4e-06 | 101.1622 | 101.1452 | 123.9562 | 21.2612 | 0.001933 | 0.001765 | -152.4633 | 0.2412 | 8.4011 | 0.003572 | 0.005047 | 0.004883 | 4.2e-05 | 0.001101 | 0.27383 | 1.3486 | 2.548 | 98.83 | 20.87 | 38.169 | 68.333 | 107.827 | 64.3003 | 0.009539 | 1.0462 | 0.02491 | 23.8095 | 0.001506 | 0.14349 | -0.000274 | -0.020245 | -0.020245 | -0.29131
2019-03-21 16:11 | 1553199046.8784 | 1256 | 0.2095 | 0.00046 | -4e-06 | 101.0931 | 101.2068 | 123.7051 | 21.2612 | 0.001967 | 0.001859 | -152.5012 | 0.24154 | 8.4029 | 0.003612 | 0.00487 | 0.004731 | 4.3e-05 | 0.0011 | 0.27256 | 1.3458 | 2.548 | 98.83 | 20.87 | 38.166 | 69.1582 | 108.8087 | 66.5652 | 0.009539 | 1.0392 | 0.024802 | 23.8651 | -0.010057 | 0.14369 | -0.000168 | -0.012257 | -0.012257 | -0.48182
2019-03-21 16:11 | 1553199047.8784 | 1257 | 0.2095 | 0.000465 | -3e-06 | 101.1622 | 101.2695 | 123.2481 | 21.2612 | 0.001873 | 0.001829 | -152.485 | 0.24138 | 8.4001 | 0.003585 | 0.004812 | 0.004657 | 4.2e-05 | 0.001103 | 0.27041 | 1.3543 | 2.5479 | 98.83 | 20.87 | 38.162 | 66.8836 | 108.1633 | 65.8348 | 0.009538 | 1.0415 | 0.024782 | 23.7953 | -0.00948 | 0.1436 | -0.000191 | -0.013983 | -0.013983 | -0.42208
2019-03-21 16:11 | 1553199048.8795 | 1258 | 0.2095 | 0.000459 | -4e-06 | 101.1503 | 101.2419 | 123.1467 | 21.2612 | 0.001979 | 0.001792 | -152.5258 | 0.24032 | 8.4028 | 0.00361 | 0.004788 | 0.004591 | 4.2e-05 | 0.001104 | 0.26825 | 1.3651 | 2.5479 | 98.83 | 20.87 | 38.16 | 69.4459 | 108.7693 | 64.9428 | 0.009537 | 1.0442 | 0.024805 | 23.7542 | -0.008097 | 0.14297 | -0.000191 | -0.013976 | -0.013976 | -0.42043
2019-03-21 16:11 | 1553199049.8795 | 1259 | 0.20951 | 0.000459 | -4e-06 | 101.1616 | 101.3032 | 122.8887 | 21.2612 | 0.001929 | 0.001809 | -152.5293 | 0.24145 | 8.4014 | 0.003618 | 0.004922 | 0.004766 | 4.2e-05 | 0.001108 | 0.27776 | 1.3438 | 2.548 | 98.83 | 20.87 | 38.156 | 68.2418 | 108.9505 | 65.3556 | 0.009536 | 1.043 | 0.024764 | 23.7438 | -0.012511 | 0.14364 | -0.000215 | -0.01577 | -0.01577 | -0.37436
2019-03-21 16:11 | 1553199050.8797 | 1260 | 0.20949 | 0.000461 | -4e-06 | 101.156 | 101.1935 | 122.996 | 21.2612 | 0.001932 | 0.001859 | -152.5497 | 0.24209 | 8.4008 | 0.003577 | 0.004292 | 0.004184 | 4.3e-05 | 0.001112 | 0.30165 | 1.3104 | 2.5479 | 98.83 | 20.87 | 38.151 | 68.3104 | 107.9618 | 66.5621 | 0.009535 | 1.0392 | 0.02473 | 23.7965 | -0.003323 | 0.14402 | -8.1e-05 | -0.005792 | -0.005792 | -1.022
2019-03-21 16:11 | 1553199051.8796 | 1261 | 0.2095 | 0.000462 | -4e-06 | 101.2334 | 101.2586 | 123.2293 | 21.2612 | 0.001944 | 0.001852 | -152.5606 | 0.24244 | 8.4005 | 0.003583 | 0.004649 | 0.004449 | 4.2e-05 | 0.001112 | 0.2912 | 1.3188 | 2.5481 | 98.83 | 20.87 | 38.148 | 68.5943 | 108.1058 | 66.3882 | 0.009534 | 1.0398 | 0.02476 | 23.8129 | -0.002224 | 0.14423 | -0.000166 | -0.012119 | -0.012119 | -0.48914
2019-03-21 16:11 | 1553199052.8794 | 1262 | 0.2095 | 0.000458 | -4e-06 | 101.1595 | 101.1987 | 123.3086 | 21.2612 | 0.001958 | 0.001854 | -152.5463 | 0.242 | 8.4031 | 0.003508 | 0.004299 | 0.00418 | 4.2e-05 | 0.001113 | 0.29231 | 1.3235 | 2.5481 | 98.83 | 20.87 | 38.144 | 68.9428 | 106.2852 | 66.4303 | 0.009533 | 1.0397 | 0.024767 | 23.8221 | -0.003458 | 0.14396 | -0.000176 | -0.012894 | -0.012894 | -0.45888
2019-03-21 16:11 | 1553199053.8789 | 1263 | 0.20949 | 0.000459 | -4e-06 | 101.1551 | 101.2416 | 123.433 | 21.2612 | 0.001951 | 0.001799 | -152.5494 | 0.24179 | 8.4018 | 0.003502 | 0.004815 | 0.004686 | 4.2e-05 | 0.001113 | 0.26952 | 1.3657 | 2.548 | 98.83 | 20.87 | 38.142 | 68.7635 | 106.1397 | 65.1052 | 0.009533 | 1.0437 | 0.024828 | 23.7875 | -0.007646 | 0.14384 | -0.000114 | -0.008221 | -0.008221 | -0.7191
2019-03-21 16:11 | 1553199054.8788 | 1264 | 0.2095 | 0.000463 | -4e-06 | 101.1518 | 101.3059 | 123.6872 | 21.2612 | 0.001921 | 0.001797 | -152.5517 | 0.24118 | 8.4001 | 0.003455 | 0.004445 | 0.00423 | 4.2e-05 | 0.001114 | 0.24403 | 1.3043 | 2.5481 | 98.83 | 20.87 | 38.14 | 68.0591 | 105.0092 | 65.0631 | 0.009532 | 1.0439 | 0.024855 | 23.8105 | -0.013621 | 0.14348 | -0.000194 | -0.014227 | -0.014227 | -0.41447
2019-03-21 16:11 | 1553199055.8788 | 1265 | 0.2095 | 0.00046 | -4e-06 | 101.1113 | 101.3172 | 123.8783 | 21.2612 | 0.001915 | 0.001811 | -152.5714 | 0.24124 | 8.4018 | 0.003528 | 0.004474 | 0.004319 | 4.5e-05 | 0.001112 | 0.28512 | 1.3059 | 2.548 | 98.83 | 20.87 | 38.138 | 67.9094 | 106.7583 | 65.397 | 0.009532 | 1.0428 | 0.024862 | 23.8407 | -0.018197 | 0.14351 | -0.000166 | -0.012122 | -0.012122 | -0.4866
2019-03-21 16:11 | 1553199056.879 | 1266 | 0.20951 | 0.000459 | -4e-06 | 101.1542 | 101.2637 | 123.8886 | 21.2612 | 0.00194 | 0.001803 | -152.5432 | 0.24138 | 8.4024 | 0.003465 | 0.005126 | 0.005011 | 4.4e-05 | 0.001114 | 0.25802 | 1.2981 | 2.5478 | 98.83 | 20.87 | 38.139 | 68.496 | 105.238 | 65.212 | 0.009532 | 1.0434 | 0.024869 | 23.8351 | -0.00968 | 0.1436 | -0.000227 | -0.016749 | -0.016749 | -0.35237
2019-03-21 16:11 | 1553199057.8804 | 1267 | 0.20949 | 0.000465 | -3e-06 | 101.1697 | 101.4119 | 124.0529 | 21.2612 | 0.001919 | 0.001778 | -152.5592 | 0.24136 | 8.402 | 0.003468 | 0.004863 | 0.004666 | 4.5e-05 | 0.001113 | 0.27417 | 1.3079 | 2.548 | 98.83 | 20.87 | 38.14 | 68.0125 | 105.3095 | 64.6002 | 0.009532 | 1.0453 | 0.024908 | 23.8294 | -0.021387 | 0.14359 | -0.000116 | -0.008438 | -0.008438 | -0.69937
2019-03-21 16:11 | 1553199058.8792 | 1268 | 0.20949 | 0.000459 | -3e-06 | 101.1849 | 101.3826 | 123.8912 | 21.2612 | 0.001969 | 0.001807 | -152.529 | 0.24086 | 8.4041 | 0.003555 | 0.005368 | 0.005211 | 4.6e-05 | 0.001117 | 0.2625 | 1.2852 | 2.5479 | 98.83 | 20.87 | 38.14 | 69.2089 | 107.426 | 65.3128 | 0.009532 | 1.0431 | 0.024866 | 23.8389 | -0.01746 | 0.14329 | -0.00012 | -0.008764 | -0.008764 | -0.67195
2019-03-21 16:11 | 1553199059.8801 | 1269 | 0.20949 | 0.000461 | -3e-06 | 101.1488 | 101.2726 | 123.8377 | 21.2612 | 0.001864 | 0.001773 | -152.5324 | 0.24136 | 8.4029 | 0.00361 | 0.004623 | 0.004485 | 4.7e-05 | 0.001116 | 0.28119 | 1.3231 | 2.5479 | 98.83 | 20.87 | 38.141 | 66.6761 | 108.7504 | 64.4899 | 0.009532 | 1.0456 | 0.024891 | 23.8048 | -0.010937 | 0.14359 | -0.000146 | -0.010679 | -0.010679 | -0.55265
2019-03-21 16:11 | 1553199060.8804 | 1270 | 0.20955 | 0.000462 | -4e-06 | 101.1133 | 101.1479 | 123.5694 | 21.2612 | 0.001955 | 0.001758 | -152.5138 | 0.24149 | 8.4056 | 0.003522 | 0.004969 | 0.004775 | 4.4e-05 | 0.001116 | 0.29645 | 1.3585 | 2.5479 | 98.83 | 20.87 | 38.142 | 68.8587 | 106.6229 | 64.1261 | 0.009533 | 1.0468 | 0.024877 | 23.7662 | -0.003056 | 0.14366 | -0.000484 | -0.035957 | -0.035957 | -0.16421
2019-03-21 16:11 | 1553199061.8782 | 1271 | 0.20949 | 0.000458 | -4e-06 | 101.1247 | 101.3523 | 123.2504 | 21.2612 | 0.00191 | 0.001806 | -152.5175 | 0.2406 | 8.4033 | 0.003535 | 0.004792 | 0.004592 | 4.5e-05 | 0.001119 | 0.30071 | 1.3411 | 2.548 | 98.83 | 20.87 | 38.143 | 67.7776 | 106.9435 | 65.2832 | 0.009533 | 1.0432 | 0.024803 | 23.7762 | -0.020107 | 0.14313 | -8.4e-05 | -0.005995 | -0.005995 | -0.98126
2019-03-21 16:11 | 1553199062.8783 | 1272 | 0.20948 | 0.000461 | -4e-06 | 101.0713 | 101.2395 | 123.2762 | 21.2612 | 0.001926 | 0.001839 | -152.5254 | 0.24139 | 8.4031 | 0.003471 | 0.005251 | 0.005128 | 4.8e-05 | 0.001118 | 0.27024 | 1.3194 | 2.5479 | 98.83 | 20.87 | 38.147 | 68.1606 | 105.3937 | 66.0823 | 0.009534 | 1.0407 | 0.024776 | 23.8067 | -0.014871 | 0.1436 | -6.6e-05 | -0.004673 | -0.004673 | -1.2631
2019-03-21 16:11 | 1553199063.8784 | 1273 | 0.20948 | 0.000464 | -3e-06 | 101.1649 | 101.3482 | 123.111 | 21.2612 | 0.001888 | 0.001793 | -152.5172 | 0.24096 | 8.4012 | 0.003431 | 0.005053 | 0.004871 | 4.5e-05 | 0.001119 | 0.28998 | 1.3407 | 2.5479 | 98.83 | 20.87 | 38.148 | 67.251 | 104.4058 | 64.9685 | 0.009534 | 1.0441 | 0.0248 | 23.7517 | -0.01619 | 0.14335 | -6.1e-05 | -0.004372 | -0.004372 | -1.3477
2019-03-21 16:11 | 1553199064.879 | 1274 | 0.20953 | 0.000458 | -4e-06 | 101.0824 | 101.1813 | 122.9002 | 21.2612 | 0.00195 | 0.001791 | -152.518 | 0.2412 | 8.4006 | 0.00342 | 0.00494 | 0.004733 | 4.3e-05 | 0.001122 | 0.25874 | 1.32 | 2.548 | 98.83 | 20.87 | 38.149 | 68.7395 | 104.155 | 64.9334 | 0.009534 | 1.0443 | 0.02478 | 23.7301 | -0.008749 | 0.14349 | -0.000381 | -0.028112 | -0.028112 | -0.20978
2019-03-21 16:11 | 1553199065.8787 | 1275 |  | 0.00046 |  | 101.1214 | 101.185 | 122.9314 | 21.2612 | 0.001949 | 0.001813 | -152.5169 | 0.2416 | 8.4008 | 0.003432 | 0.005017 | 0.004875 | 4.2e-05 | 0.001121 | 0.28512 | 1.3604 | 2.5479 | 98.83 | 20.87 | 38.152 | 68.719 | 104.4461 | 65.4501 | 0.009535 | 1.0427 | 0.024765 | 23.7512 | -0.005627 | 0.14373
2019-03-21 16:11 | 1553199066.8787 | 1276 |  | 0.000462 |  | 100.9888 | 101.0573 | 123.0617 | 21.2612 | 0.001904 | 0.001824 | -152.5241 | 0.24199 | 8.4014 | 0.003429 | 0.004735 | 0.004457 | 4.5e-05 | 0.001119 | 0.26792 | 1.3075 | 2.548 | 98.83 | 20.87 | 38.154 | 67.6466 | 104.369 | 65.7179 | 0.009536 | 1.0418 | 0.024768 | 23.7732 | -0.006065 | 0.14396
2019-03-21 16:11 | 1553199067.8787 | 1277 |  |  |  | 101.0326 | 101.2232 | 123.1603 | 21.2612 | 0.00192 | 0.001788 | -152.5242 | 0.24113 | 8.4025 | 0.00338 | 0.004199 | 0.004052 | 4.8e-05 | 0.00112 | 0.2975 | 1.3222 | 2.548 | 98.83 | 20.87 | 38.155 | 68.0178 | 103.1804 | 64.8475 | 0.009536 | 1.0445 | 0.02481 | 23.7522 | -0.016855 | 0.14345
2019-03-21 16:11 | 1553199068.8787 | 1278 |  |  |  | 101.0472 | 101.1489 | 123.1988 | 21.2612 | 0.001938 | 0.001815 | -152.5497 | 0.24164 | 8.3997 | 0.003334 | 0.005086 | 0.004896 | 4.1e-05 | 0.001117 | 0.27433 | 1.3304 | 2.548 | 98.83 | 20.87 | 38.159 | 68.4556 | 102.0705 | 65.5016 | 0.009537 | 1.0425 | 0.02479 | 23.7789 | -0.009 | 0.14375
2019-03-21 16:11 | 1553199069.8787 | 1279 |  |  |  | 101.0767 | 101.1299 | 123.2384 | 21.2612 | 0.001933 | 0.001842 | -152.5735 | 0.24146 | 8.4022 | 0.003477 | 0.004496 | 0.004345 | 4.5e-05 | 0.001117 | 0.28512 | 1.3549 | 2.5481 | 98.83 | 20.87 | 38.16 | 68.3284 | 105.5329 | 66.1463 | 0.009537 | 1.0405 | 0.02477 | 23.8053 | -0.004702 | 0.14365
2019-03-21 16:11 | 1553199070.8787 | 1280 |  |  |  | 101.0713 | 101.1666 |  | 21.2612 | 0.001937 | 0.001758 | -152.5795 | 0.24149 | 8.4021 | 0.003449 | 0.004858 | 0.004616 | 4.7e-05 | 0.001113 | 0.26792 | 1.3632 | 2.548 | 98.83 | 20.87 | 38.162 | 68.4431 | 104.8511 | 64.13 | 0.009538 | 1.0467 |  |  | -0.00843 | 0.14366
2019-03-21 16:11 | 1553199071.8787 | 1281 |  |  |  | 101.0993 | 101.1438 |  | 21.2612 | 0.001942 |  | -152.5666 | 0.24163 | 8.4007 | 0.003402 | 0.00467 | 0.004466 | 4.6e-05 | 0.001115 | 0.27815 | 1.3527 | 2.5478 | 98.83 | 20.87 | 38.162 | 68.5592 | 103.7095 |  | 0.009538 |  |  |  | -0.003934 | 0.14375
2019-03-21 16:11 | 1553199072.8787 | 1282 |  |  |  | 101.0073 | 100.9504 |  | 21.2612 | 0.001848 |  |  | 0.2414 | 8.4034 | 0.003441 | 0.005513 | 0.0053 | 5e-05 | 0.001117 | 0.26609 | 1.3314 | 2.5478 | 98.83 | 20.87 | 38.164 | 66.2843 | 104.6678 |  | 0.009538 |  |  |  | 0.005042 | 0.14361
2019-03-21 16:11 | 1553199073.8787 | 1283 |  |  |  | 100.9921 | 101.1925 |  | 21.2612 |  |  |  | 0.24114 | 8.4035 | 0.003449 | 0.004417 | 0.004334 | 4.6e-05 | 0.001115 | 0.27295 | 1.3181 | 2.5479 | 98.83 | 20.87 | 38.166 |  | 104.8424 |  | 0.009539 |  |  |  | -0.017734 | 0.14346
2019-03-21 16:11 | 1553199074.8787 | 1284 |  |  |  | 100.9468 | 101.0015 |  | 21.2612 |  |  |  | 0.24189 | 8.4017 | 0.003467 | 0.004517 | 0.004334 | 4.9e-05 | 0.001116 | 0.25962 | 1.2979 | 2.5479 | 98.83 | 20.87 | 38.167 |  | 105.2906 |  | 0.009539 |  |  |  | -0.004841 | 0.1439
2019-03-21 16:11 | 1553199075.8787 | 1285 |  |  |  | 100.9817 | 101.0028 |  | 21.2612 |  |  |  | 0.24111 | 8.401 | 0.003394 | 0.004589 | 0.004343 | 5.1e-05 | 0.00112 | 0.27848 | 1.3143 | 2.5479 | 98.83 | 20.87 | 38.168 |  | 103.5246 |  | 0.009539 |  |  |  | -0.001874 | 0.14343
2019-03-21 16:11 | 1553199076.8787 | 1286 |  |  |  | 100.97 | 101.1714 |  | 21.2612 |  |  |  | 0.24088 | 8.3981 | 0.003398 | 0.004904 | 0.004759 | 5.1e-05 | 0.001116 | 0.24729 | 1.2873 | 2.548 | 98.83 | 20.87 | 38.168 |  | 103.6244 |  | 0.009539 |  |  |  | -0.01782 | 0.1433
2019-03-21 16:11 | 1553199077.8787

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-21 15:50 | 0 | 0 | 0 | 0 | 0 | 46.6 | 0.010205
2019-03-21 15:50 | 1 | 0 | 0 | 0 | 0 | 46.8011 | 0.005648
2019-03-21 15:50 | 2 | 0 | 0 | 0 | 0 | 46.7029 | -0.003137
2019-03-21 15:50 | 3 | 0 | 0 | 0 | 0 | 46.5182 | 0.002943
2019-03-21 15:50 | 4 | 0 | 0 | 0 | 0 | 46.7535 | 0.0006
2019-03-21 15:50 | 5 | 0 | 0 | 0 | 0 | 46.5078 | 0.015626
2019-03-21 15:50 | 6 | 0 | 0 | 0 | 0 | 46.4988 | 0.092805
2019-03-21 15:50 | 7 | 0 | 0 | 0 | 0 | 46.4369 | 0.19584
2019-03-21 15:50 | 8 | 0 | 0 | 0 | 0 | 46.3883 | 0.59433
2019-03-21 15:50 | 9 | 0 | 0 | 0 | 0 | 46.3062 | 0.95476
2019-03-21 15:50 | 10 | 0 | 0 | 0 | 0 | 46.2558 | 1.3173
2019-03-21 15:50 | 11 | 0 | 0 | 0 | 0 | 46.158 | 2.0635
2019-03-21 15:50 | 12 | 0 | 0 | 0 | 0 | 46.0804 | 2.7554
2019-03-21 15:50 | 13 | 0 | 0 | 0 | 0 | 45.9504 | 3.1414
2019-03-21 15:50 | 14 | 0 | 0 | 0 | 0 | 45.8394 | 3.3391
2019-03-21 15:50 | 15 | 0 | 0 | 0 | 0 | 45.7271 | 3.3701
2019-03-21 15:50 | 16 | 0.10084 | 0 | 0 | 0 | 45.644 | 3.4085
2019-03-21 15:50 | 17 | 0.11222 | 0 | 0 | 0 | 45.5241 | 3.3831
2019-03-21 15:50 | 18 | 0.094963 | 0 | 0 | 0 | 45.4255 | 3.3897
2019-03-21 15:50 | 19 | 0.11736 | 0 | 0 | 0 | 45.3257 | 3.3568
2019-03-21 15:50 | 20 | 0.11815 | 0 | 0 | 0 | 45.1962 | 3.1954
2019-03-21 15:50 | 21 | 0.083742 | 0 | 0 | 0 | 45.098 | 3.0798
2019-03-21 15:50 | 22 | 0.081792 | 0 | 0 | 0 | 45.0218 | 2.888
2019-03-21 15:50 | 23 | 0.072685 | 0 | 0 | 0 | 44.9354 | 2.7865
2019-03-21 15:50 | 24 | 0.059436 | 0 | 0 | 0 | 44.8766 | 2.4736
2019-03-21 15:50 | 25 | 0.079376 | 0 | 0 | 0 | 44.809 | 2.2699
2019-03-21 15:50 | 26 | 0.070035 | 0 | 0 | 0 | 44.7241 | 2.1508
2019-03-21 15:50 | 27 | 0.066254 | 0 | 0 | 0 | 44.6673 | 2.0379
2019-03-21 15:50 | 28 | 0.099382 | 0 | 0 | 0 | 44.5836 | 2.0016
2019-03-21 15:50 | 29 | 0.079891 | 0 | 0 | 0 | 44.4803 | 1.9786
2019-03-21 15:50 | 30 | 0.08008 | 0 | 0 | 0 | 44.4201 | 1.9305
2019-03-21 15:50 | 31 | 0.088934 | 0 | 0 | 0 | 44.3179 | 1.8867
2019-03-21 15:50 | 32 | 0.081574 | 0 | 0 | 0 | 44.2456 | 1.8324
2019-03-21 15:50 | 33 | 0.089702 | 0 | 0 | 0 | 44.1519 | 1.7912
2019-03-21 15:50 | 34 | 0.075712 | 0 | 0 | 0 | 44.0714 | 1.7413
2019-03-21 15:50 | 35 | 0.064197 | 0 | 0 | 0 | 44.0002 | 1.68
2019-03-21 15:50 | 36 | 0.065842 | 0 | 0 | 0 | 43.94 | 1.5821
2019-03-21 15:50 | 37 | 0.053755 | 0 | 0 | 0 | 43.8718 | 1.5096
2019-03-21 15:50 | 38 | 0.03154 | 0 | 0 | 0 | 43.8344 | 1.4102
2019-03-21 15:50 | 39 | 0.044809 | 0 | 0 | 0 | 43.8002 | 1.3303
2019-03-21 15:50 | 40 | 0.05022 | 0 | 0 | 0 | 43.7456 | 1.2916
2019-03-21 15:50 | 41 | 0.058291 | 0 | 0 | 0 | 43.6994 | 1.1849
2019-03-21 15:51 | 42 | 0.060637 | 0 | 0 | 0 | 43.6308 | 1.1371
2019-03-21 15:51 | 43 | 0.052452 | 0 | 0 | 0 | 43.5809 | 1.1099
2019-03-21 15:51 | 44 | 0.048897 | 0 | 0 | 0 | 43.525 | 1.1102
2019-03-21 15:51 | 45 | 0.044199 | 0 | 0 | 0 | 43.4829 | 1.0675
2019-03-21 15:51 | 46 | 0.049592 | 0 | 0 | 0 | 43.4335 | 1.0721
2019-03-21 15:51 | 47 | 0.075893 | 0 | 0 | 0 | 43.3792 | 1.0409
2019-03-21 15:51 | 48 | 0.065593 | 0 | 0 | 0 | 43.2911 | 1.0114
2019-03-21 15:51 | 49 | 0.033379 | 0 | 0 | 0 | 43.2537 | 1.0095
2019-03-21 15:51 | 50 | 0.03641 | 0 | 0 | 0 | 43.2161 | 0.99502
2019-03-21 15:51 | 51 | 0.038385 | 0 | 0 | 0 | 43.1806 | 0.98826
2019-03-21 15:51 | 52 | 0.02305 | 0 | 0 | 0 | 43.143 | 0.99465
2019-03-21 15:51 | 53 | 0.029385 | 0 | 0 | 0 | 43.129 | 0.98646
2019-03-21 15:51 | 54 | 0.049949 | 0 | 0 | 0 | 43.0804 | 0.98544
2019-03-21 15:51 | 55 | 0.055688 | 0 | 0 | 0 | 43.0327 | 0.99786
2019-03-21 15:51 | 56 | 0.054192 | 0 | 0 | 0 | 42.9717 | 0.98321
2019-03-21 15:51 | 57 | 0.040194 | 0 | 0 | 0 | 42.9274 | 0.98977
2019-03-21 15:51 | 58 | 0.036783 | 0 | 0 | 0 | 42.8884 | 0.95707
2019-03-21 15:51 | 59 | 0.05374 | 0 | 0 | 0 | 42.8491 | 0.9549
2019-03-21 15:51 | 60 | 0.044901 | 0 | 0 | 0 | 42.7864 | 0.956
2019-03-21 15:51 | 61 | 0.05221 | 0 | 0 | 0 | 42.7559 | 0.93051
2019-03-21 15:51 | 62 | 0.061265 | 0 | 0 | 0 | 42.682 | 0.91335
2019-03-21 15:51 | 63 | 0.038205 | 0 | 0 | 0 | 42.6407 | 0.9091
2019-03-21 15:51 | 64 | 0.0432 | 0 | 0 | 0 | 42.5996 | 0.90795
2019-03-21 15:51 | 65 | 0.04591 | 0 | 0 | 0 | 42.5546 | 0.90833
2019-03-21 15:51 | 66 | 0.03147 | 0 | 0 | 0 | 42.5116 | 0.90792
2019-03-21 15:51 | 67 | 0.029712 | 0 | 0 | 0 | 42.4882 | 0.90096
2019-03-21 15:51 | 68 | 0.049654 | 0 | 0 | 0 | 42.4465 | 0.91836
2019-03-21 15:51 | 69 | 0.065422 | 0 | 0 | 0 | 42.3901 | 0.90925
2019-03-21 15:51 | 70 | 0.058302 | 0 | 0 | 0 | 42.3222 | 0.90145
2019-03-21 15:51 | 71 | 0.026088 | 0 | 0 | 0 | 42.2792 | 0.903
2019-03-21 15:51 | 72 | 0.038972 | 0.011082 | 2.4893 | 22.6579 | 42.2592 | 0.88302
2019-03-21 15:51 | 73 | 0.059206 | 0.0069009 | 1.6278 | 14.7739 | 42.1989 | 0.87471
2019-03-21 15:51 | 74 | 0.047958 | 0.0085384 | 2.0007 | 18.0263 | 42.1479 | 0.8645
2019-03-21 15:51 | 75 | 0.057094 | 0.0067782 | 1.6729 | 15.634 | 42.0989 | 0.89261
2019-03-21 15:51 | 76 | 0.062773 | 0.0065212 | 1.5104 | 13.7716 | 42.0346 | 0.86448
2019-03-21 15:51 | 77 | 0.050516 | 0.0081241 | 1.8717 | 17.3038 | 41.9779 | 0.87412
2019-03-21 15:51 | 78 | 0.04434 | 0.008738 | 2.121 | 19.3063 | 41.9325 | 0.85604
2019-03-21 15:51 | 79 | 0.038799 | 0.0099638 | 2.4029 | 22.0582 | 41.8885 | 0.85583
2019-03-21 15:51 | 80 | 0.045797 | 0.0084512 | 2.0288 | 18.617 | 41.8515 | 0.85261
2019-03-21 15:51 | 81 | 0.059845 | 0.0068409 | 1.5426 | 14.2406 | 41.7954 | 0.85223
2019-03-21 15:51 | 82 | 0.048192 | 0.0080108 | 1.8979 | 17.8864 | 41.7379 | 0.86199
2019-03-21 15:51 | 83 | 0.042951 | 0.0095342 | 2.1258 | 20.0951 | 41.6979 | 0.86309
2019-03-21 15:51 | 84 | 0.049309 | 0.0078453 | 1.8398 | 17.5534 | 41.6497 | 0.86554
2019-03-21 15:51 | 85 | 0.031785 | 0.012164 | 2.8436 | 27.1726 | 41.6046 | 0.86369
2019-03-21 15:51 | 86 | 0.028441 | 0.013585 | 3.1621 | 30.4444 | 41.5829 | 0.86588
2019-03-21 15:51 | 87 | 0.032116 | 0.012025 | 2.7889 | 27.1786 | 41.5453 | 0.87288
2019-03-21 15:51 | 88 | 0.043773 | 0.0088121 | 2.034 | 19.9145 | 41.5165 | 0.87171
2019-03-21 15:51 | 89 | 0.051536 | 0.0074955 | 1.728 | 17.1359 | 41.459 | 0.88311
2019-03-21 15:51 | 90 | 0.042994 | 0.0095145 | 2.0708 | 20.5002 | 41.4177 | 0.88138
2019-03-21 15:51 | 91 | 0.035691 | 0.011509 | 2.506 | 25.1829 | 41.3731 | 0.8988
2019-03-21 15:51 | 92 | 0.03049 | 0.013448 | 2.9198 | 29.6453 | 41.3455 | 0.9039
2019-03-21 15:51 | 93 | 0.034322 | 0.012601 | 2.592 | 26.23 | 41.3096 | 0.90026
2019-03-21 15:51 | 94 | 0.04047 | 0.010691 | 2.1991 | 22.201 | 41.2759 | 0.89848
2019-03-21 15:51 | 95 | 0.049059 | 0.0088056 | 1.8105 | 18.4664 | 41.228 | 0.90594
2019-03-21 15:51 | 96 | 0.057168 | 0.0079343 | 1.5535 | 16.0762 | 41.1784 | 0.91904
2019-03-21 15:51 | 97 | 0.047857 | 0.0090247 | 1.8593 | 19.2531 | 41.1183 | 0.92141
2019-03-21 15:51 | 98 | 0.035813 | 0.012068 | 2.4924 | 25.4823 | 41.084 | 0.91261
2019-03-21 15:51 | 99 | 0.0183 | 0.023642 | 4.8766 | 49.7406 | 41.0476 | 0.91023
2019-03-21 15:51 | 100 | 0.018356 | 0.023548 | 4.8728 | 50.8767 | 41.0425 | 0.9339
2019-03-21 15:52 | 101 | 0.036938 | 0.011683 | 2.4148 | 25.0647 | 41.0055 | 0.92585
2019-03-21 15:52 | 102 | 0.0622 | 0.0069392 | 1.4365 | 14.8664 | 40.9673 | 0.92469
2019-03-21 15:52 | 103 | 0.05481 | 0.0078913 | 1.633 | 16.8873 | 40.8892 | 0.92559
2019-03-21 15:52 | 104 | 0.047819 | 0.0090235 | 1.8732 | 19.4286 | 40.8584 | 0.92906
2019-03-21 15:52 | 105 | 0.045792 | 0.0094425 | 1.9626 | 20.4275 | 40.7925 | 0.93542
2019-03-21 15:52 | 106 | 0.033125 | 0.01301 | 2.704 | 28.6982 | 40.7681 | 0.95064
2019-03-21 15:52 | 107 | 0.066219 | 0.0065211 | 1.3575 | 14.3696 | 40.7161 | 0.95155
2019-03-21 15:52 | 108 | 0.057917 | 0.0074299 | 1.5486 | 16.2792 | 40.6447 | 0.94284
2019-03-21 15:52 | 109 | 0.052339 | 0.0086449 | 1.7137 | 18.4605 | 40.6001 | 0.9662
2019-03-21 15:52 | 110 | 0.069884 | 0.0064795 | 1.287 | 13.8473 | 40.5351 | 0.9677
2019-03-21 15:52 | 111 | 0.058617 | 0.0081116 | 1.5399 | 16.6587 | 40.4673 | 0.97648
2019-03-21 15:52 | 112 | 0.037094 | 0.013442 | 2.4388 | 26.3162 | 40.4204 | 0.97619
2019-03-21 15:52 | 113 | 0.045725 | 0.011382 | 1.9838 | 21.7002 | 40.3858 | 0.99224
2019-03-21 15:52 | 114 | 0.052263 | 0.010404 | 1.744 | 19.0984 | 40.3286 | 0.99814
2019-03-21 15:52 | 115 | 0.058948 | 0.0092048 | 1.5489 | 17.0426 | 40.2813 | 1.0046
2019-03-21 15:52 | 116 | 0.065202 | 0.008691 | 1.4111 | 15.5409 | 40.2117 | 1.0133
2019-03-21 15:52 | 117 | 0.038802 | 0.014543 | 2.3694 | 26.3056 | 40.1583 | 1.0207
2019-03-21 15:52 | 118 | 0.044954 | 0.013113 | 2.0662 | 22.8702 | 40.1267 | 1.0281
2019-03-21 15:52 | 119 | 0.053349 | 0.01104 | 1.7476 | 19.3947 | 40.0671 | 1.0347
2019-03-21 15:52 | 120 | 0.05539 | 0.011057 | 1.6943 | 18.8728 | 40.0213 | 1.0454
2019-03-21 15:52 | 121 | 0.066157 | 0.0092598 | 1.4244 | 15.9337 | 39.9548 | 1.0541
2019-03-21 15:52 | 122 | 0.057499 | 0.010627 | 1.6409 | 18.1783 | 39.8935 | 1.0452
2019-03-21 15:52 | 123 | 0.057885 | 0.010536 | 1.6342 | 18.2415 | 39.8385 | 1.0559
2019-03-21 15:52 | 124 | 0.044262 | 0.013803 | 2.1482 | 23.6804 | 39.7807 | 1.0481
2019-03-21 15:52 | 125 | 0.044322 | 0.013786 | 2.1576 | 23.4223 | 39.747 | 1.0381
2019-03-21 15:52 | 126 | 0.037306 | 0.016381 | 2.5666 | 27.9999 | 39.693 | 1.0446
2019-03-21 15:52 | 127 | 0.042211 | 0.015018 | 2.2791 | 25.0609 | 39.6693 | 1.0578
2019-03-21 15:52 | 128 | 0.056288 | 0.011618 | 1.7049 | 18.7073 | 39.6064 | 1.053
2019-03-21 15:52 | 129 | 0.057351 | 0.01176 | 1.6731 | 18.4132 | 39.5605 | 1.056
2019-03-21 15:52 | 130 | 0.031204 | 0.022391 | 3.0929 | 34.0375 | 39.498 | 1.0621
2019-03-21 15:52 | 131 | 0.036513 | 0.019149 | 2.6548 | 29.0447 | 39.4902 | 1.0605
2019-03-21 15:52 | 132 | 0.066179 | 0.010906 | 1.468 | 15.7581 | 39.4189 | 1.0429
2019-03-21 15:52 | 133 | 0.062484 | 0.01119 | 1.5582 | 16.6824 | 39.3657 | 1.0424
2019-03-21 15:52 | 134 | 0.058235 | 0.011981 | 1.6732 | 17.916 | 39.2956 | 1.0433
2019-03-21 15:52 | 135 | 0.029881 | 0.023358 | 3.268 | 35.2744 | 39.2539 | 1.054
2019-03-21 15:52 | 136 | 0.051537 | 0.013566 | 1.9062 | 20.5519 | 39.2237 | 1.0592
2019-03-21 15:52 | 137 | 0.078908 | 0.0088487 | 1.2461 | 13.4346 | 39.1492 | 1.0601
2019-03-21 15:52 | 138 | 0.061102 | 0.011805 | 1.6127 | 17.3864 | 39.0765 | 1.0623
2019-03-21 15:52 | 139 | 0.062325 | 0.011562 | 1.5802 | 17.0221 | 39.0232 | 1.0609
2019-03-21 15:52 | 140 | 0.07891 | 0.0094153 | 1.2527 | 13.2451 | 38.9486 | 1.0452
2019-03-21 15:52 | 141 | 0.059387 | 0.012502 | 1.6657 | 17.6344 | 38.8738 | 1.0472
2019-03-21 15:52 | 142 | 0.049085 | 0.015575 | 2.0178 | 21.4042 | 38.828 | 1.0506
2019-03-21 15:52 | 143 | 0.055096 | 0.014724 | 1.8073 | 18.8761 | 38.7712 | 1.04
2019-03-21 15:52 | 144 | 0.061762 | 0.013144 | 1.6174 | 17.1676 | 38.717 | 1.0603
2019-03-21 15:52 | 145 | 0.076726 | 0.010858 | 1.3009 | 13.4763 | 38.6465 | 1.034
2019-03-21 15:52 | 146 | 0.058074 | 0.014719 | 1.7201 | 18.0069 | 38.5715 | 1.0457
2019-03-21 15:52 | 147 | 0.047143 | 0.018598 | 2.1207 | 21.5694 | 38.5285 | 1.0169
2019-03-21 15:52 | 148 | 0.065884 | 0.013676 | 1.5183 | 15.741 | 38.4702 | 1.0371
2019-03-21 15:52 | 149 | 0.063621 | 0.0145 | 1.5722 | 15.8183 | 38.4015 | 1.0064
2019-03-21 15:52 | 150 | 0.049584 | 0.018592 | 2.016 | 20.1903 | 38.346 | 1.0011
2019-03-21 15:52 | 151 | 0.054596 | 0.017305 | 1.8291 | 18.5079 | 38.2982 | 1.0105
2019-03-21 15:52 | 152 | 0.049477 | 0.019056 | 2.0149 | 20.0042 | 38.2387 | 0.98975
2019-03-21 15:52 | 153 | 0.046731 | 0.020199 | 2.1357 | 21.4404 | 38.1986 | 1.0019
2019-03-21 15:52 | 154 | 0.05389 | 0.017935 | 1.8497 | 18.4309 | 38.1428 | 0.99324
2019-03-21 15:52 | 155 | 0.056428 | 0.016783 | 1.7638 | 17.6201 | 38.0917 | 0.99427
2019-03-21 15:52 | 156 | 0.059864 | 0.016181 | 1.6605 | 16.6795 | 38.0301 | 0.9985
2019-03-21 15:52 | 157 | 0.055347 | 0.017881 | 1.7874 | 17.9959 | 37.9739 | 0.99602
2019-03-21 15:52 | 158 | 0.04989 | 0.019847 | 1.9805 | 19.701 | 37.9194 | 0.98289
2019-03-21 15:52 | 159 | 0.060842 | 0.016631 | 1.6214 | 16.0717 | 37.8702 | 0.97783
2019-03-21 15:52 | 160 | 0.067764 | 0.015579 | 1.4522 | 14.2204 | 37.7985 | 0.96363
2019-03-21 15:52 | 161 | 0.06481 | 0.01699 | 1.5147 | 14.8351 | 37.7366 | 0.96147
2019-03-21 15:53 | 162 | 0.084182 | 0.013633 | 1.167 | 11.5177 | 37.6644 | 0.96958
2019-03-21 15:53 | 163 | 0.073129 | 0.015982 | 1.3354 | 13.3051 | 37.5754 | 0.97298
2019-03-21 15:53 | 164 | 0.058792 | 0.020672 | 1.6603 | 16.7341 | 37.5191 | 0.98383
2019-03-21 15:53 | 165 | 0.07139 | 0.017629 | 1.3643 | 13.6186 | 37.4516 | 0.97223
2019-03-21 15:53 | 166 | 0.071516 | 0.017566 | 1.359 | 13.5639 | 37.379 | 0.97004
2019-03-21 15:53 | 167 | 0.058397 | 0.02189 | 1.662 | 16.661 | 37.3118 | 0.97296
2019-03-21 15:53 | 168 | 0.060872 | 0.021051 | 1.5937 | 16.0642 | 37.2593 | 0.97785
2019-03-21 15:53 | 169 | 0.044136 | 0.028512 | 2.193 | 21.4006 | 37.1943 | 0.94454
2019-03-21 15:53 | 170 | 0.033 | 0.037501 | 2.9401 | 28.6032 | 37.1692 | 0.9439
2019-03-21 15:53 | 171 | 0.051788 | 0.023914 | 1.8714 | 19.1622 | 37.1211 | 0.99237
2019-03-21 15:53 | 172 | 0.050384 | 0.024617 | 1.9208 | 19.8953 | 37.0695 | 1.0024
2019-03-21 15:53 | 173 | 0.06015 | 0.021778 | 1.6061 | 15.1389 | 37.0181 | 0.9106
2019-03-21 15:53 | 174 | 0.066314 | 0.020381 | 1.4503 | 14.5208 | 36.95 | 0.96293
2019-03-21 15:53 | 175 | 0.068753 | 0.020343 | 1.4004 | 14.3129 | 36.8864 | 0.98405
2019-03-21 15:53 | 176 | 0.076541 | 0.018583 | 1.2566 | 13.6266 | 36.8114 | 1.043
2019-03-21 15:53 | 177 | 0.077799 | 0.01854 | 1.2364 | 13.5609 | 36.7347 | 1.055
2019-03-21 15:53 | 178 | 0.084436 | 0.016873 | 1.1473 | 12.3601 | 36.6548 | 1.0436
2019-03-21 15:53 | 179 | 0.08506 | 0.016456 | 1.144 | 11.6177 | 36.567 | 0.98819
2019-03-21 15:53 | 180 | 0.091026 | 0.01538 | 1.0708 | 11.6194 | 36.4837 | 1.0577
2019-03-21 15:53 | 181 | 0.091703 | 0.014988 | 1.0664 | 11.3813 | 36.3867 | 1.0437
2019-03-21 15:53 | 182 | 0.078392 | 0.016956 | 1.2523 | 13.0293 | 36.3042 | 1.0214
2019-03-21 15:53 | 183 | 0.057267 | 0.022366 | 1.7129 | 18.3693 | 36.2315 | 1.052
2019-03-21 15:53 | 184 | 0.065133 | 0.018989 | 1.5126 | 16.0853 | 36.1828 | 1.0477
2019-03-21 15:53 | 185 | 0.064734 | 0.018018 | 1.5181 | 16.3693 | 36.1029 | 1.0596
2019-03-21 15:53 | 186 | 0.045285 | 0.025289 | 2.1779 | 23.0799 | 36.0575 | 1.0452
2019-03-21 15:53 | 187 | 0.055398 | 0.020251 | 1.7808 | 19.3898 | 36.0058 | 1.0742
2019-03-21 15:53 | 188 | 0.053395 | 0.020174 | 1.8559 | 19.1202 | 35.9488 | 1.0209
2019-03-21 15:53 | 189 | 0.054543 | 0.019312 | 1.8183 | 19.2642 | 35.8983 | 1.0507
2019-03-21 15:53 | 190 | 0.062745 | 0.01642 | 1.5824 | 16.9765 | 35.8385 | 1.0652
2019-03-21 15:53 | 191 | 0.049234 | 0.02003 | 2.0216 | 21.5592 | 35.7775 | 1.0615
2019-03-21 15:53 | 192 | 0.057428 | 0.017173 | 1.7369 | 18.6125 | 35.7356 | 1.0689
2019-03-21 15:53 | 193 | 0.050155 | 0.019223 | 1.9936 | 21.2167 | 35.666 | 1.0641
2019-03-21 15:53 | 194 | 0.040019 | 0.023516 | 2.5032 | 27.2182 | 35.6354 | 1.0893
2019-03-21 15:53 | 195 | 0.067239 | 0.013664 | 1.4931 | 16.1017 | 35.5774 | 1.0827
2019-03-21 15:53 | 196 | 0.065851 | 0.01361 | 1.5276 | 16.277 | 35.5074 | 1.0719
2019-03-21 15:53 | 197 | 0.052187 | 0.016758 | 1.9339 | 20.7818 | 35.4493 | 1.0845
2019-03-21 15:53 | 198 | 0.048778 | 0.017438 | 2.0769 | 22.0695 | 35.4007 | 1.0765
2019-03-21 15:53 | 199 | 0.053267 | 0.015575 | 1.9058 | 20.218 | 35.3496 | 1.077
2019-03-21 15:53 | 200 | 0.05474 | 0.014352 | 1.8584 | 19.3219 | 35.2949 | 1.0577
2019-03-21 15:53 | 201 | 0.044748 | 0.017033 | 2.2728 | 23.302 | 35.2429 | 1.0427
2019-03-21 15:53 | 202 | 0.038892 | 0.018956 | 2.6015 | 26.7177 | 35.2038 | 1.0391
2019-03-21 15:53 | 203 | 0.066644 | 0.010739 | 1.5172 | 15.6641 | 35.157 | 1.0439
2019-03-21 15:53 | 204 | 0.080111 | 0.0089163 | 1.2549 | 12.8504 | 35.0745 | 1.0295
2019-03-21 15:53 | 205 | 0.041113 | 0.016823 | 2.4381 | 24.9602 | 35.009 | 1.0262
2019-03-21 15:53 | 206 | 0.053467 | 0.012549 | 1.8767 | 19.1298 | 34.9809 | 1.0228
2019-03-21 15:53 | 207 | 0.068398 | 0.0094793 | 1.4634 | 14.9598 | 34.9008 | 1.0232
2019-03-21 15:53 | 208 | 0.051841 | 0.012915 | 1.9199 | 19.2829 | 34.8512 | 0.99964
2019-03-21 15:53 | 209 | 0.057808 | 0.011614 | 1.7192 | 16.6442 | 34.7923 | 0.96217
2019-03-21 15:53 | 210 | 0.066617 | 0.0097382 | 1.486 | 14.9851 | 34.7343 | 0.99826
2019-03-21 15:53 | 211 | 0.077053 | 0.0084343 | 1.2815 | 12.1245 | 34.6595 | 0.93423
2019-03-21 15:53 | 212 | 0.052807 | 0.012337 | 1.8651 | 17.3325 | 34.5884 | 0.91527
2019-03-21 15:53 | 213 | 0.046931 | 0.013395 | 2.0847 | 19.9444 | 34.5498 | 0.93601
2019-03-21 15:53 | 214 | 0.055845 | 0.011264 | 1.7434 | 17.3039 | 34.4907 | 0.96634
2019-03-21 15:53 | 215 | 0.051721 | 0.011704 | 1.8678 | 17.7345 | 34.4412 | 0.91724
2019-03-21 15:53 | 216 | 0.037101 | 0.016317 | 2.5917 | 25.9536 | 34.3899 | 0.96292
2019-03-21 15:53 | 217 | 0.035948 | 0.016238 | 2.6607 | 25.444 | 34.3638 | 0.91466
2019-03-21 15:53 | 218 | 0.042911 | 0.013091 | 2.2201 | 21.0939 | 34.3162 | 0.90515
2019-03-21 15:53 | 219 | 0.028989 | 0.020177 | 3.2755 | 31.8987 | 34.2827 | 0.92472
2019-03-21 15:53 | 220 | 0.022723 | 0.025624 | 4.1289 | 40.051 | 34.2565 | 0.91006
2019-03-21 15:54 | 221 | 0.025161 | 0.023177 | 3.7331 | 36.5224 | 34.2344 | 0.91895
2019-03-21 15:54 | 222 | 0.042454 | 0.014266 | 2.1987 | 22.3637 | 34.2026 | 0.94943
2019-03-21 15:54 | 223 | 0.044343 | 0.013137 | 2.0945 | 20.3595 | 34.1532 | 0.90281
2019-03-21 15:54 | 224 | 0.040307 | 0.015012 | 2.2944 | 21.8753 | 34.1158 | 0.88172
2019-03-21 15:54 | 225 | 0.037101 | 0.016323 | 2.4833 | 23.4003 | 34.0724 | 0.86818
2019-03-21 15:54 | 226 | 0.039783 | 0.014654 | 2.3098 | 22.1005 | 34.0401 | 0.87923
2019-03-21 15:54 | 227 | 0.042128 | 0.013327 | 2.1746 | 20.1505 | 33.9927 | 0.8489
2019-03-21 15:54 | 228 | 0.042657 | 0.013156 | 2.131 | 19.3869 | 33.9556 | 0.82698
2019-03-21 15:54 | 229 | 0.067816 | 0.0082847 | 1.3357 | 12.4486 | 33.9022 | 0.84422
2019-03-21 15:54 | 230 | 0.059737 | 0.0094043 | 1.5091 | 14.4977 | 33.8278 | 0.86606
2019-03-21 15:54 | 231 | 0.050696 | 0.011536 | 1.7703 | 16.7015 | 33.7836 | 0.84669
2019-03-21 15:54 | 232 | 0.052848 | 0.011082 | 1.6892 | 15.5495 | 33.7238 | 0.82176
2019-03-21 15:54 | 233 | 0.052723 | 0.011518 | 1.686 | 15.5782 | 33.6787 | 0.82132
2019-03-21 15:54 | 234 | 0.031156 | 0.019505 | 2.8301 | 27.7289 | 33.6226 | 0.86392
2019-03-21 15:54 | 235 | 0.051344 | 0.012283 | 1.7116 | 15.9147 | 33.6065 | 0.81713
2019-03-21 15:54 | 236 | 0.070093 | 0.0093265 | 1.2457 | 12.1132 | 33.5202 | 0.84905
2019-03-21 15:54 | 237 | 0.04943 | 0.013654 | 1.7544 | 17.9625 | 33.4758 | 0.88789
2019-03-21 15:54 | 238 | 0.046904 | 0.014383 | 1.8405 | 18.2786 | 33.4184 | 0.85733
2019-03-21 15:54 | 239 | 0.027668 | 0.025177 | 3.1039 | 29.2836 | 33.3851 | 0.81021
2019-03-21 15:54 | 240 | 0.029126 | 0.024713 | 2.938 | 27.5842 | 33.3574 | 0.80341
2019-03-21 15:54 | 241 | 0.05734 | 0.012919 | 1.4857 | 14.3499 | 33.3204 | 0.82283
2019-03-21 15:54 | 242 | 0.054108 | 0.013725 | 1.5725 | 15.0431 | 33.2499 | 0.81395
2019-03-21 15:54 | 243 | 0.050894 | 0.015037 | 1.6659 | 15.896 | 33.213 | 0.80901
2019-03-21 15:54 | 244 | 0.053032 | 0.014433 | 1.5903 | 14.7527 | 33.1476 | 0.78237
2019-03-21 15:54 | 245 | 0.029277 | 0.026175 | 2.8793 | 26.7568 | 33.1124 | 0.78336
2019-03-21 15:54 | 246 | 0.035983 | 0.02127 | 2.3239 | 22.3923 | 33.0819 | 0.80575
2019-03-21 15:54 | 247 | 0.046307 | 0.01654 | 1.8064 | 16.7904 | 33.0388 | 0.77751
2019-03-21 15:54 | 248 | 0.05232 | 0.014633 | 1.5873 | 14.7014 | 32.9906 | 0.76918
2019-03-21 15:54 | 249 | 0.046453 | 0.016986 | 1.7807 | 15.313 | 32.9375 | 0.71134
2019-03-21 15:54 | 250 | 0.038596 | 0.020486 | 2.1357 | 20.1907 | 32.8987 | 0.77928
2019-03-21 15:54 | 251 | 0.023321 | 0.034839 | 3.5144 | 31.907 | 32.8616 | 0.74412
2019-03-21 15:54 | 252 | 0.027167 | 0.03073 | 3.0005 | 28.1907 | 32.8471 | 0.76585
2019-03-21 15:54 | 253 | 0.03933 | 0.0212 | 2.0646 | 19.8078 | 32.8044 | 0.77904
2019-03-21 15:54 | 254 | 0.056598 | 0.014746 | 1.4292 | 13.2205 | 32.7667 | 0.74825
2019-03-21 15:54 | 255 | 0.081399 | 0.010256 | 0.99092 | 9.2058 | 32.6903 | 0.74935
2019-03-21 15:54 | 256 | 0.064365 | 0.013001 | 1.2534 | 11.1871 | 32.6139 | 0.72006
2019-03-21 15:54 | 257 | 0.062622 | 0.013711 | 1.2791 | 11.714 | 32.5589 | 0.73356
2019-03-21 15:54 | 258 | 0.06254 | 0.013724 | 1.273 | 10.7626 | 32.4878 | 0.6731
2019-03-21 15:54 | 259 | 0.066831 | 0.01285 | 1.1882 | 11.0337 | 32.4328 | 0.7374
2019-03-21 15:54 | 260 | 0.066036 | 0.012993 | 1.1972 | 10.5612 | 32.3558 | 0.69742
2019-03-21 15:54 | 261 | 0.048758 | 0.017574 | 1.6068 | 15.9313 | 32.3046 | 0.77678
2019-03-21 15:54 | 262 | 0.052718 | 0.016223 | 1.4773 | 13.2589 | 32.2539 | 0.69898
2019-03-21 15:54 | 263 | 0.037761 | 0.023222 | 2.0576 | 18.7071 | 32.2034 | 0.7064
2019-03-21 15:54 | 264 | 0.019431 | 0.045191 | 3.9805 | 35.4267 | 32.1792 | 0.68839
2019-03-21 15:54 | 265 | 0.019558 | 0.045994 | 3.9337 | 35.9175 | 32.1598 | 0.70247
2019-03-21 15:54 | 266 | 0.025335 | 0.038174 | 3.0189 | 27.5016 | 32.1384 | 0.69676
2019-03-21 15:54 | 267 | 0.018792 | 0.053797 | 4.0484 | 38.1672 | 32.1118 | 0.71724
2019-03-21 15:54 | 268 | 0.019121 | 0.056565 | 3.9795 | 37.4727 | 32.0991 | 0.71651
2019-03-21 15:54 | 269 | 0.027134 | 0.042374 | 2.7993 | 26.684 | 32.071 | 0.72406
2019-03-21 15:54 | 270 | 0.048867 | 0.025347 | 1.5455 | 14.6634 | 32.0413 | 0.71656
2019-03-21 15:54 | 271 | 0.067344 | 0.019774 | 1.1192 | 10.8317 | 31.9745 | 0.72945
2019-03-21 15:54 | 272 | 0.057109 | 0.0245 | 1.3165 | 12.7702 | 31.9136 | 0.72929
2019-03-21 15:54 | 273 | 0.060455 | 0.02391 | 1.2434 | 12.1519 | 31.858 | 0.73464
2019-03-21 15:54 | 274 | 0.053754 | 0.027699 | 1.3922 | 13.2556 | 31.7949 | 0.71255
2019-03-21 15:54 | 275 | 0.047668 | 0.031198 | 1.5666 | 15.2805 | 31.75 | 0.7284
2019-03-21 15:54 | 276 | 0.063566 | 0.023717 | 1.1704 | 11.3646 | 31.6946 | 0.72241
2019-03-21 15:54 | 277 | 0.057206 | 0.026405 | 1.3024 | 12.4309 | 31.628 | 0.71113
2019-03-21 15:54 | 278 | 0.050195 | 0.031365 | 1.4751 | 14.5938 | 31.5811 | 0.73254
2019-03-21 15:54 | 279 | 0.040246 | 0.041272 | 1.8317 | 18.2478 | 31.5281 | 0.7344
2019-03-21 15:54 | 280 | 0.040524 | 0.044396 | 1.8165 | 18.1022 | 31.4984 | 0.73358
2019-03-21 15:54 | 281 | 0.029741 | 0.065808 | 2.4653 | 25.2371 | 31.4494 | 0.75058
2019-03-21 15:55 | 282 | 0.023804 | 0.088838 | 3.0846 | 31.9235 | 31.4376 | 0.7599
2019-03-21 15:55 | 283 | 0.025855 | 0.086138 | 2.8371 | 29.7128 | 31.3988 | 0.76821
2019-03-21 15:55 | 284 | 0.055769 | 0.042265 | 1.3113 | 13.9451 | 31.3791 | 0.7777
2019-03-21 15:55 | 285 | 0.069865 | 0.035008 | 1.0483 | 11.1796 | 31.2914 | 0.78106
2019-03-21 15:55 | 286 | 0.05 | 0.050509 | 1.4631 | 15.4483 | 31.2475 | 0.77241
2019-03-21 15:55 | 287 | 0.060435 | 0.043326 | 1.211 | 12.726 | 31.1848 | 0.7691
2019-03-21 15:55 | 288 | 0.067438 | 0.039723 | 1.0867 | 12.2249 | 31.1265 | 0.82442
2019-03-21 15:55 | 289 | 0.082044 | 0.032767 | 0.88981 | 10.6702 | 31.0484 | 0.87542
2019-03-21 15:55 | 290 | 0.080736 | 0.033231 | 0.90414 | 13.0203 | 30.9658 | 1.0512
2019-03-21 15:55 | 291 | 0.093619 | 0.028317 | 0.78205 | 12.8982 | 30.8836 | 1.2075
2019-03-21 15:55 | 292 | 0.1133 | 0.02352 | 0.65368 | 12.067 | 30.7766 | 1.3672
2019-03-21 15:55 | 293 | 0.12935 | 0.020535 | 0.58671 | 12.188 | 30.6586 | 1.5765
2019-03-21 15:55 | 294 | 0.11871 | 0.022298 | 0.66403 | 14.1608 | 30.5252 | 1.681
2019-03-21 15:55 | 295 | 0.091389 | 0.029336 | 0.90888 | 20.1346 | 30.4246 | 1.8401
2019-03-21 15:55 | 296 | 0.11892 | 0.023329 | 0.74035 | 17.5723 | 30.3292 | 2.0897
2019-03-21 15:55 | 297 | 0.14448 | 0.019889 | 0.66975 | 16.0502 | 30.1869 | 2.3189
2019-03-21 15:55 | 298 | 0.12654 | 0.023881 | 0.85105 | 20.1869 | 30.0509 | 2.5545
2019-03-21 15:55 | 299 | 0.12061 | 0.026208 | 0.98282 | 22.2657 | 29.932 | 2.6856
2019-03-21 15:55 | 300 | 0.11928 | 0.027714 | 1.089 | 23.5252 | 29.8083 | 2.8062
2019-03-21 15:55 | 301 | 0.11536 | 0.029619 | 1.2195 | 25.1017 | 29.6936 | 2.8957
2019-03-21 15:55 | 302 | 0.12251 | 0.028747 | 1.2268 | 24.3019 | 29.5756 | 2.9772
2019-03-21 15:55 | 303 | 0.10693 | 0.033475 | 1.5153 | 28.3007 | 29.4542 | 3.0261
2019-03-21 15:55 | 304 | 0.08665 | 0.041612 | 1.957 | 34.9261 | 29.3627 | 3.0263
2019-03-21 15:55 | 305 | 0.099203 | 0.036178 | 1.7836 | 30.3235 | 29.2735 | 3.0082
2019-03-21 15:55 | 306 | 0.086089 | 0.041298 | 2.1424 | 34.2849 | 29.1698 | 2.9516
2019-03-21 15:55 | 307 | 0.067069 | 0.051764 | 2.8434 | 42.7364 | 29.1022 | 2.8663
2019-03-21 15:55 | 308 | 0.092526 | 0.036647 | 2.1215 | 29.2958 | 29.0253 | 2.7106
2019-03-21 15:55 | 309 | 0.095087 | 0.034632 | 2.1102 | 26.91 | 28.9223 | 2.5588
2019-03-21 15:55 | 310 | 0.074806 | 0.042604 | 2.716 | 32.0483 | 28.8411 | 2.3974
2019-03-21 15:55 | 311 | 0.072506 | 0.04284 | 2.8194 | 30.5397 | 28.7692 | 2.2143
2019-03-21 15:55 | 312 | 0.058267 | 0.05142 | 3.5011 | 35.13 | 28.6982 | 2.0469
2019-03-21 15:55 | 313 | 0.056943 | 0.051195 | 3.5591 | 34.831 | 28.6489 | 1.9834
2019-03-21 15:55 | 314 | 0.082363 | 0.034345 | 2.4341 | 23.1826 | 28.5783 | 1.9094
2019-03-21 15:55 | 315 | 0.069197 | 0.039663 | 2.8531 | 26.9188 | 28.4932 | 1.8627
2019-03-21 15:55 | 316 | 0.053706 | 0.049828 | 3.6057 | 33.7651 | 28.4408 | 1.8134
2019-03-21 15:55 | 317 | 0.066171 | 0.039128 | 2.8701 | 26.7741 | 28.3797 | 1.7717
2019-03-21 15:55 | 318 | 0.047286 | 0.054318 | 3.933 | 36.1576 | 28.3147 | 1.7097
2019-03-21 15:55 | 319 | 0.058321 | 0.043311 | 3.1207 | 28.2867 | 28.2782 | 1.6497
2019-03-21 15:55 | 320 | 0.077927 | 0.032106 | 2.2867 | 20.919 | 28.1958 | 1.6301
2019-03-21 15:55 | 321 | 0.067923 | 0.036857 | 2.5728 | 23.9721 | 28.129 | 1.6283
2019-03-21 15:55 | 322 | 0.075611 | 0.033105 | 2.2648 | 21.5946 | 28.0567 | 1.6328
2019-03-21 15:55 | 323 | 0.073879 | 0.033914 | 2.2767 | 22.0214 | 27.9799 | 1.6269
2019-03-21 15:55 | 324 | 0.066746 | 0.038101 | 2.4739 | 24.2941 | 27.9099 | 1.6215
2019-03-21 15:55 | 325 | 0.08084 | 0.031413 | 2.0073 | 20.262 | 27.8415 | 1.638
2019-03-21 15:55 | 326 | 0.08664 | 0.029777 | 1.8459 | 18.8707 | 27.7504 | 1.635
2019-03-21 15:55 | 327 | 0.06834 | 0.037711 | 2.3101 | 23.806 | 27.6737 | 1.6269
2019-03-21 15:55 | 328 | 0.078706 | 0.03274 | 1.989 | 20.5796 | 27.6077 | 1.6198
2019-03-21 15:55 | 329 | 0.071208 | 0.036505 | 2.1775 | 22.8613 | 27.5203 | 1.6279
2019-03-21 15:55 | 330 | 0.054317 | 0.047854 | 2.8329 | 29.8947 | 27.4671 | 1.6238
2019-03-21 15:55 | 331 | 0.075418 | 0.034464 | 2.0271 | 21.6802 | 27.4032 | 1.6351
2019-03-21 15:55 | 332 | 0.063953 | 0.04101 | 2.3748 | 25.7388 | 27.3229 | 1.6461
2019-03-21 15:55 | 333 | 0.067098 | 0.039757 | 2.2511 | 25.0765 | 27.272 | 1.6826
2019-03-21 15:55 | 334 | 0.086046 | 0.031237 | 1.7491 | 19.5947 | 27.1852 | 1.686
2019-03-21 15:55 | 335 | 0.074203 | 0.037129 | 2.0259 | 23.0753 | 27.1068 | 1.7123
2019-03-21 15:55 | 336 | 0.078038 | 0.035814 | 1.9214 | 22.1501 | 27.0336 | 1.7285
2019-03-21 15:55 | 337 | 0.088848 | 0.031996 | 1.6913 | 19.7554 | 26.9491 | 1.7552
2019-03-21 15:55 | 338 | 0.08093 | 0.035877 | 1.8537 | 22.0167 | 26.8602 | 1.7818
2019-03-21 15:55 | 339 | 0.078878 | 0.037124 | 1.9086 | 22.7984 | 26.7859 | 1.7983
2019-03-21 15:55 | 340 | 0.093375 | 0.032026 | 1.6167 | 19.2264 | 26.6989 | 1.7953
2019-03-21 15:56 | 341 | 0.079352 | 0.037682 | 1.907 | 22.7573 | 26.6055 | 1.8058
2019-03-21 15:56 | 342 | 0.079558 | 0.038123 | 1.908 | 22.8907 | 26.5373 | 1.8211
2019-03-21 15:56 | 343 | 0.085462 | 0.035427 | 1.7806 | 21.2577 | 26.4453 | 1.8167
2019-03-21 15:56 | 344 | 0.064339 | 0.047749 | 2.3729 | 28.4181 | 26.3725 | 1.8284
2019-03-21 15:56 | 345 | 0.062712 | 0.049738 | 2.4475 | 29.4127 | 26.3123 | 1.8445
2019-03-21 15:56 | 346 | 0.065791 | 0.048415 | 2.3434 | 28.1764 | 26.2452 | 1.8538
2019-03-21 15:56 | 347 | 0.069938 | 0.046404 | 2.2116 | 26.8377 | 26.1798 | 1.877
2019-03-21 15:56 | 348 | 0.08885 | 0.036985 | 1.7455 | 21.8871 | 26.1024 | 1.9447
2019-03-21 15:56 | 349 | 0.079021 | 0.042488 | 1.9783 | 24.9465 | 26.009 | 1.9713
2019-03-21 15:56 | 350 | 0.066169 | 0.051663 | 2.3725 | 30.1964 | 25.9449 | 1.9981
2019-03-21 15:56 | 351 | 0.089745 | 0.038662 | 1.7687 | 22.2055 | 25.8678 | 1.9928
2019-03-21 15:56 | 352 | 0.10599 | 0.033045 | 1.5033 | 18.9063 | 25.7668 | 2.0038
2019-03-21 15:56 | 353 | 0.10123 | 0.035064 | 1.5901 | 20.2522 | 25.6613 | 2.0502
2019-03-21 15:56 | 354 | 0.096895 | 0.037356 | 1.6755 | 21.482 | 25.5647 | 2.0815
2019-03-21 15:56 | 355 | 0.099988 | 0.036727 | 1.6327 | 20.9338 | 25.4655 | 2.0931
2019-03-21 15:56 | 356 | 0.10961 | 0.034095 | 1.5018 | 19.197 | 25.363 | 2.1041
2019-03-21 15:56 | 357 | 0.11275 | 0.033659 | 1.4701 | 19.503 | 25.2482 | 2.1989
2019-03-21 15:56 | 358 | 0.099348 | 0.038889 | 1.6871 | 22.337 | 25.1417 | 2.2191
2019-03-21 15:56 | 359 | 0.092066 | 0.042533 | 1.8345 | 24.4042 | 25.0479 | 2.2468
2019-03-21 15:56 | 360 | 0.10839 | 0.036436 | 1.5725 | 21.3925 | 24.9522 | 2.3188
2019-03-21 15:56 | 361 | 0.098476 | 0.040301 | 1.7529 | 24.0438 | 24.8371 | 2.3677
2019-03-21 15:56 | 362 | 0.081834 | 0.04889 | 2.1346 | 29.178 | 24.7566 | 2.3877
2019-03-21 15:56 | 363 | 0.10133 | 0.040203 | 1.7407 | 23.8724 | 24.6646 | 2.419
2019-03-21 15:56 | 364 | 0.11849 | 0.034905 | 1.508 | 20.6753 | 24.554 | 2.4498
2019-03-21 15:56 | 365 | 0.11688 | 0.035679 | 1.5446 | 21.1119 | 24.4326 | 2.4676
2019-03-21 15:56 | 366 | 0.10653 | 0.039569 | 1.7193 | 23.3653 | 24.3228 | 2.4891
2019-03-21 15:56 | 367 | 0.1063 | 0.039898 | 1.7473 | 23.8432 | 24.2171 | 2.5346
2019-03-21 15:56 | 368 | 0.11066 | 0.038521 | 1.7013 | 23.2082 | 24.1084 | 2.5682
2019-03-21 15:56 | 369 | 0.12577 | 0.034032 | 1.5153 | 20.7552 | 23.9935 | 2.6103
2019-03-21 15:56 | 370 | 0.12233 | 0.034972 | 1.578 | 21.3783 | 23.8614 | 2.6153
2019-03-21 15:56 | 371 | 0.11612 | 0.036744 | 1.6788 | 22.6511 | 23.7498 | 2.6303
2019-03-21 15:56 | 372 | 0.12041 | 0.035542 | 1.6364 | 21.9897 | 23.6274 | 2.6478
2019-03-21 15:56 | 373 | 0.098433 | 0.043737 | 2.0258 | 27.1453 | 23.5149 | 2.672
2019-03-21 15:56 | 374 | 0.093071 | 0.046313 | 2.1586 | 29.1909 | 23.4264 | 2.7168
2019-03-21 15:56 | 375 | 0.10937 | 0.03998 | 1.8567 | 25.1076 | 23.3232 | 2.746
2019-03-21 15:56 | 376 | 0.11176 | 0.039415 | 1.8322 | 25.0666 | 23.2105 | 2.8014
2019-03-21 15:56 | 377 | 0.11252 | 0.039111 | 1.8438 | 25.1817 | 23.1006 | 2.8335
2019-03-21 15:56 | 378 | 0.10811 | 0.040746 | 1.9474 | 26.4884 | 22.9869 | 2.8637
2019-03-21 15:56 | 379 | 0.10184 | 0.043231 | 2.0926 | 28.1489 | 22.8844 | 2.8667
2019-03-21 15:56 | 380 | 0.11384 | 0.038337 | 1.8965 | 25.0613 | 22.7787 | 2.853
2019-03-21 15:56 | 381 | 0.12505 | 0.03455 | 1.7459 | 22.7577 | 22.6568 | 2.8458
2019-03-21 15:56 | 382 | 0.12265 | 0.035195 | 1.7964 | 23.1422 | 22.5321 | 2.8385
2019-03-21 15:56 | 383 | 0.11623 | 0.037199 | 1.9153 | 24.7531 | 22.4126 | 2.8771
2019-03-21 15:56 | 384 | 0.12245 | 0.035126 | 1.8322 | 23.7028 | 22.2967 | 2.9024
2019-03-21 15:56 | 385 | 0.12564 | 0.03403 | 1.7943 | 23.0476 | 22.1682 | 2.8957
2019-03-21 15:56 | 386 | 0.12789 | 0.033349 | 1.7709 | 22.5389 | 22.0462 | 2.8824
2019-03-21 15:56 | 387 | 0.11591 | 0.036927 | 1.9627 | 24.8714 | 21.9158 | 2.883
2019-03-21 15:56 | 388 | 0.11026 | 0.038461 | 2.0675 | 25.9148 | 21.8127 | 2.8574
2019-03-21 15:56 | 389 | 0.12282 | 0.034898 | 1.8667 | 23.3941 | 21.6912 | 2.8732
2019-03-21 15:56 | 390 | 0.11265 | 0.038224 | 2.044 | 25.576 | 21.5722 | 2.8811
2019-03-21 15:56 | 391 | 0.10409 | 0.041348 | 2.2212 | 27.9212 | 21.4657 | 2.9064
2019-03-21 15:56 | 392 | 0.10746 | 0.040172 | 2.1561 | 27.0627 | 21.3606 | 2.9082
2019-03-21 15:56 | 393 | 0.12425 | 0.03474 | 1.8731 | 23.4742 | 21.2476 | 2.9166
2019-03-21 15:56 | 394 | 0.12044 | 0.035657 | 1.9396 | 24.3035 | 21.117 | 2.927
2019-03-21 15:56 | 395 | 0.11043 | 0.038468 | 2.1208 | 26.6438 | 21.0084 | 2.9424
2019-03-21 15:56 | 396 | 0.11922 | 0.035138 | 1.9757 | 25.0393 | 20.8919 | 2.9851
2019-03-21 15:56 | 397 | 0.12094 | 0.033993 | 1.9601 | 25.189 | 20.7713 | 3.0464
2019-03-21 15:56 | 398 | 0.11825 | 0.034282 | 2.0156 | 26.1384 | 20.6507 | 3.0909
2019-03-21 15:56 | 399 | 0.1339 | 0.029771 | 1.7898 | 23.332 | 20.5307 | 3.1242
2019-03-21 15:56 | 400 | 0.13643 | 0.028999 | 1.7663 | 22.9814 | 20.3862 | 3.1354
2019-03-21 15:56 | 401 | 0.12401 | 0.031842 | 1.9538 | 25.2427 | 20.2616 | 3.1303
2019-03-21 15:57 | 402 | 0.1251 | 0.031687 | 1.9509 | 25.2729 | 20.135 | 3.1617
2019-03-21 15:57 | 403 | 0.13132 | 0.030365 | 1.8759 | 24.0636 | 20.0097 | 3.16
2019-03-21 15:57 | 404 | 0.13767 | 0.029392 | 1.8116 | 23.073 | 19.8727 | 3.1765
2019-03-21 15:57 | 405 | 0.12741 | 0.031821 | 1.9672 | 24.876 | 19.7388 | 3.1694
2019-03-21 15:57 | 406 | 0.10839 | 0.037817 | 2.3301 | 29.0549 | 19.6198 | 3.1491
2019-03-21 15:57 | 407 | 0.11538 | 0.035489 | 2.1975 | 27.2967 | 19.5155 | 3.1495
2019-03-21 15:57 | 408 | 0.13233 | 0.030924 | 1.9247 | 23.7447 | 19.3867 | 3.1421
2019-03-21 15:57 | 409 | 0.1178 | 0.034667 | 2.1635 | 26.5286 | 19.258 | 3.125
2019-03-21 15:57 | 410 | 0.11772 | 0.034835 | 2.1704 | 26.5232 | 19.148 | 3.1223
2019-03-21 15:57 | 411 | 0.12854 | 0.031904 | 1.9924 | 24.2606 | 19.0199 | 3.1186
2019-03-21 15:57 | 412 | 0.12766 | 0.032214 | 2.006 | 24.7054 | 18.8939 | 3.1538
2019-03-21 15:57 | 413 | 0.11243 | 0.036729 | 2.2776 | 28.3729 | 18.7679 | 3.19
2019-03-21 15:57 | 414 | 0.11554 | 0.035816 | 2.2161 | 27.8687 | 18.6644 | 3.2199
2019-03-21 15:57 | 415 | 0.13199 | 0.03111 | 1.9397 | 24.3429 | 18.5331 | 3.2131
2019-03-21 15:57 | 416 | 0.14241 | 0.02866 | 1.8035 | 23.0303 | 18.402 | 3.2797
2019-03-21 15:57 | 417 | 0.1396 | 0.029165 | 1.8539 | 23.6257 | 18.2522 | 3.2981
2019-03-21 15:57 | 418 | 0.12319 | 0.032995 | 2.1079 | 26.7794 | 18.1257 | 3.299
2019-03-21 15:57 | 419 | 0.14044 | 0.029012 | 1.8549 | 23.6504 | 17.9981 | 3.3215
2019-03-21 15:57 | 420 | 0.14411 | 0.028294 | 1.8205 | 23.1341 | 17.8478 | 3.3339
2019-03-21 15:57 | 421 | 0.1293 | 0.031492 | 2.0372 | 26.0249 | 17.7142 | 3.3651
2019-03-21 15:57 | 422 | 0.13767 | 0.029575 | 1.9245 | 24.4934 | 17.584 | 3.372
2019-03-21 15:57 | 423 | 0.1444 | 0.028004 | 1.842 | 23.3305 | 17.439 | 3.3689
2019-03-21 15:57 | 424 | 0.13655 | 0.029473 | 1.9485 | 24.6673 | 17.2984 | 3.3682
2019-03-21 15:57 | 425 | 0.14629 | 0.02741 | 1.8315 | 23.4403 | 17.1625 | 3.429
2019-03-21 15:57 | 426 | 0.13723 | 0.029017 | 1.9595 | 25.4233 | 17.0098 | 3.4889
2019-03-21 15:57 | 427 | 0.13401 | 0.029561 | 2.0191 | 26.5823 | 16.8861 | 3.5622
2019-03-21 15:57 | 428 | 0.16444 | 0.023876 | 1.6597 | 21.9037 | 16.7345 | 3.6017
2019-03-21 15:57 | 429 | 0.14894 | 0.026281 | 1.8404 | 24.4604 | 16.5678 | 3.6432
2019-03-21 15:57 | 430 | 0.13511 | 0.028902 | 2.0398 | 26.9646 | 16.4367 | 3.6431
2019-03-21 15:57 | 431 | 0.14805 | 0.026828 | 1.8809 | 24.9026 | 16.2913 | 3.6868
2019-03-21 15:57 | 432 | 0.14472 | 0.027635 | 1.9354 | 25.3428 | 16.1439 | 3.6675
2019-03-21 15:57 | 433 | 0.14355 | 0.027916 | 1.9736 | 25.5619 | 16.0014 | 3.6694
2019-03-21 15:57 | 434 | 0.15569 | 0.025857 | 1.835 | 23.526 | 15.8541 | 3.6627
2019-03-21 15:57 | 435 | 0.14232 | 0.028309 | 2.0237 | 25.7745 | 15.696 | 3.6682
2019-03-21 15:57 | 436 | 0.13416 | 0.030133 | 2.1565 | 27.2641 | 15.5682 | 3.6577
2019-03-21 15:57 | 437 | 0.15076 | 0.026749 | 1.9236 | 24.2372 | 15.4222 | 3.6539
2019-03-21 15:57 | 438 | 0.13267 | 0.03063 | 2.2006 | 27.7061 | 15.2742 | 3.6757
2019-03-21 15:57 | 439 | 0.13311 | 0.030439 | 2.1965 | 27.5485 | 15.1527 | 3.6669
2019-03-21 15:57 | 440 | 0.14953 | 0.02721 | 1.971 | 24.5788 | 15.0039 | 3.6752
2019-03-21 15:57 | 441 | 0.1512 | 0.026926 | 1.9569 | 24.0677 | 14.857 | 3.639
2019-03-21 15:57 | 442 | 0.14936 | 0.027399 | 1.9881 | 23.96 | 14.703 | 3.5786
2019-03-21 15:57 | 443 | 0.14351 | 0.028541 | 2.0751 | 24.7806 | 14.5598 | 3.5564
2019-03-21 15:57 | 444 | 0.12223 | 0.033417 | 2.4307 | 28.7421 | 14.4188 | 3.5131
2019-03-21 15:57 | 445 | 0.1383 | 0.029742 | 2.1511 | 25.4311 | 14.3061 | 3.5172
2019-03-21 15:57 | 446 | 0.16376 | 0.025113 | 1.8146 | 21.3899 | 14.1398 | 3.5029
2019-03-21 15:57 | 447 | 0.15356 | 0.026799 | 1.933 | 22.8864 | 13.9873 | 3.5144
2019-03-21 15:57 | 448 | 0.14292 | 0.028509 | 2.0731 | 24.6337 | 13.8339 | 3.5207
2019-03-21 15:57 | 449 | 0.13198 | 0.030616 | 2.2349 | 26.6674 | 13.7011 | 3.5195
2019-03-21 15:57 | 450 | 0.14185 | 0.028353 | 2.0765 | 25.0351 | 13.565 | 3.5513
2019-03-21 15:57 | 451 | 0.13943 | 0.029122 | 2.1092 | 25.8427 | 13.4199 | 3.6031
2019-03-21 15:57 | 452 | 0.13611 | 0.029847 | 2.1615 | 26.552 | 13.2865 | 3.614
2019-03-21 15:57 | 453 | 0.13708 | 0.030015 | 2.1438 | 26.4084 | 13.1465 | 3.62
2019-03-21 15:57 | 454 | 0.14466 | 0.028798 | 2.042 | 25.2414 | 13.0112 | 3.6515
2019-03-21 15:57 | 455 | 0.12883 | 0.032749 | 2.2945 | 28.1259 | 12.8627 | 3.6234
2019-03-21 15:57 | 456 | 0.1181 | 0.036361 | 2.5066 | 30.3788 | 12.7524 | 3.5878
2019-03-21 15:57 | 457 | 0.12529 | 0.034678 | 2.3745 | 28.321 | 12.6219 | 3.5484
2019-03-21 15:57 | 458 | 0.13069 | 0.033553 | 2.279 | 27.0207 | 12.5016 | 3.5314
2019-03-21 15:57 | 459 | 0.14397 | 0.030736 | 2.0698 | 24.4448 | 12.3598 | 3.5192
2019-03-21 15:57 | 460 | 0.11295 | 0.039426 | 2.6409 | 31.1904 | 12.2242 | 3.5231
2019-03-21 15:58 | 461 | 0.099616 | 0.04446 | 2.9872 | 35.2979 | 12.1298 | 3.5163
2019-03-21 15:58 | 462 | 0.11596 | 0.038047 | 2.5629 | 30.2908 | 12.0167 | 3.5125
2019-03-21 15:58 | 463 | 0.13961 | 0.031397 | 2.1234 | 24.5456 | 11.8956 | 3.4268
2019-03-21 15:58 | 464 | 0.14348 | 0.030341 | 2.0562 | 23.1733 | 11.7427 | 3.325
2019-03-21 15:58 | 465 | 0.13413 | 0.03274 | 2.1926 | 24.5571 | 11.6127 | 3.2938
2019-03-21 15:58 | 466 | 0.12513 | 0.035296 | 2.3443 | 26.262 | 11.4747 | 3.2863
2019-03-21 15:58 | 467 | 0.11844 | 0.037269 | 2.463 | 27.6399 | 11.3611 | 3.2736
2019-03-21 15:58 | 468 | 0.13729 | 0.032114 | 2.1094 | 23.8973 | 11.2317 | 3.2808
2019-03-21 15:58 | 469 | 0.13686 | 0.031883 | 2.102 | 23.8753 | 11.0905 | 3.2677
2019-03-21 15:58 | 470 | 0.14157 | 0.030607 | 2.0248 | 23.4091 | 10.9573 | 3.3141
2019-03-21 15:58 | 471 | 0.14518 | 0.029367 | 1.9707 | 23.1671 | 10.808 | 3.3635
2019-03-21 15:58 | 472 | 0.13552 | 0.030901 | 2.1057 | 25.2408 | 10.6704 | 3.4205
2019-03-21 15:58 | 473 | 0.12435 | 0.033419 | 2.296 | 27.7732 | 10.5377 | 3.4535
2019-03-21 15:58 | 474 | 0.11373 | 0.036506 | 2.5204 | 30.3588 | 10.4213 | 3.4528
2019-03-21 15:58 | 475 | 0.11521 | 0.036093 | 2.4931 | 29.3657 | 10.3069 | 3.3832
2019-03-21 15:58 | 476 | 0.12422 | 0.033789 | 2.3142 | 26.7051 | 10.1889 | 3.3174
2019-03-21 15:58 | 477 | 0.12302 | 0.03447 | 2.3355 | 26.1999 | 10.061 | 3.2232
2019-03-21 15:58 | 478 | 0.11525 | 0.037073 | 2.4947 | 27.7721 | 9.9448 | 3.2008
2019-03-21 15:58 | 479 | 0.11018 | 0.038901 | 2.5956 | 29.0065 | 9.8304 | 3.1961
2019-03-21 15:58 | 480 | 0.087344 | 0.049278 | 3.2524 | 36.5301 | 9.7273 | 3.1907
2019-03-21 15:58 | 481 | 0.11969 | 0.035737 | 2.357 | 26.472 | 9.6431 | 3.1685
2019-03-21 15:58 | 482 | 0.12042 | 0.035272 | 2.3223 | 26.1776 | 9.4946 | 3.1524
2019-03-21 15:58 | 483 | 0.10572 | 0.040356 | 2.6298 | 29.5952 | 9.4057 | 3.1287
2019-03-21 15:58 | 484 | 0.14019 | 0.030213 | 1.9676 | 22.2907 | 9.2725 | 3.125
2019-03-21 15:58 | 485 | 0.13338 | 0.031616 | 2.059 | 22.9018 | 9.1345 | 3.0547
2019-03-21 15:58 | 486 | 0.11977 | 0.035111 | 2.2854 | 25.3122 | 9.008 | 3.0317
2019-03-21 15:58 | 487 | 0.12787 | 0.032945 | 2.1316 | 23.7252 | 8.8902 | 3.0337
2019-03-21 15:58 | 488 | 0.12273 | 0.034137 | 2.207 | 24.4141 | 8.7551 | 2.9964
2019-03-21 15:58 | 489 | 0.11057 | 0.038067 | 2.4338 | 27.0139 | 8.6463 | 2.9869
2019-03-21 15:58 | 490 | 0.11999 | 0.03508 | 2.2302 | 24.4035 | 8.5294 | 2.9281
2019-03-21 15:58 | 491 | 0.10487 | 0.040139 | 2.5347 | 27.6037 | 8.4112 | 2.8947
2019-03-21 15:58 | 492 | 0.11602 | 0.036493 | 2.2786 | 24.8317 | 8.314 | 2.881
2019-03-21 15:58 | 493 | 0.12039 | 0.035139 | 2.1786 | 23.7843 | 8.1815 | 2.8634
2019-03-21 15:58 | 494 | 0.075035 | 0.056514 | 3.4781 | 38.143 | 8.0841 | 2.862
2019-03-21 15:58 | 495 | 0.095139 | 0.04409 | 2.719 | 29.9349 | 8.0164 | 2.848
2019-03-21 15:58 | 496 | 0.11196 | 0.036714 | 2.2925 | 25.2159 | 7.8942 | 2.8232
2019-03-21 15:58 | 497 | 0.079108 | 0.051559 | 3.2262 | 35.405 | 7.804 | 2.8008
2019-03-21 15:58 | 498 | 0.080149 | 0.050475 | 3.1667 | 34.9658 | 7.7283 | 2.8025
2019-03-21 15:58 | 499 | 0.09613 | 0.041792 | 2.6296 | 29.2545 | 7.6389 | 2.8122
2019-03-21 15:58 | 500 | 0.11764 | 0.034003 | 2.1354 | 23.755 | 7.535 | 2.7945
2019-03-21 15:58 | 501 | 0.11348 | 0.035383 | 2.2066 | 24.5364 | 7.4107 | 2.7844
2019-03-21 15:58 | 502 | 0.083451 | 0.048431 | 2.9864 | 33.1583 | 7.3145 | 2.7671
2019-03-21 15:58 | 503 | 0.082598 | 0 | 0 | 0 | 7.2368 | 2.708
2019-03-21 15:58 | 504 | 0.096229 | 0 | 0 | 0 | 7.1452 | 2.6656
2019-03-21 15:58 | 505 | 0.096397 | 0 | 0 | 0 | 7.0472 | 2.6032
2019-03-21 15:58 | 506 | 0.097152 | 0 | 0 | 0 | 6.952 | 2.528
2019-03-21 15:58 | 507 | 0.11989 | 0 | 0 | 0 | 6.8481 | 2.5159
2019-03-21 15:58 | 508 | 0.11721 | 0 | 0 | 0 | 6.7184 | 2.4956
2019-03-21 15:58 | 509 | 0.099238 | 0 | 0 | 0 | 6.6175 | 2.4829
2019-03-21 15:58 | 510 | 0.10979 | 0 | 0 | 0 | 6.5135 | 2.4781
2019-03-21 15:58 | 511 | 0.10985 | 0 | 0 | 0 | 6.3998 | 2.449
2019-03-21 15:58 | 512 | 0.10932 | 0 | 0 | 0 | 6.2943 | 2.3963
2019-03-21 15:58 | 513 | 0.10852 | 0 | 0 | 0 | 6.1821 | 2.3767
2019-03-21 15:58 | 514 | 0.078996 | 0 | 0 | 0 | 6.0839 | 2.372
2019-03-21 15:58 | 515 | 0.074165 | 0 | 0 | 0 | 6.0183 | 2.2921
2019-03-21 15:58 | 516 | 0.081684 | 0 | 0 | 0 | 5.9323 | 2.3121
2019-03-21 15:58 | 517 | 0.065488 | 0 | 0 | 0 | 5.8596 | 2.2962
2019-03-21 15:58 | 518 | 0.080461 | 0 | 0 | 0 | 5.7944 | 2.2705
2019-03-21 15:58 | 519 | 0.087222 | 0 | 0 | 0 | 5.7012 | 2.2608
2019-03-21 15:58 | 520 | 0.04888 | 0 | 0 | 0 | 5.6307 | 2.2216
2019-03-21 15:58 | 521 | 0.041441 | 0 | 0 | 0 | 5.596 | 2.1777
2019-03-21 15:59 | 522 | 0.066289 | 0 | 0 | 0 | 5.5393 | 2.137
2019-03-21 15:59 | 523 | 0.06863 | 0 | 0 | 0 | 5.4679 | 2.1205
2019-03-21 15:59 | 524 | 0.079333 | 0 | 0 | 0 | 5.4009 | 2.0852
2019-03-21 15:59 | 525 | 0.076348 | 0 | 0 | 0 | 5.3123 | 2.0311
2019-03-21 15:59 | 526 | 0.079845 | 0 | 0 | 0 | 5.247 | 1.9823
2019-03-21 15:59 | 527 | 0.087179 | 0 | 0 | 0 | 5.1526 | 1.9283
2019-03-21 15:59 | 528 | 0.05588 | 0 | 0 | 0 | 5.0811 | 1.9705
2019-03-21 15:59 | 529 | 0.063959 | 0 | 0 | 0 | 5.0316 | 1.9195
2019-03-21 15:59 | 530 | 0.082724 | 0 | 0 | 0 | 4.95 | 1.8867
2019-03-21 15:59 | 531 | 0.080495 | 0 | 0 | 0 | 4.871 | 1.8626
2019-03-21 15:59 | 532 | 0.085003 | 0 | 0 | 0 | 4.7884 | 1.8222
2019-03-21 15:59 | 533 | 0.071224 | 0 | 0 | 0 | 4.7048 | 1.795
2019-03-21 15:59 | 534 | 0.080885 | 0 | 0 | 0 | 4.6407 | 1.7805
2019-03-21 15:59 | 535 | 0.085065 | 0 | 0 | 0 | 4.5445 | 1.7683
2019-03-21 15:59 | 536 | 0.064071 | 0 | 0 | 0 | 4.4765 | 1.7221
2019-03-21 15:59 | 537 | 0.064493 | 0 | 0 | 0 | 4.4116 | 1.7145
2019-03-21 15:59 | 538 | 0.064836 | 0 | 0 | 0 | 4.3468 | 1.6527
2019-03-21 15:59 | 539 | 0.069099 | 0 | 0 | 0 | 4.2809 | 1.6141
2019-03-21 15:59 | 540 | 0.075137 | 0 | 0 | 0 | 4.2088 | 1.57
2019-03-21 15:59 | 541 | 0.058585 | 0 | 0 | 0 | 4.1365 | 1.5451
2019-03-21 15:59 | 542 | 0.035404 | 0 | 0 | 0 | 4.0931 | 1.5451
2019-03-21 15:59 | 543 | 0.042933 | 0 | 0 | 0 | 4.0582 | 1.5436
2019-03-21 15:59 | 544 | 0.04563 | 0 | 0 | 0 | 4.0075 | 1.5185
2019-03-21 15:59 | 545 | 0.046272 | 0 | 0 | 0 | 3.9675 | 1.479
2019-03-21 15:59 | 546 | 0.048598 | 0 | 0 | 0 | 3.9147 | 1.4434
2019-03-21 15:59 | 547 | 0.04729 | 0 | 0 | 0 | 3.8711 | 1.4159
2019-03-21 15:59 | 548 | 0.054863 | 0 | 0 | 0 | 3.8191 | 1.4008
2019-03-21 15:59 | 549 | 0.022785 | 0 | 0 | 0 | 3.7701 | 1.3731
2019-03-21 15:59 | 550 | 0.021757 | 0 | 0 | 0 | 3.7658 | 1.3384
2019-03-21 15:59 | 551 | 0.057909 | 0 | 0 | 0 | 3.7176 | 1.295
2019-03-21 15:59 | 552 | 0.052019 | 0 | 0 | 0 | 3.6598 | 1.2768
2019-03-21 15:59 | 553 | 0.03448 | 0 | 0 | 0 | 3.6169 | 1.234
2019-03-21 15:59 | 554 | 0.042943 | 0 | 0 | 0 | 3.5848 | 1.2194
2019-03-21 15:59 | 555 | 0.045853 | 0 | 0 | 0 | 3.5319 | 1.2297
2019-03-21 15:59 | 556 | 0.041109 | 0 | 0 | 0 | 3.4952 | 1.2052
2019-03-21 15:59 | 557 | 0.037793 | 0 | 0 | 0 | 3.4499 | 1.2094
2019-03-21 15:59 | 558 | 0.026698 | 0 | 0 | 0 | 3.4214 | 1.1779
2019-03-21 15:59 | 559 | 0.02041 | 0 | 0 | 0 | 3.3953 | 1.1521
2019-03-21 15:59 | 560 | 0.021699 | 0 | 0 | 0 | 3.3788 | 1.1127
2019-03-21 15:59 | 561 | 0.0099383 | 0 | 0 | 0 | 3.3544 | 1.0857
2019-03-21 15:59 | 562 | 0.014432 | 0 | 0 | 0 | 3.3546 | 1.0597
2019-03-21 15:59 | 563 | 0.042586 | 0 | 0 | 0 | 3.3202 | 1.0101
2019-03-21 15:59 | 564 | 0.030323 | 0 | 0 | 0 | 3.2791 | 0.97606
2019-03-21 15:59 | 565 | 0.011785 | 0 | 0 | 0 | 3.2617 | 0.91076
2019-03-21 15:59 | 566 | 0.017021 | 0 | 0 | 0 | 3.2501 | 0.86321
2019-03-21 15:59 | 567 | 0.0071393 | 0 | 0 | 0 | 3.23 | 0.83968
2019-03-21 15:59 | 568 | 0.02373 | 0 | 0 | 0 | 3.2294 | 0.82933
2019-03-21 15:59 | 569 | 0.046169 | 0 | 0 | 0 | 3.1816 | 0.74523
2019-03-21 15:59 | 570 | 0.021627 | 0 | 0 | 0 | 3.1482 | 0.68039
2019-03-21 15:59 | 571 | 0.018089 | 0 | 0 | 0 | 3.1342 | 0.58157
2019-03-21 15:59 | 572 | 0.027815 | 0 | 0 | 0 | 3.1091 | 0.5358
2019-03-21 15:59 | 573 | 0.0026002 | 0 | 0 | 0 | 3.0862 | 0.5058
2019-03-21 15:59 | 574 | 0.0017044 | 0 | 0 | 0 | 3.0983 | 0.47283
2019-03-21 15:59 | 575 | 0.015158 | 0 | 0 | 0 | 3.0789 | 0.47092
2019-03-21 15:59 | 576 | 0.014782 | 0 | 0 | 0 | 3.071 | 0.44477
2019-03-21 15:59 | 577 | 0.017533 | 0 | 0 | 0 | 3.0496 | 0.42793
2019-03-21 15:59 | 578 | -0.0029803 | 0 | 0 | 0 | 3.0414 | 0.38194
2019-03-21 15:59 | 579 | -0.010997 | 0 | 0 | 0 | 3.0524 | 0.38284
2019-03-21 15:59 | 580 | 0.0032934 | 0 | 0 | 0 | 3.0575 | 0.34933
2019-03-21 15:59 | 581 | 0.01454 | 0 | 0 | 0 | 3.0463 | 0.3472
2019-03-21 16:00 | 582 | 0.0090934 | 0 | 0 | 0 | 3.0321 | 0.33515
2019-03-21 16:00 | 583 | 0.023627 | 0 | 0 | 0 | 3.0243 | 0.32119
2019-03-21 16:00 | 584 | 0.015233 | 0 | 0 | 0 | 2.9907 | 0.31226
2019-03-21 16:00 | 585 | -0.011046 | 0 | 0 | 0 | 2.9983 | 0.29052
2019-03-21 16:00 | 586 | -0.014171 | 0 | 0 | 0 | 3.0069 | 0.28418
2019-03-21 16:00 | 587 | 0.0008415 | 0 | 0 | 0 | 3.021 | 0.27512
2019-03-21 16:00 | 588 | 0.028084 | 0 | 0 | 0 | 3.0024 | 0.28015
2019-03-21 16:00 | 589 | 0.025715 | 0 | 0 | 0 | 2.9723 | 0.25961
2019-03-21 16:00 | 590 | 0.0099331 | 0 | 0 | 0 | 2.9547 | 0.26368
2019-03-21 16:00 | 591 | 0.013846 | 0 | 0 | 0 | 2.9482 | 0.26673
2019-03-21 16:00 | 592 | 0.011016 | 0 | 0 | 0 | 2.9285 | 0.24962
2019-03-21 16:00 | 593 | -0.0065328 | 0 | 0 | 0 | 2.929 | 0.25993
2019-03-21 16:00 | 594 | 0.0091301 | 0 | 0 | 0 | 2.9335 | 0.24319
2019-03-21 16:00 | 595 | 0.026436 | 0 | 0 | 0 | 2.9101 | 0.2128
2019-03-21 16:00 | 596 | 0.019504 | 0 | 0 | 0 | 2.8869 | 0.23378
2019-03-21 16:00 | 597 | 0.0088743 | 0 | 0 | 0 | 2.8729 | 0.22615
2019-03-21 16:00 | 598 | -0.0049518 | 0 | 0 | 0 | 2.8697 | 0.22681
2019-03-21 16:00 | 599 | -0.0037008 | 0 | 0 | 0 | 2.8789 | 0.20845
2019-03-21 16:00 | 600 | 0.0029343 | 0 | 0 | 0 | 2.8758 | 0.21674
2019-03-21 16:00 | 601 | -0.0097288 | 0 | 0 | 0 | 2.8773 | 0.19868
2019-03-21 16:00 | 602 | -0.0077272 | 0 | 0 | 0 | 2.8914 | 0.19364
2019-03-21 16:00 | 603 | 0.0201 | 0 | 0 | 0 | 2.8872 | 0.19923
2019-03-21 16:00 | 604 | -0.00034858 | 0 | 0 | 0 | 2.8623 | 0.2017
2019-03-21 16:00 | 605 | -0.0096245 | 0 | 0 | 0 | 2.8858 | 0.17872
2019-03-21 16:00 | 606 | 8.3167e-05 | 0 | 0 | 0 | 2.8766 | 0.19126
2019-03-21 16:00 | 607 | 0.0060212 | 0 | 0 | 0 | 2.8862 | 0.19994

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062478 | 0.062678 | 0 | 0.2095 | 2019-03-21 09:44 | Group A | 0.1 | 0.22 | Do Nothing | 0.20947 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025146 | 0.025271 | 0 | 0.0261 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | 0.00046112 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002726 | 0.002739 | 0 | 0.0011 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | -2.9489e-06 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.6721 | 183.6734 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 99.3463 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4138 | 209.9622 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 99.2818 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -16.237 | 1447 | 0 |  | 2019-03-21 09:27 | None | -10 | 1000 | Do Nothing | 123.9589 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2612 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018798 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017579 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -152.1519 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.24092 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4022 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0036254 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0049378 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.004768 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 5.1102e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011864 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.26503 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3284 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5479 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 98.92 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.878 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 38.5564 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 67.059 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 109.137 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.1261 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0096321 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0468 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.02492 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.807 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN