## Cone Summary
Date: | 2019-03-29
Time: | 11:20
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | +PP2-B2-117up-B2_75-r3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0407 |  | Start Test | 0 | 2019-03-29 11:35
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 6 | 2019-03-29 11:35
Sample Area | 1 | m^2 | Wetting | 158 | 2019-03-29 11:38
Cone Heat Flux | 75 | kW/m^2 | Dripping | 175 | 2019-03-29 11:38
Cone Orientation |  |  | Bottom Ignition | 287 | 2019-03-29 11:40
Sample Grid |  |  | Flame Out | 547 | 2019-03-29 11:44
Sample Thickness | 106 | mm
Sample Initial Mass | 47.7 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 23.7876 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.4896 | kW/m2 | Average Heat Release Rate | 2.0272 | kW/m2
Total Heat Release | 1.0514 | MJ/m2 | Average Mass Loss Rate | 0.09017 | g/m2s
Peak Mass Loss Rate | 0.18765 | g/m2s | Effective Heat of Combustion | 22.5254 | MJ/kg
Total Mass Lost | 44.6608 | g | Average Specific Extinction Area | 75.3997 | m2/kg
Time to Ignition | 6 | s | Peak CO Yield | 0.08327 | kg/kg
Time to Peak Heat Release | 414 | s | Peak CO2 Yield | 5.9233 | kg/kg
Time to Flame Out | 547 | s | Average CO Yield | 0.024686 | kg/kg
Average Heat Release Rate | 1.9398 | kW/m2 | Average CO2 Yield | 1.9172 | kg/kg
Average Mass Loss Rate | 0.082412 | g/m2s
Effective Heat of Combustion | 23.5416 | MJ/kg
Dripping | 175 | s | Time to dripping
HRRD | 0.93759 | kW | HRR at dripping time
THRD | 0.21053 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.2458 | kW | AHRR from Ignition to Dripping
Wetting | 158 | s | Time to wetting
HRRW | 0.94804 | kW | HRR at wetting time
THRW | 0.19457 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.2468 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-29 11:20 | 1553872831.8475 | -888 | 0.20827 | 0.000516 | 0 | 101.7827 | 101.87 | 129.237 | 20.8308 | 0.00155 | 0.001431 | -152.1124 | 0.22496 | 8.4027 | 0.002076 | 0.003091 | 0.002812 | 7e-05 | 0.000619 | 0.26648 | 1.3355 | 2.5472 | 100.3 | 20.83 | 31.831 | 58.7091 | 71.3643 | 55.8272 | 0.007819 | 1.0732 | 0.02551 | 23.7706 | -0.007669 | 0.13383 | -8.1e-05 | -0.006162 | -0.006162 | -0.884
2019-03-29 11:21 | 1553872832.8474 | -887 | 0.20827 | 0.000516 | 0 | 101.7827 | 101.87 | 129.237 | 20.8308 | 0.00155 | 0.001431 | -152.1124 | 0.22496 | 8.4027 | 0.002076 | 0.003091 | 0.002812 | 7e-05 | 0.000619 | 0.26648 | 1.3355 | 2.5472 | 100.3 | 20.83 | 31.83 | 58.7091 | 71.3643 | 55.8272 | 0.007819 | 1.0732 | 0.02551 | 23.7706 | -0.007669 | 0.13383 | -8.1e-05 | -0.006161 | -0.006161 | -0.88411
2019-03-29 11:21 | 1553872833.8475 | -886 | 0.20824 | 0.00052 | 0 | 101.7794 | 101.8769 | 129.237 | 20.8313 | 0.001519 | 0.001436 | -152.1124 | 0.22582 | 8.4018 | 0.002155 | 0.003247 | 0.003114 | 7e-05 | 0.000618 | 0.26305 | 1.3532 | 2.5473 | 100.3 | 20.83 | 31.83 | 57.9544 | 73.2483 | 55.9428 | 0.007819 | 1.0728 | 0.025505 | 23.7748 | -0.008557 | 0.13434 | -8.7e-05 | -0.006592 | -0.006592 | -0.82948
2019-03-29 11:21 | 1553872834.8476 | -885 | 0.20824 | 0.000518 | 0 | 101.8462 | 101.8666 | 129.237 | 20.8319 | 0.001551 | 0.001356 | -152.1124 | 0.22548 | 8.4009 | 0.002169 | 0.003654 | 0.003439 | 6.9e-05 | 0.000614 | 0.25393 | 1.3327 | 2.5474 | 100.3 | 20.83 | 31.831 | 58.7265 | 73.5989 | 54.0157 | 0.007819 | 1.0791 | 0.02558 | 23.7051 | -0.001794 | 0.13414 | -8.3e-05 | -0.006329 | -0.006329 | -0.86264
2019-03-29 11:21 | 1553872835.847 | -884 | 0.20825 | 0.000519 | 0 | 101.7908 | 101.9096 | 129.237 | 20.8316 | 0.001547 | 0.001376 | -152.1124 | 0.22497 | 8.4009 | 0.002066 | 0.003614 | 0.003394 | 7.3e-05 | 0.000614 | 0.27273 | 1.3423 | 2.5473 | 100.3 | 20.83 | 31.831 | 58.6206 | 71.1156 | 54.5041 | 0.007819 | 1.0775 | 0.025561 | 23.7228 | -0.010434 | 0.13384 | -8.5e-05 | -0.006459 | -0.006459 | -0.84336
2019-03-29 11:21 | 1553872836.8466 | -883 | 0.20823 | 0.000519 | -1e-06 | 101.8588 | 101.8267 | 129.237 | 20.8317 | 0.00155 | 0.001452 | -152.1124 | 0.2255 | 8.4039 | 0.002151 | 0.003329 | 0.003195 | 7.3e-05 | 0.000612 | 0.2521 | 1.3677 | 2.5474 | 100.3 | 20.83 | 31.833 | 58.6965 | 73.1521 | 56.3269 | 0.00782 | 1.0715 | 0.02549 | 23.7887 | 0.002817 | 0.13415 | -8.4e-05 | -0.006348 | -0.006348 | -0.86013
2019-03-29 11:21 | 1553872837.8467 | -882 | 0.20826 | 0.00052 | -1e-06 | 101.8109 | 101.7909 | 129.237 | 20.8322 | 0.001535 | 0.001412 | -152.1124 | 0.22629 | 8.4047 | 0.002119 | 0.00347 | 0.003274 | 7.4e-05 | 0.00062 | 0.25608 | 1.3515 | 2.5474 | 100.3 | 20.84 | 31.835 | 58.3308 | 72.3791 | 55.359 | 0.007825 | 1.0747 | 0.025528 | 23.7537 | 0.001759 | 0.13462 | -8.5e-05 | -0.006467 | -0.006467 | -0.84719
2019-03-29 11:21 | 1553872838.847 | -881 | 0.20827 | 0.00052 | 0 | 101.8016 | 102.0243 | 129.237 | 20.8328 | 0.001553 | 0.001411 | -152.1124 | 0.22553 | 8.4023 | 0.002089 | 0.003225 | 0.003062 | 6.9e-05 | 0.000621 | 0.26233 | 1.3363 | 2.5474 | 100.3 | 20.84 | 31.836 | 58.7773 | 71.6812 | 55.3557 | 0.007825 | 1.0747 | 0.025528 | 23.7536 | -0.019546 | 0.13417 | -8.5e-05 | -0.006479 | -0.006479 | -0.84283
2019-03-29 11:21 | 1553872839.8458 | -880 | 0.20826 | 0.000518 | -1e-06 | 101.7636 | 101.725 | 129.237 | 20.8338 | 0.001516 | 0.001476 | -152.1124 | 0.22634 | 8.4038 | 0.002097 | 0.003632 | 0.003379 | 7.1e-05 | 0.000626 | 0.24419 | 1.3318 | 2.5474 | 100.3 | 20.84 | 31.841 | 57.8854 | 71.86 | 56.9248 | 0.007827 | 1.0696 | 0.025467 | 23.8102 | 0.003393 | 0.13465 | -8.3e-05 | -0.00626 | -0.00626 | -0.87542
2019-03-29 11:21 | 1553872840.8468 | -879 | 0.20823 | 0.000519 | -1e-06 | 101.8016 | 101.8499 | 129.237 | 20.8337 | 0.001529 | 0.001457 | -152.1124 | 0.22554 | 8.3994 | 0.002098 | 0.003958 | 0.00381 | 6.9e-05 | 0.000624 | 0.25658 | 1.3425 | 2.5474 | 100.3 | 20.84 | 31.844 | 58.1908 | 71.8787 | 56.459 | 0.007827 | 1.0711 | 0.025485 | 23.7934 | -0.004243 | 0.13417 | -8.3e-05 | -0.006283 | -0.006283 | -0.86908
2019-03-29 11:21 | 1553872841.8468 | -878 | 0.20824 | 0.000517 | 0 | 101.6708 | 101.8472 | 129.237 | 20.8332 | 0.001528 | 0.001443 | -152.1124 | 0.22625 | 8.4019 | 0.002048 | 0.004149 | 0.003875 | 6.9e-05 | 0.000623 | 0.26432 | 1.3163 | 2.5474 | 100.3 | 20.84 | 31.847 | 58.1784 | 70.6774 | 56.1137 | 0.007828 | 1.0722 | 0.025499 | 23.781 | -0.015504 | 0.13459 | -8.2e-05 | -0.006262 | -0.006262 | -0.87485
2019-03-29 11:21 | 1553872842.8458 | -877 | 0.20823 | 0.000514 | 0 | 101.8537 | 101.9472 | 129.0474 | 20.8327 | 0.001581 | 0.001371 | -152.1262 | 0.22459 | 8.4021 | 0.002066 | 0.003442 | 0.003241 | 7e-05 | 0.000624 | 0.23308 | 1.3399 | 2.5474 | 100.3 | 20.84 | 31.851 | 59.4429 | 71.1183 | 54.3844 | 0.007829 | 1.0779 | 0.025547 | 23.701 | -0.008206 | 0.13361 | -7.8e-05 | -0.005947 | -0.005947 | -0.91439
2019-03-29 11:21 | 1553872843.8459 | -876 | 0.20821 | 0.000517 | -1e-06 | 101.7627 | 101.8137 | 128.566 | 20.8335 | 0.001494 | 0.001403 | -152.1417 | 0.22544 | 8.3996 | 0.002074 | 0.002859 | 0.002709 | 7.2e-05 | 0.00062 | 0.25011 | 1.3755 | 2.5475 | 100.3 | 20.84 | 31.852 | 57.3593 | 71.2992 | 55.1637 | 0.007829 | 1.0753 | 0.025469 | 23.6849 | -0.004485 | 0.13411 | -8.1e-05 | -0.006134 | -0.006134 | -0.88993
2019-03-29 11:21 | 1553872844.8459 | -875 | 0.20822 | 0.000518 | 0 | 101.7504 | 101.867 | 128.3686 | 20.8331 | 0.001582 | 0.001402 | -152.1307 | 0.2256 | 8.4014 | 0.002082 | 0.002383 | 0.002182 | 7.4e-05 | 0.000624 | 0.25962 | 1.3259 | 2.5474 | 100.3 | 20.84 | 31.854 | 59.4771 | 71.4953 | 55.1179 | 0.00783 | 1.0755 | 0.025451 | 23.665 | -0.01024 | 0.13421 | -8.3e-05 | -0.006319 | -0.006319 | -0.86441
2019-03-29 11:21 | 1553872845.8459 | -874 | 0.20822 | 0.000517 | -1e-06 | 101.7992 | 101.7936 | 128.1662 | 20.8338 | 0.001496 | 0.001417 | -152.1225 | 0.22498 | 8.4024 | 0.002071 | 0.003647 | 0.003345 | 7.1e-05 | 0.000625 | 0.24563 | 1.3749 | 2.5474 | 100.3 | 20.84 | 31.858 | 57.3901 | 71.248 | 55.504 | 0.007831 | 1.0742 | 0.025416 | 23.6603 | 0.000493 | 0.13384 | -8.1e-05 | -0.006142 | -0.006142 | -0.88686
2019-03-29 11:21 | 1553872846.846 | -873 | 0.20821 | 0.000519 | -1e-06 | 101.7776 | 101.856 | 128.1388 | 20.8343 | 0.001549 | 0.001418 | -152.122 | 0.22523 | 8.4042 | 0.002057 | 0.002745 | 0.002583 | 7.4e-05 | 0.000623 | 0.23922 | 1.3789 | 2.5476 | 100.3 | 20.84 | 31.859 | 58.6788 | 70.8921 | 55.521 | 0.007831 | 1.0742 | 0.025413 | 23.6584 | -0.006887 | 0.13399 | -8.4e-05 | -0.006367 | -0.006367 | -0.85653
2019-03-29 11:21 | 1553872847.847 | -872 | 0.20825 | 0.000517 | 0 | 101.6537 | 101.869 | 128.0376 | 20.835 | 0.001553 | 0.001459 | -152.1317 | 0.22545 | 8.4036 | 0.002083 | 0.003578 | 0.003347 | 7.3e-05 | 0.000622 | 0.23612 | 1.3671 | 2.5474 | 100.3 | 20.84 | 31.862 | 58.7858 | 71.5361 | 56.5063 | 0.007832 | 1.071 | 0.025365 | 23.6845 | -0.018923 | 0.13412 | -8.3e-05 | -0.006267 | -0.006267 | -0.87106
2019-03-29 11:21 | 1553872848.8459 | -871 | 0.20825 | 0.000519 | 0 | 101.7103 | 101.6475 | 128.0188 | 20.834 | 0.001553 | 0.001413 | -152.1195 | 0.22567 | 8.4033 | 0.00207 | 0.002862 | 0.002686 | 7.5e-05 | 0.00062 | 0.25984 | 1.338 | 2.5473 | 100.3 | 20.84 | 31.863 | 58.7716 | 71.2148 | 55.3992 | 0.007832 | 1.0746 | 0.025406 | 23.6429 | 0.005523 | 0.13425 | -8.4e-05 | -0.006363 | -0.006363 | -0.85877
2019-03-29 11:21 | 1553872849.8461 | -870 | 0.20826 | 0.000516 | -1e-06 | 101.7025 | 101.7407 | 127.9198 | 20.8349 | 0.001527 | 0.001442 | -152.1129 | 0.22527 | 8.4024 | 0.002063 | 0.003646 | 0.003452 | 7.2e-05 | 0.000622 | 0.24298 | 1.3473 | 2.5474 | 100.3 | 20.84 | 31.866 | 58.156 | 71.0472 | 56.1038 | 0.007833 | 1.0723 | 0.025369 | 23.6591 | -0.003357 | 0.13401 | -8e-05 | -0.006058 | -0.006058 | -0.90031
2019-03-29 11:21 | 1553872850.847 | -869 | 0.20826 | 0.000519 | 0 | 101.6765 | 101.7014 | 127.7979 | 20.8362 | 0.001511 | 0.001411 | -152.1118 | 0.22612 | 8.4016 | 0.002091 | 0.003603 | 0.003426 | 7e-05 | 0.000622 | 0.25354 | 1.3184 | 2.5473 | 100.3 | 20.84 | 31.869 | 57.758 | 71.7158 | 55.3489 | 0.007834 | 1.0747 | 0.025386 | 23.6207 | -0.002196 | 0.13452 | -8.4e-05 | -0.006376 | -0.006376 | -0.8587
2019-03-29 11:21 | 1553872851.8471 | -868 | 0.20826 | 0.000514 | 0 | 101.6603 | 101.5789 | 127.8379 | 20.8361 | 0.001494 | 0.001455 | -152.1117 | 0.22505 | 8.404 | 0.00197 | 0.003609 | 0.003447 | 7.5e-05 | 0.000625 | 0.26648 | 1.3192 | 2.5473 | 100.3 | 20.84 | 31.871 | 57.3453 | 68.811 | 56.4128 | 0.007834 | 1.0713 | 0.025349 | 23.6626 | 0.007164 | 0.13388 | -7.8e-05 | -0.00591 | -0.00591 | -0.92207
2019-03-29 11:21 | 1553872852.8462 | -867 | 0.20826 | 0.000515 | 0 | 101.6998 | 101.5932 | 127.6165 | 20.8354 | 0.001513 | 0.001462 | -152.106 | 0.22485 | 8.4015 | 0.002049 | 0.003525 | 0.003323 | 7.3e-05 | 0.000618 | 0.26665 | 1.3243 | 2.5475 | 100.3 | 20.84 | 31.872 | 57.8001 | 70.7014 | 56.5687 | 0.007834 | 1.0707 | 0.025321 | 23.6477 | 0.009378 | 0.13377 | -8e-05 | -0.006024 | -0.006024 | -0.90372
2019-03-29 11:21 | 1553872853.8466 | -866 | 0.20823 | 0.000515 | 0 | 101.6498 | 101.6349 | 127.4393 | 20.8361 | 0.001498 | 0.001452 | -152.0956 | 0.22522 | 8.4048 | 0.00206 | 0.003179 | 0.002952 | 7.4e-05 | 0.000616 | 0.26034 | 1.3269 | 2.5474 | 100.3 | 20.84 | 31.872 | 57.4479 | 70.989 | 56.3454 | 0.007834 | 1.0715 | 0.025312 | 23.6233 | 0.001315 | 0.13398 | -8e-05 | -0.006042 | -0.006042 | -0.90253
2019-03-29 11:21 | 1553872854.8462 | -865 | 0.20824 | 0.000517 | 0 | 101.666 | 101.7571 | 127.4609 | 20.8366 | 0.001505 | 0.001391 | -152.1072 | 0.22567 | 8.3996 | 0.00204 | 0.003465 | 0.003282 | 7.4e-05 | 0.000611 | 0.24746 | 1.3436 | 2.5475 | 100.3 | 20.84 | 31.872 | 57.6105 | 70.5021 | 54.8695 | 0.007834 | 1.0763 | 0.025371 | 23.5723 | -0.008009 | 0.13425 | -8.2e-05 | -0.006227 | -0.006227 | -0.87745
2019-03-29 11:21 | 1553872855.8453 | -864 | 0.20827 | 0.000521 | 0 | 101.5543 | 101.5513 | 127.5163 | 20.8361 | 0.001547 | 0.001414 | -152.1099 | 0.22489 | 8.4031 | 0.002031 | 0.003105 | 0.002909 | 7.3e-05 | 0.000612 | 0.24347 | 1.3351 | 2.5474 | 100.3 | 20.84 | 31.872 | 58.6234 | 70.278 | 55.4145 | 0.007834 | 1.0745 | 0.025355 | 23.597 | 0.000271 | 0.13379 | -8.7e-05 | -0.006579 | -0.006579 | -0.82759
2019-03-29 11:21 | 1553872856.8455 | -863 | 0.2083 | 0.000521 | 0 | 101.5735 | 101.4561 | 127.6846 | 20.837 | 0.001484 | 0.001421 | -152.1136 | 0.22468 | 8.4013 | 0.00207 | 0.003623 | 0.003466 | 7.1e-05 | 0.000611 | 0.25465 | 1.3333 | 2.5474 | 100.3 | 20.85 | 31.872 | 57.1027 | 71.2105 | 55.5839 | 0.007839 | 1.074 | 0.025365 | 23.6187 | 0.010349 | 0.13366 | -8.7e-05 | -0.006571 | -0.006571 | -0.82783
2019-03-29 11:21 | 1553872857.8464 | -862 | 0.20829 | 0.000517 | 0 | 101.5175 | 101.4625 | 127.5557 | 20.8375 | 0.001533 | 0.001421 | -152.1041 | 0.22552 | 8.4012 | 0.002131 | 0.003747 | 0.003639 | 7.2e-05 | 0.000606 | 0.26178 | 1.3811 | 2.5473 | 100.3 | 20.85 | 31.872 | 58.3078 | 72.6947 | 55.5931 | 0.007839 | 1.0739 | 0.025352 | 23.6071 | 0.004846 | 0.13416 | -8.2e-05 | -0.006237 | -0.006237 | -0.87542
2019-03-29 11:21 | 1553872858.8454 | -861 | 0.20828 | 0.000517 | 0 | 101.4616 | 101.3827 | 127.7754 | 20.8379 | 0.001555 | 0.001424 | -152.1379 | 0.22549 | 8.404 | 0.002006 | 0.003623 | 0.003479 | 7.3e-05 | 0.00061 | 0.27527 | 1.4009 | 2.5474 | 100.3 | 20.85 | 31.871 | 58.823 | 69.6776 | 55.6578 | 0.007839 | 1.0737 | 0.025372 | 23.6297 | 0.006956 | 0.13414 | -8.3e-05 | -0.006289 | -0.006289 | -0.86805
2019-03-29 11:21 | 1553872859.8464 | -860 | 0.20827 | 0.000519 | 0 | 101.3924 | 101.4223 | 127.8512 | 20.8366 | 0.001474 | 0.001423 | -152.1556 | 0.22504 | 8.4026 | 0.002074 | 0.00336 | 0.003224 | 7e-05 | 0.000615 | 0.38488 | 1.4122 | 2.5475 | 100.3 | 20.85 | 31.871 | 56.8745 | 71.3107 | 55.6352 | 0.007839 | 1.0738 | 0.02538 | 23.6359 | -0.002634 | 0.13387 | -8.5e-05 | -0.006414 | -0.006414 | -0.84956
2019-03-29 11:21 | 1553872860.8468 | -859 | 0.20829 | 0.000519 | 0 | 101.3873 | 101.4683 | 128.1449 | 20.8374 | 0.00151 | 0.001447 | -152.1451 | 0.226 | 8.4046 | 0.002066 | 0.003697 | 0.003537 | 7.2e-05 | 0.000612 | 0.36138 | 1.4504 | 2.5476 | 100.3 | 20.85 | 31.871 | 57.7467 | 71.1282 | 56.2315 | 0.007839 | 1.0718 | 0.025386 | 23.6845 | -0.007145 | 0.13445 | -8.5e-05 | -0.006409 | -0.006409 | -0.85383
2019-03-29 11:21 | 1553872861.8465 | -858 | 0.20827 | 0.00052 | 0 | 101.3759 | 101.354 | 128.2909 | 20.8384 | 0.001544 | 0.001449 | -152.1292 | 0.22634 | 8.4008 | 0.002089 | 0.003407 | 0.003157 | 6.9e-05 | 0.000616 | 0.3653 | 1.3983 | 2.5475 | 100.3 | 20.85 | 31.871 | 58.5701 | 71.6789 | 56.2583 | 0.007839 | 1.0718 | 0.025399 | 23.6989 | 0.001936 | 0.13465 | -8.5e-05 | -0.006453 | -0.006453 | -0.8493
2019-03-29 11:21 | 1553872862.8465 | -857 | 0.20826 | 0.000517 | 0 | 101.2772 | 101.2649 | 128.7556 | 20.838 | 0.001534 | 0.001425 | -152.1324 | 0.22595 | 8.4023 | 0.002067 | 0.003054 | 0.002887 | 6.9e-05 | 0.000621 | 0.52214 | 1.3933 | 2.5476 | 100.3 | 20.85 | 31.873 | 58.3264 | 71.1437 | 55.6922 | 0.007839 | 1.0736 | 0.025467 | 23.7214 | 0.00108 | 0.13442 | -8.2e-05 | -0.006205 | -0.006205 | -0.8816
2019-03-29 11:21 | 1553872863.8465 | -856 | 0.20827 | 0.000516 | 0 | 101.3068 | 101.2206 | 129.1723 | 20.839 | 0.00154 | 0.00146 | -152.1291 | 0.22496 | 8.4032 | 0.002068 | 0.003416 | 0.003268 | 7.1e-05 | 0.000622 | 1.2379 | 1.3407 | 2.5473 | 100.3 | 20.85 | 31.874 | 58.4751 | 71.1602 | 56.5275 | 0.00784 | 1.0709 | 0.025476 | 23.7899 | 0.007617 | 0.13383 | -8.1e-05 | -0.006189 | -0.006189 | -0.88015
2019-03-29 11:21 | 1553872864.8469 | -855 | 0.20829 | 0.000521 | 0 | 101.1751 | 101.3526 | 129.5128 | 20.8394 | 0.001579 | 0.001453 | -152.1124 | 0.22515 | 8.4021 | 0.002 | 0.003017 | 0.002877 | 7e-05 | 0.000623 | 3.8642 | 1.32 | 2.5473 | 100.3 | 20.86 | 31.876 | 59.4041 | 69.5435 | 56.3682 | 0.007845 | 1.0714 | 0.025516 | 23.8155 | -0.015683 | 0.13394 | -8.7e-05 | -0.006602 | -0.006602 | -0.82576
2019-03-29 11:21 | 1553872865.8466 | -854 | 0.20827 | 0.000517 | 0 | 101.1458 | 101.0598 | 129.6502 | 20.8394 | 0.001563 | 0.001474 | -152.099 | 0.22541 | 8.4037 | 0.001991 | 0.002851 | 0.002672 | 7e-05 | 0.000625 | 9.4988 | 1.2994 | 2.5472 | 100.3 | 20.86 | 31.881 | 59.017 | 69.3174 | 56.8749 | 0.007846 | 1.0698 | 0.02551 | 23.8464 | 0.007603 | 0.1341 | -8.2e-05 | -0.006239 | -0.006239 | -0.87484
2019-03-29 11:21 | 1553872866.8466 | -853 | 0.20824 | 0.00052 | 0 | 101.0308 | 100.9841 | 129.5809 | 20.8395 | 0.001532 | 0.001447 | -152.094 | 0.22555 | 8.4024 | 0.002045 | 0.003008 | 0.002826 | 6.8e-05 | 0.000625 | 35.9329 | 1.333 | 2.5474 | 100.3 | 20.86 | 31.884 | 58.2856 | 70.6126 | 56.2133 | 0.007847 | 1.0719 | 0.025529 | 23.8162 | 0.004141 | 0.13418 | -8.6e-05 | -0.006571 | -0.006571 | -0.83112
2019-03-29 11:21 | 1553872867.847 | -852 | 0.20823 | 0.000518 | 0 | 101.02 | 101.0004 | 129.6485 | 20.8412 | 0.001521 | 0.001444 | -152.0905 | 0.22603 | 8.4016 | 0.002046 | 0.003783 | 0.003576 | 7e-05 | 0.00062 | 59.3122 | 1.3246 | 2.5473 | 100.3 | 20.86 | 31.887 | 58.0101 | 70.6486 | 56.1567 | 0.007848 | 1.0721 | 0.025537 | 23.8203 | 0.001736 | 0.13446 | -8.3e-05 | -0.006344 | -0.006344 | -0.86259
2019-03-29 11:21 | 1553872868.8457 | -851 | 0.20824 | 0.00052 | 0 | 100.9027 | 100.8864 | 129.4511 | 20.8405 | 0.001554 | 0.001405 | -152.0668 | 0.22556 | 8.4033 | 0.001969 | 0.002837 | 0.002685 | 6.6e-05 | 0.000622 | 66.7002 | 1.3458 | 2.5474 | 100.3 | 20.86 | 31.895 | 58.7983 | 68.7905 | 55.2183 | 0.00785 | 1.0752 | 0.025554 | 23.7683 | 0.001441 | 0.13419 | -8.6e-05 | -0.006537 | -0.006537 | -0.8355
2019-03-29 11:21 | 1553872869.8457 | -850 | 0.20825 | 0.00052 | 0 | 100.8509 | 100.8131 | 129.6605 | 20.8405 | 0.001534 | 0.001436 | -152.0687 | 0.22523 | 8.4018 | 0.002023 | 0.002698 | 0.002471 | 6.7e-05 | 0.000623 | 71.4773 | 1.3847 | 2.5475 | 100.3 | 20.86 | 31.899 | 58.3271 | 70.0827 | 55.9525 | 0.007851 | 1.0728 | 0.025547 | 23.8141 | 0.003358 | 0.13399 | -8.5e-05 | -0.006495 | -0.006495 | -0.83962
2019-03-29 11:21 | 1553872870.8459 | -849 | 0.20824 | 0.00052 | 0 | 100.677 | 100.5875 | 129.6173 | 20.8409 | 0.001501 | 0.001446 | -152.0743 | 0.22523 | 8.4026 | 0.002043 | 0.00316 | 0.003004 | 6.9e-05 | 0.000623 | 73.4899 | 1.3598 | 2.5475 | 100.3 | 20.86 | 31.904 | 57.5293 | 70.5775 | 56.1973 | 0.007852 | 1.072 | 0.025533 | 23.8189 | 0.007958 | 0.13399 | -8.6e-05 | -0.006518 | -0.006518 | -0.83667
2019-03-29 11:21 | 1553872871.846 | -848 | 0.20823 | 0.00052 | 0 | 100.61 | 100.5428 | 129.4826 | 20.8402 | 0.001541 | 0.001413 | -152.0802 | 0.22515 | 8.4015 | 0.002033 | 0.002857 | 0.002621 | 7e-05 | 0.000626 | 74.9481 | 1.352 | 2.5475 | 100.3 | 20.86 | 31.914 | 58.484 | 70.3217 | 55.3943 | 0.007854 | 1.0746 | 0.025551 | 23.7775 | 0.005975 | 0.13394 | -8.6e-05 | -0.006556 | -0.006556 | -0.83156
2019-03-29 11:21 | 1553872872.8458 | -847 | 0.20823 | 0.000516 | 0 | 100.5674 | 100.4609 | 129.5912 | 20.8419 | 0.001525 | 0.001403 | -152.0801 | 0.22601 | 8.4027 | 0.002026 | 0.003495 | 0.003293 | 6.9e-05 | 0.000619 | 75.0545 | 1.3556 | 2.5475 | 100.3 | 20.86 | 31.917 | 58.1056 | 70.172 | 55.1547 | 0.007855 | 1.0754 | 0.025571 | 23.7788 | 0.009486 | 0.13445 | -8e-05 | -0.006127 | -0.006127 | -0.89309
2019-03-29 11:21 | 1553872873.8457 | -846 | 0.20825 | 0.00052 | 0 | 100.4064 | 100.3291 | 129.4715 | 20.8411 | 0.001522 | 0.001433 | -152.0981 | 0.22588 | 8.4036 | 0.002029 | 0.0032 | 0.002997 | 6.7e-05 | 0.000622 | 75.1494 | 1.3999 | 2.5475 | 100.3 | 20.86 | 31.925 | 58.0288 | 70.2484 | 55.8959 | 0.007857 | 1.0729 | 0.02553 | 23.7946 | 0.006888 | 0.13438 | -8.6e-05 | -0.006513 | -0.006513 | -0.83978
2019-03-29 11:21 | 1553872874.8449 | -845 | 0.20824 | 0.000519 | 0 | 100.3531 | 100.0984 | 129.2276 | 20.8411 | 0.001492 | 0.001446 | -152.1157 | 0.22551 | 8.402 | 0.001997 | 0.002993 | 0.002944 | 6.8e-05 | 0.000613 | 75.0943 | 1.3749 | 2.5474 | 100.3 | 20.87 | 31.929 | 57.3126 | 69.4743 | 56.1999 | 0.007863 | 1.0719 | 0.025494 | 23.7832 | 0.022732 | 0.13416 | -8.4e-05 | -0.006379 | -0.006379 | -0.856
2019-03-29 11:21 | 1553872875.8449 | -844 | 0.2082 | 0.000519 | 0 | 100.2334 | 100.0561 | 129.2037 | 20.8411 | 0.001564 | 0.00145 | -152.123 | 0.22513 | 8.3995 | 0.002052 | 0.003187 | 0.002946 | 6.7e-05 | 0.000616 | 75.1176 | 1.3294 | 2.5474 | 100.3 | 20.87 | 31.933 | 59.0523 | 70.7872 | 56.2875 | 0.007864 | 1.0717 | 0.025489 | 23.7842 | 0.015836 | 0.13393 | -8.4e-05 | -0.006352 | -0.006352 | -0.85813
2019-03-29 11:21 | 1553872876.8448 | -843 | 0.20823 | 0.000516 | -1e-06 | 100.0394 | 99.9496 | 129.1439 | 20.8425 | 0.001517 | 0.001468 | -152.1166 | 0.225 | 8.4033 | 0.001981 | 0.003395 | 0.003183 | 6.7e-05 | 0.000618 | 75.2456 | 1.3801 | 2.5475 | 100.3 | 20.86 | 31.94 | 57.9089 | 69.0942 | 56.7243 | 0.007861 | 1.0702 | 0.025466 | 23.7944 | 0.008035 | 0.13386 | -8.1e-05 | -0.006117 | -0.006117 | -0.89062
2019-03-29 11:21 | 1553872877.8449 | -842 | 0.20821 | 0.000519 | -1e-06 | 100.0331 | 99.8493 | 129.1204 | 20.8423 | 0.001583 | 0.001472 | -152.1192 | 0.22538 | 8.4014 | 0.002068 | 0.002691 | 0.002493 | 6.7e-05 | 0.000619 | 75.2409 | 1.3306 | 2.5474 | 100.3 | 20.87 | 31.943 | 59.5009 | 71.1868 | 56.8221 | 0.007866 | 1.0699 | 0.02546 | 23.7958 | 0.016457 | 0.13408 | -8.4e-05 | -0.006386 | -0.006386 | -0.85456
2019-03-29 11:21 | 1553872878.8461 | -841 | 0.20824 | 0.000519 | -1e-06 | 99.8616 | 99.7636 | 129.1882 | 20.8428 | 0.001575 | 0.001475 | -152.126 | 0.22485 | 8.4018 | 0.002079 | 0.003684 | 0.00354 | 6.7e-05 | 0.000622 | 75.3414 | 1.347 | 2.5475 | 100.3 | 20.87 | 31.951 | 59.3199 | 71.4307 | 56.89 | 0.007868 | 1.0697 | 0.025464 | 23.8045 | 0.008784 | 0.13377 | -8.4e-05 | -0.006341 | -0.006341 | -0.85854
2019-03-29 11:21 | 1553872879.845 | -840 | 0.20825 | 0.000515 | 0 | 99.6802 | 99.4807 | 129.0878 | 20.8429 | 0.001561 | 0.001404 | -152.1025 | 0.22509 | 8.4027 | 0.002005 | 0.003599 | 0.003285 | 6.6e-05 | 0.000622 | 75.3988 | 1.2972 | 2.5472 | 100.3 | 20.87 | 31.955 | 58.9663 | 69.6591 | 55.1899 | 0.007869 | 1.0752 | 0.02552 | 23.7339 | 0.017924 | 0.13391 | -7.9e-05 | -0.006057 | -0.006057 | -0.89976
2019-03-29 11:21 | 1553872880.8449 | -839 | 0.20824 | 0.000517 | 0 | 99.5976 | 99.3872 | 129.3697 | 20.8434 | 0.001593 | 0.001448 | -152.0958 | 0.22536 | 8.4019 | 0.002033 | 0.003342 | 0.003205 | 6.8e-05 | 0.000627 | 75.4623 | 1.3133 | 2.5473 | 100.3 | 20.87 | 31.958 | 59.7474 | 70.3357 | 56.2398 | 0.00787 | 1.0718 | 0.025507 | 23.7977 | 0.01892 | 0.13407 | -8.2e-05 | -0.006213 | -0.006213 | -0.87828
2019-03-29 11:21 | 1553872881.845 | -838 | 0.20823 | 0.000514 | 0 | 99.5066 | 99.1503 | 129.4826 | 20.844 | 0.001593 | 0.001418 | -152.095 | 0.22568 | 8.4037 | 0.002082 | 0.00298 | 0.002736 | 6.5e-05 | 0.000626 | 75.4605 | 1.3395 | 2.5475 | 100.3 | 20.87 | 31.964 | 59.748 | 71.5034 | 55.5343 | 0.007871 | 1.0741 | 0.025545 | 23.7826 | 0.032086 | 0.13426 | -7.8e-05 | -0.005954 | -0.005954 | -0.91767
2019-03-29 11:21 | 1553872882.846 | -837 | 0.20824 | 0.000517 | 0 | 99.2621 | 98.9858 | 129.2866 | 20.8431 | 0.001554 | 0.001448 | -152.0915 | 0.22543 | 8.4034 | 0.002005 | 0.003748 | 0.003478 | 6.3e-05 | 0.000626 | 75.4681 | 1.3479 | 2.5474 | 100.3 | 20.87 | 31.967 | 58.818 | 69.6687 | 56.2527 | 0.007872 | 1.0718 | 0.025498 | 23.7905 | 0.024928 | 0.13411 | -8.2e-05 | -0.006195 | -0.006195 | -0.88101
2019-03-29 11:21 | 1553872883.8443 | -836 | 0.20824 | 0.000517 | 0 | 99.0738 | 98.7176 | 129.2521 | 20.8444 | 0.001561 | 0.001442 | -152.0828 | 0.22592 | 8.4037 | 0.00212 | 0.003455 | 0.003307 | 6.3e-05 | 0.000625 | 75.4455 | 1.3539 | 2.5474 | 100.3 | 20.87 | 31.973 | 58.9718 | 72.4253 | 56.1123 | 0.007874 | 1.0722 | 0.0255 | 23.7823 | 0.032219 | 0.1344 | -8.1e-05 | -0.006206 | -0.006206 | -0.88141
2019-03-29 11:21 | 1553872884.8442 | -835 | 0.20823 | 0.000514 | -1e-06 | 98.8436 | 98.5957 | 129.299 | 20.8436 | 0.00159 | 0.001451 | -152.0673 | 0.22564 | 8.4021 | 0.002028 | 0.003627 | 0.003478 | 6.6e-05 | 0.000628 | 75.4501 | 1.3385 | 2.5473 | 100.3 | 20.87 | 31.976 | 59.6703 | 70.2088 | 56.3245 | 0.007874 | 1.0715 | 0.025497 | 23.7943 | 0.02246 | 0.13423 | -7.8e-05 | -0.00587 | -0.00587 | -0.93073
2019-03-29 11:21 | 1553872885.8442 | -834 | 0.20824 | 0.000511 | 0 | 98.7388 | 98.4681 | 129.3283 | 20.8455 | 0.001542 | 0.001446 | -152.0715 | 0.22575 | 8.4026 | 0.002006 | 0.003006 | 0.002862 | 6.4e-05 | 0.000627 | 75.4749 | 1.3325 | 2.5474 | 100.3 | 20.87 | 31.98 | 58.5293 | 69.6984 | 56.216 | 0.007875 | 1.0719 | 0.025504 | 23.793 | 0.024561 | 0.1343 | -7.3e-05 | -0.00557 | -0.00557 | -0.98127
2019-03-29 11:21 | 1553872886.8442 | -833 | 0.20821 | 0.000513 | -1e-06 | 98.5386 | 98.3022 | 129.5635 | 20.8449 | 0.001568 | 0.001438 | -152.0866 | 0.22571 | 8.4006 | 0.001992 | 0.003702 | 0.003432 | 6.3e-05 | 0.000621 | 75.3805 | 1.3192 | 2.5472 | 100.3 | 20.87 | 31.988 | 59.1428 | 69.3546 | 56.0154 | 0.007877 | 1.0725 | 0.025535 | 23.8074 | 0.021482 | 0.13428 | -7.6e-05 | -0.005757 | -0.005757 | -0.94924
2019-03-29 11:21 | 1553872887.8442 | -832 | 0.20827 | 0.000515 | 0 | 98.3689 | 98.0712 | 129.6898 | 20.8448 | 0.00156 | 0.001387 | -152.087 | 0.2262 | 8.4021 | 0.002041 | 0.003351 | 0.003267 | 6.3e-05 | 0.00062 | 75.4341 | 1.3091 | 2.5472 | 100.3 | 20.87 | 31.992 | 58.9448 | 70.5367 | 54.7891 | 0.007878 | 1.0766 | 0.025595 | 23.7746 | 0.027111 | 0.13456 | -7.9e-05 | -0.005995 | -0.005995 | -0.91354
2019-03-29 11:21 | 1553872888.8442 | -831 | 0.20835 | 0.000512 | -1e-06 | 98.159 | 97.6077 | 129.7018 | 20.8457 | 0.001533 | 0.00144 | -152.0741 | 0.22655 | 8.4016 | 0.001982 | 0.003541 | 0.003377 | 6.3e-05 | 0.000617 | 75.4612 | 1.3317 | 2.5475 | 100.3 | 20.87 | 31.996 | 58.2962 | 69.1073 | 56.0546 | 0.007879 | 1.0724 | 0.025547 | 23.8215 | 0.050382 | 0.13478 | -7.5e-05 | -0.005696 | -0.005696 | -0.96304
2019-03-29 11:21 | 1553872889.8453 | -830 | 0.20819 | 0.000514 | -1e-06 | 98.0091 | 97.5493 | 129.7307 | 20.8453 | 0.001527 | 0.001466 | -152.0719 | 0.22592 | 8.4 | 0.002036 | 0.003396 | 0.003203 | 6.6e-05 | 0.000619 | 75.4749 | 1.3781 | 2.5474 | 100.3 | 20.88 | 32 | 58.1699 | 70.407 | 56.6974 | 0.007885 | 1.0703 | 0.025525 | 23.8474 | 0.042056 | 0.1344 | -7.7e-05 | -0.005833 | -0.005833 | -0.93772
2019-03-29 11:21 | 1553872890.8454 | -829 | 0.20835 | 0.000512 | 0 | 97.7448 | 97.3329 | 129.533 | 20.8456 | 0.00151 | 0.001443 | -152.0919 | 0.22675 | 8.4019 | 0.002017 | 0.002895 | 0.00269 | 6.9e-05 | 0.000622 | 75.4133 | 1.3077 | 2.5473 | 100.3 | 20.88 | 32.004 | 57.7428 | 69.9442 | 56.1362 | 0.007886 | 1.0722 | 0.025527 | 23.809 | 0.037763 | 0.13489 | -7.6e-05 | -0.005831 | -0.005831 | -0.94154
2019-03-29 11:22 | 1553872891.8443 | -828 | 0.20816 | 0.000511 | 0 | 97.4876 | 96.9722 | 129.5692 | 20.8469 | 0.001535 | 0.001421 | -152.0885 | 0.22644 | 8.4018 | 0.002068 | 0.003431 | 0.003221 | 6.6e-05 | 0.000628 | 75.317 | 1.3314 | 2.5474 | 100.3 | 20.88 | 32.005 | 58.3528 | 71.1874 | 55.5967 | 0.007886 | 1.0739 | 0.025551 | 23.7928 | 0.047415 | 0.13471 | -7.5e-05 | -0.005746 | -0.005746 | -0.9541
2019-03-29 11:22 | 1553872892.8454 | -827 | 0.2083 | 0.00051 | -1e-06 | 97.3586 | 96.7862 | 129.3504 | 20.8471 | 0.001561 | 0.001445 | -152.0773 | 0.22587 | 8.4028 | 0.00205 | 0.003278 | 0.003096 | 7.1e-05 | 0.000629 | 75.3884 | 1.4057 | 2.5474 | 100.3 | 20.88 | 32.007 | 58.9868 | 70.7541 | 56.1893 | 0.007887 | 1.072 | 0.025507 | 23.7941 | 0.052744 | 0.13437 | -7.3e-05 | -0.005514 | -0.005514 | -0.99179
2019-03-29 11:22 | 1553872893.8454 | -826 | 0.20824 | 0.000513 | 0 | 97.1802 | 96.7586 | 129.2788 | 20.8469 | 0.001469 | 0.001458 | -152.0666 | 0.22644 | 8.4028 | 0.002115 | 0.003536 | 0.003397 | 6.9e-05 | 0.000633 | 75.4553 | 1.4394 | 2.5477 | 100.3 | 20.88 | 32.009 | 56.7677 | 72.3107 | 56.4969 | 0.007887 | 1.071 | 0.025488 | 23.7986 | 0.038894 | 0.13471 | -7.7e-05 | -0.005832 | -0.005832 | -0.94015
2019-03-29 11:22 | 1553872894.8454 | -825 | 0.20823 | 0.000515 | 0 | 96.9479 | 96.4074 | 129.3222 | 20.8468 | 0.001532 | 0.001517 | -152.0822 | 0.22681 | 8.4016 | 0.002047 | 0.002692 | 0.002379 | 7.1e-05 | 0.000632 | 75.4289 | 1.3481 | 2.5473 | 100.3 | 20.88 | 32.01 | 58.2691 | 70.6728 | 57.9187 | 0.007888 | 1.0664 | 0.025437 | 23.8539 | 0.050012 | 0.13493 | -8e-05 | -0.006049 | -0.006049 | -0.90791
2019-03-29 11:22 | 1553872895.8455 | -824 | 0.20831 | 0.000519 | 0 | 96.6627 | 95.9797 | 129.4358 | 20.8486 | 0.001578 | 0.001415 | -152.0796 | 0.22649 | 8.4017 | 0.002091 | 0.003268 | 0.003137 | 6.9e-05 | 0.000634 | 75.4264 | 1.341 | 2.5472 | 100.3 | 20.88 | 32.011 | 59.3904 | 71.7253 | 55.4643 | 0.007888 | 1.0743 | 0.025543 | 23.7758 | 0.063418 | 0.13474 | -8.5e-05 | -0.006457 | -0.006457 | -0.8493
2019-03-29 11:22 | 1553872896.8455 | -823 | 0.20815 | 0.000513 | -1e-06 | 96.4897 | 95.7992 | 129.525 | 20.8474 | 0.001563 | 0.001452 | -152.0608 | 0.22634 | 8.4021 | 0.00208 | 0.003087 | 0.002929 | 7e-05 | 0.00064 | 75.4049 | 1.3207 | 2.5474 | 100.3 | 20.87 | 32.012 | 59.0327 | 71.4568 | 56.3439 | 0.007883 | 1.0715 | 0.025518 | 23.8158 | 0.064235 | 0.13465 | -7.6e-05 | -0.005712 | -0.005712 | -0.95939
2019-03-29 11:22 | 1553872897.8447 | -822 | 0.20826 | 0.000513 | 0 | 96.3128 | 95.5835 | 129.6166 | 20.8476 | 0.001617 | 0.001403 | -152.0646 | 0.22545 | 8.4017 | 0.00203 | 0.003892 | 0.003716 | 7.2e-05 | 0.00064 | 75.3817 | 1.3238 | 2.5474 | 100.3 | 20.87 | 32.011 | 60.3395 | 70.2602 | 55.1634 | 0.007883 | 1.0753 | 0.025573 | 23.7815 | 0.067983 | 0.13412 | -7.7e-05 | -0.005878 | -0.005878 | -0.92866
2019-03-29 11:22 | 1553872898.8446 | -821 | 0.20825 | 0.000514 | 0 | 96.0766 | 95.3013 | 129.6248 | 20.8479 | 0.001594 | 0.001448 | -152.0658 | 0.22582 | 8.4012 | 0.002045 | 0.002241 | 0.002057 | 6.8e-05 | 0.000642 | 75.3536 | 1.315 | 2.5473 | 100.3 | 20.87 | 32.013 | 59.7713 | 70.6309 | 56.2574 | 0.007884 | 1.0718 | 0.025531 | 23.8218 | 0.072473 | 0.13434 | -7.8e-05 | -0.005912 | -0.005912 | -0.92486
2019-03-29 11:22 | 1553872899.8456 | -820 | 0.2083 | 0.000513 | 0 | 95.929 | 95.3159 | 129.9485 | 20.8477 | 0.001519 | 0.001436 | -152.0608 | 0.22603 | 8.4036 | 0.00201 | 0.003202 | 0.003046 | 6.9e-05 | 0.000645 | 75.3701 | 1.3363 | 2.5474 | 100.3 | 20.87 | 32.016 | 57.9687 | 69.7879 | 55.954 | 0.007884 | 1.0727 | 0.025575 | 23.8405 | 0.057347 | 0.13447 | -7.7e-05 | -0.005905 | -0.005905 | -0.92681
2019-03-29 11:22 | 1553872900.8447 | -819 | 0.20826 | 0.000513 | 0 | 95.7416 | 94.9651 | 130.1083 | 20.8492 | 0.001557 | 0.00144 | -152.0583 | 0.2254 | 8.4035 | 0.002074 | 0.003422 | 0.003217 | 7.1e-05 | 0.000647 | 75.4446 | 1.3536 | 2.5474 | 100.3 | 20.87 | 32.02 | 58.8811 | 71.323 | 56.0588 | 0.007885 | 1.0724 | 0.025587 | 23.859 | 0.072843 | 0.13409 | -7.7e-05 | -0.005895 | -0.005895 | -0.92575
2019-03-29 11:22 | 1553872901.8447 | -818 | 0.20824 | 0.000518 | 0 | 95.5659 | 94.7357 | 129.9311 | 20.8483 | 0.001532 | 0.001431 | -152.054 | 0.22653 | 8.4017 | 0.002044 | 0.002948 | 0.002846 | 7e-05 | 0.000652 | 71.4724 | 1.363 | 2.5475 | 100.3 | 20.87 | 32.031 | 58.2761 | 70.6079 | 55.841 | 0.007888 | 1.0731 | 0.025578 | 23.8349 | 0.07804 | 0.13476 | -8.3e-05 | -0.006314 | -0.006314 | -0.86872
2019-03-29 11:22 | 1553872902.8438 | -817 | 0.20828 | 0.000515 | 0 | 95.3926 | 94.5965 | 130.0849 | 20.8488 | 0.00158 | 0.001426 | -152.0533 | 0.2255 | 8.4011 | 0.002088 | 0.003199 | 0.003004 | 6.9e-05 | 0.000655 | 49.8974 | 1.3476 | 2.5474 | 100.3 | 20.87 | 32.038 | 59.4516 | 71.6617 | 55.7131 | 0.00789 | 1.0735 | 0.025598 | 23.8443 | 0.074961 | 0.13415 | -7.9e-05 | -0.00602 | -0.00602 | -0.90702
2019-03-29 11:22 | 1553872903.8438 | -816 | 0.20826 | 0.000515 | 0 | 95.2217 | 94.4627 | 130.0898 | 20.8484 | 0.001562 | 0.001488 | -152.0577 | 0.22612 | 8.4014 | 0.001981 | 0.003654 | 0.003506 | 6.9e-05 | 0.000657 | 44.4723 | 1.3264 | 2.5474 | 100.3 | 20.87 | 32.054 | 59.003 | 69.0984 | 57.2272 | 0.007894 | 1.0686 | 0.025539 | 23.8996 | 0.071579 | 0.13452 | -7.9e-05 | -0.006056 | -0.006056 | -0.90406
2019-03-29 11:22 | 1553872904.8439 | -815 | 0.20823 | 0.000512 | 0 | 95.0349 | 94.2829 | 129.8267 | 20.8493 | 0.001508 | 0.001496 | -152.0411 | 0.22706 | 8.4017 | 0.002078 | 0.003649 | 0.003519 | 6.9e-05 | 0.000653 | 41.4181 | 1.345 | 2.5474 | 100.3 | 20.87 | 32.063 | 57.701 | 71.4182 | 57.4223 | 0.007896 | 1.068 | 0.025506 | 23.8825 | 0.071063 | 0.13508 | -7.6e-05 | -0.005764 | -0.005764 | -0.95386
2019-03-29 11:22 | 1553872905.8439 | -814 | 0.20821 | 0.000514 | 0 | 94.8242 | 94.0242 | 129.6354 | 20.8492 | 0.001585 | 0.001446 | -152.0285 | 0.22592 | 8.4021 | 0.002023 | 0.00289 | 0.002756 | 6.8e-05 | 0.000649 | 41.2834 | 1.3501 | 2.5474 | 100.3 | 20.87 | 32.082 | 59.5663 | 70.1075 | 56.1967 | 0.007901 | 1.072 | 0.025535 | 23.8206 | 0.075784 | 0.1344 | -7.9e-05 | -0.006033 | -0.006033 | -0.90671
2019-03-29 11:22 | 1553872906.8448 | -813 | 0.20823 | 0.000513 | 0 | 94.6461 | 93.7392 | 129.7591 | 20.8496 | 0.001542 | 0.001482 | -152.0392 | 0.22564 | 8.4005 | 0.002068 | 0.00291 | 0.002641 | 6.8e-05 | 0.00065 | 36.399 | 1.3551 | 2.5474 | 100.3 | 20.87 | 32.093 | 58.5246 | 71.1869 | 57.0721 | 0.007903 | 1.0691 | 0.025513 | 23.8636 | 0.086119 | 0.13423 | -7.7e-05 | -0.005824 | -0.005824 | -0.93802
2019-03-29 11:22 | 1553872907.8438 | -812 | 0.20823 | 0.000516 | 0 | 94.526 | 93.5658 | 129.8814 | 20.8501 | 0.001581 | 0.001454 | -152.0339 | 0.22659 | 8.4011 | 0.0021 | 0.003075 | 0.00287 | 6.8e-05 | 0.000651 | 32.511 | 1.3256 | 2.5474 | 100.3 | 20.87 | 32.103 | 59.4702 | 71.9585 | 56.4068 | 0.007906 | 1.0713 | 0.025551 | 23.8508 | 0.091327 | 0.1348 | -8e-05 | -0.006126 | -0.006126 | -0.89557
2019-03-29 11:22 | 1553872908.8449 | -811 | 0.2082 | 0.000514 | 0 | 94.4665 | 93.6474 | 129.5844 | 20.8506 | 0.00152 | 0.001503 | -152.0602 | 0.22661 | 8.4003 | 0.002026 | 0.003619 | 0.003442 | 6.9e-05 | 0.00065 | 54.4166 | 1.3572 | 2.5474 | 100.3 | 20.87 | 32.122 | 58.0067 | 70.1752 | 57.592 | 0.00791 | 1.0674 | 0.025476 | 23.8663 | 0.077894 | 0.13481 | -7.8e-05 | -0.005905 | -0.005905 | -0.92914
2019-03-29 11:22 | 1553872909.8449 | -810 | 0.20824 | 0.000512 | -1e-06 | 94.3321 | 93.2515 | 129.1171 | 20.8509 | 0.001557 | 0.001476 | -152.0738 | 0.22537 | 8.4045 | 0.002075 | 0.003225 | 0.003041 | 7e-05 | 0.000652 | 75.1665 | 1.4072 | 2.5475 | 100.3 | 20.87 | 32.131 | 58.881 | 71.3636 | 56.9394 | 0.007913 | 1.0695 | 0.025455 | 23.7997 | 0.10305 | 0.13408 | -7.5e-05 | -0.005661 | -0.005661 | -0.96387
2019-03-29 11:22 | 1553872910.8449 | -809 | 0.20825 | 0.000513 | 0 | 94.2707 | 93.3774 | 128.8861 | 20.8503 | 0.001575 | 0.001469 | -152.062 | 0.22622 | 8.4016 | 0.002026 | 0.003469 | 0.00334 | 6.9e-05 | 0.000653 | 75.4471 | 1.3221 | 2.5474 | 100.3 | 20.87 | 32.139 | 59.3287 | 70.184 | 56.7578 | 0.007915 | 1.0701 | 0.025439 | 23.7719 | 0.085161 | 0.13458 | -7.6e-05 | -0.005765 | -0.005765 | -0.95011
2019-03-29 11:22 | 1553872911.8451 | -808 | 0.20827 | 0.00051 | -1e-06 | 94.2096 | 93.2204 | 128.9204 | 20.8512 | 0.001526 | 0.001483 | -152.0668 | 0.22643 | 8.4006 | 0.002027 | 0.00388 | 0.003636 | 6.5e-05 | 0.000656 | 75.4559 | 1.311 | 2.5472 | 100.3 | 20.87 | 32.151 | 58.1451 | 70.2046 | 57.0954 | 0.007917 | 1.069 | 0.025429 | 23.7872 | 0.094416 | 0.1347 | -7.1e-05 | -0.005392 | -0.005392 | -1.0167
2019-03-29 11:22 | 1553872912.844 | -807 | 0.20824 | 0.000514 | 0 | 94.0055 | 93.0959 | 128.9267 | 20.8509 | 0.001571 | 0.001453 | -152.0654 | 0.22594 | 8.403 | 0.002007 | 0.003811 | 0.003669 | 6.5e-05 | 0.000656 | 75.5546 | 1.3154 | 2.5473 | 100.3 | 20.87 | 32.155 | 59.2278 | 69.7145 | 56.3739 | 0.007918 | 1.0714 | 0.025458 | 23.7618 | 0.086974 | 0.13441 | -7.8e-05 | -0.005933 | -0.005933 | -0.92202
2019-03-29 11:22 | 1553872913.8443 | -806 | 0.20827 | 0.000511 | -1e-06 | 93.8942 | 92.8617 | 128.9929 | 20.8514 | 0.001562 | 0.001485 | -152.0826 | 0.22537 | 8.4041 | 0.002072 | 0.003379 | 0.003223 | 6.9e-05 | 0.000656 | 75.5818 | 1.3814 | 2.5474 | 100.3 | 20.87 | 32.162 | 59.0122 | 71.2807 | 57.1628 | 0.00792 | 1.0688 | 0.025434 | 23.7963 | 0.098895 | 0.13408 | -7.4e-05 | -0.005574 | -0.005574 | -0.97899
2019-03-29 11:22 | 1553872914.8441 | -805 | 0.20825 | 0.000516 | 0 | 93.807 | 92.7017 | 128.9122 | 20.8525 | 0.001509 | 0.001446 | -152.0917 | 0.22674 | 8.4015 | 0.002062 | 0.00362 | 0.003459 | 6.8e-05 | 0.000657 | 75.5963 | 1.3338 | 2.5474 | 100.3 | 20.87 | 32.166 | 57.7362 | 71.0368 | 56.2218 | 0.007921 | 1.0719 | 0.025462 | 23.7549 | 0.10602 | 0.13489 | -8e-05 | -0.00605 | -0.00605 | -0.90746
2019-03-29 11:22 | 1553872915.8431 | -804 | 0.20826 | 0.00051 | 0 | 93.6032 | 92.5747 | 128.8676 | 20.8523 | 0.001544 | 0.001464 | -152.0963 | 0.22522 | 8.4036 | 0.002062 | 0.003475 | 0.003255 | 7e-05 | 0.000658 | 75.5065 | 1.3564 | 2.5475 | 100.3 | 20.87 | 32.169 | 58.5673 | 71.0447 | 56.6509 | 0.007922 | 1.0705 | 0.025441 | 23.7663 | 0.098825 | 0.13398 | -7.3e-05 | -0.005509 | -0.005509 | -0.98985
2019-03-29 11:22 | 1553872916.8432 | -803 | 0.20826 | 0.000511 | 0 | 93.4505 | 92.4829 | 128.3582 | 20.8535 | 0.001572 | 0.001488 | -152.0857 | 0.22565 | 8.401 | 0.002045 | 0.003289 | 0.003149 | 6.9e-05 | 0.000655 | 75.4661 | 1.37 | 2.5475 | 100.3 | 20.87 | 32.174 | 59.2561 | 70.6332 | 57.22 | 0.007923 | 1.0686 | 0.025369 | 23.7397 | 0.0931 | 0.13424 | -7.4e-05 | -0.005642 | -0.005642 | -0.9684
2019-03-29 11:22 | 1553872917.8441 | -802 | 0.20825 | 0.000511 | 0 | 93.3221 | 92.2522 | 128.274 | 20.8528 | 0.001524 | 0.001451 | -152.0678 | 0.22599 | 8.404 | 0.002038 | 0.003188 | 0.002945 | 7e-05 | 0.000656 | 75.5239 | 1.3578 | 2.5475 | 100.3 | 20.87 | 32.177 | 58.0818 | 70.458 | 56.3265 | 0.007924 | 1.0715 | 0.025395 | 23.6998 | 0.10314 | 0.13444 | -7.4e-05 | -0.005636 | -0.005636 | -0.97083
2019-03-29 11:22 | 1553872918.8442 | -801 | 0.20825 | 0.000512 | -1e-06 | 93.1799 | 92.2437 | 128.5343 | 20.8538 | 0.001586 | 0.001496 | -152.0556 | 0.22666 | 8.4028 | 0.002049 | 0.00395 | 0.003735 | 6.8e-05 | 0.000657 | 75.4686 | 1.3132 | 2.5473 | 100.3 | 20.87 | 32.183 | 59.5825 | 70.7272 | 57.4181 | 0.007925 | 1.068 | 0.025379 | 23.7631 | 0.090331 | 0.13484 | -7.4e-05 | -0.005522 | -0.005522 | -0.9938
2019-03-29 11:22 | 1553872919.8442 | -800 | 0.20825 | 0.000512 | -1e-06 | 93.1192 | 91.8942 | 128.7587 | 20.854 | 0.001543 | 0.001519 | -152.065 | 0.22581 | 8.402 | 0.002068 | 0.003361 | 0.003216 | 6.9e-05 | 0.000662 | 75.4411 | 1.3583 | 2.5474 | 100.3 | 20.87 | 32.186 | 58.5477 | 71.1881 | 57.964 | 0.007926 | 1.0662 | 0.02538 | 23.8035 | 0.11845 | 0.13433 | -7.5e-05 | -0.005611 | -0.005611 | -0.97436
2019-03-29 11:22 | 1553872920.8433 | -799 | 0.20825 | 0.000515 | -1e-06 | 93.0075 | 91.8338 | 128.6969 | 20.8535 | 0.001576 | 0.001446 | -152.0708 | 0.22638 | 8.4023 | 0.002108 | 0.003957 | 0.003684 | 6.5e-05 | 0.000657 | 75.4548 | 1.4441 | 2.5476 | 100.3 | 20.87 | 32.189 | 59.3585 | 72.1395 | 56.2039 | 0.007927 | 1.0719 | 0.025442 | 23.7345 | 0.1136 | 0.13468 | -7.8e-05 | -0.00588 | -0.00588 | -0.93223
2019-03-29 11:22 | 1553872921.8433 | -798 | 0.20826 | 0.000515 | 0 | 92.8947 | 91.7553 | 128.8242 | 20.8537 | 0.001585 | 0.001489 | -152.0731 | 0.22625 | 8.4018 | 0.002004 | 0.003363 | 0.003159 | 6.6e-05 | 0.000664 | 75.4853 | 1.3742 | 2.5475 | 100.3 | 20.87 | 32.194 | 59.5674 | 69.6564 | 57.2512 | 0.007928 | 1.0685 | 0.025414 | 23.7839 | 0.11039 | 0.13459 | -7.9e-05 | -0.005989 | -0.005989 | -0.91463
2019-03-29 11:22 | 1553872922.8433 | -797 | 0.20825 | 0.000514 | 0 | 92.8327 | 91.829 | 128.5599 | 20.854 | 0.001551 | 0.001474 | -152.085 | 0.22683 | 8.4016 | 0.00204 | 0.003442 | 0.003253 | 6.9e-05 | 0.000667 | 75.477 | 1.3356 | 2.5475 | 100.29 | 20.87 | 32.196 | 58.7348 | 70.5159 | 56.8876 | 0.007929 | 1.0697 | 0.025402 | 23.7464 | 0.097237 | 0.13494 | -7.7e-05 | -0.005805 | -0.005805 | -0.94617
2019-03-29 11:22 | 1553872923.8434 | -796 | 0.20825 | 0.000514 | 0 | 92.7612 | 91.6778 | 128.9103 | 20.8542 | 0.001608 | 0.001494 | -152.0742 | 0.22675 | 8.4018 | 0.00213 | 0.003781 | 0.003649 | 7.1e-05 | 0.000672 | 75.4356 | 1.3464 | 2.5475 | 100.29 | 20.87 | 32.2 | 60.1091 | 72.671 | 57.3801 | 0.00793 | 1.0681 | 0.025417 | 23.7965 | 0.10508 | 0.1349 | -7.8e-05 | -0.005895 | -0.005895 | -0.93128
2019-03-29 11:22 | 1553872924.8434 | -795 | 0.20825 | 0.000518 | 0 | 92.6546 | 91.4338 | 129.4218 | 20.8542 | 0.001543 | 0.001481 | -152.0822 | 0.22571 | 8.4047 | 0.002135 | 0.003183 | 0.002965 | 6.8e-05 | 0.000675 | 75.4742 | 1.3889 | 2.5474 | 100.29 | 20.87 | 32.207 | 58.5621 | 72.8027 | 57.0615 | 0.007932 | 1.0691 | 0.02548 | 23.8322 | 0.11864 | 0.13427 | -8.3e-05 | -0.006306 | -0.006306 | -0.86669
2019-03-29 11:22 | 1553872925.8444 | -794 | 0.20827 | 0.000513 | 0 | 92.5169 | 91.228 | 129.5816 | 20.8555 | 0.001614 | 0.001417 | -152.0987 | 0.22554 | 8.4049 | 0.002106 | 0.004001 | 0.003853 | 6.9e-05 | 0.00068 | 75.4654 | 1.3569 | 2.5475 | 100.29 | 20.87 | 32.21 | 60.2768 | 72.1072 | 55.5213 | 0.007933 | 1.0742 | 0.025556 | 23.7912 | 0.12549 | 0.13417 | -7.6e-05 | -0.005789 | -0.005789 | -0.94327
2019-03-29 11:22 | 1553872926.8436 | -793 | 0.20826 | 0.000513 | 0 | 92.3155 | 91.0792 | 129.4312 | 20.8561 | 0.001562 | 0.001459 | -152.1156 | 0.22534 | 8.402 | 0.00207 | 0.003313 | 0.003155 | 6.9e-05 | 0.000688 | 75.4071 | 1.3345 | 2.5472 | 100.29 | 20.87 | 32.216 | 59.0035 | 71.2446 | 56.5167 | 0.007934 | 1.0709 | 0.025502 | 23.8134 | 0.12059 | 0.13405 | -7.6e-05 | -0.005779 | -0.005779 | -0.94412
2019-03-29 11:22 | 1553872927.8445 | -792 | 0.20824 | 0.000513 | 0 | 92.1586 | 91.113 | 129.2903 | 20.8558 | 0.001588 | 0.001416 | -152.1288 | 0.22634 | 8.4018 | 0.002123 | 0.002993 | 0.002754 | 6.6e-05 | 0.000695 | 75.4304 | 1.3234 | 2.5473 | 100.29 | 20.87 | 32.219 | 59.6486 | 72.4996 | 55.4815 | 0.007935 | 1.0743 | 0.025528 | 23.763 | 0.10207 | 0.13465 | -7.6e-05 | -0.005829 | -0.005829 | -0.94015
2019-03-29 11:22 | 1553872928.8435 | -791 | 0.20828 | 0.000513 | 0 | 92.0317 | 90.67 | 129.1028 | 20.855 | 0.00152 | 0.001486 | -152.1143 | 0.22651 | 8.4031 | 0.002065 | 0.00348 | 0.003285 | 6.7e-05 | 0.000692 | 75.3607 | 1.369 | 2.5473 | 100.29 | 20.87 | 32.221 | 57.9985 | 71.1102 | 57.1813 | 0.007936 | 1.0688 | 0.025444 | 23.8071 | 0.13334 | 0.13475 | -7.6e-05 | -0.005814 | -0.005814 | -0.94336
2019-03-29 11:22 | 1553872929.8446 | -790 | 0.20828 | 0.000516 | 0 | 91.8719 | 90.5427 | 128.9345 | 20.8564 | 0.001572 | 0.001524 | -152.0918 | 0.22569 | 8.4036 | 0.002108 | 0.003856 | 0.003667 | 6.6e-05 | 0.000697 | 75.426 | 1.3108 | 2.5473 | 100.29 | 20.87 | 32.225 | 59.2432 | 72.1501 | 58.1086 | 0.007937 | 1.0658 | 0.025392 | 23.825 | 0.13035 | 0.13426 | -8e-05 | -0.006046 | -0.006046 | -0.90384
2019-03-29 11:22 | 1553872930.8436 | -789 | 0.20827 | 0.000515 | 0 | 91.7345 | 90.3935 | 129.1016 | 20.8567 | 0.001539 | 0.001446 | -152.0999 | 0.22614 | 8.4036 | 0.002083 | 0.003066 | 0.002876 | 6.6e-05 | 0.000699 | 75.4893 | 1.3617 | 2.5473 | 100.29 | 20.87 | 32.227 | 58.4674 | 71.5617 | 56.2182 | 0.007937 | 1.0719 | 0.025481 | 23.7723 | 0.13171 | 0.13453 | -7.9e-05 | -0.006022 | -0.006022 | -0.90929
2019-03-29 11:22 | 1553872931.8427 | -788 | 0.20825 | 0.000513 | 0 | 91.598 | 90.3529 | 128.9798 | 20.8565 | 0.001573 | 0.001514 | -152.1072 | 0.22638 | 8.4015 | 0.002085 | 0.003547 | 0.00333 | 6.9e-05 | 0.000692 | 75.3518 | 1.3964 | 2.5475 | 100.29 | 20.87 | 32.229 | 59.2778 | 71.6038 | 57.8646 | 0.007937 | 1.0666 | 0.025406 | 23.8204 | 0.12241 | 0.13467 | -7.6e-05 | -0.005766 | -0.005766 | -0.95064
2019-03-29 11:22 | 1553872932.8438 | -787 | 0.20827 | 0.000515 | 0 | 91.4663 | 90.1171 | 129.0671 | 20.8571 | 0.001556 | 0.001468 | -152.1018 | 0.22642 | 8.4026 | 0.002148 | 0.00329 | 0.003153 | 6.7e-05 | 0.000695 | 75.3934 | 1.3863 | 2.5475 | 100.29 | 20.87 | 32.229 | 58.8808 | 73.1118 | 56.7566 | 0.007937 | 1.0701 | 0.025457 | 23.7885 | 0.13292 | 0.1347 | -7.9e-05 | -0.00598 | -0.00598 | -0.9167
2019-03-29 11:22 | 1553872933.8427 | -786 | 0.20823 | 0.000515 | 0 | 91.301 | 89.9092 | 128.397 | 20.8567 | 0.001609 | 0.00149 | -152.0791 | 0.22662 | 8.4028 | 0.002029 | 0.003028 | 0.002717 | 6.7e-05 | 0.000699 | 75.418 | 1.3417 | 2.5473 | 100.29 | 20.87 | 32.229 | 60.1537 | 70.2496 | 57.2824 | 0.007937 | 1.0684 | 0.025371 | 23.7456 | 0.1374 | 0.13482 | -7.9e-05 | -0.005955 | -0.005955 | -0.92136
2019-03-29 11:22 | 1553872934.8427 | -785 | 0.20825 | 0.000514 | 0 | 91.2591 | 89.8536 | 128.1571 | 20.8572 | 0.001635 | 0.001516 | -152.0554 | 0.22523 | 8.4004 | 0.002063 | 0.003097 | 0.002927 | 6.6e-05 | 0.000699 | 75.4469 | 1.35 | 2.5475 | 100.29 | 20.87 | 32.227 | 60.7748 | 71.0719 | 57.9169 | 0.007937 | 1.0664 | 0.025323 | 23.7461 | 0.13883 | 0.13399 | -7.8e-05 | -0.005851 | -0.005851 | -0.93204
2019-03-29 11:22 | 1553872935.8428 | -784 | 0.20826 | 0.00051 | 0 | 91.0074 | 89.6041 | 128.112 | 20.8568 | 0.001608 | 0.001512 | -152.0467 | 0.22591 | 8.4052 | 0.002043 | 0.003668 | 0.003554 | 6.6e-05 | 0.000693 | 75.4236 | 1.3484 | 2.5474 | 100.29 | 20.87 | 32.223 | 60.1217 | 70.5974 | 57.8 | 0.007936 | 1.0668 | 0.025323 | 23.7378 | 0.13899 | 0.13439 | -7.2e-05 | -0.005419 | -0.005419 | -1.0094
2019-03-29 11:22 | 1553872936.8438 | -783 | 0.20825 | 0.000515 | 0 | 90.7976 | 89.3239 | 128.5214 | 20.858 | 0.001624 | 0.001461 | -152.0365 | 0.22592 | 8.4028 | 0.002005 | 0.003175 | 0.003011 | 6.3e-05 | 0.000699 | 75.4236 | 1.3302 | 2.5473 | 100.29 | 20.87 | 32.216 | 60.5169 | 69.6683 | 56.5806 | 0.007934 | 1.0707 | 0.02541 | 23.7318 | 0.14636 | 0.1344 | -7.9e-05 | -0.005988 | -0.005988 | -0.91352
2019-03-29 11:22 | 1553872937.8438 | -782 | 0.20828 | 0.000512 | -1e-06 | 90.7374 | 89.3881 | 128.6547 | 20.858 | 0.001572 | 0.001433 | -152.0423 | 0.22604 | 8.4033 | 0.002013 | 0.003558 | 0.003314 | 6e-05 | 0.000696 | 75.4847 | 1.338 | 2.5473 | 100.29 | 20.87 | 32.211 | 59.2501 | 69.8805 | 55.9123 | 0.007933 | 1.0729 | 0.025449 | 23.7201 | 0.13401 | 0.13447 | -7.4e-05 | -0.005601 | -0.005601 | -0.9772
2019-03-29 11:22 | 1553872938.8439 | -781 | 0.20827 | 0.000515 | -1e-06 | 90.5608 | 89.0161 | 128.8976 | 20.8591 | 0.001601 | 0.001465 | -152.0601 | 0.22575 | 8.4006 | 0.002079 | 0.003211 | 0.003033 | 6.2e-05 | 0.000692 | 75.4417 | 1.3257 | 2.5474 | 100.29 | 20.87 | 32.203 | 59.9668 | 71.4465 | 56.6661 | 0.007931 | 1.0704 | 0.025444 | 23.7696 | 0.15388 | 0.1343 | -7.8e-05 | -0.005866 | -0.005866 | -0.93175
2019-03-29 11:22 | 1553872939.8429 | -780 | 0.20825 | 0.000511 | 0 | 90.521 | 89.1325 | 128.8206 | 20.8579 | 0.001618 | 0.001457 | -152.0765 | 0.22513 | 8.3992 | 0.002032 | 0.00312 | 0.00291 | 6.4e-05 | 0.000698 | 75.4534 | 1.3433 | 2.5475 | 100.29 | 20.87 | 32.191 | 60.3684 | 70.3217 | 56.4755 | 0.007928 | 1.0711 | 0.025443 | 23.7557 | 0.13826 | 0.13393 | -7.4e-05 | -0.005615 | -0.005615 | -0.97071
2019-03-29 11:22 | 1553872940.844 | -779 | 0.20826 | 0.000513 | 0 | 90.383 | 88.856 | 128.6711 | 20.8598 | 0.001563 | 0.001476 | -152.0564 | 0.22536 | 8.4046 | 0.001998 | 0.003529 | 0.003358 | 6.6e-05 | 0.000698 | 75.4257 | 1.3565 | 2.5474 | 100.29 | 20.87 | 32.184 | 59.0291 | 69.5143 | 56.9449 | 0.007926 | 1.0695 | 0.025411 | 23.7588 | 0.1524 | 0.13406 | -7.6e-05 | -0.005774 | -0.005774 | -0.94504
2019-03-29 11:22 | 1553872941.8439 | -778 | 0.20828 | 0.000511 | 0 | 90.3408 | 88.7229 | 128.6601 | 20.8588 | 0.001596 | 0.001504 | -152.0662 | 0.22582 | 8.4025 | 0.002099 | 0.003482 | 0.003267 | 6.9e-05 | 0.000698 | 75.4614 | 1.3296 | 2.5473 | 100.29 | 20.87 | 32.176 | 59.8345 | 71.9405 | 57.6126 | 0.007924 | 1.0674 | 0.025384 | 23.7818 | 0.16163 | 0.13434 | -7.3e-05 | -0.005549 | -0.005549 | -0.98541
2019-03-29 11:22 | 1553872942.8441 | -777 | 0.20829 | 0.000513 | 0 | 90.2989 | 88.7076 | 129.0324 | 20.8588 | 0.001574 | 0.001477 | -152.0784 | 0.22627 | 8.4014 | 0.002024 | 0.004032 | 0.003891 | 6.5e-05 | 0.000701 | 75.4629 | 1.3234 | 2.5474 | 100.29 | 20.87 | 32.164 | 59.3041 | 70.1324 | 56.9548 | 0.007921 | 1.0695 | 0.025446 | 23.7924 | 0.15903 | 0.13461 | -7.6e-05 | -0.005749 | -0.005749 | -0.95305
2019-03-29 11:22 | 1553872943.843 | -776 | 0.20828 | 0.000514 | 0 | 90.2474 | 88.8714 | 129.5232 | 20.8594 | 0.001608 | 0.001524 | -152.0902 | 0.22598 | 8.4006 | 0.00206 | 0.003291 | 0.003099 | 6.8e-05 | 0.000699 | 75.5029 | 1.3408 | 2.5474 | 100.29 | 20.87 | 32.158 | 60.1368 | 70.9953 | 58.0913 | 0.00792 | 1.0658 | 0.02545 | 23.8787 | 0.13743 | 0.13444 | -7.7e-05 | -0.005826 | -0.005826 | -0.9392
2019-03-29 11:22 | 1553872944.844 | -775 | 0.20826 | 0.000516 | 0 | 90.2651 | 88.7714 | 129.4574 | 20.8599 | 0.001601 | 0.00151 | -152.1022 | 0.22637 | 8.4033 | 0.002 | 0.003282 | 0.003026 | 6.6e-05 | 0.000701 | 75.4906 | 1.3355 | 2.5474 | 100.29 | 20.87 | 32.147 | 59.9638 | 69.5548 | 57.7647 | 0.007917 | 1.0669 | 0.025457 | 23.8608 | 0.14925 | 0.13467 | -7.9e-05 | -0.005997 | -0.005997 | -0.91396
2019-03-29 11:22 | 1553872945.8441 | -774 | 0.20828 | 0.000513 | 0 | 90.1893 | 88.597 | 129.2833 | 20.8597 | 0.001604 | 0.001509 | -152.1019 | 0.22576 | 8.4011 | 0.002125 | 0.003364 | 0.003223 | 6.3e-05 | 0.000696 | 75.4847 | 1.3562 | 2.5475 | 100.29 | 20.87 | 32.141 | 60.0371 | 72.5572 | 57.744 | 0.007916 | 1.0669 | 0.02544 | 23.844 | 0.15933 | 0.13431 | -7.6e-05 | -0.005764 | -0.005764 | -0.94832
2019-03-29 11:22 | 1553872946.8441 | -773 | 0.2083 | 0.000512 | 0 | 90.2857 | 88.7083 | 128.9713 | 20.8599 | 0.001601 | 0.001459 | -152.1192 | 0.22612 | 8.4024 | 0.002056 | 0.003465 | 0.003361 | 6.5e-05 | 0.000702 | 75.4871 | 1.3769 | 2.5475 | 100.29 | 20.87 | 32.137 | 59.9582 | 70.9124 | 56.5306 | 0.007915 | 1.0709 | 0.025456 | 23.7715 | 0.15766 | 0.13452 | -7.4e-05 | -0.005659 | -0.005659 | -0.96751
2019-03-29 11:22 | 1553872947.8421 | -772 | 0.20827 | 0.000513 | 0 | 90.3091 | 88.799 | 128.7368 | 20.861 | 0.001658 | 0.001539 | -152.1193 | 0.22512 | 8.4045 | 0.002054 | 0.003786 | 0.003629 | 6.6e-05 | 0.000694 | 75.4535 | 1.343 | 2.5474 | 100.29 | 20.88 | 32.13 | 61.3381 | 70.8456 | 58.4614 | 0.007918 | 1.0646 | 0.025359 | 23.8194 | 0.15082 | 0.13392 | -7.6e-05 | -0.005733 | -0.005733 | -0.95071
2019-03-29 11:22 | 1553872948.8422 | -771 | 0.20828 | 0.000514 | 0 | 90.3088 | 88.6717 | 128.8164 | 20.8613 | 0.001593 | 0.001487 | -152.1016 | 0.22592 | 8.4024 | 0.002103 | 0.002836 | 0.002725 | 6.4e-05 | 0.000693 | 75.507 | 1.3553 | 2.5474 | 100.29 | 20.88 | 32.126 | 59.7703 | 72.034 | 57.2118 | 0.007917 | 1.0687 | 0.025415 | 23.7818 | 0.16362 | 0.1344 | -7.7e-05 | -0.005817 | -0.005817 | -0.94038
2019-03-29 11:22 | 1553872949.8433 | -770 | 0.20827 | 0.000507 | -1e-06 | 90.286 | 88.8024 | 128.9579 | 20.8609 | 0.001604 | 0.00149 | -152.0888 | 0.22566 | 8.4019 | 0.002089 | 0.003709 | 0.003411 | 6.5e-05 | 0.000691 | 75.4223 | 1.3587 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.0406 | 71.7084 | 57.292 | 0.007912 | 1.0684 | 0.025426 | 23.7977 | 0.1482 | 0.13425 | -6.8e-05 | -0.005126 | -0.005126 | -1.0659
2019-03-29 11:22 | 1553872950.8423 | -769 | 0.20829 | 0.000515 | 0 | 90.3944 | 88.8461 | 129.1375 | 20.8608 | 0.001567 | 0.001531 | -152.098 | 0.22545 | 8.4051 | 0.00204 | 0.003716 | 0.003542 | 6.5e-05 | 0.000689 | 75.4377 | 1.3797 | 2.5475 | 100.29 | 20.87 | 32.124 | 59.1294 | 70.5311 | 58.2625 | 0.007912 | 1.0653 | 0.025406 | 23.8492 | 0.15452 | 0.13412 | -7.8e-05 | -0.00592 | -0.00592 | -0.92202
2019-03-29 11:22 | 1553872951.8424 | -768 | 0.20828 | 0.000514 | -1e-06 | 90.4788 | 88.9591 | 129.4206 | 20.8619 | 0.001618 | 0.001477 | -152.0803 | 0.22707 | 8.4045 | 0.002023 | 0.002482 | 0.002349 | 6.4e-05 | 0.00069 | 75.4243 | 1.3438 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.3746 | 70.1035 | 56.9759 | 0.007912 | 1.0694 | 0.025483 | 23.829 | 0.15151 | 0.13508 | -7.7e-05 | -0.005788 | -0.005788 | -0.94985
2019-03-29 11:23 | 1553872952.8424 | -767 | 0.20826 | 0.00051 | -1e-06 | 90.5889 | 89.1379 | 129.4722 | 20.8624 | 0.001573 | 0.001479 | -152.0726 | 0.22624 | 8.4031 | 0.002043 | 0.003376 | 0.003157 | 6.3e-05 | 0.000682 | 75.3762 | 1.3304 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.2936 | 70.5951 | 57.0162 | 0.007912 | 1.0693 | 0.025487 | 23.8352 | 0.14443 | 0.13459 | -7.1e-05 | -0.005382 | -0.005382 | -1.0178
2019-03-29 11:23 | 1553872953.8423 | -766 | 0.2083 | 0.000509 | -1e-06 | 90.6335 | 89.1335 | 129.1288 | 20.8621 | 0.00162 | 0.001494 | -152.0759 | 0.22715 | 8.4043 | 0.002123 | 0.003461 | 0.00333 | 6.6e-05 | 0.000689 | 75.4094 | 1.3484 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.4271 | 72.5126 | 57.3769 | 0.007912 | 1.0681 | 0.025439 | 23.8166 | 0.14928 | 0.13513 | -7.1e-05 | -0.005348 | -0.005348 | -1.0284
2019-03-29 11:23 | 1553872954.8434 | -765 | 0.20828 | 0.000511 | 0 | 90.671 | 89.2253 | 129.084 | 20.8624 | 0.001588 | 0.001502 | -152.0752 | 0.22611 | 8.4014 | 0.002147 | 0.003087 | 0.002953 | 6.9e-05 | 0.000692 | 75.3712 | 1.381 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.6377 | 73.1048 | 57.5792 | 0.007912 | 1.0675 | 0.025427 | 23.8197 | 0.14376 | 0.13451 | -7.4e-05 | -0.005596 | -0.005596 | -0.97836
2019-03-29 11:23 | 1553872955.8434 | -764 | 0.20827 | 0.000513 | -1e-06 | 90.7263 | 89.2529 | 129.3816 | 20.8624 | 0.001657 | 0.001468 | -152.0821 | 0.22659 | 8.4031 | 0.002091 | 0.002949 | 0.002745 | 6.4e-05 | 0.000688 | 75.4307 | 1.4204 | 2.5476 | 100.29 | 20.87 | 32.124 | 61.3085 | 71.7419 | 56.7469 | 0.007912 | 1.0702 | 0.025488 | 23.8171 | 0.14645 | 0.1348 | -7.5e-05 | -0.005681 | -0.005681 | -0.96578
2019-03-29 11:23 | 1553872956.8434 | -763 | 0.20828 | 0.000517 | 0 | 90.7718 | 89.3471 | 129.8971 | 20.8619 | 0.00161 | 0.001518 | -152.0681 | 0.22709 | 8.4046 | 0.002087 | 0.003755 | 0.003556 | 6.3e-05 | 0.000693 | 75.3873 | 1.3717 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.173 | 71.6441 | 57.9686 | 0.007912 | 1.0662 | 0.025492 | 23.9087 | 0.1415 | 0.1351 | -8.1e-05 | -0.00615 | -0.00615 | -0.89408
2019-03-29 11:23 | 1553872957.8424 | -762 | 0.20827 | 0.000507 | 0 | 90.8107 | 89.2434 | 130.0757 | 20.8635 | 0.00163 | 0.00147 | -152.0674 | 0.22716 | 8.4001 | 0.002095 | 0.002895 | 0.002745 | 6.5e-05 | 0.000694 | 75.3878 | 1.348 | 2.5472 | 100.29 | 20.87 | 32.124 | 60.6685 | 71.8513 | 56.7943 | 0.007912 | 1.07 | 0.025555 | 23.8826 | 0.15573 | 0.13514 | -6.8e-05 | -0.005203 | -0.005203 | -1.057
2019-03-29 11:23 | 1553872958.8435 | -761 | 0.20828 | 0.00051 | 0 | 90.8685 | 89.4802 | 129.6957 | 20.8642 | 0.001537 | 0.001459 | -152.0712 | 0.22619 | 8.3999 | 0.002123 | 0.002865 | 0.002646 | 6.2e-05 | 0.000692 | 75.3916 | 1.3368 | 2.5473 | 100.29 | 20.87 | 32.124 | 58.43 | 72.5209 | 56.5347 | 0.007912 | 1.0709 | 0.025527 | 23.8383 | 0.13771 | 0.13456 | -7.2e-05 | -0.005505 | -0.005505 | -0.99486
2019-03-29 11:23 | 1553872959.8425 | -760 | 0.20827 | 0.000513 | 0 | 90.9374 | 89.4144 | 129.4596 | 20.8637 | 0.001545 | 0.001499 | -152.0678 | 0.22626 | 8.4019 | 0.002079 | 0.003711 | 0.003529 | 6.6e-05 | 0.000694 | 75.4293 | 1.4028 | 2.5474 | 100.29 | 20.87 | 32.124 | 58.601 | 71.4532 | 57.4919 | 0.007912 | 1.0678 | 0.025467 | 23.8512 | 0.15107 | 0.1346 | -7.7e-05 | -0.005847 | -0.005847 | -0.93691
2019-03-29 11:23 | 1553872960.8436 | -759 | 0.20826 | 0.000516 | 0 | 90.9365 | 89.3997 | 129.5304 | 20.8643 | 0.00161 | 0.00152 | -152.065 | 0.22712 | 8.4019 | 0.002047 | 0.003224 | 0.003033 | 6.6e-05 | 0.000691 | 75.4203 | 1.354 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.1878 | 70.7023 | 58.0194 | 0.007912 | 1.0661 | 0.025454 | 23.8767 | 0.15245 | 0.13512 | -8e-05 | -0.00614 | -0.00614 | -0.89565
2019-03-29 11:23 | 1553872961.8428 | -758 | 0.2083 | 0.000513 | 0 | 91.0269 | 89.6809 | 129.4858 | 20.8641 | 0.001627 | 0.001545 | -152.0598 | 0.22632 | 8.4011 | 0.002086 | 0.003363 | 0.003063 | 6.2e-05 | 0.000689 | 75.4666 | 1.3296 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.5784 | 71.6226 | 58.6177 | 0.007912 | 1.0641 | 0.025427 | 23.8942 | 0.13324 | 0.13464 | -7.6e-05 | -0.005764 | -0.005764 | -0.95074
2019-03-29 11:23 | 1553872962.8426 | -757 | 0.20825 | 0.000512 | 0 | 91.0753 | 89.6321 | 129.2192 | 20.8658 | 0.001626 | 0.001481 | -152.0466 | 0.22687 | 8.4032 | 0.002054 | 0.003472 | 0.00339 | 6.6e-05 | 0.000683 | 75.4507 | 1.3476 | 2.5472 | 100.29 | 20.87 | 32.124 | 60.5536 | 70.8618 | 57.0736 | 0.007912 | 1.0691 | 0.02546 | 23.8139 | 0.14288 | 0.13497 | -7.4e-05 | -0.005655 | -0.005655 | -0.97131
2019-03-29 11:23 | 1553872963.8418 | -756 | 0.20826 | 0.000515 | 0 | 91.0628 | 89.7253 | 129.1833 | 20.865 | 0.0016 | 0.001506 | -152.0411 | 0.22728 | 8.4039 | 0.002086 | 0.003601 | 0.003412 | 6.7e-05 | 0.000682 | 75.4602 | 1.3688 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.9359 | 71.6266 | 57.6818 | 0.007912 | 1.0671 | 0.025433 | 23.8326 | 0.13235 | 0.13521 | -7.9e-05 | -0.005971 | -0.005971 | -0.92167
2019-03-29 11:23 | 1553872964.8418 | -755 | 0.20824 | 0.000512 | 0 | 91.1592 | 89.9526 | 129.3903 | 20.8659 | 0.001598 | 0.001473 | -152.0395 | 0.22695 | 8.3998 | 0.002177 | 0.002858 | 0.002664 | 6.1e-05 | 0.000681 | 75.4395 | 1.3298 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.882 | 73.8087 | 56.8639 | 0.007912 | 1.0698 | 0.025485 | 23.8221 | 0.11918 | 0.13501 | -7.5e-05 | -0.005667 | -0.005667 | -0.96964
2019-03-29 11:23 | 1553872965.8421 | -754 | 0.20824 | 0.000515 | 0 | 91.3274 | 89.8724 | 129.4556 | 20.8652 | 0.00163 | 0.001449 | -152.0306 | 0.227 | 8.402 | 0.002072 | 0.003191 | 0.002988 | 6.4e-05 | 0.000681 | 75.4471 | 1.3383 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.6678 | 71.2972 | 56.2932 | 0.007912 | 1.0716 | 0.025513 | 23.8075 | 0.14365 | 0.13504 | -7.8e-05 | -0.005943 | -0.005943 | -0.92479
2019-03-29 11:23 | 1553872966.8418 | -753 | 0.20825 | 0.000512 | 0 | 91.4061 | 90.0867 | 129.5152 | 20.8661 | 0.001602 | 0.001547 | -152.0189 | 0.22739 | 8.403 | 0.002049 | 0.003857 | 0.00367 | 6.6e-05 | 0.000683 | 75.423 | 1.2974 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.9962 | 70.746 | 58.6518 | 0.007912 | 1.064 | 0.025428 | 23.8981 | 0.13005 | 0.13527 | -7.4e-05 | -0.005628 | -0.005628 | -0.97821
2019-03-29 11:23 | 1553872967.8429 | -752 | 0.20826 | 0.000515 | 0 | 91.4702 | 90.2154 | 129.4668 | 20.8658 | 0.001614 | 0.001458 | -152.0239 | 0.22659 | 8.4025 | 0.002088 | 0.003313 | 0.003104 | 6.5e-05 | 0.000681 | 75.4087 | 1.3036 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.2861 | 71.683 | 56.5185 | 0.007912 | 1.0709 | 0.025506 | 23.8167 | 0.12355 | 0.1348 | -7.8e-05 | -0.005947 | -0.005947 | -0.92246
2019-03-29 11:23 | 1553872968.8429 | -751 | 0.20825 | 0.000512 | 0 | 91.5597 | 90.3014 | 129.644 | 20.8664 | 0.00162 | 0.001444 | -152.0137 | 0.22683 | 8.402 | 0.002114 | 0.003632 | 0.003395 | 6.6e-05 | 0.000683 | 75.4401 | 1.4048 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.4236 | 72.3157 | 56.1649 | 0.007912 | 1.0721 | 0.025537 | 23.8202 | 0.12377 | 0.13494 | -7.5e-05 | -0.005756 | -0.005756 | -0.95416
2019-03-29 11:23 | 1553872969.8419 | -750 | 0.20826 | 0.000514 | 0 | 91.4974 | 90.2553 | 129.8102 | 20.8668 | 0.001581 | 0.001481 | -152.0012 | 0.22613 | 8.402 | 0.002009 | 0.003845 | 0.003695 | 6.6e-05 | 0.000682 | 75.4413 | 1.3829 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.484 | 69.7852 | 57.0699 | 0.007912 | 1.0691 | 0.025518 | 23.8682 | 0.12226 | 0.13453 | -7.7e-05 | -0.005841 | -0.005841 | -0.93741
2019-03-29 11:23 | 1553872970.8419 | -749 | 0.20826 | 0.000513 | 0 | 91.5917 | 90.3171 | 129.4776 | 20.8669 | 0.001655 | 0.001534 | -152.0086 | 0.22648 | 8.4029 | 0.002094 | 0.003759 | 0.00365 | 6.6e-05 | 0.00068 | 75.3801 | 1.4138 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.2627 | 71.8164 | 58.3434 | 0.007912 | 1.065 | 0.025436 | 23.8835 | 0.12535 | 0.13473 | -7.6e-05 | -0.005743 | -0.005743 | -0.95477
2019-03-29 11:23 | 1553872971.8434 | -748 | 0.20823 | 0.000516 | 0 | 91.5402 | 90.3212 | 129.1021 | 20.8675 | 0.001567 | 0.001492 | -152.028 | 0.22727 | 8.4031 | 0.002033 | 0.00307 | 0.002834 | 6.1e-05 | 0.000683 | 75.5188 | 1.3575 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.1554 | 70.3678 | 57.3446 | 0.007912 | 1.0682 | 0.025438 | 23.8129 | 0.11991 | 0.1352 | -7.9e-05 | -0.006037 | -0.006037 | -0.91157
2019-03-29 11:23 | 1553872972.843 | -747 | 0.20827 | 0.000511 | 0 | 91.6001 | 90.2683 | 129.2974 | 20.8667 | 0.001598 | 0.001514 | -152.0434 | 0.22687 | 8.4019 | 0.002152 | 0.003208 | 0.003108 | 6.3e-05 | 0.000681 | 75.514 | 1.3367 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.8797 | 73.2249 | 57.8707 | 0.007912 | 1.0665 | 0.025437 | 23.8499 | 0.131 | 0.13497 | -7.3e-05 | -0.005573 | -0.005573 | -0.98569
2019-03-29 11:23 | 1553872973.8421 | -746 | 0.20824 | 0.000514 | 0 | 91.6605 | 90.4526 | 129.4912 | 20.8676 | 0.001582 | 0.001485 | -152.0457 | 0.22632 | 8.401 | 0.002147 | 0.003884 | 0.003683 | 6.1e-05 | 0.000682 | 75.4912 | 1.3254 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.5137 | 73.0987 | 57.1692 | 0.007912 | 1.0688 | 0.025483 | 23.8425 | 0.11866 | 0.13464 | -7.7e-05 | -0.005883 | -0.005883 | -0.93147
2019-03-29 11:23 | 1553872974.8431 | -745 | 0.20823 | 0.000512 | 0 | 91.7611 | 90.3454 | 129.2985 | 20.868 | 0.001655 | 0.001541 | -152.0538 | 0.22693 | 8.4026 | 0.002155 | 0.002879 | 0.002666 | 6.3e-05 | 0.00069 | 75.4413 | 1.3107 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.2708 | 73.2971 | 58.5121 | 0.007912 | 1.0645 | 0.025412 | 23.8731 | 0.13907 | 0.135 | -7.4e-05 | -0.005622 | -0.005622 | -0.97735
2019-03-29 11:23 | 1553872975.8421 | -744 | 0.20826 | 0.000514 | 0 | 91.7168 | 90.4888 | 128.78 | 20.8681 | 0.00163 | 0.001444 | -152.0529 | 0.22615 | 8.4018 | 0.002116 | 0.003016 | 0.002814 | 6.4e-05 | 0.0007 | 75.4327 | 1.3203 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.6628 | 72.3476 | 56.1822 | 0.007912 | 1.072 | 0.025451 | 23.7413 | 0.12057 | 0.13454 | -7.8e-05 | -0.005896 | -0.005896 | -0.92865
2019-03-29 11:23 | 1553872976.8431 | -743 | 0.20825 | 0.000512 | 0 | 91.7833 | 90.4539 | 128.0982 | 20.8683 | 0.001608 | 0.001532 | -152.067 | 0.22596 | 8.4019 | 0.0021 | 0.003067 | 0.002914 | 6.4e-05 | 0.000703 | 75.3656 | 1.3165 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.1308 | 71.9598 | 58.2925 | 0.007912 | 1.0652 | 0.025302 | 23.7541 | 0.13049 | 0.13442 | -7.4e-05 | -0.005596 | -0.005596 | -0.97772
2019-03-29 11:23 | 1553872977.8421 | -742 | 0.20823 | 0.000509 | 0 | 91.7695 | 90.512 | 128.1376 | 20.8687 | 0.001601 | 0.001487 | -152.058 | 0.22642 | 8.4023 | 0.002157 | 0.003539 | 0.003375 | 6.3e-05 | 0.000699 | 75.4498 | 1.3245 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.9605 | 73.3367 | 57.2072 | 0.007912 | 1.0687 | 0.025348 | 23.7189 | 0.12342 | 0.1347 | -7.1e-05 | -0.005427 | -0.005427 | -1.0101
2019-03-29 11:23 | 1553872978.8412 | -741 | 0.20831 | 0.000517 | 0 | 91.7318 | 90.5079 | 128.0336 | 20.8688 | 0.0016 | 0.001485 | -152.0626 | 0.22712 | 8.4011 | 0.002055 | 0.003619 | 0.003518 | 6.5e-05 | 0.000703 | 75.4514 | 1.3513 | 2.5475 | 100.29 | 20.87 | 32.124 | 59.9303 | 70.8992 | 57.162 | 0.007912 | 1.0688 | 0.025339 | 23.7076 | 0.12014 | 0.13512 | -8.2e-05 | -0.006224 | -0.006224 | -0.88358
2019-03-29 11:23 | 1553872979.8416 | -740 | 0.20825 | 0.000511 | 0 | 91.7548 | 90.5556 | 128.3071 | 20.8695 | 0.001673 | 0.001533 | -152.0854 | 0.22687 | 8.4055 | 0.002037 | 0.003575 | 0.003307 | 6.6e-05 | 0.000704 | 75.4083 | 1.3519 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.6935 | 70.4499 | 58.3382 | 0.007912 | 1.065 | 0.025321 | 23.7751 | 0.11767 | 0.13497 | -7.3e-05 | -0.005518 | -0.005518 | -0.99547
2019-03-29 11:23 | 1553872980.8413 | -739 | 0.20826 | 0.000508 | 0 | 91.7916 | 90.3973 | 128.3552 | 20.8694 | 0.001619 | 0.001529 | -152.0728 | 0.22619 | 8.4021 | 0.002054 | 0.003178 | 0.00302 | 6.4e-05 | 0.000706 | 75.4304 | 1.3329 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.407 | 70.8733 | 58.2301 | 0.007912 | 1.0654 | 0.02533 | 23.7757 | 0.13691 | 0.13456 | -7e-05 | -0.005281 | -0.005281 | -1.0371
2019-03-29 11:23 | 1553872981.8413 | -738 | 0.20824 | 0.000511 | 0 | 91.7743 | 90.5202 | 128.3451 | 20.8692 | 0.001622 | 0.001497 | -152.0521 | 0.22637 | 8.4012 | 0.002071 | 0.003061 | 0.002841 | 6.2e-05 | 0.000705 | 75.3988 | 1.3751 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.4842 | 71.2809 | 57.4473 | 0.007912 | 1.0679 | 0.025359 | 23.7467 | 0.12307 | 0.13467 | -7.4e-05 | -0.005583 | -0.005583 | -0.98179
2019-03-29 11:23 | 1553872982.8417 | -737 | 0.20826 | 0.000514 | 0 | 91.7872 | 90.4304 | 127.997 | 20.8704 | 0.001634 | 0.001517 | -152.0578 | 0.22614 | 8.4021 | 0.00207 | 0.003412 | 0.003264 | 6.3e-05 | 0.00071 | 75.4702 | 1.3801 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.7557 | 71.2523 | 57.9353 | 0.007912 | 1.0663 | 0.025306 | 23.732 | 0.1332 | 0.13453 | -7.7e-05 | -0.005847 | -0.005847 | -0.93647
2019-03-29 11:23 | 1553872983.8427 | -736 | 0.20825 | 0.000516 | 0 | 91.7773 | 90.4782 | 128.2052 | 20.8706 | 0.001595 | 0.001523 | -152.0597 | 0.2259 | 8.4016 | 0.00207 | 0.003545 | 0.00337 | 6.5e-05 | 0.000712 | 75.4078 | 1.3949 | 2.5475 | 100.29 | 20.87 | 32.124 | 59.8122 | 71.2587 | 58.0984 | 0.007912 | 1.0658 | 0.02532 | 23.7571 | 0.12751 | 0.13439 | -7.9e-05 | -0.005937 | -0.005937 | -0.92127
2019-03-29 11:23 | 1553872984.8424 | -735 | 0.20828 | 0.000512 | 0 | 91.8093 | 90.514 | 128.1904 | 20.8707 | 0.001611 | 0.001509 | -152.0472 | 0.22608 | 8.4031 | 0.002038 | 0.002875 | 0.002677 | 6e-05 | 0.00071 | 75.4408 | 1.348 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.2126 | 70.4838 | 57.7564 | 0.007912 | 1.0669 | 0.025332 | 23.7435 | 0.12709 | 0.13449 | -7.4e-05 | -0.005584 | -0.005584 | -0.98035
2019-03-29 11:23 | 1553872985.8418 | -734 | 0.20828 | 0.00051 | 0 | 91.821 | 90.6686 | 128.3477 | 20.8711 | 0.001617 | 0.001535 | -152.0311 | 0.22631 | 8.4021 | 0.002039 | 0.003153 | 0.002954 | 6.5e-05 | 0.000709 | 75.4311 | 1.3994 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.3561 | 70.5092 | 58.3743 | 0.007912 | 1.0649 | 0.025324 | 23.7802 | 0.11297 | 0.13463 | -7.1e-05 | -0.005373 | -0.005373 | -1.0198
2019-03-29 11:23 | 1553872986.8425 | -733 | 0.20828 | 0.000513 | 0 | 91.8826 | 90.6577 | 128.7223 | 20.8708 | 0.001677 | 0.001557 | -152.0241 | 0.22722 | 8.4025 | 0.001991 | 0.003379 | 0.003124 | 6.2e-05 | 0.000713 | 75.3934 | 1.361 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.794 | 69.3502 | 58.9014 | 0.007912 | 1.0632 | 0.025341 | 23.8338 | 0.12005 | 0.13517 | -7.6e-05 | -0.005719 | -0.005719 | -0.96202
2019-03-29 11:23 | 1553872987.8424 | -732 | 0.20826 | 0.000515 | 0 | 91.9266 | 90.4963 | 128.9023 | 20.871 | 0.001572 | 0.001528 | -152.0112 | 0.2269 | 8.4026 | 0.00207 | 0.00342 | 0.003257 | 6.4e-05 | 0.000712 | 75.3102 | 1.3771 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.2613 | 71.2413 | 58.1989 | 0.007912 | 1.0655 | 0.025385 | 23.8252 | 0.14027 | 0.13498 | -7.7e-05 | -0.005846 | -0.005846 | -0.93981
2019-03-29 11:23 | 1553872988.8415 | -731 | 0.20827 | 0.000515 | 0 | 91.9093 | 90.6638 | 129.1328 | 20.8713 | 0.001686 | 0.001528 | -152.0153 | 0.22802 | 8.4023 | 0.002033 | 0.003637 | 0.003511 | 6.2e-05 | 0.000712 | 75.4541 | 1.381 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.0102 | 70.3707 | 58.2155 | 0.007912 | 1.0654 | 0.025407 | 23.8471 | 0.12204 | 0.13565 | -7.8e-05 | -0.005896 | -0.005896 | -0.93638
2019-03-29 11:23 | 1553872989.8415 | -730 | 0.20828 | 0.000507 | -1e-06 | 91.9063 | 90.6051 | 128.796 | 20.8716 | 0.0016 | 0.001505 | -152.0073 | 0.22687 | 8.4034 | 0.002006 | 0.003687 | 0.003487 | 6.4e-05 | 0.000707 | 75.4431 | 1.3305 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.9532 | 69.7065 | 57.6562 | 0.007912 | 1.0672 | 0.025396 | 23.7959 | 0.12754 | 0.13496 | -6.8e-05 | -0.005131 | -0.005131 | -1.0706
2019-03-29 11:23 | 1553872990.8429 | -729 | 0.20828 | 0.000514 | 0 | 91.9446 | 90.6345 | 129.0162 | 20.873 | 0.001699 | 0.001524 | -151.9974 | 0.22624 | 8.4012 | 0.002038 | 0.002599 | 0.002481 | 6.3e-05 | 0.000708 | 75.4743 | 1.3222 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3297 | 70.4891 | 58.1194 | 0.007912 | 1.0657 | 0.0254 | 23.8329 | 0.12837 | 0.13459 | -7.6e-05 | -0.005779 | -0.005779 | -0.94785
2019-03-29 11:23 | 1553872991.8425 | -728 | 0.20828 | 0.00051 | 0 | 91.953 | 90.6584 | 129.2847 | 20.8728 | 0.001639 | 0.001558 | -152.0118 | 0.22625 | 8.4015 | 0.001984 | 0.003534 | 0.003349 | 6.2e-05 | 0.000707 | 75.4469 | 1.3077 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.8965 | 69.1954 | 58.94 | 0.007912 | 1.0631 | 0.025395 | 23.8872 | 0.12683 | 0.13459 | -7.2e-05 | -0.005425 | -0.005425 | -1.0099
2019-03-29 11:23 | 1553872992.8416 | -727 | 0.20828 | 0.000515 | 0 | 91.9689 | 90.7761 | 129.4955 | 20.8728 | 0.001684 | 0.00155 | -152.0056 | 0.22661 | 8.4037 | 0.002036 | 0.003244 | 0.003071 | 6.5e-05 | 0.000704 | 75.4214 | 1.3595 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.9631 | 70.4406 | 58.7468 | 0.007912 | 1.0637 | 0.025423 | 23.8997 | 0.11676 | 0.13481 | -7.8e-05 | -0.005951 | -0.005951 | -0.92202
2019-03-29 11:23 | 1553872993.8421 | -726 | 0.20828 | 0.000512 | 0 | 91.9081 | 90.8311 | 129.342 | 20.873 | 0.001614 | 0.001499 | -152.0002 | 0.22657 | 8.4031 | 0.002005 | 0.003089 | 0.002892 | 6.5e-05 | 0.000701 | 75.4088 | 1.3456 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.2791 | 69.6806 | 57.505 | 0.007912 | 1.0677 | 0.025455 | 23.8408 | 0.10544 | 0.13478 | -7.5e-05 | -0.005725 | -0.005725 | -0.95812
2019-03-29 11:23 | 1553872994.8418 | -725 | 0.20829 | 0.000514 | -1e-06 | 91.8776 | 90.5584 | 129.4251 | 20.8735 | 0.001708 | 0.001501 | -152.0067 | 0.22666 | 8.4042 | 0.002105 | 0.003105 | 0.0029 | 6.2e-05 | 0.000697 | 75.4239 | 1.3494 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.5474 | 72.0895 | 57.5523 | 0.007912 | 1.0676 | 0.025462 | 23.8502 | 0.12936 | 0.13484 | -7.7e-05 | -0.005796 | -0.005796 | -0.94679
2019-03-29 11:23 | 1553872995.8407 | -724 | 0.20831 | 0.000518 | 0 | 91.797 | 90.5748 | 129.7229 | 20.8737 | 0.001602 | 0.001507 | -152.0223 | 0.22715 | 8.4024 | 0.002073 | 0.003289 | 0.003198 | 6.7e-05 | 0.000699 | 75.3717 | 1.3319 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.9998 | 71.3217 | 57.7059 | 0.007912 | 1.0671 | 0.025485 | 23.8832 | 0.1199 | 0.13513 | -8.1e-05 | -0.006183 | -0.006183 | -0.88946
2019-03-29 11:23 | 1553872996.8421 | -723 | 0.20832 | 0.000508 | 0 | 91.7201 | 90.5734 | 129.5471 | 20.8746 | 0.001644 | 0.001549 | -152.0207 | 0.22704 | 8.4019 | 0.0021 | 0.003144 | 0.002949 | 6.5e-05 | 0.000705 | 75.3861 | 1.3796 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.9958 | 71.9673 | 58.7155 | 0.007912 | 1.0638 | 0.025429 | 23.9034 | 0.11253 | 0.13507 | -6.9e-05 | -0.005219 | -0.005219 | -1.0533
2019-03-29 11:23 | 1553872997.8408 | -722 | 0.2083 | 0.000514 | 0 | 91.6881 | 90.3963 | 129.0678 | 20.8739 | 0.001629 | 0.00152 | -152.0324 | 0.22661 | 8.4031 | 0.002114 | 0.00344 | 0.003296 | 6.3e-05 | 0.000704 | 75.3852 | 1.3489 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.6381 | 72.3205 | 58.0255 | 0.007912 | 1.066 | 0.025408 | 23.8343 | 0.12692 | 0.13481 | -7.7e-05 | -0.005871 | -0.005871 | -0.93461
2019-03-29 11:23 | 1553872998.8408 | -721 | 0.20828 | 0.000513 | 0 | 91.6073 | 90.3096 | 129.2237 | 20.8759 | 0.001615 | 0.001486 | -152.0286 | 0.22732 | 8.4008 | 0.002115 | 0.003394 | 0.003269 | 6.6e-05 | 0.000703 | 75.4279 | 1.3335 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.3217 | 72.3461 | 57.1884 | 0.007912 | 1.0687 | 0.025456 | 23.8185 | 0.12761 | 0.13523 | -7.6e-05 | -0.005809 | -0.005809 | -0.94754
2019-03-29 11:23 | 1553872999.8412 | -720 | 0.20823 | 0.000513 | 0 | 91.5821 | 90.2195 | 129.395 | 20.8745 | 0.001592 | 0.001502 | -152.034 | 0.22724 | 8.4001 | 0.002107 | 0.002871 | 0.002592 | 6.5e-05 | 0.000702 | 75.425 | 1.34 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.7651 | 72.1474 | 57.5885 | 0.007912 | 1.0674 | 0.025457 | 23.8487 | 0.13409 | 0.13519 | -7.6e-05 | -0.005836 | -0.005836 | -0.94286
2019-03-29 11:23 | 1553873000.8418 | -719 | 0.20827 | 0.000515 | 0 | 91.5926 | 90.3423 | 129.2582 | 20.8754 | 0.001651 | 0.001512 | -152.05 | 0.2272 | 8.4029 | 0.002162 | 0.003351 | 0.003212 | 6.3e-05 | 0.000695 | 75.4302 | 1.3482 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.1874 | 73.4751 | 57.8234 | 0.007912 | 1.0667 | 0.025435 | 23.8446 | 0.12294 | 0.13516 | -7.9e-05 | -0.006021 | -0.006021 | -0.91362
2019-03-29 11:23 | 1553873001.8409 | -718 | 0.20828 | 0.000518 | 0 | 91.6399 | 90.3369 | 128.9772 | 20.8756 | 0.00159 | 0.001553 | -152.0345 | 0.22692 | 8.402 | 0.002142 | 0.0037 | 0.0035 | 6.5e-05 | 0.000692 | 75.4241 | 1.3413 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.7155 | 72.9842 | 58.8127 | 0.007912 | 1.0635 | 0.025369 | 23.8542 | 0.12809 | 0.13499 | -8.2e-05 | -0.006231 | -0.006231 | -0.8817
2019-03-29 11:23 | 1553873002.8409 | -717 | 0.20831 | 0.000509 | 0 | 91.6474 | 90.2615 | 128.8631 | 20.8752 | 0.00164 | 0.001575 | -152.0206 | 0.22707 | 8.4017 | 0.002115 | 0.00332 | 0.00312 | 6.7e-05 | 0.000692 | 75.4009 | 1.3728 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.902 | 72.3399 | 59.3389 | 0.007912 | 1.0618 | 0.025338 | 23.8625 | 0.13629 | 0.13508 | -7e-05 | -0.005265 | -0.005265 | -1.0442
2019-03-29 11:23 | 1553873003.841 | -716 | 0.20826 | 0.000512 | -1e-06 | 91.5666 | 90.1277 | 128.9596 | 20.8764 | 0.001685 | 0.001548 | -152.0258 | 0.22692 | 8.4022 | 0.002131 | 0.003403 | 0.003261 | 7e-05 | 0.000696 | 75.3963 | 1.3528 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.9996 | 72.7231 | 58.6947 | 0.007912 | 1.0639 | 0.025372 | 23.8483 | 0.14167 | 0.13499 | -7.4e-05 | -0.005542 | -0.005542 | -0.99139
2019-03-29 11:23 | 1553873004.8414 | -715 | 0.20828 | 0.000514 | 0 | 91.6492 | 90.3696 | 128.689 | 20.876 | 0.001693 | 0.001508 | -152.0325 | 0.22739 | 8.4029 | 0.00211 | 0.003688 | 0.003461 | 6.9e-05 | 0.000691 | 75.3979 | 1.3621 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.1976 | 72.2222 | 57.7208 | 0.007912 | 1.067 | 0.025383 | 23.7883 | 0.12576 | 0.13527 | -7.6e-05 | -0.005748 | -0.005748 | -0.95774
2019-03-29 11:23 | 1553873005.841 | -714 | 0.20828 | 0.000512 | 0 | 91.6255 | 90.5215 | 128.5955 | 20.8772 | 0.00161 | 0.001496 | -152.03 | 0.22722 | 8.4008 | 0.002137 | 0.003959 | 0.003779 | 7.1e-05 | 0.000692 | 75.4383 | 1.3631 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.19 | 72.8742 | 57.4314 | 0.007912 | 1.068 | 0.025384 | 23.7693 | 0.10843 | 0.13517 | -7.4e-05 | -0.005582 | -0.005582 | -0.98559
2019-03-29 11:23 | 1553873006.8416 | -713 | 0.20827 | 0.000512 | 0 | 91.6369 | 90.0417 | 128.7772 | 20.8767 | 0.001589 | 0.001551 | -152.0195 | 0.22679 | 8.4022 | 0.002093 | 0.002851 | 0.002677 | 6.7e-05 | 0.00069 | 75.4149 | 1.3998 | 2.5476 | 100.29 | 20.87 | 32.124 | 59.6892 | 71.8041 | 58.7786 | 0.007912 | 1.0636 | 0.025351 | 23.8345 | 0.15708 | 0.13492 | -7.4e-05 | -0.005614 | -0.005614 | -0.97815
2019-03-29 11:23 | 1553873007.8421 | -712 | 0.20827 | 0.000512 | 0 | 91.6552 | 90.2526 | 129.0115 | 20.8776 | 0.001632 | 0.001512 | -152.0229 | 0.22754 | 8.4011 | 0.002121 | 0.003297 | 0.00316 | 6.9e-05 | 0.000697 | 75.4491 | 1.3376 | 2.5475 | 100.29 | 20.87 | 32.124 | 60.7229 | 72.4746 | 57.8381 | 0.007912 | 1.0666 | 0.02541 | 23.8223 | 0.13793 | 0.13536 | -7.4e-05 | -0.005578 | -0.005578 | -0.98771
2019-03-29 11:23 | 1553873008.8402 | -711 | 0.20827 | 0.000515 | 0 | 91.633 | 90.2993 | 128.6023 | 20.8779 | 0.001694 | 0.001573 | -152.034 | 0.22696 | 8.4032 | 0.002112 | 0.00367 | 0.00351 | 7.1e-05 | 0.000699 | 75.3881 | 1.3287 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2041 | 72.2644 | 59.3071 | 0.007912 | 1.0619 | 0.025313 | 23.8372 | 0.13114 | 0.13502 | -7.8e-05 | -0.00589 | -0.00589 | -0.93293
2019-03-29 11:23 | 1553873009.8401 | -710 | 0.20825 | 0.000513 | 0 | 91.6279 | 90.3283 | 128.4993 | 20.8782 | 0.001568 | 0.0015 | -152.0422 | 0.22645 | 8.4025 | 0.002118 | 0.003716 | 0.003528 | 6.9e-05 | 0.000704 | 75.3817 | 1.34 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.1909 | 72.4113 | 57.5359 | 0.007912 | 1.0676 | 0.025371 | 23.7641 | 0.12777 | 0.13472 | -7.6e-05 | -0.005745 | -0.005745 | -0.95441
2019-03-29 11:23 | 1553873010.8402 | -709 | 0.20825 | 0.000515 | 0 | 91.6665 | 90.3096 | 128.4592 | 20.8776 | 0.001666 | 0.001541 | -152.0267 | 0.22621 | 8.4033 | 0.00216 | 0.003326 | 0.003172 | 6.7e-05 | 0.000703 | 75.3502 | 1.3778 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.5319 | 73.428 | 58.5299 | 0.007912 | 1.0644 | 0.025329 | 23.7961 | 0.1334 | 0.13457 | -7.8e-05 | -0.005861 | -0.005861 | -0.93445
2019-03-29 11:24 | 1553873011.8403 | -708 | 0.20825 | 0.000512 | 0 | 91.6144 | 90.4713 | 128.9648 | 20.8776 | 0.001614 | 0.001562 | -152.0447 | 0.22626 | 8.4026 | 0.002109 | 0.003735 | 0.003537 | 6.6e-05 | 0.000711 | 75.3957 | 1.3461 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.2819 | 72.2043 | 59.0362 | 0.007912 | 1.0628 | 0.025359 | 23.8611 | 0.11231 | 0.1346 | -7.4e-05 | -0.005651 | -0.005651 | -0.9695
2019-03-29 11:24 | 1553873012.8403 | -707 | 0.20824 | 0.000512 | 0 | 91.5282 | 90.1577 | 129.0314 | 20.8784 | 0.001622 | 0.001552 | -152.0488 | 0.22614 | 8.4027 | 0.002104 | 0.003459 | 0.00332 | 6.7e-05 | 0.000712 | 75.4573 | 1.3131 | 2.5472 | 100.29 | 20.87 | 32.124 | 60.4726 | 72.0818 | 58.7952 | 0.007912 | 1.0636 | 0.025375 | 23.8586 | 0.13495 | 0.13453 | -7.5e-05 | -0.005705 | -0.005705 | -0.95969
2019-03-29 11:24 | 1553873013.8417 | -706 | 0.20825 | 0.000517 | 1e-06 | 91.5435 | 90.199 | 128.9993 | 20.8787 | 0.001601 | 0.001558 | -152.0429 | 0.22663 | 8.4002 | 0.002039 | 0.003572 | 0.003371 | 6.9e-05 | 0.000711 | 75.3988 | 1.3203 | 2.5474 | 100.29 | 20.87 | 32.124 | 59.9719 | 70.5041 | 58.9293 | 0.007912 | 1.0631 | 0.025367 | 23.8604 | 0.13234 | 0.13482 | -8.1e-05 | -0.006179 | -0.006179 | -0.88801
2019-03-29 11:24 | 1553873014.8413 | -705 | 0.20827 | 0.000515 | 0 | 91.522 | 90.2881 | 129.3084 | 20.8792 | 0.001655 | 0.001533 | -152.0356 | 0.22625 | 8.4018 | 0.002023 | 0.002771 | 0.002587 | 6.9e-05 | 0.000712 | 75.4341 | 1.3281 | 2.5476 | 100.29 | 20.87 | 32.124 | 61.2764 | 70.1296 | 58.3368 | 0.007912 | 1.065 | 0.02542 | 23.8677 | 0.12141 | 0.1346 | -7.8e-05 | -0.00589 | -0.00589 | -0.93008
2019-03-29 11:24 | 1553873015.8414 | -704 | 0.20829 | 0.000512 | 0 | 91.5555 | 90.3045 | 129.4905 | 20.8791 | 0.001701 | 0.0015 | -152.041 | 0.22662 | 8.4017 | 0.002089 | 0.003472 | 0.003209 | 6.8e-05 | 0.000713 | 75.4263 | 1.3539 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3841 | 71.709 | 57.5421 | 0.007912 | 1.0676 | 0.025468 | 23.8558 | 0.12306 | 0.13481 | -7.4e-05 | -0.005685 | -0.005685 | -0.96509
2019-03-29 11:24 | 1553873016.8407 | -703 | 0.20828 | 0.000511 | 0 | 91.5959 | 90.2499 | 129.3021 | 20.8802 | 0.001688 | 0.001548 | -152.0476 | 0.22739 | 8.4013 | 0.002078 | 0.00334 | 0.00311 | 6.6e-05 | 0.000714 | 75.4196 | 1.3053 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.0664 | 71.4504 | 58.6875 | 0.007912 | 1.0639 | 0.025406 | 23.8797 | 0.13242 | 0.13528 | -7.3e-05 | -0.005528 | -0.005528 | -0.9959
2019-03-29 11:24 | 1553873017.8405 | -702 | 0.20825 | 0.000512 | 0 | 91.604 | 90.2949 | 129.1924 | 20.8793 | 0.001549 | 0.001519 | -152.0453 | 0.22751 | 8.4005 | 0.002154 | 0.003179 | 0.003059 | 6.4e-05 | 0.000719 | 75.4153 | 1.3364 | 2.5474 | 100.29 | 20.87 | 32.124 | 58.7326 | 73.2854 | 58.0074 | 0.007912 | 1.0661 | 0.025421 | 23.8451 | 0.12874 | 0.13535 | -7.5e-05 | -0.005677 | -0.005677 | -0.97042
2019-03-29 11:24 | 1553873018.8405 | -701 | 0.20825 | 0.000515 | 0 | 91.639 | 90.3686 | 129.1777 | 20.8801 | 0.001641 | 0.001504 | -152.0469 | 0.22704 | 8.4024 | 0.002104 | 0.003175 | 0.002905 | 6.8e-05 | 0.000723 | 75.3949 | 1.3554 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.9419 | 72.0654 | 57.634 | 0.007912 | 1.0673 | 0.025434 | 23.8303 | 0.12486 | 0.13507 | -7.8e-05 | -0.005909 | -0.005909 | -0.93032
2019-03-29 11:24 | 1553873019.8415 | -700 | 0.20825 | 0.000514 | 0 | 91.6734 | 90.2621 | 129.1941 | 20.8816 | 0.001614 | 0.001515 | -152.0277 | 0.22819 | 8.403 | 0.002138 | 0.003296 | 0.00314 | 6.8e-05 | 0.000725 | 75.3757 | 1.3162 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.2921 | 72.898 | 57.9048 | 0.007912 | 1.0664 | 0.025425 | 23.8416 | 0.13877 | 0.13575 | -7.6e-05 | -0.005811 | -0.005811 | -0.95079
2019-03-29 11:24 | 1553873020.8405 | -699 | 0.20827 | 0.000512 | 0 | 91.6812 | 90.5376 | 129.43 | 20.8817 | 0.001719 | 0.001556 | -152.0297 | 0.22738 | 8.4014 | 0.002114 | 0.00292 | 0.002707 | 6.7e-05 | 0.000721 | 75.3949 | 1.326 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.821 | 72.3142 | 58.8922 | 0.007912 | 1.0633 | 0.025411 | 23.8989 | 0.11228 | 0.13527 | -7.5e-05 | -0.005658 | -0.005658 | -0.97301
2019-03-29 11:24 | 1553873021.8416 | -698 | 0.20824 | 0.000514 | 0 | 91.7027 | 90.4417 | 128.8427 | 20.881 | 0.001591 | 0.00154 | -152.0097 | 0.22706 | 8.4015 | 0.002139 | 0.003004 | 0.002855 | 7e-05 | 0.000718 | 75.4348 | 1.3148 | 2.5473 | 100.29 | 20.87 | 32.124 | 59.7269 | 72.9263 | 58.5052 | 0.007912 | 1.0645 | 0.025368 | 23.8307 | 0.12386 | 0.13508 | -7.7e-05 | -0.005826 | -0.005826 | -0.94363
2019-03-29 11:24 | 1553873022.8415 | -697 | 0.20825 | 0.000509 | 0 | 91.5983 | 90.2444 | 128.9305 | 20.8814 | 0.001642 | 0.001541 | -152.0181 | 0.22716 | 8.3999 | 0.002155 | 0.003079 | 0.00293 | 6.8e-05 | 0.000719 | 75.4272 | 1.3464 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.9649 | 73.3094 | 58.5408 | 0.007912 | 1.0644 | 0.025375 | 23.8401 | 0.13319 | 0.13514 | -7e-05 | -0.005294 | -0.005294 | -1.039
2019-03-29 11:24 | 1553873023.8415 | -696 | 0.20825 | 0.000512 | 0 | 91.6971 | 90.5567 | 128.9537 | 20.882 | 0.001646 | 0.001584 | -152.023 | 0.22744 | 8.402 | 0.002056 | 0.003864 | 0.003639 | 6.6e-05 | 0.000726 | 75.4411 | 1.3509 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.0609 | 70.922 | 59.5769 | 0.007912 | 1.0611 | 0.025338 | 23.8795 | 0.11194 | 0.1353 | -7.4e-05 | -0.005653 | -0.005653 | -0.9741
2019-03-29 11:24 | 1553873024.84 | -695 | 0.20826 | 0.000512 | -1e-06 | 91.7237 | 90.3908 | 129.0685 | 20.882 | 0.001657 | 0.001496 | -152.0207 | 0.22633 | 8.4015 | 0.002119 | 0.003666 | 0.003536 | 6.9e-05 | 0.000725 | 75.3886 | 1.308 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.3196 | 72.4392 | 57.4479 | 0.007912 | 1.0679 | 0.02543 | 23.8135 | 0.13093 | 0.13464 | -7.3e-05 | -0.005509 | -0.005509 | -0.99477
2019-03-29 11:24 | 1553873025.8396 | -694 | 0.20828 | 0.000514 | 0 | 91.8321 | 90.5301 | 128.4702 | 20.8825 | 0.001662 | 0.001555 | -152.013 | 0.22742 | 8.4013 | 0.002092 | 0.003557 | 0.003339 | 6.7e-05 | 0.00072 | 75.3937 | 1.2911 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.4506 | 71.7864 | 58.8626 | 0.007912 | 1.0633 | 0.025317 | 23.8091 | 0.12773 | 0.13529 | -7.7e-05 | -0.005784 | -0.005784 | -0.95199
2019-03-29 11:24 | 1553873026.8406 | -693 | 0.20824 | 0.000512 | 0 | 91.7689 | 90.358 | 128.6108 | 20.8826 | 0.001623 | 0.001565 | -152.0051 | 0.22688 | 8.403 | 0.002143 | 0.003496 | 0.00332 | 6.7e-05 | 0.000726 | 75.3575 | 1.3319 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.5174 | 73.0105 | 59.1161 | 0.007912 | 1.0625 | 0.025322 | 23.8312 | 0.13858 | 0.13497 | -7.4e-05 | -0.005581 | -0.005581 | -0.98437
2019-03-29 11:24 | 1553873027.8407 | -692 | 0.20823 | 0.000514 | 0 | 91.8264 | 90.5045 | 128.7699 | 20.8828 | 0.001668 | 0.001544 | -152.0024 | 0.22681 | 8.4034 | 0.002073 | 0.003719 | 0.003498 | 6.6e-05 | 0.000724 | 75.3514 | 1.3092 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.6055 | 71.3455 | 58.6116 | 0.007912 | 1.0642 | 0.025356 | 23.8278 | 0.1297 | 0.13493 | -7.6e-05 | -0.005776 | -0.005776 | -0.95076
2019-03-29 11:24 | 1553873028.8411 | -691 | 0.20824 | 0.000516 | 0 | 91.8348 | 90.4785 | 128.8765 | 20.883 | 0.001712 | 0.001566 | -151.993 | 0.22765 | 8.4033 | 0.002147 | 0.003789 | 0.003569 | 6.6e-05 | 0.000726 | 75.3669 | 1.3561 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.6551 | 73.1168 | 59.129 | 0.007912 | 1.0625 | 0.025347 | 23.8563 | 0.13308 | 0.13543 | -7.9e-05 | -0.005988 | -0.005988 | -0.92054
2019-03-29 11:24 | 1553873029.8398 | -690 | 0.20824 | 0.000515 | 0 | 91.8174 | 90.5167 | 128.8272 | 20.883 | 0.001672 | 0.001504 | -152.0058 | 0.22773 | 8.4039 | 0.002179 | 0.003548 | 0.003393 | 6.8e-05 | 0.000728 | 75.3647 | 1.3368 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.7004 | 73.8909 | 57.6431 | 0.007912 | 1.0673 | 0.025399 | 23.7983 | 0.12761 | 0.13548 | -7.7e-05 | -0.005825 | -0.005825 | -0.94655
2019-03-29 11:24 | 1553873030.8412 | -689 | 0.20826 | 0.000516 | 0 | 91.8701 | 90.5703 | 128.634 | 20.884 | 0.001669 | 0.001572 | -152.0119 | 0.22752 | 8.4027 | 0.00216 | 0.003719 | 0.003525 | 6.6e-05 | 0.000732 | 75.4318 | 1.3336 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.617 | 73.4327 | 59.2833 | 0.007912 | 1.062 | 0.025317 | 23.8393 | 0.12745 | 0.13535 | -7.9e-05 | -0.005942 | -0.005942 | -0.92708
2019-03-29 11:24 | 1553873031.8399 | -688 | 0.20823 | 0.000511 | 0 | 91.9733 | 90.6775 | 128.9075 | 20.8831 | 0.00169 | 0.001538 | -152.0192 | 0.22746 | 8.4035 | 0.002101 | 0.00393 | 0.003689 | 6.2e-05 | 0.000731 | 75.3855 | 1.3121 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.1263 | 72.0107 | 58.4679 | 0.007912 | 1.0646 | 0.025375 | 23.8354 | 0.12692 | 0.13531 | -7.3e-05 | -0.005501 | -0.005501 | -1.0012
2019-03-29 11:24 | 1553873032.8398 | -687 | 0.20824 | 0.000515 | 0 | 92.0482 | 90.7908 | 128.9603 | 20.8849 | 0.001682 | 0.001566 | -152.0084 | 0.22787 | 8.4026 | 0.002155 | 0.003763 | 0.003623 | 6.4e-05 | 0.000727 | 75.394 | 1.3579 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.9397 | 73.3051 | 59.1372 | 0.007912 | 1.0625 | 0.025355 | 23.8643 | 0.12302 | 0.13556 | -7.8e-05 | -0.005903 | -0.005903 | -0.93469
2019-03-29 11:24 | 1553873033.8409 | -686 | 0.20826 | 0.000513 | 0 | 92.1359 | 90.8795 | 128.6929 | 20.8842 | 0.001669 | 0.001476 | -152.0167 | 0.22802 | 8.4014 | 0.002155 | 0.004036 | 0.003886 | 7e-05 | 0.000732 | 75.3694 | 1.3697 | 2.5476 | 100.29 | 20.87 | 32.124 | 61.6093 | 73.3027 | 56.9585 | 0.007912 | 1.0695 | 0.025412 | 23.7613 | 0.12281 | 0.13565 | -7.6e-05 | -0.005757 | -0.005757 | -0.95904
2019-03-29 11:24 | 1553873034.8409 | -685 | 0.20825 | 0.000512 | 0 | 92.343 | 91.2273 | 128.3979 | 20.8855 | 0.001698 | 0.001529 | -152.0102 | 0.22775 | 8.4026 | 0.002165 | 0.003849 | 0.003636 | 6.5e-05 | 0.000731 | 75.3261 | 1.3877 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3301 | 73.5449 | 58.238 | 0.007912 | 1.0654 | 0.025334 | 23.78 | 0.10873 | 0.13549 | -7.5e-05 | -0.005659 | -0.005659 | -0.97443
2019-03-29 11:24 | 1553873035.841 | -684 | 0.20826 | 0.000515 | 0 | 92.3873 | 91.1799 | 128.7521 | 20.885 | 0.001634 | 0.001521 | -152.0237 | 0.22695 | 8.4018 | 0.002188 | 0.003638 | 0.003473 | 6.7e-05 | 0.000733 | 75.4045 | 1.3481 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.7802 | 74.1064 | 58.0583 | 0.007912 | 1.0659 | 0.025376 | 23.8063 | 0.11767 | 0.13501 | -7.8e-05 | -0.005932 | -0.005932 | -0.92639
2019-03-29 11:24 | 1553873036.8409 | -683 | 0.20826 | 0.000514 | 0 | 92.5355 | 91.3225 | 128.7462 | 20.8852 | 0.001642 | 0.001526 | -152.0251 | 0.22707 | 8.4051 | 0.002231 | 0.003377 | 0.003229 | 6.9e-05 | 0.000732 | 75.3679 | 1.323 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.9608 | 75.1238 | 58.175 | 0.007912 | 1.0656 | 0.025371 | 23.8099 | 0.11802 | 0.13509 | -7.7e-05 | -0.005847 | -0.005847 | -0.94024
2019-03-29 11:24 | 1553873037.8411 | -682 | 0.20825 | 0.000518 | 0 | 92.6334 | 91.4154 | 128.7002 | 20.8857 | 0.001649 | 0.001646 | -152.0282 | 0.22744 | 8.4018 | 0.002168 | 0.003372 | 0.003201 | 6.5e-05 | 0.000737 | 75.3498 | 1.3665 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.131 | 73.6091 | 61.0777 | 0.007912 | 1.0563 | 0.025256 | 23.9097 | 0.11839 | 0.13531 | -8.2e-05 | -0.006178 | -0.006178 | -0.89142
2019-03-29 11:24 | 1553873038.8415 | -681 | 0.20824 | 0.000517 | 0 | 92.7244 | 91.5597 | 128.6523 | 20.8864 | 0.001642 | 0.001528 | -152.0164 | 0.22739 | 8.4032 | 0.002101 | 0.003287 | 0.003166 | 6.7e-05 | 0.000736 | 75.394 | 1.3495 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.9589 | 72.0177 | 58.2153 | 0.007912 | 1.0654 | 0.02536 | 23.8027 | 0.11306 | 0.13528 | -8e-05 | -0.006067 | -0.006067 | -0.90744
2019-03-29 11:24 | 1553873039.8411 | -680 | 0.20826 | 0.000516 | 0 | 92.7738 | 91.6041 | 128.7167 | 20.8865 | 0.001737 | 0.001592 | -152.0122 | 0.22764 | 8.4024 | 0.002107 | 0.003375 | 0.003185 | 6.6e-05 | 0.000743 | 75.4104 | 1.3313 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.2513 | 72.163 | 59.7645 | 0.007912 | 1.0605 | 0.025307 | 23.8642 | 0.11349 | 0.13542 | -8e-05 | -0.006043 | -0.006043 | -0.91207
2019-03-29 11:24 | 1553873040.8392 | -679 | 0.20827 | 0.000515 | 0 | 92.8261 | 91.7444 | 128.5918 | 20.8858 | 0.001661 | 0.001592 | -152.0116 | 0.22766 | 8.4017 | 0.002075 | 0.003279 | 0.00303 | 6.8e-05 | 0.000743 | 75.3717 | 1.3252 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.4343 | 71.395 | 59.7684 | 0.007912 | 1.0605 | 0.025295 | 23.8528 | 0.10485 | 0.13543 | -7.8e-05 | -0.005887 | -0.005887 | -0.93631
2019-03-29 11:24 | 1553873041.8394 | -678 | 0.20825 | 0.000517 | 0 | 92.898 | 91.6778 | 128.6077 | 20.8884 | 0.001652 | 0.001594 | -152.0003 | 0.22722 | 8.4021 | 0.002132 | 0.002492 | 0.0023 | 6.7e-05 | 0.000743 | 75.3981 | 1.3361 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.2203 | 72.748 | 59.8062 | 0.007912 | 1.0603 | 0.025295 | 23.8556 | 0.11826 | 0.13517 | -8.1e-05 | -0.006176 | -0.006176 | -0.89079
2019-03-29 11:24 | 1553873042.8403 | -677 | 0.20822 | 0.000515 | 0 | 93.0162 | 91.8867 | 128.159 | 20.8873 | 0.00167 | 0.001574 | -152.0092 | 0.22791 | 8.4003 | 0.0021 | 0.003296 | 0.003093 | 6.7e-05 | 0.000741 | 75.3753 | 1.3017 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.642 | 71.9946 | 59.3375 | 0.007912 | 1.0618 | 0.025269 | 23.7972 | 0.10928 | 0.13559 | -7.8e-05 | -0.0059 | -0.0059 | -0.93531
2019-03-29 11:24 | 1553873043.8403 | -676 | 0.20822 | 0.000513 | 0 | 93.0584 | 91.944 | 128.0207 | 20.8883 | 0.001618 | 0.001528 | -151.9872 | 0.228 | 8.4007 | 0.002144 | 0.003406 | 0.003208 | 6.8e-05 | 0.000744 | 75.3735 | 1.3371 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.3892 | 73.0347 | 58.2219 | 0.007912 | 1.0654 | 0.025297 | 23.7444 | 0.10776 | 0.13564 | -7.6e-05 | -0.005768 | -0.005768 | -0.95713
2019-03-29 11:24 | 1553873044.8393 | -675 | 0.20824 | 0.000514 | 0 | 93.2081 | 91.9164 | 128.4582 | 20.8889 | 0.001666 | 0.001508 | -151.9872 | 0.22719 | 8.402 | 0.002117 | 0.003545 | 0.003339 | 6.9e-05 | 0.000745 | 75.2768 | 1.3141 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.5397 | 72.3897 | 57.7326 | 0.007912 | 1.067 | 0.025359 | 23.7674 | 0.12483 | 0.13515 | -7.7e-05 | -0.005848 | -0.005848 | -0.94063
2019-03-29 11:24 | 1553873045.8407 | -674 | 0.20824 | 0.000512 | 0 | 93.2162 | 91.9614 | 128.5852 | 20.8884 | 0.001669 | 0.001563 | -151.9836 | 0.22704 | 8.4027 | 0.002104 | 0.003006 | 0.002831 | 7.1e-05 | 0.000745 | 75.343 | 1.3167 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.6307 | 72.0929 | 59.06 | 0.007912 | 1.0627 | 0.025321 | 23.8268 | 0.12122 | 0.13507 | -7.5e-05 | -0.005662 | -0.005662 | -0.97097
2019-03-29 11:24 | 1553873046.8394 | -673 | 0.20827 | 0.000517 | 0 | 93.3895 | 92.0887 | 128.3456 | 20.8887 | 0.001665 | 0.00156 | -151.9949 | 0.22747 | 8.4012 | 0.002139 | 0.003417 | 0.003156 | 6.7e-05 | 0.000743 | 75.3529 | 1.338 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.5192 | 72.9375 | 58.9952 | 0.007912 | 1.0629 | 0.0253 | 23.8023 | 0.12546 | 0.13532 | -8e-05 | -0.00606 | -0.00606 | -0.90888
2019-03-29 11:24 | 1553873047.8404 | -672 | 0.20826 | 0.000515 | 0 | 93.5104 | 92.6208 | 128.2721 | 20.8887 | 0.00169 | 0.001542 | -152.0056 | 0.22644 | 8.4002 | 0.002059 | 0.00336 | 0.00328 | 6.8e-05 | 0.000745 | 75.3864 | 1.3679 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.1411 | 71.0057 | 58.5595 | 0.007912 | 1.0643 | 0.025309 | 23.7798 | 0.085503 | 0.13471 | -7.7e-05 | -0.005845 | -0.005845 | -0.93808
2019-03-29 11:24 | 1553873048.8405 | -671 | 0.20829 | 0.000517 | 0 | 93.5412 | 92.514 | 128.3019 | 20.8889 | 0.001703 | 0.001554 | -152.027 | 0.22758 | 8.4013 | 0.00216 | 0.003338 | 0.003154 | 7e-05 | 0.000741 | 75.3963 | 1.3496 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.4326 | 73.442 | 58.8493 | 0.007912 | 1.0634 | 0.025301 | 23.793 | 0.098774 | 0.13539 | -8.1e-05 | -0.006105 | -0.006105 | -0.90254
2019-03-29 11:24 | 1553873049.8395 | -670 | 0.20825 | 0.000514 | 0 | 93.7538 | 92.8126 | 128.2348 | 20.8886 | 0.001639 | 0.001551 | -152.0392 | 0.22807 | 8.4014 | 0.002114 | 0.003914 | 0.003693 | 6.4e-05 | 0.000741 | 75.3769 | 1.3334 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.9126 | 72.3286 | 58.7723 | 0.007912 | 1.0636 | 0.025298 | 23.784 | 0.090247 | 0.13568 | -7.6e-05 | -0.005784 | -0.005784 | -0.95482
2019-03-29 11:24 | 1553873050.8395 | -669 | 0.20827 | 0.000511 | 0 | 93.9034 | 92.7699 | 128.3174 | 20.8893 | 0.001665 | 0.001567 | -152.0292 | 0.22776 | 8.4004 | 0.002203 | 0.003381 | 0.003236 | 6.6e-05 | 0.000737 | 75.3486 | 1.3812 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.5213 | 74.4746 | 59.168 | 0.007912 | 1.0624 | 0.025291 | 23.8058 | 0.10863 | 0.13549 | -7.2e-05 | -0.00544 | -0.00544 | -1.0137
2019-03-29 11:24 | 1553873051.8406 | -668 | 0.2083 | 0.000515 | 0 | 93.9447 | 93.0006 | 127.9871 | 20.8893 | 0.00165 | 0.001614 | -152.0313 | 0.22785 | 8.4028 | 0.002196 | 0.003386 | 0.003248 | 6.5e-05 | 0.000738 | 75.3204 | 1.3646 | 2.5471 | 100.29 | 20.87 | 32.124 | 61.1657 | 74.2967 | 60.2957 | 0.007912 | 1.0588 | 0.025215 | 23.8155 | 0.090343 | 0.13555 | -7.9e-05 | -0.005937 | -0.005937 | -0.92916
2019-03-29 11:24 | 1553873052.8394 | -667 | 0.20827 | 0.000513 | -1e-06 | 93.9893 | 93.161 | 128.2665 | 20.8907 | 0.001639 | 0.001551 | -152.0376 | 0.22824 | 8.4033 | 0.002177 | 0.003746 | 0.0036 | 6.7e-05 | 0.000738 | 75.3697 | 1.354 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.8944 | 73.8476 | 58.7782 | 0.007912 | 1.0636 | 0.0253 | 23.7872 | 0.079175 | 0.13578 | -7.5e-05 | -0.005606 | -0.005606 | -0.98583
2019-03-29 11:24 | 1553873053.8407 | -666 | 0.20826 | 0.000518 | 0 | 94.1438 | 93.1999 | 128.4573 | 20.8907 | 0.001679 | 0.001509 | -152.0378 | 0.22752 | 8.4042 | 0.002094 | 0.00353 | 0.003328 | 6.7e-05 | 0.000743 | 75.4347 | 1.3624 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.8697 | 71.8419 | 57.7626 | 0.007912 | 1.0669 | 0.025358 | 23.7684 | 0.090127 | 0.13535 | -8.1e-05 | -0.006134 | -0.006134 | -0.89802
2019-03-29 11:24 | 1553873054.8396 | -665 | 0.20825 | 0.000516 | 0 | 94.2839 | 93.2255 | 128.338 | 20.8903 | 0.001738 | 0.00153 | -152.0432 | 0.22702 | 8.4009 | 0.002158 | 0.003185 | 0.003 | 6.7e-05 | 0.000746 | 75.3702 | 1.3388 | 2.5472 | 100.29 | 20.87 | 32.124 | 63.2887 | 73.3951 | 58.2741 | 0.007912 | 1.0652 | 0.025327 | 23.7757 | 0.10097 | 0.13506 | -8e-05 | -0.006034 | -0.006034 | -0.91102
2019-03-29 11:24 | 1553873055.8396 | -664 | 0.20827 | 0.000515 | 0 | 94.4524 | 93.3941 | 127.8738 | 20.8916 | 0.001669 | 0.001616 | -152.0312 | 0.22761 | 8.4026 | 0.002209 | 0.003976 | 0.003793 | 6.7e-05 | 0.000745 | 75.3647 | 1.3058 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.6298 | 74.6022 | 60.3558 | 0.007912 | 1.0586 | 0.025202 | 23.8071 | 0.10078 | 0.1354 | -7.8e-05 | -0.005877 | -0.005877 | -0.93779
2019-03-29 11:24 | 1553873056.839 | -663 | 0.20827 | 0.000515 | 0 | 94.512 | 93.6406 | 127.9167 | 20.8913 | 0.001655 | 0.001597 | -152.0393 | 0.22774 | 8.4039 | 0.002144 | 0.003548 | 0.003383 | 6.7e-05 | 0.000745 | 75.3719 | 1.3154 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.2778 | 73.0429 | 59.8894 | 0.007912 | 1.0601 | 0.025224 | 23.7944 | 0.082853 | 0.13548 | -7.7e-05 | -0.005788 | -0.005788 | -0.95266
2019-03-29 11:24 | 1553873057.8387 | -662 | 0.20826 | 0.000516 | 0 | 94.6057 | 93.5935 | 128.239 | 20.8921 | 0.001704 | 0.001527 | -152.0166 | 0.22652 | 8.4035 | 0.002105 | 0.003837 | 0.003716 | 6.7e-05 | 0.000746 | 75.3598 | 1.3574 | 2.5472 | 100.29 | 20.87 | 32.124 | 62.4741 | 72.1091 | 58.2093 | 0.007912 | 1.0654 | 0.025319 | 23.7642 | 0.096215 | 0.13476 | -7.9e-05 | -0.00597 | -0.00597 | -0.91867
2019-03-29 11:24 | 1553873058.9083 | -661 | 0.20826 | 0.000513 | 0 | 94.7218 | 93.7651 | 128.5408 | 20.8919 | 0.001715 | 0.001526 | -152.0009 | 0.22696 | 8.4039 | 0.002097 | 0.003373 | 0.003138 | 6.7e-05 | 0.000747 | 75.4309 | 1.3424 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.7435 | 71.9094 | 58.1777 | 0.007912 | 1.0655 | 0.02535 | 23.791 | 0.090798 | 0.13502 | -7.5e-05 | -0.005698 | -0.005698 | -0.96442
2019-03-29 11:24 | 1553873059.8518 | -660 | 0.20823 | 0.000514 | 0 | 94.7601 | 93.7928 | 128.4268 | 20.8923 | 0.001663 | 0.001587 | -151.9808 | 0.22718 | 8.4044 | 0.002119 | 0.00307 | 0.002902 | 6.5e-05 | 0.000747 | 75.3934 | 1.3683 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.4727 | 72.45 | 59.6497 | 0.007912 | 1.0608 | 0.025283 | 23.8332 | 0.091779 | 0.13515 | -7.7e-05 | -0.005795 | -0.005795 | -0.94924
2019-03-29 11:24 | 1553873060.8528 | -659 | 0.20826 | 0.000517 | 0 | 94.911 | 94.046 | 128.446 | 20.8919 | 0.001672 | 0.0016 | -151.994 | 0.22763 | 8.4027 | 0.002157 | 0.003288 | 0.003081 | 6.8e-05 | 0.000748 | 75.3694 | 1.3454 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.6934 | 73.3623 | 59.9566 | 0.007912 | 1.0599 | 0.025273 | 23.846 | 0.081891 | 0.13542 | -8e-05 | -0.006012 | -0.006012 | -0.91681
2019-03-29 11:24 | 1553873061.8529 | -658 | 0.20827 | 0.000512 | 0 | 94.9894 | 94.1443 | 128.7577 | 20.8936 | 0.001675 | 0.001541 | -151.9977 | 0.22798 | 8.4015 | 0.00214 | 0.002629 | 0.002434 | 7e-05 | 0.000757 | 75.3715 | 1.3416 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.7653 | 72.9485 | 58.5534 | 0.007912 | 1.0643 | 0.025357 | 23.8246 | 0.079935 | 0.13562 | -7.4e-05 | -0.005604 | -0.005604 | -0.98491
2019-03-29 11:24 | 1553873062.853 | -657 | 0.2083 | 0.000516 | 0 | 95.0499 | 94.1893 | 128.8298 | 20.8928 | 0.001618 | 0.001566 | -151.9901 | 0.2282 | 8.4011 | 0.002221 | 0.003344 | 0.003101 | 7e-05 | 0.000758 | 75.4035 | 1.3449 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.3891 | 74.8975 | 59.1494 | 0.007912 | 1.0624 | 0.025342 | 23.8527 | 0.081348 | 0.13576 | -7.9e-05 | -0.00597 | -0.00597 | -0.92552
2019-03-29 11:24 | 1553873063.8523 | -656 | 0.20825 | 0.000513 | 0 | 95.1834 | 94.4784 | 128.9112 | 20.8929 | 0.001663 | 0.001496 | -152.0034 | 0.22753 | 8.4057 | 0.002151 | 0.00392 | 0.003776 | 6.6e-05 | 0.000758 | 75.3425 | 1.3361 | 2.5472 | 100.29 | 20.87 | 32.124 | 61.4841 | 73.2302 | 57.4603 | 0.007912 | 1.0679 | 0.025415 | 23.7995 | 0.066493 | 0.13536 | -7.5e-05 | -0.005711 | -0.005711 | -0.96457
2019-03-29 11:24 | 1553873064.852 | -655 | 0.20827 | 0.000515 | 0 | 95.3055 | 94.4593 | 128.4359 | 20.8935 | 0.001697 | 0.001606 | -152.0162 | 0.22757 | 8.4013 | 0.002193 | 0.003535 | 0.003391 | 6.9e-05 | 0.000761 | 75.317 | 1.3619 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.2989 | 74.23 | 60.1152 | 0.007912 | 1.0594 | 0.025266 | 23.8508 | 0.079771 | 0.13538 | -7.8e-05 | -0.005891 | -0.005891 | -0.93533
2019-03-29 11:24 | 1553873065.852 | -654 | 0.20826 | 0.000519 | 0 | 95.3432 | 94.5668 | 128.9084 | 20.8938 | 0.001682 | 0.001651 | -152.0171 | 0.22718 | 8.4014 | 0.002136 | 0.003483 | 0.00331 | 6.8e-05 | 0.000765 | 75.3575 | 1.3809 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.9351 | 72.8661 | 61.1983 | 0.007912 | 1.0559 | 0.025272 | 23.9334 | 0.073136 | 0.13515 | -8.3e-05 | -0.006263 | -0.006263 | -0.87823
2019-03-29 11:24 | 1553873066.8521 | -653 | 0.20825 | 0.000517 | 0 | 95.4534 | 94.6716 | 129.1769 | 20.8941 | 0.001653 | 0.001615 | -152.0072 | 0.22833 | 8.4045 | 0.002149 | 0.003786 | 0.003608 | 7.1e-05 | 0.000766 | 75.4076 | 1.3619 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.2328 | 73.1612 | 60.3363 | 0.007912 | 1.0587 | 0.025331 | 23.9274 | 0.07356 | 0.13583 | -8.1e-05 | -0.006149 | -0.006149 | -0.89901
2019-03-29 11:24 | 1553873067.8521 | -652 | 0.20824 | 0.000512 | 0 | 95.4668 | 94.7231 | 129.3344 | 20.8941 | 0.00166 | 0.001578 | -152.0087 | 0.2281 | 8.4032 | 0.002209 | 0.003023 | 0.002872 | 6.8e-05 | 0.000769 | 75.4695 | 1.3642 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.4166 | 74.6181 | 59.4354 | 0.007912 | 1.0615 | 0.02538 | 23.9096 | 0.069954 | 0.1357 | -7.4e-05 | -0.005613 | -0.005613 | -0.9839
2019-03-29 11:24 | 1553873068.8532 | -651 | 0.2082 | 0.000513 | 0 | 95.5527 | 94.9999 | 129.0309 | 20.8948 | 0.001675 | 0.001583 | -152.0241 | 0.22757 | 8.4016 | 0.002161 | 0.003896 | 0.003736 | 6.7e-05 | 0.000768 | 75.4216 | 1.3311 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.7658 | 73.4672 | 59.5503 | 0.007912 | 1.0612 | 0.025346 | 23.8857 | 0.051901 | 0.13538 | -7.5e-05 | -0.005697 | -0.005697 | -0.96718
2019-03-29 11:24 | 1553873069.8525 | -650 | 0.20824 | 0.000514 | 0 | 95.5755 | 94.7098 | 129.1138 | 20.8946 | 0.001741 | 0.001573 | -152.0153 | 0.22782 | 8.4023 | 0.002184 | 0.003435 | 0.003289 | 6.5e-05 | 0.00077 | 75.4252 | 1.3236 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.3677 | 74.0098 | 59.3045 | 0.007912 | 1.0619 | 0.025364 | 23.8845 | 0.081386 | 0.13553 | -7.6e-05 | -0.005796 | -0.005796 | -0.95166
2019-03-29 11:24 | 1553873070.8525 | -649 | 0.20821 | 0.000512 | 0 | 95.6258 | 95.0361 | 129.1744 | 20.8948 | 0.001624 | 0.001513 | -152.0107 | 0.22745 | 8.4032 | 0.00211 | 0.00396 | 0.00383 | 6.9e-05 | 0.000769 | 75.4022 | 1.3229 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.5512 | 72.2405 | 57.8706 | 0.007912 | 1.0665 | 0.025425 | 23.8385 | 0.05533 | 0.13531 | -7.5e-05 | -0.005698 | -0.005698 | -0.96655
2019-03-29 11:24 | 1553873071.8513 | -648 | 0.20823 | 0.000513 | 0 | 95.6836 | 94.8381 | 129.0802 | 20.8947 | 0.001731 | 0.001582 | -152.0132 | 0.22707 | 8.4047 | 0.002183 | 0.003488 | 0.003336 | 7e-05 | 0.000771 | 75.3672 | 1.3348 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.1148 | 73.9919 | 59.5394 | 0.007912 | 1.0612 | 0.025352 | 23.8898 | 0.079382 | 0.13508 | -7.5e-05 | -0.00569 | -0.00569 | -0.96616
2019-03-29 11:25 | 1553873072.8522 | -647 | 0.20824 | 0.000515 | 0 | 95.7485 | 95.1774 | 128.8826 | 20.8961 | 0.00162 | 0.001598 | -152.0025 | 0.22775 | 8.4043 | 0.002172 | 0.003185 | 0.002937 | 6.8e-05 | 0.000765 | 75.3782 | 1.3087 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.457 | 73.7243 | 59.9146 | 0.007912 | 1.06 | 0.025318 | 23.885 | 0.053514 | 0.13549 | -7.7e-05 | -0.005801 | -0.005801 | -0.95066
2019-03-29 11:25 | 1553873073.8523 | -646 | 0.20823 | 0.000511 | 0 | 95.8988 | 95.104 | 129.2314 | 20.8959 | 0.001753 | 0.001579 | -152.0068 | 0.22794 | 8.4032 | 0.002175 | 0.003374 | 0.003135 | 7.2e-05 | 0.000761 | 75.3992 | 1.35 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.658 | 73.809 | 59.4701 | 0.007912 | 1.0614 | 0.025369 | 23.9013 | 0.074439 | 0.1356 | -7.2e-05 | -0.005439 | -0.005439 | -1.0148
2019-03-29 11:25 | 1553873074.8527 | -645 | 0.20823 | 0.000512 | 0 | 96.0032 | 95.2354 | 129.6802 | 20.8962 | 0.001695 | 0.001546 | -152.0043 | 0.22774 | 8.4023 | 0.002234 | 0.003689 | 0.003531 | 6.8e-05 | 0.000762 | 75.4383 | 1.3778 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.2462 | 75.2155 | 58.6749 | 0.007912 | 1.064 | 0.025443 | 23.9142 | 0.071828 | 0.13548 | -7.4e-05 | -0.005599 | -0.005599 | -0.98487
2019-03-29 11:25 | 1553873075.8519 | -644 | 0.20825 | 0.00051 | 0 | 96.1095 | 95.2938 | 129.3967 | 20.8975 | 0.001673 | 0.001583 | -152.0039 | 0.22746 | 8.4032 | 0.002284 | 0.003201 | 0.003068 | 6.8e-05 | 0.000761 | 75.4061 | 1.3684 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.7388 | 76.4086 | 59.5505 | 0.007912 | 1.0612 | 0.025382 | 23.9195 | 0.076243 | 0.13532 | -7.1e-05 | -0.005374 | -0.005374 | -1.0248
2019-03-29 11:25 | 1553873076.8524 | -643 | 0.20823 | 0.000513 | 0 | 96.2906 | 95.8323 | 128.9988 | 20.8966 | 0.001695 | 0.001541 | -151.9975 | 0.22864 | 8.4033 | 0.002195 | 0.003394 | 0.003211 | 6.8e-05 | 0.000762 | 75.4237 | 1.3446 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2591 | 74.2698 | 58.5564 | 0.007912 | 1.0643 | 0.025381 | 23.847 | 0.042673 | 0.13602 | -7.6e-05 | -0.005751 | -0.005751 | -0.96269
2019-03-29 11:25 | 1553873077.8513 | -642 | 0.20824 | 0.000511 | -1e-06 | 96.3256 | 95.748 | 128.665 | 20.897 | 0.001682 | 0.001548 | -152.0048 | 0.228 | 8.4031 | 0.002309 | 0.003512 | 0.003309 | 7.1e-05 | 0.000767 | 75.4783 | 1.3059 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.9493 | 77.0083 | 58.7186 | 0.007912 | 1.0638 | 0.025342 | 23.8219 | 0.053797 | 0.13564 | -7.1e-05 | -0.005379 | -0.005379 | -1.0263
2019-03-29 11:25 | 1553873078.8514 | -641 | 0.20825 | 0.000515 | -1e-06 | 96.5001 | 95.8726 | 128.5939 | 20.8976 | 0.001666 | 0.001578 | -151.9978 | 0.22773 | 8.4039 | 0.002198 | 0.004424 | 0.00415 | 6.9e-05 | 0.000764 | 75.4009 | 1.3475 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.5614 | 74.3604 | 59.4304 | 0.007912 | 1.0615 | 0.025308 | 23.8409 | 0.058358 | 0.13548 | -7.7e-05 | -0.00576 | -0.00576 | -0.95733
2019-03-29 11:25 | 1553873079.8514 | -640 | 0.20825 | 0.000517 | 0 | 96.689 | 96.033 | 128.5167 | 20.8979 | 0.001699 | 0.001581 | -152.0001 | 0.22757 | 8.4031 | 0.002262 | 0.004001 | 0.003888 | 7e-05 | 0.000757 | 75.4237 | 1.3403 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.354 | 75.8835 | 59.5049 | 0.007912 | 1.0613 | 0.025297 | 23.8364 | 0.060895 | 0.13538 | -7.9e-05 | -0.005983 | -0.005983 | -0.921
2019-03-29 11:25 | 1553873080.8529 | -639 | 0.20826 | 0.000511 | 0 | 96.7677 | 96.0306 | 128.6218 | 20.8984 | 0.00165 | 0.001507 | -152.0146 | 0.22748 | 8.405 | 0.002249 | 0.003867 | 0.003652 | 7e-05 | 0.000758 | 75.2759 | 1.3562 | 2.5476 | 100.29 | 20.87 | 32.124 | 61.1809 | 75.5739 | 57.7286 | 0.007912 | 1.067 | 0.025376 | 23.7824 | 0.068397 | 0.13533 | -7.2e-05 | -0.005468 | -0.005468 | -1.0073
2019-03-29 11:25 | 1553873081.8526 | -638 | 0.20827 | 0.000513 | 0 | 96.8863 | 96.3917 | 128.6493 | 20.8982 | 0.001729 | 0.001581 | -152.0175 | 0.22686 | 8.4001 | 0.002159 | 0.003619 | 0.003402 | 7e-05 | 0.000762 | 75.3028 | 1.3905 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.0799 | 73.4114 | 59.509 | 0.007912 | 1.0613 | 0.02531 | 23.8488 | 0.045778 | 0.13496 | -7.5e-05 | -0.005689 | -0.005689 | -0.96552
2019-03-29 11:25 | 1553873082.8516 | -637 | 0.20828 | 0.00051 | 0 | 96.9183 | 96.334 | 128.7796 | 20.8979 | 0.001614 | 0.001551 | -152.0178 | 0.22749 | 8.4023 | 0.002234 | 0.003093 | 0.002973 | 6.9e-05 | 0.000765 | 75.3539 | 1.3827 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.3094 | 75.2195 | 58.786 | 0.007912 | 1.0636 | 0.025351 | 23.835 | 0.054088 | 0.13533 | -7.1e-05 | -0.005385 | -0.005385 | -1.0229
2019-03-29 11:25 | 1553873083.8519 | -636 | 0.20828 | 0.000513 | 0 | 97.0347 | 96.4159 | 128.2005 | 20.8983 | 0.001726 | 0.001553 | -151.9988 | 0.22738 | 8.4007 | 0.002168 | 0.003554 | 0.003408 | 6.9e-05 | 0.000766 | 75.294 | 1.4049 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.0066 | 73.629 | 58.8452 | 0.007912 | 1.0634 | 0.025291 | 23.7834 | 0.057226 | 0.13527 | -7.5e-05 | -0.005676 | -0.005676 | -0.97001
2019-03-29 11:25 | 1553873084.8526 | -635 | 0.20832 | 0.000516 | 0 | 97.1278 | 96.5514 | 128.047 | 20.8981 | 0.001701 | 0.001591 | -151.9937 | 0.2282 | 8.4018 | 0.002244 | 0.003501 | 0.0032 | 6.9e-05 | 0.00076 | 75.3102 | 1.3791 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3932 | 75.4521 | 59.7445 | 0.007912 | 1.0605 | 0.025242 | 23.8014 | 0.053242 | 0.13575 | -7.9e-05 | -0.005927 | -0.005927 | -0.93226
2019-03-29 11:25 | 1553873085.853 | -634 | 0.20834 | 0.000511 | 0 | 97.1661 | 96.5913 | 127.9261 | 20.8992 | 0.001667 | 0.001563 | -152 | 0.22783 | 8.4018 | 0.00226 | 0.003115 | 0.002898 | 6.8e-05 | 0.000759 | 75.4017 | 1.3354 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.5894 | 75.8521 | 59.0831 | 0.007912 | 1.0626 | 0.025255 | 23.7665 | 0.053071 | 0.13554 | -7.2e-05 | -0.005442 | -0.005442 | -1.0137
2019-03-29 11:25 | 1553873086.852 | -633 | 0.2083 | 0.000518 | 0 | 97.2688 | 96.736 | 128.4012 | 20.8992 | 0.001672 | 0.001572 | -151.9968 | 0.22843 | 8.4007 | 0.002277 | 0.003834 | 0.003686 | 6.9e-05 | 0.000751 | 75.3947 | 1.3553 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.7123 | 76.2613 | 59.3068 | 0.007912 | 1.0619 | 0.025294 | 23.8186 | 0.049127 | 0.13589 | -8.1e-05 | -0.006146 | -0.006146 | -0.89986
2019-03-29 11:25 | 1553873087.8517 | -632 | 0.20825 | 0.000513 | 0 | 97.308 | 96.6534 | 128.3059 | 20.8992 | 0.00171 | 0.001554 | -152.0004 | 0.22836 | 8.4024 | 0.002207 | 0.003856 | 0.003654 | 7e-05 | 0.000746 | 75.3473 | 1.3619 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.6319 | 74.57 | 58.8702 | 0.007912 | 1.0633 | 0.025301 | 23.7941 | 0.060373 | 0.13585 | -7.5e-05 | -0.005699 | -0.005699 | -0.97026
2019-03-29 11:25 | 1553873088.8507 | -631 | 0.20826 | 0.000512 | 0 | 97.3529 | 96.749 | 128.3702 | 20.9004 | 0.001688 | 0.001594 | -151.9989 | 0.2277 | 8.4027 | 0.002173 | 0.003783 | 0.003591 | 6.6e-05 | 0.000748 | 75.3627 | 1.3388 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.0908 | 73.758 | 59.8255 | 0.007912 | 1.0603 | 0.025271 | 23.8343 | 0.05566 | 0.13546 | -7.3e-05 | -0.005557 | -0.005557 | -0.99216
2019-03-29 11:25 | 1553873089.852 | -630 | 0.20828 | 0.000511 | 0 | 97.4679 | 97.033 | 128.4 | 20.8992 | 0.001756 | 0.001599 | -152.0067 | 0.22837 | 8.4026 | 0.00221 | 0.004058 | 0.003789 | 6.9e-05 | 0.000743 | 75.4049 | 1.3302 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.7361 | 74.6346 | 59.9567 | 0.007912 | 1.0599 | 0.025269 | 23.8417 | 0.040001 | 0.13586 | -7.2e-05 | -0.005451 | -0.005451 | -1.0143
2019-03-29 11:25 | 1553873090.8508 | -629 | 0.20826 | 0.00051 | 0 | 97.5343 | 97.0364 | 127.8986 | 20.9001 | 0.001729 | 0.001558 | -151.9995 | 0.22857 | 8.402 | 0.002196 | 0.002983 | 0.002835 | 6.9e-05 | 0.000744 | 75.4002 | 1.354 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.0732 | 74.3051 | 58.968 | 0.007912 | 1.063 | 0.025257 | 23.7598 | 0.045783 | 0.13598 | -7.1e-05 | -0.005318 | -0.005318 | -1.0407
2019-03-29 11:25 | 1553873091.8512 | -628 | 0.20824 | 0.00051 | 0 | 97.496 | 96.9999 | 128.0998 | 20.9001 | 0.00169 | 0.001579 | -151.9948 | 0.22779 | 8.4036 | 0.002176 | 0.003698 | 0.003521 | 6.9e-05 | 0.000745 | 75.4182 | 1.2934 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.1358 | 73.8162 | 59.4639 | 0.007912 | 1.0614 | 0.025258 | 23.7962 | 0.045635 | 0.13551 | -7.1e-05 | -0.005358 | -0.005358 | -1.0294
2019-03-29 11:25 | 1553873092.8519 | -627 | 0.20825 | 0.000513 | 0 | 97.6154 | 96.9555 | 128.1432 | 20.9008 | 0.001692 | 0.001562 | -151.9951 | 0.22842 | 8.4011 | 0.002143 | 0.003598 | 0.003449 | 7.4e-05 | 0.000742 | 75.4044 | 1.3151 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.1929 | 73.0243 | 59.0634 | 0.007912 | 1.0627 | 0.025277 | 23.7859 | 0.060678 | 0.13589 | -7.5e-05 | -0.005646 | -0.005646 | -0.97951
2019-03-29 11:25 | 1553873093.8519 | -626 | 0.20827 | 0.000514 | 0 | 97.593 | 97.208 | 128.396 | 20.9013 | 0.00166 | 0.001563 | -152.0018 | 0.22795 | 8.4006 | 0.00227 | 0.003131 | 0.002942 | 6.7e-05 | 0.000736 | 75.4006 | 1.3314 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.4187 | 76.0865 | 59.0701 | 0.007912 | 1.0627 | 0.025302 | 23.8096 | 0.035352 | 0.1356 | -7.6e-05 | -0.005766 | -0.005766 | -0.9572
2019-03-29 11:25 | 1553873094.8523 | -625 | 0.20825 | 0.000513 | 0 | 97.7142 | 97.1053 | 128.4791 | 20.9018 | 0.001705 | 0.001564 | -151.9793 | 0.22742 | 8.4021 | 0.002264 | 0.002434 | 0.002292 | 7e-05 | 0.000738 | 75.469 | 1.3795 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.4953 | 75.9381 | 59.0982 | 0.007912 | 1.0626 | 0.025309 | 23.8183 | 0.055913 | 0.13529 | -7.5e-05 | -0.005671 | -0.005671 | -0.97107
2019-03-29 11:25 | 1553873095.8509 | -624 | 0.20824 | 0.000509 | 0 | 97.8202 | 97.2985 | 128.772 | 20.9018 | 0.001686 | 0.001582 | -151.9825 | 0.2277 | 8.4026 | 0.002297 | 0.004471 | 0.004221 | 7.4e-05 | 0.000736 | 75.4279 | 1.3539 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.0368 | 76.7395 | 59.5352 | 0.007912 | 1.0612 | 0.025321 | 23.8612 | 0.047832 | 0.13546 | -7e-05 | -0.005299 | -0.005299 | -1.0403
2019-03-29 11:25 | 1553873096.8522 | -623 | 0.20827 | 0.000512 | 0 | 97.9477 | 97.4746 | 128.7397 | 20.9017 | 0.001666 | 0.001574 | -151.9889 | 0.22773 | 8.4019 | 0.002162 | 0.00357 | 0.0034 | 7e-05 | 0.000734 | 75.4309 | 1.3209 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.5755 | 73.4802 | 59.3413 | 0.007912 | 1.0618 | 0.025326 | 23.8512 | 0.04331 | 0.13548 | -7.3e-05 | -0.005525 | -0.005525 | -0.99792
2019-03-29 11:25 | 1553873097.852 | -622 | 0.20825 | 0.000513 | 0 | 97.9923 | 97.6091 | 128.7317 | 20.902 | 0.001634 | 0.001614 | -151.9793 | 0.22807 | 8.4025 | 0.002189 | 0.003106 | 0.002931 | 6.9e-05 | 0.000737 | 75.3335 | 1.3324 | 2.5474 | 100.29 | 20.87 | 32.124 | 60.7843 | 74.1455 | 60.3144 | 0.007912 | 1.0587 | 0.025288 | 23.8853 | 0.035051 | 0.13568 | -7.4e-05 | -0.005547 | -0.005547 | -0.99557
2019-03-29 11:25 | 1553873098.8521 | -621 | 0.20826 | 0.000515 | 0 | 98.1617 | 97.8725 | 128.5467 | 20.903 | 0.001675 | 0.001522 | -151.9682 | 0.22809 | 8.4018 | 0.002272 | 0.004133 | 0.003933 | 7e-05 | 0.00074 | 75.308 | 1.328 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.7777 | 76.1264 | 58.0969 | 0.007912 | 1.0658 | 0.025354 | 23.7887 | 0.026391 | 0.13569 | -7.7e-05 | -0.005848 | -0.005848 | -0.94431
2019-03-29 11:25 | 1553873099.8512 | -620 | 0.20827 | 0.000512 | -1e-06 | 98.2578 | 97.8722 | 128.2364 | 20.9025 | 0.001677 | 0.001575 | -151.9726 | 0.22742 | 8.4028 | 0.002212 | 0.003227 | 0.003126 | 6.9e-05 | 0.000747 | 75.4424 | 1.3819 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.8217 | 74.686 | 59.3748 | 0.007912 | 1.0617 | 0.025275 | 23.8057 | 0.035173 | 0.13529 | -7.3e-05 | -0.005449 | -0.005449 | -1.0105
2019-03-29 11:25 | 1553873100.8511 | -619 | 0.20827 | 0.000513 | 0 | 98.2686 | 97.8848 | 128.4265 | 20.9029 | 0.001649 | 0.001612 | -151.9809 | 0.2285 | 8.4029 | 0.002213 | 0.003416 | 0.003165 | 7e-05 | 0.000742 | 75.4365 | 1.3552 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.157 | 74.7193 | 60.2731 | 0.007912 | 1.0589 | 0.025259 | 23.8555 | 0.035 | 0.13594 | -7.5e-05 | -0.005666 | -0.005666 | -0.97647
2019-03-29 11:25 | 1553873101.8512 | -618 | 0.20831 | 0.000513 | 0 | 98.4075 | 98.0292 | 128.1817 | 20.9029 | 0.001684 | 0.001616 | -151.9813 | 0.22792 | 8.4017 | 0.002193 | 0.002951 | 0.002819 | 6.9e-05 | 0.000745 | 75.399 | 1.3422 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.001 | 74.243 | 60.3683 | 0.007912 | 1.0585 | 0.025232 | 23.8362 | 0.034451 | 0.13559 | -7.4e-05 | -0.005589 | -0.005589 | -0.9873
2019-03-29 11:25 | 1553873102.8503 | -617 | 0.20829 | 0.000514 | 0 | 98.4602 | 97.9971 | 128.0249 | 20.9022 | 0.001664 | 0.001593 | -151.9904 | 0.22789 | 8.4037 | 0.002183 | 0.003792 | 0.003585 | 6.9e-05 | 0.000746 | 75.4208 | 1.3755 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.5074 | 73.9984 | 59.8157 | 0.007912 | 1.0603 | 0.025237 | 23.8019 | 0.042166 | 0.13557 | -7.5e-05 | -0.005668 | -0.005668 | -0.9735
2019-03-29 11:25 | 1553873103.8502 | -616 | 0.20824 | 0.000514 | 0 | 98.5515 | 98.3643 | 127.8921 | 20.9038 | 0.001709 | 0.00157 | -152.0003 | 0.22778 | 8.4035 | 0.002166 | 0.003187 | 0.003082 | 6.8e-05 | 0.000742 | 75.4394 | 1.3557 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.6012 | 73.5968 | 59.2471 | 0.007912 | 1.0621 | 0.025246 | 23.7692 | 0.017 | 0.13551 | -7.5e-05 | -0.005682 | -0.005682 | -0.97066
2019-03-29 11:25 | 1553873104.8502 | -615 | 0.20828 | 0.000514 | 0 | 98.6778 | 98.1964 | 127.724 | 20.904 | 0.001679 | 0.001554 | -152.02 | 0.22789 | 8.4036 | 0.002213 | 0.003462 | 0.003229 | 6.9e-05 | 0.000745 | 75.409 | 1.394 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.8802 | 74.7125 | 58.8608 | 0.007912 | 1.0634 | 0.025244 | 23.7398 | 0.043739 | 0.13557 | -7.5e-05 | -0.005656 | -0.005656 | -0.9755
2019-03-29 11:25 | 1553873105.8506 | -614 | 0.20828 | 0.000515 | 0 | 98.685 | 98.1278 | 127.4041 | 20.9036 | 0.001749 | 0.001604 | -152.0171 | 0.22789 | 8.4033 | 0.002183 | 0.003221 | 0.00304 | 7.4e-05 | 0.000745 | 75.4156 | 1.3451 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.5748 | 74.0012 | 60.0812 | 0.007912 | 1.0595 | 0.025166 | 23.7535 | 0.050641 | 0.13557 | -7.7e-05 | -0.005764 | -0.005764 | -0.95732
2019-03-29 11:25 | 1553873106.8514 | -613 | 0.20824 | 0.000513 | -1e-06 | 98.7433 | 98.438 | 127.1074 | 20.9049 | 0.001742 | 0.001628 | -152.0217 | 0.22815 | 8.4036 | 0.002208 | 0.003461 | 0.003276 | 7.3e-05 | 0.000741 | 75.458 | 1.3185 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.4027 | 74.5997 | 60.6408 | 0.007912 | 1.0577 | 0.025116 | 23.7458 | 0.027696 | 0.13573 | -7.4e-05 | -0.005478 | -0.005478 | -1.0084
2019-03-29 11:25 | 1553873107.8503 | -612 | 0.20825 | 0.000517 | 0 | 98.7825 | 98.5872 | 127.0461 | 20.9043 | 0.001716 | 0.001578 | -152.0144 | 0.22898 | 8.401 | 0.0022 | 0.00318 | 0.002994 | 7e-05 | 0.000737 | 75.4589 | 1.3543 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.7789 | 74.4113 | 59.4446 | 0.007912 | 1.0615 | 0.025155 | 23.6975 | 0.017706 | 0.13622 | -7.9e-05 | -0.005936 | -0.005936 | -0.93402
2019-03-29 11:25 | 1553873108.8509 | -611 | 0.20827 | 0.000509 | 0 | 98.9199 | 98.5104 | 127.251 | 20.9051 | 0.001749 | 0.00158 | -152.0304 | 0.22786 | 8.4029 | 0.00223 | 0.003489 | 0.003323 | 6.9e-05 | 0.000736 | 75.4306 | 1.3688 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.5577 | 75.1321 | 59.504 | 0.007912 | 1.0613 | 0.025173 | 23.7187 | 0.037108 | 0.13555 | -7e-05 | -0.005249 | -0.005249 | -1.051
2019-03-29 11:25 | 1553873109.8503 | -610 | 0.20823 | 0.000513 | 0 | 99.0004 | 98.5752 | 127.6041 | 20.9046 | 0.001704 | 0.001594 | -152.0311 | 0.22792 | 8.4033 | 0.00213 | 0.003658 | 0.003489 | 6.9e-05 | 0.000743 | 75.4008 | 1.3123 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.4747 | 72.7262 | 59.8313 | 0.007912 | 1.0603 | 0.025195 | 23.7633 | 0.0385 | 0.13559 | -7.4e-05 | -0.005574 | -0.005574 | -0.99007
2019-03-29 11:25 | 1553873110.8515 | -609 | 0.20827 | 0.000512 | 0 | 98.999 | 98.8431 | 127.6151 | 20.9044 | 0.001691 | 0.001602 | -152.0227 | 0.22839 | 8.4027 | 0.002159 | 0.003707 | 0.003467 | 7e-05 | 0.000745 | 75.4619 | 1.3314 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.1626 | 73.4119 | 60.014 | 0.007912 | 1.0597 | 0.025189 | 23.7708 | 0.014088 | 0.13587 | -7.4e-05 | -0.005553 | -0.005553 | -0.99587
2019-03-29 11:25 | 1553873111.8504 | -608 | 0.20822 | 0.000511 | 0 | 99.1726 | 98.7773 | 127.6308 | 20.9056 | 0.001678 | 0.00157 | -152.026 | 0.22811 | 8.4019 | 0.002187 | 0.00334 | 0.003159 | 7e-05 | 0.000747 | 75.4706 | 1.3865 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.8505 | 74.1045 | 59.2496 | 0.007912 | 1.0621 | 0.02522 | 23.745 | 0.035722 | 0.1357 | -7.2e-05 | -0.005372 | -0.005372 | -1.0281
2019-03-29 11:25 | 1553873112.8504 | -607 | 0.20823 | 0.000513 | 0 | 99.2369 | 98.9019 | 127.4121 | 20.9063 | 0.001696 | 0.001622 | -152.0265 | 0.22862 | 8.4029 | 0.002174 | 0.003699 | 0.003486 | 7.2e-05 | 0.000757 | 75.4512 | 1.3574 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2919 | 73.7756 | 60.5109 | 0.007912 | 1.0581 | 0.025151 | 23.7696 | 0.030252 | 0.13601 | -7.5e-05 | -0.005617 | -0.005617 | -0.98547
2019-03-29 11:25 | 1553873113.8516 | -606 | 0.20827 | 0.00051 | 0 | 99.2977 | 99.209 | 127.4555 | 20.9064 | 0.001691 | 0.001624 | -152.0338 | 0.22792 | 8.4038 | 0.002119 | 0.003161 | 0.002942 | 7.1e-05 | 0.000761 | 75.4182 | 1.3656 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.1732 | 72.4621 | 60.5632 | 0.007912 | 1.0579 | 0.025153 | 23.7755 | 0.007991 | 0.13559 | -7.1e-05 | -0.005375 | -0.005375 | -1.0266
2019-03-29 11:25 | 1553873114.8505 | -605 | 0.20823 | 0.00051 | 0 | 99.3905 | 99.2076 | 127.5552 | 20.9067 | 0.001746 | 0.001553 | -152.0253 | 0.22782 | 8.4004 | 0.002185 | 0.003962 | 0.003701 | 7e-05 | 0.000768 | 75.4279 | 1.403 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.4859 | 74.0432 | 58.844 | 0.007912 | 1.0634 | 0.025228 | 23.7235 | 0.016469 | 0.13553 | -7.1e-05 | -0.005378 | -0.005378 | -1.0257
2019-03-29 11:25 | 1553873115.8516 | -604 | 0.20823 | 0.000512 | 0 | 99.5069 | 99.3605 | 127.6665 | 20.9065 | 0.001683 | 0.001575 | -152.0118 | 0.22827 | 8.4036 | 0.00217 | 0.003223 | 0.003161 | 6.7e-05 | 0.000765 | 75.4375 | 1.364 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.9763 | 73.7001 | 59.3662 | 0.007912 | 1.0617 | 0.025219 | 23.7525 | 0.013167 | 0.1358 | -7.4e-05 | -0.005556 | -0.005556 | -0.9948
2019-03-29 11:25 | 1553873116.851 | -603 | 0.20824 | 0.000511 | 0 | 99.6072 | 99.4018 | 127.6008 | 20.9063 | 0.001679 | 0.001542 | -152.009 | 0.22839 | 8.402 | 0.002234 | 0.003331 | 0.00307 | 7.1e-05 | 0.000774 | 75.3895 | 1.3257 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.8762 | 75.2283 | 58.5898 | 0.007912 | 1.0642 | 0.025242 | 23.7186 | 0.01846 | 0.13587 | -7.3e-05 | -0.005527 | -0.005527 | -1.0005
2019-03-29 11:25 | 1553873117.8497 | -602 | 0.20826 | 0.000514 | -1e-06 | 99.6443 | 99.3779 | 127.5468 | 20.907 | 0.001723 | 0.001631 | -152.0087 | 0.22875 | 8.4051 | 0.002105 | 0.003139 | 0.002956 | 7.4e-05 | 0.000771 | 75.4103 | 1.3371 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.934 | 72.1367 | 60.7209 | 0.007912 | 1.0574 | 0.025156 | 23.7896 | 0.023942 | 0.13608 | -7.5e-05 | -0.005571 | -0.005571 | -0.9942
2019-03-29 11:25 | 1553873118.8507 | -601 | 0.20823 | 0.000512 | -1e-06 | 99.7347 | 99.538 | 127.5351 | 20.9071 | 0.001715 | 0.001558 | -152.0037 | 0.22823 | 8.4033 | 0.002153 | 0.003605 | 0.003348 | 7.2e-05 | 0.000769 | 75.3796 | 1.3156 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.752 | 73.2804 | 58.9719 | 0.007912 | 1.063 | 0.025221 | 23.7262 | 0.017658 | 0.13578 | -7.3e-05 | -0.005483 | -0.005483 | -1.0079
2019-03-29 11:25 | 1553873119.85 | -600 | 0.20823 | 0.000512 | -1e-06 | 99.7952 | 99.5172 | 127.3914 | 20.9066 | 0.001782 | 0.001597 | -152.0112 | 0.22828 | 8.4038 | 0.002168 | 0.003443 | 0.003304 | 7.2e-05 | 0.000768 | 75.3392 | 1.3243 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.3704 | 73.6432 | 59.906 | 0.007912 | 1.06 | 0.025171 | 23.7461 | 0.02495 | 0.13581 | -7.3e-05 | -0.005433 | -0.005433 | -1.0173
2019-03-29 11:25 | 1553873120.8502 | -599 | 0.20826 | 0.000512 | 0 | 99.8574 | 99.5227 | 127.4621 | 20.9068 | 0.001698 | 0.0016 | -151.9993 | 0.22811 | 8.401 | 0.002156 | 0.003878 | 0.003734 | 7.2e-05 | 0.000769 | 75.4315 | 1.3439 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3369 | 73.3417 | 59.9703 | 0.007912 | 1.0598 | 0.025176 | 23.755 | 0.030038 | 0.1357 | -7.3e-05 | -0.005471 | -0.005471 | -1.0094
2019-03-29 11:25 | 1553873121.8508 | -598 | 0.20823 | 0.000512 | -1e-06 | 99.955 | 99.7988 | 127.3884 | 20.9083 | 0.001729 | 0.001603 | -152.0052 | 0.22728 | 8.4021 | 0.002207 | 0.003204 | 0.003045 | 6.7e-05 | 0.00077 | 75.397 | 1.327 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.0929 | 74.5788 | 60.0617 | 0.007912 | 1.0595 | 0.025165 | 23.7514 | 0.013994 | 0.13521 | -7.3e-05 | -0.005424 | -0.005424 | -1.0145
2019-03-29 11:25 | 1553873122.851 | -597 | 0.20823 | 0.000511 | -1e-06 | 99.9975 | 99.7919 | 127.4334 | 20.9084 | 0.001632 | 0.001614 | -151.992 | 0.22755 | 8.4034 | 0.002132 | 0.003008 | 0.002834 | 6.9e-05 | 0.000769 | 75.3617 | 1.3344 | 2.5473 | 100.29 | 20.87 | 32.124 | 60.7466 | 72.7844 | 60.3213 | 0.007912 | 1.0587 | 0.02516 | 23.7648 | 0.018408 | 0.13537 | -7.1e-05 | -0.005297 | -0.005297 | -1.0402
2019-03-29 11:25 | 1553873123.8503 | -596 | 0.20823 | 0.00051 | 0 | 100.0397 | 99.811 | 127.2778 | 20.9085 | 0.001677 | 0.001584 | -151.9717 | 0.22823 | 8.4024 | 0.002155 | 0.003662 | 0.003501 | 6.7e-05 | 0.000772 | 75.4146 | 1.3242 | 2.5473 | 100.29 | 20.87 | 32.124 | 61.8251 | 73.3325 | 59.5894 | 0.007912 | 1.061 | 0.025172 | 23.7242 | 0.02047 | 0.13578 | -7e-05 | -0.005273 | -0.005273 | -1.0479
2019-03-29 11:25 | 1553873124.8509 | -595 | 0.20824 | 0.000513 | -1e-06 | 100.1394 | 99.9213 | 127.3304 | 20.9086 | 0.00174 | 0.001586 | -151.9721 | 0.22836 | 8.4032 | 0.002174 | 0.003591 | 0.003436 | 6.6e-05 | 0.000767 | 75.422 | 1.3619 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.3509 | 73.7979 | 59.6327 | 0.007912 | 1.0609 | 0.025176 | 23.7307 | 0.019504 | 0.13585 | -7.4e-05 | -0.005482 | -0.005482 | -1.0085
2019-03-29 11:25 | 1553873125.85 | -594 | 0.20822 | 0.000516 | 0 | 100.1681 | 99.895 | 127.4252 | 20.9081 | 0.001673 | 0.001602 | -151.9808 | 0.22798 | 8.4045 | 0.002127 | 0.002944 | 0.002767 | 7e-05 | 0.000774 | 75.4322 | 1.364 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.7381 | 72.6539 | 60.035 | 0.007912 | 1.0596 | 0.02517 | 23.7539 | 0.024424 | 0.13562 | -7.7e-05 | -0.005784 | -0.005784 | -0.9543
2019-03-29 11:25 | 1553873126.8514 | -593 | 0.2082 | 0.000516 | -1e-06 | 100.2415 | 99.9947 | 127.237 | 20.9084 | 0.001725 | 0.00153 | -151.9877 | 0.228 | 8.4017 | 0.002169 | 0.003383 | 0.003154 | 6.8e-05 | 0.000769 | 75.4171 | 1.3175 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.0024 | 73.6738 | 58.2962 | 0.007912 | 1.0652 | 0.025217 | 23.6743 | 0.022051 | 0.13564 | -7.8e-05 | -0.005803 | -0.005803 | -0.95135
2019-03-29 11:25 | 1553873127.85 | -592 | 0.20822 | 0.000514 | -1e-06 | 100.2966 | 100.2394 | 127.1065 | 20.9095 | 0.001733 | 0.001576 | -151.9973 | 0.22742 | 8.4015 | 0.00222 | 0.00358 | 0.003412 | 6.8e-05 | 0.000776 | 75.4591 | 1.3478 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.1814 | 74.8995 | 59.3998 | 0.007912 | 1.0616 | 0.025162 | 23.7015 | 0.005102 | 0.13529 | -7.5e-05 | -0.005587 | -0.005587 | -0.98553
2019-03-29 11:25 | 1553873128.8502 | -591 | 0.20826 | 0.000512 | 0 | 100.3226 | 100.3868 | 126.9086 | 20.9085 | 0.001701 | 0.00158 | -151.9916 | 0.22827 | 8.4006 | 0.002145 | 0.003289 | 0.003197 | 6.4e-05 | 0.000776 | 75.3937 | 1.3782 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.412 | 73.0887 | 59.4877 | 0.007912 | 1.0614 | 0.025139 | 23.6862 | -0.005722 | 0.1358 | -7.2e-05 | -0.005407 | -0.005407 | -1.0222
2019-03-29 11:25 | 1553873129.8501 | -590 | 0.20829 | 0.000513 | 0 | 100.3765 | 100.1121 | 126.8013 | 20.9102 | 0.001702 | 0.001606 | -151.9837 | 0.22834 | 8.404 | 0.00219 | 0.003532 | 0.003244 | 6.8e-05 | 0.000779 | 75.4514 | 1.3414 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.4316 | 74.1822 | 60.1377 | 0.007912 | 1.0593 | 0.025104 | 23.6993 | 0.023593 | 0.13584 | -7.5e-05 | -0.005592 | -0.005592 | -0.98872
2019-03-29 11:25 | 1553873130.85 | -589 | 0.20827 | 0.000509 | -1e-06 | 100.3852 | 100.3049 | 126.5389 | 20.9096 | 0.001708 | 0.001588 | -151.9854 | 0.2278 | 8.4006 | 0.002236 | 0.002962 | 0.002803 | 6.4e-05 | 0.000777 | 75.4586 | 1.3419 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.5864 | 75.2871 | 59.6836 | 0.007912 | 1.0607 | 0.025095 | 23.6586 | 0.007155 | 0.13552 | -6.9e-05 | -0.005125 | -0.005125 | -1.0761
2019-03-29 11:26 | 1553873131.8509 | -588 | 0.2083 | 0.000515 | 0 | 100.4989 | 100.3646 | 126.7337 | 20.9109 | 0.001711 | 0.001631 | -151.964 | 0.22814 | 8.4 | 0.002188 | 0.003298 | 0.003164 | 6.8e-05 | 0.000777 | 75.4609 | 1.3138 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.6541 | 74.137 | 60.7295 | 0.007912 | 1.0574 | 0.025075 | 23.714 | 0.01196 | 0.13572 | -7.7e-05 | -0.005747 | -0.005747 | -0.96109
2019-03-29 11:26 | 1553873132.8511 | -587 | 0.20822 | 0.000516 | 0 | 100.4633 | 100.1912 | 126.9337 | 20.9113 | 0.001719 | 0.001625 | -151.9748 | 0.22844 | 8.4036 | 0.002186 | 0.003693 | 0.003518 | 6.5e-05 | 0.000781 | 75.4289 | 1.3295 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.842 | 74.075 | 60.5735 | 0.007912 | 1.0579 | 0.025101 | 23.7272 | 0.024254 | 0.1359 | -7.7e-05 | -0.005766 | -0.005766 | -0.95932
2019-03-29 11:26 | 1553873133.8492 | -586 | 0.20824 | 0.000512 | 0 | 100.5139 | 100.3523 | 126.9198 | 20.9117 | 0.001747 | 0.001574 | -151.9826 | 0.22744 | 8.4049 | 0.002254 | 0.00379 | 0.003577 | 6.6e-05 | 0.000783 | 75.4387 | 1.3655 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.5297 | 75.7233 | 59.3588 | 0.007912 | 1.0618 | 0.025145 | 23.6826 | 0.014387 | 0.1353 | -7.2e-05 | -0.0054 | -0.0054 | -1.0197
2019-03-29 11:26 | 1553873134.8492 | -585 | 0.20826 | 0.000516 | 0 | 100.6138 | 100.5711 | 126.9428 | 20.9117 | 0.001699 | 0.001593 | -151.9851 | 0.22739 | 8.4032 | 0.002234 | 0.003214 | 0.003059 | 6.7e-05 | 0.000781 | 75.3658 | 1.3889 | 2.5476 | 100.29 | 20.87 | 32.124 | 62.358 | 75.2353 | 59.8264 | 0.007912 | 1.0603 | 0.02513 | 23.7014 | 0.003802 | 0.13527 | -7.8e-05 | -0.005857 | -0.005857 | -0.93993
2019-03-29 11:26 | 1553873135.8492 | -584 | 0.20829 | 0.000511 | 0 | 100.612 | 100.553 | 126.8473 | 20.9115 | 0.00169 | 0.001646 | -151.9753 | 0.22903 | 8.4024 | 0.002195 | 0.003823 | 0.00364 | 6.7e-05 | 0.000782 | 75.3916 | 1.3616 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.1601 | 74.2933 | 61.0825 | 0.007912 | 1.0563 | 0.025073 | 23.7372 | 0.005251 | 0.13625 | -7.1e-05 | -0.005353 | -0.005353 | -1.0359
2019-03-29 11:26 | 1553873136.8492 | -583 | 0.20828 | 0.000516 | -1e-06 | 100.6965 | 100.5578 | 127.1217 | 20.9114 | 0.001714 | 0.001592 | -151.9599 | 0.22829 | 8.4024 | 0.002207 | 0.003994 | 0.00381 | 6.7e-05 | 0.000783 | 75.3581 | 1.3415 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7203 | 74.5779 | 59.7801 | 0.007912 | 1.0604 | 0.025149 | 23.7165 | 0.012327 | 0.13581 | -7.7e-05 | -0.005747 | -0.005747 | -0.96185
2019-03-29 11:26 | 1553873137.8503 | -582 | 0.2083 | 0.000514 | 0 | 100.7447 | 100.6083 | 127.052 | 20.9131 | 0.001708 | 0.001609 | -151.9561 | 0.22798 | 8.4045 | 0.00223 | 0.003943 | 0.003779 | 6.8e-05 | 0.000779 | 75.3565 | 1.329 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.596 | 75.1307 | 60.2015 | 0.007912 | 1.0591 | 0.025127 | 23.725 | 0.012115 | 0.13563 | -7.5e-05 | -0.005597 | -0.005597 | -0.98624
2019-03-29 11:26 | 1553873138.8493 | -581 | 0.20828 | 0.000511 | 0 | 100.7764 | 100.6578 | 127.2738 | 20.9122 | 0.001713 | 0.00159 | -151.9512 | 0.22783 | 8.4 | 0.002135 | 0.0032 | 0.003014 | 6.8e-05 | 0.000769 | 75.3771 | 1.3207 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.7062 | 72.8527 | 59.7338 | 0.007912 | 1.0606 | 0.025166 | 23.729 | 0.010533 | 0.13554 | -7.2e-05 | -0.005392 | -0.005392 | -1.0231
2019-03-29 11:26 | 1553873139.8498 | -580 | 0.20828 | 0.000512 | 0 | 100.7952 | 100.6643 | 127.5111 | 20.9137 | 0.001718 | 0.001566 | -151.9375 | 0.22718 | 8.4022 | 0.002252 | 0.003157 | 0.003026 | 6.6e-05 | 0.000768 | 75.4173 | 1.3173 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.82 | 75.672 | 59.1644 | 0.007912 | 1.0624 | 0.025211 | 23.7308 | 0.01163 | 0.13515 | -7.2e-05 | -0.005419 | -0.005419 | -1.015
2019-03-29 11:26 | 1553873140.8504 | -579 | 0.2083 | 0.000513 | 0 | 100.8446 | 100.7168 | 127.6822 | 20.913 | 0.001713 | 0.001581 | -151.9486 | 0.22824 | 8.4017 | 0.002217 | 0.00368 | 0.003549 | 6.8e-05 | 0.000769 | 75.426 | 1.3105 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.7069 | 74.8327 | 59.5367 | 0.007912 | 1.0612 | 0.025214 | 23.76 | 0.011343 | 0.13578 | -7.4e-05 | -0.005558 | -0.005558 | -0.99425
2019-03-29 11:26 | 1553873141.8504 | -578 | 0.20829 | 0.000515 | 0 | 100.8614 | 100.9284 | 127.6395 | 20.9137 | 0.001716 | 0.001585 | -151.9647 | 0.22817 | 8.402 | 0.002241 | 0.003709 | 0.003621 | 6.8e-05 | 0.000773 | 75.3064 | 1.3413 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7849 | 75.4033 | 59.622 | 0.007912 | 1.0609 | 0.025207 | 23.7591 | -0.005942 | 0.13574 | -7.7e-05 | -0.005841 | -0.005841 | -0.94581
2019-03-29 11:26 | 1553873142.8498 | -577 | 0.2083 | 0.000509 | 0 | 100.8895 | 100.8469 | 127.4447 | 20.9137 | 0.001644 | 0.001566 | -151.9579 | 0.22811 | 8.405 | 0.002217 | 0.003437 | 0.003243 | 6.7e-05 | 0.000772 | 75.3742 | 1.3152 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.0512 | 74.8194 | 59.177 | 0.007912 | 1.0623 | 0.025204 | 23.7251 | 0.003784 | 0.1357 | -6.8e-05 | -0.005125 | -0.005125 | -1.0778
2019-03-29 11:26 | 1553873143.8506 | -576 | 0.20829 | 0.000513 | 0 | 100.9308 | 100.9052 | 127.5398 | 20.9133 | 0.001749 | 0.001577 | -151.9467 | 0.22768 | 8.4034 | 0.002205 | 0.002743 | 0.002574 | 7.4e-05 | 0.00078 | 75.3826 | 1.3613 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.5625 | 74.5416 | 59.4252 | 0.007912 | 1.0616 | 0.025204 | 23.7428 | 0.002271 | 0.13545 | -7.5e-05 | -0.005614 | -0.005614 | -0.98191
2019-03-29 11:26 | 1553873144.8505 | -575 | 0.20833 | 0.000513 | -1e-06 | 100.9745 | 100.9602 | 127.3691 | 20.9138 | 0.001706 | 0.001583 | -151.9676 | 0.22845 | 8.4023 | 0.002241 | 0.003466 | 0.003338 | 7e-05 | 0.000781 | 75.3375 | 1.3736 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.5327 | 75.4002 | 59.5828 | 0.007912 | 1.061 | 0.025181 | 23.7325 | 0.001274 | 0.1359 | -7.3e-05 | -0.005467 | -0.005467 | -1.0117
2019-03-29 11:26 | 1553873145.8506 | -574 | 0.2083 | 0.000514 | 0 | 101.0144 | 100.7844 | 127.6426 | 20.9138 | 0.001665 | 0.00164 | -151.9788 | 0.22847 | 8.4038 | 0.002218 | 0.003541 | 0.003378 | 7.2e-05 | 0.000782 | 75.3523 | 1.3412 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.5437 | 74.8585 | 60.9496 | 0.007912 | 1.0567 | 0.025157 | 23.8067 | 0.020385 | 0.13592 | -7.5e-05 | -0.005637 | -0.005637 | -0.98137
2019-03-29 11:26 | 1553873146.8496 | -573 | 0.2083 | 0.000512 | 0 | 101.0347 | 100.8817 | 127.8893 | 20.9149 | 0.001725 | 0.001633 | -151.9902 | 0.22818 | 8.4041 | 0.002265 | 0.003368 | 0.003144 | 7e-05 | 0.000782 | 75.3104 | 1.323 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.9945 | 75.9851 | 60.7779 | 0.007912 | 1.0573 | 0.025187 | 23.8236 | 0.013559 | 0.13575 | -7.3e-05 | -0.0055 | -0.0055 | -1.0046
2019-03-29 11:26 | 1553873147.8496 | -572 | 0.20834 | 0.000511 | 0 | 100.9721 | 101.0052 | 128.269 | 20.9151 | 0.001694 | 0.001654 | -152.0018 | 0.22877 | 8.4031 | 0.002255 | 0.003141 | 0.002982 | 6.9e-05 | 0.000783 | 75.312 | 1.3068 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2482 | 75.7429 | 61.2925 | 0.007912 | 1.0556 | 0.025205 | 23.8773 | -0.002929 | 0.1361 | -7.1e-05 | -0.005365 | -0.005365 | -1.0324
2019-03-29 11:26 | 1553873148.8497 | -571 | 0.20828 | 0.000514 | -1e-06 | 101.1155 | 101.0919 | 128.3824 | 20.9158 | 0.001729 | 0.001554 | -151.9864 | 0.22847 | 8.4024 | 0.002208 | 0.004043 | 0.003865 | 7.2e-05 | 0.000788 | 75.3591 | 1.3131 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.1064 | 74.6173 | 58.888 | 0.007912 | 1.0633 | 0.025308 | 23.8018 | 0.00209 | 0.13592 | -7.5e-05 | -0.005637 | -0.005637 | -0.98136
2019-03-29 11:26 | 1553873149.8496 | -570 | 0.20829 | 0.000511 | 0 | 101.0721 | 100.9192 | 128.8098 | 20.9159 | 0.001695 | 0.001608 | -151.9959 | 0.22811 | 8.4018 | 0.002212 | 0.003496 | 0.00333 | 7e-05 | 0.000787 | 75.3353 | 1.3324 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2747 | 74.6984 | 60.1785 | 0.007912 | 1.0592 | 0.025301 | 23.8877 | 0.013543 | 0.1357 | -7.1e-05 | -0.00535 | -0.00535 | -1.0324
2019-03-29 11:26 | 1553873150.8487 | -569 | 0.20831 | 0.00051 | -1e-06 | 101.1568 | 101.1496 | 129.1058 | 20.9164 | 0.001716 | 0.001611 | -151.9773 | 0.22821 | 8.402 | 0.002254 | 0.003543 | 0.00334 | 6.9e-05 | 0.000787 | 75.3814 | 1.3411 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.7914 | 75.7271 | 60.2602 | 0.007912 | 1.0589 | 0.025327 | 23.9181 | 0.000642 | 0.13576 | -6.9e-05 | -0.005185 | -0.005185 | -1.0657
2019-03-29 11:26 | 1553873151.849 | -568 | 0.2083 | 0.000513 | -1e-06 | 101.1517 | 101.1001 | 129.3591 | 20.9168 | 0.001702 | 0.001651 | -151.9706 | 0.22671 | 8.4027 | 0.002142 | 0.003867 | 0.003679 | 6.7e-05 | 0.000793 | 75.3817 | 1.3065 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.435 | 73.0335 | 61.2159 | 0.007912 | 1.0559 | 0.025315 | 23.9758 | 0.004569 | 0.13487 | -7.3e-05 | -0.005522 | -0.005522 | -0.99409
2019-03-29 11:26 | 1553873152.8498 | -567 | 0.20824 | 0.000512 | -1e-06 | 101.1883 | 101.0666 | 129.8377 | 20.917 | 0.001714 | 0.001625 | -151.9819 | 0.22701 | 8.4019 | 0.002125 | 0.003364 | 0.003094 | 7.2e-05 | 0.000799 | 75.3864 | 1.2989 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7328 | 72.6087 | 60.601 | 0.007912 | 1.0578 | 0.025385 | 23.998 | 0.010757 | 0.13505 | -7.2e-05 | -0.005461 | -0.005461 | -1.0065
2019-03-29 11:26 | 1553873153.8497 | -566 | 0.20824 | 0.000513 | 0 | 101.2248 | 101.2189 | 130.0616 | 20.9171 | 0.001697 | 0.001589 | -151.9756 | 0.2278 | 8.4029 | 0.002263 | 0.003303 | 0.003162 | 6.7e-05 | 0.000804 | 75.3949 | 1.3173 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3211 | 75.9248 | 59.7316 | 0.007912 | 1.0606 | 0.02544 | 23.9874 | 0.000523 | 0.13552 | -7.4e-05 | -0.005606 | -0.005606 | -0.98389
2019-03-29 11:26 | 1553873154.8498 | -565 | 0.20825 | 0.000509 | -1e-06 | 101.1745 | 101.111 | 130.405 | 20.9169 | 0.001782 | 0.001586 | -151.9565 | 0.22838 | 8.4007 | 0.002207 | 0.003334 | 0.003186 | 6.7e-05 | 0.000807 | 75.3954 | 1.3297 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.3764 | 74.5958 | 59.6603 | 0.007912 | 1.0608 | 0.025477 | 24.0165 | 0.005614 | 0.13586 | -6.9e-05 | -0.005202 | -0.005202 | -1.063
2019-03-29 11:26 | 1553873155.8491 | -564 | 0.20825 | 0.000513 | -1e-06 | 101.2433 | 101.2315 | 130.2475 | 20.9174 | 0.001778 | 0.001581 | -151.9403 | 0.22828 | 8.4027 | 0.002238 | 0.00351 | 0.003388 | 7e-05 | 0.000814 | 75.3471 | 1.3496 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.2794 | 75.3256 | 59.5246 | 0.007912 | 1.0612 | 0.025467 | 23.9971 | 0.001047 | 0.13581 | -7.3e-05 | -0.005548 | -0.005548 | -0.99625
2019-03-29 11:26 | 1553873156.8489 | -563 | 0.20826 | 0.00051 | 0 | 101.238 | 101.2943 | 130.1642 | 20.917 | 0.001699 | 0.001596 | -151.9223 | 0.2277 | 8.4023 | 0.002251 | 0.003329 | 0.003139 | 6.9e-05 | 0.000816 | 75.3513 | 1.3615 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3851 | 75.6395 | 59.8957 | 0.007912 | 1.0601 | 0.025444 | 24.0028 | -0.004975 | 0.13546 | -7e-05 | -0.005312 | -0.005312 | -1.0378
2019-03-29 11:26 | 1553873157.8499 | -562 | 0.20825 | 0.000512 | -1e-06 | 101.3005 | 101.2485 | 130.1602 | 20.9178 | 0.001738 | 0.001588 | -151.9213 | 0.22778 | 8.4016 | 0.002188 | 0.003602 | 0.003345 | 6.9e-05 | 0.000815 | 75.3805 | 1.3443 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.32 | 74.1313 | 59.7089 | 0.007912 | 1.0606 | 0.025451 | 23.9957 | 0.004589 | 0.1355 | -7.2e-05 | -0.005421 | -0.005421 | -1.0173
2019-03-29 11:26 | 1553873158.8502 | -561 | 0.20829 | 0.000512 | -1e-06 | 101.2721 | 101.2666 | 130.3135 | 20.9178 | 0.001683 | 0.001624 | -151.9335 | 0.22827 | 8.4039 | 0.002161 | 0.003245 | 0.003057 | 7e-05 | 0.000806 | 75.3798 | 1.3318 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.9866 | 73.4912 | 60.569 | 0.007912 | 1.0579 | 0.025433 | 24.0408 | 0.00048 | 0.1358 | -7.3e-05 | -0.005479 | -0.005479 | -1.0087
2019-03-29 11:26 | 1553873159.8499 | -560 | 0.20826 | 0.000511 | -1e-06 | 101.2984 | 101.0748 | 130.056 | 20.9179 | 0.001734 | 0.001563 | -151.9217 | 0.22785 | 8.4015 | 0.002195 | 0.002732 | 0.002583 | 7e-05 | 0.000799 | 75.4166 | 1.3702 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.2209 | 74.2981 | 59.0934 | 0.007912 | 1.0626 | 0.025464 | 23.9639 | 0.019764 | 0.13555 | -7e-05 | -0.005316 | -0.005316 | -1.0377
2019-03-29 11:26 | 1553873160.85 | -559 | 0.20826 | 0.000514 | 0 | 101.3218 | 101.1455 | 130.1543 | 20.9196 | 0.001699 | 0.001612 | -151.9354 | 0.22774 | 8.4042 | 0.002151 | 0.00313 | 0.003002 | 6.8e-05 | 0.000797 | 75.4106 | 1.3872 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3642 | 73.2481 | 60.2751 | 0.007912 | 1.0588 | 0.025429 | 24.0155 | 0.015576 | 0.13548 | -7.5e-05 | -0.005689 | -0.005689 | -0.96925
2019-03-29 11:26 | 1553873161.8496 | -558 | 0.20826 | 0.000511 | 0 | 101.3301 | 101.2646 | 130.0522 | 20.9193 | 0.001699 | 0.001597 | -151.933 | 0.22878 | 8.4033 | 0.002163 | 0.003107 | 0.002879 | 7e-05 | 0.000792 | 75.4309 | 1.3537 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3647 | 73.5292 | 59.9254 | 0.007912 | 1.06 | 0.025432 | 23.9935 | 0.005789 | 0.1361 | -7.1e-05 | -0.005364 | -0.005364 | -1.0326
2019-03-29 11:26 | 1553873162.8491 | -557 | 0.20825 | 0.000514 | -1e-06 | 101.3523 | 101.3366 | 129.667 | 20.9188 | 0.001693 | 0.001605 | -151.9153 | 0.22842 | 8.4009 | 0.002253 | 0.003427 | 0.003209 | 6.9e-05 | 0.00079 | 75.4295 | 1.3203 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2281 | 75.6833 | 60.1196 | 0.007912 | 1.0593 | 0.025387 | 23.965 | 0.001385 | 0.13589 | -7.4e-05 | -0.005599 | -0.005599 | -0.98769
2019-03-29 11:26 | 1553873163.8502 | -556 | 0.20825 | 0.000511 | -1e-06 | 101.3795 | 101.282 | 129.6175 | 20.9187 | 0.001667 | 0.001635 | -151.9247 | 0.22889 | 8.401 | 0.00224 | 0.003371 | 0.003175 | 6.5e-05 | 0.000793 | 75.4104 | 1.3132 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.5961 | 75.3713 | 60.8325 | 0.007912 | 1.0571 | 0.025355 | 23.986 | 0.00861 | 0.13617 | -7.1e-05 | -0.005382 | -0.005382 | -1.0298
2019-03-29 11:26 | 1553873164.8482 | -555 | 0.20826 | 0.000513 | 0 | 101.3613 | 101.2315 | 129.4579 | 20.9195 | 0.001742 | 0.001658 | -151.9106 | 0.2282 | 8.402 | 0.002182 | 0.004198 | 0.004099 | 6.4e-05 | 0.00079 | 75.3686 | 1.3515 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.4068 | 73.9964 | 61.3866 | 0.007912 | 1.0553 | 0.025318 | 23.9911 | 0.011461 | 0.13576 | -7.4e-05 | -0.005574 | -0.005574 | -0.9912
2019-03-29 11:26 | 1553873165.8482 | -554 | 0.20825 | 0.000512 | 0 | 101.3903 | 101.2878 | 129.6229 | 20.9201 | 0.001725 | 0.001641 | -151.938 | 0.22823 | 8.4019 | 0.002175 | 0.003051 | 0.002914 | 6.6e-05 | 0.000795 | 75.3419 | 1.3489 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.009 | 73.8286 | 60.9799 | 0.007912 | 1.0566 | 0.02535 | 23.9918 | 0.009048 | 0.13578 | -7.2e-05 | -0.00541 | -0.00541 | -1.0215
2019-03-29 11:26 | 1553873166.8492 | -553 | 0.20826 | 0.000512 | 0 | 101.3981 | 101.1427 | 129.7227 | 20.9197 | 0.001706 | 0.001601 | -151.9586 | 0.22736 | 8.4019 | 0.002197 | 0.003037 | 0.002795 | 7e-05 | 0.000792 | 75.3719 | 1.3576 | 2.5476 | 100.29 | 20.87 | 32.124 | 62.5526 | 74.3606 | 60.0055 | 0.007912 | 1.0597 | 0.025397 | 23.966 | 0.022553 | 0.13526 | -7.3e-05 | -0.005586 | -0.005586 | -0.98547
2019-03-29 11:26 | 1553873167.8493 | -552 | 0.20822 | 0.000514 | -1e-06 | 101.346 | 101.2902 | 129.8809 | 20.9196 | 0.001727 | 0.001656 | -151.9361 | 0.22827 | 8.4014 | 0.002177 | 0.004241 | 0.004076 | 6.4e-05 | 0.000792 | 75.3878 | 1.3688 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.061 | 73.8671 | 61.3342 | 0.007912 | 1.0555 | 0.025362 | 24.0284 | 0.004928 | 0.1358 | -7.4e-05 | -0.005585 | -0.005585 | -0.98955
2019-03-29 11:26 | 1553873168.8494 | -551 | 0.20826 | 0.000516 | 0 | 101.3059 | 101.2328 | 129.6729 | 20.9207 | 0.001736 | 0.001634 | -151.9445 | 0.22842 | 8.4012 | 0.002198 | 0.003294 | 0.003187 | 6.6e-05 | 0.000789 | 75.3853 | 1.3865 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.276 | 74.3765 | 60.8171 | 0.007912 | 1.0571 | 0.025361 | 23.9906 | 0.006452 | 0.13589 | -7.8e-05 | -0.005897 | -0.005897 | -0.93783
2019-03-29 11:26 | 1553873169.8486 | -550 | 0.20826 | 0.000511 | 0 | 101.3376 | 101.4509 | 129.7255 | 20.9198 | 0.001784 | 0.001604 | -151.9428 | 0.22852 | 8.4012 | 0.002163 | 0.00422 | 0.004024 | 6.9e-05 | 0.000792 | 75.3552 | 1.3974 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.4247 | 73.5235 | 60.0987 | 0.007912 | 1.0594 | 0.025394 | 23.9696 | -0.009995 | 0.13595 | -7.1e-05 | -0.005338 | -0.005338 | -1.0365
2019-03-29 11:26 | 1553873170.8484 | -549 | 0.20826 | 0.000514 | 0 | 101.302 | 101.1806 | 129.668 | 20.9214 | 0.001698 | 0.001642 | -151.9361 | 0.22889 | 8.4018 | 0.002197 | 0.00401 | 0.003863 | 7.2e-05 | 0.000795 | 75.3927 | 1.4043 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3448 | 74.3623 | 60.9946 | 0.007912 | 1.0566 | 0.025354 | 23.9965 | 0.010723 | 0.13617 | -7.5e-05 | -0.005671 | -0.005671 | -0.97735
2019-03-29 11:26 | 1553873171.8488 | -548 | 0.20825 | 0.000512 | 0 | 101.2511 | 101.1909 | 129.6565 | 20.9213 | 0.001694 | 0.001602 | -151.9291 | 0.22792 | 8.4026 | 0.002149 | 0.003856 | 0.003639 | 7e-05 | 0.000793 | 75.3216 | 1.3335 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2462 | 73.1953 | 60.0343 | 0.007912 | 1.0596 | 0.025389 | 23.9609 | 0.005324 | 0.13559 | -7.3e-05 | -0.005547 | -0.005547 | -0.99492
2019-03-29 11:26 | 1553873172.8494 | -547 | 0.20822 | 0.000512 | 0 | 101.3469 | 101.298 | 129.9102 | 20.9223 | 0.001768 | 0.001588 | -151.9259 | 0.22828 | 8.4015 | 0.00217 | 0.003687 | 0.00358 | 6.8e-05 | 0.0008 | 75.361 | 1.3415 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.0435 | 73.696 | 59.7008 | 0.007912 | 1.0607 | 0.025427 | 23.9723 | 0.004314 | 0.13581 | -7.3e-05 | -0.005541 | -0.005541 | -0.99744
2019-03-29 11:26 | 1553873173.8498 | -546 | 0.20825 | 0.000514 | 0 | 101.3529 | 101.326 | 130.2262 | 20.9212 | 0.001707 | 0.001646 | -151.9284 | 0.22806 | 8.4043 | 0.00226 | 0.004464 | 0.00421 | 7.2e-05 | 0.000806 | 75.369 | 1.3594 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.5611 | 75.8665 | 61.1122 | 0.007912 | 1.0562 | 0.025404 | 24.0523 | 0.002372 | 0.13567 | -7.6e-05 | -0.005797 | -0.005797 | -0.95251
2019-03-29 11:26 | 1553873174.8486 | -545 | 0.20825 | 0.000511 | 0 | 101.3011 | 101.3171 | 130.2609 | 20.9217 | 0.001698 | 0.001615 | -151.938 | 0.22888 | 8.4028 | 0.002253 | 0.003742 | 0.0036 | 7e-05 | 0.000809 | 75.3717 | 1.3175 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.3482 | 75.6853 | 60.3586 | 0.007912 | 1.0586 | 0.025436 | 24.0284 | -0.001416 | 0.13616 | -7.1e-05 | -0.005381 | -0.005381 | -1.03
2019-03-29 11:26 | 1553873175.8494 | -544 | 0.20826 | 0.000514 | 0 | 101.3227 | 101.1916 | 130.3236 | 20.9217 | 0.001685 | 0.001623 | -151.9216 | 0.22794 | 8.401 | 0.002222 | 0.003204 | 0.003063 | 6.8e-05 | 0.000809 | 75.3287 | 1.3322 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.0302 | 74.9457 | 60.5422 | 0.007912 | 1.058 | 0.025435 | 24.0408 | 0.011581 | 0.1356 | -7.5e-05 | -0.005731 | -0.005731 | -0.96299
2019-03-29 11:26 | 1553873176.8489 | -543 | 0.20826 | 0.000513 | 0 | 101.2679 | 101.2404 | 130.1257 | 20.9222 | 0.00171 | 0.001604 | -151.9214 | 0.22894 | 8.4013 | 0.002205 | 0.0032 | 0.00296 | 7e-05 | 0.000809 | 75.3331 | 1.3185 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.6417 | 74.5392 | 60.093 | 0.007912 | 1.0594 | 0.025433 | 24.0063 | 0.002432 | 0.1362 | -7.5e-05 | -0.005689 | -0.005689 | -0.97437
2019-03-29 11:26 | 1553873177.8496 | -542 | 0.20826 | 0.000512 | 0 | 101.3343 | 101.1168 | 129.89 | 20.9228 | 0.001719 | 0.0016 | -151.9301 | 0.22775 | 8.4014 | 0.002201 | 0.003745 | 0.003578 | 6.9e-05 | 0.000813 | 75.3412 | 1.2962 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.8554 | 74.4438 | 59.9827 | 0.007912 | 1.0598 | 0.025414 | 23.9806 | 0.019221 | 0.13549 | -7.2e-05 | -0.005427 | -0.005427 | -1.0161
2019-03-29 11:26 | 1553873178.8496 | -541 | 0.20828 | 0.000514 | 0 | 101.2903 | 101.3895 | 129.8762 | 20.9232 | 0.00176 | 0.001603 | -151.9235 | 0.22802 | 8.4026 | 0.002187 | 0.003497 | 0.003253 | 6.8e-05 | 0.000811 | 75.4072 | 1.3344 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.8533 | 74.106 | 60.0714 | 0.007912 | 1.0595 | 0.025409 | 23.9825 | -0.008753 | 0.13565 | -7.5e-05 | -0.005714 | -0.005714 | -0.9662
2019-03-29 11:26 | 1553873179.8489 | -540 | 0.20832 | 0.000512 | 0 | 101.2778 | 101.2383 | 129.6466 | 20.9233 | 0.001684 | 0.001629 | -151.9244 | 0.228 | 8.4012 | 0.002175 | 0.003618 | 0.003444 | 6.7e-05 | 0.000808 | 75.3904 | 1.3173 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.0257 | 73.8279 | 60.6852 | 0.007912 | 1.0575 | 0.025363 | 23.9834 | 0.003485 | 0.13564 | -7.2e-05 | -0.005502 | -0.005502 | -1.0033
2019-03-29 11:26 | 1553873180.8477 | -539 | 0.20828 | 0.000514 | 0 | 101.2936 | 101.2554 | 129.4307 | 20.9233 | 0.001742 | 0.001612 | -151.9238 | 0.22834 | 8.4009 | 0.002247 | 0.003858 | 0.003601 | 6.8e-05 | 0.000802 | 75.3758 | 1.3184 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.4223 | 75.5466 | 60.2874 | 0.007912 | 1.0588 | 0.025357 | 23.9491 | 0.003378 | 0.13584 | -7.5e-05 | -0.005664 | -0.005664 | -0.97605
2019-03-29 11:26 | 1553873181.8488 | -538 | 0.2083 | 0.000517 | -1e-06 | 101.3361 | 101.2373 | 129.5281 | 20.9235 | 0.001688 | 0.001581 | -151.9306 | 0.22808 | 8.4002 | 0.002273 | 0.003659 | 0.003607 | 6.8e-05 | 0.000804 | 75.3448 | 1.3344 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.1094 | 76.1813 | 59.5291 | 0.007912 | 1.0612 | 0.025396 | 23.9309 | 0.008729 | 0.13569 | -7.8e-05 | -0.005884 | -0.005884 | -0.93859
2019-03-29 11:26 | 1553873182.8479 | -537 | 0.20827 | 0.000514 | 0 | 101.2371 | 101.0567 | 129.468 | 20.9245 | 0.001708 | 0.00165 | -151.9404 | 0.22867 | 8.402 | 0.0023 | 0.003544 | 0.003342 | 6.7e-05 | 0.000799 | 75.3127 | 1.3414 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.5931 | 76.8292 | 61.1999 | 0.007912 | 1.0559 | 0.025326 | 23.9853 | 0.015945 | 0.13604 | -7.5e-05 | -0.005657 | -0.005657 | -0.97874
2019-03-29 11:26 | 1553873183.8488 | -536 | 0.20826 | 0.000513 | 0 | 101.2775 | 101.3547 | 129.0838 | 20.9241 | 0.001671 | 0.001628 | -151.9471 | 0.22923 | 8.4037 | 0.002271 | 0.00318 | 0.003027 | 6.7e-05 | 0.000798 | 75.296 | 1.3589 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.7044 | 76.1266 | 60.6594 | 0.007912 | 1.0576 | 0.025309 | 23.9303 | -0.006818 | 0.13637 | -7.4e-05 | -0.005609 | -0.005609 | -0.98942
2019-03-29 11:26 | 1553873184.8479 | -535 | 0.20824 | 0.000511 | 0 | 101.2224 | 101.3093 | 128.812 | 20.924 | 0.001723 | 0.001638 | -151.9284 | 0.22886 | 8.4032 | 0.002251 | 0.003781 | 0.003689 | 6.8e-05 | 0.000801 | 75.2407 | 1.3225 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.9526 | 75.6619 | 60.9227 | 0.007912 | 1.0568 | 0.025273 | 23.9146 | -0.007677 | 0.13615 | -7.2e-05 | -0.005474 | -0.005474 | -1.0124
2019-03-29 11:26 | 1553873185.8478 | -534 | 0.20826 | 0.000512 | 0 | 101.2933 | 101.0888 | 128.507 | 20.9242 | 0.001716 | 0.001585 | -151.9317 | 0.22837 | 8.4019 | 0.00226 | 0.004102 | 0.003904 | 6.8e-05 | 0.0008 | 75.2565 | 1.3292 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7913 | 75.8625 | 59.64 | 0.007912 | 1.0609 | 0.025291 | 23.8403 | 0.018076 | 0.13586 | -7.3e-05 | -0.005508 | -0.005508 | -1.0039
2019-03-29 11:26 | 1553873186.8479 | -533 | 0.20828 | 0.000515 | 0 | 101.1954 | 101.0523 | 128.4782 | 20.9248 | 0.001751 | 0.001568 | -151.9313 | 0.22908 | 8.4017 | 0.002258 | 0.003586 | 0.003385 | 7e-05 | 0.000802 | 75.184 | 1.3562 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.6284 | 75.8318 | 59.2206 | 0.007912 | 1.0622 | 0.025305 | 23.8226 | 0.012661 | 0.13628 | -7.6e-05 | -0.005784 | -0.005784 | -0.95892
2019-03-29 11:26 | 1553873187.8489 | -532 | 0.20828 | 0.000513 | 0 | 101.1694 | 101.312 | 128.3564 | 20.9252 | 0.00172 | 0.001586 | -151.9421 | 0.22855 | 8.4016 | 0.002305 | 0.0033 | 0.003188 | 6.6e-05 | 0.000799 | 75.1991 | 1.3274 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.8907 | 76.9422 | 59.6496 | 0.007912 | 1.0608 | 0.025276 | 23.8267 | -0.012601 | 0.13596 | -7.3e-05 | -0.005511 | -0.005511 | -1.0042
2019-03-29 11:26 | 1553873188.8479 | -531 | 0.20827 | 0.000512 | 0 | 101.1326 | 101.0898 | 128.3711 | 20.9257 | 0.001729 | 0.001621 | -151.9566 | 0.22934 | 8.4042 | 0.002273 | 0.003609 | 0.003452 | 6.9e-05 | 0.000797 | 75.1469 | 1.2893 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.1066 | 76.1818 | 60.5078 | 0.007912 | 1.0581 | 0.025245 | 23.8588 | 0.003781 | 0.13643 | -7.4e-05 | -0.00557 | -0.00557 | -0.99698
2019-03-29 11:26 | 1553873189.8484 | -530 | 0.20825 | 0.000514 | 0 | 101.1712 | 101.1223 | 128.4366 | 20.9257 | 0.001744 | 0.001624 | -151.9859 | 0.22845 | 8.4022 | 0.002279 | 0.003618 | 0.003358 | 7e-05 | 0.0008 | 75.0907 | 1.3193 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.4597 | 76.3144 | 60.568 | 0.007912 | 1.0579 | 0.025249 | 23.867 | 0.004327 | 0.1359 | -7.5e-05 | -0.005696 | -0.005696 | -0.9711
2019-03-29 11:26 | 1553873190.8494 | -529 | 0.20827 | 0.000517 | 0 | 101.1637 | 101.0683 | 128.4002 | 20.9266 | 0.001714 | 0.001615 | -152.0019 | 0.22772 | 8.4051 | 0.00225 | 0.003278 | 0.00317 | 6.9e-05 | 0.000801 | 75.0708 | 1.315 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.7373 | 75.6419 | 60.3499 | 0.007912 | 1.0586 | 0.025254 | 23.8558 | 0.008436 | 0.13547 | -7.9e-05 | -0.005945 | -0.005945 | -0.92742
2019-03-29 11:26 | 1553873191.8479 | -528 | 0.20827 | 0.000516 | 0 | 101.1634 | 101.1711 | 128.3139 | 20.926 | 0.001718 | 0.001614 | -151.9995 | 0.22875 | 8.4011 | 0.002326 | 0.002849 | 0.002672 | 6.9e-05 | 0.000806 | 75.0523 | 1.3157 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.8484 | 77.4489 | 60.3377 | 0.007912 | 1.0586 | 0.025246 | 23.8474 | -0.000677 | 0.13608 | -7.7e-05 | -0.005814 | -0.005814 | -0.95257
2019-03-29 11:27 | 1553873192.8482 | -527 | 0.20828 | 0.000515 | 0 | 101.1293 | 101.0294 | 128.1479 | 20.9265 | 0.001764 | 0.001639 | -152.0017 | 0.22847 | 8.4024 | 0.002347 | 0.003397 | 0.003229 | 6.9e-05 | 0.000804 | 75.0193 | 1.3326 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.939 | 77.9552 | 60.9478 | 0.007912 | 1.0567 | 0.025207 | 23.8537 | 0.008836 | 0.13592 | -7.7e-05 | -0.005791 | -0.005791 | -0.95531
2019-03-29 11:27 | 1553873193.8494 | -526 | 0.20826 | 0.000514 | 0 | 101.1769 | 101.0834 | 128.3 | 20.9269 | 0.0018 | 0.001593 | -151.9923 | 0.22843 | 8.4035 | 0.002244 | 0.00406 | 0.003849 | 6.9e-05 | 0.00081 | 75.0347 | 1.3099 | 2.5475 | 100.29 | 20.87 | 32.124 | 64.8187 | 75.4855 | 59.8279 | 0.007912 | 1.0603 | 0.025264 | 23.8279 | 0.008272 | 0.13589 | -7.5e-05 | -0.005704 | -0.005704 | -0.96969
2019-03-29 11:27 | 1553873194.8482 | -525 | 0.20827 | 0.000515 | 0 | 101.0371 | 100.7458 | 128.338 | 20.9271 | 0.001753 | 0.001631 | -152.0183 | 0.22834 | 8.4036 | 0.002301 | 0.003613 | 0.003403 | 6.9e-05 | 0.00081 | 75.0382 | 1.3281 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.6773 | 76.845 | 60.7547 | 0.007912 | 1.0573 | 0.025233 | 23.8645 | 0.025823 | 0.13584 | -7.6e-05 | -0.005788 | -0.005788 | -0.95513
2019-03-29 11:27 | 1553873195.8484 | -524 | 0.20825 | 0.000513 | 0 | 101.0919 | 101.0759 | 128.0747 | 20.9282 | 0.001715 | 0.001612 | -152.0118 | 0.22811 | 8.4017 | 0.002355 | 0.003566 | 0.003355 | 7.1e-05 | 0.000809 | 75.0292 | 1.3453 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7708 | 78.1478 | 60.2977 | 0.007912 | 1.0588 | 0.025224 | 23.8237 | 0.001418 | 0.1357 | -7.4e-05 | -0.005615 | -0.005615 | -0.98355
2019-03-29 11:27 | 1553873196.8472 | -523 | 0.20825 | 0.000516 | 0 | 101.0027 | 101.0622 | 127.9881 | 20.9283 | 0.001711 | 0.001616 | -152.0007 | 0.22895 | 8.4027 | 0.002272 | 0.004494 | 0.004354 | 6.9e-05 | 0.000811 | 75.05 | 1.408 | 2.5477 | 100.29 | 20.87 | 32.124 | 62.6811 | 76.1696 | 60.3734 | 0.007912 | 1.0585 | 0.025213 | 23.8184 | -0.00527 | 0.1362 | -7.7e-05 | -0.005819 | -0.005819 | -0.95271
2019-03-29 11:27 | 1553873197.8483 | -522 | 0.20826 | 0.000511 | 0 | 101.0066 | 100.9086 | 127.9914 | 20.9286 | 0.00174 | 0.001605 | -151.9855 | 0.22917 | 8.4023 | 0.002301 | 0.003409 | 0.003298 | 7.1e-05 | 0.000808 | 75.0541 | 1.3868 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.3798 | 76.8558 | 60.1197 | 0.007912 | 1.0593 | 0.025222 | 23.8096 | 0.008677 | 0.13633 | -7.1e-05 | -0.005332 | -0.005332 | -1.0407
2019-03-29 11:27 | 1553873198.8473 | -521 | 0.20825 | 0.000512 | 0 | 100.9246 | 100.8649 | 127.7318 | 20.9285 | 0.001703 | 0.00157 | -151.988 | 0.2284 | 8.4036 | 0.002202 | 0.003793 | 0.003578 | 6.8e-05 | 0.000805 | 75.1 | 1.357 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.4892 | 74.4751 | 59.2702 | 0.007912 | 1.062 | 0.025229 | 23.7551 | 0.005285 | 0.13587 | -7.2e-05 | -0.005393 | -0.005393 | -1.0254
2019-03-29 11:27 | 1553873199.8473 | -520 | 0.20822 | 0.000515 | 0 | 100.9623 | 100.9189 | 127.7991 | 20.9284 | 0.001736 | 0.001613 | -151.9661 | 0.23005 | 8.4028 | 0.002302 | 0.003492 | 0.003376 | 6.9e-05 | 0.000803 | 75.067 | 1.3629 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.2781 | 76.8915 | 60.3063 | 0.007912 | 1.0587 | 0.025196 | 23.7984 | 0.003846 | 0.13685 | -7.6e-05 | -0.005661 | -0.005661 | -0.98392
2019-03-29 11:27 | 1553873200.8487 | -519 | 0.20822 | 0.000512 | -1e-06 | 100.9686 | 100.8172 | 127.831 | 20.9289 | 0.001725 | 0.0016 | -151.9586 | 0.22869 | 8.406 | 0.002277 | 0.003679 | 0.003438 | 7e-05 | 0.000794 | 75.1221 | 1.3222 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.005 | 76.2761 | 60.0065 | 0.007912 | 1.0597 | 0.025211 | 23.7906 | 0.013422 | 0.13605 | -7.1e-05 | -0.005301 | -0.005301 | -1.0445
2019-03-29 11:27 | 1553873201.8477 | -518 | 0.20824 | 0.000514 | 0 | 100.9802 | 101.0165 | 127.6205 | 20.9285 | 0.001721 | 0.001598 | -151.9752 | 0.22896 | 8.4021 | 0.002243 | 0.003421 | 0.003308 | 6.7e-05 | 0.000788 | 75.0979 | 1.3395 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.9048 | 75.4698 | 59.9577 | 0.007912 | 1.0599 | 0.025192 | 23.7693 | -0.00321 | 0.13621 | -7.4e-05 | -0.005586 | -0.005586 | -0.99237
2019-03-29 11:27 | 1553873202.8484 | -517 | 0.20824 | 0.000511 | 0 | 100.847 | 100.9694 | 127.712 | 20.9295 | 0.001748 | 0.001606 | -151.968 | 0.22795 | 8.402 | 0.00222 | 0.003454 | 0.003206 | 6.7e-05 | 0.000781 | 75.1241 | 1.3213 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.5602 | 74.9174 | 60.1447 | 0.007912 | 1.0593 | 0.025194 | 23.7845 | -0.010845 | 0.13561 | -7.2e-05 | -0.005428 | -0.005428 | -1.0167
2019-03-29 11:27 | 1553873203.8474 | -516 | 0.20827 | 0.000516 | 0 | 100.9066 | 100.8066 | 127.4191 | 20.9293 | 0.001767 | 0.001666 | -151.9601 | 0.22848 | 8.4032 | 0.002258 | 0.003966 | 0.003772 | 6.7e-05 | 0.000778 | 75.1354 | 1.312 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.0113 | 75.8213 | 61.5954 | 0.007912 | 1.0547 | 0.02511 | 23.8089 | 0.00887 | 0.13592 | -7.7e-05 | -0.005747 | -0.005747 | -0.96255
2019-03-29 11:27 | 1553873204.8475 | -515 | 0.20824 | 0.000514 | 0 | 100.9159 | 100.8441 | 127.2764 | 20.9297 | 0.001709 | 0.00164 | -151.9587 | 0.22891 | 8.4 | 0.002254 | 0.004067 | 0.003962 | 6.8e-05 | 0.000777 | 75.1824 | 1.3502 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.6348 | 75.7243 | 60.9634 | 0.007912 | 1.0567 | 0.02512 | 23.773 | 0.006361 | 0.13618 | -7.5e-05 | -0.005628 | -0.005628 | -0.98476
2019-03-29 11:27 | 1553873205.8475 | -514 | 0.20826 | 0.000516 | 0 | 100.856 | 100.8469 | 127.8036 | 20.9298 | 0.001712 | 0.001549 | -151.9547 | 0.22923 | 8.4035 | 0.002274 | 0.004126 | 0.003912 | 6.9e-05 | 0.000788 | 75.1327 | 1.3409 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.7013 | 76.2155 | 58.7692 | 0.007912 | 1.0636 | 0.025255 | 23.7439 | 0.000812 | 0.13637 | -7.7e-05 | -0.005784 | -0.005784 | -0.95956
2019-03-29 11:27 | 1553873206.8475 | -513 | 0.20827 | 0.000516 | -1e-06 | 100.8009 | 100.7718 | 127.9867 | 20.9301 | 0.001703 | 0.001673 | -151.9595 | 0.22871 | 8.402 | 0.002302 | 0.002977 | 0.002826 | 6.9e-05 | 0.000788 | 75.1277 | 1.3464 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.4876 | 76.8753 | 61.751 | 0.007912 | 1.0542 | 0.02516 | 23.8674 | 0.002588 | 0.13606 | -7.6e-05 | -0.005699 | -0.005699 | -0.97169
2019-03-29 11:27 | 1553873207.8476 | -512 | 0.2082 | 0.000512 | 0 | 100.7952 | 100.6789 | 128.1841 | 20.931 | 0.001695 | 0.001637 | -151.9564 | 0.22887 | 8.4023 | 0.002288 | 0.003598 | 0.00346 | 6.6e-05 | 0.000792 | 75.1365 | 1.4108 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3002 | 76.5393 | 60.8943 | 0.007912 | 1.0569 | 0.025212 | 23.8552 | 0.010327 | 0.13615 | -7.2e-05 | -0.005377 | -0.005377 | -1.0305
2019-03-29 11:27 | 1553873208.8486 | -511 | 0.20827 | 0.00051 | 0 | 100.7662 | 100.6544 | 128.7899 | 20.9298 | 0.001688 | 0.001595 | -151.9415 | 0.22823 | 8.4023 | 0.002284 | 0.002529 | 0.002342 | 6.8e-05 | 0.000793 | 75.1889 | 1.4084 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.1288 | 76.4448 | 59.8903 | 0.007912 | 1.0601 | 0.02531 | 23.8755 | 0.009933 | 0.13578 | -7e-05 | -0.005247 | -0.005247 | -1.0531
2019-03-29 11:27 | 1553873209.8486 | -510 | 0.20822 | 0.000514 | 0 | 100.7982 | 100.7523 | 128.9303 | 20.9303 | 0.001815 | 0.001613 | -151.9513 | 0.22895 | 8.4035 | 0.002389 | 0.003913 | 0.003708 | 7e-05 | 0.000796 | 75.1708 | 1.3802 | 2.5476 | 100.29 | 20.87 | 32.124 | 65.1785 | 78.9668 | 60.3176 | 0.007912 | 1.0587 | 0.025307 | 23.9039 | 0.004076 | 0.1362 | -7.6e-05 | -0.005735 | -0.005735 | -0.96656
2019-03-29 11:27 | 1553873210.8487 | -509 | 0.20825 | 0.000514 | 0 | 100.7216 | 100.6356 | 128.9903 | 20.9315 | 0.001715 | 0.001612 | -151.9529 | 0.22815 | 8.4008 | 0.002339 | 0.003924 | 0.003807 | 6.8e-05 | 0.0008 | 75.2057 | 1.335 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.7757 | 77.7731 | 60.286 | 0.007912 | 1.0588 | 0.025314 | 23.9083 | 0.007641 | 0.13573 | -7.5e-05 | -0.00567 | -0.00567 | -0.97428
2019-03-29 11:27 | 1553873211.8466 | -508 | 0.20823 | 0.000515 | 0 | 100.7007 | 100.6711 | 128.9995 | 20.931 | 0.001735 | 0.001636 | -151.9537 | 0.22795 | 8.4024 | 0.002369 | 0.003779 | 0.003637 | 6.6e-05 | 0.000802 | 75.1562 | 1.372 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.2618 | 78.4992 | 60.8702 | 0.007912 | 1.057 | 0.025293 | 23.9301 | 0.002626 | 0.13561 | -7.5e-05 | -0.005699 | -0.005699 | -0.96853
2019-03-29 11:27 | 1553873212.8478 | -507 | 0.20825 | 0.000516 | 0 | 100.7572 | 100.6376 | 128.7777 | 20.9326 | 0.00174 | 0.001582 | -151.9444 | 0.22841 | 8.4017 | 0.002294 | 0.003801 | 0.003615 | 6.9e-05 | 0.000799 | 75.19 | 1.3321 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.375 | 76.6944 | 59.5731 | 0.007912 | 1.0611 | 0.025321 | 23.863 | 0.010622 | 0.13588 | -7.8e-05 | -0.00593 | -0.00593 | -0.93264
2019-03-29 11:27 | 1553873213.8479 | -506 | 0.20817 | 0.000515 | 0 | 100.6713 | 100.4571 | 128.5831 | 20.9327 | 0.001683 | 0.001647 | -151.9537 | 0.22871 | 8.4034 | 0.002287 | 0.004016 | 0.003778 | 6.9e-05 | 0.000799 | 75.176 | 1.299 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.0112 | 76.5302 | 61.1407 | 0.007912 | 1.0561 | 0.025242 | 23.9011 | 0.019053 | 0.13606 | -7.6e-05 | -0.005755 | -0.005755 | -0.96229
2019-03-29 11:27 | 1553873214.8472 | -505 | 0.20827 | 0.000514 | 0 | 100.5935 | 100.6018 | 128.393 | 20.9331 | 0.001716 | 0.001633 | -151.9368 | 0.229 | 8.4009 | 0.002287 | 0.003531 | 0.003443 | 6.9e-05 | 0.000797 | 75.1822 | 1.2925 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.7898 | 76.5157 | 60.7925 | 0.007912 | 1.0572 | 0.025237 | 23.871 | -0.00074 | 0.13623 | -7.5e-05 | -0.005654 | -0.005654 | -0.98066
2019-03-29 11:27 | 1553873215.8468 | -504 | 0.20825 | 0.000513 | 0 | 100.6737 | 100.5455 | 127.9749 | 20.9328 | 0.001769 | 0.001568 | -151.9368 | 0.2287 | 8.4026 | 0.002277 | 0.003315 | 0.003043 | 7.1e-05 | 0.000804 | 75.2452 | 1.284 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.0624 | 76.2737 | 59.2229 | 0.007912 | 1.0622 | 0.025255 | 23.776 | 0.011399 | 0.13605 | -7.4e-05 | -0.005585 | -0.005585 | -0.9915
2019-03-29 11:27 | 1553873216.8469 | -503 | 0.20819 | 0.000512 | 0 | 100.592 | 100.5103 | 127.7832 | 20.9331 | 0.001732 | 0.001642 | -151.9418 | 0.22828 | 8.4009 | 0.002226 | 0.003925 | 0.003778 | 7.1e-05 | 0.000805 | 75.2307 | 1.3288 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.1796 | 75.0651 | 61.0224 | 0.007912 | 1.0565 | 0.025168 | 23.8224 | 0.007263 | 0.1358 | -7.3e-05 | -0.005449 | -0.005449 | -1.0144
2019-03-29 11:27 | 1553873217.8472 | -502 | 0.2082 | 0.00051 | 0 | 100.5576 | 100.3479 | 127.7907 | 20.9345 | 0.001672 | 0.001592 | -151.9705 | 0.22894 | 8.3985 | 0.002235 | 0.003572 | 0.003307 | 6.8e-05 | 0.000802 | 75.2402 | 1.3102 | 2.5474 | 100.29 | 20.87 | 32.124 | 61.741 | 75.2792 | 59.8102 | 0.007912 | 1.0603 | 0.025214 | 23.7799 | 0.01867 | 0.1362 | -7e-05 | -0.005279 | -0.005279 | -1.05
2019-03-29 11:27 | 1553873218.8479 | -501 | 0.20819 | 0.000514 | 0 | 100.5474 | 100.4636 | 127.3947 | 20.9344 | 0.001794 | 0.001629 | -151.9684 | 0.22923 | 8.4005 | 0.002285 | 0.003997 | 0.003796 | 6.6e-05 | 0.000809 | 75.2429 | 1.3284 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.6674 | 76.4717 | 60.6938 | 0.007912 | 1.0575 | 0.025142 | 23.7745 | 0.007459 | 0.13637 | -7.4e-05 | -0.005535 | -0.005535 | -1.0028
2019-03-29 11:27 | 1553873219.8479 | -500 | 0.20827 | 0.000511 | 0 | 100.4965 | 100.308 | 127.0762 | 20.9353 | 0.001747 | 0.001622 | -151.9571 | 0.22873 | 8.4028 | 0.002312 | 0.003804 | 0.003615 | 7.2e-05 | 0.00081 | 75.1487 | 1.3363 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.555 | 77.1228 | 60.5342 | 0.007912 | 1.058 | 0.025116 | 23.7391 | 0.016797 | 0.13607 | -7.1e-05 | -0.005283 | -0.005283 | -1.0483
2019-03-29 11:27 | 1553873220.847 | -499 | 0.20823 | 0.000511 | 1e-06 | 100.5477 | 100.4714 | 127.0487 | 20.9341 | 0.001738 | 0.001648 | -151.9413 | 0.22908 | 8.4026 | 0.002274 | 0.003377 | 0.003087 | 7.5e-05 | 0.000805 | 75.1067 | 1.3449 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.3288 | 76.2135 | 61.16 | 0.007912 | 1.056 | 0.02509 | 23.7588 | 0.006786 | 0.13628 | -7.2e-05 | -0.005436 | -0.005436 | -1.0204
2019-03-29 11:27 | 1553873221.847 | -498 | 0.20826 | 0.000514 | 1e-06 | 100.5025 | 100.4858 | 126.9572 | 20.9347 | 0.001705 | 0.001624 | -151.9273 | 0.22856 | 8.401 | 0.002262 | 0.003495 | 0.003341 | 7.2e-05 | 0.000798 | 75.1713 | 1.3107 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.5272 | 75.9223 | 60.5782 | 0.007912 | 1.0579 | 0.025103 | 23.7295 | 0.001489 | 0.13597 | -7.6e-05 | -0.005739 | -0.005739 | -0.96431
2019-03-29 11:27 | 1553873222.8476 | -497 | 0.20824 | 0.000516 | 0 | 100.4447 | 100.3923 | 127.2311 | 20.935 | 0.00175 | 0.001661 | -151.9344 | 0.22892 | 8.4014 | 0.002278 | 0.002973 | 0.002751 | 7e-05 | 0.000792 | 75.132 | 1.3451 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.6188 | 76.2992 | 61.4689 | 0.007912 | 1.0551 | 0.025097 | 23.7868 | 0.004673 | 0.13619 | -7.7e-05 | -0.005768 | -0.005768 | -0.96087
2019-03-29 11:27 | 1553873223.8481 | -496 | 0.20825 | 0.000514 | 0 | 100.4884 | 100.3776 | 127.9179 | 20.9362 | 0.001752 | 0.001647 | -151.9183 | 0.2283 | 8.401 | 0.002296 | 0.003709 | 0.003484 | 7.2e-05 | 0.000794 | 75.1521 | 1.3357 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.6767 | 76.7549 | 61.1497 | 0.007912 | 1.0561 | 0.025176 | 23.8395 | 0.009872 | 0.13582 | -7.4e-05 | -0.005582 | -0.005582 | -0.99027
2019-03-29 11:27 | 1553873224.8481 | -495 | 0.20825 | 0.000513 | 0 | 100.4447 | 100.1667 | 128.4059 | 20.9367 | 0.001784 | 0.00163 | -151.9283 | 0.2281 | 8.4026 | 0.002223 | 0.003535 | 0.003427 | 7.2e-05 | 0.000795 | 75.1462 | 1.3494 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.4307 | 74.9914 | 60.7272 | 0.007912 | 1.0574 | 0.02524 | 23.8698 | 0.024795 | 0.1357 | -7.3e-05 | -0.005514 | -0.005514 | -1.0016
2019-03-29 11:27 | 1553873225.8481 | -494 | 0.20824 | 0.000512 | 0 | 100.4366 | 100.2104 | 128.5904 | 20.9355 | 0.001689 | 0.001662 | -151.9406 | 0.22879 | 8.4007 | 0.00223 | 0.003657 | 0.003437 | 7.2e-05 | 0.000793 | 75.1623 | 1.3535 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.1452 | 75.1524 | 61.4928 | 0.007912 | 1.055 | 0.025229 | 23.9144 | 0.020175 | 0.13611 | -7.2e-05 | -0.005409 | -0.005409 | -1.0241
2019-03-29 11:27 | 1553873226.8463 | -493 | 0.20826 | 0.000517 | 0 | 100.3298 | 100.253 | 128.4592 | 20.9374 | 0.001725 | 0.001616 | -151.9277 | 0.22867 | 8.4047 | 0.002326 | 0.003835 | 0.003656 | 7.4e-05 | 0.000789 | 75.132 | 1.3306 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.0275 | 77.4732 | 60.4042 | 0.007912 | 1.0584 | 0.025258 | 23.8633 | 0.006847 | 0.13604 | -7.9e-05 | -0.005924 | -0.005924 | -0.93458
2019-03-29 11:27 | 1553873227.8462 | -492 | 0.20825 | 0.00051 | 0 | 100.3711 | 100.1602 | 128.3221 | 20.9369 | 0.001764 | 0.001666 | -151.9021 | 0.22799 | 8.4025 | 0.002233 | 0.003537 | 0.00336 | 7.4e-05 | 0.000787 | 75.1713 | 1.2942 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.9616 | 75.2403 | 61.609 | 0.007912 | 1.0546 | 0.025199 | 23.8936 | 0.018815 | 0.13563 | -7e-05 | -0.005262 | -0.005262 | -1.049
2019-03-29 11:27 | 1553873228.8476 | -491 | 0.20825 | 0.000514 | 0 | 100.3474 | 100.1042 | 128.4425 | 20.9374 | 0.001666 | 0.001629 | -151.9177 | 0.22889 | 8.4022 | 0.002254 | 0.003251 | 0.003051 | 6.9e-05 | 0.000784 | 75.1594 | 1.3144 | 2.5475 | 100.29 | 20.87 | 32.124 | 61.5986 | 75.7419 | 60.7069 | 0.007912 | 1.0575 | 0.025245 | 23.8725 | 0.021707 | 0.13617 | -7.4e-05 | -0.005577 | -0.005577 | -0.9937
2019-03-29 11:27 | 1553873229.8472 | -490 | 0.20823 | 0.000514 | 0 | 100.3106 | 100.1455 | 128.6117 | 20.9373 | 0.001728 | 0.00165 | -151.9305 | 0.2291 | 8.4018 | 0.00223 | 0.003174 | 0.003019 | 7.2e-05 | 0.000786 | 75.1544 | 1.2885 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.1015 | 75.1604 | 61.2107 | 0.007912 | 1.0559 | 0.025242 | 23.9063 | 0.014735 | 0.13629 | -7.5e-05 | -0.005631 | -0.005631 | -0.98508
2019-03-29 11:27 | 1553873230.8463 | -489 | 0.20826 | 0.000508 | 0 | 100.2786 | 99.868 | 128.3676 | 20.9372 | 0.00177 | 0.001617 | -151.9393 | 0.22851 | 8.4037 | 0.002236 | 0.003902 | 0.003673 | 7e-05 | 0.000787 | 75.1379 | 1.3036 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.0972 | 75.3028 | 60.4274 | 0.007912 | 1.0584 | 0.025248 | 23.8556 | 0.036696 | 0.13594 | -6.7e-05 | -0.005057 | -0.005057 | -1.0941
2019-03-29 11:27 | 1553873231.8467 | -488 | 0.20824 | 0.000515 | -1e-06 | 100.2777 | 100.095 | 128.3026 | 20.9373 | 0.001756 | 0.001618 | -151.9406 | 0.22847 | 8.4019 | 0.002225 | 0.003691 | 0.003458 | 7e-05 | 0.000795 | 75.1817 | 1.3054 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.7738 | 75.0465 | 60.4379 | 0.007912 | 1.0583 | 0.025241 | 23.8499 | 0.016312 | 0.13591 | -7.5e-05 | -0.005628 | -0.005628 | -0.98285
2019-03-29 11:27 | 1553873232.8463 | -487 | 0.20821 | 0.000519 | 0 | 100.2591 | 99.9448 | 128.0233 | 20.9388 | 0.001749 | 0.001589 | -151.9461 | 0.2283 | 8.3997 | 0.002191 | 0.002973 | 0.002866 | 7.1e-05 | 0.000797 | 75.1944 | 1.3581 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.5881 | 74.2326 | 59.7535 | 0.007912 | 1.0605 | 0.025239 | 23.7995 | 0.028085 | 0.13582 | -8.1e-05 | -0.00606 | -0.00606 | -0.91214
2019-03-29 11:27 | 1553873233.8473 | -486 | 0.20823 | 0.00051 | 0 | 100.1723 | 100.1028 | 127.6003 | 20.9383 | 0.001724 | 0.001591 | -151.9743 | 0.22897 | 8.4016 | 0.002218 | 0.002516 | 0.002318 | 7.2e-05 | 0.000792 | 75.1503 | 1.3459 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.9884 | 74.8652 | 59.7898 | 0.007912 | 1.0604 | 0.025196 | 23.7614 | 0.006207 | 0.13622 | -7e-05 | -0.005222 | -0.005222 | -1.0617
2019-03-29 11:27 | 1553873234.8475 | -485 | 0.20822 | 0.000514 | 0 | 100.2235 | 99.8592 | 127.5515 | 20.9379 | 0.001757 | 0.001623 | -151.9837 | 0.22767 | 8.402 | 0.002153 | 0.003617 | 0.003351 | 6.9e-05 | 0.000789 | 75.1078 | 1.3592 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.7972 | 73.3089 | 60.5603 | 0.007912 | 1.0579 | 0.025162 | 23.7844 | 0.032576 | 0.13544 | -7.4e-05 | -0.005589 | -0.005589 | -0.98634
2019-03-29 11:27 | 1553873235.8478 | -484 | 0.20825 | 0.000512 | 0 | 100.1678 | 99.9831 | 127.7421 | 20.9396 | 0.001747 | 0.001643 | -151.9782 | 0.22794 | 8.4045 | 0.002266 | 0.003129 | 0.002921 | 6.9e-05 | 0.000785 | 75.2047 | 1.3993 | 2.5476 | 100.29 | 20.87 | 32.124 | 63.5606 | 76.0285 | 61.0427 | 0.007912 | 1.0564 | 0.025163 | 23.8193 | 0.016516 | 0.1356 | -7.1e-05 | -0.005347 | -0.005347 | -1.0322
2019-03-29 11:27 | 1553873236.8465 | -483 | 0.20822 | 0.000516 | -1e-06 | 100.1101 | 100.0332 | 127.9944 | 20.9396 | 0.001751 | 0.001627 | -151.9871 | 0.22939 | 8.4031 | 0.002179 | 0.003441 | 0.003246 | 7.2e-05 | 0.000786 | 75.1284 | 1.3589 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.6371 | 73.9402 | 60.6502 | 0.007912 | 1.0577 | 0.025203 | 23.8288 | 0.006869 | 0.13647 | -7.6e-05 | -0.005711 | -0.005711 | -0.97249
2019-03-29 11:27 | 1553873237.8475 | -482 | 0.20824 | 0.000515 | 0 | 100.1179 | 99.8551 | 128.0493 | 20.9397 | 0.001748 | 0.001616 | -151.9743 | 0.22924 | 8.4032 | 0.002236 | 0.004062 | 0.003897 | 6.9e-05 | 0.000784 | 75.1427 | 1.3706 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.5764 | 75.303 | 60.4081 | 0.007912 | 1.0584 | 0.025217 | 23.8253 | 0.023508 | 0.13637 | -7.5e-05 | -0.005663 | -0.005663 | -0.9802
2019-03-29 11:27 | 1553873238.8479 | -481 | 0.20828 | 0.000513 | 0 | 100.0685 | 100.0158 | 127.8667 | 20.9391 | 0.001694 | 0.001637 | -151.9501 | 0.22969 | 8.4022 | 0.002244 | 0.002698 | 0.002518 | 7.1e-05 | 0.000783 | 75.1108 | 1.368 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.2666 | 75.4942 | 60.9039 | 0.007912 | 1.0569 | 0.025181 | 23.826 | 0.004707 | 0.13664 | -7.3e-05 | -0.005517 | -0.005517 | -1.008
2019-03-29 11:27 | 1553873239.8466 | -480 | 0.20826 | 0.000515 | 0 | 100.1086 | 99.7793 | 127.7977 | 20.9396 | 0.001705 | 0.001636 | -151.943 | 0.22886 | 8.401 | 0.002246 | 0.0026 | 0.002483 | 7.3e-05 | 0.000781 | 75.1639 | 1.33 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.5306 | 75.5437 | 60.8849 | 0.007912 | 1.0569 | 0.025174 | 23.8189 | 0.029468 | 0.13615 | -7.5e-05 | -0.005658 | -0.005658 | -0.97932
2019-03-29 11:27 | 1553873240.8467 | -479 | 0.20827 | 0.000518 | 0 | 100.0382 | 99.8899 | 128.0939 | 20.9398 | 0.001741 | 0.001653 | -151.9345 | 0.22886 | 8.4044 | 0.002337 | 0.003847 | 0.003631 | 7e-05 | 0.00078 | 75.0992 | 1.3973 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.4015 | 77.7374 | 61.2876 | 0.007912 | 1.0556 | 0.025188 | 23.8608 | 0.013274 | 0.13615 | -8e-05 | -0.005998 | -0.005998 | -0.92378
2019-03-29 11:27 | 1553873241.8467 | -478 | 0.20828 | 0.000512 | 0 | 100.0643 | 99.78 | 128.3439 | 20.9398 | 0.001784 | 0.001625 | -151.935 | 0.22887 | 8.4033 | 0.002284 | 0.003427 | 0.003251 | 7.2e-05 | 0.000788 | 75.0631 | 1.3472 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.4393 | 76.4513 | 60.6131 | 0.007912 | 1.0578 | 0.025238 | 23.86 | 0.025448 | 0.13615 | -7.2e-05 | -0.00542 | -0.00542 | -1.0225
2019-03-29 11:27 | 1553873242.847 | -477 | 0.20828 | 0.000512 | 0 | 100.002 | 99.9063 | 128.3174 | 20.9399 | 0.001696 | 0.001649 | -151.9191 | 0.22979 | 8.4028 | 0.002272 | 0.004196 | 0.003991 | 7.3e-05 | 0.00079 | 75.144 | 1.3386 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3245 | 76.1621 | 61.1962 | 0.007912 | 1.0559 | 0.025214 | 23.8784 | 0.008568 | 0.1367 | -7.2e-05 | -0.0054 | -0.0054 | -1.0303
2019-03-29 11:27 | 1553873243.8458 | -476 | 0.20827 | 0.000518 | -1e-06 | 100.0421 | 99.6506 | 128.2416 | 20.9397 | 0.001725 | 0.001637 | -151.895 | 0.22864 | 8.4024 | 0.002266 | 0.003344 | 0.003225 | 7e-05 | 0.00079 | 75.1442 | 1.2985 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.0281 | 76.0387 | 60.8934 | 0.007912 | 1.0569 | 0.025218 | 23.8605 | 0.03507 | 0.13602 | -7.9e-05 | -0.005919 | -0.005919 | -0.93525
2019-03-29 11:27 | 1553873244.8457 | -475 | 0.20827 | 0.000517 | 0 | 100.0119 | 99.8704 | 128.1421 | 20.9401 | 0.001831 | 0.001616 | -151.869 | 0.22802 | 8.4013 | 0.002305 | 0.003315 | 0.003145 | 7.4e-05 | 0.000789 | 75.1259 | 1.3327 | 2.5475 | 100.29 | 20.87 | 32.124 | 65.5628 | 76.9597 | 60.3856 | 0.007912 | 1.0585 | 0.025227 | 23.8331 | 0.012661 | 0.13565 | -7.7e-05 | -0.00582 | -0.00582 | -0.94868
2019-03-29 11:27 | 1553873245.846 | -474 | 0.20825 | 0.000511 | -1e-06 | 99.9389 | 99.6356 | 127.9062 | 20.9407 | 0.00172 | 0.001648 | -151.8705 | 0.2282 | 8.4014 | 0.002259 | 0.004251 | 0.004072 | 7.1e-05 | 0.00079 | 75.132 | 1.3227 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.9048 | 75.8528 | 61.1767 | 0.007912 | 1.056 | 0.025174 | 23.8394 | 0.02718 | 0.13576 | -7e-05 | -0.005203 | -0.005203 | -1.0621
2019-03-29 11:27 | 1553873246.8468 | -473 | 0.20826 | 0.000517 | 0 | 99.9559 | 99.7592 | 127.7172 | 20.9405 | 0.00171 | 0.001623 | -151.8654 | 0.22818 | 8.4026 | 0.00228 | 0.002904 | 0.002713 | 7.4e-05 | 0.000791 | 75.0981 | 1.33 | 2.5476 | 100.29 | 20.87 | 32.124 | 62.6509 | 76.3569 | 60.5776 | 0.007912 | 1.0579 | 0.025178 | 23.8004 | 0.017623 | 0.13574 | -7.8e-05 | -0.005857 | -0.005857 | -0.94333
2019-03-29 11:27 | 1553873247.8458 | -472 | 0.20824 | 0.000511 | 0 | 99.9227 | 99.7616 | 127.3224 | 20.9407 | 0.001819 | 0.001632 | -151.8955 | 0.22866 | 8.3992 | 0.002323 | 0.003633 | 0.003492 | 7.1e-05 | 0.000789 | 75.1411 | 1.3407 | 2.5473 | 100.29 | 20.87 | 32.124 | 65.2793 | 77.4014 | 60.7734 | 0.007912 | 1.0573 | 0.025132 | 23.7706 | 0.014435 | 0.13603 | -7e-05 | -0.005244 | -0.005244 | -1.0557
2019-03-29 11:27 | 1553873248.8461 | -471 | 0.20819 | 0.000517 | -1e-06 | 99.9499 | 99.6704 | 127.102 | 20.9414 | 0.001729 | 0.00167 | -151.8946 | 0.22908 | 8.4019 | 0.002278 | 0.003785 | 0.003682 | 6.9e-05 | 0.000787 | 75.0988 | 1.3223 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.1126 | 76.3126 | 61.6886 | 0.007912 | 1.0544 | 0.025076 | 23.7825 | 0.025047 | 0.13628 | -7.7e-05 | -0.005734 | -0.005734 | -0.96725
2019-03-29 11:27 | 1553873249.8476 | -470 | 0.20823 | 0.000516 | 0 | 99.911 | 99.7704 | 126.9565 | 20.9412 | 0.0018 | 0.001633 | -151.8936 | 0.22859 | 8.4023 | 0.002256 | 0.003783 | 0.003588 | 7e-05 | 0.000784 | 75.0781 | 1.3732 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.8326 | 75.7977 | 60.8137 | 0.007912 | 1.0571 | 0.025094 | 23.7378 | 0.012595 | 0.13599 | -7.8e-05 | -0.005827 | -0.005827 | -0.94981
2019-03-29 11:27 | 1553873250.8459 | -469 | 0.20825 | 0.000512 | 0 | 99.9329 | 99.894 | 126.756 | 20.9415 | 0.001743 | 0.001623 | -151.8977 | 0.22895 | 8.4033 | 0.002275 | 0.003396 | 0.003187 | 6.9e-05 | 0.000778 | 75.1254 | 1.3754 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.4633 | 76.2432 | 60.5646 | 0.007912 | 1.0579 | 0.025084 | 23.7102 | 0.003481 | 0.1362 | -7.2e-05 | -0.005343 | -0.005343 | -1.0375
2019-03-29 11:28 | 1553873251.847 | -468 | 0.20823 | 0.000515 | -1e-06 | 99.8981 | 99.4797 | 126.6931 | 20.9417 | 0.001804 | 0.001672 | -151.8971 | 0.22843 | 8.4031 | 0.002325 | 0.003391 | 0.003212 | 7.2e-05 | 0.000785 | 75.0907 | 1.3803 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.9266 | 77.443 | 61.7488 | 0.007912 | 1.0542 | 0.025033 | 23.7464 | 0.037549 | 0.13589 | -7.5e-05 | -0.005573 | -0.005573 | -0.99237
2019-03-29 11:28 | 1553873252.846 | -467 | 0.20824 | 0.000513 | 0 | 99.9029 | 99.6776 | 126.6546 | 20.942 | 0.001741 | 0.001642 | -151.9238 | 0.22924 | 8.4014 | 0.002355 | 0.003823 | 0.003623 | 7.2e-05 | 0.000785 | 75.1392 | 1.3611 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.4029 | 78.1559 | 61.0325 | 0.007912 | 1.0564 | 0.025056 | 23.7174 | 0.020197 | 0.13637 | -7.3e-05 | -0.00545 | -0.00545 | -1.0184
2019-03-29 11:28 | 1553873253.8464 | -466 | 0.20824 | 0.000512 | -1e-06 | 99.8868 | 99.7646 | 126.6816 | 20.9431 | 0.001758 | 0.001648 | -151.9379 | 0.22898 | 8.402 | 0.002354 | 0.004215 | 0.004055 | 7.3e-05 | 0.000787 | 75.1383 | 1.3616 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.8163 | 78.1328 | 61.1673 | 0.007912 | 1.056 | 0.025054 | 23.7247 | 0.010944 | 0.13622 | -7.1e-05 | -0.005262 | -0.005262 | -1.0536
2019-03-29 11:28 | 1553873254.8461 | -465 | 0.20826 | 0.000511 | 0 | 99.8975 | 99.639 | 126.6255 | 20.9416 | 0.001725 | 0.001589 | -151.9575 | 0.22909 | 8.4019 | 0.002347 | 0.003374 | 0.003137 | 7.3e-05 | 0.000788 | 75.1708 | 1.3384 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.0183 | 77.9861 | 59.7585 | 0.007912 | 1.0605 | 0.025101 | 23.6694 | 0.023176 | 0.13628 | -7e-05 | -0.005265 | -0.005265 | -1.0535
2019-03-29 11:28 | 1553873255.847 | -464 | 0.20822 | 0.000511 | 0 | 99.8431 | 99.6711 | 126.744 | 20.9431 | 0.001761 | 0.001638 | -151.9609 | 0.22947 | 8.4011 | 0.00226 | 0.003395 | 0.003199 | 7.3e-05 | 0.000787 | 75.2364 | 1.3767 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.8921 | 75.8948 | 60.9179 | 0.007912 | 1.0568 | 0.025069 | 23.7217 | 0.015418 | 0.13651 | -7e-05 | -0.005215 | -0.005215 | -1.0653
2019-03-29 11:28 | 1553873256.8461 | -463 | 0.20825 | 0.000515 | 0 | 99.8808 | 99.6837 | 127.0191 | 20.9433 | 0.001711 | 0.00161 | -151.955 | 0.22792 | 8.402 | 0.002269 | 0.003388 | 0.003306 | 7.2e-05 | 0.00079 | 75.158 | 1.3553 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.6957 | 76.1092 | 60.2598 | 0.007912 | 1.0589 | 0.025121 | 23.724 | 0.017663 | 0.13559 | -7.5e-05 | -0.005646 | -0.005646 | -0.97747
2019-03-29 11:28 | 1553873257.8453 | -462 | 0.20823 | 0.000515 | 0 | 99.832 | 99.6697 | 127.4294 | 20.9438 | 0.001753 | 0.001661 | -151.9549 | 0.22861 | 8.4032 | 0.002281 | 0.003628 | 0.003365 | 7e-05 | 0.000794 | 75.2456 | 1.3514 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.6889 | 76.3795 | 61.4813 | 0.007912 | 1.055 | 0.025116 | 23.8058 | 0.014548 | 0.136 | -7.5e-05 | -0.0056 | -0.0056 | -0.98851
2019-03-29 11:28 | 1553873258.8451 | -461 | 0.20825 | 0.000511 | 0 | 99.8068 | 99.5701 | 127.7512 | 20.9431 | 0.001763 | 0.001616 | -151.9632 | 0.22879 | 8.4024 | 0.00233 | 0.0036 | 0.003424 | 7.2e-05 | 0.000794 | 75.1736 | 1.3215 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.9349 | 77.5575 | 60.4059 | 0.007912 | 1.0584 | 0.025188 | 23.7975 | 0.021243 | 0.13611 | -7e-05 | -0.005283 | -0.005283 | -1.0486
2019-03-29 11:28 | 1553873259.8461 | -460 | 0.20824 | 0.000514 | 0 | 99.8146 | 99.5742 | 127.9587 | 20.9435 | 0.001757 | 0.001636 | -151.9545 | 0.22791 | 8.4002 | 0.002299 | 0.003497 | 0.003304 | 7.1e-05 | 0.000791 | 75.2287 | 1.3758 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.7917 | 76.8119 | 60.8846 | 0.007912 | 1.0569 | 0.02519 | 23.8339 | 0.021572 | 0.13559 | -7.4e-05 | -0.005614 | -0.005614 | -0.98295
2019-03-29 11:28 | 1553873260.8466 | -459 | 0.20826 | 0.000513 | 0 | 99.8015 | 99.5107 | 128.0676 | 20.9443 | 0.001758 | 0.00162 | -151.9626 | 0.22839 | 8.4018 | 0.002294 | 0.004054 | 0.003915 | 7.3e-05 | 0.000793 | 75.1993 | 1.3479 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.8293 | 76.7145 | 60.5001 | 0.007912 | 1.0581 | 0.025216 | 23.8303 | 0.026096 | 0.13587 | -7.3e-05 | -0.005485 | -0.005485 | -1.0082
2019-03-29 11:28 | 1553873261.8456 | -458 | 0.20826 | 0.00051 | 0 | 99.7988 | 99.5858 | 127.9841 | 20.9448 | 0.001753 | 0.001604 | -151.9649 | 0.22882 | 8.4045 | 0.002412 | 0.00398 | 0.003834 | 7.2e-05 | 0.000795 | 75.2239 | 1.3648 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.6891 | 79.5331 | 60.1229 | 0.007912 | 1.0593 | 0.025222 | 23.809 | 0.019108 | 0.13612 | -6.9e-05 | -0.005203 | -0.005203 | -1.0647
2019-03-29 11:28 | 1553873262.8455 | -457 | 0.20827 | 0.000513 | -1e-06 | 99.7596 | 99.6308 | 128.3843 | 20.9447 | 0.00172 | 0.001636 | -151.9507 | 0.22783 | 8.4007 | 0.00227 | 0.003737 | 0.003532 | 7e-05 | 0.000801 | 75.1864 | 1.3406 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.8954 | 76.1317 | 60.8913 | 0.007912 | 1.0569 | 0.025232 | 23.8737 | 0.011548 | 0.13554 | -7.3e-05 | -0.005437 | -0.005437 | -1.0146
2019-03-29 11:28 | 1553873263.8466 | -456 | 0.20824 | 0.000517 | 0 | 99.8054 | 99.6168 | 128.5509 | 20.9453 | 0.001704 | 0.001643 | -151.9505 | 0.22836 | 8.4005 | 0.002309 | 0.003242 | 0.003101 | 6.9e-05 | 0.000802 | 75.1478 | 1.3359 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.5132 | 77.0676 | 61.0562 | 0.007912 | 1.0564 | 0.025242 | 23.8951 | 0.016909 | 0.13585 | -7.8e-05 | -0.005887 | -0.005887 | -0.93915
2019-03-29 11:28 | 1553873264.8465 | -455 | 0.20825 | 0.000514 | 0 | 99.8461 | 99.5909 | 128.8183 | 20.9451 | 0.001762 | 0.001619 | -151.9274 | 0.22804 | 8.4036 | 0.002267 | 0.004038 | 0.003786 | 7e-05 | 0.000808 | 75.1654 | 1.317 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.908 | 76.0581 | 60.4727 | 0.007912 | 1.0582 | 0.02529 | 23.899 | 0.022886 | 0.13566 | -7.4e-05 | -0.005635 | -0.005635 | -0.97977
2019-03-29 11:28 | 1553873265.8463 | -454 | 0.20821 | 0.000515 | 0 | 99.791 | 99.4281 | 128.9723 | 20.9458 | 0.00179 | 0.001687 | -151.9209 | 0.22796 | 8.4017 | 0.002282 | 0.003793 | 0.003616 | 7e-05 | 0.000807 | 75.1995 | 1.3042 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.5856 | 76.4141 | 62.1168 | 0.007912 | 1.053 | 0.025243 | 23.9722 | 0.032584 | 0.13561 | -7.5e-05 | -0.005659 | -0.005659 | -0.97529
2019-03-29 11:28 | 1553873266.8454 | -453 | 0.20828 | 0.000516 | 0 | 99.7793 | 99.5807 | 129.0828 | 20.9453 | 0.001777 | 0.001667 | -151.9195 | 0.22826 | 8.4046 | 0.002304 | 0.002906 | 0.002712 | 6.7e-05 | 0.000804 | 75.2057 | 1.2879 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.2751 | 76.9414 | 61.6188 | 0.007912 | 1.0546 | 0.025273 | 23.9646 | 0.017824 | 0.13579 | -7.7e-05 | -0.00578 | -0.00578 | -0.95619
2019-03-29 11:28 | 1553873267.8465 | -452 | 0.20822 | 0.000515 | 0 | 99.7347 | 99.4687 | 129.2152 | 20.9467 | 0.00177 | 0.001685 | -151.9061 | 0.22873 | 8.4018 | 0.002278 | 0.003606 | 0.003471 | 6.9e-05 | 0.000807 | 75.2099 | 1.3415 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.1085 | 76.3229 | 62.0655 | 0.007912 | 1.0532 | 0.025269 | 23.9929 | 0.023886 | 0.13607 | -7.5e-05 | -0.005672 | -0.005672 | -0.97638
2019-03-29 11:28 | 1553873268.8456 | -451 | 0.20825 | 0.00051 | 0 | 99.7811 | 99.5953 | 129.2028 | 20.9456 | 0.001772 | 0.001636 | -151.9161 | 0.22795 | 8.4008 | 0.002239 | 0.003619 | 0.003396 | 7.1e-05 | 0.00081 | 75.1205 | 1.3232 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.1592 | 75.3803 | 60.8813 | 0.007912 | 1.0569 | 0.025313 | 23.9493 | 0.016667 | 0.1356 | -7e-05 | -0.005312 | -0.005312 | -1.039
2019-03-29 11:28 | 1553873269.8467 | -450 | 0.20824 | 0.000511 | 0 | 99.7889 | 99.4889 | 129.2455 | 20.9468 | 0.001778 | 0.001628 | -151.9229 | 0.22813 | 8.4024 | 0.002258 | 0.003632 | 0.003526 | 6.9e-05 | 0.00081 | 75.1338 | 1.3219 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.3009 | 75.8414 | 60.6917 | 0.007912 | 1.0575 | 0.025324 | 23.9465 | 0.026933 | 0.13571 | -7e-05 | -0.005304 | -0.005304 | -1.0414
2019-03-29 11:28 | 1553873270.8465 | -449 | 0.20824 | 0.000515 | 0 | 99.7039 | 99.6462 | 129.4086 | 20.947 | 0.001774 | 0.001649 | -151.9285 | 0.22799 | 8.4032 | 0.002222 | 0.003639 | 0.003374 | 6.7e-05 | 0.000816 | 75.1182 | 1.3523 | 2.5475 | 100.29 | 20.87 | 32.124 | 64.2027 | 74.9695 | 61.2069 | 0.007912 | 1.0559 | 0.02532 | 23.9801 | 0.005176 | 0.13563 | -7.6e-05 | -0.005769 | -0.005769 | -0.95686
2019-03-29 11:28 | 1553873271.8456 | -448 | 0.20827 | 0.000511 | 0 | 99.7934 | 99.4134 | 129.7175 | 20.9476 | 0.001759 | 0.00164 | -151.9142 | 0.22827 | 8.4029 | 0.002256 | 0.00414 | 0.003939 | 7e-05 | 0.000817 | 75.0722 | 1.3296 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.8488 | 75.8018 | 60.9707 | 0.007912 | 1.0566 | 0.02536 | 24.0002 | 0.034119 | 0.1358 | -7.1e-05 | -0.00539 | -0.00539 | -1.0254
2019-03-29 11:28 | 1553873272.8456 | -447 | 0.20827 | 0.000513 | 0 | 99.7401 | 99.5629 | 129.7778 | 20.946 | 0.001767 | 0.001623 | -151.9207 | 0.2291 | 8.4026 | 0.002206 | 0.003828 | 0.003648 | 7e-05 | 0.00082 | 75.1198 | 1.3554 | 2.5475 | 100.29 | 20.87 | 32.124 | 64.0268 | 74.5921 | 60.5683 | 0.007912 | 1.0579 | 0.025381 | 23.9913 | 0.015902 | 0.13629 | -7.3e-05 | -0.005596 | -0.005596 | -0.99131
2019-03-29 11:28 | 1553873273.845 | -446 | 0.20825 | 0.000514 | 0 | 99.7685 | 99.6401 | 130.0196 | 20.9478 | 0.00173 | 0.001621 | -151.9057 | 0.22824 | 8.4035 | 0.002194 | 0.003719 | 0.003533 | 7.2e-05 | 0.000822 | 75.1442 | 1.3169 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.1462 | 74.3154 | 60.53 | 0.007912 | 1.058 | 0.025406 | 24.0123 | 0.011526 | 0.13578 | -7.5e-05 | -0.005672 | -0.005672 | -0.97432
2019-03-29 11:28 | 1553873274.8447 | -445 | 0.20828 | 0.000514 | 0 | 99.7539 | 99.537 | 129.9698 | 20.9469 | 0.001698 | 0.001579 | -151.9192 | 0.2282 | 8.4028 | 0.002274 | 0.003372 | 0.003227 | 7e-05 | 0.000822 | 75.1702 | 1.3561 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.3665 | 76.2165 | 59.5213 | 0.007912 | 1.0612 | 0.025439 | 23.9714 | 0.019468 | 0.13576 | -7.4e-05 | -0.00558 | -0.00558 | -0.99023
2019-03-29 11:28 | 1553873275.8451 | -444 | 0.20824 | 0.000508 | 0 | 99.7751 | 99.5919 | 129.8598 | 20.9485 | 0.001743 | 0.001645 | -151.9191 | 0.22856 | 8.4022 | 0.002154 | 0.00389 | 0.003738 | 7e-05 | 0.000826 | 75.1772 | 1.3588 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.4545 | 73.3417 | 61.1084 | 0.007912 | 1.0562 | 0.025368 | 24.0183 | 0.016437 | 0.13597 | -6.6e-05 | -0.004992 | -0.004992 | -1.1086
2019-03-29 11:28 | 1553873276.8447 | -443 | 0.20826 | 0.000513 | 0 | 99.756 | 99.5274 | 129.5793 | 20.9479 | 0.001753 | 0.001671 | -151.9283 | 0.22841 | 8.4047 | 0.002219 | 0.003929 | 0.003693 | 6.5e-05 | 0.000825 | 75.1907 | 1.3337 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.6992 | 74.9023 | 61.7353 | 0.007912 | 1.0542 | 0.025317 | 24.0148 | 0.020514 | 0.13588 | -7.2e-05 | -0.005439 | -0.005439 | -1.0168
2019-03-29 11:28 | 1553873277.845 | -442 | 0.20827 | 0.000513 | 0 | 99.7033 | 99.6162 | 129.4687 | 20.9474 | 0.001771 | 0.001677 | -151.923 | 0.22893 | 8.4029 | 0.002197 | 0.003829 | 0.003695 | 6.7e-05 | 0.000826 | 75.1539 | 1.3227 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.1406 | 74.3861 | 61.8849 | 0.007912 | 1.0538 | 0.025301 | 24.01 | 0.007818 | 0.13619 | -7.3e-05 | -0.005471 | -0.005471 | -1.0131
2019-03-29 11:28 | 1553873278.8451 | -441 | 0.20825 | 0.000513 | -1e-06 | 99.7619 | 99.5223 | 129.4135 | 20.9485 | 0.001749 | 0.001639 | -151.9139 | 0.22919 | 8.403 | 0.002226 | 0.003597 | 0.003368 | 6.7e-05 | 0.000826 | 75.2253 | 1.3074 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.5958 | 75.0839 | 60.949 | 0.007912 | 1.0567 | 0.025331 | 23.9713 | 0.021511 | 0.13635 | -7.2e-05 | -0.005371 | -0.005371 | -1.0332
2019-03-29 11:28 | 1553873279.8448 | -440 | 0.20826 | 0.000515 | -1e-06 | 99.7455 | 99.4776 | 129.726 | 20.9495 | 0.001788 | 0.001658 | -151.9054 | 0.22878 | 8.4031 | 0.002238 | 0.003224 | 0.003147 | 6.7e-05 | 0.000829 | 75.1415 | 1.3351 | 2.5475 | 100.29 | 20.87 | 32.124 | 64.5361 | 75.3717 | 61.411 | 0.007912 | 1.0553 | 0.025344 | 24.0168 | 0.024054 | 0.1361 | -7.4e-05 | -0.005576 | -0.005576 | -0.99348
2019-03-29 11:28 | 1553873280.8458 | -439 | 0.20822 | 0.000511 | -1e-06 | 99.7153 | 99.5363 | 129.868 | 20.9502 | 0.00172 | 0.001651 | -151.8884 | 0.22882 | 8.4029 | 0.002151 | 0.002922 | 0.002698 | 6.7e-05 | 0.000832 | 75.1363 | 1.3381 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.9077 | 73.2739 | 61.2399 | 0.007912 | 1.0558 | 0.025364 | 24.0238 | 0.016066 | 0.13613 | -7e-05 | -0.005258 | -0.005258 | -1.0538
2019-03-29 11:28 | 1553873281.8452 | -438 | 0.20821 | 0.000512 | -1e-06 | 99.7395 | 99.4626 | 129.8349 | 20.949 | 0.001696 | 0.001635 | -151.8912 | 0.22828 | 8.4014 | 0.002159 | 0.003311 | 0.003083 | 6.7e-05 | 0.000831 | 75.1995 | 1.3283 | 2.5475 | 100.29 | 20.87 | 32.124 | 62.3288 | 73.4725 | 60.8728 | 0.007912 | 1.057 | 0.025375 | 24.0075 | 0.024868 | 0.1358 | -7.1e-05 | -0.005379 | -0.005379 | -1.0275
2019-03-29 11:28 | 1553873282.846 | -437 | 0.20827 | 0.000511 | 0 | 99.7898 | 99.5448 | 129.9065 | 20.9497 | 0.001807 | 0.001592 | -151.8863 | 0.229 | 8.403 | 0.0022 | 0.0039 | 0.003743 | 6.6e-05 | 0.000826 | 75.1885 | 1.3238 | 2.5473 | 100.29 | 20.87 | 32.124 | 65.014 | 74.456 | 59.8267 | 0.007912 | 1.0603 | 0.025422 | 23.9765 | 0.021983 | 0.13623 | -7e-05 | -0.005341 | -0.005341 | -1.0382
2019-03-29 11:28 | 1553873283.8462 | -436 | 0.20822 | 0.000513 | -1e-06 | 99.6919 | 99.4435 | 129.8734 | 20.9501 | 0.001775 | 0.001674 | -151.8989 | 0.22889 | 8.4034 | 0.00221 | 0.003883 | 0.003683 | 6.9e-05 | 0.000827 | 75.242 | 1.3237 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.244 | 74.6957 | 61.8085 | 0.007912 | 1.054 | 0.025343 | 24.0447 | 0.022317 | 0.13617 | -7.2e-05 | -0.005415 | -0.005415 | -1.0235
2019-03-29 11:28 | 1553873284.846 | -435 | 0.20826 | 0.000512 | -1e-06 | 99.7608 | 99.4554 | 129.9098 | 20.9507 | 0.001766 | 0.001659 | -151.9066 | 0.22811 | 8.4021 | 0.002242 | 0.004223 | 0.0041 | 6.8e-05 | 0.000833 | 75.2185 | 1.3139 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.0126 | 75.4555 | 61.4521 | 0.007912 | 1.0551 | 0.02536 | 24.0353 | 0.027418 | 0.1357 | -7e-05 | -0.00529 | -0.00529 | -1.044
2019-03-29 11:28 | 1553873285.845 | -434 | 0.2082 | 0.000515 | 0 | 99.7688 | 99.4592 | 129.9353 | 20.9505 | 0.001755 | 0.001605 | -151.8956 | 0.22888 | 8.402 | 0.002211 | 0.003321 | 0.00319 | 6.8e-05 | 0.000834 | 75.1402 | 1.3198 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.7487 | 74.7218 | 60.1387 | 0.007912 | 1.0593 | 0.025413 | 23.9904 | 0.027805 | 0.13616 | -7.4e-05 | -0.005627 | -0.005627 | -0.98491
2019-03-29 11:28 | 1553873286.8454 | -433 | 0.20821 | 0.000507 | 0 | 99.7087 | 99.5302 | 130.0433 | 20.9502 | 0.001777 | 0.001672 | -151.8929 | 0.22858 | 8.4029 | 0.002141 | 0.003353 | 0.003102 | 6.8e-05 | 0.000831 | 75.0974 | 1.3251 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.2737 | 73.0423 | 61.7674 | 0.007912 | 1.0541 | 0.025361 | 24.059 | 0.016027 | 0.13598 | -6.6e-05 | -0.004949 | -0.004949 | -1.1183
2019-03-29 11:28 | 1553873287.8451 | -432 | 0.20806 | 0.000515 | 0 | 99.7206 | 99.4646 | 129.96 | 20.9513 | 0.001783 | 0.001612 | -151.895 | 0.22848 | 8.4023 | 0.002238 | 0.003838 | 0.003696 | 6.5e-05 | 0.000832 | 75.1828 | 1.3455 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.4341 | 75.3547 | 60.317 | 0.007912 | 1.0587 | 0.025408 | 23.9991 | 0.022992 | 0.13592 | -7.5e-05 | -0.005682 | -0.005682 | -0.97357
2019-03-29 11:28 | 1553873288.8452 | -431 | 0.20825 | 0.000512 | 0 | 99.7625 | 99.5421 | 130.0187 | 20.9515 | 0.001833 | 0.001629 | -151.8965 | 0.22951 | 8.401 | 0.002196 | 0.003745 | 0.003524 | 6.8e-05 | 0.00084 | 75.1838 | 1.3466 | 2.5473 | 100.29 | 20.87 | 32.124 | 65.6355 | 74.3665 | 60.7151 | 0.007912 | 1.0574 | 0.025399 | 24.0189 | 0.019786 | 0.13654 | -7.1e-05 | -0.005362 | -0.005362 | -1.0364
2019-03-29 11:28 | 1553873289.8455 | -430 | 0.20826 | 0.000516 | -1e-06 | 99.7802 | 99.4315 | 129.6267 | 20.9516 | 0.001717 | 0.001682 | -151.9028 | 0.22866 | 8.4025 | 0.002164 | 0.003175 | 0.003149 | 6.7e-05 | 0.000835 | 75.1128 | 1.3211 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.8297 | 73.5997 | 61.9951 | 0.007912 | 1.0534 | 0.025312 | 24.0285 | 0.031311 | 0.13603 | -7.5e-05 | -0.005668 | -0.005668 | -0.97674
2019-03-29 11:28 | 1553873290.8452 | -429 | 0.20829 | 0.000515 | -1e-06 | 99.7395 | 99.538 | 129.392 | 20.9527 | 0.001736 | 0.001638 | -151.9114 | 0.22923 | 8.4043 | 0.002208 | 0.003388 | 0.003167 | 6.7e-05 | 0.000842 | 75.1492 | 1.3119 | 2.5475 | 100.29 | 20.87 | 32.124 | 63.3093 | 74.642 | 60.9381 | 0.007912 | 1.0567 | 0.025329 | 23.9689 | 0.018087 | 0.13637 | -7.4e-05 | -0.005565 | -0.005565 | -0.99736
2019-03-29 11:28 | 1553873291.8453 | -428 | 0.20825 | 0.000513 | -1e-06 | 99.7347 | 99.5496 | 128.9446 | 20.9525 | 0.001757 | 0.001613 | -151.919 | 0.22925 | 8.4045 | 0.002169 | 0.003827 | 0.003545 | 6.8e-05 | 0.000836 | 75.1591 | 1.3194 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.8153 | 73.7142 | 60.3273 | 0.007912 | 1.0587 | 0.025308 | 23.9055 | 0.016615 | 0.13638 | -7.2e-05 | -0.005411 | -0.005411 | -1.0259
2019-03-29 11:28 | 1553873292.8452 | -427 | 0.20825 | 0.000514 | 0 | 99.7437 | 99.482 | 128.6533 | 20.9524 | 0.001772 | 0.001643 | -151.9155 | 0.22839 | 8.4019 | 0.002239 | 0.003791 | 0.003767 | 6.9e-05 | 0.000833 | 75.186 | 1.2917 | 2.5472 | 100.29 | 20.87 | 32.124 | 64.1659 | 75.3995 | 61.06 | 0.007912 | 1.0564 | 0.025252 | 23.9047 | 0.023494 | 0.13587 | -7.4e-05 | -0.005573 | -0.005573 | -0.9923
2019-03-29 11:28 | 1553873293.8452 | -426 | 0.20823 | 0.000513 | 0 | 99.7733 | 99.65 | 128.0113 | 20.9524 | 0.001712 | 0.001595 | -151.9153 | 0.22955 | 8.4026 | 0.002152 | 0.003466 | 0.003214 | 6.8e-05 | 0.000831 | 75.161 | 1.3621 | 2.5473 | 100.29 | 20.87 | 32.124 | 62.7289 | 73.2987 | 59.8959 | 0.007912 | 1.06 | 0.025233 | 23.8035 | 0.011066 | 0.13656 | -7.2e-05 | -0.005434 | -0.005434 | -1.0227
2019-03-29 11:28 | 1553873294.8443 | -425 | 0.20829 | 0.000514 | 0 | 99.7817 | 99.294 | 127.8261 | 20.9534 | 0.001763 | 0.001682 | -151.9075 | 0.22861 | 8.4018 | 0.002216 | 0.003271 | 0.002907 | 6.9e-05 | 0.00083 | 75.1598 | 1.3845 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.948 | 74.85 | 61.9906 | 0.007912 | 1.0534 | 0.025136 | 23.8609 | 0.043826 | 0.136 | -7.4e-05 | -0.005505 | -0.005505 | -1.0055
2019-03-29 11:28 | 1553873295.8457 | -424 | 0.20821 | 0.000509 | 0 | 99.7563 | 99.5489 | 127.5567 | 20.9522 | 0.001695 | 0.001676 | -151.9009 | 0.23018 | 8.4055 | 0.002227 | 0.003921 | 0.003772 | 6.9e-05 | 0.000826 | 75.184 | 1.376 | 2.5474 | 100.29 | 20.87 | 32.124 | 62.3052 | 75.0937 | 61.8631 | 0.007912 | 1.0538 | 0.025114 | 23.8312 | 0.018609 | 0.13693 | -6.7e-05 | -0.005062 | -0.005062 | -1.1011
2019-03-29 11:28 | 1553873296.8454 | -423 | 0.20828 | 0.000515 | 0 | 99.7832 | 99.3626 | 127.532 | 20.9537 | 0.001799 | 0.001649 | -151.8827 | 0.22901 | 8.4047 | 0.002199 | 0.003253 | 0.003125 | 6.7e-05 | 0.000826 | 75.1245 | 1.4036 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.8228 | 74.4412 | 61.2152 | 0.007912 | 1.0559 | 0.025136 | 23.8059 | 0.037783 | 0.13624 | -7.5e-05 | -0.005634 | -0.005634 | -0.98413
2019-03-29 11:28 | 1553873297.8454 | -422 | 0.20825 | 0.000513 | 0 | 99.7407 | 99.5305 | 127.797 | 20.9536 | 0.001732 | 0.001647 | -151.8906 | 0.22854 | 8.4044 | 0.002188 | 0.003165 | 0.002854 | 6.9e-05 | 0.000826 | 75.1383 | 1.4053 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.2016 | 74.1674 | 61.1572 | 0.007912 | 1.0561 | 0.025164 | 23.8285 | 0.018869 | 0.13596 | -7.2e-05 | -0.005428 | -0.005428 | -1.0195
2019-03-29 11:28 | 1553873298.8448 | -421 | 0.20827 | 0.000511 | 0 | 99.788 | 99.4298 | 127.9097 | 20.9533 | 0.001863 | 0.001687 | -151.8943 | 0.22908 | 8.402 | 0.002287 | 0.003155 | 0.003033 | 6.8e-05 | 0.00082 | 75.1214 | 1.3625 | 2.5474 | 100.29 | 20.87 | 32.124 | 66.3584 | 76.5454 | 62.1211 | 0.007912 | 1.053 | 0.025139 | 23.8734 | 0.032162 | 0.13628 | -7e-05 | -0.005279 | -0.005279 | -1.0507
2019-03-29 11:28 | 1553873299.8445 | -420 | 0.20825 | 0.000515 | 0 | 99.7164 | 99.5117 | 127.7947 | 20.9534 | 0.001786 | 0.001629 | -151.8915 | 0.2299 | 8.4018 | 0.002279 | 0.002971 | 0.002829 | 6.9e-05 | 0.000817 | 75.1293 | 1.3335 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.504 | 76.3451 | 60.7277 | 0.007912 | 1.0574 | 0.02518 | 23.813 | 0.018382 | 0.13676 | -7.5e-05 | -0.00565 | -0.00565 | -0.98513
2019-03-29 11:28 | 1553873300.8455 | -419 | 0.20825 | 0.000512 | 0 | 99.7078 | 99.4837 | 127.7871 | 20.9559 | 0.001773 | 0.00168 | -151.8777 | 0.22878 | 8.4019 | 0.002201 | 0.003547 | 0.003283 | 6.6e-05 | 0.000819 | 75.1279 | 1.3551 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.1863 | 74.4738 | 61.9588 | 0.007912 | 1.0535 | 0.025133 | 23.8562 | 0.020119 | 0.1361 | -7.2e-05 | -0.005457 | -0.005457 | -1.015
2019-03-29 11:28 | 1553873301.8459 | -418 | 0.20827 | 0.000512 | 0 | 99.7464 | 99.5418 | 127.8808 | 20.9537 | 0.00175 | 0.001716 | -151.8664 | 0.22985 | 8.4029 | 0.002247 | 0.004028 | 0.003867 | 6.9e-05 | 0.000818 | 75.1542 | 1.361 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.6306 | 75.5924 | 62.823 | 0.007912 | 1.0508 | 0.02511 | 23.8956 | 0.018367 | 0.13674 | -7.1e-05 | -0.005318 | -0.005318 | -1.0466
2019-03-29 11:28 | 1553873302.8456 | -417 | 0.20827 | 0.000514 | 0 | 99.7108 | 99.4943 | 127.5794 | 20.9546 | 0.001792 | 0.001635 | -151.8657 | 0.22798 | 8.4036 | 0.002228 | 0.003527 | 0.003342 | 7e-05 | 0.000829 | 75.1112 | 1.3252 | 2.5474 | 100.29 | 20.87 | 32.124 | 64.6379 | 75.1241 | 60.8664 | 0.007912 | 1.057 | 0.025154 | 23.7979 | 0.019436 | 0.13562 | -7.4e-05 | -0.005581 | -0.005581 | -0.98911
2019-03-29 11:28 | 1553873303.8445 | -416 | 0.20828 | 0.000511 | 0 | 99.7515 | 99.4629 | 127.6088 | 20.955 | 0.001858 | 0.001653 | -151.8597 | 0.22802 | 8.4028 | 0.002233 | 0.00337 | 0.003175 | 6.7e-05 | 0.000828 | 75.0875 | 1.359 | 2.5474 | 100.29 | 20.87 | 32.124 | 66.2439 | 75.2569 | 61.307 | 0.007912 | 1.0556 | 0.02514 | 23.8163 | 0.025911 | 0.13565 | -7.1e-05 | -0.005343 | -0.005343 | -1.0332
2019-03-29 11:28 | 1553873304.8436 | -415 | 0.2083 | 0.000515 | 0 | 99.7269 | 99.6769 | 127.6097 | 20.9549 | 0.001793 | 0.001649 | -151.8668 | 0.2284 | 8.4034 | 0.002194 | 0.003865 | 0.003763 | 7e-05 | 0.000827 | 75.1227 | 1.3683 | 2.5476 | 100.29 | 20.87 | 32.124 | 64.6685 | 74.3044 | 61.2115 | 0.007912 | 1.0559 | 0.025144 | 23.813 | 0.004487 | 0.13587 | -7.5e-05 | -0.005654 | -0.005654 | -0.97801
2019-03-29 11:28 | 1553873305.8438 | -414 | 0.20825 | 0.000512 | 0 | 99.8125 | 99.5974 | 127.5386 | 20.9546 | 0.001847 | 0.001689 | -151.8934 | 0.22837 | 8.4028 | 0.002227 | 0.003113 | 0.00291 | 6.8e-05 | 0.000827 | 75.141 | 1.372 | 2.5474 | 100.29 | 20.87 | 32.124 | 65.9632 | 75.1053 | 62.1615 | 0.007912 | 1.0529 | 0.025101 | 23.8401 | 0.0193 | 0.13586 | -7.1e-05 | -0.005354 | -0.005354 | -1.0328
2019-03-29 11:28 | 1553873306.8437 | -413 | 0.20827 | 0.000511 | 0 | 99.7383 | 99.4322 | 127.4125 | 20.9553 | 0.001775 | 0.001699 | -151.9123 | 0.2277 | 8.4021 | 0.00223 | 0.003763 | 0.003617 | 6.7e-05 | 0.000829 | 75.1261 | 1.3547 | 2.5473 | 100.29 | 20.87 | 32.124 | 64.2295 | 75.1786 | 62.4122 | 0.007912 | 1.0521 | 0.025079 | 23.8373 | 0.027492 | 0.13546 | -7.1e-05 | -0.005309 | -0.005309 | -1.0385
2019-03-29 11:28 | 1553873307.8443 | -412 | 0.20826 | 0.00051 | 0 | 99.7138 | 99.3278 | 127.1572 | 20.9563 | 0.001758 | 0.001608 | -151.9102 | 0.22842 | 8.4022 | 0.002272 | 0.003525 | 0.003416 | 6.8e-05 | 0.000829 | 75.1173 | 1.3187 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.8409 | 76.1951 | 60.2091 | 0.007912 | 1.0591 | 0.025137 | 23.7351 | 0.03469 | 0.13589 | -7e-05 | -0.005278 | -0.005278 | -1.0479
2019-03-29 11:28 | 1553873308.8448 | -411 | 0.20827 | 0.00051 | 0 | 99.7853 | 99.4291 | 127.0731 | 20.9553 | 0.001749 | 0.001638 | -151.9012 | 0.22885 | 8.4011 | 0.002256 | 0.003736 | 0.003524 | 6.7e-05 | 0.000829 | 75.1378 | 1.3124 | 2.5474 | 100.29 | 20.87 | 32.124 | 63.618 | 75.7926 | 60.9437 | 0.007912 | 1.0567 | 0.025101 | 23.7534 | 0.031982 | 0.13614 | -6.9e-05 | -0.005178 | -0.005178 | -1.0702
2019-03-29 11:28 | 1553873309.8442 | -410 | 0.20828 | 0.000517 | 0 | 99.7569 | 99.6332 | 127.2191 | 20.9571 | 0.001749 | 0.001606 | -151.9133 | 0.23011 | 8.4014 | 0.002248 | 0.003729 | 0.003521 | 6.8e-05 | 0.000837 | 75.1526 | 1.3362 | 2.5473 | 100.29 | 20.87 | 32.124 | 63.6083 | 75.6059 | 60.167 | 0.007912 | 1.0592 | 0.025144 | 23.7393 | 0.011092 | 0.13689 | -7.7e-05 | -0.005786 | -0.005786 | -0.96293
2019-03-29 11:28 | 1553873310.8441 | -409 | 0.20822 | 0.00051 | 1e-06 | 99.72 | 99.5701 | 127.3233 | 20.9559 | 0.00177 | 0.001704 | -151.9309 | 0.22833 | 8.4014 | 0.002294 | 0.003423 | 0.003246 | 7e-05 | 0.000844 | 75.1506 | 1.4015 | 2.5476 | 100.3 | 21.02 | 31.856 | 64.131 | 76.7132 | 62.5392 | 0.007918 | 1.0517 | 0.025066 | 23.8334 | 0.01346 | 0.13583 | -7e-05 | -0.005299 | -0.005299 | -1.0432
2019-03-29 11:28 | 1553873311.8449 | -408 | 0.20825 | 0.000513 | 0 | 99.8012 | 99.4145 | 127.6766 | 20.9563 | 0.001715 | 0.001656 | -151.9496 | 0.22827 | 8.4019 | 0.00218 | 0.003612 | 0.003394 | 6.8e-05 | 0.000848 | 75.1489 | 1.3857 | 2.5474 | 100.3 | 21.09 | 31.723 | 62.795 | 73.9658 | 61.3754 | 0.007919 | 1.0554 | 0.025144 | 23.8251 | 0.034724 | 0.1358 | -7.3e-05 | -0.005516 | -0.005516 | -1.0021
2019-03-29 11:29 | 1553873312.8448 | -407 | 0.20824 | 0.000513 | 0 | 99.761 | 99.5517 | 128.131 | 20.9566 | 0.001789 | 0.001623 | -151.9469 | 0.22904 | 8.4019 | 0.002216 | 0.003145 | 0.003024 | 6.7e-05 | 0.00085 | 75.1897 | 1.3926 | 2.5474 | 100.3 | 21.16 | 31.592 | 64.5687 | 74.8469 | 60.5699 | 0.00792 | 1.0579 | 0.025219 | 23.8387 | 0.018793 | 0.13625 | -7.3e-05 | -0.005473 | -0.005473 | -1.0133
2019-03-29 11:29 | 1553873313.8449 | -406 | 0.20986 | 0.000515 | 0 | 99.7742 | 99.5329 | 128.2702 | 20.9568 | 0.001748 | 0.001621 | -151.9594 | 0.22947 | 8.4024 | 0.002254 | 0.003879 | 0.003706 | 6.7e-05 | 0.000851 | 75.1598 | 1.3306 | 2.5475 | 100.3 | 21.3 | 31.33 | 63.5891 | 75.7607 | 60.5293 | 0.007922 | 1.058 | 0.025234 | 23.8502 | 0.021661 | 0.13651 | -7.5e-05 | -0.00571 | -0.00571 | -0.9731
2019-03-29 11:29 | 1553873314.845 | -405 | 0.21035 | 0.00051 | 0 | 99.7563 | 99.5715 | 128.108 | 20.9579 | 0.001789 | 0.00164 | -151.9511 | 0.22897 | 8.4023 | 0.002245 | 0.00385 | 0.003687 | 6.8e-05 | 0.000844 | 75.1749 | 1.3245 | 2.5473 | 100.3 | 21.37 | 31.199 | 64.5848 | 75.5514 | 60.9863 | 0.007923 | 1.0566 | 0.025201 | 23.8514 | 0.016586 | 0.13621 | -6.9e-05 | -0.00523 | -0.00523 | -1.06
2019-03-29 11:29 | 1553873315.844 | -404 | 0.21042 | 0.000517 | 0 | 99.7877 | 99.5424 | 128.1376 | 20.9581 | 0.001732 | 0.001593 | -151.9372 | 0.2292 | 8.4047 | 0.002231 | 0.004025 | 0.003887 | 7.2e-05 | 0.000838 | 75.196 | 1.3153 | 2.5473 | 100.3 | 21.51 | 30.939 | 63.2038 | 75.2077 | 59.8582 | 0.007925 | 1.0602 | 0.025247 | 23.8139 | 0.022009 | 0.13635 | -7.7e-05 | -0.005887 | -0.005887 | -0.9427
2019-03-29 11:29 | 1553873316.844 | -403 | 0.20899 | 0.000516 | 0 | 99.7578 | 99.78 | 128.0829 | 20.958 | 0.001784 | 0.001633 | -151.9323 | 0.22945 | 8.4031 | 0.002254 | 0.003678 | 0.003502 | 6.8e-05 | 0.000839 | 75.1465 | 1.2795 | 2.5473 | 100.3 | 21.58 | 30.81 | 64.4545 | 75.7486 | 60.8223 | 0.007926 | 1.0571 | 0.025205 | 23.8432 | -0.001993 | 0.1365 | -7.6e-05 | -0.005758 | -0.005758 | -0.96484
2019-03-29 11:29 | 1553873317.844 | -402 | 0.20842 | 0.000515 | 0 | 99.7506 | 99.4636 | 128.2195 | 20.958 | 0.001763 | 0.001662 | -151.9325 | 0.2296 | 8.4011 | 0.002246 | 0.003287 | 0.003114 | 6.8e-05 | 0.000841 | 75.1418 | 1.2935 | 2.5474 | 100.3 | 21.64 | 30.68 | 63.9532 | 75.5646 | 61.5257 | 0.007921 | 1.0549 | 0.025192 | 23.881 | 0.025769 | 0.13659 | -7.5e-05 | -0.005676 | -0.005676 | -0.97944
2019-03-29 11:29 | 1553873318.8445 | -401 | 0.20857 | 0.000521 | 1e-06 | 99.7706 | 99.4018 | 127.9303 | 20.9587 | 0.001748 | 0.001635 | -151.938 | 0.22904 | 8.4013 | 0.002261 | 0.003912 | 0.003668 | 6.8e-05 | 0.000844 | 75.0938 | 1.3094 | 2.5472 | 100.3 | 21.78 | 30.421 | 63.5886 | 75.9269 | 60.8791 | 0.007922 | 1.0569 | 0.025188 | 23.831 | 0.033124 | 0.13625 | -8.4e-05 | -0.006339 | -0.006339 | -0.8748
2019-03-29 11:29 | 1553873319.8431 | -400 | 0.2087 | 0.00053 | 0 | 99.7266 | 99.6312 | 127.8118 | 20.9594 | 0.001812 | 0.001679 | -151.9385 | 0.22984 | 8.4019 | 0.002297 | 0.003442 | 0.00341 | 6.9e-05 | 0.000846 | 75.161 | 1.3477 | 2.5475 | 100.3 | 21.85 | 30.292 | 65.1257 | 76.7812 | 61.9271 | 0.007922 | 1.0536 | 0.025137 | 23.8573 | 0.008564 | 0.13673 | -9.5e-05 | -0.007186 | -0.007186 | -0.7744
2019-03-29 11:29 | 1553873320.8445 | -399 | 0.2089 | 0.000536 | 0 | 99.7796 | 99.5223 | 127.6411 | 20.9586 | 0.001738 | 0.001614 | -151.9205 | 0.22925 | 8.4022 | 0.002278 | 0.003119 | 0.002817 | 5.7e-05 | 0.00085 | 75.0672 | 1.3656 | 2.5475 | 100.3 | 21.91 | 30.165 | 63.3606 | 76.3307 | 60.3764 | 0.007918 | 1.0585 | 0.025178 | 23.7862 | 0.023095 | 0.13638 | -0.000102 | -0.007698 | -0.007698 | -0.7211
2019-03-29 11:29 | 1553873321.8435 | -398 | 0.20894 | 0.000544 | 0 | 99.7997 | 99.5305 | 127.5313 | 20.9591 | 0.001752 | 0.00171 | -151.9116 | 0.2286 | 8.4038 | 0.002391 | 0.003399 | 0.003299 | 6.9e-05 | 0.000854 | 75.0724 | 1.3317 | 2.5473 | 100.3 | 21.97 | 30.039 | 63.6982 | 79.0423 | 62.6908 | 0.007914 | 1.0512 | 0.02508 | 23.8583 | 0.024156 | 0.136 | -0.000113 | -0.00849 | -0.00849 | -0.65192
2019-03-29 11:29 | 1553873322.8442 | -397 | 0.20903 | 0.000536 | 0 | 99.7602 | 99.4701 | 127.1682 | 20.9603 | 0.001745 | 0.001645 | -151.9095 | 0.22975 | 8.4044 | 0.002258 | 0.002896 | 0.002656 | 7.2e-05 | 0.00086 | 75.0915 | 1.3903 | 2.5475 | 100.3 | 21.97 | 30.042 | 63.5299 | 75.8499 | 61.1129 | 0.007915 | 1.0562 | 0.025104 | 23.7682 | 0.026044 | 0.13668 | -0.000101 | -0.007604 | -0.007604 | -0.73157
2019-03-29 11:29 | 1553873323.8436 | -396 | 0.20913 | 0.000536 | -1e-06 | 99.7637 | 99.6155 | 127.5191 | 20.9602 | 0.001788 | 0.001651 | -151.8817 | 0.22916 | 8.4066 | 0.002247 | 0.003528 | 0.003419 | 7.1e-05 | 0.000861 | 75.1191 | 1.375 | 2.5474 | 100.3 | 21.96 | 30.046 | 64.5599 | 75.5942 | 61.2706 | 0.007911 | 1.0557 | 0.025132 | 23.8066 | 0.013302 | 0.13633 | -0.000101 | -0.007584 | -0.007584 | -0.73164
2019-03-29 11:29 | 1553873324.8439 | -395 | 0.20919 | 0.000533 | -1e-06 | 99.8335 | 99.4496 | 127.7724 | 20.961 | 0.001769 | 0.001648 | -151.8874 | 0.22942 | 8.4019 | 0.002178 | 0.0038 | 0.003587 | 7.2e-05 | 0.000864 | 75.1228 | 1.3677 | 2.5474 | 100.3 | 21.96 | 30.047 | 64.1022 | 73.9245 | 61.1916 | 0.007911 | 1.0559 | 0.02516 | 23.8274 | 0.034459 | 0.13648 | -9.8e-05 | -0.007358 | -0.007358 | -0.75491
2019-03-29 11:29 | 1553873325.8443 | -394 | 0.20921 | 0.000524 | 0 | 99.7266 | 99.5752 | 128.0543 | 20.9603 | 0.001772 | 0.001674 | -151.8798 | 0.23014 | 8.4012 | 0.002237 | 0.00342 | 0.003265 | 6.9e-05 | 0.000866 | 75.1302 | 1.3952 | 2.5475 | 100.3 | 21.95 | 30.048 | 64.1672 | 75.3526 | 61.8246 | 0.007907 | 1.0539 | 0.025164 | 23.8763 | 0.013591 | 0.13691 | -8.7e-05 | -0.006556 | -0.006556 | -0.84994
2019-03-29 11:29 | 1553873326.8433 | -393 | 0.20926 | 0.000518 | 0 | 99.7904 | 99.4517 | 128.0876 | 20.9614 | 0.001783 | 0.00169 | -151.8843 | 0.23022 | 8.4017 | 0.00215 | 0.003758 | 0.003557 | 7.1e-05 | 0.000867 | 75.1361 | 1.356 | 2.5474 | 100.3 | 21.94 | 30.05 | 64.4344 | 73.273 | 62.1946 | 0.007902 | 1.0528 | 0.025154 | 23.8926 | 0.030412 | 0.13696 | -7.9e-05 | -0.005988 | -0.005988 | -0.93087
2019-03-29 11:29 | 1553873327.8444 | -392 | 0.20927 | 0.000515 | 0 | 99.7985 | 99.4957 | 127.6534 | 20.9607 | 0.001794 | 0.001682 | -151.8892 | 0.22938 | 8.4052 | 0.002184 | 0.003359 | 0.003239 | 7.3e-05 | 0.000871 | 75.0807 | 1.3855 | 2.5474 | 100.3 | 21.94 | 30.052 | 64.7141 | 74.076 | 62.0119 | 0.007903 | 1.0534 | 0.025118 | 23.8455 | 0.027178 | 0.13646 | -7.5e-05 | -0.005659 | -0.005659 | -0.98137
2019-03-29 11:29 | 1553873328.8433 | -391 | 0.20934 | 0.000507 | 0 | 99.8122 | 99.5605 | 127.7529 | 20.962 | 0.00174 | 0.001631 | -151.8738 | 0.22967 | 8.4008 | 0.002229 | 0.004537 | 0.004348 | 7e-05 | 0.000864 | 75.1298 | 1.3432 | 2.5474 | 100.3 | 21.93 | 30.055 | 63.4004 | 75.1514 | 60.7869 | 0.007899 | 1.0572 | 0.025174 | 23.8112 | 0.022584 | 0.13663 | -6.5e-05 | -0.004875 | -0.004875 | -1.1406
2019-03-29 11:29 | 1553873329.8448 | -390 | 0.20933 | 0.000502 | 0 | 99.8332 | 99.5971 | 127.5057 | 20.9611 | 0.001761 | 0.001647 | -151.8595 | 0.22902 | 8.403 | 0.002238 | 0.003721 | 0.00362 | 7.1e-05 | 0.000862 | 75.1123 | 1.3545 | 2.5475 | 100.3 | 21.92 | 30.057 | 63.919 | 75.3868 | 61.1715 | 0.007895 | 1.056 | 0.025135 | 23.8019 | 0.021182 | 0.13624 | -5.9e-05 | -0.004424 | -0.004424 | -1.2533
2019-03-29 11:29 | 1553873330.8444 | -389 | 0.20934 | 0.000492 | 0 | 99.7904 | 99.4117 | 127.2921 | 20.962 | 0.00171 | 0.001651 | -151.8806 | 0.23005 | 8.4003 | 0.002202 | 0.003446 | 0.003171 | 7.1e-05 | 0.000866 | 75.1959 | 1.3967 | 2.5473 | 100.3 | 21.92 | 30.06 | 62.6826 | 74.5014 | 61.2716 | 0.007895 | 1.0557 | 0.02511 | 23.7855 | 0.034004 | 0.13686 | -4.5e-05 | -0.003385 | -0.003385 | -1.6456
2019-03-29 11:29 | 1553873331.8434 | -388 | 0.20934 | 0.00049 | 0 | 99.8739 | 99.5875 | 127.0022 | 20.9618 | 0.001801 | 0.00161 | -151.8761 | 0.22919 | 8.4029 | 0.002257 | 0.003736 | 0.003541 | 7.2e-05 | 0.000855 | 75.1223 | 1.3543 | 2.5474 | 100.3 | 21.91 | 30.065 | 64.8721 | 75.835 | 60.2646 | 0.007892 | 1.0589 | 0.025119 | 23.7226 | 0.025686 | 0.13634 | -4.3e-05 | -0.003268 | -0.003268 | -1.6982
2019-03-29 11:29 | 1553873332.8436 | -387 | 0.20933 | 0.000487 | 0 | 99.8263 | 99.6267 | 127.0346 | 20.962 | 0.001745 | 0.00167 | -151.8969 | 0.2298 | 8.4029 | 0.002215 | 0.003346 | 0.003194 | 7.1e-05 | 0.000847 | 75.1381 | 1.3664 | 2.5473 | 100.3 | 21.91 | 30.067 | 63.5167 | 74.8218 | 61.7116 | 0.007892 | 1.0543 | 0.025068 | 23.777 | 0.017898 | 0.13671 | -3.9e-05 | -0.002893 | -0.002893 | -1.9232
2019-03-29 11:29 | 1553873333.8436 | -386 | 0.20934 | 0.000478 | -1e-06 | 99.7826 | 99.4308 | 126.5882 | 20.9627 | 0.001817 | 0.001705 | -151.9158 | 0.22948 | 8.4033 | 0.002215 | 0.003378 | 0.003165 | 7.3e-05 | 0.000842 | 75.088 | 1.3986 | 2.5475 | 100.3 | 21.9 | 30.073 | 65.2617 | 74.8131 | 62.5568 | 0.007889 | 1.0516 | 0.024993 | 23.7651 | 0.031587 | 0.13652 | -2.8e-05 | -0.002047 | -0.002047 | -2.7138
2019-03-29 11:29 | 1553873334.8436 | -385 | 0.20935 | 0.000479 | -1e-06 | 99.8628 | 99.5667 | 126.3251 | 20.9627 | 0.001781 | 0.00165 | -151.9155 | 0.22925 | 8.4021 | 0.002185 | 0.003966 | 0.003743 | 6.9e-05 | 0.000836 | 75.1211 | 1.3712 | 2.5474 | 100.3 | 21.89 | 30.076 | 64.4013 | 74.1109 | 61.2292 | 0.007885 | 1.0558 | 0.025016 | 23.6934 | 0.026564 | 0.13638 | -2.8e-05 | -0.002099 | -0.002099 | -2.6438
2019-03-29 11:29 | 1553873335.8436 | -384 | 0.20935 | 0.000476 | -1e-06 | 99.8374 | 99.5947 | 125.8375 | 20.9627 | 0.001754 | 0.001669 | -151.9203 | 0.22951 | 8.401 | 0.002187 | 0.003322 | 0.003152 | 7.1e-05 | 0.000836 | 75.1225 | 1.3746 | 2.5475 | 100.3 | 21.89 | 30.078 | 63.7299 | 74.1411 | 61.6873 | 0.007886 | 1.0544 | 0.024951 | 23.6639 | 0.021771 | 0.13654 | -2.4e-05 | -0.001793 | -0.001793 | -3.0992
2019-03-29 11:29 | 1553873336.8433 | -383 | 0.20936 | 0.000471 | 0 | 99.7967 | 99.7322 | 125.6645 | 20.9639 | 0.001715 | 0.001647 | -151.9175 | 0.22941 | 8.4018 | 0.00227 | 0.003487 | 0.003324 | 7.4e-05 | 0.000832 | 75.1444 | 1.3967 | 2.5475 | 100.3 | 21.88 | 30.085 | 62.7954 | 76.1364 | 61.1585 | 0.007883 | 1.056 | 0.024953 | 23.6289 | 0.00578 | 0.13648 | -1.9e-05 | -0.001444 | -0.001444 | -3.8479
2019-03-29 11:29 | 1553873337.8437 | -382 | 0.20935 | 0.000475 | 0 | 99.7676 | 99.6363 | 125.8061 | 20.9636 | 0.00177 | 0.001618 | -151.9184 | 0.23003 | 8.4047 | 0.002225 | 0.00392 | 0.00374 | 6.9e-05 | 0.000836 | 75.1812 | 1.3218 | 2.5476 | 100.3 | 21.88 | 30.089 | 64.1392 | 75.0728 | 60.4774 | 0.007884 | 1.0582 | 0.024993 | 23.6181 | 0.011782 | 0.13684 | -2.4e-05 | -0.001796 | -0.001796 | -3.1006
2019-03-29 11:29 | 1553873338.8431 | -381 | 0.20936 | 0.00047 | 0 | 99.7551 | 99.5882 | 125.7044 | 20.9639 | 0.001776 | 0.00165 | -151.9375 | 0.22881 | 8.4019 | 0.002238 | 0.003667 | 0.003482 | 7.1e-05 | 0.000835 | 75.1856 | 1.325 | 2.5475 | 100.3 | 21.87 | 30.096 | 64.2745 | 75.388 | 61.2344 | 0.007881 | 1.0558 | 0.024954 | 23.6354 | 0.014976 | 0.13612 | -1.9e-05 | -0.001397 | -0.001397 | -3.9643
2019-03-29 11:29 | 1553873339.8438 | -380 | 0.20936 | 0.000472 | 0 | 99.7697 | 99.6578 | 125.8136 | 20.9634 | 0.001747 | 0.001698 | -151.9422 | 0.23 | 8.402 | 0.002262 | 0.004207 | 0.004005 | 7.3e-05 | 0.00083 | 75.1914 | 1.3046 | 2.5475 | 100.3 | 21.87 | 30.101 | 63.5744 | 75.9502 | 62.3943 | 0.007882 | 1.0522 | 0.024922 | 23.6866 | 0.01004 | 0.13683 | -2e-05 | -0.001504 | -0.001504 | -3.7027
2019-03-29 11:29 | 1553873340.8437 | -379 | 0.20939 | 0.000473 | -1e-06 | 99.8529 | 99.7332 | 126.0237 | 20.9641 | 0.001761 | 0.001689 | -151.9281 | 0.22915 | 8.4018 | 0.002276 | 0.003989 | 0.003843 | 7e-05 | 0.000827 | 75.1577 | 1.3362 | 2.5474 | 100.3 | 21.87 | 30.105 | 63.907 | 76.2815 | 62.1778 | 0.007883 | 1.0528 | 0.024951 | 23.6987 | 0.01073 | 0.13632 | -2.1e-05 | -0.001501 | -0.001501 | -3.6956
2019-03-29 11:29 | 1553873341.8442 | -378 | 0.20941 | 0.000467 | 0 | 99.8823 | 99.4889 | 126.28 | 20.9657 | 0.001798 | 0.00162 | -151.9289 | 0.22887 | 8.4016 | 0.002254 | 0.003848 | 0.003699 | 7e-05 | 0.000826 | 75.1349 | 1.381 | 2.5474 | 100.3 | 21.86 | 30.11 | 64.8149 | 75.7655 | 60.5192 | 0.00788 | 1.0581 | 0.025038 | 23.6641 | 0.0353 | 0.13616 | -1.4e-05 | -0.001051 | -0.001051 | -5.2729
2019-03-29 11:29 | 1553873342.8439 | -377 | 0.20939 | 0.000469 | 0 | 99.852 | 99.6769 | 126.418 | 20.9657 | 0.001793 | 0.001649 | -151.916 | 0.22921 | 8.4031 | 0.002237 | 0.003575 | 0.003329 | 6.8e-05 | 0.000827 | 75.1332 | 1.3561 | 2.5475 | 100.3 | 21.86 | 30.118 | 64.687 | 75.3516 | 61.2056 | 0.007882 | 1.0559 | 0.025026 | 23.7013 | 0.015702 | 0.13636 | -1.7e-05 | -0.001222 | -0.001222 | -4.541
2019-03-29 11:29 | 1553873343.8439 | -376 | 0.20942 | 0.000466 | 0 | 99.8904 | 99.6206 | 126.8637 | 20.9649 | 0.001771 | 0.001679 | -151.9083 | 0.22905 | 8.4013 | 0.002258 | 0.003705 | 0.00355 | 6.9e-05 | 0.000834 | 75.1665 | 1.3755 | 2.5475 | 100.3 | 21.85 | 30.123 | 64.1637 | 75.8596 | 61.9283 | 0.007878 | 1.0536 | 0.025043 | 23.7687 | 0.024188 | 0.13626 | -1.4e-05 | -0.001062 | -0.001062 | -5.2225
2019-03-29 11:29 | 1553873344.8429 | -375 | 0.20944 | 0.000466 | 0 | 99.8925 | 99.8247 | 127.371 | 20.9653 | 0.001803 | 0.001682 | -151.8964 | 0.2284 | 8.404 | 0.00218 | 0.003514 | 0.003307 | 6.9e-05 | 0.000833 | 75.1205 | 1.3927 | 2.5475 | 100.3 | 21.85 | 30.129 | 64.9347 | 73.9746 | 62 | 0.00788 | 1.0534 | 0.025091 | 23.8187 | 0.00607 | 0.13587 | -1.2e-05 | -0.000918 | -0.000918 | -6.0228
2019-03-29 11:29 | 1553873345.844 | -374 | 0.20944 | 0.000467 | -1e-06 | 99.92 | 99.682 | 127.7843 | 20.9654 | 0.001782 | 0.001628 | -151.8849 | 0.22871 | 8.401 | 0.002278 | 0.003638 | 0.00343 | 6.8e-05 | 0.000832 | 75.1765 | 1.3457 | 2.5474 | 100.3 | 21.84 | 30.142 | 64.4297 | 76.3366 | 60.7091 | 0.007878 | 1.0575 | 0.02518 | 23.8114 | 0.021327 | 0.13606 | -1.4e-05 | -0.001013 | -0.001013 | -5.4677
2019-03-29 11:29 | 1553873346.844 | -373 | 0.20944 | 0.000462 | 0 | 99.9134 | 99.5285 | 127.9543 | 20.9655 | 0.001832 | 0.001615 | -151.8874 | 0.22884 | 8.4024 | 0.002222 | 0.003822 | 0.003674 | 6.9e-05 | 0.000829 | 75.1402 | 1.3429 | 2.5474 | 100.3 | 21.84 | 30.148 | 65.6264 | 74.9881 | 60.3869 | 0.00788 | 1.0585 | 0.025209 | 23.8157 | 0.034529 | 0.13614 | -8e-06 | -0.000554 | -0.000554 | -10.0023
2019-03-29 11:29 | 1553873347.8434 | -372 | 0.20941 | 0.000469 | 0 | 99.9673 | 99.8629 | 128.0834 | 20.9655 | 0.001763 | 0.001664 | -151.8724 | 0.22863 | 8.4007 | 0.002203 | 0.003901 | 0.003669 | 6.6e-05 | 0.000839 | 75.1203 | 1.3402 | 2.5474 | 100.3 | 21.84 | 30.156 | 63.973 | 74.5492 | 61.586 | 0.007882 | 1.0547 | 0.025176 | 23.8705 | 0.009343 | 0.13601 | -1.6e-05 | -0.001183 | -0.001183 | -4.6776
2019-03-29 11:29 | 1553873348.843 | -371 | 0.20947 | 0.000468 | 0 | 99.9544 | 99.694 | 128.1841 | 20.9665 | 0.001761 | 0.001662 | -151.8585 | 0.22886 | 8.4028 | 0.002305 | 0.003584 | 0.00335 | 6.7e-05 | 0.00084 | 75.0929 | 1.3616 | 2.5474 | 100.3 | 21.83 | 30.17 | 63.917 | 76.9815 | 61.5221 | 0.007881 | 1.0549 | 0.025188 | 23.8776 | 0.023336 | 0.13615 | -1.5e-05 | -0.001169 | -0.001169 | -4.7406
2019-03-29 11:29 | 1553873349.8441 | -370 | 0.20944 | 0.000469 | 0 | 99.9553 | 99.6807 | 128.2601 | 20.966 | 0.001746 | 0.001638 | -151.8452 | 0.22847 | 8.4016 | 0.002318 | 0.004041 | 0.003926 | 6.5e-05 | 0.00084 | 75.1238 | 1.3361 | 2.5474 | 100.3 | 21.83 | 30.176 | 63.5528 | 77.2928 | 60.9426 | 0.007882 | 1.0567 | 0.025218 | 23.864 | 0.024611 | 0.13591 | -1.6e-05 | -0.001201 | -0.001201 | -4.6063
2019-03-29 11:29 | 1553873350.8441 | -369 | 0.20944 | 0.000469 | -1e-06 | 99.9885 | 99.7332 | 128.0463 | 20.9674 | 0.001776 | 0.001652 | -151.8396 | 0.22858 | 8.4041 | 0.00234 | 0.004032 | 0.003895 | 6.7e-05 | 0.000845 | 75.1499 | 1.3526 | 2.5474 | 100.3 | 21.82 | 30.188 | 64.281 | 77.8281 | 61.2908 | 0.007881 | 1.0556 | 0.025184 | 23.8565 | 0.022868 | 0.13598 | -1.7e-05 | -0.00122 | -0.00122 | -4.5347
2019-03-29 11:29 | 1553873351.8425 | -368 | 0.20943 | 0.000468 | 0 | 99.9849 | 99.8271 | 128.0536 | 20.9672 | 0.001735 | 0.001681 | -151.831 | 0.22897 | 8.4025 | 0.002386 | 0.003297 | 0.003063 | 6.6e-05 | 0.000844 | 75.1 | 1.3457 | 2.5474 | 100.3 | 21.82 | 30.193 | 63.2868 | 78.9387 | 61.9814 | 0.007882 | 1.0535 | 0.025158 | 23.8818 | 0.014134 | 0.13621 | -1.6e-05 | -0.001175 | -0.001175 | -4.72
2019-03-29 11:29 | 1553873352.8425 | -367 | 0.20944 | 0.000464 | -1e-06 | 100.014 | 99.8616 | 127.9097 | 20.9669 | 0.001822 | 0.00167 | -151.8251 | 0.22825 | 8.4013 | 0.002313 | 0.00339 | 0.003287 | 6.7e-05 | 0.000844 | 75.1415 | 1.3123 | 2.5473 | 100.3 | 21.81 | 30.2 | 65.3726 | 77.183 | 61.7194 | 0.007879 | 1.0543 | 0.025154 | 23.8591 | 0.013643 | 0.13579 | -1e-05 | -0.000678 | -0.000678 | -8.1554
2019-03-29 11:29 | 1553873353.8422 | -366 | 0.20945 | 0.000467 | 0 | 100.0957 | 99.8909 | 127.7285 | 20.9671 | 0.001736 | 0.001685 | -151.8384 | 0.22918 | 8.4001 | 0.002314 | 0.004005 | 0.003817 | 6.6e-05 | 0.000835 | 75.1594 | 1.3323 | 2.5475 | 100.3 | 21.81 | 30.214 | 63.3048 | 77.2033 | 62.0742 | 0.007883 | 1.0532 | 0.025123 | 23.8548 | 0.018319 | 0.13634 | -1.4e-05 | -0.001058 | -0.001058 | -5.2465
2019-03-29 11:29 | 1553873354.8422 | -365 | 0.20943 | 0.000467 | 0 | 100.0661 | 99.9233 | 127.5334 | 20.9672 | 0.001725 | 0.001684 | -151.8452 | 0.22833 | 8.4021 | 0.002262 | 0.003768 | 0.003575 | 6.6e-05 | 0.000837 | 75.1577 | 1.3155 | 2.5475 | 100.3 | 21.8 | 30.219 | 63.0541 | 75.9486 | 62.0597 | 0.007879 | 1.0532 | 0.025104 | 23.836 | 0.012768 | 0.13583 | -1.4e-05 | -0.001027 | -0.001027 | -5.3819
2019-03-29 11:29 | 1553873355.8439 | -364 | 0.20943 | 0.000472 | 0 | 100.0451 | 99.9875 | 127.7315 | 20.9669 | 0.001835 | 0.001682 | -151.8523 | 0.22848 | 8.4033 | 0.002295 | 0.003565 | 0.003349 | 6.9e-05 | 0.000849 | 75.1779 | 1.3315 | 2.5474 | 100.3 | 21.8 | 30.225 | 65.691 | 76.7488 | 62.0218 | 0.007881 | 1.0533 | 0.025125 | 23.8532 | 0.005154 | 0.13593 | -2e-05 | -0.001562 | -0.001562 | -3.5428
2019-03-29 11:29 | 1553873356.8432 | -363 | 0.20945 | 0.000471 | 0 | 100.0499 | 99.865 | 127.8073 | 20.9671 | 0.001802 | 0.001629 | -151.8467 | 0.22882 | 8.4006 | 0.002402 | 0.004051 | 0.003905 | 6.7e-05 | 0.000848 | 75.132 | 1.3489 | 2.5475 | 100.3 | 21.8 | 30.236 | 64.9068 | 79.3255 | 60.7316 | 0.007884 | 1.0574 | 0.025181 | 23.8143 | 0.016549 | 0.13613 | -1.8e-05 | -0.001391 | -0.001391 | -3.9844
2019-03-29 11:29 | 1553873357.8433 | -362 | 0.20944 | 0.000469 | 0 | 100.1831 | 100.0885 | 127.6907 | 20.9676 | 0.001749 | 0.0017 | -151.8631 | 0.22848 | 8.4028 | 0.002289 | 0.00332 | 0.003186 | 6.7e-05 | 0.000848 | 75.1422 | 1.3362 | 2.5474 | 100.3 | 21.79 | 30.241 | 63.6362 | 76.6138 | 62.4506 | 0.00788 | 1.052 | 0.025105 | 23.8646 | 0.00845 | 0.13592 | -1.7e-05 | -0.001244 | -0.001244 | -4.4479
2019-03-29 11:29 | 1553873358.8436 | -361 | 0.20942 | 0.000469 | -1e-06 | 100.1705 | 100.1086 | 127.5484 | 20.9689 | 0.001759 | 0.001632 | -151.8721 | 0.2288 | 8.4025 | 0.002325 | 0.003835 | 0.003683 | 6.6e-05 | 0.000846 | 75.0834 | 1.3899 | 2.5476 | 100.3 | 21.79 | 30.251 | 63.878 | 77.4625 | 60.8069 | 0.007883 | 1.0572 | 0.025153 | 23.7929 | 0.005528 | 0.13612 | -1.6e-05 | -0.001148 | -0.001148 | -4.8243
2019-03-29 11:29 | 1553873359.8651 | -360 | 0.20945 | 0.000467 | 0 | 100.184 | 100.1421 | 127.2867 | 20.9677 | 0.001796 | 0.001656 | -151.8819 | 0.22884 | 8.4043 | 0.00231 | 0.003689 | 0.003473 | 6.8e-05 | 0.000848 | 75.104 | 1.3538 | 2.5474 | 100.3 | 21.79 | 30.255 | 64.7567 | 77.1222 | 61.3928 | 0.007884 | 1.0553 | 0.025105 | 23.7893 | 0.003743 | 0.13614 | -1.4e-05 | -0.001035 | -0.001035 | -5.3533
2019-03-29 11:29 | 1553873360.8573 | -359 | 0.20944 | 0.000467 | 0 | 100.1888 | 100.2134 | 127.4766 | 20.9685 | 0.001787 | 0.001621 | -151.8926 | 0.22846 | 8.4004 | 0.002316 | 0.003611 | 0.003459 | 6.9e-05 | 0.000845 | 75.1008 | 1.3278 | 2.5473 | 100.3 | 21.78 | 30.265 | 64.5422 | 77.2609 | 60.5354 | 0.007881 | 1.058 | 0.025156 | 23.7765 | -0.002198 | 0.13591 | -1.4e-05 | -0.0011 | -0.0011 | -5.0268
2019-03-29 11:29 | 1553873361.8574 | -358 | 0.20945 | 0.000467 | 0 | 100.254 | 100.1015 | 127.1867 | 20.9685 | 0.00178 | 0.001663 | -151.9119 | 0.22912 | 8.401 | 0.002291 | 0.004083 | 0.003955 | 6.9e-05 | 0.000847 | 75.1024 | 1.2995 | 2.5474 | 100.3 | 21.78 | 30.271 | 64.3718 | 76.6567 | 61.5631 | 0.007883 | 1.0548 | 0.025089 | 23.786 | 0.013622 | 0.1363 | -1.5e-05 | -0.001089 | -0.001089 | -5.0944
2019-03-29 11:29 | 1553873362.8573 | -357 | 0.20948 | 0.000466 | 0 | 100.2576 | 100.1312 | 127.1142 | 20.9688 | 0.001786 | 0.001675 | -151.917 | 0.228 | 8.4013 | 0.002216 | 0.00356 | 0.003424 | 6.8e-05 | 0.000845 | 75.147 | 1.3254 | 2.5474 | 100.3 | 21.78 | 30.276 | 64.5157 | 74.8548 | 61.8454 | 0.007884 | 1.0539 | 0.025071 | 23.7892 | 0.01129 | 0.13564 | -1.3e-05 | -0.000967 | -0.000967 | -5.709
2019-03-29 11:29 | 1553873363.8578 | -356 | 0.20944 | 0.000471 | 0 | 100.3142 | 100.3602 | 127.2142 | 20.969 | 0.001826 | 0.001595 | -151.9079 | 0.22878 | 8.4042 | 0.002229 | 0.003834 | 0.00358 | 7e-05 | 0.000845 | 75.092 | 1.3191 | 2.5474 | 100.3 | 21.78 | 30.283 | 65.4817 | 75.1621 | 59.9151 | 0.007886 | 1.06 | 0.025153 | 23.7299 | -0.004097 | 0.1361 | -1.9e-05 | -0.001392 | -0.001392 | -3.9783
2019-03-29 11:29 | 1553873364.8565 | -355 | 0.20943 | 0.000469 | 0 | 100.3636 | 100.1885 | 127.2304 | 20.9697 | 0.001756 | 0.001699 | -151.9107 | 0.22864 | 8.4032 | 0.00232 | 0.003624 | 0.003446 | 6.7e-05 | 0.000845 | 75.082 | 1.3307 | 2.5474 | 100.3 | 21.77 | 30.285 | 63.8062 | 77.3625 | 62.4244 | 0.007882 | 1.0521 | 0.025061 | 23.8207 | 0.015619 | 0.13602 | -1.7e-05 | -0.001245 | -0.001245 | -4.4466
2019-03-29 11:29 | 1553873365.8565 | -354 | 0.20952 | 0.000467 | 0 | 100.3549 | 100.268 | 126.8602 | 20.9691 | 0.001793 | 0.001687 | -151.9219 | 0.2291 | 8.4032 | 0.002349 | 0.003976 | 0.003795 | 6.5e-05 | 0.000844 | 75.1399 | 1.3062 | 2.5475 | 100.3 | 21.77 | 30.288 | 64.6917 | 78.052 | 62.1436 | 0.007883 | 1.0529 | 0.025035 | 23.776 | 0.007748 | 0.13629 | -1.4e-05 | -0.001053 | -0.001053 | -5.2664
2019-03-29 11:29 | 1553873366.8559 | -353 | 0.20934 | 0.000468 | 0 | 100.3588 | 100.3271 | 126.8088 | 20.9701 | 0.00179 | 0.001676 | -151.92 | 0.22865 | 8.4018 | 0.002254 | 0.003227 | 0.003123 | 6.6e-05 | 0.000851 | 75.0794 | 1.3662 | 2.5475 | 100.3 | 21.77 | 30.292 | 64.6193 | 75.7666 | 61.8804 | 0.007884 | 1.0538 | 0.02504 | 23.7619 | 0.00283 | 0.13602 | -1.5e-05 | -0.001087 | -0.001087 | -5.0934
2019-03-29 11:29 | 1553873367.8556 | -352 | 0.20944 | 0.000465 | -1e-06 | 100.4229 | 100.065 | 126.7778 | 20.97 | 0.001857 | 0.001702 | -151.9053 | 0.22847 | 8.4032 | 0.002288 | 0.004437 | 0.004278 | 6.7e-05 | 0.000852 | 75.1139 | 1.3667 | 2.5475 | 100.3 | 21.76 | 30.293 | 66.2201 | 76.574 | 62.496 | 0.007879 | 1.0518 | 0.025014 | 23.7808 | 0.031935 | 0.13592 | -1.1e-05 | -0.000757 | -0.000757 | -7.304
2019-03-29 11:29 | 1553873368.8566 | -351 | 0.20943 | 0.000461 | -1e-06 | 100.3693 | 100.266 | 126.7804 | 20.9694 | 0.001793 | 0.001683 | -151.9048 | 0.22852 | 8.4017 | 0.002319 | 0.004182 | 0.003927 | 6.9e-05 | 0.000857 | 75.1096 | 1.3213 | 2.5473 | 100.3 | 21.76 | 30.296 | 64.6888 | 77.3319 | 62.05 | 0.00788 | 1.0532 | 0.02503 | 23.7652 | 0.009211 | 0.13595 | -6e-06 | -0.000408 | -0.000408 | -13.5675
2019-03-29 11:29 | 1553873369.8556 | -350 | 0.20946 | 0.000465 | 0 | 100.334 | 100.2008 | 127.0259 | 20.9704 | 0.001771 | 0.001668 | -151.9039 | 0.2292 | 8.4028 | 0.002281 | 0.004143 | 0.004043 | 6.8e-05 | 0.000854 | 75.0538 | 1.3364 | 2.5474 | 100.3 | 21.76 | 30.296 | 64.1622 | 76.4154 | 61.6893 | 0.00788 | 1.0544 | 0.025068 | 23.7754 | 0.01188 | 0.13635 | -1.1e-05 | -0.000843 | -0.000843 | -6.5799
2019-03-29 11:29 | 1553873370.8567 | -349 | 0.20938 | 0.000466 | 0 | 100.4058 | 100.3455 | 126.8438 | 20.9705 | 0.001751 | 0.001615 | -151.8971 | 0.22839 | 8.4039 | 0.002391 | 0.003862 | 0.003639 | 6.9e-05 | 0.000856 | 75.0905 | 1.3821 | 2.5475 | 100.3 | 21.76 | 30.296 | 63.6811 | 79.0512 | 60.406 | 0.00788 | 1.0584 | 0.025098 | 23.7128 | 0.005374 | 0.13587 | -1.2e-05 | -0.000889 | -0.000889 | -6.2185
2019-03-29 11:29 | 1553873371.8561 | -348 | 0.20947 | 0.00046 | 0 | 100.4477 | 100.3909 | 126.9123 | 20.9708 | 0.001779 | 0.001675 | -151.8868 | 0.22861 | 8.3996 | 0.002367 | 0.003541 | 0.003366 | 6.7e-05 | 0.000863 | 75.0272 | 1.3391 | 2.5475 | 100.3 | 21.75 | 30.296 | 64.3577 | 78.4753 | 61.8428 | 0.007875 | 1.0539 | 0.025051 | 23.7702 | 0.005061 | 0.136 | -5e-06 | -0.000347 | -0.000347 | -15.9378
2019-03-29 11:30 | 1553873372.8567 | -347 | 0.20945 | 0.000469 | 0 | 100.46 | 100.3687 | 127.0973 | 20.9718 | 0.001709 | 0.001651 | -151.8891 | 0.22841 | 8.4027 | 0.002393 | 0.003921 | 0.003824 | 6.8e-05 | 0.000867 | 75.0389 | 1.3676 | 2.5475 | 100.3 | 21.75 | 30.295 | 62.6711 | 79.0991 | 61.2686 | 0.007875 | 1.0557 | 0.025091 | 23.7672 | 0.008131 | 0.13588 | -1.7e-05 | -0.001259 | -0.001259 | -4.3916
2019-03-29 11:30 | 1553873373.8557 | -346 | 0.20948 | 0.000465 | 0 | 100.5148 | 100.283 | 126.8656 | 20.9719 | 0.001733 | 0.001644 | -151.895 | 0.2291 | 8.4011 | 0.002318 | 0.003906 | 0.003762 | 6.9e-05 | 0.000871 | 75.1044 | 1.3116 | 2.5474 | 100.3 | 21.74 | 30.292 | 63.243 | 77.3139 | 61.1022 | 0.007869 | 1.0562 | 0.025074 | 23.7396 | 0.020644 | 0.13629 | -1.1e-05 | -0.000814 | -0.000814 | -6.8143
2019-03-29 11:30 | 1553873374.8562 | -345 | 0.20944 | 0.000466 | 0 | 100.4564 | 100.597 | 127.0945 | 20.9716 | 0.001772 | 0.001625 | -151.8901 | 0.22895 | 8.4042 | 0.00239 | 0.003738 | 0.003522 | 6.5e-05 | 0.000873 | 75.0405 | 1.311 | 2.5474 | 100.3 | 21.74 | 30.291 | 64.1728 | 79.0358 | 60.6447 | 0.007869 | 1.0577 | 0.025114 | 23.7447 | -0.012513 | 0.1362 | -1.2e-05 | -0.000907 | -0.000907 | -6.1146
2019-03-29 11:30 | 1553873375.8568 | -344 | 0.20944 | 0.000462 | 0 | 100.469 | 100.2882 | 127.3391 | 20.9718 | 0.001787 | 0.001676 | -151.8803 | 0.22953 | 8.4031 | 0.002372 | 0.003972 | 0.00376 | 6.8e-05 | 0.000876 | 75.066 | 1.3187 | 2.5475 | 100.3 | 21.74 | 30.29 | 64.5445 | 78.6089 | 61.8609 | 0.007869 | 1.0538 | 0.025093 | 23.8108 | 0.016111 | 0.13654 | -8e-06 | -0.000624 | -0.000624 | -8.9118
2019-03-29 11:30 | 1553873376.8563 | -343 | 0.20945 | 0.000463 | 0 | 100.574 | 100.5049 | 127.4276 | 20.973 | 0.001725 | 0.001724 | -151.8794 | 0.2284 | 8.4021 | 0.002315 | 0.003223 | 0.003033 | 6.8e-05 | 0.000878 | 75.0649 | 1.3585 | 2.5474 | 100.3 | 21.74 | 30.288 | 63.0652 | 77.2427 | 63.0385 | 0.007868 | 1.0501 | 0.025057 | 23.8609 | 0.006152 | 0.13587 | -9e-06 | -0.000708 | -0.000708 | -7.8109
2019-03-29 11:30 | 1553873377.8563 | -342 | 0.20946 | 0.000467 | 0 | 100.5372 | 100.4424 | 127.4104 | 20.9723 | 0.001779 | 0.001648 | -151.8909 | 0.22899 | 8.4033 | 0.002343 | 0.003517 | 0.003319 | 6.6e-05 | 0.000878 | 75.076 | 1.3929 | 2.5475 | 100.3 | 21.74 | 30.288 | 64.3521 | 77.9068 | 61.1949 | 0.007868 | 1.0559 | 0.025125 | 23.7938 | 0.008437 | 0.13623 | -1.4e-05 | -0.001031 | -0.001031 | -5.3781
2019-03-29 11:30 | 1553873378.857 | -341 | 0.20945 | 0.000465 | 0 | 100.5552 | 100.4684 | 127.5386 | 20.9733 | 0.001772 | 0.001685 | -151.8766 | 0.22878 | 8.403 | 0.002295 | 0.003387 | 0.003209 | 7e-05 | 0.000889 | 75.0995 | 1.3955 | 2.5476 | 100.3 | 21.73 | 30.287 | 64.1956 | 76.7621 | 62.0883 | 0.007863 | 1.0531 | 0.025104 | 23.8375 | 0.007726 | 0.1361 | -1.1e-05 | -0.000797 | -0.000797 | -6.9513
2019-03-29 11:30 | 1553873379.8559 | -340 | 0.20947 | 0.000462 | -1e-06 | 100.5734 | 100.5049 | 127.4827 | 20.9737 | 0.001776 | 0.001676 | -151.8698 | 0.22966 | 8.4033 | 0.002358 | 0.003215 | 0.003019 | 6.6e-05 | 0.000893 | 75.0566 | 1.3314 | 2.5473 | 100.3 | 21.73 | 30.285 | 64.2871 | 78.2785 | 61.8832 | 0.007863 | 1.0538 | 0.025106 | 23.825 | 0.006099 | 0.13663 | -7e-06 | -0.00047 | -0.00047 | -11.8293
2019-03-29 11:30 | 1553873380.8573 | -339 | 0.20947 | 0.000461 | -1e-06 | 100.6109 | 100.4025 | 127.4998 | 20.9747 | 0.00177 | 0.001622 | -151.8675 | 0.22836 | 8.4027 | 0.002336 | 0.003872 | 0.003756 | 6.6e-05 | 0.000892 | 75.0742 | 1.3332 | 2.5474 | 100.3 | 21.73 | 30.285 | 64.1406 | 77.7518 | 60.5641 | 0.007863 | 1.0579 | 0.025157 | 23.7797 | 0.018543 | 0.13585 | -5e-06 | -0.00035 | -0.00035 | -15.8012
2019-03-29 11:30 | 1553873381.857 | -338 | 0.20946 | 0.000465 | -1e-06 | 100.6219 | 100.3711 | 127.2452 | 20.9739 | 0.001776 | 0.001672 | -151.8642 | 0.22906 | 8.4025 | 0.002357 | 0.003731 | 0.003459 | 6.7e-05 | 0.000891 | 75.0464 | 1.3145 | 2.5474 | 100.3 | 21.72 | 30.284 | 64.2709 | 78.257 | 61.781 | 0.007857 | 1.0541 | 0.025086 | 23.7992 | 0.022325 | 0.13627 | -1.1e-05 | -0.000743 | -0.000743 | -7.4612
2019-03-29 11:30 | 1553873382.856 | -337 | 0.20947 | 0.000468 | 0 | 100.6566 | 100.6608 | 127.14 | 20.9736 | 0.001778 | 0.001653 | -151.8797 | 0.22819 | 8.4035 | 0.002259 | 0.00354 | 0.003359 | 6.8e-05 | 0.0009 | 75.0801 | 1.3356 | 2.5475 | 100.3 | 21.72 | 30.285 | 64.3333 | 75.901 | 61.3188 | 0.007858 | 1.0555 | 0.025093 | 23.7729 | -0.000373 | 0.13575 | -1.5e-05 | -0.001086 | -0.001086 | -5.0867
2019-03-29 11:30 | 1553873383.8551 | -336 | 0.20946 | 0.000462 | 0 | 100.6558 | 100.6342 | 127.2445 | 20.9746 | 0.00178 | 0.001625 | -151.8915 | 0.2286 | 8.4031 | 0.002297 | 0.003428 | 0.003287 | 6.6e-05 | 0.000894 | 75.1235 | 1.3544 | 2.5477 | 100.3 | 21.72 | 30.288 | 64.3896 | 76.8046 | 60.6351 | 0.007859 | 1.0577 | 0.025129 | 23.7584 | 0.001913 | 0.136 | -7e-06 | -0.000469 | -0.000469 | -11.791
2019-03-29 11:30 | 1553873384.8555 | -335 | 0.20946 | 0.000466 | -1e-06 | 100.6875 | 100.5489 | 126.6708 | 20.9746 | 0.001799 | 0.001626 | -151.906 | 0.2282 | 8.4018 | 0.002282 | 0.00353 | 0.003362 | 6.1e-05 | 0.000895 | 75.1277 | 1.3346 | 2.5475 | 100.3 | 21.72 | 30.294 | 64.8407 | 76.4358 | 60.6741 | 0.00786 | 1.0576 | 0.025071 | 23.7061 | 0.012318 | 0.13575 | -1.2e-05 | -0.000844 | -0.000844 | -6.5497
2019-03-29 11:30 | 1553873385.8566 | -334 | 0.20947 | 0.000463 | -1e-06 | 100.7096 | 100.6731 | 126.9736 | 20.9756 | 0.001803 | 0.001673 | -151.9181 | 0.2293 | 8.4035 | 0.002301 | 0.004095 | 0.003846 | 6.4e-05 | 0.000891 | 75.0864 | 1.3589 | 2.5474 | 100.3 | 21.71 | 30.297 | 64.9261 | 76.9146 | 61.8161 | 0.007856 | 1.054 | 0.025058 | 23.775 | 0.003242 | 0.13641 | -7e-06 | -0.000519 | -0.000519 | -10.7012
2019-03-29 11:30 | 1553873386.8552 | -333 | 0.20946 | 0.000459 | 0 | 100.745 | 100.5902 | 126.9471 | 20.9756 | 0.001783 | 0.001716 | -151.9123 | 0.2299 | 8.4008 | 0.002327 | 0.003885 | 0.003778 | 6.6e-05 | 0.000882 | 75.1218 | 1.3722 | 2.5477 | 100.3 | 21.71 | 30.305 | 64.4506 | 77.5258 | 62.8482 | 0.007858 | 1.0507 | 0.025017 | 23.8091 | 0.01375 | 0.13677 | -3e-06 | -0.000208 | -0.000208 | -26.8012
2019-03-29 11:30 | 1553873387.8558 | -332 | 0.20945 | 0.000465 | 0 | 100.7494 | 100.7127 | 127.2632 | 20.9755 | 0.001766 | 0.00173 | -151.9002 | 0.23011 | 8.4019 | 0.002301 | 0.003566 | 0.003228 | 6.4e-05 | 0.000885 | 75.1418 | 1.3559 | 2.5474 | 100.3 | 21.7 | 30.309 | 64.0437 | 76.9114 | 63.1675 | 0.007854 | 1.0497 | 0.025036 | 23.8501 | 0.00326 | 0.13689 | -1e-05 | -0.000739 | -0.000739 | -7.5417
2019-03-29 11:30 | 1553873388.8553 | -331 | 0.20951 | 0.000463 | 0 | 100.6899 | 100.751 | 127.3815 | 20.9758 | 0.001809 | 0.001644 | -151.9069 | 0.22951 | 8.4039 | 0.002311 | 0.003934 | 0.003841 | 6.6e-05 | 0.000882 | 75.1297 | 1.3969 | 2.5475 | 100.3 | 21.7 | 30.314 | 65.0713 | 77.1392 | 61.1075 | 0.007856 | 1.0562 | 0.025125 | 23.788 | -0.005424 | 0.13653 | -8e-06 | -0.00061 | -0.00061 | -9.1064
2019-03-29 11:30 | 1553873389.8552 | -330 | 0.20946 | 0.000464 | 0 | 100.8279 | 100.6687 | 127.6834 | 20.9769 | 0.001827 | 0.001661 | -151.8998 | 0.22899 | 8.4012 | 0.002292 | 0.003265 | 0.003108 | 6.1e-05 | 0.00088 | 75.1072 | 1.3128 | 2.5474 | 100.3 | 21.7 | 30.324 | 65.5064 | 76.6891 | 61.5159 | 0.007858 | 1.0549 | 0.025139 | 23.8307 | 0.014131 | 0.13622 | -1e-05 | -0.000709 | -0.000709 | -7.8189
2019-03-29 11:30 | 1553873390.8553 | -329 | 0.20949 | 0.000465 | -1e-06 | 100.7824 | 100.6236 | 127.7947 | 20.9757 | 0.001753 | 0.001657 | -151.8961 | 0.22931 | 8.4018 | 0.002344 | 0.004213 | 0.003938 | 6.6e-05 | 0.000881 | 75.0647 | 1.336 | 2.5473 | 100.3 | 21.69 | 30.331 | 63.7213 | 77.943 | 61.4114 | 0.007855 | 1.0552 | 0.025154 | 23.8374 | 0.014098 | 0.13642 | -1e-05 | -0.000723 | -0.000723 | -7.6757
2019-03-29 11:30 | 1553873391.8554 | -328 | 0.20949 | 0.000463 | -1e-06 | 100.7647 | 100.7861 | 127.9259 | 20.9765 | 0.001759 | 0.001695 | -151.8967 | 0.22831 | 8.4004 | 0.002277 | 0.00363 | 0.003522 | 6.5e-05 | 0.00088 | 75.0898 | 1.2931 | 2.5474 | 100.3 | 21.69 | 30.344 | 63.8777 | 76.3279 | 62.3319 | 0.007859 | 1.0524 | 0.025133 | 23.8824 | -0.001899 | 0.13583 | -8e-06 | -0.000561 | -0.000561 | -9.862
2019-03-29 11:30 | 1553873392.8554 | -327 | 0.20948 | 0.00046 | -1e-06 | 100.8099 | 100.7045 | 127.8806 | 20.9763 | 0.001767 | 0.001671 | -151.8756 | 0.22892 | 8.4026 | 0.002266 | 0.003703 | 0.003503 | 6.4e-05 | 0.000873 | 75.1522 | 1.3226 | 2.5474 | 100.3 | 21.69 | 30.351 | 64.0649 | 76.0617 | 61.7552 | 0.00786 | 1.0542 | 0.02515 | 23.8576 | 0.009354 | 0.13619 | -5e-06 | -0.000311 | -0.000311 | -17.8444
2019-03-29 11:30 | 1553873393.8554 | -326 | 0.20946 | 0.000464 | -2e-06 | 100.8425 | 100.8509 | 127.6259 | 20.9773 | 0.001766 | 0.001698 | -151.8556 | 0.22862 | 8.4036 | 0.00229 | 0.003898 | 0.003713 | 6.4e-05 | 0.000872 | 75.1482 | 1.3477 | 2.5472 | 100.3 | 21.68 | 30.358 | 64.0565 | 76.6381 | 62.4116 | 0.007857 | 1.0521 | 0.0251 | 23.8572 | -0.000746 | 0.13601 | -8e-06 | -0.000532 | -0.000532 | -10.4114
2019-03-29 11:30 | 1553873394.8565 | -325 | 0.20947 | 0.000459 | -1e-06 | 100.8084 | 100.7793 | 127.7109 | 20.9771 | 0.001805 | 0.001717 | -151.8304 | 0.22926 | 8.4003 | 0.002277 | 0.003114 | 0.00289 | 6.5e-05 | 0.000873 | 75.1212 | 1.3151 | 2.5475 | 100.3 | 21.68 | 30.373 | 64.9923 | 76.3214 | 62.8583 | 0.007861 | 1.0507 | 0.025092 | 23.881 | 0.002585 | 0.13639 | -3e-06 | -0.000173 | -0.000173 | -32.0619
2019-03-29 11:30 | 1553873395.8559 | -324 | 0.20948 | 0.00046 | -1e-06 | 100.8279 | 100.8414 | 127.352 | 20.9767 | 0.001765 | 0.001661 | -151.8062 | 0.22889 | 8.4027 | 0.002309 | 0.003709 | 0.003529 | 6.7e-05 | 0.000876 | 75.1413 | 1.2799 | 2.5474 | 100.3 | 21.68 | 30.38 | 64.0277 | 77.0973 | 61.5265 | 0.007863 | 1.0549 | 0.025106 | 23.8001 | -0.0012 | 0.13617 | -4e-06 | -0.000301 | -0.000301 | -18.3894
2019-03-29 11:30 | 1553873396.8559 | -323 | 0.20947 | 0.000457 | 0 | 100.8437 | 100.7404 | 127.4057 | 20.9776 | 0.001802 | 0.001615 | -151.8191 | 0.22826 | 8.4026 | 0.002355 | 0.003727 | 0.003512 | 6.4e-05 | 0.000884 | 75.1522 | 1.3295 | 2.5474 | 100.3 | 21.68 | 30.386 | 64.9013 | 78.212 | 60.4199 | 0.007865 | 1.0584 | 0.025153 | 23.7658 | 0.009173 | 0.13579 | -1e-06 | -5.9e-05 | -5.9e-05 | -94.2815
2019-03-29 11:30 | 1553873397.8555 | -322 | 0.20947 | 0.000459 | 0 | 100.8348 | 100.7653 | 127.2344 | 20.9778 | 0.001757 | 0.001664 | -151.8369 | 0.22897 | 8.402 | 0.002234 | 0.003663 | 0.003543 | 6.5e-05 | 0.000889 | 75.1157 | 1.3154 | 2.5474 | 100.3 | 21.68 | 30.399 | 63.8343 | 75.2862 | 61.5893 | 0.007868 | 1.0547 | 0.025092 | 23.7914 | 0.006164 | 0.13622 | -3e-06 | -0.000181 | -0.000181 | -30.7131
2019-03-29 11:30 | 1553873398.8559 | -321 | 0.20944 | 0.00046 | 0 | 100.7839 | 101.0011 | 127.1255 | 20.9777 | 0.001827 | 0.00172 | -151.8513 | 0.22941 | 8.4029 | 0.002273 | 0.003519 | 0.003307 | 6.6e-05 | 0.000894 | 75.1257 | 1.3304 | 2.5474 | 100.3 | 21.68 | 30.405 | 65.5072 | 76.2276 | 62.9323 | 0.00787 | 1.0505 | 0.025032 | 23.8288 | -0.01926 | 0.13648 | -4e-06 | -0.000307 | -0.000307 | -18.0756
2019-03-29 11:30 | 1553873399.855 | -320 | 0.20947 | 0.000457 | 0 | 100.8907 | 100.8018 | 126.9224 | 20.978 | 0.001766 | 0.001717 | -151.8502 | 0.22821 | 8.4027 | 0.002307 | 0.00426 | 0.004221 | 6.5e-05 | 0.000899 | 75.1577 | 1.3405 | 2.5475 | 100.3 | 21.67 | 30.417 | 64.0493 | 77.0566 | 62.867 | 0.007868 | 1.0507 | 0.025014 | 23.8075 | 0.007887 | 0.13576 | 0 | -1e-06 | -1e-06 | -4296.0576
2019-03-29 11:30 | 1553873400.8556 | -319 | 0.20953 | 0.000457 | 0 | 100.806 | 100.8322 | 127.0008 | 20.9783 | 0.001811 | 0.001662 | -151.8646 | 0.22913 | 8.4009 | 0.002319 | 0.003328 | 0.00316 | 6.5e-05 | 0.000902 | 75.2239 | 1.3669 | 2.5475 | 100.3 | 21.67 | 30.421 | 65.1207 | 77.3288 | 61.5423 | 0.007869 | 1.0548 | 0.025071 | 23.7679 | -0.002321 | 0.13631 | 0 | 1.8e-05 | 1.8e-05 | 305.6301
2019-03-29 11:30 | 1553873401.8557 | -318 | 0.20944 | 0.000458 | -1e-06 | 100.8069 | 100.8066 | 126.9231 | 20.9779 | 0.001758 | 0.00165 | -151.8783 | 0.22834 | 8.4026 | 0.002248 | 0.003694 | 0.003556 | 6.5e-05 | 0.000898 | 75.0394 | 1.3458 | 2.5473 | 100.3 | 21.67 | 30.424 | 63.8626 | 75.6322 | 61.2613 | 0.00787 | 1.0557 | 0.025074 | 23.7506 | 3e-05 | 0.13584 | 0 | 4.7e-05 | 4.7e-05 | 116.5002
2019-03-29 11:30 | 1553873402.8552 | -317 | 0.20953 | 0.000459 | 0 | 100.8108 | 100.6544 | 126.8027 | 20.9796 | 0.001804 | 0.001653 | -151.8798 | 0.22949 | 8.4014 | 0.002285 | 0.003578 | 0.003421 | 6.5e-05 | 0.000898 | 75.1385 | 1.3735 | 2.5474 | 100.3 | 21.67 | 30.428 | 64.954 | 76.5288 | 61.3333 | 0.007871 | 1.0555 | 0.025059 | 23.7419 | 0.013891 | 0.13652 | -4e-06 | -0.000281 | -0.000281 | -19.7486
2019-03-29 11:30 | 1553873403.8551 | -316 | 0.20947 | 0.000456 | -1e-06 | 100.8273 | 100.7909 | 126.9051 | 20.9792 | 0.001768 | 0.001646 | -151.8797 | 0.23018 | 8.4019 | 0.002259 | 0.003169 | 0.002933 | 6.7e-05 | 0.000892 | 75.1342 | 1.3437 | 2.5474 | 100.3 | 21.67 | 30.428 | 64.1067 | 75.9002 | 61.1508 | 0.007871 | 1.0561 | 0.025076 | 23.745 | 0.003228 | 0.13694 | 1e-06 | 0.000152 | 0.000152 | 36.7635
2019-03-29 11:30 | 1553873404.8561 | -315 | 0.20947 | 0.000457 | -1e-06 | 100.885 | 100.6308 | 127.0327 | 20.9795 | 0.00171 | 0.00168 | -151.8986 | 0.22889 | 8.4023 | 0.002257 | 0.003508 | 0.003364 | 6.7e-05 | 0.000894 | 75.1605 | 1.3556 | 2.5475 | 100.3 | 21.66 | 30.428 | 62.7106 | 75.8428 | 61.9675 | 0.007866 | 1.0535 | 0.025058 | 23.7859 | 0.022568 | 0.13617 | 0 | 4.5e-05 | 4.5e-05 | 122.9429
2019-03-29 11:30 | 1553873405.8548 | -314 | 0.20946 | 0.000456 | -1e-06 | 100.9009 | 100.7294 | 126.8377 | 20.9801 | 0.001766 | 0.001662 | -151.9043 | 0.22809 | 8.4044 | 0.002291 | 0.003444 | 0.003268 | 6.6e-05 | 0.000884 | 75.0651 | 1.3668 | 2.5474 | 100.3 | 21.66 | 30.426 | 64.0481 | 76.6672 | 61.543 | 0.007865 | 1.0548 | 0.025055 | 23.7526 | 0.015212 | 0.13569 | 2e-06 | 0.000209 | 0.000209 | 26.4327
2019-03-29 11:30 | 1553873406.8552 | -313 | 0.20946 | 0.000459 | -1e-06 | 100.88 | 100.6639 | 126.706 | 20.9797 | 0.001818 | 0.001649 | -151.9136 | 0.2288 | 8.4038 | 0.0023 | 0.003209 | 0.002955 | 6.5e-05 | 0.00088 | 75.1117 | 1.3499 | 2.5473 | 100.3 | 21.66 | 30.424 | 65.3009 | 76.8805 | 61.2217 | 0.007865 | 1.0558 | 0.025054 | 23.7289 | 0.019175 | 0.13611 | -2e-06 | -7.8e-05 | -7.8e-05 | -71.2772
2019-03-29 11:30 | 1553873407.857 | -312 | 0.20947 | 0.000458 | -1e-06 | 100.8503 | 100.8895 | 126.8253 | 20.9811 | 0.001813 | 0.001661 | -151.8933 | 0.22937 | 8.4035 | 0.002292 | 0.003467 | 0.003327 | 6.7e-05 | 0.000874 | 75.1071 | 1.3611 | 2.5475 | 100.3 | 21.65 | 30.419 | 65.1843 | 76.6963 | 61.5231 | 0.007859 | 1.0549 | 0.025055 | 23.7507 | -0.003476 | 0.13645 | 0 | 2.8e-05 | 2.8e-05 | 199.187
2019-03-29 11:30 | 1553873408.8559 | -311 | 0.20946 | 0.000457 | -1e-06 | 100.8351 | 100.7612 | 127.26 | 20.9812 | 0.001774 | 0.001703 | -151.8841 | 0.2291 | 8.4026 | 0.002294 | 0.003478 | 0.003287 | 6.4e-05 | 0.000875 | 75.1062 | 1.3821 | 2.5474 | 100.3 | 21.65 | 30.417 | 64.2328 | 76.737 | 62.5366 | 0.007858 | 1.0517 | 0.02506 | 23.8274 | 0.006554 | 0.13629 | 0 | 5.3e-05 | 5.3e-05 | 105.2372
2019-03-29 11:30 | 1553873409.8553 | -310 | 0.20945 | 0.000459 | 0 | 100.8889 | 100.6728 | 127.2529 | 20.9806 | 0.001794 | 0.001645 | -151.8844 | 0.2286 | 8.4031 | 0.002203 | 0.002989 | 0.002739 | 6.4e-05 | 0.000877 | 75.1189 | 1.3882 | 2.5476 | 100.3 | 21.64 | 30.411 | 64.7303 | 74.5492 | 61.1411 | 0.007852 | 1.0561 | 0.025111 | 23.7772 | 0.019182 | 0.136 | -2e-06 | -0.000161 | -0.000161 | -34.4123
2019-03-29 11:30 | 1553873410.8563 | -309 | 0.20944 | 0.000461 | 0 | 100.9159 | 100.71 | 127.2154 | 20.9814 | 0.001767 | 0.001629 | -151.8818 | 0.22895 | 8.4043 | 0.002241 | 0.003649 | 0.003513 | 6.6e-05 | 0.000882 | 75.0645 | 1.3961 | 2.5475 | 100.3 | 21.64 | 30.411 | 64.0689 | 75.4759 | 60.759 | 0.007852 | 1.0573 | 0.025122 | 23.7601 | 0.018266 | 0.1362 | -5e-06 | -0.000353 | -0.000353 | -15.6933
2019-03-29 11:30 | 1553873411.8556 | -308 | 0.20947 | 0.000458 | -1e-06 | 100.8354 | 100.737 | 127.4506 | 20.9825 | 0.001779 | 0.001632 | -151.8801 | 0.22944 | 8.403 | 0.002285 | 0.003947 | 0.003822 | 6.6e-05 | 0.000894 | 74.9917 | 1.3665 | 2.5475 | 100.3 | 21.64 | 30.412 | 64.3676 | 76.5216 | 60.8273 | 0.007852 | 1.0571 | 0.025142 | 23.7845 | 0.008732 | 0.13649 | -1e-06 | -8e-06 | -8e-06 | -659.3183
2019-03-29 11:30 | 1553873412.8571 | -307 | 0.20947 | 0.000456 | -1e-06 | 100.853 | 100.7496 | 127.6916 | 20.9818 | 0.001781 | 0.001661 | -151.8862 | 0.22953 | 8.4028 | 0.002229 | 0.003557 | 0.003332 | 6.7e-05 | 0.000894 | 75.0398 | 1.349 | 2.5474 | 100.3 | 21.63 | 30.417 | 64.409 | 75.1886 | 61.5315 | 0.007849 | 1.0549 | 0.02514 | 23.832 | 0.009177 | 0.13654 | 1e-06 | 0.000122 | 0.000122 | 45.4756
2019-03-29 11:30 | 1553873413.854 | -306 | 0.20949 | 0.000456 | -1e-06 | 100.8261 | 100.8277 | 127.7573 | 20.9824 | 0.001766 | 0.001611 | -151.8945 | 0.23001 | 8.4022 | 0.00223 | 0.002915 | 0.002763 | 6.8e-05 | 0.000899 | 75.0733 | 1.4107 | 2.5475 | 100.3 | 21.63 | 30.421 | 64.0458 | 75.2032 | 60.3184 | 0.00785 | 1.0587 | 0.025192 | 23.7949 | -0.000148 | 0.13684 | 1e-06 | 0.000174 | 0.000174 | 31.9309
2019-03-29 11:30 | 1553873414.8551 | -305 | 0.20945 | 0.000457 | -1e-06 | 100.783 | 100.6346 | 127.5006 | 20.9823 | 0.001823 | 0.001709 | -151.884 | 0.23 | 8.4016 | 0.002263 | 0.003167 | 0.002946 | 6.4e-05 | 0.000897 | 75.0801 | 1.3611 | 2.5475 | 100.3 | 21.63 | 30.426 | 65.4275 | 75.992 | 62.6664 | 0.007851 | 1.0513 | 0.025078 | 23.8545 | 0.01318 | 0.13683 | -1e-06 | -1.1e-05 | -1.1e-05 | -502.2834
2019-03-29 11:30 | 1553873415.8542 | -304 | 0.20947 | 0.000462 | -1e-06 | 100.8644 | 100.7953 | 127.5292 | 20.9829 | 0.001793 | 0.001765 | -151.8882 | 0.23005 | 8.4021 | 0.002237 | 0.003629 | 0.003444 | 7e-05 | 0.0009 | 75.1076 | 1.357 | 2.5474 | 100.3 | 21.62 | 30.437 | 64.6989 | 75.3776 | 64.0376 | 0.007849 | 1.047 | 0.02503 | 23.9059 | 0.006127 | 0.13685 | -5e-06 | -0.000339 | -0.000339 | -16.4218
2019-03-29 11:30 | 1553873416.8541 | -303 | 0.20946 | 0.000457 | -1e-06 | 100.7417 | 100.766 | 127.6524 | 20.9834 | 0.001797 | 0.001699 | -151.8682 | 0.2297 | 8.4018 | 0.002187 | 0.003512 | 0.003334 | 6.7e-05 | 0.000907 | 75.0844 | 1.341 | 2.5474 | 100.3 | 21.62 | 30.443 | 64.8079 | 74.1675 | 62.441 | 0.007851 | 1.052 | 0.025102 | 23.8607 | -0.002158 | 0.13664 | 1e-06 | 0.000129 | 0.000129 | 43.1588
2019-03-29 11:30 | 1553873417.8552 | -302 | 0.20951 | 0.000458 | -1e-06 | 100.8689 | 100.6905 | 127.5205 | 20.9831 | 0.001848 | 0.00167 | -151.8725 | 0.23007 | 8.4022 | 0.002203 | 0.003894 | 0.003702 | 6.8e-05 | 0.000908 | 75.0435 | 1.3139 | 2.5475 | 100.3 | 21.61 | 30.457 | 66.0205 | 74.5436 | 61.7375 | 0.007849 | 1.0542 | 0.025115 | 23.8234 | 0.015828 | 0.13687 | -1e-06 | 2e-05 | 2e-05 | 275.7094
2019-03-29 11:30 | 1553873418.8552 | -301 | 0.20949 | 0.000456 | -1e-06 | 100.8626 | 100.8131 | 127.4456 | 20.9843 | 0.001795 | 0.001682 | -151.8737 | 0.22978 | 8.4022 | 0.002232 | 0.0037 | 0.003563 | 6.2e-05 | 0.000908 | 75.0888 | 1.3872 | 2.5475 | 100.3 | 21.61 | 30.465 | 64.758 | 75.2512 | 62.0248 | 0.007851 | 1.0533 | 0.025097 | 23.8266 | 0.004393 | 0.1367 | 2e-06 | 0.000205 | 0.000205 | 27.1739
2019-03-29 11:30 | 1553873419.8542 | -300 | 0.2095 | 0.000459 | -1e-06 | 100.7737 | 100.712 | 127.4822 | 20.9836 | 0.001793 | 0.001703 | -151.8832 | 0.22949 | 8.4009 | 0.002212 | 0.00312 | 0.002895 | 6.2e-05 | 0.000908 | 75.1148 | 1.2712 | 2.5473 | 100.3 | 21.61 | 30.471 | 64.7065 | 74.7719 | 62.5372 | 0.007853 | 1.0517 | 0.025081 | 23.8482 | 0.005473 | 0.13652 | -2e-06 | -8.8e-05 | -8.8e-05 | -63.31
2019-03-29 11:30 | 1553873420.8543 | -299 | 0.20947 | 0.000457 | -1e-06 | 100.7743 | 100.8206 | 127.5454 | 20.9841 | 0.001737 | 0.001631 | -151.8714 | 0.22995 | 8.404 | 0.002284 | 0.003447 | 0.003199 | 6.4e-05 | 0.000915 | 75.1063 | 1.2997 | 2.5474 | 100.3 | 21.6 | 30.486 | 63.3642 | 76.5005 | 60.8101 | 0.007852 | 1.0571 | 0.025152 | 23.7927 | -0.004107 | 0.1368 | 0 | 6.1e-05 | 6.1e-05 | 91.1647
2019-03-29 11:30 | 1553873421.8553 | -298 | 0.20948 | 0.000455 | -1e-06 | 100.7916 | 100.7059 | 127.4437 | 20.9844 | 0.001771 | 0.001665 | -151.8637 | 0.22911 | 8.4046 | 0.002323 | 0.003558 | 0.003412 | 6.6e-05 | 0.000907 | 75.1088 | 1.3199 | 2.5474 | 100.3 | 21.6 | 30.494 | 64.1618 | 77.4486 | 61.6239 | 0.007854 | 1.0546 | 0.025112 | 23.8122 | 0.007613 | 0.1363 | 3e-06 | 0.000262 | 0.000262 | 21.1844
2019-03-29 11:30 | 1553873422.8543 | -297 | 0.20951 | 0.000456 | -1e-06 | 100.8222 | 100.7564 | 127.1813 | 20.9854 | 0.001788 | 0.001612 | -151.8628 | 0.22912 | 8.403 | 0.002317 | 0.003119 | 0.00287 | 6.8e-05 | 0.000914 | 75.1 | 1.3103 | 2.5474 | 100.3 | 21.6 | 30.506 | 64.5913 | 77.2898 | 60.3509 | 0.007857 | 1.0586 | 0.025134 | 23.7424 | 0.005837 | 0.13631 | 2e-06 | 0.000216 | 0.000216 | 25.7156
2019-03-29 11:30 | 1553873423.8554 | -296 | 0.2095 | 0.000459 | -1e-06 | 100.7806 | 100.5267 | 127.1959 | 20.9848 | 0.001771 | 0.001705 | -151.8658 | 0.23064 | 8.401 | 0.002386 | 0.003597 | 0.003474 | 6.9e-05 | 0.000905 | 75.0732 | 1.3522 | 2.5474 | 100.3 | 21.59 | 30.513 | 64.1653 | 78.9652 | 62.5728 | 0.007854 | 1.0516 | 0.025052 | 23.8227 | 0.022558 | 0.13721 | -2e-06 | -0.000142 | -0.000142 | -39.2738
2019-03-29 11:30 | 1553873424.8553 | -295 | 0.20952 | 0.000459 | -1e-06 | 100.8294 | 100.6721 | 127.4191 | 20.9857 | 0.0018 | 0.001684 | -151.8606 | 0.22996 | 8.4026 | 0.002348 | 0.00409 | 0.003959 | 6.4e-05 | 0.000904 | 75.0706 | 1.3462 | 2.5473 | 100.3 | 21.59 | 30.518 | 64.864 | 78.0424 | 62.0822 | 0.007855 | 1.0531 | 0.025092 | 23.8262 | 0.013961 | 0.1368 | -2e-06 | -6.7e-05 | -6.7e-05 | -83.6142
2019-03-29 11:30 | 1553873425.8548 | -294 | 0.20951 | 0.000458 | -1e-06 | 100.7905 | 100.7909 | 127.582 | 20.9859 | 0.001827 | 0.001653 | -151.8532 | 0.22954 | 8.4015 | 0.00225 | 0.003862 | 0.003671 | 6.6e-05 | 0.0009 | 75.039 | 1.3219 | 2.5474 | 100.3 | 21.58 | 30.527 | 65.5191 | 75.6939 | 61.3221 | 0.007853 | 1.0555 | 0.025137 | 23.8143 | -3.8e-05 | 0.13655 | -1e-06 | -2.4e-05 | -2.4e-05 | -228.1083
2019-03-29 11:30 | 1553873426.8545 | -293 | 0.20947 | 0.000458 | -1e-06 | 100.8383 | 100.8076 | 127.3492 | 20.9858 | 0.001753 | 0.001682 | -151.8609 | 0.23042 | 8.4019 | 0.002349 | 0.003929 | 0.003764 | 5.9e-05 | 0.000896 | 75.0432 | 1.3421 | 2.5474 | 100.3 | 21.58 | 30.531 | 63.7468 | 78.062 | 62.0308 | 0.007854 | 1.0533 | 0.025087 | 23.8178 | 0.002727 | 0.13708 | -1e-06 | -4.7e-05 | -4.7e-05 | -117.4833
2019-03-29 11:30 | 1553873427.8551 | -292 | 0.2095 | 0.000456 | -1e-06 | 100.7934 | 100.7206 | 127.082 | 20.986 | 0.001814 | 0.001657 | -151.8702 | 0.2297 | 8.4032 | 0.002348 | 0.003997 | 0.003739 | 6.3e-05 | 0.000896 | 75.0595 | 1.3912 | 2.5474 | 100.3 | 21.58 | 30.535 | 65.2047 | 78.0489 | 61.4262 | 0.007855 | 1.0552 | 0.025084 | 23.7713 | 0.006469 | 0.13664 | 1e-06 | 0.000131 | 0.000131 | 42.5613
2019-03-29 11:30 | 1553873428.8539 | -291 | 0.20947 | 0.000458 | -1e-06 | 100.8979 | 100.7393 | 126.7077 | 20.9863 | 0.001773 | 0.001678 | -151.8513 | 0.22967 | 8.4039 | 0.002305 | 0.003436 | 0.003294 | 6.6e-05 | 0.000895 | 75.0527 | 1.3499 | 2.5475 | 100.3 | 21.57 | 30.542 | 64.2197 | 77.0194 | 61.9431 | 0.007852 | 1.0536 | 0.025027 | 23.7546 | 0.014068 | 0.13663 | -1e-06 | 0 | 0 | 29340.6654
2019-03-29 11:30 | 1553873429.8539 | -290 | 0.20949 | 0.000455 | -1e-06 | 100.8719 | 100.71 | 126.6694 | 20.987 | 0.00183 | 0.001653 | -151.8517 | 0.22971 | 8.3999 | 0.002292 | 0.0039 | 0.003671 | 6.5e-05 | 0.000895 | 74.9962 | 1.3157 | 2.5474 | 100.3 | 21.57 | 30.545 | 65.5904 | 76.7044 | 61.3388 | 0.007853 | 1.0555 | 0.025046 | 23.7296 | 0.014365 | 0.13666 | 3e-06 | 0.000262 | 0.000262 | 21.2302
2019-03-29 11:30 | 1553873430.8536 | -289 | 0.20947 | 0.000458 | 0 | 100.806 | 100.8998 | 126.8661 | 20.9869 | 0.001841 | 0.001682 | -151.8614 | 0.23019 | 8.4028 | 0.002266 | 0.003696 | 0.003456 | 6.5e-05 | 0.00089 | 75.0269 | 1.3185 | 2.5474 | 100.3 | 21.57 | 30.549 | 65.858 | 76.0806 | 62.0273 | 0.007854 | 1.0533 | 0.02504 | 23.7725 | -0.008313 | 0.13694 | -2e-06 | -0.000101 | -0.000101 | -55.3352
2019-03-29 11:30 | 1553873431.8549 | -288 | 0.2095 | 0.000457 | -1e-06 | 100.8398 | 100.7568 | 126.841 | 20.9881 | 0.001756 | 0.001659 | -151.8547 | 0.22953 | 8.4018 | 0.002226 | 0.003863 | 0.003633 | 6.6e-05 | 0.00089 | 75.0685 | 1.3575 | 2.5474 | 100.3 | 21.57 | 30.554 | 63.8224 | 75.1044 | 61.4846 | 0.007855 | 1.055 | 0.025058 | 23.7508 | 0.007373 | 0.13655 | 1e-06 | 0.000134 | 0.000134 | 41.4015
2019-03-29 11:31 | 1553873432.8548 | -287 | 0.20952 | 0.000458 | -1e-06 | 100.8868 | 100.9056 | 127.0809 | 20.9877 | 0.001824 | 0.001688 | -151.8373 | 0.23006 | 8.403 | 0.002319 | 0.003265 | 0.003119 | 6.5e-05 | 0.000885 | 75.0697 | 1.3453 | 2.5474 | 100.3 | 21.56 | 30.562 | 65.4474 | 77.3535 | 62.1663 | 0.007852 | 1.0529 | 0.025056 | 23.7975 | -0.001659 | 0.13686 | -1e-06 | 7e-06 | 7e-06 | 814.1585
2019-03-29 11:31 | 1553873433.8542 | -286 | 0.20951 | 0.000459 | -1e-06 | 100.9078 | 100.6496 | 126.6663 | 20.9883 | 0.001829 | 0.001653 | -151.8385 | 0.22981 | 8.4046 | 0.002416 | 0.003582 | 0.00344 | 6.4e-05 | 0.000886 | 75.0141 | 1.2873 | 2.5475 | 100.3 | 21.56 | 30.566 | 65.5613 | 79.6844 | 61.3393 | 0.007853 | 1.0555 | 0.025046 | 23.7293 | 0.022917 | 0.13671 | -2e-06 | -0.000108 | -0.000108 | -51.7572
2019-03-29 11:31 | 1553873434.8537 | -285 | 0.20951 | 0.000457 | -1e-06 | 100.8743 | 100.8352 | 126.883 | 20.9886 | 0.001807 | 0.001721 | -151.8386 | 0.22988 | 8.4025 | 0.002253 | 0.003353 | 0.003113 | 6.4e-05 | 0.000886 | 75.1 | 1.321 | 2.5475 | 100.3 | 21.56 | 30.573 | 65.0386 | 75.7527 | 62.9764 | 0.007855 | 1.0503 | 0.025006 | 23.8077 | 0.00346 | 0.13676 | 1e-06 | 9.3e-05 | 9.3e-05 | 59.6648
2019-03-29 11:31 | 1553873435.8537 | -284 | 0.2095 | 0.000455 | 0 | 100.8961 | 100.7707 | 126.7572 | 20.9876 | 0.001806 | 0.001672 | -151.865 | 0.22949 | 8.4022 | 0.002297 | 0.004158 | 0.003942 | 6.3e-05 | 0.000887 | 75.0676 | 1.3095 | 2.5475 | 100.3 | 21.55 | 30.575 | 65.0227 | 76.8196 | 61.7975 | 0.007851 | 1.054 | 0.025038 | 23.7541 | 0.011121 | 0.13653 | 2e-06 | 0.000124 | 0.000124 | 44.8821
2019-03-29 11:31 | 1553873436.8547 | -283 | 0.20948 | 0.000461 | 0 | 100.8958 | 100.9175 | 127.0363 | 20.9881 | 0.001798 | 0.00168 | -151.8552 | 0.22957 | 8.4015 | 0.002273 | 0.003549 | 0.003397 | 6.5e-05 | 0.000887 | 75.0573 | 1.3214 | 2.5474 | 100.3 | 21.55 | 30.578 | 64.8257 | 76.2454 | 61.9822 | 0.007852 | 1.0535 | 0.025058 | 23.7868 | -0.001922 | 0.13657 | -5e-06 | -0.000386 | -0.000386 | -14.4015
2019-03-29 11:31 | 1553873437.8548 | -282 | 0.20949 | 0.000459 | 0 | 100.8785 | 100.7417 | 127.447 | 20.9877 | 0.001786 | 0.001701 | -151.8496 | 0.22915 | 8.4008 | 0.002361 | 0.003305 | 0.00317 | 6.4e-05 | 0.000886 | 75.0732 | 1.3363 | 2.5474 | 100.3 | 21.55 | 30.586 | 64.5334 | 78.3521 | 62.4866 | 0.007854 | 1.0519 | 0.02508 | 23.8431 | 0.012131 | 0.13632 | -3e-06 | -0.000195 | -0.000195 | -28.3929
2019-03-29 11:31 | 1553873438.8552 | -281 | 0.2095 | 0.000456 | 0 | 100.9266 | 100.823 | 127.5348 | 20.9891 | 0.001837 | 0.001652 | -151.8701 | 0.22867 | 8.4009 | 0.00231 | 0.00313 | 0.002994 | 6.4e-05 | 0.000884 | 75.1195 | 1.3363 | 2.5473 | 100.3 | 21.55 | 30.589 | 65.7704 | 77.1344 | 61.3221 | 0.007854 | 1.0555 | 0.025132 | 23.8099 | 0.009194 | 0.13604 | 2e-06 | 0.000163 | 0.000163 | 33.8673
2019-03-29 11:31 | 1553873439.8538 | -280 | 0.2095 | 0.000459 | -1e-06 | 100.8749 | 100.9083 | 127.3151 | 20.989 | 0.001818 | 0.00168 | -151.8707 | 0.22893 | 8.4024 | 0.002209 | 0.003442 | 0.003165 | 6.5e-05 | 0.000886 | 75.1107 | 1.3227 | 2.5473 | 100.3 | 21.55 | 30.593 | 65.3183 | 74.706 | 61.9754 | 0.007855 | 1.0535 | 0.025086 | 23.8126 | -0.002963 | 0.13619 | -2e-06 | -0.000113 | -0.000113 | -48.9228
2019-03-29 11:31 | 1553873440.854 | -279 | 0.20952 | 0.000458 | 0 | 100.8144 | 100.8032 | 126.9041 | 20.989 | 0.001759 | 0.001664 | -151.8701 | 0.22888 | 8.4064 | 0.002286 | 0.00337 | 0.003178 | 6.3e-05 | 0.000884 | 75.0563 | 1.2986 | 2.5473 | 100.3 | 21.54 | 30.603 | 63.878 | 76.5558 | 61.612 | 0.007853 | 1.0546 | 0.025059 | 23.7613 | 0.000996 | 0.13616 | -2e-06 | -9e-05 | -9e-05 | -61.2751
2019-03-29 11:31 | 1553873441.854 | -278 | 0.20953 | 0.000457 | -1e-06 | 100.8728 | 100.7482 | 127.3858 | 20.9891 | 0.001757 | 0.001669 | -151.875 | 0.22999 | 8.403 | 0.002266 | 0.003877 | 0.003618 | 6.2e-05 | 0.000887 | 75.0954 | 1.3134 | 2.5474 | 100.3 | 21.54 | 30.608 | 63.8297 | 76.0696 | 61.7176 | 0.007854 | 1.0543 | 0.025103 | 23.8101 | 0.011051 | 0.13682 | 0 | 6.3e-05 | 6.3e-05 | 88.331
2019-03-29 11:31 | 1553873442.8554 | -277 | 0.20948 | 0.000459 | -1e-06 | 100.8464 | 100.8397 | 127.8414 | 20.9898 | 0.001844 | 0.001684 | -151.8859 | 0.22977 | 8.4048 | 0.002248 | 0.003513 | 0.003323 | 6.3e-05 | 0.000886 | 75.1044 | 1.3565 | 2.5474 | 100.3 | 21.54 | 30.618 | 65.9404 | 75.6383 | 62.0778 | 0.007857 | 1.0532 | 0.025134 | 23.8654 | 0.000598 | 0.13669 | -3e-06 | -0.000156 | -0.000156 | -35.6625
2019-03-29 11:31 | 1553873443.854 | -276 | 0.20951 | 0.000459 | -1e-06 | 100.8177 | 100.852 | 128.3214 | 20.99 | 0.001814 | 0.001676 | -151.8749 | 0.22944 | 8.4047 | 0.002314 | 0.003904 | 0.003665 | 6.6e-05 | 0.00089 | 75.1313 | 1.3088 | 2.5474 | 100.3 | 21.54 | 30.623 | 65.2101 | 77.2249 | 61.8786 | 0.007858 | 1.0538 | 0.025189 | 23.9031 | -0.003041 | 0.13649 | -2e-06 | -0.000123 | -0.000123 | -45.1648
2019-03-29 11:31 | 1553873444.853 | -275 | 0.20947 | 0.000457 | 0 | 100.8093 | 100.7704 | 127.9953 | 20.9894 | 0.001797 | 0.001694 | -151.8771 | 0.22982 | 8.4023 | 0.002293 | 0.003888 | 0.003709 | 6.7e-05 | 0.00089 | 75.0907 | 1.3434 | 2.5474 | 100.3 | 21.53 | 30.628 | 64.806 | 76.7277 | 62.3312 | 0.007855 | 1.0524 | 0.02514 | 23.8888 | 0.003452 | 0.13672 | 0 | 4.9e-05 | 4.9e-05 | 112.4916
2019-03-29 11:31 | 1553873445.853 | -274 | 0.2095 | 0.000462 | 0 | 100.8838 | 100.7864 | 127.7965 | 20.9908 | 0.001732 | 0.00166 | -151.8599 | 0.22949 | 8.4022 | 0.002293 | 0.003469 | 0.003303 | 6.4e-05 | 0.000888 | 75.1126 | 1.3445 | 2.5475 | 100.3 | 21.53 | 30.636 | 63.2301 | 76.7322 | 61.4974 | 0.007857 | 1.055 | 0.025151 | 23.8406 | 0.00864 | 0.13652 | -6e-06 | -0.000399 | -0.000399 | -13.9425
2019-03-29 11:31 | 1553873446.8531 | -273 | 0.20952 | 0.000455 | 0 | 100.7812 | 100.7998 | 127.771 | 20.9911 | 0.001829 | 0.00169 | -151.8576 | 0.23012 | 8.4034 | 0.002305 | 0.003788 | 0.00353 | 6.5e-05 | 0.000881 | 75.0895 | 1.3407 | 2.5474 | 100.3 | 21.52 | 30.64 | 65.5655 | 77.0241 | 62.2212 | 0.007853 | 1.0527 | 0.025122 | 23.864 | -0.001649 | 0.1369 | 3e-06 | 0.000231 | 0.000231 | 24.1619
2019-03-29 11:31 | 1553873447.8531 | -272 | 0.20947 | 0.00046 | -1e-06 | 100.8521 | 100.8742 | 127.3571 | 20.9917 | 0.001833 | 0.001664 | -151.8571 | 0.22962 | 8.402 | 0.002291 | 0.004462 | 0.004297 | 6.4e-05 | 0.000881 | 75.0486 | 1.315 | 2.5475 | 100.3 | 21.52 | 30.651 | 65.6793 | 76.6667 | 61.5975 | 0.007856 | 1.0547 | 0.025104 | 23.8031 | -0.001955 | 0.1366 | -4e-06 | -0.000249 | -0.000249 | -22.3455
2019-03-29 11:31 | 1553873448.8531 | -271 | 0.20943 | 0.000466 | -1e-06 | 100.8392 | 100.711 | 127.2018 | 20.9921 | 0.001792 | 0.001681 | -151.8702 | 0.22967 | 8.4032 | 0.002321 | 0.003681 | 0.00354 | 6.7e-05 | 0.000879 | 75.094 | 1.3826 | 2.5473 | 100.3 | 21.52 | 30.655 | 64.675 | 77.4109 | 62.0181 | 0.007857 | 1.0533 | 0.025073 | 23.8036 | 0.011381 | 0.13663 | -1.1e-05 | -0.000817 | -0.000817 | -6.8058
2019-03-29 11:31 | 1553873449.8533 | -270 | 0.20949 | 0.000459 | -1e-06 | 100.8922 | 100.9335 | 127.5207 | 20.9926 | 0.001797 | 0.00169 | -151.8893 | 0.23042 | 8.4019 | 0.002284 | 0.003737 | 0.003506 | 6.6e-05 | 0.000877 | 75.1698 | 1.3765 | 2.5474 | 100.3 | 21.51 | 30.661 | 64.816 | 76.5133 | 62.2201 | 0.007854 | 1.0527 | 0.025097 | 23.8406 | -0.003662 | 0.13708 | -1e-06 | -5.3e-05 | -5.3e-05 | -104.8655
2019-03-29 11:31 | 1553873450.8532 | -269 | 0.2095 | 0.000458 | -1e-06 | 100.8913 | 100.9482 | 127.8874 | 20.9921 | 0.001858 | 0.001692 | -151.8707 | 0.23047 | 8.4047 | 0.002337 | 0.003571 | 0.003368 | 6.6e-05 | 0.000877 | 75.1211 | 1.3324 | 2.5474 | 100.3 | 21.51 | 30.674 | 66.2789 | 77.7912 | 62.2816 | 0.007857 | 1.0525 | 0.025131 | 23.877 | -0.005042 | 0.13711 | 0 | 3.5e-05 | 3.5e-05 | 158.0826
2019-03-29 11:31 | 1553873451.8543 | -268 | 0.20945 | 0.000456 | -1e-06 | 100.8144 | 100.6032 | 127.4818 | 20.9926 | 0.001826 | 0.001729 | -151.8503 | 0.2299 | 8.4017 | 0.00232 | 0.004062 | 0.003983 | 6.7e-05 | 0.000874 | 75.0983 | 1.3291 | 2.5475 | 100.3 | 21.51 | 30.68 | 65.5029 | 77.3763 | 63.167 | 0.007858 | 1.0497 | 0.025058 | 23.8705 | 0.01876 | 0.13677 | 1e-06 | 0.000145 | 0.000145 | 38.4997
2019-03-29 11:31 | 1553873452.8547 | -267 | 0.20947 | 0.000457 | -1e-06 | 100.8509 | 100.7315 | 126.8926 | 20.9923 | 0.001768 | 0.001703 | -151.8475 | 0.22993 | 8.4023 | 0.002281 | 0.003904 | 0.003725 | 6.3e-05 | 0.000876 | 75.1522 | 1.3573 | 2.5474 | 100.3 | 21.5 | 30.691 | 64.117 | 76.4391 | 62.5519 | 0.007856 | 1.0517 | 0.025023 | 23.7935 | 0.010598 | 0.13678 | 1e-06 | 0.000129 | 0.000129 | 43.3085
2019-03-29 11:31 | 1553873453.8533 | -266 | 0.20945 | 0.000459 | -1e-06 | 100.8279 | 100.7684 | 126.8156 | 20.9931 | 0.001845 | 0.001666 | -151.8392 | 0.22945 | 8.4011 | 0.002273 | 0.003123 | 0.003012 | 6.5e-05 | 0.000886 | 75.0932 | 1.3128 | 2.5474 | 100.3 | 21.5 | 30.697 | 65.971 | 76.2387 | 61.6442 | 0.007858 | 1.0545 | 0.025049 | 23.7541 | 0.005281 | 0.1365 | -1e-06 | -5.3e-05 | -5.3e-05 | -104.2713
2019-03-29 11:31 | 1553873454.8547 | -265 | 0.2095 | 0.000458 | -1e-06 | 100.8009 | 100.8383 | 126.6694 | 20.9929 | 0.001756 | 0.001695 | -151.8436 | 0.23001 | 8.4018 | 0.002237 | 0.003603 | 0.003388 | 6.6e-05 | 0.000891 | 75.0703 | 1.3406 | 2.5473 | 100.3 | 21.5 | 30.703 | 63.8138 | 75.3779 | 62.3557 | 0.00786 | 1.0523 | 0.025008 | 23.7657 | -0.003317 | 0.13684 | -1e-06 | 1.1e-05 | 1.1e-05 | 499.5868
2019-03-29 11:31 | 1553873455.8543 | -264 | 0.20953 | 0.000455 | -1e-06 | 100.9039 | 100.6632 | 126.7109 | 20.9936 | 0.001823 | 0.001649 | -151.8595 | 0.22943 | 8.4015 | 0.002225 | 0.003432 | 0.003286 | 6.7e-05 | 0.000889 | 75.1444 | 1.341 | 2.5474 | 100.31 | 21.5 | 30.714 | 65.4242 | 75.0838 | 61.2553 | 0.007862 | 1.0557 | 0.025053 | 23.7305 | 0.021358 | 0.13649 | 4e-06 | 0.000316 | 0.000316 | 17.5765
2019-03-29 11:31 | 1553873456.8544 | -263 | 0.20952 | 0.000459 | -1e-06 | 100.8467 | 100.6748 | 126.6821 | 20.9933 | 0.001815 | 0.001652 | -151.8755 | 0.22962 | 8.4023 | 0.002301 | 0.003942 | 0.003765 | 6.5e-05 | 0.000886 | 75.1112 | 1.3233 | 2.5474 | 100.31 | 21.49 | 30.721 | 65.2376 | 76.9311 | 61.3161 | 0.007859 | 1.0555 | 0.025048 | 23.73 | 0.015258 | 0.1366 | -2e-06 | -6.6e-05 | -6.6e-05 | -84.6474
2019-03-29 11:31 | 1553873457.8544 | -262 | 0.20952 | 0.000455 | -1e-06 | 100.7734 | 100.711 | 127.0142 | 20.9944 | 0.001795 | 0.001648 | -151.8821 | 0.22975 | 8.4029 | 0.002353 | 0.003688 | 0.003471 | 6.7e-05 | 0.000884 | 75.0977 | 1.3105 | 2.5474 | 100.31 | 21.49 | 30.726 | 64.7562 | 78.1731 | 61.2202 | 0.00786 | 1.0559 | 0.025085 | 23.7577 | 0.005538 | 0.13668 | 3e-06 | 0.0003 | 0.0003 | 18.5476
2019-03-29 11:31 | 1553873458.8548 | -261 | 0.20953 | 0.000462 | -1e-06 | 100.8656 | 100.9997 | 127.1032 | 20.9945 | 0.001799 | 0.001731 | -151.8599 | 0.22945 | 8.4017 | 0.002311 | 0.004134 | 0.003924 | 6.4e-05 | 0.000882 | 75.0498 | 1.3454 | 2.5474 | 100.31 | 21.49 | 30.735 | 64.8453 | 77.173 | 63.2072 | 0.007862 | 1.0496 | 0.025019 | 23.8365 | -0.01189 | 0.1365 | -6e-06 | -0.000367 | -0.000367 | -15.1315
2019-03-29 11:31 | 1553873459.8535 | -260 | 0.20951 | 0.000462 | -1e-06 | 100.8081 | 100.7393 | 127.0893 | 20.9945 | 0.001775 | 0.001685 | -151.8464 | 0.22939 | 8.4023 | 0.002199 | 0.003742 | 0.003598 | 6.8e-05 | 0.000882 | 75.1004 | 1.3796 | 2.5475 | 100.31 | 21.49 | 30.74 | 64.2792 | 74.4762 | 62.1065 | 0.007863 | 1.0531 | 0.025059 | 23.7962 | 0.006103 | 0.13647 | -6e-06 | -0.000419 | -0.000419 | -13.2608
2019-03-29 11:31 | 1553873460.8535 | -259 | 0.20949 | 0.000455 | -1e-06 | 100.8401 | 100.726 | 126.8964 | 20.994 | 0.001799 | 0.001712 | -151.8676 | 0.22901 | 8.4024 | 0.002265 | 0.003441 | 0.003331 | 6.6e-05 | 0.000881 | 75.0645 | 1.3338 | 2.5473 | 100.31 | 21.49 | 30.749 | 64.8564 | 76.0603 | 62.7637 | 0.007866 | 1.051 | 0.025015 | 23.8014 | 0.010127 | 0.13624 | 3e-06 | 0.000324 | 0.000324 | 17.1125
2019-03-29 11:31 | 1553873461.8526 | -258 | 0.20948 | 0.000462 | 0 | 100.8234 | 100.6871 | 126.7687 | 20.9947 | 0.001801 | 0.001681 | -151.8714 | 0.22853 | 8.402 | 0.002271 | 0.003742 | 0.003616 | 6.4e-05 | 0.000881 | 75.0774 | 1.3641 | 2.5473 | 100.31 | 21.49 | 30.753 | 64.9102 | 76.2028 | 62.0206 | 0.007867 | 1.0533 | 0.02503 | 23.7631 | 0.012095 | 0.13595 | -6e-06 | -0.000431 | -0.000431 | -12.8453
2019-03-29 11:31 | 1553873462.8536 | -257 | 0.2095 | 0.000459 | -1e-06 | 100.8111 | 100.7411 | 126.5126 | 20.9952 | 0.001797 | 0.001698 | -151.8699 | 0.23031 | 8.4022 | 0.002263 | 0.003623 | 0.003413 | 6.7e-05 | 0.000886 | 75.109 | 1.3715 | 2.5475 | 100.3 | 21.48 | 30.757 | 64.8194 | 76.0147 | 62.4197 | 0.007864 | 1.0521 | 0.02499 | 23.7532 | 0.006218 | 0.13701 | -2e-06 | -0.000135 | -0.000135 | -41.2438
2019-03-29 11:31 | 1553873463.8536 | -256 | 0.20949 | 0.000456 | -1e-06 | 100.8404 | 100.5725 | 126.5762 | 20.9959 | 0.001798 | 0.001719 | -151.8859 | 0.2287 | 8.4014 | 0.002286 | 0.003907 | 0.00367 | 6.5e-05 | 0.00089 | 75.0478 | 1.3857 | 2.5475 | 100.3 | 21.48 | 30.766 | 64.8357 | 76.5547 | 62.9399 | 0.007866 | 1.0504 | 0.024977 | 23.7776 | 0.023802 | 0.13605 | 2e-06 | 0.000196 | 0.000196 | 28.3036
2019-03-29 11:31 | 1553873464.8537 | -255 | 0.20949 | 0.000455 | -1e-06 | 100.7396 | 100.6523 | 126.7152 | 20.9953 | 0.001789 | 0.001677 | -151.8952 | 0.23019 | 8.4039 | 0.002271 | 0.003835 | 0.0037 | 6.4e-05 | 0.00089 | 75.1015 | 1.3413 | 2.5474 | 100.3 | 21.48 | 30.772 | 64.6212 | 76.2111 | 61.93 | 0.007868 | 1.0536 | 0.025028 | 23.7549 | 0.00775 | 0.13694 | 3e-06 | 0.000302 | 0.000302 | 18.4621
2019-03-29 11:31 | 1553873465.8527 | -254 | 0.2095 | 0.000459 | -1e-06 | 100.7779 | 100.5759 | 127.1708 | 20.9956 | 0.001753 | 0.001746 | -151.8732 | 0.22817 | 8.405 | 0.002207 | 0.003213 | 0.003169 | 6.4e-05 | 0.000892 | 75.081 | 1.3668 | 2.5473 | 100.3 | 21.47 | 30.782 | 63.7602 | 74.661 | 63.5806 | 0.007865 | 1.0485 | 0.025012 | 23.8561 | 0.017947 | 0.13574 | -1e-06 | -4.3e-05 | -4.3e-05 | -128.7153
2019-03-29 11:31 | 1553873466.8541 | -253 | 0.20949 | 0.000455 | -1e-06 | 100.8413 | 100.7864 | 127.0539 | 20.9958 | 0.001792 | 0.001647 | -151.8723 | 0.22878 | 8.4014 | 0.002276 | 0.003378 | 0.003105 | 6.3e-05 | 0.000887 | 75.0579 | 1.3347 | 2.5473 | 100.3 | 21.47 | 30.788 | 64.6825 | 76.3302 | 61.2035 | 0.007867 | 1.0559 | 0.025089 | 23.7608 | 0.00487 | 0.1361 | 3e-06 | 0.000268 | 0.000268 | 20.658
2019-03-29 11:31 | 1553873467.8544 | -252 | 0.20951 | 0.00046 | -1e-06 | 100.721 | 100.5492 | 126.9419 | 20.9957 | 0.001798 | 0.001687 | -151.8528 | 0.2296 | 8.4015 | 0.002256 | 0.003948 | 0.003714 | 6.4e-05 | 0.000887 | 75.0344 | 1.3949 | 2.5474 | 100.3 | 21.47 | 30.795 | 64.8277 | 75.8429 | 62.1678 | 0.007869 | 1.0529 | 0.025042 | 23.7845 | 0.015266 | 0.13659 | -3e-06 | -0.000136 | -0.000136 | -40.8498
2019-03-29 11:31 | 1553873468.8541 | -251 | 0.2095 | 0.000456 | -1e-06 | 100.7719 | 100.8281 | 127.0452 | 20.9968 | 0.001765 | 0.001717 | -151.8526 | 0.22865 | 8.403 | 0.002331 | 0.004171 | 0.004074 | 6.6e-05 | 0.000887 | 75.1004 | 1.3581 | 2.5473 | 100.3 | 21.47 | 30.81 | 64.0403 | 77.6486 | 62.8737 | 0.007872 | 1.0507 | 0.025026 | 23.8192 | -0.004986 | 0.13603 | 2e-06 | 0.00016 | 0.00016 | 34.6345
2019-03-29 11:31 | 1553873469.8528 | -250 | 0.2095 | 0.000459 | 0 | 100.7899 | 100.8049 | 126.8957 | 20.9973 | 0.001796 | 0.001732 | -151.8491 | 0.22956 | 8.4006 | 0.002264 | 0.003829 | 0.003622 | 6.7e-05 | 0.000887 | 75.019 | 1.3546 | 2.5474 | 100.3 | 21.47 | 30.819 | 64.7768 | 76.0299 | 63.2458 | 0.007875 | 1.0495 | 0.024997 | 23.8184 | -0.001333 | 0.13657 | -3e-06 | -0.000158 | -0.000158 | -35.097
2019-03-29 11:31 | 1553873470.8528 | -249 | 0.20948 | 0.000458 | -1e-06 | 100.7213 | 100.554 | 126.6537 | 20.9971 | 0.001759 | 0.001681 | -151.844 | 0.22971 | 8.4037 | 0.002305 | 0.003473 | 0.003286 | 6.8e-05 | 0.000889 | 75.0362 | 1.3027 | 2.5473 | 100.31 | 21.46 | 30.833 | 63.8907 | 77.0299 | 62.0105 | 0.007873 | 1.0534 | 0.025019 | 23.7519 | 0.014867 | 0.13666 | -1e-06 | -1.1e-05 | -1.1e-05 | -488.6543
2019-03-29 11:31 | 1553873471.854 | -248 | 0.20949 | 0.00046 | 0 | 100.791 | 100.7315 | 126.6652 | 20.9973 | 0.001772 | 0.001676 | -151.8505 | 0.22917 | 8.4011 | 0.00231 | 0.003765 | 0.003572 | 6.6e-05 | 0.000893 | 75.0306 | 1.3695 | 2.5472 | 100.31 | 21.46 | 30.842 | 64.2133 | 77.1517 | 61.8905 | 0.007875 | 1.0537 | 0.025025 | 23.7488 | 0.005286 | 0.13634 | -4e-06 | -0.000241 | -0.000241 | -23.0427
2019-03-29 11:31 | 1553873472.8528 | -247 | 0.2095 | 0.00046 | -1e-06 | 100.7467 | 100.6465 | 127.0558 | 20.9978 | 0.001766 | 0.001692 | -151.8506 | 0.23103 | 8.4009 | 0.002271 | 0.003069 | 0.002866 | 6.8e-05 | 0.000899 | 75.059 | 1.3812 | 2.5475 | 100.31 | 21.46 | 30.85 | 64.0747 | 76.1971 | 62.2791 | 0.007877 | 1.0525 | 0.025049 | 23.7991 | 0.008903 | 0.13744 | -3e-06 | -0.000158 | -0.000158 | -35.3132
2019-03-29 11:31 | 1553873473.8533 | -246 | 0.20949 | 0.000463 | -1e-06 | 100.7357 | 100.7537 | 127.1503 | 20.9974 | 0.001791 | 0.001648 | -151.8428 | 0.22986 | 8.4018 | 0.002306 | 0.00366 | 0.003455 | 6.5e-05 | 0.000904 | 75.0527 | 1.3397 | 2.5474 | 100.31 | 21.45 | 30.866 | 64.6701 | 77.0396 | 61.2169 | 0.007876 | 1.0559 | 0.025098 | 23.7703 | -0.001599 | 0.13675 | -6e-06 | -0.000427 | -0.000427 | -13.0274
2019-03-29 11:31 | 1553873474.8543 | -245 | 0.20948 | 0.000458 | -1e-06 | 100.7902 | 100.6158 | 127.1466 | 20.9987 | 0.001754 | 0.001625 | -151.8231 | 0.22965 | 8.4018 | 0.002325 | 0.003452 | 0.003226 | 6.4e-05 | 0.000904 | 75.0466 | 1.3244 | 2.5472 | 100.31 | 21.45 | 30.874 | 63.7712 | 77.4913 | 60.6773 | 0.007878 | 1.0576 | 0.025118 | 23.7507 | 0.015486 | 0.13662 | 0 | 5.4e-05 | 5.4e-05 | 102.2724
2019-03-29 11:31 | 1553873475.852 | -244 | 0.20949 | 0.000459 | -1e-06 | 100.7518 | 100.6144 | 127.0173 | 20.999 | 0.001765 | 0.00165 | -151.8203 | 0.23046 | 8.4015 | 0.002323 | 0.003714 | 0.003587 | 6.8e-05 | 0.0009 | 75.0392 | 1.3544 | 2.5473 | 100.31 | 21.45 | 30.881 | 64.0393 | 77.4652 | 61.2826 | 0.00788 | 1.0557 | 0.025083 | 23.7602 | 0.012207 | 0.1371 | -1e-06 | -4.4e-05 | -4.4e-05 | -127.8696
2019-03-29 11:31 | 1553873476.8521 | -243 | 0.20947 | 0.000459 | 0 | 100.7366 | 100.6953 | 127.3926 | 20.9994 | 0.001784 | 0.001671 | -151.8259 | 0.22977 | 8.4029 | 0.002319 | 0.003281 | 0.003048 | 6.6e-05 | 0.000897 | 75.0706 | 1.3796 | 2.5474 | 100.31 | 21.45 | 30.892 | 64.5089 | 77.351 | 61.7676 | 0.007883 | 1.0541 | 0.025101 | 23.8125 | 0.003663 | 0.13669 | -2e-06 | -9.6e-05 | -9.6e-05 | -58.2392
2019-03-29 11:31 | 1553873477.8526 | -242 | 0.20951 | 0.000461 | -1e-06 | 100.7291 | 100.5759 | 127.4306 | 20.9993 | 0.001836 | 0.001695 | -151.8422 | 0.22873 | 8.4031 | 0.002289 | 0.003831 | 0.003661 | 6.5e-05 | 0.000894 | 75.0997 | 1.3338 | 2.5474 | 100.31 | 21.45 | 30.898 | 65.742 | 76.6351 | 62.3502 | 0.007884 | 1.0523 | 0.025083 | 23.8368 | 0.013616 | 0.13607 | -4e-06 | -0.000237 | -0.000237 | -23.3458
2019-03-29 11:31 | 1553873478.8521 | -241 | 0.20952 | 0.000461 | -1e-06 | 100.7303 | 100.5533 | 127.3144 | 20.9989 | 0.001727 | 0.001674 | -151.8593 | 0.22928 | 8.4012 | 0.002268 | 0.003529 | 0.00337 | 6.4e-05 | 0.00089 | 75.0502 | 1.3299 | 2.5473 | 100.31 | 21.44 | 30.909 | 63.1224 | 76.1259 | 61.8614 | 0.007882 | 1.0538 | 0.02509 | 23.8085 | 0.015726 | 0.1364 | -5e-06 | -0.000303 | -0.000303 | -18.3512
2019-03-29 11:31 | 1553873479.8525 | -240 | 0.20947 | 0.000461 | -1e-06 | 100.6486 | 100.6766 | 127.2832 | 21.0008 | 0.001821 | 0.001653 | -151.8691 | 0.22879 | 8.4011 | 0.002262 | 0.00316 | 0.002917 | 6.9e-05 | 0.000893 | 75.0263 | 1.3429 | 2.5475 | 100.31 | 21.44 | 30.914 | 65.3797 | 75.9989 | 61.3351 | 0.007884 | 1.0555 | 0.025107 | 23.7869 | -0.002486 | 0.13611 | -4e-06 | -0.000214 | -0.000214 | -25.8656
2019-03-29 11:31 | 1553873480.8535 | -239 | 0.20944 | 0.000465 | -1e-06 | 100.6243 | 100.4257 | 127.3337 | 21.0006 | 0.001793 | 0.001671 | -151.8692 | 0.22907 | 8.3994 | 0.002325 | 0.003516 | 0.003345 | 6.8e-05 | 0.000893 | 75.0333 | 1.328 | 2.5473 | 100.31 | 21.44 | 30.919 | 64.7129 | 77.5063 | 61.7858 | 0.007885 | 1.0541 | 0.025095 | 23.8076 | 0.017673 | 0.13627 | -9e-06 | -0.000671 | -0.000671 | -8.2674
2019-03-29 11:31 | 1553873481.8522 | -238 | 0.2095 | 0.000462 | -1e-06 | 100.648 | 100.597 | 127.3618 | 21.0008 | 0.001779 | 0.001711 | -151.8629 | 0.22919 | 8.4018 | 0.002268 | 0.003508 | 0.003296 | 6.7e-05 | 0.0009 | 74.9872 | 1.3343 | 2.5473 | 100.31 | 21.44 | 30.928 | 64.3768 | 76.1299 | 62.7555 | 0.007887 | 1.051 | 0.025061 | 23.8447 | 0.004528 | 0.13635 | -5e-06 | -0.000301 | -0.000301 | -18.4272
2019-03-29 11:31 | 1553873482.8532 | -237 | 0.20947 | 0.000466 | 0 | 100.6857 | 100.6243 | 127.1478 | 21.0001 | 0.001786 | 0.001669 | -151.8739 | 0.22947 | 8.4032 | 0.002338 | 0.003804 | 0.003598 | 6.5e-05 | 0.000905 | 75.0274 | 1.3141 | 2.5475 | 100.31 | 21.43 | 30.933 | 64.5524 | 77.8129 | 61.7289 | 0.007884 | 1.0542 | 0.025079 | 23.7882 | 0.005452 | 0.13651 | -1.1e-05 | -0.000837 | -0.000837 | -6.6389
2019-03-29 11:31 | 1553873483.8536 | -236 | 0.20946 | 0.000469 | -1e-06 | 100.6327 | 100.6875 | 126.967 | 21.0011 | 0.001845 | 0.001679 | -151.8693 | 0.22993 | 8.4016 | 0.00237 | 0.003695 | 0.003564 | 6.6e-05 | 0.000906 | 75.0457 | 1.3124 | 2.5474 | 100.31 | 21.43 | 30.941 | 65.9677 | 78.5919 | 61.9746 | 0.007886 | 1.0535 | 0.025052 | 23.78 | -0.004867 | 0.13678 | -1.4e-05 | -0.00097 | -0.00097 | -5.7394
2019-03-29 11:31 | 1553873484.8523 | -235 | 0.2095 | 0.000471 | -1e-06 | 100.6935 | 100.6455 | 126.8661 | 21.0008 | 0.001823 | 0.001717 | -151.8815 | 0.2296 | 8.4017 | 0.002358 | 0.003791 | 0.003578 | 6.5e-05 | 0.000907 | 75.0706 | 1.3008 | 2.5475 | 100.31 | 21.43 | 30.946 | 65.4454 | 78.2908 | 62.8806 | 0.007887 | 1.0506 | 0.025008 | 23.8027 | 0.004263 | 0.13659 | -1.7e-05 | -0.001221 | -0.001221 | -4.5522
2019-03-29 11:31 | 1553873485.8524 | -234 | 0.20947 | 0.000464 | -1e-06 | 100.6109 | 100.6213 | 126.9126 | 21.0019 | 0.001736 | 0.001651 | -151.8939 | 0.22979 | 8.4035 | 0.002331 | 0.003183 | 0.003045 | 6.7e-05 | 0.00091 | 75.0351 | 1.3449 | 2.5474 | 100.31 | 21.43 | 30.951 | 63.3495 | 77.6535 | 61.2941 | 0.007888 | 1.0556 | 0.025072 | 23.7508 | -0.000925 | 0.1367 | -8e-06 | -0.000585 | -0.000585 | -9.5062
2019-03-29 11:31 | 1553873486.8524 | -233 | 0.20949 | 0.00047 | -1e-06 | 100.6037 | 100.4704 | 126.8586 | 21.0021 | 0.001824 | 0.001618 | -151.8941 | 0.23005 | 8.4045 | 0.002279 | 0.003727 | 0.003573 | 7.2e-05 | 0.000913 | 75.0223 | 1.3334 | 2.5474 | 100.31 | 21.42 | 30.96 | 65.4561 | 76.3993 | 60.5047 | 0.007886 | 1.0581 | 0.025096 | 23.7177 | 0.011855 | 0.13686 | -1.6e-05 | -0.001163 | -0.001163 | -4.7877
2019-03-29 11:31 | 1553873487.8527 | -232 | 0.20947 | 0.00047 | 0 | 100.5624 | 100.4916 | 127.0158 | 21.003 | 0.001836 | 0.001688 | -151.8932 | 0.23008 | 8.4048 | 0.00233 | 0.003582 | 0.003365 | 6.7e-05 | 0.000911 | 75.0425 | 1.3315 | 2.5474 | 100.31 | 21.42 | 30.966 | 65.7502 | 77.6202 | 62.1968 | 0.007887 | 1.0528 | 0.025048 | 23.7925 | 0.006299 | 0.13687 | -1.5e-05 | -0.001141 | -0.001141 | -4.8834
2019-03-29 11:31 | 1553873488.8524 | -231 | 0.20948 | 0.000466 | -1e-06 | 100.5671 | 100.537 | 126.9593 | 21.0031 | 0.001765 | 0.001696 | -151.8895 | 0.22956 | 8.4043 | 0.00236 | 0.003826 | 0.003693 | 6.8e-05 | 0.000907 | 75.0796 | 1.339 | 2.5474 | 100.31 | 21.42 | 30.972 | 64.0385 | 78.3438 | 62.3898 | 0.007889 | 1.0522 | 0.025035 | 23.794 | 0.002685 | 0.13656 | -1e-05 | -0.000735 | -0.000735 | -7.5571
2019-03-29 11:31 | 1553873489.8524 | -230 | 0.20955 | 0.000466 | -1e-06 | 100.5932 | 100.3994 | 127.128 | 21.0026 | 0.001834 | 0.001674 | -151.888 | 0.22978 | 8.4023 | 0.002361 | 0.004033 | 0.003893 | 6.6e-05 | 0.000911 | 75.078 | 1.3597 | 2.5474 | 100.31 | 21.42 | 30.983 | 65.7 | 78.3589 | 61.8542 | 0.007892 | 1.0539 | 0.025072 | 23.7908 | 0.017246 | 0.1367 | -1.1e-05 | -0.000739 | -0.000739 | -7.5317
2019-03-29 11:31 | 1553873490.8538 | -229 | 0.2095 | 0.000463 | -1e-06 | 100.5627 | 100.3527 | 127.2301 | 21.0027 | 0.001817 | 0.00168 | -151.8881 | 0.22992 | 8.4037 | 0.002395 | 0.003701 | 0.003477 | 7e-05 | 0.000914 | 75.0825 | 1.3448 | 2.5474 | 100.31 | 21.42 | 30.989 | 65.2974 | 79.1884 | 62.0058 | 0.007893 | 1.0534 | 0.025076 | 23.8058 | 0.018697 | 0.13678 | -6e-06 | -0.00041 | -0.00041 | -13.5858
2019-03-29 11:32 | 1553873491.8516 | -228 | 0.20952 | 0.000464 | 0 | 100.5869 | 100.3837 | 127.2125 | 21.0036 | 0.001787 | 0.001726 | -151.902 | 0.22919 | 8.4035 | 0.002306 | 0.003959 | 0.003696 | 6.8e-05 | 0.000908 | 75.0521 | 1.3296 | 2.5475 | 100.31 | 21.42 | 30.999 | 64.5637 | 77.0571 | 63.1113 | 0.007896 | 1.0499 | 0.025033 | 23.8434 | 0.018086 | 0.13634 | -9e-06 | -0.000651 | -0.000651 | -8.5272
2019-03-29 11:32 | 1553873492.8526 | -227 | 0.20951 | 0.000467 | -1e-06 | 100.6037 | 100.4772 | 127.2546 | 21.003 | 0.001802 | 0.001692 | -151.896 | 0.22962 | 8.3992 | 0.002388 | 0.003378 | 0.003251 | 7.3e-05 | 0.000908 | 75.1085 | 1.3584 | 2.5475 | 100.31 | 21.42 | 31.005 | 64.9465 | 79.0265 | 62.2936 | 0.007897 | 1.0525 | 0.025068 | 23.8183 | 0.011248 | 0.1366 | -1.1e-05 | -0.000817 | -0.000817 | -6.806
2019-03-29 11:32 | 1553873493.8516 | -226 | 0.2095 | 0.000465 | -1e-06 | 100.5573 | 100.4895 | 127.2215 | 21.0046 | 0.001869 | 0.001711 | -151.8979 | 0.22822 | 8.4037 | 0.002418 | 0.003579 | 0.003399 | 6.6e-05 | 0.00091 | 75.1 | 1.3499 | 2.5475 | 100.31 | 21.42 | 31.018 | 66.5392 | 79.753 | 62.749 | 0.007901 | 1.051 | 0.025048 | 23.8313 | 0.006028 | 0.13577 | -9e-06 | -0.000653 | -0.000653 | -8.464
2019-03-29 11:32 | 1553873494.853 | -225 | 0.20951 | 0.000467 | -1e-06 | 100.5172 | 100.5332 | 127.3877 | 21.0042 | 0.001822 | 0.001705 | -151.9018 | 0.22984 | 8.4045 | 0.00239 | 0.003555 | 0.003301 | 7e-05 | 0.000911 | 75.0592 | 1.3955 | 2.5475 | 100.31 | 21.42 | 31.024 | 65.4127 | 79.0674 | 62.598 | 0.007902 | 1.0515 | 0.02507 | 23.8415 | -0.001428 | 0.13673 | -1.2e-05 | -0.00085 | -0.00085 | -6.5444
2019-03-29 11:32 | 1553873495.8517 | -224 | 0.20951 | 0.000465 | -1e-06 | 100.5243 | 100.4124 | 127.236 | 21.0047 | 0.001827 | 0.001682 | -151.9059 | 0.23022 | 8.3996 | 0.002369 | 0.0034 | 0.003193 | 6.7e-05 | 0.000912 | 75.086 | 1.3511 | 2.5474 | 100.31 | 21.41 | 31.031 | 65.5453 | 78.5714 | 62.0454 | 0.007899 | 1.0533 | 0.025076 | 23.8077 | 0.009967 | 0.13696 | -8e-06 | -0.000555 | -0.000555 | -10.0444
2019-03-29 11:32 | 1553873496.8523 | -223 | 0.20951 | 0.000464 | -1e-06 | 100.5827 | 100.4981 | 127.202 | 21.0045 | 0.001818 | 0.001686 | -151.9005 | 0.22908 | 8.4019 | 0.002424 | 0.003796 | 0.003599 | 6.9e-05 | 0.000911 | 75.08 | 1.3358 | 2.5473 | 100.31 | 21.41 | 31.042 | 65.325 | 79.8832 | 62.1482 | 0.007902 | 1.0529 | 0.025068 | 23.8082 | 0.007532 | 0.13628 | -8e-06 | -0.000597 | -0.000597 | -9.2929
2019-03-29 11:32 | 1553873497.852 | -222 | 0.20952 | 0.000463 | 0 | 100.5621 | 100.3697 | 127.2034 | 21.0052 | 0.001831 | 0.001697 | -151.8759 | 0.22923 | 8.4037 | 0.002414 | 0.0034 | 0.003168 | 6.8e-05 | 0.000912 | 75.0748 | 1.3185 | 2.5474 | 100.31 | 21.41 | 31.048 | 65.6386 | 79.6453 | 62.4113 | 0.007903 | 1.0521 | 0.025059 | 23.8177 | 0.017123 | 0.13637 | -7e-06 | -0.000518 | -0.000518 | -10.7047
2019-03-29 11:32 | 1553873498.8528 | -221 | 0.20952 | 0.000461 | -1e-06 | 100.4744 | 100.3793 | 127.2355 | 21.0053 | 0.001847 | 0.001693 | -151.8638 | 0.22993 | 8.4023 | 0.002427 | 0.003823 | 0.003712 | 6.6e-05 | 0.000914 | 75.0521 | 1.3293 | 2.5474 | 100.31 | 21.41 | 31.052 | 66.0093 | 79.9652 | 62.3209 | 0.007904 | 1.0524 | 0.025065 | 23.8175 | 0.008467 | 0.13678 | -4e-06 | -0.000272 | -0.000272 | -20.502
2019-03-29 11:32 | 1553873499.8518 | -220 | 0.2095 | 0.000461 | -1e-06 | 100.5228 | 100.3318 | 127.3055 | 21.0066 | 0.00183 | 0.001709 | -151.8624 | 0.22964 | 8.401 | 0.002463 | 0.003682 | 0.003555 | 6.8e-05 | 0.000916 | 75.0915 | 1.3324 | 2.5475 | 100.31 | 21.41 | 31.062 | 65.6133 | 80.8242 | 62.7002 | 0.007907 | 1.0512 | 0.025058 | 23.8375 | 0.017011 | 0.13661 | -4e-06 | -0.000257 | -0.000257 | -21.6617
2019-03-29 11:32 | 1553873500.8522 | -219 | 0.20949 | 0.00046 | -1e-06 | 100.4857 | 100.354 | 127.3205 | 21.0069 | 0.001742 | 0.00166 | -151.855 | 0.22978 | 8.4034 | 0.00242 | 0.0036 | 0.003385 | 7.1e-05 | 0.000922 | 75.0624 | 1.3468 | 2.5475 | 100.31 | 21.41 | 31.066 | 63.4988 | 79.7851 | 61.5232 | 0.007908 | 1.0549 | 0.025103 | 23.7971 | 0.011731 | 0.1367 | -3e-06 | -0.000215 | -0.000215 | -25.9088
2019-03-29 11:32 | 1553873501.8522 | -218 | 0.20953 | 0.000461 | -1e-06 | 100.4579 | 100.4438 | 127.4799 | 21.0063 | 0.001819 | 0.001641 | -151.8494 | 0.22981 | 8.4044 | 0.002415 | 0.00416 | 0.003984 | 6.7e-05 | 0.000921 | 75.1006 | 1.3688 | 2.5474 | 100.31 | 21.41 | 31.075 | 65.3399 | 79.6612 | 61.0598 | 0.00791 | 1.0564 | 0.025137 | 23.7955 | 0.001256 | 0.13672 | -3e-06 | -0.000196 | -0.000196 | -28.3532
2019-03-29 11:32 | 1553873502.8519 | -217 | 0.2095 | 0.000458 | -1e-06 | 100.548 | 100.3281 | 127.6397 | 21.0067 | 0.001798 | 0.001659 | -151.8475 | 0.23091 | 8.4027 | 0.002407 | 0.003306 | 0.003083 | 6.7e-05 | 0.000922 | 75.1195 | 1.3279 | 2.5473 | 100.31 | 21.41 | 31.079 | 64.8492 | 79.4879 | 61.4861 | 0.007911 | 1.055 | 0.025136 | 23.8256 | 0.019583 | 0.13737 | 0 | 2e-05 | 2e-05 | 281.1081
2019-03-29 11:32 | 1553873503.8519 | -216 | 0.20952 | 0.000457 | 0 | 100.5076 | 100.4114 | 127.5902 | 21.0071 | 0.001763 | 0.001664 | -151.8504 | 0.22966 | 8.4025 | 0.002336 | 0.003613 | 0.003445 | 6.5e-05 | 0.000921 | 75.0735 | 1.3281 | 2.5475 | 100.31 | 21.41 | 31.083 | 63.9908 | 77.763 | 61.6125 | 0.007912 | 1.0546 | 0.025127 | 23.8254 | 0.008567 | 0.13663 | 1e-06 | 8.3e-05 | 8.3e-05 | 67.1217
2019-03-29 11:32 | 1553873504.8524 | -215 | 0.20955 | 0.000457 | -1e-06 | 100.4827 | 100.437 | 127.1924 | 21.0066 | 0.0018 | 0.001699 | -151.8541 | 0.22926 | 8.4022 | 0.002365 | 0.003461 | 0.003291 | 6.8e-05 | 0.000914 | 75.0406 | 1.3386 | 2.5475 | 100.31 | 21.4 | 31.09 | 64.8954 | 78.461 | 62.4551 | 0.007909 | 1.052 | 0.025056 | 23.8182 | 0.004075 | 0.13639 | 1e-06 | 0.000117 | 0.000117 | 47.5172
2019-03-29 11:32 | 1553873505.853 | -214 | 0.20955 | 0.000458 | 0 | 100.5178 | 100.4673 | 126.9734 | 21.008 | 0.001819 | 0.001691 | -151.8367 | 0.22925 | 8.405 | 0.002415 | 0.004102 | 0.003838 | 6.9e-05 | 0.000911 | 75.0853 | 1.3094 | 2.5474 | 100.31 | 21.4 | 31.093 | 65.3572 | 79.6646 | 62.2727 | 0.00791 | 1.0525 | 0.025041 | 23.7912 | 0.004488 | 0.13638 | -1e-06 | -6.8e-05 | -6.8e-05 | -81.1353
2019-03-29 11:32 | 1553873506.852 | -213 | 0.20949 | 0.000462 | -1e-06 | 100.5558 | 100.4114 | 126.7459 | 21.0076 | 0.001765 | 0.001654 | -151.8193 | 0.22877 | 8.4022 | 0.002345 | 0.00385 | 0.003662 | 6.6e-05 | 0.000906 | 75.0817 | 1.3058 | 2.5474 | 100.31 | 21.4 | 31.105 | 64.0469 | 77.9862 | 61.3776 | 0.007913 | 1.0554 | 0.025052 | 23.7381 | 0.012854 | 0.1361 | -5e-06 | -0.000363 | -0.000363 | -15.2649
2019-03-29 11:32 | 1553873507.851 | -212 | 0.20952 | 0.00046 | -1e-06 | 100.5294 | 100.2957 | 126.4898 | 21.0079 | 0.001814 | 0.001638 | -151.8334 | 0.23001 | 8.4018 | 0.00238 | 0.003788 | 0.003628 | 6.4e-05 | 0.000902 | 75.0625 | 1.3081 | 2.5474 | 100.31 | 21.4 | 31.112 | 65.2298 | 78.8231 | 61.0042 | 0.007915 | 1.0565 | 0.025041 | 23.7009 | 0.020823 | 0.13683 | -2e-06 | -5.4e-05 | -5.4e-05 | -104.0901
2019-03-29 11:32 | 1553873508.851 | -211 | 0.20949 | 0.000458 | -1e-06 | 100.4699 | 100.481 | 126.7032 | 21.0076 | 0.001767 | 0.001703 | -151.8315 | 0.22997 | 8.402 | 0.002385 | 0.003858 | 0.003713 | 6.4e-05 | 0.0009 | 75.0634 | 1.3397 | 2.5474 | 100.31 | 21.4 | 31.119 | 64.0985 | 78.9532 | 62.5459 | 0.007917 | 1.0517 | 0.025004 | 23.7756 | -0.00099 | 0.13681 | 0 | 5.6e-05 | 5.6e-05 | 100.1508
2019-03-29 11:32 | 1553873509.8525 | -210 | 0.20947 | 0.000461 | -1e-06 | 100.4621 | 100.5673 | 126.5645 | 21.0085 | 0.001854 | 0.001659 | -151.8142 | 0.23038 | 8.4033 | 0.002376 | 0.003406 | 0.003278 | 6.5e-05 | 0.000901 | 75.0541 | 1.3517 | 2.5474 | 100.31 | 21.4 | 31.136 | 66.1863 | 78.7442 | 61.5028 | 0.007921 | 1.055 | 0.02503 | 23.7256 | -0.009366 | 0.13706 | -4e-06 | -0.000262 | -0.000262 | -21.2994
2019-03-29 11:32 | 1553873510.8511 | -209 | 0.2095 | 0.000461 | -1e-06 | 100.4774 | 100.4813 | 126.7175 | 21.0088 | 0.001879 | 0.001726 | -151.8038 | 0.22956 | 8.4023 | 0.002392 | 0.003633 | 0.003444 | 6.7e-05 | 0.000902 | 75.0482 | 1.3733 | 2.5475 | 100.31 | 21.4 | 31.144 | 66.7852 | 79.1148 | 63.1241 | 0.007923 | 1.0499 | 0.024984 | 23.7974 | -0.000355 | 0.13657 | -4e-06 | -0.00026 | -0.00026 | -21.3556
2019-03-29 11:32 | 1553873511.8511 | -208 | 0.20949 | 0.00046 | -1e-06 | 100.5109 | 100.3595 | 126.7663 | 21.0085 | 0.001797 | 0.001643 | -151.7954 | 0.23038 | 8.4021 | 0.00243 | 0.004101 | 0.003973 | 6.6e-05 | 0.000906 | 75.0023 | 1.3484 | 2.5475 | 100.31 | 21.39 | 31.152 | 64.8229 | 80.0407 | 61.1208 | 0.00792 | 1.0562 | 0.025064 | 23.7309 | 0.013482 | 0.13705 | -2e-06 | -0.000129 | -0.000129 | -43.2764
2019-03-29 11:32 | 1553873512.8517 | -207 | 0.20949 | 0.000459 | -1e-06 | 100.4983 | 100.3776 | 126.8072 | 21.0092 | 0.001736 | 0.001679 | -151.811 | 0.22936 | 8.4009 | 0.002384 | 0.004002 | 0.003807 | 6.4e-05 | 0.00091 | 74.9693 | 1.3389 | 2.5474 | 100.31 | 21.39 | 31.164 | 63.3447 | 78.928 | 61.9919 | 0.007923 | 1.0534 | 0.025035 | 23.7657 | 0.010751 | 0.13644 | -1e-06 | -3.2e-05 | -3.2e-05 | -174.1232
2019-03-29 11:32 | 1553873513.8521 | -206 | 0.20948 | 0.000459 | -1e-06 | 100.4986 | 100.3602 | 127.1698 | 21.0096 | 0.001806 | 0.001671 | -151.8158 | 0.22949 | 8.4019 | 0.002359 | 0.003174 | 0.002887 | 6.2e-05 | 0.000911 | 75.0252 | 1.3713 | 2.5474 | 100.31 | 21.39 | 31.171 | 65.0284 | 78.3174 | 61.7907 | 0.007925 | 1.0541 | 0.025079 | 23.7925 | 0.012329 | 0.13653 | -2e-06 | -0.000121 | -0.000121 | -46.0235
2019-03-29 11:32 | 1553873514.8523 | -205 | 0.20949 | 0.000459 | -1e-06 | 100.4381 | 100.224 | 127.5313 | 21.009 | 0.001826 | 0.001672 | -151.8058 | 0.22975 | 8.4009 | 0.002412 | 0.00366 | 0.003466 | 6.3e-05 | 0.000908 | 75.0333 | 1.3459 | 2.5474 | 100.31 | 21.39 | 31.183 | 65.5126 | 79.6045 | 61.8159 | 0.007928 | 1.054 | 0.025113 | 23.8272 | 0.01909 | 0.13668 | -1e-06 | -3.1e-05 | -3.1e-05 | -179.4883
2019-03-29 11:32 | 1553873515.8513 | -204 | 0.2095 | 0.000458 | 0 | 100.4486 | 100.2315 | 127.5855 | 21.0107 | 0.001837 | 0.001713 | -151.8119 | 0.22995 | 8.4012 | 0.002334 | 0.004225 | 0.004094 | 6.5e-05 | 0.000916 | 75.0618 | 1.3181 | 2.5474 | 100.31 | 21.39 | 31.189 | 65.7819 | 77.7189 | 62.7987 | 0.007929 | 1.0509 | 0.025082 | 23.8672 | 0.019353 | 0.1368 | -1e-06 | -3.2e-05 | -3.2e-05 | -173.5304
2019-03-29 11:32 | 1553873516.8527 | -203 | 0.20943 | 0.000462 | -1e-06 | 100.5306 | 100.4124 | 127.5257 | 21.0097 | 0.001809 | 0.001644 | -151.8156 | 0.23061 | 8.4022 | 0.00236 | 0.004457 | 0.00423 | 6.6e-05 | 0.00092 | 75.0719 | 1.3162 | 2.5473 | 100.31 | 21.39 | 31.194 | 65.1113 | 78.3509 | 61.1272 | 0.007931 | 1.0561 | 0.025139 | 23.8021 | 0.010526 | 0.13719 | -5e-06 | -0.000352 | -0.000352 | -15.8688
2019-03-29 11:32 | 1553873517.8514 | -202 | 0.20955 | 0.000455 | -1e-06 | 100.4573 | 100.3441 | 127.5616 | 21.0101 | 0.001765 | 0.001663 | -151.8249 | 0.23053 | 8.4024 | 0.002336 | 0.003807 | 0.003643 | 6.5e-05 | 0.000919 | 75.0708 | 1.3106 | 2.5472 | 100.31 | 21.39 | 31.201 | 64.0525 | 77.7813 | 61.5966 | 0.007933 | 1.0547 | 0.025124 | 23.8222 | 0.010081 | 0.13714 | 3e-06 | 0.000299 | 0.000299 | 18.6986
2019-03-29 11:32 | 1553873518.8528 | -201 | 0.20946 | 0.000459 | -1e-06 | 100.4477 | 100.3769 | 127.6322 | 21.0107 | 0.001811 | 0.001661 | -151.8274 | 0.23081 | 8.404 | 0.002329 | 0.004084 | 0.003891 | 6.4e-05 | 0.000924 | 74.9999 | 1.3194 | 2.5475 | 100.31 | 21.39 | 31.203 | 65.1543 | 77.6158 | 61.5425 | 0.007933 | 1.0548 | 0.025133 | 23.8269 | 0.006308 | 0.13731 | -2e-06 | -9.3e-05 | -9.3e-05 | -60.2202
2019-03-29 11:32 | 1553873519.8528 | -200 | 0.20952 | 0.000462 | -1e-06 | 100.5803 | 100.5704 | 127.6292 | 21.011 | 0.001849 | 0.001715 | -151.8301 | 0.22996 | 8.4009 | 0.002299 | 0.003477 | 0.003306 | 6.4e-05 | 0.000921 | 74.996 | 1.3278 | 2.5474 | 100.31 | 21.38 | 31.202 | 66.0625 | 76.8851 | 62.8388 | 0.007928 | 1.0508 | 0.025085 | 23.8727 | 0.000882 | 0.1368 | -5e-06 | -0.000334 | -0.000334 | -16.6617
2019-03-29 11:32 | 1553873520.8514 | -199 | 0.20948 | 0.000459 | -1e-06 | 100.6132 | 100.5093 | 127.3569 | 21.011 | 0.001829 | 0.001684 | -151.835 | 0.23043 | 8.4006 | 0.002321 | 0.003584 | 0.003409 | 6e-05 | 0.00092 | 75.0294 | 1.3431 | 2.5475 | 100.31 | 21.38 | 31.201 | 65.599 | 77.4283 | 62.0928 | 0.007928 | 1.0531 | 0.025086 | 23.8207 | 0.009244 | 0.13708 | -2e-06 | -5.6e-05 | -5.6e-05 | -99.2692
2019-03-29 11:32 | 1553873521.8515 | -198 | 0.20952 | 0.000464 | -1e-06 | 100.5387 | 100.3458 | 127.2644 | 21.0109 | 0.001781 | 0.001709 | -151.8494 | 0.23022 | 8.4011 | 0.002319 | 0.003463 | 0.003215 | 6.3e-05 | 0.000918 | 75.1079 | 1.3358 | 2.5475 | 100.31 | 21.38 | 31.197 | 64.4411 | 77.3621 | 62.7154 | 0.007927 | 1.0512 | 0.025053 | 23.8342 | 0.017175 | 0.13696 | -7e-06 | -0.000497 | -0.000497 | -11.205
2019-03-29 11:32 | 1553873522.8511 | -197 | 0.20948 | 0.000458 | -1e-06 | 100.5504 | 100.3086 | 127.2179 | 21.0114 | 0.001783 | 0.00172 | -151.8328 | 0.23031 | 8.4019 | 0.002348 | 0.003628 | 0.003463 | 6.5e-05 | 0.000916 | 75.1056 | 1.3034 | 2.5474 | 100.31 | 21.38 | 31.195 | 64.4868 | 78.0656 | 62.9766 | 0.007926 | 1.0503 | 0.025039 | 23.8391 | 0.021531 | 0.13701 | 1e-06 | 0.000102 | 0.000102 | 54.5594
2019-03-29 11:32 | 1553873523.8516 | -196 | 0.20949 | 0.000457 | -1e-06 | 100.5142 | 100.5629 | 126.8184 | 21.0113 | 0.001797 | 0.001687 | -151.8126 | 0.22997 | 8.4033 | 0.002364 | 0.003623 | 0.003434 | 6.5e-05 | 0.000916 | 75.0886 | 1.3538 | 2.5475 | 100.31 | 21.38 | 31.19 | 64.8148 | 78.4437 | 62.1736 | 0.007925 | 1.0529 | 0.02503 | 23.7732 | -0.004336 | 0.13681 | 1e-06 | 0.000146 | 0.000146 | 38.1385
2019-03-29 11:32 | 1553873524.8507 | -195 | 0.20939 | 0.000459 | -1e-06 | 100.4842 | 100.4107 | 126.7222 | 21.0122 | 0.001814 | 0.001688 | -151.8131 | 0.23012 | 8.4019 | 0.00236 | 0.003405 | 0.003305 | 6.2e-05 | 0.000913 | 75.0611 | 1.3547 | 2.5475 | 100.31 | 21.38 | 31.186 | 65.2364 | 78.3501 | 62.2081 | 0.007924 | 1.0527 | 0.025019 | 23.7654 | 0.006549 | 0.1369 | -2e-06 | -9.9e-05 | -9.9e-05 | -56.5049
2019-03-29 11:32 | 1553873525.8519 | -194 | 0.20944 | 0.000463 | -1e-06 | 100.4624 | 100.1848 | 126.8344 | 21.0121 | 0.00181 | 0.001715 | -151.8114 | 0.2303 | 8.4051 | 0.002333 | 0.003919 | 0.003653 | 6.2e-05 | 0.00091 | 75.0705 | 1.3327 | 2.5474 | 100.31 | 21.38 | 31.184 | 65.1393 | 77.7116 | 62.861 | 0.007923 | 1.0507 | 0.025006 | 23.799 | 0.024753 | 0.137 | -6e-06 | -0.000429 | -0.000429 | -12.9938
2019-03-29 11:32 | 1553873526.851 | -193 | 0.20949 | 0.000462 | -1e-06 | 100.4983 | 100.4083 | 126.73 | 21.0121 | 0.001774 | 0.001714 | -151.803 | 0.23033 | 8.4025 | 0.002321 | 0.003915 | 0.003765 | 6.1e-05 | 0.000909 | 75.0121 | 1.3134 | 2.5475 | 100.31 | 21.38 | 31.178 | 64.283 | 77.4099 | 62.8305 | 0.007922 | 1.0508 | 0.024996 | 23.7881 | 0.008014 | 0.13702 | -4e-06 | -0.000263 | -0.000263 | -21.1955
2019-03-29 11:32 | 1553873527.8517 | -192 | 0.20943 | 0.000456 | -1e-06 | 100.4579 | 100.4673 | 126.7696 | 21.0133 | 0.001822 | 0.001733 | -151.7936 | 0.23043 | 8.4017 | 0.00235 | 0.003463 | 0.003351 | 6.2e-05 | 0.000908 | 75.0947 | 1.3357 | 2.5475 | 100.31 | 21.38 | 31.175 | 65.4201 | 78.1088 | 63.284 | 0.007921 | 1.0494 | 0.024983 | 23.8079 | -0.000841 | 0.13708 | 2e-06 | 0.000226 | 0.000226 | 24.7245
2019-03-29 11:32 | 1553873528.8506 | -191 | 0.2095 | 0.00046 | -1e-06 | 100.5648 | 100.2189 | 126.6863 | 21.0133 | 0.001774 | 0.001659 | -151.7828 | 0.22984 | 8.4035 | 0.002275 | 0.004037 | 0.003788 | 6.1e-05 | 0.000908 | 75.1036 | 1.3358 | 2.5474 | 100.31 | 21.37 | 31.171 | 64.264 | 76.3121 | 61.5036 | 0.007915 | 1.055 | 0.025042 | 23.737 | 0.030816 | 0.13673 | -3e-06 | -0.000156 | -0.000156 | -35.6734
2019-03-29 11:32 | 1553873529.8507 | -190 | 0.20958 | 0.000462 | -1e-06 | 100.4606 | 100.3527 | 126.9328 | 21.0139 | 0.001841 | 0.001694 | -151.7734 | 0.2297 | 8.4014 | 0.00233 | 0.003752 | 0.003554 | 6.2e-05 | 0.000904 | 75.127 | 1.3398 | 2.5475 | 100.31 | 21.37 | 31.168 | 65.8921 | 77.6425 | 62.3362 | 0.007914 | 1.0523 | 0.025035 | 23.7897 | 0.009614 | 0.13665 | -5e-06 | -0.000354 | -0.000354 | -15.7225
2019-03-29 11:32 | 1553873530.8511 | -189 | 0.20916 | 0.000462 | -1e-06 | 100.4908 | 100.3049 | 127.2062 | 21.0137 | 0.001816 | 0.001696 | -151.7669 | 0.2305 | 8.4024 | 0.002398 | 0.00308 | 0.002879 | 6.3e-05 | 0.000908 | 75.1108 | 1.3369 | 2.5474 | 100.31 | 21.37 | 31.166 | 65.2887 | 79.27 | 62.404 | 0.007914 | 1.0521 | 0.025059 | 23.8177 | 0.016565 | 0.13712 | -5e-06 | -0.000338 | -0.000338 | -16.5258
2019-03-29 11:32 | 1553873531.8509 | -188 | 0.20959 | 0.00046 | -1e-06 | 100.4699 | 100.4155 | 127.3109 | 21.0146 | 0.0018 | 0.001717 | -151.7771 | 0.22958 | 8.4034 | 0.002427 | 0.003964 | 0.003779 | 6.2e-05 | 0.000908 | 75.0737 | 1.3643 | 2.5474 | 100.31 | 21.37 | 31.165 | 64.9063 | 79.9601 | 62.9088 | 0.007914 | 1.0505 | 0.025051 | 23.8454 | 0.004845 | 0.13658 | -3e-06 | -0.000146 | -0.000146 | -37.9684
2019-03-29 11:32 | 1553873532.8508 | -187 | 0.20938 | 0.000457 | -1e-06 | 100.4342 | 100.1878 | 127.1006 | 21.0148 | 0.0018 | 0.001693 | -151.7691 | 0.23081 | 8.4018 | 0.002376 | 0.00431 | 0.004124 | 6.2e-05 | 0.00091 | 75.069 | 1.329 | 2.5474 | 100.31 | 21.37 | 31.163 | 64.8985 | 78.7425 | 62.334 | 0.007913 | 1.0523 | 0.025051 | 23.8053 | 0.021973 | 0.13731 | 1e-06 | 0.000157 | 0.000157 | 35.6586
2019-03-29 11:32 | 1553873533.8518 | -186 | 0.20979 | 0.00046 | -1e-06 | 100.445 | 100.3909 | 127.1353 | 21.0153 | 0.001786 | 0.001693 | -151.7824 | 0.23159 | 8.4037 | 0.002387 | 0.003439 | 0.003311 | 6.2e-05 | 0.00091 | 75.1121 | 1.3401 | 2.5474 | 100.31 | 21.37 | 31.162 | 64.5696 | 79.0142 | 62.3337 | 0.007913 | 1.0523 | 0.025055 | 23.8085 | 0.004822 | 0.13777 | -2e-06 | -0.000117 | -0.000117 | -47.8461
2019-03-29 11:32 | 1553873534.8513 | -185 | 0.20938 | 0.000458 | -1e-06 | 100.4705 | 100.2909 | 127.225 | 21.015 | 0.001775 | 0.001628 | -151.7753 | 0.2307 | 8.401 | 0.002316 | 0.003503 | 0.003289 | 6.4e-05 | 0.000909 | 75.1209 | 1.3499 | 2.5474 | 100.31 | 21.36 | 31.161 | 64.2959 | 77.299 | 60.7524 | 0.007908 | 1.0573 | 0.025123 | 23.7607 | 0.016001 | 0.13724 | 1e-06 | 9.9e-05 | 9.9e-05 | 56.2008
2019-03-29 11:32 | 1553873535.8519 | -184 | 0.20945 | 0.000458 | -1e-06 | 100.4968 | 100.3472 | 126.8541 | 21.0152 | 0.001838 | 0.001672 | -151.7778 | 0.23067 | 8.4024 | 0.002345 | 0.00396 | 0.003797 | 6.1e-05 | 0.000906 | 75.1008 | 1.3226 | 2.5475 | 100.31 | 21.36 | 31.161 | 65.8034 | 77.9977 | 61.8212 | 0.007908 | 1.054 | 0.025046 | 23.764 | 0.013325 | 0.13723 | 0 | 6.2e-05 | 6.2e-05 | 90.1284
2019-03-29 11:32 | 1553873536.8519 | -183 | 0.20951 | 0.000458 | -1e-06 | 100.4489 | 100.3144 | 127.0234 | 21.0166 | 0.001849 | 0.001656 | -151.7925 | 0.23084 | 8.4018 | 0.002411 | 0.003781 | 0.003691 | 6.1e-05 | 0.000906 | 75.1238 | 1.3339 | 2.5475 | 100.31 | 21.36 | 31.162 | 66.0694 | 79.5792 | 61.4302 | 0.007908 | 1.0552 | 0.025078 | 23.766 | 0.011982 | 0.13733 | 0 | 3.7e-05 | 3.7e-05 | 150.2061
2019-03-29 11:32 | 1553873537.8519 | -182 | 0.20957 | 0.000458 | -1e-06 | 100.4546 | 100.3882 | 127.1095 | 21.0157 | 0.001784 | 0.001749 | -151.8104 | 0.23046 | 8.403 | 0.002392 | 0.003958 | 0.003672 | 6.2e-05 | 0.000906 | 75.0911 | 1.3336 | 2.5474 | 100.31 | 21.36 | 31.164 | 64.52 | 79.1355 | 63.683 | 0.007909 | 1.0481 | 0.025002 | 23.8539 | 0.005918 | 0.1371 | 0 | 7e-05 | 7e-05 | 79.7887
2019-03-29 11:32 | 1553873538.8511 | -181 | 0.20941 | 0.000462 | -1e-06 | 100.4827 | 100.2155 | 127.2949 | 21.0159 | 0.001792 | 0.001685 | -151.8148 | 0.23001 | 8.4042 | 0.002353 | 0.002499 | 0.002302 | 6.4e-05 | 0.000908 | 75.1277 | 1.3398 | 2.5475 | 100.31 | 21.36 | 31.164 | 64.7044 | 78.1863 | 62.1217 | 0.007909 | 1.053 | 0.025078 | 23.8159 | 0.023823 | 0.13684 | -5e-06 | -0.000315 | -0.000315 | -17.6589
2019-03-29 11:32 | 1553873539.8512 | -180 | 0.20955 | 0.000458 | -1e-06 | 100.4525 | 100.4144 | 126.9494 | 21.0169 | 0.001823 | 0.001713 | -151.8373 | 0.23063 | 8.4023 | 0.002437 | 0.003225 | 0.002953 | 6.4e-05 | 0.000906 | 75.1257 | 1.3343 | 2.5474 | 100.31 | 21.35 | 31.166 | 65.4508 | 80.2166 | 62.8079 | 0.007904 | 1.0509 | 0.025019 | 23.8079 | 0.00339 | 0.1372 | 1e-06 | 0.000101 | 0.000101 | 55.4957
2019-03-29 11:32 | 1553873540.8511 | -179 | 0.20948 | 0.000463 | -1e-06 | 100.4759 | 100.2349 | 126.8001 | 21.0173 | 0.001831 | 0.001679 | -151.8694 | 0.23082 | 8.4011 | 0.002369 | 0.003324 | 0.00316 | 6.4e-05 | 0.000905 | 75.1121 | 1.345 | 2.5474 | 100.31 | 21.35 | 31.168 | 65.6486 | 78.5749 | 61.9802 | 0.007905 | 1.0535 | 0.025035 | 23.7646 | 0.021474 | 0.13731 | -7e-06 | -0.000463 | -0.000463 | -12.062
2019-03-29 11:32 | 1553873541.8512 | -178 | 0.20947 | 0.000456 | -1e-06 | 100.4229 | 100.2342 | 126.5973 | 21.0171 | 0.001828 | 0.001718 | -151.8533 | 0.23066 | 8.4054 | 0.002397 | 0.00339 | 0.003256 | 6.1e-05 | 0.00091 | 75.1067 | 1.3622 | 2.5474 | 100.31 | 21.35 | 31.172 | 65.5782 | 79.2531 | 62.9191 | 0.007906 | 1.0505 | 0.02498 | 23.7788 | 0.016817 | 0.13722 | 3e-06 | 0.000283 | 0.000283 | 19.745
2019-03-29 11:32 | 1553873542.8516 | -177 | 0.20946 | 0.000458 | 0 | 100.4235 | 100.2377 | 126.7025 | 21.0176 | 0.0018 | 0.001662 | -151.87 | 0.23059 | 8.4021 | 0.002351 | 0.003423 | 0.003219 | 6.3e-05 | 0.000905 | 75.0286 | 1.3657 | 2.5475 | 100.31 | 21.34 | 31.172 | 64.9068 | 78.1442 | 61.5914 | 0.007901 | 1.0547 | 0.02504 | 23.7417 | 0.016566 | 0.13718 | -1e-06 | -2.3e-05 | -2.3e-05 | -247.3306
2019-03-29 11:32 | 1553873543.8512 | -176 | 0.20945 | 0.000461 | -1e-06 | 100.4283 | 100.2919 | 126.7243 | 21.0174 | 0.001826 | 0.001659 | -151.8715 | 0.23099 | 8.4027 | 0.002397 | 0.003918 | 0.003764 | 6.5e-05 | 0.000913 | 75.0708 | 1.2899 | 2.5473 | 100.31 | 21.34 | 31.171 | 65.5224 | 79.2423 | 61.5076 | 0.007901 | 1.0549 | 0.025045 | 23.7407 | 0.012151 | 0.13742 | -4e-06 | -0.000286 | -0.000286 | -19.5636
2019-03-29 11:32 | 1553873544.8503 | -175 | 0.20948 | 0.000459 | -1e-06 | 100.4169 | 100.3557 | 126.6739 | 21.0175 | 0.001719 | 0.001698 | -151.8767 | 0.23099 | 8.402 | 0.002322 | 0.003983 | 0.003751 | 6.2e-05 | 0.000911 | 75.0297 | 1.3571 | 2.5475 | 100.31 | 21.34 | 31.174 | 62.9529 | 77.4519 | 62.4483 | 0.007901 | 1.052 | 0.025005 | 23.7693 | 0.005448 | 0.13742 | 0 | 2.9e-05 | 2.9e-05 | 192.0368
2019-03-29 11:32 | 1553873545.8513 | -174 | 0.20945 | 0.000462 | -1e-06 | 100.4354 | 100.2646 | 127.1398 | 21.0178 | 0.001802 | 0.001714 | -151.8794 | 0.22984 | 8.4035 | 0.002353 | 0.003702 | 0.003546 | 6.5e-05 | 0.000905 | 75.068 | 1.3687 | 2.5475 | 100.31 | 21.34 | 31.174 | 64.9506 | 78.2028 | 62.8291 | 0.007901 | 1.0508 | 0.025037 | 23.8265 | 0.015226 | 0.13673 | -4e-06 | -0.000288 | -0.000288 | -19.3397
2019-03-29 11:32 | 1553873546.8518 | -173 | 0.2094 | 0.000459 | -1e-06 | 100.4618 | 100.4322 | 126.8952 | 21.0188 | 0.001776 | 0.001755 | -151.8656 | 0.23024 | 8.4018 | 0.002286 | 0.004109 | 0.003901 | 6.7e-05 | 0.000913 | 75.0423 | 1.3855 | 2.5475 | 100.31 | 21.34 | 31.174 | 64.3184 | 76.5839 | 63.8321 | 0.007901 | 1.0477 | 0.024975 | 23.8391 | 0.002636 | 0.13697 | -1e-06 | -9e-06 | -9e-06 | -637.3422
2019-03-29 11:32 | 1553873547.8505 | -172 | 0.20945 | 0.000461 | -1e-06 | 100.4672 | 100.3752 | 126.6016 | 21.0204 | 0.001863 | 0.001729 | -151.8493 | 0.23085 | 8.4039 | 0.002335 | 0.004466 | 0.004406 | 6.5e-05 | 0.000906 | 75.0396 | 1.3361 | 2.5474 | 100.31 | 21.34 | 31.178 | 66.4158 | 77.7688 | 63.2026 | 0.007902 | 1.0496 | 0.02497 | 23.7893 | 0.008193 | 0.13733 | -3e-06 | -0.000184 | -0.000184 | -30.3604
2019-03-29 11:32 | 1553873548.8506 | -171 | 0.20947 | 0.000454 | 0 | 100.4809 | 100.4619 | 126.2176 | 21.0188 | 0.001791 | 0.001704 | -151.853 | 0.23045 | 8.4013 | 0.002311 | 0.003822 | 0.003556 | 6.2e-05 | 0.000908 | 75.0521 | 1.3488 | 2.5473 | 100.31 | 21.34 | 31.178 | 64.683 | 77.1913 | 62.5878 | 0.007902 | 1.0516 | 0.024955 | 23.7314 | 0.001697 | 0.13709 | 5e-06 | 0.000417 | 0.000417 | 13.3669
2019-03-29 11:32 | 1553873549.8519 | -170 | 0.20951 | 0.000458 | -1e-06 | 100.4759 | 100.3086 | 126.4976 | 21.0197 | 0.001821 | 0.001692 | -151.8428 | 0.23033 | 8.4018 | 0.002363 | 0.003493 | 0.003313 | 6.7e-05 | 0.000911 | 75.0141 | 1.3274 | 2.5474 | 100.31 | 21.33 | 31.18 | 65.4051 | 78.4338 | 62.2966 | 0.007898 | 1.0525 | 0.024993 | 23.7474 | 0.014898 | 0.13702 | 1e-06 | 0.000128 | 0.000128 | 43.4982
2019-03-29 11:32 | 1553873550.8509 | -169 | 0.20946 | 0.000461 | -1e-06 | 100.4127 | 100.4523 | 126.58 | 21.0199 | 0.001763 | 0.001708 | -151.8218 | 0.23145 | 8.4031 | 0.002303 | 0.003768 | 0.003595 | 6.5e-05 | 0.000915 | 75.0171 | 1.3312 | 2.5474 | 100.31 | 21.33 | 31.181 | 64.0196 | 77.0028 | 62.696 | 0.007898 | 1.0512 | 0.024987 | 23.7693 | -0.003529 | 0.13769 | 0.000173 | 0.013034 | 0.013034 | 0.42995
2019-03-29 11:33 | 1553873551.8516 | -168 | 0.20948 | 0.000457 | -1e-06 | 100.4627 | 100.2261 | 126.6868 | 21.0201 | 0.001758 | 0.001675 | -151.8265 | 0.23118 | 8.4028 | 0.00239 | 0.003332 | 0.00313 | 6.8e-05 | 0.000913 | 75.0306 | 1.3507 | 2.5475 | 100.31 | 21.33 | 31.181 | 63.905 | 79.0868 | 61.8845 | 0.007898 | 1.0538 | 0.025027 | 23.7506 | 0.021093 | 0.13753 | 5e-05 | 0.00383 | 0.00383 | 1.4615
2019-03-29 11:33 | 1553873552.8505 | -167 | 0.20946 | 0.000462 | -1e-06 | 100.4543 | 100.3574 | 126.5645 | 21.0201 | 0.001787 | 0.001644 | -151.8176 | 0.23088 | 8.404 | 0.002429 | 0.003604 | 0.003404 | 6.3e-05 | 0.000911 | 75.0401 | 1.3422 | 2.5474 | 100.31 | 21.33 | 31.181 | 64.5883 | 80.0224 | 61.1492 | 0.007898 | 1.0561 | 0.025043 | 23.7131 | 0.008628 | 0.13735 | 0.000143 | 0.010805 | 0.010805 | 0.51735
2019-03-29 11:33 | 1553873553.85 | -166 | 0.20952 | 0.000463 | -1e-06 | 100.4651 | 100.2056 | 126.5915 | 21.0198 | 0.001759 | 0.001708 | -151.8127 | 0.23052 | 8.4051 | 0.002395 | 0.00364 | 0.003433 | 6.7e-05 | 0.000909 | 75.0855 | 1.3373 | 2.5474 | 100.31 | 21.33 | 31.181 | 63.9096 | 79.2034 | 62.6943 | 0.007898 | 1.0512 | 0.024988 | 23.7703 | 0.023134 | 0.13714 | -0.000222 | -0.016603 | -0.016603 | -0.33618
2019-03-29 11:33 | 1553873554.8506 | -165 | 0.20945 | 0.000458 | -1e-06 | 100.4229 | 100.2783 | 126.7389 | 21.0202 | 0.001873 | 0.001645 | -151.8055 | 0.23023 | 8.4014 | 0.00235 | 0.003727 | 0.003526 | 6.5e-05 | 0.000911 | 75.0688 | 1.3374 | 2.5475 | 100.31 | 21.33 | 31.184 | 66.651 | 78.1172 | 61.1673 | 0.007899 | 1.056 | 0.025059 | 23.73 | 0.012889 | 0.13697 | 0.000236 | 0.017818 | 0.017818 | 0.31286
2019-03-29 11:33 | 1553873555.8514 | -164 | 0.2095 | 0.00046 | -1e-06 | 100.395 | 100.4762 | 126.5572 | 21.0207 | 0.001789 | 0.001723 | -151.8017 | 0.23077 | 8.4036 | 0.002365 | 0.003473 | 0.003329 | 6.5e-05 | 0.000911 | 75.0972 | 1.3464 | 2.5474 | 100.31 | 21.32 | 31.185 | 64.6294 | 78.4754 | 63.0559 | 0.007894 | 1.0501 | 0.024971 | 23.7799 | -0.00723 | 0.13728 | -7.9e-05 | -0.005898 | -0.005898 | -0.94739
2019-03-29 11:33 | 1553873556.8501 | -163 | 0.20944 | 0.00046 | -1e-06 | 100.4774 | 100.3011 | 126.887 | 21.0212 | 0.001795 | 0.001742 | -151.7996 | 0.23078 | 8.404 | 0.002379 | 0.003712 | 0.003564 | 6.8e-05 | 0.000907 | 75.0432 | 1.3673 | 2.5474 | 100.31 | 21.32 | 31.185 | 64.7938 | 78.815 | 63.4987 | 0.007894 | 1.0487 | 0.024987 | 23.8265 | 0.015701 | 0.13729 | 0.00026 | 0.01964 | 0.01964 | 0.28452
2019-03-29 11:33 | 1553873557.851 | -162 | 0.2095 | 0.000459 | -1e-06 | 100.4342 | 100.238 | 126.7379 | 21.0222 | 0.001758 | 0.001711 | -151.803 | 0.23161 | 8.4039 | 0.002314 | 0.003301 | 0.003064 | 6.8e-05 | 0.000909 | 75.0392 | 1.356 | 2.5475 | 100.31 | 21.32 | 31.188 | 63.897 | 77.2561 | 62.7693 | 0.007895 | 1.051 | 0.024999 | 23.7867 | 0.017495 | 0.13778 | -5.4e-05 | -0.004 | -0.004 | -1.4021
2019-03-29 11:33 | 1553873558.8508 | -161 | 0.20951 | 0.000456 | -1e-06 | 100.4881 | 100.265 | 126.6703 | 21.0217 | 0.001832 | 0.001693 | -151.8078 | 0.2301 | 8.3998 | 0.002359 | 0.003886 | 0.003711 | 6.6e-05 | 0.000904 | 74.9915 | 1.3171 | 2.5474 | 100.31 | 21.32 | 31.188 | 65.6796 | 78.3423 | 62.3393 | 0.007895 | 1.0523 | 0.025009 | 23.7652 | 0.019887 | 0.13689 | -0.000105 | -0.007866 | -0.007866 | -0.70824
2019-03-29 11:33 | 1553873559.8507 | -160 | 0.20949 | 0.000459 | -1e-06 | 100.4986 | 100.4486 | 126.6886 | 21.0222 | 0.001855 | 0.001713 | -151.8233 | 0.23094 | 8.4006 | 0.002393 | 0.003421 | 0.00327 | 6.7e-05 | 0.000901 | 74.9872 | 1.3582 | 2.5474 | 100.31 | 21.32 | 31.191 | 66.2207 | 79.1628 | 62.8224 | 0.007896 | 1.0508 | 0.024993 | 23.784 | 0.004454 | 0.13738 | 1.7e-05 | 0.001335 | 0.001335 | 4.188
2019-03-29 11:33 | 1553873560.8508 | -159 | 0.2095 | 0.000458 | -1e-06 | 100.5249 | 100.324 | 126.9292 | 21.0209 | 0.001776 | 0.001676 | -151.8454 | 0.22997 | 8.4008 | 0.002332 | 0.003885 | 0.003661 | 6.5e-05 | 0.000895 | 75.0304 | 1.3376 | 2.5475 | 100.31 | 21.32 | 31.193 | 64.3268 | 77.6991 | 61.9282 | 0.007896 | 1.0536 | 0.02505 | 23.7748 | 0.017898 | 0.13681 | -6e-05 | -0.004479 | -0.004479 | -1.2433
2019-03-29 11:33 | 1553873561.8508 | -158 | 0.20952 | 0.00046 | -1e-06 | 100.4666 | 100.5909 | 127.0687 | 21.0226 | 0.001833 | 0.001684 | -151.8562 | 0.2297 | 8.4026 | 0.002383 | 0.003811 | 0.003717 | 6.7e-05 | 0.000893 | 75.0263 | 1.3221 | 2.5475 | 100.31 | 21.32 | 31.194 | 65.7108 | 78.909 | 62.1268 | 0.007897 | 1.053 | 0.025056 | 23.795 | -0.011061 | 0.13664 | -0.000201 | -0.015082 | -0.015082 | -0.36875
2019-03-29 11:33 | 1553873562.8513 | -157 | 0.20951 | 0.00046 | -1e-06 | 100.5504 | 100.4254 | 127.1327 | 21.0239 | 0.001845 | 0.001667 | -151.8455 | 0.23064 | 8.4041 | 0.002398 | 0.003627 | 0.003424 | 6.6e-05 | 0.000894 | 75.0565 | 1.3166 | 2.5475 | 100.3 | 21.32 | 31.198 | 65.9938 | 79.2806 | 61.6953 | 0.007899 | 1.0544 | 0.025078 | 23.7856 | 0.011129 | 0.13721 | -0.000151 | -0.011358 | -0.011358 | -0.49166
2019-03-29 11:33 | 1553873563.8499 | -156 | 0.20949 | 0.00046 | 0 | 100.6153 | 100.6008 | 126.971 | 21.0228 | 0.001774 | 0.001687 | -151.8553 | 0.22954 | 8.4023 | 0.002343 | 0.004002 | 0.003832 | 6.4e-05 | 0.00089 | 75.1175 | 1.3158 | 2.5474 | 100.3 | 21.32 | 31.2 | 64.2795 | 77.9638 | 62.1778 | 0.007899 | 1.0528 | 0.025044 | 23.7876 | 0.001294 | 0.13655 | -1.7e-05 | -0.001279 | -0.001279 | -4.3444
2019-03-29 11:33 | 1553873564.8514 | -155 | 0.20949 | 0.000461 | 0 | 100.5875 | 100.4673 | 126.8328 | 21.0227 | 0.001781 | 0.001741 | -151.852 | 0.23105 | 8.4023 | 0.002355 | 0.003877 | 0.00374 | 6.6e-05 | 0.000889 | 75.0706 | 1.3674 | 2.5474 | 100.3 | 21.32 | 31.203 | 64.4502 | 78.2364 | 63.4838 | 0.0079 | 1.0488 | 0.024982 | 23.8209 | 0.010692 | 0.13745 | -6e-06 | -0.000392 | -0.000392 | -14.2646
2019-03-29 11:33 | 1553873565.8509 | -154 | 0.20947 | 0.00046 | -1e-06 | 100.5543 | 100.6028 | 127.3048 | 21.0231 | 0.001806 | 0.001687 | -151.8436 | 0.23089 | 8.4048 | 0.002339 | 0.004291 | 0.00407 | 6.9e-05 | 0.000887 | 75.0418 | 1.3564 | 2.5474 | 100.3 | 21.31 | 31.209 | 65.0541 | 77.8728 | 62.1913 | 0.007896 | 1.0528 | 0.025077 | 23.8193 | -0.004318 | 0.13736 | 8.4e-05 | 0.006386 | 0.006386 | 0.87548
2019-03-29 11:33 | 1553873566.851 | -153 | 0.20949 | 0.000455 | -1e-06 | 100.6489 | 100.652 | 127.3107 | 21.0231 | 0.001716 | 0.001703 | -151.8452 | 0.23162 | 8.4034 | 0.002389 | 0.004519 | 0.004409 | 6.6e-05 | 0.000887 | 75.0409 | 1.3363 | 2.5473 | 100.3 | 21.31 | 31.211 | 62.8867 | 79.0631 | 62.5679 | 0.007897 | 1.0516 | 0.025063 | 23.8333 | -0.000276 | 0.13779 | -3e-05 | -0.002157 | -0.002157 | -2.6004
2019-03-29 11:33 | 1553873567.85 | -152 | 0.2095 | 0.000458 | -1e-06 | 100.7168 | 100.5144 | 127.5987 | 21.0242 | 0.001825 | 0.001714 | -151.8507 | 0.23084 | 8.4014 | 0.002302 | 0.003737 | 0.003591 | 6.8e-05 | 0.000896 | 75.0699 | 1.3211 | 2.5474 | 100.3 | 21.31 | 31.217 | 65.5148 | 76.9649 | 62.8282 | 0.007899 | 1.0508 | 0.025082 | 23.8695 | 0.017991 | 0.13733 | -9.6e-05 | -0.007166 | -0.007166 | -0.77994
2019-03-29 11:33 | 1553873568.85 | -151 | 0.20948 | 0.000459 | -1e-06 | 100.6908 | 100.6615 | 127.4083 | 21.0241 | 0.001853 | 0.00171 | -151.8382 | 0.23083 | 8.4025 | 0.002312 | 0.003903 | 0.003698 | 6.8e-05 | 0.000898 | 75.0848 | 1.303 | 2.5474 | 100.3 | 21.31 | 31.218 | 66.1906 | 77.2103 | 62.7383 | 0.007899 | 1.0511 | 0.025067 | 23.8485 | 0.002598 | 0.13732 | 5e-05 | 0.00384 | 0.00384 | 1.4555
2019-03-29 11:33 | 1553873569.8495 | -150 | 0.20951 | 0.000457 | -1e-06 | 100.7303 | 100.6052 | 126.8715 | 21.0248 | 0.001829 | 0.001682 | -151.816 | 0.23061 | 8.399 | 0.002353 | 0.003999 | 0.003782 | 6.7e-05 | 0.000903 | 75.0609 | 1.3163 | 2.5474 | 100.3 | 21.31 | 31.221 | 65.5999 | 78.1901 | 62.068 | 0.0079 | 1.0532 | 0.025039 | 23.7744 | 0.011112 | 0.13719 | -0.000119 | -0.00888 | -0.00888 | -0.62877
2019-03-29 11:33 | 1553873570.8498 | -149 | 0.20946 | 0.000458 | 0 | 100.6725 | 100.5786 | 126.4398 | 21.0243 | 0.001804 | 0.001696 | -151.7911 | 0.23092 | 8.4021 | 0.002329 | 0.003513 | 0.003425 | 6.9e-05 | 0.000907 | 75.0556 | 1.3118 | 2.5475 | 100.3 | 21.3 | 31.225 | 65.0107 | 77.6128 | 62.3972 | 0.007896 | 1.0521 | 0.024984 | 23.7456 | 0.008348 | 0.13737 | 0.000191 | 0.014407 | 0.014407 | 0.38808
2019-03-29 11:33 | 1553873571.8492 | -148 | 0.20951 | 0.000456 | -1e-06 | 100.736 | 100.7492 | 126.5797 | 21.0257 | 0.001814 | 0.001682 | -151.7743 | 0.23059 | 8.4049 | 0.002344 | 0.003134 | 0.002997 | 6.8e-05 | 0.000907 | 75.0563 | 1.3404 | 2.5473 | 100.3 | 21.3 | 31.227 | 65.2541 | 77.9891 | 62.0734 | 0.007896 | 1.0532 | 0.02501 | 23.7472 | -0.001178 | 0.13718 | -0.000112 | -0.008373 | -0.008373 | -0.66679
2019-03-29 11:33 | 1553873572.8492 | -147 | 0.2095 | 0.00046 | -1e-06 | 100.7479 | 100.6892 | 126.9335 | 21.026 | 0.001827 | 0.001701 | -151.7819 | 0.22992 | 8.4038 | 0.002355 | 0.003715 | 0.003583 | 6.8e-05 | 0.000909 | 75.0453 | 1.3942 | 2.5475 | 100.3 | 21.3 | 31.23 | 65.5534 | 78.2607 | 62.5339 | 0.007897 | 1.0517 | 0.025027 | 23.7967 | 0.005219 | 0.13678 | -5.1e-05 | -0.003759 | -0.003759 | -1.4809
2019-03-29 11:33 | 1553873573.8506 | -146 | 0.20948 | 0.000452 | -1e-06 | 100.7518 | 100.5851 | 127.2785 | 21.0257 | 0.001789 | 0.001689 | -151.7742 | 0.23059 | 8.4022 | 0.002386 | 0.003627 | 0.003485 | 6.5e-05 | 0.000908 | 75.1119 | 1.3582 | 2.5475 | 100.3 | 21.3 | 31.231 | 64.6469 | 78.9909 | 62.2408 | 0.007897 | 1.0526 | 0.025072 | 23.8186 | 0.014816 | 0.13718 | 3.9e-05 | 0.00299 | 0.00299 | 1.8674
2019-03-29 11:33 | 1553873574.8493 | -145 | 0.2095 | 0.000455 | -1e-06 | 100.7734 | 100.7223 | 127.1337 | 21.0263 | 0.001822 | 0.001668 | -151.7828 | 0.23017 | 8.4024 | 0.00237 | 0.003281 | 0.003077 | 6.4e-05 | 0.000909 | 75.1187 | 1.3666 | 2.5474 | 100.3 | 21.3 | 31.231 | 65.4334 | 78.6132 | 61.7334 | 0.007897 | 1.0542 | 0.025077 | 23.7871 | 0.004537 | 0.13693 | -6.2e-05 | -0.004576 | -0.004576 | -1.2179
2019-03-29 11:33 | 1553873575.8492 | -144 | 0.20949 | 0.000456 | -1e-06 | 100.7261 | 100.6731 | 126.7426 | 21.0261 | 0.001818 | 0.001661 | -151.8004 | 0.23182 | 8.403 | 0.002376 | 0.003865 | 0.003766 | 6.6e-05 | 0.000913 | 75.1072 | 1.3433 | 2.5474 | 100.3 | 21.3 | 31.227 | 65.3428 | 78.7457 | 61.5691 | 0.007896 | 1.0548 | 0.025045 | 23.7446 | 0.004704 | 0.13791 | -4e-06 | -0.000249 | -0.000249 | -22.5511
2019-03-29 11:33 | 1553873576.8494 | -143 | 0.20949 | 0.000453 | -1e-06 | 100.7919 | 100.7868 | 126.4227 | 21.0257 | 0.001831 | 0.001715 | -151.8031 | 0.23096 | 8.4013 | 0.002406 | 0.003495 | 0.003269 | 6.9e-05 | 0.000908 | 75.0999 | 1.3766 | 2.5474 | 100.3 | 21.3 | 31.225 | 65.6536 | 79.4813 | 62.8571 | 0.007896 | 1.0507 | 0.024965 | 23.7602 | 0.000458 | 0.1374 | -3.5e-05 | -0.002541 | -0.002541 | -2.2007
2019-03-29 11:33 | 1553873577.8498 | -142 | 0.20953 | 0.000455 | -1e-06 | 100.8204 | 100.7404 | 126.3335 | 21.0266 | 0.001818 | 0.001716 | -151.7919 | 0.23085 | 8.4005 | 0.002377 | 0.003766 | 0.003547 | 6.9e-05 | 0.000907 | 75.0931 | 1.3751 | 2.5474 | 100.3 | 21.3 | 31.22 | 65.3432 | 78.7733 | 62.89 | 0.007894 | 1.0506 | 0.024955 | 23.753 | 0.007101 | 0.13733 | -0.000283 | -0.021181 | -0.021181 | -0.26389
2019-03-29 11:33 | 1553873578.8493 | -141 | 0.20949 | 0.000453 | -1e-06 | 100.8847 | 100.9438 | 126.3131 | 21.0269 | 0.00178 | 0.001694 | -151.7855 | 0.23153 | 8.404 | 0.002371 | 0.003721 | 0.003557 | 6.8e-05 | 0.000907 | 75.0805 | 1.3597 | 2.5474 | 100.3 | 21.3 | 31.217 | 64.4185 | 78.6249 | 62.3544 | 0.007894 | 1.0523 | 0.024973 | 23.7322 | -0.005233 | 0.13774 | -2.3e-05 | -0.001672 | -0.001672 | -3.3525
2019-03-29 11:33 | 1553873579.8495 | -140 | 0.20949 | 0.000462 | -1e-06 | 100.8 | 100.7684 | 126.4436 | 21.0261 | 0.001789 | 0.001695 | -151.7887 | 0.23112 | 8.4042 | 0.002417 | 0.004109 | 0.003956 | 6.5e-05 | 0.000908 | 75.1146 | 1.3792 | 2.5474 | 100.3 | 21.3 | 31.215 | 64.6456 | 79.7506 | 62.3857 | 0.007893 | 1.0522 | 0.024985 | 23.7455 | 0.002811 | 0.13749 | 6e-06 | 0.000523 | 0.000523 | 10.7077
2019-03-29 11:33 | 1553873580.8499 | -139 | 0.20951 | 0.000458 | -1e-06 | 100.8868 | 100.8387 | 126.387 | 21.0279 | 0.001765 | 0.001682 | -151.7946 | 0.23056 | 8.4015 | 0.002332 | 0.003562 | 0.003348 | 6.3e-05 | 0.000911 | 75.1318 | 1.3584 | 2.5474 | 100.3 | 21.29 | 31.213 | 64.0719 | 77.7027 | 62.0787 | 0.007888 | 1.0531 | 0.024991 | 23.7293 | 0.004272 | 0.13716 | -0.000133 | -0.009983 | -0.009983 | -0.55917
2019-03-29 11:33 | 1553873581.8505 | -138 | 0.20953 | 0.000455 | -1e-06 | 100.8365 | 100.7608 | 126.1842 | 21.0285 | 0.001795 | 0.001711 | -151.799 | 0.23072 | 8.402 | 0.002393 | 0.003772 | 0.003604 | 7e-05 | 0.000913 | 75.0686 | 1.3447 | 2.5474 | 100.3 | 21.29 | 31.213 | 64.7821 | 79.155 | 62.7621 | 0.007888 | 1.051 | 0.024945 | 23.7345 | 0.006717 | 0.13726 | -0.000242 | -0.018102 | -0.018102 | -0.3086
2019-03-29 11:33 | 1553873582.8496 | -137 | 0.2095 | 0.00046 | -1e-06 | 100.9165 | 100.8803 | 126.2394 | 21.0275 | 0.001899 | 0.001699 | -151.8064 | 0.23042 | 8.4032 | 0.002428 | 0.003037 | 0.002903 | 6.6e-05 | 0.000912 | 75.094 | 1.3501 | 2.5475 | 100.3 | 21.29 | 31.214 | 67.291 | 79.9945 | 62.4901 | 0.007888 | 1.0519 | 0.024961 | 23.73 | 0.003206 | 0.13708 | -7e-05 | -0.005208 | -0.005208 | -1.0713
2019-03-29 11:33 | 1553873583.8505 | -136 | 0.20951 | 0.000459 | -1e-06 | 100.8952 | 100.6885 | 125.9615 | 21.0288 | 0.001862 | 0.001689 | -151.8056 | 0.23083 | 8.4029 | 0.002401 | 0.003248 | 0.003059 | 6.9e-05 | 0.000909 | 75.0653 | 1.4057 | 2.5474 | 100.3 | 21.29 | 31.216 | 66.3958 | 79.361 | 62.2469 | 0.007889 | 1.0526 | 0.024942 | 23.6953 | 0.018345 | 0.13732 | -0.00013 | -0.009632 | -0.009632 | -0.58028
2019-03-29 11:33 | 1553873584.8492 | -135 | 0.20949 | 0.000458 | 0 | 100.9365 | 100.8499 | 126.1453 | 21.0285 | 0.001795 | 0.001663 | -151.8181 | 0.23051 | 8.4031 | 0.00242 | 0.004194 | 0.003982 | 7.1e-05 | 0.00091 | 75.082 | 1.3862 | 2.5475 | 100.3 | 21.29 | 31.221 | 64.7939 | 79.8231 | 61.6253 | 0.00789 | 1.0546 | 0.024983 | 23.6906 | 0.007677 | 0.13713 | 2e-06 | 0.000154 | 0.000154 | 36.3559
2019-03-29 11:33 | 1553873585.8497 | -134 | 0.2095 | 0.000452 | -1e-06 | 100.9712 | 100.9032 | 126.445 | 21.0294 | 0.001839 | 0.001698 | -151.8147 | 0.23072 | 8.3995 | 0.002349 | 0.004012 | 0.003895 | 6.8e-05 | 0.000913 | 75.0721 | 1.3525 | 2.5474 | 100.3 | 21.29 | 31.229 | 65.8586 | 78.1198 | 62.449 | 0.007892 | 1.052 | 0.024982 | 23.7479 | 0.006033 | 0.13726 | -5.7e-05 | -0.004246 | -0.004246 | -1.3156
2019-03-29 11:33 | 1553873586.8487 | -133 | 0.2095 | 0.000458 | -1e-06 | 100.9964 | 100.9557 | 126.6886 | 21.0286 | 0.001765 | 0.001737 | -151.8103 | 0.23146 | 8.402 | 0.002374 | 0.004075 | 0.003861 | 7.1e-05 | 0.000915 | 75.0426 | 1.3507 | 2.5474 | 100.3 | 21.29 | 31.235 | 64.0813 | 78.7051 | 63.4035 | 0.007893 | 1.049 | 0.024971 | 23.8045 | 0.003602 | 0.13769 | -0.000101 | -0.007498 | -0.007498 | -0.74744
2019-03-29 11:33 | 1553873587.8503 | -132 | 0.20949 | 0.000454 | -1e-06 | 100.918 | 100.967 | 126.9356 | 21.0293 | 0.001816 | 0.001695 | -151.8073 | 0.23126 | 8.4018 | 0.00231 | 0.003282 | 0.003164 | 7.2e-05 | 0.00092 | 75.0134 | 1.3614 | 2.5474 | 100.3 | 21.29 | 31.242 | 65.2966 | 77.1748 | 62.3871 | 0.007895 | 1.0522 | 0.025033 | 23.7917 | -0.004344 | 0.13758 | -1.4e-05 | -0.000993 | -0.000993 | -5.6362
2019-03-29 11:33 | 1553873588.8501 | -131 | 0.2095 | 0.000458 | -1e-06 | 100.9638 | 100.8943 | 127.1435 | 21.0289 | 0.00181 | 0.001698 | -151.8234 | 0.23031 | 8.4029 | 0.002362 | 0.00372 | 0.003469 | 6.9e-05 | 0.000924 | 75.0724 | 1.3871 | 2.5475 | 100.3 | 21.28 | 31.257 | 65.1636 | 78.4247 | 62.4587 | 0.007894 | 1.052 | 0.025051 | 23.8137 | 0.006157 | 0.13701 | -9.4e-05 | -0.007017 | -0.007017 | -0.79463
2019-03-29 11:33 | 1553873589.8502 | -130 | 0.20951 | 0.000458 | -1e-06 | 100.9913 | 101.0605 | 127.2358 | 21.0294 | 0.001809 | 0.001716 | -151.831 | 0.23117 | 8.4017 | 0.002361 | 0.003289 | 0.003184 | 7.4e-05 | 0.000928 | 75.0321 | 1.3526 | 2.5474 | 100.3 | 21.28 | 31.265 | 65.122 | 78.3901 | 62.8817 | 0.007896 | 1.0506 | 0.025044 | 23.8374 | -0.006126 | 0.13752 | -0.000112 | -0.00837 | -0.00837 | -0.66869
2019-03-29 11:33 | 1553873590.8486 | -129 | 0.20951 | 0.000456 | -1e-06 | 101.0461 | 100.8704 | 127.3508 | 21.0299 | 0.001846 | 0.001663 | -151.8347 | 0.23113 | 8.4039 | 0.002364 | 0.003285 | 0.003088 | 7.3e-05 | 0.000929 | 75.0143 | 1.3691 | 2.5473 | 100.3 | 21.28 | 31.28 | 66.03 | 78.4749 | 61.6071 | 0.0079 | 1.0546 | 0.025103 | 23.8029 | 0.015564 | 0.1375 | -0.000149 | -0.011192 | -0.011192 | -0.50002
2019-03-29 11:33 | 1553873591.8486 | -128 | 0.20952 | 0.000459 | -1e-06 | 101.035 | 101.0909 | 127.3236 | 21.0302 | 0.00186 | 0.001684 | -151.8381 | 0.23034 | 8.4024 | 0.002375 | 0.003651 | 0.003464 | 7.1e-05 | 0.000928 | 75.0864 | 1.3568 | 2.5475 | 100.3 | 21.28 | 31.287 | 66.3471 | 78.7364 | 62.1185 | 0.007902 | 1.053 | 0.025081 | 23.8185 | -0.004945 | 0.13703 | -0.00019 | -0.014251 | -0.014251 | -0.39136
2019-03-29 11:33 | 1553873592.8498 | -127 | 0.20951 | 0.000458 | -1e-06 | 101.1323 | 101.0045 | 127.1072 | 21.0299 | 0.001822 | 0.001727 | -151.8246 | 0.23056 | 8.402 | 0.002486 | 0.003852 | 0.003636 | 6.9e-05 | 0.000928 | 75.0888 | 1.3552 | 2.5475 | 100.3 | 21.28 | 31.294 | 65.4417 | 81.394 | 63.1609 | 0.007903 | 1.0498 | 0.025021 | 23.8352 | 0.011307 | 0.13716 | -0.000132 | -0.009905 | -0.009905 | -0.56358
2019-03-29 11:33 | 1553873593.8498 | -126 | 0.20951 | 0.000459 | -1e-06 | 101.1233 | 101.0028 | 127.3238 | 21.0309 | 0.001869 | 0.001716 | -151.828 | 0.22971 | 8.4026 | 0.002493 | 0.003926 | 0.003722 | 7e-05 | 0.000924 | 75.0697 | 1.3345 | 2.5474 | 100.3 | 21.28 | 31.304 | 66.5654 | 81.5589 | 62.887 | 0.007906 | 1.0506 | 0.025053 | 23.8458 | 0.010664 | 0.13666 | -0.000124 | -0.009342 | -0.009342 | -0.59538
2019-03-29 11:33 | 1553873594.8498 | -125 | 0.20952 | 0.000454 | -1e-06 | 101.1254 | 101.1369 | 127.3088 | 21.0307 | 0.001829 | 0.001714 | -151.8027 | 0.23053 | 8.4016 | 0.002501 | 0.003926 | 0.003775 | 7e-05 | 0.000921 | 75.0677 | 1.3507 | 2.5475 | 100.3 | 21.28 | 31.31 | 65.6227 | 81.7506 | 62.8565 | 0.007907 | 1.0507 | 0.025052 | 23.8433 | -0.001021 | 0.13714 | -0.000174 | -0.013018 | -0.013018 | -0.42877
2019-03-29 11:33 | 1553873595.8488 | -124 | 0.2095 | 0.000458 | -1e-06 | 101.2176 | 101.181 | 127.0879 | 21.0307 | 0.00179 | 0.00168 | -151.7961 | 0.23117 | 8.4036 | 0.002448 | 0.003732 | 0.003478 | 7.2e-05 | 0.000919 | 74.9872 | 1.32 | 2.5473 | 100.3 | 21.28 | 31.314 | 64.6649 | 80.4801 | 62.0174 | 0.007908 | 1.0533 | 0.025062 | 23.7929 | 0.003237 | 0.13752 | -8.5e-05 | -0.006368 | -0.006368 | -0.87889
2019-03-29 11:33 | 1553873596.8502 | -123 | 0.20951 | 0.000459 | -1e-06 | 101.1544 | 101.0632 | 127.0518 | 21.0311 | 0.001796 | 0.001702 | -151.8099 | 0.23038 | 8.4021 | 0.002455 | 0.003624 | 0.003584 | 7.1e-05 | 0.000915 | 75.0127 | 1.3096 | 2.5474 | 100.3 | 21.28 | 31.321 | 64.8239 | 80.6684 | 62.5663 | 0.00791 | 1.0516 | 0.025038 | 23.809 | 0.008069 | 0.13705 | -0.000145 | -0.010844 | -0.010844 | -0.5144
2019-03-29 11:33 | 1553873597.8499 | -122 | 0.20948 | 0.000465 | -1e-06 | 101.2646 | 101.1233 | 127.1574 | 21.0321 | 0.001852 | 0.00173 | -151.813 | 0.23052 | 8.4035 | 0.002424 | 0.004339 | 0.004102 | 6.8e-05 | 0.000909 | 75.0044 | 1.3481 | 2.5474 | 100.3 | 21.28 | 31.326 | 66.1595 | 79.9093 | 63.2274 | 0.007911 | 1.0496 | 0.025024 | 23.8423 | 0.012489 | 0.13714 | 5.9e-05 | 0.004518 | 0.004518 | 1.2355
2019-03-29 11:33 | 1553873598.85 | -121 | 0.20951 | 0.000462 | -1e-06 | 101.2377 | 101.2458 | 127.2306 | 21.0312 | 0.001842 | 0.00171 | -151.8034 | 0.23066 | 8.4005 | 0.002471 | 0.003592 | 0.00346 | 6.7e-05 | 0.000906 | 75.0446 | 1.356 | 2.5474 | 100.3 | 21.28 | 31.334 | 65.9142 | 81.0327 | 62.7429 | 0.007914 | 1.0511 | 0.025049 | 23.832 | -0.000722 | 0.13722 | -0.000116 | -0.008693 | -0.008693 | -0.64246
2019-03-29 11:33 | 1553873599.8503 | -120 | 0.20951 | 0.000462 | 0 | 101.3053 | 101.1936 | 127.2801 | 21.0327 | 0.001825 | 0.001772 | -151.8048 | 0.23083 | 8.4027 | 0.002354 | 0.003836 | 0.003665 | 6.9e-05 | 0.000909 | 75.0468 | 1.3273 | 2.5475 | 100.3 | 21.28 | 31.339 | 65.513 | 78.2274 | 64.2321 | 0.007915 | 1.0464 | 0.024998 | 23.8894 | 0.009868 | 0.13732 | -0.000109 | -0.008192 | -0.008192 | -0.6822
2019-03-29 11:33 | 1553873600.8491 | -119 | 0.20947 | 0.000461 | -1e-06 | 101.2712 | 101.3086 | 127.3776 | 21.0325 | 0.001798 | 0.001736 | -151.7972 | 0.23077 | 8.4017 | 0.002438 | 0.003877 | 0.003726 | 6.7e-05 | 0.000908 | 75.0455 | 1.3444 | 2.5474 | 100.3 | 21.28 | 31.346 | 64.8604 | 80.2418 | 63.3768 | 0.007917 | 1.0491 | 0.02504 | 23.8682 | -0.003306 | 0.13728 | 0.000111 | 0.008408 | 0.008408 | 0.66453
2019-03-29 11:33 | 1553873601.8481 | -118 | 0.20947 | 0.00046 | -1e-06 | 101.3194 | 101.2929 | 127.0905 | 21.0319 | 0.001844 | 0.00174 | -151.7887 | 0.23152 | 8.4008 | 0.002395 | 0.003846 | 0.003607 | 7.1e-05 | 0.000915 | 75.0014 | 1.3715 | 2.5474 | 100.3 | 21.27 | 31.352 | 65.9681 | 79.2214 | 63.4638 | 0.007913 | 1.0488 | 0.025008 | 23.8444 | 0.002336 | 0.13773 | 0.000141 | 0.010681 | 0.010681 | 0.52483
2019-03-29 11:33 | 1553873602.8484 | -117 | 0.20947 | 0.000464 | -1e-06 | 101.3544 | 101.2782 | 126.9306 | 21.0324 | 0.001824 | 0.001682 | -151.789 | 0.23008 | 8.4019 | 0.002386 | 0.003584 | 0.003391 | 7.2e-05 | 0.000916 | 75.0315 | 1.3544 | 2.5471 | 100.3 | 21.27 | 31.358 | 65.4877 | 79.0092 | 62.0816 | 0.007915 | 1.0531 | 0.025044 | 23.7804 | 0.006723 | 0.13687 | 0.000119 | 0.00902 | 0.00902 | 0.61761
2019-03-29 11:33 | 1553873603.8485 | -116 | 0.20953 | 0.000462 | -2e-06 | 101.4119 | 101.3915 | 126.8783 | 21.0327 | 0.001863 | 0.001719 | -151.7854 | 0.23046 | 8.4033 | 0.002357 | 0.003852 | 0.00375 | 6.9e-05 | 0.000923 | 75.0493 | 1.3741 | 2.5473 | 100.3 | 21.27 | 31.369 | 66.4308 | 78.3068 | 62.9725 | 0.007917 | 1.0503 | 0.025006 | 23.8071 | 0.001792 | 0.1371 | -0.000288 | -0.021529 | -0.021529 | -0.25917
2019-03-29 11:33 | 1553873604.8482 | -115 | 0.20955 | 0.000457 | -1e-06 | 101.3954 | 101.4663 | 126.9358 | 21.0324 | 0.001808 | 0.001742 | -151.7989 | 0.23178 | 8.4006 | 0.002356 | 0.003797 | 0.003567 | 6.9e-05 | 0.00092 | 75.0905 | 1.3797 | 2.5474 | 100.3 | 21.27 | 31.375 | 65.1037 | 78.2676 | 63.5297 | 0.007919 | 1.0486 | 0.024991 | 23.8322 | -0.006251 | 0.13788 | -0.000351 | -0.02629 | -0.02629 | -0.21346
2019-03-29 11:33 | 1553873605.8482 | -114 | 0.20951 | 0.00046 | -1e-06 | 101.4214 | 101.466 | 126.6699 | 21.0327 | 0.001776 | 0.001732 | -151.7906 | 0.2304 | 8.3996 | 0.002346 | 0.004191 | 0.004067 | 6.8e-05 | 0.000921 | 75.0859 | 1.3624 | 2.5474 | 100.3 | 21.27 | 31.38 | 64.3469 | 78.0349 | 63.2764 | 0.00792 | 1.0494 | 0.024974 | 23.7983 | -0.003924 | 0.13706 | -0.000118 | -0.008779 | -0.008779 | -0.63542
2019-03-29 11:33 | 1553873606.8498 | -113 | 0.20952 | 0.000458 | -1e-06 | 101.4825 | 101.4595 | 126.6722 | 21.0327 | 0.001833 | 0.001685 | -151.7845 | 0.23138 | 8.404 | 0.002358 | 0.003638 | 0.003393 | 6.7e-05 | 0.000918 | 75.088 | 1.3452 | 2.5475 | 100.3 | 21.27 | 31.388 | 65.7042 | 78.3381 | 62.1592 | 0.007922 | 1.0529 | 0.025016 | 23.759 | 0.002031 | 0.13765 | -0.000177 | -0.013222 | -0.013222 | -0.42372
2019-03-29 11:33 | 1553873607.8483 | -112 | 0.20951 | 0.000458 | -1e-06 | 101.475 | 101.5356 | 126.5128 | 21.0329 | 0.001804 | 0.001735 | -151.7855 | 0.23153 | 8.4021 | 0.002399 | 0.003402 | 0.003294 | 6.8e-05 | 0.000922 | 75.1016 | 1.3592 | 2.5475 | 100.3 | 21.27 | 31.391 | 65.0034 | 79.3049 | 63.3631 | 0.007923 | 1.0491 | 0.024955 | 23.7866 | -0.005336 | 0.13774 | -0.000109 | -0.008123 | -0.008123 | -0.69013
2019-03-29 11:33 | 1553873608.8482 | -111 | 0.20949 | 0.000455 | -1e-06 | 101.4529 | 101.5414 | 126.648 | 21.0335 | 0.001824 | 0.001702 | -151.7907 | 0.23182 | 8.3999 | 0.002381 | 0.003514 | 0.003273 | 6.9e-05 | 0.000926 | 75.1806 | 1.3335 | 2.5474 | 100.3 | 21.26 | 31.395 | 65.4894 | 78.8876 | 62.5546 | 0.007919 | 1.0517 | 0.024999 | 23.7707 | -0.007799 | 0.13791 | -2e-05 | -0.001488 | -0.001488 | -3.7729
2019-03-29 11:33 | 1553873609.8493 | -110 | 0.20952 | 0.00046 | -2e-06 | 101.4882 | 101.597 | 126.825 | 21.0337 | 0.001808 | 0.001719 | -151.7899 | 0.23169 | 8.4016 | 0.002384 | 0.004029 | 0.003797 | 6.8e-05 | 0.000926 | 75.0798 | 1.3411 | 2.5474 | 100.3 | 21.26 | 31.395 | 65.1011 | 78.9416 | 62.9788 | 0.007919 | 1.0503 | 0.025 | 23.8023 | -0.009585 | 0.13783 | -0.000192 | -0.0143 | -0.0143 | -0.39229
2019-03-29 11:33 | 1553873610.8483 | -109 | 0.20949 | 0.000455 | -1e-06 | 101.5648 | 101.5096 | 127.0604 | 21.0337 | 0.001835 | 0.001675 | -151.8049 | 0.23018 | 8.4012 | 0.002442 | 0.003683 | 0.00349 | 7.1e-05 | 0.000928 | 75.0929 | 1.3123 | 2.5473 | 100.3 | 21.26 | 31.395 | 65.7552 | 80.3388 | 61.9001 | 0.007919 | 1.0537 | 0.025064 | 23.7861 | 0.004861 | 0.13693 | -5e-06 | -0.000319 | -0.000319 | -17.4918
2019-03-29 11:34 | 1553873611.8493 | -108 | 0.20952 | 0.000458 | -1e-06 | 101.5244 | 101.5062 | 127.5407 | 21.0346 | 0.001808 | 0.001726 | -151.8058 | 0.23077 | 8.403 | 0.002407 | 0.003495 | 0.003359 | 6.7e-05 | 0.000927 | 75.0875 | 1.3476 | 2.5474 | 100.3 | 21.26 | 31.394 | 65.1198 | 79.4993 | 63.132 | 0.007919 | 1.0498 | 0.025065 | 23.8748 | 0.001603 | 0.13728 | -0.000198 | -0.01482 | -0.01482 | -0.37702
2019-03-29 11:34 | 1553873612.8487 | -107 | 0.2095 | 0.000458 | -2e-06 | 101.5912 | 101.6533 | 127.7663 | 21.0342 | 0.001778 | 0.00171 | -151.8003 | 0.23105 | 8.4023 | 0.002445 | 0.00333 | 0.003067 | 6.8e-05 | 0.000928 | 75.0941 | 1.392 | 2.5475 | 100.3 | 21.26 | 31.39 | 64.3859 | 80.4088 | 62.7591 | 0.007918 | 1.051 | 0.025101 | 23.8827 | -0.005471 | 0.13745 | -7e-05 | -0.005153 | -0.005153 | -1.0857
2019-03-29 11:34 | 1553873613.8494 | -106 | 0.20953 | 0.000462 | -1e-06 | 101.5648 | 101.7185 | 127.5876 | 21.0342 | 0.001807 | 0.001715 | -151.7905 | 0.23088 | 8.4039 | 0.00242 | 0.003971 | 0.003855 | 6.7e-05 | 0.00093 | 75.1052 | 1.4024 | 2.5475 | 100.3 | 21.26 | 31.386 | 65.0829 | 79.8279 | 62.8686 | 0.007917 | 1.0507 | 0.025079 | 23.8699 | -0.013524 | 0.13735 | -0.000266 | -0.019993 | -0.019993 | -0.2796
2019-03-29 11:34 | 1553873614.8484 | -105 | 0.20954 | 0.000458 | -1e-06 | 101.61 | 101.7031 | 127.3426 | 21.0348 | 0.001892 | 0.001758 | -151.7892 | 0.23068 | 8.4018 | 0.002368 | 0.004134 | 0.003922 | 7e-05 | 0.000933 | 75.0592 | 1.3756 | 2.5474 | 100.3 | 21.26 | 31.383 | 67.1257 | 78.5788 | 63.9047 | 0.007916 | 1.0474 | 0.025017 | 23.8837 | -0.008194 | 0.13723 | -0.000299 | -0.022398 | -0.022398 | -0.24936
2019-03-29 11:34 | 1553873615.8475 | -104 | 0.20951 | 0.00046 | -1e-06 | 101.7396 | 101.8601 | 127.4975 | 21.0356 | 0.001818 | 0.001696 | -151.8009 | 0.23066 | 8.4025 | 0.002414 | 0.003426 | 0.003262 | 6.9e-05 | 0.000938 | 75.0504 | 1.347 | 2.5473 | 100.3 | 21.26 | 31.38 | 65.3629 | 79.6702 | 62.4221 | 0.007915 | 1.0521 | 0.025087 | 23.8456 | -0.010588 | 0.13722 | -0.000108 | -0.008087 | -0.008087 | -0.69057
2019-03-29 11:34 | 1553873616.8475 | -103 | 0.20951 | 0.000461 | -1e-06 | 101.6857 | 101.7349 | 127.2752 | 21.0353 | 0.00186 | 0.001658 | -151.8004 | 0.23057 | 8.401 | 0.002358 | 0.003805 | 0.003775 | 7.1e-05 | 0.000936 | 75.0895 | 1.3541 | 2.5475 | 100.3 | 21.26 | 31.379 | 66.3701 | 78.3226 | 61.4968 | 0.007915 | 1.055 | 0.0251 | 23.7919 | -0.004321 | 0.13717 | -0.000133 | -0.009979 | -0.009979 | -0.55943
2019-03-29 11:34 | 1553873617.8475 | -102 | 0.20952 | 0.000459 | -1e-06 | 101.7178 | 101.784 | 127.3635 | 21.0354 | 0.001835 | 0.001763 | -151.8158 | 0.23057 | 8.4025 | 0.002321 | 0.003716 | 0.003516 | 6.8e-05 | 0.000938 | 75.0873 | 1.3283 | 2.5474 | 100.3 | 21.26 | 31.377 | 65.7514 | 77.4443 | 64.0316 | 0.007915 | 1.047 | 0.025014 | 23.8901 | -0.005824 | 0.13717 | -0.000198 | -0.014853 | -0.014853 | -0.37586
2019-03-29 11:34 | 1553873618.8475 | -101 | 0.2095 | 0.00046 | 0 | 101.6819 | 101.6359 | 127.0722 | 21.0358 | 0.001771 | 0.00177 | -151.8256 | 0.23129 | 8.4013 | 0.002347 | 0.003841 | 0.003624 | 6.7e-05 | 0.000942 | 75.0993 | 1.339 | 2.5474 | 100.3 | 21.26 | 31.378 | 64.232 | 78.0715 | 64.2 | 0.007915 | 1.0465 | 0.024979 | 23.8687 | 0.004043 | 0.13759 | -5.6e-05 | -0.004191 | -0.004191 | -1.3363
2019-03-29 11:34 | 1553873619.8487 | -100 | 0.20953 | 0.00046 | -1e-06 | 101.8088 | 101.8595 | 127.1611 | 21.0353 | 0.001842 | 0.001707 | -151.8205 | 0.23075 | 8.4021 | 0.002399 | 0.00309 | 0.002929 | 6.8e-05 | 0.000947 | 75.1011 | 1.3266 | 2.5475 | 100.31 | 21.26 | 31.379 | 65.9307 | 79.309 | 62.6821 | 0.007914 | 1.0513 | 0.025044 | 23.8233 | -0.004451 | 0.13728 | -0.000252 | -0.018906 | -0.018906 | -0.29552
2019-03-29 11:34 | 1553873620.848 | -99 | 0.20951 | 0.000462 | -1e-06 | 101.6932 | 101.6871 | 126.8534 | 21.0354 | 0.001853 | 0.001722 | -151.7985 | 0.22972 | 8.4046 | 0.002412 | 0.004079 | 0.003943 | 6.7e-05 | 0.000947 | 75.0419 | 1.3463 | 2.5474 | 100.31 | 21.26 | 31.381 | 66.197 | 79.6356 | 63.0406 | 0.007915 | 1.0501 | 0.025001 | 23.8072 | 0.000539 | 0.13666 | -0.000108 | -0.008081 | -0.008081 | -0.68827
2019-03-29 11:34 | 1553873621.8491 | -98 | 0.20953 | 0.000455 | 0 | 101.7247 | 101.8127 | 126.6884 | 21.0367 | 0.001802 | 0.001677 | -151.8082 | 0.23183 | 8.3999 | 0.002356 | 0.003594 | 0.003409 | 6.9e-05 | 0.000949 | 75.074 | 1.3828 | 2.5475 | 100.31 | 21.26 | 31.384 | 64.9783 | 78.2819 | 61.957 | 0.007916 | 1.0535 | 0.025025 | 23.7533 | -0.007737 | 0.13792 | -0.000221 | -0.016619 | -0.016619 | -0.33776
2019-03-29 11:34 | 1553873622.8476 | -97 | 0.20953 | 0.000458 | 0 | 101.7178 | 101.8557 | 126.5523 | 21.0355 | 0.001848 | 0.001739 | -151.8265 | 0.23125 | 8.4038 | 0.002291 | 0.003528 | 0.003351 | 6.8e-05 | 0.000947 | 75.052 | 1.3551 | 2.5474 | 100.31 | 21.26 | 31.388 | 66.0706 | 76.7257 | 63.4531 | 0.007917 | 1.0488 | 0.024956 | 23.7935 | -0.01212 | 0.13757 | -0.000265 | -0.01989 | -0.01989 | -0.28151
2019-03-29 11:34 | 1553873623.8482 | -96 | 0.20955 | 0.000461 | -1e-06 | 101.6747 | 101.8926 | 126.8281 | 21.0359 | 0.001837 | 0.001781 | -151.8323 | 0.23121 | 8.4033 | 0.002344 | 0.003926 | 0.003744 | 7e-05 | 0.000955 | 75.0624 | 1.3405 | 2.5475 | 100.31 | 21.26 | 31.395 | 65.8103 | 77.9827 | 64.4732 | 0.007918 | 1.0457 | 0.024945 | 23.8555 | -0.019147 | 0.13755 | -0.000354 | -0.026533 | -0.026533 | -0.21099
2019-03-29 11:34 | 1553873624.8492 | -95 | 0.20953 | 0.000463 | -1e-06 | 101.7058 | 101.8567 | 126.9076 | 21.0362 | 0.001883 | 0.001689 | -151.8188 | 0.23144 | 8.4015 | 0.002279 | 0.003811 | 0.003621 | 7.2e-05 | 0.000962 | 75.0066 | 1.3331 | 2.5474 | 100.31 | 21.26 | 31.399 | 66.9234 | 76.4355 | 62.2409 | 0.007919 | 1.0526 | 0.025036 | 23.7839 | -0.013263 | 0.13768 | -0.000269 | -0.020213 | -0.020213 | -0.27724
2019-03-29 11:34 | 1553873625.8478 | -94 | 0.20949 | 0.000461 | -1e-06 | 101.728 | 101.8035 | 126.6316 | 21.0366 | 0.001878 | 0.001714 | -151.8295 | 0.23116 | 8.4026 | 0.002353 | 0.003659 | 0.00351 | 7e-05 | 0.000966 | 75.0166 | 1.3208 | 2.5473 | 100.31 | 21.26 | 31.405 | 66.7903 | 78.2047 | 62.8442 | 0.007921 | 1.0507 | 0.024986 | 23.7794 | -0.006638 | 0.13752 | -4.1e-05 | -0.003033 | -0.003033 | -1.8456
2019-03-29 11:34 | 1553873626.8488 | -93 | 0.20951 | 0.00046 | -1e-06 | 101.7262 | 101.7339 | 126.5283 | 21.0373 | 0.001783 | 0.001728 | -151.831 | 0.23022 | 8.403 | 0.00237 | 0.00389 | 0.003668 | 7.2e-05 | 0.000972 | 75.0757 | 1.3393 | 2.5473 | 100.31 | 21.26 | 31.408 | 64.5152 | 78.6225 | 63.1987 | 0.007922 | 1.0496 | 0.024963 | 23.7822 | -0.000677 | 0.13695 | -0.00013 | -0.009684 | -0.009684 | -0.57562
2019-03-29 11:34 | 1553873627.8482 | -92 | 0.20952 | 0.000455 | 0 | 101.7887 | 101.7018 | 126.1593 | 21.0367 | 0.001873 | 0.001717 | -151.8206 | 0.23114 | 8.4042 | 0.002307 | 0.004049 | 0.003904 | 7.3e-05 | 0.000968 | 75.0615 | 1.3416 | 2.5473 | 100.31 | 21.26 | 31.409 | 66.6875 | 77.0954 | 62.9319 | 0.007922 | 1.0505 | 0.024936 | 23.7381 | 0.007643 | 0.1375 | -0.000213 | -0.015913 | -0.015913 | -0.35168
2019-03-29 11:34 | 1553873628.8479 | -91 | 0.2095 | 0.000459 | -1e-06 | 101.6861 | 101.7915 | 126.3596 | 21.0372 | 0.001826 | 0.001756 | -151.8076 | 0.23094 | 8.4032 | 0.002302 | 0.00353 | 0.003314 | 7.3e-05 | 0.000965 | 75.0112 | 1.3926 | 2.5475 | 100.31 | 21.26 | 31.412 | 65.5385 | 76.976 | 63.8725 | 0.007923 | 1.0475 | 0.024921 | 23.7902 | -0.009274 | 0.13739 | -5.3e-05 | -0.003917 | -0.003917 | -1.4277
2019-03-29 11:34 | 1553873629.8486 | -90 | 0.20955 | 0.000458 | -1e-06 | 101.6816 | 101.6973 | 126.0373 | 21.0372 | 0.001848 | 0.001651 | -151.7976 | 0.2304 | 8.4013 | 0.002326 | 0.003664 | 0.003607 | 7.2e-05 | 0.000963 | 75.0021 | 1.3479 | 2.5476 | 100.31 | 21.26 | 31.413 | 66.0824 | 77.552 | 61.3322 | 0.007923 | 1.0555 | 0.024984 | 23.6701 | -0.001388 | 0.13707 | -0.000363 | -0.027252 | -0.027252 | -0.2047
2019-03-29 11:34 | 1553873630.8489 | -89 | 0.2095 | 0.000459 | -1e-06 | 101.7669 | 101.8615 | 126.0176 | 21.037 | 0.001858 | 0.001706 | -151.8178 | 0.23114 | 8.4015 | 0.002364 | 0.00347 | 0.003338 | 7.2e-05 | 0.000958 | 75.0722 | 1.3633 | 2.5475 | 100.31 | 21.25 | 31.413 | 66.3123 | 78.4645 | 62.6587 | 0.007918 | 1.0513 | 0.024932 | 23.7151 | -0.008314 | 0.13751 | -0.0001 | -0.00744 | -0.00744 | -0.75225
2019-03-29 11:34 | 1553873631.8484 | -88 | 0.2095 | 0.000461 | -1e-06 | 101.6367 | 101.7455 | 125.8955 | 21.0375 | 0.001847 | 0.001747 | -151.7974 | 0.23081 | 8.4025 | 0.002318 | 0.00383 | 0.003542 | 6.8e-05 | 0.000956 | 74.9956 | 1.3452 | 2.5475 | 100.31 | 21.25 | 31.414 | 66.0601 | 77.3654 | 63.6363 | 0.007918 | 1.0483 | 0.024884 | 23.7381 | -0.00957 | 0.13731 | -8.4e-05 | -0.006196 | -0.006196 | -0.90202
2019-03-29 11:34 | 1553873632.8484 | -87 | 0.20952 | 0.000461 | -1e-06 | 101.7609 | 101.812 | 126.1943 | 21.0376 | 0.001862 | 0.001691 | -151.7878 | 0.22997 | 8.402 | 0.002299 | 0.003953 | 0.003904 | 7e-05 | 0.000957 | 75.059 | 1.3397 | 2.5473 | 100.31 | 21.25 | 31.413 | 66.4222 | 76.9265 | 62.3039 | 0.007918 | 1.0524 | 0.024963 | 23.7192 | -0.004493 | 0.13681 | -0.000168 | -0.012563 | -0.012563 | -0.44322
2019-03-29 11:34 | 1553873633.8471 | -86 | 0.2095 | 0.00046 | -1e-06 | 101.6409 | 101.7564 | 126.1852 | 21.0382 | 0.001821 | 0.001676 | -151.788 | 0.23044 | 8.4031 | 0.002277 | 0.004147 | 0.003926 | 7.4e-05 | 0.000958 | 75.0154 | 1.3372 | 2.5474 | 100.31 | 21.25 | 31.411 | 65.4341 | 76.3814 | 61.938 | 0.007918 | 1.0536 | 0.024976 | 23.7054 | -0.010162 | 0.13709 | -6.4e-05 | -0.004805 | -0.004805 | -1.1612
2019-03-29 11:34 | 1553873634.8481 | -85 | 0.20953 | 0.000457 | 0 | 101.7313 | 101.7008 | 126.5729 | 21.0382 | 0.001849 | 0.001701 | -151.8017 | 0.22993 | 8.4027 | 0.00229 | 0.004078 | 0.003889 | 7.1e-05 | 0.000952 | 75.0233 | 1.3332 | 2.5475 | 100.31 | 21.25 | 31.409 | 66.0928 | 76.7034 | 62.5506 | 0.007917 | 1.0517 | 0.024991 | 23.7635 | 0.002681 | 0.13678 | -0.000244 | -0.018349 | -0.018349 | -0.3034
2019-03-29 11:34 | 1553873635.8484 | -84 | 0.2095 | 0.000461 | -1e-06 | 101.625 | 101.7591 | 126.9567 | 21.0381 | 0.001817 | 0.00172 | -151.79 | 0.23072 | 8.4003 | 0.002257 | 0.003317 | 0.003184 | 7.5e-05 | 0.00095 | 75.1016 | 1.313 | 2.5474 | 100.31 | 21.25 | 31.409 | 65.3388 | 75.9028 | 63.005 | 0.007917 | 1.0502 | 0.025012 | 23.8156 | -0.011798 | 0.13725 | -5.5e-05 | -0.004067 | -0.004067 | -1.3736
2019-03-29 11:34 | 1553873636.8482 | -83 | 0.20953 | 0.000461 | -1e-06 | 101.6801 | 101.8048 | 127.566 | 21.0389 | 0.001867 | 0.001667 | -151.7886 | 0.23003 | 8.4014 | 0.00235 | 0.003108 | 0.002962 | 7.2e-05 | 0.000946 | 75.0981 | 1.3341 | 2.5474 | 100.31 | 21.25 | 31.409 | 66.5249 | 78.1531 | 61.7207 | 0.007917 | 1.0543 | 0.02512 | 23.827 | -0.01097 | 0.13684 | -0.000247 | -0.018577 | -0.018577 | -0.29981
2019-03-29 11:34 | 1553873637.8477 | -82 | 0.20951 | 0.000461 | -1e-06 | 101.6594 | 101.5796 | 127.8853 | 21.0385 | 0.001844 | 0.001701 | -151.7681 | 0.23073 | 8.4019 | 0.002311 | 0.003598 | 0.003343 | 7.4e-05 | 0.000944 | 75.0581 | 1.313 | 2.5474 | 100.31 | 21.25 | 31.41 | 65.969 | 77.1955 | 62.5304 | 0.007917 | 1.0517 | 0.025121 | 23.8857 | 0.007025 | 0.13726 | -0.000107 | -0.007993 | -0.007993 | -0.69896
2019-03-29 11:34 | 1553873638.8473 | -81 | 0.20952 | 0.00046 | 0 | 101.6016 | 101.5919 | 127.7299 | 21.0391 | 0.00188 | 0.00172 | -151.7599 | 0.23065 | 8.4026 | 0.002341 | 0.004185 | 0.00396 | 7.4e-05 | 0.000942 | 75.0749 | 1.3597 | 2.5475 | 100.3 | 21.25 | 31.41 | 66.8355 | 77.9209 | 62.9896 | 0.007918 | 1.0503 | 0.025089 | 23.8875 | 0.000859 | 0.13721 | -0.000199 | -0.014965 | -0.014965 | -0.37317
2019-03-29 11:34 | 1553873639.8472 | -80 | 0.20952 | 0.000461 | -1e-06 | 101.6046 | 101.5342 | 127.8494 | 21.0394 | 0.001849 | 0.001687 | -151.7591 | 0.23031 | 8.4016 | 0.002381 | 0.0036 | 0.003446 | 7e-05 | 0.000945 | 75.2246 | 1.3686 | 2.5474 | 100.3 | 21.25 | 31.41 | 66.0939 | 78.8819 | 62.2048 | 0.007918 | 1.0528 | 0.02513 | 23.8707 | 0.006202 | 0.13701 | -0.000163 | -0.012245 | -0.012245 | -0.4554
2019-03-29 11:34 | 1553873640.8473 | -79 | 0.20952 | 0.000457 | 0 | 101.6304 | 101.4973 | 127.84 | 21.0391 | 0.001779 | 0.001691 | -151.7345 | 0.23027 | 8.4019 | 0.002339 | 0.004241 | 0.004024 | 7.3e-05 | 0.00095 | 74.053 | 1.4012 | 2.5475 | 100.3 | 21.25 | 31.41 | 64.4132 | 77.889 | 62.2942 | 0.007918 | 1.0525 | 0.025126 | 23.873 | 0.011716 | 0.13699 | -0.000155 | -0.011665 | -0.011665 | -0.47797
2019-03-29 11:34 | 1553873641.8473 | -78 | 0.20949 | 0.00046 | -1e-06 | 101.4852 | 101.6581 | 127.7623 | 21.0397 | 0.001839 | 0.001718 | -151.7462 | 0.23066 | 8.4007 | 0.002401 | 0.003601 | 0.003369 | 7.3e-05 | 0.00095 | 73.4253 | 1.381 | 2.5473 | 100.3 | 21.24 | 31.408 | 65.8695 | 79.3678 | 62.9488 | 0.007913 | 1.0504 | 0.025094 | 23.889 | -0.015225 | 0.13722 | -1.4e-05 | -0.001002 | -0.001002 | -5.576
2019-03-29 11:34 | 1553873642.8477 | -77 | 0.20948 | 0.000464 | -1e-06 | 101.5382 | 101.5772 | 127.663 | 21.0398 | 0.001875 | 0.001743 | -151.7371 | 0.23031 | 8.4022 | 0.002358 | 0.003871 | 0.003764 | 7.1e-05 | 0.00094 | 62.1078 | 1.3291 | 2.5473 | 100.3 | 21.24 | 31.408 | 66.7326 | 78.328 | 63.5582 | 0.007913 | 1.0485 | 0.025061 | 23.9014 | -0.003437 | 0.13701 | 5.1e-05 | 0.003878 | 0.003878 | 1.4379
2019-03-29 11:34 | 1553873643.8484 | -76 | 0.2095 | 0.000461 | -1e-06 | 101.4292 | 101.5339 | 127.5355 | 21.0399 | 0.001865 | 0.001748 | -151.7341 | 0.231 | 8.4023 | 0.00243 | 0.003483 | 0.003333 | 7.4e-05 | 0.00094 | 51.1223 | 1.3256 | 2.5474 | 100.3 | 21.24 | 31.408 | 66.4883 | 80.0595 | 63.6706 | 0.007913 | 1.0482 | 0.025044 | 23.8934 | -0.009223 | 0.13742 | -7.6e-05 | -0.005657 | -0.005657 | -0.98867
2019-03-29 11:34 | 1553873644.8484 | -75 | 0.2095 | 0.00046 | 0 | 101.452 | 101.3953 | 127.1658 | 21.0398 | 0.001873 | 0.001725 | -151.7258 | 0.22962 | 8.4012 | 0.002429 | 0.003526 | 0.003279 | 7.4e-05 | 0.000937 | 26.4425 | 1.3828 | 2.5475 | 100.3 | 21.24 | 31.409 | 66.6905 | 80.0345 | 63.1176 | 0.007913 | 1.0499 | 0.025029 | 23.8392 | 0.004998 | 0.1366 | -7e-05 | -0.005247 | -0.005247 | -1.0595
2019-03-29 11:34 | 1553873645.8478 | -74 | 0.2095 | 0.000462 | -1e-06 | 101.3718 | 101.411 | 127.2844 | 21.0398 | 0.00179 | 0.001723 | -151.7154 | 0.23003 | 8.3994 | 0.002412 | 0.003734 | 0.003613 | 7.2e-05 | 0.000933 | 16.0478 | 1.3558 | 2.5474 | 100.3 | 21.24 | 31.409 | 64.6737 | 79.6392 | 63.0745 | 0.007913 | 1.05 | 0.025042 | 23.8488 | -0.003463 | 0.13685 | -9.1e-05 | -0.006831 | -0.006831 | -0.81538
2019-03-29 11:34 | 1553873646.8465 | -73 | 0.20947 | 0.000458 | -1e-06 | 101.2987 | 101.2943 | 126.9076 | 21.0407 | 0.00186 | 0.001759 | -151.7129 | 0.23043 | 8.4042 | 0.002481 | 0.003655 | 0.003452 | 7.2e-05 | 0.00093 | 4.3691 | 1.3406 | 2.5475 | 100.31 | 21.24 | 31.411 | 66.3761 | 81.2854 | 63.9376 | 0.007913 | 1.0473 | 0.024973 | 23.844 | 0.000392 | 0.13708 | 8.5e-05 | 0.006401 | 0.006401 | 0.87157
2019-03-29 11:34 | 1553873647.8465 | -72 | 0.20948 | 0.000459 | -1e-06 | 101.2173 | 101.2247 | 127.1126 | 21.0408 | 0.001823 | 0.001767 | -151.7076 | 0.23097 | 8.4033 | 0.002475 | 0.004208 | 0.004015 | 7.5e-05 | 0.000929 | 1.0513 | 1.3829 | 2.5475 | 100.31 | 21.24 | 31.412 | 65.4833 | 81.159 | 64.1415 | 0.007913 | 1.0467 | 0.024985 | 23.8705 | -0.00065 | 0.1374 | 4.7e-05 | 0.003574 | 0.003574 | 1.5647
2019-03-29 11:34 | 1553873648.8469 | -71 | 0.2095 | 0.000462 | -1e-06 | 101.1978 | 101.0711 | 127.4475 | 21.0408 | 0.001845 | 0.001692 | -151.7022 | 0.2304 | 8.4009 | 0.002462 | 0.003398 | 0.003224 | 7.3e-05 | 0.000928 | 0.74197 | 1.3456 | 2.5475 | 100.31 | 21.24 | 31.412 | 66.007 | 80.8346 | 62.3203 | 0.007913 | 1.0524 | 0.025086 | 23.8373 | 0.011211 | 0.13706 | -7.2e-05 | -0.005412 | -0.005412 | -1.0307
2019-03-29 11:34 | 1553873649.8466 | -70 | 0.20949 | 0.000458 | -1e-06 | 101.1517 | 101.0141 | 127.4581 | 21.0415 | 0.001797 | 0.00168 | -151.6977 | 0.23006 | 8.4009 | 0.002433 | 0.003738 | 0.003487 | 7.5e-05 | 0.000921 | 0.5916 | 1.3432 | 2.5474 | 100.31 | 21.24 | 31.415 | 64.8441 | 80.139 | 62.031 | 0.007914 | 1.0533 | 0.025098 | 23.828 | 0.012181 | 0.13686 | 2.1e-05 | 0.001616 | 0.001616 | 3.446
2019-03-29 11:34 | 1553873650.8477 | -69 | 0.20952 | 0.000459 | -1e-06 | 100.9617 | 100.8632 | 127.4876 | 21.0416 | 0.001869 | 0.001722 | -151.6894 | 0.23026 | 8.4053 | 0.002457 | 0.003902 | 0.003742 | 7.2e-05 | 0.000918 | 0.42542 | 1.353 | 2.5475 | 100.31 | 21.24 | 31.417 | 66.5775 | 80.7262 | 63.0435 | 0.007914 | 1.0501 | 0.025063 | 23.8667 | 0.008725 | 0.13698 | -0.000184 | -0.01383 | -0.01383 | -0.40312
2019-03-29 11:34 | 1553873651.8481 | -68 | 0.20944 | 0.00046 | -1e-06 | 100.9368 | 100.8469 | 127.6874 | 21.0417 | 0.001756 | 0.001696 | -151.6912 | 0.23065 | 8.4035 | 0.002441 | 0.003287 | 0.003075 | 7.1e-05 | 0.000912 | 0.43371 | 1.3513 | 2.5475 | 100.3 | 21.24 | 31.422 | 63.8558 | 80.3279 | 62.4281 | 0.007916 | 1.0521 | 0.025106 | 23.8635 | 0.007976 | 0.13721 | 0.000292 | 0.022104 | 0.022104 | 0.25265
2019-03-29 11:34 | 1553873652.8476 | -67 | 0.2095 | 0.00046 | -1e-06 | 100.7973 | 100.867 | 127.5184 | 21.0425 | 0.0018 | 0.001704 | -151.6966 | 0.23056 | 8.4033 | 0.002333 | 0.003445 | 0.003244 | 7.4e-05 | 0.000911 | 0.41214 | 1.3412 | 2.5474 | 100.3 | 21.24 | 31.426 | 64.9153 | 77.7394 | 62.618 | 0.007917 | 1.0515 | 0.025082 | 23.8545 | -0.006179 | 0.13716 | -4.7e-05 | -0.003475 | -0.003475 | -1.6066
2019-03-29 11:34 | 1553873653.8483 | -66 | 0.20949 | 0.000459 | -1e-06 | 100.7913 | 100.5448 | 127.3172 | 21.0426 | 0.001831 | 0.001704 | -151.7028 | 0.23044 | 8.4022 | 0.00249 | 0.003281 | 0.00313 | 6.9e-05 | 0.000908 | 0.39832 | 1.3682 | 2.5475 | 100.3 | 21.24 | 31.43 | 65.6833 | 81.5187 | 62.6213 | 0.007918 | 1.0514 | 0.025062 | 23.8358 | 0.021905 | 0.13709 | -1.2e-05 | -0.000831 | -0.000831 | -6.7164
2019-03-29 11:34 | 1553873654.8471 | -65 | 0.20927 | 0.00046 | -1e-06 | 100.5938 | 100.6274 | 127.278 | 21.0423 | 0.001782 | 0.001766 | -151.7193 | 0.23138 | 8.3997 | 0.002445 | 0.003976 | 0.003776 | 7e-05 | 0.00091 | 0.38256 | 1.3607 | 2.5475 | 100.3 | 21.24 | 31.44 | 64.4842 | 80.4205 | 64.097 | 0.007921 | 1.0468 | 0.025003 | 23.8844 | -0.002988 | 0.13765 | 0.001321 | 0.099327 | 0.099327 | 0.056402
2019-03-29 11:34 | 1553873655.8482 | -64 | 0.20941 | 0.000456 | -1e-06 | 100.5118 | 100.3318 | 126.8903 | 21.0424 | 0.001835 | 0.001708 | -151.7279 | 0.23067 | 8.4017 | 0.002433 | 0.003504 | 0.003316 | 7.1e-05 | 0.000902 | 0.3788 | 1.3398 | 2.5475 | 100.3 | 21.24 | 31.446 | 65.7682 | 80.1367 | 62.7193 | 0.007922 | 1.0511 | 0.025016 | 23.7992 | 0.016026 | 0.13723 | 0.000467 | 0.035201 | 0.035201 | 0.15866
2019-03-29 11:34 | 1553873656.8467 | -63 | 0.20933 | 0.000457 | -1e-06 | 100.3573 | 100.2516 | 127.0773 | 21.0435 | 0.001865 | 0.001714 | -151.7208 | 0.23031 | 8.4036 | 0.002426 | 0.003537 | 0.003379 | 7.2e-05 | 0.0009 | 0.35867 | 1.3579 | 2.5475 | 100.3 | 21.24 | 31.458 | 66.4998 | 79.9631 | 62.8526 | 0.007925 | 1.0507 | 0.02503 | 23.8215 | 0.009423 | 0.13701 | 0.000988 | 0.074383 | 0.074383 | 0.074969
2019-03-29 11:34 | 1553873657.8474 | -62 | 0.20953 | 0.00046 | -1e-06 | 100.2056 | 100.138 | 126.8011 | 21.0427 | 0.001839 | 0.001688 | -151.7218 | 0.23005 | 8.4018 | 0.002498 | 0.003355 | 0.003261 | 7.1e-05 | 0.000902 | 0.37449 | 1.3065 | 2.5472 | 100.3 | 21.24 | 31.461 | 65.8559 | 81.6991 | 62.2386 | 0.007926 | 1.0526 | 0.025025 | 23.7738 | 0.006033 | 0.13686 | -0.000224 | -0.016797 | -0.016797 | -0.33162
2019-03-29 11:34 | 1553873658.8472 | -61 | 0.20952 | 0.000459 | 0 | 100.0454 | 99.9203 | 126.5499 | 21.0432 | 0.001837 | 0.001737 | -151.7278 | 0.23039 | 8.4016 | 0.002435 | 0.003425 | 0.003187 | 6.8e-05 | 0.000905 | 0.36785 | 1.3219 | 2.5474 | 100.3 | 21.24 | 31.465 | 65.8285 | 80.193 | 63.4014 | 0.007927 | 1.049 | 0.024957 | 23.7914 | 0.011197 | 0.13706 | -0.000163 | -0.012187 | -0.012187 | -0.45774
2019-03-29 11:34 | 1553873659.8479 | -60 | 0.20947 | 0.00046 | -1e-06 | 99.9239 | 99.682 | 126.7372 | 21.0438 | 0.001731 | 0.001723 | -151.7349 | 0.23109 | 8.4006 | 0.002441 | 0.003418 | 0.003236 | 7e-05 | 0.000911 | 0.35187 | 1.33 | 2.5475 | 100.3 | 21.24 | 31.471 | 63.2613 | 80.3221 | 63.0738 | 0.007929 | 1.05 | 0.024988 | 23.7974 | 0.021675 | 0.13748 | 0.000101 | 0.007659 | 0.007659 | 0.73053
2019-03-29 11:34 | 1553873660.8677 | -59 | 0.2095 | 0.000463 | 0 | 99.6581 | 99.4128 | 126.8004 | 21.0441 | 0.001842 | 0.00172 | -151.7674 | 0.23079 | 8.4047 | 0.002453 | 0.003634 | 0.003493 | 6.8e-05 | 0.000915 | 0.35889 | 1.3295 | 2.5474 | 100.3 | 21.24 | 31.472 | 65.9462 | 80.6282 | 63.0046 | 0.007929 | 1.0502 | 0.024997 | 23.8009 | 0.022045 | 0.1373 | -9e-05 | -0.006734 | -0.006734 | -0.82976
2019-03-29 11:34 | 1553873661.861 | -58 | 0.20949 | 0.000461 | -1e-06 | 99.5069 | 99.2524 | 126.9985 | 21.0432 | 0.001886 | 0.001696 | -151.7624 | 0.2319 | 8.4012 | 0.002446 | 0.003517 | 0.003327 | 6.8e-05 | 0.000919 | 0.34324 | 1.3512 | 2.5474 | 100.3 | 21.24 | 31.473 | 66.9886 | 80.4487 | 62.4256 | 0.007929 | 1.0521 | 0.025038 | 23.799 | 0.022912 | 0.13796 | -6e-06 | -0.00041 | -0.00041 | -13.6876
2019-03-29 11:34 | 1553873662.8597 | -57 | 0.20948 | 0.000464 | -1e-06 | 99.2692 | 98.9411 | 127.086 | 21.0449 | 0.001867 | 0.001738 | -151.7498 | 0.23169 | 8.4023 | 0.002465 | 0.003541 | 0.003317 | 7.1e-05 | 0.00092 | 0.33085 | 1.3027 | 2.5475 | 100.3 | 21.24 | 31.475 | 66.5283 | 80.9142 | 63.442 | 0.00793 | 1.0489 | 0.025009 | 23.8432 | 0.029617 | 0.13783 | 5.7e-05 | 0.00434 | 0.00434 | 1.2925
2019-03-29 11:34 | 1553873663.8586 | -56 | 0.20951 | 0.00046 | 0 | 99.0792 | 98.5991 | 127.3494 | 21.0443 | 0.001823 | 0.001699 | -151.7441 | 0.23052 | 8.4014 | 0.002399 | 0.004063 | 0.003946 | 6.5e-05 | 0.000918 | 0.35745 | 1.3571 | 2.5474 | 100.3 | 21.24 | 31.475 | 65.4805 | 79.3177 | 62.4901 | 0.00793 | 1.0519 | 0.02507 | 23.8341 | 0.043442 | 0.13714 | -0.000142 | -0.01066 | -0.01066 | -0.52359
2019-03-29 11:34 | 1553873664.859 | -55 | 0.20947 | 0.00046 | 0 | 98.8529 | 98.5008 | 127.7092 | 21.0448 | 0.001878 | 0.001652 | -151.7354 | 0.23114 | 8.4021 | 0.002396 | 0.003759 | 0.003601 | 6.7e-05 | 0.000918 | 0.34396 | 1.336 | 2.5474 | 100.3 | 21.24 | 31.47 | 66.7976 | 79.2526 | 61.3591 | 0.007928 | 1.0554 | 0.025148 | 23.8275 | 0.031911 | 0.13751 | 8.7e-05 | 0.00663 | 0.00663 | 0.84418
2019-03-29 11:34 | 1553873665.8586 | -54 | 0.2095 | 0.000462 | -1e-06 | 98.5952 | 98.2367 | 127.9343 | 21.0456 | 0.00184 | 0.001696 | -151.7342 | 0.23161 | 8.4016 | 0.002453 | 0.003556 | 0.003378 | 6.9e-05 | 0.00092 | 0.35977 | 1.3784 | 2.5474 | 100.3 | 21.24 | 31.468 | 65.8847 | 80.6272 | 62.4265 | 0.007928 | 1.0521 | 0.02513 | 23.8865 | 0.03258 | 0.13779 | -6.2e-05 | -0.004631 | -0.004631 | -1.211
2019-03-29 11:34 | 1553873666.8601 | -53 | 0.20949 | 0.000461 | -1e-06 | 98.388 | 97.7469 | 127.751 | 21.0451 | 0.001828 | 0.001699 | -151.7354 | 0.2315 | 8.4028 | 0.002368 | 0.00488 | 0.004602 | 6.5e-05 | 0.000919 | 0.32676 | 1.379 | 2.5475 | 100.3 | 21.24 | 31.46 | 65.6085 | 78.5793 | 62.5026 | 0.007926 | 1.0518 | 0.025109 | 23.8721 | 0.058472 | 0.13772 | 2.3e-05 | 0.001814 | 0.001814 | 3.0899
2019-03-29 11:34 | 1553873667.8591 | -52 | 0.20946 | 0.000459 | -1e-06 | 98.1114 | 97.479 | 127.6174 | 21.0453 | 0.00181 | 0.001699 | -151.7597 | 0.23041 | 8.4014 | 0.002443 | 0.003293 | 0.00318 | 6.5e-05 | 0.000922 | 0.34722 | 1.4135 | 2.5475 | 100.3 | 21.24 | 31.454 | 65.1802 | 80.3851 | 62.5059 | 0.007924 | 1.0518 | 0.025096 | 23.8598 | 0.057842 | 0.13707 | 0.000151 | 0.011481 | 0.011481 | 0.48593
2019-03-29 11:34 | 1553873668.8601 | -51 | 0.20948 | 0.000457 | -1e-06 | 97.8127 | 97.2906 | 127.6555 | 21.0457 | 0.001805 | 0.001778 | -151.7776 | 0.23088 | 8.4016 | 0.00239 | 0.003745 | 0.003509 | 6.7e-05 | 0.000926 | 0.32676 | 1.4114 | 2.5475 | 100.3 | 21.24 | 31.45 | 65.0582 | 79.1006 | 64.389 | 0.007923 | 1.0459 | 0.025029 | 23.9302 | 0.047869 | 0.13735 | 6.5e-05 | 0.004944 | 0.004944 | 1.1306
2019-03-29 11:34 | 1553873669.8598 | -50 | 0.20947 | 0.000458 | -1e-06 | 97.4529 | 96.8842 | 127.4635 | 21.0455 | 0.001794 | 0.001713 | -151.7847 | 0.23211 | 8.4011 | 0.002441 | 0.003743 | 0.003632 | 6.7e-05 | 0.000929 | 0.32549 | 1.3723 | 2.5475 | 100.3 | 21.24 | 31.443 | 64.7903 | 80.3402 | 62.8358 | 0.007922 | 1.0508 | 0.025068 | 23.8571 | 0.052354 | 0.13809 | 0.00014 | 0.010625 | 0.010625 | 0.52896
2019-03-29 11:34 | 1553873670.8588 | -49 | 0.20947 | 0.000461 | 0 | 97.1108 | 96.5825 | 127.3726 | 21.0454 | 0.00187 | 0.001702 | -151.77 | 0.23148 | 8.4035 | 0.002455 | 0.003661 | 0.003497 | 6.8e-05 | 0.000923 | 0.32045 | 1.3665 | 2.5474 | 100.3 | 21.24 | 31.44 | 66.6219 | 80.6809 | 62.5771 | 0.007921 | 1.0516 | 0.025069 | 23.8394 | 0.048794 | 0.13771 | 0.000107 | 0.008077 | 0.008077 | 0.6939
2019-03-29 11:34 | 1553873671.8598 | -48 | 0.20948 | 0.000459 | -1e-06 | 96.8779 | 96.2135 | 127.2933 | 21.0456 | 0.001867 | 0.001735 | -151.7716 | 0.23088 | 8.4017 | 0.002456 | 0.003191 | 0.003042 | 6.5e-05 | 0.000931 | 0.34562 | 1.3245 | 2.5475 | 100.3 | 21.24 | 31.439 | 66.5463 | 80.7005 | 63.3568 | 0.007921 | 1.0491 | 0.025032 | 23.8596 | 0.061551 | 0.13735 | 5.9e-05 | 0.004467 | 0.004467 | 1.2516
2019-03-29 11:35 | 1553873672.8589 | -47 | 0.20949 | 0.000461 | -1e-06 | 96.4199 | 95.8303 | 127.3008 | 21.0468 | 0.001816 | 0.0017 | -151.7798 | 0.23057 | 8.404 | 0.002357 | 0.003643 | 0.003472 | 6.5e-05 | 0.000914 | 0.33035 | 1.3379 | 2.5473 | 100.31 | 21.24 | 31.437 | 65.3059 | 78.3059 | 62.5252 | 0.007919 | 1.0517 | 0.025064 | 23.8308 | 0.054868 | 0.13716 | 2e-06 | 0.000189 | 0.000189 | 29.5153
2019-03-29 11:35 | 1553873673.8602 | -46 | 0.20951 | 0.000454 | -1e-06 | 96.0978 | 95.4043 | 127.1844 | 21.046 | 0.001829 | 0.001709 | -151.7769 | 0.23128 | 8.4022 | 0.002402 | 0.00373 | 0.003574 | 6.4e-05 | 0.000913 | 0.35369 | 1.363 | 2.5474 | 100.31 | 21.24 | 31.436 | 65.6257 | 79.4065 | 62.7416 | 0.007919 | 1.0511 | 0.025044 | 23.8276 | 0.064783 | 0.13759 | -0.000103 | -0.007718 | -0.007718 | -0.72558
2019-03-29 11:35 | 1553873674.859 | -45 | 0.20952 | 0.000462 | -1e-06 | 95.6626 | 94.9463 | 127.3205 | 21.0477 | 0.00175 | 0.001727 | -151.7801 | 0.23102 | 8.404 | 0.002408 | 0.003704 | 0.003471 | 6.4e-05 | 0.000904 | 0.31924 | 1.3259 | 2.5474 | 100.31 | 21.24 | 31.436 | 63.722 | 79.5448 | 63.1761 | 0.007919 | 1.0497 | 0.025042 | 23.8558 | 0.067225 | 0.13744 | -0.000176 | -0.013162 | -0.013162 | -0.425
2019-03-29 11:35 | 1553873675.8589 | -44 | 0.2095 | 0.000461 | -1e-06 | 95.3109 | 94.5644 | 127.0773 | 21.0468 | 0.001806 | 0.001778 | -151.7812 | 0.23114 | 8.4013 | 0.002438 | 0.003241 | 0.00311 | 6.3e-05 | 0.000897 | 0.32189 | 1.3197 | 2.5473 | 100.31 | 21.24 | 31.437 | 65.0655 | 80.2618 | 64.4057 | 0.007919 | 1.0459 | 0.024972 | 23.8765 | 0.070329 | 0.13751 | -6.7e-05 | -0.00498 | -0.00498 | -1.1239
2019-03-29 11:35 | 1553873676.8604 | -43 | 0.20949 | 0.00046 | -1e-06 | 95.0223 | 94.1194 | 126.9931 | 21.048 | 0.001784 | 0.001715 | -151.7922 | 0.23148 | 8.4047 | 0.002418 | 0.003915 | 0.003698 | 6.6e-05 | 0.000906 | 0.36984 | 1.3208 | 2.5473 | 100.3 | 21.24 | 31.438 | 64.5451 | 79.7961 | 62.8865 | 0.00792 | 1.0506 | 0.02502 | 23.8148 | 0.085402 | 0.13771 | 7e-06 | 0.00055 | 0.00055 | 10.1819
2019-03-29 11:35 | 1553873677.8601 | -42 | 0.20948 | 0.000459 | 0 | 94.714 | 93.7508 | 126.9825 | 21.0473 | 0.001854 | 0.001718 | -151.7978 | 0.23106 | 8.3998 | 0.002487 | 0.003929 | 0.003722 | 6.3e-05 | 0.000909 | 0.33356 | 1.3391 | 2.5475 | 100.3 | 21.24 | 31.443 | 66.2303 | 81.4437 | 62.9499 | 0.007922 | 1.0504 | 0.025017 | 23.8161 | 0.091431 | 0.13746 | 3.5e-05 | 0.002645 | 0.002645 | 2.1151
2019-03-29 11:35 | 1553873678.8581 | -41 | 0.20948 | 0.00046 | -1e-06 | 94.4018 | 93.4324 | 127.1698 | 21.0475 | 0.001842 | 0.001696 | -151.8012 | 0.23116 | 8.4008 | 0.002469 | 0.003777 | 0.003631 | 6.4e-05 | 0.000912 | 0.31492 | 1.3194 | 2.5473 | 100.3 | 21.24 | 31.447 | 65.9512 | 80.998 | 62.4399 | 0.007923 | 1.052 | 0.025054 | 23.8155 | 0.09233 | 0.13752 | 6.7e-05 | 0.00512 | 0.00512 | 1.0931
2019-03-29 11:35 | 1553873679.8591 | -40 | 0.20949 | 0.000459 | -1e-06 | 94.0797 | 92.9498 | 127.1203 | 21.0478 | 0.00188 | 0.001697 | -151.8005 | 0.22998 | 8.4031 | 0.002493 | 0.004016 | 0.003866 | 6.2e-05 | 0.000912 | 0.34274 | 1.2861 | 2.5473 | 100.3 | 21.23 | 31.456 | 66.8606 | 81.5824 | 62.4631 | 0.00792 | 1.0519 | 0.025049 | 23.8117 | 0.10808 | 0.13682 | 1.2e-05 | 0.000923 | 0.000923 | 6.0338
2019-03-29 11:35 | 1553873680.8595 | -39 | 0.20949 | 0.000462 | -1e-06 | 93.8292 | 92.8126 | 127.1503 | 21.0475 | 0.001826 | 0.001761 | -151.7883 | 0.23097 | 8.4009 | 0.002492 | 0.003227 | 0.00303 | 6.3e-05 | 0.000919 | 0.33716 | 1.3433 | 2.5475 | 100.3 | 21.24 | 31.461 | 65.56 | 81.5532 | 63.9966 | 0.007926 | 1.0472 | 0.024994 | 23.8689 | 0.09744 | 0.1374 | -3.9e-05 | -0.002915 | -0.002915 | -1.9182
2019-03-29 11:35 | 1553873681.8592 | -38 | 0.2095 | 0.00046 | -1e-06 | 93.5092 | 92.4082 | 126.8546 | 21.0475 | 0.001827 | 0.001766 | -151.7948 | 0.23061 | 8.4022 | 0.002499 | 0.00375 | 0.003551 | 6.5e-05 | 0.000921 | 0.3314 | 1.3215 | 2.5473 | 100.3 | 21.24 | 31.466 | 65.5914 | 81.727 | 64.1122 | 0.007927 | 1.0468 | 0.024961 | 23.8452 | 0.10594 | 0.13719 | -6.9e-05 | -0.005105 | -0.005105 | -1.0937
2019-03-29 11:35 | 1553873682.8592 | -37 | 0.20947 | 0.000456 | -1e-06 | 93.1731 | 91.9754 | 126.6133 | 21.0484 | 0.001868 | 0.001698 | -151.8032 | 0.23008 | 8.4036 | 0.002386 | 0.002674 | 0.002514 | 6.5e-05 | 0.00092 | 0.31293 | 1.3158 | 2.5474 | 100.3 | 21.24 | 31.475 | 66.5598 | 79.0143 | 62.48 | 0.00793 | 1.0519 | 0.024998 | 23.7648 | 0.11572 | 0.13687 | 0.000141 | 0.010626 | 0.010626 | 0.52426
2019-03-29 11:35 | 1553873683.8586 | -36 | 0.20952 | 0.000463 | -1e-06 | 92.7773 | 91.6225 | 126.4366 | 21.0478 | 0.001805 | 0.00172 | -151.8089 | 0.23071 | 8.4026 | 0.002422 | 0.004175 | 0.003947 | 6.8e-05 | 0.000917 | 0.31564 | 1.3349 | 2.5474 | 100.3 | 21.24 | 31.482 | 65.0561 | 79.8918 | 63.0137 | 0.007931 | 1.0502 | 0.024961 | 23.7671 | 0.11203 | 0.13725 | -0.000158 | -0.011765 | -0.011765 | -0.4748
2019-03-29 11:35 | 1553873684.8583 | -35 | 0.20953 | 0.000456 | -1e-06 | 92.5121 | 91.1546 | 126.0608 | 21.0492 | 0.001857 | 0.001747 | -151.8112 | 0.23156 | 8.4004 | 0.002441 | 0.004201 | 0.004071 | 6.7e-05 | 0.000917 | 0.34108 | 1.3652 | 2.5475 | 100.3 | 21.23 | 31.488 | 66.3078 | 80.3444 | 63.6531 | 0.007928 | 1.0482 | 0.0249 | 23.7543 | 0.13223 | 0.13775 | -0.000239 | -0.017871 | -0.017871 | -0.31372
2019-03-29 11:35 | 1553873685.8593 | -34 | 0.20958 | 0.000457 | -1e-06 | 92.0449 | 90.9167 | 126.2058 | 21.0486 | 0.001855 | 0.001666 | -151.8104 | 0.23126 | 8.4036 | 0.002447 | 0.00356 | 0.003346 | 6.8e-05 | 0.000917 | 0.33035 | 1.3477 | 2.5475 | 100.3 | 21.23 | 31.501 | 66.255 | 80.4789 | 61.6988 | 0.007931 | 1.0543 | 0.024987 | 23.6989 | 0.11031 | 0.13758 | -0.00057 | -0.042788 | -0.042788 | -0.13087
2019-03-29 11:35 | 1553873686.8593 | -33 | 0.20948 | 0.000459 | -1e-06 | 91.6282 | 90.4062 | 126.234 | 21.0486 | 0.001835 | 0.001677 | -151.8247 | 0.23109 | 8.4015 | 0.002414 | 0.003847 | 0.003746 | 6.6e-05 | 0.000907 | 0.34324 | 1.3251 | 2.5474 | 100.3 | 21.23 | 31.507 | 65.7843 | 79.6883 | 61.9637 | 0.007933 | 1.0535 | 0.02498 | 23.7109 | 0.1201 | 0.13747 | 7.8e-05 | 0.005931 | 0.005931 | 0.94338
2019-03-29 11:35 | 1553873687.8584 | -32 | 0.20947 | 0.00046 | -1e-06 | 91.222 | 89.8151 | 126.5356 | 21.0497 | 0.001836 | 0.00176 | -151.8157 | 0.2312 | 8.4018 | 0.002493 | 0.003484 | 0.003259 | 6.7e-05 | 0.000903 | 0.32117 | 1.3192 | 2.5474 | 100.3 | 21.23 | 31.52 | 65.7934 | 81.5967 | 63.9675 | 0.007936 | 1.0472 | 0.024935 | 23.8101 | 0.13903 | 0.13754 | 0.000136 | 0.010232 | 0.010232 | 0.54708
2019-03-29 11:35 | 1553873688.8583 | -31 | 0.20949 | 0.000458 | -1e-06 | 90.7937 | 89.3004 | 126.5495 | 21.0486 | 0.001809 | 0.001735 | -151.784 | 0.23211 | 8.4029 | 0.002448 | 0.003986 | 0.003778 | 7.1e-05 | 0.000907 | 0.34595 | 1.3223 | 2.5474 | 100.3 | 21.23 | 31.526 | 65.1592 | 80.5093 | 63.3566 | 0.007938 | 1.0491 | 0.024959 | 23.7898 | 0.14833 | 0.13808 | -3.6e-05 | -0.002642 | -0.002642 | -2.1271
2019-03-29 11:35 | 1553873689.8584 | -30 | 0.20948 | 0.000459 | -1e-06 | 90.4635 | 88.9642 | 126.5549 | 21.0491 | 0.001841 | 0.001736 | -151.7887 | 0.2313 | 8.4033 | 0.002505 | 0.004419 | 0.004287 | 6.9e-05 | 0.0009 | 0.32333 | 1.3348 | 2.5473 | 100.3 | 21.23 | 31.532 | 65.92 | 81.8877 | 63.3835 | 0.007939 | 1.0491 | 0.024959 | 23.7913 | 0.14948 | 0.1376 | 6.7e-05 | 0.00508 | 0.00508 | 1.1023
2019-03-29 11:35 | 1553873690.8585 | -29 | 0.20947 | 0.000459 | 0 | 90.3839 | 88.8502 | 126.6255 | 21.0483 | 0.001836 | 0.001719 | -151.7974 | 0.23125 | 8.3996 | 0.002506 | 0.003841 | 0.003696 | 6.9e-05 | 0.000901 | 0.33467 | 1.3533 | 2.5474 | 100.3 | 21.23 | 31.542 | 65.8059 | 81.8876 | 62.9783 | 0.007942 | 1.0503 | 0.024981 | 23.7836 | 0.15308 | 0.13757 | 0.000144 | 0.010803 | 0.010803 | 0.51831
2019-03-29 11:35 | 1553873691.8595 | -28 | 0.20948 | 0.000462 | -1e-06 | 90.4075 | 88.9635 | 126.8483 | 21.0497 | 0.001854 | 0.001681 | -151.7997 | 0.23062 | 8.4018 | 0.002462 | 0.00411 | 0.003888 | 6.9e-05 | 0.000901 | 0.32211 | 1.3345 | 2.5474 | 100.3 | 21.23 | 31.547 | 66.2367 | 80.8393 | 62.0782 | 0.007943 | 1.0531 | 0.025036 | 23.7726 | 0.14402 | 0.1372 | 3.6e-05 | 0.002774 | 0.002774 | 2.0132
2019-03-29 11:35 | 1553873692.8599 | -27 | 0.20953 | 0.000461 | -1e-06 | 90.5431 | 88.9902 | 127.044 | 21.0491 | 0.001889 | 0.001698 | -151.7998 | 0.23034 | 8.4017 | 0.002473 | 0.003836 | 0.003729 | 7.1e-05 | 0.000905 | 0.33085 | 1.3536 | 2.5474 | 100.3 | 21.23 | 31.556 | 67.0764 | 81.1164 | 62.4808 | 0.007945 | 1.0519 | 0.02504 | 23.8052 | 0.15475 | 0.13703 | -0.000248 | -0.018588 | -0.018588 | -0.30004
2019-03-29 11:35 | 1553873693.8585 | -26 | 0.20953 | 0.000462 | -1e-06 | 90.7964 | 89.3297 | 127.2663 | 21.049 | 0.00181 | 0.001763 | -151.8046 | 0.23047 | 8.4022 | 0.002436 | 0.003706 | 0.003518 | 6.8e-05 | 0.000903 | 0.32117 | 1.345 | 2.5474 | 100.3 | 21.23 | 31.562 | 65.1721 | 80.2189 | 64.0434 | 0.007947 | 1.047 | 0.025004 | 23.8814 | 0.14566 | 0.13711 | -0.000222 | -0.016618 | -0.016618 | -0.3358
2019-03-29 11:35 | 1553873694.8582 | -25 | 0.20947 | 0.00046 | -1e-06 | 91.1131 | 89.7482 | 127.4226 | 21.0491 | 0.001871 | 0.001668 | -151.7961 | 0.23056 | 8.4015 | 0.002453 | 0.004016 | 0.003838 | 7.3e-05 | 0.000906 | 0.30757 | 1.3424 | 2.5473 | 100.3 | 21.23 | 31.566 | 66.6377 | 80.6362 | 61.7595 | 0.007948 | 1.0542 | 0.025105 | 23.815 | 0.13501 | 0.13716 | 0.000146 | 0.011105 | 0.011105 | 0.50268
2019-03-29 11:35 | 1553873695.8591 | -24 | 0.20949 | 0.000459 | -1e-06 | 91.454 | 90.1219 | 127.5231 | 21.0497 | 0.00184 | 0.001717 | -151.7915 | 0.23084 | 8.4011 | 0.002443 | 0.003718 | 0.003557 | 7.1e-05 | 0.000911 | 0.32709 | 1.3041 | 2.5473 | 100.3 | 21.23 | 31.573 | 65.8947 | 80.377 | 62.9405 | 0.007949 | 1.0504 | 0.02507 | 23.8664 | 0.13125 | 0.13733 | -2.5e-05 | -0.001814 | -0.001814 | -3.0811
2019-03-29 11:35 | 1553873696.858 | -23 | 0.2095 | 0.000456 | -1e-06 | 91.8009 | 90.4434 | 127.7808 | 21.0491 | 0.001843 | 0.001753 | -151.7737 | 0.23184 | 8.4012 | 0.002356 | 0.00411 | 0.003889 | 7.2e-05 | 0.00091 | 0.34003 | 1.3088 | 2.5474 | 100.3 | 21.23 | 31.575 | 65.9754 | 78.2887 | 63.8038 | 0.00795 | 1.0478 | 0.025064 | 23.9211 | 0.13326 | 0.13792 | -5.1e-05 | -0.003798 | -0.003798 | -1.4781
2019-03-29 11:35 | 1553873697.8577 | -22 | 0.2095 | 0.000462 | -2e-06 | 92.238 | 90.9505 | 127.9432 | 21.051 | 0.001812 | 0.001701 | -151.7703 | 0.2314 | 8.4005 | 0.002434 | 0.003501 | 0.003362 | 7e-05 | 0.000908 | 0.32012 | 1.3054 | 2.5473 | 100.3 | 21.23 | 31.58 | 65.2164 | 80.1638 | 62.5442 | 0.007951 | 1.0517 | 0.025126 | 23.8916 | 0.12573 | 0.13766 | -5e-05 | -0.003705 | -0.003705 | -1.5121
2019-03-29 11:35 | 1553873698.8577 | -21 | 0.2095 | 0.00046 | -1e-06 | 92.584 | 91.4409 | 127.9451 | 21.0498 | 0.001813 | 0.0017 | -151.7714 | 0.23062 | 8.3996 | 0.002487 | 0.004145 | 0.003974 | 7.1e-05 | 0.000911 | 0.34291 | 1.2985 | 2.5473 | 100.3 | 21.23 | 31.582 | 65.2534 | 81.4446 | 62.5172 | 0.007952 | 1.0518 | 0.025128 | 23.8908 | 0.11112 | 0.1372 | -5.8e-05 | -0.004344 | -0.004344 | -1.2855
2019-03-29 11:35 | 1553873699.8587 | -20 | 0.2095 | 0.00046 | -1e-06 | 92.9896 | 91.7109 | 128.15 | 21.0504 | 0.001831 | 0.001651 | -151.7693 | 0.23103 | 8.403 | 0.00241 | 0.003804 | 0.003682 | 7.1e-05 | 0.00091 | 0.31708 | 1.3151 | 2.5474 | 100.3 | 21.23 | 31.584 | 65.6754 | 79.5824 | 61.359 | 0.007952 | 1.0554 | 0.025191 | 23.8686 | 0.12385 | 0.13744 | -7.7e-05 | -0.005743 | -0.005743 | -0.97412
2019-03-29 11:35 | 1553873700.8589 | -19 | 0.20952 | 0.000457 | -1e-06 | 93.2913 | 92.2358 | 128.2432 | 21.0504 | 0.001822 | 0.001699 | -151.7786 | 0.23142 | 8.4017 | 0.002436 | 0.003786 | 0.003666 | 7.4e-05 | 0.00092 | 0.31238 | 1.3533 | 2.5475 | 100.3 | 21.23 | 31.585 | 65.4575 | 80.2202 | 62.4952 | 0.007952 | 1.0518 | 0.025158 | 23.9178 | 0.10178 | 0.13767 | -0.000166 | -0.012513 | -0.012513 | -0.44779
2019-03-29 11:35 | 1553873701.8578 | -18 | 0.2095 | 0.000453 | -1e-06 | 93.5876 | 92.5126 | 128.3681 | 21.0505 | 0.0018 | 0.001732 | -151.7527 | 0.2305 | 8.4051 | 0.002398 | 0.003078 | 0.002886 | 6.9e-05 | 0.000925 | 0.32532 | 1.3694 | 2.5475 | 100.3 | 21.23 | 31.586 | 64.9231 | 79.3002 | 63.2968 | 0.007953 | 1.0493 | 0.02514 | 23.958 | 0.10334 | 0.13712 | -7e-05 | -0.005194 | -0.005194 | -1.0744
2019-03-29 11:35 | 1553873702.8583 | -17 | 0.20951 | 0.00046 | -1e-06 | 93.9112 | 92.858 | 128.4758 | 21.0508 | 0.001838 | 0.001706 | -151.755 | 0.2313 | 8.4029 | 0.0024 | 0.004141 | 0.003958 | 7.2e-05 | 0.000925 | 0.31061 | 1.3281 | 2.5474 | 100.3 | 21.23 | 31.59 | 65.8592 | 79.3473 | 62.6839 | 0.007954 | 1.0513 | 0.025173 | 23.9462 | 0.10088 | 0.1376 | -9.9e-05 | -0.007395 | -0.007395 | -0.75728
2019-03-29 11:35 | 1553873703.8593 | -16 | 0.20953 | 0.000458 | -1e-06 | 94.3396 | 93.3344 | 128.7831 | 21.0509 | 0.0019 | 0.001751 | -151.7507 | 0.2309 | 8.4037 | 0.002528 | 0.003858 | 0.003685 | 7.3e-05 | 0.000927 | 0.29894 | 1.3131 | 2.5475 | 100.3 | 21.23 | 31.592 | 67.337 | 82.437 | 63.7539 | 0.007954 | 1.0479 | 0.025163 | 24.013 | 0.095811 | 0.13736 | -0.000225 | -0.016981 | -0.016981 | -0.32923
2019-03-29 11:35 | 1553873704.8589 | -15 | 0.2095 | 0.000456 | -1e-06 | 94.5299 | 93.6146 | 128.7967 | 21.0507 | 0.001784 | 0.001729 | -151.7624 | 0.23103 | 8.4029 | 0.002545 | 0.003724 | 0.003496 | 7.2e-05 | 0.000922 | 0.31868 | 1.3436 | 2.5472 | 100.3 | 21.23 | 31.594 | 64.5414 | 82.8375 | 63.2288 | 0.007955 | 1.0495 | 0.025184 | 23.9955 | 0.087031 | 0.13744 | -5.8e-05 | -0.004348 | -0.004348 | -1.2865
2019-03-29 11:35 | 1553873705.8593 | -14 | 0.20948 | 0.000459 | -2e-06 | 94.8125 | 93.9091 | 128.7455 | 21.0513 | 0.001829 | 0.001718 | -151.7498 | 0.23191 | 8.4028 | 0.002548 | 0.004562 | 0.004442 | 7.1e-05 | 0.000921 | 0.3131 | 1.3942 | 2.5474 | 100.3 | 21.23 | 31.595 | 65.6403 | 82.9007 | 62.9624 | 0.007955 | 1.0504 | 0.025189 | 23.9813 | 0.085631 | 0.13796 | 6.2e-05 | 0.004817 | 0.004817 | 1.1657
2019-03-29 11:35 | 1553873706.859 | -13 | 0.20951 | 0.000461 | -2e-06 | 95.028 | 94.1303 | 128.9345 | 21.0518 | 0.001859 | 0.00166 | -151.7548 | 0.2314 | 8.4052 | 0.002578 | 0.003744 | 0.003566 | 7.3e-05 | 0.000918 | 0.30613 | 1.3256 | 2.5473 | 100.3 | 21.23 | 31.595 | 66.3479 | 83.6391 | 61.5543 | 0.007955 | 1.0548 | 0.025261 | 23.9485 | 0.0849 | 0.13766 | -0.000114 | -0.008533 | -0.008533 | -0.65663
2019-03-29 11:35 | 1553873707.8594 | -12 | 0.2095 | 0.000455 | -1e-06 | 95.2507 | 94.3702 | 128.9828 | 21.0516 | 0.001798 | 0.001656 | -151.7709 | 0.23132 | 8.4032 | 0.00257 | 0.004112 | 0.003975 | 7.1e-05 | 0.000915 | 0.31166 | 1.3508 | 2.5475 | 100.3 | 21.23 | 31.594 | 64.8789 | 83.433 | 61.4736 | 0.007955 | 1.0551 | 0.025269 | 23.9501 | 0.083068 | 0.13761 | -3.6e-05 | -0.002676 | -0.002676 | -2.093
2019-03-29 11:35 | 1553873708.858 | -11 | 0.2095 | 0.000462 | -1e-06 | 95.5081 | 94.7572 | 128.9223 | 21.0517 | 0.001816 | 0.001707 | -140.514 | 0.2307 | 8.4017 | 0.0026 | 0.004213 | 0.004008 | 7.1e-05 | 0.00092 | 0.32012 | 1.3287 | 2.5475 | 100.3 | 21.23 | 31.594 | 65.3207 | 84.1631 | 62.6949 | 0.007955 | 1.0512 | 0.025217 | 23.9882 | 0.070601 | 0.13724 | -5.4e-05 | -0.004051 | -0.004051 | -1.3789
2019-03-29 11:35 | 1553873709.8581 | -10 | 0.2095 | 0.000469 | -1e-06 | 95.7045 | 94.7897 | 128.9481 | 21.0528 | 0.001801 | 0.001698 | -107.4993 | 0.23083 | 8.4004 | 0.002632 | 0.004333 | 0.004219 | 7.3e-05 | 0.000922 | 0.31581 | 1.3331 | 2.5474 | 100.3 | 21.23 | 31.592 | 64.966 | 84.9415 | 62.4896 | 0.007954 | 1.0519 | 0.025227 | 23.9832 | 0.085912 | 0.13732 | -7.1e-05 | -0.005282 | -0.005282 | -1.0581
2019-03-29 11:35 | 1553873710.8571 | -9 | 0.20949 | 0.00046 | -1e-06 | 95.8536 | 94.9299 | 129.1622 | 21.0521 | 0.001791 | 0.001753 | -74.4476 | 0.23084 | 8.4027 | 0.002586 | 0.004514 | 0.004345 | 7.1e-05 | 0.000914 | 0.32875 | 1.3374 | 2.5475 | 100.3 | 21.23 | 31.589 | 64.7302 | 83.8187 | 63.7972 | 0.007953 | 1.0478 | 0.025199 | 24.0499 | 0.086607 | 0.13733 | -3.3e-05 | -0.002416 | -0.002416 | -2.3135
2019-03-29 11:35 | 1553873711.8581 | -8 | 0.20949 | 0.000461 | -1e-06 | 95.9362 | 95.104 | 129.1051 | 21.0528 | 0.00182 | 0.001629 | -41.4004 | 0.23165 | 8.4034 | 0.002566 | 0.003796 | 0.003606 | 7.3e-05 | 0.000914 | 0.32156 | 1.3191 | 2.5473 | 100.3 | 21.23 | 31.588 | 65.4184 | 83.3514 | 60.8128 | 0.007953 | 1.0571 | 0.025306 | 23.9378 | 0.077928 | 0.13781 | 5e-06 | 0.000427 | 0.000427 | 13.1225
2019-03-29 11:35 | 1553873712.8581 | -7 | 0.20951 | 0.000462 | -2e-06 | 96.0544 | 95.4187 | 128.9004 | 21.0525 | 0.001816 | 0.001613 | -8.3418 | 0.23094 | 8.4038 | 0.002569 | 0.003832 | 0.003688 | 7.1e-05 | 0.000912 | 0.31686 | 1.3307 | 2.5473 | 100.3 | 21.23 | 31.586 | 65.3232 | 83.4284 | 60.4309 | 0.007953 | 1.0583 | 0.0253 | 23.9052 | 0.059399 | 0.13738 | -0.000137 | -0.010361 | -0.010361 | -0.53969
2019-03-29 11:35 | 1553873713.8572 | -6 | 0.20951 | 0.000463 | -1e-06 | 96.2385 | 95.6245 | 128.65 | 21.0528 | 0.001768 | 0.001673 | 24.7208 | 0.23957 | 8.4022 | 0.002557 | 0.003507 | 0.003301 | 7.1e-05 | 0.000911 | 0.30359 | 1.3345 | 2.5474 | 100.3 | 21.23 | 31.585 | 64.1611 | 83.1235 | 61.8686 | 0.007952 | 1.0538 | 0.025221 | 23.9333 | 0.057254 | 0.14252 | -0.000122 | -0.009206 | -0.009206 | -0.63008
2019-03-29 11:35 | 1553873714.8572 | -5 | 0.20951 | 0.000461 | -1e-06 | 96.3672 | 95.721 | 128.5622 | 21.0526 | 0.001784 | 0.001666 | 178.3401 | 0.25613 | 8.4029 | 0.002611 | 0.003936 | 0.003862 | 7.3e-05 | 0.000912 | 0.31885 | 1.3112 | 2.5475 | 100.3 | 21.23 | 31.584 | 64.5457 | 84.424 | 61.7107 | 0.007952 | 1.0543 | 0.025218 | 23.9195 | 0.060179 | 0.15237 | -0.000121 | -0.009098 | -0.009098 | -0.68167
2019-03-29 11:35 | 1553873715.8582 | -4 | 0.20947 | 0.000465 | -1e-06 | 96.5304 | 95.6801 | 128.5636 | 21.0528 | 0.001786 | 0.001614 | 178.6966 | 0.25641 | 8.4023 | 0.002625 | 0.004178 | 0.00403 | 7.3e-05 | 0.000914 | 0.32388 | 1.3301 | 2.5475 | 100.3 | 21.23 | 31.581 | 64.5935 | 84.7729 | 60.4625 | 0.007951 | 1.0582 | 0.025266 | 23.875 | 0.079136 | 0.15254 | 0.00011 | 0.008458 | 0.008458 | 0.734
2019-03-29 11:35 | 1553873716.8584 | -3 | 0.20947 | 0.000467 | -2e-06 | 96.6603 | 96.0357 | 128.0613 | 21.0533 | 0.001808 | 0.001659 | 178.8502 | 0.2565 | 8.4028 | 0.002604 | 0.004156 | 0.004079 | 7.4e-05 | 0.000916 | 0.31421 | 1.3705 | 2.5476 | 100.3 | 21.23 | 31.58 | 65.1397 | 84.2505 | 61.5324 | 0.007951 | 1.0549 | 0.025176 | 23.8665 | 0.057982 | 0.15259 | 9.3e-05 | 0.007125 | 0.007125 | 0.87162
2019-03-29 11:35 | 1553873717.8573 | -2 | 0.20951 | 0.000464 | -1e-06 | 96.7959 | 95.9296 | 127.4933 | 21.0527 | 0.001801 | 0.001669 | 178.7737 | 0.25652 | 8.4007 | 0.002628 | 0.004626 | 0.004562 | 7.4e-05 | 0.000915 | 0.2975 | 1.3464 | 2.5474 | 100.3 | 21.23 | 31.578 | 64.9565 | 84.8291 | 61.7943 | 0.007951 | 1.054 | 0.02511 | 23.8229 | 0.080412 | 0.1526 | -0.000113 | -0.008443 | -0.008443 | -0.73567
2019-03-29 11:35 | 1553873718.8583 | -1 | 0.20953 | 0.000459 | -1e-06 | 96.9839 | 96.3579 | 127.3884 | 21.053 | 0.001782 | 0.001673 | 178.836 | 0.2564 | 8.3999 | 0.002577 | 0.003473 | 0.00339 | 7e-05 | 0.000916 | 0.33069 | 1.3649 | 2.5519 | 100.3 | 21.23 | 31.578 | 64.4939 | 83.6058 | 61.8719 | 0.007951 | 1.0538 | 0.025097 | 23.8158 | 0.057918 | 0.15253 | -0.000246 | -0.018468 | -0.018468 | -0.33615
2019-03-29 11:35 | 1553873719.8574 | 0 | 0.20948 | 0.000464 | -1e-06 | 92.5322 | 95.4122 | 126.7466 | 21.0413 | 0.00179 | 0.001652 | 178.9422 | 0.25725 | 8.4032 | 0.002605 | 0.004185 | 0.004446 | 7.3e-05 | 0.000919 | 0.29479 | 1.34 | 2.7175 | 100.3 | 21.23 | 31.578 | 64.6932 | 84.2764 | 61.3518 | 0.007951 | 1.0554 | 0.025053 | 23.7373 | -0.27415 | 0.15304 | 6e-05 | 0.004607 | 0.004607 | 1.3519
2019-03-29 11:35 | 1553873720.8584 | 1 | 0.20949 | 0.000464 | -1e-06 | 81.2157 | 90.8693 | 125.5174 | 21.0366 | 0.001801 | 0.001673 | 178.9458 | 0.25696 | 8.4044 | 0.002622 | 0.004028 | 0.003704 | 6.9e-05 | 0.000928 | 0.29894 | 1.339 | 2.5474 | 100.3 | 21.23 | 31.581 | 64.944 | 84.6713 | 61.8668 | 0.007951 | 1.0538 | 0.024912 | 23.6401 | -1.0046 | 0.15286 | 2.6e-05 | 0.001976 | 0.001976 | 3.148
2019-03-29 11:35 | 1553873721.8575 | 2 | 0.20946 | 0.000468 | -1e-06 | 76.7101 | 90.027 | 124.1194 | 21.0354 | 0.001744 | 0.001642 | 178.996 | 0.25686 | 8.4023 | 0.002595 | 0.003707 | 0.004048 | 7.2e-05 | 0.00093 | 0.3298 | 1.3453 | 2.5474 | 100.3 | 21.23 | 31.583 | 63.564 | 84.018 | 61.1209 | 0.007952 | 1.0562 | 0.024801 | 23.4819 | -1.4318 | 0.15281 | 0.000144 | 0.010774 | 0.010774 | 0.57723
2019-03-29 11:35 | 1553873722.8585 | 3 | 0.20944 | 0.000471 | 0 | 77.6827 | 92.6686 | 122.8589 | 21.0417 | 0.001753 | 0.001703 | 178.7595 | 0.25647 | 8.4026 | 0.002572 | 0.003924 | 0.003945 | 6.8e-05 | 0.00093 | 0.29999 | 1.3199 | 2.5475 | 100.3 | 21.23 | 31.588 | 63.7965 | 83.4717 | 62.5797 | 0.007953 | 1.0516 | 0.024621 | 23.4133 | -1.5778 | 0.15257 | 0.000271 | 0.02009 | 0.02009 | 0.30909
2019-03-29 11:35 | 1553873723.8575 | 4 | 0.20944 | 0.00047 | 2e-06 | 81.5533 | 94.6712 | 121.3154 | 21.0427 | 0.001804 | 0.001761 | 178.9263 | 0.25637 | 8.4026 | 0.002542 | 0.004082 | 0.004045 | 6.8e-05 | 0.000931 | 0.32228 | 1.3177 | 2.5474 | 100.3 | 21.23 | 31.589 | 65.0342 | 82.7714 | 63.9802 | 0.007953 | 1.0472 | 0.024415 | 23.3142 | -1.3341 | 0.15251 | 0.000277 | 0.020219 | 0.020219 | 0.307
2019-03-29 11:35 | 1553873724.8585 | 5 | 0.20939 | 0.000468 | 5e-06 | 87.2457 | 95.7183 | 119.4554 | 21.0429 | 0.001803 | 0.001805 | 179.108 | 0.25611 | 8.4028 | 0.002618 | 0.003743 | 0.003707 | 6.8e-05 | 0.00093 | 0.31868 | 1.3251 | 2.5474 | 100.3 | 21.23 | 31.591 | 64.9978 | 84.5885 | 65.0476 | 0.007954 | 1.0439 | 0.024189 | 23.1714 | -0.82899 | 0.15236 | 0.000557 | 0.040215 | 0.040215 | 0.1542
2019-03-29 11:35 | 1553873725.8565 | 6 | 0.2091 | 0.000479 | 1e-05 | 91.6857 | 95.9118 | 117.4051 | 21.0392 | 0.001911 | 0.001795 | 178.8906 | 0.25627 | 8.4042 | 0.002505 | 0.004317 | 0.00425 | 6.6e-05 | 0.000927 | 0.33019 | 1.3277 | 1.7276 | 100.3 | 21.23 | 31.592 | 67.5864 | 81.8726 | 64.8029 | 0.007954 | 1.0447 | 0.023989 | 22.9634 | -0.40307 | 0.15246 | 0.002293 | 0.16482 | 0.16482 | 0.037648
2019-03-29 11:35 | 1553873726.8576 | 7 | 0.208 | 0.000508 | 1.7e-05 | 94.6793 | 95.4149 | 115.5325 | 21.0533 | 0.001854 | 0.001923 | 178.8477 | 0.2559 | 8.4029 | 0.002559 | 0.003855 | 0.003807 | 6.2e-05 | 0.000926 | 0.28512 | 1.337 | 2.5474 | 100.3 | 21.23 | 31.592 | 66.2298 | 83.1692 | 67.8924 | 0.007954 | 1.0352 | 0.023689 | 22.8834 | -0.069226 | 0.15224 | 0.008889 | 0.63164 | 0.63164 | 0.009809
2019-03-29 11:35 | 1553873727.8576 | 8 | 0.2068 | 0.000608 | 2.8e-05 | 96.3412 | 93.3757 | 113.4861 | 21.0533 | 0.001944 | 0.00195 | 178.8051 | 0.25614 | 8.402 | 0.002585 | 0.003734 | 0.003612 | 6.4e-05 | 0.000925 | 0.28998 | 1.3115 | 2.5474 | 100.3 | 21.23 | 31.592 | 68.4049 | 83.8122 | 68.5436 | 0.007954 | 1.0332 | 0.023456 | 22.7015 | 0.27965 | 0.15238 | 0.015977 | 1.1235 | 1.1235 | 0.00552
2019-03-29 11:35 | 1553873728.8577 | 9 | 0.20565 | 0.001055 | 4.1e-05 | 97.1646 | 88.9222 | 111.4654 | 21.054 | 0.002056 | 0.00198 | 178.7342 | 0.25592 | 8.403 | 0.002636 | 0.004419 | 0.004455 | 6.5e-05 | 0.000925 | 0.27903 | 1.337 | 2.5474 | 100.3 | 21.23 | 31.585 | 71.0833 | 85.023 | 69.2681 | 0.007952 | 1.031 | 0.023221 | 22.5223 | 0.79289 | 0.15225 | 0.022279 | 1.5498 | 1.5498 | 0.003998
2019-03-29 11:35 | 1553873729.8577 | 10 | 0.20464 | 0.001715 | 5.4e-05 | 97.5795 | 84.8793 | 110.1411 | 21.0548 | 0.002093 | 0.002087 | 178.6457 | 0.25712 | 8.4036 | 0.002549 | 0.004009 | 0.003907 | 6.6e-05 | 0.000929 | 0.30397 | 1.3407 | 2.5474 | 100.3 | 21.23 | 31.58 | 71.9722 | 82.9353 | 71.8319 | 0.007951 | 1.0234 | 0.022997 | 22.4717 | 1.2472 | 0.15296 | 0.027508 | 1.8938 | 1.8938 | 0.003287
2019-03-29 11:35 | 1553873730.8567 | 11 | 0.20382 | 0.002471 | 7e-05 | 97.836 | 82.0852 | 109.2463 | 21.054 | 0.002148 | 0.002137 | 178.5441 | 0.25651 | 8.4022 | 0.002516 | 0.004174 | 0.004107 | 6.4e-05 | 0.000924 | 0.29789 | 1.3419 | 2.5474 | 100.3 | 21.23 | 31.567 | 73.3139 | 82.146 | 73.0409 | 0.007948 | 1.0198 | 0.022863 | 22.4195 | 1.5701 | 0.1526 | 0.031463 | 2.1522 | 2.1522 | 0.002886
2019-03-29 11:36 | 1553873731.8577 | 12 | 0.20263 | 0.003328 | 8.3e-05 | 98.01 | 78.1273 | 108.9149 | 21.0541 | 0.002194 | 0.002185 | 178.4251 | 0.25594 | 8.4018 | 0.002597 | 0.003813 | 0.003686 | 6.4e-05 | 0.000925 | 0.29623 | 1.3316 | 2.5474 | 100.3 | 21.23 | 31.558 | 74.4145 | 84.082 | 74.1978 | 0.007946 | 1.0164 | 0.022791 | 22.4228 | 2.028 | 0.15226 | 0.037527 | 2.5572 | 2.5572 | 0.002423
2019-03-29 11:36 | 1553873732.8568 | 13 | 0.20197 | 0.004255 | 9.4e-05 | 98.0597 | 74.3513 | 108.6091 | 21.0552 | 0.002265 | 0.002273 | 178.3287 | 0.25728 | 8.4021 | 0.002549 | 0.004194 | 0.004067 | 6.5e-05 | 0.000926 | 0.29607 | 1.3622 | 2.5474 | 100.3 | 21.23 | 31.55 | 76.1103 | 82.9309 | 76.2989 | 0.007944 | 1.0103 | 0.02269 | 22.4589 | 2.4756 | 0.15305 | 0.040327 | 2.7348 | 2.7348 | 0.002278
2019-03-29 11:36 | 1553873733.8568 | 14 | 0.20112 | 0.005177 | 0.000104 | 98.186 | 72.2223 | 108.2441 | 21.0551 | 0.002383 | 0.00235 | 178.2183 | 0.25704 | 8.4027 | 0.002714 | 0.003795 | 0.003683 | 6.5e-05 | 0.000922 | 0.27383 | 1.3353 | 2.5475 | 100.3 | 21.23 | 31.532 | 78.9547 | 86.9138 | 78.1541 | 0.007939 | 1.005 | 0.022592 | 22.4806 | 2.747 | 0.15291 | 0.044237 | 2.9857 | 2.9857 | 0.002084
2019-03-29 11:36 | 1553873734.8579 | 15 | 0.20048 | 0.005857 | 0.00011 | 98.221 | 70.1681 | 108.407 | 21.0558 | 0.00238 | 0.002354 | 178.1033 | 0.25628 | 8.4023 | 0.002668 | 0.00452 | 0.004401 | 6.7e-05 | 0.000921 | 0.31094 | 1.3202 | 2.5474 | 100.3 | 21.23 | 31.523 | 78.8735 | 85.806 | 78.2458 | 0.007937 | 1.0047 | 0.022606 | 22.5004 | 3.0083 | 0.15246 | 0.047214 | 3.1877 | 3.1877 | 0.001947
2019-03-29 11:36 | 1553873735.8569 | 16 | 0.20045 | 0.006281 | 0.000115 | 98.2336 | 69.2084 | 109.1421 | 21.0559 | 0.002483 | 0.002371 | 178.0085 | 0.2561 | 8.4047 | 0.002743 | 0.003872 | 0.003791 | 6.7e-05 | 0.000922 | 0.30309 | 1.3159 | 2.5473 | 100.3 | 21.23 | 31.506 | 81.3529 | 87.6009 | 78.6655 | 0.007933 | 1.0035 | 0.022669 | 22.5901 | 3.1326 | 0.15235 | 0.046871 | 3.1732 | 3.1732 | 0.001954
2019-03-29 11:36 | 1553873736.8579 | 17 | 0.20046 | 0.006499 | 0.000117 | 98.3347 | 69.2429 | 109.9158 | 21.0565 | 0.002434 | 0.002405 | 177.9043 | 0.25665 | 8.4033 | 0.002813 | 0.003306 | 0.003182 | 6.7e-05 | 0.000919 | 0.30702 | 1.3403 | 2.5475 | 100.3 | 21.23 | 31.496 | 80.1863 | 89.2956 | 79.4737 | 0.00793 | 1.0012 | 0.022723 | 22.696 | 3.1374 | 0.15268 | 0.046572 | 3.1604 | 3.1604 | 0.001966
2019-03-29 11:36 | 1553873737.858 | 18 | 0.2005 | 0.006705 | 0.000118 | 98.338 | 69.4006 | 111.1185 | 21.0562 | 0.002541 | 0.002392 | 177.8078 | 0.25684 | 8.4019 | 0.00277 | 0.004266 | 0.004218 | 6.9e-05 | 0.00092 | 0.32068 | 1.3172 | 2.5474 | 100.3 | 21.23 | 31.487 | 82.7494 | 88.2578 | 79.1733 | 0.007928 | 1.0021 | 0.022857 | 22.8101 | 3.1173 | 0.15279 | 0.046103 | 3.1471 | 3.1471 | 0.001976
2019-03-29 11:36 | 1553873738.858 | 19 | 0.20073 | 0.006866 | 0.000116 | 98.3066 | 70.4258 | 111.9644 | 21.0559 | 0.002573 | 0.002419 | 177.71 | 0.25721 | 8.4011 | 0.002846 | 0.004064 | 0.004041 | 6.8e-05 | 0.00092 | 0.29789 | 1.3788 | 2.5474 | 100.3 | 21.22 | 31.473 | 83.512 | 90.0866 | 79.8091 | 0.007919 | 1.0002 | 0.022923 | 22.9174 | 2.9833 | 0.15301 | 0.044525 | 3.0485 | 3.0485 | 0.002043
2019-03-29 11:36 | 1553873739.8567 | 20 | 0.20114 | 0.006962 | 0.000113 | 98.4212 | 72.1507 | 112.2853 | 21.0561 | 0.002572 | 0.002377 | 177.5874 | 0.25568 | 8.4033 | 0.002828 | 0.004436 | 0.004333 | 6.9e-05 | 0.000928 | 0.31365 | 1.3646 | 2.5474 | 100.3 | 21.22 | 31.468 | 83.4833 | 89.66 | 78.8005 | 0.007918 | 1.0031 | 0.022989 | 22.9174 | 2.7773 | 0.1521 | 0.041987 | 2.8835 | 2.8835 | 0.002147
2019-03-29 11:36 | 1553873740.8561 | 21 | 0.20146 | 0.006989 | 0.00011 | 98.4185 | 73.3206 | 112.6761 | 21.057 | 0.00257 | 0.00249 | 177.4996 | 0.25637 | 8.4028 | 0.002889 | 0.003852 | 0.003796 | 6.7e-05 | 0.000922 | 0.28567 | 1.402 | 2.5476 | 100.3 | 21.22 | 31.46 | 83.4414 | 91.1201 | 81.5155 | 0.007916 | 0.99543 | 0.02294 | 23.0457 | 2.6332 | 0.15251 | 0.040026 | 2.7435 | 2.7435 | 0.002262
2019-03-29 11:36 | 1553873741.8565 | 22 | 0.20186 | 0.006943 | 0.000108 | 98.558 | 74.7158 | 112.8226 | 21.0563 | 0.002599 | 0.002493 | 177.4243 | 0.25709 | 8.4013 | 0.00287 | 0.004764 | 0.004672 | 6.8e-05 | 0.000919 | 0.26825 | 1.3495 | 2.5475 | 100.3 | 21.22 | 31.459 | 84.1443 | 90.6591 | 81.599 | 0.007916 | 0.9952 | 0.022953 | 23.0634 | 2.4772 | 0.15294 | 0.037654 | 2.5826 | 2.5826 | 0.00241
2019-03-29 11:36 | 1553873742.8565 | 23 | 0.20226 | 0.006833 | 0.000107 | 98.5092 | 75.0837 | 113.2378 | 21.057 | 0.002614 | 0.002498 | 177.3388 | 0.25704 | 8.3994 | 0.002848 | 0.004374 | 0.004234 | 6.8e-05 | 0.000922 | 0.29518 | 1.3355 | 2.5475 | 100.3 | 21.22 | 31.457 | 84.5 | 90.1275 | 81.706 | 0.007915 | 0.9949 | 0.022991 | 23.1092 | 2.4289 | 0.15291 | 0.035419 | 2.4338 | 2.4338 | 0.002557
2019-03-29 11:36 | 1553873743.8571 | 24 | 0.20243 | 0.006675 | 0.000105 | 98.4937 | 75.1905 | 113.9103 | 21.057 | 0.002615 | 0.002517 | 177.2482 | 0.25732 | 8.402 | 0.002883 | 0.004263 | 0.004196 | 6.6e-05 | 0.000924 | 0.29031 | 1.3292 | 2.5474 | 100.3 | 21.22 | 31.456 | 84.5336 | 90.9764 | 82.1831 | 0.007915 | 0.99356 | 0.023044 | 23.1933 | 2.4147 | 0.15308 | 0.03458 | 2.3817 | 2.3817 | 0.002616
2019-03-29 11:36 | 1553873744.8562 | 25 | 0.20264 | 0.006502 | 0.000103 | 98.5577 | 75.4086 | 114.3736 | 21.0575 | 0.002704 | 0.002512 | 177.1763 | 0.2553 | 8.4035 | 0.002931 | 0.004341 | 0.004178 | 6.9e-05 | 0.000933 | 0.30127 | 1.337 | 2.5473 | 100.3 | 21.22 | 31.456 | 86.6678 | 92.1299 | 82.0501 | 0.007915 | 0.99393 | 0.023095 | 23.2361 | 2.3946 | 0.15188 | 0.03354 | 2.3154 | 2.3154 | 0.00267
2019-03-29 11:36 | 1553873745.8572 | 26 | 0.20286 | 0.006325 | 0.0001 | 98.5218 | 75.5936 | 114.7363 | 21.0572 | 0.002702 | 0.002511 | 177.0798 | 0.25633 | 8.4023 | 0.00293 | 0.004354 | 0.004304 | 6.6e-05 | 0.000938 | 0.33483 | 1.3233 | 2.5473 | 100.3 | 21.22 | 31.456 | 86.6223 | 92.1203 | 82.0297 | 0.007915 | 0.99399 | 0.023132 | 23.2722 | 2.3695 | 0.15249 | 0.032472 | 2.2455 | 2.2455 | 0.002764
2019-03-29 11:36 | 1553873746.8577 | 27 | 0.20299 | 0.006145 | 9.8e-05 | 98.5122 | 75.2352 | 114.7107 | 21.0578 | 0.002726 | 0.002549 | 176.995 | 0.25679 | 8.4009 | 0.002887 | 0.004554 | 0.004409 | 6.8e-05 | 0.000935 | 0.29999 | 1.3375 | 2.5473 | 100.3 | 21.22 | 31.458 | 87.2093 | 91.0841 | 82.9467 | 0.007916 | 0.99143 | 0.0231 | 23.2997 | 2.4111 | 0.15276 | 0.031885 | 2.202 | 2.202 | 0.002824
2019-03-29 11:36 | 1553873747.8563 | 28 | 0.20319 | 0.005983 | 9.5e-05 | 98.5613 | 75.2564 | 114.2246 | 21.0569 | 0.002717 | 0.002606 | 176.9153 | 0.25632 | 8.4028 | 0.002835 | 0.0047 | 0.004653 | 6.8e-05 | 0.000939 | 0.29502 | 1.3386 | 2.5474 | 100.3 | 21.22 | 31.459 | 86.9796 | 89.8157 | 84.3024 | 0.007916 | 0.98767 | 0.023007 | 23.2945 | 2.413 | 0.15249 | 0.03093 | 2.1277 | 2.1277 | 0.002917
2019-03-29 11:36 | 1553873748.8574 | 29 | 0.2034 | 0.005828 | 9.3e-05 | 98.462 | 76.5314 | 113.8411 | 21.0587 | 0.002734 | 0.002594 | 176.8531 | 0.2569 | 8.4039 | 0.00292 | 0.004122 | 0.004012 | 6.8e-05 | 0.000941 | 0.28119 | 1.354 | 2.5474 | 100.3 | 21.22 | 31.459 | 87.3942 | 91.8777 | 84.0256 | 0.007916 | 0.98844 | 0.022978 | 23.2463 | 2.2537 | 0.15283 | 0.029868 | 2.0521 | 2.0521 | 0.003031
2019-03-29 11:36 | 1553873749.8574 | 30 | 0.20357 | 0.005686 | 9.1e-05 | 98.4943 | 77.0021 | 114.03 | 21.0581 | 0.002715 | 0.002651 | 176.7935 | 0.25628 | 8.3992 | 0.002924 | 0.004082 | 0.004086 | 6.6e-05 | 0.000945 | 0.29087 | 1.3489 | 2.5474 | 100.3 | 21.22 | 31.459 | 86.9355 | 91.971 | 85.4015 | 0.007916 | 0.98464 | 0.022952 | 23.3104 | 2.2018 | 0.15246 | 0.028976 | 1.9888 | 1.9888 | 0.00312
2019-03-29 11:36 | 1553873750.8574 | 31 | 0.2037 | 0.00556 | 8.9e-05 | 98.4221 | 78.1399 | 114.475 | 21.0584 | 0.002755 | 0.002607 | 176.6958 | 0.25652 | 8.4018 | 0.002944 | 0.004119 | 0.003946 | 6.9e-05 | 0.000949 | 0.30215 | 1.328 | 2.5474 | 100.3 | 21.22 | 31.459 | 87.8907 | 92.4554 | 84.3336 | 0.007916 | 0.98759 | 0.023031 | 23.321 | 2.0641 | 0.1526 | 0.028366 | 1.9537 | 1.9537 | 0.003179
2019-03-29 11:36 | 1553873751.8574 | 32 | 0.20372 | 0.005437 | 8.6e-05 | 98.4988 | 79.1873 | 114.8353 | 21.0585 | 0.002779 | 0.002625 | 176.605 | 0.25682 | 8.4036 | 0.002952 | 0.004534 | 0.004431 | 6.5e-05 | 0.000949 | 0.30757 | 1.3565 | 2.5473 | 100.3 | 21.22 | 31.457 | 88.4835 | 92.6525 | 84.759 | 0.007915 | 0.98641 | 0.023054 | 23.3716 | 1.9519 | 0.15278 | 0.028413 | 1.9591 | 1.9591 | 0.003174
2019-03-29 11:36 | 1553873752.8568 | 33 | 0.2037 | 0.005325 | 8.4e-05 | 98.45 | 79.6412 | 115.0473 | 21.0584 | 0.002771 | 0.002621 | 176.5268 | 0.25665 | 8.4042 | 0.002903 | 0.004462 | 0.004339 | 6.8e-05 | 0.000952 | 0.29551 | 1.3686 | 2.5475 | 100.3 | 21.22 | 31.456 | 88.2774 | 91.4717 | 84.6769 | 0.007915 | 0.98664 | 0.023078 | 23.3904 | 1.8964 | 0.15268 | 0.028683 | 1.9799 | 1.9799 | 0.003139
2019-03-29 11:36 | 1553873753.8578 | 34 | 0.20373 | 0.005231 | 8.2e-05 | 98.4841 | 79.2235 | 115.1142 | 21.0586 | 0.00283 | 0.002604 | 176.4447 | 0.25741 | 8.4021 | 0.002992 | 0.004531 | 0.004492 | 7.1e-05 | 0.000954 | 0.30381 | 1.378 | 2.5474 | 100.3 | 21.21 | 31.453 | 89.6969 | 93.6009 | 84.2567 | 0.007909 | 0.9878 | 0.023098 | 23.3835 | 1.9465 | 0.15313 | 0.028629 | 1.978 | 1.978 | 0.003151
2019-03-29 11:36 | 1553873754.8575 | 35 | 0.2038 | 0.005156 | 7.8e-05 | 98.438 | 79.7614 | 115.2332 | 21.0594 | 0.002812 | 0.002636 | 176.3548 | 0.25674 | 8.4021 | 0.002969 | 0.004529 | 0.004409 | 6.8e-05 | 0.000947 | 0.29695 | 1.3687 | 2.5473 | 100.3 | 21.21 | 31.452 | 89.275 | 93.0516 | 85.0306 | 0.007909 | 0.98566 | 0.023085 | 23.4209 | 1.8818 | 0.15273 | 0.028308 | 1.9549 | 1.9549 | 0.00318
2019-03-29 11:36 | 1553873755.8569 | 36 | 0.20396 | 0.005093 | 7.5e-05 | 98.4263 | 80.4607 | 115.2238 | 21.0597 | 0.002833 | 0.00268 | 176.2901 | 0.25581 | 8.403 | 0.003 | 0.004297 | 0.004269 | 6.3e-05 | 0.000951 | 0.32964 | 1.3827 | 2.5476 | 100.3 | 21.21 | 31.452 | 89.776 | 93.8007 | 86.1041 | 0.007909 | 0.98272 | 0.02305 | 23.455 | 1.8027 | 0.15218 | 0.027382 | 1.8883 | 1.8883 | 0.00328
2019-03-29 11:36 | 1553873756.856 | 37 | 0.20428 | 0.005025 | 7.2e-05 | 98.4114 | 80.9586 | 115.8766 | 21.0606 | 0.002832 | 0.002639 | 176.2194 | 0.25631 | 8.4028 | 0.003044 | 0.004198 | 0.00407 | 6.5e-05 | 0.000946 | 0.32444 | 1.3922 | 2.5476 | 100.3 | 21.21 | 31.45 | 89.7511 | 94.8565 | 85.1013 | 0.007909 | 0.98547 | 0.023147 | 23.4885 | 1.7461 | 0.15248 | 0.025556 | 1.7701 | 1.7701 | 0.003506
2019-03-29 11:36 | 1553873757.8567 | 38 | 0.20452 | 0.004961 | 6.8e-05 | 98.3964 | 82.0937 | 116.8575 | 21.0602 | 0.002815 | 0.002613 | 176.1703 | 0.25634 | 8.4016 | 0.002941 | 0.004996 | 0.004874 | 6.5e-05 | 0.000948 | 0.30702 | 1.3512 | 2.5475 | 100.3 | 21.21 | 31.449 | 89.3329 | 92.3736 | 84.4797 | 0.007908 | 0.98718 | 0.023265 | 23.5673 | 1.6202 | 0.1525 | 0.024199 | 1.6849 | 1.6849 | 0.003684
2019-03-29 11:36 | 1553873758.8557 | 39 | 0.20483 | 0.00488 | 6.5e-05 | 98.4344 | 82.12 | 117.5955 | 21.06 | 0.002821 | 0.002635 | 176.0929 | 0.25594 | 8.4032 | 0.002958 | 0.0045 | 0.004471 | 6.5e-05 | 0.000948 | 0.28655 | 1.3474 | 2.5474 | 100.3 | 21.21 | 31.447 | 89.4859 | 92.7832 | 85.0202 | 0.007908 | 0.98569 | 0.023321 | 23.6594 | 1.6208 | 0.15226 | 0.022494 | 1.5701 | 1.5701 | 0.003947
2019-03-29 11:36 | 1553873759.8557 | 40 | 0.20499 | 0.004783 | 6e-05 | 98.3455 | 83.1555 | 117.9412 | 21.0609 | 0.002824 | 0.002609 | 176.0415 | 0.25677 | 8.401 | 0.002961 | 0.004307 | 0.004203 | 6.5e-05 | 0.000948 | 0.30502 | 1.3189 | 2.5474 | 100.3 | 21.21 | 31.446 | 89.5633 | 92.8702 | 84.3845 | 0.007908 | 0.98744 | 0.023376 | 23.6731 | 1.5007 | 0.15275 | 0.021615 | 1.5125 | 1.5125 | 0.00411
2019-03-29 11:36 | 1553873760.8556 | 41 | 0.20519 | 0.004668 | 5.6e-05 | 98.3114 | 83.8002 | 118.3515 | 21.0612 | 0.002787 | 0.002625 | 175.9938 | 0.2569 | 8.4041 | 0.00295 | 0.004501 | 0.004414 | 6.2e-05 | 0.00095 | 0.31022 | 1.3402 | 2.5475 | 100.3 | 21.21 | 31.446 | 88.6597 | 92.5951 | 84.7671 | 0.007908 | 0.98639 | 0.023404 | 23.727 | 1.4285 | 0.15283 | 0.020562 | 1.4409 | 1.4409 | 0.004317
2019-03-29 11:36 | 1553873761.8557 | 42 | 0.20544 | 0.004532 | 5.3e-05 | 98.3623 | 83.7452 | 118.6475 | 21.061 | 0.002794 | 0.002622 | 175.9431 | 0.25704 | 8.403 | 0.002915 | 0.004932 | 0.004789 | 6.4e-05 | 0.000953 | 0.32709 | 1.3048 | 2.5474 | 100.3 | 21.21 | 31.446 | 88.8343 | 91.7627 | 84.7107 | 0.007908 | 0.98654 | 0.023435 | 23.7547 | 1.439 | 0.15291 | 0.019242 | 1.3503 | 1.3503 | 0.004609
2019-03-29 11:36 | 1553873762.8561 | 43 | 0.20556 | 0.00439 | 5e-05 | 98.4087 | 84.9142 | 118.9942 | 21.0606 | 0.002777 | 0.002576 | 175.888 | 0.25558 | 8.4037 | 0.002863 | 0.004138 | 0.00401 | 6.9e-05 | 0.00095 | 0.30143 | 1.3295 | 2.5475 | 100.3 | 21.21 | 31.447 | 88.4347 | 90.5021 | 83.5874 | 0.007908 | 0.98965 | 0.023506 | 23.752 | 1.3192 | 0.15204 | 0.018672 | 1.3144 | 1.3144 | 0.004708
2019-03-29 11:36 | 1553873763.8562 | 44 | 0.20575 | 0.00424 | 4.7e-05 | 98.3332 | 85.4708 | 119.3636 | 21.0622 | 0.002773 | 0.00264 | 175.8319 | 0.25637 | 8.4046 | 0.002808 | 0.004276 | 0.004287 | 6.6e-05 | 0.000946 | 0.28655 | 1.3698 | 2.5473 | 100.3 | 21.21 | 31.447 | 88.3248 | 89.166 | 85.1442 | 0.007908 | 0.98535 | 0.023491 | 23.8407 | 1.2539 | 0.15252 | 0.017736 | 1.2479 | 1.2479 | 0.004974
2019-03-29 11:36 | 1553873764.8559 | 45 | 0.20587 | 0.004093 | 4.4e-05 | 98.3494 | 85.9602 | 119.5005 | 21.0616 | 0.002768 | 0.00261 | 175.7958 | 0.25652 | 8.4014 | 0.002878 | 0.004118 | 0.003854 | 6.5e-05 | 0.000948 | 0.32836 | 1.3597 | 2.5474 | 100.3 | 21.21 | 31.447 | 88.2203 | 90.8548 | 84.4186 | 0.007908 | 0.98735 | 0.023529 | 23.8302 | 1.2043 | 0.1526 | 0.017218 | 1.2135 | 1.2135 | 0.005118
2019-03-29 11:36 | 1553873765.8568 | 46 | 0.20598 | 0.003953 | 4.2e-05 | 98.3027 | 86.4592 | 119.8117 | 21.0609 | 0.002747 | 0.002599 | 175.7428 | 0.257 | 8.402 | 0.002789 | 0.004416 | 0.0044 | 6.5e-05 | 0.000945 | 0.28998 | 1.3421 | 2.5475 | 100.3 | 21.21 | 31.447 | 87.7095 | 88.7172 | 84.1464 | 0.007908 | 0.9881 | 0.023568 | 23.8522 | 1.1483 | 0.15289 | 0.016703 | 1.1793 | 1.1793 | 0.005276
2019-03-29 11:36 | 1553873766.8573 | 47 | 0.20606 | 0.003815 | 3.9e-05 | 98.3443 | 86.7438 | 119.7789 | 21.0622 | 0.002731 | 0.00256 | 175.6989 | 0.25648 | 8.4019 | 0.002852 | 0.004499 | 0.004336 | 6.8e-05 | 0.000938 | 0.28047 | 1.3248 | 2.5474 | 100.3 | 21.21 | 31.449 | 87.3291 | 90.2279 | 83.2039 | 0.007908 | 0.99072 | 0.023596 | 23.8174 | 1.1227 | 0.15258 | 0.016437 | 1.1621 | 1.1621 | 0.005344
2019-03-29 11:36 | 1553873767.8569 | 48 | 0.20611 | 0.003694 | 3.8e-05 | 98.3069 | 87.3943 | 119.3451 | 21.0611 | 0.0028 | 0.002635 | 175.6247 | 0.25521 | 8.4018 | 0.002715 | 0.004368 | 0.004258 | 6.2e-05 | 0.000938 | 0.29662 | 1.2939 | 2.5474 | 100.3 | 21.21 | 31.45 | 88.9798 | 86.9461 | 85.015 | 0.007909 | 0.98571 | 0.023494 | 23.8346 | 1.0525 | 0.15182 | 0.016289 | 1.1467 | 1.1467 | 0.005389
2019-03-29 11:36 | 1553873768.8569 | 49 | 0.20616 | 0.003577 | 3.7e-05 | 98.341 | 87.4496 | 119.4017 | 21.0613 | 0.002795 | 0.002617 | 175.59 | 0.25701 | 8.3985 | 0.002746 | 0.003864 | 0.003749 | 6.5e-05 | 0.000942 | 0.32676 | 1.3166 | 2.5475 | 100.3 | 21.21 | 31.453 | 88.8728 | 87.6959 | 84.5697 | 0.007909 | 0.98693 | 0.023514 | 23.8254 | 1.0499 | 0.1529 | 0.016116 | 1.1356 | 1.1356 | 0.00548
2019-03-29 11:36 | 1553873769.8564 | 50 | 0.20618 | 0.003485 | 3.7e-05 | 98.3461 | 87.8991 | 119.3376 | 21.0616 | 0.002773 | 0.002578 | 175.5171 | 0.25602 | 8.4039 | 0.002795 | 0.004532 | 0.004454 | 6.1e-05 | 0.000939 | 0.31205 | 1.2887 | 2.5472 | 100.3 | 21.21 | 31.459 | 88.3334 | 88.8546 | 83.6486 | 0.007911 | 0.98948 | 0.023538 | 23.7883 | 1.0045 | 0.15231 | 0.016086 | 1.1346 | 1.1346 | 0.005464
2019-03-29 11:36 | 1553873770.8561 | 51 | 0.20621 | 0.003402 | 3.5e-05 | 98.3192 | 87.8533 | 119.2118 | 21.0615 | 0.002821 | 0.002584 | 175.4764 | 0.25635 | 8.4013 | 0.002747 | 0.003616 | 0.003512 | 6.7e-05 | 0.000933 | 0.3011 | 1.3185 | 2.5474 | 100.3 | 21.21 | 31.461 | 89.481 | 87.7188 | 83.7857 | 0.007911 | 0.9891 | 0.023521 | 23.7804 | 1.0067 | 0.1525 | 0.016045 | 1.131 | 1.131 | 0.005488
2019-03-29 11:36 | 1553873771.8554 | 52 | 0.20625 | 0.003335 | 3.5e-05 | 98.3536 | 88.1438 | 119.1174 | 21.0618 | 0.002749 | 0.002563 | 175.4333 | 0.25608 | 8.4015 | 0.002697 | 0.003923 | 0.003807 | 6.3e-05 | 0.000925 | 0.30071 | 1.319 | 2.5474 | 100.3 | 21.21 | 31.47 | 87.7597 | 86.512 | 83.2907 | 0.007914 | 0.99047 | 0.023528 | 23.7545 | 0.98032 | 0.15234 | 0.015863 | 1.1185 | 1.1185 | 0.005544
2019-03-29 11:36 | 1553873772.8552 | 53 | 0.20632 | 0.003275 | 3.5e-05 | 98.3686 | 88.0404 | 119.1367 | 21.0623 | 0.002788 | 0.002559 | 175.3523 | 0.25671 | 8.4015 | 0.002751 | 0.004418 | 0.004283 | 6.6e-05 | 0.000923 | 0.30718 | 1.2838 | 2.5473 | 100.3 | 21.21 | 31.474 | 88.6865 | 87.8031 | 83.1898 | 0.007915 | 0.99076 | 0.023533 | 23.753 | 0.99218 | 0.15272 | 0.015551 | 1.0968 | 1.0968 | 0.005667
2019-03-29 11:36 | 1553873773.8562 | 54 | 0.2063 | 0.00323 | 3.4e-05 | 98.3826 | 88.2847 | 119.2566 | 21.0631 | 0.002797 | 0.002603 | 175.3146 | 0.25688 | 8.3997 | 0.002672 | 0.004334 | 0.004219 | 6.6e-05 | 0.000923 | 0.30469 | 1.2744 | 2.5473 | 100.3 | 21.21 | 31.478 | 88.9201 | 85.9036 | 84.2531 | 0.007916 | 0.98781 | 0.02351 | 23.8004 | 0.96867 | 0.15282 | 0.015739 | 1.1089 | 1.1089 | 0.005609
2019-03-29 11:36 | 1553873774.8565 | 55 | 0.20632 | 0.003191 | 3.3e-05 | 98.3656 | 88.0673 | 119.3075 | 21.063 | 0.002799 | 0.002546 | 175.257 | 0.25598 | 8.4018 | 0.00273 | 0.004075 | 0.004005 | 6.6e-05 | 0.000922 | 0.29319 | 1.3133 | 2.5473 | 100.3 | 21.2 | 31.486 | 88.9559 | 87.2947 | 82.8793 | 0.007913 | 0.99162 | 0.023561 | 23.7597 | 0.98917 | 0.15228 | 0.015631 | 1.1038 | 1.1038 | 0.005615
2019-03-29 11:36 | 1553873775.8563 | 56 | 0.20638 | 0.003156 | 3.3e-05 | 98.4443 | 88.5666 | 118.9533 | 21.0634 | 0.002848 | 0.002559 | 175.1982 | 0.25646 | 8.401 | 0.002673 | 0.004049 | 0.004033 | 6.8e-05 | 0.000922 | 0.27992 | 1.3319 | 2.5475 | 100.3 | 21.2 | 31.491 | 90.1368 | 85.9265 | 83.1885 | 0.007914 | 0.99076 | 0.023515 | 23.7347 | 0.94576 | 0.15257 | 0.015336 | 1.0808 | 1.0808 | 0.005745
2019-03-29 11:36 | 1553873776.8552 | 57 | 0.20644 | 0.003113 | 3.3e-05 | 98.5407 | 88.7291 | 118.8738 | 21.0629 | 0.00277 | 0.002554 | 175.1456 | 0.25687 | 8.4011 | 0.002688 | 0.003798 | 0.003692 | 6.7e-05 | 0.000917 | 0.3079 | 1.345 | 2.5474 | 100.3 | 21.2 | 31.496 | 88.2559 | 86.2797 | 83.0663 | 0.007915 | 0.9911 | 0.023512 | 23.7227 | 0.93812 | 0.15281 | 0.014989 | 1.0562 | 1.0562 | 0.005888
2019-03-29 11:36 | 1553873777.8554 | 58 | 0.20645 | 0.003084 | 3.3e-05 | 98.5383 | 89.0765 | 119.0979 | 21.0633 | 0.002768 | 0.002582 | 175.0956 | 0.2563 | 8.4034 | 0.002616 | 0.003919 | 0.00379 | 6.7e-05 | 0.000914 | 0.31188 | 1.3043 | 2.5474 | 100.3 | 21.2 | 31.506 | 88.215 | 84.5506 | 83.7455 | 0.007918 | 0.98921 | 0.023511 | 23.7677 | 0.90295 | 0.15247 | 0.014971 | 1.0549 | 1.0549 | 0.005883
2019-03-29 11:36 | 1553873778.8567 | 59 | 0.20648 | 0.003045 | 3.2e-05 | 98.6311 | 89.3529 | 119.3125 | 21.0636 | 0.002769 | 0.00257 | 175.0351 | 0.25694 | 8.4012 | 0.002649 | 0.003902 | 0.003914 | 6.5e-05 | 0.000916 | 0.28783 | 1.3378 | 2.5475 | 100.3 | 21.2 | 31.51 | 88.234 | 85.3616 | 83.4603 | 0.007919 | 0.99 | 0.023542 | 23.7796 | 0.88365 | 0.15286 | 0.014872 | 1.0494 | 1.0494 | 0.005929
2019-03-29 11:36 | 1553873779.8564 | 60 | 0.20651 | 0.003008 | 3.2e-05 | 98.6511 | 89.4427 | 119.3423 | 21.0634 | 0.002817 | 0.002573 | 174.9999 | 0.25676 | 8.4025 | 0.002681 | 0.003909 | 0.003807 | 6.4e-05 | 0.000912 | 0.29192 | 1.341 | 2.5474 | 100.3 | 21.2 | 31.517 | 89.3914 | 86.131 | 83.5332 | 0.007921 | 0.9898 | 0.023542 | 23.785 | 0.87649 | 0.15274 | 0.01474 | 1.0401 | 1.0401 | 0.005977
2019-03-29 11:36 | 1553873780.8553 | 61 | 0.20653 | 0.002984 | 3.2e-05 | 98.723 | 89.9113 | 118.9428 | 21.0631 | 0.002782 | 0.00253 | 174.9391 | 0.25773 | 8.4015 | 0.002683 | 0.004214 | 0.004086 | 6.8e-05 | 0.000908 | 0.29518 | 1.3484 | 2.5473 | 100.3 | 21.2 | 31.521 | 88.5552 | 86.1699 | 82.4885 | 0.007922 | 0.99271 | 0.023537 | 23.7103 | 0.83626 | 0.15332 | 0.014622 | 1.0315 | 1.0315 | 0.006049
2019-03-29 11:36 | 1553873781.8558 | 62 | 0.20654 | 0.002957 | 3.1e-05 | 98.802 | 90.0915 | 119.0848 | 21.0629 | 0.002785 | 0.002566 | 174.8895 | 0.25665 | 8.4025 | 0.002657 | 0.004168 | 0.004043 | 6.9e-05 | 0.000909 | 0.31365 | 1.3225 | 2.5473 | 100.3 | 21.2 | 31.527 | 88.6355 | 85.5507 | 83.355 | 0.007923 | 0.9903 | 0.023523 | 23.7533 | 0.82551 | 0.15268 | 0.014631 | 1.0316 | 1.0316 | 0.006024
2019-03-29 11:36 | 1553873782.8555 | 63 | 0.20658 | 0.00293 | 3.1e-05 | 98.8382 | 90.0939 | 119.0916 | 21.065 | 0.002795 | 0.002506 | 174.8706 | 0.25633 | 8.402 | 0.002716 | 0.004084 | 0.004038 | 6.6e-05 | 0.000909 | 0.30287 | 1.3345 | 2.5475 | 100.3 | 21.2 | 31.529 | 88.8719 | 86.968 | 81.9149 | 0.007924 | 0.99431 | 0.023571 | 23.706 | 0.82855 | 0.15249 | 0.014383 | 1.0162 | 1.0162 | 0.006108
2019-03-29 11:36 | 1553873783.8555 | 64 | 0.20653 | 0.002907 | 3.1e-05 | 98.8583 | 90.6587 | 119.3303 | 21.0646 | 0.002783 | 0.002528 | 174.8119 | 0.25588 | 8.4013 | 0.002639 | 0.004711 | 0.004651 | 6.7e-05 | 0.000905 | 0.31902 | 1.3331 | 2.5473 | 100.3 | 21.2 | 31.53 | 88.5896 | 85.106 | 82.4556 | 0.007924 | 0.9928 | 0.023577 | 23.7478 | 0.77446 | 0.15222 | 0.014721 | 1.0403 | 1.0403 | 0.005955
2019-03-29 11:36 | 1553873784.8558 | 65 | 0.2066 | 0.002887 | 3.1e-05 | 98.9127 | 90.8434 | 119.1608 | 21.0641 | 0.00276 | 0.002565 | 174.7615 | 0.25589 | 8.402 | 0.002645 | 0.004295 | 0.004176 | 6.3e-05 | 0.000901 | 0.30596 | 1.2974 | 2.5473 | 100.3 | 21.2 | 31.532 | 88.0244 | 85.2623 | 83.3367 | 0.007924 | 0.99035 | 0.023531 | 23.7603 | 0.76119 | 0.15223 | 0.014367 | 1.0133 | 1.0133 | 0.006114
2019-03-29 11:36 | 1553873785.8567 | 66 | 0.20658 | 0.002871 | 3.1e-05 | 99.0145 | 90.9345 | 119.4045 | 21.0644 | 0.002735 | 0.002567 | 174.7115 | 0.25689 | 8.4033 | 0.002642 | 0.004497 | 0.004391 | 6.5e-05 | 0.000894 | 0.28119 | 1.3093 | 2.5474 | 100.3 | 21.2 | 31.533 | 87.4265 | 85.1892 | 83.3721 | 0.007925 | 0.99025 | 0.023554 | 23.7858 | 0.76143 | 0.15282 | 0.014457 | 1.0206 | 1.0206 | 0.006094
2019-03-29 11:36 | 1553873786.8561 | 67 | 0.2066 | 0.00285 | 3.2e-05 | 98.9705 | 90.8577 | 119.4653 | 21.0651 | 0.002835 | 0.002514 | 174.6624 | 0.25589 | 8.4011 | 0.002705 | 0.004004 | 0.003988 | 6.7e-05 | 0.000894 | 0.30807 | 1.3572 | 2.5474 | 100.3 | 21.2 | 31.533 | 89.833 | 86.701 | 82.1186 | 0.007925 | 0.99374 | 0.023601 | 23.75 | 0.76501 | 0.15223 | 0.014398 | 1.0185 | 1.0185 | 0.006083
2019-03-29 11:36 | 1553873787.8545 | 68 | 0.2066 | 0.002832 | 3.2e-05 | 98.9975 | 91.0945 | 119.3798 | 21.065 | 0.002752 | 0.002532 | 174.623 | 0.25651 | 8.4023 | 0.002648 | 0.004278 | 0.004119 | 6.6e-05 | 0.00089 | 0.31133 | 1.3002 | 2.5474 | 100.3 | 21.2 | 31.531 | 87.8401 | 85.3421 | 82.5512 | 0.007924 | 0.99253 | 0.023579 | 23.7559 | 0.74415 | 0.1526 | 0.014432 | 1.0199 | 1.0199 | 0.00609
2019-03-29 11:36 | 1553873788.8556 | 69 | 0.20661 | 0.002812 | 3.1e-05 | 99.0687 | 91.531 | 119.1331 | 21.0653 | 0.002755 | 0.002576 | 174.5708 | 0.25668 | 8.4045 | 0.002655 | 0.00404 | 0.004048 | 6.6e-05 | 0.000885 | 0.28816 | 1.3155 | 2.5474 | 100.3 | 21.2 | 31.531 | 87.9013 | 85.4874 | 83.593 | 0.007924 | 0.98964 | 0.02352 | 23.7661 | 0.70783 | 0.1527 | 0.014359 | 1.0122 | 1.0122 | 0.00614
2019-03-29 11:36 | 1553873789.8562 | 70 | 0.20662 | 0.002796 | 3.2e-05 | 99.1438 | 91.3952 | 119.2582 | 21.064 | 0.002796 | 0.002534 | 174.5147 | 0.25674 | 8.4016 | 0.002641 | 0.004571 | 0.004475 | 6.6e-05 | 0.000885 | 0.31277 | 1.2985 | 2.5473 | 100.3 | 21.2 | 31.529 | 88.8873 | 85.16 | 82.5986 | 0.007924 | 0.9924 | 0.023565 | 23.7454 | 0.72789 | 0.15273 | 0.014341 | 1.0129 | 1.0129 | 0.006137
2019-03-29 11:36 | 1553873790.8551 | 71 | 0.20663 | 0.002785 | 3.1e-05 | 99.1184 | 91.3195 | 119.412 | 21.0663 | 0.002792 | 0.002526 | 174.4694 | 0.25696 | 8.4038 | 0.002627 | 0.003826 | 0.003725 | 6.8e-05 | 0.000882 | 0.29679 | 1.3083 | 2.5474 | 100.3 | 21.2 | 31.528 | 88.8028 | 84.8242 | 82.4036 | 0.007923 | 0.99295 | 0.023587 | 23.7542 | 0.73302 | 0.15287 | 0.014277 | 1.0094 | 1.0094 | 0.006164
2019-03-29 11:36 | 1553873791.8548 | 72 | 0.20661 | 0.002773 | 3.1e-05 | 99.2429 | 91.6345 | 119.181 | 21.065 | 0.002769 | 0.002498 | 174.4198 | 0.25666 | 8.4037 | 0.002618 | 0.00382 | 0.003749 | 6.3e-05 | 0.000876 | 0.29535 | 1.3394 | 2.5473 | 100.3 | 21.2 | 31.525 | 88.251 | 84.5967 | 81.7324 | 0.007923 | 0.99482 | 0.023586 | 23.7088 | 0.71344 | 0.15269 | 0.014388 | 1.0172 | 1.0172 | 0.006109
2019-03-29 11:37 | 1553873792.8563 | 73 | 0.20665 | 0.002768 | 3e-05 | 99.3264 | 91.9717 | 119.0759 | 21.0653 | 0.002774 | 0.00253 | 174.3558 | 0.25655 | 8.4009 | 0.002568 | 0.004431 | 0.004328 | 6.8e-05 | 0.000879 | 0.30359 | 1.3404 | 2.5475 | 100.3 | 21.2 | 31.523 | 88.362 | 83.4165 | 82.4813 | 0.007922 | 0.99273 | 0.023551 | 23.7233 | 0.68812 | 0.15262 | 0.014198 | 1.0023 | 1.0023 | 0.006198
2019-03-29 11:37 | 1553873793.8547 | 74 | 0.20666 | 0.00276 | 3e-05 | 99.3884 | 92.0829 | 118.7519 | 21.065 | 0.00281 | 0.002514 | 174.3065 | 0.25665 | 8.4015 | 0.002574 | 0.004044 | 0.00398 | 6.6e-05 | 0.000883 | 0.30901 | 1.3571 | 2.5475 | 100.3 | 21.2 | 31.52 | 89.2304 | 83.5403 | 82.1099 | 0.007921 | 0.99377 | 0.023531 | 23.6787 | 0.68288 | 0.15268 | 0.014121 | 0.99601 | 0.99601 | 0.006239
2019-03-29 11:37 | 1553873794.8558 | 75 | 0.20665 | 0.002759 | 3e-05 | 99.3872 | 92.3293 | 119.1212 | 21.0665 | 0.002818 | 0.002481 | 174.259 | 0.2564 | 8.4022 | 0.002572 | 0.00324 | 0.003158 | 6.8e-05 | 0.000892 | 0.29767 | 1.3557 | 2.5475 | 100.3 | 21.19 | 31.514 | 89.4115 | 83.4927 | 81.3157 | 0.007915 | 0.99599 | 0.023594 | 23.6889 | 0.65887 | 0.15254 | 0.014172 | 1.0023 | 1.0023 | 0.006194
2019-03-29 11:37 | 1553873795.8548 | 76 | 0.20667 | 0.002749 | 3.1e-05 | 99.4581 | 92.114 | 119.0418 | 21.0661 | 0.002842 | 0.002547 | 174.2151 | 0.25599 | 8.4017 | 0.0025 | 0.003723 | 0.003593 | 6.5e-05 | 0.000888 | 0.31492 | 1.3383 | 2.5476 | 100.3 | 21.19 | 31.512 | 90.005 | 81.77 | 82.8963 | 0.007915 | 0.99157 | 0.023534 | 23.7338 | 0.68614 | 0.15229 | 0.014096 | 0.99431 | 0.99431 | 0.006234
2019-03-29 11:37 | 1553873796.8558 | 77 | 0.20668 | 0.00274 | 3e-05 | 99.474 | 92.5047 | 118.7465 | 21.0666 | 0.00274 | 0.002516 | 174.1553 | 0.25724 | 8.4031 | 0.002607 | 0.004511 | 0.004448 | 6.9e-05 | 0.000886 | 0.30773 | 1.3806 | 2.5475 | 100.3 | 21.19 | 31.51 | 87.5442 | 84.3368 | 82.1551 | 0.007914 | 0.99364 | 0.023529 | 23.6796 | 0.6497 | 0.15303 | 0.014033 | 0.98972 | 0.98972 | 0.006293
2019-03-29 11:37 | 1553873797.8563 | 78 | 0.2067 | 0.002737 | 3.1e-05 | 99.6311 | 92.2242 | 118.8447 | 21.0669 | 0.002804 | 0.002426 | 174.1195 | 0.25653 | 8.4036 | 0.002463 | 0.003789 | 0.003719 | 7.2e-05 | 0.000885 | 0.27041 | 1.3181 | 2.5474 | 100.3 | 21.2 | 31.51 | 89.0923 | 80.888 | 79.9913 | 0.007919 | 0.99973 | 0.023611 | 23.6172 | 0.69098 | 0.15261 | 0.013932 | 0.98601 | 0.98601 | 0.006299
2019-03-29 11:37 | 1553873798.8551 | 79 | 0.20671 | 0.002727 | 3e-05 | 99.6081 | 92.4392 | 119.0167 | 21.0672 | 0.002728 | 0.002429 | 174.0682 | 0.25659 | 8.4038 | 0.002529 | 0.004235 | 0.004186 | 6.9e-05 | 0.000884 | 0.3298 | 1.3248 | 2.5474 | 100.3 | 21.2 | 31.512 | 87.2651 | 82.4785 | 80.0743 | 0.007919 | 0.99949 | 0.023625 | 23.637 | 0.66809 | 0.15265 | 0.013867 | 0.98197 | 0.98197 | 0.006327
2019-03-29 11:37 | 1553873799.856 | 80 | 0.20671 | 0.002715 | 3e-05 | 99.5871 | 92.3754 | 119.3399 | 21.0666 | 0.002768 | 0.002557 | 174.0216 | 0.25674 | 8.4028 | 0.002489 | 0.004359 | 0.004214 | 7e-05 | 0.000881 | 0.28926 | 1.3409 | 2.5473 | 100.3 | 21.19 | 31.516 | 88.2307 | 81.4968 | 83.1363 | 0.007916 | 0.9909 | 0.023555 | 23.7715 | 0.67238 | 0.15273 | 0.013877 | 0.97978 | 0.97978 | 0.006344
2019-03-29 11:37 | 1553873800.8555 | 81 | 0.20671 | 0.002714 | 3e-05 | 99.7652 | 92.6873 | 119.8939 | 21.0669 | 0.002769 | 0.002523 | 173.9735 | 0.25642 | 8.4036 | 0.002454 | 0.003926 | 0.003898 | 6.7e-05 | 0.000882 | 0.30342 | 1.3123 | 2.5475 | 100.3 | 21.2 | 31.52 | 88.2442 | 80.6667 | 82.3206 | 0.007921 | 0.99318 | 0.023637 | 23.7993 | 0.65821 | 0.15255 | 0.013884 | 0.98373 | 0.98373 | 0.006311
2019-03-29 11:37 | 1553873801.8556 | 82 | 0.20673 | 0.002704 | 3.1e-05 | 99.732 | 92.5518 | 120.1237 | 21.0675 | 0.002766 | 0.002439 | 173.9252 | 0.25553 | 8.4051 | 0.002502 | 0.003785 | 0.003682 | 6.5e-05 | 0.000889 | 0.29983 | 1.2938 | 2.5474 | 100.3 | 21.19 | 31.529 | 88.1831 | 81.8259 | 80.3077 | 0.007919 | 0.99883 | 0.023727 | 23.7546 | 0.66832 | 0.15202 | 0.013786 | 0.98045 | 0.98045 | 0.00631
2019-03-29 11:37 | 1553873802.8551 | 83 | 0.20671 | 0.002699 | 3.1e-05 | 99.7793 | 92.8471 | 119.9927 | 21.067 | 0.002805 | 0.002509 | 173.8805 | 0.25654 | 8.4029 | 0.002476 | 0.003923 | 0.003875 | 6.4e-05 | 0.000889 | 0.30901 | 1.3803 | 2.5474 | 100.3 | 21.19 | 31.533 | 89.1011 | 81.1976 | 82.0009 | 0.00792 | 0.99407 | 0.023657 | 23.7984 | 0.64407 | 0.15261 | 0.013898 | 0.98555 | 0.98555 | 0.006302
2019-03-29 11:37 | 1553873803.8541 | 84 | 0.20671 | 0.002702 | 3e-05 | 99.8625 | 92.8508 | 119.8662 | 21.0678 | 0.002752 | 0.00251 | 173.8444 | 0.25626 | 8.4032 | 0.002497 | 0.004018 | 0.00386 | 6.6e-05 | 0.000885 | 0.32101 | 1.4064 | 2.5476 | 100.3 | 21.19 | 31.539 | 87.8382 | 81.7039 | 82.0087 | 0.007921 | 0.99405 | 0.023645 | 23.7861 | 0.65117 | 0.15245 | 0.013926 | 0.98696 | 0.98696 | 0.006287
2019-03-29 11:37 | 1553873804.8545 | 85 | 0.20672 | 0.002695 | 3e-05 | 99.7979 | 93.0358 | 119.603 | 21.0676 | 0.002744 | 0.002497 | 173.8097 | 0.2572 | 8.4032 | 0.002413 | 0.003572 | 0.003522 | 6.7e-05 | 0.000882 | 0.30901 | 1.352 | 2.5474 | 100.3 | 21.19 | 31.553 | 87.6388 | 79.6763 | 81.6913 | 0.007925 | 0.99494 | 0.023629 | 23.7494 | 0.62757 | 0.15301 | 0.013833 | 0.9798 | 0.9798 | 0.006356
2019-03-29 11:37 | 1553873805.8545 | 86 | 0.20674 | 0.002692 | 3e-05 | 99.9493 | 93.2221 | 119.7699 | 21.0683 | 0.00276 | 0.002435 | 173.7654 | 0.25613 | 8.4046 | 0.002487 | 0.00341 | 0.003279 | 6.3e-05 | 0.00088 | 0.32731 | 1.3722 | 2.5476 | 100.3 | 21.19 | 31.561 | 88.0291 | 81.4719 | 80.2133 | 0.007927 | 0.9991 | 0.023695 | 23.7164 | 0.62324 | 0.15237 | 0.013747 | 0.97638 | 0.97638 | 0.006352
2019-03-29 11:37 | 1553873806.8543 | 87 | 0.20674 | 0.002685 | 3e-05 | 99.9894 | 93.2559 | 119.8537 | 21.0682 | 0.002835 | 0.002452 | 173.7295 | 0.25583 | 8.4012 | 0.002563 | 0.003599 | 0.003582 | 6.3e-05 | 0.000884 | 0.31598 | 1.3704 | 2.5475 | 100.3 | 21.2 | 31.578 | 89.8375 | 83.2861 | 80.6218 | 0.007936 | 0.99795 | 0.02369 | 23.7384 | 0.62359 | 0.15219 | 0.013754 | 0.97672 | 0.97672 | 0.006342
2019-03-29 11:37 | 1553873807.8552 | 88 | 0.20674 | 0.002683 | 3e-05 | 99.9344 | 93.644 | 119.88 | 21.0676 | 0.002708 | 0.002482 | 173.6738 | 0.25754 | 8.4034 | 0.002511 | 0.003587 | 0.003487 | 6.5e-05 | 0.000883 | 0.29645 | 1.3691 | 2.5475 | 100.3 | 21.2 | 31.587 | 86.7854 | 82.046 | 81.3489 | 0.007938 | 0.9959 | 0.023668 | 23.7654 | 0.58152 | 0.15321 | 0.013763 | 0.97639 | 0.97639 | 0.006386
2019-03-29 11:37 | 1553873808.8552 | 89 | 0.20676 | 0.002681 | 2.9e-05 | 100.0463 | 93.8272 | 119.9662 | 21.0684 | 0.002744 | 0.002498 | 173.6281 | 0.25746 | 8.4022 | 0.002468 | 0.003774 | 0.003608 | 6.3e-05 | 0.00089 | 0.3418 | 1.3479 | 2.5475 | 100.3 | 21.19 | 31.595 | 87.6475 | 80.9957 | 81.7309 | 0.007935 | 0.99483 | 0.023664 | 23.7867 | 0.57404 | 0.15316 | 0.013634 | 0.96711 | 0.96711 | 0.006446
2019-03-29 11:37 | 1553873809.8544 | 90 | 0.20674 | 0.002677 | 2.9e-05 | 100.1014 | 93.7747 | 119.6291 | 21.0686 | 0.002736 | 0.002494 | 173.581 | 0.25654 | 8.4011 | 0.00238 | 0.003942 | 0.003896 | 6.4e-05 | 0.000893 | 0.31902 | 1.3453 | 2.5475 | 100.3 | 21.19 | 31.609 | 87.4445 | 78.896 | 81.6414 | 0.007939 | 0.99508 | 0.023633 | 23.7503 | 0.58398 | 0.15261 | 0.013762 | 0.97499 | 0.97499 | 0.006371
2019-03-29 11:37 | 1553873810.8543 | 91 | 0.20675 | 0.00267 | 2.9e-05 | 100.1427 | 94.104 | 119.0362 | 21.0675 | 0.00277 | 0.002515 | 173.5409 | 0.25623 | 8.4047 | 0.002497 | 0.004455 | 0.004394 | 6.1e-05 | 0.000887 | 0.33196 | 1.3327 | 2.5474 | 100.3 | 21.19 | 31.616 | 88.2587 | 81.6981 | 82.13 | 0.007941 | 0.99371 | 0.023559 | 23.7077 | 0.55631 | 0.15243 | 0.013732 | 0.96975 | 0.96975 | 0.006397
2019-03-29 11:37 | 1553873811.8547 | 92 | 0.20675 | 0.002665 | 2.9e-05 | 100.2005 | 94.3023 | 118.6055 | 21.0674 | 0.002783 | 0.002496 | 173.4848 | 0.25623 | 8.4027 | 0.0025 | 0.003604 | 0.003528 | 6.5e-05 | 0.000879 | 0.30735 | 1.3541 | 2.5475 | 100.3 | 21.19 | 31.628 | 88.5851 | 81.7636 | 81.6793 | 0.007944 | 0.99497 | 0.023531 | 23.6497 | 0.54264 | 0.15243 | 0.013694 | 0.96596 | 0.96596 | 0.006423
2019-03-29 11:37 | 1553873812.8553 | 93 | 0.20674 | 0.002661 | 2.9e-05 | 100.2328 | 94.1825 | 118.8846 | 21.0686 | 0.002726 | 0.00248 | 173.4366 | 0.25595 | 8.402 | 0.002511 | 0.003585 | 0.003514 | 6.3e-05 | 0.000877 | 0.28008 | 1.3694 | 2.5474 | 100.3 | 21.19 | 31.633 | 87.2111 | 82.0501 | 81.2951 | 0.007945 | 0.99605 | 0.023571 | 23.6647 | 0.5569 | 0.15226 | 0.013805 | 0.97545 | 0.97545 | 0.006353
2019-03-29 11:37 | 1553873813.8558 | 94 | 0.20675 | 0.002659 | 2.8e-05 | 100.2747 | 94.4149 | 118.9498 | 21.0694 | 0.002778 | 0.002513 | 173.3962 | 0.25607 | 8.4042 | 0.002476 | 0.004244 | 0.00414 | 6.3e-05 | 0.000874 | 0.3183 | 1.3665 | 2.5475 | 100.3 | 21.19 | 31.638 | 88.4544 | 81.2017 | 82.0814 | 0.007946 | 0.99385 | 0.023552 | 23.6974 | 0.53859 | 0.15233 | 0.013717 | 0.96843 | 0.96843 | 0.006402
2019-03-29 11:37 | 1553873814.8555 | 95 | 0.20677 | 0.002654 | 2.7e-05 | 100.2924 | 94.6893 | 119.0477 | 21.0685 | 0.002795 | 0.002442 | 173.345 | 0.25679 | 8.4025 | 0.00252 | 0.004122 | 0.00411 | 6.1e-05 | 0.000871 | 0.3235 | 1.3233 | 2.5473 | 100.3 | 21.19 | 31.646 | 88.8753 | 82.2599 | 80.3827 | 0.007948 | 0.99862 | 0.023618 | 23.6504 | 0.51421 | 0.15276 | 0.013588 | 0.96206 | 0.96206 | 0.006463
2019-03-29 11:37 | 1553873815.8559 | 96 | 0.20676 | 0.002655 | 2.7e-05 | 100.3058 | 94.9183 | 118.889 | 21.0691 | 0.002716 | 0.002559 | 173.2988 | 0.25735 | 8.4015 | 0.002474 | 0.003787 | 0.003688 | 6.4e-05 | 0.000867 | 0.31598 | 1.3467 | 2.5473 | 100.3 | 21.19 | 31.65 | 86.9587 | 81.1437 | 83.1909 | 0.007949 | 0.99075 | 0.023509 | 23.7283 | 0.4938 | 0.1531 | 0.013693 | 0.96504 | 0.96504 | 0.006457
2019-03-29 11:37 | 1553873816.8555 | 97 | 0.20676 | 0.002644 | 2.7e-05 | 100.4642 | 94.8958 | 119.2416 | 21.0691 | 0.002792 | 0.002552 | 173.2441 | 0.25583 | 8.403 | 0.002465 | 0.003409 | 0.003352 | 6e-05 | 0.000861 | 0.29839 | 1.3612 | 2.5475 | 100.3 | 21.19 | 31.654 | 88.8101 | 80.929 | 83.0247 | 0.00795 | 0.99121 | 0.023549 | 23.758 | 0.51003 | 0.15219 | 0.013708 | 0.96778 | 0.96778 | 0.0064
2019-03-29 11:37 | 1553873817.8549 | 98 | 0.20679 | 0.002642 | 2.7e-05 | 100.4711 | 95.2327 | 119.1672 | 21.0702 | 0.002763 | 0.002467 | 173.1774 | 0.25654 | 8.4041 | 0.002494 | 0.003726 | 0.00361 | 6.1e-05 | 0.000854 | 0.31222 | 1.3504 | 2.5476 | 100.3 | 21.19 | 31.66 | 88.1051 | 81.6343 | 80.9895 | 0.007952 | 0.99691 | 0.023609 | 23.6826 | 0.47895 | 0.15261 | 0.013516 | 0.9567 | 0.9567 | 0.006492
2019-03-29 11:37 | 1553873818.8536 | 99 | 0.20682 | 0.002637 | 2.7e-05 | 100.5788 | 95.3866 | 118.8308 | 21.0693 | 0.002757 | 0.002517 | 173.1226 | 0.25626 | 8.4012 | 0.002488 | 0.004175 | 0.004027 | 6.1e-05 | 0.000851 | 0.31382 | 1.3428 | 2.5475 | 100.3 | 21.19 | 31.663 | 87.9459 | 81.4917 | 82.1794 | 0.007952 | 0.99357 | 0.023537 | 23.6888 | 0.47409 | 0.15245 | 0.013327 | 0.94036 | 0.94036 | 0.006598
2019-03-29 11:37 | 1553873819.8546 | 100 | 0.20682 | 0.002626 | 2.7e-05 | 100.6692 | 95.4306 | 118.6965 | 21.0696 | 0.002764 | 0.002515 | 173.0615 | 0.25572 | 8.4016 | 0.002437 | 0.004249 | 0.004148 | 5.9e-05 | 0.000847 | 0.30685 | 1.3582 | 2.5476 | 100.3 | 21.19 | 31.664 | 88.1226 | 80.2732 | 82.1407 | 0.007953 | 0.99368 | 0.023525 | 23.6742 | 0.478 | 0.15213 | 0.013358 | 0.94207 | 0.94207 | 0.006572
2019-03-29 11:37 | 1553873820.8536 | 101 | 0.20683 | 0.002621 | 2.7e-05 | 100.6956 | 95.6808 | 118.6956 | 21.0691 | 0.002791 | 0.00251 | 172.9882 | 0.25592 | 8.3996 | 0.002401 | 0.003397 | 0.00332 | 6.2e-05 | 0.000842 | 0.3095 | 1.3521 | 2.5474 | 100.3 | 21.19 | 31.665 | 88.7822 | 79.3956 | 82.0218 | 0.007953 | 0.99401 | 0.023528 | 23.6701 | 0.45693 | 0.15225 | 0.013316 | 0.93926 | 0.93926 | 0.006597
2019-03-29 11:37 | 1553873821.8541 | 102 | 0.2068 | 0.002612 | 2.7e-05 | 100.7863 | 95.8647 | 118.9066 | 21.0698 | 0.002726 | 0.002485 | 172.9489 | 0.2556 | 8.4024 | 0.002405 | 0.003713 | 0.003634 | 6e-05 | 0.000848 | 0.33107 | 1.3593 | 2.5474 | 100.3 | 21.19 | 31.666 | 87.2036 | 79.4828 | 81.4102 | 0.007953 | 0.99573 | 0.02357 | 23.6707 | 0.4478 | 0.15206 | 0.013457 | 0.95091 | 0.95091 | 0.006508
2019-03-29 11:37 | 1553873822.8537 | 103 | 0.20684 | 0.002606 | 2.7e-05 | 100.7943 | 96.5579 | 118.6735 | 21.0696 | 0.002771 | 0.002481 | 172.9136 | 0.25673 | 8.4056 | 0.002382 | 0.004261 | 0.004087 | 6.3e-05 | 0.000848 | 0.30309 | 1.369 | 2.5473 | 100.3 | 21.18 | 31.669 | 88.2826 | 78.9508 | 81.3234 | 0.007949 | 0.99597 | 0.023549 | 23.6446 | 0.38407 | 0.15273 | 0.01326 | 0.93619 | 0.93619 | 0.00664
2019-03-29 11:37 | 1553873823.8537 | 104 | 0.20679 | 0.002601 | 2.8e-05 | 100.8231 | 96.1596 | 118.9632 | 21.0701 | 0.002786 | 0.00254 | 172.8511 | 0.2562 | 8.4026 | 0.002417 | 0.003958 | 0.003929 | 6.2e-05 | 0.000844 | 0.31542 | 1.3605 | 2.5475 | 100.3 | 21.18 | 31.671 | 88.6674 | 79.777 | 82.7427 | 0.00795 | 0.992 | 0.023531 | 23.7208 | 0.4236 | 0.15241 | 0.013554 | 0.95612 | 0.95612 | 0.006488
2019-03-29 11:37 | 1553873824.8541 | 105 | 0.20678 | 0.002603 | 2.7e-05 | 100.8937 | 96.4514 | 119.0104 | 21.0702 | 0.002778 | 0.002507 | 172.7747 | 0.25723 | 8.4026 | 0.00242 | 0.003878 | 0.003763 | 6.2e-05 | 0.00085 | 0.33035 | 1.3323 | 2.5474 | 100.3 | 21.18 | 31.673 | 88.4701 | 79.8434 | 81.9449 | 0.00795 | 0.99423 | 0.023562 | 23.6989 | 0.40276 | 0.15303 | 0.013599 | 0.96057 | 0.96057 | 0.006484
2019-03-29 11:37 | 1553873825.8548 | 106 | 0.20679 | 0.002603 | 2.8e-05 | 100.97 | 96.3674 | 118.7867 | 21.0705 | 0.002743 | 0.00251 | 172.73 | 0.25636 | 8.4049 | 0.002355 | 0.003371 | 0.003309 | 6.2e-05 | 0.000851 | 0.28854 | 1.347 | 2.5473 | 100.3 | 21.18 | 31.676 | 87.6151 | 78.2952 | 82.0061 | 0.007951 | 0.99406 | 0.023538 | 23.6787 | 0.41731 | 0.15251 | 0.013553 | 0.95632 | 0.95632 | 0.006491
2019-03-29 11:37 | 1553873826.8549 | 107 | 0.20682 | 0.002611 | 2.8e-05 | 100.9246 | 96.0081 | 118.4245 | 21.0704 | 0.002803 | 0.002557 | 172.6779 | 0.25607 | 8.4028 | 0.002365 | 0.004358 | 0.004252 | 6.4e-05 | 0.000853 | 0.3027 | 1.3473 | 2.5475 | 100.3 | 21.18 | 31.676 | 89.0574 | 78.526 | 83.1423 | 0.007951 | 0.99089 | 0.023465 | 23.6803 | 0.4467 | 0.15234 | 0.013389 | 0.94184 | 0.94184 | 0.006583
2019-03-29 11:37 | 1553873827.8549 | 108 | 0.20682 | 0.002608 | 2.7e-05 | 101.0512 | 96.4473 | 118.3745 | 21.0707 | 0.002777 | 0.002518 | 172.6363 | 0.25612 | 8.4022 | 0.002408 | 0.004053 | 0.003997 | 6.2e-05 | 0.000858 | 0.31349 | 1.3382 | 2.5473 | 100.3 | 21.18 | 31.678 | 88.4463 | 79.5663 | 82.2106 | 0.007951 | 0.99348 | 0.02349 | 23.6444 | 0.41709 | 0.15237 | 0.013352 | 0.94035 | 0.94035 | 0.006595
2019-03-29 11:37 | 1553873828.8545 | 109 | 0.20681 | 0.002612 | 2.6e-05 | 101.0506 | 96.6275 | 118.6557 | 21.071 | 0.00275 | 0.002488 | 172.5936 | 0.25635 | 8.4038 | 0.002454 | 0.003479 | 0.003394 | 6.1e-05 | 0.00086 | 0.31996 | 1.3341 | 2.5474 | 100.3 | 21.17 | 31.68 | 87.7827 | 80.6668 | 81.4941 | 0.007947 | 0.99549 | 0.023542 | 23.6486 | 0.40034 | 0.1525 | 0.013403 | 0.94602 | 0.94602 | 0.006561
2019-03-29 11:37 | 1553873829.8543 | 110 | 0.20683 | 0.002607 | 2.6e-05 | 101.088 | 96.9456 | 119.016 | 21.0706 | 0.002778 | 0.002448 | 172.5729 | 0.25613 | 8.4018 | 0.00243 | 0.003868 | 0.00377 | 6.1e-05 | 0.000859 | 0.31078 | 1.3543 | 2.5474 | 100.3 | 21.17 | 31.681 | 88.4656 | 80.0838 | 80.5198 | 0.007947 | 0.99823 | 0.02361 | 23.6519 | 0.37425 | 0.15237 | 0.013345 | 0.94468 | 0.94468 | 0.006565
2019-03-29 11:37 | 1553873830.855 | 111 | 0.20681 | 0.002607 | 2.6e-05 | 101.1083 | 97.2613 | 119.1597 | 21.0708 | 0.002781 | 0.00248 | 172.5543 | 0.25702 | 8.4013 | 0.002459 | 0.003513 | 0.003453 | 6.1e-05 | 0.000861 | 0.32228 | 1.3601 | 2.5475 | 100.3 | 21.17 | 31.685 | 88.5243 | 80.7843 | 81.3052 | 0.007948 | 0.99602 | 0.023598 | 23.6924 | 0.34697 | 0.1529 | 0.013415 | 0.94916 | 0.94916 | 0.006556
2019-03-29 11:37 | 1553873831.854 | 112 | 0.20681 | 0.002607 | 2.6e-05 | 101.1407 | 97.2879 | 119.0622 | 21.0713 | 0.002734 | 0.002551 | 172.5112 | 0.25695 | 8.4014 | 0.002428 | 0.004288 | 0.004066 | 6.5e-05 | 0.000867 | 0.32532 | 1.3731 | 2.5475 | 100.3 | 21.17 | 31.688 | 87.4003 | 80.0588 | 83.0052 | 0.007949 | 0.99127 | 0.023532 | 23.7394 | 0.34738 | 0.15286 | 0.013423 | 0.947 | 0.947 | 0.00657
2019-03-29 11:37 | 1553873832.8554 | 113 | 0.20681 | 0.002602 | 2.7e-05 | 101.1335 | 97.2555 | 118.7681 | 21.0711 | 0.002762 | 0.002558 | 172.4718 | 0.25626 | 8.4013 | 0.002461 | 0.003738 | 0.003652 | 6.2e-05 | 0.000869 | 0.29518 | 1.3613 | 2.5475 | 100.3 | 21.17 | 31.693 | 88.0677 | 80.8399 | 83.1711 | 0.00795 | 0.99081 | 0.023498 | 23.7156 | 0.34973 | 0.15245 | 0.013419 | 0.94534 | 0.94534 | 0.006563
2019-03-29 11:37 | 1553873833.854 | 114 | 0.20683 | 0.002602 | 2.7e-05 | 101.235 | 97.4002 | 118.2658 | 21.0712 | 0.002772 | 0.002533 | 172.4433 | 0.25641 | 8.3991 | 0.002379 | 0.004052 | 0.003969 | 6.3e-05 | 0.000871 | 0.31222 | 1.3477 | 2.5476 | 100.3 | 21.17 | 31.695 | 88.3117 | 78.8658 | 82.5659 | 0.007951 | 0.99249 | 0.023468 | 23.6453 | 0.3454 | 0.15254 | 0.01333 | 0.93788 | 0.93788 | 0.00662
2019-03-29 11:37 | 1553873834.8541 | 115 | 0.2068 | 0.002601 | 2.8e-05 | 101.2326 | 97.2179 | 118.3571 | 21.071 | 0.002812 | 0.002508 | 172.4019 | 0.257 | 8.4018 | 0.002487 | 0.003944 | 0.0039 | 6.3e-05 | 0.000868 | 0.28495 | 1.34 | 2.5474 | 100.3 | 21.17 | 31.697 | 89.2752 | 81.4534 | 81.9721 | 0.007951 | 0.99415 | 0.023496 | 23.6347 | 0.36194 | 0.15289 | 0.01353 | 0.95303 | 0.95303 | 0.006529
2019-03-29 11:37 | 1553873835.8541 | 116 | 0.20677 | 0.002598 | 2.8e-05 | 101.1948 | 96.9108 | 118.8477 | 21.0716 | 0.002774 | 0.002486 | 172.359 | 0.25557 | 8.4026 | 0.002447 | 0.003883 | 0.003826 | 6.4e-05 | 0.000876 | 0.30143 | 1.347 | 2.5474 | 100.3 | 21.17 | 31.703 | 88.3782 | 80.4957 | 81.4518 | 0.007953 | 0.99561 | 0.023562 | 23.6663 | 0.38692 | 0.15204 | 0.013658 | 0.96475 | 0.96475 | 0.006414
2019-03-29 11:37 | 1553873836.8532 | 117 | 0.20678 | 0.002605 | 2.8e-05 | 101.3041 | 96.8104 | 118.9761 | 21.0726 | 0.002734 | 0.002511 | 172.3286 | 0.2567 | 8.4019 | 0.002402 | 0.003719 | 0.003594 | 6.1e-05 | 0.000876 | 0.30685 | 1.3011 | 2.5473 | 100.3 | 21.17 | 31.706 | 87.4087 | 79.4237 | 82.0328 | 0.007954 | 0.99398 | 0.023556 | 23.6984 | 0.40583 | 0.15271 | 0.013611 | 0.96117 | 0.96117 | 0.006467
2019-03-29 11:37 | 1553873837.8541 | 118 | 0.20676 | 0.002607 | 2.8e-05 | 101.1595 | 97.0186 | 118.9254 | 21.0727 | 0.00279 | 0.002518 | 172.2977 | 0.25653 | 8.4006 | 0.002463 | 0.004275 | 0.004142 | 6.6e-05 | 0.000871 | 0.29767 | 1.3523 | 2.5474 | 100.3 | 21.17 | 31.713 | 88.7638 | 80.8804 | 82.2046 | 0.007955 | 0.9935 | 0.023545 | 23.6991 | 0.37384 | 0.15261 | 0.013747 | 0.97035 | 0.97035 | 0.006401
2019-03-29 11:37 | 1553873838.8532 | 119 | 0.20676 | 0.002609 | 2.7e-05 | 101.1805 | 97.1616 | 118.6526 | 21.0726 | 0.002706 | 0.002487 | 172.253 | 0.2569 | 8.4032 | 0.002406 | 0.003469 | 0.003377 | 6.6e-05 | 0.000874 | 0.30845 | 1.3364 | 2.5475 | 100.3 | 21.17 | 31.717 | 86.7444 | 79.5206 | 81.456 | 0.007956 | 0.9956 | 0.023543 | 23.647 | 0.36252 | 0.15283 | 0.013718 | 0.96825 | 0.96825 | 0.006424
2019-03-29 11:37 | 1553873839.8533 | 120 | 0.20676 | 0.002616 | 2.8e-05 | 101.2317 | 96.9845 | 118.4801 | 21.0723 | 0.00277 | 0.002506 | 172.1962 | 0.25635 | 8.4039 | 0.0024 | 0.003814 | 0.00381 | 6.5e-05 | 0.000865 | 0.28241 | 1.3333 | 2.5476 | 100.3 | 21.17 | 31.719 | 88.2745 | 79.3793 | 81.9251 | 0.007957 | 0.99428 | 0.02351 | 23.6454 | 0.38337 | 0.1525 | 0.013737 | 0.96821 | 0.96821 | 0.006411
2019-03-29 11:37 | 1553873840.8532 | 121 | 0.20677 | 0.002619 | 2.8e-05 | 101.2894 | 97.3422 | 118.3876 | 21.0726 | 0.002774 | 0.002554 | 172.1462 | 0.25644 | 8.4012 | 0.002401 | 0.004247 | 0.004118 | 6.5e-05 | 0.000873 | 0.3011 | 1.3466 | 2.5474 | 100.3 | 21.16 | 31.725 | 88.3669 | 79.4052 | 83.0694 | 0.007953 | 0.99109 | 0.023463 | 23.6742 | 0.35554 | 0.15255 | 0.013659 | 0.96075 | 0.96075 | 0.006463
2019-03-29 11:37 | 1553873841.8544 | 122 | 0.20674 | 0.002628 | 2.9e-05 | 101.2326 | 97.2142 | 118.4937 | 21.0726 | 0.002744 | 0.002518 | 172.1023 | 0.25729 | 8.4034 | 0.002396 | 0.003853 | 0.003665 | 6.8e-05 | 0.000878 | 0.33539 | 1.3518 | 2.5474 | 100.3 | 21.16 | 31.728 | 87.6398 | 79.2844 | 82.203 | 0.007954 | 0.99351 | 0.023502 | 23.656 | 0.36229 | 0.15306 | 0.013816 | 0.97338 | 0.97338 | 0.0064
2019-03-29 11:37 | 1553873842.8534 | 123 | 0.20676 | 0.002632 | 2.9e-05 | 101.2763 | 97.0562 | 118.8526 | 21.0729 | 0.002764 | 0.002554 | 172.0274 | 0.25596 | 8.402 | 0.002442 | 0.00446 | 0.00445 | 6.7e-05 | 0.000871 | 0.31581 | 1.3761 | 2.5474 | 100.3 | 21.16 | 31.731 | 88.1195 | 80.3941 | 83.0676 | 0.007955 | 0.99109 | 0.023509 | 23.7206 | 0.3807 | 0.15227 | 0.013673 | 0.96358 | 0.96358 | 0.006432
2019-03-29 11:37 | 1553873843.8534 | 124 | 0.20672 | 0.002635 | 3e-05 | 101.2882 | 97.1667 | 119.2101 | 21.0731 | 0.002776 | 0.002543 | 171.9581 | 0.25581 | 8.4034 | 0.002425 | 0.003465 | 0.003345 | 6.7e-05 | 0.000867 | 0.32278 | 1.3509 | 2.5476 | 100.3 | 21.16 | 31.735 | 88.4298 | 79.988 | 82.8159 | 0.007956 | 0.9918 | 0.023553 | 23.7479 | 0.37157 | 0.15218 | 0.013912 | 0.98217 | 0.98217 | 0.006306
2019-03-29 11:37 | 1553873844.8534 | 125 | 0.20675 | 0.002638 | 3.1e-05 | 101.2499 | 97.1429 | 118.9456 | 21.0729 | 0.002796 | 0.002465 | 171.9164 | 0.25645 | 8.4029 | 0.002443 | 0.004213 | 0.004175 | 6.7e-05 | 0.000869 | 0.29502 | 1.3515 | 2.5473 | 100.3 | 21.15 | 31.738 | 88.9092 | 80.4112 | 80.9383 | 0.007952 | 0.99706 | 0.023589 | 23.6589 | 0.37039 | 0.15256 | 0.013774 | 0.9739 | 0.9739 | 0.006376
2019-03-29 11:37 | 1553873845.8544 | 126 | 0.20674 | 0.002648 | 3.1e-05 | 101.1954 | 96.9613 | 118.4991 | 21.0727 | 0.002748 | 0.002554 | 171.8355 | 0.25627 | 8.4034 | 0.002381 | 0.003669 | 0.003589 | 6.5e-05 | 0.000873 | 0.3079 | 1.332 | 2.5474 | 100.3 | 21.15 | 31.74 | 87.7336 | 78.9149 | 83.0845 | 0.007952 | 0.99105 | 0.023474 | 23.6859 | 0.38231 | 0.15246 | 0.013785 | 0.9699 | 0.9699 | 0.006398
2019-03-29 11:37 | 1553873846.8534 | 127 | 0.20669 | 0.002652 | 3.1e-05 | 101.2805 | 97.0135 | 118.2874 | 21.0735 | 0.002758 | 0.002527 | 171.7773 | 0.25601 | 8.4033 | 0.002448 | 0.003847 | 0.003843 | 6.2e-05 | 0.000871 | 0.31006 | 1.3227 | 2.5475 | 100.3 | 21.15 | 31.74 | 87.9872 | 80.5413 | 82.4363 | 0.007952 | 0.99285 | 0.023474 | 23.6432 | 0.385 | 0.1523 | 0.014081 | 0.99073 | 0.99073 | 0.006257
2019-03-29 11:37 | 1553873847.8545 | 128 | 0.20671 | 0.00266 | 3.2e-05 | 101.199 | 96.947 | 118.3374 | 21.0725 | 0.002829 | 0.002513 | 171.7151 | 0.2566 | 8.4051 | 0.002409 | 0.003831 | 0.003794 | 6.4e-05 | 0.000863 | 0.31996 | 1.3191 | 2.5474 | 100.3 | 21.15 | 31.74 | 89.6924 | 79.5868 | 82.1015 | 0.007952 | 0.99379 | 0.02349 | 23.637 | 0.38395 | 0.15265 | 0.013939 | 0.98133 | 0.98133 | 0.006331
2019-03-29 11:37 | 1553873848.8537 | 129 | 0.20669 | 0.002669 | 3.3e-05 | 101.2314 | 97.1299 | 118.6135 | 21.0732 | 0.00274 | 0.002538 | 171.6623 | 0.25566 | 8.402 | 0.002365 | 0.003853 | 0.003704 | 6.7e-05 | 0.000861 | 0.27273 | 1.2893 | 2.5474 | 100.3 | 21.15 | 31.737 | 87.5576 | 78.5341 | 82.6965 | 0.007951 | 0.99213 | 0.023498 | 23.6844 | 0.36994 | 0.15209 | 0.014043 | 0.98891 | 0.98891 | 0.006259
2019-03-29 11:37 | 1553873849.8535 | 130 | 0.20669 | 0.002673 | 3.4e-05 | 101.3002 | 97.2347 | 118.2989 | 21.0735 | 0.002742 | 0.00255 | 171.6077 | 0.25665 | 8.4032 | 0.002324 | 0.004013 | 0.003986 | 6.5e-05 | 0.000859 | 0.2995 | 1.2919 | 2.5473 | 100.3 | 21.15 | 31.736 | 87.5923 | 77.5497 | 82.9862 | 0.007951 | 0.99132 | 0.023457 | 23.6626 | 0.36638 | 0.15268 | 0.014091 | 0.99054 | 0.99054 | 0.006273
2019-03-29 11:37 | 1553873850.8526 | 131 | 0.20667 | 0.002681 | 3.5e-05 | 101.2918 | 97.4248 | 118.1762 | 21.0729 | 0.002795 | 0.002581 | 171.5416 | 0.25674 | 8.4031 | 0.002391 | 0.003622 | 0.003516 | 6.4e-05 | 0.000859 | 0.32278 | 1.2999 | 2.5474 | 100.3 | 21.14 | 31.73 | 88.8678 | 79.1638 | 83.7221 | 0.007945 | 0.98928 | 0.023421 | 23.6747 | 0.34817 | 0.15273 | 0.014173 | 0.99468 | 0.99468 | 0.006249
2019-03-29 11:38 | 1553873851.8526 | 132 | 0.20669 | 0.002682 | 3.7e-05 | 101.2667 | 97.4558 | 118.1999 | 21.0737 | 0.002743 | 0.002557 | 171.4666 | 0.25613 | 8.4045 | 0.002325 | 0.003835 | 0.003763 | 6.6e-05 | 0.000855 | 0.31006 | 1.3367 | 2.5475 | 100.3 | 21.14 | 31.727 | 87.6291 | 77.579 | 83.1449 | 0.007944 | 0.99088 | 0.023442 | 23.658 | 0.3431 | 0.15237 | 0.014046 | 0.98662 | 0.98662 | 0.006286
2019-03-29 11:38 | 1553873852.8537 | 133 | 0.20669 | 0.002689 | 3.7e-05 | 101.2577 | 97.3934 | 118.2177 | 21.0728 | 0.002755 | 0.002537 | 171.4263 | 0.25739 | 8.4016 | 0.002341 | 0.003483 | 0.003383 | 6.8e-05 | 0.000858 | 0.29822 | 1.3419 | 2.5474 | 100.3 | 21.14 | 31.723 | 87.9093 | 77.9483 | 82.6657 | 0.007943 | 0.99221 | 0.02346 | 23.6438 | 0.34804 | 0.15312 | 0.014027 | 0.98594 | 0.98594 | 0.006321
2019-03-29 11:38 | 1553873853.8538 | 134 | 0.20667 | 0.002692 | 3.8e-05 | 101.2924 | 97.3708 | 118.125 | 21.0734 | 0.002758 | 0.00249 | 171.3895 | 0.25605 | 8.4024 | 0.00234 | 0.003903 | 0.003866 | 6.4e-05 | 0.000864 | 0.29175 | 1.3212 | 2.5475 | 100.3 | 21.14 | 31.712 | 87.9787 | 77.9434 | 81.5401 | 0.00794 | 0.99536 | 0.023488 | 23.5971 | 0.35318 | 0.15233 | 0.014151 | 0.99585 | 0.99585 | 0.006225
2019-03-29 11:38 | 1553873854.8528 | 135 | 0.20668 | 0.002696 | 3.8e-05 | 101.2011 | 97.5131 | 117.9727 | 21.0746 | 0.002812 | 0.002548 | 171.3331 | 0.25637 | 8.4003 | 0.002369 | 0.003899 | 0.003785 | 6.4e-05 | 0.000855 | 0.28263 | 1.3443 | 2.5475 | 100.3 | 21.14 | 31.707 | 89.2836 | 78.6385 | 82.9365 | 0.007939 | 0.99146 | 0.023426 | 23.6283 | 0.33205 | 0.15252 | 0.014116 | 0.9908 | 0.9908 | 0.006265
2019-03-29 11:38 | 1553873855.8528 | 136 | 0.20669 | 0.002701 | 3.9e-05 | 101.2532 | 97.7261 | 117.975 | 21.0739 | 0.002804 | 0.002494 | 171.2879 | 0.25544 | 8.4037 | 0.002342 | 0.003298 | 0.003242 | 6.3e-05 | 0.000857 | 0.30381 | 1.3198 | 2.5475 | 100.3 | 21.14 | 31.693 | 89.0852 | 77.9741 | 81.6459 | 0.007936 | 0.99507 | 0.023469 | 23.5857 | 0.31714 | 0.15196 | 0.014041 | 0.98725 | 0.98725 | 0.006265
2019-03-29 11:38 | 1553873856.8527 | 137 | 0.2067 | 0.002703 | 3.9e-05 | 101.1763 | 97.3777 | 118.1327 | 21.0736 | 0.002749 | 0.002535 | 171.2352 | 0.25618 | 8.405 | 0.002381 | 0.003771 | 0.003644 | 6.4e-05 | 0.000855 | 0.30541 | 1.3351 | 2.5474 | 100.3 | 21.14 | 31.688 | 87.7587 | 78.9088 | 82.6305 | 0.007934 | 0.99231 | 0.023452 | 23.6342 | 0.34229 | 0.1524 | 0.013968 | 0.98147 | 0.98147 | 0.00632
2019-03-29 11:38 | 1553873857.8538 | 138 | 0.20673 | 0.002709 | 3.9e-05 | 101.2682 | 97.4869 | 118.3273 | 21.0745 | 0.002799 | 0.002537 | 171.1828 | 0.25581 | 8.4025 | 0.002395 | 0.004247 | 0.004172 | 6.3e-05 | 0.000853 | 0.32477 | 1.3867 | 2.5476 | 100.3 | 21.14 | 31.682 | 88.9639 | 79.2581 | 82.6664 | 0.007933 | 0.99221 | 0.023471 | 23.6548 | 0.34038 | 0.15218 | 0.013797 | 0.97017 | 0.97017 | 0.006384
2019-03-29 11:38 | 1553873858.8539 | 139 | 0.2067 | 0.002713 | 3.9e-05 | 101.3113 | 97.7422 | 117.8752 | 21.0736 | 0.002746 | 0.002541 | 171.1201 | 0.25693 | 8.4032 | 0.002413 | 0.003605 | 0.003589 | 6.3e-05 | 0.000856 | 0.27992 | 1.3458 | 2.5475 | 100.3 | 21.14 | 31.669 | 87.7079 | 79.689 | 82.774 | 0.00793 | 0.99191 | 0.023422 | 23.6131 | 0.32079 | 0.15285 | 0.013974 | 0.98058 | 0.98058 | 0.006344
2019-03-29 11:38 | 1553873859.8528 | 140 | 0.2067 | 0.002716 | 3.8e-05 | 101.2622 | 97.7196 | 117.844 | 21.0728 | 0.002744 | 0.002648 | 171.0738 | 0.25623 | 8.4023 | 0.002464 | 0.003583 | 0.003502 | 6.6e-05 | 0.000855 | 0.29048 | 1.3273 | 2.5473 | 100.3 | 21.14 | 31.664 | 87.6571 | 80.9125 | 85.3347 | 0.007928 | 0.98483 | 0.023335 | 23.6948 | 0.31852 | 0.15243 | 0.013955 | 0.97564 | 0.97564 | 0.006359
2019-03-29 11:38 | 1553873860.8538 | 141 | 0.20671 | 0.002718 | 3.8e-05 | 101.302 | 97.9636 | 118.21 | 21.075 | 0.002815 | 0.00253 | 171.0311 | 0.2558 | 8.4015 | 0.002361 | 0.003878 | 0.003732 | 6.1e-05 | 0.000851 | 0.28102 | 1.3536 | 2.5473 | 100.3 | 21.13 | 31.65 | 89.3629 | 78.4468 | 82.5048 | 0.00792 | 0.99266 | 0.023464 | 23.6377 | 0.29973 | 0.15218 | 0.013902 | 0.97734 | 0.97734 | 0.006337
2019-03-29 11:38 | 1553873861.8543 | 142 | 0.20674 | 0.002717 | 3.8e-05 | 101.3044 | 98.0479 | 118.5165 | 21.0727 | 0.00279 | 0.002565 | 170.9878 | 0.25602 | 8.4007 | 0.002441 | 0.003534 | 0.003435 | 6.3e-05 | 0.000852 | 0.2756 | 1.3664 | 2.5475 | 100.3 | 21.13 | 31.643 | 88.7433 | 80.3665 | 83.3535 | 0.007918 | 0.9903 | 0.023467 | 23.6966 | 0.29225 | 0.15231 | 0.013685 | 0.96218 | 0.96218 | 0.006443
2019-03-29 11:38 | 1553873862.8533 | 143 | 0.20673 | 0.002714 | 3.8e-05 | 101.27 | 98.378 | 118.5144 | 21.0741 | 0.002837 | 0.002617 | 170.9353 | 0.25622 | 8.4018 | 0.002332 | 0.003489 | 0.003444 | 6.7e-05 | 0.000856 | 0.31421 | 1.3313 | 2.5475 | 100.3 | 21.13 | 31.637 | 89.876 | 77.7423 | 84.6017 | 0.007917 | 0.98684 | 0.023426 | 23.7378 | 0.25915 | 0.15243 | 0.013778 | 0.96698 | 0.96698 | 0.006416
2019-03-29 11:38 | 1553873863.8529 | 144 | 0.20669 | 0.002707 | 3.8e-05 | 101.2353 | 98.5145 | 118.509 | 21.0729 | 0.002794 | 0.002583 | 170.8675 | 0.25667 | 8.4017 | 0.002366 | 0.004307 | 0.004187 | 6.1e-05 | 0.00086 | 0.30845 | 1.3663 | 2.5475 | 100.3 | 21.13 | 31.624 | 88.8427 | 78.558 | 83.7798 | 0.007913 | 0.98912 | 0.023452 | 23.71 | 0.24368 | 0.15269 | 0.01405 | 0.98723 | 0.98723 | 0.006295
2019-03-29 11:38 | 1553873864.8534 | 145 | 0.2067 | 0.002706 | 3.9e-05 | 101.279 | 98.5141 | 118.8217 | 21.075 | 0.002838 | 0.002583 | 170.8432 | 0.25569 | 8.4027 | 0.002374 | 0.004052 | 0.003969 | 6e-05 | 0.00086 | 0.31421 | 1.3588 | 2.5474 | 100.3 | 21.13 | 31.619 | 89.9051 | 78.7501 | 83.7763 | 0.007912 | 0.98913 | 0.023483 | 23.7411 | 0.24757 | 0.15211 | 0.013932 | 0.98017 | 0.98017 | 0.006316
2019-03-29 11:38 | 1553873865.8531 | 146 | 0.20669 | 0.002702 | 4e-05 | 101.235 | 98.5128 | 119.1981 | 21.0745 | 0.002838 | 0.002573 | 170.7998 | 0.25718 | 8.4052 | 0.002446 | 0.00409 | 0.003938 | 6.2e-05 | 0.000861 | 0.2943 | 1.3267 | 2.5473 | 100.3 | 21.13 | 31.615 | 89.9202 | 80.4805 | 83.5307 | 0.007911 | 0.98981 | 0.023528 | 23.7705 | 0.24381 | 0.153 | 0.014053 | 0.99057 | 0.99057 | 0.006286
2019-03-29 11:38 | 1553873866.8522 | 147 | 0.2067 | 0.002702 | 4.1e-05 | 101.2796 | 98.3698 | 118.9721 | 21.0741 | 0.002796 | 0.002594 | 170.7249 | 0.25625 | 8.4017 | 0.002462 | 0.00393 | 0.003857 | 6.1e-05 | 0.000855 | 0.29479 | 1.3677 | 2.5475 | 100.3 | 21.12 | 31.607 | 88.8964 | 80.8654 | 84.0382 | 0.007904 | 0.9884 | 0.023489 | 23.7648 | 0.26074 | 0.15244 | 0.013972 | 0.98313 | 0.98313 | 0.006311
2019-03-29 11:38 | 1553873867.8521 | 148 | 0.2067 | 0.002698 | 4.3e-05 | 101.2517 | 98.7496 | 118.77 | 21.0747 | 0.002827 | 0.002549 | 170.6563 | 0.25633 | 8.4006 | 0.002408 | 0.004483 | 0.0044 | 6.3e-05 | 0.000859 | 0.3079 | 1.3532 | 2.5475 | 100.3 | 21.12 | 31.604 | 89.6524 | 79.5647 | 82.9624 | 0.007904 | 0.99139 | 0.023505 | 23.7089 | 0.22381 | 0.15249 | 0.013981 | 0.98429 | 0.98429 | 0.006305
2019-03-29 11:38 | 1553873868.8532 | 149 | 0.20672 | 0.002692 | 4.4e-05 | 101.2781 | 98.651 | 119.059 | 21.0748 | 0.002855 | 0.002557 | 170.6033 | 0.25665 | 8.4016 | 0.002461 | 0.00431 | 0.00421 | 6.3e-05 | 0.000858 | 0.28423 | 1.3164 | 2.5473 | 100.3 | 21.12 | 31.598 | 90.333 | 80.8511 | 83.1449 | 0.007902 | 0.99088 | 0.023527 | 23.7438 | 0.23508 | 0.15268 | 0.013845 | 0.97562 | 0.97562 | 0.006369
2019-03-29 11:38 | 1553873869.8523 | 150 | 0.20677 | 0.002692 | 4.6e-05 | 101.2957 | 98.4626 | 119.3498 | 21.0751 | 0.002826 | 0.002571 | 170.5474 | 0.2559 | 8.4028 | 0.002357 | 0.003745 | 0.003643 | 6.3e-05 | 0.000856 | 0.28943 | 1.3186 | 2.5475 | 100.3 | 21.12 | 31.596 | 89.6207 | 78.3434 | 83.4909 | 0.007902 | 0.98992 | 0.023545 | 23.7843 | 0.25373 | 0.15223 | 0.013532 | 0.95409 | 0.95409 | 0.006494
2019-03-29 11:38 | 1553873870.8532 | 151 | 0.20674 | 0.002693 | 4.7e-05 | 101.2466 | 98.465 | 119.0451 | 21.0755 | 0.00275 | 0.002541 | 170.4913 | 0.25573 | 8.4029 | 0.002383 | 0.003651 | 0.003456 | 6.2e-05 | 0.000859 | 0.29695 | 1.3413 | 2.5475 | 100.3 | 21.12 | 31.594 | 87.7846 | 78.9642 | 82.7705 | 0.007901 | 0.99192 | 0.023538 | 23.7299 | 0.24918 | 0.15213 | 0.013756 | 0.96958 | 0.96958 | 0.006386
2019-03-29 11:38 | 1553873871.8522 | 152 | 0.20675 | 0.0027 | 4.7e-05 | 101.2553 | 98.5862 | 118.7841 | 21.0761 | 0.002833 | 0.002568 | 170.4535 | 0.25649 | 8.4004 | 0.002437 | 0.003532 | 0.003519 | 6.1e-05 | 0.000859 | 0.31277 | 1.3656 | 2.8622 | 100.3 | 21.12 | 31.592 | 89.7999 | 80.2795 | 83.4068 | 0.007901 | 0.99015 | 0.023491 | 23.7251 | 0.23895 | 0.15258 | 0.013647 | 0.95996 | 0.95996 | 0.006469
2019-03-29 11:38 | 1553873872.8522 | 153 | 0.20674 | 0.002705 | 4.7e-05 | 101.3221 | 98.4346 | 118.9787 | 21.0759 | 0.0028 | 0.002569 | 170.4032 | 0.25646 | 8.4028 | 0.00236 | 0.004364 | 0.004233 | 6e-05 | 0.000854 | 0.32604 | 1.3387 | 2.5474 | 100.3 | 21.12 | 31.593 | 88.9925 | 78.4199 | 83.4424 | 0.007901 | 0.99005 | 0.023509 | 23.7457 | 0.2586 | 0.15257 | 0.013726 | 0.9663 | 0.9663 | 0.006426
2019-03-29 11:38 | 1553873873.8523 | 154 | 0.20672 | 0.002701 | 4.7e-05 | 101.2514 | 98.5073 | 118.8627 | 21.076 | 0.002826 | 0.002538 | 170.354 | 0.25704 | 8.4027 | 0.002367 | 0.003145 | 0.003018 | 6.3e-05 | 0.00085 | 0.28871 | 1.3191 | 2.5474 | 100.3 | 21.11 | 31.596 | 89.6209 | 78.5774 | 82.6852 | 0.007897 | 0.99216 | 0.023523 | 23.7089 | 0.24576 | 0.15291 | 0.013846 | 0.97529 | 0.97529 | 0.006381
2019-03-29 11:38 | 1553873874.8524 | 155 | 0.20677 | 0.002701 | 4.8e-05 | 101.3254 | 98.5657 | 118.6947 | 21.0754 | 0.002769 | 0.002542 | 170.2817 | 0.25673 | 8.4034 | 0.002399 | 0.00435 | 0.004356 | 6.2e-05 | 0.000852 | 0.29358 | 1.2929 | 2.5475 | 100.3 | 21.11 | 31.597 | 88.258 | 79.3524 | 82.7976 | 0.007897 | 0.99185 | 0.023503 | 23.6959 | 0.24699 | 0.15273 | 0.013513 | 0.95094 | 0.95094 | 0.006537
2019-03-29 11:38 | 1553873875.8537 | 156 | 0.20675 | 0.002702 | 5e-05 | 101.2622 | 98.9981 | 118.7078 | 21.0762 | 0.002806 | 0.002523 | 170.229 | 0.25685 | 8.4017 | 0.002421 | 0.00418 | 0.004008 | 6.1e-05 | 0.000853 | 0.30198 | 1.3272 | 2.5474 | 100.3 | 21.11 | 31.6 | 89.1327 | 79.8764 | 82.3313 | 0.007898 | 0.99315 | 0.023519 | 23.6816 | 0.20226 | 0.1528 | 0.013671 | 0.96262 | 0.96262 | 0.006461
2019-03-29 11:38 | 1553873876.8524 | 157 | 0.20677 | 0.002697 | 5.1e-05 | 101.287 | 98.9588 | 119.1059 | 21.0765 | 0.002804 | 0.002573 | 170.177 | 0.25668 | 8.4027 | 0.002399 | 0.003586 | 0.003526 | 6e-05 | 0.00086 | 0.28998 | 1.2989 | 2.5473 | 100.3 | 21.11 | 31.604 | 89.1057 | 79.3542 | 83.5264 | 0.007899 | 0.98982 | 0.023519 | 23.7612 | 0.208 | 0.1527 | 0.013564 | 0.95507 | 0.95507 | 0.006507
2019-03-29 11:38 | 1553873877.8535 | 158 | 0.20678 | 0.002693 | 5.1e-05 | 101.2655 | 98.9861 | 119.185 | 21.0753 | 0.002825 | 0.002574 | 170.1211 | 0.25721 | 8.4011 | 0.002396 | 0.004121 | 0.004033 | 6.2e-05 | 0.000858 | 0.29518 | 1.3555 | 2.5475 | 100.3 | 21.11 | 31.607 | 89.5958 | 79.2884 | 83.5659 | 0.007899 | 0.98971 | 0.023526 | 23.7704 | 0.20363 | 0.15302 | 0.013461 | 0.94804 | 0.94804 | 0.006569
2019-03-29 11:38 | 1553873878.8524 | 159 | 0.20678 | 0.002687 | 5.1e-05 | 101.3107 | 99.2145 | 118.9921 | 21.0747 | 0.002828 | 0.002547 | 170.0703 | 0.25633 | 8.4035 | 0.002419 | 0.004165 | 0.004093 | 6e-05 | 0.000858 | 0.26969 | 1.354 | 2.5474 | 100.3 | 21.11 | 31.609 | 89.6711 | 79.8282 | 82.9132 | 0.0079 | 0.99152 | 0.023528 | 23.7294 | 0.18701 | 0.15249 | 0.013481 | 0.94951 | 0.94951 | 0.006537
2019-03-29 11:38 | 1553873879.8525 | 160 | 0.20678 | 0.002679 | 5.2e-05 | 101.2583 | 99.4588 | 119.0421 | 21.075 | 0.002836 | 0.002531 | 170.0307 | 0.25662 | 8.4018 | 0.00243 | 0.003895 | 0.003791 | 6.3e-05 | 0.000858 | 0.29231 | 1.3448 | 2.5475 | 100.3 | 21.1 | 31.609 | 89.8668 | 80.0927 | 82.5196 | 0.007895 | 0.99262 | 0.023546 | 23.7212 | 0.16038 | 0.15266 | 0.013506 | 0.95195 | 0.95195 | 0.006527
2019-03-29 11:38 | 1553873880.8526 | 161 | 0.2068 | 0.002676 | 5.3e-05 | 101.3119 | 99.4097 | 119.1094 | 21.0755 | 0.002766 | 0.002585 | 169.9669 | 0.25622 | 8.4023 | 0.002367 | 0.003373 | 0.003309 | 5.9e-05 | 0.000858 | 0.27721 | 1.3858 | 2.5474 | 100.3 | 21.1 | 31.609 | 88.1728 | 78.597 | 83.8212 | 0.007895 | 0.989 | 0.02351 | 23.7713 | 0.16954 | 0.15243 | 0.013409 | 0.94364 | 0.94364 | 0.006574
2019-03-29 11:38 | 1553873881.8526 | 162 | 0.20683 | 0.00267 | 5.3e-05 | 101.311 | 99.436 | 118.9423 | 21.0758 | 0.002805 | 0.00253 | 169.9017 | 0.25659 | 8.403 | 0.002398 | 0.003726 | 0.003577 | 6.1e-05 | 0.000859 | 0.29679 | 1.3573 | 2.5474 | 100.3 | 21.1 | 31.61 | 89.1104 | 79.3288 | 82.4922 | 0.007895 | 0.9927 | 0.023537 | 23.7104 | 0.16709 | 0.15265 | 0.013229 | 0.93201 | 0.93201 | 0.006666
2019-03-29 11:38 | 1553873882.8516 | 163 | 0.2068 | 0.002664 | 5.4e-05 | 101.252 | 99.4756 | 118.5128 | 21.0759 | 0.002857 | 0.002589 | 169.8327 | 0.2569 | 8.402 | 0.0024 | 0.004225 | 0.00409 | 6.5e-05 | 0.00086 | 0.30093 | 1.3345 | 2.5475 | 100.3 | 21.1 | 31.61 | 90.3756 | 79.3796 | 83.913 | 0.007895 | 0.98875 | 0.023448 | 23.7148 | 0.15832 | 0.15283 | 0.013394 | 0.94002 | 0.94002 | 0.006617
2019-03-29 11:38 | 1553873883.8527 | 164 | 0.2068 | 0.002661 | 5.6e-05 | 101.2547 | 99.5814 | 118.6541 | 21.0768 | 0.002778 | 0.002593 | 169.7575 | 0.25685 | 8.4006 | 0.002459 | 0.003575 | 0.003514 | 6.2e-05 | 0.000853 | 0.30309 | 1.3297 | 2.5474 | 100.3 | 21.1 | 31.61 | 88.4692 | 80.7892 | 84.0136 | 0.007895 | 0.98847 | 0.023459 | 23.7322 | 0.14905 | 0.1528 | 0.013378 | 0.93924 | 0.93924 | 0.006621
2019-03-29 11:38 | 1553873884.8531 | 165 | 0.20678 | 0.002657 | 5.7e-05 | 101.3143 | 99.567 | 118.7437 | 21.0761 | 0.002809 | 0.002589 | 169.6837 | 0.25626 | 8.4013 | 0.002396 | 0.004189 | 0.004117 | 6.3e-05 | 0.000849 | 0.27848 | 1.3339 | 2.5475 | 100.3 | 21.1 | 31.609 | 89.2046 | 79.2837 | 83.9271 | 0.007895 | 0.98871 | 0.02347 | 23.7383 | 0.1556 | 0.15245 | 0.013507 | 0.94868 | 0.94868 | 0.00654
2019-03-29 11:38 | 1553873885.8517 | 166 | 0.20679 | 0.002655 | 5.9e-05 | 101.2553 | 99.5803 | 118.3679 | 21.0761 | 0.002845 | 0.002607 | 169.6219 | 0.25681 | 8.402 | 0.002464 | 0.00392 | 0.003911 | 6.3e-05 | 0.000854 | 0.27743 | 1.2974 | 2.5475 | 100.3 | 21.1 | 31.609 | 90.0812 | 80.9257 | 84.3665 | 0.007895 | 0.98749 | 0.023419 | 23.7153 | 0.1492 | 0.15278 | 0.013434 | 0.94141 | 0.94141 | 0.006605
2019-03-29 11:38 | 1553873886.8527 | 167 | 0.20681 | 0.00265 | 5.9e-05 | 101.2966 | 99.7281 | 118.3848 | 21.0758 | 0.002775 | 0.002586 | 169.5687 | 0.2563 | 8.4031 | 0.002465 | 0.003891 | 0.0038 | 5.8e-05 | 0.00085 | 0.31686 | 1.3148 | 2.5474 | 100.3 | 21.09 | 31.61 | 88.3958 | 80.9458 | 83.8434 | 0.00789 | 0.98894 | 0.023438 | 23.6997 | 0.13958 | 0.15247 | 0.01332 | 0.93409 | 0.93409 | 0.006644
2019-03-29 11:38 | 1553873887.8532 | 168 | 0.2068 | 0.002655 | 6e-05 | 101.2793 | 99.537 | 118.4175 | 21.0765 | 0.002824 | 0.002555 | 169.5145 | 0.25661 | 8.4017 | 0.002457 | 0.003944 | 0.003882 | 6.4e-05 | 0.000852 | 0.32532 | 1.3205 | 2.5474 | 100.3 | 21.09 | 31.612 | 89.5761 | 80.743 | 83.0983 | 0.007891 | 0.99101 | 0.023465 | 23.6782 | 0.15521 | 0.15266 | 0.013412 | 0.94161 | 0.94161 | 0.006598
2019-03-29 11:38 | 1553873888.8528 | 169 | 0.20682 | 0.002652 | 5.9e-05 | 101.3334 | 99.7158 | 118.286 | 21.0763 | 0.002858 | 0.002599 | 169.4588 | 0.25525 | 8.4047 | 0.002439 | 0.003927 | 0.003795 | 5.8e-05 | 0.000856 | 0.30917 | 1.3411 | 2.5475 | 100.3 | 21.09 | 31.614 | 90.4041 | 80.3186 | 84.1661 | 0.007891 | 0.98805 | 0.023417 | 23.7005 | 0.14394 | 0.15185 | 0.013286 | 0.9309 | 0.9309 | 0.006639
2019-03-29 11:38 | 1553873889.8528 | 170 | 0.20679 | 0.002654 | 5.8e-05 | 101.2517 | 99.5196 | 117.933 | 21.0769 | 0.002819 | 0.002597 | 169.3727 | 0.25584 | 8.4026 | 0.002429 | 0.003696 | 0.003614 | 6.1e-05 | 0.000852 | 0.30757 | 1.3684 | 2.5474 | 100.3 | 21.09 | 31.615 | 89.4623 | 80.0885 | 84.1083 | 0.007892 | 0.98821 | 0.023384 | 23.6632 | 0.15434 | 0.1522 | 0.013474 | 0.94283 | 0.94283 | 0.00657
2019-03-29 11:38 | 1553873890.8534 | 171 | 0.20682 | 0.002646 | 5.8e-05 | 101.2975 | 99.6588 | 118.0032 | 21.0759 | 0.002854 | 0.002626 | 169.3072 | 0.2563 | 8.4012 | 0.002445 | 0.004347 | 0.004196 | 5.7e-05 | 0.000841 | 0.28871 | 1.338 | 2.5473 | 100.3 | 21.09 | 31.62 | 90.3045 | 80.4492 | 84.8218 | 0.007893 | 0.98624 | 0.023368 | 23.6938 | 0.14588 | 0.15247 | 0.013304 | 0.93031 | 0.93031 | 0.006671
2019-03-29 11:38 | 1553873891.8532 | 172 | 0.20682 | 0.002636 | 5.8e-05 | 101.2391 | 99.7732 | 118.4055 | 21.0754 | 0.002831 | 0.002594 | 169.2284 | 0.25661 | 8.404 | 0.002461 | 0.003509 | 0.003402 | 6.1e-05 | 0.000838 | 0.29623 | 1.3259 | 2.5474 | 100.3 | 21.09 | 31.623 | 89.7359 | 80.8393 | 84.0443 | 0.007894 | 0.98838 | 0.023433 | 23.7084 | 0.13047 | 0.15266 | 0.013301 | 0.93267 | 0.93267 | 0.006662
2019-03-29 11:38 | 1553873892.8528 | 173 | 0.20681 | 0.002625 | 5.9e-05 | 101.3044 | 99.8179 | 118.5726 | 21.0764 | 0.002817 | 0.002577 | 169.1619 | 0.25603 | 8.4053 | 0.002465 | 0.004677 | 0.004557 | 6.3e-05 | 0.000832 | 0.2808 | 1.3651 | 2.5475 | 100.3 | 21.09 | 31.626 | 89.4055 | 80.9301 | 83.6295 | 0.007894 | 0.98953 | 0.023463 | 23.7113 | 0.13222 | 0.15232 | 0.013363 | 0.93817 | 0.93817 | 0.006608
2019-03-29 11:38 | 1553873893.853 | 174 | 0.20684 | 0.002618 | 5.9e-05 | 101.2876 | 99.7288 | 118.4445 | 21.0766 | 0.00284 | 0.002572 | 169.1204 | 0.25525 | 8.4043 | 0.002455 | 0.003821 | 0.003821 | 6e-05 | 0.000836 | 0.28136 | 1.358 | 2.9558 | 100.3 | 21.09 | 31.634 | 89.9738 | 80.6969 | 83.5025 | 0.007896 | 0.98989 | 0.023455 | 23.6943 | 0.13873 | 0.15185 | 0.013204 | 0.92661 | 0.92661 | 0.00667
2019-03-29 11:38 | 1553873894.853 | 175 | 0.20682 | 0.002613 | 5.8e-05 | 101.3469 | 99.6418 | 118.6958 | 21.0771 | 0.002808 | 0.00253 | 169.0728 | 0.25582 | 8.4013 | 0.002444 | 0.003128 | 0.003037 | 6.1e-05 | 0.00084 | 0.30574 | 1.3335 | 2.5772 | 100.3 | 21.09 | 31.637 | 89.2008 | 80.4418 | 82.4997 | 0.007897 | 0.99268 | 0.023513 | 23.6861 | 0.15177 | 0.15219 | 0.013326 | 0.93759 | 0.93759 | 0.006606
2019-03-29 11:38 | 1553873895.853 | 176 | 0.20682 | 0.002612 | 5.9e-05 | 101.3083 | 99.4783 | 118.9141 | 21.0771 | 0.00282 | 0.002607 | 169.0048 | 0.2559 | 8.4027 | 0.002453 | 0.004259 | 0.004079 | 5.8e-05 | 0.000847 | 0.32012 | 1.3085 | 2.5476 | 100.3 | 21.09 | 31.642 | 89.489 | 80.6477 | 84.3501 | 0.007898 | 0.98754 | 0.023473 | 23.7694 | 0.16305 | 0.15224 | 0.013311 | 0.93488 | 0.93488 | 0.006628
2019-03-29 11:38 | 1553873896.8515 | 177 | 0.20685 | 0.002607 | 5.9e-05 | 101.3098 | 99.2503 | 118.8463 | 21.0776 | 0.002767 | 0.002589 | 168.9443 | 0.25653 | 8.4013 | 0.002492 | 0.003797 | 0.003702 | 6.3e-05 | 0.000844 | 0.27256 | 1.3535 | 2.5475 | 100.3 | 21.08 | 31.644 | 88.2108 | 81.5971 | 83.9178 | 0.007894 | 0.98874 | 0.023481 | 23.7483 | 0.1837 | 0.15261 | 0.013175 | 0.92563 | 0.92563 | 0.00671
2019-03-29 11:38 | 1553873897.8521 | 178 | 0.2068 | 0.002607 | 5.8e-05 | 101.2876 | 99.3711 | 118.3912 | 21.0768 | 0.002899 | 0.002605 | 168.8896 | 0.2561 | 8.4021 | 0.002465 | 0.003853 | 0.003728 | 6.1e-05 | 0.000851 | 0.29574 | 1.3714 | 2.5474 | 100.3 | 21.08 | 31.647 | 91.3795 | 80.9459 | 84.3047 | 0.007895 | 0.98767 | 0.023423 | 23.7156 | 0.17087 | 0.15235 | 0.013436 | 0.94176 | 0.94176 | 0.006584
2019-03-29 11:38 | 1553873898.8511 | 179 | 0.20685 | 0.002612 | 5.7e-05 | 101.3529 | 99.1397 | 118.5862 | 21.0776 | 0.002813 | 0.002608 | 168.8223 | 0.25593 | 8.4007 | 0.002483 | 0.003954 | 0.003861 | 6.1e-05 | 0.000854 | 0.3183 | 1.3248 | 2.5475 | 100.3 | 21.08 | 31.646 | 89.3022 | 81.3858 | 84.3725 | 0.007895 | 0.98748 | 0.02344 | 23.7374 | 0.19748 | 0.15225 | 0.013141 | 0.9218 | 0.9218 | 0.006722
2019-03-29 11:38 | 1553873899.8522 | 180 | 0.20683 | 0.002611 | 5.6e-05 | 101.308 | 99.3605 | 118.8346 | 21.0772 | 0.002723 | 0.002622 | 168.7567 | 0.2567 | 8.4023 | 0.00246 | 0.00468 | 0.004679 | 6.1e-05 | 0.00086 | 0.31078 | 1.3744 | 2.5475 | 100.3 | 21.08 | 31.646 | 87.143 | 80.8115 | 84.7194 | 0.007895 | 0.98652 | 0.023453 | 23.7737 | 0.17362 | 0.15271 | 0.013285 | 0.93249 | 0.93249 | 0.006665
2019-03-29 11:38 | 1553873900.8523 | 181 | 0.20682 | 0.002608 | 5.6e-05 | 101.2595 | 99.2145 | 118.9991 | 21.0772 | 0.002776 | 0.002589 | 168.7122 | 0.25618 | 8.4036 | 0.002463 | 0.003584 | 0.003426 | 6.3e-05 | 0.000856 | 0.30215 | 1.328 | 2.5474 | 100.3 | 21.08 | 31.645 | 88.4229 | 80.9011 | 83.9205 | 0.007894 | 0.98873 | 0.023496 | 23.7636 | 0.18249 | 0.1524 | 0.013335 | 0.9377 | 0.9377 | 0.006615
2019-03-29 11:38 | 1553873901.8512 | 182 | 0.20683 | 0.00261 | 5.6e-05 | 101.3161 | 99.4066 | 118.9813 | 21.0777 | 0.002812 | 0.002557 | 168.6491 | 0.25687 | 8.4008 | 0.002484 | 0.004357 | 0.004324 | 6.3e-05 | 0.000862 | 0.31669 | 1.3667 | 2.5475 | 100.3 | 21.08 | 31.645 | 89.2897 | 81.4015 | 83.1634 | 0.007894 | 0.99083 | 0.023519 | 23.7366 | 0.17018 | 0.15281 | 0.013306 | 0.93658 | 0.93658 | 0.006641
2019-03-29 11:38 | 1553873902.8514 | 183 | 0.20683 | 0.002607 | 5.6e-05 | 101.2849 | 99.1786 | 118.9489 | 21.0778 | 0.002868 | 0.002576 | 168.6127 | 0.25695 | 8.4017 | 0.002518 | 0.004031 | 0.003919 | 5.9e-05 | 0.00086 | 0.30862 | 1.3435 | 2.5474 | 100.3 | 21.08 | 31.646 | 90.6501 | 82.2237 | 83.6174 | 0.007895 | 0.98957 | 0.023501 | 23.7485 | 0.18797 | 0.15286 | 0.013271 | 0.93338 | 0.93338 | 0.006665
2019-03-29 11:38 | 1553873903.8525 | 184 | 0.20683 | 0.002611 | 5.6e-05 | 101.2403 | 99.0858 | 118.8153 | 21.0777 | 0.002858 | 0.002605 | 168.5398 | 0.25685 | 8.4015 | 0.002497 | 0.003817 | 0.003847 | 6.1e-05 | 0.000859 | 0.30143 | 1.3231 | 2.5474 | 100.3 | 21.08 | 31.647 | 90.4055 | 81.7055 | 84.3024 | 0.007895 | 0.98767 | 0.023465 | 23.758 | 0.19241 | 0.1528 | 0.013267 | 0.93168 | 0.93168 | 0.006675
2019-03-29 11:38 | 1553873904.8514 | 185 | 0.20679 | 0.002609 | 5.6e-05 | 101.2861 | 98.6643 | 118.4348 | 21.0779 | 0.002799 | 0.002646 | 168.4804 | 0.25661 | 8.3998 | 0.002466 | 0.004381 | 0.00427 | 6.4e-05 | 0.000859 | 0.30055 | 1.3462 | 2.5474 | 100.3 | 21.08 | 31.647 | 88.9728 | 80.9674 | 85.2983 | 0.007895 | 0.98493 | 0.023395 | 23.7529 | 0.23458 | 0.15266 | 0.013537 | 0.94786 | 0.94786 | 0.006555
2019-03-29 11:38 | 1553873905.8527 | 186 | 0.20679 | 0.002617 | 5.5e-05 | 101.2966 | 98.6066 | 118.5728 | 21.0788 | 0.002826 | 0.002583 | 168.4344 | 0.25609 | 8.4017 | 0.002468 | 0.004207 | 0.004131 | 6.1e-05 | 0.000859 | 0.27831 | 1.3579 | 2.5474 | 100.3 | 21.08 | 31.646 | 89.6159 | 81.0041 | 83.7879 | 0.007895 | 0.98909 | 0.023458 | 23.7166 | 0.24074 | 0.15235 | 0.013502 | 0.948 | 0.948 | 0.006541
2019-03-29 11:38 | 1553873906.8513 | 187 | 0.20679 | 0.00262 | 5.5e-05 | 101.2864 | 98.1305 | 118.7437 | 21.0784 | 0.002853 | 0.002631 | 168.3743 | 0.25734 | 8.4019 | 0.002391 | 0.003797 | 0.003712 | 6.4e-05 | 0.000862 | 0.28241 | 1.3767 | 2.5475 | 100.3 | 21.07 | 31.646 | 90.2812 | 79.1715 | 84.9452 | 0.00789 | 0.9859 | 0.023437 | 23.7721 | 0.28313 | 0.15309 | 0.013533 | 0.94925 | 0.94925 | 0.006564
2019-03-29 11:38 | 1553873907.8524 | 188 | 0.2068 | 0.002627 | 5.5e-05 | 101.293 | 98.1241 | 119.2094 | 21.0785 | 0.00285 | 0.002628 | 168.2999 | 0.25707 | 8.4033 | 0.002385 | 0.003915 | 0.003811 | 6.6e-05 | 0.000865 | 0.30502 | 1.3605 | 2.5473 | 100.3 | 21.07 | 31.647 | 90.1967 | 79.0125 | 84.858 | 0.00789 | 0.98614 | 0.023486 | 23.8158 | 0.2843 | 0.15293 | 0.013416 | 0.94302 | 0.94302 | 0.0066
2019-03-29 11:38 | 1553873908.8524 | 189 | 0.20678 | 0.002624 | 5.5e-05 | 101.3206 | 98.1828 | 119.0451 | 21.0786 | 0.002874 | 0.002562 | 168.2402 | 0.25629 | 8.4029 | 0.00245 | 0.00386 | 0.003756 | 6e-05 | 0.000865 | 0.29374 | 1.3712 | 2.5474 | 100.3 | 21.07 | 31.648 | 90.78 | 80.5821 | 83.2813 | 0.00789 | 0.9905 | 0.023521 | 23.7469 | 0.28138 | 0.15247 | 0.013589 | 0.95661 | 0.95661 | 0.006487
2019-03-29 11:38 | 1553873909.853 | 190 | 0.20678 | 0.002628 | 5.5e-05 | 101.2712 | 97.8787 | 119.0855 | 21.0792 | 0.002884 | 0.002591 | 168.1731 | 0.25642 | 8.4037 | 0.002436 | 0.004419 | 0.00436 | 6.5e-05 | 0.000866 | 0.30807 | 1.3822 | 2.5475 | 100.3 | 21.07 | 31.65 | 91.015 | 80.2551 | 83.9823 | 0.007891 | 0.98856 | 0.023502 | 23.7743 | 0.30477 | 0.15254 | 0.013565 | 0.95417 | 0.95417 | 0.006507
2019-03-29 11:38 | 1553873910.8519 | 191 | 0.20681 | 0.002624 | 5.4e-05 | 101.3047 | 98.0636 | 118.8414 | 21.0799 | 0.002836 | 0.002604 | 168.0904 | 0.25686 | 8.4037 | 0.002446 | 0.00402 | 0.003992 | 6.4e-05 | 0.000863 | 0.32012 | 1.3887 | 2.5476 | 100.3 | 21.07 | 31.653 | 89.8723 | 80.4913 | 84.2851 | 0.007892 | 0.98772 | 0.023468 | 23.76 | 0.29084 | 0.15281 | 0.013389 | 0.94044 | 0.94044 | 0.006613
2019-03-29 11:38 | 1553873911.8516 | 192 | 0.20679 | 0.002626 | 5.4e-05 | 101.2748 | 97.8097 | 118.6454 | 21.0788 | 0.002856 | 0.00258 | 168.0327 | 0.25631 | 8.401 | 0.00241 | 0.003526 | 0.003429 | 6.3e-05 | 0.000855 | 0.28208 | 1.3689 | 2.5475 | 100.3 | 21.07 | 31.656 | 90.3512 | 79.6311 | 83.7042 | 0.007892 | 0.98933 | 0.023468 | 23.7211 | 0.31139 | 0.15248 | 0.013478 | 0.94674 | 0.94674 | 0.006555
2019-03-29 11:39 | 1553873912.8519 | 193 | 0.20677 | 0.002629 | 5.3e-05 | 101.2697 | 97.7285 | 118.4238 | 21.08 | 0.002881 | 0.002593 | 167.9477 | 0.25642 | 8.4022 | 0.002435 | 0.003991 | 0.003898 | 6e-05 | 0.000859 | 0.32372 | 1.3332 | 2.5473 | 100.3 | 21.07 | 31.658 | 90.9447 | 80.2121 | 84.03 | 0.007893 | 0.98842 | 0.023435 | 23.7097 | 0.31837 | 0.15255 | 0.01361 | 0.95474 | 0.95474 | 0.006503
2019-03-29 11:39 | 1553873913.8515 | 194 | 0.20679 | 0.002628 | 5.2e-05 | 101.2409 | 97.6875 | 118.4682 | 21.0793 | 0.002872 | 0.002595 | 167.8703 | 0.25655 | 8.4021 | 0.002351 | 0.003861 | 0.003869 | 6.3e-05 | 0.000861 | 0.32278 | 1.3318 | 2.5474 | 100.3 | 21.07 | 31.66 | 90.725 | 78.193 | 84.062 | 0.007893 | 0.98834 | 0.023439 | 23.7153 | 0.31958 | 0.15262 | 0.013486 | 0.94619 | 0.94619 | 0.006565
2019-03-29 11:39 | 1553873914.8516 | 195 | 0.20678 | 0.002631 | 5.1e-05 | 101.2799 | 97.6394 | 118.428 | 21.079 | 0.002878 | 0.002583 | 167.7921 | 0.25673 | 8.4009 | 0.002382 | 0.004425 | 0.004306 | 6.2e-05 | 0.000868 | 0.32947 | 1.3406 | 2.5474 | 100.3 | 21.07 | 31.659 | 90.879 | 78.9474 | 83.7833 | 0.007893 | 0.98911 | 0.023444 | 23.702 | 0.32743 | 0.15273 | 0.013548 | 0.95084 | 0.95084 | 0.006537
2019-03-29 11:39 | 1553873915.8507 | 196 | 0.20686 | 0.002629 | 5e-05 | 101.2739 | 97.8677 | 118.0766 | 21.0806 | 0.002882 | 0.002613 | 167.7116 | 0.25601 | 8.4034 | 0.002372 | 0.003753 | 0.003617 | 6.5e-05 | 0.000867 | 0.32101 | 1.3621 | 2.5473 | 100.3 | 21.07 | 31.659 | 90.9812 | 78.7144 | 84.5041 | 0.007893 | 0.98711 | 0.023385 | 23.6907 | 0.30601 | 0.1523 | 0.0131 | 0.91717 | 0.91717 | 0.006759
2019-03-29 11:39 | 1553873916.8507 | 197 | 0.20682 | 0.002625 | 4.8e-05 | 101.2697 | 97.9087 | 117.8537 | 21.0794 | 0.002839 | 0.002633 | 167.6406 | 0.25674 | 8.4008 | 0.002389 | 0.004243 | 0.004155 | 6.5e-05 | 0.00087 | 0.3131 | 1.3059 | 2.5474 | 100.3 | 21.07 | 31.658 | 89.9297 | 79.1162 | 84.9822 | 0.007893 | 0.9858 | 0.023348 | 23.6841 | 0.30189 | 0.15273 | 0.013341 | 0.93262 | 0.93262 | 0.006665
2019-03-29 11:39 | 1553873917.8508 | 198 | 0.20684 | 0.002617 | 4.7e-05 | 101.2703 | 98.1626 | 117.9116 | 21.0793 | 0.002862 | 0.002668 | 167.5793 | 0.25698 | 8.4014 | 0.002418 | 0.004357 | 0.004302 | 6.2e-05 | 0.000876 | 0.32748 | 1.3592 | 2.5473 | 100.3 | 21.07 | 31.658 | 90.4923 | 79.8053 | 85.8257 | 0.007893 | 0.98348 | 0.023326 | 23.7178 | 0.27878 | 0.15288 | 0.013235 | 0.92442 | 0.92442 | 0.006731
2019-03-29 11:39 | 1553873918.8518 | 199 | 0.20684 | 0.002612 | 4.6e-05 | 101.2056 | 98.2616 | 117.9159 | 21.0792 | 0.002831 | 0.002644 | 167.5036 | 0.25643 | 8.3999 | 0.002364 | 0.003653 | 0.003586 | 6e-05 | 0.000875 | 0.30055 | 1.3695 | 2.5474 | 100.3 | 21.07 | 31.657 | 89.759 | 78.5273 | 85.2512 | 0.007893 | 0.98506 | 0.023345 | 23.6993 | 0.26405 | 0.15255 | 0.013248 | 0.92617 | 0.92617 | 0.006704
2019-03-29 11:39 | 1553873919.8522 | 200 | 0.20686 | 0.002606 | 4.5e-05 | 101.2891 | 98.1756 | 117.5584 | 21.0785 | 0.002834 | 0.002585 | 167.4505 | 0.25599 | 8.4025 | 0.002387 | 0.003556 | 0.003419 | 6.3e-05 | 0.000881 | 0.29518 | 1.338 | 2.5474 | 100.3 | 21.07 | 31.655 | 89.8203 | 79.0755 | 83.8328 | 0.007892 | 0.98897 | 0.023356 | 23.6164 | 0.27926 | 0.15229 | 0.013103 | 0.91648 | 0.91648 | 0.006763
2019-03-29 11:39 | 1553873920.8519 | 201 | 0.20687 | 0.002599 | 4.3e-05 | 101.2047 | 97.9039 | 117.6163 | 21.079 | 0.00283 | 0.002642 | 167.3987 | 0.25641 | 8.4029 | 0.002354 | 0.003767 | 0.003741 | 6.2e-05 | 0.000879 | 0.29446 | 1.3193 | 2.5474 | 100.3 | 21.07 | 31.653 | 89.7237 | 78.2676 | 85.1978 | 0.007892 | 0.9852 | 0.023317 | 23.6674 | 0.29659 | 0.15254 | 0.013048 | 0.91123 | 0.91123 | 0.006813
2019-03-29 11:39 | 1553873921.8509 | 202 | 0.20687 | 0.002593 | 4.2e-05 | 101.2023 | 98.1244 | 117.6708 | 21.0789 | 0.00279 | 0.002655 | 167.3391 | 0.25616 | 8.4027 | 0.002337 | 0.002931 | 0.002812 | 6.2e-05 | 0.000888 | 0.29264 | 1.3248 | 2.5474 | 100.3 | 21.07 | 31.649 | 88.772 | 77.8615 | 85.5117 | 0.007891 | 0.98434 | 0.023312 | 23.6832 | 0.27626 | 0.15239 | 0.013051 | 0.91135 | 0.91135 | 0.006805
2019-03-29 11:39 | 1553873922.8522 | 203 | 0.20687 | 0.002583 | 4.1e-05 | 101.1948 | 98.2384 | 117.7668 | 21.0799 | 0.002871 | 0.002624 | 167.277 | 0.25669 | 8.4018 | 0.00237 | 0.004206 | 0.004185 | 5.9e-05 | 0.000889 | 0.29214 | 1.3452 | 2.5472 | 100.3 | 21.07 | 31.647 | 90.7116 | 78.6693 | 84.7751 | 0.00789 | 0.98637 | 0.023346 | 23.6685 | 0.26521 | 0.15271 | 0.013096 | 0.91585 | 0.91585 | 0.006786
2019-03-29 11:39 | 1553873923.8523 | 204 | 0.20689 | 0.002578 | 4e-05 | 101.2098 | 98.0285 | 117.899 | 21.0798 | 0.002907 | 0.002653 | 167.2439 | 0.2567 | 8.4037 | 0.002411 | 0.004421 | 0.004235 | 6e-05 | 0.000885 | 0.26609 | 1.3099 | 2.5475 | 100.3 | 21.07 | 31.644 | 91.5761 | 79.6366 | 85.475 | 0.007889 | 0.98444 | 0.023336 | 23.705 | 0.28567 | 0.15271 | 0.013005 | 0.90915 | 0.90915 | 0.006836
2019-03-29 11:39 | 1553873924.851 | 205 | 0.20688 | 0.002578 | 3.9e-05 | 101.1999 | 98.0582 | 118.16 | 21.0801 | 0.00285 | 0.002635 | 167.1968 | 0.25635 | 8.4024 | 0.002404 | 0.004074 | 0.003995 | 6.4e-05 | 0.000887 | 0.28174 | 1.3181 | 2.5473 | 100.3 | 21.07 | 31.639 | 90.2063 | 79.4873 | 85.0414 | 0.007888 | 0.98563 | 0.023376 | 23.7168 | 0.28208 | 0.1525 | 0.013054 | 0.91417 | 0.91417 | 0.00679
2019-03-29 11:39 | 1553873925.852 | 206 | 0.20689 | 0.002575 | 3.9e-05 | 101.1993 | 97.8104 | 118.4182 | 21.0797 | 0.002903 | 0.002636 | 167.1353 | 0.257 | 8.4017 | 0.002438 | 0.003837 | 0.003767 | 5.7e-05 | 0.000886 | 0.30093 | 1.316 | 2.5473 | 100.3 | 21.07 | 31.638 | 91.4846 | 80.2969 | 85.0604 | 0.007888 | 0.98558 | 0.023401 | 23.7434 | 0.30466 | 0.15289 | 0.013009 | 0.91201 | 0.91201 | 0.006823
2019-03-29 11:39 | 1553873926.8524 | 207 | 0.20689 | 0.002569 | 3.8e-05 | 101.1398 | 97.6602 | 118.2877 | 21.0803 | 0.002829 | 0.002637 | 167.0916 | 0.25502 | 8.4007 | 0.002376 | 0.003576 | 0.003454 | 6.1e-05 | 0.000879 | 0.2855 | 1.3692 | 2.5474 | 100.3 | 21.07 | 31.637 | 89.6961 | 78.8045 | 85.0837 | 0.007888 | 0.98552 | 0.023387 | 23.731 | 0.31314 | 0.15171 | 0.012966 | 0.9085 | 0.9085 | 0.006796
2019-03-29 11:39 | 1553873927.8501 | 208 | 0.20693 | 0.002561 | 3.8e-05 | 101.0718 | 97.6732 | 118.0989 | 21.0798 | 0.002864 | 0.002641 | 167.046 | 0.25631 | 8.403 | 0.002426 | 0.003217 | 0.003154 | 6.2e-05 | 0.000874 | 0.31542 | 1.3406 | 2.5474 | 100.3 | 21.07 | 31.636 | 90.5451 | 80.0139 | 85.183 | 0.007887 | 0.98524 | 0.023365 | 23.7154 | 0.30594 | 0.15248 | 0.012777 | 0.89437 | 0.89437 | 0.006939
2019-03-29 11:39 | 1553873928.851 | 209 | 0.20692 | 0.002553 | 3.9e-05 | 101.1006 | 97.5667 | 118.2989 | 21.0801 | 0.002841 | 0.002678 | 167.0065 | 0.25622 | 8.4026 | 0.002385 | 0.003718 | 0.003596 | 6e-05 | 0.000879 | 0.30502 | 1.3342 | 2.5474 | 100.3 | 21.06 | 31.636 | 89.9921 | 79.0155 | 86.0598 | 0.007882 | 0.98284 | 0.023357 | 23.7645 | 0.31824 | 0.15243 | 0.012803 | 0.89584 | 0.89584 | 0.006925
2019-03-29 11:39 | 1553873929.8515 | 210 | 0.20691 | 0.002545 | 3.9e-05 | 101.1284 | 97.5886 | 118.6773 | 21.0804 | 0.002884 | 0.002637 | 166.9466 | 0.25626 | 8.4027 | 0.002444 | 0.00356 | 0.003466 | 5.9e-05 | 0.00088 | 0.30342 | 1.3662 | 2.5473 | 100.3 | 21.06 | 31.636 | 91.0326 | 80.445 | 85.0761 | 0.007882 | 0.98554 | 0.023426 | 23.7698 | 0.3187 | 0.15245 | 0.012906 | 0.90575 | 0.90575 | 0.00685
2019-03-29 11:39 | 1553873930.8502 | 211 | 0.20691 | 0.002536 | 3.9e-05 | 101.0892 | 97.5234 | 118.7857 | 21.0795 | 0.002893 | 0.002611 | 166.9067 | 0.2562 | 8.4018 | 0.002478 | 0.003998 | 0.003881 | 6e-05 | 0.000884 | 0.33539 | 1.3544 | 2.5475 | 100.3 | 21.06 | 31.638 | 91.2375 | 81.2622 | 84.4446 | 0.007883 | 0.98728 | 0.023457 | 23.7597 | 0.32121 | 0.15241 | 0.012893 | 0.90608 | 0.90608 | 0.006846
2019-03-29 11:39 | 1553873931.8502 | 212 | 0.2069 | 0.00253 | 3.9e-05 | 101.0368 | 97.6248 | 118.8831 | 21.0799 | 0.002893 | 0.002651 | 166.859 | 0.25634 | 8.4004 | 0.002483 | 0.004332 | 0.004164 | 6.4e-05 | 0.000882 | 0.28385 | 1.3634 | 2.5475 | 100.3 | 21.06 | 31.641 | 91.2495 | 81.3732 | 85.4213 | 0.007884 | 0.98459 | 0.023435 | 23.8019 | 0.30728 | 0.1525 | 0.012977 | 0.9111 | 0.9111 | 0.006812
2019-03-29 11:39 | 1553873932.8512 | 213 | 0.2069 | 0.00253 | 3.8e-05 | 101.0598 | 97.6674 | 119.4078 | 21.08 | 0.002855 | 0.002652 | 166.8153 | 0.25649 | 8.405 | 0.002363 | 0.00382 | 0.003793 | 6.3e-05 | 0.000885 | 0.31094 | 1.3372 | 2.5475 | 100.3 | 21.06 | 31.644 | 90.3295 | 78.4899 | 85.4509 | 0.007884 | 0.98451 | 0.023486 | 23.8554 | 0.30541 | 0.15258 | 0.012959 | 0.9118 | 0.9118 | 0.006811
2019-03-29 11:39 | 1553873933.8516 | 214 | 0.20691 | 0.002521 | 3.9e-05 | 101.0299 | 97.5684 | 119.6237 | 21.0811 | 0.002859 | 0.00261 | 166.7608 | 0.25681 | 8.4028 | 0.002411 | 0.003193 | 0.003072 | 6.5e-05 | 0.000885 | 0.29695 | 1.3379 | 2.5474 | 100.3 | 21.06 | 31.648 | 90.4283 | 79.6542 | 84.429 | 0.007885 | 0.98732 | 0.023541 | 23.8429 | 0.31183 | 0.15277 | 0.012946 | 0.913 | 0.913 | 0.00681
2019-03-29 11:39 | 1553873934.8503 | 215 | 0.20691 | 0.002519 | 3.8e-05 | 100.9754 | 97.376 | 119.6422 | 21.0812 | 0.002848 | 0.002609 | 166.7246 | 0.25655 | 8.4039 | 0.00246 | 0.00376 | 0.003769 | 6.6e-05 | 0.000882 | 0.29551 | 1.3223 | 2.5475 | 100.3 | 21.06 | 31.654 | 90.1697 | 80.831 | 84.4097 | 0.007887 | 0.98738 | 0.023543 | 23.8441 | 0.32467 | 0.15262 | 0.012963 | 0.91436 | 0.91436 | 0.006793
2019-03-29 11:39 | 1553873935.8503 | 216 | 0.20692 | 0.002513 | 3.9e-05 | 100.9608 | 97.2514 | 119.5697 | 21.0813 | 0.002859 | 0.002609 | 166.6828 | 0.2567 | 8.3997 | 0.002404 | 0.003307 | 0.003101 | 6.3e-05 | 0.000886 | 0.31022 | 1.3078 | 2.5474 | 100.3 | 21.06 | 31.657 | 90.4302 | 79.4696 | 84.4167 | 0.007888 | 0.98735 | 0.023536 | 23.8371 | 0.33482 | 0.15271 | 0.012875 | 0.90781 | 0.90781 | 0.006846
2019-03-29 11:39 | 1553873936.8508 | 217 | 0.20689 | 0.00251 | 3.9e-05 | 100.9736 | 97.1203 | 119.7082 | 21.0809 | 0.002901 | 0.002676 | 166.6133 | 0.25573 | 8.4026 | 0.002414 | 0.003794 | 0.003742 | 6.2e-05 | 0.000881 | 0.31277 | 1.3091 | 2.5474 | 100.3 | 21.06 | 31.662 | 91.4249 | 79.7242 | 86.0116 | 0.007889 | 0.98297 | 0.023497 | 23.904 | 0.34802 | 0.15213 | 0.013076 | 0.92043 | 0.92043 | 0.006727
2019-03-29 11:39 | 1553873937.8517 | 218 | 0.20689 | 0.002511 | 3.9e-05 | 100.9063 | 96.521 | 119.6638 | 21.0817 | 0.002857 | 0.002613 | 166.5543 | 0.25592 | 8.4027 | 0.002333 | 0.003706 | 0.00357 | 6.6e-05 | 0.000883 | 0.28943 | 1.2882 | 2.5472 | 100.3 | 21.06 | 31.662 | 90.3774 | 77.7823 | 84.4999 | 0.007889 | 0.98713 | 0.023542 | 23.8492 | 0.39742 | 0.15225 | 0.013085 | 0.9229 | 0.9229 | 0.006714
2019-03-29 11:39 | 1553873938.8514 | 219 | 0.2069 | 0.002509 | 3.9e-05 | 100.8716 | 96.7442 | 119.6561 | 21.0814 | 0.002893 | 0.002639 | 166.4918 | 0.25572 | 8.4033 | 0.002353 | 0.003199 | 0.003135 | 6.4e-05 | 0.000883 | 0.30198 | 1.3423 | 2.5474 | 100.3 | 21.06 | 31.661 | 91.2408 | 78.252 | 85.1387 | 0.007889 | 0.98537 | 0.02352 | 23.8697 | 0.37368 | 0.15213 | 0.013006 | 0.91643 | 0.91643 | 0.006756
2019-03-29 11:39 | 1553873939.8505 | 220 | 0.20689 | 0.002506 | 4e-05 | 100.8443 | 96.879 | 119.6451 | 21.0812 | 0.0029 | 0.002618 | 166.4565 | 0.2558 | 8.4022 | 0.002324 | 0.003677 | 0.003506 | 6.4e-05 | 0.000882 | 0.31006 | 1.339 | 2.5475 | 100.3 | 21.06 | 31.662 | 91.4057 | 77.5644 | 84.6305 | 0.007889 | 0.98677 | 0.023536 | 23.8517 | 0.35881 | 0.15218 | 0.013092 | 0.92309 | 0.92309 | 0.00671
2019-03-29 11:39 | 1553873940.8519 | 221 | 0.20689 | 0.002501 | 4e-05 | 100.7444 | 96.7517 | 119.4815 | 21.082 | 0.002918 | 0.002649 | 166.404 | 0.25728 | 8.4014 | 0.00239 | 0.003657 | 0.003566 | 6e-05 | 0.00088 | 0.2975 | 1.3539 | 2.5474 | 100.3 | 21.06 | 31.663 | 91.8575 | 79.1428 | 85.378 | 0.007889 | 0.98471 | 0.023495 | 23.8603 | 0.3617 | 0.15305 | 0.013052 | 0.91862 | 0.91862 | 0.006781
2019-03-29 11:39 | 1553873941.8506 | 222 | 0.20695 | 0.002494 | 4.1e-05 | 100.7884 | 96.7685 | 119.2812 | 21.0817 | 0.002858 | 0.002625 | 166.3539 | 0.25616 | 8.4025 | 0.002344 | 0.003458 | 0.003381 | 6.2e-05 | 0.000874 | 0.28512 | 1.314 | 2.5474 | 100.3 | 21.06 | 31.664 | 90.4118 | 78.0355 | 84.7964 | 0.007889 | 0.98631 | 0.023495 | 23.8209 | 0.36406 | 0.15239 | 0.012703 | 0.89397 | 0.89397 | 0.006938
2019-03-29 11:39 | 1553873942.8505 | 223 | 0.20694 | 0.002486 | 4.1e-05 | 100.7638 | 97.0801 | 118.7928 | 21.0819 | 0.002878 | 0.002615 | 166.3085 | 0.25732 | 8.4023 | 0.002312 | 0.004108 | 0.003993 | 6.2e-05 | 0.000882 | 0.31022 | 1.3408 | 2.5474 | 100.3 | 21.05 | 31.66 | 90.8863 | 77.2625 | 84.5647 | 0.007884 | 0.98695 | 0.023454 | 23.7644 | 0.33313 | 0.15308 | 0.012804 | 0.89957 | 0.89957 | 0.006926
2019-03-29 11:39 | 1553873943.8499 | 224 | 0.20694 | 0.002483 | 4.1e-05 | 100.765 | 97.2401 | 118.2358 | 21.0821 | 0.002866 | 0.002627 | 166.266 | 0.25649 | 8.4033 | 0.00229 | 0.003604 | 0.003486 | 6.3e-05 | 0.000878 | 0.30989 | 1.3169 | 2.5473 | 100.3 | 21.05 | 31.659 | 90.5825 | 76.7304 | 84.8334 | 0.007883 | 0.98621 | 0.02339 | 23.7175 | 0.31849 | 0.15258 | 0.012768 | 0.89455 | 0.89455 | 0.006942
2019-03-29 11:39 | 1553873944.8506 | 225 | 0.20697 | 0.002475 | 4.2e-05 | 100.7216 | 97.3166 | 118.1468 | 21.0828 | 0.002859 | 0.002628 | 166.1991 | 0.25661 | 8.4055 | 0.002268 | 0.003972 | 0.00385 | 6.3e-05 | 0.000878 | 0.29048 | 1.3341 | 2.5474 | 100.3 | 21.05 | 31.657 | 90.416 | 76.2188 | 84.874 | 0.007883 | 0.98609 | 0.02338 | 23.71 | 0.30761 | 0.15266 | 0.012617 | 0.88357 | 0.88357 | 0.007032
2019-03-29 11:39 | 1553873945.85 | 226 | 0.20699 | 0.002472 | 4.2e-05 | 100.6528 | 97.3821 | 117.8067 | 21.083 | 0.002841 | 0.002669 | 166.1596 | 0.25648 | 8.4035 | 0.002261 | 0.0045 | 0.004394 | 6.1e-05 | 0.000877 | 0.28871 | 1.2982 | 2.5474 | 100.3 | 21.05 | 31.656 | 89.9848 | 76.0342 | 85.8426 | 0.007883 | 0.98343 | 0.023315 | 23.7078 | 0.29548 | 0.15258 | 0.012532 | 0.8751 | 0.8751 | 0.007096
2019-03-29 11:39 | 1553873946.8507 | 227 | 0.20697 | 0.002472 | 4.2e-05 | 100.7015 | 96.9039 | 117.7461 | 21.0823 | 0.002898 | 0.002621 | 166.1314 | 0.25621 | 8.4008 | 0.002295 | 0.003488 | 0.003415 | 6.3e-05 | 0.000872 | 0.28854 | 1.3553 | 2.5475 | 100.3 | 21.05 | 31.655 | 91.3558 | 76.8615 | 84.7032 | 0.007882 | 0.98657 | 0.023346 | 23.6641 | 0.34384 | 0.15242 | 0.012614 | 0.882 | 0.882 | 0.007033
2019-03-29 11:39 | 1553873947.8507 | 228 | 0.20701 | 0.002458 | 4.2e-05 | 100.6375 | 96.879 | 117.6593 | 21.0828 | 0.002832 | 0.002638 | 166.065 | 0.25655 | 8.403 | 0.002287 | 0.003659 | 0.003538 | 6.2e-05 | 0.000873 | 0.29695 | 1.3249 | 2.5473 | 100.3 | 21.05 | 31.656 | 89.7712 | 76.6772 | 85.0971 | 0.007883 | 0.98548 | 0.023325 | 23.6684 | 0.34044 | 0.15262 | 0.012412 | 0.86706 | 0.86706 | 0.007164
2019-03-29 11:39 | 1553873948.8507 | 229 | 0.20703 | 0.002447 | 4.3e-05 | 100.5558 | 96.8545 | 117.6429 | 21.0832 | 0.002898 | 0.00261 | 166.0085 | 0.25664 | 8.4033 | 0.002297 | 0.003807 | 0.003701 | 6.1e-05 | 0.00087 | 0.28694 | 1.3278 | 2.5474 | 100.3 | 21.05 | 31.656 | 91.3739 | 76.9114 | 84.4388 | 0.007883 | 0.98729 | 0.023345 | 23.645 | 0.33545 | 0.15267 | 0.012277 | 0.85833 | 0.85833 | 0.007239
2019-03-29 11:39 | 1553873949.8498 | 230 | 0.20702 | 0.002442 | 4.3e-05 | 100.5408 | 96.8367 | 117.652 | 21.083 | 0.002849 | 0.002649 | 165.9569 | 0.25641 | 8.4035 | 0.002354 | 0.003548 | 0.00347 | 6.2e-05 | 0.00087 | 0.29015 | 1.3208 | 2.5475 | 100.3 | 21.05 | 31.663 | 90.1973 | 78.2685 | 85.3805 | 0.007884 | 0.9847 | 0.023315 | 23.677 | 0.33576 | 0.15254 | 0.012377 | 0.86419 | 0.86419 | 0.007184
2019-03-29 11:39 | 1553873950.8508 | 231 | 0.20704 | 0.002433 | 4.3e-05 | 100.5315 | 97.108 | 117.7349 | 21.0824 | 0.002895 | 0.002635 | 165.9128 | 0.25648 | 8.4021 | 0.002348 | 0.003179 | 0.002992 | 6.3e-05 | 0.000864 | 0.28407 | 1.2915 | 2.5473 | 100.3 | 21.05 | 31.666 | 91.2981 | 78.129 | 85.0318 | 0.007885 | 0.98566 | 0.023334 | 23.6738 | 0.3099 | 0.15258 | 0.012236 | 0.85507 | 0.85507 | 0.007263
2019-03-29 11:39 | 1553873951.8511 | 232 | 0.20705 | 0.002428 | 4.4e-05 | 100.4872 | 96.9985 | 117.8668 | 21.0822 | 0.002906 | 0.002623 | 165.8556 | 0.25707 | 8.4051 | 0.002343 | 0.004727 | 0.004694 | 6.4e-05 | 0.000861 | 0.27865 | 1.3356 | 2.5472 | 100.3 | 21.05 | 31.669 | 91.5574 | 78.0096 | 84.7454 | 0.007886 | 0.98645 | 0.023357 | 23.6776 | 0.31606 | 0.15293 | 0.012183 | 0.85216 | 0.85216 | 0.007304
2019-03-29 11:39 | 1553873952.8511 | 233 | 0.20708 | 0.002413 | 4.3e-05 | 100.4723 | 97.0626 | 118.0952 | 21.0822 | 0.002891 | 0.002612 | 165.7997 | 0.25735 | 8.402 | 0.002362 | 0.002873 | 0.002783 | 6e-05 | 0.000859 | 0.30735 | 1.3263 | 2.5474 | 100.3 | 21.05 | 31.676 | 91.1868 | 78.4695 | 84.491 | 0.007888 | 0.98715 | 0.023388 | 23.6921 | 0.30881 | 0.1531 | 0.012041 | 0.84328 | 0.84328 | 0.007389
2019-03-29 11:39 | 1553873953.8508 | 234 | 0.20708 | 0.002405 | 4.4e-05 | 100.5237 | 96.9882 | 118.6083 | 21.0839 | 0.002885 | 0.002576 | 165.7396 | 0.25558 | 8.4016 | 0.002266 | 0.003566 | 0.00349 | 6.1e-05 | 0.000863 | 0.32388 | 1.3654 | 2.5474 | 100.3 | 21.05 | 31.681 | 91.0516 | 76.1607 | 83.6211 | 0.007889 | 0.98956 | 0.023467 | 23.7146 | 0.32025 | 0.15204 | 0.012066 | 0.84792 | 0.84792 | 0.007298
2019-03-29 11:39 | 1553873954.851 | 235 | 0.20712 | 0.00239 | 4.4e-05 | 100.4884 | 96.9702 | 118.6203 | 21.0833 | 0.002865 | 0.00267 | 165.6996 | 0.25712 | 8.4004 | 0.002253 | 0.003604 | 0.003524 | 6.3e-05 | 0.00086 | 0.28119 | 1.3258 | 2.5473 | 100.3 | 21.06 | 31.685 | 90.5625 | 75.8609 | 85.878 | 0.007895 | 0.98334 | 0.023394 | 23.7907 | 0.31878 | 0.15296 | 0.011837 | 0.8292 | 0.8292 | 0.007508
2019-03-29 11:39 | 1553873955.851 | 236 | 0.20712 | 0.002383 | 4.4e-05 | 100.4295 | 96.7978 | 118.8116 | 21.0829 | 0.002868 | 0.002622 | 165.6428 | 0.25601 | 8.4023 | 0.002259 | 0.004056 | 0.003841 | 6e-05 | 0.000861 | 0.28528 | 1.3493 | 2.5475 | 100.3 | 21.06 | 31.693 | 90.6325 | 75.9928 | 84.7187 | 0.007897 | 0.98652 | 0.023451 | 23.7714 | 0.32944 | 0.1523 | 0.011855 | 0.83249 | 0.83249 | 0.007446
2019-03-29 11:39 | 1553873956.8509 | 237 | 0.20716 | 0.002371 | 4.5e-05 | 100.3938 | 97.4364 | 118.9036 | 21.0838 | 0.002896 | 0.002592 | 165.5792 | 0.25789 | 8.4053 | 0.002279 | 0.003664 | 0.003681 | 6.2e-05 | 0.000863 | 0.27632 | 1.3345 | 2.5474 | 100.3 | 21.06 | 31.697 | 91.3229 | 76.4677 | 84.0041 | 0.007898 | 0.9885 | 0.023484 | 23.7569 | 0.26745 | 0.15342 | 0.011609 | 0.81629 | 0.81629 | 0.00765
2019-03-29 11:39 | 1553873957.85 | 238 | 0.20715 | 0.002362 | 4.4e-05 | 100.3855 | 97.4616 | 119.0388 | 21.0829 | 0.00284 | 0.002615 | 165.564 | 0.25692 | 8.4027 | 0.002316 | 0.003506 | 0.003402 | 6e-05 | 0.000858 | 0.28816 | 1.3113 | 2.5473 | 100.3 | 21.06 | 31.707 | 89.9674 | 77.3788 | 84.5643 | 0.0079 | 0.98695 | 0.023479 | 23.789 | 0.26439 | 0.15284 | 0.011696 | 0.82224 | 0.82224 | 0.007565
2019-03-29 11:39 | 1553873958.8495 | 239 | 0.20713 | 0.002356 | 4.5e-05 | 100.436 | 97.5268 | 118.9698 | 21.0836 | 0.002873 | 0.002641 | 165.5348 | 0.25696 | 8.4001 | 0.002256 | 0.003613 | 0.003548 | 6.1e-05 | 0.000862 | 0.32637 | 1.3102 | 2.5473 | 100.3 | 21.06 | 31.71 | 90.7551 | 75.932 | 85.1877 | 0.007901 | 0.98523 | 0.023451 | 23.8028 | 0.26292 | 0.15286 | 0.011784 | 0.82741 | 0.82741 | 0.007519
2019-03-29 11:39 | 1553873959.8491 | 240 | 0.20714 | 0.002342 | 4.6e-05 | 100.3732 | 97.2487 | 118.7994 | 21.0834 | 0.002839 | 0.002605 | 165.4751 | 0.25655 | 8.4047 | 0.002299 | 0.003935 | 0.003783 | 6.2e-05 | 0.000858 | 0.30663 | 1.3652 | 2.5474 | 100.3 | 21.06 | 31.714 | 89.957 | 76.9608 | 84.3117 | 0.007902 | 0.98764 | 0.023463 | 23.7567 | 0.28286 | 0.15262 | 0.011762 | 0.82628 | 0.82628 | 0.007518
2019-03-29 11:39 | 1553873960.8495 | 241 | 0.20714 | 0.00233 | 4.6e-05 | 100.4393 | 97.1234 | 118.9383 | 21.0837 | 0.002866 | 0.002612 | 165.4274 | 0.25631 | 8.4009 | 0.002268 | 0.003441 | 0.003417 | 6.2e-05 | 0.000845 | 0.27168 | 1.3325 | 2.5473 | 100.3 | 21.06 | 31.723 | 90.6053 | 76.2182 | 84.472 | 0.007904 | 0.9872 | 0.023472 | 23.7759 | 0.30028 | 0.15248 | 0.011776 | 0.82748 | 0.82748 | 0.0075
2019-03-29 11:39 | 1553873961.9319 | 242 | 0.20714 | 0.002317 | 4.7e-05 | 100.3489 | 97.3572 | 119.0775 | 21.0832 | 0.00288 | 0.002629 | 165.3879 | 0.25612 | 8.4055 | 0.002354 | 0.003747 | 0.003644 | 6.2e-05 | 0.000843 | 0.28871 | 1.3229 | 2.5475 | 100.3 | 21.06 | 31.726 | 90.9322 | 78.2866 | 84.896 | 0.007905 | 0.98603 | 0.023471 | 23.8039 | 0.27072 | 0.15237 | 0.011782 | 0.82786 | 0.82786 | 0.007491
2019-03-29 11:39 | 1553873962.8667 | 243 | 0.20717 | 0.00232 | 4.7e-05 | 100.3376 | 97.4452 | 119.4385 | 21.0843 | 0.002852 | 0.002561 | 165.3343 | 0.25618 | 8.4034 | 0.002235 | 0.004089 | 0.003969 | 6.1e-05 | 0.000839 | 0.30613 | 1.3439 | 2.5473 | 100.3 | 21.06 | 31.733 | 90.2573 | 75.4252 | 83.2594 | 0.007907 | 0.99056 | 0.023561 | 23.7854 | 0.26162 | 0.1524 | 0.011633 | 0.82049 | 0.82049 | 0.00756
2019-03-29 11:39 | 1553873963.8669 | 244 | 0.20725 | 0.002308 | 4.9e-05 | 100.3142 | 97.6162 | 119.5432 | 21.085 | 0.002888 | 0.00261 | 165.2834 | 0.25576 | 8.4029 | 0.002317 | 0.003878 | 0.003734 | 5.7e-05 | 0.000847 | 0.32388 | 1.3244 | 2.5474 | 100.3 | 21.06 | 31.736 | 91.1238 | 77.3972 | 84.4454 | 0.007907 | 0.98728 | 0.023532 | 23.8354 | 0.24386 | 0.15215 | 0.011159 | 0.7859 | 0.7859 | 0.00788
2019-03-29 11:39 | 1553873964.8668 | 245 | 0.20723 | 0.002301 | 5e-05 | 100.2618 | 97.5971 | 119.5143 | 21.0846 | 0.002799 | 0.002648 | 165.2303 | 0.25584 | 8.4011 | 0.002326 | 0.003823 | 0.003845 | 5.9e-05 | 0.00084 | 0.30342 | 1.3872 | 2.5474 | 100.3 | 21.06 | 31.738 | 88.9933 | 77.5987 | 85.3479 | 0.007908 | 0.98479 | 0.0235 | 23.8626 | 0.24094 | 0.1522 | 0.011244 | 0.79078 | 0.79078 | 0.007834
2019-03-29 11:39 | 1553873965.8669 | 246 | 0.20727 | 0.002291 | 5e-05 | 100.2717 | 97.5043 | 119.7533 | 21.0844 | 0.002875 | 0.002613 | 165.1741 | 0.25646 | 8.4041 | 0.002258 | 0.003674 | 0.00351 | 6e-05 | 0.000839 | 0.31116 | 1.3673 | 2.5473 | 100.3 | 21.05 | 31.742 | 90.817 | 75.9797 | 84.5133 | 0.007904 | 0.98709 | 0.023551 | 23.8586 | 0.25033 | 0.15257 | 0.011016 | 0.77633 | 0.77633 | 0.007998
2019-03-29 11:39 | 1553873966.8683 | 247 | 0.20725 | 0.002275 | 5.2e-05 | 100.2436 | 97.2589 | 119.6566 | 21.0849 | 0.0029 | 0.002605 | 165.1421 | 0.25626 | 8.4006 | 0.002303 | 0.003785 | 0.003711 | 5.8e-05 | 0.000838 | 0.27815 | 1.3853 | 2.5474 | 100.3 | 21.05 | 31.744 | 91.4114 | 77.0495 | 84.3094 | 0.007904 | 0.98765 | 0.023548 | 23.8422 | 0.27036 | 0.15245 | 0.011162 | 0.78646 | 0.78646 | 0.007889
2019-03-29 11:39 | 1553873967.867 | 248 | 0.20728 | 0.002266 | 5.2e-05 | 100.2139 | 97.3408 | 119.7143 | 21.0846 | 0.002858 | 0.002584 | 165.0915 | 0.25616 | 8.4041 | 0.002255 | 0.003531 | 0.003366 | 6.1e-05 | 0.000841 | 0.28495 | 1.3825 | 2.5475 | 100.3 | 21.06 | 31.747 | 90.4098 | 75.9142 | 83.8199 | 0.00791 | 0.98901 | 0.02357 | 23.8316 | 0.26019 | 0.15239 | 0.010976 | 0.77407 | 0.77407 | 0.008012
2019-03-29 11:39 | 1553873968.867 | 249 | 0.2073 | 0.002251 | 5.2e-05 | 100.2181 | 97.5548 | 119.8859 | 21.0846 | 0.002861 | 0.002641 | 165.0419 | 0.25617 | 8.3986 | 0.002247 | 0.003589 | 0.00357 | 5.8e-05 | 0.000843 | 0.29031 | 1.3822 | 2.5474 | 100.3 | 21.06 | 31.749 | 90.4826 | 75.7225 | 85.18 | 0.007911 | 0.98525 | 0.023542 | 23.894 | 0.24092 | 0.1524 | 0.01091 | 0.76843 | 0.76843 | 0.008072
2019-03-29 11:39 | 1553873969.8674 | 250 | 0.20729 | 0.002241 | 5.3e-05 | 100.1657 | 97.2906 | 120.1185 | 21.0845 | 0.002834 | 0.002616 | 165.0075 | 0.25643 | 8.4045 | 0.00225 | 0.00337 | 0.003162 | 5.9e-05 | 0.00084 | 0.30973 | 1.413 | 2.5475 | 100.3 | 21.06 | 31.749 | 89.8362 | 75.7918 | 84.5843 | 0.007911 | 0.98689 | 0.023584 | 23.8973 | 0.2605 | 0.15255 | 0.010976 | 0.77446 | 0.77446 | 0.008017
2019-03-29 11:39 | 1553873970.8674 | 251 | 0.20732 | 0.002225 | 5.3e-05 | 100.1978 | 97.6367 | 119.9415 | 21.0852 | 0.002877 | 0.002665 | 164.9446 | 0.25726 | 8.4008 | 0.00226 | 0.004598 | 0.004602 | 6.3e-05 | 0.000836 | 0.28385 | 1.3682 | 2.5473 | 100.3 | 21.06 | 31.749 | 90.8552 | 76.0218 | 85.7627 | 0.007911 | 0.98365 | 0.023528 | 23.919 | 0.2316 | 0.15304 | 0.010829 | 0.76226 | 0.76226 | 0.008172
2019-03-29 11:40 | 1553873971.8681 | 252 | 0.2073 | 0.002214 | 5.4e-05 | 100.1616 | 97.7968 | 119.9117 | 21.0855 | 0.002847 | 0.002655 | 164.8972 | 0.25579 | 8.4031 | 0.002304 | 0.003377 | 0.00331 | 5.9e-05 | 0.000842 | 0.29717 | 1.3194 | 2.5474 | 100.3 | 21.06 | 31.748 | 90.1498 | 77.0781 | 85.5195 | 0.00791 | 0.98432 | 0.023533 | 23.9079 | 0.21371 | 0.15217 | 0.010929 | 0.76945 | 0.76945 | 0.008049
2019-03-29 11:40 | 1553873972.867 | 253 | 0.20733 | 0.002203 | 5.3e-05 | 100.2256 | 97.8268 | 119.5629 | 21.0853 | 0.002881 | 0.002588 | 164.867 | 0.25626 | 8.4039 | 0.002277 | 0.004171 | 0.00403 | 5.9e-05 | 0.000841 | 0.28119 | 1.2986 | 2.5475 | 100.3 | 21.06 | 31.745 | 90.9506 | 76.4216 | 83.9111 | 0.00791 | 0.98875 | 0.023552 | 23.8195 | 0.21668 | 0.15245 | 0.010791 | 0.76031 | 0.76031 | 0.008161
2019-03-29 11:40 | 1553873973.8675 | 254 | 0.20732 | 0.002193 | 5.4e-05 | 100.1774 | 97.5654 | 119.3446 | 21.0852 | 0.002894 | 0.002628 | 164.8174 | 0.25658 | 8.4029 | 0.002292 | 0.003033 | 0.002958 | 5.9e-05 | 0.000849 | 0.2912 | 1.3402 | 2.5473 | 100.3 | 21.06 | 31.743 | 91.2672 | 76.7911 | 84.8754 | 0.007909 | 0.98609 | 0.023498 | 23.8299 | 0.23632 | 0.15264 | 0.010833 | 0.76146 | 0.76146 | 0.008159
2019-03-29 11:40 | 1553873974.8672 | 255 | 0.20735 | 0.002182 | 5.5e-05 | 100.1565 | 97.4009 | 119.4446 | 21.0854 | 0.002864 | 0.002605 | 164.7778 | 0.25631 | 8.4018 | 0.002339 | 0.003352 | 0.003314 | 5.7e-05 | 0.000851 | 0.31454 | 1.3213 | 2.5474 | 100.3 | 21.06 | 31.739 | 90.552 | 77.9214 | 84.306 | 0.007908 | 0.98766 | 0.023527 | 23.8209 | 0.24954 | 0.15248 | 0.010681 | 0.75164 | 0.75164 | 0.008256
2019-03-29 11:40 | 1553873975.8662 | 256 | 0.20737 | 0.002169 | 5.4e-05 | 100.1286 | 97.6974 | 119.6357 | 21.0856 | 0.002873 | 0.002579 | 164.7435 | 0.25648 | 8.4026 | 0.002348 | 0.00331 | 0.003151 | 5.8e-05 | 0.000855 | 0.31061 | 1.3166 | 2.5473 | 100.3 | 21.06 | 31.737 | 90.7548 | 78.133 | 83.6937 | 0.007908 | 0.98936 | 0.023566 | 23.8195 | 0.21986 | 0.15258 | 0.010594 | 0.74674 | 0.74674 | 0.008316
2019-03-29 11:40 | 1553873976.8666 | 257 | 0.20738 | 0.002164 | 5.5e-05 | 100.096 | 97.2231 | 119.8462 | 21.086 | 0.002894 | 0.002642 | 164.6873 | 0.25771 | 8.4025 | 0.002344 | 0.004294 | 0.004267 | 6.1e-05 | 0.000857 | 0.30574 | 1.3091 | 2.5473 | 100.3 | 21.06 | 31.735 | 91.2797 | 78.039 | 85.2133 | 0.007907 | 0.98516 | 0.023537 | 23.8912 | 0.26048 | 0.15331 | 0.010512 | 0.74003 | 0.74003 | 0.008432
2019-03-29 11:40 | 1553873977.8663 | 258 | 0.20738 | 0.00215 | 5.7e-05 | 100.0924 | 97.0903 | 120.1323 | 21.0846 | 0.002817 | 0.002631 | 164.6206 | 0.25735 | 8.4033 | 0.00234 | 0.003637 | 0.003532 | 5.9e-05 | 0.000853 | 0.32189 | 1.3004 | 2.5473 | 100.3 | 21.06 | 31.73 | 89.4284 | 77.9372 | 84.9286 | 0.007906 | 0.98594 | 0.023574 | 23.9102 | 0.27238 | 0.1531 | 0.010523 | 0.7419 | 0.7419 | 0.008399
2019-03-29 11:40 | 1553873978.8669 | 259 | 0.20736 | 0.002142 | 5.8e-05 | 100.1116 | 97.0422 | 120.1974 | 21.0843 | 0.002868 | 0.002586 | 164.5841 | 0.25611 | 8.4028 | 0.002332 | 0.003747 | 0.003651 | 6.1e-05 | 0.000856 | 0.28816 | 1.3138 | 2.5473 | 100.3 | 21.06 | 31.727 | 90.648 | 77.749 | 83.8517 | 0.007905 | 0.98892 | 0.023616 | 23.8807 | 0.27853 | 0.15236 | 0.010675 | 0.75384 | 0.75384 | 0.008226
2019-03-29 11:40 | 1553873979.8673 | 260 | 0.20739 | 0.002131 | 6e-05 | 100.1493 | 96.691 | 120.0469 | 21.0856 | 0.002825 | 0.002629 | 164.5338 | 0.25616 | 8.401 | 0.002296 | 0.00398 | 0.003873 | 5.7e-05 | 0.000856 | 0.27223 | 1.3135 | 2.5473 | 100.3 | 21.06 | 31.721 | 89.6099 | 76.8827 | 84.8843 | 0.007904 | 0.98607 | 0.023567 | 23.9002 | 0.31433 | 0.15239 | 0.010496 | 0.73955 | 0.73955 | 0.008387
2019-03-29 11:40 | 1553873980.8663 | 261 | 0.20739 | 0.00212 | 6.2e-05 | 100.1616 | 96.6278 | 120.2861 | 21.0862 | 0.002868 | 0.002594 | 164.4904 | 0.25666 | 8.4012 | 0.002315 | 0.003804 | 0.003725 | 5.9e-05 | 0.000859 | 0.28838 | 1.3214 | 2.5473 | 100.3 | 21.06 | 31.718 | 90.6419 | 77.3515 | 84.0572 | 0.007903 | 0.98835 | 0.023618 | 23.8964 | 0.32127 | 0.15269 | 0.01052 | 0.74272 | 0.74272 | 0.008367
2019-03-29 11:40 | 1553873981.8679 | 262 | 0.20738 | 0.002108 | 6.4e-05 | 100.1008 | 96.8193 | 119.9671 | 21.0858 | 0.002874 | 0.002597 | 164.4404 | 0.25739 | 8.4046 | 0.002334 | 0.004028 | 0.003999 | 5.7e-05 | 0.000855 | 0.30021 | 1.3401 | 2.5473 | 100.3 | 21.06 | 31.715 | 90.7878 | 77.7988 | 84.1254 | 0.007902 | 0.98816 | 0.023584 | 23.8669 | 0.29813 | 0.15312 | 0.010578 | 0.74562 | 0.74562 | 0.008358
2019-03-29 11:40 | 1553873982.8664 | 263 | 0.2074 | 0.002098 | 6.7e-05 | 100.0945 | 97.1664 | 119.9417 | 21.0848 | 0.002815 | 0.002664 | 164.383 | 0.25638 | 8.4027 | 0.002372 | 0.003733 | 0.003602 | 5.7e-05 | 0.000852 | 0.31852 | 1.2885 | 2.5473 | 100.3 | 21.06 | 31.71 | 89.3653 | 78.7255 | 85.7296 | 0.007901 | 0.98374 | 0.023529 | 23.9179 | 0.26556 | 0.15252 | 0.010487 | 0.73733 | 0.73733 | 0.008419
2019-03-29 11:40 | 1553873983.8664 | 264 | 0.20744 | 0.002091 | 6.8e-05 | 100.0493 | 97.2934 | 119.8256 | 21.086 | 0.002839 | 0.002613 | 164.333 | 0.2567 | 8.4054 | 0.002415 | 0.003622 | 0.003489 | 6e-05 | 0.00085 | 0.30541 | 1.3082 | 2.5474 | 100.3 | 21.06 | 31.708 | 89.9453 | 79.758 | 84.5197 | 0.0079 | 0.98707 | 0.023557 | 23.866 | 0.24984 | 0.15271 | 0.010275 | 0.72313 | 0.72313 | 0.008595
2019-03-29 11:40 | 1553873984.8674 | 265 | 0.20746 | 0.002083 | 6.9e-05 | 100.1241 | 97.4951 | 119.8002 | 21.0855 | 0.00282 | 0.002587 | 164.2848 | 0.25666 | 8.4016 | 0.002347 | 0.004304 | 0.004149 | 5.7e-05 | 0.000849 | 0.28191 | 1.312 | 2.5473 | 100.3 | 21.07 | 31.704 | 89.4776 | 78.1173 | 83.873 | 0.007904 | 0.98886 | 0.023576 | 23.8419 | 0.23801 | 0.15269 | 0.010135 | 0.71372 | 0.71372 | 0.008707
2019-03-29 11:40 | 1553873985.8675 | 266 | 0.20745 | 0.002083 | 7.1e-05 | 100.084 | 97.6633 | 119.3977 | 21.0854 | 0.002772 | 0.002573 | 164.222 | 0.25655 | 8.4017 | 0.002348 | 0.003999 | 0.00396 | 5.7e-05 | 0.000854 | 0.28456 | 1.3366 | 2.5474 | 100.3 | 21.07 | 31.702 | 88.3432 | 78.1359 | 83.5367 | 0.007904 | 0.98979 | 0.023548 | 23.7906 | 0.219 | 0.15262 | 0.010216 | 0.71853 | 0.71853 | 0.008645
2019-03-29 11:40 | 1553873986.8675 | 267 | 0.20741 | 0.002076 | 7.2e-05 | 100.0978 | 97.6971 | 119.3033 | 21.0862 | 0.002832 | 0.002571 | 164.1716 | 0.2569 | 8.4022 | 0.002373 | 0.003711 | 0.003666 | 5.7e-05 | 0.000853 | 0.31277 | 1.3293 | 2.5474 | 100.3 | 21.07 | 31.7 | 89.7753 | 78.7322 | 83.5025 | 0.007903 | 0.98989 | 0.02354 | 23.78 | 0.21714 | 0.15283 | 0.010447 | 0.73454 | 0.73454 | 0.008468
2019-03-29 11:40 | 1553873987.8675 | 268 | 0.20741 | 0.002079 | 7.4e-05 | 100.075 | 97.8265 | 118.6665 | 21.0866 | 0.002893 | 0.002588 | 164.1452 | 0.25726 | 8.4017 | 0.002363 | 0.004055 | 0.003979 | 5.8e-05 | 0.000852 | 0.30093 | 1.313 | 2.5474 | 100.3 | 21.07 | 31.695 | 91.2451 | 78.4901 | 83.9039 | 0.007902 | 0.98877 | 0.023463 | 23.7298 | 0.20327 | 0.15304 | 0.010463 | 0.73312 | 0.73312 | 0.008496
2019-03-29 11:40 | 1553873988.8665 | 269 | 0.20742 | 0.002078 | 7.6e-05 | 100.1346 | 97.7312 | 118.6233 | 21.0858 | 0.002835 | 0.00254 | 164.1004 | 0.25678 | 8.4046 | 0.002379 | 0.00387 | 0.00385 | 6.1e-05 | 0.000853 | 0.30254 | 1.3584 | 2.5474 | 100.3 | 21.07 | 31.693 | 89.8469 | 78.8894 | 82.7493 | 0.007901 | 0.99198 | 0.023497 | 23.6871 | 0.2173 | 0.15276 | 0.01034 | 0.72546 | 0.72546 | 0.00857
2019-03-29 11:40 | 1553873989.8676 | 270 | 0.20744 | 0.002077 | 7.8e-05 | 100.0879 | 97.7538 | 118.3858 | 21.0871 | 0.002861 | 0.002584 | 164.0469 | 0.2562 | 8.4057 | 0.002369 | 0.003795 | 0.003692 | 5.7e-05 | 0.00086 | 0.29103 | 1.3303 | 2.5474 | 100.3 | 21.07 | 31.691 | 90.4874 | 78.6566 | 83.8005 | 0.007901 | 0.98906 | 0.023439 | 23.6983 | 0.21107 | 0.15241 | 0.010251 | 0.71726 | 0.71726 | 0.008649
2019-03-29 11:40 | 1553873990.8667 | 271 | 0.20744 | 0.002075 | 8e-05 | 100.1427 | 97.4309 | 118.1806 | 21.0858 | 0.002922 | 0.002597 | 163.9955 | 0.25588 | 8.402 | 0.002395 | 0.003972 | 0.00389 | 5.6e-05 | 0.000859 | 0.29911 | 1.2889 | 2.5473 | 100.3 | 21.06 | 31.687 | 91.9581 | 79.2683 | 84.1284 | 0.007895 | 0.98815 | 0.023408 | 23.6887 | 0.24555 | 0.15222 | 0.01028 | 0.71822 | 0.71822 | 0.008626
2019-03-29 11:40 | 1553873991.8666 | 272 | 0.20744 | 0.002068 | 8.3e-05 | 100.0433 | 97.6469 | 118.5273 | 21.0866 | 0.002822 | 0.002544 | 163.9492 | 0.25732 | 8.4035 | 0.002411 | 0.003785 | 0.003688 | 5.6e-05 | 0.000854 | 0.29551 | 1.3195 | 2.5474 | 100.3 | 21.06 | 31.684 | 89.535 | 79.6471 | 82.8616 | 0.007894 | 0.99167 | 0.023484 | 23.6813 | 0.21686 | 0.15308 | 0.010244 | 0.71782 | 0.71782 | 0.00868
2019-03-29 11:40 | 1553873992.8658 | 273 | 0.20744 | 0.002065 | 8.5e-05 | 100.1023 | 97.7234 | 118.1778 | 21.0866 | 0.002816 | 0.002636 | 163.9177 | 0.25652 | 8.4025 | 0.002386 | 0.003593 | 0.003452 | 5.5e-05 | 0.000855 | 0.29446 | 1.3123 | 2.5473 | 100.3 | 21.06 | 31.68 | 89.4015 | 79.0538 | 85.0672 | 0.007893 | 0.98556 | 0.023377 | 23.7195 | 0.21513 | 0.1526 | 0.010228 | 0.71337 | 0.71337 | 0.008707
2019-03-29 11:40 | 1553873993.8658 | 274 | 0.20744 | 0.002063 | 8.6e-05 | 100.0298 | 97.7176 | 118.3107 | 21.0874 | 0.002816 | 0.00259 | 163.886 | 0.25722 | 8.4018 | 0.002395 | 0.00367 | 0.003617 | 6.1e-05 | 0.000848 | 0.30663 | 1.3162 | 2.5474 | 100.3 | 21.06 | 31.677 | 89.3935 | 79.2689 | 83.9562 | 0.007893 | 0.98863 | 0.023427 | 23.696 | 0.20919 | 0.15302 | 0.010227 | 0.71474 | 0.71474 | 0.008714
2019-03-29 11:40 | 1553873994.8668 | 275 | 0.20743 | 0.002052 | 8.7e-05 | 100.0188 | 97.4514 | 118.5515 | 21.0873 | 0.002808 | 0.002657 | 163.8331 | 0.25657 | 8.4025 | 0.00235 | 0.003436 | 0.003307 | 5.7e-05 | 0.000847 | 0.27632 | 1.3811 | 2.5475 | 100.3 | 21.06 | 31.673 | 89.1974 | 78.2 | 85.572 | 0.007892 | 0.98418 | 0.023397 | 23.7737 | 0.2326 | 0.15264 | 0.010335 | 0.72137 | 0.72137 | 0.008612
2019-03-29 11:40 | 1553873995.8658 | 276 | 0.20744 | 0.002054 | 8.9e-05 | 100.0855 | 97.5172 | 118.8503 | 21.0862 | 0.002846 | 0.002589 | 163.8006 | 0.25552 | 8.4049 | 0.002418 | 0.003477 | 0.003421 | 5.8e-05 | 0.000842 | 0.25912 | 1.3673 | 2.5476 | 100.3 | 21.06 | 31.67 | 90.1269 | 79.8137 | 83.9363 | 0.007891 | 0.98868 | 0.023481 | 23.7493 | 0.23252 | 0.15201 | 0.010249 | 0.71775 | 0.71775 | 0.00862
2019-03-29 11:40 | 1553873996.8658 | 277 | 0.20743 | 0.002048 | 9e-05 | 99.9897 | 97.5974 | 118.7102 | 21.0865 | 0.002841 | 0.002617 | 163.7775 | 0.25724 | 8.4002 | 0.002346 | 0.003478 | 0.003472 | 5.9e-05 | 0.000844 | 0.31797 | 1.3339 | 2.5475 | 100.3 | 21.06 | 31.668 | 89.9959 | 78.0824 | 84.5973 | 0.00789 | 0.98686 | 0.023445 | 23.7572 | 0.2166 | 0.15303 | 0.010309 | 0.72076 | 0.72076 | 0.008642
2019-03-29 11:40 | 1553873997.8669 | 278 | 0.20745 | 0.002046 | 9.2e-05 | 100.04 | 97.3067 | 119.0491 | 21.0868 | 0.002821 | 0.002652 | 163.7328 | 0.25583 | 8.4059 | 0.002316 | 0.003917 | 0.003857 | 5.8e-05 | 0.000846 | 0.29391 | 1.3515 | 2.5475 | 100.3 | 21.06 | 31.663 | 89.5203 | 77.3663 | 85.4466 | 0.007889 | 0.98452 | 0.023451 | 23.8194 | 0.24779 | 0.15219 | 0.010231 | 0.71538 | 0.71538 | 0.008659
2019-03-29 11:40 | 1553873998.8663 | 279 | 0.20743 | 0.002041 | 9.4e-05 | 100.0472 | 96.8268 | 118.8698 | 21.0869 | 0.002798 | 0.002648 | 163.6827 | 0.25629 | 8.4041 | 0.002393 | 0.003364 | 0.003209 | 5.8e-05 | 0.000849 | 0.32621 | 1.3347 | 2.5474 | 100.3 | 21.06 | 31.661 | 88.9634 | 79.2225 | 85.3539 | 0.007889 | 0.98477 | 0.023436 | 23.7983 | 0.29265 | 0.15246 | 0.01034 | 0.72249 | 0.72249 | 0.008589
2019-03-29 11:40 | 1553873999.867 | 280 | 0.20744 | 0.002031 | 9.6e-05 | 99.9616 | 96.8865 | 119.1784 | 21.0875 | 0.002825 | 0.002605 | 163.6457 | 0.25733 | 8.4021 | 0.002364 | 0.003685 | 0.003591 | 6.1e-05 | 0.000848 | 0.3126 | 1.3312 | 2.5474 | 100.3 | 21.06 | 31.659 | 89.6118 | 78.5176 | 84.3126 | 0.007888 | 0.98764 | 0.023501 | 23.7946 | 0.27948 | 0.15308 | 0.01029 | 0.72082 | 0.72082 | 0.008644
2019-03-29 11:40 | 1553874000.867 | 281 | 0.20744 | 0.002031 | 9.7e-05 | 99.9466 | 97.0196 | 118.981 | 21.0868 | 0.002771 | 0.002567 | 163.5822 | 0.25626 | 8.4026 | 0.002316 | 0.003983 | 0.003825 | 5.8e-05 | 0.000846 | 0.30287 | 1.293 | 2.5473 | 100.3 | 21.06 | 31.656 | 88.3126 | 77.3656 | 83.4064 | 0.007887 | 0.99015 | 0.023511 | 23.7447 | 0.26586 | 0.15245 | 0.010257 | 0.71873 | 0.71873 | 0.008633
2019-03-29 11:40 | 1553874001.8674 | 282 | 0.20744 | 0.002027 | 0.0001 | 99.9652 | 96.9367 | 118.8414 | 21.0869 | 0.002806 | 0.002582 | 163.5243 | 0.25626 | 8.4022 | 0.002387 | 0.003861 | 0.003844 | 6.1e-05 | 0.000844 | 0.31116 | 1.3771 | 2.5472 | 100.3 | 21.06 | 31.653 | 89.1588 | 79.0822 | 83.7684 | 0.007887 | 0.98915 | 0.023485 | 23.7428 | 0.27517 | 0.15245 | 0.01027 | 0.71875 | 0.71875 | 0.008632
2019-03-29 11:40 | 1553874002.867 | 283 | 0.20743 | 0.002021 | 0.000101 | 99.9652 | 97.2169 | 118.9388 | 21.0876 | 0.002903 | 0.002594 | 163.4531 | 0.25512 | 8.4034 | 0.002328 | 0.003935 | 0.003853 | 5.8e-05 | 0.000844 | 0.28296 | 1.3549 | 2.5474 | 100.3 | 21.06 | 31.649 | 91.5 | 77.6717 | 84.0554 | 0.007886 | 0.98835 | 0.023485 | 23.7621 | 0.24935 | 0.15177 | 0.010365 | 0.72535 | 0.72535 | 0.008516
2019-03-29 11:40 | 1553874003.8661 | 284 | 0.20741 | 0.002023 | 0.000102 | 99.9583 | 97.335 | 118.2499 | 21.0873 | 0.002822 | 0.002639 | 163.3725 | 0.25502 | 8.4029 | 0.002203 | 0.003945 | 0.003876 | 6e-05 | 0.000842 | 0.30469 | 1.3289 | 2.5473 | 100.3 | 21.06 | 31.647 | 89.5319 | 74.6635 | 85.1256 | 0.007885 | 0.9854 | 0.023382 | 23.7286 | 0.23788 | 0.15171 | 0.010497 | 0.73134 | 0.73134 | 0.008443
2019-03-29 11:40 | 1553874004.8661 | 285 | 0.20737 | 0.002022 | 0.000103 | 99.9939 | 97.2603 | 117.5023 | 21.087 | 0.002827 | 0.002623 | 163.2828 | 0.25586 | 8.4032 | 0.002312 | 0.003481 | 0.003351 | 5.9e-05 | 0.000837 | 0.2677 | 1.3484 | 2.5474 | 100.3 | 21.06 | 31.645 | 89.6629 | 77.2731 | 84.7453 | 0.007885 | 0.98645 | 0.023321 | 23.641 | 0.24793 | 0.15221 | 0.010713 | 0.74452 | 0.74452 | 0.008321
2019-03-29 11:40 | 1553874005.8653 | 286 | 0.20716 | 0.002027 | 0.000104 | 100.0107 | 97.5456 | 116.7688 | 21.0872 | 0.002863 | 0.002669 | 163.1954 | 0.25549 | 8.4033 | 0.002304 | 0.003843 | 0.003761 | 6.1e-05 | 0.000836 | 0.27345 | 1.3276 | 2.5474 | 100.3 | 21.06 | 31.64 | 90.5241 | 77.0829 | 85.8599 | 0.007883 | 0.98339 | 0.023212 | 23.6037 | 0.22323 | 0.15199 | 0.011962 | 0.82796 | 0.82796 | 0.007471
2019-03-29 11:40 | 1553874006.8656 | 287 | 0.20678 | 0.002056 | 0.000106 | 99.908 | 98.0776 | 116.384 | 21.0882 | 0.002939 | 0.002636 | 163.0905 | 0.25546 | 8.4029 | 0.002265 | 0.003681 | 0.003628 | 5.7e-05 | 0.000834 | 0.27472 | 1.3239 | 1.3506 | 100.3 | 21.06 | 31.638 | 92.3678 | 76.1383 | 85.0617 | 0.007883 | 0.98558 | 0.023199 | 23.5386 | 0.16539 | 0.15197 | 0.014231 | 0.98525 | 0.98525 | 0.006278
2019-03-29 11:40 | 1553874007.8662 | 288 | 0.20597 | 0.002138 | 0.000108 | 99.9847 | 98.5858 | 115.8057 | 21.0875 | 0.002871 | 0.002718 | 162.9951 | 0.25627 | 8.4019 | 0.002251 | 0.003601 | 0.003503 | 5.5e-05 | 0.00083 | 0.26432 | 1.3226 | 2.5474 | 100.3 | 21.06 | 31.637 | 90.7308 | 75.8031 | 87.0389 | 0.007883 | 0.98017 | 0.023078 | 23.5447 | 0.12602 | 0.15246 | 0.018979 | 1.3084 | 1.3084 | 0.004743
2019-03-29 11:40 | 1553874008.8653 | 289 | 0.2052 | 0.002276 | 0.000112 | 99.9829 | 98.638 | 115.0494 | 21.0868 | 0.00293 | 0.0027 | 162.8844 | 0.25536 | 8.4007 | 0.002319 | 0.003434 | 0.003337 | 5.8e-05 | 0.000836 | 0.27649 | 1.3053 | 2.5474 | 100.3 | 21.06 | 31.639 | 92.1319 | 77.4451 | 86.6002 | 0.007883 | 0.98136 | 0.023016 | 23.4534 | 0.12112 | 0.15191 | 0.023435 | 1.6118 | 1.6118 | 0.003836
2019-03-29 11:40 | 1553874009.8653 | 290 | 0.20455 | 0.00268 | 0.000114 | 99.9802 | 98.0534 | 114.2459 | 21.0868 | 0.002888 | 0.002715 | 162.7606 | 0.25608 | 8.4013 | 0.002361 | 0.003675 | 0.003656 | 5.6e-05 | 0.000842 | 0.31349 | 1.3916 | 2.5475 | 100.3 | 21.06 | 31.641 | 91.1311 | 78.442 | 86.9617 | 0.007884 | 0.98038 | 0.022924 | 23.3831 | 0.17406 | 0.15234 | 0.026828 | 1.8379 | 1.8379 | 0.003374
2019-03-29 11:40 | 1553874010.8663 | 291 | 0.204 | 0.003163 | 0.000117 | 99.9236 | 97.7688 | 113.7627 | 21.0864 | 0.002878 | 0.00274 | 162.6627 | 0.25592 | 8.402 | 0.002319 | 0.004214 | 0.004051 | 5.8e-05 | 0.000843 | 0.26985 | 1.3601 | 2.5475 | 100.3 | 21.06 | 31.647 | 90.8946 | 77.4524 | 87.569 | 0.007885 | 0.97873 | 0.022857 | 23.3533 | 0.195 | 0.15224 | 0.029496 | 2.0146 | 2.0146 | 0.003076
2019-03-29 11:40 | 1553874011.8663 | 292 | 0.20363 | 0.003591 | 0.000121 | 99.9826 | 97.1732 | 113.0747 | 21.087 | 0.002918 | 0.002724 | 162.5483 | 0.25493 | 8.403 | 0.002312 | 0.004114 | 0.004087 | 5.6e-05 | 0.00085 | 0.28102 | 1.3526 | 2.5474 | 100.3 | 21.06 | 31.649 | 91.8551 | 77.2658 | 87.189 | 0.007886 | 0.97976 | 0.022799 | 23.2703 | 0.25492 | 0.15166 | 0.031173 | 2.1236 | 2.1236 | 0.002907
2019-03-29 11:40 | 1553874012.8664 | 293 | 0.20337 | 0.00391 | 0.000125 | 99.9014 | 96.9555 | 112.3517 | 21.0868 | 0.00294 | 0.002771 | 162.4306 | 0.25623 | 8.4031 | 0.002369 | 0.003982 | 0.003876 | 5.4e-05 | 0.000851 | 0.27096 | 1.3445 | 2.5475 | 100.3 | 21.06 | 31.652 | 92.3706 | 78.6487 | 88.3213 | 0.007886 | 0.97669 | 0.022691 | 23.2322 | 0.26773 | 0.15243 | 0.032377 | 2.1949 | 2.1949 | 0.002827
2019-03-29 11:40 | 1553874013.8674 | 294 | 0.20319 | 0.004243 | 0.000129 | 99.9802 | 96.6514 | 112.2461 | 21.0872 | 0.003017 | 0.002758 | 162.3365 | 0.2558 | 8.4007 | 0.002379 | 0.004258 | 0.004248 | 5.8e-05 | 0.000847 | 0.26698 | 1.3625 | 2.5475 | 100.3 | 21.05 | 31.661 | 94.2462 | 78.8922 | 88.0054 | 0.007884 | 0.97754 | 0.02269 | 23.2111 | 0.30287 | 0.15218 | 0.032993 | 2.2364 | 2.2364 | 0.002769
2019-03-29 11:40 | 1553874014.867 | 295 | 0.20304 | 0.004442 | 0.000132 | 99.9517 | 96.1238 | 111.9534 | 21.0878 | 0.003009 | 0.002833 | 162.2485 | 0.25505 | 8.4026 | 0.002362 | 0.004265 | 0.004086 | 6.1e-05 | 0.000841 | 0.26272 | 1.3278 | 2.5475 | 100.3 | 21.05 | 31.665 | 94.0541 | 78.4665 | 89.8033 | 0.007885 | 0.9727 | 0.022604 | 23.2385 | 0.34929 | 0.15173 | 0.033672 | 2.2736 | 2.2736 | 0.002716
2019-03-29 11:40 | 1553874015.8664 | 296 | 0.20294 | 0.004635 | 0.000137 | 99.9254 | 95.9886 | 111.6135 | 21.0872 | 0.002998 | 0.00287 | 162.1373 | 0.25564 | 8.4021 | 0.002337 | 0.003806 | 0.003689 | 5.8e-05 | 0.000848 | 0.30807 | 1.2826 | 2.5473 | 100.3 | 21.05 | 31.676 | 93.7711 | 77.8665 | 90.6961 | 0.007888 | 0.97032 | 0.022542 | 23.2317 | 0.35952 | 0.15208 | 0.034052 | 2.2926 | 2.2926 | 0.0027
2019-03-29 11:40 | 1553874016.8669 | 297 | 0.2028 | 0.004812 | 0.000143 | 99.8691 | 95.5552 | 111.1713 | 21.0875 | 0.002943 | 0.00283 | 162.0345 | 0.25614 | 8.4006 | 0.002324 | 0.004032 | 0.003949 | 5.8e-05 | 0.000854 | 0.28208 | 1.3309 | 2.5473 | 100.3 | 21.05 | 31.682 | 92.4594 | 77.5618 | 89.7344 | 0.007889 | 0.97289 | 0.022527 | 23.155 | 0.39496 | 0.15238 | 0.034643 | 2.3305 | 2.3305 | 0.002661
2019-03-29 11:40 | 1553874017.8659 | 298 | 0.20268 | 0.004978 | 0.000148 | 99.8877 | 95.4487 | 111.0828 | 21.0877 | 0.003087 | 0.002835 | 161.9236 | 0.25646 | 8.4029 | 0.002352 | 0.004071 | 0.004005 | 5.6e-05 | 0.00086 | 0.28495 | 1.3689 | 2.5474 | 100.3 | 21.05 | 31.689 | 95.9235 | 78.2403 | 89.8571 | 0.007891 | 0.97256 | 0.022514 | 23.1497 | 0.40659 | 0.15257 | 0.035168 | 2.3642 | 2.3642 | 0.002626
2019-03-29 11:40 | 1553874018.8666 | 299 | 0.20248 | 0.005134 | 0.000153 | 99.9029 | 95.3357 | 111.4213 | 21.0884 | 0.003125 | 0.002899 | 161.8153 | 0.25633 | 8.4014 | 0.00244 | 0.003745 | 0.003617 | 5.8e-05 | 0.000855 | 0.27184 | 1.3488 | 2.5474 | 100.3 | 21.05 | 31.701 | 96.8476 | 80.3498 | 91.3938 | 0.007894 | 0.96846 | 0.022501 | 23.2339 | 0.41856 | 0.15249 | 0.036172 | 2.4299 | 2.4299 | 0.002554
2019-03-29 11:40 | 1553874019.8655 | 300 | 0.20217 | 0.005271 | 0.000158 | 99.8667 | 94.9467 | 111.9619 | 21.0879 | 0.003178 | 0.002891 | 161.7239 | 0.2562 | 8.4002 | 0.002424 | 0.004231 | 0.004187 | 5.6e-05 | 0.000855 | 0.27505 | 1.361 | 2.5475 | 100.3 | 21.05 | 31.706 | 98.1191 | 79.979 | 91.201 | 0.007895 | 0.96897 | 0.022562 | 23.2841 | 0.45189 | 0.15241 | 0.037823 | 2.5473 | 2.5473 | 0.002435
2019-03-29 11:40 | 1553874020.8647 | 301 | 0.20197 | 0.005416 | 0.000165 | 99.891 | 94.3378 | 112.282 | 21.0882 | 0.003152 | 0.002945 | 161.612 | 0.25571 | 8.4016 | 0.002434 | 0.004076 | 0.003945 | 5.6e-05 | 0.000851 | 0.31526 | 1.3523 | 2.5475 | 100.3 | 21.05 | 31.713 | 97.4884 | 80.1961 | 92.5117 | 0.007897 | 0.9655 | 0.022553 | 23.3592 | 0.5116 | 0.15212 | 0.03887 | 2.6165 | 2.6165 | 0.002366
2019-03-29 11:40 | 1553874021.8661 | 302 | 0.20183 | 0.005562 | 0.00017 | 99.8757 | 94.4747 | 112.4355 | 21.0885 | 0.003127 | 0.002909 | 161.5037 | 0.25518 | 8.4034 | 0.002383 | 0.004194 | 0.004085 | 5.7e-05 | 0.000843 | 0.30093 | 1.3071 | 2.5473 | 100.3 | 21.05 | 31.716 | 96.8862 | 78.9793 | 91.6253 | 0.007897 | 0.96784 | 0.022596 | 23.3469 | 0.49727 | 0.15181 | 0.039464 | 2.6612 | 2.6612 | 0.002322
2019-03-29 11:40 | 1553874022.8657 | 303 | 0.20186 | 0.005711 | 0.000175 | 99.8838 | 94.7139 | 112.8855 | 21.0881 | 0.003201 | 0.002968 | 161.4079 | 0.25593 | 8.4047 | 0.002442 | 0.003947 | 0.003828 | 5.7e-05 | 0.000835 | 0.29479 | 1.3278 | 2.5473 | 100.3 | 21.05 | 31.718 | 98.6754 | 80.3947 | 93.0552 | 0.007898 | 0.96406 | 0.022597 | 23.4393 | 0.47537 | 0.15225 | 0.039144 | 2.6395 | 2.6395 | 0.002348
2019-03-29 11:40 | 1553874023.8657 | 304 | 0.20186 | 0.00585 | 0.000178 | 99.8317 | 94.8682 | 112.989 | 21.0885 | 0.00317 | 0.002953 | 161.2866 | 0.25645 | 8.4036 | 0.002472 | 0.003495 | 0.00339 | 5.7e-05 | 0.000841 | 0.30934 | 1.3195 | 2.5474 | 100.3 | 21.05 | 31.723 | 97.9403 | 81.1313 | 92.6873 | 0.007899 | 0.96503 | 0.022619 | 23.4383 | 0.45615 | 0.15256 | 0.03893 | 2.6274 | 2.6274 | 0.002363
2019-03-29 11:40 | 1553874024.8647 | 305 | 0.20185 | 0.00597 | 0.000179 | 99.8586 | 94.5347 | 113.1791 | 21.0878 | 0.003155 | 0.002929 | 161.1794 | 0.25531 | 8.4007 | 0.00247 | 0.003613 | 0.003672 | 5.8e-05 | 0.000838 | 0.28799 | 1.2911 | 2.5472 | 100.3 | 21.05 | 31.725 | 97.5595 | 81.0666 | 92.1102 | 0.0079 | 0.96656 | 0.022656 | 23.4395 | 0.49006 | 0.15188 | 0.038889 | 2.6289 | 2.6289 | 0.002351
2019-03-29 11:40 | 1553874025.8662 | 306 | 0.20198 | 0.006063 | 0.000179 | 99.8197 | 94.4064 | 113.3453 | 21.0881 | 0.003207 | 0.002973 | 161.0648 | 0.25593 | 8.4014 | 0.00248 | 0.003929 | 0.003706 | 5.8e-05 | 0.000846 | 0.27505 | 1.3412 | 2.5474 | 100.3 | 21.05 | 31.728 | 98.8155 | 81.315 | 93.1845 | 0.0079 | 0.96372 | 0.022639 | 23.4912 | 0.49872 | 0.15225 | 0.037994 | 2.5665 | 2.5665 | 0.002414
2019-03-29 11:40 | 1553874026.8648 | 307 | 0.20215 | 0.006131 | 0.000179 | 99.8467 | 94.6508 | 113.5444 | 21.0891 | 0.003209 | 0.003003 | 160.9656 | 0.25571 | 8.4013 | 0.002489 | 0.004083 | 0.003925 | 6.2e-05 | 0.000852 | 0.27721 | 1.2993 | 2.5473 | 100.3 | 21.04 | 31.733 | 98.8738 | 81.538 | 93.8966 | 0.007897 | 0.96185 | 0.022637 | 23.5346 | 0.47801 | 0.15212 | 0.036855 | 2.4893 | 2.4893 | 0.002487
2019-03-29 11:40 | 1553874027.8648 | 308 | 0.20228 | 0.006156 | 0.000179 | 99.9167 | 95.1456 | 113.7467 | 21.0882 | 0.003207 | 0.002986 | 160.856 | 0.25611 | 8.4036 | 0.002571 | 0.003507 | 0.003465 | 6.3e-05 | 0.000858 | 0.27489 | 1.3122 | 2.5473 | 100.3 | 21.04 | 31.734 | 98.8275 | 83.4918 | 93.492 | 0.007897 | 0.96292 | 0.02267 | 23.5426 | 0.43764 | 0.15236 | 0.03605 | 2.4384 | 2.4384 | 0.002543
2019-03-29 11:40 | 1553874028.8649 | 309 | 0.20224 | 0.006157 | 0.000178 | 99.8966 | 95.5896 | 113.5148 | 21.0893 | 0.003251 | 0.003015 | 160.7618 | 0.25573 | 8.4056 | 0.002506 | 0.004079 | 0.003936 | 6e-05 | 0.000859 | 0.31188 | 1.3545 | 2.5474 | 100.3 | 21.04 | 31.739 | 99.8961 | 81.943 | 94.1824 | 0.007898 | 0.96111 | 0.022625 | 23.5407 | 0.3942 | 0.15214 | 0.036302 | 2.4507 | 2.4507 | 0.002527
2019-03-29 11:40 | 1553874029.8659 | 310 | 0.20242 | 0.006136 | 0.000178 | 99.8769 | 95.5893 | 113.2723 | 21.0893 | 0.003225 | 0.002958 | 160.6579 | 0.25601 | 8.4018 | 0.002508 | 0.003862 | 0.003792 | 6.3e-05 | 0.00086 | 0.24762 | 1.3515 | 2.5475 | 100.3 | 21.04 | 31.74 | 99.2733 | 81.9906 | 92.8131 | 0.007899 | 0.9647 | 0.022643 | 23.4717 | 0.39246 | 0.1523 | 0.035262 | 2.3824 | 2.3824 | 0.002602
2019-03-29 11:40 | 1553874030.8649 | 311 | 0.2025 | 0.006098 | 0.000178 | 99.9164 | 95.9989 | 113.2841 | 21.089 | 0.003201 | 0.003011 | 160.5618 | 0.25601 | 8.4021 | 0.002452 | 0.004373 | 0.004128 | 6e-05 | 0.000865 | 0.2521 | 1.3389 | 2.5473 | 100.3 | 21.04 | 31.742 | 98.6941 | 80.6381 | 94.0883 | 0.007899 | 0.96135 | 0.022605 | 23.5138 | 0.35776 | 0.1523 | 0.034836 | 2.3496 | 2.3496 | 0.002638
2019-03-29 11:40 | 1553874031.866 | 312 | 0.20256 | 0.006052 | 0.000178 | 99.9314 | 96.3511 | 113.4289 | 21.0905 | 0.003262 | 0.002933 | 160.4903 | 0.25581 | 8.4026 | 0.002568 | 0.004227 | 0.004219 | 5.9e-05 | 0.000867 | 0.26305 | 1.3505 | 2.5474 | 100.3 | 21.04 | 31.745 | 100.1476 | 83.4215 | 92.2254 | 0.0079 | 0.96625 | 0.022677 | 23.469 | 0.32634 | 0.15218 | 0.034528 | 2.3362 | 2.3362 | 0.002651
2019-03-29 11:41 | 1553874032.8658 | 313 | 0.20264 | 0.006001 | 0.000181 | 99.9544 | 96.3159 | 113.4854 | 21.0887 | 0.003227 | 0.003037 | 160.3796 | 0.2566 | 8.4034 | 0.002573 | 0.003818 | 0.003687 | 6e-05 | 0.000869 | 0.28744 | 1.4002 | 2.5475 | 100.3 | 21.04 | 31.747 | 99.3095 | 83.5556 | 94.7328 | 0.0079 | 0.95967 | 0.022605 | 23.5553 | 0.33167 | 0.15265 | 0.034085 | 2.2988 | 2.2988 | 0.002703
2019-03-29 11:41 | 1553874033.865 | 314 | 0.20261 | 0.005952 | 0.000183 | 99.9377 | 96.4913 | 113.4908 | 21.0885 | 0.003254 | 0.003026 | 160.2952 | 0.25581 | 8.4005 | 0.002486 | 0.003866 | 0.003708 | 5.9e-05 | 0.000876 | 0.27848 | 1.351 | 2.5475 | 100.3 | 21.04 | 31.755 | 99.9549 | 81.4541 | 94.4463 | 0.007902 | 0.96042 | 0.022615 | 23.5467 | 0.31389 | 0.15218 | 0.03437 | 2.3189 | 2.3189 | 0.002671
2019-03-29 11:41 | 1553874034.8651 | 315 | 0.20261 | 0.005907 | 0.000186 | 99.9326 | 96.705 | 113.3392 | 21.0892 | 0.00332 | 0.003041 | 160.1888 | 0.25569 | 8.4031 | 0.002518 | 0.004328 | 0.004257 | 6.2e-05 | 0.000884 | 0.2808 | 1.3393 | 2.5474 | 100.3 | 21.04 | 31.759 | 101.563 | 82.2355 | 94.8201 | 0.007903 | 0.95944 | 0.022588 | 23.5429 | 0.29366 | 0.15211 | 0.034356 | 2.315 | 2.315 | 0.002674
2019-03-29 11:41 | 1553874035.8651 | 316 | 0.20261 | 0.00586 | 0.000188 | 99.9386 | 96.778 | 113.0735 | 21.0899 | 0.003302 | 0.003038 | 160.1068 | 0.25557 | 8.4017 | 0.002482 | 0.004344 | 0.004301 | 6e-05 | 0.000882 | 0.29048 | 1.327 | 2.5474 | 100.3 | 21.04 | 31.763 | 101.1209 | 81.3627 | 94.7466 | 0.007904 | 0.95963 | 0.022564 | 23.513 | 0.28744 | 0.15204 | 0.034415 | 2.3164 | 2.3164 | 0.002671
2019-03-29 11:41 | 1553874036.8656 | 317 | 0.20255 | 0.005836 | 0.000191 | 99.8826 | 96.8295 | 113.0709 | 21.0898 | 0.00329 | 0.003049 | 159.9897 | 0.25577 | 8.4009 | 0.002517 | 0.003836 | 0.003747 | 6.2e-05 | 0.000883 | 0.27113 | 1.296 | 2.5474 | 100.3 | 21.04 | 31.772 | 100.8382 | 82.193 | 95.0079 | 0.007907 | 0.95895 | 0.022556 | 23.521 | 0.27767 | 0.15216 | 0.034834 | 2.3435 | 2.3435 | 0.002643
2019-03-29 11:41 | 1553874037.8646 | 318 | 0.20242 | 0.005822 | 0.000194 | 99.9146 | 96.7879 | 112.4836 | 21.0889 | 0.003329 | 0.003027 | 159.8828 | 0.25608 | 8.4006 | 0.002541 | 0.004053 | 0.003881 | 5.8e-05 | 0.00088 | 0.27632 | 1.3055 | 2.5473 | 100.3 | 21.04 | 31.776 | 101.774 | 82.7742 | 94.4858 | 0.007908 | 0.96031 | 0.022513 | 23.4432 | 0.28438 | 0.15234 | 0.035605 | 2.3908 | 2.3908 | 0.002593
2019-03-29 11:41 | 1553874038.8641 | 319 | 0.2023 | 0.005822 | 0.000196 | 99.9481 | 96.7354 | 112.2238 | 21.0908 | 0.003244 | 0.003044 | 159.7834 | 0.25552 | 8.404 | 0.002486 | 0.004492 | 0.004396 | 5.9e-05 | 0.00087 | 0.27831 | 1.2939 | 2.5474 | 100.3 | 21.04 | 31.785 | 99.7264 | 81.4664 | 94.8844 | 0.00791 | 0.95927 | 0.022475 | 23.4288 | 0.29224 | 0.15201 | 0.036349 | 2.4365 | 2.4365 | 0.002539
2019-03-29 11:41 | 1553874039.8652 | 320 | 0.20218 | 0.005839 | 0.000197 | 99.9191 | 97.0074 | 111.6689 | 21.0899 | 0.003306 | 0.003044 | 159.6901 | 0.25661 | 8.4029 | 0.002481 | 0.004059 | 0.003973 | 5.8e-05 | 0.00086 | 0.29927 | 1.3452 | 2.5474 | 100.3 | 21.04 | 31.79 | 101.2325 | 81.3293 | 94.9031 | 0.007911 | 0.95922 | 0.022418 | 23.3714 | 0.26453 | 0.15266 | 0.037033 | 2.4761 | 2.4761 | 0.002509
2019-03-29 11:41 | 1553874040.8654 | 321 | 0.20211 | 0.005878 | 0.000199 | 99.9784 | 97.1783 | 110.9366 | 21.0904 | 0.003374 | 0.003115 | 159.5751 | 0.25533 | 8.403 | 0.002556 | 0.003442 | 0.003351 | 6e-05 | 0.000858 | 0.2688 | 1.328 | 2.5474 | 100.3 | 21.04 | 31.795 | 102.8664 | 83.1531 | 96.6014 | 0.007912 | 0.95482 | 0.022293 | 23.3483 | 0.25408 | 0.15189 | 0.037385 | 2.4856 | 2.4856 | 0.002487
2019-03-29 11:41 | 1553874041.8643 | 322 | 0.202 | 0.005921 | 0.0002 | 99.9412 | 97.1859 | 110.6669 | 21.0902 | 0.003428 | 0.00312 | 159.4373 | 0.25605 | 8.4026 | 0.002544 | 0.003662 | 0.00359 | 6.1e-05 | 0.000856 | 0.2959 | 1.313 | 2.5474 | 100.3 | 21.03 | 31.804 | 104.181 | 82.8571 | 96.7209 | 0.00791 | 0.95451 | 0.022263 | 23.3237 | 0.25006 | 0.15232 | 0.038034 | 2.5253 | 2.5253 | 0.002455
2019-03-29 11:41 | 1553874042.866 | 323 | 0.20186 | 0.005984 | 0.0002 | 100.0074 | 97.4207 | 110.2289 | 21.0909 | 0.003382 | 0.003142 | 159.3368 | 0.25609 | 8.4024 | 0.002485 | 0.003648 | 0.003506 | 5.7e-05 | 0.000853 | 0.30342 | 1.3152 | 2.5473 | 100.3 | 21.03 | 31.809 | 103.0737 | 81.4446 | 97.2648 | 0.007911 | 0.95311 | 0.022202 | 23.2946 | 0.2344 | 0.15235 | 0.038766 | 2.5669 | 2.5669 | 0.002416
2019-03-29 11:41 | 1553874043.8654 | 324 | 0.20172 | 0.006054 | 0.000198 | 100.0855 | 97.2374 | 109.8808 | 21.0901 | 0.003366 | 0.00314 | 159.2267 | 0.25547 | 8.4021 | 0.002438 | 0.00426 | 0.004086 | 5.9e-05 | 0.000856 | 0.2896 | 1.3089 | 2.5473 | 100.3 | 21.03 | 31.815 | 102.6673 | 80.3025 | 97.2099 | 0.007912 | 0.95325 | 0.022169 | 23.2561 | 0.25823 | 0.15198 | 0.039541 | 2.6144 | 2.6144 | 0.002366
2019-03-29 11:41 | 1553874044.8647 | 325 | 0.20149 | 0.006116 | 0.000197 | 100.0152 | 97.3292 | 109.659 | 21.0912 | 0.003413 | 0.003182 | 159.1104 | 0.25566 | 8.4049 | 0.002436 | 0.003939 | 0.003808 | 6e-05 | 0.000857 | 0.31238 | 1.3363 | 2.5473 | 100.3 | 21.03 | 31.822 | 103.8237 | 80.2528 | 98.2275 | 0.007914 | 0.95064 | 0.022116 | 23.2645 | 0.2435 | 0.15209 | 0.040824 | 2.6929 | 2.6929 | 0.002299
2019-03-29 11:41 | 1553874045.8644 | 326 | 0.2015 | 0.006194 | 0.000195 | 100.0313 | 97.4708 | 109.5815 | 21.0903 | 0.00342 | 0.003117 | 158.9993 | 0.25604 | 8.4021 | 0.002562 | 0.003946 | 0.003926 | 6.1e-05 | 0.000854 | 0.28744 | 1.2896 | 2.5473 | 100.3 | 21.03 | 31.825 | 103.9738 | 83.2927 | 96.6467 | 0.007915 | 0.9547 | 0.022155 | 23.2067 | 0.23194 | 0.15232 | 0.040686 | 2.6886 | 2.6886 | 0.002306
2019-03-29 11:41 | 1553874046.8654 | 327 | 0.20142 | 0.006285 | 0.000193 | 100.0613 | 97.677 | 109.3834 | 21.0911 | 0.00346 | 0.003221 | 158.8983 | 0.25618 | 8.4019 | 0.002602 | 0.004181 | 0.004052 | 6.2e-05 | 0.000856 | 0.29822 | 1.3611 | 2.5474 | 100.3 | 21.03 | 31.83 | 104.9646 | 84.2469 | 99.1794 | 0.007916 | 0.94821 | 0.02206 | 23.265 | 0.21572 | 0.1524 | 0.041045 | 2.7007 | 2.7007 | 0.002297
2019-03-29 11:41 | 1553874047.8645 | 328 | 0.20129 | 0.006367 | 0.000192 | 100.1592 | 97.9862 | 109.4888 | 21.0906 | 0.003412 | 0.003208 | 158.793 | 0.25585 | 8.4 | 0.00257 | 0.003874 | 0.003778 | 6.1e-05 | 0.00086 | 0.30453 | 1.3613 | 2.5474 | 100.3 | 21.03 | 31.832 | 103.7979 | 83.4699 | 98.8627 | 0.007917 | 0.94901 | 0.02208 | 23.2663 | 0.19619 | 0.15221 | 0.041694 | 2.746 | 2.746 | 0.002256
2019-03-29 11:41 | 1553874048.8647 | 329 | 0.20112 | 0.006452 | 0.000193 | 100.1262 | 97.9421 | 109.843 | 21.0913 | 0.003463 | 0.003216 | 158.6552 | 0.25551 | 8.4023 | 0.002549 | 0.00428 | 0.004213 | 6.2e-05 | 0.000872 | 0.30646 | 1.3299 | 2.5475 | 100.3 | 21.03 | 31.835 | 105.0377 | 82.9746 | 99.0367 | 0.007917 | 0.94857 | 0.022111 | 23.3094 | 0.19727 | 0.152 | 0.042623 | 2.8109 | 2.8109 | 0.002201
2019-03-29 11:41 | 1553874049.8646 | 330 | 0.20092 | 0.006533 | 0.000196 | 100.1047 | 97.9916 | 110.2674 | 21.0915 | 0.003481 | 0.003215 | 158.517 | 0.25631 | 8.4034 | 0.002665 | 0.004365 | 0.004293 | 6.3e-05 | 0.000878 | 0.29015 | 1.319 | 2.5474 | 100.3 | 21.03 | 31.837 | 105.4595 | 85.7611 | 99.0361 | 0.007918 | 0.94857 | 0.022153 | 23.3543 | 0.19083 | 0.15248 | 0.043703 | 2.8875 | 2.8875 | 0.002149
2019-03-29 11:41 | 1553874050.8646 | 331 | 0.20072 | 0.00662 | 0.0002 | 100.0894 | 97.8732 | 110.4444 | 21.0915 | 0.003555 | 0.003256 | 158.4128 | 0.25604 | 8.4012 | 0.002674 | 0.003765 | 0.003663 | 6.4e-05 | 0.000875 | 0.26576 | 1.3602 | 2.5474 | 100.3 | 21.03 | 31.839 | 107.2488 | 85.9743 | 100.0175 | 0.007918 | 0.94608 | 0.022142 | 23.4039 | 0.20028 | 0.15232 | 0.044817 | 2.9593 | 2.9593 | 0.002095
2019-03-29 11:41 | 1553874051.8636 | 332 | 0.20071 | 0.006718 | 0.000204 | 100.1274 | 97.9759 | 110.0871 | 21.0921 | 0.003487 | 0.003237 | 158.3174 | 0.25637 | 8.4013 | 0.002602 | 0.004177 | 0.00409 | 6.1e-05 | 0.000879 | 0.27721 | 1.3355 | 2.5475 | 100.3 | 21.03 | 31.842 | 105.6168 | 84.2384 | 99.5647 | 0.007919 | 0.94723 | 0.022119 | 23.3518 | 0.19429 | 0.15252 | 0.044734 | 2.9506 | 2.9506 | 0.002104
2019-03-29 11:41 | 1553874052.8636 | 333 | 0.20068 | 0.006827 | 0.000207 | 100.1924 | 97.7869 | 109.9376 | 21.0922 | 0.003504 | 0.003284 | 158.1944 | 0.25584 | 8.402 | 0.002663 | 0.004387 | 0.004378 | 6e-05 | 0.000875 | 0.27793 | 1.352 | 2.5475 | 100.3 | 21.03 | 31.843 | 106.0299 | 85.7141 | 100.7041 | 0.007919 | 0.94434 | 0.022071 | 23.3716 | 0.21737 | 0.1522 | 0.044762 | 2.9458 | 2.9458 | 0.002103
2019-03-29 11:41 | 1553874053.8651 | 334 | 0.20073 | 0.006929 | 0.000209 | 100.1343 | 97.4538 | 109.678 | 21.0916 | 0.003572 | 0.003257 | 158.0499 | 0.25512 | 8.402 | 0.002597 | 0.003792 | 0.00362 | 5.9e-05 | 0.000881 | 0.26969 | 1.3446 | 2.5474 | 100.3 | 21.03 | 31.843 | 107.6646 | 84.1264 | 100.0333 | 0.007919 | 0.94604 | 0.022064 | 23.323 | 0.24271 | 0.15177 | 0.04436 | 2.9184 | 2.9184 | 0.002117
2019-03-29 11:41 | 1553874054.8638 | 335 | 0.20081 | 0.00702 | 0.000211 | 100.1945 | 97.1469 | 109.543 | 21.0922 | 0.003555 | 0.003228 | 157.9328 | 0.2552 | 8.4028 | 0.002677 | 0.003809 | 0.003704 | 5.6e-05 | 0.000879 | 0.27721 | 1.3589 | 2.5474 | 100.3 | 21.02 | 31.843 | 107.2533 | 86.0592 | 99.3468 | 0.007914 | 0.94778 | 0.022071 | 23.2872 | 0.27628 | 0.15182 | 0.043741 | 2.8786 | 2.8786 | 0.002147
2019-03-29 11:41 | 1553874055.8637 | 336 | 0.20083 | 0.007081 | 0.000211 | 100.1642 | 97.3896 | 109.6261 | 21.0922 | 0.003557 | 0.003256 | 157.8015 | 0.25656 | 8.4021 | 0.002585 | 0.004299 | 0.0042 | 6e-05 | 0.000883 | 0.27936 | 1.3054 | 2.5473 | 100.3 | 21.03 | 31.842 | 107.3031 | 83.8312 | 100.0064 | 0.007919 | 0.94611 | 0.02206 | 23.3167 | 0.25127 | 0.15263 | 0.043565 | 2.8655 | 2.8655 | 0.002168
2019-03-29 11:41 | 1553874056.8651 | 337 | 0.20081 | 0.00712 | 0.00021 | 100.2858 | 97.4374 | 109.7041 | 21.0925 | 0.003598 | 0.003262 | 157.6628 | 0.25566 | 8.4036 | 0.002664 | 0.003933 | 0.003787 | 6.4e-05 | 0.000881 | 0.26665 | 1.3147 | 2.5473 | 100.3 | 21.03 | 31.842 | 108.3131 | 85.7363 | 100.1653 | 0.007919 | 0.9457 | 0.022063 | 23.3299 | 0.25773 | 0.15209 | 0.043647 | 2.8713 | 2.8713 | 0.002156
2019-03-29 11:41 | 1553874057.8648 | 338 | 0.20081 | 0.007149 | 0.000209 | 100.2154 | 97.535 | 109.7465 | 21.092 | 0.003526 | 0.003295 | 157.5493 | 0.25611 | 8.4015 | 0.002612 | 0.004409 | 0.004342 | 6.5e-05 | 0.000871 | 0.27201 | 1.3397 | 2.5475 | 100.3 | 21.03 | 31.841 | 106.5642 | 84.4943 | 100.9551 | 0.007919 | 0.94371 | 0.022044 | 23.3591 | 0.2425 | 0.15236 | 0.043586 | 2.8649 | 2.8649 | 0.002164
2019-03-29 11:41 | 1553874058.8649 | 339 | 0.20071 | 0.007177 | 0.000208 | 100.2412 | 97.6241 | 109.494 | 21.092 | 0.003574 | 0.003301 | 157.4487 | 0.25521 | 8.4035 | 0.002639 | 0.003806 | 0.003757 | 6.2e-05 | 0.000875 | 0.27455 | 1.3544 | 2.5474 | 100.3 | 21.03 | 31.84 | 107.7131 | 85.1489 | 101.1058 | 0.007919 | 0.94333 | 0.022014 | 23.3369 | 0.23663 | 0.15183 | 0.044162 | 2.8988 | 2.8988 | 0.002132
2019-03-29 11:41 | 1553874059.8639 | 340 | 0.20061 | 0.007197 | 0.000207 | 100.2609 | 97.9432 | 109.1172 | 21.0922 | 0.003624 | 0.003326 | 157.3445 | 0.25622 | 8.4032 | 0.002639 | 0.003746 | 0.003616 | 6.3e-05 | 0.000876 | 0.29789 | 1.3931 | 2.5474 | 100.3 | 21.03 | 31.84 | 108.9465 | 85.146 | 101.7206 | 0.007919 | 0.94178 | 0.021958 | 23.3159 | 0.2092 | 0.15243 | 0.044748 | 2.9299 | 2.9299 | 0.002117
2019-03-29 11:41 | 1553874060.8649 | 341 | 0.20062 | 0.007229 | 0.000207 | 100.2184 | 97.8087 | 108.9712 | 21.0934 | 0.003593 | 0.003323 | 157.2137 | 0.2561 | 8.4009 | 0.00261 | 0.003269 | 0.003138 | 6.3e-05 | 0.000871 | 0.29175 | 1.3308 | 2.5476 | 100.3 | 21.02 | 31.839 | 108.1848 | 84.4557 | 101.641 | 0.007913 | 0.94198 | 0.021946 | 23.2978 | 0.2177 | 0.15235 | 0.044658 | 2.9224 | 2.9224 | 0.002122
2019-03-29 11:41 | 1553874061.8639 | 342 | 0.20062 | 0.007269 | 0.000207 | 100.2888 | 97.7978 | 109.1337 | 21.0927 | 0.003641 | 0.003318 | 157.0735 | 0.25555 | 8.4009 | 0.002668 | 0.003892 | 0.003749 | 6.3e-05 | 0.000871 | 0.26753 | 1.3319 | 2.5473 | 100.3 | 21.02 | 31.84 | 109.3513 | 85.8301 | 101.5108 | 0.007914 | 0.94231 | 0.021966 | 23.3111 | 0.22497 | 0.15202 | 0.044599 | 2.9212 | 2.9212 | 0.002118
2019-03-29 11:41 | 1553874062.864 | 343 | 0.20062 | 0.007308 | 0.000206 | 100.2981 | 97.9497 | 108.9792 | 21.0934 | 0.003669 | 0.003319 | 156.9581 | 0.25559 | 8.4039 | 0.002629 | 0.004382 | 0.004305 | 6.5e-05 | 0.000872 | 0.32587 | 1.3208 | 2.5474 | 100.3 | 21.02 | 31.84 | 110.0262 | 84.8937 | 101.5294 | 0.007914 | 0.94226 | 0.02195 | 23.2952 | 0.21192 | 0.15205 | 0.044574 | 2.9175 | 2.9175 | 0.002121
2019-03-29 11:41 | 1553874063.864 | 344 | 0.20066 | 0.007337 | 0.000205 | 100.2771 | 97.9111 | 109.3555 | 21.0941 | 0.003672 | 0.003347 | 156.8502 | 0.25604 | 8.4044 | 0.002655 | 0.004443 | 0.004347 | 6.8e-05 | 0.000874 | 0.29391 | 1.2878 | 2.5473 | 100.3 | 21.02 | 31.842 | 110.1034 | 85.5358 | 102.2085 | 0.007914 | 0.94056 | 0.021968 | 23.3565 | 0.21357 | 0.15232 | 0.044295 | 2.9017 | 2.9017 | 0.002136
2019-03-29 11:41 | 1553874064.865 | 345 | 0.20063 | 0.007353 | 0.000202 | 100.3816 | 97.894 | 109.2675 | 21.093 | 0.00362 | 0.003354 | 156.7279 | 0.25547 | 8.4009 | 0.002592 | 0.003511 | 0.003395 | 6.6e-05 | 0.000873 | 0.28296 | 1.3079 | 2.5474 | 100.3 | 21.02 | 31.843 | 108.8431 | 84.0114 | 102.396 | 0.007914 | 0.94009 | 0.021954 | 23.3529 | 0.22445 | 0.15198 | 0.044476 | 2.9117 | 2.9117 | 0.002124
2019-03-29 11:41 | 1553874065.8641 | 346 | 0.20058 | 0.007359 | 0.0002 | 100.3501 | 97.8418 | 109.3135 | 21.0934 | 0.00367 | 0.003403 | 156.6067 | 0.25587 | 8.4021 | 0.002646 | 0.004195 | 0.004148 | 6.4e-05 | 0.000872 | 0.30038 | 1.3487 | 2.5474 | 100.3 | 21.02 | 31.844 | 110.0476 | 85.3189 | 103.5695 | 0.007915 | 0.93716 | 0.021924 | 23.3943 | 0.22642 | 0.15221 | 0.044734 | 2.9248 | 2.9248 | 0.002118
2019-03-29 11:41 | 1553874066.864 | 347 | 0.20052 | 0.007367 | 0.000196 | 100.3115 | 97.9845 | 109.079 | 21.0934 | 0.003738 | 0.003441 | 156.4931 | 0.25626 | 8.4038 | 0.002752 | 0.004398 | 0.004284 | 6.4e-05 | 0.00087 | 0.30663 | 1.3415 | 2.5473 | 100.3 | 21.02 | 31.847 | 111.7179 | 87.8496 | 104.5019 | 0.007915 | 0.93484 | 0.021874 | 23.3981 | 0.20994 | 0.15245 | 0.045074 | 2.9404 | 2.9404 | 0.00211
2019-03-29 11:41 | 1553874067.8631 | 348 | 0.20043 | 0.007375 | 0.000194 | 100.3157 | 98.1305 | 108.7684 | 21.0941 | 0.00368 | 0.003446 | 156.3664 | 0.25536 | 8.4024 | 0.002759 | 0.004148 | 0.004088 | 6.4e-05 | 0.000864 | 0.30237 | 1.3443 | 2.5476 | 100.3 | 21.02 | 31.849 | 110.2943 | 88.0395 | 104.6259 | 0.007916 | 0.93454 | 0.021839 | 23.3686 | 0.19699 | 0.15192 | 0.045639 | 2.9726 | 2.9726 | 0.00208
2019-03-29 11:41 | 1553874068.8641 | 349 | 0.20034 | 0.007393 | 0.000193 | 100.4522 | 98.5711 | 108.4912 | 21.093 | 0.003709 | 0.003395 | 156.2489 | 0.25631 | 8.4007 | 0.00267 | 0.003658 | 0.003589 | 6.5e-05 | 0.000859 | 0.29518 | 1.3194 | 2.5473 | 100.3 | 21.02 | 31.854 | 111.0064 | 85.892 | 103.3702 | 0.007917 | 0.93765 | 0.021847 | 23.3 | 0.16908 | 0.15248 | 0.046166 | 3.0081 | 3.0081 | 0.002063
2019-03-29 11:41 | 1553874069.8633 | 350 | 0.2002 | 0.007426 | 0.000193 | 100.4064 | 98.5838 | 108.7332 | 21.0937 | 0.003758 | 0.003426 | 156.1144 | 0.25475 | 8.4029 | 0.002732 | 0.004004 | 0.003887 | 6.3e-05 | 0.000858 | 0.28727 | 1.3135 | 2.5473 | 100.3 | 21.02 | 31.857 | 112.1896 | 87.3785 | 104.1445 | 0.007918 | 0.93573 | 0.021849 | 23.3499 | 0.16386 | 0.15155 | 0.046957 | 3.0599 | 3.0599 | 0.002016
2019-03-29 11:41 | 1553874070.8631 | 351 | 0.20006 | 0.007476 | 0.000192 | 100.4205 | 98.5152 | 108.8468 | 21.0944 | 0.003807 | 0.003457 | 156.0011 | 0.2562 | 8.4016 | 0.002767 | 0.004336 | 0.004239 | 6.6e-05 | 0.000861 | 0.28766 | 1.3382 | 2.5474 | 100.3 | 21.02 | 31.86 | 113.3964 | 88.231 | 104.882 | 0.007919 | 0.9339 | 0.021839 | 23.3849 | 0.17134 | 0.15241 | 0.047693 | 3.1064 | 3.1064 | 0.001997
2019-03-29 11:41 | 1553874071.8636 | 352 | 0.19994 | 0.007537 | 0.000191 | 100.4702 | 98.5981 | 108.5858 | 21.0937 | 0.003745 | 0.003457 | 155.8861 | 0.25521 | 8.4038 | 0.00272 | 0.00395 | 0.003906 | 6.5e-05 | 0.000862 | 0.30702 | 1.3419 | 2.5474 | 100.3 | 21.02 | 31.865 | 111.8793 | 87.0849 | 104.8743 | 0.00792 | 0.93392 | 0.021813 | 23.3566 | 0.16824 | 0.15183 | 0.048389 | 3.1479 | 3.1479 | 0.001963
2019-03-29 11:41 | 1553874072.8633 | 353 | 0.19986 | 0.007602 | 0.000191 | 100.4782 | 98.2759 | 108.4063 | 21.0933 | 0.003793 | 0.003473 | 155.7543 | 0.25626 | 8.4011 | 0.002821 | 0.00408 | 0.003964 | 6.5e-05 | 0.000862 | 0.2672 | 1.3331 | 2.5475 | 100.3 | 21.01 | 31.867 | 113.0382 | 89.5158 | 105.2789 | 0.007916 | 0.93292 | 0.021784 | 23.3498 | 0.19823 | 0.15245 | 0.048786 | 3.1694 | 3.1694 | 0.001958
2019-03-29 11:41 | 1553874073.8642 | 354 | 0.19984 | 0.007681 | 0.000191 | 100.4881 | 98.2002 | 108.164 | 21.0942 | 0.003813 | 0.003483 | 155.6305 | 0.25501 | 8.4003 | 0.002752 | 0.004335 | 0.004257 | 6.4e-05 | 0.000864 | 0.26792 | 1.3568 | 2.5474 | 100.3 | 21.02 | 31.873 | 113.5464 | 87.8504 | 105.5182 | 0.007922 | 0.93233 | 0.021752 | 23.3311 | 0.20601 | 0.15171 | 0.048794 | 3.1653 | 3.1653 | 0.001951
2019-03-29 11:41 | 1553874074.8643 | 355 | 0.19984 | 0.00775 | 0.000192 | 100.5112 | 98.5203 | 108.0943 | 21.0954 | 0.003794 | 0.003462 | 155.5123 | 0.25572 | 8.4018 | 0.00276 | 0.004096 | 0.003961 | 6.4e-05 | 0.000858 | 0.26753 | 1.3575 | 2.5476 | 100.3 | 21.02 | 31.875 | 113.0696 | 88.0437 | 104.9954 | 0.007922 | 0.93362 | 0.02176 | 23.3074 | 0.17895 | 0.15213 | 0.048706 | 3.1608 | 3.1608 | 0.001959
2019-03-29 11:41 | 1553874075.8633 | 356 | 0.19984 | 0.007819 | 0.000193 | 100.4938 | 98.1227 | 107.5789 | 21.0944 | 0.003799 | 0.003461 | 155.3963 | 0.25549 | 8.4024 | 0.002773 | 0.00449 | 0.004438 | 6.3e-05 | 0.000861 | 0.29518 | 1.3265 | 2.5475 | 100.3 | 21.01 | 31.878 | 113.2013 | 88.3715 | 104.9709 | 0.007918 | 0.93368 | 0.021709 | 23.2511 | 0.21357 | 0.15199 | 0.048627 | 3.1482 | 3.1482 | 0.001965
2019-03-29 11:41 | 1553874076.8644 | 357 | 0.19984 | 0.007878 | 0.000194 | 100.4974 | 98.2008 | 107.2967 | 21.0945 | 0.003798 | 0.003531 | 155.2785 | 0.25539 | 8.4021 | 0.002769 | 0.004293 | 0.004073 | 6.1e-05 | 0.000866 | 0.27776 | 1.329 | 2.5473 | 100.3 | 21.01 | 31.885 | 113.1697 | 88.2584 | 106.6879 | 0.00792 | 0.92946 | 0.021632 | 23.2732 | 0.20677 | 0.15193 | 0.048514 | 3.1296 | 3.1296 | 0.001976
2019-03-29 11:41 | 1553874077.8634 | 358 | 0.19978 | 0.007921 | 0.000195 | 100.5273 | 98.0408 | 107.2439 | 21.0945 | 0.003794 | 0.003487 | 155.1555 | 0.25462 | 8.4015 | 0.002728 | 0.00382 | 0.003766 | 6.3e-05 | 0.000866 | 0.26576 | 1.3356 | 2.5474 | 100.3 | 21.01 | 31.889 | 113.0719 | 87.2823 | 105.6192 | 0.007921 | 0.93209 | 0.021657 | 23.2347 | 0.22403 | 0.15147 | 0.048822 | 3.1531 | 3.1531 | 0.001955
2019-03-29 11:41 | 1553874078.8635 | 359 | 0.19977 | 0.007951 | 0.000197 | 100.5764 | 98.064 | 107.3465 | 21.0943 | 0.003866 | 0.003513 | 155.0193 | 0.25531 | 8.4026 | 0.00269 | 0.004323 | 0.004212 | 6.4e-05 | 0.000874 | 0.2844 | 1.3298 | 2.5473 | 100.3 | 21.01 | 31.893 | 114.8297 | 86.3641 | 106.2406 | 0.007922 | 0.93056 | 0.021649 | 23.2649 | 0.22628 | 0.15188 | 0.048851 | 3.1537 | 3.1537 | 0.00196
2019-03-29 11:41 | 1553874079.8639 | 360 | 0.19975 | 0.007983 | 0.000198 | 100.607 | 98.2613 | 106.9622 | 21.0947 | 0.003799 | 0.00355 | 154.9117 | 0.25535 | 8.402 | 0.0027 | 0.004045 | 0.003935 | 6.2e-05 | 0.000881 | 0.30685 | 1.3619 | 2.5474 | 100.3 | 21.01 | 31.901 | 113.1826 | 86.6065 | 107.1322 | 0.007924 | 0.92838 | 0.021585 | 23.2505 | 0.21102 | 0.15191 | 0.048949 | 3.1506 | 3.1506 | 0.001962
2019-03-29 11:41 | 1553874080.865 | 361 | 0.1997 | 0.008005 | 0.000199 | 100.6126 | 98.1148 | 106.7759 | 21.0945 | 0.003937 | 0.003537 | 154.7817 | 0.25514 | 8.4038 | 0.002645 | 0.004753 | 0.004655 | 6.6e-05 | 0.000881 | 0.27848 | 1.3395 | 2.5473 | 100.3 | 21.01 | 31.906 | 116.5492 | 85.2779 | 106.8156 | 0.007925 | 0.92915 | 0.021575 | 23.2206 | 0.22486 | 0.15178 | 0.049241 | 3.1679 | 3.1679 | 0.00195
2019-03-29 11:41 | 1553874081.8637 | 362 | 0.19956 | 0.008028 | 0.000199 | 100.6028 | 98.1374 | 106.6637 | 21.0948 | 0.004001 | 0.003599 | 154.647 | 0.25564 | 8.4045 | 0.002697 | 0.00425 | 0.004151 | 5.7e-05 | 0.00088 | 0.30071 | 1.3393 | 2.5474 | 100.3 | 21.01 | 31.915 | 118.1113 | 86.5465 | 108.3358 | 0.007927 | 0.92545 | 0.021521 | 23.2547 | 0.22193 | 0.15208 | 0.050023 | 3.21 | 3.21 | 0.001928
2019-03-29 11:41 | 1553874082.8645 | 363 | 0.19941 | 0.008061 | 0.000198 | 100.6399 | 97.9592 | 106.9054 | 21.0944 | 0.003923 | 0.003623 | 154.5238 | 0.25599 | 8.4048 | 0.002667 | 0.003639 | 0.003589 | 6e-05 | 0.000875 | 0.28639 | 1.3096 | 2.5474 | 100.3 | 21.01 | 31.92 | 116.2133 | 85.8192 | 108.9027 | 0.007929 | 0.92407 | 0.021529 | 23.2984 | 0.24148 | 0.15229 | 0.050898 | 3.2673 | 3.2673 | 0.001897
2019-03-29 11:41 | 1553874083.8626 | 364 | 0.19936 | 0.0081 | 0.000195 | 100.7183 | 98.0067 | 106.9526 | 21.0942 | 0.003933 | 0.003615 | 154.3996 | 0.25528 | 8.4029 | 0.002739 | 0.004536 | 0.004428 | 5.9e-05 | 0.000874 | 0.27489 | 1.3106 | 2.5474 | 100.3 | 21.01 | 31.924 | 116.472 | 87.5522 | 108.7221 | 0.00793 | 0.92451 | 0.021539 | 23.298 | 0.24412 | 0.15187 | 0.051167 | 3.2862 | 3.2862 | 0.001881
2019-03-29 11:41 | 1553874084.8642 | 365 | 0.19931 | 0.00815 | 0.000194 | 100.6842 | 98.0766 | 107.2596 | 21.0947 | 0.003912 | 0.003621 | 154.2505 | 0.25591 | 8.4034 | 0.002689 | 0.004051 | 0.003833 | 6.2e-05 | 0.000877 | 0.29535 | 1.3097 | 2.5473 | 100.3 | 21.01 | 31.931 | 115.9422 | 86.3394 | 108.8739 | 0.007931 | 0.92414 | 0.021566 | 23.336 | 0.2347 | 0.15224 | 0.051358 | 3.3026 | 3.3026 | 0.001876
2019-03-29 11:41 | 1553874085.8637 | 366 | 0.19914 | 0.008209 | 0.000194 | 100.7189 | 98.3234 | 107.8854 | 21.094 | 0.003962 | 0.003583 | 154.1201 | 0.25581 | 8.4011 | 0.002667 | 0.004335 | 0.004314 | 6.2e-05 | 0.000874 | 0.30845 | 1.3157 | 2.5473 | 100.3 | 21.01 | 31.933 | 117.1568 | 85.8073 | 107.9403 | 0.007932 | 0.92641 | 0.021655 | 23.3754 | 0.21531 | 0.15218 | 0.052307 | 3.3774 | 3.3774 | 0.001834
2019-03-29 11:41 | 1553874086.8628 | 367 | 0.19909 | 0.008269 | 0.000195 | 100.6941 | 98.3913 | 107.9338 | 21.0944 | 0.003947 | 0.003592 | 153.984 | 0.25558 | 8.4016 | 0.002725 | 0.003926 | 0.003779 | 5.8e-05 | 0.000869 | 0.25299 | 1.3437 | 2.5474 | 100.3 | 21.01 | 31.938 | 116.7985 | 87.2151 | 108.1567 | 0.007933 | 0.92588 | 0.021654 | 23.3873 | 0.20693 | 0.15204 | 0.052521 | 3.3909 | 3.3909 | 0.001825
2019-03-29 11:41 | 1553874087.8637 | 368 | 0.19909 | 0.00833 | 0.000199 | 100.7563 | 98.4681 | 108.3225 | 21.0956 | 0.003986 | 0.003594 | 153.8471 | 0.25506 | 8.4013 | 0.002742 | 0.004636 | 0.004513 | 5.9e-05 | 0.000879 | 0.30359 | 1.3464 | 2.5474 | 100.3 | 21.01 | 31.94 | 117.7579 | 87.6215 | 108.2122 | 0.007934 | 0.92575 | 0.021691 | 23.4311 | 0.20548 | 0.15173 | 0.05246 | 3.3927 | 3.3927 | 0.00182
2019-03-29 11:41 | 1553874088.8633 | 369 | 0.19909 | 0.008399 | 0.000201 | 100.7345 | 98.4094 | 108.2713 | 21.0945 | 0.003963 | 0.003609 | 153.7047 | 0.25486 | 8.4016 | 0.002692 | 0.0039 | 0.003725 | 5.6e-05 | 0.00088 | 0.28385 | 1.325 | 2.5474 | 100.3 | 21.01 | 31.94 | 117.2 | 86.4133 | 108.5703 | 0.007934 | 0.92488 | 0.021676 | 23.4365 | 0.20887 | 0.15162 | 0.052377 | 3.3848 | 3.3848 | 0.001823
2019-03-29 11:41 | 1553874089.8639 | 370 | 0.19908 | 0.00845 | 0.000204 | 100.7779 | 98.3309 | 108.2154 | 21.096 | 0.003993 | 0.003659 | 153.5634 | 0.25607 | 8.4036 | 0.002652 | 0.00501 | 0.004912 | 6e-05 | 0.000878 | 0.25625 | 1.3153 | 2.5475 | 100.3 | 21.01 | 31.942 | 117.9317 | 85.4549 | 109.793 | 0.007934 | 0.92193 | 0.021636 | 23.468 | 0.21986 | 0.15233 | 0.052383 | 3.3787 | 3.3787 | 0.001835
2019-03-29 11:41 | 1553874090.8628 | 371 | 0.19913 | 0.0085 | 0.000205 | 100.8452 | 97.9845 | 108.4323 | 21.0951 | 0.004019 | 0.003619 | 153.4429 | 0.25529 | 8.3996 | 0.002731 | 0.004037 | 0.003946 | 5.8e-05 | 0.000884 | 0.29502 | 1.3259 | 2.5474 | 100.3 | 21.01 | 31.942 | 118.5592 | 87.3673 | 108.8097 | 0.007934 | 0.9243 | 0.021685 | 23.4613 | 0.25741 | 0.15187 | 0.052035 | 3.3639 | 3.3639 | 0.001837
2019-03-29 11:42 | 1553874091.8628 | 372 | 0.19925 | 0.008524 | 0.000206 | 100.7967 | 98.0527 | 108.3772 | 21.0942 | 0.003987 | 0.00366 | 153.3163 | 0.25536 | 8.404 | 0.002684 | 0.003823 | 0.003766 | 5.9e-05 | 0.000886 | 0.29264 | 1.3466 | 2.5474 | 100.3 | 21.01 | 31.943 | 117.781 | 86.2321 | 109.8116 | 0.007934 | 0.92188 | 0.021651 | 23.4861 | 0.24688 | 0.15191 | 0.051305 | 3.3116 | 3.3116 | 0.001867
2019-03-29 11:42 | 1553874092.864 | 373 | 0.19924 | 0.008537 | 0.000206 | 100.8365 | 98.194 | 108.2483 | 21.0955 | 0.003953 | 0.003656 | 153.1747 | 0.25491 | 8.4001 | 0.002707 | 0.003661 | 0.003596 | 5.6e-05 | 0.00089 | 0.29087 | 1.3562 | 2.5475 | 100.3 | 21.01 | 31.942 | 116.9515 | 86.7774 | 109.7113 | 0.007934 | 0.92212 | 0.021641 | 23.469 | 0.23753 | 0.15165 | 0.051295 | 3.3094 | 3.3094 | 0.001865
2019-03-29 11:42 | 1553874093.8639 | 374 | 0.19925 | 0.008531 | 0.000206 | 100.8695 | 98.0565 | 107.8692 | 21.0957 | 0.004074 | 0.00374 | 153.0551 | 0.25551 | 8.4006 | 0.002703 | 0.003948 | 0.003841 | 6.2e-05 | 0.000883 | 0.28943 | 1.3408 | 2.5474 | 100.3 | 21.01 | 31.942 | 119.895 | 86.6871 | 111.7669 | 0.007934 | 0.9172 | 0.021546 | 23.4907 | 0.25299 | 0.15201 | 0.051233 | 3.2908 | 3.2908 | 0.00188
2019-03-29 11:42 | 1553874094.8643 | 375 | 0.19926 | 0.008525 | 0.000205 | 100.8713 | 98.4408 | 107.5472 | 21.0955 | 0.003994 | 0.003707 | 152.9446 | 0.25599 | 8.404 | 0.002735 | 0.003907 | 0.003786 | 5.8e-05 | 0.000887 | 0.28351 | 1.3475 | 2.5475 | 100.3 | 21.01 | 31.943 | 117.9519 | 87.4543 | 110.9639 | 0.007934 | 0.91912 | 0.021536 | 23.4311 | 0.21816 | 0.15229 | 0.051207 | 3.2877 | 3.2877 | 0.001885
2019-03-29 11:42 | 1553874095.8643 | 376 | 0.19926 | 0.008514 | 0.000204 | 100.9093 | 98.5442 | 107.1413 | 21.0954 | 0.004046 | 0.003709 | 152.7867 | 0.25606 | 8.4052 | 0.00271 | 0.004348 | 0.004297 | 5.6e-05 | 0.000889 | 0.2959 | 1.3526 | 2.5475 | 100.3 | 21.01 | 31.942 | 119.2214 | 86.849 | 111.0025 | 0.007934 | 0.91902 | 0.021494 | 23.3881 | 0.21214 | 0.15233 | 0.051251 | 3.2843 | 3.2843 | 0.001888
2019-03-29 11:42 | 1553874096.864 | 377 | 0.19914 | 0.008502 | 0.000201 | 100.973 | 98.3787 | 106.7989 | 21.0949 | 0.004054 | 0.003761 | 152.6403 | 0.25493 | 8.4021 | 0.002736 | 0.003923 | 0.003912 | 5.3e-05 | 0.000889 | 0.29856 | 1.3472 | 2.5474 | 100.3 | 21.01 | 31.942 | 119.424 | 87.4725 | 112.2737 | 0.007934 | 0.91599 | 0.021424 | 23.3893 | 0.23282 | 0.15166 | 0.051933 | 3.3172 | 3.3172 | 0.001861
2019-03-29 11:42 | 1553874097.8631 | 378 | 0.1989 | 0.008502 | 0.000198 | 100.9006 | 98.5503 | 106.182 | 21.0955 | 0.004192 | 0.003833 | 152.4859 | 0.25529 | 8.4043 | 0.002601 | 0.00447 | 0.004253 | 5.7e-05 | 0.000895 | 0.2912 | 1.3532 | 2.5474 | 100.3 | 21.01 | 31.943 | 122.7884 | 84.2355 | 114.0235 | 0.007934 | 0.91185 | 0.021314 | 23.3745 | 0.21081 | 0.15187 | 0.05338 | 3.392 | 3.392 | 0.001822
2019-03-29 11:42 | 1553874098.8631 | 379 | 0.19866 | 0.00853 | 0.000195 | 100.9434 | 98.61 | 105.5844 | 21.096 | 0.004141 | 0.003793 | 152.3528 | 0.25581 | 8.4046 | 0.002656 | 0.003983 | 0.003912 | 5.8e-05 | 0.000893 | 0.28313 | 1.3849 | 2.5476 | 100.3 | 21.01 | 31.942 | 121.5427 | 85.5422 | 113.0551 | 0.007934 | 0.91414 | 0.021281 | 23.2795 | 0.20919 | 0.15218 | 0.054807 | 3.4771 | 3.4771 | 0.001781
2019-03-29 11:42 | 1553874099.8622 | 380 | 0.19847 | 0.008584 | 0.000192 | 100.9796 | 98.7036 | 105.4081 | 21.0958 | 0.004151 | 0.003801 | 152.2053 | 0.25547 | 8.4012 | 0.002703 | 0.004142 | 0.004109 | 5.3e-05 | 0.000895 | 0.29407 | 1.3379 | 2.5474 | 100.3 | 21.01 | 31.941 | 121.784 | 86.6731 | 113.2485 | 0.007934 | 0.91368 | 0.021258 | 23.2658 | 0.20392 | 0.15198 | 0.055893 | 3.5422 | 3.5422 | 0.001746
2019-03-29 11:42 | 1553874100.8626 | 381 | 0.19839 | 0.008664 | 0.000189 | 101.0422 | 98.8561 | 104.8279 | 21.0967 | 0.004159 | 0.00388 | 152.0399 | 0.25552 | 8.4035 | 0.002654 | 0.003534 | 0.003482 | 5.6e-05 | 0.000897 | 0.30414 | 1.3276 | 2.5474 | 100.3 | 21.01 | 31.939 | 121.9792 | 85.5092 | 115.1672 | 0.007933 | 0.90917 | 0.021147 | 23.2592 | 0.19564 | 0.15201 | 0.05627 | 3.5475 | 3.5475 | 0.001744
2019-03-29 11:42 | 1553874101.8623 | 382 | 0.19812 | 0.008755 | 0.000186 | 101.0443 | 98.7036 | 104.5014 | 21.0964 | 0.004222 | 0.003822 | 151.8953 | 0.25462 | 8.4007 | 0.00273 | 0.003698 | 0.003598 | 5.8e-05 | 0.000887 | 0.26808 | 1.3381 | 2.5474 | 100.3 | 21.01 | 31.937 | 123.5268 | 87.3348 | 113.7659 | 0.007933 | 0.91246 | 0.021152 | 23.1811 | 0.20964 | 0.15147 | 0.057768 | 3.6427 | 3.6427 | 0.001692
2019-03-29 11:42 | 1553874102.8636 | 383 | 0.19775 | 0.008857 | 0.000183 | 101.0356 | 98.4363 | 104.1681 | 21.0965 | 0.004204 | 0.003837 | 151.7552 | 0.2556 | 8.402 | 0.002655 | 0.00413 | 0.004008 | 5.8e-05 | 0.000885 | 0.28423 | 1.3591 | 2.5474 | 100.3 | 21.01 | 31.933 | 123.0985 | 85.5233 | 114.1298 | 0.007932 | 0.9116 | 0.021108 | 23.1549 | 0.23312 | 0.15205 | 0.059822 | 3.7642 | 3.7642 | 0.001644
2019-03-29 11:42 | 1553874103.8626 | 384 | 0.19764 | 0.008984 | 0.00018 | 101.0482 | 98.8029 | 103.9101 | 21.0964 | 0.004282 | 0.003854 | 151.5715 | 0.25516 | 8.4048 | 0.002732 | 0.004188 | 0.004059 | 5.7e-05 | 0.000878 | 0.286 | 1.3497 | 2.5473 | 100.3 | 21.01 | 31.929 | 124.9922 | 87.3856 | 114.5302 | 0.007931 | 0.91066 | 0.021071 | 23.1382 | 0.20099 | 0.1518 | 0.060303 | 3.7878 | 3.7878 | 0.001631
2019-03-29 11:42 | 1553874104.8634 | 385 | 0.19756 | 0.009111 | 0.000178 | 101.0814 | 98.9653 | 103.8569 | 21.0965 | 0.004199 | 0.003821 | 151.4101 | 0.25597 | 8.4018 | 0.002771 | 0.003939 | 0.00381 | 5.6e-05 | 0.000878 | 0.28783 | 1.3923 | 2.5474 | 100.3 | 21.01 | 31.925 | 122.9523 | 88.3151 | 113.7266 | 0.00793 | 0.91255 | 0.021088 | 23.1083 | 0.18924 | 0.15228 | 0.060688 | 3.815 | 3.815 | 0.001625
2019-03-29 11:42 | 1553874105.8623 | 386 | 0.19752 | 0.00925 | 0.000177 | 101.0488 | 99.0418 | 103.6479 | 21.0965 | 0.004261 | 0.003864 | 151.2686 | 0.25613 | 8.4023 | 0.002832 | 0.004371 | 0.004337 | 5.7e-05 | 0.000869 | 0.29214 | 1.3526 | 2.5475 | 100.3 | 21.01 | 31.923 | 124.4905 | 89.7877 | 114.7773 | 0.007929 | 0.91008 | 0.021038 | 23.1164 | 0.17944 | 0.15237 | 0.060744 | 3.8096 | 3.8096 | 0.001628
2019-03-29 11:42 | 1553874106.8633 | 387 | 0.19745 | 0.00937 | 0.000175 | 101.0598 | 98.935 | 104.0599 | 21.0968 | 0.004303 | 0.003961 | 151.1148 | 0.25602 | 8.4033 | 0.002912 | 0.00428 | 0.004164 | 5.9e-05 | 0.00087 | 0.31542 | 1.3253 | 2.5474 | 100.3 | 21.01 | 31.922 | 125.5194 | 91.725 | 117.1572 | 0.007929 | 0.90453 | 0.021015 | 23.2332 | 0.19007 | 0.15231 | 0.061036 | 3.8238 | 3.8238 | 0.001621
2019-03-29 11:42 | 1553874107.8638 | 388 | 0.19728 | 0.009487 | 0.000175 | 101.1386 | 99.023 | 103.7456 | 21.0968 | 0.004305 | 0.0039 | 150.967 | 0.25577 | 8.4009 | 0.002956 | 0.005259 | 0.005142 | 6.1e-05 | 0.000865 | 0.29662 | 1.3211 | 2.5474 | 100.3 | 21.01 | 31.919 | 125.5591 | 92.787 | 115.665 | 0.007928 | 0.908 | 0.021024 | 23.1537 | 0.18908 | 0.15216 | 0.061893 | 3.8788 | 3.8788 | 0.001597
2019-03-29 11:42 | 1553874108.8624 | 389 | 0.19713 | 0.009585 | 0.000175 | 101.1098 | 99.065 | 103.5076 | 21.0966 | 0.004308 | 0.003976 | 150.7994 | 0.25626 | 8.4006 | 0.00294 | 0.003703 | 0.003648 | 6.2e-05 | 0.000859 | 0.26344 | 1.3141 | 2.5475 | 100.3 | 21.01 | 31.918 | 125.64 | 92.3815 | 117.5164 | 0.007928 | 0.9037 | 0.02095 | 23.1821 | 0.18275 | 0.15245 | 0.062673 | 3.9137 | 3.9137 | 0.001585
2019-03-29 11:42 | 1553874109.8627 | 390 | 0.19711 | 0.009685 | 0.000176 | 101.1053 | 98.9124 | 103.1762 | 21.0963 | 0.004363 | 0.003968 | 150.6195 | 0.25552 | 8.4013 | 0.002963 | 0.004401 | 0.004302 | 5.9e-05 | 0.000856 | 0.2975 | 1.3469 | 2.5474 | 100.3 | 21.01 | 31.918 | 126.9972 | 92.9451 | 117.3236 | 0.007928 | 0.90415 | 0.020921 | 23.1393 | 0.19614 | 0.15201 | 0.062664 | 3.9078 | 3.9078 | 0.001583
2019-03-29 11:42 | 1553874110.8635 | 391 | 0.19708 | 0.009779 | 0.000174 | 101.0347 | 98.9438 | 102.7353 | 21.0973 | 0.00429 | 0.003977 | 150.4562 | 0.25663 | 8.4016 | 0.003024 | 0.004069 | 0.003974 | 6e-05 | 0.000859 | 0.28783 | 1.3873 | 2.5475 | 100.3 | 21.01 | 31.919 | 125.1993 | 94.4266 | 117.5338 | 0.007928 | 0.90366 | 0.020871 | 23.096 | 0.18705 | 0.15267 | 0.062725 | 3.9023 | 3.9023 | 0.001592
2019-03-29 11:42 | 1553874111.8639 | 392 | 0.19693 | 0.009873 | 0.00017 | 101.1326 | 98.9879 | 102.863 | 21.0972 | 0.004376 | 0.00399 | 150.2721 | 0.25549 | 8.4025 | 0.002993 | 0.004235 | 0.00414 | 6.2e-05 | 0.000863 | 0.29302 | 1.3749 | 2.5474 | 100.3 | 21.01 | 31.919 | 127.3112 | 93.6769 | 117.853 | 0.007928 | 0.90292 | 0.020875 | 23.1198 | 0.19173 | 0.15199 | 0.063496 | 3.951 | 3.951 | 0.001566
2019-03-29 11:42 | 1553874112.8636 | 393 | 0.19673 | 0.009951 | 0.000168 | 101.1062 | 98.7653 | 103.0083 | 21.0979 | 0.004413 | 0.003999 | 150.0977 | 0.25526 | 8.4007 | 0.003035 | 0.004478 | 0.004305 | 6.1e-05 | 0.000866 | 0.29192 | 1.3733 | 2.5475 | 100.3 | 21.01 | 31.923 | 128.2196 | 94.6721 | 118.0699 | 0.007929 | 0.90242 | 0.020884 | 23.1425 | 0.20953 | 0.15185 | 0.064577 | 4.0199 | 4.0199 | 0.001537
2019-03-29 11:42 | 1553874113.8616 | 394 | 0.19627 | 0.01005 | 0.000164 | 101.0856 | 98.8814 | 102.7417 | 21.0987 | 0.004529 | 0.00403 | 149.9572 | 0.25632 | 8.4017 | 0.003045 | 0.005009 | 0.004943 | 6.4e-05 | 0.000874 | 0.31365 | 1.3212 | 2.5474 | 100.3 | 21.01 | 31.927 | 131.0655 | 94.925 | 118.8361 | 0.00793 | 0.90066 | 0.020837 | 23.1352 | 0.1972 | 0.15249 | 0.067209 | 4.1736 | 4.1736 | 0.001487
2019-03-29 11:42 | 1553874114.8616 | 395 | 0.19602 | 0.010181 | 0.000164 | 101.2035 | 98.9916 | 102.2443 | 21.0985 | 0.004444 | 0.004088 | 149.7867 | 0.25618 | 8.4009 | 0.003058 | 0.004775 | 0.004726 | 6e-05 | 0.000867 | 0.30735 | 1.3128 | 2.5473 | 100.3 | 21.01 | 31.931 | 128.9756 | 95.2505 | 120.254 | 0.007931 | 0.89741 | 0.020749 | 23.1208 | 0.19766 | 0.1524 | 0.068574 | 4.24 | 4.24 | 0.001463
2019-03-29 11:42 | 1553874115.8617 | 396 | 0.19601 | 0.010331 | 0.000163 | 101.1287 | 98.9763 | 101.9612 | 21.0976 | 0.004461 | 0.004098 | 149.6331 | 0.25579 | 8.3998 | 0.003032 | 0.004006 | 0.003905 | 5.8e-05 | 0.000865 | 0.29535 | 1.3282 | 2.5474 | 100.3 | 21.01 | 31.939 | 129.384 | 94.6131 | 120.4943 | 0.007933 | 0.89686 | 0.020714 | 23.0958 | 0.19243 | 0.15217 | 0.068471 | 4.2265 | 4.2265 | 0.001465
2019-03-29 11:42 | 1553874116.8616 | 397 | 0.19606 | 0.010487 | 0.000164 | 101.0598 | 98.879 | 101.5305 | 21.0985 | 0.004463 | 0.004095 | 149.4579 | 0.25599 | 8.4044 | 0.003031 | 0.004998 | 0.004864 | 6.1e-05 | 0.000861 | 0.26504 | 1.3319 | 2.5475 | 100.3 | 21.01 | 31.943 | 129.4503 | 94.5875 | 120.4334 | 0.007934 | 0.897 | 0.020672 | 23.0452 | 0.19514 | 0.15229 | 0.067984 | 4.188 | 4.188 | 0.00148
2019-03-29 11:42 | 1553874117.8617 | 398 | 0.19609 | 0.01063 | 0.000167 | 101.1532 | 98.95 | 101.4559 | 21.0983 | 0.004535 | 0.004122 | 149.2854 | 0.2553 | 8.4046 | 0.002957 | 0.004466 | 0.004397 | 6e-05 | 0.000856 | 0.2656 | 1.313 | 2.5474 | 100.3 | 21.01 | 31.949 | 131.2142 | 92.8024 | 121.0854 | 0.007936 | 0.89552 | 0.020647 | 23.0558 | 0.19698 | 0.15188 | 0.067641 | 4.1618 | 4.1618 | 0.001485
2019-03-29 11:42 | 1553874118.8631 | 399 | 0.19611 | 0.010743 | 0.00017 | 101.1547 | 98.8326 | 101.8948 | 21.0977 | 0.004502 | 0.004128 | 149.1254 | 0.25634 | 8.4035 | 0.003004 | 0.005544 | 0.005354 | 5.8e-05 | 0.000857 | 0.28854 | 1.3518 | 2.5475 | 100.3 | 21.01 | 31.952 | 130.3996 | 93.9378 | 121.2331 | 0.007937 | 0.89518 | 0.020688 | 23.11 | 0.20773 | 0.1525 | 0.067365 | 4.1529 | 4.1529 | 0.001495
2019-03-29 11:42 | 1553874119.8618 | 400 | 0.19612 | 0.010806 | 0.000172 | 101.1368 | 98.8247 | 102.0685 | 21.0981 | 0.004576 | 0.004152 | 148.966 | 0.25584 | 8.4009 | 0.003078 | 0.004753 | 0.004655 | 6.1e-05 | 0.000857 | 0.262 | 1.3109 | 2.5473 | 100.3 | 21.01 | 31.955 | 132.2131 | 95.7114 | 121.8037 | 0.007937 | 0.89389 | 0.02069 | 23.1464 | 0.20685 | 0.1522 | 0.067214 | 4.1441 | 4.1441 | 0.001495
2019-03-29 11:42 | 1553874120.8628 | 401 | 0.19611 | 0.010844 | 0.000173 | 101.1296 | 98.8496 | 102.5959 | 21.0981 | 0.004551 | 0.004118 | 148.8033 | 0.25538 | 8.403 | 0.003148 | 0.004796 | 0.004679 | 5.6e-05 | 0.00086 | 0.30021 | 1.2981 | 2.5474 | 100.3 | 21.01 | 31.959 | 131.6181 | 97.4143 | 120.9866 | 0.007938 | 0.89574 | 0.020765 | 23.1821 | 0.20396 | 0.15192 | 0.067215 | 4.1591 | 4.1591 | 0.001487
2019-03-29 11:42 | 1553874121.8618 | 402 | 0.19611 | 0.01087 | 0.000173 | 101.1499 | 98.8926 | 102.503 | 21.0994 | 0.004553 | 0.004123 | 148.6506 | 0.25582 | 8.4039 | 0.003123 | 0.004421 | 0.004278 | 5.7e-05 | 0.000863 | 0.29407 | 1.3452 | 2.5474 | 100.3 | 21.01 | 31.962 | 131.6528 | 96.8146 | 121.1151 | 0.007939 | 0.89545 | 0.020752 | 23.1753 | 0.20187 | 0.15219 | 0.067229 | 4.1574 | 4.1574 | 0.00149
2019-03-29 11:42 | 1553874122.8622 | 403 | 0.19597 | 0.010884 | 0.000172 | 101.0877 | 98.7421 | 102.0945 | 21.0982 | 0.004597 | 0.004164 | 148.5015 | 0.25694 | 8.4028 | 0.003229 | 0.005085 | 0.004979 | 6.7e-05 | 0.000867 | 0.27903 | 1.3816 | 2.5475 | 100.3 | 21.01 | 31.966 | 132.7517 | 99.3603 | 122.1119 | 0.00794 | 0.89319 | 0.020685 | 23.1584 | 0.20999 | 0.15285 | 0.068028 | 4.1929 | 4.1929 | 0.001484
2019-03-29 11:42 | 1553874123.8624 | 404 | 0.19586 | 0.010915 | 0.000171 | 101.1236 | 98.821 | 102.1678 | 21.0992 | 0.004581 | 0.004174 | 148.3251 | 0.25533 | 8.4029 | 0.003167 | 0.005201 | 0.005028 | 6.3e-05 | 0.000858 | 0.30359 | 1.3264 | 2.5474 | 100.3 | 21.01 | 31.973 | 132.3445 | 97.8826 | 122.3465 | 0.007942 | 0.89266 | 0.020686 | 23.1735 | 0.20603 | 0.1519 | 0.068647 | 4.2312 | 4.2312 | 0.001461
2019-03-29 11:42 | 1553874124.8619 | 405 | 0.19575 | 0.010957 | 0.000169 | 101.1455 | 98.8974 | 102.1239 | 21.0988 | 0.004628 | 0.004215 | 148.1573 | 0.25521 | 8.4016 | 0.003384 | 0.004516 | 0.004424 | 6.7e-05 | 0.000856 | 0.3183 | 1.3264 | 2.5473 | 100.3 | 21.01 | 31.976 | 133.4965 | 103.1136 | 123.355 | 0.007943 | 0.89039 | 0.020655 | 23.1981 | 0.20104 | 0.15183 | 0.069223 | 4.2602 | 4.2602 | 0.00145
2019-03-29 11:42 | 1553874125.863 | 406 | 0.19572 | 0.011002 | 0.000166 | 101.1344 | 98.9165 | 102.0281 | 21.0985 | 0.004594 | 0.004234 | 147.9825 | 0.25565 | 8.4063 | 0.003359 | 0.004241 | 0.004169 | 6.3e-05 | 0.000854 | 0.3011 | 1.3295 | 2.5473 | 100.3 | 21.01 | 31.981 | 132.6737 | 102.5232 | 123.8356 | 0.007944 | 0.88931 | 0.020633 | 23.2012 | 0.19834 | 0.15209 | 0.069349 | 4.2635 | 4.2635 | 0.001452
2019-03-29 11:42 | 1553874126.8634 | 407 | 0.19568 | 0.011059 | 0.000162 | 101.0542 | 98.9943 | 101.9253 | 21.0996 | 0.00468 | 0.004265 | 147.8167 | 0.25617 | 8.4015 | 0.003477 | 0.00568 | 0.005545 | 6.4e-05 | 0.000851 | 0.32068 | 1.3624 | 2.5475 | 100.3 | 21.01 | 31.984 | 134.7958 | 105.3629 | 124.5962 | 0.007945 | 0.88762 | 0.020603 | 23.2118 | 0.1842 | 0.1524 | 0.069558 | 4.2703 | 4.2703 | 0.001452
2019-03-29 11:42 | 1553874127.8631 | 408 | 0.19557 | 0.011117 | 0.000158 | 101.0793 | 98.8756 | 102.1222 | 21.0991 | 0.00464 | 0.00422 | 147.6472 | 0.25573 | 8.4027 | 0.003516 | 0.005258 | 0.005109 | 6.5e-05 | 0.000845 | 0.29336 | 1.3752 | 2.5475 | 100.3 | 21.01 | 31.991 | 133.8123 | 106.3248 | 123.4824 | 0.007946 | 0.89011 | 0.020652 | 23.2016 | 0.19717 | 0.15214 | 0.070126 | 4.3153 | 4.3153 | 0.001435
2019-03-29 11:42 | 1553874128.8631 | 409 | 0.19544 | 0.011178 | 0.000156 | 101.0685 | 98.9865 | 102.184 | 21.1003 | 0.004675 | 0.004268 | 147.477 | 0.25597 | 8.4014 | 0.003583 | 0.005046 | 0.004977 | 6.4e-05 | 0.000848 | 0.28711 | 1.4087 | 2.5474 | 100.3 | 21.01 | 31.995 | 134.6612 | 107.9591 | 124.6628 | 0.007947 | 0.88747 | 0.020627 | 23.2431 | 0.18618 | 0.15228 | 0.070831 | 4.3535 | 4.3535 | 0.001424
2019-03-29 11:42 | 1553874129.8631 | 410 | 0.19538 | 0.011244 | 0.000155 | 101.0634 | 99.1213 | 102.1873 | 21.0996 | 0.004721 | 0.004255 | 147.3103 | 0.25496 | 8.402 | 0.003682 | 0.004644 | 0.004601 | 6.4e-05 | 0.000849 | 0.32693 | 1.3315 | 2.5475 | 100.3 | 21.01 | 31.999 | 135.8008 | 110.3408 | 124.3457 | 0.007948 | 0.88817 | 0.020636 | 23.2342 | 0.17356 | 0.15168 | 0.07109 | 4.3711 | 4.3711 | 0.001412
2019-03-29 11:42 | 1553874130.8612 | 411 | 0.19526 | 0.01131 | 0.000155 | 101.0515 | 99.1158 | 101.5981 | 21.1003 | 0.00466 | 0.004293 | 147.1453 | 0.25599 | 8.4021 | 0.00366 | 0.004708 | 0.004609 | 6.5e-05 | 0.00085 | 0.30917 | 1.3436 | 2.5473 | 100.3 | 21.01 | 32.009 | 134.3023 | 109.8154 | 125.2683 | 0.007951 | 0.88612 | 0.020553 | 23.194 | 0.17299 | 0.15229 | 0.071749 | 4.3936 | 4.3936 | 0.001411
2019-03-29 11:42 | 1553874131.8622 | 412 | 0.19514 | 0.011371 | 0.000157 | 101.0311 | 99.1568 | 101.366 | 21.1008 | 0.004757 | 0.004301 | 146.9837 | 0.25515 | 8.4039 | 0.003661 | 0.005038 | 0.004999 | 6.4e-05 | 0.000852 | 0.27865 | 1.3153 | 2.5473 | 100.3 | 21.01 | 32.015 | 136.6826 | 109.8472 | 125.4747 | 0.007952 | 0.88566 | 0.020524 | 23.1735 | 0.1675 | 0.15179 | 0.072396 | 4.4267 | 4.4267 | 0.001396
2019-03-29 11:42 | 1553874132.8615 | 413 | 0.19497 | 0.01143 | 0.00016 | 101.0607 | 98.9606 | 101.3622 | 21.1003 | 0.004696 | 0.004338 | 146.798 | 0.25521 | 8.4022 | 0.003642 | 0.005326 | 0.005199 | 6.4e-05 | 0.00085 | 0.30596 | 1.3213 | 2.5473 | 100.3 | 21.01 | 32.022 | 135.178 | 109.3872 | 126.3792 | 0.007954 | 0.88365 | 0.0205 | 23.1994 | 0.18784 | 0.15183 | 0.073318 | 4.4774 | 4.4774 | 0.00138
2019-03-29 11:42 | 1553874133.8612 | 414 | 0.19494 | 0.011503 | 0.000164 | 101.0311 | 98.6793 | 101.435 | 21.0999 | 0.004792 | 0.004289 | 146.6153 | 0.25565 | 8.4004 | 0.003734 | 0.004996 | 0.004961 | 6.4e-05 | 0.000845 | 0.30287 | 1.3431 | 2.5473 | 100.3 | 21.01 | 32.035 | 137.5546 | 111.6209 | 125.1707 | 0.007957 | 0.88633 | 0.020539 | 23.1726 | 0.21067 | 0.15209 | 0.073384 | 4.4896 | 4.4896 | 0.001379
2019-03-29 11:42 | 1553874134.8623 | 415 | 0.19493 | 0.01158 | 0.000168 | 101.0533 | 98.5595 | 101.6307 | 21.1002 | 0.00479 | 0.004359 | 146.4503 | 0.25581 | 8.4018 | 0.003694 | 0.005049 | 0.004964 | 6.3e-05 | 0.000846 | 0.29734 | 1.3155 | 2.5473 | 100.3 | 21.01 | 32.043 | 137.5099 | 110.6409 | 126.9007 | 0.007959 | 0.8825 | 0.020514 | 23.2452 | 0.2235 | 0.15218 | 0.073351 | 4.482 | 4.482 | 0.001382
2019-03-29 11:42 | 1553874135.8613 | 416 | 0.19494 | 0.01165 | 0.00017 | 100.9859 | 98.5807 | 101.3427 | 21.1013 | 0.00483 | 0.004371 | 146.2764 | 0.25533 | 8.4037 | 0.003608 | 0.005176 | 0.005029 | 6.3e-05 | 0.000847 | 0.30596 | 1.3439 | 2.5474 | 100.3 | 21.01 | 32.06 | 138.4791 | 108.5549 | 127.1863 | 0.007964 | 0.88187 | 0.020478 | 23.2206 | 0.21561 | 0.15189 | 0.073243 | 4.4673 | 4.4673 | 0.001384
2019-03-29 11:42 | 1553874136.8622 | 417 | 0.19495 | 0.011709 | 0.000172 | 101.1221 | 98.7698 | 101.4641 | 21.1015 | 0.004823 | 0.004327 | 146.1079 | 0.25469 | 8.4029 | 0.003596 | 0.004865 | 0.004751 | 6.5e-05 | 0.000842 | 0.31653 | 1.3195 | 2.5473 | 100.3 | 21.01 | 32.068 | 138.3062 | 108.2644 | 126.102 | 0.007965 | 0.88427 | 0.020518 | 23.203 | 0.21053 | 0.15151 | 0.073126 | 4.4689 | 4.4689 | 0.00138
2019-03-29 11:42 | 1553874137.8614 | 418 | 0.19501 | 0.01175 | 0.000172 | 101.0257 | 98.5459 | 101.3653 | 21.101 | 0.004829 | 0.004351 | 145.9416 | 0.25475 | 8.4026 | 0.003615 | 0.004508 | 0.004374 | 6.3e-05 | 0.000843 | 0.26537 | 1.3319 | 2.5474 | 100.3 | 21.01 | 32.076 | 138.4539 | 108.7248 | 126.7049 | 0.007967 | 0.88293 | 0.020492 | 23.2092 | 0.2223 | 0.15155 | 0.072694 | 4.4371 | 4.4371 | 0.00139
2019-03-29 11:42 | 1553874138.8628 | 419 | 0.19504 | 0.011784 | 0.000173 | 100.9709 | 98.4411 | 100.6149 | 21.1001 | 0.004839 | 0.004362 | 145.7663 | 0.25545 | 8.4015 | 0.003712 | 0.005349 | 0.005228 | 6.2e-05 | 0.000842 | 0.29894 | 1.374 | 2.5475 | 100.3 | 21.01 | 32.092 | 138.6998 | 111.0796 | 126.9669 | 0.007971 | 0.88236 | 0.020409 | 23.1307 | 0.22696 | 0.15197 | 0.072482 | 4.4063 | 4.4063 | 0.001404
2019-03-29 11:42 | 1553874139.8618 | 420 | 0.19501 | 0.011801 | 0.000175 | 101.0359 | 98.7763 | 100.3097 | 21.1015 | 0.004848 | 0.004418 | 145.5998 | 0.25606 | 8.3991 | 0.003714 | 0.005439 | 0.005296 | 6.1e-05 | 0.000848 | 0.28495 | 1.3687 | 2.5474 | 100.3 | 21.01 | 32.099 | 138.9362 | 111.1206 | 128.3347 | 0.007973 | 0.87935 | 0.020344 | 23.135 | 0.20231 | 0.15233 | 0.0726 | 4.3992 | 4.3992 | 0.001409
2019-03-29 11:42 | 1553874140.8629 | 421 | 0.19507 | 0.011808 | 0.000177 | 100.9668 | 98.9162 | 100.5106 | 21.1022 | 0.004836 | 0.00443 | 145.442 | 0.25502 | 8.4031 | 0.003691 | 0.005633 | 0.005611 | 6.1e-05 | 0.000854 | 0.31725 | 1.3327 | 2.5472 | 100.3 | 21.01 | 32.114 | 138.6411 | 110.569 | 128.6358 | 0.007977 | 0.87869 | 0.020357 | 23.1669 | 0.18353 | 0.15171 | 0.072269 | 4.3818 | 4.3818 | 0.001409
2019-03-29 11:42 | 1553874141.8626 | 422 | 0.19508 | 0.011806 | 0.000179 | 101.0362 | 98.8148 | 100.7224 | 21.1015 | 0.004911 | 0.004418 | 145.2775 | 0.25508 | 8.4016 | 0.003635 | 0.005067 | 0.00489 | 6.5e-05 | 0.000859 | 0.31006 | 1.3275 | 2.5473 | 100.3 | 21.01 | 32.12 | 140.4959 | 109.2219 | 128.3417 | 0.007978 | 0.87933 | 0.020385 | 23.1828 | 0.19885 | 0.15175 | 0.07221 | 4.3844 | 4.3844 | 0.001409
2019-03-29 11:42 | 1553874142.8615 | 423 | 0.19515 | 0.011823 | 0.00018 | 100.9757 | 98.7118 | 100.697 | 21.1013 | 0.004901 | 0.004402 | 145.1199 | 0.25556 | 8.4034 | 0.003651 | 0.005347 | 0.005281 | 6e-05 | 0.000863 | 0.30143 | 1.3327 | 2.5475 | 100.3 | 21.01 | 32.125 | 140.2543 | 109.5955 | 127.936 | 0.00798 | 0.88022 | 0.020393 | 23.1681 | 0.20283 | 0.15203 | 0.071749 | 4.3581 | 4.3581 | 0.00142
2019-03-29 11:42 | 1553874143.8619 | 424 | 0.1953 | 0.011822 | 0.00018 | 100.9955 | 98.8377 | 100.8979 | 21.1014 | 0.004922 | 0.00444 | 144.9751 | 0.25553 | 8.4018 | 0.003618 | 0.005109 | 0.005023 | 6.2e-05 | 0.000863 | 0.32764 | 1.2959 | 2.5474 | 100.3 | 21.01 | 32.131 | 140.7752 | 108.8073 | 128.8829 | 0.007981 | 0.87815 | 0.020389 | 23.2186 | 0.19317 | 0.15202 | 0.070903 | 4.3062 | 4.3062 | 0.001437
2019-03-29 11:42 | 1553874144.8619 | 425 | 0.19537 | 0.011814 | 0.000179 | 101.0191 | 98.5725 | 101.1108 | 21.1013 | 0.004876 | 0.004462 | 144.8338 | 0.25493 | 8.4005 | 0.003665 | 0.004737 | 0.004686 | 6.6e-05 | 0.000871 | 0.29607 | 1.3092 | 2.5474 | 100.3 | 21.01 | 32.134 | 139.6208 | 109.9465 | 129.4108 | 0.007982 | 0.877 | 0.020398 | 23.2583 | 0.2193 | 0.15166 | 0.070493 | 4.2832 | 4.2832 | 0.001441
2019-03-29 11:42 | 1553874145.8606 | 426 | 0.19553 | 0.011783 | 0.000178 | 101.0359 | 98.4879 | 101.5563 | 21.1014 | 0.004951 | 0.004444 | 144.6765 | 0.25488 | 8.4016 | 0.00357 | 0.004827 | 0.004771 | 5.7e-05 | 0.000879 | 0.2959 | 1.3569 | 2.5475 | 100.3 | 21.01 | 32.136 | 141.4893 | 107.6342 | 128.9802 | 0.007982 | 0.87794 | 0.020453 | 23.297 | 0.22847 | 0.15163 | 0.069546 | 4.2375 | 4.2375 | 0.001456
2019-03-29 11:42 | 1553874146.8616 | 427 | 0.19579 | 0.011746 | 0.000178 | 100.9521 | 98.7237 | 101.2932 | 21.1011 | 0.004969 | 0.004459 | 144.5393 | 0.25567 | 8.3996 | 0.0035 | 0.005122 | 0.004987 | 6e-05 | 0.000883 | 0.31741 | 1.3185 | 2.5474 | 100.3 | 21.01 | 32.138 | 141.9192 | 105.9239 | 129.3388 | 0.007983 | 0.87716 | 0.020418 | 23.2772 | 0.19965 | 0.1521 | 0.068078 | 4.1413 | 4.1413 | 0.001495
2019-03-29 11:42 | 1553874147.8617 | 428 | 0.19584 | 0.011691 | 0.000178 | 100.9629 | 98.4937 | 101.4873 | 21.1022 | 0.004971 | 0.004403 | 144.3963 | 0.25495 | 8.4028 | 0.003469 | 0.004595 | 0.004475 | 6.3e-05 | 0.000881 | 0.31653 | 1.3126 | 2.5474 | 100.3 | 21.01 | 32.137 | 141.983 | 105.191 | 127.9743 | 0.007983 | 0.88014 | 0.020472 | 23.26 | 0.22147 | 0.15167 | 0.067834 | 4.1375 | 4.1375 | 0.001492
2019-03-29 11:42 | 1553874148.8607 | 429 | 0.19586 | 0.011614 | 0.000179 | 100.9829 | 98.6483 | 102.3591 | 21.1023 | 0.004893 | 0.004453 | 144.2486 | 0.25416 | 8.4055 | 0.003526 | 0.005005 | 0.004944 | 6.2e-05 | 0.000884 | 0.31996 | 1.3189 | 2.5474 | 100.3 | 21.01 | 32.135 | 140.0484 | 106.5727 | 129.1908 | 0.007982 | 0.87748 | 0.020529 | 23.3951 | 0.20922 | 0.1512 | 0.067782 | 4.1457 | 4.1457 | 0.001484
2019-03-29 11:42 | 1553874149.8607 | 430 | 0.19587 | 0.01153 | 0.000181 | 101.0093 | 98.7653 | 102.9208 | 21.1014 | 0.004916 | 0.004445 | 144.0984 | 0.25505 | 8.4031 | 0.003479 | 0.004736 | 0.004615 | 6.2e-05 | 0.000885 | 0.29336 | 1.3271 | 2.5474 | 100.3 | 21.01 | 32.133 | 140.6202 | 105.4285 | 129.0034 | 0.007982 | 0.87789 | 0.02059 | 23.4537 | 0.20094 | 0.15173 | 0.067863 | 4.1629 | 4.1629 | 0.001483
2019-03-29 11:42 | 1553874150.8611 | 431 | 0.19617 | 0.011456 | 0.000183 | 101.006 | 98.7234 | 102.9795 | 21.1021 | 0.00497 | 0.004423 | 143.923 | 0.25499 | 8.4025 | 0.003442 | 0.004186 | 0.004102 | 6e-05 | 0.000883 | 0.30934 | 1.3227 | 2.5475 | 100.3 | 21.01 | 32.132 | 141.9445 | 104.5199 | 128.4524 | 0.007981 | 0.87909 | 0.02061 | 23.4443 | 0.20445 | 0.15169 | 0.066154 | 4.0624 | 4.0624 | 0.00152
2019-03-29 11:42 | 1553874151.8608 | 432 | 0.19628 | 0.011359 | 0.000184 | 100.968 | 98.5903 | 102.7776 | 21.1024 | 0.004946 | 0.004435 | 143.7647 | 0.25477 | 8.402 | 0.003448 | 0.004444 | 0.00427 | 6.1e-05 | 0.000882 | 0.28064 | 1.3317 | 2.5474 | 100.3 | 21.01 | 32.13 | 141.3699 | 104.6799 | 128.7483 | 0.007981 | 0.87845 | 0.020582 | 23.4299 | 0.21316 | 0.15156 | 0.065627 | 4.0247 | 4.0247 | 0.001533
2019-03-29 11:43 | 1553874152.8619 | 433 | 0.19636 | 0.011262 | 0.000185 | 101.0021 | 98.6268 | 102.764 | 21.102 | 0.004975 | 0.004434 | 143.6126 | 0.2554 | 8.4007 | 0.003421 | 0.00474 | 0.004604 | 6.2e-05 | 0.000883 | 0.30845 | 1.3216 | 2.5475 | 100.3 | 21.01 | 32.126 | 142.0808 | 104.0167 | 128.7304 | 0.00798 | 0.87848 | 0.020581 | 23.4279 | 0.21286 | 0.15193 | 0.065217 | 3.9995 | 3.9995 | 0.001546
2019-03-29 11:43 | 1553874153.8619 | 434 | 0.19642 | 0.011164 | 0.000186 | 101.0323 | 98.8254 | 102.6257 | 21.1021 | 0.004993 | 0.00451 | 143.4613 | 0.25476 | 8.4024 | 0.00345 | 0.004481 | 0.004397 | 6.4e-05 | 0.000878 | 0.30093 | 1.3758 | 2.5476 | 100.3 | 21.01 | 32.125 | 142.5105 | 104.7181 | 130.613 | 0.00798 | 0.87439 | 0.020519 | 23.4669 | 0.19755 | 0.15156 | 0.065005 | 3.9745 | 3.9745 | 0.001552
2019-03-29 11:43 | 1553874154.8613 | 435 | 0.19656 | 0.011073 | 0.000186 | 100.965 | 98.7503 | 102.2999 | 21.1025 | 0.004993 | 0.004469 | 143.2716 | 0.25496 | 8.4038 | 0.003317 | 0.00488 | 0.004658 | 6.2e-05 | 0.000876 | 0.32012 | 1.3514 | 2.5473 | 100.3 | 21.01 | 32.123 | 142.5263 | 101.5022 | 129.6074 | 0.007979 | 0.87657 | 0.020512 | 23.4004 | 0.19838 | 0.15168 | 0.064295 | 3.9299 | 3.9299 | 0.001571
2019-03-29 11:43 | 1553874155.861 | 436 | 0.1966 | 0.010979 | 0.000187 | 100.9494 | 98.6657 | 102.015 | 21.1026 | 0.004986 | 0.004467 | 143.1253 | 0.25627 | 8.4017 | 0.003299 | 0.005096 | 0.004975 | 6.1e-05 | 0.000863 | 0.30431 | 1.3016 | 2.5475 | 100.3 | 21.01 | 32.12 | 142.3465 | 101.0644 | 129.5461 | 0.007978 | 0.8767 | 0.020485 | 23.366 | 0.20467 | 0.15245 | 0.064124 | 3.9143 | 3.9143 | 0.001585
2019-03-29 11:43 | 1553874156.8611 | 437 | 0.19658 | 0.010902 | 0.000185 | 100.9578 | 98.7466 | 102.1131 | 21.1037 | 0.00503 | 0.004508 | 142.9667 | 0.25529 | 8.4033 | 0.003378 | 0.004995 | 0.004931 | 6.3e-05 | 0.00086 | 0.28998 | 1.3071 | 2.5473 | 100.3 | 21.01 | 32.116 | 143.4279 | 102.9738 | 130.5652 | 0.007977 | 0.87449 | 0.020469 | 23.4068 | 0.19808 | 0.15187 | 0.064388 | 3.9274 | 3.9274 | 0.001574
2019-03-29 11:43 | 1553874157.862 | 438 | 0.19636 | 0.010831 | 0.000182 | 100.9826 | 98.7022 | 101.7148 | 21.1033 | 0.005026 | 0.004472 | 142.8018 | 0.25516 | 8.4027 | 0.003406 | 0.004683 | 0.004557 | 5.8e-05 | 0.000863 | 0.30829 | 1.3309 | 2.5473 | 100.3 | 21.01 | 32.109 | 143.3295 | 103.6689 | 129.6616 | 0.007976 | 0.87645 | 0.020452 | 23.3349 | 0.2043 | 0.1518 | 0.065733 | 4.0059 | 4.0059 | 0.001542
2019-03-29 11:43 | 1553874158.861 | 439 | 0.19609 | 0.010781 | 0.000179 | 100.947 | 98.378 | 101.6617 | 21.103 | 0.004996 | 0.004459 | 142.6552 | 0.25579 | 8.4037 | 0.003415 | 0.004584 | 0.004554 | 6.4e-05 | 0.000866 | 0.31902 | 1.3209 | 2.5474 | 100.3 | 21.01 | 32.105 | 142.6037 | 103.8815 | 129.3566 | 0.007975 | 0.87712 | 0.020454 | 23.32 | 0.23058 | 0.15217 | 0.067398 | 4.1075 | 4.1075 | 0.001508
2019-03-29 11:43 | 1553874159.861 | 440 | 0.19601 | 0.010765 | 0.000178 | 101.0353 | 98.6097 | 101.4826 | 21.1033 | 0.005067 | 0.00444 | 142.49 | 0.25503 | 8.4024 | 0.003376 | 0.005101 | 0.004932 | 6.5e-05 | 0.000873 | 0.31741 | 1.3418 | 2.5474 | 100.3 | 21.01 | 32.101 | 144.3459 | 102.9286 | 128.8777 | 0.007974 | 0.87816 | 0.020449 | 23.2856 | 0.21735 | 0.15172 | 0.067944 | 4.1395 | 4.1395 | 0.001492
2019-03-29 11:43 | 1553874160.8601 | 441 | 0.19597 | 0.010783 | 0.000177 | 101.0509 | 98.6237 | 101.5897 | 21.1032 | 0.005048 | 0.004432 | 142.3506 | 0.25506 | 8.4025 | 0.003367 | 0.00413 | 0.004079 | 6.1e-05 | 0.000875 | 0.30541 | 1.3727 | 2.5474 | 100.3 | 21.01 | 32.092 | 143.8942 | 102.7091 | 128.6861 | 0.007971 | 0.87858 | 0.020464 | 23.2923 | 0.21746 | 0.15174 | 0.06812 | 4.1534 | 4.1534 | 0.001487
2019-03-29 11:43 | 1553874161.8601 | 442 | 0.1959 | 0.010832 | 0.000176 | 101.0302 | 98.4578 | 101.8265 | 21.1039 | 0.005005 | 0.004471 | 142.2023 | 0.25488 | 8.4028 | 0.003348 | 0.005126 | 0.004952 | 6.4e-05 | 0.000874 | 0.31974 | 1.3467 | 2.5474 | 100.3 | 21.01 | 32.088 | 142.8094 | 102.2629 | 129.639 | 0.00797 | 0.8765 | 0.020464 | 23.3471 | 0.23069 | 0.15163 | 0.068483 | 4.1754 | 4.1754 | 0.001478
2019-03-29 11:43 | 1553874162.8603 | 443 | 0.19606 | 0.010882 | 0.000178 | 100.994 | 98.5493 | 102.2539 | 21.1034 | 0.005052 | 0.004459 | 142.0631 | 0.25456 | 8.4026 | 0.003395 | 0.004942 | 0.004928 | 6.4e-05 | 0.000873 | 0.32084 | 1.3504 | 2.5475 | 100.3 | 21.01 | 32.08 | 143.9947 | 103.3968 | 129.3578 | 0.007968 | 0.87711 | 0.020514 | 23.3879 | 0.21918 | 0.15144 | 0.067491 | 4.1252 | 4.1252 | 0.001494
2019-03-29 11:43 | 1553874163.8612 | 444 | 0.19606 | 0.010928 | 0.00018 | 100.9886 | 98.4776 | 102.2579 | 21.1033 | 0.00498 | 0.004421 | 141.9062 | 0.25523 | 8.4021 | 0.003335 | 0.004828 | 0.004645 | 6.3e-05 | 0.000876 | 0.28423 | 1.3481 | 2.5474 | 100.3 | 21.02 | 32.077 | 142.2033 | 101.9357 | 128.4081 | 0.007973 | 0.87919 | 0.020538 | 23.3607 | 0.22521 | 0.15184 | 0.067445 | 4.1272 | 4.1272 | 0.001497
2019-03-29 11:43 | 1553874164.8603 | 445 | 0.19635 | 0.010953 | 0.000182 | 101.0578 | 98.4493 | 102.4002 | 21.1045 | 0.00505 | 0.004479 | 141.7655 | 0.25532 | 8.4011 | 0.003349 | 0.00475 | 0.004693 | 6.3e-05 | 0.000875 | 0.29319 | 1.3355 | 2.5474 | 100.3 | 21.02 | 32.075 | 143.9414 | 102.2848 | 129.8511 | 0.007972 | 0.87604 | 0.020516 | 23.4189 | 0.2339 | 0.15189 | 0.065698 | 4.0163 | 4.0163 | 0.001539
2019-03-29 11:43 | 1553874165.8612 | 446 | 0.1968 | 0.010949 | 0.000183 | 101.0263 | 98.4872 | 102.4757 | 21.1037 | 0.005076 | 0.004444 | 141.6083 | 0.25595 | 8.4033 | 0.003355 | 0.004793 | 0.004688 | 5.9e-05 | 0.000877 | 0.30613 | 1.312 | 2.5474 | 100.3 | 21.02 | 32.07 | 144.5906 | 102.4229 | 128.9786 | 0.007971 | 0.87794 | 0.020546 | 23.4022 | 0.22768 | 0.15226 | 0.063001 | 3.8576 | 3.8576 | 0.001606
2019-03-29 11:43 | 1553874166.8603 | 447 | 0.19691 | 0.010916 | 0.000184 | 101.1074 | 98.2609 | 102.6856 | 21.1041 | 0.005062 | 0.004507 | 141.4633 | 0.25702 | 8.4013 | 0.00331 | 0.004717 | 0.004732 | 6.1e-05 | 0.000876 | 0.31996 | 1.3348 | 2.5473 | 100.3 | 21.02 | 32.068 | 144.244 | 101.3359 | 130.5241 | 0.00797 | 0.87458 | 0.020527 | 23.4711 | 0.25543 | 0.1529 | 0.062398 | 3.8174 | 3.8174 | 0.00163
2019-03-29 11:43 | 1553874167.8603 | 448 | 0.19692 | 0.010845 | 0.000182 | 100.9928 | 98.3903 | 103.2217 | 21.1041 | 0.005014 | 0.004546 | 141.3234 | 0.2567 | 8.405 | 0.003263 | 0.004567 | 0.00448 | 6.3e-05 | 0.000874 | 0.29214 | 1.336 | 2.5474 | 100.3 | 21.02 | 32.069 | 143.0401 | 100.1953 | 131.4915 | 0.007971 | 0.87249 | 0.020556 | 23.5605 | 0.23352 | 0.15271 | 0.062441 | 3.8255 | 3.8255 | 0.001625
2019-03-29 11:43 | 1553874168.8614 | 449 | 0.19691 | 0.010754 | 0.000181 | 101.0377 | 98.435 | 103.1555 | 21.1042 | 0.005025 | 0.004469 | 141.1676 | 0.25605 | 8.4023 | 0.003305 | 0.004116 | 0.003961 | 6.3e-05 | 0.000878 | 0.30862 | 1.3463 | 2.5474 | 100.3 | 21.02 | 32.072 | 143.3046 | 101.2157 | 129.5995 | 0.007971 | 0.87659 | 0.020598 | 23.4978 | 0.23343 | 0.15232 | 0.062591 | 3.8425 | 3.8425 | 0.001613
2019-03-29 11:43 | 1553874169.8604 | 450 | 0.19696 | 0.010663 | 0.000179 | 101.0224 | 98.5309 | 103.0891 | 21.104 | 0.005019 | 0.004492 | 141.0325 | 0.25514 | 8.4044 | 0.003276 | 0.004894 | 0.004723 | 5.9e-05 | 0.000881 | 0.30685 | 1.3443 | 2.5475 | 100.3 | 21.02 | 32.074 | 143.1799 | 100.5108 | 130.1666 | 0.007972 | 0.87536 | 0.020577 | 23.5068 | 0.22337 | 0.15178 | 0.062374 | 3.8254 | 3.8254 | 0.001615
2019-03-29 11:43 | 1553874170.8614 | 451 | 0.197 | 0.010572 | 0.000179 | 100.9988 | 98.6906 | 103.2672 | 21.104 | 0.005062 | 0.004501 | 140.8851 | 0.25576 | 8.4033 | 0.003328 | 0.005008 | 0.004891 | 6.4e-05 | 0.000885 | 0.29247 | 1.3109 | 2.5474 | 100.3 | 21.02 | 32.082 | 144.2238 | 101.7675 | 130.3906 | 0.007974 | 0.87487 | 0.020589 | 23.5336 | 0.20679 | 0.15215 | 0.062251 | 3.8201 | 3.8201 | 0.001621
2019-03-29 11:43 | 1553874171.8605 | 452 | 0.19701 | 0.010502 | 0.000179 | 101.0631 | 98.7571 | 103.1696 | 21.104 | 0.005068 | 0.004452 | 140.7224 | 0.25544 | 8.401 | 0.003298 | 0.003562 | 0.003455 | 6.2e-05 | 0.000889 | 0.30879 | 1.3323 | 2.5475 | 100.3 | 21.02 | 32.086 | 144.3907 | 101.042 | 129.1774 | 0.007975 | 0.87751 | 0.02061 | 23.4871 | 0.20646 | 0.15196 | 0.062307 | 3.8275 | 3.8275 | 0.001616
2019-03-29 11:43 | 1553874172.8605 | 453 | 0.19709 | 0.010454 | 0.00018 | 101.023 | 98.8001 | 102.9189 | 21.1042 | 0.005044 | 0.004475 | 140.5645 | 0.25583 | 8.4008 | 0.003365 | 0.004492 | 0.00446 | 6.2e-05 | 0.000892 | 0.2907 | 1.3091 | 2.5474 | 100.3 | 21.02 | 32.092 | 143.7875 | 102.6575 | 129.7335 | 0.007976 | 0.8763 | 0.020571 | 23.4748 | 0.19901 | 0.15219 | 0.061839 | 3.7916 | 3.7916 | 0.001634
2019-03-29 11:43 | 1553874173.8615 | 454 | 0.19726 | 0.010404 | 0.000182 | 101.1182 | 98.8466 | 103.3027 | 21.1046 | 0.005066 | 0.004451 | 140.4172 | 0.25544 | 8.4028 | 0.00328 | 0.004608 | 0.004515 | 6.3e-05 | 0.000894 | 0.29214 | 1.3611 | 2.5474 | 100.3 | 21.02 | 32.103 | 144.3319 | 100.603 | 129.1604 | 0.007979 | 0.87755 | 0.020624 | 23.5018 | 0.20323 | 0.15196 | 0.060925 | 3.7453 | 3.7453 | 0.001651
2019-03-29 11:43 | 1553874174.8615 | 455 | 0.19732 | 0.010352 | 0.000183 | 101.1012 | 98.6305 | 103.3184 | 21.1037 | 0.004985 | 0.00445 | 140.298 | 0.257 | 8.4022 | 0.003266 | 0.004362 | 0.004199 | 6.1e-05 | 0.000893 | 0.30917 | 1.382 | 2.5475 | 100.3 | 21.02 | 32.108 | 142.3258 | 100.2616 | 129.1276 | 0.00798 | 0.87762 | 0.020626 | 23.5026 | 0.22129 | 0.15289 | 0.060634 | 3.7278 | 3.7278 | 0.001669
2019-03-29 11:43 | 1553874175.8605 | 456 | 0.19752 | 0.010285 | 0.000184 | 101.0949 | 98.479 | 103.8223 | 21.1051 | 0.005051 | 0.004486 | 140.1639 | 0.2552 | 8.4011 | 0.003301 | 0.004526 | 0.004492 | 5.8e-05 | 0.000886 | 0.26017 | 1.318 | 2.5474 | 100.3 | 21.02 | 32.116 | 143.9675 | 101.1256 | 130.024 | 0.007982 | 0.87567 | 0.020653 | 23.5861 | 0.23449 | 0.15182 | 0.059482 | 3.662 | 3.662 | 0.001687
2019-03-29 11:43 | 1553874176.8606 | 457 | 0.19768 | 0.010193 | 0.000185 | 101.1455 | 98.5305 | 103.5954 | 21.1048 | 0.00505 | 0.004471 | 140.0196 | 0.25584 | 8.4024 | 0.003303 | 0.005098 | 0.005048 | 6.1e-05 | 0.000886 | 0.28982 | 1.3556 | 2.5475 | 100.3 | 21.02 | 32.121 | 143.9392 | 101.1575 | 129.6584 | 0.007984 | 0.87646 | 0.02064 | 23.5496 | 0.23429 | 0.1522 | 0.058676 | 3.6102 | 3.6102 | 0.001716
2019-03-29 11:43 | 1553874177.8606 | 458 | 0.19769 | 0.010092 | 0.000185 | 101.0667 | 98.3346 | 103.3707 | 21.1054 | 0.004995 | 0.004508 | 139.8834 | 0.25565 | 8.404 | 0.003293 | 0.005173 | 0.005027 | 6.3e-05 | 0.000883 | 0.30685 | 1.3741 | 2.5475 | 100.3 | 21.02 | 32.126 | 142.5732 | 100.9153 | 130.5629 | 0.007985 | 0.8745 | 0.020595 | 23.5504 | 0.24512 | 0.15209 | 0.058726 | 3.6053 | 3.6053 | 0.001717
2019-03-29 11:43 | 1553874178.8597 | 459 | 0.19765 | 0.009988 | 0.000185 | 101.0985 | 98.5845 | 103.1989 | 21.1054 | 0.005004 | 0.004512 | 139.75 | 0.25583 | 8.4033 | 0.003398 | 0.004866 | 0.004776 | 6e-05 | 0.00088 | 0.30055 | 1.3635 | 2.5474 | 100.3 | 21.02 | 32.133 | 142.7884 | 103.4728 | 130.6491 | 0.007987 | 0.87431 | 0.020575 | 23.5334 | 0.22524 | 0.15219 | 0.059105 | 3.625 | 3.625 | 0.001709
2019-03-29 11:43 | 1553874179.8597 | 460 | 0.19757 | 0.009909 | 0.000185 | 101.1718 | 98.5162 | 103.5172 | 21.1053 | 0.004995 | 0.004505 | 139.604 | 0.256 | 8.4023 | 0.003359 | 0.004594 | 0.004395 | 6.2e-05 | 0.000877 | 0.31006 | 1.3568 | 2.5474 | 100.3 | 21.02 | 32.137 | 142.5821 | 102.5323 | 130.491 | 0.007988 | 0.87465 | 0.020611 | 23.565 | 0.23792 | 0.1523 | 0.05963 | 3.6635 | 3.6635 | 0.001692
2019-03-29 11:43 | 1553874180.861 | 461 | 0.19746 | 0.009849 | 0.000185 | 101.1128 | 98.1974 | 103.3109 | 21.1059 | 0.005011 | 0.004452 | 139.4697 | 0.25643 | 8.4014 | 0.003454 | 0.005116 | 0.004979 | 6.2e-05 | 0.000875 | 0.29767 | 1.36 | 2.5475 | 100.3 | 21.01 | 32.141 | 142.9601 | 104.828 | 129.184 | 0.007984 | 0.87749 | 0.020624 | 23.5034 | 0.26169 | 0.15255 | 0.060357 | 3.7103 | 3.7103 | 0.001673
2019-03-29 11:43 | 1553874181.8611 | 462 | 0.19736 | 0.009827 | 0.000185 | 101.1257 | 98.2155 | 103.6533 | 21.1054 | 0.005022 | 0.004457 | 139.3464 | 0.25584 | 8.4043 | 0.003572 | 0.004731 | 0.004626 | 6.1e-05 | 0.000873 | 0.2907 | 1.3511 | 2.5475 | 100.3 | 21.01 | 32.147 | 143.2463 | 107.6774 | 129.3103 | 0.007985 | 0.87722 | 0.020655 | 23.546 | 0.26118 | 0.1522 | 0.060993 | 3.7549 | 3.7549 | 0.00165
2019-03-29 11:43 | 1553874182.8608 | 463 | 0.19738 | 0.009825 | 0.000186 | 101.117 | 98.2841 | 103.8104 | 21.1061 | 0.005027 | 0.004448 | 139.2188 | 0.2559 | 8.4018 | 0.003576 | 0.004904 | 0.004803 | 6.2e-05 | 0.000874 | 0.28943 | 1.3262 | 2.5475 | 100.3 | 21.01 | 32.149 | 143.358 | 107.7821 | 129.0933 | 0.007986 | 0.87769 | 0.020676 | 23.5575 | 0.25417 | 0.15223 | 0.060895 | 3.7527 | 3.7527 | 0.001651
2019-03-29 11:43 | 1553874183.8598 | 464 | 0.19739 | 0.009837 | 0.000187 | 101.1272 | 98.0118 | 103.7005 | 21.1064 | 0.005004 | 0.004448 | 139.1017 | 0.25573 | 8.4024 | 0.003574 | 0.005149 | 0.005021 | 6.3e-05 | 0.000867 | 0.28711 | 1.3033 | 2.5475 | 100.3 | 21.01 | 32.152 | 142.8111 | 107.7381 | 129.0921 | 0.007986 | 0.87769 | 0.020665 | 23.545 | 0.27989 | 0.15213 | 0.060801 | 3.7449 | 3.7449 | 0.001653
2019-03-29 11:43 | 1553874184.8598 | 465 | 0.19741 | 0.009865 | 0.000187 | 101.1236 | 98.2032 | 103.7308 | 21.106 | 0.005013 | 0.004487 | 138.9522 | 0.25559 | 8.4035 | 0.003648 | 0.004479 | 0.004349 | 6.1e-05 | 0.000868 | 0.30381 | 1.3166 | 2.5473 | 100.3 | 21.01 | 32.152 | 143.0316 | 109.541 | 130.0296 | 0.007986 | 0.87565 | 0.020644 | 23.5758 | 0.26211 | 0.15205 | 0.060648 | 3.7317 | 3.7317 | 0.001658
2019-03-29 11:43 | 1553874185.8603 | 466 | 0.19745 | 0.009877 | 0.000186 | 101.0964 | 98.138 | 103.7418 | 21.1066 | 0.00508 | 0.00449 | 138.811 | 0.25648 | 8.4038 | 0.003666 | 0.005114 | 0.005005 | 6.5e-05 | 0.000862 | 0.29374 | 1.3496 | 2.5474 | 100.3 | 21.01 | 32.154 | 144.6787 | 109.9784 | 130.1165 | 0.007987 | 0.87546 | 0.020643 | 23.5796 | 0.26565 | 0.15258 | 0.060427 | 3.718 | 3.718 | 0.00167
2019-03-29 11:43 | 1553874186.8609 | 467 | 0.19749 | 0.009883 | 0.000184 | 101.1068 | 98.3954 | 104.0859 | 21.1059 | 0.004991 | 0.004463 | 138.6761 | 0.25567 | 8.403 | 0.003662 | 0.005208 | 0.005172 | 6.4e-05 | 0.000863 | 0.31061 | 1.3569 | 2.5474 | 100.3 | 21.01 | 32.156 | 142.4668 | 109.8625 | 129.4409 | 0.007987 | 0.87693 | 0.020695 | 23.5989 | 0.24315 | 0.1521 | 0.060149 | 3.7103 | 3.7103 | 0.001668
2019-03-29 11:43 | 1553874187.86 | 468 | 0.19753 | 0.009878 | 0.000183 | 101.0994 | 98.4469 | 104.2146 | 21.1057 | 0.005003 | 0.004431 | 138.5235 | 0.2557 | 8.4025 | 0.003635 | 0.005381 | 0.005167 | 6.4e-05 | 0.000859 | 0.29103 | 1.363 | 2.5475 | 100.3 | 21.01 | 32.156 | 142.7775 | 109.2096 | 128.6553 | 0.007987 | 0.87865 | 0.020728 | 23.5904 | 0.2378 | 0.15212 | 0.059943 | 3.7036 | 3.7036 | 0.001672
2019-03-29 11:43 | 1553874188.86 | 469 | 0.19758 | 0.009872 | 0.000182 | 101.1071 | 98.3868 | 104.1869 | 21.1063 | 0.005076 | 0.004454 | 138.3824 | 0.25467 | 8.4037 | 0.00366 | 0.00499 | 0.004961 | 6.5e-05 | 0.00086 | 0.29247 | 1.3279 | 2.5473 | 100.3 | 21.01 | 32.155 | 144.5739 | 109.8283 | 129.2235 | 0.007987 | 0.87741 | 0.02071 | 23.604 | 0.24395 | 0.15151 | 0.059662 | 3.6833 | 3.6833 | 0.001674
2019-03-29 11:43 | 1553874189.8604 | 470 | 0.19767 | 0.009859 | 0.00018 | 101.1236 | 97.9862 | 103.8695 | 21.1073 | 0.005005 | 0.004449 | 138.261 | 0.25554 | 8.403 | 0.003667 | 0.005188 | 0.005085 | 6.6e-05 | 0.000861 | 0.26521 | 1.3718 | 2.5474 | 100.3 | 21.01 | 32.154 | 142.8242 | 109.9854 | 129.1105 | 0.007987 | 0.87765 | 0.020682 | 23.5647 | 0.28191 | 0.15202 | 0.05915 | 3.6468 | 3.6468 | 0.001697
2019-03-29 11:43 | 1553874190.8601 | 471 | 0.1978 | 0.009841 | 0.000179 | 101.1254 | 98.1514 | 104.434 | 21.1072 | 0.005027 | 0.004468 | 138.1127 | 0.2554 | 8.4001 | 0.003716 | 0.005386 | 0.005283 | 6.5e-05 | 0.000865 | 0.29192 | 1.3584 | 2.5475 | 100.3 | 21.01 | 32.155 | 143.3814 | 111.1806 | 129.569 | 0.007987 | 0.87665 | 0.020726 | 23.6421 | 0.267 | 0.15194 | 0.058364 | 3.6063 | 3.6063 | 0.001715
2019-03-29 11:43 | 1553874191.8591 | 472 | 0.19799 | 0.009802 | 0.000177 | 101.0907 | 98.3913 | 104.6427 | 21.1064 | 0.004986 | 0.004409 | 137.9821 | 0.25465 | 8.402 | 0.003685 | 0.005218 | 0.005134 | 6.1e-05 | 0.000871 | 0.29662 | 1.3892 | 2.5474 | 100.3 | 21.01 | 32.153 | 142.3606 | 110.4223 | 128.1194 | 0.007987 | 0.87982 | 0.020784 | 23.6231 | 0.24209 | 0.15149 | 0.057305 | 3.5511 | 3.5511 | 0.001736
2019-03-29 11:43 | 1553874192.8601 | 473 | 0.19806 | 0.009753 | 0.000175 | 101.0682 | 98.0793 | 105.1107 | 21.1066 | 0.004994 | 0.004434 | 137.8467 | 0.25643 | 8.4025 | 0.003712 | 0.005518 | 0.005418 | 6.3e-05 | 0.000872 | 0.32101 | 1.3234 | 2.5474 | 100.3 | 21.01 | 32.153 | 142.5408 | 111.0964 | 128.7357 | 0.007987 | 0.87847 | 0.020815 | 23.694 | 0.26851 | 0.15255 | 0.056907 | 3.5318 | 3.5318 | 0.001758
2019-03-29 11:43 | 1553874193.8605 | 474 | 0.19807 | 0.009695 | 0.000173 | 101.0868 | 98.1763 | 105.4192 | 21.1064 | 0.005007 | 0.004411 | 137.715 | 0.25503 | 8.4 | 0.003685 | 0.005208 | 0.005138 | 6.2e-05 | 0.000872 | 0.31116 | 1.3425 | 2.5474 | 100.3 | 21.01 | 32.153 | 142.8641 | 110.4279 | 128.1783 | 0.007987 | 0.87969 | 0.02086 | 23.7123 | 0.26131 | 0.15172 | 0.056954 | 3.5425 | 3.5425 | 0.001743
2019-03-29 11:43 | 1553874194.8598 | 475 | 0.19811 | 0.009625 | 0.000172 | 101.1353 | 98.6527 | 106.0572 | 21.1065 | 0.005005 | 0.004419 | 137.599 | 0.25577 | 8.403 | 0.003659 | 0.005084 | 0.004925 | 6.3e-05 | 0.000874 | 0.30829 | 1.3415 | 2.5474 | 100.3 | 21.01 | 32.152 | 142.8145 | 109.7883 | 128.3698 | 0.007986 | 0.87927 | 0.020918 | 23.7896 | 0.2223 | 0.15216 | 0.056774 | 3.5412 | 3.5412 | 0.001749
2019-03-29 11:43 | 1553874195.8593 | 476 | 0.19819 | 0.009563 | 0.000172 | 101.094 | 98.4097 | 106.2923 | 21.1077 | 0.005004 | 0.004373 | 137.4863 | 0.25598 | 8.4019 | 0.00365 | 0.0052 | 0.005082 | 6.2e-05 | 0.000875 | 0.30541 | 1.3344 | 2.5474 | 100.3 | 21.01 | 32.153 | 142.802 | 109.5804 | 127.242 | 0.007987 | 0.88175 | 0.02097 | 23.7825 | 0.24071 | 0.15228 | 0.056409 | 3.5273 | 3.5273 | 0.001757
2019-03-29 11:43 | 1553874196.8596 | 477 | 0.19837 | 0.009505 | 0.000175 | 101.0715 | 98.2828 | 106.6747 | 21.1071 | 0.004941 | 0.004367 | 137.3694 | 0.25515 | 8.4033 | 0.003617 | 0.004709 | 0.004588 | 6.1e-05 | 0.00088 | 0.31349 | 1.329 | 2.5474 | 100.3 | 21.01 | 32.154 | 141.2479 | 108.779 | 127.0909 | 0.007987 | 0.88208 | 0.021012 | 23.8207 | 0.25027 | 0.15179 | 0.055356 | 3.4684 | 3.4684 | 0.001781
2019-03-29 11:43 | 1553874197.8606 | 478 | 0.19882 | 0.00942 | 0.000177 | 101.1449 | 98.2118 | 107.1435 | 21.1077 | 0.004958 | 0.004351 | 137.2655 | 0.25544 | 8.4018 | 0.003592 | 0.004767 | 0.004684 | 6.5e-05 | 0.000883 | 0.33157 | 1.3276 | 2.5475 | 100.3 | 21.01 | 32.155 | 141.6634 | 108.1847 | 126.6924 | 0.007987 | 0.88296 | 0.021068 | 23.8611 | 0.26322 | 0.15196 | 0.052793 | 3.3171 | 3.3171 | 0.001864
2019-03-29 11:43 | 1553874198.8603 | 479 | 0.19905 | 0.009319 | 0.00018 | 101.0808 | 97.766 | 107.6225 | 21.1083 | 0.004931 | 0.004339 | 137.1654 | 0.25575 | 8.4021 | 0.003574 | 0.004783 | 0.004703 | 6.3e-05 | 0.000884 | 0.31852 | 1.3152 | 2.5474 | 100.3 | 21.01 | 32.156 | 141.0052 | 107.7315 | 126.412 | 0.007987 | 0.88358 | 0.021123 | 23.906 | 0.29824 | 0.15214 | 0.051535 | 3.2465 | 3.2465 | 0.001907
2019-03-29 11:43 | 1553874199.8607 | 480 | 0.19921 | 0.009185 | 0.000182 | 101.1281 | 97.6227 | 108.1741 | 21.1082 | 0.004934 | 0.004372 | 137.0543 | 0.25588 | 8.405 | 0.003538 | 0.006149 | 0.006023 | 6.2e-05 | 0.000882 | 0.27455 | 1.3053 | 2.5474 | 100.3 | 21.01 | 32.157 | 141.066 | 106.8589 | 127.2082 | 0.007988 | 0.88182 | 0.021156 | 23.9911 | 0.31555 | 0.15222 | 0.05078 | 3.2039 | 3.2039 | 0.001934
2019-03-29 11:43 | 1553874200.8607 | 481 | 0.1995 | 0.009034 | 0.000185 | 101.1493 | 97.45 | 108.4138 | 21.108 | 0.004848 | 0.004324 | 136.9746 | 0.25485 | 8.4042 | 0.003519 | 0.005049 | 0.004899 | 6.1e-05 | 0.00088 | 0.30071 | 1.3436 | 2.5476 | 100.3 | 21.01 | 32.157 | 138.9498 | 106.4129 | 126.0336 | 0.007988 | 0.88442 | 0.02121 | 23.9824 | 0.33326 | 0.15161 | 0.049203 | 3.1124 | 3.1124 | 0.001983
2019-03-29 11:43 | 1553874201.8604 | 482 | 0.19969 | 0.008881 | 0.000188 | 101.1889 | 97.7657 | 108.7379 | 21.1079 | 0.004862 | 0.004331 | 136.8789 | 0.25495 | 8.4039 | 0.00349 | 0.004784 | 0.00482 | 5.9e-05 | 0.000879 | 0.28296 | 1.3638 | 2.5473 | 100.3 | 21 | 32.159 | 139.2929 | 105.7055 | 126.202 | 0.007983 | 0.88405 | 0.021238 | 24.0233 | 0.30783 | 0.15167 | 0.048275 | 3.0576 | 3.0576 | 0.002019
2019-03-29 11:43 | 1553874202.8596 | 483 | 0.19998 | 0.008731 | 0.000188 | 101.1071 | 97.7002 | 108.9891 | 21.1095 | 0.004862 | 0.004328 | 136.7751 | 0.2553 | 8.4033 | 0.003433 | 0.004757 | 0.004692 | 5.9e-05 | 0.000883 | 0.28816 | 1.3374 | 2.5473 | 100.3 | 21 | 32.161 | 139.3002 | 104.3085 | 126.144 | 0.007984 | 0.88417 | 0.021264 | 24.0493 | 0.3066 | 0.15188 | 0.046707 | 2.9621 | 2.9621 | 0.002087
2019-03-29 11:43 | 1553874203.8605 | 484 | 0.2003 | 0.008582 | 0.000189 | 101.1308 | 97.5852 | 109.1726 | 21.1078 | 0.004867 | 0.004273 | 136.661 | 0.25598 | 8.4038 | 0.003496 | 0.004602 | 0.004445 | 6e-05 | 0.000883 | 0.28672 | 1.3761 | 2.5475 | 100.3 | 21 | 32.164 | 139.4147 | 105.8369 | 124.799 | 0.007984 | 0.88716 | 0.021318 | 24.0289 | 0.31922 | 0.15228 | 0.044979 | 2.8598 | 2.8598 | 0.002167
2019-03-29 11:43 | 1553874204.8605 | 485 | 0.20046 | 0.008417 | 0.000189 | 101.0634 | 97.3674 | 109.3172 | 21.1081 | 0.004778 | 0.004233 | 136.5501 | 0.25533 | 8.4043 | 0.003377 | 0.005496 | 0.005439 | 6.2e-05 | 0.000886 | 0.32444 | 1.3609 | 2.5474 | 100.3 | 21 | 32.166 | 137.2048 | 102.9638 | 123.8062 | 0.007985 | 0.88938 | 0.021358 | 24.0148 | 0.33325 | 0.1519 | 0.044216 | 2.8167 | 2.8167 | 0.002195
2019-03-29 11:43 | 1553874205.8596 | 486 | 0.20066 | 0.008258 | 0.000189 | 101.141 | 97.3497 | 109.6224 | 21.1083 | 0.004799 | 0.004265 | 136.457 | 0.25529 | 8.402 | 0.00338 | 0.004218 | 0.004097 | 6.2e-05 | 0.000889 | 0.28313 | 1.3392 | 2.5474 | 100.3 | 21 | 32.168 | 137.7335 | 103.0262 | 124.5999 | 0.007985 | 0.88761 | 0.021367 | 24.0724 | 0.34173 | 0.15187 | 0.043182 | 2.752 | 2.752 | 0.002246
2019-03-29 11:43 | 1553874206.8596 | 487 | 0.2009 | 0.008089 | 0.000189 | 101.1634 | 97.551 | 109.8017 | 21.1082 | 0.004814 | 0.004234 | 136.3526 | 0.25611 | 8.4025 | 0.003387 | 0.004756 | 0.004641 | 5.9e-05 | 0.000896 | 0.32587 | 1.3275 | 2.5473 | 100.3 | 21 | 32.172 | 138.0965 | 103.2061 | 123.8237 | 0.007986 | 0.88934 | 0.021405 | 24.0685 | 0.32524 | 0.15236 | 0.04196 | 2.679 | 2.679 | 0.002315
2019-03-29 11:43 | 1553874207.86 | 488 | 0.2011 | 0.007911 | 0.000188 | 101.0967 | 97.6493 | 110.028 | 21.1086 | 0.004775 | 0.004174 | 136.2674 | 0.256 | 8.404 | 0.003301 | 0.005179 | 0.005026 | 5.8e-05 | 0.000897 | 0.30237 | 1.3422 | 2.5474 | 100.3 | 21 | 32.173 | 137.1507 | 101.121 | 122.3692 | 0.007987 | 0.89261 | 0.021467 | 24.0491 | 0.31033 | 0.1523 | 0.041017 | 2.6264 | 2.6264 | 0.00236
2019-03-29 11:43 | 1553874208.8588 | 489 | 0.20111 | 0.007736 | 0.000187 | 101.0952 | 97.5053 | 110.4615 | 21.1085 | 0.004756 | 0.004195 | 136.1769 | 0.2556 | 8.4006 | 0.003294 | 0.004309 | 0.004215 | 5.8e-05 | 0.000893 | 0.31277 | 1.3384 | 2.5474 | 100.3 | 21 | 32.176 | 136.6693 | 100.9569 | 122.8799 | 0.007987 | 0.89146 | 0.021495 | 24.112 | 0.3234 | 0.15206 | 0.041169 | 2.6396 | 2.6396 | 0.002345
2019-03-29 11:43 | 1553874209.8587 | 490 | 0.20114 | 0.007567 | 0.000188 | 101.1374 | 97.194 | 110.6343 | 21.109 | 0.004771 | 0.004243 | 136.0899 | 0.25569 | 8.4025 | 0.00331 | 0.005241 | 0.005051 | 6.1e-05 | 0.000896 | 0.3115 | 1.3444 | 2.5476 | 100.3 | 21 | 32.178 | 137.0502 | 101.3389 | 124.0504 | 0.007988 | 0.88883 | 0.02148 | 24.1665 | 0.35573 | 0.15211 | 0.041178 | 2.6383 | 2.6383 | 0.002346
2019-03-29 11:43 | 1553874210.8591 | 491 | 0.20134 | 0.007426 | 0.000189 | 101.1266 | 97.5036 | 110.8988 | 21.1092 | 0.00475 | 0.00415 | 136.0054 | 0.25523 | 8.4042 | 0.003239 | 0.004382 | 0.004245 | 5.8e-05 | 0.000891 | 0.30989 | 1.3141 | 2.5474 | 100.3 | 21 | 32.18 | 136.5187 | 99.6174 | 121.7781 | 0.007988 | 0.89395 | 0.021567 | 24.1261 | 0.32633 | 0.15184 | 0.040121 | 2.581 | 2.581 | 0.002394
2019-03-29 11:44 | 1553874211.8601 | 492 | 0.20156 | 0.007301 | 0.000191 | 101.0733 | 97.4295 | 110.8955 | 21.1088 | 0.004747 | 0.00413 | 135.91 | 0.25648 | 8.4037 | 0.003279 | 0.004546 | 0.004446 | 5.9e-05 | 0.00089 | 0.28584 | 1.3499 | 2.5472 | 100.3 | 21 | 32.181 | 136.4598 | 100.5805 | 121.279 | 0.007989 | 0.89508 | 0.021581 | 24.1104 | 0.32842 | 0.15258 | 0.038976 | 2.5089 | 2.5089 | 0.002475
2019-03-29 11:44 | 1553874212.8588 | 493 | 0.20171 | 0.00718 | 0.00019 | 101.1143 | 96.9736 | 110.9495 | 21.1095 | 0.004757 | 0.004113 | 135.8442 | 0.25591 | 8.4015 | 0.003189 | 0.005037 | 0.005007 | 5.8e-05 | 0.000894 | 0.30215 | 1.3364 | 2.5475 | 100.3 | 21 | 32.182 | 136.7049 | 98.4022 | 120.8751 | 0.007989 | 0.896 | 0.021597 | 24.104 | 0.374 | 0.15224 | 0.03824 | 2.4634 | 2.4634 | 0.002515
2019-03-29 11:44 | 1553874213.8598 | 494 | 0.20192 | 0.007055 | 0.000188 | 101.0805 | 96.7917 | 111.2532 | 21.1091 | 0.004695 | 0.004147 | 135.7672 | 0.25601 | 8.4016 | 0.003165 | 0.004412 | 0.004336 | 5.8e-05 | 0.000893 | 0.31813 | 1.3517 | 2.5475 | 100.3 | 21 | 32.183 | 135.1611 | 97.8279 | 121.7056 | 0.007989 | 0.89411 | 0.021604 | 24.1624 | 0.3878 | 0.1523 | 0.037111 | 2.3915 | 2.3915 | 0.002592
2019-03-29 11:44 | 1553874214.8598 | 495 | 0.20213 | 0.006925 | 0.000186 | 101.1021 | 96.5705 | 111.7419 | 21.11 | 0.004682 | 0.004042 | 135.7034 | 0.25505 | 8.4026 | 0.003087 | 0.005015 | 0.004903 | 5.7e-05 | 0.000888 | 0.26913 | 1.3648 | 2.5475 | 100.3 | 21 | 32.182 | 134.8358 | 95.9602 | 119.1415 | 0.007989 | 0.89996 | 0.021722 | 24.1366 | 0.41017 | 0.15173 | 0.036047 | 2.3357 | 2.3357 | 0.002644
2019-03-29 11:44 | 1553874215.8602 | 496 | 0.20244 | 0.006789 | 0.000183 | 101.079 | 96.4647 | 111.9051 | 21.1094 | 0.004681 | 0.004102 | 135.6212 | 0.25492 | 8.4017 | 0.003109 | 0.004532 | 0.004384 | 5.6e-05 | 0.000886 | 0.30773 | 1.3861 | 2.5475 | 100.3 | 21 | 32.182 | 134.831 | 96.4797 | 120.6082 | 0.007989 | 0.8966 | 0.021697 | 24.1993 | 0.41794 | 0.15165 | 0.034337 | 2.2224 | 2.2224 | 0.002777
2019-03-29 11:44 | 1553874216.8593 | 497 | 0.20282 | 0.006636 | 0.000181 | 101.0748 | 96.2555 | 112.3043 | 21.1098 | 0.004616 | 0.004014 | 135.5719 | 0.25605 | 8.4036 | 0.003146 | 0.004135 | 0.004073 | 5.9e-05 | 0.000886 | 0.30414 | 1.3626 | 2.5475 | 100.3 | 21 | 32.178 | 133.2101 | 97.3803 | 118.4548 | 0.007988 | 0.90153 | 0.021796 | 24.1761 | 0.43698 | 0.15233 | 0.03225 | 2.0968 | 2.0968 | 0.002957
2019-03-29 11:44 | 1553874217.8599 | 498 | 0.20299 | 0.006477 | 0.000177 | 101.079 | 96.0442 | 112.7878 | 21.11 | 0.004559 | 0.00398 | 135.5159 | 0.25555 | 8.402 | 0.00311 | 0.00522 | 0.005154 | 5.7e-05 | 0.000882 | 0.29048 | 1.3254 | 2.5474 | 100.3 | 21 | 32.177 | 131.8114 | 96.5155 | 117.6137 | 0.007988 | 0.90348 | 0.021866 | 24.202 | 0.45701 | 0.15202 | 0.031429 | 2.0502 | 2.0502 | 0.003018
2019-03-29 11:44 | 1553874218.8593 | 499 | 0.20318 | 0.006314 | 0.000172 | 101.0425 | 96.4896 | 113.1052 | 21.1099 | 0.004566 | 0.004014 | 135.4498 | 0.25651 | 8.4036 | 0.00311 | 0.004433 | 0.004368 | 5.7e-05 | 0.000885 | 0.2907 | 1.3317 | 2.5473 | 100.3 | 21 | 32.176 | 132.0003 | 96.5146 | 118.4558 | 0.007987 | 0.90153 | 0.021873 | 24.2622 | 0.41239 | 0.1526 | 0.030518 | 1.9915 | 1.9915 | 0.003119
2019-03-29 11:44 | 1553874219.86 | 500 | 0.20328 | 0.006137 | 0.000168 | 101.1131 | 96.462 | 113.6279 | 21.1104 | 0.004566 | 0.004027 | 135.3839 | 0.25538 | 8.4027 | 0.003111 | 0.004209 | 0.004139 | 5.8e-05 | 0.000885 | 0.27903 | 1.3115 | 2.5473 | 100.3 | 21 | 32.173 | 131.9808 | 96.5323 | 118.7668 | 0.007987 | 0.90082 | 0.021915 | 24.3278 | 0.42121 | 0.15192 | 0.030107 | 1.9687 | 1.9687 | 0.003141
2019-03-29 11:44 | 1553874220.859 | 501 | 0.20348 | 0.005969 | 0.000163 | 101.0832 | 96.6661 | 113.6793 | 21.1099 | 0.004515 | 0.003955 | 135.3198 | 0.25598 | 8.4022 | 0.003122 | 0.004812 | 0.004705 | 5.7e-05 | 0.000889 | 0.2912 | 1.3534 | 2.5477 | 100.3 | 21 | 32.17 | 130.7239 | 96.7865 | 117.0163 | 0.007986 | 0.90486 | 0.021969 | 24.2789 | 0.39965 | 0.15228 | 0.029137 | 1.9101 | 1.9101 | 0.003245
2019-03-29 11:44 | 1553874221.859 | 502 | 0.20364 | 0.005806 | 0.000158 | 101.0874 | 96.7135 | 114.0063 | 21.1097 | 0.004531 | 0.003948 | 135.2754 | 0.25566 | 8.4024 | 0.003074 | 0.005034 | 0.00494 | 5.7e-05 | 0.000893 | 0.30237 | 1.3115 | 2.5472 | 100.3 | 21 | 32.168 | 131.1252 | 95.6302 | 116.8436 | 0.007985 | 0.90526 | 0.022005 | 24.3084 | 0.39564 | 0.15209 | 0.028396 | 1.8648 | 1.8648 | 0.003319
2019-03-29 11:44 | 1553874222.8594 | 503 | 0.20386 | 0.005633 | 0.000155 | 101.1499 | 97.0917 | 114.2687 | 21.1099 | 0.004505 | 0.003923 | 135.1955 | 0.25588 | 8.4024 | 0.003135 | 0.004519 | 0.004437 | 6e-05 | 0.000897 | 0.29983 | 1.3214 | 2.5475 | 100.3 | 21 | 32.164 | 130.4931 | 97.1101 | 116.2278 | 0.007984 | 0.90669 | 0.022048 | 24.3172 | 0.36627 | 0.15222 | 0.027283 | 1.7952 | 1.7952 | 0.003451
2019-03-29 11:44 | 1553874223.8581 | 504 | 0.20419 | 0.005466 | 0.000152 | 101.0595 | 97.5333 | 114.3243 | 21.1102 | 0.004473 | 0.003871 | 135.1396 | 0.25588 | 8.4033 | 0.003071 | 0.004032 | 0.003978 | 5.9e-05 | 0.000892 | 0.30663 | 1.3421 | 2.5474 | 100.3 | 21 | 32.162 | 129.6947 | 95.5728 | 114.9655 | 0.007984 | 0.90964 | 0.022089 | 24.2836 | 0.31768 | 0.15222 | 0.025537 | 1.6835 | 1.6835 | 0.00368
2019-03-29 11:44 | 1553874224.8593 | 505 | 0.20445 | 0.005294 | 0.00015 | 101.1191 | 97.7432 | 114.2018 | 21.1118 | 0.004407 | 0.003895 | 135.0744 | 0.25475 | 8.4008 | 0.003067 | 0.003579 | 0.003506 | 5.6e-05 | 0.000889 | 0.286 | 1.3967 | 2.5476 | 100.3 | 21 | 32.158 | 128.0818 | 95.4616 | 115.5571 | 0.007983 | 0.90826 | 0.022061 | 24.2891 | 0.30372 | 0.15155 | 0.024146 | 1.5895 | 1.5895 | 0.00388
2019-03-29 11:44 | 1553874225.8581 | 506 | 0.20458 | 0.00512 | 0.000148 | 101.2236 | 98.0408 | 114.3391 | 21.11 | 0.004405 | 0.003817 | 135.0358 | 0.25601 | 8.4014 | 0.003046 | 0.003931 | 0.003735 | 5.5e-05 | 0.000888 | 0.28031 | 1.4042 | 2.5474 | 100.3 | 20.99 | 32.156 | 128.037 | 94.972 | 113.648 | 0.007978 | 0.91274 | 0.022128 | 24.244 | 0.28576 | 0.1523 | 0.023593 | 1.5579 | 1.5579 | 0.003979
2019-03-29 11:44 | 1553874226.8582 | 507 | 0.20472 | 0.004958 | 0.000146 | 101.1203 | 97.8903 | 114.164 | 21.1107 | 0.004412 | 0.003846 | 134.996 | 0.25587 | 8.402 | 0.002961 | 0.004085 | 0.004005 | 5.8e-05 | 0.000885 | 0.28263 | 1.3844 | 2.5474 | 100.3 | 20.99 | 32.154 | 128.2043 | 92.9155 | 114.353 | 0.007977 | 0.91108 | 0.022091 | 24.2475 | 0.29037 | 0.15222 | 0.022955 | 1.5133 | 1.5133 | 0.004094
2019-03-29 11:44 | 1553874227.8586 | 508 | 0.20486 | 0.004793 | 0.000144 | 101.1712 | 97.9241 | 113.9368 | 21.1105 | 0.004345 | 0.003757 | 134.9531 | 0.25598 | 8.4041 | 0.003026 | 0.004417 | 0.004334 | 5.8e-05 | 0.000879 | 0.26698 | 1.3447 | 2.5474 | 100.3 | 20.99 | 32.151 | 126.5479 | 94.4778 | 112.199 | 0.007976 | 0.91617 | 0.022131 | 24.1559 | 0.29179 | 0.15228 | 0.022351 | 1.476 | 1.476 | 0.004199
2019-03-29 11:44 | 1553874228.8583 | 509 | 0.20505 | 0.004637 | 0.000143 | 101.1071 | 98.2855 | 113.8249 | 21.1105 | 0.004345 | 0.003834 | 134.9132 | 0.25593 | 8.401 | 0.002983 | 0.00495 | 0.004783 | 5.4e-05 | 0.000887 | 0.31116 | 1.3398 | 2.5477 | 100.3 | 20.99 | 32.149 | 126.5494 | 93.4424 | 114.051 | 0.007976 | 0.91179 | 0.022067 | 24.202 | 0.25317 | 0.15225 | 0.021399 | 1.4089 | 1.4089 | 0.004398
2019-03-29 11:44 | 1553874229.8593 | 510 | 0.20527 | 0.004491 | 0.000142 | 101.1122 | 98.8111 | 113.7134 | 21.1105 | 0.00427 | 0.003794 | 134.8952 | 0.25584 | 8.4007 | 0.002953 | 0.004645 | 0.004603 | 5.3e-05 | 0.000883 | 0.31045 | 1.3488 | 2.5476 | 100.3 | 20.99 | 32.145 | 124.7223 | 92.7087 | 113.0859 | 0.007975 | 0.91407 | 0.022084 | 24.1599 | 0.20592 | 0.1522 | 0.020273 | 1.3356 | 1.3356 | 0.004638
2019-03-29 11:44 | 1553874230.8593 | 511 | 0.20549 | 0.004344 | 0.000142 | 101.1194 | 99.3257 | 113.9018 | 21.1109 | 0.004327 | 0.003716 | 134.8548 | 0.25556 | 8.4022 | 0.002895 | 0.004595 | 0.004437 | 6e-05 | 0.000878 | 0.27599 | 1.3159 | 2.5473 | 100.3 | 20.99 | 32.142 | 126.124 | 91.3297 | 111.1818 | 0.007974 | 0.9186 | 0.022157 | 24.1203 | 0.16009 | 0.15204 | 0.019122 | 1.2637 | 1.2637 | 0.004897
2019-03-29 11:44 | 1553874231.8587 | 512 | 0.20569 | 0.004205 | 0.000143 | 101.1431 | 99.608 | 113.815 | 21.1116 | 0.004265 | 0.003743 | 134.8213 | 0.25592 | 8.402 | 0.002948 | 0.004343 | 0.004165 | 5.7e-05 | 0.000873 | 0.29031 | 1.3314 | 2.5474 | 100.3 | 20.99 | 32.14 | 124.6007 | 92.6003 | 111.8479 | 0.007974 | 0.91701 | 0.022129 | 24.132 | 0.1368 | 0.15224 | 0.018061 | 1.1917 | 1.1917 | 0.0052
2019-03-29 11:44 | 1553874232.8584 | 513 | 0.20587 | 0.00407 | 0.000144 | 101.1772 | 99.7605 | 114.0284 | 21.1116 | 0.0042 | 0.003715 | 134.8102 | 0.25618 | 8.4037 | 0.003032 | 0.004113 | 0.004022 | 5.6e-05 | 0.000877 | 0.27865 | 1.3252 | 2.5474 | 100.3 | 20.99 | 32.138 | 123.0149 | 94.626 | 111.1587 | 0.007973 | 0.91865 | 0.02217 | 24.133 | 0.12612 | 0.1524 | 0.017162 | 1.1341 | 1.1341 | 0.005469
2019-03-29 11:44 | 1553874233.8588 | 514 | 0.20601 | 0.003922 | 0.000146 | 101.2257 | 99.8844 | 114.4849 | 21.1108 | 0.00421 | 0.00369 | 134.7973 | 0.255 | 8.4016 | 0.002916 | 0.004494 | 0.004317 | 5.5e-05 | 0.00088 | 0.28672 | 1.3304 | 2.5473 | 100.3 | 20.99 | 32.136 | 123.2524 | 91.8339 | 110.5616 | 0.007973 | 0.92008 | 0.022231 | 24.1624 | 0.11931 | 0.1517 | 0.016475 | 1.0914 | 1.0914 | 0.005657
2019-03-29 11:44 | 1553874234.8584 | 515 | 0.2061 | 0.003777 | 0.000148 | 101.1715 | 99.9148 | 114.8053 | 21.1114 | 0.004229 | 0.003674 | 134.7853 | 0.25581 | 8.4047 | 0.002982 | 0.004364 | 0.004309 | 5.7e-05 | 0.000878 | 0.28888 | 1.3173 | 2.5473 | 100.3 | 20.99 | 32.134 | 123.7141 | 93.4247 | 110.1582 | 0.007972 | 0.92105 | 0.022274 | 24.1835 | 0.1118 | 0.15218 | 0.016153 | 1.0719 | 1.0719 | 0.005778
2019-03-29 11:44 | 1553874235.8585 | 516 | 0.20622 | 0.003634 | 0.00015 | 101.1269 | 99.9138 | 114.7992 | 21.1117 | 0.004191 | 0.003686 | 134.7438 | 0.25597 | 8.4017 | 0.002947 | 0.004747 | 0.004701 | 5.7e-05 | 0.000874 | 0.33179 | 1.3428 | 2.5475 | 100.3 | 20.99 | 32.133 | 122.7909 | 92.5746 | 110.4686 | 0.007972 | 0.9203 | 0.022265 | 24.1926 | 0.10795 | 0.15228 | 0.015587 | 1.0336 | 1.0336 | 0.005996
2019-03-29 11:44 | 1553874236.8585 | 517 | 0.20642 | 0.003492 | 0.000152 | 101.1299 | 99.9452 | 115.2264 | 21.1112 | 0.004136 | 0.003626 | 134.6898 | 0.25493 | 8.4014 | 0.002985 | 0.004074 | 0.003992 | 5.8e-05 | 0.000873 | 0.32803 | 1.3773 | 2.5475 | 100.3 | 20.99 | 32.131 | 121.4275 | 93.4876 | 108.9989 | 0.007971 | 0.92384 | 0.022349 | 24.1911 | 0.1054 | 0.15166 | 0.014573 | 0.96942 | 0.96942 | 0.006367
2019-03-29 11:44 | 1553874237.8596 | 518 | 0.20655 | 0.003356 | 0.000155 | 101.0853 | 100.0865 | 115.3395 | 21.111 | 0.004164 | 0.003576 | 134.6516 | 0.25508 | 8.4031 | 0.00293 | 0.004239 | 0.004106 | 6e-05 | 0.00087 | 0.3027 | 1.3219 | 2.5475 | 100.3 | 20.99 | 32.129 | 122.1253 | 92.1663 | 107.798 | 0.007971 | 0.92675 | 0.022395 | 24.1649 | 0.088821 | 0.15175 | 0.01396 | 0.93007 | 0.93007 | 0.006641
2019-03-29 11:44 | 1553874238.8579 | 519 | 0.20677 | 0.003222 | 0.00016 | 101.1838 | 100.1179 | 115.6512 | 21.1112 | 0.004077 | 0.003595 | 134.633 | 0.25551 | 8.401 | 0.002932 | 0.004508 | 0.004416 | 5.7e-05 | 0.000869 | 0.29927 | 1.3695 | 2.5475 | 100.3 | 20.99 | 32.128 | 120.0008 | 92.2229 | 108.2496 | 0.007971 | 0.92566 | 0.022412 | 24.2119 | 0.094725 | 0.152 | 0.012754 | 0.84944 | 0.84944 | 0.007283
2019-03-29 11:44 | 1553874239.8587 | 520 | 0.20687 | 0.003097 | 0.000165 | 101.1736 | 100.265 | 115.92 | 21.1103 | 0.004081 | 0.003581 | 134.5924 | 0.25568 | 8.4028 | 0.002937 | 0.003859 | 0.003771 | 6e-05 | 0.000874 | 0.3027 | 1.3532 | 2.5474 | 100.3 | 20.99 | 32.126 | 120.0808 | 92.33 | 107.909 | 0.00797 | 0.92649 | 0.022448 | 24.2292 | 0.080694 | 0.1521 | 0.012294 | 0.81952 | 0.81952 | 0.007554
2019-03-29 11:44 | 1553874240.8576 | 521 | 0.20695 | 0.002967 | 0.00017 | 101.0604 | 100.3711 | 116.085 | 21.1108 | 0.004063 | 0.003605 | 134.573 | 0.25525 | 8.4038 | 0.002932 | 0.004064 | 0.003887 | 5.7e-05 | 0.000873 | 0.31045 | 1.3238 | 2.5475 | 100.3 | 20.99 | 32.126 | 119.6506 | 92.2069 | 108.4911 | 0.00797 | 0.92507 | 0.022447 | 24.265 | 0.061221 | 0.15185 | 0.011997 | 0.79925 | 0.79925 | 0.007733
2019-03-29 11:44 | 1553874241.8589 | 522 | 0.2071 | 0.002841 | 0.000174 | 101.129 | 100.1547 | 116.63 | 21.1111 | 0.004042 | 0.00352 | 134.5608 | 0.25521 | 8.4024 | 0.002892 | 0.004483 | 0.00441 | 5.9e-05 | 0.00088 | 0.28224 | 1.3702 | 2.5474 | 100.3 | 20.99 | 32.124 | 119.1354 | 91.2595 | 106.4404 | 0.00797 | 0.93007 | 0.02256 | 24.2564 | 0.086589 | 0.15182 | 0.011242 | 0.75195 | 0.75195 | 0.008218
2019-03-29 11:44 | 1553874242.8591 | 523 | 0.20727 | 0.00272 | 0.000178 | 101.0838 | 100.2861 | 116.9378 | 21.1126 | 0.004053 | 0.003558 | 134.5405 | 0.25657 | 8.4045 | 0.002913 | 0.005038 | 0.004868 | 5.9e-05 | 0.000877 | 0.28584 | 1.3499 | 2.5475 | 100.3 | 20.99 | 32.124 | 119.4026 | 91.7637 | 107.3529 | 0.00797 | 0.92784 | 0.022563 | 24.3176 | 0.070864 | 0.15263 | 0.010395 | 0.6945 | 0.6945 | 0.008945
2019-03-29 11:44 | 1553874243.8587 | 524 | 0.20732 | 0.002606 | 0.000181 | 101.094 | 100.3168 | 117.1234 | 21.1122 | 0.003959 | 0.003542 | 134.5229 | 0.2561 | 8.4003 | 0.002856 | 0.004504 | 0.004379 | 5.8e-05 | 0.000877 | 0.29518 | 1.3405 | 2.5475 | 100.3 | 20.99 | 32.123 | 117.1066 | 90.389 | 106.9705 | 0.007969 | 0.92877 | 0.022592 | 24.3246 | 0.069025 | 0.15235 | 0.010213 | 0.68291 | 0.68291 | 0.00908
2019-03-29 11:44 | 1553874244.8587 | 525 | 0.20746 | 0.002494 | 0.000183 | 101.0904 | 100.5131 | 117.0624 | 21.1119 | 0.003997 | 0.003501 | 134.5076 | 0.25559 | 8.4036 | 0.002877 | 0.004199 | 0.004139 | 5.7e-05 | 0.000876 | 0.29623 | 1.3319 | 2.5473 | 100.3 | 20.99 | 32.121 | 118.0394 | 90.8859 | 105.9784 | 0.007969 | 0.9312 | 0.022616 | 24.2865 | 0.051227 | 0.15205 | 0.009513 | 0.63593 | 0.63593 | 0.009731
2019-03-29 11:44 | 1553874245.8581 | 526 | 0.20755 | 0.002387 | 0.000186 | 101.0613 | 100.3615 | 117.3093 | 21.1117 | 0.003968 | 0.003452 | 134.5039 | 0.25523 | 8.4026 | 0.002811 | 0.004424 | 0.004343 | 5.8e-05 | 0.000874 | 0.30469 | 1.3408 | 2.5475 | 100.3 | 20.99 | 32.119 | 117.3285 | 89.3002 | 104.77 | 0.007968 | 0.93418 | 0.022676 | 24.2734 | 0.062152 | 0.15184 | 0.009102 | 0.60954 | 0.60954 | 0.010138
2019-03-29 11:44 | 1553874246.8588 | 527 | 0.20765 | 0.002285 | 0.000187 | 101.0482 | 100.2721 | 117.3741 | 21.1114 | 0.003903 | 0.003461 | 134.4901 | 0.25476 | 8.4052 | 0.002816 | 0.004818 | 0.004748 | 5.7e-05 | 0.00088 | 0.29535 | 1.3523 | 2.5474 | 100.3 | 20.99 | 32.116 | 115.7544 | 89.4216 | 105.0079 | 0.007968 | 0.93359 | 0.022675 | 24.2877 | 0.06896 | 0.15156 | 0.008599 | 0.57513 | 0.57513 | 0.010725
2019-03-29 11:44 | 1553874247.8585 | 528 | 0.20777 | 0.002194 | 0.00019 | 101.0634 | 100.308 | 117.6612 | 21.1114 | 0.0039 | 0.003515 | 134.4736 | 0.25502 | 8.4017 | 0.002802 | 0.004323 | 0.004209 | 5.8e-05 | 0.000883 | 0.30093 | 1.3035 | 2.5474 | 100.3 | 20.99 | 32.113 | 115.6635 | 89.0803 | 106.3078 | 0.007967 | 0.93039 | 0.022664 | 24.3592 | 0.067115 | 0.15171 | 0.007993 | 0.53351 | 0.53351 | 0.011574
2019-03-29 11:44 | 1553874248.8578 | 529 | 0.20782 | 0.002099 | 0.000192 | 101.0164 | 100.5209 | 117.8421 | 21.112 | 0.003914 | 0.00339 | 134.4598 | 0.25528 | 8.3994 | 0.002833 | 0.00438 | 0.004331 | 6e-05 | 0.00088 | 0.27599 | 1.308 | 2.5474 | 100.3 | 20.98 | 32.111 | 116.0053 | 89.8403 | 103.2827 | 0.007961 | 0.93787 | 0.022772 | 24.2805 | 0.043984 | 0.15187 | 0.007821 | 0.52417 | 0.52417 | 0.011792
2019-03-29 11:44 | 1553874249.8579 | 530 | 0.20794 | 0.002008 | 0.000193 | 100.9581 | 100.1633 | 118.025 | 21.1114 | 0.003924 | 0.003355 | 134.4449 | 0.25515 | 8.404 | 0.002807 | 0.004477 | 0.004346 | 6.1e-05 | 0.000878 | 0.30198 | 1.3482 | 2.5475 | 100.3 | 20.98 | 32.107 | 116.2508 | 89.1934 | 102.4266 | 0.00796 | 0.94001 | 0.022816 | 24.2717 | 0.070698 | 0.15179 | 0.007217 | 0.48373 | 0.48373 | 0.012771
2019-03-29 11:44 | 1553874250.8583 | 531 | 0.20802 | 0.00192 | 0.000193 | 101.0027 | 100.2697 | 118.1621 | 21.1122 | 0.003876 | 0.00337 | 134.4283 | 0.25556 | 8.4015 | 0.002716 | 0.003875 | 0.003805 | 6e-05 | 0.000877 | 0.29894 | 1.3176 | 2.5475 | 100.3 | 20.98 | 32.106 | 115.0828 | 87.0176 | 102.7942 | 0.00796 | 0.93909 | 0.022818 | 24.2977 | 0.065144 | 0.15203 | 0.006822 | 0.45673 | 0.45673 | 0.013548
2019-03-29 11:44 | 1553874251.859 | 532 | 0.20809 | 0.001842 | 0.000193 | 100.9706 | 100.0725 | 117.9757 | 21.112 | 0.003806 | 0.003352 | 134.4153 | 0.25492 | 8.4045 | 0.00274 | 0.004552 | 0.004377 | 6e-05 | 0.000876 | 0.31332 | 1.3146 | 2.5474 | 100.3 | 20.98 | 32.105 | 113.3908 | 87.5876 | 102.3471 | 0.00796 | 0.94021 | 0.022813 | 24.2641 | 0.079922 | 0.15165 | 0.006522 | 0.43611 | 0.43611 | 0.014153
2019-03-29 11:44 | 1553874252.859 | 533 | 0.20818 | 0.001759 | 0.000192 | 100.9272 | 100.3257 | 117.6438 | 21.1127 | 0.003785 | 0.003315 | 134.3895 | 0.2556 | 8.4014 | 0.002731 | 0.003725 | 0.003602 | 6e-05 | 0.000878 | 0.29806 | 1.3138 | 2.5475 | 100.3 | 20.98 | 32.102 | 112.8677 | 87.3764 | 101.4737 | 0.007959 | 0.9424 | 0.022808 | 24.2017 | 0.05347 | 0.15206 | 0.006076 | 0.40549 | 0.40549 | 0.015263
2019-03-29 11:44 | 1553874253.8584 | 534 | 0.20824 | 0.001684 | 0.00019 | 100.9509 | 100.2632 | 117.4365 | 21.1131 | 0.00376 | 0.003354 | 134.3671 | 0.25562 | 8.4028 | 0.002706 | 0.003777 | 0.003633 | 5.8e-05 | 0.00088 | 0.28385 | 1.3406 | 2.5474 | 100.3 | 20.98 | 32.102 | 112.2624 | 86.7664 | 102.4055 | 0.007959 | 0.94006 | 0.022759 | 24.2104 | 0.061136 | 0.15207 | 0.005814 | 0.3868 | 0.3868 | 0.016001
2019-03-29 11:44 | 1553874254.8582 | 535 | 0.20829 | 0.00161 | 0.000188 | 100.8767 | 100.3568 | 117.4678 | 21.113 | 0.003766 | 0.003305 | 134.367 | 0.25607 | 8.4043 | 0.002781 | 0.004307 | 0.004272 | 6.1e-05 | 0.000875 | 0.28783 | 1.3401 | 2.5475 | 100.3 | 20.98 | 32.102 | 112.4193 | 88.5771 | 101.2158 | 0.007959 | 0.94305 | 0.022798 | 24.1753 | 0.046217 | 0.15233 | 0.005578 | 0.37144 | 0.37144 | 0.016692
2019-03-29 11:44 | 1553874255.8581 | 536 | 0.20831 | 0.001544 | 0.000184 | 100.8596 | 100.1653 | 117.6997 | 21.1123 | 0.003827 | 0.003309 | 134.3763 | 0.25585 | 8.4031 | 0.002712 | 0.003586 | 0.003483 | 5.8e-05 | 0.000874 | 0.28423 | 1.3249 | 2.5473 | 100.3 | 20.98 | 32.102 | 113.8906 | 86.9224 | 101.3249 | 0.007959 | 0.94277 | 0.022818 | 24.2026 | 0.061785 | 0.15221 | 0.005557 | 0.3705 | 0.3705 | 0.01672
2019-03-29 11:44 | 1553874256.8571 | 537 | 0.20837 | 0.001484 | 0.000181 | 100.8354 | 100.1441 | 117.7922 | 21.1116 | 0.003699 | 0.003325 | 134.3832 | 0.25509 | 8.403 | 0.00269 | 0.003688 | 0.003535 | 6.1e-05 | 0.000872 | 0.30359 | 1.3362 | 2.5474 | 100.3 | 20.98 | 32.102 | 110.7706 | 86.3958 | 101.6941 | 0.007959 | 0.94185 | 0.022815 | 24.2241 | 0.061525 | 0.15175 | 0.00528 | 0.35169 | 0.35169 | 0.017562
2019-03-29 11:44 | 1553874257.8586 | 538 | 0.2084 | 0.001428 | 0.000176 | 100.8497 | 100.1991 | 118.2078 | 21.1124 | 0.003699 | 0.003256 | 134.3804 | 0.25596 | 8.4011 | 0.002686 | 0.004157 | 0.004031 | 5.8e-05 | 0.000877 | 0.28943 | 1.3468 | 2.5474 | 100.3 | 20.97 | 32.105 | 110.7667 | 86.2825 | 100.0456 | 0.007955 | 0.94601 | 0.022906 | 24.2134 | 0.057892 | 0.15227 | 0.005156 | 0.34484 | 0.34484 | 0.017971
2019-03-29 11:44 | 1553874258.8582 | 539 | 0.20846 | 0.00138 | 0.000172 | 100.8536 | 100.224 | 118.2405 | 21.1126 | 0.003716 | 0.003286 | 134.3696 | 0.25547 | 8.4054 | 0.002739 | 0.004331 | 0.004289 | 5.8e-05 | 0.000882 | 0.31996 | 1.352 | 2.5475 | 100.3 | 20.97 | 32.107 | 111.1993 | 87.5701 | 100.7521 | 0.007956 | 0.94422 | 0.022888 | 24.2396 | 0.056014 | 0.15198 | 0.004897 | 0.32699 | 0.32699 | 0.018917
2019-03-29 11:44 | 1553874259.8582 | 540 | 0.20849 | 0.001333 | 0.000167 | 100.7683 | 100.0554 | 118.0766 | 21.1131 | 0.003613 | 0.003211 | 134.3788 | 0.25547 | 8.4014 | 0.002713 | 0.004282 | 0.004141 | 5.7e-05 | 0.000884 | 0.27903 | 1.3694 | 2.5475 | 100.3 | 20.97 | 32.111 | 108.6981 | 86.9427 | 98.9412 | 0.007957 | 0.94881 | 0.022927 | 24.1641 | 0.063504 | 0.15198 | 0.004747 | 0.31752 | 0.31752 | 0.019481
2019-03-29 11:44 | 1553874260.8586 | 541 | 0.20857 | 0.001288 | 0.000162 | 100.7668 | 100.1223 | 118.036 | 21.1113 | 0.003687 | 0.003194 | 134.3842 | 0.25586 | 8.4026 | 0.002731 | 0.0045 | 0.00442 | 5.9e-05 | 0.000887 | 0.29894 | 1.3259 | 2.5474 | 100.3 | 20.97 | 32.114 | 110.4824 | 87.3713 | 98.5465 | 0.007957 | 0.94982 | 0.022935 | 24.1471 | 0.057393 | 0.15221 | 0.004373 | 0.29214 | 0.29214 | 0.021205
2019-03-29 11:44 | 1553874261.8572 | 542 | 0.20858 | 0.001246 | 0.000158 | 100.803 | 100.1216 | 118.0576 | 21.1117 | 0.003647 | 0.003243 | 134.3647 | 0.2553 | 8.4025 | 0.002684 | 0.004725 | 0.004564 | 5.9e-05 | 0.00089 | 0.30021 | 1.3417 | 2.5474 | 100.3 | 20.97 | 32.115 | 109.5119 | 86.235 | 99.7273 | 0.007958 | 0.94681 | 0.022901 | 24.1877 | 0.060669 | 0.15188 | 0.004366 | 0.29147 | 0.29147 | 0.021207
2019-03-29 11:44 | 1553874262.9463 | 543 | 0.20862 | 0.001214 | 0.000153 | 100.7237 | 100.2984 | 117.8839 | 21.1123 | 0.00366 | 0.003191 | 134.373 | 0.25491 | 8.4022 | 0.002646 | 0.003521 | 0.003404 | 5.6e-05 | 0.000888 | 0.31078 | 1.3245 | 2.5474 | 100.3 | 20.97 | 32.118 | 109.8396 | 85.338 | 98.4644 | 0.007958 | 0.95003 | 0.022923 | 24.1289 | 0.037848 | 0.15165 | 0.004163 | 0.27806 | 0.27806 | 0.022196
2019-03-29 11:44 | 1553874263.8917 | 544 | 0.20863 | 0.001175 | 0.000148 | 100.7791 | 100.1981 | 117.9738 | 21.1121 | 0.003615 | 0.003205 | 134.373 | 0.25568 | 8.4023 | 0.002642 | 0.00382 | 0.003713 | 5.9e-05 | 0.000891 | 0.31188 | 1.3378 | 2.5474 | 100.3 | 20.97 | 32.12 | 108.7375 | 85.2317 | 98.8015 | 0.007959 | 0.94917 | 0.022922 | 24.149 | 0.051716 | 0.15211 | 0.004166 | 0.27853 | 0.27853 | 0.022226
2019-03-29 11:44 | 1553874264.8908 | 545 | 0.20864 | 0.001147 | 0.000145 | 100.6773 | 100.022 | 117.8668 | 21.1124 | 0.003615 | 0.00318 | 134.3302 | 0.25529 | 8.4026 | 0.00262 | 0.004385 | 0.004287 | 5.7e-05 | 0.000893 | 0.29856 | 1.3005 | 2.5473 | 100.3 | 20.96 | 32.123 | 108.7314 | 84.7129 | 98.1995 | 0.007955 | 0.95071 | 0.02293 | 24.1185 | 0.058414 | 0.15187 | 0.004136 | 0.27678 | 0.27678 | 0.022332
2019-03-29 11:44 | 1553874265.8917 | 546 | 0.20869 | 0.00112 | 0.000141 | 100.6225 | 100.0363 | 117.8741 | 21.112 | 0.003556 | 0.003161 | 134.3066 | 0.25514 | 8.4007 | 0.002628 | 0.004434 | 0.004358 | 5.9e-05 | 0.000893 | 0.27704 | 1.3231 | 2.5475 | 100.3 | 20.96 | 32.123 | 107.312 | 84.9037 | 97.7305 | 0.007955 | 0.95191 | 0.022945 | 24.104 | 0.052264 | 0.15178 | 0.003851 | 0.25752 | 0.25752 | 0.023989
2019-03-29 11:44 | 1553874266.8907 | 547 | 0.20874 | 0.001096 | 0.000138 | 100.6378 | 100.0629 | 117.8558 | 21.1124 | 0.003558 | 0.003193 | 134.2814 | 0.25559 | 8.4027 | 0.002633 | 0.00378 | 0.003654 | 5.6e-05 | 0.000892 | 0.33019 | 1.3223 | 2.5474 | 100.3 | 20.96 | 32.123 | 107.3527 | 85.0102 | 98.5038 | 0.007955 | 0.94993 | 0.022919 | 24.1273 | 0.051241 | 0.15205 | 0.003583 | 0.23887 | 0.23887 | 0.025906
2019-03-29 11:44 | 1553874267.8921 | 548 | 0.20872 | 0.001074 | 0.000136 | 100.5839 | 99.7465 | 117.7166 | 21.113 | 0.003512 | 0.003145 | 134.2567 | 0.25537 | 8.4039 | 0.002697 | 0.00352 | 0.003472 | 5.4e-05 | 0.000893 | 0.30237 | 1.3292 | 2.5475 | 100.3 | 20.96 | 32.125 | 106.2354 | 86.5437 | 97.3561 | 0.007955 | 0.95287 | 0.022941 | 24.0758 | 0.074776 | 0.15192 | 0.00373 | 0.24938 | 0.24938 | 0.024794
2019-03-29 11:44 | 1553874268.8907 | 549 | 0.20872 | 0.001053 | 0.000133 | 100.6246 | 99.937 | 117.8975 | 21.1128 | 0.003546 | 0.00311 | 134.2537 | 0.25482 | 8.4047 | 0.00257 | 0.004446 | 0.004325 | 5.7e-05 | 0.000888 | 0.3183 | 1.3874 | 2.5473 | 100.3 | 20.96 | 32.126 | 107.0546 | 83.4939 | 96.5112 | 0.007955 | 0.95505 | 0.022985 | 24.0668 | 0.061335 | 0.15159 | 0.003779 | 0.25343 | 0.25343 | 0.024345
2019-03-29 11:44 | 1553874269.8908 | 550 | 0.20872 | 0.001036 | 0.000131 | 100.5031 | 100.0465 | 118.0156 | 21.112 | 0.003532 | 0.003096 | 134.2312 | 0.25569 | 8.4008 | 0.002607 | 0.004069 | 0.003949 | 5.9e-05 | 0.000883 | 0.29391 | 1.332 | 2.5475 | 100.3 | 20.96 | 32.125 | 106.7223 | 84.391 | 96.1792 | 0.007955 | 0.95591 | 0.023007 | 24.068 | 0.040726 | 0.15211 | 0.003812 | 0.25609 | 0.25609 | 0.024175
2019-03-29 11:44 | 1553874270.8898 | 551 | 0.20869 | 0.001017 | 0.000129 | 100.5432 | 100.0438 | 118.0332 | 21.1129 | 0.003476 | 0.003105 | 134.2207 | 0.25584 | 8.4049 | 0.002614 | 0.004199 | 0.004112 | 5.8e-05 | 0.000881 | 0.3095 | 1.3687 | 2.5474 | 100.3 | 20.96 | 32.124 | 105.3721 | 84.5698 | 96.3915 | 0.007955 | 0.95536 | 0.023002 | 24.0767 | 0.044539 | 0.1522 | 0.004011 | 0.26993 | 0.26993 | 0.022949
2019-03-29 11:44 | 1553874271.8909 | 552 | 0.20871 | 0.001005 | 0.000127 | 100.507 | 99.8547 | 118.0891 | 21.1128 | 0.003484 | 0.003095 | 134.2145 | 0.25536 | 8.4033 | 0.002604 | 0.004049 | 0.004018 | 5.8e-05 | 0.000882 | 0.29839 | 1.3675 | 2.5474 | 100.3 | 20.96 | 32.124 | 105.5694 | 84.3115 | 96.1392 | 0.007955 | 0.95601 | 0.023015 | 24.0742 | 0.058236 | 0.15192 | 0.003921 | 0.26396 | 0.26396 | 0.023424
2019-03-29 11:45 | 1553874272.891 | 553 | 0.20875 | 0.000988 | 0.000124 | 100.4217 | 99.8261 | 118.3977 | 21.1121 | 0.003439 | 0.003099 | 134.2007 | 0.2559 | 8.4033 | 0.002578 | 0.004124 | 0.004013 | 5.9e-05 | 0.00088 | 0.2975 | 1.3206 | 2.5474 | 100.3 | 20.96 | 32.126 | 104.4604 | 83.6813 | 96.2498 | 0.007955 | 0.95573 | 0.023042 | 24.1093 | 0.053209 | 0.15223 | 0.003653 | 0.24587 | 0.24587 | 0.0252
2019-03-29 11:45 | 1553874273.89 | 554 | 0.20882 | 0.000969 | 0.000122 | 100.4082 | 99.7578 | 118.6379 | 21.1131 | 0.003476 | 0.003055 | 134.1953 | 0.25529 | 8.4032 | 0.002688 | 0.003093 | 0.002995 | 5.6e-05 | 0.000875 | 0.3115 | 1.3264 | 2.5474 | 100.3 | 20.96 | 32.127 | 105.373 | 86.3284 | 95.1723 | 0.007956 | 0.95852 | 0.023099 | 24.0985 | 0.058128 | 0.15187 | 0.003298 | 0.2219 | 0.2219 | 0.027856
2019-03-29 11:45 | 1553874274.89 | 555 | 0.2088 | 0.000952 | 0.000119 | 100.395 | 99.9489 | 118.9052 | 21.1141 | 0.003448 | 0.003053 | 134.2284 | 0.25584 | 8.4021 | 0.002569 | 0.003809 | 0.003685 | 5.8e-05 | 0.000878 | 0.28854 | 1.3303 | 2.5473 | 100.3 | 20.96 | 32.129 | 104.6819 | 83.4757 | 95.1262 | 0.007956 | 0.95864 | 0.023126 | 24.1241 | 0.039834 | 0.1522 | 0.003427 | 0.23132 | 0.23132 | 0.02678
2019-03-29 11:45 | 1553874275.8914 | 556 | 0.2088 | 0.000935 | 0.000118 | 100.3175 | 99.8312 | 119.2904 | 21.1136 | 0.003441 | 0.003063 | 134.2192 | 0.25622 | 8.4038 | 0.002603 | 0.003687 | 0.003607 | 5.9e-05 | 0.00088 | 0.28888 | 1.316 | 2.5473 | 100.3 | 20.97 | 32.13 | 104.5122 | 84.2983 | 95.3709 | 0.007961 | 0.95801 | 0.023156 | 24.1712 | 0.043468 | 0.15243 | 0.003464 | 0.2343 | 0.2343 | 0.026478
2019-03-29 11:45 | 1553874276.8911 | 557 | 0.20883 | 0.000921 | 0.000116 | 100.2756 | 99.8988 | 119.5474 | 21.1137 | 0.003432 | 0.003058 | 134.2235 | 0.25488 | 8.4032 | 0.002626 | 0.003784 | 0.003743 | 5.7e-05 | 0.000877 | 0.29048 | 1.3148 | 2.5474 | 100.3 | 20.96 | 32.132 | 104.2859 | 84.857 | 95.2605 | 0.007957 | 0.95829 | 0.023185 | 24.1936 | 0.033678 | 0.15163 | 0.003301 | 0.22335 | 0.22335 | 0.02763
2019-03-29 11:45 | 1553874277.8911 | 558 | 0.20885 | 0.000908 | 0.000113 | 100.2678 | 99.9302 | 119.2798 | 21.1134 | 0.003404 | 0.003007 | 134.2429 | 0.25624 | 8.4008 | 0.002562 | 0.004515 | 0.004386 | 5.8e-05 | 0.000883 | 0.29302 | 1.3348 | 2.5474 | 100.3 | 20.96 | 32.135 | 103.6086 | 83.304 | 94.0314 | 0.007958 | 0.9615 | 0.023197 | 24.1261 | 0.030173 | 0.15244 | 0.003186 | 0.2156 | 0.2156 | 0.028776
2019-03-29 11:45 | 1553874278.8901 | 559 | 0.20887 | 0.000893 | 0.000112 | 100.2675 | 99.8691 | 119.2639 | 21.1139 | 0.003373 | 0.003012 | 134.2411 | 0.25566 | 8.4011 | 0.002527 | 0.004291 | 0.004212 | 5.8e-05 | 0.000882 | 0.27417 | 1.3202 | 2.5474 | 100.3 | 20.96 | 32.137 | 102.8659 | 82.4694 | 94.1328 | 0.007958 | 0.96124 | 0.023193 | 24.1279 | 0.035616 | 0.15209 | 0.003089 | 0.20885 | 0.20885 | 0.029639
2019-03-29 11:45 | 1553874279.8911 | 560 | 0.20885 | 0.000872 | 0.000109 | 100.2636 | 99.8465 | 119.2503 | 21.1133 | 0.003392 | 0.003025 | 134.2634 | 0.25531 | 8.4034 | 0.002572 | 0.003522 | 0.00344 | 5.6e-05 | 0.000881 | 0.2975 | 1.3143 | 2.5475 | 100.3 | 20.96 | 32.138 | 103.3399 | 83.5427 | 94.446 | 0.007958 | 0.96042 | 0.023181 | 24.1368 | 0.037286 | 0.15188 | 0.003256 | 0.22057 | 0.22057 | 0.028026
2019-03-29 11:45 | 1553874280.8902 | 561 | 0.20892 | 0.000864 | 0.000107 | 100.2331 | 99.623 | 119.1176 | 21.1135 | 0.003359 | 0.002995 | 134.2634 | 0.25606 | 8.4026 | 0.002533 | 0.004514 | 0.004401 | 5.9e-05 | 0.000884 | 0.30237 | 1.3134 | 2.5474 | 100.3 | 20.96 | 32.144 | 102.5228 | 82.6114 | 93.7342 | 0.00796 | 0.96228 | 0.023191 | 24.1 | 0.05461 | 0.15233 | 0.002814 | 0.19001 | 0.19001 | 0.03263
2019-03-29 11:45 | 1553874281.8912 | 562 | 0.20891 | 0.000853 | 0.000104 | 100.2047 | 99.8704 | 119.1702 | 21.1138 | 0.003323 | 0.002978 | 134.255 | 0.25596 | 8.4018 | 0.00251 | 0.003872 | 0.003793 | 5.8e-05 | 0.000881 | 0.28313 | 1.3275 | 2.5474 | 100.3 | 20.96 | 32.146 | 101.6672 | 82.0503 | 93.3214 | 0.00796 | 0.96336 | 0.023209 | 24.0917 | 0.029884 | 0.15227 | 0.002921 | 0.19774 | 0.19774 | 0.031341
2019-03-29 11:45 | 1553874282.8912 | 563 | 0.20891 | 0.00084 | 0.000103 | 100.2502 | 99.7974 | 119.4791 | 21.1141 | 0.003343 | 0.00297 | 134.2664 | 0.25567 | 8.4034 | 0.002593 | 0.003905 | 0.003842 | 5.6e-05 | 0.000889 | 0.28888 | 1.3271 | 2.5474 | 100.3 | 20.96 | 32.152 | 102.137 | 84.0648 | 93.1227 | 0.007962 | 0.96389 | 0.023245 | 24.1164 | 0.040488 | 0.1521 | 0.002911 | 0.19748 | 0.19748 | 0.031347
2019-03-29 11:45 | 1553874283.8913 | 564 | 0.20893 | 0.000828 | 0.000101 | 100.1439 | 99.5909 | 119.5185 | 21.1139 | 0.003286 | 0.002956 | 134.2637 | 0.25563 | 8.4013 | 0.002523 | 0.004001 | 0.003927 | 5.5e-05 | 0.000896 | 0.31045 | 1.3505 | 2.5474 | 100.3 | 20.96 | 32.154 | 100.7722 | 82.3755 | 92.7987 | 0.007962 | 0.96474 | 0.02326 | 24.1097 | 0.049528 | 0.15207 | 0.002806 | 0.1903 | 0.1903 | 0.032525
2019-03-29 11:45 | 1553874284.8913 | 565 | 0.20892 | 0.00082 | 0.0001 | 100.1607 | 99.5861 | 119.301 | 21.1136 | 0.003268 | 0.003028 | 134.2439 | 0.25553 | 8.4011 | 0.002553 | 0.003997 | 0.003845 | 5.9e-05 | 0.00089 | 0.28744 | 1.3389 | 2.5475 | 100.3 | 20.96 | 32.156 | 100.331 | 83.0984 | 94.5216 | 0.007963 | 0.96022 | 0.023184 | 24.1444 | 0.051454 | 0.15202 | 0.002918 | 0.19759 | 0.19759 | 0.031313
2019-03-29 11:45 | 1553874285.8893 | 566 | 0.20891 | 0.00081 | 9.9e-05 | 100.0628 | 99.709 | 119.0007 | 21.1135 | 0.003271 | 0.002977 | 134.2535 | 0.25496 | 8.4026 | 0.00252 | 0.004025 | 0.003959 | 5.6e-05 | 0.000889 | 0.28064 | 1.3591 | 2.5474 | 100.3 | 20.96 | 32.16 | 100.3921 | 82.2973 | 93.3023 | 0.007964 | 0.96341 | 0.023193 | 24.074 | 0.03168 | 0.15168 | 0.003004 | 0.20369 | 0.20369 | 0.030306
2019-03-29 11:45 | 1553874286.8898 | 567 | 0.2089 | 0.000799 | 9.8e-05 | 100.0999 | 99.7554 | 118.809 | 21.1133 | 0.003301 | 0.002984 | 134.2521 | 0.25577 | 8.402 | 0.002549 | 0.003608 | 0.003493 | 5.9e-05 | 0.000882 | 0.28655 | 1.3691 | 2.5475 | 100.3 | 20.96 | 32.162 | 101.1132 | 83.0036 | 93.4576 | 0.007964 | 0.96301 | 0.02317 | 24.0596 | 0.030834 | 0.15216 | 0.00305 | 0.20672 | 0.20672 | 0.029957
2019-03-29 11:45 | 1553874287.8894 | 568 | 0.20895 | 0.00079 | 9.8e-05 | 100.052 | 99.6025 | 119.0416 | 21.1139 | 0.003224 | 0.002929 | 134.2436 | 0.2553 | 8.4037 | 0.002583 | 0.003898 | 0.003828 | 5.6e-05 | 0.00088 | 0.32477 | 1.3323 | 2.5475 | 100.3 | 20.96 | 32.164 | 99.2534 | 83.8064 | 92.1318 | 0.007965 | 0.9665 | 0.023234 | 24.0396 | 0.040274 | 0.15188 | 0.00275 | 0.18637 | 0.18637 | 0.033168
2019-03-29 11:45 | 1553874288.8894 | 569 | 0.20891 | 0.000786 | 9.7e-05 | 100.0475 | 99.4756 | 118.8334 | 21.1133 | 0.003269 | 0.002907 | 134.2195 | 0.2552 | 8.4016 | 0.002522 | 0.003326 | 0.003206 | 5.7e-05 | 0.000876 | 0.30237 | 1.348 | 2.5474 | 100.3 | 20.96 | 32.164 | 100.351 | 82.343 | 91.6142 | 0.007965 | 0.96787 | 0.02323 | 24.0015 | 0.051281 | 0.15182 | 0.003006 | 0.20427 | 0.20427 | 0.030249
2019-03-29 11:45 | 1553874289.8897 | 570 | 0.20892 | 0.00078 | 9.7e-05 | 100.043 | 99.5216 | 119.109 | 21.1142 | 0.003276 | 0.002959 | 134.19 | 0.25511 | 8.4006 | 0.002524 | 0.004361 | 0.004275 | 5.7e-05 | 0.000875 | 0.28766 | 1.324 | 2.5474 | 100.3 | 20.96 | 32.165 | 100.5238 | 82.3931 | 92.8774 | 0.007965 | 0.96453 | 0.023217 | 24.0709 | 0.046738 | 0.15177 | 0.002957 | 0.20074 | 0.20074 | 0.03077
2019-03-29 11:45 | 1553874290.8895 | 571 | 0.20897 | 0.000775 | 9.6e-05 | 99.9712 | 99.4175 | 119.4052 | 21.1137 | 0.003223 | 0.002931 | 134.1771 | 0.25564 | 8.4012 | 0.002622 | 0.004809 | 0.004673 | 5.7e-05 | 0.000877 | 0.30862 | 1.3639 | 2.5474 | 100.3 | 20.96 | 32.166 | 99.2381 | 84.744 | 92.1877 | 0.007965 | 0.96635 | 0.023268 | 24.0781 | 0.049673 | 0.15208 | 0.002652 | 0.17991 | 0.17991 | 0.034403
2019-03-29 11:45 | 1553874291.8911 | 572 | 0.20899 | 0.000775 | 9.6e-05 | 99.9511 | 99.4486 | 119.5451 | 21.1139 | 0.003207 | 0.002889 | 134.1711 | 0.25564 | 8.4028 | 0.002583 | 0.003915 | 0.003858 | 5.8e-05 | 0.000881 | 0.30198 | 1.3663 | 2.5474 | 100.3 | 20.97 | 32.165 | 98.8502 | 83.8102 | 91.1739 | 0.00797 | 0.96904 | 0.023314 | 24.0588 | 0.045085 | 0.15208 | 0.002515 | 0.17069 | 0.17069 | 0.036263
2019-03-29 11:45 | 1553874292.89 | 573 | 0.20902 | 0.000768 | 9.5e-05 | 99.9284 | 99.5397 | 119.3254 | 21.1138 | 0.003143 | 0.002871 | 134.1471 | 0.25579 | 8.4032 | 0.002569 | 0.003655 | 0.003481 | 5.4e-05 | 0.000886 | 0.25946 | 1.3218 | 2.5474 | 100.29 | 20.97 | 32.163 | 97.3028 | 83.4823 | 90.7439 | 0.00797 | 0.97019 | 0.023306 | 24.0225 | 0.034856 | 0.15217 | 0.002352 | 0.15922 | 0.15922 | 0.038898
2019-03-29 11:45 | 1553874293.8896 | 574 | 0.20903 | 0.000762 | 9.5e-05 | 99.9508 | 99.51 | 119.2641 | 21.1139 | 0.003248 | 0.002935 | 134.1488 | 0.25574 | 8.3998 | 0.002526 | 0.00417 | 0.004031 | 5.9e-05 | 0.000886 | 0.28512 | 1.3252 | 2.5474 | 100.29 | 20.97 | 32.159 | 99.8423 | 82.445 | 92.2778 | 0.007969 | 0.96611 | 0.023251 | 24.0669 | 0.039534 | 0.15214 | 0.002328 | 0.1572 | 0.1572 | 0.03939
2019-03-29 11:45 | 1553874294.8908 | 575 | 0.20904 | 0.000756 | 9.3e-05 | 99.9257 | 99.5285 | 119.6237 | 21.1141 | 0.003156 | 0.002924 | 134.1462 | 0.25648 | 8.4029 | 0.002548 | 0.004166 | 0.004099 | 5.8e-05 | 0.00089 | 0.30093 | 1.3253 | 2.5474 | 100.29 | 20.97 | 32.156 | 97.6195 | 82.9773 | 92.0223 | 0.007968 | 0.96679 | 0.023295 | 24.0947 | 0.035627 | 0.15258 | 0.002246 | 0.15186 | 0.15186 | 0.040891
2019-03-29 11:45 | 1553874295.8896 | 576 | 0.20903 | 0.00075 | 9.3e-05 | 99.8577 | 99.4971 | 119.6674 | 21.1137 | 0.003181 | 0.002875 | 134.1409 | 0.25624 | 8.4043 | 0.002497 | 0.003894 | 0.003876 | 5.5e-05 | 0.000893 | 0.28584 | 1.3153 | 2.5473 | 100.3 | 20.97 | 32.151 | 98.2291 | 81.739 | 90.85 | 0.007966 | 0.96991 | 0.023336 | 24.0604 | 0.032365 | 0.15244 | 0.002332 | 0.15816 | 0.15816 | 0.039227
2019-03-29 11:45 | 1553874296.8901 | 577 | 0.20902 | 0.000741 | 9.2e-05 | 99.8341 | 99.5087 | 119.6911 | 21.114 | 0.003193 | 0.00293 | 134.1455 | 0.25531 | 8.4013 | 0.002505 | 0.004409 | 0.004236 | 6.1e-05 | 0.000896 | 0.30574 | 1.3089 | 2.5473 | 100.3 | 20.97 | 32.149 | 98.5139 | 81.9492 | 92.1577 | 0.007966 | 0.96643 | 0.023297 | 24.1059 | 0.029204 | 0.15188 | 0.002419 | 0.16405 | 0.16405 | 0.037681
2019-03-29 11:45 | 1553874297.8902 | 578 | 0.20905 | 0.000737 | 9.2e-05 | 99.8748 | 99.2909 | 119.3761 | 21.1135 | 0.003161 | 0.002937 | 134.1431 | 0.25512 | 8.4014 | 0.002565 | 0.003898 | 0.003879 | 5.6e-05 | 0.000891 | 0.29479 | 1.3225 | 2.5475 | 100.3 | 20.97 | 32.147 | 97.7516 | 83.3924 | 92.3283 | 0.007965 | 0.96598 | 0.023261 | 24.0798 | 0.052444 | 0.15177 | 0.002216 | 0.14963 | 0.14963 | 0.041283
2019-03-29 11:45 | 1553874298.8898 | 579 | 0.20904 | 0.000731 | 9.2e-05 | 99.852 | 99.5353 | 119.4524 | 21.114 | 0.00313 | 0.002854 | 134.1573 | 0.25521 | 8.4006 | 0.002543 | 0.00459 | 0.004544 | 5.5e-05 | 0.000894 | 0.28351 | 1.3273 | 2.5475 | 100.3 | 20.97 | 32.14 | 96.9857 | 82.8469 | 90.3431 | 0.007964 | 0.97126 | 0.023332 | 24.022 | 0.02842 | 0.15182 | 0.00229 | 0.15524 | 0.15524 | 0.039804
2019-03-29 11:45 | 1553874299.8898 | 580 | 0.20902 | 0.000725 | 9.1e-05 | 99.7808 | 99.307 | 119.4155 | 21.1137 | 0.003153 | 0.002861 | 134.1727 | 0.25624 | 8.4028 | 0.002541 | 0.003903 | 0.003799 | 5.2e-05 | 0.000899 | 0.32173 | 1.3888 | 2.5475 | 100.3 | 20.97 | 32.137 | 97.5423 | 82.8107 | 90.5164 | 0.007963 | 0.97079 | 0.023322 | 24.024 | 0.042578 | 0.15244 | 0.00245 | 0.16642 | 0.16642 | 0.03728
2019-03-29 11:45 | 1553874300.8892 | 581 | 0.20907 | 0.000721 | 9.1e-05 | 99.8128 | 99.4885 | 119.3789 | 21.1145 | 0.003158 | 0.00289 | 134.1768 | 0.25568 | 8.401 | 0.002494 | 0.003769 | 0.003667 | 5.6e-05 | 0.000895 | 0.27743 | 1.3764 | 2.5474 | 100.3 | 20.97 | 32.134 | 97.6593 | 81.6743 | 91.2043 | 0.007962 | 0.96896 | 0.023297 | 24.0431 | 0.02911 | 0.1521 | 0.002115 | 0.14285 | 0.14285 | 0.043336
2019-03-29 11:45 | 1553874301.8893 | 582 | 0.20905 | 0.000721 | 9e-05 | 99.7922 | 99.4087 | 119.208 | 21.1142 | 0.003126 | 0.002822 | 134.1829 | 0.25588 | 8.4016 | 0.002491 | 0.004029 | 0.003952 | 5.5e-05 | 0.000898 | 0.30525 | 1.3482 | 2.5474 | 100.3 | 20.97 | 32.129 | 96.9042 | 81.6093 | 89.5775 | 0.007961 | 0.97331 | 0.023332 | 23.9721 | 0.034442 | 0.15222 | 0.00228 | 0.15466 | 0.15466 | 0.040058
2019-03-29 11:45 | 1553874302.8889 | 583 | 0.20903 | 0.000718 | 9e-05 | 99.7164 | 99.2551 | 119.1866 | 21.1142 | 0.003139 | 0.002835 | 134.18 | 0.25527 | 8.4047 | 0.002547 | 0.004066 | 0.004004 | 5.6e-05 | 0.0009 | 0.31205 | 1.3482 | 2.5474 | 100.3 | 20.97 | 32.126 | 97.2215 | 82.9554 | 89.8713 | 0.00796 | 0.97252 | 0.023321 | 23.9797 | 0.041481 | 0.15186 | 0.002385 | 0.1619 | 0.1619 | 0.038178
2019-03-29 11:45 | 1553874303.8889 | 584 | 0.20913 | 0.000711 | 9e-05 | 99.7913 | 99.3541 | 118.9439 | 21.1141 | 0.003172 | 0.002872 | 134.1763 | 0.25573 | 8.4008 | 0.002541 | 0.003993 | 0.003839 | 5.7e-05 | 0.000898 | 0.31509 | 1.3748 | 2.5475 | 100.3 | 20.97 | 32.12 | 98.0029 | 82.8112 | 90.7621 | 0.007959 | 0.97014 | 0.023268 | 23.9846 | 0.039276 | 0.15213 | 0.001757 | 0.11765 | 0.11765 | 0.052627
2019-03-29 11:45 | 1553874304.8903 | 585 | 0.20907 | 0.000712 | 9e-05 | 99.7697 | 99.4971 | 118.8036 | 21.1144 | 0.003153 | 0.002839 | 134.1658 | 0.2561 | 8.4019 | 0.002551 | 0.003843 | 0.003751 | 5.8e-05 | 0.000897 | 0.28982 | 1.3674 | 2.5475 | 100.3 | 20.97 | 32.117 | 97.5509 | 83.053 | 89.9676 | 0.007958 | 0.97226 | 0.02328 | 23.9443 | 0.024479 | 0.15235 | 0.002174 | 0.14691 | 0.14691 | 0.042208
2019-03-29 11:45 | 1553874305.8901 | 586 | 0.20905 | 0.000707 | 9e-05 | 99.788 | 99.3483 | 118.544 | 21.1148 | 0.003078 | 0.00289 | 134.1554 | 0.25596 | 8.4023 | 0.00247 | 0.004339 | 0.004241 | 6.1e-05 | 0.000898 | 0.28888 | 1.3137 | 2.5474 | 100.3 | 20.97 | 32.114 | 95.7377 | 81.1085 | 91.2147 | 0.007957 | 0.96893 | 0.023215 | 23.9592 | 0.039503 | 0.15227 | 0.002249 | 0.15174 | 0.15174 | 0.040842
2019-03-29 11:45 | 1553874306.889 | 587 | 0.20906 | 0.000706 | 8.9e-05 | 99.7581 | 99.3192 | 118.6994 | 21.1141 | 0.003074 | 0.002862 | 134.1308 | 0.25501 | 8.4019 | 0.002466 | 0.004437 | 0.00439 | 6e-05 | 0.0009 | 0.29662 | 1.3446 | 2.5474 | 100.3 | 20.97 | 32.112 | 95.6422 | 80.9915 | 90.5339 | 0.007957 | 0.97075 | 0.023252 | 23.9525 | 0.039432 | 0.15171 | 0.002219 | 0.14993 | 0.14993 | 0.041182
2019-03-29 11:45 | 1553874307.8901 | 588 | 0.20903 | 0.000709 | 8.9e-05 | 99.7682 | 99.4667 | 119.0442 | 21.1156 | 0.003076 | 0.002804 | 134.1309 | 0.25572 | 8.4003 | 0.002437 | 0.003714 | 0.003688 | 5.8e-05 | 0.000901 | 0.3027 | 1.3926 | 2.5474 | 100.3 | 20.97 | 32.111 | 95.6813 | 80.3044 | 89.1438 | 0.007957 | 0.97447 | 0.02333 | 23.9413 | 0.027076 | 0.15213 | 0.002377 | 0.16152 | 0.16152 | 0.038333
2019-03-29 11:45 | 1553874308.89 | 589 | 0.20907 | 0.00071 | 8.8e-05 | 99.7467 | 99.4684 | 118.8942 | 21.1156 | 0.003083 | 0.0028 | 134.1202 | 0.25559 | 8.4035 | 0.002464 | 0.003876 | 0.003748 | 5.8e-05 | 0.000901 | 0.31365 | 1.3889 | 2.5475 | 100.3 | 20.97 | 32.106 | 95.87 | 80.9602 | 89.0383 | 0.007955 | 0.97476 | 0.023319 | 23.9228 | 0.02499 | 0.15205 | 0.002147 | 0.14533 | 0.14533 | 0.042582
2019-03-29 11:45 | 1553874309.8905 | 590 | 0.20905 | 0.000708 | 8.8e-05 | 99.7371 | 99.4244 | 118.8104 | 21.1154 | 0.003041 | 0.002841 | 134.1098 | 0.25585 | 8.4013 | 0.002534 | 0.003633 | 0.003562 | 6e-05 | 0.000897 | 0.30757 | 1.3462 | 2.5475 | 100.3 | 20.97 | 32.104 | 94.8435 | 82.6432 | 90.0352 | 0.007955 | 0.97208 | 0.023279 | 23.9472 | 0.028091 | 0.15221 | 0.00228 | 0.15443 | 0.15443 | 0.040114
2019-03-29 11:45 | 1553874310.8895 | 591 | 0.20905 | 0.000712 | 8.7e-05 | 99.7485 | 99.2421 | 118.89 | 21.1158 | 0.003005 | 0.002802 | 134.0881 | 0.25493 | 8.4015 | 0.002411 | 0.004189 | 0.004102 | 6.2e-05 | 0.000895 | 0.28102 | 1.3223 | 2.5474 | 100.3 | 20.97 | 32.104 | 93.9861 | 79.6855 | 89.0821 | 0.007955 | 0.97464 | 0.023317 | 23.9238 | 0.045522 | 0.15166 | 0.002273 | 0.15427 | 0.15427 | 0.040012
2019-03-29 11:45 | 1553874311.8905 | 592 | 0.20902 | 0.000706 | 8.7e-05 | 99.6871 | 99.4496 | 119.2033 | 21.1163 | 0.003038 | 0.00277 | 134.0726 | 0.25638 | 8.4035 | 0.002491 | 0.003839 | 0.003808 | 6.2e-05 | 0.000898 | 0.28727 | 1.3539 | 2.5473 | 100.3 | 20.97 | 32.102 | 94.7683 | 81.5932 | 88.3266 | 0.007954 | 0.97668 | 0.023372 | 23.9303 | 0.021335 | 0.15252 | 0.002474 | 0.16876 | 0.16876 | 0.036785
2019-03-29 11:45 | 1553874312.8892 | 593 | 0.20902 | 0.000719 | 8.6e-05 | 99.7045 | 99.3175 | 119.6441 | 21.1157 | 0.003026 | 0.002815 | 134.0807 | 0.2559 | 8.402 | 0.002417 | 0.003362 | 0.003233 | 6e-05 | 0.000898 | 0.2756 | 1.3707 | 2.5474 | 100.3 | 20.97 | 32.103 | 94.4759 | 79.8368 | 89.4033 | 0.007955 | 0.97377 | 0.023381 | 24.0102 | 0.03478 | 0.15224 | 0.002473 | 0.16876 | 0.16876 | 0.036714
2019-03-29 11:45 | 1553874313.8906 | 594 | 0.209 | 0.000714 | 8.6e-05 | 99.7018 | 99.4431 | 119.7291 | 21.1159 | 0.003087 | 0.00284 | 134.0892 | 0.25579 | 8.4055 | 0.002526 | 0.003683 | 0.003561 | 6.2e-05 | 0.000895 | 0.29856 | 1.3568 | 2.5473 | 100.3 | 20.97 | 32.105 | 95.9635 | 82.4407 | 90.0074 | 0.007955 | 0.97215 | 0.023369 | 24.0387 | 0.023234 | 0.15217 | 0.00257 | 0.17555 | 0.17555 | 0.03528
2019-03-29 11:45 | 1553874314.8893 | 595 | 0.20903 | 0.000717 | 8.4e-05 | 99.7263 | 99.4602 | 119.6547 | 21.1158 | 0.003081 | 0.002815 | 134.0846 | 0.25572 | 8.4028 | 0.002531 | 0.004079 | 0.004071 | 5.9e-05 | 0.000895 | 0.31492 | 1.3528 | 2.5474 | 100.3 | 20.97 | 32.106 | 95.8005 | 82.5702 | 89.3869 | 0.007955 | 0.97382 | 0.023382 | 24.0107 | 0.023901 | 0.15213 | 0.002404 | 0.16405 | 0.16405 | 0.037741
2019-03-29 11:45 | 1553874315.8893 | 596 | 0.20902 | 0.000722 | 8.3e-05 | 99.7147 | 99.2868 | 119.9136 | 21.1164 | 0.003047 | 0.002779 | 134.0742 | 0.25611 | 8.4036 | 0.002443 | 0.00429 | 0.004149 | 6.2e-05 | 0.000892 | 0.27096 | 1.3191 | 2.5475 | 100.3 | 20.97 | 32.107 | 94.9837 | 80.4455 | 88.539 | 0.007956 | 0.9761 | 0.023435 | 24.0085 | 0.038459 | 0.15236 | 0.002419 | 0.16549 | 0.16549 | 0.03747
2019-03-29 11:45 | 1553874316.8897 | 597 | 0.20901 | 0.000725 | 8.2e-05 | 99.7066 | 99.5633 | 119.7566 | 21.1149 | 0.003014 | 0.002749 | 134.0896 | 0.25641 | 8.4009 | 0.002483 | 0.003985 | 0.003929 | 6.2e-05 | 0.000891 | 0.26034 | 1.3043 | 2.5474 | 100.3 | 20.97 | 32.109 | 94.184 | 81.3994 | 87.8003 | 0.007956 | 0.9781 | 0.023443 | 23.9683 | 0.012865 | 0.15254 | 0.002496 | 0.17108 | 0.17108 | 0.036289
2019-03-29 11:45 | 1553874317.8884 | 598 | 0.20901 | 0.000729 | 8.2e-05 | 99.7293 | 99.5066 | 119.7864 | 21.1162 | 0.003053 | 0.002791 | 134.0852 | 0.25567 | 8.4017 | 0.002443 | 0.004096 | 0.004023 | 6.1e-05 | 0.00089 | 0.28816 | 1.3072 | 2.5472 | 100.3 | 20.97 | 32.109 | 95.1283 | 80.4619 | 88.8099 | 0.007956 | 0.97537 | 0.023414 | 24.0048 | 0.019997 | 0.1521 | 0.00252 | 0.17255 | 0.17255 | 0.035875
2019-03-29 11:45 | 1553874318.8884 | 599 | 0.20902 | 0.00073 | 8.2e-05 | 99.7012 | 99.3807 | 119.6465 | 21.1173 | 0.003023 | 0.002819 | 134.0832 | 0.25607 | 8.4024 | 0.00252 | 0.004334 | 0.004217 | 6.2e-05 | 0.000892 | 0.27903 | 1.3231 | 2.5474 | 100.3 | 20.97 | 32.109 | 94.4092 | 82.2955 | 89.4845 | 0.007956 | 0.97356 | 0.023378 | 24.0131 | 0.0288 | 0.15233 | 0.002449 | 0.16733 | 0.16733 | 0.037052
2019-03-29 11:45 | 1553874319.8884 | 600 | 0.209 | 0.000732 | 8.1e-05 | 99.7799 | 99.4093 | 119.6279 | 21.1162 | 0.00296 | 0.002771 | 134.0771 | 0.25493 | 8.3995 | 0.002453 | 0.00379 | 0.003744 | 6.1e-05 | 0.000887 | 0.27367 | 1.2953 | 2.5475 | 100.3 | 20.97 | 32.11 | 92.8927 | 80.6864 | 88.3405 | 0.007956 | 0.97664 | 0.023413 | 23.9733 | 0.03328 | 0.15166 | 0.002576 | 0.17657 | 0.17657 | 0.034958
2019-03-29 11:45 | 1553874320.8895 | 601 | 0.20899 | 0.000731 | 8.1e-05 | 99.7401 | 99.4889 | 119.5833 | 21.1168 | 0.002997 | 0.002799 | 134.0895 | 0.25635 | 8.4031 | 0.002436 | 0.003644 | 0.003608 | 6.6e-05 | 0.00089 | 0.30807 | 1.3594 | 2.5474 | 100.3 | 20.97 | 32.111 | 93.7852 | 80.2916 | 89.0183 | 0.007957 | 0.97481 | 0.023387 | 23.9913 | 0.022558 | 0.1525 | 0.002589 | 0.17731 | 0.17731 | 0.035005
2019-03-29 11:45 | 1553874321.8898 | 602 | 0.209 | 0.00073 | 8e-05 | 99.7018 | 99.3807 | 119.511 | 21.1155 | 0.002915 | 0.002784 | 134.1012 | 0.25612 | 8.4036 | 0.002495 | 0.004172 | 0.003977 | 6.3e-05 | 0.000891 | 0.29999 | 1.3432 | 2.5474 | 100.3 | 20.97 | 32.112 | 91.8167 | 81.694 | 88.6569 | 0.007957 | 0.97578 | 0.023392 | 23.9721 | 0.028854 | 0.15237 | 0.002557 | 0.17508 | 0.17508 | 0.035419
2019-03-29 11:45 | 1553874322.8886 | 603 | 0.20901 | 0.00073 | 8e-05 | 99.7542 | 99.5097 | 119.2817 | 21.1159 | 0.003008 | 0.002759 | 134.1006 | 0.25566 | 8.4019 | 0.002534 | 0.004003 | 0.003899 | 6.6e-05 | 0.000889 | 0.27599 | 1.3426 | 2.5474 | 100.3 | 20.97 | 32.113 | 94.0542 | 82.6436 | 88.0472 | 0.007957 | 0.97743 | 0.023389 | 23.9289 | 0.021948 | 0.15209 | 0.002514 | 0.17207 | 0.17207 | 0.035974
2019-03-29 11:45 | 1553874323.8885 | 604 | 0.20905 | 0.000733 | 7.9e-05 | 99.6808 | 99.3544 | 119.1909 | 21.1174 | 0.002995 | 0.002794 | 134.0847 | 0.25581 | 8.4044 | 0.002527 | 0.003958 | 0.003797 | 6.3e-05 | 0.000889 | 0.29607 | 1.3329 | 2.5474 | 100.3 | 20.97 | 32.114 | 93.7418 | 82.4602 | 88.9057 | 0.007957 | 0.97511 | 0.023352 | 23.9482 | 0.029339 | 0.15218 | 0.002259 | 0.15399 | 0.15399 | 0.040223
2019-03-29 11:45 | 1553874324.889 | 605 | 0.20908 | 0.000731 | 7.9e-05 | 99.6569 | 99.382 | 119.2467 | 21.1171 | 0.002999 | 0.002733 | 134.1111 | 0.25555 | 8.4015 | 0.002551 | 0.004076 | 0.004018 | 6.1e-05 | 0.000887 | 0.30934 | 1.3357 | 2.5474 | 100.3 | 20.97 | 32.113 | 93.8424 | 83.0497 | 87.4316 | 0.007957 | 0.9791 | 0.023405 | 23.905 | 0.024702 | 0.15203 | 0.002095 | 0.14278 | 0.14278 | 0.043336
2019-03-29 11:45 | 1553874325.889 | 606 | 0.20908 | 0.000733 | 7.8e-05 | 99.6593 | 99.3981 | 119.1848 | 21.1169 | 0.003023 | 0.002741 | 134.1162 | 0.25665 | 8.4054 | 0.002592 | 0.003665 | 0.003536 | 6.1e-05 | 0.000888 | 0.28368 | 1.3342 | 2.5475 | 100.3 | 20.97 | 32.113 | 94.4073 | 84.04 | 87.607 | 0.007957 | 0.97862 | 0.023394 | 23.9046 | 0.023473 | 0.15268 | 0.002077 | 0.14152 | 0.14152 | 0.04391
2019-03-29 11:45 | 1553874326.8886 | 607 | 0.20907 | 0.000728 | 7.7e-05 | 99.706 | 99.3834 | 119.22 | 21.117 | 0.002953 | 0.002731 | 134.1099 | 0.25466 | 8.4004 | 0.002557 | 0.004269 | 0.004196 | 6.4e-05 | 0.000893 | 0.32117 | 1.3321 | 2.5474 | 100.3 | 20.97 | 32.111 | 92.7157 | 83.1858 | 87.3768 | 0.007957 | 0.97925 | 0.023405 | 23.9005 | 0.028984 | 0.1515 | 0.002131 | 0.14542 | 0.14542 | 0.042401
2019-03-29 11:45 | 1553874327.8887 | 608 | 0.20908 | 0.000723 | 7.7e-05 | 99.7452 | 99.367 | 119.032 | 21.1167 | 0.002951 | 0.002748 | 134.0899 | 0.25415 | 8.4052 | 0.002631 | 0.004159 | 0.004033 | 6.3e-05 | 0.000895 | 0.27959 | 1.32 | 2.5473 | 100.3 | 20.97 | 32.109 | 92.6694 | 84.9723 | 87.7936 | 0.007956 | 0.97812 | 0.023373 | 23.8955 | 0.033975 | 0.15119 | 0.002121 | 0.14452 | 0.14452 | 0.04258
2019-03-29 11:45 | 1553874328.8901 | 609 | 0.20909 | 0.000721 | 7.7e-05 | 99.6997 | 99.4824 | 119.4324 | 21.1172 | 0.002957 | 0.002716 | 134.0893 | 0.25491 | 8.4009 | 0.002575 | 0.003893 | 0.003796 | 6.3e-05 | 0.000896 | 0.26288 | 1.2973 | 2.5473 | 100.3 | 20.97 | 32.107 | 92.8263 | 83.6136 | 87.01 | 0.007956 | 0.98025 | 0.023437 | 23.9096 | 0.019517 | 0.15165 | 0.002041 | 0.13934 | 0.13934 | 0.044295
2019-03-29 11:45 | 1553874329.8887 | 610 | 0.20907 | 0.000715 | 7.6e-05 | 99.7329 | 99.4698 | 119.5042 | 21.1176 | 0.002979 | 0.002761 | 134.0915 | 0.25525 | 8.4013 | 0.00254 | 0.004047 | 0.00394 | 5.9e-05 | 0.000896 | 0.2975 | 1.3278 | 2.5474 | 100.3 | 20.97 | 32.106 | 93.3446 | 82.7777 | 88.1013 | 0.007955 | 0.97728 | 0.023409 | 23.953 | 0.023632 | 0.15185 | 0.002171 | 0.14829 | 0.14829 | 0.041676
2019-03-29 11:45 | 1553874330.8888 | 611 | 0.20911 | 0.000717 | 7.6e-05 | 99.7302 | 99.4244 | 119.6584 | 21.1187 | 0.002934 | 0.002682 | 134.0847 | 0.25545 | 8.4034 | 0.002576 | 0.003816 | 0.003661 | 6.2e-05 | 0.000899 | 0.28838 | 1.296 | 2.5473 | 100.3 | 20.97 | 32.103 | 92.2731 | 83.6393 | 86.2035 | 0.007955 | 0.98245 | 0.023486 | 23.9054 | 0.027474 | 0.15197 | 0.001919 | 0.13103 | 0.13103 | 0.047203
2019-03-29 11:46 | 1553874331.8898 | 612 | 0.20908 | 0.000712 | 7.6e-05 | 99.7167 | 99.3892 | 119.6472 | 21.1176 | 0.002955 | 0.002749 | 134.0808 | 0.25573 | 8.4007 | 0.002533 | 0.004417 | 0.004352 | 6.1e-05 | 0.000898 | 0.28711 | 1.3244 | 2.5473 | 100.3 | 20.97 | 32.101 | 92.7742 | 82.6172 | 87.821 | 0.007954 | 0.97804 | 0.023432 | 23.958 | 0.029428 | 0.15213 | 0.002083 | 0.14228 | 0.14228 | 0.043519
2019-03-29 11:46 | 1553874332.8888 | 613 | 0.2091 | 0.000707 | 7.5e-05 | 99.8152 | 99.3445 | 119.6882 | 21.118 | 0.002954 | 0.002738 | 134.0794 | 0.25504 | 8.4006 | 0.002504 | 0.003819 | 0.003648 | 6.4e-05 | 0.000893 | 0.27544 | 1.3824 | 2.5475 | 100.3 | 20.97 | 32.099 | 92.7504 | 81.9218 | 87.5356 | 0.007954 | 0.97882 | 0.023445 | 23.9527 | 0.042282 | 0.15172 | 0.001999 | 0.13648 | 0.13648 | 0.045245
2019-03-29 11:46 | 1553874333.8893 | 614 | 0.20907 | 0.0007 | 7.5e-05 | 99.7129 | 99.4558 | 119.8908 | 21.1168 | 0.002937 | 0.002735 | 134.0732 | 0.25563 | 8.4031 | 0.002483 | 0.003883 | 0.003713 | 6.1e-05 | 0.000887 | 0.27129 | 1.3572 | 2.5475 | 100.3 | 20.97 | 32.097 | 92.3378 | 81.4145 | 87.464 | 0.007953 | 0.97901 | 0.023467 | 23.9706 | 0.023092 | 0.15207 | 0.002188 | 0.14996 | 0.14996 | 0.041274
2019-03-29 11:46 | 1553874334.8879 | 615 | 0.20912 | 0.000699 | 7.5e-05 | 99.811 | 99.4366 | 119.9669 | 21.1175 | 0.002926 | 0.002739 | 134.0499 | 0.25564 | 8.3995 | 0.002516 | 0.003671 | 0.003524 | 6.2e-05 | 0.000887 | 0.27184 | 1.3306 | 2.5474 | 100.3 | 20.97 | 32.095 | 92.0625 | 82.2004 | 87.5679 | 0.007953 | 0.97873 | 0.023472 | 23.9816 | 0.033614 | 0.15208 | 0.001898 | 0.12947 | 0.12947 | 0.047807
2019-03-29 11:46 | 1553874335.888 | 616 | 0.20911 | 0.000697 | 7.5e-05 | 99.7069 | 99.5779 | 120.0077 | 21.1179 | 0.002945 | 0.002736 | 134.0471 | 0.25651 | 8.4007 | 0.002473 | 0.003739 | 0.003637 | 6e-05 | 0.000884 | 0.28296 | 1.3171 | 2.5475 | 100.3 | 20.97 | 32.093 | 92.5329 | 81.1848 | 87.5043 | 0.007952 | 0.9789 | 0.023478 | 23.9836 | 0.011574 | 0.1526 | 0.001967 | 0.13447 | 0.13447 | 0.046187
2019-03-29 11:46 | 1553874336.8883 | 617 | 0.20911 | 0.000696 | 7.5e-05 | 99.8146 | 99.6732 | 119.8291 | 21.1177 | 0.003002 | 0.002747 | 134.0493 | 0.25555 | 8.4041 | 0.002507 | 0.003822 | 0.003712 | 6.1e-05 | 0.000881 | 0.2943 | 1.3767 | 2.5475 | 100.3 | 20.97 | 32.09 | 93.9112 | 81.9848 | 87.7616 | 0.007951 | 0.9782 | 0.023452 | 23.9743 | 0.012686 | 0.15202 | 0.00195 | 0.1331 | 0.1331 | 0.046487
2019-03-29 11:46 | 1553874337.8894 | 618 | 0.20908 | 0.000693 | 7.4e-05 | 99.7778 | 99.5513 | 119.8889 | 21.1168 | 0.002897 | 0.002755 | 134.0671 | 0.25566 | 8.4034 | 0.002546 | 0.00401 | 0.004009 | 6.4e-05 | 0.000883 | 0.28982 | 1.3354 | 2.5474 | 100.3 | 20.97 | 32.089 | 91.3689 | 82.9291 | 87.9443 | 0.007951 | 0.97771 | 0.023452 | 23.9863 | 0.020327 | 0.15209 | 0.002129 | 0.14572 | 0.14572 | 0.04248
2019-03-29 11:46 | 1553874338.889 | 619 | 0.20906 | 0.00069 | 7.4e-05 | 99.7494 | 99.5015 | 119.8096 | 21.1178 | 0.002871 | 0.002725 | 134.0622 | 0.25559 | 8.4014 | 0.002498 | 0.003636 | 0.003533 | 6.1e-05 | 0.000877 | 0.28368 | 1.2931 | 2.5473 | 100.3 | 20.97 | 32.088 | 90.743 | 81.7743 | 87.2262 | 0.007951 | 0.97966 | 0.023467 | 23.9545 | 0.022255 | 0.15205 | 0.002265 | 0.15543 | 0.15543 | 0.039815
2019-03-29 11:46 | 1553874339.888 | 620 | 0.20911 | 0.000695 | 7.4e-05 | 99.7757 | 99.6431 | 119.4901 | 21.1185 | 0.002952 | 0.002735 | 134.053 | 0.25585 | 8.4016 | 0.002526 | 0.003746 | 0.003676 | 6e-05 | 0.000883 | 0.28943 | 1.2914 | 2.5474 | 100.3 | 20.97 | 32.087 | 92.7078 | 82.4425 | 87.4774 | 0.007951 | 0.97898 | 0.023428 | 23.9309 | 0.011894 | 0.15221 | 0.001947 | 0.1328 | 0.1328 | 0.046646
2019-03-29 11:46 | 1553874340.8895 | 621 | 0.20909 | 0.000696 | 7.3e-05 | 99.7533 | 99.453 | 119.1136 | 21.1181 | 0.002939 | 0.002755 | 134.057 | 0.25628 | 8.405 | 0.002542 | 0.003373 | 0.00327 | 6.3e-05 | 0.000873 | 0.30541 | 1.364 | 2.5476 | 100.3 | 20.96 | 32.087 | 92.3892 | 82.8259 | 87.9635 | 0.007946 | 0.97766 | 0.023375 | 23.9093 | 0.026962 | 0.15246 | 0.002048 | 0.13959 | 0.13959 | 0.044453
2019-03-29 11:46 | 1553874341.8896 | 622 | 0.20914 | 0.00069 | 7.3e-05 | 99.785 | 99.623 | 118.7564 | 21.1183 | 0.00292 | 0.002725 | 134.0512 | 0.25612 | 8.4022 | 0.002506 | 0.004364 | 0.004267 | 6.3e-05 | 0.00088 | 0.2907 | 1.3364 | 2.5475 | 100.3 | 20.96 | 32.086 | 91.924 | 81.9604 | 87.2289 | 0.007945 | 0.97965 | 0.023364 | 23.8491 | 0.014534 | 0.15237 | 0.001771 | 0.12015 | 0.12015 | 0.051616
2019-03-29 11:46 | 1553874342.8892 | 623 | 0.20911 | 0.000695 | 7.2e-05 | 99.8694 | 99.5377 | 118.5825 | 21.1189 | 0.002879 | 0.002701 | 134.0551 | 0.25523 | 8.4016 | 0.002586 | 0.004146 | 0.004041 | 6e-05 | 0.000882 | 0.26056 | 1.3821 | 2.5474 | 100.3 | 20.96 | 32.084 | 90.9474 | 83.8946 | 86.6579 | 0.007945 | 0.9812 | 0.023365 | 23.8127 | 0.029761 | 0.15183 | 0.001956 | 0.13318 | 0.13318 | 0.046402
2019-03-29 11:46 | 1553874343.8886 | 624 | 0.20918 | 0.000688 | 7.2e-05 | 99.823 | 99.4001 | 118.4646 | 21.1184 | 0.002865 | 0.002683 | 134.072 | 0.25539 | 8.4038 | 0.002529 | 0.003887 | 0.003779 | 6.5e-05 | 0.000883 | 0.27616 | 1.3741 | 2.5476 | 100.3 | 20.96 | 32.083 | 90.6022 | 82.529 | 86.2337 | 0.007945 | 0.98236 | 0.023367 | 23.7869 | 0.037972 | 0.15193 | 0.001531 | 0.10338 | 0.10338 | 0.059817
2019-03-29 11:46 | 1553874344.8882 | 625 | 0.20909 | 0.000686 | 7.2e-05 | 99.8904 | 99.5916 | 118.2677 | 21.1187 | 0.002867 | 0.002729 | 134.0726 | 0.25542 | 8.4019 | 0.002554 | 0.003866 | 0.003696 | 6.6e-05 | 0.000884 | 0.27273 | 1.3567 | 2.5475 | 100.3 | 20.96 | 32.081 | 90.6454 | 83.126 | 87.3347 | 0.007944 | 0.97936 | 0.023312 | 23.8035 | 0.026793 | 0.15195 | 0.002086 | 0.14196 | 0.14196 | 0.043562
2019-03-29 11:46 | 1553874345.8892 | 626 | 0.20916 | 0.000684 | 7.2e-05 | 99.835 | 99.5551 | 118.0452 | 21.1197 | 0.002841 | 0.002676 | 134.0731 | 0.25663 | 8.4039 | 0.00247 | 0.00443 | 0.004389 | 6.5e-05 | 0.000891 | 0.28224 | 1.3449 | 2.5475 | 100.3 | 20.96 | 32.079 | 90.0243 | 81.0952 | 86.0553 | 0.007944 | 0.98285 | 0.023332 | 23.7388 | 0.025113 | 0.15267 | 0.001656 | 0.11195 | 0.11195 | 0.055502
2019-03-29 11:46 | 1553874346.8883 | 627 | 0.20913 | 0.000684 | 7.3e-05 | 99.8257 | 99.5585 | 118.1036 | 21.1187 | 0.002911 | 0.00268 | 134.06 | 0.25598 | 8.4032 | 0.00248 | 0.004001 | 0.00394 | 6.5e-05 | 0.000891 | 0.30613 | 1.3577 | 2.5475 | 100.3 | 20.96 | 32.077 | 91.7062 | 81.3391 | 86.1505 | 0.007943 | 0.98259 | 0.023334 | 23.7479 | 0.023975 | 0.15228 | 0.00186 | 0.12624 | 0.12624 | 0.049095
2019-03-29 11:46 | 1553874347.8887 | 628 | 0.20919 | 0.000675 | 7.2e-05 | 99.9359 | 99.7489 | 118.0663 | 21.1193 | 0.002888 | 0.002699 | 134.0571 | 0.25568 | 8.4036 | 0.002437 | 0.004836 | 0.00466 | 6.4e-05 | 0.000893 | 0.3027 | 1.3802 | 2.5475 | 100.3 | 20.96 | 32.077 | 91.1595 | 80.3204 | 86.6022 | 0.007943 | 0.98136 | 0.023316 | 23.759 | 0.016746 | 0.15211 | 0.001503 | 0.10116 | 0.10116 | 0.061195
2019-03-29 11:46 | 1553874348.8878 | 629 | 0.20916 | 0.000679 | 7.2e-05 | 99.8778 | 99.5111 | 118.0142 | 21.1193 | 0.002817 | 0.002655 | 134.0601 | 0.2556 | 8.4031 | 0.002466 | 0.004098 | 0.00407 | 6.3e-05 | 0.000899 | 0.32278 | 1.3465 | 2.5474 | 100.3 | 20.96 | 32.081 | 89.4457 | 81.0129 | 85.5555 | 0.007944 | 0.98422 | 0.023345 | 23.7192 | 0.032904 | 0.15205 | 0.001654 | 0.11185 | 0.11185 | 0.055329
2019-03-29 11:46 | 1553874349.8884 | 630 | 0.20907 | 0.000678 | 7.2e-05 | 99.9347 | 99.567 | 118.256 | 21.1192 | 0.002898 | 0.002676 | 134.0652 | 0.25584 | 8.3996 | 0.002488 | 0.003802 | 0.003682 | 6e-05 | 0.000897 | 0.29286 | 1.3229 | 2.5474 | 100.3 | 20.96 | 32.084 | 91.4057 | 81.5463 | 86.0487 | 0.007945 | 0.98287 | 0.023353 | 23.7598 | 0.032966 | 0.1522 | 0.002193 | 0.14975 | 0.14975 | 0.041366
2019-03-29 11:46 | 1553874350.8874 | 631 | 0.2091 | 0.000671 | 7.3e-05 | 99.9538 | 99.8407 | 118.3916 | 21.1194 | 0.00283 | 0.002671 | 134.0602 | 0.25432 | 8.4023 | 0.002461 | 0.004293 | 0.004209 | 6.5e-05 | 0.000898 | 0.29391 | 1.3672 | 2.5475 | 100.3 | 20.95 | 32.089 | 89.7599 | 80.8813 | 85.9357 | 0.007941 | 0.98318 | 0.02337 | 23.7697 | 0.010125 | 0.15129 | 0.002073 | 0.14139 | 0.14139 | 0.043551
2019-03-29 11:46 | 1553874351.8874 | 632 | 0.20922 | 0.000673 | 7.3e-05 | 99.9867 | 99.7769 | 118.4578 | 21.1187 | 0.002852 | 0.002656 | 134.0599 | 0.25611 | 8.403 | 0.002462 | 0.003963 | 0.003935 | 6.4e-05 | 0.000895 | 0.26106 | 1.382 | 2.5475 | 100.3 | 20.95 | 32.099 | 90.2991 | 80.9132 | 85.5666 | 0.007944 | 0.98419 | 0.023388 | 23.7641 | 0.018791 | 0.15236 | 0.001309 | 0.087757 | 0.087757 | 0.07066
2019-03-29 11:46 | 1553874352.8879 | 633 | 0.20921 | 0.000673 | 7.3e-05 | 99.9966 | 99.8322 | 118.3093 | 21.1183 | 0.002843 | 0.002708 | 134.0254 | 0.25614 | 8.4031 | 0.002553 | 0.004252 | 0.004141 | 6.2e-05 | 0.000897 | 0.26576 | 1.3292 | 2.5475 | 100.3 | 20.95 | 32.104 | 90.0784 | 83.0919 | 86.8277 | 0.007945 | 0.98074 | 0.023333 | 23.7909 | 0.014719 | 0.15238 | 0.001406 | 0.09434 | 0.09434 | 0.065739
2019-03-29 11:46 | 1553874353.8875 | 634 | 0.2091 | 0.000671 | 7.2e-05 | 99.9787 | 99.6759 | 118.3508 | 21.1198 | 0.00283 | 0.002664 | 133.998 | 0.25514 | 8.4022 | 0.002532 | 0.003681 | 0.003568 | 6.3e-05 | 0.000905 | 0.29319 | 1.3159 | 2.5475 | 100.3 | 20.95 | 32.109 | 89.7641 | 82.5904 | 85.7683 | 0.007946 | 0.98364 | 0.023371 | 23.76 | 0.027128 | 0.15178 | 0.002061 | 0.14059 | 0.14059 | 0.04394
2019-03-29 11:46 | 1553874354.8885 | 635 | 0.20909 | 0.000671 | 7.3e-05 | 100.0233 | 99.7749 | 118.4973 | 21.1193 | 0.002895 | 0.002647 | 133.9901 | 0.25539 | 8.4008 | 0.002554 | 0.003955 | 0.003848 | 6e-05 | 0.000909 | 0.30143 | 1.3162 | 2.5475 | 100.3 | 20.95 | 32.123 | 91.3393 | 83.1297 | 85.3517 | 0.00795 | 0.98478 | 0.023399 | 23.7609 | 0.022241 | 0.15193 | 0.002109 | 0.14411 | 0.14411 | 0.04291
2019-03-29 11:46 | 1553874355.8888 | 636 | 0.20893 | 0.00067 | 7.2e-05 | 100.0224 | 99.7223 | 118.6395 | 21.1197 | 0.002854 | 0.002687 | 133.9893 | 0.25703 | 8.4021 | 0.002463 | 0.003806 | 0.003694 | 6.3e-05 | 0.000908 | 0.29087 | 1.342 | 2.5476 | 100.3 | 20.95 | 32.131 | 90.3465 | 80.9431 | 86.3168 | 0.007952 | 0.98214 | 0.023382 | 23.8072 | 0.026873 | 0.15291 | 0.003062 | 0.21097 | 0.21097 | 0.029498
2019-03-29 11:46 | 1553874356.888 | 637 | 0.20917 | 0.000666 | 7.3e-05 | 100.0556 | 99.9793 | 118.9963 | 21.1197 | 0.002862 | 0.002665 | 133.9885 | 0.25528 | 8.4043 | 0.002461 | 0.003688 | 0.003625 | 6e-05 | 0.000908 | 0.27223 | 1.3953 | 2.5474 | 100.3 | 20.95 | 32.147 | 90.5269 | 80.8979 | 85.7823 | 0.007956 | 0.9836 | 0.023434 | 23.8252 | 0.006823 | 0.15187 | 0.00161 | 0.10919 | 0.10919 | 0.056609
2019-03-29 11:46 | 1553874357.8887 | 638 | 0.20907 | 0.000669 | 7.2e-05 | 100.0502 | 99.7523 | 119.036 | 21.1195 | 0.002811 | 0.002692 | 133.9678 | 0.2557 | 8.4035 | 0.002497 | 0.003412 | 0.003237 | 5.9e-05 | 0.000913 | 0.26001 | 1.3618 | 2.5475 | 100.3 | 20.95 | 32.156 | 89.2976 | 81.7634 | 86.4354 | 0.007958 | 0.98181 | 0.023417 | 23.8509 | 0.026669 | 0.15212 | 0.002228 | 0.15266 | 0.15266 | 0.040557
2019-03-29 11:46 | 1553874358.8891 | 639 | 0.20913 | 0.00067 | 7.1e-05 | 100.0457 | 99.7544 | 119.3988 | 21.1194 | 0.00284 | 0.00267 | 133.963 | 0.25657 | 8.4035 | 0.002522 | 0.004493 | 0.004429 | 6.3e-05 | 0.000917 | 0.29966 | 1.331 | 2.5472 | 100.3 | 20.95 | 32.164 | 90.006 | 82.3436 | 85.9044 | 0.00796 | 0.98326 | 0.02347 | 23.8695 | 0.026084 | 0.15264 | 0.001856 | 0.12677 | 0.12677 | 0.049002
2019-03-29 11:46 | 1553874359.8887 | 640 | 0.20924 | 0.000669 | 7.2e-05 | 100.1199 | 99.9196 | 119.466 | 21.1193 | 0.002868 | 0.00265 | 133.9571 | 0.25573 | 8.4019 | 0.002484 | 0.003844 | 0.003727 | 5.9e-05 | 0.000918 | 0.30735 | 1.3338 | 2.5475 | 100.3 | 20.95 | 32.175 | 90.6719 | 81.431 | 85.4306 | 0.007963 | 0.98456 | 0.023492 | 23.8605 | 0.017919 | 0.15213 | 0.001217 | 0.081737 | 0.081737 | 0.075753
2019-03-29 11:46 | 1553874360.8891 | 641 | 0.20913 | 0.000667 | 7.1e-05 | 100.1205 | 99.8578 | 119.7256 | 21.1187 | 0.002868 | 0.002646 | 133.9533 | 0.2558 | 8.4013 | 0.002527 | 0.003447 | 0.003367 | 6.2e-05 | 0.000921 | 0.2756 | 1.3318 | 2.5476 | 100.3 | 20.95 | 32.18 | 90.6728 | 82.4715 | 85.341 | 0.007964 | 0.98481 | 0.023521 | 23.8834 | 0.023504 | 0.15218 | 0.001874 | 0.12836 | 0.12836 | 0.04825
2019-03-29 11:46 | 1553874361.8878 | 642 | 0.20919 | 0.000666 | 7e-05 | 100.1406 | 99.7254 | 119.9004 | 21.1192 | 0.002768 | 0.002642 | 133.9433 | 0.25632 | 8.4026 | 0.002545 | 0.004963 | 0.004884 | 5.8e-05 | 0.000919 | 0.2891 | 1.3437 | 2.5474 | 100.3 | 20.95 | 32.185 | 88.2716 | 82.8995 | 85.2323 | 0.007965 | 0.98511 | 0.023541 | 23.8972 | 0.037165 | 0.15249 | 0.001543 | 0.10508 | 0.10508 | 0.059059
2019-03-29 11:46 | 1553874362.8888 | 643 | 0.20912 | 0.000665 | 7e-05 | 100.1355 | 99.8189 | 119.9892 | 21.1198 | 0.002778 | 0.002631 | 133.9702 | 0.25577 | 8.4013 | 0.002412 | 0.003416 | 0.003319 | 5.9e-05 | 0.000916 | 0.29015 | 1.3756 | 2.5475 | 100.3 | 20.95 | 32.187 | 88.5111 | 79.7082 | 84.9847 | 0.007966 | 0.98579 | 0.023558 | 23.8978 | 0.028327 | 0.15216 | 0.001953 | 0.13418 | 0.13418 | 0.046153
2019-03-29 11:46 | 1553874363.8868 | 644 | 0.20913 | 0.000666 | 7e-05 | 100.2151 | 99.9063 | 120.0223 | 21.1195 | 0.002842 | 0.002657 | 133.9723 | 0.25562 | 8.3997 | 0.002551 | 0.004039 | 0.00389 | 5.8e-05 | 0.000917 | 0.28456 | 1.3591 | 2.5474 | 100.3 | 20.95 | 32.189 | 90.0422 | 83.0433 | 85.5946 | 0.007966 | 0.98411 | 0.023541 | 23.9214 | 0.027611 | 0.15207 | 0.001902 | 0.13048 | 0.13048 | 0.047436
2019-03-29 11:46 | 1553874364.8873 | 645 | 0.20918 | 0.00067 | 7e-05 | 100.2313 | 100.1503 | 120.0105 | 21.1206 | 0.002868 | 0.002637 | 133.984 | 0.25638 | 8.4022 | 0.002545 | 0.004015 | 0.003966 | 5.7e-05 | 0.000918 | 0.27223 | 1.3222 | 2.5474 | 100.3 | 20.95 | 32.192 | 90.6895 | 82.9032 | 85.1314 | 0.007967 | 0.98539 | 0.023555 | 23.9048 | 0.007233 | 0.15252 | 0.001554 | 0.10593 | 0.10593 | 0.058604
2019-03-29 11:46 | 1553874365.8869 | 646 | 0.20918 | 0.000669 | 7e-05 | 100.2867 | 99.9428 | 120.1866 | 21.1198 | 0.002805 | 0.002647 | 133.9736 | 0.25516 | 8.4023 | 0.002516 | 0.004162 | 0.004078 | 5.6e-05 | 0.000913 | 0.2896 | 1.3254 | 2.5474 | 100.3 | 20.95 | 32.192 | 89.1633 | 82.2152 | 85.3561 | 0.007967 | 0.98477 | 0.023565 | 23.9299 | 0.030725 | 0.1518 | 0.001574 | 0.10737 | 0.10737 | 0.057539
2019-03-29 11:46 | 1553874366.8869 | 647 | 0.20913 | 0.000663 | 7.1e-05 | 100.2454 | 99.9824 | 120.0556 | 21.121 | 0.002859 | 0.00262 | 133.9581 | 0.25622 | 8.3972 | 0.002516 | 0.004057 | 0.003942 | 5.3e-05 | 0.000907 | 0.28136 | 1.297 | 2.5474 | 100.3 | 20.95 | 32.191 | 90.4546 | 82.2063 | 84.7136 | 0.007967 | 0.98654 | 0.023574 | 23.8954 | 0.023497 | 0.15242 | 0.001867 | 0.12812 | 0.12812 | 0.048422
2019-03-29 11:46 | 1553874367.887 | 648 | 0.20917 | 0.000664 | 7.1e-05 | 100.3513 | 100.0312 | 120.1063 | 21.1207 | 0.002819 | 0.002664 | 133.9588 | 0.25462 | 8.4027 | 0.002509 | 0.003362 | 0.003263 | 5.8e-05 | 0.000902 | 0.29031 | 1.3529 | 2.5475 | 100.3 | 20.95 | 32.19 | 89.4916 | 82.0336 | 85.7692 | 0.007966 | 0.98364 | 0.023544 | 23.9356 | 0.028582 | 0.15147 | 0.001639 | 0.11185 | 0.11185 | 0.055118
2019-03-29 11:46 | 1553874368.8871 | 649 | 0.20915 | 0.000663 | 7e-05 | 100.3352 | 100.1769 | 119.6742 | 21.1199 | 0.00279 | 0.002669 | 133.9539 | 0.25624 | 8.4023 | 0.002474 | 0.003759 | 0.003577 | 6.1e-05 | 0.000907 | 0.28744 | 1.2956 | 2.5474 | 100.3 | 20.96 | 32.189 | 88.7948 | 81.2094 | 85.8816 | 0.007971 | 0.98333 | 0.023498 | 23.8963 | 0.01412 | 0.15244 | 0.001775 | 0.1213 | 0.1213 | 0.051146
2019-03-29 11:46 | 1553874369.887 | 650 | 0.20916 | 0.000663 | 7.1e-05 | 100.3678 | 100.238 | 119.5411 | 21.1203 | 0.002802 | 0.002584 | 133.9495 | 0.25571 | 8.4013 | 0.002536 | 0.004136 | 0.003991 | 5.7e-05 | 0.000907 | 0.26648 | 1.3212 | 2.5475 | 100.3 | 20.96 | 32.184 | 89.0787 | 82.6923 | 83.8366 | 0.00797 | 0.98896 | 0.023552 | 23.8149 | 0.011575 | 0.15212 | 0.00169 | 0.1155 | 0.1155 | 0.053606
2019-03-29 11:46 | 1553874370.8881 | 651 | 0.20914 | 0.000665 | 7.1e-05 | 100.3145 | 100.1687 | 119.4728 | 21.1208 | 0.002802 | 0.002595 | 133.9419 | 0.25627 | 8.4043 | 0.002429 | 0.003739 | 0.003655 | 6e-05 | 0.000908 | 0.26698 | 1.2979 | 2.5475 | 100.3 | 20.96 | 32.182 | 89.0939 | 80.1193 | 84.1203 | 0.007969 | 0.98817 | 0.023536 | 23.8175 | 0.01301 | 0.15245 | 0.001784 | 0.12206 | 0.12206 | 0.050834
2019-03-29 11:46 | 1553874371.8871 | 652 | 0.20914 | 0.000665 | 7.1e-05 | 100.4106 | 100.2428 | 119.7828 | 21.1199 | 0.002771 | 0.002664 | 133.9368 | 0.2564 | 8.4016 | 0.002463 | 0.003659 | 0.003565 | 5.8e-05 | 0.000907 | 0.27776 | 1.3405 | 2.5476 | 100.3 | 20.96 | 32.176 | 88.3458 | 80.9473 | 85.7605 | 0.007968 | 0.98366 | 0.023512 | 23.9031 | 0.014962 | 0.15253 | 0.001831 | 0.12522 | 0.12522 | 0.049575
2019-03-29 11:46 | 1553874372.8882 | 653 | 0.20912 | 0.000663 | 7.1e-05 | 100.3881 | 100.2977 | 119.9023 | 21.1202 | 0.002822 | 0.002678 | 133.9406 | 0.25566 | 8.4004 | 0.00245 | 0.003666 | 0.003601 | 5.9e-05 | 0.000904 | 0.28241 | 1.3738 | 2.5474 | 100.3 | 20.96 | 32.172 | 89.5614 | 80.6239 | 86.0981 | 0.007967 | 0.98273 | 0.023513 | 23.9263 | 0.008061 | 0.15209 | 0.001914 | 0.13113 | 0.13113 | 0.047205
2019-03-29 11:46 | 1553874373.8881 | 654 | 0.20912 | 0.000665 | 7.1e-05 | 100.4609 | 100.3069 | 119.9145 | 21.1206 | 0.002819 | 0.002641 | 133.9368 | 0.25603 | 8.4015 | 0.00252 | 0.003968 | 0.00385 | 6e-05 | 0.000907 | 0.29142 | 1.3732 | 2.5474 | 100.3 | 20.96 | 32.169 | 89.4947 | 82.315 | 85.2212 | 0.007966 | 0.98514 | 0.023543 | 23.8983 | 0.013718 | 0.15231 | 0.001929 | 0.13236 | 0.13236 | 0.046835
2019-03-29 11:46 | 1553874374.8872 | 655 | 0.20915 | 0.00067 | 7.1e-05 | 100.4953 | 100.3684 | 119.9241 | 21.1199 | 0.0028 | 0.002626 | 133.9352 | 0.25627 | 8.4038 | 0.002446 | 0.003583 | 0.003393 | 5.8e-05 | 0.000899 | 0.26576 | 1.3603 | 2.5475 | 100.3 | 20.96 | 32.161 | 89.0377 | 80.5325 | 84.8429 | 0.007964 | 0.98618 | 0.023556 | 23.8866 | 0.011306 | 0.15246 | 0.001749 | 0.11971 | 0.11971 | 0.051836
2019-03-29 11:46 | 1553874375.8872 | 656 | 0.20916 | 0.000668 | 7e-05 | 100.5336 | 100.2967 | 119.9535 | 21.1205 | 0.002846 | 0.002627 | 133.9492 | 0.25567 | 8.404 | 0.002486 | 0.003501 | 0.00335 | 5.6e-05 | 0.000899 | 0.28655 | 1.3421 | 2.5474 | 100.3 | 20.96 | 32.157 | 90.1455 | 81.4969 | 84.887 | 0.007963 | 0.98606 | 0.023558 | 23.891 | 0.021104 | 0.1521 | 0.001702 | 0.11641 | 0.11641 | 0.053177
2019-03-29 11:46 | 1553874376.8883 | 657 | 0.20916 | 0.00067 | 7.1e-05 | 100.5411 | 100.3394 | 119.9356 | 21.1204 | 0.002679 | 0.002664 | 133.9624 | 0.25629 | 8.4027 | 0.002528 | 0.004324 | 0.004207 | 5.8e-05 | 0.000895 | 0.28136 | 1.3471 | 2.5474 | 100.3 | 20.96 | 32.151 | 86.1192 | 82.4989 | 85.7641 | 0.007962 | 0.98365 | 0.023527 | 23.9185 | 0.017967 | 0.15246 | 0.001682 | 0.11479 | 0.11479 | 0.054056
2019-03-29 11:46 | 1553874377.8873 | 658 | 0.20916 | 0.000671 | 7e-05 | 100.5674 | 100.3568 | 119.7995 | 21.1205 | 0.002782 | 0.002593 | 133.9744 | 0.25618 | 8.4029 | 0.002481 | 0.004388 | 0.004293 | 5.8e-05 | 0.000895 | 0.28672 | 1.3444 | 2.5474 | 100.3 | 20.96 | 32.141 | 88.614 | 81.3628 | 84.0662 | 0.007959 | 0.98832 | 0.02357 | 23.8483 | 0.018758 | 0.1524 | 0.001698 | 0.11614 | 0.11614 | 0.053408
2019-03-29 11:46 | 1553874378.8867 | 659 | 0.20917 | 0.00067 | 7e-05 | 100.6255 | 100.4626 | 119.8066 | 21.1199 | 0.002751 | 0.002598 | 133.9766 | 0.25674 | 8.4024 | 0.002402 | 0.004216 | 0.004069 | 5.8e-05 | 0.00089 | 0.30525 | 1.3345 | 2.5473 | 100.3 | 20.96 | 32.138 | 87.8514 | 79.4783 | 84.1772 | 0.007958 | 0.98802 | 0.023567 | 23.8527 | 0.014497 | 0.15274 | 0.001644 | 0.11237 | 0.11237 | 0.055319
2019-03-29 11:46 | 1553874379.8877 | 660 | 0.20917 | 0.000666 | 7e-05 | 100.68 | 100.6557 | 119.9197 | 21.1206 | 0.002737 | 0.002594 | 133.9784 | 0.25653 | 8.4036 | 0.002481 | 0.003824 | 0.00385 | 5.9e-05 | 0.000897 | 0.29623 | 1.3216 | 2.5475 | 100.3 | 20.96 | 32.131 | 87.5226 | 81.366 | 84.0796 | 0.007957 | 0.98829 | 0.023581 | 23.8607 | 0.002156 | 0.15261 | 0.001663 | 0.11375 | 0.11375 | 0.054605
2019-03-29 11:46 | 1553874380.8864 | 661 | 0.20917 | 0.000671 | 7e-05 | 100.742 | 100.4486 | 120.1061 | 21.121 | 0.002729 | 0.002569 | 133.9942 | 0.25569 | 8.4013 | 0.002474 | 0.003484 | 0.003381 | 6.1e-05 | 0.000897 | 0.29319 | 1.3362 | 2.5475 | 100.3 | 20.96 | 32.128 | 87.3229 | 81.2087 | 83.4917 | 0.007956 | 0.98992 | 0.023619 | 23.8596 | 0.026087 | 0.15211 | 0.00161 | 0.11018 | 0.11018 | 0.05619
2019-03-29 11:46 | 1553874381.8875 | 662 | 0.20914 | 0.000668 | 6.9e-05 | 100.7252 | 100.6015 | 119.9274 | 21.1214 | 0.002788 | 0.00259 | 134.0066 | 0.25601 | 8.405 | 0.002497 | 0.004096 | 0.004059 | 5.9e-05 | 0.000889 | 0.26792 | 1.3321 | 2.5474 | 100.3 | 20.96 | 32.127 | 88.7603 | 81.7504 | 83.9913 | 0.007956 | 0.98853 | 0.023585 | 23.8585 | 0.010994 | 0.1523 | 0.001838 | 0.12622 | 0.12622 | 0.049109
2019-03-29 11:46 | 1553874382.8878 | 663 | 0.20915 | 0.000668 | 7e-05 | 100.78 | 100.6909 | 119.9251 | 21.1219 | 0.002717 | 0.002588 | 133.9791 | 0.25564 | 8.402 | 0.002387 | 0.003869 | 0.003745 | 5.9e-05 | 0.000894 | 0.2677 | 1.3174 | 2.5473 | 100.3 | 20.96 | 32.121 | 87.0413 | 79.103 | 83.9529 | 0.007954 | 0.98864 | 0.023586 | 23.857 | 0.00791 | 0.15208 | 0.001748 | 0.11983 | 0.11983 | 0.051654
2019-03-29 11:46 | 1553874383.8874 | 664 | 0.20914 | 0.000669 | 6.9e-05 | 100.8285 | 100.6349 | 120.0539 | 21.121 | 0.00281 | 0.002626 | 133.9719 | 0.25629 | 8.4018 | 0.002415 | 0.003433 | 0.003386 | 6.3e-05 | 0.000892 | 0.30646 | 1.3287 | 2.5475 | 100.3 | 20.96 | 32.119 | 89.2754 | 79.7819 | 84.854 | 0.007954 | 0.98615 | 0.023569 | 23.8999 | 0.017187 | 0.15246 | 0.001805 | 0.12381 | 0.12381 | 0.050121
2019-03-29 11:46 | 1553874384.8878 | 665 | 0.20916 | 0.000668 | 6.9e-05 | 100.8647 | 100.7189 | 120.1253 | 21.1218 | 0.002763 | 0.00256 | 133.9632 | 0.2558 | 8.4044 | 0.002509 | 0.003999 | 0.003889 | 6.1e-05 | 0.000897 | 0.30596 | 1.3279 | 2.5473 | 100.3 | 20.96 | 32.115 | 88.1546 | 82.0465 | 83.274 | 0.007953 | 0.99052 | 0.023628 | 23.8542 | 0.01294 | 0.15218 | 0.001682 | 0.11539 | 0.11539 | 0.053676
2019-03-29 11:46 | 1553874385.8875 | 666 | 0.20916 | 0.000659 | 6.9e-05 | 100.9272 | 100.8817 | 120.0126 | 21.1212 | 0.002745 | 0.002632 | 133.9472 | 0.25572 | 8.4029 | 0.002505 | 0.004149 | 0.004001 | 6.3e-05 | 0.000894 | 0.30143 | 1.3717 | 2.5475 | 100.3 | 20.96 | 32.114 | 87.7135 | 81.9406 | 85.0079 | 0.007952 | 0.98572 | 0.02356 | 23.9009 | 0.004041 | 0.15213 | 0.0017 | 0.11631 | 0.11631 | 0.053234
2019-03-29 11:46 | 1553874386.8876 | 667 | 0.20914 | 0.000662 | 6.9e-05 | 100.9147 | 100.7926 | 120.2863 | 21.121 | 0.002757 | 0.002624 | 133.9529 | 0.25646 | 8.4035 | 0.002469 | 0.003576 | 0.003537 | 6.3e-05 | 0.000893 | 0.286 | 1.3414 | 2.5474 | 100.3 | 20.96 | 32.114 | 87.9993 | 81.0794 | 84.8096 | 0.007952 | 0.98627 | 0.023593 | 23.9215 | 0.010828 | 0.15257 | 0.001797 | 0.12339 | 0.12339 | 0.050324
2019-03-29 11:46 | 1553874387.8876 | 668 | 0.20916 | 0.000662 | 6.9e-05 | 100.9769 | 100.9144 | 120.7741 | 21.1217 | 0.002752 | 0.002604 | 133.9479 | 0.25587 | 8.4048 | 0.002483 | 0.00391 | 0.00384 | 6.3e-05 | 0.000892 | 0.2844 | 1.351 | 2.5475 | 100.3 | 20.96 | 32.116 | 87.8928 | 81.4247 | 84.3241 | 0.007953 | 0.98761 | 0.023657 | 23.9537 | 0.005539 | 0.15221 | 0.001698 | 0.11668 | 0.11668 | 0.053093
2019-03-29 11:46 | 1553874388.8876 | 669 | 0.20918 | 0.000668 | 6.9e-05 | 101.0191 | 100.8001 | 121.1478 | 21.1223 | 0.00273 | 0.002645 | 133.9315 | 0.25642 | 8.4027 | 0.002507 | 0.003608 | 0.003541 | 6.4e-05 | 0.000894 | 0.28744 | 1.3421 | 2.5473 | 100.3 | 20.96 | 32.118 | 87.3468 | 82.0032 | 85.3032 | 0.007953 | 0.98491 | 0.023661 | 24.0236 | 0.019416 | 0.15255 | 0.001583 | 0.10857 | 0.10857 | 0.057188
2019-03-29 11:46 | 1553874389.8867 | 670 | 0.20913 | 0.000664 | 6.9e-05 | 101.0257 | 100.8868 | 121.2264 | 21.1222 | 0.002697 | 0.002592 | 133.9224 | 0.25578 | 8.4048 | 0.002503 | 0.004516 | 0.004403 | 6.4e-05 | 0.000887 | 0.26288 | 1.3175 | 2.5473 | 100.3 | 20.96 | 32.12 | 86.5629 | 81.9047 | 84.0414 | 0.007954 | 0.98839 | 0.023711 | 23.989 | 0.01231 | 0.15216 | 0.001912 | 0.1322 | 0.1322 | 0.046847
2019-03-29 11:46 | 1553874390.8867 | 671 | 0.20915 | 0.000662 | 6.9e-05 | 101.1359 | 100.9834 | 121.1794 | 21.1218 | 0.002764 | 0.002574 | 133.909 | 0.25594 | 8.403 | 0.00245 | 0.004413 | 0.004333 | 6.4e-05 | 0.00089 | 0.30198 | 1.3495 | 2.5474 | 100.3 | 20.96 | 32.124 | 88.1819 | 80.6177 | 83.6072 | 0.007955 | 0.9896 | 0.02372 | 23.9698 | 0.013499 | 0.15226 | 0.001778 | 0.12269 | 0.12269 | 0.050511
2019-03-29 11:46 | 1553874391.887 | 672 | 0.20915 | 0.000659 | 6.8e-05 | 101.0179 | 101.0144 | 121.2313 | 21.1224 | 0.002751 | 0.00259 | 133.8933 | 0.25523 | 8.4016 | 0.002525 | 0.004119 | 0.003989 | 6.4e-05 | 0.00089 | 0.30215 | 1.3435 | 2.5475 | 100.3 | 20.96 | 32.127 | 87.8608 | 82.4338 | 84.0009 | 0.007956 | 0.9885 | 0.023712 | 23.9882 | 0.000311 | 0.15183 | 0.001778 | 0.12267 | 0.12267 | 0.050376
2019-03-29 11:47 | 1553874392.8867 | 673 | 0.20918 | 0.000661 | 6.8e-05 | 101.1859 | 100.9598 | 121.5696 | 21.1216 | 0.002793 | 0.002626 | 133.8869 | 0.25475 | 8.4049 | 0.002493 | 0.004016 | 0.003957 | 6.4e-05 | 0.000891 | 0.28263 | 1.3614 | 2.5474 | 100.3 | 20.96 | 32.136 | 88.8753 | 81.6578 | 84.8678 | 0.007958 | 0.98611 | 0.023717 | 24.0507 | 0.020004 | 0.15155 | 0.001605 | 0.11036 | 0.11036 | 0.055889
2019-03-29 11:47 | 1553874393.8868 | 674 | 0.20919 | 0.000661 | 6.8e-05 | 101.1006 | 101.2035 | 121.5543 | 21.1227 | 0.002793 | 0.002622 | 133.8922 | 0.25641 | 8.4047 | 0.002458 | 0.004273 | 0.004168 | 6.4e-05 | 0.000892 | 0.3183 | 1.344 | 2.5475 | 100.3 | 20.96 | 32.144 | 88.8639 | 80.8128 | 84.7573 | 0.00796 | 0.98642 | 0.023719 | 24.0455 | -0.009103 | 0.15254 | 0.001553 | 0.10673 | 0.10673 | 0.05817
2019-03-29 11:47 | 1553874394.8868 | 675 | 0.20917 | 0.00066 | 6.8e-05 | 101.2463 | 101.1257 | 121.6224 | 21.122 | 0.002669 | 0.002576 | 133.8965 | 0.25642 | 8.4003 | 0.002441 | 0.003993 | 0.00394 | 6e-05 | 0.000891 | 0.26056 | 1.338 | 2.5474 | 100.3 | 20.96 | 32.151 | 85.8874 | 80.4036 | 83.6603 | 0.007962 | 0.98945 | 0.023762 | 24.0154 | 0.010665 | 0.15255 | 0.001634 | 0.11266 | 0.11266 | 0.055108
2019-03-29 11:47 | 1553874395.8858 | 676 | 0.20919 | 0.000655 | 6.8e-05 | 101.2383 | 101.0284 | 121.7519 | 21.1221 | 0.002781 | 0.002609 | 133.8933 | 0.25564 | 8.4026 | 0.002456 | 0.003657 | 0.003492 | 6.6e-05 | 0.000887 | 0.2975 | 1.3277 | 2.5474 | 100.3 | 20.96 | 32.17 | 88.5783 | 80.7591 | 84.439 | 0.007966 | 0.98729 | 0.023749 | 24.0544 | 0.018558 | 0.15208 | 0.001541 | 0.10598 | 0.10598 | 0.058404
2019-03-29 11:47 | 1553874396.8881 | 677 | 0.20917 | 0.000654 | 6.8e-05 | 101.323 | 101.2363 | 121.4452 | 21.1225 | 0.002656 | 0.00258 | 133.8896 | 0.25596 | 8.4009 | 0.00247 | 0.003998 | 0.003918 | 6.4e-05 | 0.000883 | 0.28672 | 1.3585 | 2.5474 | 100.3 | 20.96 | 32.181 | 85.5805 | 81.1104 | 83.7514 | 0.007969 | 0.9892 | 0.023742 | 24.0009 | 0.007657 | 0.15227 | 0.001646 | 0.11342 | 0.11342 | 0.054641
2019-03-29 11:47 | 1553874397.8895 | 678 | 0.20917 | 0.000656 | 6.8e-05 | 101.2781 | 101.2823 | 121.3745 | 21.1232 | 0.002756 | 0.002547 | 133.8918 | 0.25585 | 8.4023 | 0.002466 | 0.003487 | 0.003351 | 6.1e-05 | 0.000881 | 0.28998 | 1.3597 | 2.5474 | 100.3 | 20.96 | 32.202 | 87.9802 | 81.0097 | 82.9642 | 0.007974 | 0.99138 | 0.023761 | 23.9675 | -0.000377 | 0.15221 | 0.001653 | 0.11403 | 0.11403 | 0.054327
2019-03-29 11:47 | 1553874398.8869 | 679 | 0.20923 | 0.00066 | 6.8e-05 | 101.3619 | 101.2602 | 121.459 | 21.1223 | 0.00271 | 0.002578 | 133.8981 | 0.25525 | 8.402 | 0.002462 | 0.003649 | 0.003594 | 5.8e-05 | 0.000882 | 0.29623 | 1.3736 | 2.5473 | 100.3 | 20.96 | 32.211 | 86.8674 | 80.916 | 83.7099 | 0.007976 | 0.98931 | 0.023744 | 24.0009 | 0.008981 | 0.15185 | 0.001281 | 0.087385 | 0.087385 | 0.070723
2019-03-29 11:47 | 1553874399.8874 | 680 | 0.20918 | 0.000653 | 6.8e-05 | 101.3257 | 101.4229 | 121.4355 | 21.1227 | 0.002725 | 0.002594 | 133.905 | 0.25642 | 8.4031 | 0.002427 | 0.003214 | 0.003162 | 6e-05 | 0.000884 | 0.30934 | 1.355 | 2.5473 | 100.3 | 20.96 | 32.221 | 87.2394 | 80.0704 | 84.0988 | 0.007979 | 0.98823 | 0.023729 | 24.0117 | -0.008584 | 0.15255 | 0.001593 | 0.10959 | 0.10959 | 0.056654
2019-03-29 11:47 | 1553874400.8874 | 681 | 0.20918 | 0.000652 | 6.8e-05 | 101.4265 | 101.3868 | 121.3513 | 21.1225 | 0.002715 | 0.002588 | 133.9126 | 0.2562 | 8.4035 | 0.002458 | 0.004314 | 0.004223 | 6.3e-05 | 0.000885 | 0.29015 | 1.338 | 2.5474 | 100.3 | 20.96 | 32.238 | 87.0098 | 80.8133 | 83.9339 | 0.007983 | 0.98869 | 0.023726 | 23.9978 | 0.003507 | 0.15241 | 0.001622 | 0.11166 | 0.11166 | 0.055554
2019-03-29 11:47 | 1553874401.8887 | 682 | 0.2092 | 0.000652 | 6.8e-05 | 101.358 | 101.298 | 121.378 | 21.123 | 0.002761 | 0.002552 | 133.9218 | 0.25635 | 8.4016 | 0.002472 | 0.003738 | 0.003666 | 6.1e-05 | 0.000887 | 0.30198 | 1.3475 | 2.5475 | 100.3 | 20.96 | 32.246 | 88.0941 | 81.1568 | 83.0725 | 0.007985 | 0.99108 | 0.023758 | 23.9715 | 0.005292 | 0.1525 | 0.001501 | 0.10316 | 0.10316 | 0.060166
2019-03-29 11:47 | 1553874402.8865 | 683 | 0.20919 | 0.00065 | 6.7e-05 | 101.3999 | 101.2086 | 121.2661 | 21.1222 | 0.002816 | 0.00262 | 133.9414 | 0.25614 | 8.4019 | 0.002452 | 0.004403 | 0.004287 | 6.1e-05 | 0.00089 | 0.30469 | 1.3518 | 2.5475 | 100.3 | 20.97 | 32.257 | 89.4212 | 80.6759 | 84.717 | 0.007993 | 0.98653 | 0.023692 | 24.0156 | 0.016888 | 0.15238 | 0.001533 | 0.1052 | 0.1052 | 0.058955
2019-03-29 11:47 | 1553874403.8872 | 684 | 0.20919 | 0.000655 | 6.8e-05 | 101.3747 | 101.4714 | 121.1665 | 21.1232 | 0.002723 | 0.002612 | 133.9446 | 0.25577 | 8.4022 | 0.002486 | 0.003706 | 0.003653 | 5.9e-05 | 0.000889 | 0.29087 | 1.3195 | 2.5473 | 100.3 | 20.97 | 32.262 | 87.1971 | 81.4916 | 84.5277 | 0.007994 | 0.98705 | 0.023689 | 23.9994 | -0.008525 | 0.15216 | 0.00153 | 0.10492 | 0.10492 | 0.059021
2019-03-29 11:47 | 1553874404.8872 | 685 | 0.20918 | 0.000652 | 6.8e-05 | 101.4439 | 101.4984 | 121.213 | 21.1225 | 0.002701 | 0.002598 | 133.9237 | 0.25613 | 8.4024 | 0.002447 | 0.003576 | 0.003502 | 5.9e-05 | 0.000892 | 0.31382 | 1.3397 | 2.5473 | 100.3 | 20.97 | 32.266 | 86.6645 | 80.5538 | 84.1804 | 0.007995 | 0.98801 | 0.023705 | 23.9924 | -0.004802 | 0.15237 | 0.001591 | 0.10937 | 0.10937 | 0.056705
2019-03-29 11:47 | 1553874405.8865 | 686 | 0.20919 | 0.000653 | 6.8e-05 | 101.4981 | 101.3694 | 121.467 | 21.1242 | 0.002762 | 0.002566 | 133.9357 | 0.25592 | 8.4039 | 0.002417 | 0.003597 | 0.003459 | 5.7e-05 | 0.000896 | 0.3126 | 1.3155 | 2.5473 | 100.3 | 20.97 | 32.27 | 88.142 | 79.8372 | 83.4146 | 0.007996 | 0.99013 | 0.023755 | 23.9918 | 0.011349 | 0.15225 | 0.001516 | 0.1042 | 0.1042 | 0.059466
2019-03-29 11:47 | 1553874406.8861 | 687 | 0.20919 | 0.000656 | 6.8e-05 | 101.5382 | 101.4451 | 121.7958 | 21.123 | 0.002709 | 0.002577 | 133.9427 | 0.25533 | 8.4044 | 0.002452 | 0.003652 | 0.003557 | 5.8e-05 | 0.000892 | 0.27024 | 1.3153 | 2.5474 | 100.3 | 20.97 | 32.27 | 86.8643 | 80.6829 | 83.6793 | 0.007996 | 0.9894 | 0.023778 | 24.0331 | 0.008201 | 0.1519 | 0.001551 | 0.10685 | 0.10685 | 0.057858
2019-03-29 11:47 | 1553874407.8884 | 688 | 0.20917 | 0.000654 | 6.9e-05 | 101.6067 | 101.5202 | 121.8137 | 21.1226 | 0.002678 | 0.002547 | 133.9395 | 0.25562 | 8.4031 | 0.002407 | 0.003769 | 0.003705 | 5.4e-05 | 0.000893 | 0.29358 | 1.3305 | 2.5474 | 100.3 | 20.97 | 32.271 | 86.1112 | 79.583 | 82.9667 | 0.007996 | 0.99138 | 0.023804 | 24.0109 | 0.007619 | 0.15207 | 0.001649 | 0.11393 | 0.11393 | 0.054325
2019-03-29 11:47 | 1553874408.8862 | 689 | 0.2092 | 0.000651 | 6.7e-05 | 101.6382 | 101.6359 | 121.4881 | 21.124 | 0.002712 | 0.002579 | 133.9522 | 0.25495 | 8.4034 | 0.00248 | 0.003345 | 0.003246 | 6e-05 | 0.000891 | 0.27129 | 1.317 | 2.5474 | 100.3 | 20.97 | 32.269 | 86.9244 | 81.3575 | 83.7279 | 0.007996 | 0.98926 | 0.023747 | 24.0044 | 0.000198 | 0.15167 | 0.001485 | 0.10202 | 0.10202 | 0.060505
2019-03-29 11:47 | 1553874409.8864 | 690 | 0.20918 | 0.000655 | 6.8e-05 | 101.5945 | 101.6202 | 121.6292 | 21.1232 | 0.00276 | 0.002555 | 133.9462 | 0.25614 | 8.4022 | 0.002515 | 0.004551 | 0.004428 | 6.1e-05 | 0.000888 | 0.28998 | 1.3662 | 2.5474 | 100.3 | 20.97 | 32.269 | 88.0707 | 82.1959 | 83.146 | 0.007996 | 0.99088 | 0.02378 | 23.9987 | -0.002267 | 0.15238 | 0.001569 | 0.10811 | 0.10811 | 0.057368
2019-03-29 11:47 | 1553874410.8867 | 691 | 0.20919 | 0.000654 | 6.8e-05 | 101.6121 | 101.7441 | 121.6024 | 21.1234 | 0.00273 | 0.002578 | 133.9232 | 0.25588 | 8.4026 | 0.002496 | 0.004013 | 0.003932 | 5.6e-05 | 0.000886 | 0.26864 | 1.3403 | 2.5474 | 100.3 | 20.97 | 32.27 | 87.3486 | 81.7374 | 83.6999 | 0.007996 | 0.98934 | 0.023759 | 24.0147 | -0.011611 | 0.15222 | 0.001531 | 0.10537 | 0.10537 | 0.058799
2019-03-29 11:47 | 1553874411.8867 | 692 | 0.20919 | 0.000657 | 6.8e-05 | 101.6217 | 101.5936 | 121.5001 | 21.1233 | 0.002696 | 0.002542 | 133.9307 | 0.25585 | 8.4015 | 0.002425 | 0.004243 | 0.004196 | 5.7e-05 | 0.000882 | 0.28727 | 1.3287 | 2.5475 | 100.3 | 20.97 | 32.27 | 86.5303 | 80.0265 | 82.8277 | 0.007996 | 0.99176 | 0.023778 | 23.9753 | 0.002474 | 0.15221 | 0.001517 | 0.10441 | 0.10441 | 0.059328
2019-03-29 11:47 | 1553874412.8865 | 693 | 0.20919 | 0.000651 | 6.8e-05 | 101.7031 | 101.6089 | 121.3757 | 21.1223 | 0.002684 | 0.002544 | 133.9125 | 0.25582 | 8.4038 | 0.002503 | 0.004012 | 0.00383 | 5.8e-05 | 0.00088 | 0.28711 | 1.3523 | 2.5474 | 100.3 | 20.97 | 32.27 | 86.257 | 81.8934 | 82.8824 | 0.007996 | 0.99161 | 0.023764 | 23.9648 | 0.008285 | 0.15219 | 0.001571 | 0.10823 | 0.10823 | 0.05723
2019-03-29 11:47 | 1553874413.8854 | 694 | 0.20918 | 0.000656 | 6.7e-05 | 101.6759 | 101.5349 | 121.49 | 21.1238 | 0.002749 | 0.002507 | 133.899 | 0.25665 | 8.4034 | 0.002526 | 0.003419 | 0.003334 | 5.7e-05 | 0.000876 | 0.27704 | 1.3547 | 2.5474 | 100.3 | 20.97 | 32.271 | 87.8277 | 82.4484 | 82.0078 | 0.007996 | 0.99405 | 0.023804 | 23.9467 | 0.012411 | 0.15268 | 0.001613 | 0.11145 | 0.11145 | 0.055756
2019-03-29 11:47 | 1553874414.8855 | 695 | 0.20923 | 0.00065 | 6.6e-05 | 101.6657 | 101.6929 | 121.3916 | 21.1239 | 0.002678 | 0.002601 | 133.9075 | 0.25589 | 8.4011 | 0.002557 | 0.004213 | 0.004078 | 5.8e-05 | 0.000872 | 0.30917 | 1.3479 | 2.5474 | 100.3 | 20.97 | 32.273 | 86.1017 | 83.2082 | 84.2599 | 0.007997 | 0.98779 | 0.023719 | 24.0127 | -0.002394 | 0.15223 | 0.001312 | 0.089577 | 0.089577 | 0.069166
2019-03-29 11:47 | 1553874415.8865 | 696 | 0.20919 | 0.000647 | 6.6e-05 | 101.651 | 101.4591 | 121.0151 | 21.1234 | 0.002683 | 0.002587 | 133.8963 | 0.25668 | 8.4037 | 0.002442 | 0.00377 | 0.003627 | 5.8e-05 | 0.000878 | 0.31116 | 1.3654 | 2.5475 | 100.3 | 20.97 | 32.277 | 86.23 | 80.4439 | 83.9192 | 0.007998 | 0.98873 | 0.023694 | 23.964 | 0.016902 | 0.1527 | 0.001516 | 0.10407 | 0.10407 | 0.059719
2019-03-29 11:47 | 1553874416.8859 | 697 | 0.20919 | 0.000646 | 6.6e-05 | 101.6468 | 101.7676 | 120.5567 | 21.1233 | 0.002726 | 0.00255 | 133.8808 | 0.25601 | 8.4029 | 0.002479 | 0.003852 | 0.00379 | 5.9e-05 | 0.000881 | 0.29214 | 1.3309 | 2.5473 | 100.3 | 20.97 | 32.28 | 87.2635 | 81.3241 | 83.0338 | 0.007998 | 0.99119 | 0.023678 | 23.8889 | -0.010625 | 0.1523 | 0.001528 | 0.10489 | 0.10489 | 0.059094
2019-03-29 11:47 | 1553874417.8856 | 698 | 0.2092 | 0.000649 | 6.6e-05 | 101.6537 | 101.5765 | 120.4553 | 21.1239 | 0.002713 | 0.00257 | 133.866 | 0.25536 | 8.4015 | 0.002493 | 0.003374 | 0.003256 | 5.7e-05 | 0.000877 | 0.27632 | 1.3245 | 2.5474 | 100.3 | 20.97 | 32.281 | 86.9501 | 81.6492 | 83.5078 | 0.007999 | 0.98987 | 0.023653 | 23.8948 | 0.006795 | 0.15191 | 0.001501 | 0.10278 | 0.10278 | 0.060156
2019-03-29 11:47 | 1553874418.8866 | 699 | 0.20921 | 0.000649 | 6.7e-05 | 101.6603 | 101.7342 | 120.5725 | 21.1234 | 0.002764 | 0.002549 | 133.8604 | 0.25677 | 8.4026 | 0.002468 | 0.003789 | 0.003719 | 5.8e-05 | 0.000872 | 0.28926 | 1.3603 | 2.5475 | 100.3 | 20.97 | 32.284 | 88.1703 | 81.0582 | 83.0065 | 0.007999 | 0.99126 | 0.023681 | 23.8896 | -0.006499 | 0.15275 | 0.001422 | 0.097268 | 0.097268 | 0.063916
2019-03-29 11:47 | 1553874419.8856 | 700 | 0.20918 | 0.000647 | 6.7e-05 | 101.6804 | 101.625 | 120.4032 | 21.1234 | 0.002655 | 0.002559 | 133.8526 | 0.25651 | 8.4013 | 0.002451 | 0.003569 | 0.003361 | 5.8e-05 | 0.000865 | 0.28385 | 1.3342 | 2.5474 | 100.3 | 20.97 | 32.287 | 85.5612 | 80.6482 | 83.2491 | 0.008 | 0.99059 | 0.023656 | 23.8809 | 0.004872 | 0.1526 | 0.001597 | 0.10966 | 0.10966 | 0.05664
2019-03-29 11:47 | 1553874420.887 | 701 | 0.20921 | 0.000652 | 6.7e-05 | 101.6426 | 101.6912 | 120.3337 | 21.124 | 0.002717 | 0.002565 | 133.8773 | 0.25547 | 8.4023 | 0.002479 | 0.004022 | 0.003969 | 5.9e-05 | 0.000865 | 0.29927 | 1.3461 | 2.5474 | 100.3 | 20.97 | 32.292 | 87.0423 | 81.3325 | 83.398 | 0.008001 | 0.99018 | 0.023644 | 23.879 | -0.004272 | 0.15198 | 0.001432 | 0.097826 | 0.097826 | 0.06323
2019-03-29 11:47 | 1553874421.8867 | 702 | 0.2092 | 0.000651 | 6.7e-05 | 101.7339 | 101.6946 | 120.1347 | 21.124 | 0.002732 | 0.002574 | 133.8698 | 0.25551 | 8.403 | 0.00248 | 0.004335 | 0.004201 | 6e-05 | 0.000865 | 0.29087 | 1.3344 | 2.5473 | 100.3 | 20.97 | 32.296 | 87.4132 | 81.36 | 83.6077 | 0.008002 | 0.98959 | 0.023618 | 23.8663 | 0.003459 | 0.152 | 0.001469 | 0.10037 | 0.10037 | 0.061639
2019-03-29 11:47 | 1553874422.8857 | 703 | 0.2092 | 0.000651 | 6.7e-05 | 101.6935 | 101.8359 | 120.1591 | 21.1242 | 0.002746 | 0.002524 | 133.8831 | 0.25593 | 8.4023 | 0.002489 | 0.003376 | 0.003338 | 5.7e-05 | 0.000851 | 0.27776 | 1.3701 | 2.5474 | 100.3 | 20.97 | 32.305 | 87.7338 | 81.5731 | 82.4154 | 0.008005 | 0.99291 | 0.02366 | 23.8288 | -0.012514 | 0.15225 | 0.001457 | 0.099658 | 0.099658 | 0.062179
2019-03-29 11:47 | 1553874423.8872 | 704 | 0.20924 | 0.00065 | 6.7e-05 | 101.6869 | 101.7438 | 120.1485 | 21.125 | 0.002685 | 0.002544 | 133.8767 | 0.25577 | 8.4044 | 0.002551 | 0.003789 | 0.003744 | 6e-05 | 0.000846 | 0.29822 | 1.3454 | 2.5474 | 100.3 | 20.97 | 32.311 | 86.2878 | 83.0619 | 82.8941 | 0.008006 | 0.99158 | 0.023643 | 23.8438 | -0.004996 | 0.15216 | 0.001272 | 0.086446 | 0.086446 | 0.071639
2019-03-29 11:47 | 1553874424.8862 | 705 | 0.20923 | 0.000653 | 6.6e-05 | 101.7046 | 101.5687 | 119.9373 | 21.1242 | 0.002609 | 0.002474 | 133.8739 | 0.2567 | 8.4027 | 0.002517 | 0.004901 | 0.004824 | 6.2e-05 | 0.000845 | 0.28423 | 1.3565 | 2.5474 | 100.3 | 20.97 | 32.318 | 84.4452 | 82.2389 | 81.1948 | 0.008008 | 0.99633 | 0.023679 | 23.7659 | 0.011963 | 0.15271 | 0.001323 | 0.090249 | 0.090249 | 0.068868
2019-03-29 11:47 | 1553874425.8858 | 706 | 0.2092 | 0.000653 | 6.7e-05 | 101.6986 | 101.5625 | 119.9988 | 21.1242 | 0.002672 | 0.002528 | 133.8679 | 0.25685 | 8.3999 | 0.002476 | 0.004606 | 0.004476 | 6.2e-05 | 0.000845 | 0.29336 | 1.371 | 2.5476 | 100.3 | 20.97 | 32.334 | 85.9646 | 81.2611 | 82.4978 | 0.008012 | 0.99268 | 0.023641 | 23.8156 | 0.011977 | 0.1528 | 0.001468 | 0.1004 | 0.1004 | 0.061943
2019-03-29 11:47 | 1553874426.8848 | 707 | 0.20923 | 0.000651 | 6.6e-05 | 101.6846 | 101.5612 | 120.3603 | 21.1249 | 0.002709 | 0.002533 | 133.8738 | 0.25654 | 8.4023 | 0.002546 | 0.004213 | 0.004097 | 5.8e-05 | 0.000845 | 0.27295 | 1.3588 | 2.5476 | 100.3 | 20.97 | 32.342 | 86.8644 | 82.9456 | 82.6194 | 0.008014 | 0.99234 | 0.023673 | 23.8556 | 0.01086 | 0.15261 | 0.001319 | 0.089957 | 0.089957 | 0.069048
2019-03-29 11:47 | 1553874427.8855 | 708 | 0.20924 | 0.000656 | 6.6e-05 | 101.6789 | 101.5086 | 120.1361 | 21.1243 | 0.002668 | 0.00255 | 133.8912 | 0.25591 | 8.4025 | 0.002584 | 0.003977 | 0.00395 | 5.9e-05 | 0.000851 | 0.30989 | 1.3308 | 2.5473 | 100.3 | 20.97 | 32.357 | 85.8764 | 83.8561 | 83.0417 | 0.008017 | 0.99117 | 0.023637 | 23.8475 | 0.01499 | 0.15224 | 0.00124 | 0.084181 | 0.084181 | 0.073606
2019-03-29 11:47 | 1553874428.8859 | 709 | 0.20923 | 0.000655 | 6.6e-05 | 101.6947 | 101.7472 | 119.827 | 21.1244 | 0.002737 | 0.002529 | 133.8717 | 0.25602 | 8.401 | 0.0025 | 0.004249 | 0.00414 | 6e-05 | 0.000843 | 0.27096 | 1.3185 | 2.5474 | 100.3 | 20.97 | 32.365 | 87.5356 | 81.8391 | 82.5317 | 0.008019 | 0.99259 | 0.023623 | 23.7997 | -0.004611 | 0.15231 | 0.001281 | 0.087071 | 0.087071 | 0.071193
2019-03-29 11:47 | 1553874429.8849 | 710 | 0.20923 | 0.000653 | 6.6e-05 | 101.6582 | 101.7362 | 119.6683 | 21.1247 | 0.002689 | 0.002553 | 133.8939 | 0.2566 | 8.4024 | 0.002456 | 0.004322 | 0.004133 | 6.1e-05 | 0.000846 | 0.28351 | 1.3202 | 2.5474 | 100.3 | 20.97 | 32.374 | 86.3748 | 80.7756 | 83.1099 | 0.008022 | 0.99098 | 0.023589 | 23.8033 | -0.006864 | 0.15265 | 0.001278 | 0.086735 | 0.086735 | 0.071631
2019-03-29 11:47 | 1553874430.8855 | 711 | 0.20922 | 0.000654 | 6.6e-05 | 101.6319 | 101.6195 | 119.4256 | 21.1248 | 0.002684 | 0.002505 | 133.8783 | 0.25571 | 8.401 | 0.002579 | 0.003625 | 0.003531 | 5.8e-05 | 0.00084 | 0.28744 | 1.3293 | 2.5474 | 100.3 | 20.98 | 32.392 | 86.2588 | 83.7374 | 81.9535 | 0.008031 | 0.9942 | 0.023603 | 23.7405 | 0.001086 | 0.15212 | 0.001383 | 0.094275 | 0.094275 | 0.065673
2019-03-29 11:47 | 1553874431.8864 | 712 | 0.20924 | 0.000652 | 6.5e-05 | 101.6585 | 101.6359 | 119.5953 | 21.1258 | 0.002678 | 0.002547 | 133.8609 | 0.25601 | 8.402 | 0.002591 | 0.003819 | 0.003641 | 6.1e-05 | 0.000841 | 0.29662 | 1.3247 | 2.5474 | 100.3 | 20.98 | 32.401 | 86.1182 | 84.0144 | 82.9519 | 0.008033 | 0.99142 | 0.023587 | 23.7907 | 0.001989 | 0.1523 | 0.001265 | 0.085834 | 0.085834 | 0.072216
2019-03-29 11:47 | 1553874432.8851 | 713 | 0.20921 | 0.000653 | 6.4e-05 | 101.6067 | 101.6799 | 119.5291 | 21.1251 | 0.002599 | 0.0025 | 133.8526 | 0.25635 | 8.4022 | 0.002564 | 0.003995 | 0.003918 | 6.2e-05 | 0.000841 | 0.31636 | 1.3377 | 2.5474 | 100.3 | 20.98 | 32.417 | 84.2135 | 83.3755 | 81.8196 | 0.008037 | 0.99458 | 0.023618 | 23.7463 | -0.006442 | 0.1525 | 0.001427 | 0.097513 | 0.097513 | 0.063652
2019-03-29 11:47 | 1553874433.886 | 714 | 0.20921 | 0.000652 | 6.4e-05 | 101.5966 | 101.6079 | 119.3413 | 21.1258 | 0.00267 | 0.002501 | 133.8616 | 0.25612 | 8.4021 | 0.002548 | 0.004512 | 0.004404 | 6.1e-05 | 0.00084 | 0.31797 | 1.3376 | 2.5473 | 100.3 | 20.98 | 32.426 | 85.9123 | 82.9799 | 81.8585 | 0.00804 | 0.99447 | 0.023598 | 23.729 | -0.001001 | 0.15237 | 0.001435 | 0.098 | 0.098 | 0.06328
2019-03-29 11:47 | 1553874434.8865 | 715 | 0.20922 | 0.000648 | 6.5e-05 | 101.6744 | 101.5967 | 119.4117 | 21.1253 | 0.00265 | 0.002538 | 133.832 | 0.25623 | 8.4022 | 0.002545 | 0.004173 | 0.004108 | 6.4e-05 | 0.000845 | 0.29717 | 1.3639 | 2.5475 | 100.3 | 20.98 | 32.434 | 85.4333 | 82.9023 | 82.7486 | 0.008042 | 0.99198 | 0.023575 | 23.7657 | 0.006839 | 0.15243 | 0.001367 | 0.093045 | 0.093045 | 0.066678
2019-03-29 11:47 | 1553874435.8855 | 716 | 0.20925 | 0.000649 | 6.5e-05 | 101.5789 | 101.6168 | 119.3786 | 21.1261 | 0.002694 | 0.002536 | 133.8301 | 0.25623 | 8.4026 | 0.002585 | 0.003536 | 0.003504 | 6.3e-05 | 0.000848 | 0.27936 | 1.3885 | 2.5475 | 100.3 | 20.98 | 32.449 | 86.5011 | 83.8743 | 82.7065 | 0.008045 | 0.9921 | 0.023573 | 23.761 | -0.003337 | 0.15243 | 0.001181 | 0.079856 | 0.079856 | 0.077689
2019-03-29 11:47 | 1553874436.8852 | 717 | 0.20923 | 0.000649 | 6.5e-05 | 101.5912 | 101.5472 | 119.3026 | 21.1264 | 0.002703 | 0.002546 | 133.8335 | 0.25525 | 8.4038 | 0.002532 | 0.003888 | 0.003789 | 6.2e-05 | 0.000848 | 0.29966 | 1.3669 | 2.5474 | 100.3 | 20.98 | 32.455 | 86.7215 | 82.592 | 82.9298 | 0.008047 | 0.99148 | 0.023558 | 23.7609 | 0.003874 | 0.15185 | 0.001284 | 0.08713 | 0.08713 | 0.07093
2019-03-29 11:47 | 1553874437.8852 | 718 | 0.20925 | 0.00065 | 6.5e-05 | 101.5977 | 101.5837 | 119.4918 | 21.1253 | 0.002692 | 0.002546 | 133.821 | 0.25553 | 8.4018 | 0.002522 | 0.00436 | 0.004229 | 6.2e-05 | 0.000842 | 0.29231 | 1.3579 | 2.5475 | 100.3 | 20.97 | 32.468 | 86.4362 | 82.3469 | 82.9405 | 0.008045 | 0.99145 | 0.023577 | 23.7801 | 0.001238 | 0.15202 | 0.001198 | 0.081082 | 0.081082 | 0.076306
2019-03-29 11:47 | 1553874438.8862 | 719 | 0.20926 | 0.000646 | 6.5e-05 | 101.5385 | 101.466 | 119.6167 | 21.1259 | 0.002676 | 0.002518 | 133.8476 | 0.25653 | 8.3995 | 0.002556 | 0.003902 | 0.003789 | 6.2e-05 | 0.000838 | 0.28622 | 1.3407 | 2.5476 | 100.3 | 20.97 | 32.475 | 86.0654 | 83.1789 | 82.268 | 0.008047 | 0.99332 | 0.023611 | 23.77 | 0.006391 | 0.15261 | 0.001125 | 0.076 | 0.076 | 0.081726
2019-03-29 11:47 | 1553874439.8856 | 720 | 0.20918 | 0.000653 | 6.5e-05 | 101.5912 | 101.5847 | 119.8535 | 21.1252 | 0.002661 | 0.002534 | 133.8285 | 0.25582 | 8.3998 | 0.00258 | 0.00366 | 0.003554 | 6.2e-05 | 0.000837 | 0.28943 | 1.3743 | 2.5474 | 100.3 | 20.98 | 32.48 | 85.7065 | 83.7423 | 82.6433 | 0.008053 | 0.99228 | 0.023622 | 23.8061 | 0.000567 | 0.15219 | 0.001617 | 0.11098 | 0.11098 | 0.055812
2019-03-29 11:47 | 1553874440.8863 | 721 | 0.20924 | 0.000648 | 6.5e-05 | 101.6537 | 101.5107 | 120.061 | 21.126 | 0.002656 | 0.00252 | 133.8197 | 0.25565 | 8.3996 | 0.002557 | 0.004235 | 0.004117 | 6.4e-05 | 0.000841 | 0.30397 | 1.3515 | 2.5475 | 100.3 | 20.98 | 32.494 | 85.5901 | 83.2088 | 82.3007 | 0.008056 | 0.99323 | 0.023654 | 23.8152 | 0.012597 | 0.15209 | 0.001267 | 0.08624 | 0.08624 | 0.071776
2019-03-29 11:47 | 1553874441.8849 | 722 | 0.20922 | 0.000642 | 6.5e-05 | 101.5166 | 101.5233 | 120.041 | 21.1261 | 0.002667 | 0.002528 | 133.8204 | 0.25584 | 8.4054 | 0.002627 | 0.004181 | 0.004079 | 6.3e-05 | 0.000837 | 0.28512 | 1.3701 | 2.5475 | 100.3 | 20.98 | 32.501 | 85.84 | 84.8754 | 82.5146 | 0.008058 | 0.99264 | 0.023645 | 23.8204 | -0.000586 | 0.1522 | 0.00139 | 0.094984 | 0.094984 | 0.065218
2019-03-29 11:47 | 1553874442.8847 | 723 | 0.20918 | 0.000641 | 6.5e-05 | 101.5226 | 101.4878 | 120.2387 | 21.1268 | 0.002674 | 0.002519 | 133.8283 | 0.25532 | 8.404 | 0.002572 | 0.003802 | 0.003622 | 6.3e-05 | 0.000838 | 0.3011 | 1.3967 | 2.5475 | 100.3 | 20.97 | 32.507 | 86.0124 | 83.5551 | 82.296 | 0.008055 | 0.99325 | 0.023672 | 23.8327 | 0.003069 | 0.15189 | 0.001615 | 0.11108 | 0.11108 | 0.055652
2019-03-29 11:47 | 1553874443.8845 | 724 | 0.20926 | 0.000641 | 6.5e-05 | 101.5421 | 101.581 | 120.6891 | 21.1258 | 0.002635 | 0.002517 | 133.8239 | 0.25518 | 8.4022 | 0.002504 | 0.00361 | 0.003544 | 6e-05 | 0.000831 | 0.28584 | 1.3867 | 2.5474 | 100.3 | 20.97 | 32.519 | 85.0694 | 81.9255 | 82.2508 | 0.008058 | 0.99337 | 0.023718 | 23.8758 | -0.003425 | 0.15181 | 0.001111 | 0.075397 | 0.075397 | 0.081947
2019-03-29 11:47 | 1553874444.8858 | 725 | 0.20922 | 0.000643 | 6.5e-05 | 101.502 | 101.4038 | 121.1226 | 21.1263 | 0.002677 | 0.002568 | 133.8506 | 0.25587 | 8.4009 | 0.002518 | 0.003939 | 0.003878 | 6e-05 | 0.000825 | 0.28888 | 1.3979 | 2.5475 | 100.3 | 20.97 | 32.523 | 86.079 | 82.2614 | 83.4579 | 0.008059 | 0.99001 | 0.02372 | 23.9592 | 0.008651 | 0.15222 | 0.001353 | 0.092641 | 0.092641 | 0.066874
2019-03-29 11:47 | 1553874445.8855 | 726 | 0.20934 | 0.000643 | 6.5e-05 | 101.4663 | 101.3311 | 121.3914 | 21.1263 | 0.002651 | 0.002513 | 133.8757 | 0.25547 | 8.3992 | 0.002525 | 0.004048 | 0.00393 | 6.2e-05 | 0.000822 | 0.31526 | 1.3402 | 2.5474 | 100.3 | 20.98 | 32.529 | 85.4546 | 82.4345 | 82.1326 | 0.008065 | 0.9937 | 0.02379 | 23.9411 | 0.011926 | 0.15198 | 0.000676 | 0.044463 | 0.044463 | 0.13912
2019-03-29 11:47 | 1553874446.8844 | 727 | 0.2092 | 0.000642 | 6.4e-05 | 101.434 | 101.2165 | 121.5573 | 21.1268 | 0.002615 | 0.002492 | 126.527 | 0.25512 | 8.4048 | 0.002475 | 0.004134 | 0.004047 | 6.4e-05 | 0.000822 | 0.30901 | 1.3493 | 2.5473 | 100.3 | 20.98 | 32.531 | 84.5858 | 81.2418 | 81.6472 | 0.008066 | 0.99506 | 0.023823 | 23.9411 | 0.019204 | 0.15177 | 0.001489 | 0.1028 | 0.1028 | 0.060086
2019-03-29 11:47 | 1553874447.8859 | 728 | 0.20921 | 0.000644 | 6.4e-05 | 101.3885 | 101.4885 | 121.9331 | 21.1264 | 0.002671 | 0.002465 | 97.9952 | 0.25618 | 8.4001 | 0.002446 | 0.004214 | 0.004142 | 6.1e-05 | 0.000816 | 0.25194 | 1.3863 | 2.5475 | 100.3 | 20.98 | 32.533 | 85.9386 | 80.5271 | 80.9972 | 0.008066 | 0.99689 | 0.023882 | 23.9561 | -0.008814 | 0.1524 | 0.001418 | 0.097951 | 0.097951 | 0.063324
2019-03-29 11:47 | 1553874448.8849 | 729 | 0.20922 | 0.000643 | 6.5e-05 | 101.2775 | 101.2093 | 122.6385 | 21.1266 | 0.002628 | 0.002574 | 69.4267 | 0.25603 | 8.4018 | 0.002446 | 0.003805 | 0.003679 | 6.2e-05 | 0.00082 | 0.28152 | 1.3596 | 2.5474 | 100.3 | 20.98 | 32.537 | 84.9071 | 80.5428 | 83.6039 | 0.008067 | 0.9896 | 0.023863 | 24.1136 | 0.006022 | 0.15232 | 0.001376 | 0.094812 | 0.094812 | 0.065384
2019-03-29 11:47 | 1553874449.8846 | 730 | 0.2092 | 0.000645 | 6.3e-05 | 101.2832 | 101.1653 | 122.7796 | 21.127 | 0.002694 | 0.002458 | 40.8678 | 0.25609 | 8.4011 | 0.002523 | 0.003571 | 0.003501 | 6.1e-05 | 0.000822 | 0.28047 | 1.3612 | 2.5475 | 100.3 | 20.98 | 32.538 | 86.4988 | 82.3868 | 80.8107 | 0.008067 | 0.99741 | 0.023971 | 24.0328 | 0.010416 | 0.15235 | 0.001484 | 0.10314 | 0.10314 | 0.060118
2019-03-29 11:47 | 1553874450.8859 | 731 | 0.20922 | 0.000646 | 6.1e-05 | 101.2098 | 100.9782 | 123.027 | 21.127 | 0.002651 | 0.002451 | 12.3255 | 0.25585 | 8.4 | 0.002418 | 0.004198 | 0.004025 | 5.8e-05 | 0.000828 | 0.30989 | 1.3043 | 2.5474 | 100.3 | 20.98 | 32.537 | 85.4615 | 79.8586 | 80.6586 | 0.008067 | 0.99784 | 0.024 | 24.0518 | 0.020488 | 0.15221 | 0.001383 | 0.096098 | 0.096098 | 0.064463
2019-03-29 11:48 | 1553874451.8847 | 732 | 0.20923 | 0.000642 | 5.7e-05 | 101.144 | 101.0813 | 123.341 | 21.1269 | 0.002638 | 0.002415 | -16.2213 | 0.25074 | 8.4021 | 0.002457 | 0.004175 | 0.004131 | 5.7e-05 | 0.000832 | 0.30646 | 1.333 | 2.5474 | 100.3 | 20.98 | 32.536 | 85.1578 | 80.7926 | 79.7963 | 0.008067 | 1.0003 | 0.02406 | 24.0531 | 0.005541 | 0.14916 | 0.001337 | 0.093282 | 0.093282 | 0.065082
2019-03-29 11:48 | 1553874452.8847 | 733 | 0.20928 | 0.000642 | 5.2e-05 | 101.1455 | 100.981 | 123.3413 | 21.1272 | 0.002613 | 0.002442 | -44.7896 | 0.23355 | 8.4006 | 0.002458 | 0.003872 | 0.00371 | 6.2e-05 | 0.000837 | 0.28528 | 1.3499 | 2.5474 | 100.3 | 20.98 | 32.533 | 84.5402 | 80.8266 | 80.4469 | 0.008066 | 0.99844 | 0.024038 | 24.0753 | 0.014557 | 0.13894 | 0.001054 | 0.073019 | 0.073019 | 0.077442
2019-03-29 11:48 | 1553874453.8857 | 734 | 0.20932 | 0.000637 | 4.7e-05 | 101.0835 | 100.9496 | 123.4701 | 21.1276 | 0.002604 | 0.002463 | -73.3452 | 0.23383 | 8.4041 | 0.002393 | 0.003906 | 0.003679 | 5.6e-05 | 0.000839 | 0.29822 | 1.347 | 2.5474 | 100.3 | 20.98 | 32.526 | 84.3299 | 79.2512 | 80.9462 | 0.008064 | 0.99703 | 0.024033 | 24.1049 | 0.011858 | 0.13911 | 0.000807 | 0.055454 | 0.055454 | 0.1021
2019-03-29 11:48 | 1553874454.8849 | 735 | 0.20936 | 0.000639 | 4.2e-05 | 101.012 | 100.8796 | 123.5955 | 21.1282 | 0.002538 | 0.002422 | -101.9032 | 0.23341 | 8.4027 | 0.002525 | 0.003788 | 0.003561 | 6.2e-05 | 0.000844 | 0.32173 | 1.3875 | 2.5475 | 100.3 | 20.98 | 32.522 | 82.7505 | 82.4339 | 79.9688 | 0.008063 | 0.99979 | 0.024079 | 24.0838 | 0.011726 | 0.13886 | 0.000597 | 0.040701 | 0.040701 | 0.13885
2019-03-29 11:48 | 1553874455.8848 | 736 | 0.20939 | 0.000628 | 3.8e-05 | 100.9189 | 100.6056 | 124.0053 | 21.1281 | 0.002623 | 0.002402 | -130.4967 | 0.23405 | 8.4017 | 0.002479 | 0.003759 | 0.003554 | 6.3e-05 | 0.000849 | 0.30431 | 1.3383 | 2.5472 | 100.3 | 20.98 | 32.516 | 84.7953 | 81.3163 | 79.4824 | 0.008062 | 1.0012 | 0.024135 | 24.1071 | 0.027812 | 0.13923 | 0.000411 | 0.027551 | 0.027551 | 0.20569
2019-03-29 11:48 | 1553874456.8848 | 737 | 0.20944 | 0.000618 | 3.4e-05 | 100.8165 | 100.6032 | 124.1687 | 21.1287 | 0.00257 | 0.002397 | -151.7253 | 0.23401 | 8.4017 | 0.002528 | 0.003697 | 0.003575 | 6.4e-05 | 0.000852 | 0.28351 | 1.3597 | 2.5475 | 100.3 | 20.98 | 32.509 | 83.5173 | 82.5063 | 79.3559 | 0.00806 | 1.0015 | 0.024156 | 24.1187 | 0.018946 | 0.13921 | 0.000147 | 0.008577 | 0.008577 | 0.66062
2019-03-29 11:48 | 1553874457.8849 | 738 | 0.20941 | 0.000604 | 3e-05 | 100.862 | 100.6291 | 124.3557 | 21.1277 | 0.002588 | 0.002379 | -151.7565 | 0.23395 | 8.4002 | 0.002492 | 0.003427 | 0.003264 | 6.2e-05 | 0.000859 | 0.28694 | 1.3709 | 2.5475 | 100.3 | 20.98 | 32.507 | 83.9578 | 81.6356 | 78.9207 | 0.00806 | 1.0028 | 0.024189 | 24.1219 | 0.020676 | 0.13918 | 0.000331 | 0.022287 | 0.022287 | 0.25416
2019-03-29 11:48 | 1553874458.8842 | 739 | 0.20943 | 0.000594 | 2.6e-05 | 100.6722 | 100.6124 | 124.3872 | 21.1292 | 0.002532 | 0.002378 | -151.7757 | 0.23369 | 8.4046 | 0.002524 | 0.004041 | 0.003867 | 6.5e-05 | 0.00086 | 0.30685 | 1.4058 | 2.5474 | 100.3 | 20.98 | 32.5 | 82.6051 | 82.4037 | 78.8948 | 0.008058 | 1.0028 | 0.024193 | 24.1241 | 0.005317 | 0.13902 | 0.000203 | 0.013167 | 0.013167 | 0.42973
2019-03-29 11:48 | 1553874459.8838 | 740 | 0.20942 | 0.000581 | 2.2e-05 | 100.7156 | 100.5967 | 124.4747 | 21.1286 | 0.002535 | 0.002394 | -151.778 | 0.23409 | 8.4013 | 0.002537 | 0.003827 | 0.003647 | 6.3e-05 | 0.000864 | 0.28998 | 1.3709 | 2.5474 | 100.3 | 20.98 | 32.497 | 82.6674 | 82.7261 | 79.2762 | 0.008057 | 1.0018 | 0.024188 | 24.1456 | 0.010568 | 0.13926 | 0.000318 | 0.021798 | 0.021798 | 0.26002
2019-03-29 11:48 | 1553874460.8843 | 741 | 0.20941 | 0.000562 | 1.9e-05 | 100.6671 | 100.667 | 124.6904 | 21.1283 | 0.002523 | 0.002294 | -151.7954 | 0.23423 | 8.4004 | 0.002516 | 0.003414 | 0.003203 | 6.4e-05 | 0.000873 | 0.29479 | 1.3667 | 2.5474 | 100.3 | 20.98 | 32.495 | 82.395 | 82.2272 | 76.8866 | 0.008057 | 1.0086 | 0.024291 | 24.0845 | 1.2e-05 | 0.13934 | 0.000389 | 0.027277 | 0.027277 | 0.20792
2019-03-29 11:48 | 1553874461.885 | 742 | 0.2094 | 0.000557 | 1.6e-05 | 100.612 | 100.4172 | 124.7731 | 21.1282 | 0.002513 | 0.00234 | -151.8016 | 0.23457 | 8.4018 | 0.002459 | 0.003938 | 0.003728 | 6.5e-05 | 0.000872 | 0.30525 | 1.3574 | 2.5475 | 100.3 | 20.98 | 32.491 | 82.1549 | 80.8502 | 77.9771 | 0.008056 | 1.0055 | 0.024262 | 24.13 | 0.017342 | 0.13954 | 0.000434 | 0.030653 | 0.030653 | 0.18528
2019-03-29 11:48 | 1553874462.884 | 743 | 0.20945 | 0.000547 | 1.3e-05 | 100.5791 | 100.3025 | 125.0897 | 21.1283 | 0.002473 | 0.002343 | -151.8022 | 0.23423 | 8.4035 | 0.002362 | 0.004045 | 0.003821 | 6.7e-05 | 0.000871 | 0.30381 | 1.38 | 2.5475 | 100.3 | 20.98 | 32.49 | 81.1792 | 78.5204 | 78.0644 | 0.008055 | 1.0052 | 0.02429 | 24.1636 | 0.024634 | 0.13934 | 0.00018 | 0.012344 | 0.012344 | 0.45942
2019-03-29 11:48 | 1553874463.8855 | 744 | 0.20945 | 0.00054 | 1.1e-05 | 100.4262 | 100.2557 | 125.2366 | 21.1285 | 0.002473 | 0.002244 | -151.8126 | 0.23364 | 8.4011 | 0.002443 | 0.003896 | 0.00372 | 6.7e-05 | 0.000866 | 0.30901 | 1.3511 | 2.5474 | 100.3 | 20.99 | 32.487 | 81.1864 | 80.4543 | 75.688 | 0.00806 | 1.0121 | 0.024386 | 24.0958 | 0.015192 | 0.13899 | 0.000168 | 0.011671 | 0.011671 | 0.48472
2019-03-29 11:48 | 1553874464.8851 | 745 | 0.20948 | 0.000528 | 9e-06 | 100.4411 | 100.3271 | 125.3218 | 21.1283 | 0.002451 | 0.002289 | -151.8157 | 0.23433 | 8.4019 | 0.002425 | 0.004138 | 0.004026 | 6.6e-05 | 0.000872 | 0.30973 | 1.367 | 2.5474 | 100.3 | 20.99 | 32.487 | 80.6647 | 80.0307 | 76.7689 | 0.00806 | 1.0089 | 0.024357 | 24.1413 | 0.010163 | 0.13941 | 4.7e-05 | 0.002892 | 0.002892 | 1.962
2019-03-29 11:48 | 1553874465.8845 | 746 | 0.20944 | 0.000524 | 8e-06 | 100.3908 | 100.2172 | 125.1531 | 21.1293 | 0.002423 | 0.00229 | -151.8112 | 0.23373 | 8.4021 | 0.002498 | 0.003792 | 0.003522 | 6.6e-05 | 0.000871 | 0.30502 | 1.3588 | 2.5473 | 100.3 | 20.99 | 32.487 | 79.9803 | 81.7744 | 76.7932 | 0.00806 | 1.0089 | 0.02434 | 24.1259 | 0.015487 | 0.13905 | 0.000244 | 0.017391 | 0.017391 | 0.32541
2019-03-29 11:48 | 1553874466.8855 | 747 | 0.20949 | 0.000516 | 6e-06 | 100.3142 | 100.1759 | 125.4258 | 21.1291 | 0.002493 | 0.002324 | -151.8078 | 0.23396 | 8.4045 | 0.002454 | 0.003844 | 0.003676 | 6.7e-05 | 0.000871 | 0.29358 | 1.2918 | 2.5473 | 100.3 | 20.99 | 32.49 | 81.6696 | 80.7347 | 77.6014 | 0.00806 | 1.0065 | 0.024338 | 24.18 | 0.012343 | 0.13918 | 6e-06 | 9.2e-05 | 9.2e-05 | 61.6562
2019-03-29 11:48 | 1553874467.8852 | 748 | 0.20947 | 0.000508 | 5e-06 | 100.3801 | 100.0015 | 125.2195 | 21.1289 | 0.002481 | 0.002286 | -151.8013 | 0.23426 | 8.4043 | 0.002569 | 0.003848 | 0.003683 | 6.7e-05 | 0.00087 | 0.29302 | 1.3795 | 2.5474 | 100.3 | 20.99 | 32.49 | 81.3887 | 83.4838 | 76.6806 | 0.00806 | 1.0092 | 0.02435 | 24.1284 | 0.033798 | 0.13936 | 0.000135 | 0.009593 | 0.009593 | 0.59128
2019-03-29 11:48 | 1553874468.8852 | 749 | 0.20949 | 0.000504 | 4e-06 | 100.2241 | 100.2455 | 125.0111 | 21.1286 | 0.002408 | 0.002296 | -151.7835 | 0.23368 | 8.4024 | 0.002447 | 0.003846 | 0.003647 | 6.7e-05 | 0.000865 | 0.25984 | 1.321 | 2.5473 | 100.3 | 20.99 | 32.492 | 79.6256 | 80.5681 | 76.9321 | 0.008061 | 1.0085 | 0.024321 | 24.117 | -0.001909 | 0.13901 | -2.3e-05 | -0.001917 | -0.001917 | -2.9511
2019-03-29 11:48 | 1553874469.8852 | 750 | 0.20946 | 0.000496 | 3e-06 | 100.1948 | 99.9216 | 125.2385 | 21.1281 | 0.00244 | 0.002301 | -151.7902 | 0.23444 | 8.402 | 0.002557 | 0.003816 | 0.003691 | 6.2e-05 | 0.000864 | 0.30773 | 1.3218 | 2.5474 | 100.3 | 20.99 | 32.493 | 80.393 | 83.2061 | 77.0501 | 0.008061 | 1.0081 | 0.024339 | 24.143 | 0.024419 | 0.13947 | 0.00022 | 0.015939 | 0.015939 | 0.35612
2019-03-29 11:48 | 1553874470.8853 | 751 | 0.20946 | 0.000494 | 2e-06 | 100.1577 | 99.8711 | 125.646 | 21.1292 | 0.002438 | 0.002251 | -151.7929 | 0.23301 | 8.3996 | 0.002499 | 0.00369 | 0.003484 | 6.4e-05 | 0.000862 | 0.30613 | 1.3232 | 2.5475 | 100.3 | 20.99 | 32.495 | 80.338 | 81.8205 | 75.8398 | 0.008062 | 1.0116 | 0.024421 | 24.1404 | 0.025627 | 0.13862 | 0.000191 | 0.013902 | 0.013902 | 0.40581
2019-03-29 11:48 | 1553874471.8833 | 752 | 0.20949 | 0.00049 | 2e-06 | 100.0319 | 99.8134 | 126.0389 | 21.1289 | 0.002378 | 0.002302 | -151.7834 | 0.23413 | 8.404 | 0.002516 | 0.003398 | 0.003251 | 6.5e-05 | 0.000859 | 0.28982 | 1.3296 | 2.5474 | 100.3 | 20.99 | 32.497 | 78.9038 | 82.2069 | 77.0859 | 0.008062 | 1.008 | 0.024416 | 24.2213 | 0.01956 | 0.13928 | 4e-05 | 0.002787 | 0.002787 | 2.0339
2019-03-29 11:48 | 1553874472.8843 | 753 | 0.20951 | 0.000487 | 1e-06 | 99.999 | 99.7667 | 125.9974 | 21.1301 | 0.002323 | 0.002241 | -151.7734 | 0.23427 | 8.4033 | 0.002511 | 0.003888 | 0.003645 | 6.4e-05 | 0.000857 | 0.28368 | 1.3326 | 2.5473 | 100.3 | 20.99 | 32.499 | 77.5861 | 82.0896 | 75.6138 | 0.008063 | 1.0123 | 0.024463 | 24.1663 | 0.020806 | 0.13937 | -0.000103 | -0.007637 | -0.007637 | -0.74277
2019-03-29 11:48 | 1553874473.8834 | 754 | 0.20949 | 0.000489 | 1e-06 | 100.0197 | 99.8165 | 125.8624 | 21.1283 | 0.002443 | 0.002212 | -151.7619 | 0.2345 | 8.4004 | 0.002523 | 0.004398 | 0.004276 | 6e-05 | 0.000851 | 0.27417 | 1.3415 | 2.5474 | 100.3 | 20.99 | 32.501 | 80.4618 | 82.377 | 74.9113 | 0.008063 | 1.0143 | 0.024475 | 24.129 | 0.018187 | 0.1395 | 7e-06 | 0.000441 | 0.000441 | 12.8719
2019-03-29 11:48 | 1553874474.8844 | 755 | 0.20946 | 0.000488 | 1e-06 | 99.967 | 99.5192 | 125.8645 | 21.1291 | 0.002298 | 0.002231 | -151.7635 | 0.23384 | 8.4043 | 0.002586 | 0.004174 | 0.004033 | 6.5e-05 | 0.000858 | 0.30309 | 1.3712 | 2.5475 | 100.3 | 21 | 32.505 | 76.9746 | 83.9021 | 75.369 | 0.008069 | 1.013 | 0.024459 | 24.1451 | 0.040152 | 0.13911 | 0.000232 | 0.017039 | 0.017039 | 0.33229
2019-03-29 11:48 | 1553874475.8838 | 756 | 0.20948 | 0.000485 | 0 | 99.8341 | 99.6612 | 125.7568 | 21.1283 | 0.00238 | 0.002217 | -151.7512 | 0.23448 | 8.4011 | 0.002549 | 0.003878 | 0.003732 | 6e-05 | 0.000861 | 0.30397 | 1.3447 | 2.5474 | 100.3 | 21 | 32.508 | 78.9477 | 83.0223 | 75.0377 | 0.00807 | 1.014 | 0.02446 | 24.1233 | 0.015502 | 0.13949 | 9.6e-05 | 0.007051 | 0.007051 | 0.80514
2019-03-29 11:48 | 1553874476.8845 | 757 | 0.20948 | 0.000484 | 0 | 99.8284 | 99.5807 | 125.6514 | 21.1291 | 0.002405 | 0.002143 | -151.7466 | 0.23292 | 8.4014 | 0.002505 | 0.00355 | 0.003323 | 6.1e-05 | 0.000865 | 0.26416 | 1.4035 | 2.5474 | 100.3 | 21 | 32.51 | 79.5504 | 81.9622 | 73.2604 | 0.00807 | 1.0192 | 0.024512 | 24.0515 | 0.022224 | 0.13856 | 7.4e-05 | 0.005486 | 0.005486 | 1.0279
2019-03-29 11:48 | 1553874477.8844 | 758 | 0.20947 | 0.00048 | 0 | 99.7796 | 99.2684 | 125.7005 | 21.1295 | 0.002383 | 0.002225 | -151.7538 | 0.23381 | 8.4055 | 0.002562 | 0.004064 | 0.003915 | 5.7e-05 | 0.000869 | 0.30021 | 1.3894 | 2.5475 | 100.3 | 21 | 32.51 | 79.0308 | 83.329 | 75.2229 | 0.00807 | 1.0134 | 0.024448 | 24.1243 | 0.045945 | 0.13909 | 0.000185 | 0.013607 | 0.013607 | 0.41604
2019-03-29 11:48 | 1553874478.8844 | 759 | 0.2095 | 0.000477 | -1e-06 | 99.7371 | 99.5315 | 125.9894 | 21.1299 | 0.002397 | 0.002177 | -151.7698 | 0.23446 | 8.4011 | 0.002505 | 0.004376 | 0.004195 | 6e-05 | 0.000875 | 0.27383 | 1.3472 | 2.5475 | 100.3 | 21 | 32.51 | 79.3664 | 81.9622 | 74.0782 | 0.00807 | 1.0168 | 0.024516 | 24.1123 | 0.018455 | 0.13948 | -1.3e-05 | -0.000937 | -0.000937 | -6.059
2019-03-29 11:48 | 1553874479.8849 | 760 | 0.20949 | 0.000477 | -1e-06 | 99.7099 | 99.3425 | 125.8892 | 21.13 | 0.002335 | 0.002173 | -151.7609 | 0.23407 | 8.4036 | 0.002524 | 0.003909 | 0.003829 | 6.3e-05 | 0.000876 | 0.27776 | 1.3347 | 2.5474 | 100.3 | 21 | 32.511 | 77.8587 | 82.4116 | 73.9884 | 0.008071 | 1.017 | 0.02451 | 24.0996 | 0.03302 | 0.13925 | 2.5e-05 | 0.001911 | 0.001911 | 2.9663
2019-03-29 11:48 | 1553874480.8845 | 761 | 0.20947 | 0.000473 | -1e-06 | 99.6934 | 99.4636 | 125.4578 | 21.1295 | 0.002318 | 0.002189 | -151.7417 | 0.23359 | 8.4025 | 0.002555 | 0.00402 | 0.003799 | 6.2e-05 | 0.000879 | 0.30254 | 1.313 | 2.5473 | 100.3 | 21 | 32.509 | 77.4498 | 83.1639 | 74.3724 | 0.00807 | 1.0159 | 0.024454 | 24.0715 | 0.020641 | 0.13896 | 0.000179 | 0.013192 | 0.013192 | 0.42873
2019-03-29 11:48 | 1553874481.8846 | 762 | 0.20943 | 0.000471 | -1e-06 | 99.556 | 99.3544 | 124.9982 | 21.1297 | 0.002323 | 0.002198 | -151.7618 | 0.23457 | 8.4022 | 0.002546 | 0.004091 | 0.003871 | 6.1e-05 | 0.000881 | 0.28943 | 1.3604 | 2.5475 | 100.3 | 21 | 32.506 | 77.592 | 82.9431 | 74.5782 | 0.008069 | 1.0153 | 0.024402 | 24.0345 | 0.018132 | 0.13955 | 0.000396 | 0.029147 | 0.029147 | 0.19486
2019-03-29 11:48 | 1553874482.884 | 763 | 0.20949 | 0.000467 | -1e-06 | 99.5988 | 99.3257 | 125.0834 | 21.1293 | 0.002384 | 0.002188 | -151.7828 | 0.23437 | 8.4014 | 0.002505 | 0.003656 | 0.003446 | 6.3e-05 | 0.000877 | 0.29518 | 1.3289 | 2.5473 | 100.3 | 21 | 32.504 | 79.0563 | 81.9568 | 74.3402 | 0.008069 | 1.016 | 0.024419 | 24.0345 | 0.024558 | 0.13943 | 3.3e-05 | 0.002508 | 0.002508 | 2.2627
2019-03-29 11:48 | 1553874483.8842 | 764 | 0.20952 | 0.000472 | -1e-06 | 99.5033 | 99.2295 | 125.3664 | 21.1291 | 0.002311 | 0.002162 | -151.7833 | 0.23397 | 8.4021 | 0.002582 | 0.003399 | 0.003318 | 5.8e-05 | 0.000875 | 0.27577 | 1.3387 | 2.5473 | 100.3 | 21 | 32.501 | 77.2865 | 83.7953 | 73.7156 | 0.008068 | 1.0178 | 0.024468 | 24.04 | 0.02465 | 0.13919 | -0.000162 | -0.011888 | -0.011888 | -0.47653
2019-03-29 11:48 | 1553874484.8842 | 765 | 0.20954 | 0.000471 | -1e-06 | 99.3716 | 99.1274 | 125.6028 | 21.1299 | 0.00231 | 0.002161 | -151.7791 | 0.23422 | 8.4017 | 0.002535 | 0.004289 | 0.003989 | 6.1e-05 | 0.000871 | 0.29717 | 1.355 | 2.5473 | 100.3 | 21 | 32.494 | 77.2693 | 82.6729 | 73.7001 | 0.008066 | 1.0179 | 0.024492 | 24.0622 | 0.022006 | 0.13934 | -0.000255 | -0.018726 | -0.018726 | -0.30285
2019-03-29 11:48 | 1553874485.8847 | 766 | 0.20953 | 0.000472 | -1e-06 | 99.3767 | 99.1329 | 125.7087 | 21.1297 | 0.002306 | 0.00217 | -151.8 | 0.23369 | 8.4022 | 0.002469 | 0.003792 | 0.003712 | 6.2e-05 | 0.000867 | 0.28368 | 1.3202 | 2.5474 | 100.3 | 21 | 32.49 | 77.1655 | 81.0854 | 73.9034 | 0.008065 | 1.0173 | 0.024495 | 24.0793 | 0.021971 | 0.13902 | -0.000179 | -0.013128 | -0.013128 | -0.43099
2019-03-29 11:48 | 1553874486.8848 | 767 | 0.20953 | 0.000469 | -1e-06 | 99.1456 | 98.7961 | 125.7868 | 21.13 | 0.002289 | 0.002154 | -151.795 | 0.23387 | 8.4018 | 0.002521 | 0.00371 | 0.003567 | 6.5e-05 | 0.000872 | 0.29789 | 1.3503 | 2.5475 | 100.3 | 21 | 32.483 | 76.7612 | 82.328 | 73.5116 | 0.008064 | 1.0184 | 0.024517 | 24.0732 | 0.031593 | 0.13913 | -0.000199 | -0.014608 | -0.014608 | -0.38763
2019-03-29 11:48 | 1553874487.8843 | 768 | 0.20951 | 0.000467 | -1e-06 | 99.1441 | 98.8053 | 125.7282 | 21.1296 | 0.00229 | 0.002152 | -151.7787 | 0.23391 | 8.4003 | 0.002515 | 0.003623 | 0.00339 | 6.6e-05 | 0.00087 | 0.27527 | 1.3694 | 2.5475 | 100.3 | 21 | 32.477 | 76.7952 | 82.1951 | 73.4775 | 0.008062 | 1.0185 | 0.024512 | 24.0664 | 0.030624 | 0.13915 | -7.4e-05 | -0.005341 | -0.005341 | -1.0605
2019-03-29 11:48 | 1553874488.8828 | 769 | 0.20953 | 0.000469 | -1e-06 | 99.0561 | 98.6681 | 125.569 | 21.1299 | 0.002331 | 0.002125 | -151.7869 | 0.23459 | 8.4007 | 0.002472 | 0.003833 | 0.003671 | 6.2e-05 | 0.000871 | 0.31564 | 1.3721 | 2.5474 | 100.3 | 21 | 32.473 | 77.7697 | 81.1527 | 72.8144 | 0.008061 | 1.0205 | 0.02452 | 24.0282 | 0.035109 | 0.13956 | -0.000222 | -0.016308 | -0.016308 | -0.34828
2019-03-29 11:48 | 1553874489.8828 | 770 | 0.20954 | 0.000467 | -1e-06 | 98.9247 | 98.6319 | 125.6526 | 21.1297 | 0.002246 | 0.002169 | -151.7853 | 0.23444 | 8.4011 | 0.002524 | 0.003706 | 0.003512 | 6.6e-05 | 0.000878 | 0.27433 | 1.345 | 2.5473 | 100.3 | 21 | 32.463 | 75.7343 | 82.4003 | 73.8752 | 0.008059 | 1.0173 | 0.024491 | 24.073 | 0.026516 | 0.13947 | -0.000255 | -0.018671 | -0.018671 | -0.30403
2019-03-29 11:48 | 1553874490.8833 | 771 | 0.20953 | 0.000467 | -2e-06 | 98.8247 | 98.3561 | 125.7939 | 21.1296 | 0.002268 | 0.002127 | -151.792 | 0.23432 | 8.401 | 0.002527 | 0.004002 | 0.003857 | 6.5e-05 | 0.000878 | 0.25984 | 1.3713 | 2.5473 | 100.3 | 21 | 32.458 | 76.2673 | 82.4742 | 72.8634 | 0.008057 | 1.0203 | 0.02454 | 24.0514 | 0.042514 | 0.1394 | -0.000179 | -0.013095 | -0.013095 | -0.43325
2019-03-29 11:48 | 1553874491.883 | 772 | 0.20952 | 0.000467 | -1e-06 | 98.7206 | 98.2138 | 125.8788 | 21.1307 | 0.002286 | 0.002153 | -151.7727 | 0.23443 | 8.4031 | 0.002514 | 0.003198 | 0.002984 | 6.5e-05 | 0.000877 | 0.27383 | 1.3653 | 2.5474 | 100.3 | 21 | 32.451 | 76.6816 | 82.1635 | 73.5067 | 0.008056 | 1.0184 | 0.024526 | 24.0819 | 0.046033 | 0.13946 | -0.000131 | -0.009574 | -0.009574 | -0.59285
2019-03-29 11:48 | 1553874492.8839 | 773 | 0.20948 | 0.000463 | -1e-06 | 98.5161 | 98.0271 | 125.9878 | 21.1303 | 0.002273 | 0.002178 | -151.7728 | 0.23437 | 8.4039 | 0.002446 | 0.004191 | 0.004019 | 6.8e-05 | 0.00088 | 0.29767 | 1.3249 | 2.5474 | 100.3 | 21 | 32.442 | 76.3694 | 80.534 | 74.1043 | 0.008053 | 1.0167 | 0.024515 | 24.113 | 0.044508 | 0.13943 | 0.000138 | 0.010207 | 0.010207 | 0.55597
2019-03-29 11:48 | 1553874493.8835 | 774 | 0.20944 | 0.000466 | -2e-06 | 98.4096 | 98.0285 | 125.9295 | 21.1313 | 0.002208 | 0.00214 | -151.778 | 0.23341 | 8.3999 | 0.002477 | 0.003553 | 0.003375 | 6.8e-05 | 0.000884 | 0.31188 | 1.3583 | 2.5474 | 100.3 | 21.01 | 32.438 | 74.8152 | 81.2795 | 73.1923 | 0.008057 | 1.0194 | 0.024542 | 24.0758 | 0.034703 | 0.13886 | 0.000378 | 0.027962 | 0.027962 | 0.20211
2019-03-29 11:48 | 1553874494.884 | 775 | 0.20948 | 0.000465 | -1e-06 | 98.2791 | 97.709 | 125.7155 | 21.131 | 0.002296 | 0.002095 | -151.781 | 0.23325 | 8.4028 | 0.002502 | 0.003654 | 0.003505 | 6.3e-05 | 0.000885 | 0.30287 | 1.341 | 2.5474 | 100.3 | 21.01 | 32.431 | 76.9368 | 81.8789 | 72.0983 | 0.008056 | 1.0226 | 0.02456 | 24.0173 | 0.052029 | 0.13876 | 0.000115 | 0.00859 | 0.00859 | 0.65743
2019-03-29 11:48 | 1553874495.8841 | 776 | 0.20947 | 0.000468 | -1e-06 | 98.142 | 97.6503 | 125.5038 | 21.1309 | 0.002298 | 0.002135 | -151.7551 | 0.23388 | 8.4002 | 0.002525 | 0.003855 | 0.003753 | 6.6e-05 | 0.000885 | 0.27527 | 1.354 | 2.5475 | 100.3 | 21.01 | 32.426 | 76.9812 | 82.4281 | 73.0752 | 0.008054 | 1.0197 | 0.024504 | 24.031 | 0.044918 | 0.13914 | 0.000145 | 0.01077 | 0.01077 | 0.52581
2019-03-29 11:48 | 1553874496.8831 | 777 | 0.20949 | 0.000471 | -1e-06 | 97.9324 | 97.5974 | 125.2404 | 21.1308 | 0.002236 | 0.00211 | -151.7766 | 0.23372 | 8.4024 | 0.002568 | 0.003676 | 0.003481 | 6.6e-05 | 0.000893 | 0.30558 | 1.3509 | 2.5474 | 100.3 | 21.01 | 32.422 | 75.4913 | 83.477 | 72.4552 | 0.008053 | 1.0215 | 0.024501 | 23.9843 | 0.030648 | 0.13904 | 6.2e-05 | 0.004682 | 0.004682 | 1.2087
2019-03-29 11:48 | 1553874497.8841 | 778 | 0.20949 | 0.000468 | -1e-06 | 97.8298 | 97.2599 | 125.2845 | 21.131 | 0.00227 | 0.002103 | -151.7971 | 0.2342 | 8.4038 | 0.002526 | 0.003072 | 0.002912 | 7.1e-05 | 0.00089 | 0.30165 | 1.3126 | 2.5474 | 100.3 | 21.01 | 32.415 | 76.3061 | 82.4585 | 72.2878 | 0.008052 | 1.022 | 0.024511 | 23.9827 | 0.052254 | 0.13933 | 4.3e-05 | 0.003256 | 0.003256 | 1.7416
2019-03-29 11:48 | 1553874498.8831 | 779 | 0.20949 | 0.000472 | -1e-06 | 97.7361 | 97.2886 | 125.508 | 21.131 | 0.002296 | 0.002065 | -151.7848 | 0.23297 | 8.4028 | 0.002542 | 0.003897 | 0.003636 | 6.9e-05 | 0.000889 | 0.30862 | 1.3513 | 2.5474 | 100.3 | 21.01 | 32.411 | 76.9453 | 82.8377 | 71.3873 | 0.008051 | 1.0247 | 0.024565 | 23.9727 | 0.041047 | 0.13859 | 4e-05 | 0.003036 | 0.003036 | 1.8578
2019-03-29 11:48 | 1553874499.8836 | 780 | 0.2095 | 0.000469 | -1e-06 | 97.6316 | 97.1128 | 125.623 | 21.1314 | 0.00221 | 0.0021 | -151.7921 | 0.23318 | 8.4029 | 0.00254 | 0.004351 | 0.004111 | 7.2e-05 | 0.00089 | 0.30055 | 1.3503 | 2.5473 | 100.3 | 21.01 | 32.406 | 74.8668 | 82.796 | 72.2314 | 0.008049 | 1.0222 | 0.024546 | 24.0131 | 0.047656 | 0.13872 | -8e-06 | -0.000481 | -0.000481 | -11.731
2019-03-29 11:48 | 1553874500.8832 | 781 | 0.20948 | 0.000467 | -1e-06 | 97.4831 | 97.0016 | 125.608 | 21.1317 | 0.002236 | 0.002085 | -151.8067 | 0.23326 | 8.4024 | 0.002521 | 0.003763 | 0.003656 | 7e-05 | 0.000893 | 0.3147 | 1.368 | 2.5475 | 100.3 | 21.01 | 32.406 | 75.4837 | 82.3508 | 71.8753 | 0.008049 | 1.0232 | 0.024557 | 23.9993 | 0.044297 | 0.13877 | 9.2e-05 | 0.006851 | 0.006851 | 0.82437
2019-03-29 11:48 | 1553874501.8832 | 782 | 0.2095 | 0.000465 | -1e-06 | 97.3933 | 96.8869 | 125.7753 | 21.132 | 0.002182 | 0.002095 | -151.8151 | 0.23345 | 8.4058 | 0.002558 | 0.004 | 0.00383 | 7.2e-05 | 0.000895 | 0.31238 | 1.3674 | 2.5474 | 100.3 | 21.01 | 32.406 | 74.1987 | 83.2271 | 72.0985 | 0.008049 | 1.0226 | 0.024566 | 24.023 | 0.046634 | 0.13888 | -3.5e-05 | -0.002507 | -0.002507 | -2.2544
2019-03-29 11:48 | 1553874502.8838 | 783 | 0.20954 | 0.000471 | -2e-06 | 97.2556 | 96.5292 | 125.6352 | 21.1317 | 0.002238 | 0.002113 | -151.8165 | 0.23408 | 8.4043 | 0.002594 | 0.003235 | 0.002981 | 6.8e-05 | 0.000894 | 0.30879 | 1.3835 | 2.5475 | 100.3 | 21.01 | 32.409 | 75.5375 | 84.0871 | 72.531 | 0.00805 | 1.0213 | 0.024536 | 24.0247 | 0.06706 | 0.13925 | -0.000265 | -0.019477 | -0.019477 | -0.29099
2019-03-29 11:48 | 1553874503.8833 | 784 | 0.20955 | 0.000466 | -2e-06 | 97.159 | 96.6169 | 125.5615 | 21.1321 | 0.002261 | 0.002084 | -151.802 | 0.23412 | 8.4015 | 0.00258 | 0.003568 | 0.003427 | 7.2e-05 | 0.000892 | 0.28528 | 1.3601 | 2.5474 | 100.3 | 21.01 | 32.411 | 76.1023 | 83.7646 | 71.85 | 0.008051 | 1.0233 | 0.024553 | 23.9939 | 0.05004 | 0.13928 | -0.00031 | -0.022771 | -0.022771 | -0.24894
2019-03-29 11:48 | 1553874504.8828 | 785 | 0.20947 | 0.000464 | -1e-06 | 97.097 | 96.5074 | 125.7789 | 21.1325 | 0.002229 | 0.002037 | -151.8111 | 0.23374 | 8.4015 | 0.002553 | 0.004317 | 0.004092 | 6.8e-05 | 0.000892 | 0.32388 | 1.3721 | 2.5474 | 100.3 | 21.01 | 32.414 | 75.3223 | 83.1004 | 70.7082 | 0.008051 | 1.0267 | 0.024615 | 23.9749 | 0.054482 | 0.13905 | 0.000154 | 0.011448 | 0.011448 | 0.49438
2019-03-29 11:48 | 1553874505.8824 | 786 | 0.20951 | 0.000466 | -1e-06 | 96.9653 | 96.3988 | 125.9103 | 21.1321 | 0.002146 | 0.002108 | -151.8187 | 0.23401 | 8.4017 | 0.002548 | 0.003729 | 0.003644 | 7.2e-05 | 0.000897 | 0.31045 | 1.3746 | 2.5476 | 100.3 | 21.01 | 32.417 | 73.3272 | 83.0001 | 72.4106 | 0.008052 | 1.0217 | 0.024568 | 24.0468 | 0.052405 | 0.13922 | -6.6e-05 | -0.004851 | -0.004851 | -1.168
2019-03-29 11:48 | 1553874506.8834 | 787 | 0.20951 | 0.000466 | -1e-06 | 96.9105 | 96.3504 | 126.0206 | 21.1326 | 0.002163 | 0.00212 | -151.8183 | 0.23388 | 8.4036 | 0.002494 | 0.003835 | 0.003637 | 6.8e-05 | 0.00089 | 0.31277 | 1.3142 | 2.5473 | 100.3 | 21.01 | 32.42 | 73.7433 | 81.6949 | 72.7118 | 0.008053 | 1.0208 | 0.024568 | 24.0678 | 0.051849 | 0.13914 | -4.2e-05 | -0.003019 | -0.003019 | -1.8758
2019-03-29 11:48 | 1553874507.8838 | 788 | 0.20949 | 0.000465 | -1e-06 | 96.824 | 95.9518 | 126.1955 | 21.1332 | 0.002195 | 0.002064 | -151.8243 | 0.23388 | 8.4006 | 0.002552 | 0.004097 | 0.003934 | 6.9e-05 | 0.000885 | 0.25426 | 1.3249 | 2.5473 | 100.3 | 21.01 | 32.425 | 74.5194 | 83.0792 | 71.3699 | 0.008054 | 1.0247 | 0.024633 | 24.0377 | 0.080943 | 0.13914 | 8.8e-05 | 0.006612 | 0.006612 | 0.85648
2019-03-29 11:48 | 1553874508.8825 | 789 | 0.20953 | 0.000467 | -1e-06 | 96.768 | 96.177 | 125.8357 | 21.1324 | 0.002181 | 0.002091 | -151.8266 | 0.23338 | 8.4016 | 0.002432 | 0.003448 | 0.003249 | 6.8e-05 | 0.00089 | 0.27632 | 1.3207 | 2.5474 | 100.3 | 21.01 | 32.425 | 74.1624 | 80.2022 | 72.0038 | 0.008054 | 1.0229 | 0.024575 | 24.0255 | 0.054798 | 0.13884 | -0.00019 | -0.013933 | -0.013933 | -0.40557
2019-03-29 11:48 | 1553874509.8834 | 790 | 0.20952 | 0.000466 | -1e-06 | 96.6794 | 96.0299 | 125.612 | 21.1335 | 0.002136 | 0.002059 | -151.8211 | 0.23431 | 8.4041 | 0.002447 | 0.003171 | 0.002982 | 6.7e-05 | 0.000893 | 0.2959 | 1.3424 | 2.5473 | 100.3 | 21.02 | 32.431 | 73.0846 | 80.562 | 71.2306 | 0.008061 | 1.0252 | 0.02458 | 23.9772 | 0.060295 | 0.13939 | -0.000121 | -0.008866 | -0.008866 | -0.63986
2019-03-29 11:48 | 1553874510.8835 | 791 | 0.20951 | 0.000464 | -2e-06 | 96.5833 | 95.821 | 125.6817 | 21.1332 | 0.002202 | 0.002045 | -151.8223 | 0.23411 | 8.3994 | 0.002428 | 0.004166 | 0.003994 | 6.8e-05 | 0.000898 | 0.30071 | 1.3472 | 2.5474 | 100.3 | 21.02 | 32.434 | 74.6737 | 80.1119 | 70.8968 | 0.008061 | 1.0262 | 0.024599 | 23.9722 | 0.070877 | 0.13927 | -4.8e-05 | -0.003452 | -0.003452 | -1.6422
2019-03-29 11:49 | 1553874511.8839 | 792 | 0.20951 | 0.000467 | -1e-06 | 96.4235 | 95.7617 | 125.8918 | 21.1329 | 0.002146 | 0.002095 | -151.8265 | 0.23347 | 8.4016 | 0.00246 | 0.003721 | 0.003566 | 6.7e-05 | 0.000897 | 0.31404 | 1.3293 | 2.5474 | 100.3 | 21.02 | 32.437 | 73.3257 | 80.8849 | 72.1174 | 0.008062 | 1.0225 | 0.024576 | 24.0348 | 0.061607 | 0.13889 | -8.9e-05 | -0.006528 | -0.006528 | -0.86594
2019-03-29 11:49 | 1553874512.8837 | 793 | 0.20953 | 0.00047 | -1e-06 | 96.2705 | 95.5982 | 126.3342 | 21.133 | 0.00218 | 0.002073 | -151.83 | 0.23354 | 8.4017 | 0.002431 | 0.00393 | 0.003718 | 7e-05 | 0.000902 | 0.28152 | 1.3269 | 2.5473 | 100.3 | 21.02 | 32.442 | 74.1435 | 80.1896 | 71.5841 | 0.008063 | 1.0241 | 0.024638 | 24.0584 | 0.062689 | 0.13893 | -0.000198 | -0.014596 | -0.014596 | -0.3874
2019-03-29 11:49 | 1553874513.8837 | 794 | 0.2095 | 0.000468 | -1e-06 | 96.2047 | 95.4924 | 126.3004 | 21.1332 | 0.002142 | 0.002051 | -151.8363 | 0.23398 | 8.4032 | 0.002463 | 0.003625 | 0.003537 | 6.8e-05 | 0.000904 | 0.27273 | 1.3402 | 2.5475 | 100.3 | 21.02 | 32.443 | 73.2379 | 80.9454 | 71.0518 | 0.008064 | 1.0257 | 0.024654 | 24.0366 | 0.066473 | 0.1392 | -2.3e-05 | -0.001636 | -0.001636 | -3.4626
2019-03-29 11:49 | 1553874514.8826 | 795 | 0.20955 | 0.00047 | -1e-06 | 96.0514 | 95.2654 | 125.9246 | 21.1323 | 0.00215 | 0.002067 | -151.8268 | 0.23399 | 8.4015 | 0.002433 | 0.003772 | 0.003576 | 6.9e-05 | 0.000909 | 0.29927 | 1.3409 | 2.5474 | 100.3 | 21.02 | 32.442 | 73.4163 | 80.2185 | 71.4292 | 0.008063 | 1.0246 | 0.024604 | 24.014 | 0.073496 | 0.1392 | -0.00034 | -0.025069 | -0.025069 | -0.22599
2019-03-29 11:49 | 1553874515.8828 | 796 | 0.20953 | 0.000475 | -1e-06 | 95.9982 | 95.2511 | 125.7155 | 21.133 | 0.002134 | 0.00203 | -151.8325 | 0.23428 | 8.4029 | 0.0024 | 0.003868 | 0.003655 | 7e-05 | 0.000904 | 0.28871 | 1.3848 | 2.5476 | 100.3 | 21.02 | 32.442 | 73.046 | 79.4302 | 70.5544 | 0.008063 | 1.0272 | 0.024615 | 23.9635 | 0.069879 | 0.13937 | -0.000217 | -0.01595 | -0.01595 | -0.35564
2019-03-29 11:49 | 1553874516.8827 | 797 | 0.20951 | 0.00047 | -1e-06 | 95.8425 | 95.1378 | 125.5087 | 21.1328 | 0.002183 | 0.002031 | -151.8215 | 0.23395 | 8.4028 | 0.002427 | 0.004182 | 0.003979 | 6.9e-05 | 0.000906 | 0.27992 | 1.3854 | 2.5474 | 100.3 | 21.02 | 32.439 | 74.2134 | 80.0716 | 70.579 | 0.008063 | 1.0271 | 0.024594 | 23.9447 | 0.066011 | 0.13918 | -7.2e-05 | -0.005231 | -0.005231 | -1.0828
2019-03-29 11:49 | 1553874517.8828 | 798 | 0.20956 | 0.000468 | -1e-06 | 95.7793 | 94.9177 | 125.4566 | 21.1336 | 0.002175 | 0.002064 | -151.8173 | 0.23378 | 8.4034 | 0.002396 | 0.004035 | 0.00391 | 6.9e-05 | 0.000908 | 0.30901 | 1.376 | 2.5474 | 100.3 | 21.02 | 32.436 | 74.0239 | 79.3411 | 71.3572 | 0.008062 | 1.0248 | 0.024561 | 23.9668 | 0.080835 | 0.13908 | -0.000359 | -0.026454 | -0.026454 | -0.21397
2019-03-29 11:49 | 1553874518.8838 | 799 | 0.20954 | 0.000466 | -2e-06 | 95.5985 | 94.8422 | 125.7061 | 21.1337 | 0.002136 | 0.002012 | -151.8186 | 0.23388 | 8.3988 | 0.002457 | 0.004203 | 0.003983 | 6.3e-05 | 0.000905 | 0.27489 | 1.3304 | 2.5475 | 100.3 | 21.02 | 32.434 | 73.0956 | 80.8071 | 70.1174 | 0.008061 | 1.0285 | 0.02463 | 23.9474 | 0.071045 | 0.13913 | -0.000225 | -0.016565 | -0.016565 | -0.34185
2019-03-29 11:49 | 1553874519.8818 | 800 | 0.20954 | 0.000466 | -1e-06 | 95.4977 | 94.5955 | 126.0131 | 21.1337 | 0.002192 | 0.002031 | -151.8281 | 0.23273 | 8.4024 | 0.002376 | 0.003596 | 0.003411 | 6.8e-05 | 0.000905 | 0.29479 | 1.3433 | 2.5475 | 100.3 | 21.02 | 32.429 | 74.4296 | 78.8664 | 70.5621 | 0.00806 | 1.0272 | 0.024644 | 23.9921 | 0.084904 | 0.13845 | -0.000257 | -0.019013 | -0.019013 | -0.29637
2019-03-29 11:49 | 1553874520.8828 | 801 | 0.2095 | 0.000465 | -1e-06 | 95.4246 | 94.6095 | 125.9917 | 21.1345 | 0.002103 | 0.002059 | -151.827 | 0.23269 | 8.4046 | 0.002377 | 0.003787 | 0.003527 | 6.4e-05 | 0.000908 | 0.30663 | 1.3242 | 2.5474 | 100.3 | 21.02 | 32.427 | 72.3045 | 78.8936 | 71.2513 | 0.00806 | 1.0251 | 0.024617 | 24.0142 | 0.076737 | 0.13843 | 8e-06 | 0.000652 | 0.000652 | 8.6418
2019-03-29 11:49 | 1553874521.8829 | 802 | 0.20953 | 0.000465 | -1e-06 | 95.2606 | 94.5562 | 125.8131 | 21.1339 | 0.002203 | 0.002093 | -151.8275 | 0.23343 | 8.4028 | 0.002408 | 0.003796 | 0.003663 | 6.8e-05 | 0.000907 | 0.2959 | 1.3742 | 2.5475 | 100.3 | 21.02 | 32.423 | 74.7055 | 79.6307 | 72.0535 | 0.008059 | 1.0227 | 0.024571 | 24.0251 | 0.066383 | 0.13887 | -0.000202 | -0.014829 | -0.014829 | -0.38114
2019-03-29 11:49 | 1553874522.8828 | 803 | 0.20947 | 0.000465 | -1e-06 | 95.1783 | 94.1842 | 125.2723 | 21.1338 | 0.002179 | 0.002006 | -151.8188 | 0.23333 | 8.4024 | 0.002421 | 0.003387 | 0.003165 | 6.5e-05 | 0.000908 | 0.28008 | 1.3856 | 2.5475 | 100.3 | 21.03 | 32.418 | 74.1318 | 79.9385 | 69.9772 | 0.008062 | 1.0289 | 0.024592 | 23.9012 | 0.09391 | 0.13881 | 0.000164 | 0.012169 | 0.012169 | 0.46424
2019-03-29 11:49 | 1553874523.8819 | 804 | 0.2095 | 0.000467 | -1e-06 | 94.9801 | 94.0986 | 125.0512 | 21.1343 | 0.00211 | 0.002018 | -151.8364 | 0.2339 | 8.4028 | 0.002405 | 0.003844 | 0.003691 | 6.6e-05 | 0.000913 | 0.31205 | 1.3812 | 2.5476 | 100.3 | 21.03 | 32.415 | 72.4609 | 79.5484 | 70.2607 | 0.008062 | 1.0281 | 0.02456 | 23.8899 | 0.083407 | 0.13915 | -1.2e-05 | -0.000809 | -0.000809 | -7.0019
2019-03-29 11:49 | 1553874524.882 | 805 | 0.20949 | 0.000469 | -2e-06 | 94.8676 | 93.9481 | 125.0801 | 21.1344 | 0.002109 | 0.002019 | -151.8601 | 0.23418 | 8.4022 | 0.002444 | 0.004107 | 0.003927 | 6.2e-05 | 0.000914 | 0.28241 | 1.4059 | 2.5473 | 100.3 | 21.03 | 32.41 | 72.4555 | 80.5012 | 70.2789 | 0.00806 | 1.028 | 0.024562 | 23.8933 | 0.08712 | 0.13932 | 3.5e-05 | 0.002724 | 0.002724 | 2.0815
2019-03-29 11:49 | 1553874525.883 | 806 | 0.20948 | 0.000464 | -1e-06 | 94.7377 | 93.8146 | 125.4507 | 21.1344 | 0.002129 | 0.002072 | -151.8717 | 0.23403 | 8.403 | 0.002351 | 0.003378 | 0.003185 | 6.4e-05 | 0.000922 | 0.28694 | 1.3835 | 2.5474 | 100.3 | 21.03 | 32.407 | 72.9125 | 78.2599 | 71.5597 | 0.00806 | 1.0242 | 0.024553 | 23.9733 | 0.087577 | 0.13923 | 0.000134 | 0.009929 | 0.009929 | 0.57071
2019-03-29 11:49 | 1553874526.883 | 807 | 0.20947 | 0.000472 | -1e-06 | 94.5733 | 93.5528 | 125.6376 | 21.1345 | 0.002178 | 0.002046 | -151.8666 | 0.23329 | 8.4033 | 0.002357 | 0.003501 | 0.003318 | 6.8e-05 | 0.00092 | 0.27417 | 1.4103 | 2.5475 | 100.3 | 21.03 | 32.405 | 74.1067 | 78.3987 | 70.9293 | 0.008059 | 1.0261 | 0.024594 | 23.9692 | 0.097041 | 0.13878 | 0.00017 | 0.012619 | 0.012619 | 0.44762
2019-03-29 11:49 | 1553874527.8822 | 808 | 0.20949 | 0.00047 | -1e-06 | 94.4273 | 93.6508 | 125.8366 | 21.1342 | 0.002099 | 0.002017 | -151.8687 | 0.23405 | 8.4031 | 0.002436 | 0.0038 | 0.003613 | 6.6e-05 | 0.000925 | 0.29048 | 1.4035 | 2.5474 | 100.3 | 21.03 | 32.403 | 72.1989 | 80.2967 | 70.2349 | 0.008059 | 1.0281 | 0.024638 | 23.9639 | 0.073855 | 0.13923 | 2.2e-05 | 0.001721 | 0.001721 | 3.292
2019-03-29 11:49 | 1553874528.882 | 809 | 0.20949 | 0.000467 | -1e-06 | 94.248 | 93.1726 | 125.8239 | 21.1346 | 0.002152 | 0.002073 | -151.8691 | 0.23385 | 8.4005 | 0.00242 | 0.003903 | 0.003804 | 6.9e-05 | 0.000926 | 0.27527 | 1.4183 | 2.5474 | 100.3 | 21.03 | 32.402 | 73.4855 | 79.9081 | 71.5867 | 0.008058 | 1.0241 | 0.024589 | 24.0098 | 0.10264 | 0.13912 | 6.2e-05 | 0.004631 | 0.004631 | 1.2227
2019-03-29 11:49 | 1553874529.8821 | 810 | 0.2095 | 0.000472 | -1e-06 | 94.0785 | 93.0634 | 125.8709 | 21.1342 | 0.002183 | 0.002046 | -151.8836 | 0.2336 | 8.4013 | 0.00241 | 0.003675 | 0.003401 | 6.7e-05 | 0.000928 | 0.29717 | 1.3954 | 2.5475 | 100.3 | 21.03 | 32.401 | 74.2248 | 79.6661 | 70.9179 | 0.008058 | 1.0261 | 0.024617 | 23.991 | 0.097036 | 0.13897 | -1.7e-05 | -0.001189 | -0.001189 | -4.7584
2019-03-29 11:49 | 1553874530.8832 | 811 | 0.20948 | 0.000471 | -2e-06 | 93.9205 | 92.8337 | 125.8004 | 21.1344 | 0.002107 | 0.001972 | -151.8758 | 0.23456 | 8.4016 | 0.002304 | 0.003749 | 0.003569 | 7e-05 | 0.000934 | 0.29247 | 1.3254 | 2.5474 | 100.3 | 21.03 | 32.398 | 72.3986 | 77.1208 | 69.1397 | 0.008057 | 1.0314 | 0.024674 | 23.9222 | 0.1041 | 0.13954 | 0.00012 | 0.009024 | 0.009024 | 0.62937
2019-03-29 11:49 | 1553874531.8822 | 812 | 0.2095 | 0.000474 | -1e-06 | 93.7532 | 92.7221 | 125.5596 | 21.1343 | 0.002177 | 0.001997 | -151.8738 | 0.23243 | 8.4044 | 0.002405 | 0.003415 | 0.003243 | 6.9e-05 | 0.000929 | 0.28567 | 1.3052 | 2.5474 | 100.3 | 21.04 | 32.395 | 74.0763 | 79.5578 | 69.7546 | 0.008062 | 1.0296 | 0.024628 | 23.9208 | 0.09891 | 0.13827 | 1e-05 | 0.000758 | 0.000758 | 7.4224
2019-03-29 11:49 | 1553874532.8833 | 813 | 0.20951 | 0.000473 | -1e-06 | 93.5769 | 92.486 | 125.6296 | 21.1349 | 0.002139 | 0.002019 | -151.8725 | 0.23368 | 8.4027 | 0.002407 | 0.003503 | 0.003248 | 6.8e-05 | 0.000935 | 0.28368 | 1.349 | 2.5474 | 100.3 | 21.04 | 32.389 | 73.1569 | 79.606 | 70.2816 | 0.00806 | 1.028 | 0.024616 | 23.9458 | 0.10489 | 0.13902 | -5.2e-05 | -0.00379 | -0.00379 | -1.4928
2019-03-29 11:49 | 1553874533.8832 | 814 | 0.20949 | 0.000474 | -2e-06 | 93.482 | 92.385 | 125.6535 | 21.1344 | 0.002149 | 0.002038 | -151.866 | 0.23403 | 8.4033 | 0.002403 | 0.004314 | 0.004123 | 6.8e-05 | 0.000941 | 0.28567 | 1.3212 | 2.5474 | 100.3 | 21.04 | 32.385 | 73.3981 | 79.5125 | 70.7497 | 0.008059 | 1.0266 | 0.024602 | 23.9644 | 0.10559 | 0.13922 | 7.8e-05 | 0.005875 | 0.005875 | 0.96452
2019-03-29 11:49 | 1553874534.8827 | 815 | 0.20952 | 0.000476 | -1e-06 | 93.2913 | 92.1106 | 125.904 | 21.1345 | 0.002113 | 0.001995 | -151.8531 | 0.23375 | 8.4014 | 0.002391 | 0.003462 | 0.003333 | 6.6e-05 | 0.000943 | 0.32228 | 1.3507 | 2.5474 | 100.3 | 21.04 | 32.375 | 72.5383 | 79.2142 | 69.702 | 0.008057 | 1.0297 | 0.024664 | 23.9517 | 0.11393 | 0.13906 | -0.000143 | -0.010593 | -0.010593 | -0.53428
2019-03-29 11:49 | 1553874535.8814 | 816 | 0.20949 | 0.000474 | -1e-06 | 93.0991 | 91.9727 | 125.8195 | 21.1351 | 0.002137 | 0.002045 | -151.853 | 0.2334 | 8.4036 | 0.002469 | 0.004035 | 0.003837 | 6.6e-05 | 0.000946 | 0.28064 | 1.3399 | 2.5474 | 100.3 | 21.04 | 32.368 | 73.1251 | 81.0906 | 70.9056 | 0.008055 | 1.0261 | 0.024612 | 23.9857 | 0.10888 | 0.13885 | 2.2e-05 | 0.001683 | 0.001683 | 3.358
2019-03-29 11:49 | 1553874536.8814 | 817 | 0.20951 | 0.000476 | -1e-06 | 93.0057 | 91.7932 | 125.7361 | 21.1343 | 0.002067 | 0.002056 | -151.8598 | 0.23395 | 8.4024 | 0.002451 | 0.003739 | 0.003628 | 6.6e-05 | 0.000938 | 0.30829 | 1.3733 | 2.5476 | 100.3 | 21.04 | 32.359 | 71.4273 | 80.6643 | 71.1723 | 0.008053 | 1.0253 | 0.024595 | 23.987 | 0.11738 | 0.13918 | -5.3e-05 | -0.003826 | -0.003826 | -1.4807
2019-03-29 11:49 | 1553874537.8818 | 818 | 0.20947 | 0.000476 | -1e-06 | 92.8055 | 91.6945 | 125.604 | 21.1343 | 0.002165 | 0.002007 | -151.8674 | 0.23463 | 8.401 | 0.002542 | 0.003549 | 0.003402 | 6.4e-05 | 0.000939 | 0.30702 | 1.3453 | 2.5475 | 100.3 | 21.04 | 32.341 | 73.7814 | 82.8384 | 70.0043 | 0.008048 | 1.0288 | 0.024624 | 23.9337 | 0.10772 | 0.13958 | 0.000143 | 0.010597 | 0.010597 | 0.53606
2019-03-29 11:49 | 1553874538.8816 | 819 | 0.20952 | 0.000476 | -1e-06 | 92.6896 | 91.4423 | 125.5488 | 21.1353 | 0.002097 | 0.002028 | -151.8743 | 0.23308 | 8.4023 | 0.002461 | 0.003252 | 0.003044 | 6.6e-05 | 0.000942 | 0.3027 | 1.3128 | 2.5475 | 100.3 | 21.04 | 32.33 | 72.1501 | 80.914 | 70.4896 | 0.008045 | 1.0274 | 0.024601 | 23.9454 | 0.12118 | 0.13866 | -0.000144 | -0.010573 | -0.010573 | -0.53373
2019-03-29 11:49 | 1553874539.8815 | 820 | 0.2095 | 0.000479 | -1e-06 | 92.5977 | 91.4283 | 125.1951 | 21.1351 | 0.00209 | 0.001981 | -151.8681 | 0.2337 | 8.4046 | 0.002459 | 0.003226 | 0.003094 | 6.5e-05 | 0.000939 | 0.27168 | 1.3465 | 2.5473 | 100.3 | 21.05 | 32.318 | 71.99 | 80.8498 | 69.3742 | 0.008047 | 1.0307 | 0.024606 | 23.8728 | 0.11368 | 0.13903 | 8e-06 | 0.000688 | 0.000688 | 8.2222
2019-03-29 11:49 | 1553874540.8824 | 821 | 0.20951 | 0.000477 | -1e-06 | 92.5082 | 91.1775 | 125.3014 | 21.1354 | 0.002051 | 0.001984 | -151.8681 | 0.23316 | 8.4019 | 0.002448 | 0.003917 | 0.003733 | 6.4e-05 | 0.000939 | 0.29878 | 1.3504 | 2.5474 | 100.3 | 21.05 | 32.294 | 71.0601 | 80.5816 | 69.4391 | 0.008041 | 1.0305 | 0.024614 | 23.8852 | 0.12961 | 0.1387 | -7.3e-05 | -0.005353 | -0.005353 | -1.0545
2019-03-29 11:49 | 1553874541.8827 | 822 | 0.20952 | 0.000477 | -2e-06 | 92.2849 | 91.0171 | 125.4965 | 21.1353 | 0.002053 | 0.002006 | -151.8735 | 0.23383 | 8.4035 | 0.002515 | 0.004191 | 0.003952 | 6.7e-05 | 0.000934 | 0.28584 | 1.3277 | 2.5473 | 100.3 | 21.05 | 32.282 | 71.0892 | 82.2076 | 69.9748 | 0.008038 | 1.0289 | 0.024614 | 23.9225 | 0.12374 | 0.1391 | -0.000129 | -0.009448 | -0.009448 | -0.59922
2019-03-29 11:49 | 1553874542.8815 | 823 | 0.20951 | 0.000475 | -1e-06 | 92.1371 | 90.9724 | 125.5418 | 21.1359 | 0.002142 | 0.002 | -151.8731 | 0.2346 | 8.403 | 0.002509 | 0.004341 | 0.004137 | 6.6e-05 | 0.000937 | 0.29911 | 1.3655 | 2.5475 | 100.3 | 21.05 | 32.262 | 73.2412 | 82.0639 | 69.8186 | 0.008033 | 1.0294 | 0.024624 | 23.9213 | 0.11379 | 0.13956 | -5.5e-05 | -0.00399 | -0.00399 | -1.4237
2019-03-29 11:49 | 1553874543.8832 | 824 | 0.20951 | 0.000471 | -1e-06 | 92.0224 | 90.5993 | 125.658 | 21.1353 | 0.002126 | 0.001983 | -151.8735 | 0.23391 | 8.4033 | 0.002508 | 0.004521 | 0.004338 | 6.5e-05 | 0.000937 | 0.31509 | 1.3715 | 2.5475 | 100.3 | 21.05 | 32.252 | 72.8502 | 82.0213 | 69.4141 | 0.008031 | 1.0306 | 0.02465 | 23.9183 | 0.13941 | 0.13915 | -5.3e-05 | -0.003826 | -0.003826 | -1.4804
2019-03-29 11:49 | 1553874544.883 | 825 | 0.20948 | 0.000472 | -1e-06 | 91.8967 | 90.6741 | 125.569 | 21.1358 | 0.002058 | 0.001992 | -151.8641 | 0.23407 | 8.4035 | 0.002483 | 0.003955 | 0.003804 | 7e-05 | 0.000937 | 0.28102 | 1.3335 | 2.5475 | 100.3 | 21.05 | 32.243 | 71.2276 | 81.4415 | 69.6337 | 0.008029 | 1.0299 | 0.024633 | 23.9175 | 0.1198 | 0.13925 | 0.0001 | 0.007457 | 0.007457 | 0.76001
2019-03-29 11:49 | 1553874545.8822 | 826 | 0.20946 | 0.000473 | -1e-06 | 91.8503 | 90.529 | 125.6897 | 21.1356 | 0.002092 | 0.001965 | -151.8657 | 0.23369 | 8.4029 | 0.002471 | 0.003989 | 0.003789 | 6.9e-05 | 0.000938 | 0.30596 | 1.3532 | 2.5474 | 100.3 | 21.05 | 32.226 | 72.0311 | 81.1426 | 68.9911 | 0.008024 | 1.0319 | 0.024668 | 23.9065 | 0.1296 | 0.13902 | 0.000263 | 0.01959 | 0.01959 | 0.28883
2019-03-29 11:49 | 1553874546.8827 | 827 | 0.20951 | 0.000476 | -1e-06 | 91.6758 | 90.3843 | 125.8636 | 21.1359 | 0.002093 | 0.001957 | -151.8501 | 0.23418 | 8.4015 | 0.002507 | 0.003993 | 0.003852 | 6.9e-05 | 0.000932 | 0.30127 | 1.3718 | 2.5474 | 100.3 | 21.05 | 32.219 | 72.069 | 82.0118 | 68.783 | 0.008023 | 1.0325 | 0.024693 | 23.9158 | 0.1269 | 0.13931 | -5.1e-05 | -0.00372 | -0.00372 | -1.5243
2019-03-29 11:49 | 1553874547.8817 | 828 | 0.2095 | 0.000476 | -1e-06 | 91.586 | 90.0488 | 125.9927 | 21.1361 | 0.002126 | 0.001957 | -151.8391 | 0.23451 | 8.4033 | 0.00253 | 0.003875 | 0.003699 | 7e-05 | 0.000935 | 0.31293 | 1.3522 | 2.5475 | 100.3 | 21.06 | 32.211 | 72.8447 | 82.5632 | 68.7887 | 0.008026 | 1.0325 | 0.024705 | 23.9282 | 0.1514 | 0.13951 | 1.1e-05 | 0.000913 | 0.000913 | 6.2165
2019-03-29 11:49 | 1553874548.8827 | 829 | 0.20948 | 0.000479 | -1e-06 | 91.436 | 90.0639 | 126.0969 | 21.1364 | 0.002053 | 0.002014 | -151.8203 | 0.23252 | 8.4028 | 0.002483 | 0.003605 | 0.003412 | 7.2e-05 | 0.000929 | 0.30702 | 1.353 | 2.5475 | 100.3 | 21.06 | 32.208 | 71.1119 | 81.4273 | 70.1747 | 0.008025 | 1.0283 | 0.024666 | 23.9866 | 0.13525 | 0.13833 | 9.6e-05 | 0.007232 | 0.007232 | 0.77845
2019-03-29 11:49 | 1553874549.882 | 830 | 0.20948 | 0.000481 | -1e-06 | 91.3561 | 90.0325 | 126.3459 | 21.1361 | 0.002065 | 0.001966 | -151.8256 | 0.23282 | 8.4029 | 0.002475 | 0.003474 | 0.003292 | 6.7e-05 | 0.000929 | 0.31581 | 1.3829 | 2.5473 | 100.3 | 21.06 | 32.206 | 71.3807 | 81.2313 | 69.0148 | 0.008024 | 1.0318 | 0.024732 | 23.9697 | 0.13055 | 0.1385 | 7.6e-05 | 0.005728 | 0.005728 | 0.98408
2019-03-29 11:49 | 1553874550.8818 | 831 | 0.20949 | 0.000479 | -1e-06 | 91.1831 | 89.8465 | 126.08 | 21.1363 | 0.002085 | 0.002029 | -151.827 | 0.23323 | 8.4042 | 0.002423 | 0.003903 | 0.003694 | 6.9e-05 | 0.000927 | 0.29407 | 1.403 | 2.5475 | 100.3 | 21.06 | 32.205 | 71.8735 | 79.9971 | 70.5201 | 0.008024 | 1.0273 | 0.024652 | 23.9971 | 0.13209 | 0.13875 | 5e-05 | 0.003774 | 0.003774 | 1.4964
2019-03-29 11:49 | 1553874551.8812 | 832 | 0.20951 | 0.000482 | -1e-06 | 91.0403 | 89.7383 | 125.9824 | 21.1364 | 0.002081 | 0.001953 | -151.8302 | 0.23354 | 8.4048 | 0.002419 | 0.003049 | 0.002911 | 7e-05 | 0.00093 | 0.30646 | 1.3485 | 2.5473 | 100.3 | 21.06 | 32.206 | 71.7683 | 79.8912 | 68.7086 | 0.008024 | 1.0327 | 0.024707 | 23.9245 | 0.12885 | 0.13893 | -5.5e-05 | -0.00398 | -0.00398 | -1.4209
2019-03-29 11:49 | 1553874552.8809 | 833 | 0.20952 | 0.000484 | -1e-06 | 91.0074 | 89.5178 | 125.9751 | 21.1364 | 0.002114 | 0.001974 | -151.8389 | 0.23455 | 8.4039 | 0.002446 | 0.00426 | 0.004083 | 6.9e-05 | 0.00093 | 0.31293 | 1.3408 | 2.5474 | 100.3 | 21.06 | 32.206 | 72.5728 | 80.5548 | 69.2094 | 0.008024 | 1.0312 | 0.024689 | 23.9413 | 0.14762 | 0.13953 | -0.000117 | -0.008589 | -0.008589 | -0.6612
2019-03-29 11:49 | 1553874553.8818 | 834 | 0.20951 | 0.000484 | -1e-06 | 90.9553 | 89.6082 | 125.9873 | 21.1368 | 0.002104 | 0.001966 | -151.8261 | 0.23323 | 8.4016 | 0.002446 | 0.003192 | 0.003063 | 7e-05 | 0.000927 | 0.29247 | 1.3397 | 2.5473 | 100.3 | 21.06 | 32.209 | 72.3255 | 80.5396 | 69.0031 | 0.008025 | 1.0318 | 0.024697 | 23.9352 | 0.13346 | 0.13875 | -8.5e-05 | -0.006248 | -0.006248 | -0.90378
2019-03-29 11:49 | 1553874554.8823 | 835 | 0.20949 | 0.000487 | -1e-06 | 90.7814 | 89.3031 | 126.0643 | 21.1357 | 0.001997 | 0.00201 | -151.8275 | 0.23397 | 8.4013 | 0.002509 | 0.003351 | 0.00314 | 7.1e-05 | 0.000929 | 0.32532 | 1.3334 | 2.5472 | 100.3 | 21.06 | 32.212 | 69.7427 | 82.0551 | 70.0773 | 0.008026 | 1.0286 | 0.024666 | 23.9801 | 0.14685 | 0.13919 | 5.9e-05 | 0.004431 | 0.004431 | 1.2785
2019-03-29 11:49 | 1553874555.8812 | 836 | 0.20947 | 0.000479 | -1e-06 | 90.6635 | 89.1755 | 126.1992 | 21.1372 | 0.002027 | 0.002014 | -151.8265 | 0.23384 | 8.4018 | 0.002439 | 0.003942 | 0.003827 | 7.2e-05 | 0.000926 | 0.29574 | 1.3777 | 2.5474 | 100.3 | 21.06 | 32.223 | 70.4835 | 80.3864 | 70.1663 | 0.008029 | 1.0283 | 0.024676 | 23.996 | 0.14802 | 0.13911 | 0.000143 | 0.010726 | 0.010726 | 0.52787
2019-03-29 11:49 | 1553874556.881 | 837 | 0.20951 | 0.000484 | -1e-06 | 90.5446 | 89.1369 | 126.2265 | 21.1377 | 0.002047 | 0.001993 | -151.8289 | 0.23281 | 8.4023 | 0.002432 | 0.003998 | 0.003823 | 6.9e-05 | 0.000928 | 0.30093 | 1.4036 | 2.5475 | 100.3 | 21.06 | 32.23 | 70.9533 | 80.2122 | 69.6707 | 0.00803 | 1.0298 | 0.024697 | 23.9813 | 0.14016 | 0.1385 | -5.9e-05 | -0.004295 | -0.004295 | -1.3123
2019-03-29 11:49 | 1553874557.8823 | 838 | 0.20959 | 0.000478 | -1e-06 | 90.4698 | 88.7864 | 126.433 | 21.1374 | 0.002042 | 0.001974 | -151.8308 | 0.23338 | 8.4029 | 0.002455 | 0.003061 | 0.002876 | 6.8e-05 | 0.000919 | 0.31741 | 1.4089 | 2.5474 | 100.3 | 21.06 | 32.236 | 70.8307 | 80.7622 | 69.2104 | 0.008032 | 1.0312 | 0.024733 | 23.9848 | 0.168 | 0.13884 | -0.000581 | -0.043144 | -0.043144 | -0.13098
2019-03-29 11:49 | 1553874558.882 | 839 | 0.2095 | 0.00048 | -1e-06 | 90.31 | 88.8188 | 126.2213 | 21.1364 | 0.002103 | 0.001985 | -151.8441 | 0.23348 | 8.4016 | 0.002517 | 0.003426 | 0.003189 | 6.8e-05 | 0.000916 | 0.28191 | 1.3995 | 2.5474 | 100.3 | 21.07 | 32.25 | 72.297 | 82.2524 | 69.4745 | 0.00804 | 1.0304 | 0.024703 | 23.9739 | 0.14892 | 0.1389 | -4.8e-05 | -0.003512 | -0.003512 | -1.6096
2019-03-29 11:49 | 1553874559.8815 | 840 | 0.20947 | 0.000477 | -1e-06 | 90.2926 | 88.7236 | 126.3544 | 21.1367 | 0.002072 | 0.001976 | -151.8226 | 0.23316 | 8.4028 | 0.002489 | 0.003673 | 0.003495 | 6.8e-05 | 0.000918 | 0.29878 | 1.3751 | 2.5474 | 100.3 | 21.07 | 32.256 | 71.555 | 81.5811 | 69.2431 | 0.008042 | 1.0311 | 0.024724 | 23.9785 | 0.15679 | 0.13871 | 0.000162 | 0.012103 | 0.012103 | 0.46646
2019-03-29 11:49 | 1553874560.8812 | 841 | 0.20952 | 0.000476 | -1e-06 | 90.2483 | 88.5062 | 126.6393 | 21.1378 | 0.002054 | 0.001949 | -151.8218 | 0.23244 | 8.4008 | 0.002471 | 0.004064 | 0.00379 | 6.6e-05 | 0.00092 | 0.30558 | 1.3685 | 2.5473 | 100.3 | 21.07 | 32.266 | 71.1295 | 81.1409 | 68.6062 | 0.008044 | 1.033 | 0.024775 | 23.9832 | 0.17435 | 0.13828 | -0.000132 | -0.009798 | -0.009798 | -0.57438
2019-03-29 11:49 | 1553874561.8812 | 842 | 0.2095 | 0.000481 | -1e-06 | 90.2615 | 88.7216 | 126.5323 | 21.1385 | 0.001971 | 0.001955 | -151.8177 | 0.23312 | 8.4022 | 0.002451 | 0.003739 | 0.003629 | 7e-05 | 0.000921 | 0.30021 | 1.3851 | 2.5473 | 100.3 | 21.07 | 32.271 | 69.124 | 80.6612 | 68.7545 | 0.008046 | 1.0326 | 0.02476 | 23.9782 | 0.15391 | 0.13869 | -3.8e-05 | -0.002728 | -0.002728 | -2.0693
2019-03-29 11:49 | 1553874562.8816 | 843 | 0.20951 | 0.000482 | -1e-06 | 90.2067 | 88.5339 | 126.3861 | 21.1382 | 0.002044 | 0.001952 | -151.8138 | 0.23265 | 8.402 | 0.002469 | 0.004038 | 0.003845 | 6.8e-05 | 0.000926 | 0.32925 | 1.3818 | 2.5475 | 100.3 | 21.07 | 32.275 | 70.8923 | 81.103 | 68.6861 | 0.008047 | 1.0328 | 0.024748 | 23.962 | 0.16743 | 0.1384 | -5.6e-05 | -0.004118 | -0.004118 | -1.368
2019-03-29 11:49 | 1553874563.926 | 844 | 0.2095 | 0.000473 | -1e-06 | 90.1714 | 88.6683 | 126.4016 | 21.1383 | 0.002076 | 0.001977 | -151.8161 | 0.23395 | 8.4017 | 0.002531 | 0.003738 | 0.003548 | 6.6e-05 | 0.000918 | 0.29502 | 1.3424 | 2.5474 | 100.3 | 21.07 | 32.28 | 71.6503 | 82.5973 | 69.2814 | 0.008048 | 1.031 | 0.024728 | 23.9843 | 0.15035 | 0.13918 | -2e-06 | -8.7e-05 | -8.7e-05 | -64.955
2019-03-29 11:49 | 1553874564.8716 | 845 | 0.20952 | 0.000471 | -1e-06 | 90.1393 | 88.5618 | 126.3347 | 21.1385 | 0.00203 | 0.001955 | -151.8258 | 0.23343 | 8.4037 | 0.002446 | 0.003863 | 0.003724 | 6.8e-05 | 0.000927 | 0.31382 | 1.3129 | 2.5473 | 100.3 | 21.07 | 32.28 | 70.555 | 80.5475 | 68.7497 | 0.008048 | 1.0326 | 0.02474 | 23.9593 | 0.15792 | 0.13887 | -0.000113 | -0.008375 | -0.008375 | -0.67485
2019-03-29 11:49 | 1553874565.8698 | 846 | 0.20952 | 0.000471 | -1e-06 | 90.2743 | 88.7216 | 125.981 | 21.1379 | 0.002014 | 0.001969 | -151.8138 | 0.23362 | 8.4025 | 0.002447 | 0.003996 | 0.003802 | 6.9e-05 | 0.00093 | 0.30309 | 1.317 | 2.5473 | 100.3 | 21.07 | 32.281 | 70.1566 | 80.5584 | 69.0861 | 0.008048 | 1.0316 | 0.024694 | 23.9375 | 0.15519 | 0.13898 | -9.8e-05 | -0.00723 | -0.00723 | -0.78234
2019-03-29 11:49 | 1553874566.87 | 847 | 0.20951 | 0.000471 | -1e-06 | 90.31 | 88.899 | 125.9359 | 21.1378 | 0.002039 | 0.001941 | -151.8319 | 0.23305 | 8.4019 | 0.002438 | 0.003706 | 0.003502 | 6.5e-05 | 0.000926 | 0.29878 | 1.3601 | 2.5474 | 100.3 | 21.07 | 32.275 | 70.7609 | 80.3408 | 68.4059 | 0.008047 | 1.0336 | 0.024714 | 23.9094 | 0.14084 | 0.13864 | -6.7e-05 | -0.004867 | -0.004867 | -1.1595
2019-03-29 11:49 | 1553874567.8706 | 848 | 0.20951 | 0.000468 | -1e-06 | 90.4869 | 88.9809 | 125.4833 | 21.1378 | 0.002071 | 0.001968 | -151.8132 | 0.23448 | 8.401 | 0.002446 | 0.004061 | 0.003906 | 6.7e-05 | 0.000921 | 0.26808 | 1.3574 | 2.5474 | 100.3 | 21.07 | 32.271 | 71.5373 | 80.553 | 69.0477 | 0.008046 | 1.0317 | 0.024646 | 23.8889 | 0.15011 | 0.13949 | -4.1e-05 | -0.002983 | -0.002983 | -1.903
2019-03-29 11:49 | 1553874568.8697 | 849 | 0.20951 | 0.000464 | -1e-06 | 90.591 | 89.128 | 125.4568 | 21.1383 | 0.002069 | 0.001964 | -151.8205 | 0.23379 | 8.4016 | 0.002521 | 0.00368 | 0.003443 | 6.4e-05 | 0.000926 | 0.30431 | 1.3582 | 2.5473 | 100.3 | 21.07 | 32.26 | 71.4771 | 82.3548 | 68.9514 | 0.008043 | 1.032 | 0.024647 | 23.883 | 0.14563 | 0.13908 | -7.7e-05 | -0.005652 | -0.005652 | -1.0015
2019-03-29 11:49 | 1553874569.8707 | 850 | 0.20952 | 0.000472 | -1e-06 | 90.7338 | 89.1608 | 125.2878 | 21.139 | 0.002092 | 0.001965 | -151.8268 | 0.23393 | 8.4031 | 0.002512 | 0.003982 | 0.003776 | 6.5e-05 | 0.000922 | 0.28727 | 1.3621 | 2.5474 | 100.3 | 21.07 | 32.252 | 72.0383 | 82.1292 | 68.9826 | 0.008041 | 1.0319 | 0.024629 | 23.868 | 0.15643 | 0.13917 | -0.000117 | -0.008582 | -0.008582 | -0.65998
2019-03-29 11:49 | 1553874570.8697 | 851 | 0.2095 | 0.000466 | -1e-06 | 90.8452 | 89.3171 | 125.3798 | 21.1385 | 0.002066 | 0.001946 | -151.8295 | 0.23291 | 8.4032 | 0.00251 | 0.004512 | 0.004329 | 6.6e-05 | 0.000928 | 0.28368 | 1.3283 | 2.5473 | 100.3 | 21.07 | 32.244 | 71.4157 | 82.0759 | 68.5273 | 0.008039 | 1.0333 | 0.024655 | 23.8608 | 0.15173 | 0.13856 | -2.8e-05 | -0.001994 | -0.001994 | -2.8274
2019-03-29 11:50 | 1553874571.8707 | 852 | 0.20952 | 0.000471 | -1e-06 | 91.0026 | 89.6304 | 125.3815 | 21.1378 | 0.002032 | 0.001986 | -151.8315 | 0.23381 | 8.4019 | 0.002507 | 0.003748 | 0.003563 | 6.8e-05 | 0.000937 | 0.28888 | 1.4003 | 2.5475 | 100.3 | 21.08 | 32.223 | 70.5932 | 82.0057 | 69.486 | 0.008039 | 1.0304 | 0.02462 | 23.8944 | 0.1359 | 0.13909 | -9.6e-05 | -0.007052 | -0.007052 | -0.80275
2019-03-29 11:50 | 1553874572.8699 | 853 | 0.2095 | 0.00047 | -1e-06 | 91.067 | 89.5816 | 125.8401 | 21.1383 | 0.001994 | 0.001934 | -151.8401 | 0.2337 | 8.4044 | 0.002503 | 0.004062 | 0.003915 | 6.7e-05 | 0.000943 | 0.30397 | 1.3517 | 2.5474 | 100.3 | 21.08 | 32.213 | 69.682 | 81.9071 | 68.2299 | 0.008036 | 1.0342 | 0.024711 | 23.8942 | 0.14709 | 0.13903 | -2.8e-05 | -0.002045 | -0.002045 | -2.7664
2019-03-29 11:50 | 1553874573.8698 | 854 | 0.2095 | 0.000468 | -1e-06 | 91.2035 | 89.7035 | 126.0258 | 21.1388 | 0.002032 | 0.001945 | -151.8539 | 0.23363 | 8.402 | 0.002562 | 0.003537 | 0.003314 | 6.7e-05 | 0.000949 | 0.29789 | 1.3956 | 2.5474 | 100.3 | 21.08 | 32.192 | 70.6074 | 83.3297 | 68.5067 | 0.008031 | 1.0333 | 0.024719 | 23.9215 | 0.14833 | 0.13899 | -2.5e-05 | -0.001757 | -0.001757 | -3.2198
2019-03-29 11:50 | 1553874574.8698 | 855 | 0.20951 | 0.000464 | -1e-06 | 91.3711 | 89.8929 | 125.9183 | 21.1386 | 0.002035 | 0.001941 | -151.8613 | 0.23401 | 8.4021 | 0.002545 | 0.003798 | 0.003714 | 6.5e-05 | 0.000949 | 0.30901 | 1.4022 | 2.5474 | 100.3 | 21.08 | 32.18 | 70.6603 | 82.9277 | 68.4019 | 0.008028 | 1.0337 | 0.024712 | 23.9076 | 0.14589 | 0.13922 | -5.7e-05 | -0.004191 | -0.004191 | -1.3521
2019-03-29 11:50 | 1553874575.8703 | 856 | 0.20952 | 0.000462 | -1e-06 | 91.404 | 90.0256 | 126.0002 | 21.1396 | 0.002036 | 0.001919 | -151.8686 | 0.23402 | 8.4026 | 0.002488 | 0.003418 | 0.00317 | 6.5e-05 | 0.000946 | 0.3095 | 1.3627 | 2.5473 | 100.3 | 21.08 | 32.17 | 70.6846 | 81.5561 | 67.8935 | 0.008025 | 1.0352 | 0.024739 | 23.8976 | 0.13591 | 0.13922 | -0.000136 | -0.010093 | -0.010093 | -0.56142
2019-03-29 11:50 | 1553874576.871 | 857 | 0.20952 | 0.000469 | -1e-06 | 91.5181 | 90.1324 | 126.2072 | 21.1388 | 0.002021 | 0.00192 | -151.8566 | 0.23485 | 8.402 | 0.002537 | 0.004355 | 0.00425 | 6.4e-05 | 0.000951 | 0.29015 | 1.3289 | 2.5472 | 100.3 | 21.08 | 32.148 | 70.3423 | 82.7368 | 67.9177 | 0.00802 | 1.0351 | 0.024758 | 23.9181 | 0.13646 | 0.13971 | -0.000104 | -0.007642 | -0.007642 | -0.74408
2019-03-29 11:50 | 1553874577.871 | 858 | 0.20951 | 0.00047 | -1e-06 | 91.6082 | 90.2488 | 126.3689 | 21.139 | 0.00197 | 0.001913 | -151.8723 | 0.23341 | 8.4037 | 0.002488 | 0.003651 | 0.003473 | 6.5e-05 | 0.000955 | 0.30973 | 1.2957 | 2.5473 | 100.3 | 21.08 | 32.137 | 69.1105 | 81.5646 | 67.7479 | 0.008017 | 1.0356 | 0.02478 | 23.9274 | 0.13372 | 0.13886 | -4.5e-05 | -0.003308 | -0.003308 | -1.7083
2019-03-29 11:50 | 1553874578.8704 | 859 | 0.20949 | 0.000466 | -1e-06 | 91.7264 | 90.3638 | 126.5718 | 21.1389 | 0.001985 | 0.00198 | -151.8739 | 0.23364 | 8.4011 | 0.002553 | 0.003983 | 0.003808 | 6.1e-05 | 0.000955 | 0.28711 | 1.3582 | 2.5474 | 100.3 | 21.08 | 32.115 | 69.4651 | 83.116 | 69.3505 | 0.008012 | 1.0308 | 0.024742 | 24.0029 | 0.13386 | 0.13899 | 4.5e-05 | 0.00339 | 0.00339 | 1.669
2019-03-29 11:50 | 1553874579.87 | 860 | 0.20954 | 0.000469 | -1e-06 | 91.8641 | 90.6925 | 126.357 | 21.1392 | 0.002014 | 0.001926 | -151.8848 | 0.23431 | 8.4017 | 0.002574 | 0.003376 | 0.003236 | 6.5e-05 | 0.000949 | 0.29839 | 1.3253 | 2.5473 | 100.3 | 21.08 | 32.103 | 70.1572 | 83.6241 | 68.0438 | 0.008009 | 1.0347 | 0.024768 | 23.9367 | 0.11481 | 0.13939 | -0.000244 | -0.018068 | -0.018068 | -0.31399
2019-03-29 11:50 | 1553874580.871 | 861 | 0.20945 | 0.000475 | -1e-06 | 92.0126 | 90.684 | 126.157 | 21.139 | 0.002048 | 0.001971 | -151.884 | 0.23421 | 8.404 | 0.002464 | 0.003362 | 0.003126 | 6.3e-05 | 0.000947 | 0.31708 | 1.345 | 2.5473 | 100.3 | 21.08 | 32.091 | 70.9803 | 80.9673 | 69.1206 | 0.008006 | 1.0315 | 0.02471 | 23.9554 | 0.13009 | 0.13933 | 0.000276 | 0.020567 | 0.020567 | 0.27573
2019-03-29 11:50 | 1553874581.8704 | 862 | 0.20952 | 0.000472 | -2e-06 | 92.2032 | 90.7311 | 126.4926 | 21.1393 | 0.002032 | 0.001967 | -151.871 | 0.23422 | 8.4001 | 0.002493 | 0.004158 | 0.003964 | 6.5e-05 | 0.000939 | 0.29999 | 1.3613 | 2.5474 | 100.3 | 21.08 | 32.07 | 70.5908 | 81.6701 | 69.035 | 0.008 | 1.0317 | 0.024746 | 23.9843 | 0.14397 | 0.13933 | -9.4e-05 | -0.006914 | -0.006914 | -0.82017
2019-03-29 11:50 | 1553874582.8701 | 863 | 0.20953 | 0.000476 | -1e-06 | 92.3017 | 91.044 | 126.0049 | 21.1406 | 0.002021 | 0.001922 | -151.8532 | 0.23368 | 8.4016 | 0.00248 | 0.003719 | 0.003672 | 6.4e-05 | 0.000936 | 0.28512 | 1.3598 | 2.5474 | 100.3 | 21.09 | 32.061 | 70.3284 | 81.3674 | 67.9461 | 0.008003 | 1.035 | 0.024737 | 23.8999 | 0.12271 | 0.13902 | -0.000185 | -0.013706 | -0.013706 | -0.4128
2019-03-29 11:50 | 1553874583.87 | 864 | 0.20949 | 0.000472 | -1e-06 | 92.2876 | 91.0447 | 125.8404 | 21.1402 | 0.002041 | 0.001933 | -151.8546 | 0.23442 | 8.403 | 0.002551 | 0.004265 | 0.004031 | 6e-05 | 0.00094 | 0.30613 | 1.3447 | 2.5474 | 100.3 | 21.09 | 32.053 | 70.8178 | 83.0689 | 68.2116 | 0.008001 | 1.0342 | 0.024711 | 23.8936 | 0.12128 | 0.13945 | 5.4e-05 | 0.004101 | 0.004101 | 1.3839
2019-03-29 11:50 | 1553874584.8691 | 865 | 0.2095 | 0.000474 | -1e-06 | 92.2867 | 90.9874 | 126.4239 | 21.1403 | 0.002073 | 0.001934 | -151.8526 | 0.23445 | 8.4003 | 0.002517 | 0.003754 | 0.003611 | 6.4e-05 | 0.000939 | 0.28766 | 1.3541 | 2.5473 | 100.3 | 21.09 | 32.036 | 71.5828 | 82.2562 | 68.243 | 0.007997 | 1.0341 | 0.024767 | 23.95 | 0.12683 | 0.13947 | 2e-06 | 0.000185 | 0.000185 | 30.6915
2019-03-29 11:50 | 1553874585.8692 | 866 | 0.20949 | 0.000472 | -1e-06 | 92.2972 | 90.9423 | 126.6823 | 21.1409 | 0.002004 | 0.001957 | -151.8653 | 0.23497 | 8.4012 | 0.002537 | 0.003729 | 0.003558 | 6.5e-05 | 0.000935 | 0.28241 | 1.4051 | 2.5475 | 100.3 | 21.09 | 32.029 | 69.9305 | 82.745 | 68.7996 | 0.007995 | 1.0324 | 0.024773 | 23.994 | 0.13228 | 0.13978 | 6.5e-05 | 0.004954 | 0.004954 | 1.1483
2019-03-29 11:50 | 1553874586.8693 | 867 | 0.20951 | 0.00048 | -1e-06 | 92.3011 | 91.2218 | 126.6532 | 21.1411 | 0.001943 | 0.001924 | -151.8661 | 0.23453 | 8.4021 | 0.002541 | 0.004053 | 0.003742 | 6.3e-05 | 0.000943 | 0.30381 | 1.3827 | 2.5474 | 100.3 | 21.09 | 32.016 | 68.4606 | 82.8349 | 67.9995 | 0.007992 | 1.0349 | 0.024799 | 23.9632 | 0.1052 | 0.13952 | -8.4e-05 | -0.006205 | -0.006205 | -0.91522
2019-03-29 11:50 | 1553874587.8696 | 868 | 0.20943 | 0.000477 | -1e-06 | 92.3475 | 91.1775 | 126.7098 | 21.1413 | 0.00199 | 0.001941 | -151.8664 | 0.23371 | 8.4024 | 0.002551 | 0.004168 | 0.00399 | 6.7e-05 | 0.000946 | 0.32892 | 1.3409 | 2.5473 | 100.3 | 21.09 | 32.011 | 69.5994 | 83.0606 | 68.4045 | 0.007991 | 1.0336 | 0.02479 | 23.9827 | 0.11405 | 0.13903 | 0.00041 | 0.030607 | 0.030607 | 0.18488
2019-03-29 11:50 | 1553874588.8693 | 869 | 0.20949 | 0.000475 | -1e-06 | 92.3319 | 91.0727 | 126.6955 | 21.1418 | 0.001974 | 0.001904 | -151.8524 | 0.23409 | 8.4018 | 0.002503 | 0.00355 | 0.003365 | 6.5e-05 | 0.000948 | 0.27455 | 1.3395 | 2.5472 | 100.3 | 21.09 | 32.006 | 69.1997 | 81.9286 | 67.5338 | 0.007989 | 1.0363 | 0.02482 | 23.9508 | 0.12283 | 0.13926 | 5e-05 | 0.003772 | 0.003772 | 1.5026
2019-03-29 11:50 | 1553874589.8704 | 870 | 0.2095 | 0.000476 | -1e-06 | 92.3906 | 91.1471 | 126.6368 | 21.1415 | 0.002052 | 0.001929 | -151.8433 | 0.23412 | 8.4035 | 0.002481 | 0.004029 | 0.003901 | 6.7e-05 | 0.000949 | 0.2855 | 1.349 | 2.5473 | 100.3 | 21.09 | 31.998 | 71.0748 | 81.3988 | 68.1232 | 0.007987 | 1.0345 | 0.024793 | 23.966 | 0.1212 | 0.13928 | 3e-06 | 0.000279 | 0.000279 | 20.3498
2019-03-29 11:50 | 1553874590.8704 | 871 | 0.20945 | 0.000476 | -1e-06 | 92.4074 | 91.043 | 126.6971 | 21.1412 | 0.002027 | 0.001953 | -151.8406 | 0.23416 | 8.4028 | 0.002533 | 0.00358 | 0.003392 | 6.5e-05 | 0.000951 | 0.30541 | 1.3313 | 2.5473 | 100.3 | 21.09 | 31.996 | 70.4756 | 82.6384 | 68.7117 | 0.007987 | 1.0327 | 0.024777 | 23.9923 | 0.13305 | 0.1393 | 0.000316 | 0.023608 | 0.023608 | 0.24015
2019-03-29 11:50 | 1553874591.8698 | 872 | 0.2095 | 0.000476 | -1e-06 | 92.3286 | 90.9928 | 126.652 | 21.1407 | 0.001997 | 0.00189 | -151.8408 | 0.23378 | 8.4024 | 0.002499 | 0.003336 | 0.003221 | 6.6e-05 | 0.000954 | 0.30829 | 1.3434 | 2.5473 | 100.3 | 21.09 | 31.992 | 69.7695 | 81.8168 | 67.1949 | 0.007986 | 1.0373 | 0.024828 | 23.9348 | 0.13035 | 0.13908 | -2e-06 | -0.000103 | -0.000103 | -55.0217
2019-03-29 11:50 | 1553874592.8704 | 873 | 0.2095 | 0.000479 | -1e-06 | 92.2655 | 90.8922 | 127.1628 | 21.1415 | 0.002031 | 0.001884 | -151.8401 | 0.23384 | 8.4035 | 0.002501 | 0.00401 | 0.00377 | 6.7e-05 | 0.000956 | 0.31238 | 1.3659 | 2.5474 | 100.3 | 21.09 | 31.992 | 70.5867 | 81.8627 | 67.0333 | 0.007986 | 1.0378 | 0.024884 | 23.9773 | 0.13414 | 0.13911 | -1.9e-05 | -0.001349 | -0.001349 | -4.1963
2019-03-29 11:50 | 1553874593.8705 | 874 | 0.20953 | 0.000477 | -1e-06 | 92.1583 | 91.0276 | 127.1388 | 21.142 | 0.001991 | 0.00195 | -151.8349 | 0.23429 | 8.4025 | 0.002525 | 0.004033 | 0.00392 | 6.5e-05 | 0.000959 | 0.30934 | 1.351 | 2.5472 | 100.3 | 21.09 | 31.992 | 69.6126 | 82.4365 | 68.6182 | 0.007986 | 1.033 | 0.024824 | 24.0308 | 0.11042 | 0.13938 | -0.000171 | -0.012683 | -0.012683 | -0.44729
2019-03-29 11:50 | 1553874594.8696 | 875 | 0.20948 | 0.00047 | -1e-06 | 92.1044 | 90.7812 | 126.5541 | 21.1424 | 0.002009 | 0.001924 | -151.8468 | 0.23394 | 8.4027 | 0.002509 | 0.003389 | 0.003174 | 6.6e-05 | 0.000955 | 0.2808 | 1.3944 | 2.5474 | 100.3 | 21.09 | 31.994 | 70.0574 | 82.0655 | 68.0063 | 0.007986 | 1.0348 | 0.024789 | 23.954 | 0.12943 | 0.13917 | 0.000109 | 0.008236 | 0.008236 | 0.68771
2019-03-29 11:50 | 1553874595.8696 | 876 | 0.20948 | 0.000476 | -1e-06 | 92.0808 | 90.5608 | 126.2342 | 21.1425 | 0.002016 | 0.001833 | -151.8296 | 0.23388 | 8.4031 | 0.002498 | 0.003715 | 0.003537 | 6.5e-05 | 0.000957 | 0.30917 | 1.4137 | 2.5475 | 100.3 | 21.09 | 31.997 | 70.218 | 81.8021 | 65.8279 | 0.007987 | 1.0415 | 0.024837 | 23.8472 | 0.14889 | 0.13913 | 0.00011 | 0.008308 | 0.008308 | 0.6816
2019-03-29 11:50 | 1553874596.8696 | 877 | 0.20948 | 0.000473 | -1e-06 | 91.9458 | 90.7273 | 126.2584 | 21.1421 | 0.002015 | 0.001903 | -151.8386 | 0.23398 | 8.4023 | 0.002508 | 0.004 | 0.003788 | 6.2e-05 | 0.000962 | 0.31116 | 1.4302 | 2.5473 | 100.3 | 21.09 | 32.001 | 70.1856 | 82.0387 | 67.4996 | 0.007988 | 1.0364 | 0.024778 | 23.9083 | 0.11933 | 0.1392 | 9.8e-05 | 0.007398 | 0.007398 | 0.76581
2019-03-29 11:50 | 1553874597.869 | 878 | 0.20951 | 0.000474 | -1e-06 | 91.9428 | 90.6198 | 126.2673 | 21.1421 | 0.001996 | 0.0019 | -151.8437 | 0.23416 | 8.4028 | 0.002579 | 0.003866 | 0.003694 | 6.2e-05 | 0.000965 | 0.29231 | 1.416 | 2.5475 | 100.3 | 21.09 | 32.005 | 69.7312 | 83.7392 | 67.4287 | 0.007989 | 1.0366 | 0.024782 | 23.9066 | 0.12964 | 0.1393 | -4.9e-05 | -0.003589 | -0.003589 | -1.5796
2019-03-29 11:50 | 1553874598.8686 | 879 | 0.20972 | 0.000469 | -1e-06 | 91.8374 | 90.3857 | 126.2551 | 21.1427 | 0.002029 | 0.00191 | -151.8473 | 0.23381 | 8.4034 | 0.002549 | 0.003959 | 0.003791 | 6.1e-05 | 0.000969 | 0.30934 | 1.3642 | 2.5473 | 100.3 | 21.09 | 32.009 | 70.5357 | 83.0158 | 67.6625 | 0.00799 | 1.0359 | 0.024772 | 23.9137 | 0.14252 | 0.13909 | -0.001322 | -0.098412 | -0.098412 | -0.057524
2019-03-29 11:50 | 1553874599.8697 | 880 | 0.20946 | 0.000468 | -1e-06 | 91.7569 | 90.2853 | 126.2082 | 21.1417 | 0.002 | 0.001909 | -151.8521 | 0.23345 | 8.4043 | 0.002511 | 0.004396 | 0.004247 | 6.2e-05 | 0.000966 | 0.30182 | 1.3319 | 2.5473 | 100.3 | 21.09 | 32.013 | 69.833 | 82.1079 | 67.6414 | 0.007991 | 1.036 | 0.024768 | 23.9085 | 0.14461 | 0.13888 | 0.000251 | 0.018804 | 0.018804 | 0.30061
2019-03-29 11:50 | 1553874600.8701 | 881 | 0.20942 | 0.000467 | -1e-06 | 91.6471 | 90.1659 | 126.3424 | 21.143 | 0.001983 | 0.001872 | -151.8562 | 0.23409 | 8.4033 | 0.002429 | 0.003615 | 0.003525 | 6.1e-05 | 0.000965 | 0.29911 | 1.3162 | 2.5473 | 100.3 | 21.09 | 32.022 | 69.4177 | 80.1527 | 66.7501 | 0.007993 | 1.0387 | 0.024814 | 23.8899 | 0.14574 | 0.13926 | 0.000477 | 0.03568 | 0.03568 | 0.15885
2019-03-29 11:50 | 1553874601.8691 | 882 | 0.20938 | 0.000468 | -1e-06 | 91.5683 | 90.1021 | 126.3941 | 21.1431 | 0.002005 | 0.001925 | -151.8596 | 0.23423 | 8.4004 | 0.002498 | 0.003969 | 0.003746 | 6.3e-05 | 0.000968 | 0.29767 | 1.3402 | 2.5474 | 100.3 | 21.09 | 32.026 | 69.9528 | 81.8082 | 68.0289 | 0.007994 | 1.0348 | 0.024772 | 23.9397 | 0.14439 | 0.13934 | 0.000706 | 0.052642 | 0.052642 | 0.10773
2019-03-29 11:50 | 1553874602.8697 | 883 | 0.2095 | 0.000467 | -1e-06 | 91.4726 | 90.0789 | 126.284 | 21.1428 | 0.001996 | 0.001919 | -151.861 | 0.23505 | 8.4026 | 0.002468 | 0.003881 | 0.003667 | 6e-05 | 0.000972 | 0.28783 | 1.3758 | 2.5475 | 100.3 | 21.09 | 32.03 | 69.7364 | 81.0707 | 67.8819 | 0.007995 | 1.0352 | 0.024767 | 23.9241 | 0.13733 | 0.13983 | -1.1e-05 | -0.000773 | -0.000773 | -7.3642
2019-03-29 11:50 | 1553874603.8698 | 884 | 0.20947 | 0.000465 | -1e-06 | 91.3929 | 90.0683 | 126.4699 | 21.1436 | 0.002015 | 0.001869 | -151.8699 | 0.2337 | 8.4028 | 0.00248 | 0.003429 | 0.003253 | 5.8e-05 | 0.000974 | 0.31852 | 1.3907 | 2.5472 | 100.3 | 21.09 | 32.039 | 70.1972 | 81.3571 | 66.6835 | 0.007998 | 1.0389 | 0.024829 | 23.8996 | 0.13059 | 0.13903 | 0.000187 | 0.01404 | 0.01404 | 0.40304
2019-03-29 11:50 | 1553874604.8689 | 885 | 0.20956 | 0.000465 | -1e-06 | 91.2924 | 89.8444 | 126.9504 | 21.1434 | 0.002018 | 0.001896 | -151.8502 | 0.2332 | 8.4013 | 0.002448 | 0.003698 | 0.003508 | 5.8e-05 | 0.000973 | 0.31852 | 1.3367 | 2.5473 | 100.3 | 21.09 | 32.042 | 70.2751 | 80.6075 | 67.3239 | 0.007998 | 1.0369 | 0.024852 | 23.9675 | 0.143 | 0.13873 | -0.000356 | -0.026557 | -0.026557 | -0.21262
2019-03-29 11:50 | 1553874605.8699 | 886 | 0.20949 | 0.000462 | -2e-06 | 91.1547 | 89.7976 | 127.0877 | 21.1436 | 0.002 | 0.001939 | -151.8615 | 0.23453 | 8.4016 | 0.002444 | 0.004785 | 0.004547 | 5.8e-05 | 0.000972 | 0.32444 | 1.3322 | 2.5474 | 100.3 | 21.09 | 32.046 | 69.8425 | 80.4947 | 68.3623 | 0.007999 | 1.0338 | 0.024828 | 24.017 | 0.13416 | 0.13952 | 6.7e-05 | 0.005113 | 0.005113 | 1.1106
2019-03-29 11:50 | 1553874606.8705 | 887 | 0.2095 | 0.000464 | -1e-06 | 90.9748 | 89.4932 | 126.8572 | 21.1447 | 0.002017 | 0.0019 | -151.8648 | 0.23448 | 8.4048 | 0.00257 | 0.003774 | 0.003575 | 5.8e-05 | 0.000972 | 0.29015 | 1.3601 | 2.5473 | 100.3 | 21.09 | 32.055 | 70.249 | 83.5405 | 67.4267 | 0.008002 | 1.0366 | 0.02484 | 23.9623 | 0.14686 | 0.13949 | 0 | 4.3e-05 | 4.3e-05 | 132.7347
2019-03-29 11:50 | 1553874607.8703 | 888 | 0.20951 | 0.000463 | -1e-06 | 90.8811 | 89.5325 | 126.7391 | 21.1437 | 0.001958 | 0.00188 | -151.863 | 0.23436 | 8.4045 | 0.00255 | 0.004312 | 0.004107 | 6e-05 | 0.000978 | 0.2975 | 1.3902 | 2.5474 | 100.3 | 21.09 | 32.059 | 68.8141 | 83.0435 | 66.9402 | 0.008003 | 1.0381 | 0.024846 | 23.9341 | 0.13373 | 0.13942 | -6.6e-05 | -0.004866 | -0.004866 | -1.1661
2019-03-29 11:50 | 1553874608.8689 | 889 | 0.20953 | 0.000461 | -1e-06 | 90.7745 | 89.3106 | 126.6617 | 21.1443 | 0.001956 | 0.001876 | -151.8607 | 0.23439 | 8.4025 | 0.002486 | 0.004148 | 0.004038 | 5.6e-05 | 0.00098 | 0.29911 | 1.3638 | 2.5474 | 100.3 | 21.09 | 32.065 | 68.7678 | 81.5082 | 66.8503 | 0.008004 | 1.0384 | 0.024841 | 23.9236 | 0.14542 | 0.13944 | -0.00015 | -0.011128 | -0.011128 | -0.50998
2019-03-29 11:50 | 1553874609.869 | 890 | 0.20952 | 0.000466 | -2e-06 | 90.7449 | 89.1448 | 126.8675 | 21.1453 | 0.002029 | 0.001909 | -151.8608 | 0.23389 | 8.403 | 0.002376 | 0.003497 | 0.003409 | 6e-05 | 0.000985 | 0.29574 | 1.3361 | 2.5472 | 100.3 | 21.09 | 32.068 | 70.5233 | 78.8721 | 67.6564 | 0.008005 | 1.0359 | 0.024832 | 23.9714 | 0.15913 | 0.13914 | -0.000122 | -0.009013 | -0.009013 | -0.62834
2019-03-29 11:50 | 1553874610.8704 | 891 | 0.20952 | 0.000461 | -1e-06 | 90.5943 | 88.9693 | 126.8483 | 21.1449 | 0.002028 | 0.001886 | -151.8684 | 0.23407 | 8.4013 | 0.002442 | 0.003809 | 0.003588 | 6e-05 | 0.000986 | 0.31222 | 1.3148 | 2.5472 | 100.3 | 21.09 | 32.07 | 70.5059 | 80.464 | 67.1047 | 0.008005 | 1.0376 | 0.02485 | 23.9501 | 0.16189 | 0.13925 | -7.8e-05 | -0.005737 | -0.005737 | -0.98786
2019-03-29 11:50 | 1553874611.87 | 892 | 0.20949 | 0.00046 | -1e-06 | 90.5255 | 88.9437 | 127.079 | 21.1448 | 0.001966 | 0.00185 | -151.8719 | 0.23406 | 8.4009 | 0.002432 | 0.004419 | 0.004149 | 6e-05 | 0.000979 | 0.30453 | 1.3008 | 2.5474 | 100.3 | 21.1 | 32.071 | 69.0084 | 80.2253 | 66.2313 | 0.00801 | 1.0403 | 0.024905 | 23.9411 | 0.15767 | 0.13924 | 7.1e-05 | 0.005371 | 0.005371 | 1.0551
2019-03-29 11:50 | 1553874612.868 | 893 | 0.20952 | 0.000462 | -2e-06 | 90.3797 | 88.7673 | 126.9567 | 21.145 | 0.001958 | 0.001913 | -151.8848 | 0.23452 | 8.4029 | 0.002533 | 0.003648 | 0.003509 | 5.9e-05 | 0.000975 | 0.28838 | 1.3441 | 2.5473 | 100.3 | 21.09 | 32.072 | 68.8363 | 82.65 | 67.746 | 0.008006 | 1.0356 | 0.024838 | 23.983 | 0.16101 | 0.13951 | -0.000139 | -0.010289 | -0.010289 | -0.55186
2019-03-29 11:50 | 1553874613.8682 | 894 | 0.20951 | 0.000465 | -2e-06 | 90.2489 | 88.6209 | 126.9269 | 21.1451 | 0.001984 | 0.001918 | -151.8827 | 0.23439 | 8.4019 | 0.002506 | 0.003702 | 0.00351 | 5.9e-05 | 0.00097 | 0.28495 | 1.3473 | 2.5473 | 100.3 | 21.1 | 32.073 | 69.4618 | 81.9832 | 67.8637 | 0.008011 | 1.0353 | 0.02483 | 23.9843 | 0.16282 | 0.13944 | -4.5e-05 | -0.003212 | -0.003212 | -1.7667
2019-03-29 11:50 | 1553874614.8681 | 895 | 0.20952 | 0.000465 | -1e-06 | 90.2441 | 88.5076 | 126.5973 | 21.1454 | 0.001966 | 0.001902 | -151.8696 | 0.23459 | 8.4009 | 0.002513 | 0.003239 | 0.003077 | 5.7e-05 | 0.00097 | 0.2907 | 1.3531 | 2.5473 | 100.3 | 21.09 | 32.073 | 69.0285 | 82.1598 | 67.4709 | 0.008006 | 1.0365 | 0.024813 | 23.9393 | 0.17379 | 0.13956 | -0.000139 | -0.010294 | -0.010294 | -0.5518
2019-03-29 11:50 | 1553874615.8682 | 896 | 0.20953 | 0.000459 | -1e-06 | 90.1962 | 88.426 | 126.4227 | 21.1458 | 0.002014 | 0.001889 | -151.8641 | 0.23446 | 8.4032 | 0.002477 | 0.003472 | 0.00323 | 5.6e-05 | 0.000973 | 0.3178 | 1.38 | 2.5474 | 100.3 | 21.09 | 32.072 | 70.1642 | 81.2904 | 67.1601 | 0.008006 | 1.0374 | 0.024807 | 23.9119 | 0.17729 | 0.13948 | -0.000187 | -0.01387 | -0.01387 | -0.40928
2019-03-29 11:50 | 1553874616.8686 | 897 | 0.2095 | 0.000464 | -1e-06 | 90.2651 | 88.5451 | 126.5847 | 21.1457 | 0.002004 | 0.001895 | -151.8725 | 0.23407 | 8.4036 | 0.002468 | 0.003482 | 0.003265 | 5.7e-05 | 0.000971 | 0.30663 | 1.3479 | 2.5474 | 100.3 | 21.09 | 32.071 | 69.9429 | 81.0799 | 67.3082 | 0.008006 | 1.037 | 0.024817 | 23.9324 | 0.17208 | 0.13925 | -1.8e-05 | -0.001302 | -0.001302 | -4.3523
2019-03-29 11:50 | 1553874617.8693 | 898 | 0.20951 | 0.00046 | -1e-06 | 90.3303 | 88.5772 | 126.8468 | 21.1462 | 0.002003 | 0.001883 | -151.8665 | 0.23462 | 8.4001 | 0.002504 | 0.003518 | 0.003431 | 5.9e-05 | 0.000965 | 0.29695 | 1.3715 | 2.5474 | 100.3 | 21.09 | 32.07 | 69.909 | 81.9499 | 67.0239 | 0.008005 | 1.0378 | 0.024853 | 23.9472 | 0.1753 | 0.13958 | -6.9e-05 | -0.005044 | -0.005044 | -1.1262
2019-03-29 11:50 | 1553874618.8693 | 899 | 0.2095 | 0.000466 | -1e-06 | 90.2585 | 88.6577 | 127.2344 | 21.1466 | 0.001928 | 0.001874 | -151.8863 | 0.23416 | 8.4033 | 0.002536 | 0.003767 | 0.003499 | 5.6e-05 | 0.000958 | 0.30934 | 1.3609 | 2.5473 | 100.3 | 21.09 | 32.069 | 68.1087 | 82.7076 | 66.8113 | 0.008005 | 1.0385 | 0.024899 | 23.9762 | 0.16006 | 0.1393 | 4e-06 | 0.000351 | 0.000351 | 16.141
2019-03-29 11:50 | 1553874619.8693 | 900 | 0.20953 | 0.000463 | -1e-06 | 90.1803 | 88.5666 | 127.4456 | 21.1461 | 0.001946 | 0.001909 | -151.8723 | 0.23372 | 8.4021 | 0.002533 | 0.00364 | 0.003434 | 5.9e-05 | 0.000952 | 0.30309 | 1.344 | 2.5474 | 100.3 | 21.09 | 32.067 | 68.5426 | 82.6361 | 67.6471 | 0.008005 | 1.0359 | 0.024889 | 24.0256 | 0.16151 | 0.13904 | -0.000152 | -0.011255 | -0.011255 | -0.50281
2019-03-29 11:50 | 1553874620.8693 | 901 | 0.20947 | 0.000465 | -2e-06 | 90.2279 | 88.6588 | 127.2377 | 21.1469 | 0.001952 | 0.001851 | -151.8561 | 0.23449 | 8.4003 | 0.002558 | 0.003507 | 0.003389 | 6.1e-05 | 0.000947 | 0.29175 | 1.3239 | 2.5473 | 100.3 | 21.09 | 32.066 | 68.6728 | 83.2469 | 66.2585 | 0.008004 | 1.0402 | 0.02492 | 23.957 | 0.15693 | 0.1395 | 0.000162 | 0.012213 | 0.012213 | 0.46488
2019-03-29 11:50 | 1553874621.8697 | 902 | 0.20948 | 0.000463 | -1e-06 | 90.3983 | 88.811 | 126.8403 | 21.1465 | 0.001903 | 0.001873 | -151.8529 | 0.23386 | 8.402 | 0.002383 | 0.003974 | 0.003735 | 6.1e-05 | 0.000946 | 0.29231 | 1.3401 | 2.5473 | 100.3 | 21.09 | 32.062 | 67.5039 | 79.0285 | 66.7823 | 0.008003 | 1.0386 | 0.024861 | 23.938 | 0.15845 | 0.13913 | 0.000157 | 0.011821 | 0.011821 | 0.47901
2019-03-29 11:50 | 1553874622.8684 | 903 | 0.20949 | 0.000461 | -1e-06 | 90.436 | 88.8946 | 126.7553 | 21.1463 | 0.001987 | 0.001872 | -151.846 | 0.23439 | 8.401 | 0.002487 | 0.003621 | 0.003464 | 6.3e-05 | 0.000945 | 0.29551 | 1.357 | 2.5472 | 100.3 | 21.09 | 32.061 | 69.5278 | 81.532 | 66.7549 | 0.008003 | 1.0387 | 0.024854 | 23.9291 | 0.15376 | 0.13944 | 9e-05 | 0.006807 | 0.006807 | 0.83367
2019-03-29 11:50 | 1553874623.8694 | 904 | 0.20952 | 0.000466 | -1e-06 | 90.4357 | 88.9881 | 126.6804 | 21.1473 | 0.001951 | 0.001862 | -151.8496 | 0.2341 | 8.4026 | 0.002377 | 0.003606 | 0.003439 | 6.2e-05 | 0.000946 | 0.31404 | 1.338 | 2.5473 | 100.3 | 21.09 | 32.06 | 68.6514 | 78.9029 | 66.5108 | 0.008003 | 1.0394 | 0.024856 | 23.9134 | 0.14433 | 0.13926 | -0.000107 | -0.007919 | -0.007919 | -0.7158
2019-03-29 11:50 | 1553874624.8692 | 905 | 0.20949 | 0.000465 | -1e-06 | 90.6536 | 89.1007 | 126.7412 | 21.1463 | 0.001918 | 0.00193 | -151.8552 | 0.23436 | 8.4008 | 0.002426 | 0.00411 | 0.003953 | 6.1e-05 | 0.000945 | 0.3126 | 1.299 | 2.5473 | 100.3 | 21.09 | 32.055 | 67.872 | 80.0645 | 68.1567 | 0.008002 | 1.0344 | 0.024802 | 23.977 | 0.15454 | 0.13942 | 5.5e-05 | 0.004204 | 0.004204 | 1.3498
2019-03-29 11:50 | 1553874625.8689 | 906 | 0.20951 | 0.000462 | -1e-06 | 90.8095 | 89.3052 | 127.305 | 21.1467 | 0.001986 | 0.001855 | -151.8493 | 0.23479 | 8.4038 | 0.002375 | 0.00329 | 0.003077 | 6.1e-05 | 0.00095 | 0.29031 | 1.3101 | 2.5471 | 100.3 | 21.09 | 32.053 | 69.5047 | 78.8464 | 66.3448 | 0.008001 | 1.0399 | 0.024923 | 23.9664 | 0.14942 | 0.13967 | -4.6e-05 | -0.003395 | -0.003395 | -1.6746
2019-03-29 11:50 | 1553874626.8694 | 907 | 0.20951 | 0.000458 | -1e-06 | 90.8886 | 89.4516 | 127.3628 | 21.1458 | 0.001917 | 0.001896 | -151.8367 | 0.23472 | 8.4004 | 0.002387 | 0.003633 | 0.003495 | 6.2e-05 | 0.00095 | 0.29358 | 1.3673 | 2.5474 | 100.3 | 21.09 | 32.047 | 67.8309 | 79.1237 | 67.3262 | 0.008 | 1.0369 | 0.024893 | 24.0065 | 0.14255 | 0.13964 | -6.6e-05 | -0.004819 | -0.004819 | -1.1793
2019-03-29 11:50 | 1553874627.8689 | 908 | 0.20949 | 0.00046 | -1e-06 | 91.0047 | 89.5884 | 127.1459 | 21.1471 | 0.001922 | 0.00183 | -151.8502 | 0.23438 | 8.4011 | 0.00243 | 0.003826 | 0.003676 | 6.3e-05 | 0.000946 | 0.30574 | 1.3905 | 2.5475 | 100.3 | 21.09 | 32.045 | 67.95 | 80.1751 | 65.7528 | 0.007999 | 1.0417 | 0.024929 | 23.9305 | 0.1403 | 0.13943 | 7.2e-05 | 0.005522 | 0.005522 | 1.0278
2019-03-29 11:50 | 1553874628.869 | 909 | 0.20954 | 0.000459 | -1e-06 | 91.058 | 89.754 | 127.0724 | 21.1472 | 0.001882 | 0.001875 | -151.8315 | 0.2342 | 8.4024 | 0.002343 | 0.003771 | 0.003625 | 6.2e-05 | 0.000949 | 0.32012 | 1.3603 | 2.5475 | 100.3 | 21.09 | 32.04 | 66.9985 | 78.0719 | 66.8313 | 0.007998 | 1.0384 | 0.024882 | 23.9617 | 0.12902 | 0.13933 | -0.000231 | -0.017196 | -0.017196 | -0.32976
2019-03-29 11:50 | 1553874629.8686 | 910 | 0.20952 | 0.000463 | -1e-06 | 91.1199 | 89.7895 | 127.1205 | 21.1475 | 0.001951 | 0.001882 | -151.8362 | 0.23399 | 8.405 | 0.002327 | 0.003649 | 0.00349 | 6.1e-05 | 0.000952 | 0.27367 | 1.3483 | 2.5474 | 100.3 | 21.09 | 32.031 | 68.6548 | 77.6971 | 67.0095 | 0.007996 | 1.0379 | 0.024881 | 23.9725 | 0.13157 | 0.1392 | -0.000132 | -0.009768 | -0.009768 | -0.57999
2019-03-29 11:50 | 1553874630.8696 | 911 | 0.2095 | 0.000464 | -1e-06 | 91.2262 | 89.8168 | 127.409 | 21.1473 | 0.001928 | 0.001853 | -151.8535 | 0.23495 | 8.4019 | 0.002385 | 0.003102 | 0.002929 | 6.3e-05 | 0.000959 | 0.31564 | 1.385 | 2.5474 | 100.3 | 21.09 | 32.026 | 68.0983 | 79.0932 | 66.3134 | 0.007994 | 1.04 | 0.024934 | 23.9751 | 0.13927 | 0.13977 | 2.8e-05 | 0.002158 | 0.002158 | 2.6366
2019-03-29 11:50 | 1553874631.868 | 912 | 0.20951 | 0.000465 | -1e-06 | 91.3771 | 89.941 | 127.7043 | 21.1474 | 0.001941 | 0.001864 | -151.8572 | 0.23477 | 8.4021 | 0.002408 | 0.004098 | 0.003959 | 6.2e-05 | 0.000963 | 0.28351 | 1.3774 | 2.5475 | 100.3 | 21.09 | 32.022 | 68.4136 | 79.6484 | 66.578 | 0.007993 | 1.0392 | 0.024954 | 24.0122 | 0.14169 | 0.13966 | -3.2e-05 | -0.002293 | -0.002293 | -2.4791
2019-03-29 11:51 | 1553874632.8687 | 913 | 0.20949 | 0.000463 | -1e-06 | 91.4762 | 90.1164 | 127.828 | 21.148 | 0.001878 | 0.001829 | -151.8611 | 0.23379 | 8.4031 | 0.002335 | 0.003721 | 0.003575 | 6.1e-05 | 0.000961 | 0.29822 | 1.3569 | 2.5474 | 100.3 | 21.09 | 32.012 | 66.8939 | 77.8918 | 65.7289 | 0.007991 | 1.0418 | 0.024997 | 23.9938 | 0.13395 | 0.13908 | 8.1e-05 | 0.006149 | 0.006149 | 0.92057
2019-03-29 11:51 | 1553874633.8688 | 914 | 0.20949 | 0.000459 | -2e-06 | 91.4809 | 90.2167 | 128.0137 | 21.1483 | 0.001916 | 0.001843 | -151.8638 | 0.23395 | 8.4024 | 0.00235 | 0.003461 | 0.003308 | 6.1e-05 | 0.000962 | 0.31437 | 1.3684 | 2.5474 | 100.3 | 21.09 | 32.008 | 67.8194 | 78.2467 | 66.0731 | 0.00799 | 1.0407 | 0.025002 | 24.0234 | 0.12447 | 0.13918 | 0.000105 | 0.007967 | 0.007967 | 0.71095
2019-03-29 11:51 | 1553874634.8691 | 915 | 0.20952 | 0.000458 | -1e-06 | 91.6171 | 90.1744 | 128.123 | 21.1478 | 0.001926 | 0.001865 | -151.8681 | 0.23357 | 8.4014 | 0.002367 | 0.003944 | 0.003739 | 5.8e-05 | 0.000956 | 0.31006 | 1.3503 | 2.5474 | 100.3 | 21.09 | 31.999 | 68.0567 | 78.6571 | 66.601 | 0.007988 | 1.0391 | 0.024994 | 24.0524 | 0.14197 | 0.13895 | -0.000108 | -0.008032 | -0.008032 | -0.70413
2019-03-29 11:51 | 1553874635.8677 | 916 | 0.20947 | 0.000459 | -1e-06 | 91.6282 | 90.3509 | 127.8012 | 21.1486 | 0.001946 | 0.001837 | -151.8737 | 0.23511 | 8.4036 | 0.002347 | 0.003595 | 0.003419 | 6.1e-05 | 0.00095 | 0.29446 | 1.3645 | 2.5474 | 100.3 | 21.09 | 31.996 | 68.5356 | 78.1667 | 65.9214 | 0.007987 | 1.0412 | 0.024987 | 23.9981 | 0.12557 | 0.13986 | 0.00023 | 0.017341 | 0.017341 | 0.32828
2019-03-29 11:51 | 1553874636.8679 | 917 | 0.20949 | 0.000456 | -1e-06 | 91.6599 | 90.3072 | 127.7252 | 21.1481 | 0.001892 | 0.001886 | -151.8694 | 0.23608 | 8.4036 | 0.002336 | 0.00363 | 0.003504 | 6.2e-05 | 0.000947 | 0.29623 | 1.3769 | 2.5475 | 100.3 | 21.08 | 31.995 | 67.2513 | 77.92 | 67.1068 | 0.007982 | 1.0376 | 0.024936 | 24.0329 | 0.13299 | 0.14045 | 0.000106 | 0.008032 | 0.008032 | 0.71168
2019-03-29 11:51 | 1553874637.8682 | 918 | 0.20948 | 0.000453 | -2e-06 | 91.6782 | 90.2273 | 127.6102 | 21.1485 | 0.001952 | 0.001812 | -151.8722 | 0.23494 | 8.401 | 0.002384 | 0.003502 | 0.003271 | 5.9e-05 | 0.000937 | 0.31094 | 1.3557 | 2.5474 | 100.3 | 21.08 | 31.989 | 68.6899 | 79.0641 | 65.3101 | 0.00798 | 1.0431 | 0.024991 | 23.9585 | 0.14269 | 0.13977 | 0.000168 | 0.012751 | 0.012751 | 0.44611
2019-03-29 11:51 | 1553874638.8679 | 919 | 0.20954 | 0.000455 | -1e-06 | 91.604 | 90.258 | 127.3867 | 21.1485 | 0.001967 | 0.00181 | -151.8871 | 0.23439 | 8.4034 | 0.002351 | 0.003309 | 0.003236 | 6.3e-05 | 0.000939 | 0.31078 | 1.3938 | 2.5474 | 100.3 | 21.08 | 31.987 | 69.0471 | 78.2688 | 65.2812 | 0.00798 | 1.0432 | 0.02497 | 23.9365 | 0.1324 | 0.13944 | -0.000202 | -0.015097 | -0.015097 | -0.3759
2019-03-29 11:51 | 1553874639.8679 | 920 | 0.20952 | 0.000458 | -1e-06 | 91.524 | 90.0304 | 127.1198 | 21.1487 | 0.001956 | 0.001821 | -151.8965 | 0.23457 | 8.4016 | 0.00236 | 0.003962 | 0.003709 | 6.1e-05 | 0.000941 | 0.30381 | 1.389 | 2.5476 | 100.3 | 21.08 | 31.985 | 68.7745 | 78.4943 | 65.5352 | 0.007979 | 1.0424 | 0.024935 | 23.9204 | 0.14718 | 0.13954 | -0.000115 | -0.008557 | -0.008557 | -0.66372
2019-03-29 11:51 | 1553874640.8683 | 921 | 0.20954 | 0.000458 | -1e-06 | 91.533 | 90.1447 | 127.2187 | 21.1485 | 0.001909 | 0.001892 | -151.8946 | 0.23392 | 8.4018 | 0.002293 | 0.003583 | 0.003365 | 6.2e-05 | 0.000938 | 0.3282 | 1.3729 | 2.5474 | 100.3 | 21.08 | 31.985 | 67.6518 | 76.8781 | 67.2493 | 0.007979 | 1.0372 | 0.024881 | 23.9902 | 0.1367 | 0.13916 | -0.0002 | -0.014934 | -0.014934 | -0.37925
2019-03-29 11:51 | 1553874641.8689 | 922 | 0.20953 | 0.000458 | -1e-06 | 91.4037 | 90.0693 | 126.9839 | 21.1489 | 0.001893 | 0.001786 | -151.8935 | 0.23468 | 8.401 | 0.00232 | 0.003354 | 0.003222 | 6e-05 | 0.000935 | 0.3194 | 1.3599 | 2.5475 | 100.3 | 21.08 | 31.984 | 67.2598 | 77.5193 | 64.6883 | 0.007979 | 1.045 | 0.024953 | 23.8777 | 0.13154 | 0.13961 | -0.000162 | -0.012093 | -0.012093 | -0.46989
2019-03-29 11:51 | 1553874642.8683 | 923 | 0.20953 | 0.000456 | -2e-06 | 91.3055 | 89.97 | 127.0161 | 21.1485 | 0.001927 | 0.001824 | -151.9041 | 0.23435 | 8.4029 | 0.002335 | 0.003571 | 0.003398 | 6e-05 | 0.000936 | 0.31006 | 1.3381 | 2.5474 | 100.3 | 21.08 | 31.984 | 68.0942 | 77.8822 | 65.5997 | 0.007979 | 1.0422 | 0.024922 | 23.9129 | 0.1318 | 0.13941 | -0.000167 | -0.012445 | -0.012445 | -0.45594
2019-03-29 11:51 | 1553874643.8689 | 924 | 0.20951 | 0.000457 | -2e-06 | 91.2831 | 89.7843 | 126.9248 | 21.1489 | 0.001864 | 0.001825 | -151.9033 | 0.23388 | 8.4034 | 0.002262 | 0.003739 | 0.003548 | 6.1e-05 | 0.000933 | 0.30646 | 1.3523 | 2.5475 | 100.3 | 21.08 | 31.984 | 66.5577 | 76.1381 | 65.637 | 0.007979 | 1.0421 | 0.024912 | 23.9056 | 0.14808 | 0.13914 | -3.1e-05 | -0.002196 | -0.002196 | -2.5782
2019-03-29 11:51 | 1553874644.868 | 925 | 0.20954 | 0.000461 | -1e-06 | 91.1265 | 89.7533 | 126.9823 | 21.1494 | 0.001907 | 0.001803 | -151.9145 | 0.23366 | 8.4047 | 0.002318 | 0.003466 | 0.003315 | 5.8e-05 | 0.000934 | 0.32189 | 1.35 | 2.5473 | 100.3 | 21.08 | 31.986 | 67.6027 | 77.4802 | 65.0971 | 0.007979 | 1.0437 | 0.024937 | 23.892 | 0.13582 | 0.139 | -0.000227 | -0.016934 | -0.016934 | -0.33409
2019-03-29 11:51 | 1553874645.8673 | 926 |  | 0.000456 |  | 91.1047 | 89.7939 | 127.0644 | 21.1487 | 0.001932 | 0.001864 | -151.9074 | 0.2347 | 8.3993 | 0.002329 | 0.003511 | 0.003316 | 6.2e-05 | 0.000937 | 0.31006 | 1.3557 | 2.5474 | 100.3 | 21.08 | 31.989 | 68.1934 | 77.7389 | 66.5754 | 0.00798 | 1.0392 | 0.024891 | 23.9519 | 0.12963 | 0.13962
2019-03-29 11:51 | 1553874646.8673 | 927 |  | 0.000459 |  | 90.8745 | 89.5048 | 127.0496 | 21.1484 | 0.001942 | 0.001806 | -151.9325 | 0.23332 | 8.4009 | 0.002245 | 0.003975 | 0.003832 | 6.3e-05 | 0.000936 | 0.29662 | 1.4299 | 2.5474 | 100.3 | 21.08 | 31.992 | 68.4504 | 75.7159 | 65.1813 | 0.007981 | 1.0435 | 0.024941 | 23.9013 | 0.13584 | 0.1388
2019-03-29 11:51 | 1553874647.8673 | 928 |  |  |  | 90.7883 | 89.3113 | 127.1358 | 21.1488 | 0.001955 | 0.001834 | -151.9296 | 0.23434 | 8.4018 | 0.002257 | 0.003085 | 0.002794 | 5.9e-05 | 0.000937 | 0.28296 | 1.4016 | 2.5473 | 100.3 | 21.08 | 31.999 | 68.7627 | 76.0195 | 65.8475 | 0.007983 | 1.0414 | 0.024925 | 23.9329 | 0.14671 | 0.13941
2019-03-29 11:51 | 1553874648.8673 | 929 |  |  |  | 90.6222 | 89.2342 | 127.4304 | 21.1493 | 0.001885 | 0.00184 | -151.9256 | 0.23455 | 8.4031 | 0.002236 | 0.002756 | 0.002588 | 6.2e-05 | 0.000938 | 0.29247 | 1.4034 | 2.5475 | 100.3 | 21.08 | 32.005 | 67.0729 | 75.5078 | 65.9844 | 0.007984 | 1.041 | 0.024949 | 23.9655 | 0.13806 | 0.13954
2019-03-29 11:51 | 1553874649.8673 | 930 |  |  |  | 90.5809 | 89.0379 | 127.5367 | 21.1486 | 0.001875 | 0.001845 | -151.915 | 0.23342 | 8.4039 | 0.002262 | 0.003183 | 0.00307 | 6.1e-05 | 0.000929 | 0.30879 | 1.3797 | 2.5474 | 100.3 | 21.08 | 32.011 | 66.8421 | 76.1455 | 66.111 | 0.007986 | 1.0406 | 0.024954 | 23.98 | 0.15367 | 0.13886
2019-03-29 11:51 | 1553874650.8673 | 931 |  |  |  | 90.5416 | 88.856 |  | 21.1493 | 0.001912 | 0.001833 | -151.9126 | 0.23498 | 8.4042 | 0.0022 | 0.003901 | 0.003692 | 6.1e-05 | 0.000924 | 0.32228 | 1.361 | 2.5474 | 100.3 | 21.08 | 32.022 | 67.7322 | 74.641 | 65.8269 | 0.007988 | 1.0415 |  |  | 0.16809 | 0.13979
2019-03-29 11:51 | 1553874651.8673 | 932 |  |  |  | 90.5147 | 89.084 |  | 21.1491 | 0.001904 |  | -151.9268 | 0.234 | 8.4026 | 0.002233 | 0.002966 | 0.002813 | 6.3e-05 | 0.000926 | 0.31885 | 1.338 | 2.5473 | 100.3 | 21.08 | 32.03 | 67.5238 | 75.4382 |  | 0.00799 |  |  |  | 0.14251 | 0.13921
2019-03-29 11:51 | 1553874652.8673 | 933 |  |  |  | 90.4426 | 88.8622 |  | 21.1496 | 0.001866 |  |  | 0.23279 | 8.4006 | 0.002181 | 0.003244 | 0.003096 | 6.1e-05 | 0.000925 | 0.31404 | 1.3276 | 2.5474 | 100.3 | 21.07 | 32.046 | 66.624 | 74.1915 |  | 0.007989 |  |  |  | 0.15768 | 0.13849
2019-03-29 11:51 | 1553874653.8673 | 934 |  |  |  | 90.3135 | 88.9127 |  | 21.1491 |  |  |  | 0.23443 | 8.4035 | 0.002152 | 0.00332 | 0.003178 | 5.8e-05 | 0.000927 | 0.27992 | 1.3653 | 2.5474 | 100.3 | 21.07 | 32.056 |  | 73.4888 |  | 0.007992 |  |  |  | 0.13983 | 0.13946
2019-03-29 11:51 | 1553874654.8673 | 935 |  |  |  | 90.2887 | 88.7953 |  | 21.1491 |  |  |  | 0.23408 | 8.4037 | 0.002203 | 0.003299 | 0.003061 | 6.2e-05 | 0.000925 | 0.2677 | 1.3615 | 2.5472 | 100.3 | 21.07 | 32.065 |  | 74.7133 |  | 0.007994 |  |  |  | 0.14919 | 0.13925
2019-03-29 11:51 | 1553874655.8673 | 936 |  |  |  | 90.2624 | 88.74 |  | 21.15 |  |  |  | 0.23384 | 8.4026 | 0.002222 | 0.003486 | 0.003319 | 6.1e-05 | 0.000928 | 0.31598 | 1.3081 | 2.547 | 100.3 | 21.07 | 32.087 |  | 75.1764 |  | 0.008 |  |  |  | 0.15215 | 0.13911
2019-03-29 11:51 | 1553874656.8673 | 937 |  |  |  | 90.3016 | 88.7407 |  | 21.1498 |  |  |  | 0.23389 | 8.4006 | 0.002309 | 0.00336 | 0.00318 | 6e-05 | 0.000931 | 0.31598 | 1.3217 | 2.5472 | 100.3 | 21.07 | 32.097 |  | 77.2701 |  | 0.008002 |  |  |  | 0.15596 | 0.13914
2019-03-29 11:51 | 1553874657.8673

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-29 11:35 | 0 | 0 | 0 | 0 | 0 | 47.7 | 0.004607
2019-03-29 11:35 | 1 | 0 | 0 | 0 | 0 | 47.7036 | 0.001976
2019-03-29 11:35 | 2 | 0 | 0 | 0 | 0 | 47.7539 | 0.010774
2019-03-29 11:35 | 3 | 0 | 0 | 0 | 0 | 47.5173 | 0.02009
2019-03-29 11:35 | 4 | 0 | 0 | 0 | 0 | 47.6841 | 0.020219
2019-03-29 11:35 | 5 | 0 | 0 | 0 | 0 | 47.8658 | 0.040215
2019-03-29 11:35 | 6 | 0 | 0 | 0 | 0 | 47.6484 | 0.16482
2019-03-29 11:35 | 7 | 0 | 0 | 0 | 0 | 47.6055 | 0.63164
2019-03-29 11:35 | 8 | 0 | 0 | 0 | 0 | 47.563 | 1.1235
2019-03-29 11:35 | 9 | 0 | 0 | 0 | 0 | 47.4921 | 1.5498
2019-03-29 11:35 | 10 | 0 | 0 | 0 | 0 | 47.4036 | 1.8938
2019-03-29 11:35 | 11 | 0 | 0 | 0 | 0 | 47.3019 | 2.1522
2019-03-29 11:35 | 12 | 0 | 0 | 0 | 0 | 47.1829 | 2.5572
2019-03-29 11:36 | 13 | 0 | 0 | 0 | 0 | 47.0865 | 2.7348
2019-03-29 11:36 | 14 | 0.11557 | 0 | 0 | 0 | 46.9761 | 2.9857
2019-03-29 11:36 | 15 | 0.10446 | 0 | 0 | 0 | 46.8611 | 3.1877
2019-03-29 11:36 | 16 | 0.098469 | 0 | 0 | 0 | 46.7664 | 3.1732
2019-03-29 11:36 | 17 | 0.10104 | 0 | 0 | 0 | 46.6621 | 3.1604
2019-03-29 11:36 | 18 | 0.094401 | 0 | 0 | 0 | 46.5656 | 3.1471
2019-03-29 11:36 | 19 | 0.11321 | 0 | 0 | 0 | 46.4679 | 3.0485
2019-03-29 11:36 | 20 | 0.10833 | 0 | 0 | 0 | 46.3452 | 2.8835
2019-03-29 11:36 | 21 | 0.077781 | 0 | 0 | 0 | 46.2574 | 2.7435
2019-03-29 11:36 | 22 | 0.078929 | 0 | 0 | 0 | 46.1822 | 2.5826
2019-03-29 11:36 | 23 | 0.090475 | 0 | 0 | 0 | 46.0966 | 2.4338
2019-03-29 11:36 | 24 | 0.079636 | 0 | 0 | 0 | 46.006 | 2.3817
2019-03-29 11:36 | 25 | 0.083605 | 0 | 0 | 0 | 45.9341 | 2.3154
2019-03-29 11:36 | 26 | 0.093116 | 0 | 0 | 0 | 45.8377 | 2.2455
2019-03-29 11:36 | 27 | 0.082727 | 0 | 0 | 0 | 45.7528 | 2.202
2019-03-29 11:36 | 28 | 0.070764 | 0 | 0 | 0 | 45.6732 | 2.1277
2019-03-29 11:36 | 29 | 0.056273 | 0 | 0 | 0 | 45.6109 | 2.0521
2019-03-29 11:36 | 30 | 0.078968 | 0 | 0 | 0 | 45.5514 | 1.9888
2019-03-29 11:36 | 31 | 0.098526 | 0 | 0 | 0 | 45.4536 | 1.9537
2019-03-29 11:36 | 32 | 0.083584 | 0 | 0 | 0 | 45.3628 | 1.9591
2019-03-29 11:36 | 33 | 0.078404 | 0 | 0 | 0 | 45.2847 | 1.9799
2019-03-29 11:36 | 34 | 0.088487 | 0 | 0 | 0 | 45.2026 | 1.978
2019-03-29 11:36 | 35 | 0.077437 | 0 | 0 | 0 | 45.1126 | 1.9549
2019-03-29 11:36 | 36 | 0.067354 | 0 | 0 | 0 | 45.048 | 1.8883
2019-03-29 11:36 | 37 | 0.058102 | 0 | 0 | 0 | 44.9772 | 1.7701
2019-03-29 11:36 | 38 | 0.063654 | 0 | 0 | 0 | 44.9281 | 1.6849
2019-03-29 11:36 | 39 | 0.067027 | 0 | 0 | 0 | 44.8507 | 1.5701
2019-03-29 11:36 | 40 | 0.047119 | 0 | 0 | 0 | 44.7993 | 1.5125
2019-03-29 11:36 | 41 | 0.048541 | 0 | 0 | 0 | 44.7516 | 1.4409
2019-03-29 11:36 | 42 | 0.053084 | 0 | 0 | 0 | 44.7009 | 1.3503
2019-03-29 11:36 | 43 | 0.057605 | 0 | 0 | 0 | 44.6458 | 1.3144
2019-03-29 11:36 | 44 | 0.044725 | 0 | 0 | 0 | 44.5898 | 1.2479
2019-03-29 11:36 | 45 | 0.043636 | 0 | 0 | 0 | 44.5537 | 1.2135
2019-03-29 11:36 | 46 | 0.047373 | 0 | 0 | 0 | 44.5007 | 1.1793
2019-03-29 11:36 | 47 | 0.061623 | 0 | 0 | 0 | 44.4567 | 1.1621
2019-03-29 11:36 | 48 | 0.053782 | 0 | 0 | 0 | 44.3825 | 1.1467
2019-03-29 11:36 | 49 | 0.053171 | 0 | 0 | 0 | 44.3478 | 1.1356
2019-03-29 11:36 | 50 | 0.059777 | 0 | 0 | 0 | 44.2749 | 1.1346
2019-03-29 11:36 | 51 | 0.036044 | 0 | 0 | 0 | 44.2342 | 1.131
2019-03-29 11:36 | 52 | 0.065875 | 0 | 0 | 0 | 44.1912 | 1.1185
2019-03-29 11:36 | 53 | 0.06088 | 0 | 0 | 0 | 44.1101 | 1.0968
2019-03-29 11:36 | 54 | 0.043898 | 0 | 0 | 0 | 44.0724 | 1.1089
2019-03-29 11:36 | 55 | 0.060391 | 0 | 0 | 0 | 44.0149 | 1.1038
2019-03-29 11:36 | 56 | 0.056045 | 0 | 0 | 0 | 43.956 | 1.0808
2019-03-29 11:36 | 57 | 0.049871 | 0 | 0 | 0 | 43.9034 | 1.0562
2019-03-29 11:36 | 58 | 0.05712 | 0 | 0 | 0 | 43.8535 | 1.0549
2019-03-29 11:36 | 59 | 0.046591 | 0 | 0 | 0 | 43.793 | 1.0494
2019-03-29 11:36 | 60 | 0.046865 | 0 | 0 | 0 | 43.7578 | 1.0401
2019-03-29 11:36 | 61 | 0.059916 | 0 | 0 | 0 | 43.6969 | 1.0315
2019-03-29 11:36 | 62 | 0.029973 | 0 | 0 | 0 | 43.6473 | 1.0316
2019-03-29 11:36 | 63 | 0.036943 | 0 | 0 | 0 | 43.6284 | 1.0162
2019-03-29 11:36 | 64 | 0.057887 | 0 | 0 | 0 | 43.5697 | 1.0403
2019-03-29 11:36 | 65 | 0.049586 | 0 | 0 | 0 | 43.5194 | 1.0133
2019-03-29 11:36 | 66 | 0.050309 | 0 | 0 | 0 | 43.4693 | 1.0206
2019-03-29 11:36 | 67 | 0.043102 | 0 | 0 | 0 | 43.4203 | 1.0185
2019-03-29 11:36 | 68 | 0.044706 | 0 | 0 | 0 | 43.3808 | 1.0199
2019-03-29 11:36 | 69 | 0.05611 | 0 | 0 | 0 | 43.3286 | 1.0122
2019-03-29 11:36 | 70 | 0.050651 | 0.014389 | 1.9757 | 19.9986 | 43.2725 | 1.0129
2019-03-29 11:36 | 71 | 0.045351 | 0.015583 | 2.1999 | 22.2563 | 43.2272 | 1.0094
2019-03-29 11:36 | 72 | 0.058372 | 0.012107 | 1.7018 | 17.4265 | 43.1776 | 1.0172
2019-03-29 11:36 | 73 | 0.058009 | 0.011772 | 1.7068 | 17.2776 | 43.1137 | 1.0023
2019-03-29 11:37 | 74 | 0.047494 | 0.014366 | 2.0769 | 20.9711 | 43.0643 | 0.99601
2019-03-29 11:37 | 75 | 0.044232 | 0.015467 | 2.2352 | 22.6606 | 43.0168 | 1.0023
2019-03-29 11:37 | 76 | 0.053525 | 0.013174 | 1.8358 | 18.5765 | 42.9729 | 0.99431
2019-03-29 11:37 | 77 | 0.047801 | 0.014272 | 2.0484 | 20.7051 | 42.9132 | 0.98972
2019-03-29 11:37 | 78 | 0.041955 | 0.016862 | 2.3394 | 23.5017 | 42.8773 | 0.98601
2019-03-29 11:37 | 79 | 0.050114 | 0.013669 | 1.9525 | 19.5946 | 42.8261 | 0.98197
2019-03-29 11:37 | 80 | 0.046985 | 0.014536 | 2.0673 | 20.8529 | 42.7794 | 0.97978
2019-03-29 11:37 | 81 | 0.048605 | 0.014101 | 2.0046 | 20.2393 | 42.7313 | 0.98373
2019-03-29 11:37 | 82 | 0.047204 | 0.01506 | 2.0643 | 20.7704 | 42.6831 | 0.98045
2019-03-29 11:37 | 83 | 0.04026 | 0.017606 | 2.4087 | 24.4793 | 42.6383 | 0.98555
2019-03-29 11:37 | 84 | 0.033905 | 0.020221 | 2.8619 | 29.1095 | 42.6022 | 0.98696
2019-03-29 11:37 | 85 | 0.040091 | 0.017089 | 2.4124 | 24.4391 | 42.5675 | 0.9798
2019-03-29 11:37 | 86 | 0.039195 | 0.017529 | 2.4718 | 24.911 | 42.5232 | 0.97638
2019-03-29 11:37 | 87 | 0.045939 | 0.014953 | 2.103 | 21.2613 | 42.4874 | 0.97672
2019-03-29 11:37 | 88 | 0.052247 | 0.013135 | 1.846 | 18.6881 | 42.4316 | 0.97639
2019-03-29 11:37 | 89 | 0.04611 | 0.014385 | 2.0898 | 20.9741 | 42.386 | 0.96711
2019-03-29 11:37 | 90 | 0.042431 | 0.015611 | 2.2646 | 22.9782 | 42.3388 | 0.97499
2019-03-29 11:37 | 91 | 0.048188 | 0.013703 | 1.9826 | 20.1244 | 42.2987 | 0.96975
2019-03-29 11:37 | 92 | 0.054085 | 0.012195 | 1.761 | 17.8601 | 42.2426 | 0.96596
2019-03-29 11:37 | 93 | 0.042753 | 0.015453 | 2.2283 | 22.8162 | 42.1945 | 0.97545
2019-03-29 11:37 | 94 | 0.045579 | 0.013984 | 2.0868 | 21.247 | 42.154 | 0.96843
2019-03-29 11:37 | 95 | 0.048882 | 0.012609 | 1.9476 | 19.681 | 42.1028 | 0.96206
2019-03-29 11:37 | 96 | 0.049033 | 0.012512 | 1.9334 | 19.6816 | 42.0566 | 0.96504
2019-03-29 11:37 | 97 | 0.062386 | 0.0098505 | 1.5158 | 15.5128 | 42.0019 | 0.96778
2019-03-29 11:37 | 98 | 0.061268 | 0.010056 | 1.5462 | 15.615 | 41.9352 | 0.9567
2019-03-29 11:37 | 99 | 0.055917 | 0.010984 | 1.6859 | 16.817 | 41.8804 | 0.94036
2019-03-29 11:37 | 100 | 0.07053 | 0.0087042 | 1.3303 | 13.357 | 41.8194 | 0.94207
2019-03-29 11:37 | 101 | 0.057697 | 0.010642 | 1.6233 | 16.2792 | 41.746 | 0.93926
2019-03-29 11:37 | 102 | 0.032194 | 0.019105 | 2.9044 | 29.537 | 41.7067 | 0.95091
2019-03-29 11:37 | 103 | 0.047402 | 0.012964 | 1.9663 | 19.7499 | 41.6714 | 0.93619
2019-03-29 11:37 | 104 | 0.074361 | 0.0085637 | 1.2501 | 12.8579 | 41.6089 | 0.95612
2019-03-29 11:37 | 105 | 0.061103 | 0.010063 | 1.5245 | 15.7204 | 41.5325 | 0.96057
2019-03-29 11:37 | 106 | 0.04663 | 0.013661 | 1.9956 | 20.5086 | 41.4878 | 0.95632
2019-03-29 11:37 | 107 | 0.047352 | 0.013411 | 1.9651 | 19.89 | 41.4357 | 0.94184
2019-03-29 11:37 | 108 | 0.043138 | 0.01421 | 2.1569 | 21.7989 | 41.3941 | 0.94035
2019-03-29 11:37 | 109 | 0.031941 | 0.018522 | 2.924 | 29.6179 | 41.3514 | 0.94602
2019-03-29 11:37 | 110 | 0.015783 | 0.037592 | 5.9233 | 59.8557 | 41.3308 | 0.94468
2019-03-29 11:37 | 111 | 0.031007 | 0.019125 | 3.0134 | 30.6106 | 41.3121 | 0.94916
2019-03-29 11:37 | 112 | 0.044191 | 0.013382 | 2.1085 | 21.4296 | 41.269 | 0.947
2019-03-29 11:37 | 113 | 0.032594 | 0.018813 | 2.8491 | 29.0033 | 41.2296 | 0.94534
2019-03-29 11:37 | 114 | 0.033854 | 0.01809 | 2.7395 | 27.7035 | 41.2011 | 0.93788
2019-03-29 11:37 | 115 | 0.044255 | 0.014368 | 2.0974 | 21.5348 | 41.1598 | 0.95303
2019-03-29 11:37 | 116 | 0.036772 | 0.01734 | 2.5283 | 26.2356 | 41.1168 | 0.96475
2019-03-29 11:37 | 117 | 0.028442 | 0.022413 | 3.2768 | 33.7938 | 41.0864 | 0.96117
2019-03-29 11:37 | 118 | 0.036801 | 0.017314 | 2.5333 | 26.3675 | 41.0555 | 0.97035
2019-03-29 11:37 | 119 | 0.052509 | 0.011701 | 1.7767 | 18.4398 | 41.0109 | 0.96825
2019-03-29 11:37 | 120 | 0.054967 | 0.011575 | 1.6994 | 17.6146 | 40.954 | 0.96821
2019-03-29 11:37 | 121 | 0.043762 | 0.01451 | 2.1327 | 21.9541 | 40.904 | 0.96075
2019-03-29 11:37 | 122 | 0.059311 | 0.011107 | 1.5816 | 16.4114 | 40.8601 | 0.97338
2019-03-29 11:37 | 123 | 0.076987 | 0.008559 | 1.2207 | 12.5162 | 40.7853 | 0.96358
2019-03-29 11:37 | 124 | 0.051761 | 0.013194 | 1.8211 | 18.9752 | 40.7159 | 0.98217
2019-03-29 11:37 | 125 | 0.06092 | 0.011602 | 1.5514 | 15.9866 | 40.6743 | 0.9739
2019-03-29 11:37 | 126 | 0.072532 | 0.0096968 | 1.3016 | 13.372 | 40.5933 | 0.9699
2019-03-29 11:37 | 127 | 0.059057 | 0.011909 | 1.601 | 16.776 | 40.5351 | 0.99073
2019-03-29 11:37 | 128 | 0.057647 | 0.012603 | 1.6462 | 17.0231 | 40.4729 | 0.98133
2019-03-29 11:37 | 129 | 0.051996 | 0.014414 | 1.8319 | 19.0189 | 40.4202 | 0.98891
2019-03-29 11:37 | 130 | 0.059768 | 0.012897 | 1.5933 | 16.573 | 40.3655 | 0.99054
2019-03-29 11:37 | 131 | 0.074362 | 0.010654 | 1.2825 | 13.3762 | 40.2994 | 0.99468
2019-03-29 11:37 | 132 | 0.058715 | 0.014278 | 1.6263 | 16.8034 | 40.2244 | 0.98662
2019-03-29 11:38 | 133 | 0.034054 | 0.024636 | 2.8136 | 28.9524 | 40.1841 | 0.98594
2019-03-29 11:38 | 134 | 0.047234 | 0.018263 | 2.0331 | 21.0832 | 40.1473 | 0.99585
2019-03-29 11:38 | 135 | 0.05179 | 0.016613 | 1.8522 | 19.1311 | 40.0909 | 0.9908
2019-03-29 11:38 | 136 | 0.048025 | 0.018421 | 2.0047 | 20.5571 | 40.0457 | 0.98725
2019-03-29 11:38 | 137 | 0.052304 | 0.016901 | 1.8408 | 18.7648 | 39.993 | 0.98147
2019-03-29 11:38 | 138 | 0.05888 | 0.015026 | 1.6401 | 16.477 | 39.9406 | 0.97017
2019-03-29 11:38 | 139 | 0.05566 | 0.015862 | 1.7339 | 17.6172 | 39.878 | 0.98058
2019-03-29 11:38 | 140 | 0.043102 | 0.019884 | 2.2333 | 22.6357 | 39.8316 | 0.97564
2019-03-29 11:38 | 141 | 0.041949 | 0.020544 | 2.3091 | 23.2985 | 39.7889 | 0.97734
2019-03-29 11:38 | 142 | 0.04665 | 0.018476 | 2.0759 | 20.6254 | 39.7456 | 0.96218
2019-03-29 11:38 | 143 | 0.064509 | 0.013337 | 1.4969 | 14.9899 | 39.6932 | 0.96698
2019-03-29 11:38 | 144 | 0.045781 | 0.018814 | 2.1061 | 21.5642 | 39.6254 | 0.98723
2019-03-29 11:38 | 145 | 0.027644 | 0.032021 | 3.4913 | 35.4571 | 39.601 | 0.98017
2019-03-29 11:38 | 146 | 0.061263 | 0.014848 | 1.5761 | 16.1693 | 39.5576 | 0.99057
2019-03-29 11:38 | 147 | 0.075656 | 0.012303 | 1.2741 | 12.9947 | 39.4827 | 0.98313
2019-03-29 11:38 | 148 | 0.060007 | 0.016279 | 1.6051 | 16.403 | 39.4141 | 0.98429
2019-03-29 11:38 | 149 | 0.053151 | 0.018824 | 1.8098 | 18.3555 | 39.3611 | 0.97562
2019-03-29 11:38 | 150 | 0.057798 | 0.018111 | 1.6656 | 16.5073 | 39.3052 | 0.95409
2019-03-29 11:38 | 151 | 0.045926 | 0.023282 | 2.0963 | 21.1118 | 39.2491 | 0.96958
2019-03-29 11:38 | 152 | 0.042574 | 0.025065 | 2.2627 | 22.5483 | 39.2113 | 0.95996
2019-03-29 11:38 | 153 | 0.048869 | 0.021853 | 1.9764 | 19.7731 | 39.161 | 0.9663
2019-03-29 11:38 | 154 | 0.062295 | 0.017153 | 1.5491 | 15.6559 | 39.1118 | 0.97529
2019-03-29 11:38 | 155 | 0.064461 | 0.016915 | 1.4957 | 14.7522 | 39.0395 | 0.95094
2019-03-29 11:38 | 156 | 0.050412 | 0.022546 | 1.9146 | 19.095 | 38.9868 | 0.96262
2019-03-29 11:38 | 157 | 0.05432 | 0.021342 | 1.7735 | 17.5822 | 38.9348 | 0.95507
2019-03-29 11:38 | 158 | 0.054625 | 0.021229 | 1.7616 | 17.3555 | 38.8789 | 0.94804
2019-03-29 11:38 | 159 | 0.042738 | 0.027136 | 2.2467 | 22.217 | 38.8281 | 0.94951
2019-03-29 11:38 | 160 | 0.050657 | 0.023361 | 1.8913 | 18.7919 | 38.7885 | 0.95195
2019-03-29 11:38 | 161 | 0.066197 | 0.018193 | 1.4435 | 14.255 | 38.7247 | 0.94364
2019-03-29 11:38 | 162 | 0.066695 | 0.018078 | 1.4311 | 13.9742 | 38.6596 | 0.93201
2019-03-29 11:38 | 163 | 0.072524 | 0.016874 | 1.3082 | 12.9615 | 38.5905 | 0.94002
2019-03-29 11:38 | 164 | 0.075969 | 0.016714 | 1.248 | 12.3634 | 38.5154 | 0.93924
2019-03-29 11:38 | 165 | 0.06842 | 0.018898 | 1.3843 | 13.8656 | 38.4416 | 0.94868
2019-03-29 11:38 | 166 | 0.056445 | 0.02366 | 1.6731 | 16.6784 | 38.3797 | 0.94141
2019-03-29 11:38 | 167 | 0.052838 | 0.025295 | 1.7854 | 17.6784 | 38.3265 | 0.93409
2019-03-29 11:38 | 168 | 0.052505 | 0.025917 | 1.8021 | 17.9339 | 38.2724 | 0.94161
2019-03-29 11:38 | 169 | 0.072741 | 0.018357 | 1.2967 | 12.7975 | 38.2166 | 0.9309
2019-03-29 11:38 | 170 | 0.077219 | 0.016976 | 1.2207 | 12.2099 | 38.1306 | 0.94283
2019-03-29 11:38 | 171 | 0.071482 | 0.018326 | 1.3138 | 13.0145 | 38.065 | 0.93031
2019-03-29 11:38 | 172 | 0.075801 | 0.01733 | 1.2377 | 12.3042 | 37.9862 | 0.93267
2019-03-29 11:38 | 173 | 0.052494 | 0.025488 | 1.782 | 17.872 | 37.9198 | 0.93817
2019-03-29 11:38 | 174 | 0.040818 | 0.032767 | 2.2848 | 22.701 | 37.8782 | 0.92661
2019-03-29 11:38 | 175 | 0.058939 | 0.022364 | 1.5833 | 15.9079 | 37.8306 | 0.93759
2019-03-29 11:38 | 176 | 0.066419 | 0.020153 | 1.402 | 14.0755 | 37.7626 | 0.93488
2019-03-29 11:38 | 177 | 0.055916 | 0.023947 | 1.6628 | 16.554 | 37.7021 | 0.92563
2019-03-29 11:38 | 178 | 0.060624 | 0.021659 | 1.5298 | 15.5344 | 37.6474 | 0.94176
2019-03-29 11:38 | 179 | 0.069234 | 0.018652 | 1.3431 | 13.3143 | 37.5802 | 0.9218
2019-03-29 11:38 | 180 | 0.053421 | 0.023762 | 1.741 | 17.4556 | 37.5145 | 0.93249
2019-03-29 11:38 | 181 | 0.054275 | 0.023431 | 1.7148 | 17.2769 | 37.47 | 0.9377
2019-03-29 11:38 | 182 | 0.048258 | 0.026378 | 1.9319 | 19.4076 | 37.4069 | 0.93658
2019-03-29 11:38 | 183 | 0.053523 | 0.023765 | 1.7386 | 17.439 | 37.3705 | 0.93338
2019-03-29 11:38 | 184 | 0.070278 | 0.018072 | 1.3241 | 13.257 | 37.2977 | 0.93168
2019-03-29 11:38 | 185 | 0.050434 | 0.025107 | 1.8381 | 18.7941 | 37.2382 | 0.94786
2019-03-29 11:38 | 186 | 0.05071 | 0.024591 | 1.8387 | 18.6945 | 37.1922 | 0.948
2019-03-29 11:38 | 187 | 0.06962 | 0.017895 | 1.3396 | 13.6347 | 37.1322 | 0.94925
2019-03-29 11:38 | 188 | 0.067671 | 0.018449 | 1.3847 | 13.9353 | 37.0578 | 0.94302
2019-03-29 11:38 | 189 | 0.060899 | 0.020531 | 1.5393 | 15.7083 | 36.998 | 0.95661
2019-03-29 11:38 | 190 | 0.077614 | 0.016097 | 1.2086 | 12.2938 | 36.9309 | 0.95417
2019-03-29 11:38 | 191 | 0.069243 | 0.017689 | 1.3507 | 13.5817 | 36.8482 | 0.94044
2019-03-29 11:38 | 192 | 0.069832 | 0.01754 | 1.3403 | 13.5573 | 36.7905 | 0.94674
2019-03-29 11:38 | 193 | 0.083424 | 0.01439 | 1.1217 | 11.4445 | 36.7056 | 0.95474
2019-03-29 11:39 | 194 | 0.077041 | 0.015291 | 1.2144 | 12.2817 | 36.6281 | 0.94619
2019-03-29 11:39 | 195 | 0.08017 | 0.014414 | 1.1685 | 11.8603 | 36.5499 | 0.95084
2019-03-29 11:39 | 196 | 0.076715 | 0.014731 | 1.2172 | 11.9555 | 36.4694 | 0.91717
2019-03-29 11:39 | 197 | 0.064154 | 0.016884 | 1.451 | 14.5373 | 36.3984 | 0.93262
2019-03-29 11:39 | 198 | 0.069609 | 0.015222 | 1.3319 | 13.2803 | 36.3371 | 0.92442
2019-03-29 11:39 | 199 | 0.065683 | 0.015802 | 1.41 | 14.1006 | 36.2614 | 0.92617
2019-03-29 11:39 | 200 | 0.049902 | 0.020357 | 1.8525 | 18.3657 | 36.2084 | 0.91648
2019-03-29 11:39 | 201 | 0.055441 | 0.017479 | 1.6602 | 16.436 | 36.1565 | 0.91123
2019-03-29 11:39 | 202 | 0.063869 | 0.014817 | 1.4375 | 14.269 | 36.0969 | 0.91135
2019-03-29 11:39 | 203 | 0.046618 | 0.019845 | 1.9647 | 19.6461 | 36.0349 | 0.91585
2019-03-29 11:39 | 204 | 0.036533 | 0.024695 | 2.5011 | 24.8855 | 36.0017 | 0.90915
2019-03-29 11:39 | 205 | 0.05696 | 0.015469 | 1.6069 | 16.0492 | 35.9546 | 0.91417
2019-03-29 11:39 | 206 | 0.053621 | 0.01645 | 1.7068 | 17.0083 | 35.8931 | 0.91201
2019-03-29 11:39 | 207 | 0.043663 | 0.019672 | 2.0899 | 20.8071 | 35.8494 | 0.9085
2019-03-29 11:39 | 208 | 0.041027 | 0.020916 | 2.2152 | 21.7994 | 35.8038 | 0.89437
2019-03-29 11:39 | 209 | 0.050833 | 0.01732 | 1.7817 | 17.6233 | 35.7643 | 0.89584
2019-03-29 11:39 | 210 | 0.050939 | 0.017335 | 1.7776 | 17.7811 | 35.7045 | 0.90575
2019-03-29 11:39 | 211 | 0.042499 | 0.020805 | 2.1259 | 21.3201 | 35.6645 | 0.90608
2019-03-29 11:39 | 212 | 0.045475 | 0.019425 | 1.9802 | 20.035 | 35.6168 | 0.9111
2019-03-29 11:39 | 213 | 0.050296 | 0.01715 | 1.7943 | 18.1288 | 35.5731 | 0.9118
2019-03-29 11:39 | 214 | 0.045779 | 0.019383 | 1.9689 | 19.9435 | 35.5186 | 0.913
2019-03-29 11:39 | 215 | 0.035193 | 0.02457 | 2.5594 | 25.9815 | 35.4824 | 0.91436
2019-03-29 11:39 | 216 | 0.056949 | 0.015578 | 1.5774 | 15.9408 | 35.4406 | 0.90781
2019-03-29 11:39 | 217 | 0.066205 | 0.013378 | 1.353 | 13.9027 | 35.3712 | 0.92043
2019-03-29 11:39 | 218 | 0.062156 | 0.014277 | 1.4445 | 14.8481 | 35.3122 | 0.9229
2019-03-29 11:39 | 219 | 0.047816 | 0.018541 | 1.8744 | 19.1657 | 35.2496 | 0.91643
2019-03-29 11:39 | 220 | 0.041821 | 0.021757 | 2.142 | 22.0721 | 35.2143 | 0.92309
2019-03-29 11:39 | 221 | 0.053085 | 0.017111 | 1.6812 | 17.3047 | 35.1618 | 0.91862
2019-03-29 11:39 | 222 | 0.047828 | 0.019466 | 1.8608 | 18.6914 | 35.1117 | 0.89397
2019-03-29 11:39 | 223 | 0.041523 | 0.022383 | 2.1327 | 21.6645 | 35.0663 | 0.89957
2019-03-29 11:39 | 224 | 0.056729 | 0.016339 | 1.5549 | 15.7689 | 35.0238 | 0.89455
2019-03-29 11:39 | 225 | 0.056214 | 0.016883 | 1.5634 | 15.718 | 34.9569 | 0.88357
2019-03-29 11:39 | 226 | 0.028366 | 0.033366 | 3.086 | 30.8508 | 34.9174 | 0.8751
2019-03-29 11:39 | 227 | 0.04718 | 0.020087 | 1.8578 | 18.6944 | 34.8892 | 0.882
2019-03-29 11:39 | 228 | 0.06503 | 0.01456 | 1.339 | 13.3333 | 34.8228 | 0.86706
2019-03-29 11:39 | 229 | 0.053853 | 0.018016 | 1.6111 | 15.9382 | 34.7664 | 0.85833
2019-03-29 11:39 | 230 | 0.046355 | 0.020903 | 1.8655 | 18.6427 | 34.7147 | 0.86419
2019-03-29 11:39 | 231 | 0.05009 | 0.01936 | 1.7214 | 17.0706 | 34.6707 | 0.85507
2019-03-29 11:39 | 232 | 0.057321 | 0.017329 | 1.5026 | 14.8663 | 34.6135 | 0.85216
2019-03-29 11:39 | 233 | 0.059599 | 0.016309 | 1.4382 | 14.1494 | 34.5575 | 0.84328
2019-03-29 11:39 | 234 | 0.049013 | 0.020361 | 1.7489 | 17.2997 | 34.4974 | 0.84792
2019-03-29 11:39 | 235 | 0.046168 | 0.021549 | 1.8394 | 17.9606 | 34.4574 | 0.8292
2019-03-29 11:39 | 236 | 0.065619 | 0.015198 | 1.2935 | 12.6868 | 34.4006 | 0.83249
2019-03-29 11:39 | 237 | 0.038746 | 0.026361 | 2.1826 | 21.0677 | 34.337 | 0.81629
2019-03-29 11:39 | 238 | 0.015647 | 0.063813 | 5.383 | 52.5491 | 34.3219 | 0.82224
2019-03-29 11:39 | 239 | 0.046656 | 0.021861 | 1.7986 | 17.7343 | 34.2926 | 0.82741
2019-03-29 11:39 | 240 | 0.056882 | 0.018339 | 1.4672 | 14.5262 | 34.2329 | 0.82628
2019-03-29 11:39 | 241 | 0.041395 | 0.02521 | 2.0066 | 19.99 | 34.1852 | 0.82748
2019-03-29 11:39 | 242 | 0.04613 | 0.023113 | 1.7905 | 17.946 | 34.1458 | 0.82786
2019-03-29 11:39 | 243 | 0.053239 | 0.020104 | 1.5594 | 15.4116 | 34.0921 | 0.82049
2019-03-29 11:39 | 244 | 0.051504 | 0.021638 | 1.6016 | 15.259 | 34.0413 | 0.7859
2019-03-29 11:39 | 245 | 0.056859 | 0.019973 | 1.4444 | 13.9077 | 33.9881 | 0.79078
2019-03-29 11:39 | 246 | 0.042799 | 0.026592 | 1.9147 | 18.1388 | 33.9319 | 0.77633
2019-03-29 11:39 | 247 | 0.039373 | 0.030059 | 2.0665 | 19.9746 | 33.8999 | 0.78646
2019-03-29 11:39 | 248 | 0.05293 | 0.022381 | 1.5326 | 14.6244 | 33.8493 | 0.77407
2019-03-29 11:39 | 249 | 0.039564 | 0.029906 | 2.0343 | 19.4222 | 33.7997 | 0.76843
2019-03-29 11:39 | 250 | 0.048629 | 0.024843 | 1.6507 | 15.926 | 33.7653 | 0.77446
2019-03-29 11:39 | 251 | 0.058969 | 0.020438 | 1.3483 | 12.9263 | 33.7025 | 0.76226
2019-03-29 11:39 | 252 | 0.035889 | 0.034223 | 2.2049 | 21.4397 | 33.655 | 0.76945
2019-03-29 11:40 | 253 | 0.039256 | 0.030733 | 2.0074 | 19.3682 | 33.6249 | 0.76031
2019-03-29 11:40 | 254 | 0.046673 | 0.026276 | 1.6769 | 16.3145 | 33.5753 | 0.76146
2019-03-29 11:40 | 255 | 0.034335 | 0.036425 | 2.2708 | 21.8911 | 33.5356 | 0.75164
2019-03-29 11:40 | 256 | 0.04392 | 0.028005 | 1.7676 | 17.0023 | 33.5013 | 0.74674
2019-03-29 11:40 | 257 | 0.065788 | 0.019019 | 1.1759 | 11.2489 | 33.4452 | 0.74003
2019-03-29 11:40 | 258 | 0.051365 | 0.025284 | 1.4987 | 14.4437 | 33.3784 | 0.7419
2019-03-29 11:40 | 259 | 0.041425 | 0.031958 | 1.8547 | 18.1975 | 33.3419 | 0.75384
2019-03-29 11:40 | 260 | 0.047418 | 0.028822 | 1.6086 | 15.5964 | 33.2916 | 0.73955
2019-03-29 11:40 | 261 | 0.04554 | 0.031078 | 1.6699 | 16.3093 | 33.2483 | 0.74272
2019-03-29 11:40 | 262 | 0.054859 | 0.026592 | 1.3764 | 13.5916 | 33.1982 | 0.74562
2019-03-29 11:40 | 263 | 0.054459 | 0.027978 | 1.3767 | 13.5391 | 33.1409 | 0.73733
2019-03-29 11:40 | 264 | 0.047305 | 0.032729 | 1.5815 | 15.2867 | 33.0908 | 0.72313
2019-03-29 11:40 | 265 | 0.056353 | 0.0279 | 1.3236 | 12.6651 | 33.0426 | 0.71372
2019-03-29 11:40 | 266 | 0.059803 | 0.027021 | 1.2457 | 12.015 | 32.9799 | 0.71853
2019-03-29 11:40 | 267 | 0.035871 | 0.045667 | 2.0692 | 20.4771 | 32.9294 | 0.73454
2019-03-29 11:40 | 268 | 0.032903 | 0.051003 | 2.2517 | 22.2815 | 32.903 | 0.73312
2019-03-29 11:40 | 269 | 0.050834 | 0.033953 | 1.4588 | 14.271 | 32.8582 | 0.72546
2019-03-29 11:40 | 270 | 0.053607 | 0.032962 | 1.3793 | 13.3798 | 32.8047 | 0.71726
2019-03-29 11:40 | 271 | 0.049884 | 0.036283 | 1.4788 | 14.3978 | 32.7533 | 0.71822
2019-03-29 11:40 | 272 | 0.038419 | 0.049036 | 1.9199 | 18.6841 | 32.7071 | 0.71782
2019-03-29 11:40 | 273 | 0.028634 | 0.06707 | 2.5605 | 24.9132 | 32.6755 | 0.71337
2019-03-29 11:40 | 274 | 0.044051 | 0.044205 | 1.6664 | 16.2253 | 32.6438 | 0.71474
2019-03-29 11:40 | 275 | 0.04527 | 0.043459 | 1.6108 | 15.935 | 32.5909 | 0.72137
2019-03-29 11:40 | 276 | 0.024257 | 0.08327 | 3.0199 | 29.5898 | 32.5584 | 0.71775
2019-03-29 11:40 | 277 | 0.032616 | 0.062528 | 2.2359 | 22.0984 | 32.5354 | 0.72076
2019-03-29 11:40 | 278 | 0.050307 | 0.04145 | 1.4486 | 14.2202 | 32.4907 | 0.71538
2019-03-29 11:40 | 279 | 0.041829 | 0.050903 | 1.7368 | 17.2724 | 32.4405 | 0.72249
2019-03-29 11:40 | 280 | 0.04959 | 0.043972 | 1.4619 | 14.5355 | 32.4035 | 0.72082
2019-03-29 11:40 | 281 | 0.0618 | 0.035667 | 1.1735 | 11.63 | 32.3401 | 0.71873
2019-03-29 11:40 | 282 | 0.063358 | 0.035826 | 1.1412 | 11.3442 | 32.2821 | 0.71875
2019-03-29 11:40 | 283 | 0.076234 | 0.030073 | 0.94561 | 9.5147 | 32.2109 | 0.72535
2019-03-29 11:40 | 284 | 0.086112 | 0.026769 | 0.83429 | 8.4929 | 32.1303 | 0.73134
2019-03-29 11:40 | 285 | 0.087866 | 0.026423 | 0.81511 | 8.4734 | 32.0406 | 0.74452
2019-03-29 11:40 | 286 | 0.096764 | 0.024113 | 0.73851 | 8.5566 | 31.9532 | 0.82796
2019-03-29 11:40 | 287 | 0.10033 | 0.023689 | 0.72203 | 9.8199 | 31.8483 | 0.98525
2019-03-29 11:40 | 288 | 0.10114 | 0.023818 | 0.74093 | 12.9359 | 31.7529 | 1.3084
2019-03-29 11:40 | 289 | 0.12071 | 0.020641 | 0.65913 | 13.3529 | 31.6422 | 1.6118
2019-03-29 11:40 | 290 | 0.11056 | 0.022846 | 0.84397 | 16.6238 | 31.5184 | 1.8379
2019-03-29 11:40 | 291 | 0.10368 | 0.02493 | 1.0591 | 19.4317 | 31.4205 | 2.0146
2019-03-29 11:40 | 292 | 0.11938 | 0.022334 | 1.0416 | 17.7887 | 31.3061 | 2.1236
2019-03-29 11:40 | 293 | 0.10668 | 0.025698 | 1.2631 | 20.5748 | 31.1885 | 2.1949
2019-03-29 11:40 | 294 | 0.087182 | 0.032449 | 1.6772 | 25.6516 | 31.0943 | 2.2364
2019-03-29 11:40 | 295 | 0.099779 | 0.028902 | 1.5284 | 22.786 | 31.0063 | 2.2736
2019-03-29 11:40 | 296 | 0.10824 | 0.027576 | 1.4661 | 21.1805 | 30.8952 | 2.2926
2019-03-29 11:40 | 297 | 0.10642 | 0.029257 | 1.5471 | 21.8995 | 30.7923 | 2.3305
2019-03-29 11:40 | 298 | 0.11165 | 0.028844 | 1.5245 | 21.1745 | 30.6814 | 2.3642
2019-03-29 11:40 | 299 | 0.097888 | 0.033992 | 1.7924 | 24.8236 | 30.5732 | 2.4299
2019-03-29 11:40 | 300 | 0.10057 | 0.034258 | 1.7959 | 25.328 | 30.4817 | 2.5473
2019-03-29 11:40 | 301 | 0.11285 | 0.03187 | 1.6439 | 23.1847 | 30.3698 | 2.6165
2019-03-29 11:40 | 302 | 0.099632 | 0.037264 | 1.9159 | 26.7104 | 30.2615 | 2.6612
2019-03-29 11:40 | 303 | 0.10869 | 0.035165 | 1.8034 | 24.2846 | 30.1657 | 2.6395
2019-03-29 11:40 | 304 | 0.11574 | 0.033622 | 1.7364 | 22.7009 | 30.0444 | 2.6274
2019-03-29 11:40 | 305 | 0.11101 | 0.035308 | 1.8505 | 23.6806 | 29.9372 | 2.6289
2019-03-29 11:40 | 306 | 0.10664 | 0.036728 | 1.9549 | 24.0662 | 29.8226 | 2.5665
2019-03-29 11:40 | 307 | 0.10441 | 0.037508 | 2.0188 | 23.8402 | 29.7234 | 2.4893
2019-03-29 11:40 | 308 | 0.10199 | 0.038455 | 2.0782 | 23.9082 | 29.6138 | 2.4384
2019-03-29 11:40 | 309 | 0.098371 | 0.039569 | 2.1508 | 24.9126 | 29.5196 | 2.4507
2019-03-29 11:40 | 310 | 0.10285 | 0.037874 | 2.0516 | 23.1626 | 29.4158 | 2.3824
2019-03-29 11:40 | 311 | 0.079913 | 0.048665 | 2.6199 | 29.4015 | 29.3196 | 2.3496
2019-03-29 11:40 | 312 | 0.091206 | 0.042775 | 2.2854 | 25.6142 | 29.2481 | 2.3362
2019-03-29 11:40 | 313 | 0.098962 | 0.03996 | 2.0819 | 23.2292 | 29.1375 | 2.2988
2019-03-29 11:41 | 314 | 0.095262 | 0.041989 | 2.1461 | 24.3419 | 29.0531 | 2.3189
2019-03-29 11:41 | 315 | 0.093118 | 0.043608 | 2.1763 | 24.8609 | 28.9466 | 2.315
2019-03-29 11:41 | 316 | 0.098344 | 0.04169 | 2.0421 | 23.5536 | 28.8646 | 2.3164
2019-03-29 11:41 | 317 | 0.11556 | 0.036032 | 1.7301 | 20.2794 | 28.7476 | 2.3435
2019-03-29 11:41 | 318 | 0.10287 | 0.041036 | 1.9352 | 23.2414 | 28.6406 | 2.3908
2019-03-29 11:41 | 319 | 0.093901 | 0.045341 | 2.1164 | 25.9478 | 28.5412 | 2.4365
2019-03-29 11:41 | 320 | 0.10173 | 0.041958 | 1.9543 | 24.3393 | 28.4479 | 2.4761
2019-03-29 11:41 | 321 | 0.13135 | 0.032643 | 1.5152 | 18.9231 | 28.3329 | 2.4856
2019-03-29 11:41 | 322 | 0.12025 | 0.035787 | 1.6649 | 20.9992 | 28.1951 | 2.5253
2019-03-29 11:41 | 323 | 0.10163 | 0.042228 | 1.9854 | 25.2566 | 28.0946 | 2.5669
2019-03-29 11:41 | 324 | 0.11441 | 0.037081 | 1.7816 | 22.8504 | 27.9845 | 2.6144
2019-03-29 11:41 | 325 | 0.11509 | 0.036589 | 1.785 | 23.3984 | 27.8682 | 2.6929
2019-03-29 11:41 | 326 | 0.10525 | 0.039674 | 1.9803 | 25.546 | 27.7571 | 2.6886
2019-03-29 11:41 | 327 | 0.099581 | 0.041323 | 2.1146 | 27.1208 | 27.6561 | 2.7007
2019-03-29 11:41 | 328 | 0.12191 | 0.03361 | 1.7515 | 22.5244 | 27.5508 | 2.746
2019-03-29 11:41 | 329 | 0.14357 | 0.028728 | 1.5092 | 19.5785 | 27.413 | 2.8109
2019-03-29 11:41 | 330 | 0.12192 | 0.034422 | 1.803 | 23.6845 | 27.2748 | 2.8875
2019-03-29 11:41 | 331 | 0.094675 | 0.045209 | 2.3515 | 31.2571 | 27.1706 | 2.9593
2019-03-29 11:41 | 332 | 0.1067 | 0.040875 | 2.1153 | 27.6546 | 27.0752 | 2.9506
2019-03-29 11:41 | 333 | 0.13831 | 0.031925 | 1.6546 | 21.2979 | 26.9522 | 2.9458
2019-03-29 11:41 | 334 | 0.13143 | 0.03391 | 1.7667 | 22.2046 | 26.8077 | 2.9184
2019-03-29 11:41 | 335 | 0.12128 | 0.037114 | 1.9404 | 23.7353 | 26.6906 | 2.8786
2019-03-29 11:41 | 336 | 0.13828 | 0.032534 | 1.7157 | 20.722 | 26.5593 | 2.8655
2019-03-29 11:41 | 337 | 0.12777 | 0.035048 | 1.8673 | 22.4719 | 26.4206 | 2.8713
2019-03-29 11:41 | 338 | 0.10463 | 0.042561 | 2.2877 | 27.382 | 26.3072 | 2.8649
2019-03-29 11:41 | 339 | 0.099112 | 0.044653 | 2.4211 | 29.2483 | 26.2065 | 2.8988
2019-03-29 11:41 | 340 | 0.11702 | 0.037542 | 2.0512 | 25.0379 | 26.1024 | 2.9299
2019-03-29 11:41 | 341 | 0.13982 | 0.031402 | 1.7233 | 20.9009 | 25.9715 | 2.9224
2019-03-29 11:41 | 342 | 0.12923 | 0.034007 | 1.8766 | 22.6051 | 25.8313 | 2.9212
2019-03-29 11:41 | 343 | 0.10834 | 0.040339 | 2.2488 | 26.9291 | 25.7159 | 2.9175
2019-03-29 11:41 | 344 | 0.11458 | 0.037989 | 2.1366 | 25.3252 | 25.6081 | 2.9017
2019-03-29 11:41 | 345 | 0.12363 | 0.034669 | 1.9831 | 23.5515 | 25.4857 | 2.9117
2019-03-29 11:41 | 346 | 0.11621 | 0.036469 | 2.1086 | 25.1685 | 25.3645 | 2.9248
2019-03-29 11:41 | 347 | 0.12029 | 0.034448 | 2.0347 | 24.4448 | 25.2509 | 2.9404
2019-03-29 11:41 | 348 | 0.12174 | 0.033638 | 2.0095 | 24.4184 | 25.1242 | 2.9726
2019-03-29 11:41 | 349 | 0.12696 | 0.0321 | 1.9323 | 23.6942 | 25.0068 | 3.0081
2019-03-29 11:41 | 350 | 0.12522 | 0.032548 | 1.968 | 24.4358 | 24.8723 | 3.0599
2019-03-29 11:41 | 351 | 0.11103 | 0.036499 | 2.2333 | 27.9766 | 24.7589 | 3.1064
2019-03-29 11:41 | 352 | 0.12416 | 0.032433 | 2.0112 | 25.3546 | 24.6439 | 3.1479
2019-03-29 11:41 | 353 | 0.12961 | 0.031027 | 1.9406 | 24.453 | 24.5122 | 3.1694
2019-03-29 11:41 | 354 | 0.12057 | 0.033304 | 2.1046 | 26.2525 | 24.3884 | 3.1653
2019-03-29 11:41 | 355 | 0.11652 | 0.034655 | 2.1981 | 27.1258 | 24.2701 | 3.1608
2019-03-29 11:41 | 356 | 0.11624 | 0.034838 | 2.2179 | 27.0839 | 24.1541 | 3.1482
2019-03-29 11:41 | 357 | 0.11946 | 0.033953 | 2.1666 | 26.1971 | 24.0363 | 3.1296
2019-03-29 11:41 | 358 | 0.13244 | 0.030819 | 1.9673 | 23.8074 | 23.9133 | 3.1531
2019-03-29 11:41 | 359 | 0.12112 | 0.034033 | 2.1585 | 26.0379 | 23.7771 | 3.1537
2019-03-29 11:41 | 360 | 0.11604 | 0.035599 | 2.2554 | 27.152 | 23.6695 | 3.1506
2019-03-29 11:41 | 361 | 0.13519 | 0.030696 | 1.9404 | 23.4333 | 23.5395 | 3.1679
2019-03-29 11:41 | 362 | 0.12926 | 0.032024 | 2.0301 | 24.8339 | 23.4048 | 3.21
2019-03-29 11:41 | 363 | 0.12064 | 0.034152 | 2.1849 | 27.0837 | 23.2816 | 3.2673
2019-03-29 11:41 | 364 | 0.13827 | 0.029358 | 1.9163 | 23.7661 | 23.1574 | 3.2862
2019-03-29 11:41 | 365 | 0.14135 | 0.028608 | 1.8886 | 23.3647 | 23.0083 | 3.3026
2019-03-29 11:41 | 366 | 0.13162 | 0.030849 | 2.0512 | 25.6598 | 22.8779 | 3.3774
2019-03-29 11:41 | 367 | 0.13653 | 0.029892 | 1.9919 | 24.8364 | 22.7418 | 3.3909
2019-03-29 11:41 | 368 | 0.13979 | 0.029844 | 1.9631 | 24.269 | 22.6049 | 3.3927
2019-03-29 11:41 | 369 | 0.14406 | 0.029231 | 1.9194 | 23.4959 | 22.4625 | 3.3848
2019-03-29 11:41 | 370 | 0.13033 | 0.032733 | 2.1306 | 25.9252 | 22.3212 | 3.3787
2019-03-29 11:41 | 371 | 0.12057 | 0.035635 | 2.3219 | 27.9002 | 22.2007 | 3.3639
2019-03-29 11:41 | 372 | 0.1364 | 0.031604 | 2.055 | 24.279 | 22.0741 | 3.3116
2019-03-29 11:42 | 373 | 0.13258 | 0.032499 | 2.1164 | 24.9617 | 21.9326 | 3.3094
2019-03-29 11:42 | 374 | 0.10928 | 0.039255 | 2.5546 | 30.1131 | 21.8129 | 3.2908
2019-03-29 11:42 | 375 | 0.13442 | 0.031745 | 2.0745 | 24.4591 | 21.7024 | 3.2877
2019-03-29 11:42 | 376 | 0.15547 | 0.027258 | 1.7877 | 21.1241 | 21.5445 | 3.2843
2019-03-29 11:42 | 377 | 0.15116 | 0.027534 | 1.8302 | 21.9449 | 21.3981 | 3.3172
2019-03-29 11:42 | 378 | 0.14317 | 0.02849 | 1.9224 | 23.6921 | 21.2438 | 3.392
2019-03-29 11:42 | 379 | 0.13709 | 0.029257 | 2.0111 | 25.3636 | 21.1107 | 3.4771
2019-03-29 11:42 | 380 | 0.15941 | 0.024747 | 1.7386 | 22.2208 | 20.9631 | 3.5422
2019-03-29 11:42 | 381 | 0.15683 | 0.024631 | 1.7744 | 22.6199 | 20.7977 | 3.5475
2019-03-29 11:42 | 382 | 0.13699 | 0.027757 | 2.0531 | 26.5904 | 20.6531 | 3.6427
2019-03-29 11:42 | 383 | 0.16337 | 0.022852 | 1.738 | 23.04 | 20.513 | 3.7642
2019-03-29 11:42 | 384 | 0.17785 | 0.020611 | 1.6166 | 21.2974 | 20.3293 | 3.7878
2019-03-29 11:42 | 385 | 0.14861 | 0.024413 | 1.9636 | 25.6711 | 20.1679 | 3.815
2019-03-29 11:42 | 386 | 0.14643 | 0.024578 | 2.0184 | 26.0159 | 20.0264 | 3.8096
2019-03-29 11:42 | 387 | 0.15015 | 0.023673 | 1.9919 | 25.467 | 19.8727 | 3.8238
2019-03-29 11:42 | 388 | 0.15624 | 0.022759 | 1.9389 | 24.8256 | 19.7248 | 3.8788
2019-03-29 11:42 | 389 | 0.17675 | 0.020048 | 1.7255 | 22.1419 | 19.5572 | 3.9137
2019-03-29 11:42 | 390 | 0.17087 | 0.020828 | 1.801 | 22.8698 | 19.3774 | 3.9078
2019-03-29 11:42 | 391 | 0.17316 | 0.02027 | 1.7902 | 22.5358 | 19.214 | 3.9023
2019-03-29 11:42 | 392 | 0.18379 | 0.018662 | 1.7032 | 21.4973 | 19.0299 | 3.951
2019-03-29 11:42 | 393 | 0.15416 | 0.021996 | 2.0474 | 26.0752 | 18.8555 | 4.0199
2019-03-29 11:42 | 394 | 0.1541 | 0.021433 | 2.064 | 27.0839 | 18.715 | 4.1736
2019-03-29 11:42 | 395 | 0.16274 | 0.020209 | 1.9714 | 26.0528 | 18.5445 | 4.24
2019-03-29 11:42 | 396 | 0.16323 | 0.019992 | 1.9912 | 25.8925 | 18.3909 | 4.2265
2019-03-29 11:42 | 397 | 0.17665 | 0.018549 | 1.8639 | 23.7074 | 18.2157 | 4.188
2019-03-29 11:42 | 398 | 0.16605 | 0.02007 | 2.0075 | 25.0634 | 18.0433 | 4.1618
2019-03-29 11:42 | 399 | 0.15838 | 0.021463 | 2.1313 | 26.2217 | 17.8832 | 4.1529
2019-03-29 11:42 | 400 | 0.16181 | 0.021257 | 2.0986 | 25.6114 | 17.7239 | 4.1441
2019-03-29 11:42 | 401 | 0.15831 | 0.021932 | 2.1603 | 26.2723 | 17.5612 | 4.1591
2019-03-29 11:42 | 402 | 0.14783 | 0.023472 | 2.3175 | 28.1222 | 17.4084 | 4.1574
2019-03-29 11:42 | 403 | 0.16315 | 0.021077 | 2.0959 | 25.7004 | 17.2593 | 4.1929
2019-03-29 11:42 | 404 | 0.1738 | 0.019672 | 1.9732 | 24.3457 | 17.0829 | 4.2312
2019-03-29 11:42 | 405 | 0.17133 | 0.019692 | 2.0063 | 24.8661 | 16.9151 | 4.2602
2019-03-29 11:42 | 406 | 0.17058 | 0.019406 | 2.0212 | 24.9938 | 16.7404 | 4.2635
2019-03-29 11:42 | 407 | 0.16689 | 0.01933 | 2.0736 | 25.5875 | 16.5745 | 4.2703
2019-03-29 11:42 | 408 | 0.17042 | 0.018506 | 2.0461 | 25.3218 | 16.405 | 4.3153
2019-03-29 11:42 | 409 | 0.16865 | 0.018441 | 2.0764 | 25.8133 | 16.2348 | 4.3535
2019-03-29 11:42 | 410 | 0.16586 | 0.018639 | 2.1247 | 26.3539 | 16.0681 | 4.3711
2019-03-29 11:42 | 411 | 0.16112 | 0.01911 | 2.1912 | 27.2688 | 15.9031 | 4.3936
2019-03-29 11:42 | 412 | 0.17357 | 0.017943 | 2.0422 | 25.504 | 15.7415 | 4.4267
2019-03-29 11:42 | 413 | 0.18765 | 0.016894 | 1.8965 | 23.8597 | 15.5559 | 4.4774
2019-03-29 11:42 | 414 | 0.17291 | 0.018829 | 2.0753 | 25.9656 | 15.3732 | 4.4896
2019-03-29 11:42 | 415 | 0.16844 | 0.019775 | 2.1419 | 26.608 | 15.2081 | 4.482
2019-03-29 11:42 | 416 | 0.1721 | 0.019551 | 2.1055 | 25.9584 | 15.0342 | 4.4673
2019-03-29 11:42 | 417 | 0.16622 | 0.02052 | 2.1952 | 26.8853 | 14.8657 | 4.4689
2019-03-29 11:42 | 418 | 0.17137 | 0.019879 | 2.134 | 25.8916 | 14.6994 | 4.4371
2019-03-29 11:42 | 419 | 0.17238 | 0.019797 | 2.119 | 25.5621 | 14.5241 | 4.4063
2019-03-29 11:42 | 420 | 0.16086 | 0.021391 | 2.2667 | 27.3471 | 14.3576 | 4.3992
2019-03-29 11:42 | 421 | 0.161 | 0.021631 | 2.2676 | 27.2158 | 14.1998 | 4.3818
2019-03-29 11:42 | 422 | 0.16264 | 0.021684 | 2.2475 | 26.9575 | 14.0353 | 4.3844
2019-03-29 11:42 | 423 | 0.15092 | 0.023509 | 2.4265 | 28.8776 | 13.8778 | 4.3581
2019-03-29 11:42 | 424 | 0.14069 | 0.025212 | 2.6021 | 30.6073 | 13.7329 | 4.3062
2019-03-29 11:42 | 425 | 0.15068 | 0.02342 | 2.429 | 28.4257 | 13.5916 | 4.2832
2019-03-29 11:42 | 426 | 0.14811 | 0.023758 | 2.4714 | 28.6112 | 13.4343 | 4.2375
2019-03-29 11:42 | 427 | 0.138 | 0.025455 | 2.6396 | 30.0098 | 13.2971 | 4.1413
2019-03-29 11:42 | 428 | 0.14562 | 0.024186 | 2.4963 | 28.4133 | 13.1542 | 4.1375
2019-03-29 11:42 | 429 | 0.14727 | 0.024117 | 2.4589 | 28.1509 | 13.0064 | 4.1457
2019-03-29 11:42 | 430 | 0.16441 | 0.021908 | 2.1931 | 25.3198 | 12.8562 | 4.1629
2019-03-29 11:42 | 431 | 0.16946 | 0.021511 | 2.1162 | 23.9727 | 12.6808 | 4.0624
2019-03-29 11:42 | 432 | 0.15381 | 0.023797 | 2.3086 | 26.1669 | 12.5225 | 4.0247
2019-03-29 11:42 | 433 | 0.148 | 0.024864 | 2.3786 | 27.0228 | 12.3705 | 3.9995
2019-03-29 11:43 | 434 | 0.1741 | 0.021187 | 1.9984 | 22.8286 | 12.2191 | 3.9745
2019-03-29 11:43 | 435 | 0.17016 | 0.021671 | 2.0273 | 23.0951 | 12.0294 | 3.9299
2019-03-29 11:43 | 436 | 0.14828 | 0.024969 | 2.3036 | 26.3974 | 11.8831 | 3.9143
2019-03-29 11:43 | 437 | 0.16427 | 0.02228 | 2.0632 | 23.9084 | 11.7245 | 3.9274
2019-03-29 11:43 | 438 | 0.15472 | 0.023253 | 2.1746 | 25.8918 | 11.5597 | 4.0059
2019-03-29 11:43 | 439 | 0.15656 | 0.022603 | 2.1393 | 26.2363 | 11.413 | 4.1075
2019-03-29 11:43 | 440 | 0.15312 | 0.022976 | 2.1835 | 27.0346 | 11.2478 | 4.1395
2019-03-29 11:43 | 441 | 0.14247 | 0.024572 | 2.3524 | 29.1523 | 11.1084 | 4.1534
2019-03-29 11:43 | 442 | 0.14299 | 0.024345 | 2.3545 | 29.2011 | 10.9601 | 4.1754
2019-03-29 11:43 | 443 | 0.1486 | 0.023749 | 2.2816 | 27.7599 | 10.821 | 4.1252
2019-03-29 11:43 | 444 | 0.14891 | 0.023995 | 2.2892 | 27.7171 | 10.6641 | 4.1272
2019-03-29 11:43 | 445 | 0.14865 | 0.024278 | 2.296 | 27.0191 | 10.5234 | 4.0163
2019-03-29 11:43 | 446 | 0.15293 | 0.023763 | 2.2342 | 25.2256 | 10.3661 | 3.8576
2019-03-29 11:43 | 447 | 0.14012 | 0.026054 | 2.4289 | 27.245 | 10.2211 | 3.8174
2019-03-29 11:43 | 448 | 0.14917 | 0.02424 | 2.2698 | 25.6445 | 10.0812 | 3.8255
2019-03-29 11:43 | 449 | 0.14572 | 0.024729 | 2.3088 | 26.3694 | 9.9254 | 3.8425
2019-03-29 11:43 | 450 | 0.13819 | 0.025762 | 2.4115 | 27.6821 | 9.7904 | 3.8254
2019-03-29 11:43 | 451 | 0.15652 | 0.022758 | 2.1122 | 24.407 | 9.643 | 3.8201
2019-03-29 11:43 | 452 | 0.16251 | 0.021941 | 2.0229 | 23.5521 | 9.4802 | 3.8275
2019-03-29 11:43 | 453 | 0.15454 | 0.023158 | 2.1135 | 24.5345 | 9.3223 | 3.7916
2019-03-29 11:43 | 454 | 0.13108 | 0.027678 | 2.4863 | 28.5731 | 9.175 | 3.7453
2019-03-29 11:43 | 455 | 0.12345 | 0.029552 | 2.627 | 30.1974 | 9.0559 | 3.7278
2019-03-29 11:43 | 456 | 0.14112 | 0.026027 | 2.2861 | 25.9494 | 8.9217 | 3.662
2019-03-29 11:43 | 457 | 0.14133 | 0.026114 | 2.261 | 25.5454 | 8.7775 | 3.6102
2019-03-29 11:43 | 458 | 0.13309 | 0.027669 | 2.3719 | 27.0893 | 8.6412 | 3.6053
2019-03-29 11:43 | 459 | 0.14047 | 0.02619 | 2.222 | 25.8069 | 8.5078 | 3.625
2019-03-29 11:43 | 460 | 0.14209 | 0.025937 | 2.1831 | 25.7832 | 8.3618 | 3.6635
2019-03-29 11:43 | 461 | 0.12747 | 0.02893 | 2.4202 | 29.1073 | 8.2276 | 3.7103
2019-03-29 11:43 | 462 | 0.12543 | 0.029444 | 2.4578 | 29.9354 | 8.1042 | 3.7549
2019-03-29 11:43 | 463 | 0.12 | 0.030975 | 2.5711 | 31.2727 | 7.9766 | 3.7527
2019-03-29 11:43 | 464 | 0.13311 | 0.02806 | 2.3195 | 28.1346 | 7.8595 | 3.7449
2019-03-29 11:43 | 465 | 0.14854 | 0.025119 | 2.0824 | 25.123 | 7.7101 | 3.7317
2019-03-29 11:43 | 466 | 0.13594 | 0.0273 | 2.2781 | 27.3509 | 7.5688 | 3.718
2019-03-29 11:43 | 467 | 0.14421 | 0 | 0 | 0 | 7.4339 | 3.7103
2019-03-29 11:43 | 468 | 0.14997 | 0 | 0 | 0 | 7.2813 | 3.7036
2019-03-29 11:43 | 469 | 0.12802 | 0 | 0 | 0 | 7.1402 | 3.6833
2019-03-29 11:43 | 470 | 0.13467 | 0 | 0 | 0 | 7.0188 | 3.6468
2019-03-29 11:43 | 471 | 0.14131 | 0 | 0 | 0 | 6.8705 | 3.6063
2019-03-29 11:43 | 472 | 0.13179 | 0 | 0 | 0 | 6.7399 | 3.5511
2019-03-29 11:43 | 473 | 0.13523 | 0 | 0 | 0 | 6.6046 | 3.5318
2019-03-29 11:43 | 474 | 0.12384 | 0 | 0 | 0 | 6.4729 | 3.5425
2019-03-29 11:43 | 475 | 0.11269 | 0 | 0 | 0 | 6.3568 | 3.5412
2019-03-29 11:43 | 476 | 0.11561 | 0 | 0 | 0 | 6.2442 | 3.5273
2019-03-29 11:43 | 477 | 0.1111 | 0 | 0 | 0 | 6.1272 | 3.4684
2019-03-29 11:43 | 478 | 0.10003 | 0 | 0 | 0 | 6.0233 | 3.3171
2019-03-29 11:43 | 479 | 0.10789 | 0 | 0 | 0 | 5.9232 | 3.2465
2019-03-29 11:43 | 480 | 0.094978 | 0 | 0 | 0 | 5.8121 | 3.2039
2019-03-29 11:43 | 481 | 0.084412 | 0 | 0 | 0 | 5.7324 | 3.1124
2019-03-29 11:43 | 482 | 0.10019 | 0 | 0 | 0 | 5.6367 | 3.0576
2019-03-29 11:43 | 483 | 0.1099 | 0 | 0 | 0 | 5.5329 | 2.9621
2019-03-29 11:43 | 484 | 0.11486 | 0 | 0 | 0 | 5.4188 | 2.8598
2019-03-29 11:43 | 485 | 0.10079 | 0 | 0 | 0 | 5.3079 | 2.8167
2019-03-29 11:43 | 486 | 0.098854 | 0 | 0 | 0 | 5.2148 | 2.752
2019-03-29 11:43 | 487 | 0.095299 | 0 | 0 | 0 | 5.1104 | 2.679
2019-03-29 11:43 | 488 | 0.086571 | 0 | 0 | 0 | 5.0252 | 2.6264
2019-03-29 11:43 | 489 | 0.089394 | 0 | 0 | 0 | 4.9347 | 2.6396
2019-03-29 11:43 | 490 | 0.08451 | 0 | 0 | 0 | 4.8477 | 2.6383
2019-03-29 11:43 | 491 | 0.092172 | 0 | 0 | 0 | 4.7633 | 2.581
2019-03-29 11:43 | 492 | 0.080625 | 0 | 0 | 0 | 4.6679 | 2.5089
2019-03-29 11:44 | 493 | 0.070056 | 0 | 0 | 0 | 4.602 | 2.4634
2019-03-29 11:44 | 494 | 0.069747 | 0 | 0 | 0 | 4.525 | 2.3915
2019-03-29 11:44 | 495 | 0.074653 | 0 | 0 | 0 | 4.4613 | 2.3357
2019-03-29 11:44 | 496 | 0.066764 | 0 | 0 | 0 | 4.379 | 2.2224
2019-03-29 11:44 | 497 | 0.049083 | 0 | 0 | 0 | 4.3297 | 2.0968
2019-03-29 11:44 | 498 | 0.061581 | 0 | 0 | 0 | 4.2737 | 2.0502
2019-03-29 11:44 | 499 | 0.066999 | 0 | 0 | 0 | 4.2077 | 1.9915
2019-03-29 11:44 | 500 | 0.066638 | 0 | 0 | 0 | 4.1417 | 1.9687
2019-03-29 11:44 | 501 | 0.051095 | 0 | 0 | 0 | 4.0777 | 1.9101
2019-03-29 11:44 | 502 | 0.062496 | 0 | 0 | 0 | 4.0333 | 1.8648
2019-03-29 11:44 | 503 | 0.070118 | 0 | 0 | 0 | 3.9534 | 1.7952
2019-03-29 11:44 | 504 | 0.060801 | 0 | 0 | 0 | 3.8974 | 1.6835
2019-03-29 11:44 | 505 | 0.052552 | 0 | 0 | 0 | 3.8322 | 1.5895
2019-03-29 11:44 | 506 | 0.036721 | 0 | 0 | 0 | 3.7936 | 1.5579
2019-03-29 11:44 | 507 | 0.041717 | 0 | 0 | 0 | 3.7538 | 1.5133
2019-03-29 11:44 | 508 | 0.043484 | 0 | 0 | 0 | 3.7109 | 1.476
2019-03-29 11:44 | 509 | 0.026821 | 0 | 0 | 0 | 3.671 | 1.4089
2019-03-29 11:44 | 510 | 0.027965 | 0 | 0 | 0 | 3.653 | 1.3356
2019-03-29 11:44 | 511 | 0.040689 | 0 | 0 | 0 | 3.6126 | 1.2637
2019-03-29 11:44 | 512 | 0.021565 | 0 | 0 | 0 | 3.5791 | 1.1917
2019-03-29 11:44 | 513 | 0.010195 | 0 | 0 | 0 | 3.568 | 1.1341
2019-03-29 11:44 | 514 | 0.010138 | 0 | 0 | 0 | 3.5551 | 1.0914
2019-03-29 11:44 | 515 | 0.025673 | 0 | 0 | 0 | 3.5431 | 1.0719
2019-03-29 11:44 | 516 | 0.051494 | 0 | 0 | 0 | 3.5016 | 1.0336
2019-03-29 11:44 | 517 | 0.048733 | 0 | 0 | 0 | 3.4477 | 0.96942
2019-03-29 11:44 | 518 | 0.02529 | 0 | 0 | 0 | 3.4094 | 0.93007
2019-03-29 11:44 | 519 | 0.029723 | 0 | 0 | 0 | 3.3908 | 0.84944
2019-03-29 11:44 | 520 | 0.032415 | 0 | 0 | 0 | 3.3503 | 0.81952
2019-03-29 11:44 | 521 | 0.013374 | 0 | 0 | 0 | 3.3308 | 0.79925
2019-03-29 11:44 | 522 | 0.015902 | 0 | 0 | 0 | 3.3186 | 0.75195
2019-03-29 11:44 | 523 | 0.019822 | 0 | 0 | 0 | 3.2983 | 0.6945
2019-03-29 11:44 | 524 | 0.017187 | 0 | 0 | 0 | 3.2807 | 0.68291
2019-03-29 11:44 | 525 | 0.0084531 | 0 | 0 | 0 | 3.2654 | 0.63593
2019-03-29 11:44 | 526 | 0.0075483 | 0 | 0 | 0 | 3.2617 | 0.60954
2019-03-29 11:44 | 527 | 0.016206 | 0 | 0 | 0 | 3.2479 | 0.57513
2019-03-29 11:44 | 528 | 0.015296 | 0 | 0 | 0 | 3.2315 | 0.53351
2019-03-29 11:44 | 529 | 0.014019 | 0 | 0 | 0 | 3.2176 | 0.52417
2019-03-29 11:44 | 530 | 0.01611 | 0 | 0 | 0 | 3.2027 | 0.48373
2019-03-29 11:44 | 531 | 0.013845 | 0 | 0 | 0 | 3.1861 | 0.45673
2019-03-29 11:44 | 532 | 0.01942 | 0 | 0 | 0 | 3.1732 | 0.43611
2019-03-29 11:44 | 533 | 0.027043 | 0 | 0 | 0 | 3.1473 | 0.40549
2019-03-29 11:44 | 534 | 0.011697 | 0 | 0 | 0 | 3.1249 | 0.3868
2019-03-29 11:44 | 535 | -0.0066462 | 0 | 0 | 0 | 3.1249 | 0.37144
2019-03-29 11:44 | 536 | -0.0096366 | 0 | 0 | 0 | 3.1341 | 0.3705
2019-03-29 11:44 | 537 | -0.0025492 | 0 | 0 | 0 | 3.141 | 0.35169
2019-03-29 11:44 | 538 | 0.009266 | 0 | 0 | 0 | 3.1382 | 0.34484
2019-03-29 11:44 | 539 | 0.0011967 | 0 | 0 | 0 | 3.1274 | 0.32699
2019-03-29 11:44 | 540 | -0.011045 | 0 | 0 | 0 | 3.1366 | 0.31752
2019-03-29 11:44 | 541 | 0.0096377 | 0 | 0 | 0 | 3.142 | 0.29214
2019-03-29 11:44 | 542 | 0.0069851 | 0 | 0 | 0 | 3.1226 | 0.29147
2019-03-29 11:44 | 543 | -0.010034 | 0 | 0 | 0 | 3.1308 | 0.27806
2019-03-29 11:44 | 544 | 0.023703 | 0 | 0 | 0 | 3.1308 | 0.27853
2019-03-29 11:44 | 545 | 0.036644 | 0 | 0 | 0 | 3.088 | 0.27678
2019-03-29 11:44 | 546 | 0.022817 | 0 | 0 | 0 | 3.0644 | 0.25752
2019-03-29 11:44 | 547 | 0.026874 | 0 | 0 | 0 | 3.0392 | 0.23887

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062464 | 0.062685 | 0 | 0.2095 | 2019-03-29 10:39 | Group A | 0.1 | 0.22 | Do Nothing | 0.20949 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025175 | 0.025309 | 0 | 0.0261 | 2019-03-29 10:39 | Group A | -0.01 | 0.1 | Do Nothing | 0.000458 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002731 | 0.002746 | 0 | 0.0011 | 2019-03-29 10:39 | Group A | -0.01 | 0.1 | Do Nothing | -8.3175e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.7471 | 184.5289 | 0 | 100 | 2019-03-29 10:39 | None | 0 | 1000 | Do Nothing | 100.7851 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4621 | 210.394 | 0 | 100 | 2019-03-29 10:39 | None | 0 | 1000 | Do Nothing | 100.7128 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -15.579 | 1447 | 0 |  | 2019-03-29 10:24 | None | -10 | 1000 | Do Nothing | 126.7954 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.0259 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018104 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001698 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.8122 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23073 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4025 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0023803 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0037139 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0035411 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 6.7896e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.00090901 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 75.0596 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3484 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5474 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.3024 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 21.3016 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 31.2314 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 65.1591 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 78.856 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 62.4534 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0078979 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.052 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025012 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.7762 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN