## Cone Summary
Date: | 2019-03-20
Time: | 12:42
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B4S-Test1
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0411 |  | Start Test | 0 | 2019-03-20 12:50
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 6 | 2019-03-20 12:50
Sample Area | 1 | m^2 | Wetting | 143 | 2019-03-20 12:52
Cone Heat Flux | 75 | kW/m^2 | Dripping | 153 | 2019-03-20 12:53
Cone Orientation |  |  | Bottom Ignition | 254 | 2019-03-20 12:54
Sample Grid |  |  | Flame Out | 522 | 2019-03-20 12:58
Sample Thickness | 106 | mm
Sample Initial Mass | 45.2 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 24.1661 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 3.9516 | kW/m2 | Average Heat Release Rate | 2.2216 | kW/m2
Total Heat Release | 1.0657 | MJ/m2 | Average Mass Loss Rate | 0.095717 | g/m2s
Peak Mass Loss Rate | 0.17052 | g/m2s | Effective Heat of Combustion | 23.2108 | MJ/kg
Total Mass Lost | 44.3871 | g | Average Specific Extinction Area | 106.5024 | m2/kg
Time to Ignition | 6 | s | Peak CO Yield | 0.073415 | kg/kg
Time to Peak Heat Release | 396 | s | Peak CO2 Yield | 3.4084 | kg/kg
Time to Flame Out | 522 | s | Average CO Yield | 0.024345 | kg/kg
Average Heat Release Rate | 2.0613 | kW/m2 | Average CO2 Yield | 1.9965 | kg/kg
Average Mass Loss Rate | 0.08572 | g/m2s
Effective Heat of Combustion | 24.0089 | MJ/kg
Dripping | 153 | s | Time to dripping
HRRD | 1.3527 | kW | HRR at dripping time
THRD | 0.23835 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.6215 | kW | AHRR from Ignition to Dripping
Wetting | 143 | s | Time to wetting
HRRW | 1.4189 | kW | HRR at wetting time
THRW | 0.22447 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.5927 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-20 12:42 | 1553100120.1635 | -463 | 0.20949 | -0.001109 | -2.2e-05 | 107.8324 | 109.1773 | 129.5498 | 20.9568 | 0.001821 | 0.00172 | -152.2177 | 0.2387 | 8.4027 | 0.003429 | 0.00462 | 0.004346 | 6.3e-05 | 0.001161 | 0.31166 | 1.2989 | 2.5479 | 101 | 20.67 | 18.248 | 65.3511 | 104.0632 | 62.9183 | 0.004408 | 1.0505 | 0.025518 | 24.2909 | -0.11087 | 0.142 | 0.001999 | 0.15537 | 0.15537 | 0.037566
2019-03-20 12:42 | 1553100121.1635 | -462 | 0.20949 | -0.001109 | -2.2e-05 | 107.8324 | 109.1773 | 129.5498 | 20.9568 | 0.001821 | 0.00172 | -152.2177 | 0.2387 | 8.4027 | 0.003429 | 0.00462 | 0.004346 | 6.3e-05 | 0.001161 | 0.31166 | 1.2989 | 2.5479 | 101 | 20.67 | 18.25 | 65.3511 | 104.0632 | 62.9183 | 0.004408 | 1.0505 | 0.025518 | 24.2909 | -0.11087 | 0.142 | 0.001999 | 0.15537 | 0.15537 | 0.037566
2019-03-20 12:42 | 1553100122.1645 | -461 | 0.20951 | -0.001115 | -2.2e-05 | 107.888 | 109.1508 | 129.5498 | 20.9578 | 0.001884 | 0.001736 | -152.2177 | 0.2389 | 8.4029 | 0.003487 | 0.004701 | 0.004455 | 6.1e-05 | 0.001167 | 0.30917 | 1.3511 | 2.5479 | 101 | 20.67 | 18.254 | 66.8542 | 105.4751 | 63.3143 | 0.004409 | 1.0493 | 0.025503 | 24.3052 | -0.10408 | 0.14212 | 0.002007 | 0.15589 | 0.15589 | 0.037469
2019-03-20 12:42 | 1553100123.1651 | -460 | 0.20949 | -0.001111 | -2.2e-05 | 107.8754 | 109.0042 | 129.5498 | 20.9575 | 0.001871 | 0.001701 | -152.2177 | 0.23914 | 8.406 | 0.003508 | 0.004913 | 0.004813 | 6.2e-05 | 0.001167 | 0.33483 | 1.3482 | 2.548 | 101 | 20.67 | 18.256 | 66.5543 | 105.9895 | 62.4728 | 0.00441 | 1.0519 | 0.025535 | 24.2748 | -0.09311 | 0.14226 | 0.002002 | 0.15568 | 0.15568 | 0.037558
2019-03-20 12:42 | 1553100124.1636 | -459 | 0.20958 | -0.001114 | -2.2e-05 | 107.8246 | 109.1586 | 129.5498 | 20.9584 | 0.001858 | 0.001724 | -152.2177 | 0.23853 | 8.4009 | 0.003522 | 0.004755 | 0.004585 | 6e-05 | 0.001168 | 0.33859 | 1.339 | 2.548 | 101 | 20.67 | 18.256 | 66.2298 | 106.3186 | 63.0041 | 0.00441 | 1.0502 | 0.025515 | 24.294 | -0.10998 | 0.1419 | 0.002005 | 0.15581 | 0.15581 | 0.037431
2019-03-20 12:42 | 1553100125.1637 | -458 | 0.20955 | -0.001113 | -2.2e-05 | 107.8159 | 109.1179 | 129.5498 | 20.9579 | 0.001829 | 0.001772 | -152.2177 | 0.23887 | 8.4014 | 0.003591 | 0.004677 | 0.004457 | 5.9e-05 | 0.001162 | 0.32084 | 1.3443 | 2.5479 | 101 | 20.67 | 18.258 | 65.5433 | 107.9944 | 64.1601 | 0.00441 | 1.0466 | 0.025471 | 24.3357 | -0.10736 | 0.14211 | 0.002005 | 0.15553 | 0.15553 | 0.037552
2019-03-20 12:42 | 1553100126.1651 | -457 | 0.20952 | -0.001111 | -2.2e-05 | 107.9049 | 109.2319 | 129.5498 | 20.9586 | 0.001848 | 0.001741 | -152.2177 | 0.23839 | 8.4029 | 0.003553 | 0.004697 | 0.004485 | 6.1e-05 | 0.001162 | 0.34938 | 1.3312 | 2.548 | 101 | 20.67 | 18.258 | 66.0046 | 107.0775 | 63.4229 | 0.00441 | 1.0489 | 0.025499 | 24.3091 | -0.10933 | 0.14182 | 0.002002 | 0.15549 | 0.15549 | 0.037486
2019-03-20 12:42 | 1553100127.1638 | -456 | 0.20954 | -0.001114 | -2.2e-05 | 107.9236 | 109.384 | 129.5498 | 20.9579 | 0.001914 | 0.001751 | -152.2177 | 0.23792 | 8.4017 | 0.003492 | 0.004171 | 0.00395 | 6e-05 | 0.001166 | 0.3366 | 1.3626 | 2.548 | 101 | 20.67 | 18.257 | 67.5745 | 105.5916 | 63.6605 | 0.00441 | 1.0482 | 0.02549 | 24.3177 | -0.12022 | 0.14154 | 0.002005 | 0.15566 | 0.15566 | 0.037373
2019-03-20 12:42 | 1553100128.165 | -455 | 0.20958 | -0.001115 | -2.2e-05 | 107.8848 | 109.1549 | 129.5498 | 20.9582 | 0.001828 | 0.001742 | -152.2177 | 0.23813 | 8.4007 | 0.003538 | 0.004891 | 0.004666 | 6.1e-05 | 0.001169 | 0.35651 | 1.381 | 2.548 | 101 | 20.67 | 18.255 | 65.5094 | 106.7056 | 63.4398 | 0.00441 | 1.0489 | 0.025498 | 24.3097 | -0.10468 | 0.14166 | 0.002007 | 0.15594 | 0.15594 | 0.037336
2019-03-20 12:42 | 1553100129.1639 | -454 | 0.20953 | -0.001107 | -2.2e-05 | 107.8683 | 108.9885 | 129.5498 | 20.9589 | 0.001892 | 0.001737 | -152.2177 | 0.23908 | 8.4 | 0.003497 | 0.005027 | 0.004792 | 6e-05 | 0.001166 | 0.30021 | 1.3508 | 2.5481 | 101 | 20.67 | 18.251 | 67.0527 | 105.7279 | 63.3379 | 0.004409 | 1.0492 | 0.025502 | 24.3061 | -0.092412 | 0.14223 | 0.001997 | 0.15509 | 0.15509 | 0.037693
2019-03-20 12:42 | 1553100130.1644 | -453 | 0.2095 | -0.001109 | -2.3e-05 | 107.8463 | 109.2229 | 129.5498 | 20.9587 | 0.001793 | 0.001714 | -152.2177 | 0.23869 | 8.4014 | 0.003502 | 0.004673 | 0.004436 | 5.9e-05 | 0.001162 | 0.31116 | 1.3433 | 2.548 | 101 | 20.67 | 18.24 | 64.6715 | 105.8315 | 62.7723 | 0.004406 | 1.051 | 0.025524 | 24.2856 | -0.11345 | 0.142 | 0.002 | 0.15546 | 0.15546 | 0.037541
2019-03-20 12:42 | 1553100131.1641 | -452 | 0.20956 | -0.001111 | -2.3e-05 | 107.8874 | 109.3115 | 129.6287 | 20.9597 | 0.001902 | 0.001722 | -152.211 | 0.23817 | 8.405 | 0.003456 | 0.0047 | 0.004489 | 6.3e-05 | 0.001162 | 0.34092 | 1.3457 | 2.5481 | 101 | 20.67 | 18.24 | 67.311 | 104.7288 | 62.9733 | 0.004406 | 1.0503 | 0.025524 | 24.3003 | -0.11729 | 0.14169 | 0.002003 | 0.15572 | 0.15572 | 0.037396
2019-03-20 12:42 | 1553100132.1631 | -451 | 0.20956 | -0.001113 | -2.3e-05 | 107.8952 | 109.1123 | 129.8899 | 20.9584 | 0.001801 | 0.001744 | -152.2314 | 0.23843 | 8.4006 | 0.00347 | 0.005137 | 0.004941 | 6.1e-05 | 0.001154 | 0.34916 | 1.3304 | 2.5479 | 101 | 20.67 | 18.222 | 64.8712 | 105.073 | 63.486 | 0.004402 | 1.0487 | 0.02553 | 24.3433 | -0.10033 | 0.14184 | 0.002006 | 0.15603 | 0.15603 | 0.037363
2019-03-20 12:42 | 1553100133.1644 | -450 | 0.20955 | -0.001105 | -2.3e-05 | 107.8738 | 109.2386 | 129.9489 | 20.9578 | 0.001889 | 0.001722 | -152.2221 | 0.23821 | 8.4026 | 0.003504 | 0.004719 | 0.004483 | 5.6e-05 | 0.001149 | 0.33948 | 1.3297 | 2.548 | 101 | 20.67 | 18.211 | 66.9916 | 105.8857 | 62.9739 | 0.004399 | 1.0503 | 0.025555 | 24.3303 | -0.11245 | 0.14171 | 0.001996 | 0.15546 | 0.15546 | 0.037466
2019-03-20 12:42 | 1553100134.1632 | -449 | 0.20953 | -0.001109 | -2.3e-05 | 107.8563 | 109.0622 | 130.165 | 20.9589 | 0.001893 | 0.001752 | -152.2024 | 0.23873 | 8.4025 | 0.003453 | 0.005128 | 0.004955 | 5.9e-05 | 0.001146 | 0.33483 | 1.3093 | 2.5479 | 101 | 20.67 | 18.186 | 67.0917 | 104.6565 | 63.6802 | 0.004393 | 1.0481 | 0.02555 | 24.3761 | -0.099445 | 0.14202 | 0.002 | 0.15574 | 0.15574 | 0.03748
2019-03-20 12:42 | 1553100135.1642 | -448 | 0.20958 | -0.001116 | -2.2e-05 | 107.9372 | 109.3907 | 130.4657 | 20.9596 | 0.001912 | 0.001698 | -152.1976 | 0.23888 | 8.4014 | 0.003348 | 0.005163 | 0.004995 | 6.1e-05 | 0.001152 | 0.3027 | 1.3045 | 2.548 | 101 | 20.67 | 18.172 | 67.5442 | 102.1141 | 62.3851 | 0.004389 | 1.0522 | 0.025628 | 24.3573 | -0.11965 | 0.14211 | 0.002008 | 0.15674 | 0.15674 | 0.037263
2019-03-20 12:42 | 1553100136.1633 | -447 | 0.20956 | -0.001115 | -2.2e-05 | 107.9223 | 109.193 | 130.3401 | 20.9586 | 0.001806 | 0.001696 | -152.1939 | 0.238 | 8.4028 | 0.003458 | 0.004997 | 0.004729 | 5.8e-05 | 0.001154 | 0.35546 | 1.3238 | 2.5479 | 101 | 20.67 | 18.158 | 64.9969 | 104.7876 | 62.3472 | 0.004386 | 1.0523 | 0.025617 | 24.3442 | -0.1047 | 0.14158 | 0.002007 | 0.1566 | 0.1566 | 0.037159
2019-03-20 12:42 | 1553100137.1648 | -446 | 0.20949 | -0.001114 | -2.2e-05 | 107.8935 | 109.1926 | 130.2758 | 20.9594 | 0.001902 | 0.001716 | -152.1875 | 0.23755 | 8.4035 | 0.003488 | 0.004956 | 0.004707 | 5.8e-05 | 0.00116 | 0.36768 | 1.3051 | 2.548 | 101 | 20.67 | 18.128 | 67.3029 | 105.5083 | 62.8286 | 0.004379 | 1.0508 | 0.025593 | 24.3556 | -0.10705 | 0.14132 | 0.002006 | 0.15635 | 0.15635 | 0.037147
2019-03-20 12:42 | 1553100138.1634 | -445 | 0.20953 | -0.00111 | -2.3e-05 | 107.8926 | 109.2162 | 130.3319 | 20.9595 | 0.001833 | 0.001769 | -152.1909 | 0.23984 | 8.4033 | 0.003488 | 0.005304 | 0.005106 | 5.9e-05 | 0.001161 | 0.31868 | 1.2975 | 2.5479 | 101 | 20.67 | 18.112 | 65.6363 | 105.5083 | 64.089 | 0.004375 | 1.0469 | 0.02555 | 24.4065 | -0.10906 | 0.14268 | 0.002001 | 0.15576 | 0.15576 | 0.037649
2019-03-20 12:42 | 1553100139.1645 | -444 | 0.20952 | -0.001112 | -2.2e-05 | 107.8728 | 109.0779 | 129.9437 | 20.9605 | 0.001836 | 0.001699 | -152.1924 | 0.23882 | 8.4031 | 0.003529 | 0.004565 | 0.004477 | 5.9e-05 | 0.001167 | 0.34058 | 1.3291 | 2.5481 | 101 | 20.67 | 18.087 | 65.7082 | 106.4914 | 62.4218 | 0.004369 | 1.0521 | 0.025576 | 24.3098 | -0.099365 | 0.14208 | 0.002003 | 0.15603 | 0.15603 | 0.037425
2019-03-20 12:42 | 1553100140.1649 | -443 | 0.20957 | -0.001114 | -2.2e-05 | 107.9201 | 109.0551 | 129.4902 | 20.9606 | 0.001826 | 0.001748 | -152.1964 | 0.23882 | 8.4021 | 0.00349 | 0.00485 | 0.004589 | 6.3e-05 | 0.00117 | 0.32388 | 1.3385 | 2.548 | 101 | 20.67 | 18.078 | 65.4766 | 105.5525 | 63.5911 | 0.004367 | 1.0484 | 0.025487 | 24.3096 | -0.09358 | 0.14208 | 0.002005 | 0.1557 | 0.1557 | 0.037503
2019-03-20 12:42 | 1553100141.1636 | -442 | 0.20952 | -0.001111 | -2.2e-05 | 107.8948 | 109.0764 | 129.4909 | 20.9601 | 0.001867 | 0.001765 | -152.1969 | 0.23867 | 8.402 | 0.003578 | 0.005091 | 0.004881 | 6e-05 | 0.001176 | 0.3115 | 1.3027 | 2.5479 | 101 | 20.67 | 18.071 | 66.4671 | 107.6792 | 63.9952 | 0.004365 | 1.0472 | 0.025471 | 24.3242 | -0.097419 | 0.14199 | 0.002002 | 0.15526 | 0.15526 | 0.037586
2019-03-20 12:42 | 1553100142.1646 | -441 | 0.20955 | -0.001115 | -2.2e-05 | 107.934 | 109.1646 | 129.544 | 20.9601 | 0.001807 | 0.001731 | -152.182 | 0.2395 | 8.4021 | 0.003521 | 0.004905 | 0.004772 | 6.3e-05 | 0.001174 | 0.34164 | 1.317 | 2.548 | 101 | 20.67 | 18.063 | 65.0164 | 106.3052 | 63.1916 | 0.004363 | 1.0497 | 0.025507 | 24.3002 | -0.1014 | 0.14248 | 0.002006 | 0.15586 | 0.15586 | 0.03757
2019-03-20 12:42 | 1553100143.1637 | -440 | 0.2096 | -0.001115 | -2.2e-05 | 107.9146 | 109.3963 | 129.5017 | 20.9607 | 0.001942 | 0.001705 | -152.1824 | 0.23889 | 8.4045 | 0.003456 | 0.00558 | 0.005356 | 5.8e-05 | 0.00117 | 0.35347 | 1.316 | 2.548 | 101 | 20.67 | 18.062 | 68.2675 | 104.7378 | 62.5595 | 0.004363 | 1.0516 | 0.025527 | 24.2734 | -0.12198 | 0.14212 | 0.002008 | 0.15614 | 0.15614 | 0.037409
2019-03-20 12:42 | 1553100144.1646 | -439 | 0.20956 | -0.001114 | -2.2e-05 | 107.944 | 109.2513 | 129.6371 | 20.9604 | 0.001792 | 0.001758 | -152.2032 | 0.23856 | 8.4026 | 0.003562 | 0.004763 | 0.004496 | 5.7e-05 | 0.001163 | 0.30845 | 1.3065 | 2.5479 | 101 | 20.67 | 18.066 | 64.6614 | 107.2999 | 63.844 | 0.004364 | 1.0476 | 0.025492 | 24.3325 | -0.10768 | 0.14192 | 0.002005 | 0.15571 | 0.15571 | 0.03746
2019-03-20 12:42 | 1553100145.1631 | -438 | 0.20954 | -0.001117 | -2.2e-05 | 107.9201 | 109.2842 | 129.6886 | 20.9601 | 0.001873 | 0.001786 | -152.2161 | 0.23835 | 8.4025 | 0.003605 | 0.004748 | 0.004636 | 5.7e-05 | 0.001161 | 0.33107 | 1.3392 | 2.548 | 101 | 20.67 | 18.072 | 66.6006 | 108.3424 | 64.5081 | 0.004365 | 1.0456 | 0.025471 | 24.3613 | -0.11235 | 0.1418 | 0.002009 | 0.15588 | 0.15588 | 0.037386
2019-03-20 12:42 | 1553100146.1639 | -437 | 0.20956 | -0.001116 | -2.3e-05 | 107.965 | 108.9231 | 129.6808 | 20.9617 | 0.001866 | 0.001757 | -152.2109 | 0.23869 | 8.4021 | 0.003602 | 0.004916 | 0.004743 | 5.6e-05 | 0.001158 | 0.34307 | 1.38 | 2.5479 | 101 | 20.67 | 18.079 | 66.4329 | 108.2755 | 63.8062 | 0.004367 | 1.0477 | 0.025497 | 24.3353 | -0.079028 | 0.14199 | 0.002008 | 0.15603 | 0.15603 | 0.037404
2019-03-20 12:42 | 1553100147.1629 | -436 | 0.20956 | -0.001114 | -2.2e-05 | 107.8968 | 109.1863 | 129.7578 | 20.9605 | 0.001894 | 0.001735 | -152.2232 | 0.23889 | 8.4051 | 0.003567 | 0.005189 | 0.004962 | 5.8e-05 | 0.001153 | 0.28335 | 1.3077 | 2.5479 | 101 | 20.67 | 18.095 | 67.1135 | 107.4284 | 63.2842 | 0.004371 | 1.0494 | 0.025525 | 24.3236 | -0.10626 | 0.14212 | 0.002006 | 0.15596 | 0.15596 | 0.037452
2019-03-20 12:42 | 1553100148.1639 | -435 | 0.20956 | -0.001115 | -2.2e-05 | 107.9298 | 109.1156 | 129.8329 | 20.9608 | 0.001816 | 0.001775 | -152.2332 | 0.23915 | 8.4035 | 0.003484 | 0.005686 | 0.00544 | 6e-05 | 0.001152 | 0.32444 | 1.3034 | 2.548 | 101 | 20.67 | 18.104 | 65.2216 | 105.3971 | 64.2566 | 0.004373 | 1.0463 | 0.025495 | 24.3658 | -0.097741 | 0.14227 | 0.002007 | 0.15586 | 0.15586 | 0.037516
2019-03-20 12:42 | 1553100149.164 | -434 | 0.20954 | -0.001118 | -2.3e-05 | 107.9084 | 109.2068 | 130.3183 | 20.9611 | 0.001898 | 0.001776 | -152.2308 | 0.2387 | 8.4021 | 0.003502 | 0.005571 | 0.005325 | 6.5e-05 | 0.001157 | 0.32156 | 1.3316 | 2.548 | 101 | 20.67 | 18.114 | 67.2039 | 105.8363 | 64.2655 | 0.004375 | 1.0463 | 0.025542 | 24.4116 | -0.10698 | 0.142 | 0.002012 | 0.15653 | 0.15653 | 0.037286
2019-03-20 12:42 | 1553100150.163 | -433 | 0.20957 | -0.001123 | -2.2e-05 | 107.9171 | 109.3447 | 130.6162 | 20.9608 | 0.00193 | 0.001698 | -152.237 | 0.2388 | 8.4001 | 0.003472 | 0.005196 | 0.005016 | 6.3e-05 | 0.001154 | 0.33356 | 1.3624 | 2.548 | 101 | 20.67 | 18.134 | 67.9767 | 105.1234 | 62.3932 | 0.00438 | 1.0522 | 0.025643 | 24.3716 | -0.11755 | 0.14206 | 0.002016 | 0.15748 | 0.15748 | 0.037076
2019-03-20 12:42 | 1553100151.1635 | -432 | 0.20954 | -0.001113 | -2.2e-05 | 107.9217 | 109.1048 | 130.7744 | 20.9609 | 0.001859 | 0.00178 | -152.2312 | 0.23914 | 8.4038 | 0.003504 | 0.005233 | 0.005038 | 6.3e-05 | 0.001156 | 0.34468 | 1.3128 | 2.548 | 101 | 20.67 | 18.143 | 66.2705 | 105.8919 | 64.3707 | 0.004382 | 1.046 | 0.025583 | 24.4581 | -0.097523 | 0.14226 | 0.002004 | 0.15614 | 0.15614 | 0.037447
2019-03-20 12:43 | 1553100152.1632 | -431 | 0.20951 | -0.001112 | -2.2e-05 | 107.9353 | 109.1164 | 130.5181 | 20.9608 | 0.001848 | 0.001658 | -152.2249 | 0.23924 | 8.4016 | 0.003514 | 0.004824 | 0.004565 | 6.2e-05 | 0.001157 | 0.30901 | 1.333 | 2.5481 | 101 | 20.67 | 18.159 | 66.0012 | 106.1343 | 61.4376 | 0.004386 | 1.0552 | 0.02567 | 24.3277 | -0.097347 | 0.14232 | 0.002002 | 0.15652 | 0.15652 | 0.037372
2019-03-20 12:43 | 1553100153.1636 | -430 | 0.2095 | -0.001116 | -2.2e-05 | 107.9505 | 109.1429 | 130.538 | 20.9616 | 0.001883 | 0.001738 | -152.2298 | 0.23896 | 8.4043 | 0.003494 | 0.004275 | 0.004046 | 5.9e-05 | 0.001158 | 0.32637 | 1.3332 | 2.548 | 101 | 20.67 | 18.166 | 66.8486 | 105.6612 | 63.3461 | 0.004388 | 1.0492 | 0.025599 | 24.3989 | -0.098262 | 0.14216 | 0.002007 | 0.15645 | 0.15645 | 0.037346
2019-03-20 12:43 | 1553100154.1633 | -429 | 0.20951 | -0.001111 | -2.2e-05 | 107.9006 | 109.073 | 130.4746 | 20.9618 | 0.001826 | 0.001763 | -152.2353 | 0.23891 | 8.405 | 0.003461 | 0.00465 | 0.004402 | 6.6e-05 | 0.001157 | 0.32587 | 1.3015 | 2.548 | 101 | 20.67 | 18.172 | 65.4824 | 104.8402 | 63.9694 | 0.004389 | 1.0472 | 0.025569 | 24.4155 | -0.09666 | 0.14213 | 0.002002 | 0.15586 | 0.15586 | 0.037479
2019-03-20 12:43 | 1553100155.1643 | -428 | 0.20954 | -0.001113 | -2.2e-05 | 107.8893 | 109.1306 | 130.3859 | 20.9628 | 0.001853 | 0.001783 | -152.2183 | 0.23967 | 8.4027 | 0.003471 | 0.004639 | 0.004368 | 6.3e-05 | 0.001163 | 0.34324 | 1.2839 | 2.548 | 101 | 20.67 | 18.181 | 66.1335 | 105.1039 | 64.4341 | 0.004392 | 1.0458 | 0.025543 | 24.424 | -0.10232 | 0.14258 | 0.002004 | 0.15589 | 0.15589 | 0.037591
2019-03-20 12:43 | 1553100156.1638 | -427 | 0.20954 | -0.001107 | -2.2e-05 | 107.944 | 109.3432 | 130.6141 | 20.9618 | 0.001858 | 0.001693 | -152.2248 | 0.2383 | 8.4014 | 0.003428 | 0.004594 | 0.004414 | 6.8e-05 | 0.001163 | 0.35081 | 1.2832 | 2.548 | 101 | 20.67 | 18.186 | 66.2473 | 104.0515 | 62.2699 | 0.004393 | 1.0525 | 0.025647 | 24.3669 | -0.1152 | 0.14177 | 0.001996 | 0.15587 | 0.15587 | 0.037382
2019-03-20 12:43 | 1553100157.1634 | -426 | 0.2095 | -0.001107 | -2.2e-05 | 107.976 | 109.1108 | 130.4176 | 20.9615 | 0.001922 | 0.001689 | -152.215 | 0.23983 | 8.4043 | 0.003522 | 0.00445 | 0.004246 | 6.2e-05 | 0.001167 | 0.32516 | 1.2924 | 2.5481 | 101 | 20.67 | 18.193 | 67.7723 | 106.3255 | 62.1829 | 0.004395 | 1.0528 | 0.025631 | 24.3454 | -0.093511 | 0.14267 | 0.001996 | 0.15585 | 0.15585 | 0.037627
2019-03-20 12:43 | 1553100158.1635 | -425 | 0.20956 | -0.001106 | -2.2e-05 | 107.9527 | 109.141 | 130.0047 | 20.9611 | 0.00186 | 0.001733 | -152.1876 | 0.23901 | 8.4039 | 0.003421 | 0.004735 | 0.004597 | 6.5e-05 | 0.001168 | 0.31885 | 1.3167 | 2.548 | 101 | 20.67 | 18.196 | 66.2943 | 103.8789 | 63.2419 | 0.004395 | 1.0495 | 0.02555 | 24.3452 | -0.097921 | 0.14219 | 0.001995 | 0.1553 | 0.1553 | 0.03763
2019-03-20 12:43 | 1553100159.1645 | -424 | 0.20952 | -0.001108 | -2.2e-05 | 107.9614 | 109.2128 | 129.5315 | 20.9619 | 0.001862 | 0.001732 | -152.1781 | 0.23873 | 8.4042 | 0.003379 | 0.004937 | 0.004716 | 6.5e-05 | 0.001173 | 0.34899 | 1.3242 | 2.548 | 101 | 20.67 | 18.199 | 66.3315 | 102.8744 | 63.2034 | 0.004396 | 1.0496 | 0.025505 | 24.2995 | -0.10308 | 0.14202 | 0.001998 | 0.15519 | 0.15519 | 0.037611
2019-03-20 12:43 | 1553100160.1638 | -423 | 0.20948 | -0.00111 | -2.2e-05 | 107.9391 | 109.2322 | 129.3292 | 20.9618 | 0.001843 | 0.001739 | -152.18 | 0.23859 | 8.4024 | 0.003317 | 0.004958 | 0.004748 | 6.2e-05 | 0.001171 | 0.34363 | 1.3353 | 2.5481 | 101 | 20.67 | 18.203 | 65.8788 | 101.3735 | 63.3916 | 0.004397 | 1.049 | 0.025478 | 24.2873 | -0.10652 | 0.14194 | 0.002001 | 0.15523 | 0.15523 | 0.037582
2019-03-20 12:43 | 1553100161.1635 | -422 | 0.20953 | -0.001105 | -2.2e-05 | 107.9763 | 109.1186 | 128.7522 | 20.9615 | 0.001814 | 0.00177 | -152.1827 | 0.23805 | 8.4028 | 0.003354 | 0.00492 | 0.004716 | 6.3e-05 | 0.001171 | 0.33373 | 1.3182 | 2.548 | 101 | 20.67 | 18.204 | 65.1903 | 102.2518 | 64.1348 | 0.004397 | 1.0467 | 0.025393 | 24.2598 | -0.094127 | 0.14162 | 0.001995 | 0.15428 | 0.15428 | 0.037726
2019-03-20 12:43 | 1553100162.1648 | -421 | 0.20958 | -0.001107 | -2.3e-05 | 107.9194 | 109.1657 | 128.8227 | 20.9611 | 0.00183 | 0.00173 | -152.1937 | 0.23943 | 8.4026 | 0.003326 | 0.004826 | 0.004681 | 6.6e-05 | 0.00117 | 0.30469 | 1.3441 | 2.5481 | 101 | 20.67 | 18.205 | 65.5613 | 101.5792 | 63.1738 | 0.004397 | 1.0497 | 0.025437 | 24.2318 | -0.1027 | 0.14244 | 0.001997 | 0.15477 | 0.15477 | 0.037826
2019-03-20 12:43 | 1553100163.1637 | -420 | 0.20953 | -0.001107 | -2.2e-05 | 107.9873 | 109.1433 | 128.7788 | 20.9619 | 0.001843 | 0.001734 | -152.194 | 0.23873 | 8.4037 | 0.00332 | 0.00456 | 0.00432 | 6.5e-05 | 0.001171 | 0.32139 | 1.3173 | 2.5479 | 101 | 20.67 | 18.203 | 65.8782 | 101.4474 | 63.2598 | 0.004397 | 1.0494 | 0.025429 | 24.2308 | -0.095238 | 0.14202 | 0.001997 | 0.15465 | 0.15465 | 0.037744
2019-03-20 12:43 | 1553100164.1631 | -419 | 0.20954 | -0.001105 | -2.2e-05 | 107.899 | 109.2644 | 128.8142 | 20.961 | 0.001885 | 0.001741 | -152.1916 | 0.23917 | 8.403 | 0.003333 | 0.005253 | 0.005 | 6e-05 | 0.001167 | 0.3126 | 1.3499 | 2.548 | 101 | 20.67 | 18.201 | 66.9004 | 101.7589 | 63.4377 | 0.004397 | 1.0489 | 0.025426 | 24.2405 | -0.11247 | 0.14228 | 0.001994 | 0.15442 | 0.15442 | 0.037869
2019-03-20 12:43 | 1553100165.1629 | -418 | 0.20958 | -0.001112 | -2.3e-05 | 108.0113 | 109.2539 | 128.7137 | 20.9619 | 0.001931 | 0.001715 | -152.2143 | 0.23853 | 8.4017 | 0.003372 | 0.004968 | 0.004805 | 6.2e-05 | 0.001165 | 0.34036 | 1.3142 | 2.5481 | 101 | 20.67 | 18.195 | 67.9996 | 102.6976 | 62.8029 | 0.004395 | 1.0509 | 0.02544 | 24.2082 | -0.10232 | 0.14191 | 0.002004 | 0.15534 | 0.15534 | 0.037545
2019-03-20 12:43 | 1553100166.1639 | -417 | 0.20951 | -0.001112 | -2.2e-05 | 107.9679 | 109.344 | 128.5572 | 20.9629 | 0.001811 | 0.001736 | -152.2238 | 0.23915 | 8.4044 | 0.003389 | 0.004882 | 0.004611 | 6.3e-05 | 0.001165 | 0.3063 | 1.3507 | 2.5481 | 101 | 20.67 | 18.191 | 65.1173 | 103.1127 | 63.3154 | 0.004394 | 1.0493 | 0.025405 | 24.2119 | -0.11328 | 0.14227 | 0.002003 | 0.15498 | 0.15498 | 0.037731
2019-03-20 12:43 | 1553100167.1643 | -416 | 0.20954 | -0.001116 | -2.2e-05 | 107.9815 | 109.1268 | 128.7595 | 20.9629 | 0.001868 | 0.001768 | -152.2213 | 0.23902 | 8.4031 | 0.003369 | 0.005194 | 0.00494 | 6.5e-05 | 0.001164 | 0.3741 | 1.3103 | 2.5481 | 101 | 20.67 | 18.185 | 66.4863 | 102.628 | 64.0871 | 0.004393 | 1.0469 | 0.025396 | 24.2588 | -0.094373 | 0.14219 | 0.002008 | 0.15537 | 0.15537 | 0.037615
2019-03-20 12:43 | 1553100168.163 | -415 | 0.20951 | -0.001111 | -2.2e-05 | 107.9786 | 109.2846 | 129.0041 | 20.963 | 0.00187 | 0.001697 | -152.2295 | 0.23894 | 8.4011 | 0.003348 | 0.004993 | 0.004792 | 6.3e-05 | 0.001158 | 0.33301 | 1.3378 | 2.5481 | 101.01 | 20.67 | 18.17 | 66.5315 | 102.1066 | 62.3689 | 0.004389 | 1.0522 | 0.025485 | 24.2199 | -0.10753 | 0.14214 | 0.002001 | 0.15526 | 0.15526 | 0.037627
2019-03-20 12:43 | 1553100169.1641 | -414 | 0.20955 | -0.001116 | -2.2e-05 | 107.9466 | 109.1679 | 129.2806 | 20.9629 | 0.001871 | 0.001736 | -152.2395 | 0.2397 | 8.404 | 0.003443 | 0.004676 | 0.004491 | 6.4e-05 | 0.001161 | 0.33124 | 1.3729 | 2.548 | 101.01 | 20.67 | 18.161 | 66.5518 | 104.4285 | 63.3108 | 0.004386 | 1.0493 | 0.025477 | 24.2798 | -0.10064 | 0.1426 | 0.002007 | 0.15576 | 0.15576 | 0.037628
2019-03-20 12:43 | 1553100170.163 | -413 | 0.20958 | -0.001112 | -2.2e-05 | 108.0003 | 109.2984 | 129.6374 | 20.9628 | 0.001853 | 0.001763 | -152.2255 | 0.23904 | 8.4021 | 0.003423 | 0.004632 | 0.004424 | 6.3e-05 | 0.001162 | 0.34307 | 1.3087 | 2.548 | 101.01 | 20.67 | 18.147 | 66.1357 | 103.9299 | 63.9478 | 0.004383 | 1.0473 | 0.025488 | 24.3363 | -0.10687 | 0.14221 | 0.002003 | 0.15557 | 0.15557 | 0.037571
2019-03-20 12:43 | 1553100171.1632 | -412 | 0.20954 | -0.001117 | -2.2e-05 | 107.9631 | 109.1227 | 129.8838 | 20.9632 | 0.001795 | 0.00174 | -152.2249 | 0.23971 | 8.4022 | 0.003461 | 0.004629 | 0.004407 | 6.7e-05 | 0.001164 | 0.35795 | 1.3534 | 2.5481 | 101.01 | 20.67 | 18.14 | 64.7242 | 104.8415 | 63.4165 | 0.004381 | 1.049 | 0.025532 | 24.3402 | -0.095563 | 0.14261 | 0.002009 | 0.15619 | 0.15619 | 0.037525
2019-03-20 12:43 | 1553100172.1642 | -411 | 0.20954 | -0.001116 | -2.2e-05 | 108.0242 | 109.2651 | 129.6407 | 20.9628 | 0.001922 | 0.00172 | -152.2078 | 0.23896 | 8.4003 | 0.003336 | 0.004786 | 0.004602 | 6.4e-05 | 0.001164 | 0.34396 | 1.3377 | 2.5481 | 101.01 | 20.67 | 18.135 | 67.7781 | 101.8321 | 62.9217 | 0.00438 | 1.0505 | 0.025527 | 24.2995 | -0.10217 | 0.14216 | 0.002007 | 0.15608 | 0.15608 | 0.037434
2019-03-20 12:43 | 1553100173.1642 | -410 | 0.20953 | -0.001118 | -2.2e-05 | 107.9084 | 109.1321 | 129.6536 | 20.9642 | 0.001857 | 0.001747 | -152.182 | 0.23949 | 8.4035 | 0.003455 | 0.005077 | 0.004835 | 6.5e-05 | 0.001165 | 0.30414 | 1.3703 | 2.5479 | 101.01 | 20.67 | 18.127 | 66.216 | 104.7122 | 63.58 | 0.004378 | 1.0485 | 0.025503 | 24.3245 | -0.10086 | 0.14247 | 0.00201 | 0.15611 | 0.15611 | 0.037509
2019-03-20 12:43 | 1553100174.1637 | -409 | 0.20952 | -0.001114 | -2.2e-05 | 107.9724 | 109.2771 | 129.4057 | 20.9639 | 0.001846 | 0.001726 | -152.1769 | 0.23937 | 8.4062 | 0.003411 | 0.005003 | 0.004847 | 6.3e-05 | 0.001167 | 0.31564 | 1.3168 | 2.5479 | 101.01 | 20.67 | 18.125 | 65.9684 | 103.6432 | 63.0759 | 0.004378 | 1.05 | 0.025498 | 24.2831 | -0.10743 | 0.1424 | 0.002006 | 0.15576 | 0.15576 | 0.037575
2019-03-20 12:43 | 1553100175.1645 | -408 | 0.20953 | -0.001113 | -2.2e-05 | 107.9964 | 109.1194 | 129.348 | 20.9637 | 0.001868 | 0.00169 | -152.1631 | 0.23894 | 8.4024 | 0.003404 | 0.004617 | 0.004453 | 6.1e-05 | 0.001164 | 0.30541 | 1.3367 | 2.5479 | 101.01 | 20.67 | 18.124 | 66.4971 | 103.4819 | 62.1969 | 0.004377 | 1.0528 | 0.025526 | 24.2459 | -0.092528 | 0.14215 | 0.002004 | 0.15574 | 0.15574 | 0.037513
2019-03-20 12:43 | 1553100176.1634 | -407 | 0.20951 | -0.001117 | -2.2e-05 | 108.0135 | 109.3608 | 129.4944 | 20.9633 | 0.001846 | 0.001738 | -152.1583 | 0.23873 | 8.4021 | 0.003387 | 0.004994 | 0.004789 | 6.1e-05 | 0.00116 | 0.31885 | 1.4065 | 2.548 | 101.01 | 20.67 | 18.124 | 65.9654 | 103.054 | 63.3666 | 0.004377 | 1.0491 | 0.025496 | 24.3019 | -0.11088 | 0.14202 | 0.002008 | 0.1559 | 0.1559 | 0.037441
2019-03-20 12:43 | 1553100177.1635 | -406 | 0.20952 | -0.001115 | -2.2e-05 | 108.0989 | 109.5376 | 129.4466 | 20.9637 | 0.001881 | 0.001729 | -152.167 | 0.23939 | 8.4024 | 0.003442 | 0.004331 | 0.00419 | 6.2e-05 | 0.001158 | 0.33754 | 1.3526 | 2.548 | 101.01 | 20.67 | 18.126 | 66.7953 | 104.4034 | 63.1453 | 0.004378 | 1.0498 | 0.025499 | 24.2894 | -0.11826 | 0.14242 | 0.002006 | 0.15578 | 0.15578 | 0.037574
2019-03-20 12:43 | 1553100178.1625 | -405 | 0.20954 | -0.001115 | -2.2e-05 | 108.0025 | 109.1896 | 129.3527 | 20.9641 | 0.00187 | 0.001722 | -152.1539 | 0.2382 | 8.4013 | 0.003418 | 0.004494 | 0.004216 | 6.8e-05 | 0.001162 | 0.32549 | 1.3356 | 2.548 | 101.01 | 20.67 | 18.136 | 66.5444 | 103.8118 | 62.9644 | 0.00438 | 1.0504 | 0.025497 | 24.2741 | -0.097777 | 0.14171 | 0.002007 | 0.15587 | 0.15587 | 0.037366
2019-03-20 12:43 | 1553100179.1626 | -404 | 0.20955 | -0.001116 | -2.2e-05 | 108.0332 | 109.2532 | 129.4632 | 20.9641 | 0.001823 | 0.001753 | -152.1613 | 0.23827 | 8.4026 | 0.003334 | 0.004647 | 0.004404 | 6.3e-05 | 0.00117 | 0.34058 | 1.3227 | 2.548 | 101.01 | 20.67 | 18.143 | 65.4042 | 101.7871 | 63.7141 | 0.004382 | 1.048 | 0.025479 | 24.3115 | -0.10044 | 0.14175 | 0.002008 | 0.15588 | 0.15588 | 0.037373
2019-03-20 12:43 | 1553100180.1636 | -403 | 0.20954 | -0.001113 | -2.2e-05 | 107.9847 | 109.1732 | 129.3935 | 20.9639 | 0.001877 | 0.001749 | -152.165 | 0.23909 | 8.4012 | 0.003351 | 0.005258 | 0.005093 | 6.6e-05 | 0.001167 | 0.33339 | 1.2904 | 2.548 | 101.01 | 20.67 | 18.152 | 66.7034 | 102.1883 | 63.6148 | 0.004384 | 1.0483 | 0.025476 | 24.3014 | -0.097903 | 0.14223 | 0.002003 | 0.15545 | 0.15545 | 0.037606
2019-03-20 12:43 | 1553100181.1631 | -402 | 0.20953 | -0.001116 | -2.2e-05 | 108.095 | 109.3451 | 129.6261 | 20.9634 | 0.001843 | 0.001738 | -152.1679 | 0.23896 | 8.4021 | 0.003377 | 0.004652 | 0.004408 | 6.2e-05 | 0.001171 | 0.31813 | 1.2865 | 2.5481 | 101.01 | 20.67 | 18.17 | 65.8803 | 102.8083 | 63.3596 | 0.004389 | 1.0491 | 0.025509 | 24.314 | -0.10285 | 0.14216 | 0.002007 | 0.15596 | 0.15596 | 0.037462
2019-03-20 12:43 | 1553100182.1628 | -401 | 0.20958 | -0.001113 | -2.2e-05 | 108.0245 | 109.2782 | 129.8085 | 20.9639 | 0.001855 | 0.001684 | -152.1889 | 0.23892 | 8.4019 | 0.003452 | 0.005199 | 0.00507 | 5.9e-05 | 0.001169 | 0.33196 | 1.3479 | 2.548 | 101.01 | 20.67 | 18.178 | 66.1618 | 104.6449 | 62.069 | 0.004391 | 1.0532 | 0.025576 | 24.2844 | -0.10321 | 0.14213 | 0.002004 | 0.15617 | 0.15617 | 0.037406
2019-03-20 12:43 | 1553100183.1632 | -400 | 0.20959 | -0.001115 | -2.2e-05 | 108.0261 | 109.409 | 129.6869 | 20.9642 | 0.001852 | 0.00174 | -152.207 | 0.2388 | 8.4021 | 0.003331 | 0.004699 | 0.004473 | 6.5e-05 | 0.001174 | 0.32261 | 1.313 | 2.5481 | 101.01 | 20.67 | 18.195 | 66.106 | 101.7064 | 63.4089 | 0.004395 | 1.049 | 0.025513 | 24.3215 | -0.11378 | 0.14206 | 0.002007 | 0.15597 | 0.15597 | 0.037434
2019-03-20 12:43 | 1553100184.1639 | -399 | 0.20953 | -0.001117 | -2.3e-05 | 107.9692 | 109.3373 | 130.0315 | 20.9637 | 0.001821 | 0.001763 | -152.2207 | 0.23958 | 8.4012 | 0.003313 | 0.004548 | 0.00438 | 6.2e-05 | 0.001173 | 0.33627 | 1.2929 | 2.5481 | 101.01 | 20.67 | 18.202 | 65.3507 | 101.2652 | 63.9566 | 0.004396 | 1.0473 | 0.025526 | 24.3736 | -0.11262 | 0.14252 | 0.002009 | 0.15623 | 0.15623 | 0.037495
2019-03-20 12:43 | 1553100185.1638 | -398 | 0.20953 | -0.001117 | -2.2e-05 | 107.9598 | 109.2382 | 130.2995 | 20.9642 | 0.001771 | 0.001711 | -152.219 | 0.23961 | 8.4029 | 0.003367 | 0.004854 | 0.004596 | 6.2e-05 | 0.00118 | 0.31349 | 1.3146 | 2.5481 | 101.01 | 20.67 | 18.209 | 64.1485 | 102.5806 | 62.6982 | 0.004398 | 1.0512 | 0.0256 | 24.3531 | -0.10529 | 0.14255 | 0.002009 | 0.1566 | 0.1566 | 0.037411
2019-03-20 12:43 | 1553100186.1633 | -397 | 0.20957 | -0.001121 | -2.2e-05 | 108.0656 | 109.3492 | 130.393 | 20.9645 | 0.001884 | 0.001727 | -152.2221 | 0.23868 | 8.401 | 0.003284 | 0.004663 | 0.004464 | 6e-05 | 0.001185 | 0.35563 | 1.3086 | 2.5481 | 101.01 | 20.67 | 18.222 | 66.8608 | 100.5763 | 63.0991 | 0.004401 | 1.05 | 0.025594 | 24.3764 | -0.10562 | 0.14199 | 0.002014 | 0.15698 | 0.15698 | 0.037175
2019-03-20 12:43 | 1553100187.164 | -396 | 0.20955 | -0.001118 | -2.2e-05 | 108.0423 | 109.1395 | 130.5293 | 20.9651 | 0.001865 | 0.001709 | -152.216 | 0.23833 | 8.4036 | 0.003399 | 0.004848 | 0.004583 | 6.3e-05 | 0.001186 | 0.32189 | 1.3208 | 2.5479 | 101.01 | 20.67 | 18.228 | 66.4066 | 103.3646 | 62.6553 | 0.004403 | 1.0513 | 0.025624 | 24.373 | -0.090379 | 0.14178 | 0.00201 | 0.15686 | 0.15686 | 0.03715
2019-03-20 12:43 | 1553100188.1631 | -395 | 0.20956 | -0.001118 | -2.2e-05 | 108.0138 | 109.2606 | 130.6256 | 20.9651 | 0.001875 | 0.00173 | -152.2403 | 0.23919 | 8.403 | 0.003344 | 0.004416 | 0.004254 | 6.2e-05 | 0.001189 | 0.33124 | 1.3351 | 2.5479 | 101.01 | 20.67 | 18.238 | 66.6601 | 102.0317 | 63.1707 | 0.004405 | 1.0497 | 0.025614 | 24.4007 | -0.10265 | 0.14229 | 0.00201 | 0.15678 | 0.15678 | 0.037302
2019-03-20 12:43 | 1553100189.1631 | -394 | 0.20956 | -0.001118 | -2.2e-05 | 108.0252 | 109.3044 | 130.546 | 20.9652 | 0.001833 | 0.001742 | -152.2459 | 0.23988 | 8.4025 | 0.003348 | 0.005015 | 0.00481 | 6.4e-05 | 0.001191 | 0.32549 | 1.3222 | 2.5478 | 101.01 | 20.67 | 18.241 | 65.638 | 102.1266 | 63.4551 | 0.004406 | 1.0488 | 0.025595 | 24.4036 | -0.1053 | 0.14271 | 0.002009 | 0.15664 | 0.15664 | 0.037443
2019-03-20 12:43 | 1553100190.1646 | -393 | 0.20952 | -0.00112 | -2.2e-05 | 108.0109 | 109.2457 | 130.4415 | 20.9656 | 0.001793 | 0.001724 | -152.2532 | 0.23934 | 8.4038 | 0.003327 | 0.004585 | 0.004322 | 6.3e-05 | 0.001193 | 0.3131 | 1.3334 | 2.548 | 101.01 | 20.67 | 18.245 | 64.6852 | 101.6025 | 63.0292 | 0.004407 | 1.0502 | 0.025601 | 24.3784 | -0.10167 | 0.14238 | 0.002012 | 0.15687 | 0.15687 | 0.037304
2019-03-20 12:43 | 1553100191.1642 | -392 | 0.20949 | -0.001126 | -2.1e-05 | 108.0886 | 109.3227 | 130.2402 | 20.9661 | 0.001881 | 0.00172 | -152.2316 | 0.23863 | 8.402 | 0.003328 | 0.004913 | 0.004695 | 6.2e-05 | 0.001197 | 0.31222 | 1.3456 | 2.548 | 101.01 | 20.67 | 18.251 | 66.7929 | 101.6335 | 62.9194 | 0.004408 | 1.0505 | 0.025586 | 24.3556 | -0.10155 | 0.14196 | 0.002018 | 0.15714 | 0.15714 | 0.03713
2019-03-20 12:43 | 1553100192.1644 | -391 | 0.20916 | -0.001122 | -1.6e-05 | 108.0242 | 109.2248 | 130.2587 | 20.9658 | 0.001891 | 0.00172 | -152.2345 | 0.23877 | 8.3998 | 0.003334 | 0.004032 | 0.003884 | 6.5e-05 | 0.001198 | 0.33915 | 1.3072 | 2.5481 | 101.01 | 20.67 | 18.254 | 67.0469 | 101.7839 | 62.9378 | 0.004409 | 1.0505 | 0.025587 | 24.358 | -0.09886 | 0.14204 | 0.002006 | 0.15567 | 0.15567 | 0.037503
2019-03-20 12:43 | 1553100193.1643 | -390 | 0.20913 | -0.001119 | -1e-05 | 108.0455 | 109.353 | 130.4854 | 20.965 | 0.00182 | 0.00173 | -152.2334 | 0.239 | 8.4014 | 0.003287 | 0.004299 | 0.004077 | 6.3e-05 | 0.001197 | 0.32803 | 1.33 | 2.548 | 101.01 | 20.67 | 18.259 | 65.3207 | 100.63 | 63.173 | 0.00441 | 1.0497 | 0.0256 | 24.3877 | -0.10759 | 0.14218 | 0.001996 | 0.15456 | 0.15456 | 0.037808
2019-03-20 12:43 | 1553100194.1633 | -389 | 0.19902 | -0.001116 | -6e-06 | 108.0362 | 109.2472 | 130.1282 | 20.966 | 0.001815 | 0.00173 | -152.2123 | 0.23863 | 8.4017 | 0.003348 | 0.004356 | 0.004014 | 6.1e-05 | 0.001199 | 0.33516 | 1.3393 | 2.5481 | 101.01 | 20.67 | 18.261 | 65.2156 | 102.1204 | 63.1591 | 0.004411 | 1.0498 | 0.025566 | 24.3538 | -0.099705 | 0.14196 | 0.001961 | 0.14411 | 0.14411 | 0.040487
2019-03-20 12:43 | 1553100195.1638 | -388 | 0.17141 | -0.001096 | -4e-06 | 108.0578 | 109.2816 | 129.7963 | 20.9652 | 0.001854 | 0.001706 | -152.1978 | 0.23846 | 8.4059 | 0.003375 | 0.004758 | 0.004563 | 6.4e-05 | 0.001195 | 0.30237 | 1.2785 | 2.548 | 101.01 | 20.68 | 18.264 | 66.1499 | 102.7723 | 62.5827 | 0.004414 | 1.0516 | 0.025555 | 24.3018 | -0.10073 | 0.14186 | 0.00187 | 0.11828 | 0.11828 | 0.049296
2019-03-20 12:43 | 1553100196.1635 | -387 | 0.13596 | -0.001028 | -5e-06 | 107.9857 | 109.2214 | 129.6123 | 20.9666 | 0.00187 | 0.001725 | -152.1841 | 0.23872 | 8.4033 | 0.00343 | 0.004556 | 0.004368 | 6.2e-05 | 0.001191 | 0.3621 | 1.2902 | 2.5479 | 101.01 | 20.68 | 18.267 | 66.5281 | 104.1004 | 63.041 | 0.004415 | 1.0501 | 0.025519 | 24.3012 | -0.10177 | 0.14201 | 0.001715 | 0.086026 | 0.086026 | 0.067848
2019-03-20 12:43 | 1553100197.1636 | -386 | 0.10085 | -0.000949 | -6e-06 | 108.0381 | 109.3933 | 129.4041 | 20.9658 | 0.001861 | 0.001752 | -152.1844 | 0.23855 | 8.4037 | 0.00331 | 0.005314 | 0.005085 | 6.3e-05 | 0.001194 | 0.33931 | 1.2757 | 2.548 | 101.01 | 20.68 | 18.268 | 66.3267 | 101.1896 | 63.686 | 0.004415 | 1.0481 | 0.025475 | 24.3049 | -0.1115 | 0.14192 | 0.001562 | 0.058177 | 0.058177 | 0.10026
2019-03-20 12:43 | 1553100198.1638 | -385 | 0.082375 | -0.000895 | -8e-06 | 107.9896 | 109.2259 | 129.2747 | 20.9662 | 0.001875 | 0.001732 | -152.1777 | 0.23782 | 8.4017 | 0.003325 | 0.004736 | 0.004525 | 6e-05 | 0.001186 | 0.33069 | 1.3236 | 2.5482 | 101.01 | 20.68 | 18.268 | 66.6556 | 101.5646 | 63.2195 | 0.004415 | 1.0496 | 0.025479 | 24.276 | -0.10182 | 0.14148 | 0.001474 | 0.045061 | 0.045061 | 0.12904
2019-03-20 12:43 | 1553100199.1626 | -384 | 0.058822 | -0.000884 | -1e-05 | 108.0679 | 109.289 | 129.4379 | 20.966 | 0.001906 | 0.001727 | -152.1648 | 0.23884 | 8.4006 | 0.003273 | 0.004639 | 0.004473 | 6.1e-05 | 0.001193 | 0.31924 | 1.3344 | 2.5481 | 101.01 | 20.68 | 18.267 | 67.392 | 100.3013 | 63.0967 | 0.004415 | 1.05 | 0.0255 | 24.2868 | -0.10051 | 0.14209 | 0.001428 | 0.03148 | 0.03148 | 0.18551
2019-03-20 12:43 | 1553100200.1626 | -383 | 0.040542 | -0.000918 | -1.2e-05 | 107.9692 | 109.2939 | 129.5815 | 20.967 | 0.001829 | 0.001721 | -152.1361 | 0.23899 | 8.4021 | 0.003399 | 0.00459 | 0.004416 | 6.1e-05 | 0.001196 | 0.30237 | 1.33 | 2.5479 | 101.01 | 20.68 | 18.268 | 65.5468 | 103.3555 | 62.9554 | 0.004415 | 1.0504 | 0.02552 | 24.2952 | -0.10907 | 0.14217 | 0.001438 | 0.022179 | 0.022179 | 0.26346
2019-03-20 12:43 | 1553100201.1637 | -382 | 0.033167 | -0.000979 | -1.3e-05 | 107.9873 | 109.258 | 129.6655 | 20.9662 | 0.001883 | 0.001746 | -152.1502 | 0.23887 | 8.4021 | 0.003369 | 0.004628 | 0.004383 | 6.2e-05 | 0.001195 | 0.34108 | 1.3256 | 2.548 | 101.01 | 20.68 | 18.269 | 66.8569 | 102.6185 | 63.5646 | 0.004415 | 1.0485 | 0.025505 | 24.3251 | -0.10464 | 0.14211 | 0.001491 | 0.019013 | 0.019013 | 0.30718
2019-03-20 12:43 | 1553100202.1638 | -381 | 0.019245 | -0.001051 | -1.5e-05 | 108.0193 | 109.1878 | 129.5501 | 20.9669 | 0.001932 | 0.001743 | -152.1471 | 0.23854 | 8.399 | 0.003361 | 0.004654 | 0.004465 | 6.2e-05 | 0.001192 | 0.33284 | 1.3483 | 2.5481 | 101.01 | 20.68 | 18.269 | 68.0201 | 102.4236 | 63.4762 | 0.004415 | 1.0488 | 0.025497 | 24.3111 | -0.096231 | 0.14191 | 0.001545 | 0.011854 | 0.011854 | 0.49201
2019-03-20 12:43 | 1553100203.1628 | -380 | 0.010075 | -0.001124 | -1.6e-05 | 108.0255 | 109.4554 | 129.4841 | 20.9666 | 0.001891 | 0.001754 | -152.155 | 0.23986 | 8.4002 | 0.003363 | 0.005032 | 0.004819 | 5.9e-05 | 0.001189 | 0.30917 | 1.3604 | 2.5481 | 101.01 | 20.68 | 18.267 | 67.04 | 102.4768 | 63.7532 | 0.004415 | 1.0479 | 0.02548 | 24.3149 | -0.11762 | 0.14269 | 0.001605 | 0.006952 | 0.006952 | 0.84362
2019-03-20 12:43 | 1553100204.1629 | -379 | 0.006227 | -0.001192 | -1.7e-05 | 108.0624 | 109.2502 | 129.3315 | 20.9677 | 0.001885 | 0.001773 | -152.1455 | 0.23911 | 8.4024 | 0.003363 | 0.004652 | 0.004468 | 6.2e-05 | 0.001192 | 0.32068 | 1.3169 | 2.5479 | 101.01 | 20.68 | 18.263 | 66.9069 | 102.481 | 64.2079 | 0.004414 | 1.0465 | 0.025448 | 24.3169 | -0.097781 | 0.14225 | 0.001668 | 0.004891 | 0.004891 | 1.1953
2019-03-20 12:43 | 1553100205.1629 | -378 | 0.004072 | -0.001248 | -1.8e-05 | 107.9505 | 109.2229 | 129.4156 | 20.9682 | 0.001876 | 0.001721 | -152.1741 | 0.23913 | 8.4013 | 0.003349 | 0.004475 | 0.004277 | 6.2e-05 | 0.001189 | 0.34202 | 1.3494 | 2.548 | 101.01 | 20.68 | 18.258 | 66.6707 | 102.1382 | 62.9527 | 0.004413 | 1.0504 | 0.025503 | 24.2795 | -0.10481 | 0.14226 | 0.001722 | 0.003722 | 0.003722 | 1.5709
2019-03-20 12:43 | 1553100206.1646 | -377 | 0.002934 | -0.001301 | -1.8e-05 | 108.0523 | 109.165 | 129.3836 | 20.9677 | 0.001846 | 0.001774 | -152.1768 | 0.23911 | 8.4007 | 0.003406 | 0.004295 | 0.004076 | 6e-05 | 0.001195 | 0.32516 | 1.3558 | 2.548 | 101.01 | 20.68 | 18.245 | 65.9509 | 103.5299 | 64.2361 | 0.004409 | 1.0464 | 0.025452 | 24.3229 | -0.091632 | 0.14225 | 0.001773 | 0.003082 | 0.003082 | 1.8969
2019-03-20 12:43 | 1553100207.1631 | -376 | 0.002226 | -0.001349 | -1.9e-05 | 107.953 | 109.1855 | 129.4799 | 20.9679 | 0.001869 | 0.001727 | -152.1699 | 0.23931 | 8.4042 | 0.003373 | 0.004907 | 0.004705 | 6.4e-05 | 0.001198 | 0.32444 | 1.3358 | 2.548 | 101.01 | 20.68 | 18.238 | 66.5106 | 102.7176 | 63.107 | 0.004408 | 1.0499 | 0.025504 | 24.2912 | -0.10154 | 0.14236 | 0.001821 | 0.002739 | 0.002739 | 2.1359
2019-03-20 12:43 | 1553100208.1631 | -375 | 0.001921 | -0.001385 | -1.9e-05 | 107.9527 | 109.403 | 129.5867 | 20.9678 | 0.001815 | 0.001767 | -152.1734 | 0.239 | 8.4015 | 0.003414 | 0.004557 | 0.004342 | 5.9e-05 | 0.001197 | 0.3095 | 1.3031 | 2.548 | 101.01 | 20.68 | 18.231 | 65.2041 | 103.7156 | 64.0535 | 0.004406 | 1.047 | 0.025479 | 24.3353 | -0.11937 | 0.14218 | 0.001856 | 0.002547 | 0.002547 | 2.2939
2019-03-20 12:43 | 1553100209.1622 | -374 | 0.001661 | -0.001417 | -2e-05 | 107.9395 | 109.1687 | 129.3336 | 20.968 | 0.001869 | 0.001703 | -152.1773 | 0.23883 | 8.4028 | 0.003379 | 0.004655 | 0.004366 | 5.7e-05 | 0.001196 | 0.32549 | 1.3 | 2.5481 | 101.01 | 20.69 | 18.215 | 66.5178 | 102.8796 | 62.5199 | 0.004405 | 1.0518 | 0.025512 | 24.2562 | -0.10129 | 0.14208 | 0.001888 | 0.002459 | 0.002459 | 2.3751
2019-03-20 12:43 | 1553100210.1637 | -373 | 0.001407 | -0.001442 | -2.1e-05 | 107.9679 | 109.1922 | 129.2189 | 20.9685 | 0.001875 | 0.001726 | -152.2029 | 0.23911 | 8.4026 | 0.003336 | 0.004291 | 0.004068 | 6.1e-05 | 0.001201 | 0.30879 | 1.3489 | 2.5481 | 101.01 | 20.69 | 18.206 | 66.6689 | 101.8338 | 63.0717 | 0.004403 | 1.05 | 0.02548 | 24.2654 | -0.10086 | 0.14225 | 0.001914 | 0.002325 | 0.002325 | 2.5149
2019-03-20 12:43 | 1553100211.1628 | -372 | 0.00103 | -0.001467 | -2.1e-05 | 107.9236 | 109.2775 | 129.0921 | 20.9689 | 0.001862 | 0.001782 | -152.2096 | 0.2397 | 8.4014 | 0.003338 | 0.00445 | 0.004019 | 6.1e-05 | 0.001201 | 0.32045 | 1.3259 | 2.5481 | 101.01 | 20.68 | 18.197 | 66.3439 | 101.8782 | 64.4097 | 0.004398 | 1.0459 | 0.025417 | 24.3017 | -0.11151 | 0.1426 | 0.001937 | 0.002057 | 0.002057 | 2.8487
2019-03-20 12:44 | 1553100212.164 | -371 | 0.000857 | -0.001486 | -2.1e-05 | 107.9359 | 109.2311 | 129.2128 | 20.9695 | 0.001863 | 0.001726 | -152.1909 | 0.23961 | 8.4016 | 0.003305 | 0.004481 | 0.004094 | 6e-05 | 0.0012 | 0.31996 | 1.3114 | 2.5481 | 101.01 | 20.69 | 18.177 | 66.3748 | 101.074 | 63.0657 | 0.004396 | 1.0501 | 0.025479 | 24.2646 | -0.10669 | 0.14254 | 0.001956 | 0.001943 | 0.001943 | 3.0153
2019-03-20 12:44 | 1553100213.1625 | -370 | 0.000651 | -0.001497 | -2.1e-05 | 107.8495 | 109.2012 | 129.2705 | 20.969 | 0.001831 | 0.001755 | -152.204 | 0.23896 | 8.4022 | 0.00335 | 0.004581 | 0.004205 | 6.1e-05 | 0.001204 | 0.30431 | 1.332 | 2.548 | 101.01 | 20.69 | 18.166 | 65.6065 | 102.1775 | 63.7805 | 0.004393 | 1.0478 | 0.025458 | 24.2958 | -0.1114 | 0.14216 | 0.001967 | 0.001829 | 0.001829 | 3.1943
2019-03-20 12:44 | 1553100214.1634 | -369 | 0.000498 | -0.001512 | -2.1e-05 | 107.932 | 109.3556 | 129.6573 | 20.9695 | 0.001832 | 0.001757 | -152.2248 | 0.23866 | 8.4014 | 0.003262 | 0.0046 | 0.004258 | 6e-05 | 0.001207 | 0.33644 | 1.347 | 2.548 | 101.01 | 20.69 | 18.143 | 65.6196 | 100.0397 | 63.8319 | 0.004387 | 1.0477 | 0.025494 | 24.334 | -0.1172 | 0.14198 | 0.001982 | 0.00173 | 0.00173 | 3.3724
2019-03-20 12:44 | 1553100215.1631 | -368 | 0.000397 | -0.001522 | -2.1e-05 | 107.8625 | 109.2647 | 129.8092 | 20.9695 | 0.001885 | 0.001761 | -152.2027 | 0.23998 | 8.3993 | 0.003385 | 0.004725 | 0.004365 | 6.2e-05 | 0.001208 | 0.31852 | 1.3391 | 2.548 | 101.01 | 20.69 | 18.132 | 66.8945 | 103.0248 | 63.9067 | 0.004385 | 1.0474 | 0.025506 | 24.3509 | -0.11553 | 0.14276 | 0.001991 | 0.001622 | 0.001622 | 3.6182
2019-03-20 12:44 | 1553100216.1626 | -367 | 0.000362 | -0.001529 | -2.1e-05 | 107.9459 | 109.3829 | 129.766 | 20.9694 | 0.001878 | 0.00172 | -152.1941 | 0.23963 | 8.4011 | 0.003357 | 0.00476 | 0.004389 | 5.9e-05 | 0.00121 | 0.32173 | 1.3264 | 2.548 | 101.01 | 20.69 | 18.111 | 66.7205 | 102.3309 | 62.9288 | 0.00438 | 1.0505 | 0.025539 | 24.3115 | -0.11828 | 0.14255 | 0.001999 | 0.001633 | 0.001633 | 3.5889
2019-03-20 12:44 | 1553100217.1627 | -366 | 3.3e-05 | -0.001536 | -2.1e-05 | 107.9191 | 109.2535 | 129.6123 | 20.9698 | 0.001867 | 0.001725 | -152.2122 | 0.23835 | 8.4032 | 0.003341 | 0.004671 | 0.004261 | 6.3e-05 | 0.001211 | 0.34058 | 1.341 | 2.548 | 101.01 | 20.69 | 18.105 | 66.457 | 101.9488 | 63.0409 | 0.004378 | 1.0501 | 0.025519 | 24.3012 | -0.10992 | 0.1418 | 0.002005 | 0.001382 | 0.001382 | 4.2178
2019-03-20 12:44 | 1553100218.1635 | -365 | 0.000293 | -0.001542 | -2.1e-05 | 107.9155 | 109.2494 | 129.4005 | 20.9707 | 0.001855 | 0.001801 | -152.2128 | 0.23909 | 8.4004 | 0.003313 | 0.0043 | 0.003909 | 6.1e-05 | 0.001214 | 0.32333 | 1.2995 | 2.5482 | 101.01 | 20.69 | 18.099 | 66.1878 | 101.2752 | 64.8744 | 0.004377 | 1.0444 | 0.025429 | 24.3474 | -0.10988 | 0.14223 | 0.002011 | 0.001581 | 0.001581 | 3.6965
2019-03-20 12:44 | 1553100219.1628 | -364 | 0.000276 | -0.001545 | -2.2e-05 | 107.923 | 109.3193 | 129.3398 | 20.9699 | 0.00191 | 0.001705 | -152.2206 | 0.23936 | 8.4044 | 0.003282 | 0.005221 | 0.00487 | 5.9e-05 | 0.001214 | 0.32516 | 1.3387 | 2.548 | 101.01 | 20.69 | 18.091 | 67.4933 | 100.5249 | 62.5691 | 0.004375 | 1.0516 | 0.025511 | 24.2586 | -0.11499 | 0.14239 | 0.002015 | 0.001599 | 0.001599 | 3.6611
2019-03-20 12:44 | 1553100220.1641 | -363 | 0.00024 | -0.001546 | -2.2e-05 | 107.9307 | 109.3111 | 129.1161 | 20.9699 | 0.001874 | 0.001782 | -152.1996 | 0.23918 | 8.4019 | 0.003329 | 0.00447 | 0.00409 | 6.2e-05 | 0.001213 | 0.31996 | 1.3425 | 2.5479 | 101.01 | 20.69 | 18.09 | 66.6296 | 101.6529 | 64.4232 | 0.004375 | 1.0458 | 0.025418 | 24.3044 | -0.11367 | 0.14229 | 0.002016 | 0.001568 | 0.001568 | 3.7297
2019-03-20 12:44 | 1553100221.1628 | -362 | 0.000236 | -0.001549 | -2.2e-05 | 107.9207 | 109.1362 | 129.0517 | 20.9702 | 0.001909 | 0.00175 | -152.1862 | 0.2393 | 8.4012 | 0.003379 | 0.004956 | 0.00462 | 5.9e-05 | 0.001211 | 0.31974 | 1.3626 | 2.548 | 101.01 | 20.69 | 18.09 | 67.474 | 102.87 | 63.6593 | 0.004375 | 1.0482 | 0.025441 | 24.2709 | -0.10018 | 0.14236 | 0.002019 | 0.001575 | 0.001575 | 3.7141
2019-03-20 12:44 | 1553100222.1639 | -361 | 0.000159 | -0.001554 | -2.2e-05 | 107.9285 | 109.3152 | 129.0426 | 20.9701 | 0.001853 | 0.001775 | -152.2053 | 0.23937 | 8.402 | 0.003371 | 0.00417 | 0.00376 | 5.9e-05 | 0.001208 | 0.30735 | 1.3535 | 2.548 | 101.01 | 20.69 | 18.089 | 66.1195 | 102.6718 | 64.2492 | 0.004374 | 1.0464 | 0.025418 | 24.2913 | -0.11419 | 0.1424 | 0.002024 | 0.001523 | 0.001523 | 3.8439
2019-03-20 12:44 | 1553100223.1633 | -360 | 0.000186 | -0.001551 | -2.2e-05 | 107.8521 | 109.0857 | 129.1144 | 20.9713 | 0.001849 | 0.001735 | -152.1926 | 0.2391 | 8.4025 | 0.003414 | 0.004733 | 0.004375 | 5.9e-05 | 0.001213 | 0.30071 | 1.3428 | 2.5482 | 101.01 | 20.69 | 18.096 | 66.041 | 103.7134 | 63.304 | 0.004376 | 1.0493 | 0.02546 | 24.264 | -0.10172 | 0.14224 | 0.002021 | 0.001548 | 0.001548 | 3.7776
2019-03-20 12:44 | 1553100224.1632 | -359 | 0.000154 | -0.001559 | -2.2e-05 | 107.8997 | 109.1986 | 128.9665 | 20.9711 | 0.001808 | 0.001749 | -152.1918 | 0.239 | 8.4034 | 0.003369 | 0.005044 | 0.004665 | 5.9e-05 | 0.001211 | 0.3246 | 1.351 | 2.5482 | 101.01 | 20.69 | 18.1 | 65.0387 | 102.6399 | 63.6195 | 0.004377 | 1.0483 | 0.025434 | 24.2614 | -0.10703 | 0.14218 | 0.002028 | 0.001496 | 0.001496 | 3.9071
2019-03-20 12:44 | 1553100225.1631 | -358 | 0.000155 | -0.001556 | -2.2e-05 | 107.9424 | 109.2416 | 129.0257 | 20.9714 | 0.001856 | 0.001696 | -152.1932 | 0.23926 | 8.402 | 0.003364 | 0.00515 | 0.004839 | 6e-05 | 0.001217 | 0.30038 | 1.332 | 2.5481 | 101.01 | 20.69 | 18.104 | 66.202 | 102.5004 | 62.3489 | 0.004378 | 1.0523 | 0.025488 | 24.2212 | -0.10701 | 0.14234 | 0.002025 | 0.001512 | 0.001512 | 3.8701
2019-03-20 12:44 | 1553100226.1631 | -357 | 0.000119 | -0.001563 | -2.2e-05 | 107.9304 | 109.1792 | 129.263 | 20.9716 | 0.001829 | 0.001787 | -152.2078 | 0.23918 | 8.4017 | 0.003405 | 0.004501 | 0.004132 | 5.9e-05 | 0.001218 | 0.34363 | 1.34 | 2.5481 | 101.01 | 20.69 | 18.113 | 65.545 | 103.5141 | 64.5327 | 0.00438 | 1.0455 | 0.025429 | 24.3222 | -0.10289 | 0.14229 | 0.002033 | 0.001503 | 0.001503 | 3.8898
2019-03-20 12:44 | 1553100227.1638 | -356 | 0.000114 | -0.001562 | -2.2e-05 | 107.9123 | 109.2733 | 129.3214 | 20.9714 | 0.001837 | 0.001723 | -152.1953 | 0.23991 | 8.4019 | 0.003388 | 0.004944 | 0.004585 | 5.9e-05 | 0.001216 | 0.31188 | 1.3583 | 2.5482 | 101.01 | 20.69 | 18.119 | 65.7462 | 103.0879 | 63.0122 | 0.004382 | 1.0502 | 0.025492 | 24.2729 | -0.11211 | 0.14272 | 0.002032 | 0.001503 | 0.001503 | 3.9027
2019-03-20 12:44 | 1553100228.1633 | -355 | 9e-05 | -0.001561 | -2.2e-05 | 107.9126 | 109.0831 | 129.4301 | 20.9713 | 0.001887 | 0.001698 | -152.1867 | 0.24002 | 8.403 | 0.00341 | 0.004837 | 0.004467 | 5.9e-05 | 0.001216 | 0.33682 | 1.3095 | 2.548 | 101.01 | 20.69 | 18.132 | 66.947 | 103.6214 | 62.4058 | 0.004385 | 1.0521 | 0.025526 | 24.2611 | -0.096496 | 0.14279 | 0.00203 | 0.001461 | 0.001461 | 4.0158
2019-03-20 12:44 | 1553100229.1622 | -354 | 0.000172 | -0.001558 | -2.2e-05 | 107.879 | 109.0345 | 129.5127 | 20.9717 | 0.001836 | 0.001731 | -152.1878 | 0.23852 | 8.4021 | 0.003363 | 0.004289 | 0.003941 | 6.1e-05 | 0.001219 | 0.30182 | 1.3111 | 2.548 | 101.01 | 20.69 | 18.136 | 65.7292 | 102.4983 | 63.1983 | 0.004386 | 1.0496 | 0.025504 | 24.2975 | -0.095299 | 0.1419 | 0.002028 | 0.001546 | 0.001546 | 3.7714
2019-03-20 12:44 | 1553100230.1636 | -353 | 0.000112 | -0.001567 | -2.2e-05 | 107.8647 | 109.3731 | 129.5949 | 20.9723 | 0.0018 | 0.001734 | -152.1971 | 0.23923 | 8.4016 | 0.003337 | 0.004193 | 0.003817 | 5.7e-05 | 0.001222 | 0.32908 | 1.3167 | 2.5481 | 101.01 | 20.69 | 18.141 | 64.8666 | 101.8658 | 63.2799 | 0.004387 | 1.0494 | 0.025509 | 24.3082 | -0.12421 | 0.14232 | 0.002037 | 0.00151 | 0.00151 | 3.8736
2019-03-20 12:44 | 1553100231.1623 | -352 | 9.6e-05 | -0.001564 | -2.2e-05 | 107.9126 | 109.301 | 129.8052 | 20.9721 | 0.001861 | 0.001766 | -152.236 | 0.23897 | 8.402 | 0.003371 | 0.004943 | 0.004528 | 6e-05 | 0.001217 | 0.32748 | 1.3024 | 2.5481 | 101.01 | 20.7 | 18.143 | 66.3328 | 102.6849 | 64.032 | 0.00439 | 1.047 | 0.025501 | 24.3551 | -0.11435 | 0.14216 | 0.002035 | 0.001514 | 0.001514 | 3.8589
2019-03-20 12:44 | 1553100232.1628 | -351 | 0.000146 | -0.001566 | -2.2e-05 | 107.9712 | 109.3589 | 129.7885 | 20.9716 | 0.001837 | 0.001743 | -152.224 | 0.23982 | 8.4009 | 0.003427 | 0.004167 | 0.003808 | 6.4e-05 | 0.001219 | 0.33373 | 1.3828 | 2.5482 | 101.01 | 20.7 | 18.142 | 65.759 | 104.0266 | 63.4777 | 0.00439 | 1.0488 | 0.02552 | 24.3335 | -0.11423 | 0.14267 | 0.002035 | 0.001516 | 0.001516 | 3.8674
2019-03-20 12:44 | 1553100233.1634 | -350 | 0.000107 | -0.001568 | -2.2e-05 | 107.9304 | 109.0457 | 129.5137 | 20.9717 | 0.001877 | 0.001745 | -152.2118 | 0.23896 | 8.3997 | 0.003454 | 0.004355 | 0.003991 | 6.1e-05 | 0.001215 | 0.32637 | 1.3531 | 2.5481 | 101.01 | 20.7 | 18.14 | 66.7079 | 104.6824 | 63.5449 | 0.004389 | 1.0486 | 0.025491 | 24.3101 | -0.091956 | 0.14216 | 0.002038 | 0.001487 | 0.001487 | 3.9294
2019-03-20 12:44 | 1553100234.1635 | -349 | 0.000146 | -0.00157 | -2.2e-05 | 107.9184 | 109.2745 | 129.6456 | 20.9728 | 0.001861 | 0.00171 | -152.2104 | 0.23894 | 8.4012 | 0.003432 | 0.004237 | 0.003872 | 6.4e-05 | 0.001219 | 0.33251 | 1.4084 | 2.5482 | 101.01 | 20.7 | 18.13 | 66.3335 | 104.1545 | 62.7003 | 0.004387 | 1.0512 | 0.025536 | 24.292 | -0.11169 | 0.14214 | 0.00204 | 0.001536 | 0.001536 | 3.8042
2019-03-20 12:44 | 1553100235.163 | -348 | 0.000116 | -0.001567 | -2.2e-05 | 107.8877 | 109.0345 | 129.5275 | 20.9733 | 0.001871 | 0.001709 | -152.2175 | 0.23991 | 8.4022 | 0.003431 | 0.004815 | 0.004435 | 6.4e-05 | 0.00122 | 0.32477 | 1.325 | 2.548 | 101.01 | 20.7 | 18.124 | 66.5776 | 104.1328 | 62.6639 | 0.004386 | 1.0513 | 0.025525 | 24.2796 | -0.094575 | 0.14272 | 0.002037 | 0.001527 | 0.001527 | 3.8417
2019-03-20 12:44 | 1553100236.1632 | -347 | 4.9e-05 | -0.001564 | -2.2e-05 | 107.9178 | 109.1493 | 129.4116 | 20.9731 | 0.001847 | 0.001767 | -152.2144 | 0.24032 | 8.404 | 0.003434 | 0.004181 | 0.003747 | 6.4e-05 | 0.001216 | 0.30309 | 1.383 | 2.5479 | 101.01 | 20.7 | 18.11 | 65.9812 | 104.209 | 64.0649 | 0.004382 | 1.0469 | 0.025461 | 24.3193 | -0.10149 | 0.14297 | 0.002034 | 0.001446 | 0.001446 | 4.0636
2019-03-20 12:44 | 1553100237.1627 | -346 | 9e-05 | -0.001565 | -2.2e-05 | 107.8825 | 109.2352 | 129.4827 | 20.9735 | 0.0019 | 0.001701 | -152.1948 | 0.23905 | 8.4043 | 0.0034 | 0.004719 | 0.004344 | 5.9e-05 | 0.001218 | 0.34058 | 1.3514 | 2.548 | 101.01 | 20.7 | 18.102 | 67.2557 | 103.3854 | 62.4797 | 0.00438 | 1.0519 | 0.025528 | 24.2687 | -0.11145 | 0.14221 | 0.002035 | 0.001495 | 0.001495 | 3.9099
2019-03-20 12:44 | 1553100238.1637 | -345 | 8.5e-05 | -0.001564 | -2.2e-05 | 107.9382 | 109.113 | 129.6726 | 20.9735 | 0.001822 | 0.00172 | -152.1846 | 0.23978 | 8.4012 | 0.003446 | 0.004688 | 0.004303 | 6.5e-05 | 0.001212 | 0.32803 | 1.3115 | 2.5482 | 101.01 | 20.7 | 18.082 | 65.3822 | 104.4932 | 62.9398 | 0.004375 | 1.0504 | 0.025529 | 24.3032 | -0.09683 | 0.14265 | 0.002034 | 0.001476 | 0.001476 | 3.971
2019-03-20 12:44 | 1553100239.1641 | -344 | 3.2e-05 | -0.001571 | -2.2e-05 | 107.868 | 109.1657 | 129.6325 | 20.9727 | 0.001881 | 0.001779 | -152.1745 | 0.2393 | 8.4023 | 0.00345 | 0.004569 | 0.004229 | 6.3e-05 | 0.001209 | 0.33339 | 1.3046 | 2.5482 | 101.01 | 20.7 | 18.073 | 66.818 | 104.6071 | 64.3634 | 0.004373 | 1.046 | 0.025471 | 24.3508 | -0.10697 | 0.14236 | 0.002041 | 0.001455 | 0.001455 | 4.0206
2019-03-20 12:44 | 1553100240.1629 | -343 | 9.2e-05 | -0.001572 | -2.2e-05 | 107.8686 | 109.2543 | 129.809 | 20.973 | 0.001825 | 0.001785 | -152.1685 | 0.23848 | 8.4034 | 0.003436 | 0.004832 | 0.004411 | 6.3e-05 | 0.00121 | 0.35275 | 1.3211 | 2.548 | 101.01 | 20.7 | 18.063 | 65.4557 | 104.2622 | 64.5028 | 0.004371 | 1.0456 | 0.025483 | 24.3724 | -0.11417 | 0.14187 | 0.002042 | 0.001492 | 0.001492 | 3.9085
2019-03-20 12:44 | 1553100241.1629 | -342 | 3.8e-05 | -0.001573 | -2.3e-05 | 107.869 | 109.4094 | 129.8019 | 20.9725 | 0.001837 | 0.001716 | -152.1359 | 0.23904 | 8.4002 | 0.003486 | 0.005101 | 0.004782 | 6e-05 | 0.001213 | 0.33804 | 1.34 | 2.548 | 101.01 | 20.7 | 18.046 | 65.7435 | 105.469 | 62.8445 | 0.004367 | 1.0507 | 0.025546 | 24.3118 | -0.12683 | 0.1422 | 0.002044 | 0.001512 | 0.001512 | 3.8651
2019-03-20 12:44 | 1553100242.1633 | -341 | 7.8e-05 | -0.001573 | -2.2e-05 | 107.9686 | 109.2139 | 129.5203 | 20.9736 | 0.0018 | 0.001696 | -152.1363 | 0.23888 | 8.4024 | 0.003451 | 0.005111 | 0.004715 | 6.1e-05 | 0.001214 | 0.31526 | 1.383 | 2.5482 | 101.01 | 20.7 | 18.039 | 64.8703 | 104.6121 | 62.3597 | 0.004365 | 1.0523 | 0.025536 | 24.2679 | -0.10258 | 0.14211 | 0.002043 | 0.0015 | 0.0015 | 3.8937
2019-03-20 12:44 | 1553100243.162 | -340 | 8e-05 | -0.00157 | -2.2e-05 | 107.8871 | 109.2726 | 129.5212 | 20.9738 | 0.00185 | 0.0017 | -152.1587 | 0.23924 | 8.4012 | 0.003412 | 0.004147 | 0.003797 | 6.1e-05 | 0.001215 | 0.30486 | 1.399 | 2.5482 | 101.01 | 20.7 | 18.034 | 66.0629 | 103.6686 | 62.4465 | 0.004364 | 1.052 | 0.025533 | 24.2712 | -0.11414 | 0.14233 | 0.002039 | 0.001488 | 0.001488 | 3.9325
2019-03-20 12:44 | 1553100244.1631 | -339 | 9.3e-05 | -0.001571 | -2.2e-05 | 107.9857 | 109.2808 | 129.3461 | 20.9728 | 0.001858 | 0.00169 | -152.1629 | 0.23906 | 8.4034 | 0.003404 | 0.004717 | 0.004321 | 6.2e-05 | 0.001223 | 0.32444 | 1.3429 | 2.5481 | 101.01 | 20.69 | 18.022 | 66.2594 | 103.4901 | 62.205 | 0.004358 | 1.0528 | 0.025525 | 24.246 | -0.10664 | 0.14222 | 0.002041 | 0.001503 | 0.001503 | 3.8898
2019-03-20 12:44 | 1553100245.1631 | -338 | 5.2e-05 | -0.001572 | -2.3e-05 | 107.921 | 109.2292 | 129.3083 | 20.9738 | 0.001877 | 0.001777 | -152.1493 | 0.23966 | 8.4038 | 0.003386 | 0.004636 | 0.00422 | 6.1e-05 | 0.001224 | 0.31542 | 1.3051 | 2.5481 | 101.01 | 20.7 | 18.018 | 66.7223 | 103.0502 | 64.3052 | 0.00436 | 1.0462 | 0.025442 | 24.3183 | -0.10777 | 0.14257 | 0.002042 | 0.00149 | 0.00149 | 3.9323
2019-03-20 12:44 | 1553100246.1625 | -337 | 5.2e-05 | -0.001573 | -2.2e-05 | 107.964 | 109.2087 | 129.4313 | 20.9747 | 0.001807 | 0.001759 | -152.149 | 0.23895 | 8.4029 | 0.003372 | 0.004781 | 0.004428 | 6.1e-05 | 0.001223 | 0.33356 | 1.3191 | 2.5482 | 101.01 | 20.69 | 18.011 | 65.0338 | 102.7148 | 63.8829 | 0.004356 | 1.0475 | 0.02547 | 24.3146 | -0.10253 | 0.14215 | 0.002043 | 0.001477 | 0.001477 | 3.9568
2019-03-20 12:44 | 1553100247.1626 | -336 | 5.7e-05 | -0.001567 | -2.2e-05 | 107.9851 | 109.1582 | 129.52 | 20.9743 | 0.00182 | 0.001817 | -152.1724 | 0.2392 | 8.402 | 0.003394 | 0.00417 | 0.003727 | 6.3e-05 | 0.001221 | 0.31996 | 1.3288 | 2.5481 | 101.01 | 20.7 | 18.008 | 65.3408 | 103.2353 | 65.28 | 0.004358 | 1.0432 | 0.025426 | 24.3733 | -0.096651 | 0.1423 | 0.002037 | 0.001482 | 0.001482 | 3.9461
2019-03-20 12:44 | 1553100248.1632 | -335 | 4.9e-05 | -0.001568 | -2.3e-05 | 107.9825 | 109.3402 | 129.3435 | 20.9747 | 0.001863 | 0.001729 | -152.1783 | 0.23937 | 8.4016 | 0.003336 | 0.004251 | 0.003975 | 6.3e-05 | 0.001219 | 0.32892 | 1.3312 | 2.5482 | 101.01 | 20.7 | 18.007 | 66.3815 | 101.8273 | 63.1613 | 0.004357 | 1.0498 | 0.025488 | 24.2803 | -0.11177 | 0.1424 | 0.002038 | 0.001521 | 0.001521 | 3.8486
2019-03-20 12:44 | 1553100249.1627 | -334 | 2.2e-05 | -0.001565 | -2.3e-05 | 107.9572 | 109.2319 | 129.3029 | 20.974 | 0.001846 | 0.001757 | -152.1709 | 0.23912 | 8.4026 | 0.003412 | 0.004926 | 0.004558 | 6.3e-05 | 0.001208 | 0.33826 | 1.2929 | 2.5481 | 101.01 | 20.69 | 18.004 | 65.9578 | 103.6667 | 63.8143 | 0.004354 | 1.0477 | 0.02546 | 24.3001 | -0.10499 | 0.14225 | 0.002035 | 0.001475 | 0.001475 | 3.9627
2019-03-20 12:44 | 1553100250.1634 | -333 | 7.9e-05 | -0.001563 | -2.2e-05 | 107.9783 | 109.2083 | 129.0184 | 20.9745 | 0.001841 | 0.001747 | -152.189 | 0.23961 | 8.4034 | 0.003362 | 0.004561 | 0.004195 | 6.3e-05 | 0.001199 | 0.33876 | 1.295 | 2.5479 | 101.01 | 20.69 | 18.004 | 65.8404 | 102.4632 | 63.5945 | 0.004354 | 1.0484 | 0.02544 | 24.2654 | -0.10132 | 0.14255 | 0.002033 | 0.001496 | 0.001496 | 3.9164
2019-03-20 12:44 | 1553100251.1625 | -332 | 1.8e-05 | -0.001567 | -2.2e-05 | 107.9883 | 109.1743 | 128.695 | 20.9756 | 0.001857 | 0.001732 | -152.1761 | 0.23919 | 8.4019 | 0.003368 | 0.004014 | 0.003626 | 6.4e-05 | 0.001197 | 0.30596 | 1.3362 | 2.548 | 101.01 | 20.69 | 18.002 | 66.2325 | 102.6137 | 63.2278 | 0.004353 | 1.0495 | 0.025422 | 24.2218 | -0.0977 | 0.14229 | 0.002036 | 0.001446 | 0.001446 | 4.0441
2019-03-20 12:44 | 1553100252.1635 | -331 | -4e-05 | -0.001566 | -2.3e-05 | 107.931 | 109.5368 | 128.8102 | 20.9746 | 0.001903 | 0.001734 | -152.1726 | 0.23965 | 8.4023 | 0.003389 | 0.004269 | 0.003903 | 6e-05 | 0.001197 | 0.31526 | 1.3299 | 2.548 | 101.01 | 20.7 | 18 | 67.3308 | 103.1169 | 63.276 | 0.004356 | 1.0494 | 0.025432 | 24.2343 | -0.1321 | 0.14257 | 0.002036 | 0.001436 | 0.001436 | 4.0815
2019-03-20 12:44 | 1553100253.1634 | -330 | 7.3e-05 | -0.001567 | -2.3e-05 | 107.9818 | 109.2659 | 128.8088 | 20.9741 | 0.00192 | 0.001681 | -152.1611 | 0.23817 | 8.4035 | 0.003407 | 0.004711 | 0.004291 | 6.2e-05 | 0.001198 | 0.32637 | 1.318 | 2.5482 | 101.01 | 20.7 | 17.999 | 67.7571 | 103.5456 | 61.9835 | 0.004355 | 1.0534 | 0.02548 | 24.1876 | -0.10573 | 0.14169 | 0.002037 | 0.001501 | 0.001501 | 3.8803
2019-03-20 12:44 | 1553100254.1635 | -329 | 0.000107 | -0.001567 | -2.3e-05 | 107.9508 | 109.2767 | 128.8834 | 20.9756 | 0.001924 | 0.001764 | -152.1406 | 0.23872 | 8.4009 | 0.003433 | 0.00481 | 0.004487 | 6.3e-05 | 0.001201 | 0.31758 | 1.3144 | 2.548 | 101.01 | 20.69 | 17.994 | 67.8358 | 104.1989 | 63.9884 | 0.004351 | 1.0472 | 0.025412 | 24.2669 | -0.10919 | 0.14201 | 0.002037 | 0.001527 | 0.001527 | 3.8231
2019-03-20 12:44 | 1553100255.1629 | -328 | 8.9e-05 | -0.001566 | -2.2e-05 | 107.9792 | 109.2954 | 129.0722 | 20.9753 | 0.001849 | 0.001734 | -152.1436 | 0.23902 | 8.4008 | 0.003481 | 0.004662 | 0.004251 | 6.4e-05 | 0.001202 | 0.32139 | 1.3183 | 2.5481 | 101.01 | 20.69 | 17.992 | 66.0505 | 105.3437 | 63.2829 | 0.004351 | 1.0494 | 0.025457 | 24.2592 | -0.10837 | 0.14219 | 0.002036 | 0.001508 | 0.001508 | 3.8756
2019-03-20 12:44 | 1553100256.1627 | -327 | 0.000104 | -0.001569 | -2.3e-05 | 107.9906 | 109.3272 | 129.2393 | 20.9753 | 0.00183 | 0.001761 | -152.1439 | 0.23857 | 8.401 | 0.00342 | 0.004231 | 0.003827 | 6.5e-05 | 0.001203 | 0.32549 | 1.3387 | 2.5481 | 101.01 | 20.69 | 17.99 | 65.5884 | 103.8676 | 63.9172 | 0.00435 | 1.0474 | 0.02545 | 24.2978 | -0.11003 | 0.14193 | 0.002039 | 0.001526 | 0.001526 | 3.8219
2019-03-20 12:44 | 1553100257.1621 | -326 | 4.9e-05 | -0.001572 | -2.3e-05 | 108.0067 | 109.2154 | 129.2137 | 20.9745 | 0.001886 | 0.001724 | -152.1388 | 0.239 | 8.403 | 0.003394 | 0.005032 | 0.004675 | 6.2e-05 | 0.001206 | 0.29966 | 1.2928 | 2.548 | 101.01 | 20.69 | 17.989 | 66.9174 | 103.2365 | 63.0386 | 0.00435 | 1.0501 | 0.02548 | 24.2637 | -0.099541 | 0.14218 | 0.002042 | 0.001502 | 0.001502 | 3.8898
2019-03-20 12:44 | 1553100258.1618 | -325 | 0.00011 | -0.001573 | -2.2e-05 | 108.0294 | 109.3238 | 129.2107 | 20.9765 | 0.001851 | 0.001732 | -152.1446 | 0.23991 | 8.3987 | 0.003478 | 0.004286 | 0.003921 | 6.2e-05 | 0.001199 | 0.323 | 1.2856 | 2.5482 | 101.01 | 20.69 | 17.989 | 66.0819 | 105.2864 | 63.2374 | 0.00435 | 1.0495 | 0.025472 | 24.2706 | -0.10654 | 0.14272 | 0.002042 | 0.001506 | 0.001506 | 3.896
2019-03-20 12:44 | 1553100259.1629 | -324 | 9.5e-05 | -0.001569 | -2.3e-05 | 108.0161 | 109.3017 | 129.0266 | 20.9756 | 0.001887 | 0.001752 | -152.1502 | 0.239 | 8.4013 | 0.003468 | 0.004875 | 0.004548 | 6.4e-05 | 0.001203 | 0.31725 | 1.2729 | 2.5479 | 101.01 | 20.69 | 17.992 | 66.9636 | 105.0236 | 63.6965 | 0.004351 | 1.0481 | 0.025437 | 24.2698 | -0.10583 | 0.14218 | 0.00204 | 0.001547 | 0.001547 | 3.7784
2019-03-20 12:44 | 1553100260.1633 | -323 | 6.1e-05 | -0.001569 | -2.3e-05 | 108.0504 | 109.2214 | 129.1794 | 20.9751 | 0.001888 | 0.001742 | -152.129 | 0.23956 | 8.4015 | 0.003487 | 0.004306 | 0.003896 | 6.5e-05 | 0.001205 | 0.32908 | 1.3437 | 2.5482 | 101.01 | 20.69 | 17.995 | 66.974 | 105.4984 | 63.4647 | 0.004352 | 1.0488 | 0.025461 | 24.2759 | -0.096415 | 0.14252 | 0.002039 | 0.001494 | 0.001494 | 3.9198
2019-03-20 12:44 | 1553100261.163 | -322 | 4.5e-05 | -0.001571 | -2.3e-05 | 108.0387 | 109.2853 | 129.428 | 20.9756 | 0.001832 | 0.001747 | -152.1545 | 0.23986 | 8.4023 | 0.003485 | 0.005193 | 0.004821 | 6.3e-05 | 0.001208 | 0.33588 | 1.2904 | 2.5481 | 101.01 | 20.69 | 17.999 | 65.6245 | 105.4484 | 63.58 | 0.004353 | 1.0485 | 0.025481 | 24.3034 | -0.10261 | 0.14269 | 0.002041 | 0.001481 | 0.001481 | 3.9606
2019-03-20 12:44 | 1553100262.163 | -321 | 5.4e-05 | -0.001569 | -2.2e-05 | 108.0135 | 109.338 | 129.5088 | 20.9756 | 0.001842 | 0.001772 | -152.16 | 0.23932 | 8.4009 | 0.003463 | 0.00492 | 0.004496 | 6.3e-05 | 0.001212 | 0.31061 | 1.2865 | 2.5482 | 101.01 | 20.69 | 18.007 | 65.8814 | 104.9088 | 64.1875 | 0.004355 | 1.0466 | 0.025466 | 24.3329 | -0.10901 | 0.14237 | 0.002039 | 0.001456 | 0.001456 | 4.0184
2019-03-20 12:44 | 1553100263.162 | -320 | 8.9e-05 | -0.001569 | -2.3e-05 | 108.0274 | 109.3687 | 129.6961 | 20.9766 | 0.001869 | 0.001752 | -152.1449 | 0.23893 | 8.4026 | 0.003413 | 0.004446 | 0.004108 | 6.2e-05 | 0.00121 | 0.2995 | 1.3275 | 2.548 | 101.01 | 20.69 | 18.012 | 66.5276 | 103.7123 | 63.7013 | 0.004356 | 1.0481 | 0.025503 | 24.3329 | -0.11037 | 0.14214 | 0.002039 | 0.001515 | 0.001515 | 3.8549
2019-03-20 12:44 | 1553100264.1625 | -319 | 5e-05 | -0.001569 | -2.2e-05 | 108.0636 | 109.3447 | 129.7078 | 20.9762 | 0.001874 | 0.00175 | -152.1672 | 0.2388 | 8.4005 | 0.003466 | 0.00455 | 0.004167 | 6.6e-05 | 0.001212 | 0.29878 | 1.3261 | 2.5481 | 101.01 | 20.69 | 18.025 | 66.6459 | 104.9811 | 63.6526 | 0.004359 | 1.0482 | 0.025506 | 24.3322 | -0.10541 | 0.14206 | 0.002039 | 0.001478 | 0.001478 | 3.9501
2019-03-20 12:44 | 1553100265.1621 | -318 | 6.3e-05 | -0.001569 | -2.2e-05 | 108.008 | 109.4501 | 129.5857 | 20.9761 | 0.001819 | 0.001741 | -152.1804 | 0.23915 | 8.4047 | 0.003347 | 0.005022 | 0.004645 | 6.4e-05 | 0.001212 | 0.33754 | 1.3211 | 2.5481 | 101.01 | 20.69 | 18.034 | 65.3307 | 102.1145 | 63.4509 | 0.004361 | 1.0489 | 0.025501 | 24.3135 | -0.11864 | 0.14227 | 0.002038 | 0.001471 | 0.001471 | 3.9752
2019-03-20 12:44 | 1553100266.1626 | -317 | 4.8e-05 | -0.001569 | -2.3e-05 | 108.0442 | 109.2745 | 129.2388 | 20.9759 | 0.001916 | 0.001742 | -152.1828 | 0.23923 | 8.4004 | 0.003412 | 0.004588 | 0.004284 | 6.8e-05 | 0.001207 | 0.33284 | 1.2909 | 2.548 | 101.01 | 20.69 | 18.044 | 67.6403 | 103.6707 | 63.4679 | 0.004364 | 1.0488 | 0.025467 | 24.2815 | -0.10127 | 0.14232 | 0.002039 | 0.001479 | 0.001479 | 3.954
2019-03-20 12:44 | 1553100267.1626 | -316 | 9.1e-05 | -0.00157 | -2.2e-05 | 108.0704 | 109.252 | 129.2189 | 20.9772 | 0.00184 | 0.001752 | -152.1784 | 0.23866 | 8.4004 | 0.003347 | 0.00463 | 0.004228 | 6.8e-05 | 0.001198 | 0.31686 | 1.3313 | 2.548 | 101.01 | 20.69 | 18.068 | 65.836 | 102.1068 | 63.7011 | 0.004369 | 1.0481 | 0.025456 | 24.2881 | -0.097265 | 0.14198 | 0.00204 | 0.001502 | 0.001502 | 3.8862
2019-03-20 12:44 | 1553100268.1633 | -315 | 0.00014 | -0.001572 | -2.3e-05 | 108.0565 | 109.2614 | 129.4278 | 20.9759 | 0.001876 | 0.00178 | -152.1857 | 0.23917 | 8.4011 | 0.003359 | 0.004624 | 0.004225 | 6.7e-05 | 0.001193 | 0.33268 | 1.3274 | 2.5482 | 101.01 | 20.69 | 18.079 | 66.6893 | 102.3869 | 64.3891 | 0.004372 | 1.0459 | 0.02545 | 24.3325 | -0.099181 | 0.14228 | 0.002042 | 0.001564 | 0.001564 | 3.7384
2019-03-20 12:44 | 1553100269.1638 | -314 | 4.3e-05 | -0.001574 | -2.3e-05 | 108.0569 | 109.4363 | 129.6625 | 20.9766 | 0.001842 | 0.00179 | -152.181 | 0.23875 | 8.4025 | 0.00335 | 0.004459 | 0.004116 | 6.7e-05 | 0.001192 | 0.31708 | 1.3788 | 2.5481 | 101.01 | 20.69 | 18.091 | 65.8713 | 102.1668 | 64.6218 | 0.004375 | 1.0452 | 0.025465 | 24.363 | -0.11346 | 0.14203 | 0.002044 | 0.001502 | 0.001502 | 3.8867
2019-03-20 12:44 | 1553100270.1624 | -313 | 3.3e-05 | -0.00157 | -2.3e-05 | 108.0261 | 109.3141 | 129.9061 | 20.9771 | 0.001851 | 0.001719 | -152.1958 | 0.23898 | 8.4008 | 0.003391 | 0.004842 | 0.00448 | 6.8e-05 | 0.001188 | 0.30182 | 1.379 | 2.5481 | 101.01 | 20.69 | 18.116 | 66.0849 | 103.1746 | 62.9248 | 0.004381 | 1.0505 | 0.025553 | 24.3245 | -0.10601 | 0.14217 | 0.00204 | 0.001482 | 0.001482 | 3.9425
2019-03-20 12:44 | 1553100271.1635 | -312 | 6.7e-05 | -0.001572 | -2.2e-05 | 108.1946 | 109.2206 | 130.0871 | 20.9766 | 0.00185 | 0.00174 | -152.2036 | 0.23838 | 8.3998 | 0.003389 | 0.004866 | 0.004542 | 6.6e-05 | 0.001185 | 0.31868 | 1.368 | 2.548 | 101.01 | 20.69 | 18.127 | 66.0719 | 103.115 | 63.4115 | 0.004384 | 1.049 | 0.025552 | 24.3591 | -0.084422 | 0.14181 | 0.002042 | 0.001496 | 0.001496 | 3.8962
2019-03-20 12:45 | 1553100272.1616 | -311 | 5.6e-05 | -0.001572 | -2.2e-05 | 108.0662 | 109.5118 | 130.4223 | 20.9767 | 0.00184 | 0.001767 | -152.1977 | 0.23846 | 8.4012 | 0.003471 | 0.004795 | 0.004429 | 6.6e-05 | 0.001182 | 0.33555 | 1.3026 | 2.5481 | 101.01 | 20.69 | 18.148 | 65.8178 | 105.1207 | 64.0699 | 0.004389 | 1.0469 | 0.02556 | 24.4143 | -0.11886 | 0.14186 | 0.002042 | 0.001487 | 0.001487 | 3.9216
2019-03-20 12:45 | 1553100273.162 | -310 | 6.8e-05 | -0.001572 | -2.3e-05 | 108.1044 | 109.4613 | 130.4819 | 20.9778 | 0.001891 | 0.00175 | -152.2114 | 0.23877 | 8.4014 | 0.00348 | 0.004368 | 0.004023 | 6.5e-05 | 0.001189 | 0.335 | 1.3271 | 2.5481 | 101.01 | 20.69 | 18.158 | 67.061 | 105.3294 | 63.6626 | 0.004391 | 1.0482 | 0.025581 | 24.4051 | -0.11157 | 0.14204 | 0.002042 | 0.001513 | 0.001513 | 3.8576
2019-03-20 12:45 | 1553100274.1626 | -309 | 6.8e-05 | -0.001575 | -2.3e-05 | 108.1041 | 109.1896 | 130.6878 | 20.9773 | 0.001871 | 0.001728 | -152.1865 | 0.2392 | 8.403 | 0.003435 | 0.00485 | 0.004475 | 6.7e-05 | 0.00119 | 0.30613 | 1.3557 | 2.5481 | 101.01 | 20.69 | 18.166 | 66.5666 | 104.2317 | 63.1327 | 0.004393 | 1.0498 | 0.025622 | 24.4051 | -0.08937 | 0.1423 | 0.002045 | 0.001543 | 0.001543 | 3.7891
2019-03-20 12:45 | 1553100275.1627 | -308 | 7.1e-05 | -0.001573 | -2.3e-05 | 108.0598 | 109.307 | 130.7615 | 20.9768 | 0.001896 | 0.001737 | -152.1652 | 0.23929 | 8.4008 | 0.003405 | 0.004803 | 0.004421 | 6.6e-05 | 0.001188 | 0.31653 | 1.3703 | 2.5481 | 101.01 | 20.69 | 18.18 | 67.1661 | 103.5011 | 63.3554 | 0.004396 | 1.0492 | 0.02562 | 24.4201 | -0.10265 | 0.14235 | 0.002043 | 0.001516 | 0.001516 | 3.8592
2019-03-20 12:45 | 1553100276.1628 | -307 | 7.2e-05 | -0.001571 | -2.3e-05 | 108.1196 | 109.3316 | 130.9255 | 20.9776 | 0.001878 | 0.001798 | -152.1409 | 0.23912 | 8.404 | 0.003405 | 0.004258 | 0.003958 | 6.5e-05 | 0.001192 | 0.31581 | 1.3295 | 2.5481 | 101 | 20.69 | 18.186 | 66.7385 | 103.5192 | 64.8226 | 0.004398 | 1.0446 | 0.025581 | 24.4886 | -0.099713 | 0.14225 | 0.002041 | 0.001523 | 0.001523 | 3.8396
2019-03-20 12:45 | 1553100277.1618 | -306 | 1.3e-05 | -0.00157 | -6e-06 | 108.0889 | 109.3892 | 130.8352 | 20.9778 | 0.001845 | 0.001696 | -152.1564 | 0.23851 | 8.4024 | 0.003355 | 0.004285 | 0.003935 | 6.3e-05 | 0.001187 | 0.3314 | 1.3346 | 2.5481 | 101 | 20.69 | 18.195 | 65.9364 | 102.3006 | 62.3536 | 0.0044 | 1.0523 | 0.025666 | 24.3906 | -0.10696 | 0.14189 | 0.002023 | 0.000397 | 0.000397 | 14.6732
2019-03-20 12:45 | 1553100278.1618 | -305 | 4.9e-05 | -0.001573 | 0 | 108.1076 | 109.4322 | 130.496 | 20.9775 | 0.001845 | 0.001766 | -152.1453 | 0.23879 | 8.4044 | 0.003361 | 0.004623 | 0.004283 | 6.3e-05 | 0.001185 | 0.32764 | 1.3171 | 2.5481 | 101 | 20.69 | 18.199 | 65.9465 | 102.4536 | 64.0543 | 0.004401 | 1.047 | 0.025568 | 24.4206 | -0.10892 | 0.14205 | 0.00202 | 4.4e-05 | 4.4e-05 | 132.0803
2019-03-20 12:45 | 1553100279.1619 | -304 | 4.5e-05 | -0.00039 | 0 | 108.0585 | 109.3204 | 130.5983 | 20.9772 | 0.00187 | 0.001707 | -152.1401 | 0.23856 | 8.4025 | 0.003394 | 0.005286 | 0.004786 | 6.4e-05 | 0.001182 | 0.31725 | 1.3199 | 2.548 | 101 | 20.69 | 18.201 | 66.5484 | 103.2374 | 62.6193 | 0.004402 | 1.0515 | 0.025632 | 24.3781 | -0.10385 | 0.14192 | 0.000841 | 1.6e-05 | 1.6e-05 | 363.7008
2019-03-20 12:45 | 1553100280.1619 | -303 | 5.3e-05 | 4e-06 | 0 | 108.0649 | 109.2954 | 130.4998 | 20.9774 | 0.001858 | 0.001799 | -152.1447 | 0.23814 | 8.4023 | 0.003425 | 0.004552 | 0.004221 | 6.2e-05 | 0.001182 | 0.33716 | 1.2838 | 2.5481 | 101 | 20.69 | 18.208 | 66.2608 | 103.9843 | 64.8435 | 0.004404 | 1.0445 | 0.025538 | 24.4495 | -0.10127 | 0.14167 | 0.000446 | 1.2e-05 | 1.2e-05 | 471.8227
2019-03-20 12:45 | 1553100281.1621 | -302 | 4.8e-05 | 6e-06 | 0 | 108.0756 | 109.2035 | 130.6401 | 20.9776 | 0.001914 | 0.0017 | -152.1149 | 0.23891 | 8.4002 | 0.003388 | 0.00481 | 0.004456 | 6.4e-05 | 0.001186 | 0.2995 | 1.3101 | 2.5481 | 101 | 20.69 | 18.208 | 67.6131 | 103.0987 | 62.4478 | 0.004404 | 1.052 | 0.025643 | 24.3758 | -0.092858 | 0.14213 | 0.000445 | 3.4e-05 | 3.4e-05 | 171.1354
2019-03-20 12:45 | 1553100282.1631 | -301 | 1.3e-05 | 1e-06 | -1e-06 | 108.0442 | 109.5249 | 130.6075 | 20.9783 | 0.001863 | 0.00175 | -152.1034 | 0.23899 | 8.4002 | 0.003375 | 0.004708 | 0.004346 | 6.8e-05 | 0.001179 | 0.31564 | 1.314 | 2.548 | 101 | 20.69 | 18.21 | 66.3701 | 102.7858 | 63.6678 | 0.004404 | 1.0482 | 0.025593 | 24.417 | -0.12174 | 0.14217 | 0.00045 | 4.2e-05 | 4.2e-05 | 139.2968
2019-03-20 12:45 | 1553100283.1625 | -300 | 2.9e-05 | 2e-06 | 0 | 108.1581 | 109.4972 | 130.665 | 20.9782 | 0.001859 | 0.001754 | -152.1121 | 0.23888 | 8.3998 | 0.003378 | 0.004496 | 0.004044 | 6.6e-05 | 0.001185 | 0.31421 | 1.35 | 2.548 | 101 | 20.69 | 18.21 | 66.2827 | 102.8652 | 63.7544 | 0.004404 | 1.0479 | 0.025596 | 24.4255 | -0.11006 | 0.14211 | 0.000448 | 1.3e-05 | 1.3e-05 | 448.1489
2019-03-20 12:45 | 1553100284.1622 | -299 | 4.4e-05 | 3e-06 | 0 | 108.1354 | 109.4961 | 130.5103 | 20.978 | 0.001862 | 0.001729 | -152.1405 | 0.2393 | 8.4016 | 0.00337 | 0.004458 | 0.004153 | 6.5e-05 | 0.001182 | 0.3147 | 1.3568 | 2.5479 | 101 | 20.69 | 18.213 | 66.3433 | 102.6574 | 63.1488 | 0.004405 | 1.0498 | 0.025604 | 24.3891 | -0.11185 | 0.14236 | 0.000447 | 5e-06 | 5e-06 | 1104.6468
2019-03-20 12:45 | 1553100285.1623 | -298 | 4e-05 | -1e-06 | 0 | 108.0979 | 109.477 | 130.2885 | 20.9783 | 0.00185 | 0.001711 | -152.1515 | 0.23908 | 8.4016 | 0.003343 | 0.004331 | 0.004003 | 6.4e-05 | 0.001181 | 0.3246 | 1.3256 | 2.5481 | 101 | 20.69 | 18.22 | 66.0564 | 102.0106 | 62.7321 | 0.004407 | 1.0511 | 0.025598 | 24.3533 | -0.11339 | 0.14223 | 0.000451 | 0 | 0 | -14206.7751
2019-03-20 12:45 | 1553100286.1623 | -297 | -1.9e-05 | 2e-06 | 0 | 108.0614 | 109.4415 | 129.9667 | 20.9781 | 0.001845 | 0.001729 | -152.1615 | 0.23895 | 8.4005 | 0.00339 | 0.004661 | 0.004245 | 6.3e-05 | 0.001175 | 0.31741 | 1.3326 | 2.5481 | 101 | 20.69 | 18.223 | 65.947 | 103.1433 | 63.1568 | 0.004407 | 1.0498 | 0.02555 | 24.3386 | -0.11352 | 0.14215 | 0.000448 | -1.5e-05 | -1.5e-05 | -377.9424
2019-03-20 12:45 | 1553100287.1617 | -296 | 1e-06 | 2e-06 | 0 | 108.1853 | 109.4273 | 130.0019 | 20.9787 | 0.001841 | 0.001737 | -152.1532 | 0.23894 | 8.4003 | 0.003308 | 0.004323 | 0.003941 | 6.2e-05 | 0.001173 | 0.32853 | 1.3261 | 2.5481 | 101 | 20.69 | 18.228 | 65.8445 | 101.1584 | 63.3363 | 0.004408 | 1.0492 | 0.025547 | 24.3484 | -0.10211 | 0.14214 | 0.000448 | -4e-06 | -4e-06 | -1453.442
2019-03-20 12:45 | 1553100288.1624 | -295 | 4.9e-05 | 1e-06 | 0 | 108.1697 | 109.3911 | 130.185 | 20.979 | 0.001887 | 0.001743 | -152.1593 | 0.23954 | 8.4016 | 0.003385 | 0.004697 | 0.004349 | 6.3e-05 | 0.001177 | 0.33572 | 1.3433 | 2.5482 | 101 | 20.69 | 18.237 | 66.9621 | 103.0269 | 63.4997 | 0.004411 | 1.0487 | 0.025558 | 24.3714 | -0.10043 | 0.1425 | 0.000449 | -8e-06 | -8e-06 | -762.676
2019-03-20 12:45 | 1553100289.1625 | -294 | 5.7e-05 | 0 | 0 | 108.2092 | 109.4322 | 130.2216 | 20.9792 | 0.001855 | 0.001758 | -152.1726 | 0.2391 | 8.4002 | 0.003409 | 0.004328 | 0.003958 | 6e-05 | 0.001174 | 0.33284 | 1.3086 | 2.548 | 101 | 20.69 | 18.242 | 66.1846 | 103.6032 | 63.8608 | 0.004412 | 1.0476 | 0.025548 | 24.3879 | -0.10053 | 0.14224 | 0.000449 | -1.7e-05 | -1.7e-05 | -336.796
2019-03-20 12:45 | 1553100290.1617 | -293 | 2.4e-05 | 1e-06 | 0 | 108.2292 | 109.5865 | 130.3014 | 20.9787 | 0.001822 | 0.001718 | -152.1761 | 0.23935 | 8.4013 | 0.003442 | 0.005166 | 0.004802 | 6.2e-05 | 0.001174 | 0.30309 | 1.3271 | 2.548 | 101 | 20.69 | 18.253 | 65.3903 | 104.4016 | 62.8818 | 0.004415 | 1.0506 | 0.025593 | 24.3599 | -0.11148 | 0.14239 | 0.000449 | -1.4e-05 | -1.4e-05 | -406.5423
2019-03-20 12:45 | 1553100291.1626 | -292 | 6.6e-05 | 3e-06 | 0 | 108.2014 | 109.6067 | 130.0897 | 20.9794 | 0.001822 | 0.001686 | -152.1723 | 0.23911 | 8.4013 | 0.003396 | 0.004428 | 0.004032 | 5.9e-05 | 0.001178 | 0.31653 | 1.3547 | 2.548 | 101 | 20.69 | 18.257 | 65.3941 | 103.2955 | 62.1223 | 0.004415 | 1.053 | 0.025601 | 24.3126 | -0.11542 | 0.14225 | 0.000447 | 2e-06 | 2e-06 | 3503.6968
2019-03-20 12:45 | 1553100292.1619 | -291 | 1.7e-05 | 4e-06 | 0 | 108.1768 | 109.6938 | 129.8127 | 20.9793 | 0.001838 | 0.001739 | -152.2032 | 0.2386 | 8.4009 | 0.003457 | 0.003826 | 0.003456 | 6.3e-05 | 0.001179 | 0.32947 | 1.3104 | 2.5481 | 101 | 20.69 | 18.26 | 65.7717 | 104.7606 | 63.3869 | 0.004416 | 1.0491 | 0.025526 | 24.3325 | -0.12456 | 0.14194 | 0.000446 | -2.4e-05 | -2.4e-05 | -238.1583
2019-03-20 12:45 | 1553100293.1616 | -290 | 9e-06 | 1e-06 | 0 | 108.1688 | 109.5787 | 129.5789 | 20.978 | 0.001814 | 0.001731 | -152.1947 | 0.23927 | 8.4027 | 0.003481 | 0.004739 | 0.00432 | 6.2e-05 | 0.001184 | 0.33859 | 1.3557 | 2.548 | 101 | 20.69 | 18.268 | 65.1952 | 105.3484 | 63.1926 | 0.004418 | 1.0497 | 0.02551 | 24.3035 | -0.11583 | 0.14234 | 0.000449 | -1.2e-05 | -1.2e-05 | -497.5039
2019-03-20 12:45 | 1553100294.1629 | -289 | 2.1e-05 | 5e-06 | 0 | 108.2849 | 109.6785 | 129.5991 | 20.9798 | 0.001866 | 0.001775 | -152.1953 | 0.23932 | 8.4019 | 0.00342 | 0.00467 | 0.00434 | 6.5e-05 | 0.001188 | 0.34578 | 1.3584 | 2.5482 | 101 | 20.69 | 18.271 | 66.448 | 103.8835 | 64.2572 | 0.004419 | 1.0463 | 0.025472 | 24.3439 | -0.11438 | 0.14237 | 0.000445 | 1e-05 | 1e-05 | 593.0667
2019-03-20 12:45 | 1553100295.1628 | -288 | 2e-05 | -1e-06 | 0 | 108.2418 | 109.4946 | 129.7472 | 20.9797 | 0.001854 | 0.001796 | -152.1872 | 0.2387 | 8.403 | 0.003439 | 0.004653 | 0.00427 | 5.9e-05 | 0.001187 | 0.35762 | 1.3186 | 2.548 | 101 | 20.69 | 18.276 | 66.1631 | 104.3371 | 64.7746 | 0.00442 | 1.0447 | 0.025467 | 24.3764 | -0.10293 | 0.142 | 0.000452 | 5e-06 | 5e-06 | 1219.6057
2019-03-20 12:45 | 1553100296.1632 | -287 | -2.5e-05 | -2e-06 | 1e-06 | 108.2448 | 109.551 | 130.0949 | 20.9785 | 0.001834 | 0.001752 | -152.1934 | 0.23932 | 8.4019 | 0.003441 | 0.004858 | 0.004385 | 6.4e-05 | 0.001189 | 0.31006 | 1.3553 | 2.548 | 101 | 20.69 | 18.278 | 65.6898 | 104.3863 | 63.7077 | 0.004421 | 1.0481 | 0.025542 | 24.3705 | -0.1073 | 0.14237 | 0.000451 | -5e-05 | -5e-05 | -117.2896
2019-03-20 12:45 | 1553100297.1635 | -286 | 1.6e-05 | 4e-06 | 0 | 108.3156 | 109.3855 | 130.3061 | 20.9792 | 0.001859 | 0.001758 | -152.1796 | 0.23939 | 8.4011 | 0.003465 | 0.004385 | 0.004068 | 6.5e-05 | 0.001182 | 0.34612 | 1.3743 | 2.5482 | 101 | 20.69 | 18.279 | 66.2851 | 104.9697 | 63.8528 | 0.004421 | 1.0476 | 0.025557 | 24.3955 | -0.087915 | 0.14242 | 0.000446 | -5e-06 | -5e-06 | -1290.0096
2019-03-20 12:45 | 1553100298.1629 | -285 | 2.6e-05 | 3e-06 | 0 | 108.2441 | 109.6677 | 130.2261 | 20.9785 | 0.001925 | 0.001756 | -152.1748 | 0.23952 | 8.4022 | 0.00342 | 0.00487 | 0.004452 | 6.2e-05 | 0.001187 | 0.32173 | 1.3329 | 2.5481 | 101 | 20.69 | 18.281 | 67.8629 | 103.8634 | 63.8124 | 0.004421 | 1.0477 | 0.025551 | 24.3866 | -0.11686 | 0.14249 | 0.000447 | 6e-06 | 6e-06 | 919.798
2019-03-20 12:45 | 1553100299.1624 | -284 | -4.7e-05 | 3e-06 | 0 | 108.3023 | 109.6512 | 130.0428 | 20.9797 | 0.001878 | 0.00173 | -152.1733 | 0.23898 | 8.4019 | 0.003449 | 0.004437 | 0.004079 | 6.4e-05 | 0.001186 | 0.34252 | 1.3412 | 2.5481 | 101 | 20.69 | 18.281 | 66.7437 | 104.5778 | 63.1724 | 0.004421 | 1.0497 | 0.025557 | 24.3463 | -0.11071 | 0.14217 | 0.000447 | -3e-06 | -3e-06 | -1803.9778
2019-03-20 12:45 | 1553100300.163 | -283 | 5.5e-05 | 3e-06 | 0 | 108.2528 | 109.7615 | 129.9752 | 20.9794 | 0.001823 | 0.001731 | -152.1833 | 0.2393 | 8.3997 | 0.003451 | 0.004435 | 0.004053 | 6.4e-05 | 0.001188 | 0.31238 | 1.3474 | 2.548 | 101 | 20.69 | 18.277 | 65.4129 | 104.6373 | 63.2134 | 0.00442 | 1.0496 | 0.025549 | 24.3414 | -0.12379 | 0.14236 | 0.000447 | 2.3e-05 | 2.3e-05 | 257.3496
2019-03-20 12:45 | 1553100301.1621 | -282 | 4.7e-05 | 4e-06 | 0 | 108.271 | 109.7517 | 129.8679 | 20.9801 | 0.001822 | 0.001687 | -152.1951 | 0.23911 | 8.4018 | 0.003391 | 0.00463 | 0.004279 | 6.5e-05 | 0.001194 | 0.323 | 1.3045 | 2.548 | 101 | 20.69 | 18.271 | 65.3893 | 103.1775 | 62.1488 | 0.004419 | 1.0529 | 0.025579 | 24.2928 | -0.1215 | 0.14225 | 0.000446 | 2.2e-05 | 2.2e-05 | 262.0816
2019-03-20 12:45 | 1553100302.1631 | -281 | 6.5e-05 | 1e-06 | 0 | 108.3143 | 109.6744 | 129.8885 | 20.9793 | 0.001828 | 0.001747 | -152.192 | 0.23914 | 8.4014 | 0.003415 | 0.004172 | 0.003762 | 6.3e-05 | 0.001199 | 0.30414 | 1.2867 | 2.5479 | 101 | 20.7 | 18.266 | 65.5453 | 103.7464 | 63.5992 | 0.00442 | 1.0484 | 0.025525 | 24.3473 | -0.11162 | 0.14226 | 0.000449 | 8e-06 | 8e-06 | 703.3847
2019-03-20 12:45 | 1553100303.1626 | -280 | 1.1e-05 | 4e-06 | 0 | 108.3466 | 109.8007 | 129.9956 | 20.9803 | 0.001834 | 0.001742 | -152.1835 | 0.23846 | 8.4025 | 0.003291 | 0.004263 | 0.003943 | 6.4e-05 | 0.0012 | 0.33157 | 1.2978 | 2.5481 | 101 | 20.69 | 18.255 | 65.6752 | 100.7525 | 63.4603 | 0.004415 | 1.0488 | 0.025541 | 24.3523 | -0.11924 | 0.14186 | 0.000446 | 8e-06 | 8e-06 | 711.3541
2019-03-20 12:45 | 1553100304.1623 | -279 | 5.6e-05 | 4e-06 | 0 | 108.4097 | 109.7633 | 130.1305 | 20.9797 | 0.00188 | 0.001742 | -152.1771 | 0.23921 | 8.4027 | 0.003418 | 0.004286 | 0.003961 | 6.2e-05 | 0.0012 | 0.34379 | 1.3631 | 2.548 | 101 | 20.7 | 18.25 | 66.781 | 103.82 | 63.4629 | 0.004417 | 1.0488 | 0.025554 | 24.365 | -0.11099 | 0.1423 | 0.000446 | 1.1e-05 | 1.1e-05 | 518.9432
2019-03-20 12:45 | 1553100305.1623 | -278 | 7e-06 | 2e-06 | 1e-06 | 108.4045 | 109.7835 | 130.1594 | 20.9814 | 0.001785 | 0.001719 | -152.1833 | 0.23913 | 8.4022 | 0.003455 | 0.004092 | 0.003745 | 6.3e-05 | 0.0012 | 0.32516 | 1.3317 | 2.548 | 101 | 20.69 | 18.245 | 64.5179 | 104.7289 | 62.9265 | 0.004413 | 1.0505 | 0.025578 | 24.3483 | -0.11306 | 0.14226 | 0.000448 | -3.4e-05 | -3.4e-05 | -172.1397
2019-03-20 12:45 | 1553100306.1614 | -277 | 4.9e-05 | 1e-06 | 0 | 108.3842 | 109.7009 | 130.0796 | 20.9795 | 0.001814 | 0.001739 | -152.1728 | 0.23878 | 8.3985 | 0.003464 | 0.004542 | 0.0041 | 6.4e-05 | 0.001202 | 0.33212 | 1.3148 | 2.548 | 101 | 20.69 | 18.231 | 65.1929 | 104.9402 | 63.3886 | 0.004409 | 1.049 | 0.025552 | 24.3575 | -0.10801 | 0.14205 | 0.000448 | -5e-06 | -5e-06 | -1196.8603
2019-03-20 12:45 | 1553100307.1625 | -276 | 4e-05 | 1e-06 | 0 | 108.3754 | 109.785 | 130.0996 | 20.9804 | 0.001886 | 0.001762 | -152.1783 | 0.2387 | 8.4045 | 0.003422 | 0.004302 | 0.003905 | 6.5e-05 | 0.001202 | 0.30237 | 1.2843 | 2.5479 | 101 | 20.69 | 18.224 | 66.937 | 103.929 | 63.9634 | 0.004407 | 1.0473 | 0.025532 | 24.3802 | -0.11559 | 0.142 | 0.000449 | -2e-06 | -2e-06 | -3450.6646
2019-03-20 12:45 | 1553100308.1625 | -275 | 5.2e-05 | 2e-06 | 0 | 108.3977 | 109.8044 | 130.1547 | 20.9799 | 0.001869 | 0.001754 | -152.1786 | 0.238 | 8.4007 | 0.003462 | 0.004067 | 0.003732 | 6.5e-05 | 0.001198 | 0.32676 | 1.3071 | 2.5481 | 101 | 20.7 | 18.209 | 66.5239 | 104.9044 | 63.7473 | 0.004407 | 1.0479 | 0.025546 | 24.3775 | -0.11533 | 0.14159 | 0.000448 | -2e-06 | -2e-06 | -3866.9963
2019-03-20 12:45 | 1553100309.1625 | -274 | 4.7e-05 | 6e-06 | 0 | 108.4268 | 109.7839 | 130.3411 | 20.9806 | 0.00188 | 0.001769 | -152.1625 | 0.2388 | 8.4027 | 0.003441 | 0.00477 | 0.004371 | 6.5e-05 | 0.001199 | 0.31957 | 1.338 | 2.5482 | 101 | 20.7 | 18.202 | 66.7968 | 104.3964 | 64.1261 | 0.004405 | 1.0468 | 0.02555 | 24.4087 | -0.11125 | 0.14206 | 0.000443 | -1e-05 | -1e-05 | -595.4419
2019-03-20 12:45 | 1553100310.1619 | -273 | 6.9e-05 | 1e-06 | 0 | 108.4178 | 109.6609 | 130.0751 | 20.9799 | 0.001804 | 0.001738 | -152.1393 | 0.23881 | 8.404 | 0.003478 | 0.004687 | 0.004293 | 6.4e-05 | 0.0012 | 0.33699 | 1.3362 | 2.5481 | 101 | 20.7 | 18.195 | 64.9675 | 105.2875 | 63.3795 | 0.004403 | 1.0491 | 0.025552 | 24.3568 | -0.10198 | 0.14207 | 0.000449 | 1.5e-05 | 1.5e-05 | 377.3427
2019-03-20 12:45 | 1553100311.1626 | -272 | 3.4e-05 | 2e-06 | 0 | 108.4097 | 109.7783 | 130.1824 | 20.9808 | 0.001798 | 0.001757 | -152.1369 | 0.23867 | 8.3999 | 0.003415 | 0.004543 | 0.004132 | 6.5e-05 | 0.001199 | 0.35187 | 1.3132 | 2.548 | 101 | 20.7 | 18.195 | 64.8102 | 103.7504 | 63.8349 | 0.004403 | 1.0477 | 0.025545 | 24.3833 | -0.11221 | 0.14199 | 0.000448 | 5e-06 | 5e-06 | 1145.2061
2019-03-20 12:45 | 1553100312.1627 | -271 | 2.6e-05 | -2e-06 | 1e-06 | 108.367 | 109.6576 | 130.3416 | 20.9805 | 0.001827 | 0.00168 | -152.1173 | 0.24025 | 8.3988 | 0.0034 | 0.004548 | 0.004173 | 6.6e-05 | 0.001201 | 0.33268 | 1.353 | 2.5481 | 101 | 20.7 | 18.181 | 65.5256 | 103.4026 | 61.9818 | 0.0044 | 1.0535 | 0.025632 | 24.331 | -0.10589 | 0.14292 | 0.000451 | -4e-05 | -4e-05 | -145.2536
2019-03-20 12:45 | 1553100313.1621 | -270 | 2.5e-05 | -3e-06 | 0 | 108.4317 | 109.8841 | 129.9282 | 20.9813 | 0.001825 | 0.00173 | -152.1272 | 0.23836 | 8.4 | 0.003329 | 0.004707 | 0.004321 | 6.7e-05 | 0.001202 | 0.32372 | 1.355 | 2.5481 | 101 | 20.7 | 18.178 | 65.4631 | 101.6626 | 63.1725 | 0.004399 | 1.0497 | 0.025546 | 24.3355 | -0.11901 | 0.1418 | 0.000453 | -1.7e-05 | -1.7e-05 | -339.0287
2019-03-20 12:45 | 1553100314.1628 | -269 | 7.7e-05 | -1e-06 | 0 | 108.3838 | 109.7551 | 129.6817 | 20.9808 | 0.001817 | 0.001774 | -152.1044 | 0.23892 | 8.4013 | 0.00347 | 0.004571 | 0.004225 | 6.3e-05 | 0.001201 | 0.29717 | 1.2988 | 2.548 | 101 | 20.7 | 18.175 | 65.2689 | 105.0949 | 64.2512 | 0.004398 | 1.0464 | 0.02548 | 24.3514 | -0.11246 | 0.14214 | 0.000451 | -1.1e-05 | -1.1e-05 | -527.9909
2019-03-20 12:45 | 1553100315.1618 | -268 | 8.5e-05 | 4e-06 | 0 | 108.4834 | 109.8784 | 129.6627 | 20.9808 | 0.00186 | 0.001784 | -152.0969 | 0.23906 | 8.4034 | 0.00343 | 0.004032 | 0.003631 | 6.4e-05 | 0.001197 | 0.33588 | 1.3281 | 2.548 | 101 | 20.7 | 18.171 | 66.3196 | 104.1086 | 64.4807 | 0.004397 | 1.0457 | 0.02547 | 24.3579 | -0.11429 | 0.14222 | 0.000446 | -1.8e-05 | -1.8e-05 | -331.6184
2019-03-20 12:45 | 1553100316.1623 | -267 | 0.000104 | 2e-06 | 0 | 108.4893 | 109.8422 | 129.705 | 20.9805 | 0.001848 | 0.001725 | -152.1114 | 0.23889 | 8.4033 | 0.003504 | 0.004498 | 0.004156 | 6.2e-05 | 0.001198 | 0.31868 | 1.3075 | 2.548 | 101 | 20.7 | 18.168 | 66.0242 | 105.9089 | 63.0499 | 0.004397 | 1.0501 | 0.025528 | 24.3102 | -0.11085 | 0.14211 | 0.000448 | 2.1e-05 | 2.1e-05 | 278.5871
2019-03-20 12:45 | 1553100317.1631 | -266 | 2.5e-05 | 0 | 0 | 108.5711 | 109.7723 | 129.5015 | 20.9817 | 0.001844 | 0.001723 | -152.103 | 0.2384 | 8.4014 | 0.003462 | 0.004552 | 0.004148 | 6.7e-05 | 0.001197 | 0.335 | 1.2873 | 2.548 | 101 | 20.7 | 18.164 | 65.9238 | 104.8856 | 63.0221 | 0.004396 | 1.0502 | 0.025509 | 24.2901 | -0.098418 | 0.14182 | 0.00045 | -2.2e-05 | -2.2e-05 | -268.9928
2019-03-20 12:45 | 1553100318.162 | -265 | 5.8e-05 | -2e-06 | 0 | 108.5776 | 110.0235 | 129.7256 | 20.9819 | 0.001909 | 0.001765 | -152.0942 | 0.23976 | 8.4005 | 0.003419 | 0.004471 | 0.004134 | 6.5e-05 | 0.001202 | 0.33428 | 1.3225 | 2.548 | 101 | 20.7 | 18.163 | 67.4978 | 103.8544 | 64.0234 | 0.004395 | 1.0471 | 0.025493 | 24.3473 | -0.11833 | 0.14263 | 0.000453 | 1.5e-05 | 1.5e-05 | 386.5024
2019-03-20 12:45 | 1553100319.1611 | -264 | 4.5e-05 | 1e-06 | 0 | 108.4977 | 109.8482 | 129.709 | 20.9815 | 0.001867 | 0.001697 | -152.1132 | 0.23969 | 8.4032 | 0.003439 | 0.004192 | 0.00387 | 6.4e-05 | 0.0012 | 0.3454 | 1.3017 | 2.5479 | 101 | 20.7 | 18.163 | 66.4754 | 104.35 | 62.3805 | 0.004395 | 1.0522 | 0.025554 | 24.2864 | -0.11065 | 0.14259 | 0.00045 | 3.1e-05 | 3.1e-05 | 187.7279
2019-03-20 12:45 | 1553100320.1624 | -263 | 6.1e-05 | 3e-06 | 0 | 108.5284 | 109.813 | 129.9803 | 20.9822 | 0.001883 | 0.00173 | -152.103 | 0.23888 | 8.4011 | 0.003409 | 0.004916 | 0.004523 | 6.2e-05 | 0.001203 | 0.31686 | 1.3438 | 2.5481 | 101 | 20.7 | 18.16 | 66.8599 | 103.6235 | 63.1834 | 0.004395 | 1.0497 | 0.02555 | 24.3408 | -0.10525 | 0.14211 | 0.000447 | 3.2e-05 | 3.2e-05 | 182.3393
2019-03-20 12:45 | 1553100321.1611 | -262 | 6.3e-05 | 1e-06 | 0 | 108.5436 | 109.9992 | 129.9024 | 20.9823 | 0.001845 | 0.001729 | -152.1041 | 0.23908 | 8.4042 | 0.003451 | 0.004168 | 0.00383 | 6.6e-05 | 0.001202 | 0.34291 | 1.3515 | 2.5482 | 101 | 20.7 | 18.159 | 65.9394 | 104.6314 | 63.1578 | 0.004394 | 1.0498 | 0.025544 | 24.3326 | -0.11915 | 0.14223 | 0.000449 | 1.3e-05 | 1.3e-05 | 448.5977
2019-03-20 12:45 | 1553100322.1613 | -261 | 6.8e-05 | 1e-06 | 0 | 108.5656 | 109.8097 | 129.5167 | 20.982 | 0.001863 | 0.001757 | -152.1224 | 0.23905 | 8.4011 | 0.003424 | 0.004761 | 0.004379 | 6.5e-05 | 0.0012 | 0.33356 | 1.3143 | 2.548 | 101 | 20.7 | 18.158 | 66.3932 | 103.9754 | 63.8346 | 0.004394 | 1.0477 | 0.02548 | 24.3209 | -0.10192 | 0.14221 | 0.000449 | 1.2e-05 | 1.2e-05 | 506.5982
2019-03-20 12:45 | 1553100323.1623 | -260 | 3.4e-05 | 2e-06 | 0 | 108.639 | 109.9738 | 129.6289 | 20.9823 | 0.001811 | 0.001726 | -152.1273 | 0.23929 | 8.4032 | 0.003536 | 0.004241 | 0.003856 | 6.4e-05 | 0.001207 | 0.30326 | 1.3262 | 2.5482 | 101 | 20.7 | 18.153 | 65.1253 | 106.6993 | 63.082 | 0.004393 | 1.05 | 0.02552 | 24.3042 | -0.10922 | 0.14235 | 0.000448 | -8e-06 | -8e-06 | -764.9065
2019-03-20 12:45 | 1553100324.1617 | -259 | 8.4e-05 | -1e-06 | 0 | 108.5666 | 109.8691 | 129.401 | 20.9814 | 0.001862 | 0.001754 | -152.141 | 0.23967 | 8.4029 | 0.003411 | 0.004519 | 0.004106 | 6.3e-05 | 0.001205 | 0.32532 | 1.3247 | 2.5481 | 101 | 20.7 | 18.151 | 66.349 | 103.6588 | 63.7598 | 0.004393 | 1.0479 | 0.025471 | 24.3073 | -0.10668 | 0.14258 | 0.000451 | 5e-06 | 5e-06 | 1241.74
2019-03-20 12:45 | 1553100325.1616 | -258 | -7e-06 | 4e-06 | 0 | 108.521 | 109.9031 | 129.2689 | 20.9828 | 0.001877 | 0.001765 | -152.1618 | 0.23901 | 8.4019 | 0.003373 | 0.004576 | 0.004201 | 6.2e-05 | 0.001208 | 0.33124 | 1.3232 | 2.5481 | 101 | 20.7 | 18.148 | 66.7319 | 102.736 | 64.0337 | 0.004392 | 1.047 | 0.025448 | 24.3048 | -0.1132 | 0.14219 | 0.000446 | 4e-06 | 4e-06 | 1406.2313
2019-03-20 12:45 | 1553100326.1614 | -257 | 8.2e-05 | 1e-06 | 0 | 108.5798 | 109.9584 | 128.9403 | 20.9822 | 0.001857 | 0.001713 | -152.1568 | 0.23774 | 8.4046 | 0.003408 | 0.004423 | 0.004077 | 6.2e-05 | 0.001208 | 0.33948 | 1.3452 | 2.5481 | 101 | 20.7 | 18.147 | 66.2466 | 103.5763 | 62.7819 | 0.004392 | 1.0509 | 0.025463 | 24.2288 | -0.11285 | 0.14143 | 0.000449 | 1.5e-05 | 1.5e-05 | 394.5936
2019-03-20 12:45 | 1553100327.1619 | -256 | 8.6e-05 | 3e-06 | 0 | 108.6332 | 109.9229 | 128.6513 | 20.9822 | 0.001829 | 0.001703 | -152.1641 | 0.23881 | 8.4009 | 0.003404 | 0.004401 | 0.004037 | 6.6e-05 | 0.001207 | 0.32068 | 1.3212 | 2.548 | 101 | 20.7 | 18.144 | 65.5568 | 103.4852 | 62.54 | 0.004391 | 1.0517 | 0.025444 | 24.1929 | -0.10557 | 0.14207 | 0.000447 | 1.8e-05 | 1.8e-05 | 320.7261
2019-03-20 12:45 | 1553100328.162 | -255 | 9.4e-05 | 1e-06 | 1e-06 | 108.5882 | 110.0466 | 128.249 | 20.983 | 0.001883 | 0.001774 | -152.1769 | 0.23833 | 8.403 | 0.003451 | 0.00462 | 0.004196 | 6.3e-05 | 0.001207 | 0.31404 | 1.3537 | 2.548 | 101 | 20.7 | 18.144 | 66.8654 | 104.6306 | 64.2516 | 0.004391 | 1.0464 | 0.025339 | 24.2165 | -0.11933 | 0.14178 | 0.000448 | -2e-05 | -2e-05 | -295.2691
2019-03-20 12:45 | 1553100329.1626 | -254 | 5.4e-05 | 1e-06 | 0 | 108.6439 | 110.0167 | 128.1647 | 20.9827 | 0.001831 | 0.001739 | -152.1753 | 0.23901 | 8.4018 | 0.003511 | 0.004377 | 0.004091 | 6.3e-05 | 0.001208 | 0.33732 | 1.3004 | 2.548 | 101 | 20.71 | 18.142 | 65.6129 | 106.0746 | 63.3941 | 0.004393 | 1.049 | 0.025363 | 24.1778 | -0.11232 | 0.14219 | 0.000449 | -9e-06 | -9e-06 | -680.1776
2019-03-20 12:45 | 1553100330.163 | -253 | 7.3e-05 | 4e-06 | 0 | 108.6587 | 109.9487 | 128.0051 | 20.9831 | 0.001803 | 0.001718 | -152.1897 | 0.24025 | 8.4023 | 0.003578 | 0.00429 | 0.003912 | 6.3e-05 | 0.001211 | 0.32659 | 1.3155 | 2.5479 | 101 | 20.71 | 18.137 | 64.9387 | 107.7215 | 62.8938 | 0.004392 | 1.0506 | 0.025366 | 24.1448 | -0.10556 | 0.14292 | 0.000446 | 1.7e-05 | 1.7e-05 | 335.9972
2019-03-20 12:46 | 1553100331.163 | -252 | 6e-05 | 5e-06 | 0 | 108.5598 | 109.905 | 128.0178 | 20.9831 | 0.001875 | 0.001731 | -152.1985 | 0.23963 | 8.4031 | 0.003522 | 0.004633 | 0.004232 | 5.9e-05 | 0.001217 | 0.33843 | 1.3556 | 2.5481 | 101 | 20.71 | 18.132 | 66.6636 | 106.3537 | 63.2024 | 0.004391 | 1.0496 | 0.025356 | 24.157 | -0.11016 | 0.14255 | 0.000445 | 4e-06 | 4e-06 | 1554.5752
2019-03-20 12:46 | 1553100332.1627 | -251 | 0.000162 | 3e-06 | 1e-06 | 108.6481 | 109.8844 | 128.2157 | 20.9833 | 0.001865 | 0.001795 | -152.1929 | 0.2392 | 8.4034 | 0.003602 | 0.004707 | 0.004329 | 6.5e-05 | 0.001217 | 0.34036 | 1.3717 | 2.5481 | 101 | 20.71 | 18.129 | 66.4413 | 108.2888 | 64.7382 | 0.00439 | 1.0449 | 0.025318 | 24.2308 | -0.10121 | 0.1423 | 0.000446 | -1.3e-05 | -1.3e-05 | -456.0302
2019-03-20 12:46 | 1553100333.1622 | -250 | 5.8e-05 | 4e-06 | 0 | 108.7273 | 109.9132 | 128.4262 | 20.9822 | 0.001858 | 0.001726 | -152.2016 | 0.23989 | 8.4028 | 0.003643 | 0.004397 | 0.004058 | 6.5e-05 | 0.001221 | 0.3282 | 1.3102 | 2.5481 | 101 | 20.71 | 18.118 | 66.2647 | 109.2872 | 63.0773 | 0.004387 | 1.05 | 0.025401 | 24.191 | -0.097032 | 0.14271 | 0.000447 | 3.8e-05 | 3.8e-05 | 154.1804
2019-03-20 12:46 | 1553100334.1609 | -249 | 8.3e-05 | 2e-06 | 0 | 108.6102 | 109.9835 | 128.5246 | 20.9834 | 0.001829 | 0.001752 | -152.1994 | 0.23917 | 8.4028 | 0.003682 | 0.004646 | 0.004327 | 6.3e-05 | 0.00122 | 0.31349 | 1.3296 | 2.5481 | 101 | 20.71 | 18.112 | 65.5587 | 110.2479 | 63.725 | 0.004386 | 1.048 | 0.025386 | 24.2236 | -0.11238 | 0.14228 | 0.000447 | -1e-05 | -1e-05 | -608.0882
2019-03-20 12:46 | 1553100335.1609 | -248 | 0.000134 | 1e-06 | 0 | 108.6662 | 110.1603 | 128.6616 | 20.9837 | 0.001835 | 0.001726 | -152.1922 | 0.23875 | 8.402 | 0.003599 | 0.005008 | 0.004665 | 6.6e-05 | 0.001222 | 0.35203 | 1.3572 | 2.5481 | 101 | 20.71 | 18.101 | 65.7136 | 108.2103 | 63.0895 | 0.004383 | 1.05 | 0.025424 | 24.2136 | -0.12214 | 0.14203 | 0.000449 | 2e-05 | 2e-05 | 286.867
2019-03-20 12:46 | 1553100336.162 | -247 | 8.5e-05 | 5e-06 | 0 | 108.6817 | 110.0762 | 129.0595 | 20.9824 | 0.001816 | 0.001774 | -152.1924 | 0.23841 | 8.4039 | 0.003597 | 0.005225 | 0.004836 | 6.5e-05 | 0.001226 | 0.34108 | 1.3102 | 2.5481 | 101 | 20.71 | 18.096 | 65.244 | 108.1784 | 64.2424 | 0.004382 | 1.0464 | 0.02542 | 24.2926 | -0.11403 | 0.14183 | 0.000445 | 1.9e-05 | 1.9e-05 | 305.0213
2019-03-20 12:46 | 1553100337.1623 | -246 | 0.000106 | 2e-06 | 0 | 108.617 | 110.0676 | 129.4778 | 20.9833 | 0.00188 | 0.001734 | -152.185 | 0.23882 | 8.4033 | 0.003518 | 0.004584 | 0.004249 | 6.3e-05 | 0.001221 | 0.34307 | 1.3485 | 2.548 | 101 | 20.71 | 18.092 | 66.7854 | 106.2648 | 63.2813 | 0.004381 | 1.0494 | 0.025497 | 24.2972 | -0.11866 | 0.14208 | 0.000448 | 3.8e-05 | 3.8e-05 | 152.7314
2019-03-20 12:46 | 1553100338.1611 | -245 | 8e-05 | 1e-06 | -1e-06 | 108.6652 | 109.9229 | 129.7508 | 20.9834 | 0.001867 | 0.001766 | -152.1878 | 0.23913 | 8.4039 | 0.00355 | 0.004802 | 0.004448 | 6.3e-05 | 0.001214 | 0.34075 | 1.3335 | 2.5481 | 101 | 20.71 | 18.081 | 66.4905 | 107.027 | 64.053 | 0.004378 | 1.047 | 0.025495 | 24.3507 | -0.10293 | 0.14226 | 0.00045 | 5.5e-05 | 5.5e-05 | 106.291
2019-03-20 12:46 | 1553100339.1611 | -244 | 8.7e-05 | 3e-06 | 0 | 108.6885 | 110.0309 | 129.7684 | 20.9844 | 0.001825 | 0.001695 | -152.1695 | 0.23868 | 8.4041 | 0.003469 | 0.005599 | 0.00519 | 6.3e-05 | 0.001207 | 0.34755 | 1.3166 | 2.5481 | 101 | 20.71 | 18.075 | 65.463 | 105.0638 | 62.335 | 0.004377 | 1.0523 | 0.025562 | 24.2903 | -0.1098 | 0.14199 | 0.000447 | 2e-05 | 2e-05 | 294.1372
2019-03-20 12:46 | 1553100340.1616 | -243 | 6.5e-05 | 7e-06 | 0 | 108.7102 | 110.2245 | 129.5825 | 20.9846 | 0.001845 | 0.001751 | -152.1788 | 0.23802 | 8.4018 | 0.003486 | 0.004526 | 0.00421 | 6.1e-05 | 0.001202 | 0.3246 | 1.3539 | 2.5481 | 101 | 20.71 | 18.07 | 65.9555 | 105.4856 | 63.6934 | 0.004376 | 1.0481 | 0.025492 | 24.322 | -0.12374 | 0.14159 | 0.000443 | -1.1e-05 | -1.1e-05 | -507.1045
2019-03-20 12:46 | 1553100341.1623 | -242 | 6.2e-05 | 1e-06 | 0 | 108.8062 | 110.1517 | 129.6191 | 20.9839 | 0.001893 | 0.001748 | -152.1764 | 0.23844 | 8.4002 | 0.003394 | 0.00479 | 0.004373 | 6e-05 | 0.001199 | 0.31045 | 1.3077 | 2.5481 | 101 | 20.71 | 18.061 | 67.0957 | 103.2463 | 63.6263 | 0.004373 | 1.0483 | 0.025498 | 24.323 | -0.10993 | 0.14185 | 0.000449 | -1e-06 | -1e-06 | -3977.9606
2019-03-20 12:46 | 1553100342.1622 | -241 | 6.5e-05 | 3e-06 | 0 | 108.692 | 110.1464 | 129.6785 | 20.9844 | 0.001859 | 0.001724 | -152.1666 | 0.23824 | 8.3999 | 0.003426 | 0.004898 | 0.00453 | 6.2e-05 | 0.001196 | 0.33539 | 1.3521 | 2.5481 | 101 | 20.71 | 18.059 | 66.2909 | 104.0344 | 63.0513 | 0.004373 | 1.0501 | 0.025526 | 24.3078 | -0.11889 | 0.14173 | 0.000447 | -4e-06 | -4e-06 | -1574.4749
2019-03-20 12:46 | 1553100343.1624 | -240 | -1e-06 | 1e-06 | 0 | 108.7276 | 110.1397 | 129.7707 | 20.985 | 0.001855 | 0.001737 | -152.1357 | 0.23824 | 8.4013 | 0.003418 | 0.004853 | 0.004477 | 6.5e-05 | 0.001196 | 0.34755 | 1.3143 | 2.5482 | 101 | 20.71 | 18.057 | 66.1854 | 103.8431 | 63.3533 | 0.004373 | 1.0492 | 0.025523 | 24.3273 | -0.11542 | 0.14173 | 0.000448 | -1.9e-05 | -1.9e-05 | -304.7585
2019-03-20 12:46 | 1553100344.1624 | -239 | 0.000148 | 2e-06 | 0 | 108.7373 | 110.2077 | 130.2315 | 20.9843 | 0.001907 | 0.001741 | -152.1376 | 0.23779 | 8.4018 | 0.003432 | 0.004432 | 0.004058 | 6.3e-05 | 0.001197 | 0.31542 | 1.3219 | 2.5479 | 101 | 20.71 | 18.056 | 67.4355 | 104.1616 | 63.4402 | 0.004372 | 1.0489 | 0.025565 | 24.3736 | -0.12014 | 0.14146 | 0.000448 | 1.9e-05 | 1.9e-05 | 305.5722
2019-03-20 12:46 | 1553100345.2201 | -238 | 0.000171 | 3e-06 | 0 | 108.7923 | 110.0205 | 130.1777 | 20.984 | 0.001907 | 0.00178 | -152.1371 | 0.23843 | 8.4052 | 0.003511 | 0.005128 | 0.004812 | 6.5e-05 | 0.001195 | 0.3465 | 1.3082 | 2.5479 | 101 | 20.71 | 18.057 | 67.44 | 106.0901 | 64.3924 | 0.004373 | 1.0459 | 0.025524 | 24.403 | -0.10041 | 0.14184 | 0.000447 | 1.6e-05 | 1.6e-05 | 374.5768
2019-03-20 12:46 | 1553100346.1709 | -237 | 0.000131 | 7e-06 | 0 | 108.7632 | 110.0436 | 130.0223 | 20.9835 | 0.001848 | 0.00175 | -152.138 | 0.23882 | 8.4041 | 0.003521 | 0.004829 | 0.004443 | 6.4e-05 | 0.0012 | 0.31382 | 1.3126 | 2.5481 | 101 | 20.71 | 18.058 | 66.0226 | 106.3359 | 63.6729 | 0.004373 | 1.0482 | 0.025536 | 24.3625 | -0.10469 | 0.14208 | 0.000442 | 3e-06 | 3e-06 | 1791.7107
2019-03-20 12:46 | 1553100347.1723 | -236 | 0.00016 | 2e-06 | 0 | 108.7684 | 110.0661 | 129.8604 | 20.9848 | 0.001853 | 0.001699 | -152.1603 | 0.23865 | 8.4033 | 0.003505 | 0.004868 | 0.004508 | 6e-05 | 0.001198 | 0.34634 | 1.3087 | 2.548 | 101 | 20.71 | 18.058 | 66.1479 | 105.95 | 62.4385 | 0.004373 | 1.052 | 0.025567 | 24.3026 | -0.10608 | 0.14197 | 0.000448 | 1e-05 | 1e-05 | 578.7677
2019-03-20 12:46 | 1553100348.171 | -235 | 0.000136 | 0 | 0 | 108.7483 | 110.1763 | 129.9789 | 20.9837 | 0.001821 | 0.001711 | -152.1444 | 0.23858 | 8.4009 | 0.003441 | 0.004997 | 0.004638 | 6.6e-05 | 0.001197 | 0.31598 | 1.3031 | 2.548 | 101 | 20.71 | 18.059 | 65.3687 | 104.3876 | 62.7311 | 0.004373 | 1.0511 | 0.025567 | 24.3243 | -0.11669 | 0.14193 | 0.000449 | 1e-06 | 1e-06 | 6586.4354
2019-03-20 12:46 | 1553100349.1714 | -234 | 0.000108 | 1e-06 | 0 | 108.7871 | 110.2242 | 130.1843 | 20.9838 | 0.001863 | 0.001698 | -152.16 | 0.23878 | 8.4024 | 0.003529 | 0.004257 | 0.003841 | 6e-05 | 0.001198 | 0.36298 | 1.3269 | 2.548 | 101 | 20.71 | 18.059 | 66.3841 | 106.5149 | 62.4062 | 0.004373 | 1.0521 | 0.0256 | 24.3317 | -0.11738 | 0.14205 | 0.000448 | 1.1e-05 | 1.1e-05 | 541.6335
2019-03-20 12:46 | 1553100350.1711 | -233 | 0.00011 | 3e-06 | 0 | 108.7613 | 110.0668 | 130.7497 | 20.9846 | 0.001942 | 0.001739 | -152.1445 | 0.23878 | 8.404 | 0.003504 | 0.004722 | 0.00438 | 5.6e-05 | 0.001194 | 0.33859 | 1.3396 | 2.548 | 101 | 20.71 | 18.059 | 68.2782 | 105.9067 | 63.3943 | 0.004373 | 1.049 | 0.025618 | 24.4204 | -0.10673 | 0.14205 | 0.000446 | 6e-06 | 6e-06 | 1053.7827
2019-03-20 12:46 | 1553100351.1702 | -232 | 5.9e-05 | -1e-06 | 0 | 108.7389 | 110.2354 | 130.7678 | 20.9838 | 0.001909 | 0.001703 | -152.1459 | 0.2381 | 8.4013 | 0.003499 | 0.004525 | 0.004127 | 5.9e-05 | 0.001189 | 0.32244 | 1.3671 | 2.5482 | 101 | 20.71 | 18.059 | 67.4997 | 105.7979 | 62.5329 | 0.004373 | 1.0517 | 0.025652 | 24.3908 | -0.12225 | 0.14164 | 0.00045 | 1e-06 | 1e-06 | 8159.4212
2019-03-20 12:46 | 1553100352.1712 | -231 | 8.1e-05 | 1e-06 | 0 | 108.6995 | 110.2402 | 130.7192 | 20.9837 | 0.001865 | 0.001759 | -152.1591 | 0.23924 | 8.4027 | 0.00352 | 0.004987 | 0.004629 | 5.8e-05 | 0.001188 | 0.32748 | 1.3742 | 2.5481 | 101 | 20.71 | 18.059 | 66.4394 | 106.3087 | 63.8714 | 0.004373 | 1.0475 | 0.025597 | 24.4349 | -0.12589 | 0.14233 | 0.000448 | -6e-06 | -6e-06 | -900.6311
2019-03-20 12:46 | 1553100353.1702 | -230 | 0.000132 | -1e-06 | 0 | 108.6917 | 110.2795 | 130.4967 | 20.9839 | 0.001892 | 0.001766 | -152.1709 | 0.23901 | 8.4015 | 0.003394 | 0.004849 | 0.00454 | 5.6e-05 | 0.001184 | 0.32333 | 1.3904 | 2.5482 | 101 | 20.71 | 18.059 | 67.0715 | 103.2425 | 64.0528 | 0.004373 | 1.047 | 0.025568 | 24.4206 | -0.12972 | 0.14219 | 0.000451 | 1.9e-05 | 1.9e-05 | 305.2114
2019-03-20 12:46 | 1553100354.1707 | -229 | 0.000119 | 3e-06 | 0 | 108.8133 | 110.1898 | 130.1265 | 20.9841 | 0.001854 | 0.001733 | -152.1682 | 0.23902 | 8.4007 | 0.003427 | 0.004668 | 0.004295 | 5.7e-05 | 0.001184 | 0.33339 | 1.3982 | 2.5481 | 101 | 20.71 | 18.057 | 66.1737 | 104.0506 | 63.2563 | 0.004373 | 1.0495 | 0.025562 | 24.3571 | -0.11244 | 0.14219 | 0.000447 | 2e-05 | 2e-05 | 285.2384
2019-03-20 12:46 | 1553100355.1714 | -228 | 3.5e-05 | 1e-06 | 0 | 108.7904 | 110.1954 | 130.0862 | 20.9834 | 0.001853 | 0.001763 | -152.1552 | 0.23898 | 8.4018 | 0.003412 | 0.004577 | 0.004189 | 5.4e-05 | 0.001188 | 0.33771 | 1.3317 | 2.5481 | 101 | 20.71 | 18.056 | 66.1331 | 103.6822 | 63.9741 | 0.004372 | 1.0472 | 0.025531 | 24.3793 | -0.11478 | 0.14217 | 0.000448 | -2.1e-05 | -2.1e-05 | -277.879
2019-03-20 12:46 | 1553100356.1714 | -227 | 0.000118 | -1e-06 | 1e-06 | 108.7535 | 110.167 | 130.1258 | 20.9841 | 0.001893 | 0.001754 | -152.1554 | 0.23993 | 8.4037 | 0.003479 | 0.004815 | 0.004436 | 5.7e-05 | 0.001186 | 0.34418 | 1.3381 | 2.5479 | 101 | 20.71 | 18.055 | 67.1022 | 105.319 | 63.7702 | 0.004372 | 1.0479 | 0.025542 | 24.3757 | -0.1155 | 0.14273 | 0.00045 | -2e-05 | -2e-05 | -291.1688
2019-03-20 12:46 | 1553100357.1705 | -226 | 0.000109 | 0 | 0 | 108.682 | 110.0694 | 130.231 | 20.9832 | 0.001894 | 0.001752 | -152.1437 | 0.23946 | 8.4019 | 0.003442 | 0.004531 | 0.004167 | 5.6e-05 | 0.001187 | 0.3533 | 1.3458 | 2.5482 | 101 | 20.71 | 18.056 | 67.1326 | 104.4137 | 63.7115 | 0.004372 | 1.048 | 0.025555 | 24.3834 | -0.11346 | 0.14245 | 0.00045 | 2e-05 | 2e-05 | 294.3441
2019-03-20 12:46 | 1553100358.1706 | -225 | 8.2e-05 | 0 | 0 | 108.8027 | 110.3258 | 130.1676 | 20.9834 | 0.001873 | 0.001708 | -152.1499 | 0.23863 | 8.4023 | 0.003439 | 0.004966 | 0.004606 | 5.6e-05 | 0.001187 | 0.32278 | 1.3274 | 2.548 | 101 | 20.71 | 18.058 | 66.6138 | 104.3283 | 62.6644 | 0.004373 | 1.0513 | 0.025588 | 24.3395 | -0.12435 | 0.14196 | 0.00045 | 2.1e-05 | 2.1e-05 | 282.4197
2019-03-20 12:46 | 1553100359.1706 | -224 | 9.3e-05 | 0 | 0 | 108.7845 | 110.2634 | 130.0371 | 20.9844 | 0.001802 | 0.001762 | -152.1365 | 0.23898 | 8.4022 | 0.003439 | 0.004576 | 0.004113 | 5.8e-05 | 0.001187 | 0.3131 | 1.3281 | 2.548 | 101 | 20.71 | 18.059 | 64.9236 | 104.3354 | 63.958 | 0.004373 | 1.0473 | 0.025526 | 24.3741 | -0.12078 | 0.14217 | 0.000449 | -8e-06 | -8e-06 | -764.6396
2019-03-20 12:46 | 1553100360.171 | -223 | 0.000131 | -2e-06 | 0 | 108.7826 | 110.2866 | 129.8468 | 20.9849 | 0.001863 | 0.001747 | -152.1327 | 0.23875 | 8.4005 | 0.003382 | 0.004617 | 0.004357 | 5.4e-05 | 0.001192 | 0.31957 | 1.351 | 2.5481 | 101 | 20.71 | 18.059 | 66.3741 | 102.9557 | 63.6015 | 0.004373 | 1.0484 | 0.025521 | 24.3434 | -0.12282 | 0.14203 | 0.000452 | 2.8e-05 | 2.8e-05 | 209.5634
2019-03-20 12:46 | 1553100361.1721 | -222 | 0.005644 | -2e-06 | 0 | 108.7677 | 110.1535 | 129.7855 | 20.9838 | 0.001844 | 0.001798 | -152.122 | 0.239 | 8.403 | 0.003429 | 0.004292 | 0.003935 | 5.6e-05 | 0.001195 | 0.32278 | 1.3494 | 2.5481 | 101 | 20.71 | 18.059 | 65.9332 | 104.1085 | 64.8278 | 0.004373 | 1.0446 | 0.025469 | 24.3819 | -0.11324 | 0.14218 | 0.000454 | 0.000923 | 0.000923 | 6.3284
2019-03-20 12:46 | 1553100362.1707 | -221 | 0.047546 | -1e-06 | 0 | 108.7939 | 110.1786 | 129.9923 | 20.9835 | 0.001852 | 0.001737 | -152.1104 | 0.23948 | 8.4048 | 0.003378 | 0.004807 | 0.004389 | 5.5e-05 | 0.001194 | 0.33859 | 1.3021 | 2.5481 | 101 | 20.72 | 18.058 | 66.1168 | 102.8616 | 63.344 | 0.004375 | 1.0492 | 0.025545 | 24.3477 | -0.11312 | 0.14247 | 0.000473 | 0.008266 | 0.008266 | 0.70838
2019-03-20 12:46 | 1553100363.1721 | -220 | 0.10971 | 3e-06 | 1e-06 | 108.7933 | 110.2507 | 130.0444 | 20.9838 | 0.001885 | 0.001757 | -152.1172 | 0.23821 | 8.4026 | 0.003354 | 0.004294 | 0.004012 | 5.7e-05 | 0.001191 | 0.33212 | 1.3475 | 2.548 | 101 | 20.72 | 18.058 | 66.9225 | 102.276 | 63.8394 | 0.004375 | 1.0476 | 0.025532 | 24.3705 | -0.11903 | 0.14171 | 0.000501 | 0.02016 | 0.02016 | 0.2889
2019-03-20 12:46 | 1553100364.1719 | -219 | 0.13252 | 1.6e-05 | 0 | 108.8518 | 110.2305 | 130.0831 | 20.9841 | 0.001899 | 0.001783 | -152.0982 | 0.23869 | 8.4019 | 0.003356 | 0.004899 | 0.004467 | 5.6e-05 | 0.001187 | 0.32709 | 1.3473 | 2.5481 | 101 | 20.72 | 18.056 | 67.2388 | 102.3393 | 64.451 | 0.004375 | 1.0457 | 0.025512 | 24.3963 | -0.11258 | 0.14199 | 0.0005 | 0.024343 | 0.024343 | 0.23974
2019-03-20 12:46 | 1553100365.1709 | -218 | 0.15641 | 4.2e-05 | 0 | 108.8179 | 110.0339 | 130.166 | 20.9832 | 0.001857 | 0.001714 | -152.1113 | 0.23947 | 8.4027 | 0.003342 | 0.004335 | 0.00396 | 5.4e-05 | 0.001194 | 0.32244 | 1.3423 | 2.5482 | 101 | 20.71 | 18.054 | 66.2313 | 101.9793 | 62.7908 | 0.004372 | 1.0509 | 0.025583 | 24.344 | -0.099404 | 0.14246 | 0.000483 | 0.027848 | 0.027848 | 0.21025
2019-03-20 12:46 | 1553100366.17 | -217 | 0.16674 | 6.7e-05 | 0 | 108.8033 | 110.1345 | 130.246 | 20.9847 | 0.001865 | 0.001782 | -152.1133 | 0.24011 | 8.4008 | 0.003421 | 0.004662 | 0.004323 | 5.7e-05 | 0.001195 | 0.3126 | 1.3467 | 2.5481 | 101 | 20.71 | 18.051 | 66.4441 | 103.9019 | 64.4407 | 0.004371 | 1.0458 | 0.025529 | 24.4112 | -0.10877 | 0.14284 | 0.000459 | 0.028098 | 0.028098 | 0.20894
2019-03-20 12:46 | 1553100367.1701 | -216 | 0.17605 | 0.000112 | 0 | 108.7393 | 110.1617 | 129.6965 | 20.9841 | 0.001818 | 0.001803 | -152.1156 | 0.23857 | 8.4033 | 0.003309 | 0.004674 | 0.004256 | 5.8e-05 | 0.001187 | 0.33876 | 1.3136 | 2.5481 | 101 | 20.71 | 18.042 | 65.3121 | 101.2036 | 64.9375 | 0.004369 | 1.0442 | 0.025456 | 24.3775 | -0.11625 | 0.14192 | 0.000409 | 0.026405 | 0.026405 | 0.22091
2019-03-20 12:46 | 1553100368.171 | -215 | 0.18515 | 0.000151 | 1e-06 | 108.7684 | 110.1797 | 129.7127 | 20.9841 | 0.001915 | 0.001791 | -152.1267 | 0.23976 | 8.4027 | 0.003418 | 0.004777 | 0.004533 | 5.4e-05 | 0.001184 | 0.31349 | 1.3464 | 2.5481 | 101 | 20.71 | 18.037 | 67.624 | 103.8202 | 64.6503 | 0.004368 | 1.0451 | 0.025468 | 24.3687 | -0.11531 | 0.14264 | 0.000366 | 0.024823 | 0.024823 | 0.23617
2019-03-20 12:46 | 1553100369.1703 | -214 | 0.19125 | 0.000192 | 0 | 108.8298 | 110.1842 | 129.4468 | 20.9844 | 0.001844 | 0.001767 | -152.1471 | 0.24015 | 8.4018 | 0.003457 | 0.004495 | 0.00405 | 5.7e-05 | 0.001189 | 0.33395 | 1.3238 | 2.548 | 101 | 20.71 | 18.025 | 65.9164 | 104.7656 | 64.068 | 0.004365 | 1.0469 | 0.025464 | 24.3227 | -0.11063 | 0.14286 | 0.000318 | 0.022278 | 0.022278 | 0.26356
2019-03-20 12:46 | 1553100370.1712 | -213 | 0.19604 | 0.000242 | 0 | 108.8628 | 110.2328 | 129.1259 | 20.9849 | 0.001841 | 0.001783 | -152.1617 | 0.23868 | 8.4009 | 0.003391 | 0.004869 | 0.004492 | 5.2e-05 | 0.001188 | 0.31924 | 1.3568 | 2.5481 | 101 | 20.71 | 18.02 | 65.8474 | 103.1703 | 64.4713 | 0.004364 | 1.0457 | 0.025418 | 24.3071 | -0.11186 | 0.14199 | 0.000258 | 0.018504 | 0.018504 | 0.31538
2019-03-20 12:46 | 1553100371.1702 | -212 | 0.19862 | 0.000269 | 0 | 108.8683 | 110.1266 | 129.0975 | 20.9839 | 0.001869 | 0.001738 | -152.1585 | 0.23926 | 8.4015 | 0.003355 | 0.004899 | 0.0045 | 5.8e-05 | 0.00119 | 0.32748 | 1.3276 | 2.5481 | 101 | 20.71 | 18.013 | 66.5354 | 102.3044 | 63.3858 | 0.004362 | 1.0491 | 0.025456 | 24.2653 | -0.10279 | 0.14234 | 0.000225 | 0.016355 | 0.016355 | 0.3577
2019-03-20 12:46 | 1553100372.1712 | -211 | 0.20191 | 0.000302 | 1e-06 | 108.7855 | 110.2582 | 128.6304 | 20.9853 | 0.001892 | 0.001774 | -152.1668 | 0.23975 | 8.4029 | 0.003383 | 0.004622 | 0.004308 | 5.6e-05 | 0.001189 | 0.29159 | 1.2998 | 2.5481 | 101 | 20.71 | 18.011 | 67.0861 | 102.9939 | 64.2423 | 0.004361 | 1.0464 | 0.025377 | 24.2522 | -0.12027 | 0.14262 | 0.000185 | 0.013595 | 0.013595 | 0.43117
2019-03-20 12:46 | 1553100373.1707 | -210 | 0.20447 | 0.000329 | 0 | 108.7415 | 110.1247 | 128.7342 | 20.9849 | 0.001861 | 0.001754 | -152.1576 | 0.23988 | 8.4023 | 0.003425 | 0.004859 | 0.004455 | 5.4e-05 | 0.001196 | 0.31509 | 1.3138 | 2.5481 | 101 | 20.71 | 18.008 | 66.3375 | 104.0121 | 63.7717 | 0.004361 | 1.0479 | 0.025405 | 24.245 | -0.11306 | 0.14271 | 0.000152 | 0.011336 | 0.011336 | 0.5174
2019-03-20 12:46 | 1553100374.1714 | -209 | 0.20479 | 0.000346 | 0 | 108.7855 | 110.1976 | 128.4957 | 20.9843 | 0.001898 | 0.001695 | -152.1866 | 0.23941 | 8.4039 | 0.003415 | 0.003854 | 0.003514 | 6e-05 | 0.001202 | 0.32189 | 1.2992 | 2.548 | 101 | 20.71 | 18.004 | 67.228 | 103.749 | 62.3451 | 0.00436 | 1.0523 | 0.025436 | 24.1712 | -0.11536 | 0.14243 | 0.000131 | 0.009832 | 0.009832 | 0.5954
2019-03-20 12:46 | 1553100375.1704 | -208 | 0.20491 | 0.000369 | 1e-06 | 108.8243 | 110.1606 | 128.275 | 20.9853 | 0.00188 | 0.001739 | -152.1621 | 0.2391 | 8.4017 | 0.003411 | 0.004663 | 0.00436 | 5.7e-05 | 0.001203 | 0.28495 | 1.3185 | 2.5481 | 101 | 20.71 | 18.003 | 66.8052 | 103.6658 | 63.4011 | 0.004359 | 1.049 | 0.025374 | 24.1884 | -0.10917 | 0.14224 | 0.0001 | 0.007456 | 0.007456 | 0.78404
2019-03-20 12:46 | 1553100376.1718 | -207 | 0.2056 | 0.000383 | 1e-06 | 108.8285 | 110.1064 | 128.0976 | 20.9851 | 0.001875 | 0.001683 | -152.1571 | 0.2387 | 8.4044 | 0.003343 | 0.004801 | 0.00454 | 5.8e-05 | 0.001207 | 0.31614 | 1.313 | 2.5479 | 101 | 20.71 | 18.002 | 66.6795 | 102.0228 | 62.0576 | 0.004359 | 1.0532 | 0.025407 | 24.1234 | -0.10442 | 0.142 | 8.3e-05 | 0.006188 | 0.006188 | 0.94311
2019-03-20 12:46 | 1553100377.1716 | -206 | 0.20645 | 0.000391 | 0 | 108.8651 | 110.386 | 128.6112 | 20.9853 | 0.001856 | 0.001703 | -152.1544 | 0.23832 | 8.401 | 0.003456 | 0.004592 | 0.004393 | 6.1e-05 | 0.001211 | 0.31133 | 1.3021 | 2.548 | 101 | 20.71 | 18.004 | 66.2295 | 104.7633 | 62.528 | 0.00436 | 1.0517 | 0.02544 | 24.1887 | -0.1241 | 0.14178 | 7.3e-05 | 0.005535 | 0.005535 | 1.0527
2019-03-20 12:46 | 1553100378.1716 | -205 | 0.20715 | 0.000404 | 0 | 108.8486 | 110.1857 | 128.4678 | 20.9857 | 0.001834 | 0.001706 | -152.1422 | 0.23908 | 8.3996 | 0.003473 | 0.00495 | 0.004685 | 5.8e-05 | 0.00121 | 0.3361 | 1.3524 | 2.5482 | 101 | 20.71 | 18.006 | 65.6944 | 105.1618 | 62.6095 | 0.00436 | 1.0515 | 0.025423 | 24.1781 | -0.1092 | 0.14223 | 5.6e-05 | 0.004256 | 0.004256 | 1.3736
2019-03-20 12:46 | 1553100379.1708 | -204 | 0.20758 | 0.00041 | 0 | 108.7942 | 110.2219 | 128.8335 | 20.9869 | 0.001846 | 0.00172 | -152.1287 | 0.23942 | 8.401 | 0.003484 | 0.004815 | 0.004559 | 5.8e-05 | 0.001212 | 0.31725 | 1.3565 | 2.548 | 101 | 20.71 | 18.009 | 65.9703 | 105.4248 | 62.9507 | 0.004361 | 1.0504 | 0.025446 | 24.2248 | -0.11661 | 0.14243 | 5e-05 | 0.003778 | 0.003778 | 1.5497
2019-03-20 12:46 | 1553100380.1717 | -203 | 0.20817 | 0.000423 | 0 | 108.8757 | 110.1857 | 129.244 | 20.9853 | 0.001824 | 0.001773 | -152.1344 | 0.23842 | 8.4033 | 0.003476 | 0.004535 | 0.00431 | 5.7e-05 | 0.001215 | 0.31598 | 1.3482 | 2.5482 | 101 | 20.71 | 18.012 | 65.4521 | 105.2392 | 64.2204 | 0.004362 | 1.0465 | 0.025439 | 24.3092 | -0.10697 | 0.14183 | 3.4e-05 | 0.002557 | 0.002557 | 2.2795
2019-03-20 12:46 | 1553100381.1707 | -202 | 0.20857 | 0.000422 | 0 | 108.7858 | 110.1718 | 129.4437 | 20.9866 | 0.00188 | 0.001681 | -152.1224 | 0.23939 | 8.4036 | 0.003414 | 0.004298 | 0.00413 | 5.9e-05 | 0.001219 | 0.32875 | 1.3914 | 2.5482 | 101 | 20.71 | 18.013 | 66.7908 | 103.7371 | 62.0137 | 0.004362 | 1.0534 | 0.025542 | 24.2482 | -0.11324 | 0.14242 | 3.5e-05 | 0.002686 | 0.002686 | 2.1788
2019-03-20 12:46 | 1553100382.1699 | -201 | 0.20885 | 0.000428 | 0 | 108.889 | 110.3718 | 130.02 | 20.9852 | 0.001921 | 0.001733 | -152.1194 | 0.23957 | 8.4 | 0.003476 | 0.0047 | 0.004396 | 5.7e-05 | 0.00122 | 0.30879 | 1.3842 | 2.5481 | 101 | 20.71 | 18.015 | 67.7724 | 105.2296 | 63.252 | 0.004362 | 1.0495 | 0.025552 | 24.347 | -0.12098 | 0.14252 | 2.7e-05 | 0.002042 | 0.002042 | 2.8685
2019-03-20 12:46 | 1553100383.1712 | -200 | 0.20911 | 0.000425 | 0 | 108.7913 | 110.2059 | 130.2833 | 20.9857 | 0.001882 | 0.001759 | -152.1105 | 0.23913 | 8.4016 | 0.003414 | 0.004327 | 0.004086 | 5.9e-05 | 0.001219 | 0.33212 | 1.3388 | 2.5479 | 101 | 20.71 | 18.015 | 66.8407 | 103.7314 | 63.889 | 0.004362 | 1.0475 | 0.025553 | 24.3947 | -0.11555 | 0.14226 | 3e-05 | 0.002324 | 0.002324 | 2.5157
2019-03-20 12:46 | 1553100384.17 | -199 | 0.20913 | 0.000435 | 1e-06 | 108.8356 | 110.2126 | 130.392 | 20.9863 | 0.001882 | 0.00176 | -152.1021 | 0.23927 | 8.4005 | 0.003367 | 0.004248 | 0.004115 | 6e-05 | 0.001215 | 0.31526 | 1.3196 | 2.5473 | 101 | 20.71 | 18.014 | 66.8444 | 102.6106 | 63.9021 | 0.004362 | 1.0475 | 0.025563 | 24.4054 | -0.11245 | 0.14234 | 1.8e-05 | 0.001313 | 0.001313 | 4.454
2019-03-20 12:46 | 1553100385.1714 | -198 | 0.2094 | 0.000439 | 1e-06 | 108.7499 | 110.3894 | 130.7204 | 20.9867 | 0.001845 | 0.001752 | -152.1121 | 0.23842 | 8.4031 | 0.003377 | 0.004581 | 0.004316 | 5.6e-05 | 0.001218 | 0.29734 | 1.3012 | 2.5479 | 101 | 20.71 | 18.013 | 65.9655 | 102.8458 | 63.7219 | 0.004362 | 1.048 | 0.025602 | 24.4295 | -0.13384 | 0.14184 | 1.2e-05 | 0.000919 | 0.000919 | 6.3414
2019-03-20 12:46 | 1553100386.171 | -197 | 0.20939 | 0.000439 | 0 | 108.7797 | 110.1318 | 130.6591 | 20.9863 | 0.001868 | 0.001749 | -152.1123 | 0.23853 | 8.4034 | 0.003413 | 0.004972 | 0.004727 | 6e-05 | 0.001214 | 0.32068 | 1.3358 | 2.5481 | 101 | 20.71 | 18.012 | 66.4965 | 103.7204 | 63.6506 | 0.004362 | 1.0482 | 0.025599 | 24.4212 | -0.1105 | 0.14191 | 1.2e-05 | 0.000902 | 0.000902 | 6.4657
2019-03-20 12:46 | 1553100387.1702 | -196 | 0.20938 | 0.000444 | 1e-06 | 108.7803 | 110.2376 | 130.6713 | 20.9858 | 0.001788 | 0.001715 | -152.1175 | 0.23989 | 8.4003 | 0.003387 | 0.004884 | 0.004699 | 5.6e-05 | 0.001216 | 0.33395 | 1.3165 | 2.5482 | 101 | 20.71 | 18.008 | 64.5893 | 103.0941 | 62.8316 | 0.004361 | 1.0508 | 0.025632 | 24.3927 | -0.11903 | 0.14271 | 6e-06 | 0.000418 | 0.000418 | 14.0438
2019-03-20 12:46 | 1553100388.1702 | -195 | 0.20937 | 0.000441 | 1e-06 | 108.8153 | 110.0033 | 130.9288 | 20.9868 | 0.001856 | 0.001721 | -152.1406 | 0.23922 | 8.4034 | 0.003467 | 0.005205 | 0.004958 | 5.9e-05 | 0.001212 | 0.30287 | 1.3309 | 2.5489 | 101 | 20.71 | 18.005 | 66.2076 | 105.0133 | 62.9677 | 0.00436 | 1.0504 | 0.025652 | 24.4216 | -0.097125 | 0.14231 | 9e-06 | 0.000685 | 0.000685 | 8.5395
2019-03-20 12:46 | 1553100389.1716 | -194 | 0.20939 | 0.00044 | 0 | 108.7671 | 110.1505 | 130.65 | 20.9868 | 0.001845 | 0.001735 | -152.1457 | 0.23886 | 8.401 | 0.003473 | 0.005039 | 0.004754 | 5.7e-05 | 0.001214 | 0.33588 | 1.3139 | 2.5481 | 101 | 20.71 | 18.001 | 65.9461 | 105.1621 | 63.302 | 0.004359 | 1.0493 | 0.025611 | 24.4077 | -0.11305 | 0.1421 | 1.1e-05 | 0.000849 | 0.000849 | 6.8764
2019-03-20 12:46 | 1553100390.1703 | -193 | 0.20943 | 0.000446 | 1e-06 | 108.7538 | 110.0279 | 130.6767 | 20.9865 | 0.001889 | 0.001717 | -152.1259 | 0.23841 | 8.4052 | 0.003476 | 0.004798 | 0.004658 | 5.9e-05 | 0.001215 | 0.32836 | 1.2914 | 2.5481 | 101 | 20.71 | 17.992 | 67.0007 | 105.2389 | 62.8643 | 0.004357 | 1.0507 | 0.025631 | 24.3944 | -0.10418 | 0.14183 | 3e-06 | 0.000201 | 0.000201 | 29.0058
2019-03-20 12:47 | 1553100391.1713 | -192 | 0.2095 | 0.000443 | 0 | 108.7554 | 109.973 | 130.2502 | 20.9867 | 0.001869 | 0.001679 | -152.1427 | 0.23909 | 8.4005 | 0.003427 | 0.004955 | 0.004722 | 5.9e-05 | 0.001215 | 0.31669 | 1.3616 | 2.5479 | 101 | 20.71 | 17.987 | 66.5211 | 104.0578 | 61.9645 | 0.004356 | 1.0535 | 0.025623 | 24.3219 | -0.099582 | 0.14223 | 8e-06 | 0.000635 | 0.000635 | 9.2016
2019-03-20 12:47 | 1553100392.1707 | -191 | 0.20948 | 0.000447 | 0 | 108.563 | 109.8878 | 129.9223 | 20.988 | 0.001878 | 0.001718 | -152.1421 | 0.23942 | 8.4024 | 0.003406 | 0.004601 | 0.004375 | 6.1e-05 | 0.001219 | 0.3621 | 1.3514 | 2.548 | 101 | 20.71 | 17.975 | 66.7386 | 103.5491 | 62.8971 | 0.004353 | 1.0506 | 0.025555 | 24.325 | -0.10849 | 0.14243 | 3e-06 | 0.000186 | 0.000186 | 31.4658
2019-03-20 12:47 | 1553100393.1705 | -190 | 0.20949 | 0.000447 | 1e-06 | 108.6694 | 109.9876 | 129.5379 | 20.9876 | 0.001883 | 0.001716 | -152.1498 | 0.23902 | 8.402 | 0.003514 | 0.004562 | 0.00444 | 5.6e-05 | 0.001218 | 0.3246 | 1.2982 | 2.5481 | 101 | 20.71 | 17.968 | 66.8714 | 106.1551 | 62.8606 | 0.004351 | 1.0507 | 0.025519 | 24.2877 | -0.10785 | 0.14219 | 2e-06 | 4.8e-05 | 4.8e-05 | 120.5683
2019-03-20 12:47 | 1553100394.1715 | -189 | 0.20955 | 0.000445 | 1e-06 | 108.7493 | 110.1311 | 129.9193 | 20.9872 | 0.001915 | 0.001796 | -152.1414 | 0.23882 | 8.4041 | 0.003452 | 0.005067 | 0.004805 | 5.8e-05 | 0.001216 | 0.32964 | 1.3548 | 2.548 | 101 | 20.72 | 17.96 | 67.6302 | 104.6702 | 64.7803 | 0.004352 | 1.0447 | 0.025484 | 24.3928 | -0.11294 | 0.14208 | 5e-06 | 0.000374 | 0.000374 | 15.6187
2019-03-20 12:47 | 1553100395.1715 | -188 | 0.20956 | 0.000449 | 0 | 108.7609 | 110.2765 | 129.9383 | 20.9872 | 0.00188 | 0.001723 | -152.1429 | 0.23945 | 8.4042 | 0.003436 | 0.004711 | 0.004603 | 5.6e-05 | 0.001223 | 0.34219 | 1.3598 | 2.5481 | 101 | 20.72 | 17.946 | 66.7923 | 104.2683 | 63.0157 | 0.004348 | 1.0502 | 0.025553 | 24.3308 | -0.12378 | 0.14245 | 0 | -2e-06 | -2e-06 | -3064.124
2019-03-20 12:47 | 1553100396.17 | -187 | 0.20951 | 0.000447 | 1e-06 | 108.7331 | 110.2956 | 130.0392 | 20.9872 | 0.001846 | 0.001787 | -152.1483 | 0.23867 | 8.4045 | 0.00339 | 0.004337 | 0.004144 | 6.1e-05 | 0.001222 | 0.33284 | 1.4004 | 2.5482 | 101 | 20.72 | 17.939 | 65.9716 | 103.1561 | 64.5626 | 0.004347 | 1.0454 | 0.025504 | 24.3962 | -0.12761 | 0.14199 | 2e-06 | 0.000117 | 0.000117 | 50.0288
2019-03-20 12:47 | 1553100397.1696 | -186 | 0.20953 | 0.000443 | 0 | 108.7179 | 109.9622 | 129.5177 | 20.9875 | 0.001828 | 0.001752 | -152.138 | 0.23932 | 8.4032 | 0.003415 | 0.005132 | 0.004862 | 5.9e-05 | 0.001222 | 0.31349 | 1.3568 | 2.5481 | 101 | 20.72 | 17.925 | 65.5355 | 103.7646 | 63.7165 | 0.004343 | 1.048 | 0.025485 | 24.3167 | -0.10179 | 0.14237 | 8e-06 | 0.000605 | 0.000605 | 9.6766
2019-03-20 12:47 | 1553100398.1706 | -185 | 0.20956 | 0.000445 | 1e-06 | 108.7509 | 110.1494 | 128.8274 | 20.9877 | 0.001898 | 0.001751 | -152.1313 | 0.2384 | 8.4016 | 0.003402 | 0.00393 | 0.003709 | 5.8e-05 | 0.001219 | 0.33467 | 1.328 | 2.5481 | 101 | 20.72 | 17.918 | 67.2274 | 103.4419 | 63.702 | 0.004342 | 1.0481 | 0.025417 | 24.2513 | -0.11429 | 0.14182 | 5e-06 | 0.000345 | 0.000345 | 16.8985
2019-03-20 12:47 | 1553100399.1698 | -184 | 0.20949 | 0.00045 | 0 | 108.714 | 110.0201 | 128.683 | 20.9878 | 0.001911 | 0.00176 | -152.143 | 0.23847 | 8.4047 | 0.003442 | 0.004597 | 0.004277 | 6e-05 | 0.001222 | 0.335 | 1.3104 | 2.5481 | 101 | 20.72 | 17.913 | 67.5434 | 104.4253 | 63.9119 | 0.00434 | 1.0474 | 0.025395 | 24.2453 | -0.10682 | 0.14187 | -1e-06 | -5.6e-05 | -5.6e-05 | -104.9285
2019-03-20 12:47 | 1553100400.1698 | -183 | 0.20938 | 0.000451 | 0 | 108.6707 | 110.0642 | 128.3354 | 20.9887 | 0.001891 | 0.001733 | -152.1678 | 0.23924 | 8.4005 | 0.0034 | 0.005352 | 0.005141 | 6.1e-05 | 0.001218 | 0.32444 | 1.3287 | 2.5481 | 101 | 20.72 | 17.904 | 67.0622 | 103.4097 | 63.2687 | 0.004338 | 1.0494 | 0.025385 | 24.1894 | -0.11397 | 0.14233 | -2e-06 | -0.000136 | -0.000136 | -42.9408
2019-03-20 12:47 | 1553100401.1699 | -182 | 0.2091 | 0.000445 | 1e-06 | 108.7176 | 110.2272 | 128.4908 | 20.9885 | 0.001853 | 0.001714 | -152.1765 | 0.23911 | 8.4023 | 0.003404 | 0.005203 | 0.005048 | 5.8e-05 | 0.001206 | 0.32244 | 1.3113 | 2.548 | 101 | 20.72 | 17.904 | 66.1484 | 103.5023 | 62.7984 | 0.004338 | 1.0509 | 0.025418 | 24.1871 | -0.12334 | 0.14225 | 5e-06 | 0.000308 | 0.000308 | 18.9717
2019-03-20 12:47 | 1553100402.1703 | -181 | 0.20949 | 0.000451 | 0 | 108.6933 | 110.1946 | 128.6175 | 20.9885 | 0.001881 | 0.001767 | -152.1685 | 0.23849 | 8.4023 | 0.003498 | 0.004976 | 0.004835 | 6e-05 | 0.001207 | 0.30486 | 1.3154 | 2.5481 | 101 | 20.72 | 17.901 | 66.8295 | 105.7671 | 64.0727 | 0.004337 | 1.0469 | 0.025382 | 24.2449 | -0.1227 | 0.14188 | -3e-06 | -0.000258 | -0.000258 | -22.6198
2019-03-20 12:47 | 1553100403.17 | -180 | 0.20953 | 0.00045 | 0 | 108.7315 | 110.1902 | 128.5922 | 20.9884 | 0.001856 | 0.001758 | -152.1847 | 0.2381 | 8.4026 | 0.003445 | 0.005499 | 0.005298 | 5.9e-05 | 0.001208 | 0.3465 | 1.3121 | 2.5481 | 101 | 20.72 | 17.899 | 66.2238 | 104.4818 | 63.8565 | 0.004337 | 1.0476 | 0.025388 | 24.2347 | -0.11919 | 0.14164 | -1e-06 | -9.3e-05 | -9.3e-05 | -62.6954
2019-03-20 12:47 | 1553100404.1704 | -179 | 0.2095 | 0.000449 | 0 | 108.618 | 109.9876 | 128.2065 | 20.9881 | 0.001831 | 0.001744 | -152.1995 | 0.23912 | 8.4032 | 0.003467 | 0.005005 | 0.004841 | 6.4e-05 | 0.00121 | 0.32156 | 1.3443 | 2.5481 | 101 | 20.72 | 17.899 | 65.615 | 105.0187 | 63.5337 | 0.004337 | 1.0486 | 0.025362 | 24.1867 | -0.11208 | 0.14225 | 1e-06 | 2.8e-05 | 2.8e-05 | 206.1277
2019-03-20 12:47 | 1553100405.1701 | -178 | 0.20949 | 0.000446 | 0 | 108.7551 | 110.0631 | 127.4359 | 20.9872 | 0.001859 | 0.001773 | -152.2094 | 0.23774 | 8.4034 | 0.003522 | 0.005153 | 0.00495 | 5.8e-05 | 0.00121 | 0.32836 | 1.3073 | 2.5479 | 101 | 20.72 | 17.901 | 66.2883 | 106.3514 | 64.2129 | 0.004337 | 1.0465 | 0.02526 | 24.1383 | -0.10693 | 0.14143 | 5e-06 | 0.000343 | 0.000343 | 16.9642
2019-03-20 12:47 | 1553100406.1701 | -177 | 0.20947 | 0.00045 | 0 | 108.6481 | 110.1763 | 127.1615 | 20.9879 | 0.001843 | 0.001793 | -152.2008 | 0.23812 | 8.4024 | 0.003524 | 0.004625 | 0.004348 | 6e-05 | 0.001213 | 0.3131 | 1.3615 | 2.548 | 101 | 20.72 | 17.903 | 65.9058 | 106.3945 | 64.7045 | 0.004338 | 1.045 | 0.025215 | 24.1298 | -0.12494 | 0.14166 | -1e-06 | -0.000107 | -0.000107 | -54.2754
2019-03-20 12:47 | 1553100407.1704 | -176 | 0.2095 | 0.00045 | 0 | 108.6322 | 110.0478 | 127.3798 | 20.9879 | 0.001862 | 0.001776 | -152.1894 | 0.23893 | 8.4007 | 0.003515 | 0.006014 | 0.005821 | 6.1e-05 | 0.001217 | 0.33373 | 1.3599 | 2.5479 | 101 | 20.72 | 17.908 | 66.3662 | 106.1749 | 64.3033 | 0.004339 | 1.0462 | 0.025251 | 24.1362 | -0.1158 | 0.14214 | -1e-06 | -7.7e-05 | -7.7e-05 | -75.9058
2019-03-20 12:47 | 1553100408.1712 | -175 | 0.20952 | 0.000449 | 0 | 108.6371 | 110.1087 | 127.607 | 20.9883 | 0.001917 | 0.001733 | -152.169 | 0.23831 | 8.4012 | 0.003548 | 0.005597 | 0.005432 | 6.1e-05 | 0.001218 | 0.3402 | 1.3559 | 2.548 | 101 | 20.72 | 17.912 | 67.6944 | 106.9823 | 63.2541 | 0.00434 | 1.0495 | 0.025313 | 24.1201 | -0.12035 | 0.14177 | 0 | -5.8e-05 | -5.8e-05 | -100.2505
2019-03-20 12:47 | 1553100409.1713 | -174 | 0.20949 | 0.000448 | 1e-06 | 108.6287 | 109.8463 | 127.7009 | 20.9894 | 0.001871 | 0.001785 | -152.1491 | 0.23811 | 8.4022 | 0.003584 | 0.005069 | 0.00487 | 6.1e-05 | 0.001223 | 0.33373 | 1.358 | 2.5482 | 101 | 20.72 | 17.923 | 66.5939 | 107.8725 | 64.503 | 0.004343 | 1.0456 | 0.025276 | 24.1737 | -0.099704 | 0.14165 | 1e-06 | -9e-06 | -9e-06 | -629.1822
2019-03-20 12:47 | 1553100410.1707 | -173 | 0.20947 | 0.000448 | 0 | 108.6086 | 109.8923 | 127.4164 | 20.9888 | 0.001811 | 0.001773 | -152.1409 | 0.23841 | 8.4033 | 0.003539 | 0.004839 | 0.004626 | 6.3e-05 | 0.001225 | 0.30198 | 1.344 | 2.548 | 101 | 20.72 | 17.929 | 65.1386 | 106.7615 | 64.2238 | 0.004344 | 1.0465 | 0.025258 | 24.1368 | -0.1051 | 0.14183 | 1e-06 | 5e-05 | 5e-05 | 117.2462
2019-03-20 12:47 | 1553100411.1713 | -172 | 0.20952 | 0.000452 | 0 | 108.6603 | 109.9655 | 127.1829 | 20.9884 | 0.001871 | 0.001729 | -152.1254 | 0.23911 | 8.3997 | 0.003614 | 0.004665 | 0.004452 | 5.8e-05 | 0.001226 | 0.32693 | 1.3412 | 2.5481 | 101 | 20.72 | 17.935 | 66.5891 | 108.6004 | 63.1606 | 0.004346 | 1.0498 | 0.025275 | 24.0766 | -0.1068 | 0.14225 | -3e-06 | -0.000238 | -0.000238 | -24.6087
2019-03-20 12:47 | 1553100412.1704 | -171 | 0.20949 | 0.000448 | 1e-06 | 108.6629 | 110.1371 | 126.8601 | 20.9886 | 0.00188 | 0.001759 | -152.1384 | 0.23851 | 8.4019 | 0.003655 | 0.005053 | 0.004895 | 5.7e-05 | 0.001225 | 0.33804 | 1.3165 | 2.5481 | 101 | 20.73 | 17.945 | 66.7944 | 109.5804 | 63.8855 | 0.004351 | 1.0475 | 0.025215 | 24.072 | -0.12053 | 0.14189 | 0 | -1.1e-05 | -1.1e-05 | -554.3888
2019-03-20 12:47 | 1553100413.1694 | -170 | 0.2095 | 0.00045 | 0 | 108.6005 | 110.0474 | 126.5379 | 20.9899 | 0.001898 | 0.001774 | -152.1272 | 0.23853 | 8.4026 | 0.003618 | 0.004856 | 0.004638 | 5.7e-05 | 0.001228 | 0.3298 | 1.2951 | 2.5481 | 101 | 20.73 | 17.95 | 67.2319 | 108.6886 | 64.256 | 0.004352 | 1.0464 | 0.02517 | 24.0546 | -0.11838 | 0.1419 | -1e-06 | -0.000122 | -0.000122 | -47.8061
2019-03-20 12:47 | 1553100414.1709 | -169 | 0.20952 | 0.000452 | 0 | 108.5795 | 110.0466 | 126.5989 | 20.9894 | 0.00192 | 0.001755 | -152.1403 | 0.23925 | 8.4011 | 0.003559 | 0.004942 | 0.004745 | 6e-05 | 0.001231 | 0.29717 | 1.3178 | 2.5482 | 101 | 20.73 | 17.96 | 67.7594 | 107.244 | 63.7871 | 0.004354 | 1.0478 | 0.025193 | 24.0437 | -0.12005 | 0.14233 | -4e-06 | -0.00029 | -0.00029 | -20.1565
2019-03-20 12:47 | 1553100415.1695 | -168 | 0.20952 | 0.000451 | 1e-06 | 108.5523 | 109.9057 | 127.0679 | 20.9899 | 0.001926 | 0.001741 | -152.1407 | 0.23796 | 8.4061 | 0.00354 | 0.004409 | 0.004138 | 5.7e-05 | 0.001234 | 0.34202 | 1.3105 | 2.548 | 101 | 20.73 | 17.964 | 67.9059 | 106.7963 | 63.4526 | 0.004355 | 1.0488 | 0.025252 | 24.0762 | -0.11083 | 0.14156 | -4e-06 | -0.000348 | -0.000348 | -16.7
2019-03-20 12:47 | 1553100416.1696 | -167 | 0.20952 | 0.00045 | 0 | 108.5636 | 109.979 | 127.2387 | 20.9894 | 0.001886 | 0.001737 | -152.1429 | 0.23775 | 8.404 | 0.003507 | 0.004763 | 0.004646 | 6e-05 | 0.001235 | 0.32045 | 1.2888 | 2.548 | 101 | 20.73 | 17.967 | 66.9395 | 105.9986 | 63.3599 | 0.004356 | 1.0491 | 0.025273 | 24.0891 | -0.11586 | 0.14144 | -2e-06 | -0.000173 | -0.000173 | -33.5734
2019-03-20 12:47 | 1553100417.1708 | -166 | 0.20951 | 0.000451 | 0 | 108.5843 | 110.0616 | 127.3784 | 20.9906 | 0.001886 | 0.001746 | -152.1638 | 0.23831 | 8.402 | 0.003502 | 0.00479 | 0.004477 | 6.3e-05 | 0.001231 | 0.32659 | 1.3075 | 2.548 | 101 | 20.73 | 17.971 | 66.9539 | 105.8771 | 63.5719 | 0.004357 | 1.0485 | 0.025279 | 24.1099 | -0.12087 | 0.14177 | -2e-06 | -0.000156 | -0.000156 | -37.2453
2019-03-20 12:47 | 1553100418.1698 | -165 | 0.20951 | 0.000453 | 0 | 108.5866 | 110.0365 | 127.6314 | 20.9903 | 0.001844 | 0.00174 | -152.1592 | 0.23887 | 8.4026 | 0.003463 | 0.004395 | 0.004259 | 5.9e-05 | 0.001232 | 0.33428 | 1.3089 | 2.5481 | 101 | 20.72 | 17.972 | 65.9259 | 104.9325 | 63.4373 | 0.004355 | 1.0489 | 0.025309 | 24.129 | -0.11864 | 0.1421 | -5e-06 | -0.000402 | -0.000402 | -14.5134
2019-03-20 12:47 | 1553100419.1713 | -164 | 0.20948 | 0.000446 | 0 | 108.5355 | 109.8003 | 127.7549 | 20.9902 | 0.001881 | 0.001764 | -152.1595 | 0.23928 | 8.4047 | 0.003492 | 0.005304 | 0.004968 | 6e-05 | 0.00122 | 0.30342 | 1.3202 | 2.548 | 101 | 20.73 | 17.973 | 66.8335 | 105.6247 | 64.0019 | 0.004358 | 1.0471 | 0.0253 | 24.1609 | -0.10363 | 0.14235 | 4e-06 | 0.000258 | 0.000258 | 22.6502
2019-03-20 12:47 | 1553100420.1713 | -163 | 0.2095 | 0.000453 | 0 | 108.5349 | 109.8392 | 127.9835 | 20.99 | 0.001879 | 0.001751 | -152.1308 | 0.23904 | 8.4037 | 0.00354 | 0.0045 | 0.00433 | 5.9e-05 | 0.001217 | 0.3282 | 1.3513 | 2.548 | 101 | 20.73 | 17.97 | 66.7787 | 106.7949 | 63.6931 | 0.004357 | 1.0481 | 0.025334 | 24.1714 | -0.10685 | 0.14221 | -4e-06 | -0.000301 | -0.000301 | -19.4293
2019-03-20 12:47 | 1553100421.1711 | -162 | 0.20952 | 0.000445 | 0 | 108.5281 | 109.8276 | 128.087 | 20.9901 | 0.001871 | 0.001731 | -152.1452 | 0.23785 | 8.4025 | 0.003565 | 0.004942 | 0.004744 | 6e-05 | 0.001214 | 0.3183 | 1.3158 | 2.5481 | 101 | 20.73 | 17.967 | 66.593 | 107.4021 | 63.203 | 0.004356 | 1.0496 | 0.025363 | 24.1636 | -0.10647 | 0.1415 | 5e-06 | 0.000367 | 0.000367 | 15.8324
2019-03-20 12:47 | 1553100422.171 | -161 | 0.20951 | 0.000448 | 0 | 108.5569 | 109.951 | 128.1114 | 20.9912 | 0.001865 | 0.00172 | -152.1211 | 0.23874 | 8.4017 | 0.00351 | 0.005258 | 0.005062 | 6e-05 | 0.00121 | 0.30198 | 1.3088 | 2.548 | 101 | 20.73 | 17.96 | 66.4351 | 106.0649 | 62.95 | 0.004354 | 1.0504 | 0.025375 | 24.1568 | -0.11414 | 0.14203 | 1e-06 | 9.1e-05 | 9.1e-05 | 64.4599
2019-03-20 12:47 | 1553100423.1701 | -160 | 0.20952 | 0.000447 | 0 | 108.5908 | 109.9652 | 128.4966 | 20.9916 | 0.00187 | 0.001735 | -152.1326 | 0.23902 | 8.4032 | 0.003513 | 0.004964 | 0.004788 | 5.9e-05 | 0.001208 | 0.31598 | 1.3417 | 2.5478 | 101 | 20.73 | 17.955 | 66.571 | 106.133 | 63.3182 | 0.004353 | 1.0493 | 0.025399 | 24.2063 | -0.11249 | 0.14219 | 2e-06 | 0.000136 | 0.000136 | 43.1052
2019-03-20 12:47 | 1553100424.1703 | -159 | 0.20951 | 0.000451 | 1e-06 | 108.5772 | 109.9083 | 128.2828 | 20.9907 | 0.001879 | 0.001762 | -152.1127 | 0.23917 | 8.3996 | 0.003504 | 0.005824 | 0.005516 | 6e-05 | 0.001207 | 0.31598 | 1.3664 | 2.548 | 101 | 20.73 | 17.95 | 66.7739 | 105.9148 | 63.9634 | 0.004352 | 1.0473 | 0.025354 | 24.2094 | -0.10899 | 0.14228 | -3e-06 | -0.000264 | -0.000264 | -22.1268
2019-03-20 12:47 | 1553100425.1711 | -158 | 0.20949 | 0.000452 | 1e-06 | 108.5976 | 109.6957 | 128.1734 | 20.9915 | 0.001853 | 0.001723 | -152.1084 | 0.23869 | 8.4028 | 0.003498 | 0.005102 | 0.004909 | 6.1e-05 | 0.001213 | 0.34435 | 1.3385 | 2.5481 | 101 | 20.73 | 17.938 | 66.1636 | 105.7819 | 63.0262 | 0.004349 | 1.0502 | 0.025378 | 24.1654 | -0.089988 | 0.142 | -5e-06 | -0.000471 | -0.000471 | -12.3782
2019-03-20 12:47 | 1553100426.1702 | -157 | 0.20951 | 0.000455 | 0 | 108.4909 | 110.136 | 127.9997 | 20.9912 | 0.001865 | 0.001735 | -152.1053 | 0.2396 | 8.4021 | 0.003457 | 0.004952 | 0.004804 | 5.6e-05 | 0.001208 | 0.30596 | 1.3402 | 2.548 | 101 | 20.73 | 17.932 | 66.4473 | 104.7929 | 63.3227 | 0.004348 | 1.0493 | 0.02535 | 24.1597 | -0.13461 | 0.14254 | -7e-06 | -0.000524 | -0.000524 | -11.1703
2019-03-20 12:47 | 1553100427.1704 | -156 | 0.20949 | 0.000451 | 0 | 108.477 | 109.8897 | 127.8162 | 20.9907 | 0.001925 | 0.001739 | -152.106 | 0.23833 | 8.4059 | 0.003559 | 0.004369 | 0.00413 | 5.5e-05 | 0.001213 | 0.32333 | 1.3348 | 2.5482 | 101 | 20.73 | 17.921 | 67.8879 | 107.2657 | 63.3987 | 0.004345 | 1.049 | 0.025329 | 24.1451 | -0.11573 | 0.14178 | 9.7e-05 | 0.007387 | 0.007387 | 0.78884
2019-03-20 12:47 | 1553100428.1693 | -155 | 0.2095 | 0.000452 | 0 | 108.5067 | 109.9353 | 127.5946 | 20.9921 | 0.001861 | 0.001725 | -152.1141 | 0.23817 | 8.403 | 0.003516 | 0.005102 | 0.004773 | 5.9e-05 | 0.001216 | 0.33572 | 1.3447 | 2.548 | 101 | 20.73 | 17.917 | 66.3421 | 106.2035 | 63.0705 | 0.004344 | 1.05 | 0.025319 | 24.1123 | -0.11699 | 0.14169 | 4e-05 | 0.003038 | 0.003038 | 1.9169
2019-03-20 12:47 | 1553100429.1704 | -154 | 0.20952 | 0.000453 | 1e-06 | 108.5604 | 109.9274 | 127.6136 | 20.9924 | 0.001864 | 0.00172 | -152.1334 | 0.23866 | 8.3999 | 0.003589 | 0.005187 | 0.004965 | 6e-05 | 0.00121 | 0.31813 | 1.2908 | 2.5481 | 101 | 20.73 | 17.913 | 66.4126 | 107.9935 | 62.9567 | 0.004343 | 1.0504 | 0.025325 | 24.1101 | -0.11193 | 0.14198 | -9.2e-05 | -0.007116 | -0.007116 | -0.8201
2019-03-20 12:47 | 1553100430.1706 | -153 | 0.20949 | 0.000449 | 0 | 108.52 | 109.8317 | 127.5 | 20.9929 | 0.00187 | 0.001733 | -152.1613 | 0.23885 | 8.4024 | 0.0035 | 0.004867 | 0.004694 | 5.8e-05 | 0.001217 | 0.3079 | 1.2983 | 2.5479 | 101 | 20.73 | 17.906 | 66.5724 | 105.8233 | 63.2743 | 0.004341 | 1.0494 | 0.025302 | 24.1107 | -0.10747 | 0.14209 | 9.4e-05 | 0.007158 | 0.007158 | 0.81585
2019-03-20 12:47 | 1553100431.1705 | -152 | 0.20949 | 0.000452 | 0 | 108.5022 | 109.794 | 127.807 | 20.9922 | 0.001843 | 0.001824 | -152.1542 | 0.23884 | 8.4039 | 0.003471 | 0.004712 | 0.004516 | 5.8e-05 | 0.00122 | 0.3126 | 1.3129 | 2.5481 | 101 | 20.73 | 17.903 | 65.9146 | 105.1335 | 65.466 | 0.004341 | 1.0426 | 0.02525 | 24.2182 | -0.10586 | 0.14209 | 0.000114 | 0.008668 | 0.008668 | 0.6737
2019-03-20 12:47 | 1553100432.1695 | -151 | 0.20948 | 0.000448 | 0 | 108.5866 | 109.7865 | 128.2638 | 20.992 | 0.001961 | 0.001725 | -152.1695 | 0.2384 | 8.4024 | 0.003475 | 0.005025 | 0.004774 | 5.9e-05 | 0.001219 | 0.3095 | 1.2977 | 2.5482 | 101 | 20.73 | 17.899 | 68.7579 | 105.2174 | 63.0663 | 0.00434 | 1.0501 | 0.025385 | 24.1753 | -0.098296 | 0.14182 | 0.000135 | 0.010362 | 0.010362 | 0.56255
2019-03-20 12:47 | 1553100433.1701 | -150 | 0.20956 | 0.000452 | 0 | 108.5171 | 109.8919 | 128.3347 | 20.993 | 0.001962 | 0.001764 | -152.1713 | 0.23813 | 8.4033 | 0.003549 | 0.00498 | 0.004758 | 6.1e-05 | 0.001219 | 0.28982 | 1.3555 | 2.5481 | 101 | 20.73 | 17.898 | 68.7691 | 107.0224 | 64.0092 | 0.004339 | 1.0471 | 0.025357 | 24.2159 | -0.11261 | 0.14166 | -0.000332 | -0.025461 | -0.025461 | -0.22867
2019-03-20 12:47 | 1553100434.1697 | -149 | 0.2096 | 0.000449 | 0 | 108.5711 | 109.9487 | 128.3596 | 20.9919 | 0.001902 | 0.001758 | -152.1596 | 0.23972 | 8.4042 | 0.003496 | 0.004675 | 0.004458 | 5.8e-05 | 0.001221 | 0.32189 | 1.3188 | 2.5481 | 101 | 20.73 | 17.897 | 67.3305 | 105.7326 | 63.855 | 0.004339 | 1.0476 | 0.025365 | 24.2127 | -0.11278 | 0.14261 | -0.000551 | -0.042164 | -0.042164 | -0.13901
2019-03-20 12:47 | 1553100435.1707 | -148 | 0.20951 | 0.000451 | 0 | 108.4883 | 109.874 | 128.2964 | 20.9925 | 0.001904 | 0.001751 | -152.1669 | 0.24029 | 8.4041 | 0.003506 | 0.005313 | 0.005114 | 6e-05 | 0.001227 | 0.31636 | 1.3294 | 2.5482 | 101 | 20.73 | 17.896 | 67.3826 | 105.9748 | 63.6942 | 0.004339 | 1.0481 | 0.025365 | 24.201 | -0.11352 | 0.14295 | -2.8e-05 | -0.002156 | -0.002156 | -2.7255
2019-03-20 12:47 | 1553100436.1712 | -147 | 0.2095 | 0.000454 | 0 | 108.5135 | 109.8747 | 128.2722 | 20.9918 | 0.001877 | 0.001728 | -152.1732 | 0.23911 | 8.4038 | 0.003521 | 0.004559 | 0.004337 | 5.8e-05 | 0.001229 | 0.32908 | 1.3396 | 2.5481 | 101 | 20.73 | 17.896 | 66.7249 | 106.326 | 63.1343 | 0.004339 | 1.0498 | 0.025384 | 24.1786 | -0.1115 | 0.14225 | 3e-05 | 0.002276 | 0.002276 | 2.5684
2019-03-20 12:47 | 1553100437.1698 | -146 | 0.20955 | 0.00045 | 1e-06 | 108.5193 | 109.7962 | 128.2225 | 20.9931 | 0.0019 | 0.001739 | -152.171 | 0.23945 | 8.4024 | 0.003477 | 0.005024 | 0.004755 | 5.8e-05 | 0.001228 | 0.33019 | 1.3756 | 2.5482 | 101 | 20.73 | 17.895 | 67.288 | 105.2767 | 63.4143 | 0.004339 | 1.049 | 0.025368 | 24.184 | -0.10463 | 0.14245 | -0.000252 | -0.019354 | -0.019354 | -0.3025
2019-03-20 12:47 | 1553100438.1699 | -145 | 0.2095 | 0.000456 | 0 | 108.531 | 109.9913 | 127.9495 | 20.9929 | 0.001851 | 0.001768 | -152.1819 | 0.23883 | 8.4033 | 0.003532 | 0.004788 | 0.004585 | 5.9e-05 | 0.001231 | 0.34739 | 1.3627 | 2.548 | 101 | 20.73 | 17.893 | 66.1034 | 106.6056 | 64.1138 | 0.004338 | 1.0468 | 0.025315 | 24.1833 | -0.11955 | 0.14208 | 2e-06 | 0.000207 | 0.000207 | 28.151
2019-03-20 12:47 | 1553100439.1709 | -144 | 0.20948 | 0.000456 | 0 | 108.5853 | 109.9035 | 127.7049 | 20.9927 | 0.001882 | 0.00176 | -152.1581 | 0.23912 | 8.4032 | 0.003461 | 0.00514 | 0.00493 | 6e-05 | 0.001227 | 0.30845 | 1.318 | 2.5481 | 101 | 20.73 | 17.893 | 66.853 | 104.8736 | 63.9167 | 0.004338 | 1.0474 | 0.025298 | 24.1531 | -0.10793 | 0.14225 | 0.000161 | 0.012237 | 0.012237 | 0.47779
2019-03-20 12:47 | 1553100440.17 | -143 | 0.20948 | 0.000451 | 0 | 108.6235 | 110.0253 | 127.6906 | 20.9918 | 0.001876 | 0.001786 | -152.1531 | 0.2393 | 8.4011 | 0.003473 | 0.005098 | 0.004885 | 6.1e-05 | 0.001221 | 0.32781 | 1.3411 | 2.548 | 101 | 20.73 | 17.891 | 66.7093 | 105.1677 | 64.535 | 0.004338 | 1.0455 | 0.025274 | 24.1739 | -0.1147 | 0.14236 | 0.000171 | 0.01304 | 0.01304 | 0.44871
2019-03-20 12:47 | 1553100441.1716 | -142 | 0.20949 | 0.000454 | 1e-06 | 108.5019 | 109.8527 | 127.3901 | 20.9929 | 0.001934 | 0.001759 | -152.1698 | 0.23869 | 8.4042 | 0.003435 | 0.004846 | 0.004662 | 6.1e-05 | 0.001217 | 0.30646 | 1.3339 | 2.5481 | 101 | 20.73 | 17.889 | 68.1125 | 104.2512 | 63.8897 | 0.004337 | 1.0475 | 0.025268 | 24.1224 | -0.11067 | 0.14199 | 9.2e-05 | 0.00694 | 0.00694 | 0.84091
2019-03-20 12:47 | 1553100442.1701 | -141 | 0.2095 | 0.00045 | 0 | 108.4977 | 109.7944 | 127.0979 | 20.9922 | 0.001837 | 0.001715 | -152.18 | 0.2388 | 8.4052 | 0.003423 | 0.004323 | 0.004134 | 6.1e-05 | 0.001215 | 0.32659 | 1.307 | 2.5481 | 101 | 20.73 | 17.888 | 65.7609 | 103.9507 | 62.84 | 0.004337 | 1.0508 | 0.025278 | 24.0571 | -0.10626 | 0.14206 | 1.8e-05 | 0.001384 | 0.001384 | 4.2189
2019-03-20 12:47 | 1553100443.1704 | -140 | 0.2095 | 0.000453 | 0 | 108.5058 | 109.868 | 127.0075 | 20.9923 | 0.001871 | 0.001807 | -152.1709 | 0.23896 | 8.4049 | 0.003419 | 0.004804 | 0.004627 | 6.4e-05 | 0.001208 | 0.32587 | 1.3422 | 2.548 | 101 | 20.73 | 17.887 | 66.5826 | 103.8667 | 65.0436 | 0.004337 | 1.0439 | 0.025187 | 24.1273 | -0.11159 | 0.14216 | 2.7e-05 | 0.002079 | 0.002079 | 2.8103
2019-03-20 12:47 | 1553100444.1698 | -139 | 0.20947 | 0.000448 | 0 | 108.5003 | 109.8295 | 127.3408 | 20.9931 | 0.001821 | 0.001732 | -152.1766 | 0.23892 | 8.4034 | 0.003383 | 0.005263 | 0.005047 | 6.1e-05 | 0.001211 | 0.3235 | 1.3399 | 2.5481 | 101 | 20.73 | 17.887 | 65.381 | 102.9844 | 63.2511 | 0.004337 | 1.0495 | 0.025287 | 24.0948 | -0.10891 | 0.14214 | 0.000237 | 0.018044 | 0.018044 | 0.32376
2019-03-20 12:47 | 1553100445.1703 | -138 | 0.20947 | 0.000449 | 0 | 108.4029 | 109.7596 | 127.3174 | 20.9925 | 0.001886 | 0.001755 | -152.1741 | 0.23894 | 8.4031 | 0.003409 | 0.00431 | 0.004118 | 6.4e-05 | 0.001213 | 0.33052 | 1.3146 | 2.5479 | 101 | 20.73 | 17.888 | 66.9481 | 103.6295 | 63.7933 | 0.004337 | 1.0478 | 0.025264 | 24.112 | -0.11125 | 0.14215 | 0.00018 | 0.013731 | 0.013731 | 0.42547
2019-03-20 12:47 | 1553100446.1702 | -137 | 0.20949 | 0.00045 | 0 | 108.4973 | 109.8403 | 127.7828 | 20.991 | 0.001901 | 0.001721 | -152.1679 | 0.23918 | 8.404 | 0.003417 | 0.004844 | 0.00463 | 6.4e-05 | 0.001216 | 0.3235 | 1.31 | 2.548 | 101 | 20.73 | 17.89 | 67.3085 | 103.8112 | 62.9647 | 0.004337 | 1.0504 | 0.025342 | 24.1263 | -0.11004 | 0.14229 | 8.5e-05 | 0.006465 | 0.006465 | 0.90462
2019-03-20 12:47 | 1553100447.1694 | -136 | 0.20951 | 0.000447 | 1e-06 | 108.5387 | 109.8609 | 127.8779 | 20.9913 | 0.001886 | 0.001745 | -152.1631 | 0.23856 | 8.4032 | 0.003434 | 0.005031 | 0.004857 | 6.2e-05 | 0.001217 | 0.31741 | 1.3142 | 2.5481 | 101 | 20.73 | 17.893 | 66.9375 | 104.2164 | 63.5524 | 0.004338 | 1.0485 | 0.025329 | 24.1564 | -0.1083 | 0.14192 | -4.4e-05 | -0.003398 | -0.003398 | -1.7165
2019-03-20 12:47 | 1553100448.1705 | -135 | 0.20951 | 0.000447 | 0 | 108.4935 | 109.7764 | 128.0328 | 20.9922 | 0.001902 | 0.001816 | -152.1451 | 0.23958 | 8.4018 | 0.003328 | 0.004743 | 0.004501 | 6e-05 | 0.001221 | 0.31636 | 1.3744 | 2.5481 | 101 | 20.73 | 17.899 | 67.3284 | 101.6607 | 65.2661 | 0.00434 | 1.0432 | 0.02528 | 24.2325 | -0.10515 | 0.14253 | -4e-06 | -0.000329 | -0.000329 | -17.7845
2019-03-20 12:47 | 1553100449.1695 | -134 | 0.20952 | 0.000445 | 0 | 108.4133 | 109.8231 | 128.0976 | 20.9928 | 0.001902 | 0.001773 | -152.1433 | 0.2384 | 8.4038 | 0.003266 | 0.005037 | 0.00477 | 6.3e-05 | 0.00122 | 0.31902 | 1.3736 | 2.5482 | 101 | 20.73 | 17.903 | 67.3376 | 100.1508 | 64.2214 | 0.004341 | 1.0465 | 0.025326 | 24.2011 | -0.11557 | 0.14182 | -5.5e-05 | -0.004252 | -0.004252 | -1.3709
2019-03-20 12:47 | 1553100450.1706 | -133 | 0.20952 | 0.000448 | 0 | 108.5032 | 109.7032 | 128.1239 | 20.9924 | 0.001884 | 0.001752 | -152.1474 | 0.23802 | 8.4053 | 0.003303 | 0.00482 | 0.004626 | 6.1e-05 | 0.001211 | 0.31636 | 1.3926 | 2.548 | 101 | 20.73 | 17.909 | 66.8926 | 101.0602 | 63.7111 | 0.004342 | 1.048 | 0.025347 | 24.1853 | -0.098379 | 0.14159 | -0.000112 | -0.008608 | -0.008608 | -0.6761
2019-03-20 12:48 | 1553100451.17 | -132 | 0.2095 | 0.00045 | 0 | 108.4663 | 109.9065 | 128.115 | 20.9918 | 0.001889 | 0.00177 | -152.1343 | 0.23872 | 8.4033 | 0.00335 | 0.00413 | 0.003885 | 6.3e-05 | 0.001214 | 0.335 | 1.3741 | 2.5482 | 101 | 20.73 | 17.913 | 67.0075 | 102.1887 | 64.1618 | 0.004343 | 1.0466 | 0.025329 | 24.2006 | -0.11798 | 0.14202 | 1.1e-05 | 0.00084 | 0.00084 | 6.9457
2019-03-20 12:48 | 1553100452.1701 | -131 | 0.20954 | 0.000449 | 0 | 108.4346 | 109.6206 | 128.2143 | 20.9932 | 0.001865 | 0.001675 | -152.1251 | 0.23945 | 8.4039 | 0.003339 | 0.00464 | 0.004439 | 6.2e-05 | 0.001212 | 0.32012 | 1.3314 | 2.5481 | 101 | 20.73 | 17.919 | 66.4307 | 101.9356 | 61.8646 | 0.004344 | 1.0538 | 0.025426 | 24.1274 | -0.097296 | 0.14245 | -0.000188 | -0.014448 | -0.014448 | -0.40522
2019-03-20 12:48 | 1553100453.1708 | -130 | 0.20952 | 0.000451 | 0 | 108.4353 | 109.8545 | 128.1694 | 20.9938 | 0.001907 | 0.001755 | -152.1129 | 0.23845 | 8.4051 | 0.003317 | 0.004901 | 0.0048 | 6.3e-05 | 0.001211 | 0.335 | 1.3069 | 2.548 | 101 | 20.72 | 17.929 | 67.444 | 101.4006 | 63.7984 | 0.004344 | 1.0478 | 0.025349 | 24.1927 | -0.11631 | 0.14185 | -0.000103 | -0.007862 | -0.007862 | -0.74153
2019-03-20 12:48 | 1553100454.1708 | -129 | 0.20951 | 0.000449 | 0 | 108.4754 | 109.8527 | 128.2699 | 20.9926 | 0.001927 | 0.001773 | -152.1008 | 0.23837 | 8.4037 | 0.003292 | 0.004625 | 0.004361 | 6.2e-05 | 0.001205 | 0.31708 | 1.2935 | 2.548 | 101 | 20.72 | 17.934 | 67.9217 | 100.795 | 64.2227 | 0.004345 | 1.0465 | 0.025342 | 24.2175 | -0.11285 | 0.1418 | -2.4e-05 | -0.001795 | -0.001795 | -3.2477
2019-03-20 12:48 | 1553100455.1708 | -128 | 0.2095 | 0.000448 | 0 | 108.3712 | 109.8007 | 128.3405 | 20.9935 | 0.001858 | 0.001758 | -152.0903 | 0.23922 | 8.4014 | 0.003384 | 0.004367 | 0.004217 | 6e-05 | 0.0012 | 0.32493 | 1.3259 | 2.5481 | 101 | 20.72 | 17.937 | 66.2819 | 103.0146 | 63.8646 | 0.004346 | 1.0476 | 0.025363 | 24.2113 | -0.11721 | 0.14231 | 5.4e-05 | 0.004136 | 0.004136 | 1.4143
2019-03-20 12:48 | 1553100456.1703 | -127 | 0.2095 | 0.000448 | 1e-06 | 108.423 | 109.862 | 128.026 | 20.9931 | 0.00183 | 0.001748 | -152.0908 | 0.23865 | 8.4055 | 0.003444 | 0.004491 | 0.004314 | 6.3e-05 | 0.001194 | 0.31598 | 1.3306 | 2.5481 | 101 | 20.72 | 17.943 | 65.5917 | 104.4588 | 63.6276 | 0.004348 | 1.0483 | 0.025341 | 24.1731 | -0.11793 | 0.14197 | 1.6e-05 | 0.001158 | 0.001158 | 5.0397
2019-03-20 12:48 | 1553100457.1699 | -126 | 0.20952 | 0.000448 | 0 | 108.3641 | 109.7058 | 127.9708 | 20.9932 | 0.001838 | 0.001744 | -152.0872 | 0.23931 | 8.4014 | 0.00339 | 0.004724 | 0.004499 | 6.1e-05 | 0.001188 | 0.32388 | 1.3304 | 2.5482 | 101 | 20.72 | 17.947 | 65.7806 | 103.1714 | 63.5261 | 0.004349 | 1.0486 | 0.025339 | 24.1642 | -0.11006 | 0.14236 | -6.6e-05 | -0.005031 | -0.005031 | -1.1631
2019-03-20 12:48 | 1553100458.1689 | -125 | 0.2095 | 0.000442 | 0 | 108.4265 | 109.825 | 128.0509 | 20.9925 | 0.001896 | 0.001731 | -152.1037 | 0.23903 | 8.4035 | 0.003358 | 0.003835 | 0.00359 | 5.8e-05 | 0.001184 | 0.32012 | 1.3341 | 2.548 | 101 | 20.72 | 17.952 | 67.1938 | 102.3835 | 63.2122 | 0.00435 | 1.0496 | 0.025359 | 24.1605 | -0.11463 | 0.1422 | 3.6e-05 | 0.0028 | 0.0028 | 2.0872
2019-03-20 12:48 | 1553100459.1694 | -124 | 0.20954 | 0.000442 | 0 | 108.4773 | 109.8171 | 128.229 | 20.9938 | 0.00183 | 0.001723 | -152.0992 | 0.23888 | 8.4041 | 0.003404 | 0.004714 | 0.004544 | 6.2e-05 | 0.001181 | 0.34003 | 1.3379 | 2.548 | 101 | 20.72 | 17.954 | 65.5946 | 103.491 | 63.019 | 0.00435 | 1.0502 | 0.025384 | 24.1704 | -0.1098 | 0.14211 | -0.000194 | -0.014808 | -0.014808 | -0.39443
2019-03-20 12:48 | 1553100460.1702 | -123 | 0.20947 | 0.000445 | 0 | 108.4042 | 109.7749 | 128.4649 | 20.9932 | 0.001866 | 0.001732 | -152.0964 | 0.23936 | 8.4029 | 0.003517 | 0.00526 | 0.005072 | 6.2e-05 | 0.001181 | 0.31669 | 1.356 | 2.5481 | 101 | 20.73 | 17.954 | 66.4665 | 106.2328 | 63.2488 | 0.004353 | 1.0495 | 0.025398 | 24.2009 | -0.11239 | 0.14239 | 0.000201 | 0.015432 | 0.015432 | 0.37924
2019-03-20 12:48 | 1553100461.1702 | -122 | 0.2095 | 0.000442 | 0 | 108.4307 | 109.7058 | 128.7375 | 20.9932 | 0.001829 | 0.001793 | -152.0871 | 0.23866 | 8.4026 | 0.003433 | 0.004745 | 0.004539 | 5.8e-05 | 0.00118 | 0.34092 | 1.3513 | 2.5481 | 101 | 20.73 | 17.958 | 65.5761 | 104.1947 | 64.7026 | 0.004354 | 1.045 | 0.025371 | 24.2788 | -0.10457 | 0.14198 | 3.3e-05 | 0.002516 | 0.002516 | 2.3195
2019-03-20 12:48 | 1553100462.1705 | -121 | 0.20954 | 0.000441 | 0 | 108.4301 | 109.7196 | 128.5767 | 20.9933 | 0.001886 | 0.00178 | -152.0688 | 0.23835 | 8.402 | 0.00345 | 0.005189 | 0.004951 | 5.9e-05 | 0.001182 | 0.31404 | 1.3495 | 2.548 | 101 | 20.73 | 17.96 | 66.9472 | 104.625 | 64.4046 | 0.004354 | 1.0459 | 0.025366 | 24.2529 | -0.10575 | 0.14179 | -0.000182 | -0.013909 | -0.013909 | -0.419
2019-03-20 12:48 | 1553100463.1696 | -120 | 0.20953 | 0.000444 | 0 | 108.4705 | 109.4711 | 128.617 | 20.993 | 0.001894 | 0.001711 | -152.0843 | 0.23843 | 8.4043 | 0.003377 | 0.004581 | 0.004432 | 5.7e-05 | 0.001188 | 0.3147 | 1.3847 | 2.5481 | 101 | 20.73 | 17.963 | 67.1328 | 102.8576 | 62.7323 | 0.004355 | 1.0511 | 0.025433 | 24.1966 | -0.082128 | 0.14184 | -0.000124 | -0.00952 | -0.00952 | -0.61237
2019-03-20 12:48 | 1553100464.1693 | -119 | 0.2095 | 0.000442 | 0 | 108.3745 | 109.7562 | 128.2128 | 20.9932 | 0.001828 | 0.001776 | -152.1001 | 0.23883 | 8.4041 | 0.00347 | 0.004912 | 0.004749 | 5.7e-05 | 0.001189 | 0.31492 | 1.3578 | 2.548 | 101 | 20.73 | 17.964 | 65.5466 | 105.0889 | 64.2974 | 0.004355 | 1.0462 | 0.025334 | 24.2148 | -0.11332 | 0.14208 | 6.4e-05 | 0.004879 | 0.004879 | 1.1968
2019-03-20 12:48 | 1553100465.1694 | -118 | 0.2095 | 0.000443 | 0 | 108.4217 | 109.7873 | 128.2018 | 20.9932 | 0.001852 | 0.001726 | -152.1102 | 0.23875 | 8.4016 | 0.00346 | 0.00568 | 0.005471 | 5.8e-05 | 0.001191 | 0.29535 | 1.3545 | 2.5481 | 101 | 20.73 | 17.966 | 66.1231 | 104.8513 | 63.1082 | 0.004356 | 1.0499 | 0.025378 | 24.171 | -0.11195 | 0.14203 | 4.7e-05 | 0.003549 | 0.003549 | 1.645
2019-03-20 12:48 | 1553100466.1707 | -117 | 0.20954 | 0.000442 | 0 | 108.4301 | 109.8265 | 127.988 | 20.9933 | 0.001869 | 0.001707 | -152.1146 | 0.23926 | 8.4013 | 0.00345 | 0.004718 | 0.004502 | 5.7e-05 | 0.001196 | 0.33069 | 1.3566 | 2.5481 | 101 | 20.73 | 17.971 | 66.533 | 104.6116 | 62.6451 | 0.004357 | 1.0514 | 0.025374 | 24.1342 | -0.11446 | 0.14234 | -0.000195 | -0.014927 | -0.014927 | -0.3919
2019-03-20 12:48 | 1553100467.1693 | -116 | 0.20943 | 0.000444 | 0 | 108.3745 | 109.6157 | 128.0687 | 20.9932 | 0.001918 | 0.001691 | -152.1271 | 0.23937 | 8.4045 | 0.003501 | 0.005163 | 0.004961 | 5.3e-05 | 0.001195 | 0.34844 | 1.3564 | 2.5481 | 101 | 20.73 | 17.973 | 67.7117 | 105.8558 | 62.252 | 0.004358 | 1.0526 | 0.025397 | 24.1277 | -0.10186 | 0.1424 | 0.000452 | 0.034606 | 0.034606 | 0.16913
2019-03-20 12:48 | 1553100468.1693 | -115 | 0.20953 | 0.000444 | 0 | 108.3664 | 109.8908 | 128.0187 | 20.993 | 0.001941 | 0.001761 | -152.1264 | 0.23864 | 8.4013 | 0.003509 | 0.003848 | 0.003676 | 5.5e-05 | 0.001197 | 0.33964 | 1.3292 | 2.548 | 101 | 20.73 | 17.974 | 68.2746 | 106.0422 | 63.9459 | 0.004358 | 1.0473 | 0.025328 | 24.1838 | -0.12495 | 0.14197 | -0.000147 | -0.011231 | -0.011231 | -0.51953
2019-03-20 12:48 | 1553100469.1704 | -114 | 0.2095 | 0.000441 | 0 | 108.3505 | 109.6336 | 127.814 | 20.9924 | 0.001904 | 0.00175 | -152.1256 | 0.23924 | 8.4028 | 0.00348 | 0.004454 | 0.004244 | 5.6e-05 | 0.001197 | 0.32117 | 1.2743 | 2.548 | 101 | 20.73 | 17.975 | 67.3762 | 105.3414 | 63.6658 | 0.004358 | 1.0482 | 0.025318 | 24.1544 | -0.1053 | 0.14233 | 3.7e-05 | 0.002789 | 0.002789 | 2.0973
2019-03-20 12:48 | 1553100470.1706 | -113 | 0.20943 | 0.000442 | 0 | 108.3712 | 109.8437 | 127.6645 | 20.9932 | 0.001884 | 0.001748 | -152.1296 | 0.23939 | 8.4005 | 0.003549 | 0.005063 | 0.004819 | 5.7e-05 | 0.0012 | 0.31686 | 1.317 | 2.5479 | 101 | 20.73 | 17.975 | 66.8871 | 107.017 | 63.6228 | 0.004358 | 1.0483 | 0.025305 | 24.1388 | -0.12071 | 0.14242 | 0.000452 | 0.034549 | 0.034549 | 0.16942
2019-03-20 12:48 | 1553100471.1695 | -112 | 0.20953 | 0.000443 | 0 | 108.4146 | 109.7891 | 127.5969 | 20.9938 | 0.001905 | 0.001795 | -152.1216 | 0.23954 | 8.4027 | 0.003613 | 0.005377 | 0.00518 | 5.6e-05 | 0.001199 | 0.30558 | 1.3442 | 2.548 | 101 | 20.73 | 17.975 | 67.4119 | 108.5755 | 64.764 | 0.004358 | 1.0448 | 0.025256 | 24.1732 | -0.11269 | 0.1425 | -0.000144 | -0.011003 | -0.011003 | -0.5323
2019-03-20 12:48 | 1553100472.17 | -111 | 0.20953 | 0.000443 | 0 | 108.4233 | 109.6796 | 127.6159 | 20.9934 | 0.001869 | 0.001759 | -152.1329 | 0.23922 | 8.4025 | 0.00354 | 0.005154 | 0.004945 | 5.5e-05 | 0.001201 | 0.29247 | 1.304 | 2.5481 | 101 | 20.73 | 17.974 | 66.5291 | 106.8071 | 63.8808 | 0.004358 | 1.0475 | 0.025291 | 24.1434 | -0.10305 | 0.14231 | -0.000112 | -0.008559 | -0.008559 | -0.68337
2019-03-20 12:48 | 1553100473.1697 | -110 | 0.20954 | 0.000443 | 0 | 108.3635 | 109.6538 | 127.6291 | 20.9942 | 0.001856 | 0.00169 | -152.1198 | 0.23921 | 8.4016 | 0.003544 | 0.004304 | 0.004155 | 5.4e-05 | 0.001207 | 0.34164 | 1.3002 | 2.5481 | 101 | 20.73 | 17.971 | 66.2319 | 106.8882 | 62.2297 | 0.004357 | 1.0527 | 0.025354 | 24.0854 | -0.10588 | 0.1423 | -0.000217 | -0.016609 | -0.016609 | -0.35213
2019-03-20 12:48 | 1553100474.1688 | -109 | 0.20955 | 0.000442 | 0 | 108.4188 | 109.7286 | 127.8054 | 20.9952 | 0.001861 | 0.001768 | -152.122 | 0.23919 | 8.4011 | 0.003539 | 0.004483 | 0.004245 | 5.6e-05 | 0.001203 | 0.32139 | 1.3472 | 2.5481 | 101 | 20.73 | 17.965 | 66.3468 | 106.7709 | 64.1114 | 0.004356 | 1.0468 | 0.025301 | 24.1696 | -0.10741 | 0.14229 | -0.000236 | -0.017989 | -0.017989 | -0.3251
2019-03-20 12:48 | 1553100475.1688 | -108 | 0.20951 | 0.000443 | 0 | 108.4049 | 109.7988 | 127.7986 | 20.994 | 0.001839 | 0.001741 | -152.1065 | 0.23947 | 8.402 | 0.003488 | 0.004956 | 0.004742 | 5.7e-05 | 0.001199 | 0.31758 | 1.3207 | 2.548 | 101 | 20.73 | 17.961 | 65.821 | 105.5263 | 63.4479 | 0.004355 | 1.0489 | 0.025325 | 24.1452 | -0.11428 | 0.14246 | -3e-06 | -0.000219 | -0.000219 | -26.7252
2019-03-20 12:48 | 1553100476.1688 | -107 | 0.20954 | 0.000444 | 0 | 108.4511 | 109.6336 | 127.834 | 20.9934 | 0.001863 | 0.001771 | -152.1126 | 0.23921 | 8.4007 | 0.003493 | 0.004869 | 0.004748 | 5.5e-05 | 0.001204 | 0.31924 | 1.3571 | 2.5481 | 101 | 20.72 | 17.951 | 66.3966 | 105.6696 | 64.1713 | 0.00435 | 1.0466 | 0.025301 | 24.1744 | -0.097002 | 0.14231 | -0.000189 | -0.01442 | -0.01442 | -0.4056
2019-03-20 12:48 | 1553100477.1699 | -106 | 0.20952 | 0.000442 | 0 | 108.4838 | 109.6837 | 127.6199 | 20.9934 | 0.001922 | 0.001729 | -152.095 | 0.23885 | 8.4034 | 0.003561 | 0.004626 | 0.004374 | 5.6e-05 | 0.001208 | 0.32875 | 1.3747 | 2.5481 | 101 | 20.72 | 17.944 | 67.8039 | 107.3006 | 63.1804 | 0.004348 | 1.0497 | 0.025317 | 24.1187 | -0.098394 | 0.14209 | -9.6e-05 | -0.007313 | -0.007313 | -0.79855
2019-03-20 12:48 | 1553100478.169 | -105 | 0.20954 | 0.000438 | 0 | 108.455 | 109.5873 | 127.5204 | 20.9944 | 0.00186 | 0.001718 | -152.0901 | 0.23945 | 8.4034 | 0.00343 | 0.00444 | 0.004236 | 5.5e-05 | 0.001212 | 0.33356 | 1.3159 | 2.548 | 101 | 20.73 | 17.937 | 66.3109 | 104.144 | 62.9165 | 0.004349 | 1.0505 | 0.025317 | 24.0998 | -0.0929 | 0.14245 | -0.000194 | -0.014843 | -0.014843 | -0.39444
2019-03-20 12:48 | 1553100479.169 | -104 | 0.20954 | 0.000439 | 0 | 108.3515 | 109.563 | 127.4883 | 20.9943 | 0.001839 | 0.00176 | -152.1192 | 0.23943 | 8.4023 | 0.003415 | 0.003932 | 0.003685 | 5.5e-05 | 0.001208 | 0.32892 | 1.3549 | 2.548 | 101 | 20.72 | 17.923 | 65.8183 | 103.7643 | 63.9176 | 0.004343 | 1.0474 | 0.025277 | 24.1326 | -0.099457 | 0.14244 | -0.000197 | -0.015057 | -0.015057 | -0.38881
2019-03-20 12:48 | 1553100480.17 | -103 | 0.2095 | 0.000437 | 0 | 108.3651 | 109.7061 | 127.3904 | 20.9946 | 0.001859 | 0.001736 | -152.1017 | 0.23927 | 8.4016 | 0.003424 | 0.004404 | 0.004167 | 5.5e-05 | 0.001208 | 0.31454 | 1.3451 | 2.548 | 101 | 20.72 | 17.915 | 66.308 | 103.9925 | 63.3353 | 0.004341 | 1.0492 | 0.025289 | 24.1025 | -0.11001 | 0.14234 | 6.8e-05 | 0.005149 | 0.005149 | 1.1362
2019-03-20 12:48 | 1553100481.1691 | -102 | 0.20953 | 0.000441 | 0 | 108.4372 | 109.6927 | 127.5312 | 20.9935 | 0.001884 | 0.001732 | -152.1108 | 0.23823 | 8.4021 | 0.003446 | 0.004217 | 0.004043 | 5.6e-05 | 0.00121 | 0.32388 | 1.3584 | 2.548 | 101 | 20.72 | 17.901 | 66.899 | 104.512 | 63.2498 | 0.004337 | 1.0495 | 0.025306 | 24.1128 | -0.10297 | 0.14172 | -0.000121 | -0.009249 | -0.009249 | -0.62976
2019-03-20 12:48 | 1553100482.1691 | -101 | 0.20954 | 0.000445 | 0 | 108.4792 | 109.785 | 127.6758 | 20.9928 | 0.001902 | 0.001709 | -152.1227 | 0.23871 | 8.4018 | 0.003396 | 0.004658 | 0.004493 | 5.4e-05 | 0.001205 | 0.3413 | 1.3543 | 2.5481 | 101 | 20.72 | 17.894 | 67.3415 | 103.3164 | 62.6836 | 0.004336 | 1.0513 | 0.025342 | 24.1061 | -0.10702 | 0.14201 | -0.000216 | -0.016522 | -0.016522 | -0.35326
2019-03-20 12:48 | 1553100483.1705 | -100 | 0.2095 | 0.000439 | 1e-06 | 108.4239 | 109.7417 | 127.488 | 20.994 | 0.001801 | 0.001712 | -152.1326 | 0.23891 | 8.4027 | 0.003504 | 0.005017 | 0.004841 | 5.8e-05 | 0.001207 | 0.33516 | 1.3544 | 2.5481 | 101 | 20.72 | 17.887 | 64.9017 | 105.9229 | 62.7498 | 0.004334 | 1.051 | 0.02532 | 24.0908 | -0.10805 | 0.14213 | 3e-05 | 0.002275 | 0.002275 | 2.5671
2019-03-20 12:48 | 1553100484.1702 | -99 | 0.20952 | 0.000442 | 0 | 108.3809 | 109.853 | 127.3113 | 20.9935 | 0.001861 | 0.001737 | -152.1221 | 0.23925 | 8.4033 | 0.003389 | 0.004657 | 0.004449 | 5.8e-05 | 0.00121 | 0.31542 | 1.3254 | 2.548 | 101 | 20.72 | 17.879 | 66.3397 | 103.1465 | 63.3625 | 0.004332 | 1.0491 | 0.02528 | 24.096 | -0.12067 | 0.14233 | -5.8e-05 | -0.004391 | -0.004391 | -1.3323
2019-03-20 12:48 | 1553100485.1697 | -98 | 0.20954 | 0.000442 | 0 | 108.4566 | 109.723 | 127.1437 | 20.9941 | 0.001884 | 0.001799 | -152.1326 | 0.23913 | 8.4025 | 0.003387 | 0.004707 | 0.004483 | 5.6e-05 | 0.001203 | 0.33804 | 1.3034 | 2.5481 | 101 | 20.72 | 17.874 | 66.8927 | 103.0803 | 64.8457 | 0.004331 | 1.0445 | 0.025208 | 24.1332 | -0.10383 | 0.14226 | -0.000207 | -0.015752 | -0.015752 | -0.37117
2019-03-20 12:48 | 1553100486.1694 | -97 | 0.20953 | 0.000445 | -1e-06 | 108.4084 | 109.7895 | 127.1789 | 20.994 | 0.001851 | 0.001752 | -152.132 | 0.23991 | 8.4036 | 0.003486 | 0.00524 | 0.005006 | 5.6e-05 | 0.001205 | 0.33069 | 1.3009 | 2.548 | 101 | 20.72 | 17.869 | 66.1176 | 105.4955 | 63.7151 | 0.00433 | 1.048 | 0.025253 | 24.0961 | -0.11323 | 0.14272 | -0.000163 | -0.012369 | -0.012369 | -0.47422
2019-03-20 12:48 | 1553100487.1694 | -96 | 0.2095 | 0.000441 | 0 | 108.4631 | 109.9031 | 127.2913 | 20.9932 | 0.001868 | 0.00176 | -152.1164 | 0.23883 | 8.3993 | 0.003439 | 0.004649 | 0.004466 | 5.6e-05 | 0.001205 | 0.33085 | 1.3239 | 2.548 | 101 | 20.72 | 17.861 | 66.5164 | 104.3457 | 63.9244 | 0.004328 | 1.0474 | 0.025257 | 24.1142 | -0.11797 | 0.14208 | 3.2e-05 | 0.002451 | 0.002451 | 2.3828
2019-03-20 12:48 | 1553100488.1698 | -95 | 0.20949 | 0.000442 | 0 | 108.4553 | 109.9289 | 127.4451 | 20.9935 | 0.001867 | 0.00171 | -152.1237 | 0.23911 | 8.3995 | 0.003418 | 0.004295 | 0.004196 | 5.8e-05 | 0.001209 | 0.31116 | 1.3477 | 2.5481 | 101 | 20.72 | 17.857 | 66.4871 | 103.8469 | 62.7087 | 0.004327 | 1.0512 | 0.025318 | 24.0852 | -0.12071 | 0.14225 | 8.8e-05 | 0.006746 | 0.006746 | 0.8667
2019-03-20 12:48 | 1553100489.1685 | -94 | 0.20953 | 0.000442 | 0 | 108.434 | 109.7727 | 127.5249 | 20.9932 | 0.001833 | 0.001755 | -152.1212 | 0.2397 | 8.4036 | 0.003471 | 0.004266 | 0.004017 | 5.3e-05 | 0.001202 | 0.31797 | 1.3545 | 2.5481 | 101 | 20.72 | 17.852 | 65.6628 | 105.1204 | 63.7885 | 0.004326 | 1.0478 | 0.025285 | 24.1315 | -0.10975 | 0.1426 | -0.000138 | -0.010493 | -0.010493 | -0.55852
2019-03-20 12:48 | 1553100490.1686 | -93 | 0.20956 | 0.00044 | 0 | 108.486 | 109.9024 | 127.7014 | 20.9945 | 0.001862 | 0.001773 | -152.1318 | 0.23989 | 8.4024 | 0.003438 | 0.00484 | 0.004646 | 5.8e-05 | 0.001204 | 0.33483 | 1.3149 | 2.5479 | 101 | 20.72 | 17.851 | 66.382 | 104.3305 | 64.2301 | 0.004325 | 1.0464 | 0.025286 | 24.164 | -0.11602 | 0.14271 | -0.000311 | -0.023794 | -0.023794 | -0.24651
2019-03-20 12:48 | 1553100491.169 | -92 | 0.20952 | 0.000443 | 0 | 108.4589 | 109.7402 | 127.3807 | 20.9936 | 0.001895 | 0.001723 | -152.1267 | 0.2384 | 8.404 | 0.00342 | 0.004514 | 0.004347 | 5.6e-05 | 0.001207 | 0.30735 | 1.3032 | 2.548 | 101 | 20.72 | 17.851 | 67.1692 | 103.8964 | 63.0196 | 0.004325 | 1.0502 | 0.0253 | 24.0903 | -0.10505 | 0.14182 | -8.4e-05 | -0.006441 | -0.006441 | -0.90499
2019-03-20 12:48 | 1553100492.1687 | -91 | 0.20948 | 0.000443 | 0 | 108.5045 | 109.6927 | 127.3021 | 20.9936 | 0.001892 | 0.00177 | -152.1287 | 0.2382 | 8.4057 | 0.003368 | 0.004973 | 0.004688 | 5.6e-05 | 0.001215 | 0.36459 | 1.3031 | 2.5481 | 101 | 20.72 | 17.85 | 67.0926 | 102.6296 | 64.1502 | 0.004325 | 1.0467 | 0.025249 | 24.1233 | -0.097419 | 0.14171 | 0.000163 | 0.012414 | 0.012414 | 0.46918
2019-03-20 12:48 | 1553100493.1697 | -90 | 0.20953 | 0.000441 | 0 | 108.4876 | 109.8799 | 127.2815 | 20.9933 | 0.001842 | 0.001725 | -152.1346 | 0.23949 | 8.401 | 0.003309 | 0.005021 | 0.004842 | 5.5e-05 | 0.001212 | 0.33323 | 1.3061 | 2.5481 | 101 | 20.72 | 17.849 | 65.8875 | 101.1906 | 63.0731 | 0.004325 | 1.05 | 0.025288 | 24.0828 | -0.11406 | 0.14247 | -0.000117 | -0.008945 | -0.008945 | -0.65462
2019-03-20 12:48 | 1553100494.1698 | -89 | 0.2095 | 0.000444 | 0 | 108.4459 | 109.9154 | 127.3697 | 20.9935 | 0.001937 | 0.001728 | -152.1327 | 0.23965 | 8.4015 | 0.003351 | 0.004793 | 0.00461 | 5.6e-05 | 0.001214 | 0.31454 | 1.3033 | 2.548 | 101 | 20.72 | 17.849 | 68.1808 | 102.2274 | 63.153 | 0.004325 | 1.0498 | 0.025294 | 24.094 | -0.12039 | 0.14257 | 2.9e-05 | 0.002227 | 0.002227 | 2.6316
2019-03-20 12:48 | 1553100495.1688 | -88 | 0.2095 | 0.000441 | -1e-06 | 108.4893 | 109.89 | 127.3455 | 20.9947 | 0.001829 | 0.00176 | -152.1265 | 0.23796 | 8.4045 | 0.00333 | 0.004455 | 0.004259 | 5.7e-05 | 0.001215 | 0.33395 | 1.3019 | 2.548 | 101 | 20.72 | 17.848 | 65.5658 | 101.703 | 63.9273 | 0.004325 | 1.0474 | 0.025262 | 24.1195 | -0.11475 | 0.14156 | 4.9e-05 | 0.003799 | 0.003799 | 1.5316
2019-03-20 12:48 | 1553100496.17 | -87 | 0.2095 | 0.000446 | 1e-06 | 108.4385 | 109.8328 | 127.5082 | 20.9942 | 0.001853 | 0.001716 | -152.1227 | 0.23824 | 8.4016 | 0.003283 | 0.005122 | 0.004883 | 5.5e-05 | 0.001219 | 0.3246 | 1.3069 | 2.548 | 101 | 20.72 | 17.845 | 66.1554 | 100.5637 | 62.8523 | 0.004324 | 1.0507 | 0.025319 | 24.0964 | -0.11428 | 0.14173 | 7.3e-05 | 0.005505 | 0.005505 | 1.0581
2019-03-20 12:48 | 1553100497.1712 | -86 | 0.20955 | 0.000443 | 0 | 108.4792 | 109.8777 | 127.603 | 20.995 | 0.001901 | 0.001788 | -152.1401 | 0.23902 | 8.4048 | 0.003381 | 0.004578 | 0.004469 | 5.6e-05 | 0.001219 | 0.34562 | 1.3042 | 2.5481 | 101 | 20.73 | 17.846 | 67.3061 | 102.9421 | 64.5982 | 0.004327 | 1.0453 | 0.025262 | 24.1679 | -0.11457 | 0.14219 | -0.00022 | -0.01683 | -0.01683 | -0.34725
2019-03-20 12:48 | 1553100498.17 | -85 | 0.20949 | 0.000443 | 0 | 108.564 | 109.8601 | 127.6375 | 20.9948 | 0.001909 | 0.001766 | -152.14 | 0.23969 | 8.4032 | 0.003311 | 0.004691 | 0.00448 | 5.7e-05 | 0.001222 | 0.33516 | 1.3131 | 2.5481 | 101 | 20.73 | 17.844 | 67.5127 | 101.2494 | 64.0655 | 0.004326 | 1.0469 | 0.025286 | 24.1521 | -0.10616 | 0.14259 | 9.5e-05 | 0.007236 | 0.007236 | 0.80993
2019-03-20 12:48 | 1553100499.1704 | -84 | 0.20951 | 0.000446 | 0 | 108.5168 | 109.8717 | 127.6366 | 20.9947 | 0.001883 | 0.001748 | -152.111 | 0.24043 | 8.3995 | 0.003268 | 0.00495 | 0.004654 | 5.7e-05 | 0.001219 | 0.32516 | 1.3372 | 2.548 | 101 | 20.73 | 17.843 | 66.873 | 100.2163 | 63.6274 | 0.004326 | 1.0483 | 0.025302 | 24.1363 | -0.11099 | 0.14303 | -1.9e-05 | -0.001463 | -0.001463 | -4.0193
2019-03-20 12:48 | 1553100500.1701 | -83 | 0.20952 | 0.00044 | 0 | 108.6054 | 109.8904 | 127.5336 | 20.9931 | 0.001914 | 0.001741 | -152.1068 | 0.23909 | 8.4037 | 0.003364 | 0.004481 | 0.004256 | 5.7e-05 | 0.001215 | 0.28926 | 1.2955 | 2.5481 | 101 | 20.73 | 17.841 | 67.6163 | 102.5238 | 63.4696 | 0.004326 | 1.0488 | 0.025298 | 24.1209 | -0.10521 | 0.14224 | -5.6e-05 | -0.004285 | -0.004285 | -1.3641
2019-03-20 12:48 | 1553100501.1714 | -82 | 0.20948 | 0.000444 | 0 | 108.5339 | 109.6123 | 127.3704 | 20.9944 | 0.001891 | 0.001735 | -152.1041 | 0.23936 | 8.4042 | 0.003369 | 0.004281 | 0.004099 | 5.4e-05 | 0.001219 | 0.34346 | 1.361 | 2.5479 | 101 | 20.73 | 17.835 | 67.0567 | 102.6658 | 63.3053 | 0.004324 | 1.0493 | 0.025288 | 24.0996 | -0.088438 | 0.1424 | 0.000191 | 0.014604 | 0.014604 | 0.40075
2019-03-20 12:48 | 1553100502.1692 | -81 | 0.2095 | 0.000442 | 0 | 108.5759 | 109.8665 | 127.0888 | 20.9944 | 0.001833 | 0.001708 | -152.0967 | 0.23847 | 8.4023 | 0.003298 | 0.004314 | 0.004062 | 5.7e-05 | 0.001216 | 0.31078 | 1.3082 | 2.548 | 101 | 20.73 | 17.83 | 65.6851 | 100.9476 | 62.6674 | 0.004323 | 1.0513 | 0.025284 | 24.0501 | -0.10569 | 0.14187 | 7.6e-05 | 0.005799 | 0.005799 | 1.0054
2019-03-20 12:48 | 1553100503.1693 | -80 | 0.2095 | 0.000442 | 0 | 108.4556 | 109.7816 | 127.5218 | 20.9941 | 0.001861 | 0.001736 | -152.0838 | 0.23781 | 8.4073 | 0.003314 | 0.004848 | 0.004649 | 5.6e-05 | 0.001211 | 0.29856 | 1.3237 | 2.5481 | 101 | 20.73 | 17.824 | 66.3354 | 101.324 | 63.3457 | 0.004321 | 1.0492 | 0.025301 | 24.1153 | -0.1087 | 0.14147 | 5.5e-05 | 0.004172 | 0.004172 | 1.3938
2019-03-20 12:48 | 1553100504.1693 | -79 | 0.20952 | 0.000442 | 0 | 108.5235 | 110.1027 | 127.5798 | 20.9939 | 0.001866 | 0.001759 | -152.0934 | 0.23859 | 8.4014 | 0.003352 | 0.004815 | 0.004575 | 5.6e-05 | 0.001213 | 0.3183 | 1.348 | 2.5481 | 101 | 20.73 | 17.813 | 66.4665 | 102.2403 | 63.8969 | 0.004319 | 1.0475 | 0.025286 | 24.1406 | -0.12922 | 0.14194 | -6.4e-05 | -0.00489 | -0.00489 | -1.193
2019-03-20 12:48 | 1553100505.1688 | -78 | 0.20947 | 0.000439 | 0 | 108.5216 | 109.8414 | 128.0509 | 20.9942 | 0.00186 | 0.001744 | -152.1124 | 0.23972 | 8.4068 | 0.003412 | 0.004932 | 0.004724 | 5.4e-05 | 0.001212 | 0.31564 | 1.3209 | 2.548 | 101 | 20.73 | 17.808 | 66.3264 | 103.7036 | 63.5365 | 0.004318 | 1.0486 | 0.025347 | 24.1722 | -0.10813 | 0.14261 | 0.000223 | 0.017027 | 0.017027 | 0.34424
2019-03-20 12:48 | 1553100506.1694 | -77 | 0.20947 | 0.000443 | 0 | 108.5636 | 109.8183 | 128.1136 | 20.9938 | 0.00183 | 0.00174 | -152.1007 | 0.23933 | 8.402 | 0.003439 | 0.004494 | 0.004349 | 5.6e-05 | 0.001213 | 0.28385 | 1.3323 | 2.5481 | 101 | 20.73 | 17.794 | 65.6087 | 104.362 | 63.4344 | 0.004314 | 1.0489 | 0.025357 | 24.1744 | -0.10278 | 0.14238 | 0.000237 | 0.018197 | 0.018197 | 0.32159
2019-03-20 12:48 | 1553100507.1696 | -76 | 0.20951 | 0.000439 | 0 | 108.5604 | 109.8672 | 128.0053 | 20.9942 | 0.001827 | 0.001744 | -152.1178 | 0.23935 | 8.4014 | 0.003415 | 0.004527 | 0.004314 | 5.1e-05 | 0.001217 | 0.31332 | 1.3441 | 2.5481 | 100.99 | 20.73 | 17.788 | 65.5187 | 103.7783 | 63.5435 | 0.004313 | 1.0486 | 0.025342 | 24.1681 | -0.10703 | 0.14239 | -4e-06 | -0.000275 | -0.000275 | -21.2621
2019-03-20 12:48 | 1553100508.1684 | -75 | 0.20952 | 0.000435 | 0 | 108.5811 | 109.856 | 127.7558 | 20.9943 | 0.001947 | 0.001733 | -152.1068 | 0.23914 | 8.4027 | 0.003372 | 0.004784 | 0.004583 | 5.6e-05 | 0.001222 | 0.33107 | 1.3598 | 2.5481 | 100.99 | 20.73 | 17.783 | 68.4241 | 102.7395 | 63.2647 | 0.004312 | 1.0494 | 0.025328 | 24.1345 | -0.10441 | 0.14226 | -4.5e-05 | -0.003419 | -0.003419 | -1.71
2019-03-20 12:48 | 1553100509.1688 | -74 | 0.20951 | 0.000447 | 0 | 108.531 | 110.0022 | 127.5793 | 20.9939 | 0.001891 | 0.001714 | -152.1371 | 0.23854 | 8.4023 | 0.003317 | 0.004773 | 0.00452 | 5.4e-05 | 0.001221 | 0.32565 | 1.3293 | 2.5482 | 100.99 | 20.73 | 17.771 | 67.0711 | 101.4056 | 62.8005 | 0.004309 | 1.0509 | 0.025328 | 24.1012 | -0.12043 | 0.14191 | 1.7e-05 | 0.001326 | 0.001326 | 4.3986
2019-03-20 12:48 | 1553100510.1686 | -73 | 0.20953 | 0.000444 | 0 | 108.5873 | 109.988 | 127.5007 | 20.9937 | 0.001851 | 0.00176 | -152.1364 | 0.23944 | 8.4026 | 0.003305 | 0.004393 | 0.004193 | 5.3e-05 | 0.001219 | 0.33516 | 1.3496 | 2.5481 | 100.99 | 20.73 | 17.767 | 66.1041 | 101.1077 | 63.928 | 0.004308 | 1.0474 | 0.025277 | 24.1342 | -0.11464 | 0.14244 | -0.000136 | -0.010347 | -0.010347 | -0.56582
2019-03-20 12:48 | 1553100511.1697 | -72 | 0.20951 | 0.000445 | 0 | 108.5847 | 109.9689 | 127.5554 | 20.9935 | 0.001848 | 0.001759 | -152.1504 | 0.23926 | 8.4019 | 0.003359 | 0.004706 | 0.004488 | 5.8e-05 | 0.001217 | 0.29856 | 1.3424 | 2.5481 | 101 | 20.73 | 17.761 | 66.033 | 102.4159 | 63.8941 | 0.004306 | 1.0475 | 0.025284 | 24.1382 | -0.1133 | 0.14233 | 1.5e-05 | 0.001151 | 0.001151 | 5.0845
2019-03-20 12:49 | 1553100512.1711 | -71 | 0.20952 | 0.000441 | 0 | 108.542 | 109.8284 | 127.8011 | 20.9941 | 0.001814 | 0.001765 | -152.1572 | 0.23955 | 8.4036 | 0.003329 | 0.00475 | 0.004639 | 5.3e-05 | 0.001218 | 0.31741 | 1.3056 | 2.5482 | 101 | 20.73 | 17.759 | 65.222 | 101.6766 | 64.0252 | 0.004306 | 1.0471 | 0.025304 | 24.1661 | -0.10538 | 0.14251 | -8.7e-05 | -0.006655 | -0.006655 | -0.88006
2019-03-20 12:49 | 1553100513.1688 | -70 | 0.20949 | 0.000441 | 0 | 108.5921 | 109.8575 | 127.3268 | 20.9933 | 0.001906 | 0.001684 | -152.1477 | 0.23952 | 8.4009 | 0.003336 | 0.004242 | 0.003917 | 5.5e-05 | 0.001221 | 0.33157 | 1.3232 | 2.5481 | 101 | 20.73 | 17.756 | 67.4191 | 101.8657 | 62.0756 | 0.004305 | 1.0532 | 0.02533 | 24.0514 | -0.10363 | 0.14249 | 0.000153 | 0.011714 | 0.011714 | 0.49997
2019-03-20 12:49 | 1553100514.17 | -69 | 0.20949 | 0.000441 | 0 | 108.5798 | 109.8814 | 127.3587 | 20.9936 | 0.001828 | 0.001695 | -152.1312 | 0.23907 | 8.4019 | 0.003393 | 0.005464 | 0.005294 | 5.7e-05 | 0.001213 | 0.31542 | 1.3482 | 2.5481 | 101 | 20.73 | 17.752 | 65.5522 | 103.2425 | 62.3555 | 0.004304 | 1.0523 | 0.025322 | 24.0644 | -0.10659 | 0.14222 | 0.000115 | 0.008773 | 0.008773 | 0.66631
2019-03-20 12:49 | 1553100515.169 | -68 | 0.20951 | 0.000446 | 0 | 108.5436 | 109.8257 | 127.0089 | 20.9936 | 0.001843 | 0.001748 | -152.1237 | 0.23869 | 8.4029 | 0.003375 | 0.004516 | 0.004336 | 5.5e-05 | 0.001211 | 0.30845 | 1.34 | 2.5482 | 101 | 20.73 | 17.751 | 65.909 | 102.8103 | 63.6201 | 0.004304 | 1.0483 | 0.02524 | 24.0766 | -0.10504 | 0.142 | -1.6e-05 | -0.001201 | -0.001201 | -4.8581
2019-03-20 12:49 | 1553100516.1699 | -67 | 0.20952 | 0.000446 | 0 | 108.5727 | 110.0978 | 127.0657 | 20.9923 | 0.001848 | 0.001753 | -152.1172 | 0.23887 | 8.4023 | 0.003226 | 0.004147 | 0.003934 | 6e-05 | 0.001212 | 0.35026 | 1.3179 | 2.548 | 101 | 20.73 | 17.747 | 66.0419 | 99.2035 | 63.7468 | 0.004303 | 1.0479 | 0.025241 | 24.0865 | -0.12477 | 0.1421 | -5.7e-05 | -0.004366 | -0.004366 | -1.3378
2019-03-20 12:49 | 1553100517.17 | -66 | 0.20955 | 0.000443 | 0 | 108.5203 | 109.8355 | 127.2563 | 20.9939 | 0.001856 | 0.001752 | -152.1038 | 0.23921 | 8.402 | 0.003251 | 0.004132 | 0.003919 | 5.6e-05 | 0.001209 | 0.33301 | 1.3299 | 2.5481 | 101 | 20.73 | 17.746 | 66.2176 | 99.8013 | 63.7282 | 0.004303 | 1.048 | 0.025261 | 24.1039 | -0.10775 | 0.1423 | -0.00023 | -0.017586 | -0.017586 | -0.33258
2019-03-20 12:49 | 1553100518.1696 | -65 | 0.20953 | 0.000439 | 0 | 108.5963 | 110.0549 | 127.4784 | 20.994 | 0.001799 | 0.001718 | -152.101 | 0.23868 | 8.4022 | 0.003341 | 0.004693 | 0.004457 | 5.4e-05 | 0.00121 | 0.31205 | 1.3581 | 2.5481 | 101 | 20.73 | 17.745 | 64.8557 | 101.9734 | 62.9152 | 0.004302 | 1.0505 | 0.025313 | 24.0958 | -0.11933 | 0.14199 | -0.000123 | -0.009379 | -0.009379 | -0.62223
2019-03-20 12:49 | 1553100519.1701 | -64 | 0.20954 | 0.000441 | 1e-06 | 108.5721 | 110.0257 | 127.4737 | 20.9936 | 0.001887 | 0.00171 | -152.0994 | 0.23981 | 8.4003 | 0.003367 | 0.00489 | 0.004701 | 5.5e-05 | 0.001205 | 0.29015 | 1.3267 | 2.5481 | 101 | 20.73 | 17.743 | 66.9616 | 102.6145 | 62.7157 | 0.004302 | 1.0512 | 0.02532 | 24.0882 | -0.11896 | 0.14266 | -0.000185 | -0.014163 | -0.014163 | -0.414
2019-03-20 12:49 | 1553100520.1696 | -63 | 0.20954 | 0.000442 | 0 | 108.5607 | 109.9439 | 127.4786 | 20.9939 | 0.001897 | 0.001678 | -152.0999 | 0.23827 | 8.4012 | 0.003324 | 0.004593 | 0.004371 | 5.7e-05 | 0.001209 | 0.32261 | 1.3358 | 2.5481 | 101 | 20.73 | 17.741 | 67.2224 | 101.5532 | 61.9544 | 0.004301 | 1.0535 | 0.02535 | 24.0614 | -0.11324 | 0.14175 | -0.000195 | -0.014953 | -0.014953 | -0.3896
2019-03-20 12:49 | 1553100521.1686 | -62 | 0.20955 | 0.000444 | 0 | 108.5045 | 109.9491 | 127.512 | 20.9942 | 0.0019 | 0.001663 | -152.0868 | 0.23902 | 8.4013 | 0.003384 | 0.004525 | 0.004326 | 5.9e-05 | 0.001205 | 0.32748 | 1.3083 | 2.548 | 100.99 | 20.73 | 17.74 | 67.2789 | 103.0177 | 61.5929 | 0.004301 | 1.0547 | 0.025367 | 24.0515 | -0.1183 | 0.14219 | -0.000229 | -0.017532 | -0.017532 | -0.33335
2019-03-20 12:49 | 1553100522.1693 | -61 | 0.20952 | 0.000446 | 0 | 108.6264 | 109.8246 | 127.0728 | 20.9937 | 0.001794 | 0.001733 | -152.079 | 0.23913 | 8.4025 | 0.003363 | 0.00535 | 0.005138 | 5.5e-05 | 0.001204 | 0.33395 | 1.3157 | 2.5481 | 100.99 | 20.73 | 17.736 | 64.7375 | 102.5117 | 63.2579 | 0.004301 | 1.0495 | 0.02526 | 24.0697 | -0.098125 | 0.14226 | -9.3e-05 | -0.007121 | -0.007121 | -0.82104
2019-03-20 12:49 | 1553100523.1693 | -60 | 0.20953 | 0.00044 | 0 | 108.6093 | 109.9016 | 127.1136 | 20.9946 | 0.001804 | 0.001734 | -152.0829 | 0.23878 | 8.4025 | 0.003389 | 0.004991 | 0.004928 | 5.5e-05 | 0.001203 | 0.31061 | 1.2743 | 2.548 | 100.99 | 20.73 | 17.734 | 64.9711 | 103.142 | 63.2807 | 0.0043 | 1.0494 | 0.025263 | 24.0744 | -0.1058 | 0.14205 | -9e-05 | -0.006896 | -0.006896 | -0.84666
2019-03-20 12:49 | 1553100524.1698 | -59 | 0.20957 | 0.000443 | 0 | 108.6319 | 109.9644 | 127.0723 | 20.9952 | 0.001874 | 0.001698 | -152.1064 | 0.23901 | 8.4016 | 0.003433 | 0.005125 | 0.004816 | 5.4e-05 | 0.001201 | 0.30829 | 1.3177 | 2.5481 | 100.99 | 20.73 | 17.734 | 66.6496 | 104.2171 | 62.4235 | 0.0043 | 1.0521 | 0.025291 | 24.0398 | -0.10905 | 0.14219 | -0.000369 | -0.028202 | -0.028202 | -0.20722
2019-03-20 12:49 | 1553100525.1685 | -58 | 0.20949 | 0.000444 | 0 | 108.5607 | 109.8601 | 126.981 | 20.9944 | 0.001891 | 0.001774 | -152.102 | 0.23987 | 8.4049 | 0.003431 | 0.004813 | 0.004638 | 5.5e-05 | 0.001199 | 0.31094 | 1.3136 | 2.5481 | 100.99 | 20.73 | 17.735 | 67.0702 | 104.1528 | 64.2512 | 0.0043 | 1.0464 | 0.025214 | 24.0965 | -0.10642 | 0.1427 | 0.000152 | 0.011519 | 0.011519 | 0.50916
2019-03-20 12:49 | 1553100526.1695 | -57 | 0.20959 | 0.000445 | 1e-06 | 108.5824 | 110.0265 | 126.8451 | 20.9942 | 0.001825 | 0.001794 | -152.1362 | 0.23942 | 8.4013 | 0.003397 | 0.005243 | 0.005074 | 5.7e-05 | 0.001199 | 0.3079 | 1.3059 | 2.5481 | 100.99 | 20.73 | 17.735 | 65.4905 | 103.3335 | 64.7435 | 0.0043 | 1.0448 | 0.025182 | 24.1012 | -0.11817 | 0.14243 | -0.000492 | -0.037439 | -0.037439 | -0.15636
2019-03-20 12:49 | 1553100527.169 | -56 | 0.20951 | 0.000444 | 0 | 108.5902 | 109.9397 | 126.5374 | 20.9934 | 0.001864 | 0.0018 | -152.1492 | 0.23952 | 8.4025 | 0.003373 | 0.005071 | 0.0049 | 5.2e-05 | 0.001196 | 0.33588 | 1.3626 | 2.548 | 100.99 | 20.73 | 17.736 | 66.4286 | 102.7544 | 64.8879 | 0.004301 | 1.0444 | 0.025146 | 24.0771 | -0.11048 | 0.14249 | -8e-06 | -0.000603 | -0.000603 | -9.7175
2019-03-20 12:49 | 1553100528.1697 | -55 | 0.20954 | 0.000448 | 1e-06 | 108.5934 | 109.8377 | 126.4961 | 20.9944 | 0.001908 | 0.001734 | -152.1608 | 0.2394 | 8.4022 | 0.003378 | 0.004752 | 0.004466 | 5.3e-05 | 0.001195 | 0.34346 | 1.3276 | 2.548 | 100.99 | 20.73 | 17.738 | 67.4654 | 102.8739 | 63.2875 | 0.004301 | 1.0494 | 0.025202 | 24.0161 | -0.10191 | 0.14242 | -0.000162 | -0.012344 | -0.012344 | -0.47421
2019-03-20 12:49 | 1553100529.1687 | -54 | 0.20953 | 0.000442 | 0 | 108.5776 | 110.0208 | 126.6329 | 20.9944 | 0.001825 | 0.001698 | -152.1396 | 0.23904 | 8.4019 | 0.003481 | 0.004543 | 0.004385 | 5.4e-05 | 0.001192 | 0.32045 | 1.2974 | 2.5482 | 100.99 | 20.73 | 17.742 | 65.4789 | 105.363 | 62.4259 | 0.004302 | 1.0521 | 0.025248 | 23.9983 | -0.11811 | 0.1422 | -0.000123 | -0.009395 | -0.009395 | -0.62208
2019-03-20 12:49 | 1553100530.1691 | -53 | 0.2095 | 0.000445 | -1e-06 | 108.5704 | 109.7858 | 126.8313 | 20.9941 | 0.00184 | 0.00169 | -152.1341 | 0.23896 | 8.4033 | 0.003419 | 0.00487 | 0.004658 | 5.7e-05 | 0.001193 | 0.31636 | 1.3202 | 2.5479 | 100.99 | 20.73 | 17.742 | 65.8486 | 103.8663 | 62.2202 | 0.004302 | 1.0527 | 0.025275 | 24.0097 | -0.099567 | 0.14216 | 6.9e-05 | 0.005316 | 0.005316 | 1.099
2019-03-20 12:49 | 1553100531.1698 | -52 | 0.20953 | 0.000447 | 0 | 108.5494 | 110.0096 | 127.119 | 20.9954 | 0.001882 | 0.001742 | -152.1322 | 0.23925 | 8.404 | 0.003431 | 0.00489 | 0.004627 | 5.5e-05 | 0.001191 | 0.30901 | 1.3726 | 2.548 | 100.99 | 20.73 | 17.745 | 66.8573 | 104.1474 | 63.4798 | 0.004303 | 1.0488 | 0.025256 | 24.082 | -0.11952 | 0.14233 | -0.000135 | -0.010262 | -0.010262 | -0.57004
2019-03-20 12:49 | 1553100532.1704 | -51 | 0.20955 | 0.000441 | 0 | 108.586 | 109.9001 | 127.7509 | 20.995 | 0.001849 | 0.001702 | -152.1367 | 0.23891 | 8.4045 | 0.003475 | 0.00553 | 0.005273 | 5.5e-05 | 0.00119 | 0.3178 | 1.341 | 2.548 | 100.99 | 20.73 | 17.749 | 66.0563 | 105.2118 | 62.5141 | 0.004304 | 1.0518 | 0.025355 | 24.1071 | -0.1076 | 0.14213 | -0.000213 | -0.016309 | -0.016309 | -0.35818
2019-03-20 12:49 | 1553100533.1689 | -50 | 0.20951 | 0.000443 | 0 | 108.6749 | 109.9726 | 127.8129 | 20.9951 | 0.001878 | 0.001809 | -152.1458 | 0.23883 | 8.4008 | 0.003386 | 0.005199 | 0.005089 | 5.4e-05 | 0.001195 | 0.30862 | 1.318 | 2.548 | 100.99 | 20.73 | 17.752 | 66.7517 | 103.0774 | 65.094 | 0.004304 | 1.0438 | 0.025265 | 24.2055 | -0.10618 | 0.14208 | -1.9e-05 | -0.00147 | -0.00147 | -3.9732
2019-03-20 12:49 | 1553100534.1693 | -49 | 0.20952 | 0.000443 | 0 | 108.5698 | 110.0309 | 127.7063 | 20.9948 | 0.001844 | 0.001751 | -152.1284 | 0.23954 | 8.3999 | 0.003316 | 0.004959 | 0.004702 | 5.3e-05 | 0.001196 | 0.31868 | 1.3326 | 2.548 | 100.99 | 20.73 | 17.756 | 65.9352 | 101.3798 | 63.711 | 0.004305 | 1.048 | 0.025306 | 24.1459 | -0.11957 | 0.1425 | -4.7e-05 | -0.003635 | -0.003635 | -1.6115
2019-03-20 12:49 | 1553100535.17 | -48 | 0.20951 | 0.000446 | 0 | 108.5992 | 109.8994 | 127.6901 | 20.9952 | 0.001843 | 0.00179 | -152.1304 | 0.23954 | 8.4013 | 0.003389 | 0.004604 | 0.00439 | 5.8e-05 | 0.001197 | 0.32101 | 1.3118 | 2.5482 | 100.99 | 20.73 | 17.756 | 65.9143 | 103.1443 | 64.6304 | 0.004305 | 1.0452 | 0.02527 | 24.1773 | -0.10645 | 0.1425 | 2.4e-05 | 0.001794 | 0.001794 | 3.2639
2019-03-20 12:49 | 1553100536.169 | -47 | 0.20947 | 0.000444 | 0 | 108.6387 | 109.7346 | 127.4951 | 20.9956 | 0.001876 | 0.001694 | -152.1182 | 0.23974 | 8.4012 | 0.003347 | 0.004529 | 0.004338 | 5.6e-05 | 0.001197 | 0.30541 | 1.2969 | 2.548 | 100.99 | 20.73 | 17.756 | 66.713 | 102.1266 | 62.3192 | 0.004305 | 1.0524 | 0.025337 | 24.076 | -0.089774 | 0.14262 | 0.000219 | 0.016787 | 0.016787 | 0.34918
2019-03-20 12:49 | 1553100537.1681 | -46 | 0.20957 | 0.000442 | 0 | 108.6574 | 110.0545 | 127.6009 | 20.995 | 0.001902 | 0.001773 | -152.1216 | 0.23975 | 8.4015 | 0.003432 | 0.004911 | 0.004614 | 5.7e-05 | 0.001199 | 0.32781 | 1.2919 | 2.548 | 100.99 | 20.74 | 17.759 | 67.3343 | 104.1762 | 64.2383 | 0.004309 | 1.0464 | 0.025276 | 24.1548 | -0.11427 | 0.14263 | -0.000377 | -0.02879 | -0.02879 | -0.20361
2019-03-20 12:49 | 1553100538.1685 | -45 | 0.20953 | 0.000444 | 0 | 108.6135 | 109.8631 | 127.8281 | 20.9951 | 0.00184 | 0.001749 | -152.1069 | 0.2384 | 8.4043 | 0.003367 | 0.005211 | 0.005117 | 5.9e-05 | 0.00119 | 0.30326 | 1.3392 | 2.548 | 100.99 | 20.74 | 17.76 | 65.8402 | 102.5964 | 63.6552 | 0.004309 | 1.0482 | 0.02532 | 24.1554 | -0.10232 | 0.14182 | -0.000122 | -0.009365 | -0.009365 | -0.6224
2019-03-20 12:49 | 1553100539.1683 | -44 | 0.20954 | 0.00044 | 1e-06 | 108.5947 | 109.9483 | 127.9213 | 20.9949 | 0.001833 | 0.001718 | -152.1213 | 0.23961 | 8.4016 | 0.003305 | 0.00485 | 0.00461 | 6.1e-05 | 0.001194 | 0.31581 | 1.3191 | 2.548 | 100.99 | 20.74 | 17.761 | 65.6769 | 101.0964 | 62.907 | 0.004309 | 1.0506 | 0.025358 | 24.1373 | -0.11081 | 0.14255 | -0.000192 | -0.014759 | -0.014759 | -0.39695
2019-03-20 12:49 | 1553100540.1693 | -43 | 0.20954 | 0.000443 | 0 | 108.5672 | 109.9072 | 127.8807 | 20.9957 | 0.001862 | 0.001694 | -152.1262 | 0.23857 | 8.4021 | 0.003284 | 0.004236 | 0.004043 | 6.2e-05 | 0.001201 | 0.33251 | 1.3201 | 2.548 | 100.99 | 20.74 | 17.76 | 66.3783 | 100.6082 | 62.3404 | 0.004309 | 1.0523 | 0.025375 | 24.1131 | -0.10973 | 0.14193 | -0.000151 | -0.011584 | -0.011584 | -0.50357
2019-03-20 12:49 | 1553100541.1697 | -42 | 0.20954 | 0.000443 | 0 | 108.6154 | 109.9644 | 127.8544 | 20.9962 | 0.001842 | 0.001736 | -152.1355 | 0.23936 | 8.4008 | 0.003313 | 0.004813 | 0.004586 | 6.1e-05 | 0.001201 | 0.33284 | 1.3145 | 2.5481 | 100.99 | 20.74 | 17.759 | 65.8851 | 101.3113 | 63.3345 | 0.004309 | 1.0492 | 0.025335 | 24.1464 | -0.11041 | 0.1424 | -0.000189 | -0.014442 | -0.014442 | -0.40524
2019-03-20 12:49 | 1553100542.1694 | -41 | 0.20955 | 0.000442 | 0 | 108.6629 | 109.9801 | 127.3486 | 20.9952 | 0.001903 | 0.001738 | -152.1318 | 0.23967 | 8.401 | 0.003295 | 0.004821 | 0.004587 | 6.4e-05 | 0.001201 | 0.3569 | 1.3205 | 2.5481 | 100.99 | 20.74 | 17.755 | 67.3525 | 100.8567 | 63.3836 | 0.004308 | 1.0491 | 0.025283 | 24.1003 | -0.10777 | 0.14258 | -0.000219 | -0.016702 | -0.016702 | -0.35086
2019-03-20 12:49 | 1553100543.1684 | -40 | 0.20954 | 0.000441 | 0 | 108.5999 | 109.8257 | 127.5277 | 20.9959 | 0.001821 | 0.001763 | -152.1347 | 0.23927 | 8.405 | 0.003261 | 0.004777 | 0.004568 | 6.5e-05 | 0.001204 | 0.30807 | 1.3455 | 2.5481 | 100.99 | 20.74 | 17.752 | 65.382 | 100.0403 | 63.987 | 0.004307 | 1.0472 | 0.025278 | 24.1389 | -0.1004 | 0.14234 | -0.000198 | -0.015116 | -0.015116 | -0.38703
2019-03-20 12:49 | 1553100544.1685 | -39 | 0.20954 | 0.000444 | 0 | 108.5918 | 109.9311 | 127.3155 | 20.9963 | 0.001832 | 0.00177 | -152.1407 | 0.23908 | 8.4023 | 0.003299 | 0.004758 | 0.004563 | 6.2e-05 | 0.001204 | 0.31526 | 1.3184 | 2.5481 | 100.99 | 20.74 | 17.748 | 65.6415 | 100.9574 | 64.1585 | 0.004306 | 1.0467 | 0.02525 | 24.1249 | -0.10965 | 0.14223 | -0.000156 | -0.011896 | -0.011896 | -0.4914
2019-03-20 12:49 | 1553100545.1685 | -38 | 0.20956 | 0.000445 | 0 | 108.5834 | 109.9827 | 127.5821 | 20.9952 | 0.001823 | 0.001757 | -152.129 | 0.23902 | 8.4027 | 0.003214 | 0.004745 | 0.00448 | 6.1e-05 | 0.001205 | 0.33323 | 1.3281 | 2.5481 | 100.99 | 20.74 | 17.739 | 65.4266 | 98.9156 | 63.8342 | 0.004304 | 1.0477 | 0.025289 | 24.1385 | -0.11453 | 0.14219 | -0.00031 | -0.02365 | -0.02365 | -0.24711
2019-03-20 12:49 | 1553100546.1686 | -37 | 0.20955 | 0.000443 | 1e-06 | 108.6678 | 110.1034 | 127.6307 | 20.9957 | 0.001762 | 0.001733 | -152.118 | 0.23941 | 8.4032 | 0.003433 | 0.004959 | 0.004808 | 6.5e-05 | 0.0012 | 0.30127 | 1.3176 | 2.5481 | 100.99 | 20.74 | 17.734 | 63.9682 | 104.2105 | 63.2763 | 0.004303 | 1.0494 | 0.025315 | 24.1231 | -0.1174 | 0.14242 | -0.000234 | -0.017926 | -0.017926 | -0.32655
2019-03-20 12:49 | 1553100547.1686 | -36 | 0.20952 | 0.000442 | 0 | 108.5979 | 109.8598 | 127.8373 | 20.9949 | 0.001907 | 0.001746 | -152.1016 | 0.24034 | 8.4005 | 0.003314 | 0.004097 | 0.003822 | 6.1e-05 | 0.001206 | 0.35745 | 1.3394 | 2.5481 | 100.99 | 20.73 | 17.725 | 67.4621 | 101.317 | 63.5809 | 0.004298 | 1.0484 | 0.025324 | 24.1536 | -0.10333 | 0.14298 | -8.2e-05 | -0.006276 | -0.006276 | -0.9363
2019-03-20 12:49 | 1553100548.1684 | -35 | 0.20957 | 0.000448 | 0 | 108.6089 | 110.0919 | 127.7542 | 20.995 | 0.001865 | 0.001758 | -152.1044 | 0.23898 | 8.4049 | 0.003342 | 0.005176 | 0.005018 | 6.5e-05 | 0.0012 | 0.33373 | 1.3681 | 2.548 | 100.99 | 20.73 | 17.721 | 66.4534 | 101.9986 | 63.8706 | 0.004297 | 1.0475 | 0.025305 | 24.1561 | -0.1213 | 0.14217 | -0.000357 | -0.027302 | -0.027302 | -0.21402
2019-03-20 12:49 | 1553100549.1687 | -34 | 0.20954 | 0.000445 | 0 | 108.5643 | 110.1019 | 127.7326 | 20.9967 | 0.001873 | 0.001745 | -152.0887 | 0.23931 | 8.4029 | 0.003396 | 0.004208 | 0.00399 | 6.4e-05 | 0.001193 | 0.29966 | 1.3613 | 2.5483 | 100.99 | 20.73 | 17.717 | 66.6456 | 103.3185 | 63.563 | 0.004296 | 1.0485 | 0.025314 | 24.143 | -0.1258 | 0.14236 | -0.000154 | -0.011763 | -0.011763 | -0.49741
2019-03-20 12:49 | 1553100550.1698 | -33 | 0.20952 | 0.000442 | 1e-06 | 108.638 | 109.9495 | 127.5425 | 20.9959 | 0.001858 | 0.001771 | -152.0922 | 0.23898 | 8.404 | 0.003349 | 0.004736 | 0.00456 | 6.3e-05 | 0.001195 | 0.32709 | 1.3219 | 2.5481 | 100.99 | 20.73 | 17.71 | 66.2763 | 102.1708 | 64.1931 | 0.004294 | 1.0465 | 0.025272 | 24.1476 | -0.10733 | 0.14217 | -2.3e-05 | -0.001763 | -0.001763 | -3.3145
2019-03-20 12:49 | 1553100551.1682 | -32 | 0.20951 | 0.000445 | 0 | 108.6018 | 109.9693 | 127.4131 | 20.9965 | 0.001847 | 0.001738 | -152.0952 | 0.23868 | 8.4017 | 0.003377 | 0.004876 | 0.004662 | 6.7e-05 | 0.001195 | 0.32621 | 1.3019 | 2.548 | 100.99 | 20.73 | 17.707 | 66.007 | 102.8617 | 63.387 | 0.004293 | 1.0491 | 0.025289 | 24.1065 | -0.11192 | 0.14199 | 1.7e-05 | 0.001313 | 0.001313 | 4.4443
2019-03-20 12:49 | 1553100552.1679 | -31 | 0.20953 | 0.000444 | 1e-06 | 108.6736 | 109.9588 | 127.5326 | 20.996 | 0.001906 | 0.001787 | -152.0892 | 0.23922 | 8.4022 | 0.003371 | 0.004625 | 0.004427 | 6.5e-05 | 0.001195 | 0.30596 | 1.2969 | 2.5481 | 100.99 | 20.73 | 17.704 | 67.436 | 102.6939 | 64.5703 | 0.004293 | 1.0454 | 0.025257 | 24.1602 | -0.10516 | 0.14231 | -0.000142 | -0.01086 | -0.01086 | -0.53859
2019-03-20 12:49 | 1553100553.1681 | -30 | 0.20955 | 0.000444 | 0 | 108.6397 | 109.924 | 127.5566 | 20.9954 | 0.001849 | 0.001761 | -152.0722 | 0.23917 | 8.4049 | 0.003325 | 0.004286 | 0.004086 | 6.4e-05 | 0.001192 | 0.34075 | 1.3532 | 2.5481 | 100.99 | 20.73 | 17.703 | 66.0557 | 101.5875 | 63.9436 | 0.004293 | 1.0473 | 0.025282 | 24.14 | -0.10513 | 0.14228 | -0.000229 | -0.017433 | -0.017433 | -0.33545
2019-03-20 12:49 | 1553100554.1694 | -29 | 0.20955 | 0.000446 | 0 | 108.6238 | 109.9311 | 127.5033 | 20.9953 | 0.001889 | 0.001775 | -152.0879 | 0.23921 | 8.4028 | 0.003372 | 0.004852 | 0.00456 | 6.5e-05 | 0.00119 | 0.34235 | 1.3351 | 2.5481 | 100.99 | 20.73 | 17.703 | 67.0155 | 102.7177 | 64.281 | 0.004293 | 1.0463 | 0.025264 | 24.1471 | -0.10701 | 0.14231 | -0.000235 | -0.017884 | -0.017884 | -0.32705
2019-03-20 12:49 | 1553100555.169 | -28 | 0.20954 | 0.000443 | 0 | 108.6177 | 109.9543 | 127.1866 | 20.9954 | 0.001841 | 0.001775 | -152.1038 | 0.23939 | 8.4033 | 0.003363 | 0.005253 | 0.005073 | 6.7e-05 | 0.001195 | 0.35043 | 1.3293 | 2.5481 | 100.99 | 20.73 | 17.709 | 65.8725 | 102.519 | 64.2748 | 0.004294 | 1.0463 | 0.025233 | 24.1168 | -0.1094 | 0.14241 | -0.000147 | -0.011181 | -0.011181 | -0.52348
2019-03-20 12:49 | 1553100556.1681 | -27 | 0.20954 | 0.000445 | 0 | 108.5442 | 109.8841 | 127.1763 | 20.9958 | 0.001862 | 0.001785 | -152.1066 | 0.23933 | 8.4007 | 0.003279 | 0.004586 | 0.004326 | 6.6e-05 | 0.001191 | 0.31078 | 1.3395 | 2.5482 | 100.99 | 20.73 | 17.714 | 66.3825 | 100.4741 | 64.518 | 0.004295 | 1.0455 | 0.025223 | 24.1246 | -0.10973 | 0.14238 | -0.000153 | -0.011676 | -0.011676 | -0.50119
2019-03-20 12:49 | 1553100557.1695 | -26 | 0.20952 | 0.000446 | 0 | 108.5462 | 109.8302 | 127.0942 | 20.996 | 0.001845 | 0.001758 | -152.1034 | 0.23956 | 8.4015 | 0.003425 | 0.005117 | 0.004906 | 6.3e-05 | 0.001197 | 0.35618 | 1.3793 | 2.5481 | 100.99 | 20.73 | 17.726 | 65.9645 | 104.0097 | 63.8785 | 0.004298 | 1.0475 | 0.025239 | 24.0939 | -0.10519 | 0.14252 | -3.5e-05 | -0.002617 | -0.002617 | -2.2385
2019-03-20 12:49 | 1553100558.1692 | -25 | 0.20952 | 0.000443 | -1e-06 | 108.5106 | 109.9502 | 127.2284 | 20.9957 | 0.001854 | 0.001788 | -152.1021 | 0.23965 | 8.4024 | 0.003498 | 0.005262 | 0.00503 | 6e-05 | 0.001198 | 0.29695 | 1.378 | 2.5481 | 100.99 | 20.73 | 17.731 | 66.1756 | 105.778 | 64.585 | 0.004299 | 1.0453 | 0.025226 | 24.1319 | -0.11789 | 0.14257 | -8.1e-05 | -0.006158 | -0.006158 | -0.95158
2019-03-20 12:49 | 1553100559.1681 | -24 | 0.2095 | 0.000443 | 0 | 108.6044 | 109.7716 | 127.3056 | 20.9957 | 0.001908 | 0.001767 | -152.0989 | 0.23891 | 8.4034 | 0.003425 | 0.004648 | 0.004406 | 6.4e-05 | 0.001208 | 0.31598 | 1.3363 | 2.5482 | 100.99 | 20.73 | 17.737 | 67.4746 | 104.0147 | 64.0813 | 0.004301 | 1.0469 | 0.025252 | 24.1212 | -0.095613 | 0.14213 | 4.1e-05 | 0.003143 | 0.003143 | 1.8585
2019-03-20 12:49 | 1553100560.1687 | -23 | 0.20955 | 0.000441 | 1e-06 | 108.5818 | 109.8784 | 127.2843 | 20.9964 | 0.001914 | 0.00178 | -152.1098 | 0.23928 | 8.4015 | 0.003423 | 0.004587 | 0.0044 | 6.6e-05 | 0.00121 | 0.29175 | 1.3167 | 2.5481 | 100.99 | 20.73 | 17.748 | 67.6244 | 103.959 | 64.4122 | 0.004303 | 1.0459 | 0.025238 | 24.131 | -0.10618 | 0.14235 | -0.000222 | -0.016918 | -0.016918 | -0.3458
2019-03-20 12:49 | 1553100561.1693 | -22 | 0.20952 | 0.000442 | 0 | 108.5261 | 109.9072 | 127.3284 | 20.9961 | 0.001875 | 0.001757 | -152.1013 | 0.23874 | 8.4036 | 0.003419 | 0.004686 | 0.004513 | 6.7e-05 | 0.001216 | 0.33179 | 1.3177 | 2.5481 | 100.99 | 20.73 | 17.753 | 66.6895 | 103.8705 | 63.8584 | 0.004305 | 1.0476 | 0.025263 | 24.1154 | -0.11311 | 0.14202 | -2.8e-05 | -0.002182 | -0.002182 | -2.6753
2019-03-20 12:49 | 1553100562.1695 | -21 | 0.20954 | 0.000441 | 0 | 108.5814 | 109.9853 | 127.6211 | 20.9959 | 0.001866 | 0.001739 | -152.1039 | 0.23905 | 8.4022 | 0.003435 | 0.004814 | 0.00454 | 6.4e-05 | 0.001215 | 0.3147 | 1.3147 | 2.5478 | 100.99 | 20.73 | 17.758 | 66.4636 | 104.2562 | 63.4177 | 0.004306 | 1.049 | 0.025309 | 24.1273 | -0.11491 | 0.14221 | -0.00014 | -0.010725 | -0.010725 | -0.54496
2019-03-20 12:49 | 1553100563.1701 | -20 | 0.20952 | 0.000441 | 0 | 108.5345 | 109.9495 | 127.7758 | 20.9959 | 0.001883 | 0.001798 | -152.1007 | 0.23911 | 8.4042 | 0.003346 | 0.005356 | 0.005193 | 6.3e-05 | 0.001219 | 0.31277 | 1.3446 | 2.5478 | 100.99 | 20.73 | 17.768 | 66.8749 | 102.1041 | 64.8405 | 0.004308 | 1.0445 | 0.025271 | 24.1929 | -0.11585 | 0.14225 | -6.3e-05 | -0.004804 | -0.004804 | -1.2169
2019-03-20 12:49 | 1553100564.1689 | -19 | 0.20954 | 0.000446 | 0 | 108.5539 | 109.6374 | 128.0638 | 20.9963 | 0.001896 | 0.001752 | -152.0847 | 0.23902 | 8.4044 | 0.003377 | 0.004779 | 0.004582 | 6.1e-05 | 0.001221 | 0.34523 | 1.3501 | 2.548 | 100.99 | 20.73 | 17.772 | 67.1895 | 102.8397 | 63.7218 | 0.004309 | 1.048 | 0.025341 | 24.18 | -0.088829 | 0.14219 | -0.00017 | -0.013021 | -0.013021 | -0.44881
2019-03-20 12:49 | 1553100565.1696 | -18 | 0.20958 | 0.000445 | 0 | 108.5255 | 109.9513 | 128.4969 | 20.996 | 0.001867 | 0.001699 | -152.0766 | 0.23902 | 8.4044 | 0.003414 | 0.004011 | 0.0038 | 6.6e-05 | 0.001216 | 0.3314 | 1.307 | 2.548 | 100.99 | 20.74 | 17.779 | 66.4974 | 103.7407 | 62.4532 | 0.004314 | 1.052 | 0.025432 | 24.1752 | -0.11675 | 0.14219 | -0.000408 | -0.031327 | -0.031327 | -0.18656
2019-03-20 12:49 | 1553100566.1696 | -17 | 0.20954 | 0.000447 | 0 | 108.4184 | 109.831 | 128.8588 | 20.9955 | 0.00192 | 0.001769 | -152.0832 | 0.23911 | 8.4002 | 0.003299 | 0.004557 | 0.004312 | 6.1e-05 | 0.00122 | 0.3585 | 1.3476 | 2.5482 | 100.99 | 20.74 | 17.782 | 67.7685 | 100.9517 | 64.1406 | 0.004314 | 1.0467 | 0.025404 | 24.27 | -0.11578 | 0.14225 | -0.000157 | -0.012054 | -0.012054 | -0.48501
2019-03-20 12:49 | 1553100567.1691 | -16 | 0.20952 | 0.000443 | 0 | 108.5051 | 109.9947 | 129.1916 | 20.9954 | 0.001797 | 0.001751 | -152.0986 | 0.24021 | 8.4028 | 0.003464 | 0.004937 | 0.004819 | 6.4e-05 | 0.001215 | 0.32333 | 1.3479 | 2.548 | 100.99 | 20.74 | 17.784 | 64.808 | 104.9676 | 63.6928 | 0.004315 | 1.0481 | 0.025453 | 24.2852 | -0.12196 | 0.1429 | -2.9e-05 | -0.002182 | -0.002182 | -2.6922
2019-03-20 12:49 | 1553100568.1687 | -15 | 0.20958 | 0.000444 | 0 | 108.5074 | 109.8941 | 129.2531 | 20.9955 | 0.001928 | 0.001727 | -152.0956 | 0.23937 | 8.3998 | 0.003598 | 0.004912 | 0.00471 | 6.1e-05 | 0.001211 | 0.29966 | 1.3551 | 2.5482 | 100.99 | 20.74 | 17.784 | 67.9635 | 108.2054 | 63.1261 | 0.004315 | 1.0499 | 0.025481 | 24.2706 | -0.11359 | 0.1424 | -0.000377 | -0.02899 | -0.02899 | -0.20189
2019-03-20 12:49 | 1553100569.1689 | -14 | 0.20954 | 0.000445 | -1e-06 | 108.4851 | 109.7267 | 129.3398 | 20.9957 | 0.001853 | 0.001763 | -152.1141 | 0.24013 | 8.4017 | 0.003583 | 0.005184 | 0.004964 | 6.5e-05 | 0.001217 | 0.3194 | 1.3838 | 2.548 | 100.99 | 20.74 | 17.783 | 66.1677 | 107.8487 | 63.9993 | 0.004315 | 1.0471 | 0.025456 | 24.3102 | -0.10179 | 0.14286 | -0.000188 | -0.01445 | -0.01445 | -0.40634
2019-03-20 12:49 | 1553100570.1679 | -13 | 0.20954 | 0.000444 | 0 | 108.4475 | 109.7779 | 129.4658 | 20.9947 | 0.001851 | 0.001685 | -152.1125 | 0.23916 | 8.4046 | 0.003542 | 0.004783 | 0.0046 | 6.6e-05 | 0.001212 | 0.32261 | 1.352 | 2.5481 | 100.99 | 20.74 | 17.779 | 66.0957 | 106.8444 | 62.109 | 0.004314 | 1.0531 | 0.02554 | 24.2538 | -0.10906 | 0.14228 | -0.000177 | -0.013593 | -0.013593 | -0.43018
2019-03-20 12:50 | 1553100571.1679 | -12 | 0.20953 | 0.000445 | 0 | 108.4731 | 109.7203 | 129.2416 | 20.9957 | 0.001878 | 0.001736 | -132.1034 | 0.23885 | 8.3999 | 0.003586 | 0.004622 | 0.004356 | 5.8e-05 | 0.001211 | 0.32565 | 1.347 | 2.5481 | 100.99 | 20.74 | 17.777 | 66.7491 | 107.9212 | 63.3378 | 0.004313 | 1.0492 | 0.025472 | 24.2771 | -0.10226 | 0.14209 | -0.000103 | -0.007932 | -0.007932 | -0.73629
2019-03-20 12:50 | 1553100572.1689 | -11 | 0.20956 | 0.000443 | 0 | 108.4207 | 109.8803 | 129.3386 | 20.9965 | 0.001793 | 0.001695 | -99.5575 | 0.23937 | 8.4022 | 0.003601 | 0.005409 | 0.005266 | 6.2e-05 | 0.00121 | 0.33395 | 1.3382 | 2.548 | 100.99 | 20.74 | 17.776 | 64.7152 | 108.2734 | 62.3631 | 0.004313 | 1.0523 | 0.025518 | 24.251 | -0.11961 | 0.1424 | -0.000275 | -0.021137 | -0.021137 | -0.27689
2019-03-20 12:50 | 1553100573.169 | -10 | 0.20956 | 0.000444 | 0 | 108.4627 | 109.8142 | 129.2881 | 20.9949 | 0.001823 | 0.001714 | -66.9942 | 0.23889 | 8.4053 | 0.003632 | 0.00462 | 0.004414 | 6.2e-05 | 0.001206 | 0.34346 | 1.3199 | 2.5481 | 100.99 | 20.74 | 17.776 | 65.4342 | 109.0433 | 62.8036 | 0.004313 | 1.0509 | 0.025497 | 24.2622 | -0.11076 | 0.14211 | -0.000287 | -0.022096 | -0.022096 | -0.26433
2019-03-20 12:50 | 1553100574.1694 | -9 | 0.20952 | 0.000443 | 0 | 108.4792 | 109.8538 | 129.508 | 20.9951 | 0.001838 | 0.001701 | -34.4135 | 0.24019 | 8.4025 | 0.003695 | 0.005112 | 0.004984 | 6.3e-05 | 0.001204 | 0.36105 | 1.3347 | 2.548 | 100.99 | 20.74 | 17.776 | 65.7932 | 110.5732 | 62.4975 | 0.004313 | 1.0518 | 0.02553 | 24.2718 | -0.11262 | 0.14289 | -4.7e-05 | -0.003654 | -0.003654 | -1.6071
2019-03-20 12:50 | 1553100575.1692 | -8 | 0.20955 | 0.000439 | 1e-06 | 108.423 | 109.8538 | 129.4862 | 20.996 | 0.001861 | 0.001671 | -1.834 | 0.23989 | 8.4009 | 0.003589 | 0.005003 | 0.004776 | 6.1e-05 | 0.001197 | 0.29806 | 1.3124 | 2.5481 | 100.99 | 20.74 | 17.776 | 66.339 | 107.9988 | 61.765 | 0.004313 | 1.0541 | 0.025556 | 24.2432 | -0.11727 | 0.14271 | -0.000226 | -0.01745 | -0.01745 | -0.33614
2019-03-20 12:50 | 1553100576.1692 | -7 | 0.20954 | 0.000439 | 0 | 108.4537 | 109.9267 | 129.2071 | 20.995 | 0.001838 | 0.001664 | 30.7612 | 0.25495 | 8.4043 | 0.00347 | 0.004746 | 0.004596 | 6.3e-05 | 0.001197 | 0.32781 | 1.2752 | 2.548 | 100.99 | 20.74 | 17.779 | 65.7963 | 105.1008 | 61.599 | 0.004314 | 1.0547 | 0.025534 | 24.2111 | -0.12066 | 0.15167 | -0.000175 | -0.01348 | -0.01348 | -0.46242
2019-03-20 12:50 | 1553100577.1682 | -6 | 0.20954 | 0.000449 | 0 | 108.4773 | 109.9397 | 128.6318 | 20.9945 | 0.001816 | 0.001719 | 63.3732 | 0.26444 | 8.4032 | 0.003579 | 0.00503 | 0.004905 | 6e-05 | 0.001194 | 0.31708 | 1.2793 | 2.548 | 100.99 | 20.74 | 17.783 | 65.2582 | 107.7467 | 62.9306 | 0.004315 | 1.0505 | 0.025427 | 24.2051 | -0.11978 | 0.15732 | -0.000164 | -0.012572 | -0.012572 | -0.51431
2019-03-20 12:50 | 1553100578.1683 | -5 | 0.20953 | 0.000445 | 0 | 108.4427 | 109.6351 | 128.2762 | 20.995 | 0.001836 | 0.001693 | 173.6436 | 0.26492 | 8.4022 | 0.003565 | 0.005076 | 0.005011 | 6.1e-05 | 0.001187 | 0.33555 | 1.3122 | 2.5479 | 100.99 | 20.74 | 17.786 | 65.7362 | 107.4054 | 62.3116 | 0.004315 | 1.0524 | 0.025415 | 24.1494 | -0.097818 | 0.1576 | -0.000103 | -0.007874 | -0.007874 | -0.82261
2019-03-20 12:50 | 1553100579.1683 | -4 | 0.20952 | 0.000444 | 0 | 108.4676 | 109.7312 | 127.9582 | 20.9948 | 0.001861 | 0.001901 | 173.7469 | 0.26459 | 8.4031 | 0.003639 | 0.005295 | 0.005119 | 6.1e-05 | 0.001178 | 0.32278 | 1.343 | 2.5481 | 100.99 | 20.74 | 17.789 | 66.3549 | 109.1931 | 67.3122 | 0.004316 | 1.037 | 0.025197 | 24.2985 | -0.1036 | 0.15741 | -4.7e-05 | -0.003592 | -0.003592 | -1.8012
2019-03-20 12:50 | 1553100580.1698 | -3 | 0.20953 | 0.000436 | 0 | 108.4113 | 109.7858 | 127.8948 | 20.9948 | 0.001796 | 0.001727 | 173.694 | 0.26404 | 8.4035 | 0.003634 | 0.005059 | 0.004997 | 6.2e-05 | 0.001171 | 0.335 | 1.3492 | 2.5481 | 100.99 | 20.74 | 17.792 | 64.7907 | 109.0888 | 63.1184 | 0.004317 | 1.0499 | 0.025347 | 24.1424 | -0.11269 | 0.15708 | -0.000106 | -0.00809 | -0.00809 | -0.79804
2019-03-20 12:50 | 1553100581.1695 | -2 | 0.20953 | 0.000448 | 0 | 108.4071 | 109.7334 | 128.1154 | 20.9944 | 0.001794 | 0.001686 | 173.7722 | 0.26539 | 8.4037 | 0.003552 | 0.005037 | 0.004923 | 6.1e-05 | 0.001171 | 0.31686 | 1.3584 | 2.5481 | 100.99 | 20.74 | 17.792 | 64.7382 | 107.081 | 62.1432 | 0.004317 | 1.0529 | 0.025406 | 24.1282 | -0.10877 | 0.15788 | -0.000123 | -0.009428 | -0.009428 | -0.68823
2019-03-20 12:50 | 1553100582.1685 | -1 | 0.20952 | 0.000444 | 0 | 108.4071 | 109.72 | 127.864 | 20.9948 | 0.001836 | 0.001669 | 173.9135 | 0.26338 | 8.4016 | 0.003545 | 0.005199 | 0.005053 | 5.7e-05 | 0.001168 | 0.32676 | 1.3214 | 2.5481 | 100.99 | 20.74 | 17.807 | 65.7533 | 106.9327 | 61.7325 | 0.00432 | 1.0542 | 0.025396 | 24.0897 | -0.10767 | 0.15669 | -2.1e-05 | -0.001577 | -0.001577 | -4.0845
2019-03-20 12:50 | 1553100583.1675 | 0 | 0.20955 | 0.000441 | 0 | 108.3201 | 109.551 | 127.3951 | 20.9836 | 0.001838 | 0.001706 | 173.8562 | 0.2658 | 8.4022 | 0.003597 | 0.00612 | 0.007107 | 5.8e-05 | 0.001165 | 0.3632 | 1.31 | 2.683 | 100.99 | 20.74 | 17.812 | 65.7838 | 108.1614 | 62.6029 | 0.004322 | 1.0515 | 0.025317 | 24.0767 | -0.10107 | 0.15813 | -0.000219 | -0.016754 | -0.016754 | -0.38791
2019-03-20 12:50 | 1553100584.1676 | 1 | 0.20954 | 0.000442 | 0 | 108.2292 | 109.4886 | 126.2637 | 20.9746 | 0.001818 | 0.001739 | 173.8136 | 0.26464 | 8.4005 | 0.003596 | 0.00533 | 0.005128 | 5.9e-05 | 0.001162 | 0.31437 | 1.3333 | 2.5481 | 100.99 | 20.74 | 17.818 | 65.2841 | 108.1448 | 63.3845 | 0.004323 | 1.0491 | 0.025175 | 23.9975 | -0.10348 | 0.15743 | -0.000144 | -0.010958 | -0.010958 | -0.59047
2019-03-20 12:50 | 1553100585.1676 | 2 | 0.20953 | 0.000446 | 1e-06 | 108.3143 | 109.8336 | 125.003 | 20.9892 | 0.001773 | 0.001734 | 174.0513 | 0.26459 | 8.4014 | 0.003509 | 0.008965 | 0.018201 | 6.2e-05 | 0.001162 | 0.39495 | 1.3289 | 2.5482 | 100.99 | 20.74 | 17.827 | 64.2186 | 106.0433 | 63.2824 | 0.004325 | 1.0494 | 0.025053 | 23.8737 | -0.12459 | 0.1574 | -0.000132 | -0.010001 | -0.010001 | -0.64686
2019-03-20 12:50 | 1553100586.1687 | 3 | 0.20957 | 0.00045 | 2e-06 | 108.3001 | 109.5873 | 123.505 | 20.984 | 0.001765 | 0.001773 | 174.1117 | 0.26368 | 8.4055 | 0.00351 | 0.005293 | 0.00523 | 5.9e-05 | 0.001152 | 0.33052 | 1.3201 | 2.5485 | 100.99 | 20.74 | 17.832 | 64.0199 | 106.0705 | 64.2214 | 0.004326 | 1.0465 | 0.024867 | 23.7633 | -0.10569 | 0.15687 | -0.000345 | -0.025973 | -0.025973 | -0.24822
2019-03-20 12:50 | 1553100587.1678 | 4 | 0.20933 | 0.000445 | 5e-06 | 108.2596 | 109.6901 | 121.8309 | 20.9829 | 0.001854 | 0.00184 | 174.1046 | 0.26473 | 8.4012 | 0.003473 | 0.005385 | 0.005244 | 6.2e-05 | 0.001155 | 0.35778 | 1.3577 | 2.548 | 100.99 | 20.74 | 17.839 | 66.174 | 105.171 | 65.8269 | 0.004328 | 1.0415 | 0.02464 | 23.6578 | -0.11741 | 0.15749 | 0.00113 | 0.083699 | 0.083699 | 0.077334
2019-03-20 12:50 | 1553100588.1678 | 5 | 0.20949 | 0.000445 | 1e-05 | 108.3567 | 108.9915 | 119.8339 | 20.9801 | 0.001958 | 0.001889 | 173.9765 | 0.26428 | 8.4018 | 0.003518 | 0.005324 | 0.005173 | 6.4e-05 | 0.001158 | 0.31222 | 1.3182 | 2.5481 | 100.99 | 20.74 | 17.842 | 68.6724 | 106.243 | 66.9968 | 0.004329 | 1.0379 | 0.024395 | 23.5036 | -0.052254 | 0.15722 | 0.000154 | 0.010729 | 0.010729 | 0.6023
2019-03-20 12:50 | 1553100589.1678 | 6 | 0.20907 | 0.000458 | 2e-05 | 108.4637 | 108.0138 | 117.8193 | 20.9806 | 0.001955 | 0.001957 | 173.8697 | 0.26334 | 8.4018 | 0.003536 | 0.005047 | 0.004949 | 5.8e-05 | 0.001158 | 0.36425 | 1.3406 | 2.343 | 100.99 | 20.74 | 17.843 | 68.5832 | 106.6999 | 68.6486 | 0.004329 | 1.0329 | 0.02413 | 23.3617 | 0.03718 | 0.15666 | 0.002657 | 0.19228 | 0.19228 | 0.033487
2019-03-20 12:50 | 1553100590.1689 | 7 | 0.20818 | 0.000505 | 3.2e-05 | 108.3101 | 106.5513 | 115.7863 | 20.9909 | 0.001958 | 0.002059 | 173.8114 | 0.26387 | 8.4035 | 0.003552 | 0.005201 | 0.005071 | 5.7e-05 | 0.001161 | 0.32101 | 1.3199 | 2.0043 | 100.99 | 20.74 | 17.847 | 68.6735 | 107.0888 | 71.0978 | 0.00433 | 1.0256 | 0.023836 | 23.2421 | 0.14644 | 0.15697 | 0.007901 | 0.56614 | 0.56614 | 0.011396
2019-03-20 12:50 | 1553100591.1691 | 8 | 0.20622 | 0.000699 | 4.4e-05 | 108.3318 | 102.501 | 113.6581 | 20.9948 | 0.002084 | 0.002097 | 173.7166 | 0.26407 | 8.4055 | 0.00353 | 0.004713 | 0.004594 | 5.7e-05 | 0.001157 | 0.29878 | 1.3336 | 2.5481 | 100.99 | 20.74 | 17.848 | 71.716 | 106.55 | 72.0262 | 0.00433 | 1.0228 | 0.023584 | 23.0585 | 0.49487 | 0.15709 | 0.019463 | 1.3801 | 1.3801 | 0.004678
2019-03-20 12:50 | 1553100592.169 | 9 | 0.2047 | 0.001313 | 5.9e-05 | 108.3408 | 95.9484 | 111.7601 | 20.9947 | 0.002121 | 0.002145 | 173.6175 | 0.26425 | 8.4041 | 0.003545 | 0.005045 | 0.00497 | 5.8e-05 | 0.001155 | 0.31581 | 1.3133 | 2.548 | 100.99 | 20.74 | 17.851 | 72.5999 | 106.9231 | 73.1675 | 0.004331 | 1.0194 | 0.023348 | 22.903 | 1.0865 | 0.1572 | 0.027725 | 1.9448 | 1.9448 | 0.003322
2019-03-20 12:50 | 1553100593.1691 | 10 | 0.20323 | 0.002085 | 7.2e-05 | 108.3528 | 90.1941 | 109.8439 | 20.9941 | 0.002178 | 0.002231 | 173.528 | 0.26405 | 8.4009 | 0.003645 | 0.004651 | 0.004554 | 6.1e-05 | 0.001156 | 0.31188 | 1.3322 | 2.548 | 100.99 | 20.74 | 17.855 | 73.9635 | 109.3469 | 75.2496 | 0.004332 | 1.0133 | 0.023078 | 22.7739 | 1.6407 | 0.15708 | 0.035577 | 2.4649 | 2.4649 | 0.002619
2019-03-20 12:50 | 1553100594.1691 | 11 | 0.20217 | 0.003004 | 8.4e-05 | 108.3518 | 86.0847 | 108.9991 | 20.9942 | 0.002349 | 0.002293 | 173.3851 | 0.26323 | 8.4038 | 0.003614 | 0.005481 | 0.005391 | 5.8e-05 | 0.001152 | 0.28655 | 1.324 | 2.5479 | 100.99 | 20.74 | 17.856 | 78.0848 | 108.5847 | 76.7271 | 0.004332 | 1.0091 | 0.02294 | 22.7342 | 2.0577 | 0.1566 | 0.040777 | 2.8068 | 2.8068 | 0.002293
2019-03-20 12:50 | 1553100595.1692 | 12 | 0.2012 | 0.004033 | 9.6e-05 | 108.3091 | 82.876 | 108.0555 | 20.995 | 0.002402 | 0.002414 | 173.2649 | 0.26394 | 8.401 | 0.003722 | 0.005108 | 0.004992 | 5.8e-05 | 0.001158 | 0.29247 | 1.3101 | 2.548 | 100.99 | 20.73 | 17.856 | 79.3526 | 111.2299 | 79.6457 | 0.00433 | 1.0007 | 0.022746 | 22.7298 | 2.3939 | 0.15702 | 0.045255 | 3.087 | 3.087 | 0.00209
2019-03-20 12:50 | 1553100596.1692 | 13 | 0.20041 | 0.005068 | 0.000105 | 108.325 | 81.1908 | 107.6083 | 20.9953 | 0.002439 | 0.002475 | 173.161 | 0.2646 | 8.4028 | 0.003761 | 0.005839 | 0.00567 | 5.8e-05 | 0.001159 | 0.30718 | 1.3361 | 2.548 | 100.99 | 20.73 | 17.855 | 80.2362 | 112.1696 | 81.1142 | 0.004329 | 0.99656 | 0.022652 | 22.7299 | 2.579 | 0.15741 | 0.048734 | 3.3093 | 3.3093 | 0.001955
2019-03-20 12:50 | 1553100597.1683 | 14 | 0.19982 | 0.006028 | 0.000114 | 108.3136 | 80.0123 | 107.7893 | 20.9952 | 0.002503 | 0.002552 | 173.0302 | 0.26502 | 8.4022 | 0.003785 | 0.005308 | 0.005214 | 6.2e-05 | 0.001163 | 0.30917 | 1.3237 | 2.548 | 100.99 | 20.73 | 17.854 | 81.7714 | 112.7522 | 82.9539 | 0.004329 | 0.99141 | 0.022612 | 22.808 | 2.7089 | 0.15766 | 0.051052 | 3.4596 | 3.4596 | 0.001873
2019-03-20 12:50 | 1553100598.1673 | 15 | 0.19919 | 0.006622 | 0.000121 | 108.2871 | 78.5222 | 108.3036 | 20.9964 | 0.002616 | 0.00258 | 172.9169 | 0.26493 | 8.4014 | 0.003834 | 0.004638 | 0.004585 | 5.9e-05 | 0.001162 | 0.32333 | 1.3943 | 2.5482 | 100.99 | 20.73 | 17.853 | 84.4967 | 113.9443 | 83.618 | 0.004329 | 0.98957 | 0.022645 | 22.8837 | 2.8748 | 0.15761 | 0.054123 | 3.6718 | 3.6718 | 0.001764
2019-03-20 12:50 | 1553100599.1673 | 16 | 0.19895 | 0.007032 | 0.000126 | 108.214 | 77.6999 | 108.9162 | 20.9946 | 0.002698 | 0.002592 | 172.7774 | 0.26518 | 8.4042 | 0.003757 | 0.005207 | 0.005076 | 5.9e-05 | 0.001169 | 0.34036 | 1.3691 | 2.5479 | 100.99 | 20.73 | 17.853 | 86.4585 | 112.082 | 83.9058 | 0.004329 | 0.98877 | 0.0227 | 22.9576 | 2.9629 | 0.15776 | 0.054994 | 3.7395 | 3.7395 | 0.001734
2019-03-20 12:50 | 1553100600.1686 | 17 | 0.19893 | 0.007272 | 0.000129 | 108.2444 | 77.9814 | 109.6183 | 20.9948 | 0.002737 | 0.002671 | 172.647 | 0.26444 | 8.4014 | 0.003756 | 0.004995 | 0.00487 | 5.8e-05 | 0.001171 | 0.29983 | 1.3288 | 2.5481 | 100.99 | 20.73 | 17.855 | 87.4003 | 112.0466 | 85.8053 | 0.004329 | 0.98354 | 0.022712 | 23.0926 | 2.9331 | 0.15732 | 0.054871 | 3.733 | 3.733 | 0.001732
2019-03-20 12:50 | 1553100601.1685 | 18 | 0.19894 | 0.007507 | 0.00013 | 108.2626 | 78.4123 | 110.2096 | 20.9953 | 0.002787 | 0.002689 | 172.5504 | 0.26443 | 8.4045 | 0.00383 | 0.004893 | 0.004848 | 6.1e-05 | 0.001171 | 0.29966 | 1.3313 | 2.548 | 100.99 | 20.73 | 17.858 | 88.6152 | 113.8412 | 86.2422 | 0.00433 | 0.98234 | 0.02276 | 23.1689 | 2.8853 | 0.15731 | 0.054545 | 3.7186 | 3.7186 | 0.001739
2019-03-20 12:50 | 1553100602.1686 | 19 | 0.19906 | 0.007708 | 0.000129 | 108.2221 | 77.8808 | 110.567 | 20.995 | 0.002797 | 0.002706 | 172.4159 | 0.26364 | 8.4022 | 0.00379 | 0.005542 | 0.005439 | 5.7e-05 | 0.001166 | 0.31813 | 1.3299 | 2.5481 | 100.99 | 20.73 | 17.859 | 88.8373 | 112.8738 | 86.6485 | 0.00433 | 0.98123 | 0.022784 | 23.2196 | 2.9428 | 0.15684 | 0.053546 | 3.6546 | 3.6546 | 0.001764
2019-03-20 12:50 | 1553100603.1675 | 20 | 0.1993 | 0.00787 | 0.000127 | 108.2195 | 78.9599 | 111.2794 | 20.9947 | 0.002825 | 0.002715 | 172.2767 | 0.26457 | 8.4036 | 0.003841 | 0.004928 | 0.00483 | 5.9e-05 | 0.001166 | 0.35961 | 1.3354 | 2.5481 | 100.99 | 20.73 | 17.862 | 89.5237 | 114.1173 | 86.87 | 0.004331 | 0.98063 | 0.02285 | 23.3014 | 2.8195 | 0.15739 | 0.051967 | 3.5576 | 3.5576 | 0.001818
2019-03-20 12:50 | 1553100604.1679 | 21 | 0.19958 | 0.007976 | 0.000125 | 108.2001 | 79.8968 | 111.5456 | 20.9956 | 0.002909 | 0.002763 | 172.1768 | 0.26428 | 8.4037 | 0.003811 | 0.005032 | 0.004795 | 6e-05 | 0.001165 | 0.32084 | 1.3411 | 2.5482 | 100.99 | 20.73 | 17.862 | 91.5405 | 113.395 | 88.0389 | 0.004331 | 0.97745 | 0.02284 | 23.3671 | 2.7124 | 0.15722 | 0.050129 | 3.4308 | 3.4308 | 0.001883
2019-03-20 12:50 | 1553100605.1683 | 22 | 0.2002 | 0.008019 | 0.000123 | 108.1633 | 79.4241 | 112.15 | 20.9953 | 0.002859 | 0.002694 | 172.0665 | 0.26428 | 8.4018 | 0.003925 | 0.00488 | 0.004857 | 6e-05 | 0.001159 | 0.33157 | 1.3225 | 2.5482 | 100.99 | 20.73 | 17.862 | 90.3379 | 116.1601 | 86.3764 | 0.004331 | 0.98197 | 0.022955 | 23.3763 | 2.7624 | 0.15722 | 0.046386 | 3.1914 | 3.1914 | 0.002025
2019-03-20 12:50 | 1553100606.1687 | 23 | 0.20073 | 0.007998 | 0.000121 | 108.1616 | 80.4436 | 112.6385 | 20.9949 | 0.002889 | 0.002777 | 171.9483 | 0.26511 | 8.4021 | 0.003884 | 0.005654 | 0.005499 | 5.9e-05 | 0.001159 | 0.33754 | 1.3214 | 2.548 | 100.99 | 20.73 | 17.862 | 91.0668 | 115.1558 | 88.3562 | 0.004331 | 0.9766 | 0.022942 | 23.4916 | 2.6482 | 0.15771 | 0.043237 | 2.9738 | 2.9738 | 0.00218
2019-03-20 12:50 | 1553100607.1688 | 24 | 0.20086 | 0.007917 | 0.000117 | 108.1196 | 80.9359 | 113.0873 | 20.9939 | 0.002906 | 0.002776 | 171.8379 | 0.2639 | 8.4028 | 0.003939 | 0.005099 | 0.004977 | 5.9e-05 | 0.001158 | 0.32731 | 1.3213 | 2.5481 | 100.99 | 20.73 | 17.863 | 91.4699 | 116.4967 | 88.3333 | 0.004331 | 0.97666 | 0.022988 | 23.5376 | 2.5902 | 0.15699 | 0.042612 | 2.937 | 2.937 | 0.002197
2019-03-20 12:50 | 1553100608.1689 | 25 | 0.20172 | 0.007785 | 0.000114 | 108.1287 | 81.4606 | 113.4156 | 20.9943 | 0.002959 | 0.002802 | 171.7474 | 0.26426 | 8.4015 | 0.003989 | 0.005527 | 0.005428 | 6.2e-05 | 0.001162 | 0.30093 | 1.3537 | 2.548 | 100.99 | 20.73 | 17.865 | 92.7443 | 117.7357 | 88.9601 | 0.004332 | 0.97497 | 0.023002 | 23.5922 | 2.5331 | 0.15721 | 0.037597 | 2.5936 | 2.5936 | 0.002491
2019-03-20 12:50 | 1553100609.1691 | 26 | 0.20195 | 0.007604 | 0.00011 | 108.1613 | 83.4026 | 113.6954 | 20.9943 | 0.00293 | 0.002863 | 171.6535 | 0.26466 | 8.403 | 0.004041 | 0.005535 | 0.005399 | 6.2e-05 | 0.001165 | 0.32068 | 1.356 | 2.5479 | 100.99 | 20.73 | 17.866 | 92.0412 | 119.009 | 90.448 | 0.004332 | 0.97098 | 0.022983 | 23.6697 | 2.3251 | 0.15745 | 0.036478 | 2.5146 | 2.5146 | 0.002573
2019-03-20 12:50 | 1553100610.1679 | 27 | 0.20227 | 0.007387 | 0.000108 | 108.0594 | 84.0731 | 113.6518 | 20.9936 | 0.002961 | 0.00282 | 171.5418 | 0.2642 | 8.4051 | 0.004002 | 0.005611 | 0.005469 | 6.1e-05 | 0.00117 | 0.32372 | 1.3488 | 2.5482 | 100.99 | 20.73 | 17.868 | 92.788 | 118.051 | 89.4026 | 0.004333 | 0.97378 | 0.023011 | 23.6312 | 2.245 | 0.15718 | 0.034812 | 2.4031 | 2.4031 | 0.002688
2019-03-20 12:50 | 1553100611.169 | 28 | 0.20259 | 0.007163 | 0.000104 | 108.1426 | 83.9476 | 113.6903 | 20.9945 | 0.003021 | 0.002848 | 171.4502 | 0.26439 | 8.4035 | 0.00405 | 0.004998 | 0.004929 | 5.9e-05 | 0.00117 | 0.29806 | 1.3281 | 2.548 | 100.99 | 20.73 | 17.874 | 94.2554 | 119.2172 | 90.0698 | 0.004334 | 0.97199 | 0.022994 | 23.6569 | 2.2653 | 0.15728 | 0.033193 | 2.2899 | 2.2899 | 0.002823
2019-03-20 12:50 | 1553100612.1681 | 29 | 0.20274 | 0.006936 | 0.000101 | 108.1096 | 84.4749 | 113.9797 | 20.9941 | 0.002985 | 0.002806 | 171.3785 | 0.26499 | 8.4016 | 0.003939 | 0.005943 | 0.005766 | 5.8e-05 | 0.001174 | 0.33876 | 1.2659 | 2.5481 | 100.99 | 20.73 | 17.877 | 93.3848 | 116.5018 | 89.0647 | 0.004335 | 0.97469 | 0.023055 | 23.6542 | 2.2065 | 0.15764 | 0.03258 | 2.2538 | 2.2538 | 0.002875
2019-03-20 12:50 | 1553100613.1692 | 30 | 0.20285 | 0.006711 | 9.8e-05 | 108.1177 | 85.2314 | 114.1592 | 20.994 | 0.003051 | 0.002844 | 171.3049 | 0.26407 | 8.4014 | 0.004033 | 0.005266 | 0.005123 | 5.5e-05 | 0.001173 | 0.30773 | 1.321 | 2.5482 | 100.99 | 20.73 | 17.881 | 94.9796 | 118.7918 | 89.9905 | 0.004336 | 0.9722 | 0.023044 | 23.703 | 2.1275 | 0.1571 | 0.03215 | 2.2232 | 2.2232 | 0.002904
2019-03-20 12:50 | 1553100614.1672 | 31 | 0.20299 | 0.006519 | 9.6e-05 | 108.0617 | 86.3594 | 114.8723 | 20.994 | 0.003066 | 0.002818 | 171.2015 | 0.26413 | 8.4011 | 0.00395 | 0.005847 | 0.005792 | 5.7e-05 | 0.001173 | 0.32853 | 1.3036 | 2.5481 | 100.99 | 20.73 | 17.885 | 95.3396 | 116.7715 | 89.3583 | 0.004337 | 0.9739 | 0.023136 | 23.7563 | 2.0052 | 0.15713 | 0.031581 | 2.1927 | 2.1927 | 0.002945
2019-03-20 12:50 | 1553100615.1672 | 32 | 0.20309 | 0.006339 | 9.3e-05 | 108.096 | 87.5427 | 115.1394 | 20.9939 | 0.003054 | 0.002872 | 171.1111 | 0.26401 | 8.4027 | 0.003879 | 0.005328 | 0.005208 | 5.3e-05 | 0.00117 | 0.32659 | 1.3091 | 2.548 | 100.99 | 20.72 | 17.891 | 95.0373 | 115.0533 | 90.6625 | 0.004335 | 0.9704 | 0.023121 | 23.8266 | 1.8863 | 0.15706 | 0.031225 | 2.1668 | 2.1668 | 0.002979
2019-03-20 12:50 | 1553100616.1673 | 33 | 0.20319 | 0.006179 | 9e-05 | 108.0543 | 88.4494 | 115.7701 | 20.994 | 0.003133 | 0.002878 | 171.0317 | 0.26431 | 8.4039 | 0.003981 | 0.00539 | 0.00527 | 5.3e-05 | 0.001175 | 0.31725 | 1.3074 | 2.5481 | 100.99 | 20.72 | 17.904 | 96.9493 | 117.5389 | 90.7946 | 0.004339 | 0.97005 | 0.02318 | 23.8961 | 1.7907 | 0.15724 | 0.030782 | 2.1417 | 2.1417 | 0.003017
2019-03-20 12:50 | 1553100617.1687 | 34 | 0.2034 | 0.006038 | 8.8e-05 | 108.031 | 88.1919 | 115.9434 | 20.994 | 0.003024 | 0.002847 | 170.9592 | 0.26343 | 8.4026 | 0.00393 | 0.005303 | 0.00511 | 5.3e-05 | 0.001172 | 0.347 | 1.3354 | 2.5481 | 100.99 | 20.72 | 17.912 | 94.3079 | 116.2855 | 90.0567 | 0.004341 | 0.97202 | 0.023221 | 23.8897 | 1.8149 | 0.15671 | 0.029709 | 2.0708 | 2.0708 | 0.00311
2019-03-20 12:50 | 1553100618.1684 | 35 | 0.20364 | 0.005903 | 8.5e-05 | 108.0103 | 88.5518 | 116.3865 | 20.9942 | 0.003083 | 0.002822 | 170.8616 | 0.26383 | 8.4035 | 0.003924 | 0.005362 | 0.005205 | 5.3e-05 | 0.001175 | 0.33516 | 1.3552 | 2.5481 | 100.99 | 20.72 | 17.919 | 95.7431 | 116.1488 | 89.4414 | 0.004342 | 0.97367 | 0.023285 | 23.9151 | 1.7767 | 0.15696 | 0.02847 | 1.9902 | 1.9902 | 0.003241
2019-03-20 12:50 | 1553100619.1678 | 36 | 0.20398 | 0.005773 | 8.1e-05 | 108.086 | 89.2197 | 116.5107 | 20.9941 | 0.003033 | 0.002861 | 170.8096 | 0.26471 | 8.4013 | 0.004034 | 0.005234 | 0.00511 | 5.4e-05 | 0.00117 | 0.32101 | 1.3049 | 2.5481 | 100.99 | 20.72 | 17.934 | 94.5413 | 118.8297 | 90.3827 | 0.004346 | 0.97115 | 0.023268 | 23.9589 | 1.7158 | 0.15747 | 0.026585 | 1.8572 | 1.8572 | 0.003485
2019-03-20 12:50 | 1553100620.1677 | 37 | 0.20424 | 0.005639 | 7.8e-05 | 108.0898 | 89.6918 | 117.109 | 20.9941 | 0.003047 | 0.002814 | 170.735 | 0.26559 | 8.4012 | 0.003984 | 0.004857 | 0.004768 | 6e-05 | 0.001169 | 0.33395 | 1.316 | 2.548 | 100.99 | 20.72 | 17.94 | 94.8817 | 117.6167 | 89.2631 | 0.004347 | 0.97415 | 0.023363 | 23.9833 | 1.6689 | 0.158 | 0.025154 | 1.7647 | 1.7647 | 0.00368
2019-03-20 12:50 | 1553100621.1675 | 38 | 0.20453 | 0.005496 | 7.5e-05 | 108.0753 | 89.9496 | 117.8442 | 20.994 | 0.00305 | 0.002819 | 170.6382 | 0.26361 | 8.4025 | 0.004013 | 0.005586 | 0.005505 | 5.7e-05 | 0.001168 | 0.32587 | 1.3491 | 2.5481 | 100.99 | 20.72 | 17.951 | 94.9373 | 118.3256 | 89.366 | 0.00435 | 0.97388 | 0.023433 | 24.0618 | 1.642 | 0.15682 | 0.023606 | 1.6612 | 1.6612 | 0.00388
2019-03-20 12:50 | 1553100622.1677 | 39 | 0.20467 | 0.005343 | 7.1e-05 | 108.0174 | 90.1272 | 118.2655 | 20.9932 | 0.003036 | 0.00285 | 170.5536 | 0.26503 | 8.4041 | 0.004081 | 0.005895 | 0.005746 | 5.3e-05 | 0.001167 | 0.32139 | 1.3402 | 2.5478 | 100.99 | 20.72 | 17.955 | 94.6125 | 119.977 | 90.1124 | 0.004351 | 0.97187 | 0.023451 | 24.1296 | 1.6196 | 0.15767 | 0.02295 | 1.6165 | 1.6165 | 0.004009
2019-03-20 12:50 | 1553100623.1677 | 40 | 0.20489 | 0.00518 | 6.8e-05 | 107.9977 | 90.7069 | 118.3469 | 20.9935 | 0.003068 | 0.002822 | 170.4859 | 0.26446 | 8.4058 | 0.003991 | 0.005596 | 0.005456 | 5.6e-05 | 0.001168 | 0.28711 | 1.3248 | 2.5482 | 100.99 | 20.72 | 17.964 | 95.3729 | 117.7642 | 89.447 | 0.004353 | 0.97366 | 0.023481 | 24.1158 | 1.5606 | 0.15733 | 0.02187 | 1.5424 | 1.5424 | 0.004192
2019-03-20 12:50 | 1553100624.1677 | 41 | 0.20513 | 0.005006 | 6.4e-05 | 108.0307 | 91.8255 | 118.0733 | 20.9941 | 0.00303 | 0.002844 | 170.4282 | 0.26368 | 8.4013 | 0.004003 | 0.005707 | 0.005603 | 5.4e-05 | 0.001172 | 0.32012 | 1.3213 | 2.548 | 100.99 | 20.72 | 17.968 | 94.4725 | 118.0793 | 89.9851 | 0.004354 | 0.97222 | 0.023436 | 24.1058 | 1.4537 | 0.15686 | 0.020608 | 1.4509 | 1.4509 | 0.004443
2019-03-20 12:50 | 1553100625.1692 | 42 | 0.2053 | 0.004837 | 6.2e-05 | 108.0378 | 92.9524 | 118.4286 | 20.9936 | 0.003069 | 0.002806 | 170.3693 | 0.26346 | 8.403 | 0.003941 | 0.005877 | 0.005748 | 5.4e-05 | 0.001169 | 0.29806 | 1.3063 | 2.5481 | 100.99 | 20.72 | 17.972 | 95.3957 | 116.5535 | 89.058 | 0.004355 | 0.9747 | 0.023501 | 24.1112 | 1.3452 | 0.15673 | 0.019813 | 1.399 | 1.399 | 0.004605
2019-03-20 12:50 | 1553100626.1693 | 43 | 0.20548 | 0.004664 | 5.8e-05 | 108.031 | 93.2693 | 118.5197 | 20.9938 | 0.003044 | 0.002786 | 170.305 | 0.26432 | 8.4052 | 0.003917 | 0.005342 | 0.005232 | 5.5e-05 | 0.001169 | 0.32565 | 1.3077 | 2.548 | 100.99 | 20.72 | 17.978 | 94.8001 | 115.9694 | 88.5911 | 0.004356 | 0.97596 | 0.023525 | 24.1049 | 1.3142 | 0.15724 | 0.018944 | 1.3391 | 1.3391 | 0.004826
2019-03-20 12:50 | 1553100627.1678 | 44 | 0.20564 | 0.00451 | 5.6e-05 | 108.0174 | 93.6109 | 118.7643 | 20.9941 | 0.002982 | 0.002781 | 170.2383 | 0.26454 | 8.4035 | 0.003922 | 0.00516 | 0.005057 | 5.5e-05 | 0.001161 | 0.31797 | 1.2763 | 2.548 | 100.99 | 20.71 | 17.981 | 93.2987 | 116.0917 | 88.4679 | 0.004355 | 0.97629 | 0.023554 | 24.1257 | 1.2804 | 0.15738 | 0.018192 | 1.2876 | 1.2876 | 0.005023
2019-03-20 12:50 | 1553100628.1683 | 45 | 0.20567 | 0.004357 | 5.3e-05 | 108.0627 | 93.7466 | 118.7319 | 20.9937 | 0.003094 | 0.002762 | 170.1587 | 0.26495 | 8.404 | 0.003978 | 0.005836 | 0.005752 | 5.7e-05 | 0.001165 | 0.31564 | 1.2927 | 2.5481 | 100.99 | 20.72 | 17.983 | 95.9997 | 117.4688 | 87.9981 | 0.004358 | 0.97756 | 0.023566 | 24.1067 | 1.2712 | 0.15762 | 0.018223 | 1.2906 | 1.2906 | 0.005019
2019-03-20 12:50 | 1553100629.168 | 46 | 0.20574 | 0.004224 | 5e-05 | 108.0763 | 94.1428 | 119.0504 | 20.9938 | 0.003087 | 0.002787 | 170.0781 | 0.26456 | 8.4032 | 0.003934 | 0.005663 | 0.005515 | 5.1e-05 | 0.001167 | 0.31924 | 1.3396 | 2.548 | 100.99 | 20.72 | 17.986 | 95.8352 | 116.3898 | 88.5957 | 0.004358 | 0.97595 | 0.023578 | 24.159 | 1.2346 | 0.15739 | 0.017963 | 1.273 | 1.273 | 0.005081
2019-03-20 12:50 | 1553100630.167 | 47 | 0.20572 | 0.004112 | 4.8e-05 | 108.0611 | 94.364 | 118.9462 | 20.9942 | 0.003057 | 0.002788 | 170.027 | 0.26442 | 8.4008 | 0.003839 | 0.005202 | 0.00509 | 5.2e-05 | 0.001172 | 0.33588 | 1.3756 | 2.5481 | 100.99 | 20.72 | 17.988 | 95.1144 | 114.0749 | 88.6225 | 0.004359 | 0.97588 | 0.023567 | 24.1493 | 1.2123 | 0.1573 | 0.018229 | 1.2914 | 1.2914 | 0.005006
2019-03-20 12:50 | 1553100631.1681 | 48 | 0.20588 | 0.004 | 4.7e-05 | 108.0054 | 94.5315 | 118.8856 | 20.9931 | 0.003019 | 0.002708 | 169.9558 | 0.26369 | 8.4023 | 0.003859 | 0.00565 | 0.005545 | 5.5e-05 | 0.001173 | 0.31852 | 1.3351 | 2.548 | 100.99 | 20.72 | 17.994 | 94.2077 | 114.5607 | 86.6936 | 0.00436 | 0.98111 | 0.023624 | 24.0787 | 1.1919 | 0.15687 | 0.01742 | 1.2372 | 1.2372 | 0.005211
2019-03-20 12:51 | 1553100632.1682 | 49 | 0.20589 | 0.003892 | 4.5e-05 | 108.0255 | 95.039 | 118.6171 | 20.9945 | 0.00301 | 0.002701 | 169.9085 | 0.26443 | 8.4031 | 0.003818 | 0.005597 | 0.005451 | 5.5e-05 | 0.001174 | 0.32388 | 1.3642 | 2.5481 | 100.99 | 20.71 | 17.997 | 93.9743 | 113.5512 | 86.5411 | 0.004358 | 0.98152 | 0.023602 | 24.0464 | 1.1456 | 0.15731 | 0.017466 | 1.2394 | 1.2394 | 0.005217
2019-03-20 12:51 | 1553100633.1672 | 50 | 0.20593 | 0.003807 | 4.4e-05 | 108.042 | 95.8037 | 118.8209 | 20.9947 | 0.003056 | 0.002686 | 169.8335 | 0.26494 | 8.4028 | 0.003838 | 0.005831 | 0.005707 | 5.7e-05 | 0.001178 | 0.3194 | 1.388 | 2.5482 | 100.99 | 20.71 | 17.999 | 95.1009 | 114.0398 | 86.1769 | 0.004359 | 0.98252 | 0.023634 | 24.0548 | 1.0753 | 0.15761 | 0.017338 | 1.232 | 1.232 | 0.005258
2019-03-20 12:51 | 1553100634.1683 | 51 | 0.20594 | 0.003726 | 4.4e-05 | 108.0953 | 95.945 | 118.9525 | 20.9946 | 0.00301 | 0.002734 | 169.7935 | 0.26422 | 8.4027 | 0.00387 | 0.004813 | 0.004752 | 5.6e-05 | 0.001176 | 0.32836 | 1.3887 | 2.5481 | 100.99 | 20.71 | 18.004 | 93.9869 | 114.8215 | 87.3257 | 0.00436 | 0.97939 | 0.02361 | 24.1066 | 1.0665 | 0.15718 | 0.017375 | 1.2334 | 1.2334 | 0.005238
2019-03-20 12:51 | 1553100635.1686 | 52 | 0.20594 | 0.003653 | 4.3e-05 | 108.0983 | 95.7062 | 118.9793 | 20.9938 | 0.003001 | 0.002748 | 169.7379 | 0.26437 | 8.4031 | 0.003845 | 0.004812 | 0.004674 | 5.6e-05 | 0.001176 | 0.30541 | 1.3258 | 2.548 | 100.99 | 20.71 | 18.005 | 93.7561 | 114.2246 | 87.6592 | 0.00436 | 0.97848 | 0.023601 | 24.1205 | 1.0891 | 0.15727 | 0.017461 | 1.2391 | 1.2391 | 0.005217
2019-03-20 12:51 | 1553100636.1674 | 53 | 0.20597 | 0.003589 | 4.3e-05 | 108.1015 | 95.9816 | 119.0443 | 20.9942 | 0.003007 | 0.002773 | 169.6753 | 0.26388 | 8.4031 | 0.003854 | 0.005238 | 0.005139 | 5.4e-05 | 0.001174 | 0.31852 | 1.3118 | 2.5482 | 100.99 | 20.71 | 18.007 | 93.9176 | 114.4378 | 88.2783 | 0.004361 | 0.97681 | 0.023588 | 24.1477 | 1.0636 | 0.15698 | 0.017334 | 1.2293 | 1.2293 | 0.005248
2019-03-20 12:51 | 1553100637.1684 | 54 | 0.206 | 0.003532 | 4.3e-05 | 108.0737 | 95.9203 | 118.8936 | 20.9936 | 0.002979 | 0.002724 | 169.6273 | 0.26458 | 8.4005 | 0.003824 | 0.006167 | 0.005995 | 5.2e-05 | 0.001176 | 0.30934 | 1.2892 | 2.5479 | 100.99 | 20.71 | 18.008 | 93.2372 | 113.7084 | 87.0983 | 0.004361 | 0.98001 | 0.023611 | 24.093 | 1.067 | 0.1574 | 0.017267 | 1.2259 | 1.2259 | 0.005277
2019-03-20 12:51 | 1553100638.1684 | 55 | 0.20603 | 0.003488 | 4.3e-05 | 108.0714 | 96.3468 | 118.7969 | 20.9941 | 0.003079 | 0.002729 | 169.5876 | 0.26407 | 8.4029 | 0.00386 | 0.005896 | 0.005745 | 5.3e-05 | 0.001173 | 0.31902 | 1.2979 | 2.5481 | 100.99 | 20.71 | 18.008 | 95.646 | 114.572 | 87.216 | 0.004361 | 0.97969 | 0.023598 | 24.0872 | 1.0272 | 0.15709 | 0.017158 | 1.2174 | 1.2174 | 0.005303
2019-03-20 12:51 | 1553100639.1689 | 56 | 0.20604 | 0.003443 | 4.3e-05 | 107.9964 | 96.6267 | 118.4368 | 20.9933 | 0.00297 | 0.002714 | 169.5383 | 0.26523 | 8.4027 | 0.003908 | 0.005419 | 0.00529 | 5.7e-05 | 0.00117 | 0.31902 | 1.3179 | 2.5482 | 100.99 | 20.71 | 18.008 | 93.0132 | 115.7616 | 86.8406 | 0.004361 | 0.98071 | 0.023574 | 24.0381 | 0.99501 | 0.15779 | 0.017127 | 1.214 | 1.214 | 0.005342
2019-03-20 12:51 | 1553100640.1685 | 57 | 0.20607 | 0.003413 | 4.1e-05 | 107.9741 | 96.133 | 118.5237 | 20.9932 | 0.002992 | 0.002703 | 169.4793 | 0.26567 | 8.3998 | 0.003936 | 0.006004 | 0.005925 | 5.5e-05 | 0.001168 | 0.32189 | 1.2889 | 2.5479 | 100.99 | 20.71 | 18.008 | 93.5481 | 116.4265 | 86.5866 | 0.004361 | 0.9814 | 0.023591 | 24.0384 | 1.039 | 0.15805 | 0.016976 | 1.2043 | 1.2043 | 0.005394
2019-03-20 12:51 | 1553100641.1676 | 58 | 0.206 | 0.003389 | 4.1e-05 | 108.1578 | 96.996 | 118.284 | 20.9938 | 0.002996 | 0.002741 | 169.444 | 0.26338 | 8.4011 | 0.003911 | 0.006213 | 0.006121 | 5.1e-05 | 0.001171 | 0.33229 | 1.3224 | 2.548 | 100.99 | 20.71 | 18.006 | 93.6473 | 115.8287 | 87.5007 | 0.004361 | 0.97891 | 0.023538 | 24.0446 | 0.97425 | 0.15669 | 0.01743 | 1.2337 | 1.2337 | 0.00522
2019-03-20 12:51 | 1553100642.1681 | 59 | 0.20604 | 0.003352 | 4.1e-05 | 108.0614 | 96.722 | 118.4338 | 20.994 | 0.002969 | 0.002757 | 169.3881 | 0.26366 | 8.4043 | 0.003908 | 0.005453 | 0.005294 | 5.2e-05 | 0.001169 | 0.33555 | 1.3229 | 2.5481 | 100.99 | 20.71 | 18.005 | 92.9989 | 115.7411 | 87.8758 | 0.00436 | 0.97789 | 0.02354 | 24.0723 | 0.99158 | 0.15685 | 0.017252 | 1.2212 | 1.2212 | 0.005279
2019-03-20 12:51 | 1553100643.1678 | 60 | 0.20602 | 0.003332 | 4e-05 | 108.1096 | 96.9653 | 118.6122 | 20.9935 | 0.003 | 0.00272 | 169.3514 | 0.26454 | 8.4018 | 0.003972 | 0.005176 | 0.005119 | 5.1e-05 | 0.001167 | 0.34075 | 1.3005 | 2.5481 | 100.99 | 20.71 | 18.004 | 93.7401 | 117.3126 | 86.989 | 0.00436 | 0.9803 | 0.023587 | 24.0609 | 0.97309 | 0.15737 | 0.017402 | 1.2343 | 1.2343 | 0.00524
2019-03-20 12:51 | 1553100644.1678 | 61 | 0.20608 | 0.003318 | 4e-05 | 108.14 | 97.0633 | 118.7913 | 20.9935 | 0.002965 | 0.002696 | 169.2886 | 0.26398 | 8.4015 | 0.003896 | 0.005816 | 0.005661 | 5.5e-05 | 0.001168 | 0.30431 | 1.3379 | 2.5481 | 100.99 | 20.71 | 18.003 | 92.9015 | 115.4533 | 86.4067 | 0.00436 | 0.98189 | 0.023624 | 24.0595 | 0.96658 | 0.15704 | 0.017072 | 1.2128 | 1.2128 | 0.005322
2019-03-20 12:51 | 1553100645.1683 | 62 | 0.20601 | 0.003298 | 3.9e-05 | 108.0607 | 97.6071 | 118.7603 | 20.9937 | 0.003073 | 0.00268 | 169.2336 | 0.26396 | 8.4024 | 0.003887 | 0.004875 | 0.004805 | 5.2e-05 | 0.001166 | 0.28263 | 1.3217 | 2.5481 | 100.99 | 20.71 | 18.003 | 95.4937 | 115.2416 | 86.0236 | 0.00436 | 0.98294 | 0.023633 | 24.0436 | 0.91005 | 0.15703 | 0.017473 | 1.2418 | 1.2418 | 0.005197
2019-03-20 12:51 | 1553100646.2184 | 63 | 0.20602 | 0.00328 | 3.9e-05 | 108.1118 | 97.4512 | 118.4488 | 20.993 | 0.002962 | 0.002697 | 169.1892 | 0.26348 | 8.4028 | 0.003907 | 0.004715 | 0.004565 | 5.4e-05 | 0.001162 | 0.28208 | 1.3424 | 2.5481 | 100.99 | 20.71 | 18.001 | 92.8243 | 115.7186 | 86.4506 | 0.004359 | 0.98177 | 0.023588 | 24.0263 | 0.92857 | 0.15675 | 0.017473 | 1.2395 | 1.2395 | 0.005197
2019-03-20 12:51 | 1553100647.1707 | 64 | 0.20605 | 0.00327 | 3.9e-05 | 108.108 | 97.4008 | 118.7108 | 20.9928 | 0.003092 | 0.002685 | 169.1265 | 0.26463 | 8.4014 | 0.003876 | 0.005629 | 0.005596 | 5.1e-05 | 0.001155 | 0.33196 | 1.3363 | 2.548 | 100.99 | 20.71 | 17.997 | 95.9681 | 114.9764 | 86.1437 | 0.004358 | 0.98261 | 0.023624 | 24.0426 | 0.93288 | 0.15743 | 0.017324 | 1.2308 | 1.2308 | 0.005257
2019-03-20 12:51 | 1553100648.1706 | 65 | 0.20606 | 0.003263 | 3.9e-05 | 108.0824 | 97.1731 | 118.5066 | 20.9933 | 0.003035 | 0.002704 | 169.0625 | 0.26451 | 8.4041 | 0.003902 | 0.005317 | 0.00522 | 5.6e-05 | 0.001157 | 0.36138 | 1.3355 | 2.548 | 100.99 | 20.71 | 17.994 | 94.5858 | 115.6121 | 86.6078 | 0.004358 | 0.98134 | 0.023589 | 24.0374 | 0.9517 | 0.15736 | 0.017228 | 1.2221 | 1.2221 | 0.005292
2019-03-20 12:51 | 1553100649.1702 | 66 | 0.20605 | 0.003256 | 3.9e-05 | 108.0543 | 97.1399 | 118.6688 | 20.9935 | 0.002957 | 0.002703 | 169.0205 | 0.26475 | 8.4025 | 0.003824 | 0.005358 | 0.005199 | 5.6e-05 | 0.001157 | 0.32493 | 1.3324 | 2.5481 | 100.99 | 20.71 | 17.987 | 92.7091 | 113.7122 | 86.573 | 0.004356 | 0.98144 | 0.023606 | 24.0527 | 0.95243 | 0.1575 | 0.017308 | 1.2287 | 1.2287 | 0.005268
2019-03-20 12:51 | 1553100650.1703 | 67 | 0.20604 | 0.003251 | 4.1e-05 | 108.1209 | 97.5383 | 118.515 | 20.993 | 0.003015 | 0.002675 | 168.9821 | 0.26381 | 8.4035 | 0.003861 | 0.005045 | 0.004922 | 5.2e-05 | 0.001155 | 0.33987 | 1.2945 | 2.548 | 100.99 | 20.71 | 17.983 | 94.0932 | 114.6035 | 85.9057 | 0.004355 | 0.98326 | 0.023613 | 24.0148 | 0.92133 | 0.15694 | 0.017381 | 1.2342 | 1.2342 | 0.005226
2019-03-20 12:51 | 1553100651.1717 | 68 | 0.20604 | 0.003244 | 4.1e-05 | 108.0801 | 97.3783 | 118.476 | 20.9941 | 0.00302 | 0.002655 | 168.9171 | 0.26419 | 8.4007 | 0.003754 | 0.005115 | 0.004976 | 5.3e-05 | 0.001161 | 0.3126 | 1.2951 | 2.5481 | 100.99 | 20.71 | 17.977 | 94.2313 | 112.001 | 85.4346 | 0.004354 | 0.98455 | 0.023624 | 23.9951 | 0.93264 | 0.15716 | 0.017373 | 1.2342 | 1.2342 | 0.005234
2019-03-20 12:51 | 1553100652.1704 | 69 | 0.20602 | 0.003242 | 4e-05 | 108.0999 | 97.1616 | 118.7594 | 20.9941 | 0.002981 | 0.002692 | 168.8562 | 0.26394 | 8.4046 | 0.003755 | 0.004867 | 0.004781 | 5.4e-05 | 0.001161 | 0.30973 | 1.3139 | 2.5481 | 100.99 | 20.7 | 17.969 | 93.2799 | 112.0245 | 86.3137 | 0.004349 | 0.98215 | 0.023624 | 24.0532 | 0.95421 | 0.15702 | 0.017537 | 1.2458 | 1.2458 | 0.00518
2019-03-20 12:51 | 1553100653.1714 | 70 | 0.20603 | 0.003234 | 4e-05 | 108.0973 | 96.7254 | 118.5523 | 20.9935 | 0.003007 | 0.002676 | 168.7919 | 0.26303 | 8.402 | 0.003779 | 0.005463 | 0.005332 | 5.6e-05 | 0.001159 | 0.31902 | 1.3313 | 2.5481 | 100.99 | 20.7 | 17.965 | 93.9099 | 112.6006 | 85.9257 | 0.004348 | 0.98321 | 0.023616 | 24.0192 | 0.99424 | 0.15648 | 0.01747 | 1.2407 | 1.2407 | 0.005184
2019-03-20 12:51 | 1553100654.1719 | 71 | 0.206 | 0.003231 | 4e-05 | 108.1471 | 97.1578 | 118.45 | 20.9934 | 0.002983 | 0.002719 | 168.7407 | 0.26394 | 8.404 | 0.00375 | 0.004797 | 0.004702 | 5.3e-05 | 0.001163 | 0.32173 | 1.3196 | 2.548 | 100.99 | 20.7 | 17.954 | 93.3263 | 111.9088 | 86.9576 | 0.004345 | 0.98039 | 0.023572 | 24.0433 | 0.95846 | 0.15702 | 0.017632 | 1.2498 | 1.2498 | 0.005163
2019-03-20 12:51 | 1553100655.1706 | 72 | 0.20601 | 0.003229 | 4.1e-05 | 108.1471 | 97.0636 | 118.4127 | 20.9937 | 0.002974 | 0.002678 | 168.6797 | 0.26353 | 8.4011 | 0.003734 | 0.004912 | 0.004808 | 5.4e-05 | 0.001152 | 0.31868 | 1.3169 | 2.548 | 100.99 | 20.7 | 17.949 | 93.1143 | 111.5039 | 85.9735 | 0.004344 | 0.98308 | 0.0236 | 24.0067 | 0.96713 | 0.15677 | 0.017561 | 1.2463 | 1.2463 | 0.00517
2019-03-20 12:51 | 1553100656.1706 | 73 | 0.20603 | 0.003229 | 4.1e-05 | 108.1539 | 96.9672 | 118.5446 | 20.9935 | 0.002954 | 0.002663 | 168.6041 | 0.26357 | 8.4029 | 0.003782 | 0.005065 | 0.004894 | 5.5e-05 | 0.001151 | 0.32139 | 1.293 | 2.5479 | 100.99 | 20.7 | 17.943 | 92.6213 | 112.6875 | 85.6284 | 0.004343 | 0.98402 | 0.023625 | 24.0085 | 0.97659 | 0.1568 | 0.017475 | 1.2415 | 1.2415 | 0.005191
2019-03-20 12:51 | 1553100657.1721 | 74 | 0.206 | 0.003228 | 4.2e-05 | 108.1943 | 97.2785 | 118.2876 | 20.9938 | 0.003045 | 0.002703 | 168.57 | 0.26414 | 8.4006 | 0.003715 | 0.005363 | 0.005227 | 5.7e-05 | 0.001141 | 0.3027 | 1.3108 | 2.5481 | 100.99 | 20.7 | 17.934 | 94.8307 | 111.0455 | 86.5795 | 0.00434 | 0.98142 | 0.023568 | 24.0142 | 0.95126 | 0.15714 | 0.017634 | 1.2497 | 1.2497 | 0.005168
2019-03-20 12:51 | 1553100658.1717 | 75 | 0.20605 | 0.003227 | 4.2e-05 | 108.1283 | 96.944 | 118.5519 | 20.9941 | 0.003062 | 0.002722 | 168.5207 | 0.26377 | 8.4015 | 0.003702 | 0.005168 | 0.005115 | 5.5e-05 | 0.00114 | 0.32908 | 1.3173 | 2.5481 | 100.99 | 20.7 | 17.93 | 95.2325 | 110.7279 | 87.0379 | 0.00434 | 0.98017 | 0.023579 | 24.0564 | 0.97661 | 0.15692 | 0.017329 | 1.2286 | 1.2286 | 0.005249
2019-03-20 12:51 | 1553100659.1707 | 76 | 0.206 | 0.003228 | 4.3e-05 | 108.1164 | 97.4774 | 118.7331 | 20.9935 | 0.003006 | 0.002719 | 168.4398 | 0.26331 | 8.4013 | 0.003697 | 0.005425 | 0.005392 | 5.6e-05 | 0.001143 | 0.34451 | 1.3212 | 2.548 | 100.99 | 20.7 | 17.927 | 93.8795 | 110.6018 | 86.9593 | 0.004339 | 0.98038 | 0.0236 | 24.0721 | 0.92655 | 0.15664 | 0.017673 | 1.2541 | 1.2541 | 0.005134
2019-03-20 12:51 | 1553100660.171 | 77 | 0.20594 | 0.003228 | 4.3e-05 | 108.1746 | 97.4404 | 118.8781 | 20.9939 | 0.002973 | 0.002691 | 168.363 | 0.26376 | 8.4039 | 0.003711 | 0.005046 | 0.004951 | 6e-05 | 0.001143 | 0.32388 | 1.3414 | 2.5481 | 100.99 | 20.69 | 17.925 | 93.0849 | 110.9573 | 86.3072 | 0.004336 | 0.98216 | 0.023636 | 24.065 | 0.93476 | 0.15691 | 0.018032 | 1.2815 | 1.2815 | 0.005032
2019-03-20 12:51 | 1553100661.1708 | 78 | 0.20577 | 0.003229 | 4.3e-05 | 108.0863 | 97.2015 | 118.8673 | 20.9935 | 0.003038 | 0.002715 | 168.3061 | 0.26414 | 8.4027 | 0.003844 | 0.004808 | 0.004722 | 5.6e-05 | 0.001144 | 0.32853 | 1.3347 | 2.548 | 100.99 | 20.69 | 17.925 | 94.6449 | 114.2021 | 86.8689 | 0.004336 | 0.98063 | 0.023616 | 24.0827 | 0.9494 | 0.15714 | 0.01905 | 1.3528 | 1.3528 | 0.004774
2019-03-20 12:51 | 1553100662.1709 | 79 | 0.20609 | 0.003244 | 4.3e-05 | 108.1021 | 97.5207 | 118.5901 | 20.9941 | 0.002977 | 0.002729 | 168.2409 | 0.26493 | 8.3997 | 0.003896 | 0.005659 | 0.005515 | 6.3e-05 | 0.001147 | 0.30381 | 1.3614 | 2.5482 | 100.99 | 20.69 | 17.931 | 93.1728 | 115.45 | 87.2083 | 0.004337 | 0.97971 | 0.023578 | 24.0659 | 0.92139 | 0.15761 | 0.017098 | 1.2121 | 1.2121 | 0.005344
2019-03-20 12:51 | 1553100663.1709 | 80 | 0.20589 | 0.00325 | 4.3e-05 | 108.0756 | 97.0468 | 118.5859 | 20.9928 | 0.002973 | 0.002734 | 168.1681 | 0.2639 | 8.4012 | 0.003877 | 0.005381 | 0.005264 | 6.2e-05 | 0.001144 | 0.27959 | 1.3366 | 2.548 | 100.99 | 20.69 | 17.934 | 93.0956 | 115.0058 | 87.3281 | 0.004338 | 0.97938 | 0.023573 | 24.0695 | 0.96277 | 0.15699 | 0.018251 | 1.2937 | 1.2937 | 0.004988
2019-03-20 12:51 | 1553100664.17 | 81 | 0.20594 | 0.003257 | 4.2e-05 | 108.1429 | 97.6209 | 118.5037 | 20.9928 | 0.002986 | 0.002708 | 168.1182 | 0.26416 | 8.4014 | 0.003784 | 0.005469 | 0.005389 | 5.7e-05 | 0.001144 | 0.29679 | 1.3678 | 2.5481 | 100.99 | 20.69 | 17.939 | 93.41 | 112.7167 | 86.7128 | 0.004339 | 0.98106 | 0.023585 | 24.0406 | 0.91558 | 0.15715 | 0.017982 | 1.2753 | 1.2753 | 0.005065
2019-03-20 12:51 | 1553100665.1714 | 82 | 0.20595 | 0.003268 | 4.2e-05 | 108.129 | 97.3466 | 118.3134 | 20.9927 | 0.002997 | 0.0027 | 168.0572 | 0.26424 | 8.4005 | 0.003767 | 0.005461 | 0.005352 | 6e-05 | 0.001142 | 0.32228 | 1.3032 | 2.5481 | 100.99 | 20.69 | 17.952 | 93.6704 | 112.3084 | 86.5043 | 0.004342 | 0.98162 | 0.023573 | 24.0144 | 0.9396 | 0.15719 | 0.017874 | 1.267 | 1.267 | 0.005099
2019-03-20 12:51 | 1553100666.1701 | 83 | 0.20591 | 0.003269 | 4.1e-05 | 108.0746 | 97.4695 | 118.2385 | 20.9928 | 0.003035 | 0.002736 | 167.9983 | 0.26454 | 8.4029 | 0.003741 | 0.00541 | 0.005268 | 5.9e-05 | 0.001144 | 0.31758 | 1.2739 | 2.5481 | 100.99 | 20.69 | 17.96 | 94.5847 | 111.6705 | 87.3726 | 0.004344 | 0.97926 | 0.023537 | 24.0357 | 0.92381 | 0.15737 | 0.018126 | 1.2829 | 1.2829 | 0.005042
2019-03-20 12:51 | 1553100667.1712 | 84 | 0.20592 | 0.00328 | 4.1e-05 | 108.1225 | 97.9647 | 118.0911 | 20.9927 | 0.002929 | 0.002725 | 167.9022 | 0.26337 | 8.4019 | 0.003707 | 0.005492 | 0.005399 | 5.9e-05 | 0.001143 | 0.28943 | 1.3092 | 2.5481 | 100.99 | 20.69 | 17.974 | 92.0278 | 110.8654 | 87.1116 | 0.004347 | 0.97997 | 0.023531 | 24.012 | 0.88244 | 0.15668 | 0.018083 | 1.2796 | 1.2796 | 0.005033
2019-03-20 12:51 | 1553100668.1708 | 85 | 0.20588 | 0.003279 | 4.1e-05 | 108.0565 | 98.3034 | 117.8867 | 20.992 | 0.002997 | 0.002705 | 167.858 | 0.26421 | 8.4016 | 0.00378 | 0.004826 | 0.004714 | 5.8e-05 | 0.001147 | 0.33931 | 1.307 | 2.5479 | 100.99 | 20.69 | 17.983 | 93.6581 | 112.6266 | 86.6278 | 0.00435 | 0.98129 | 0.023526 | 23.9751 | 0.84612 | 0.15718 | 0.018318 | 1.2959 | 1.2959 | 0.004985
2019-03-20 12:51 | 1553100669.1717 | 86 | 0.20586 | 0.003282 | 4.1e-05 | 108.0336 | 98.4207 | 117.5501 | 20.993 | 0.003001 | 0.002738 | 167.8013 | 0.26514 | 8.4051 | 0.003733 | 0.005099 | 0.004979 | 5.8e-05 | 0.001149 | 0.32189 | 1.2849 | 2.548 | 100.99 | 20.69 | 17.992 | 93.7623 | 111.493 | 87.4149 | 0.004352 | 0.97915 | 0.023467 | 23.967 | 0.83355 | 0.15773 | 0.018413 | 1.2994 | 1.2994 | 0.004989
2019-03-20 12:51 | 1553100670.1713 | 87 | 0.20582 | 0.003291 | 4.2e-05 | 108.0827 | 98.1318 | 117.4588 | 20.9922 | 0.003005 | 0.002699 | 167.7582 | 0.26374 | 8.4044 | 0.003656 | 0.00522 | 0.005098 | 5.7e-05 | 0.001147 | 0.27848 | 1.3436 | 2.548 | 100.99 | 20.69 | 18.012 | 93.8585 | 109.6217 | 86.4976 | 0.004357 | 0.98164 | 0.023488 | 23.9272 | 0.86391 | 0.1569 | 0.018676 | 1.319 | 1.319 | 0.004889
2019-03-20 12:51 | 1553100671.1714 | 88 | 0.20584 | 0.003291 | 4.2e-05 | 108.129 | 97.9061 | 117.309 | 20.993 | 0.003056 | 0.00278 | 167.7107 | 0.26408 | 8.4003 | 0.003734 | 0.005371 | 0.005184 | 5.7e-05 | 0.00115 | 0.29214 | 1.345 | 2.5481 | 100.99 | 20.69 | 18.012 | 95.0936 | 111.519 | 88.4333 | 0.004357 | 0.97639 | 0.02341 | 23.9762 | 0.88834 | 0.1571 | 0.018526 | 1.3041 | 1.3041 | 0.004951
2019-03-20 12:51 | 1553100672.1714 | 89 | 0.20583 | 0.003302 | 4.3e-05 | 108.0378 | 97.8519 | 117.002 | 20.9933 | 0.003041 | 0.002708 | 167.6617 | 0.26492 | 8.4053 | 0.003686 | 0.006156 | 0.006066 | 5.6e-05 | 0.001148 | 0.3282 | 1.3349 | 2.5479 | 100.99 | 20.69 | 18.032 | 94.7277 | 110.3515 | 86.7126 | 0.004361 | 0.98106 | 0.023435 | 23.8878 | 0.88574 | 0.1576 | 0.018588 | 1.3099 | 1.3099 | 0.004945
2019-03-20 12:51 | 1553100673.1716 | 90 | 0.20583 | 0.003315 | 4.4e-05 | 108.1196 | 97.881 | 116.9001 | 20.9928 | 0.003086 | 0.002756 | 167.6226 | 0.26394 | 8.403 | 0.003758 | 0.00543 | 0.005319 | 5.4e-05 | 0.001145 | 0.28871 | 1.344 | 2.548 | 100.99 | 20.69 | 18.041 | 95.8226 | 112.0858 | 87.8535 | 0.004364 | 0.97796 | 0.023388 | 23.9152 | 0.88985 | 0.15702 | 0.018541 | 1.3038 | 1.3038 | 0.00495
2019-03-20 12:51 | 1553100674.1716 | 91 | 0.20581 | 0.003317 | 4.5e-05 | 108.0109 | 97.6437 | 116.8133 | 20.9929 | 0.003017 | 0.002694 | 167.5511 | 0.26378 | 8.4017 | 0.003729 | 0.005217 | 0.005092 | 5.4e-05 | 0.001154 | 0.31365 | 1.3281 | 2.5481 | 100.99 | 20.69 | 18.057 | 94.1402 | 111.3825 | 86.3654 | 0.004368 | 0.982 | 0.023428 | 23.857 | 0.90257 | 0.15692 | 0.018692 | 1.3166 | 1.3166 | 0.004899
2019-03-20 12:51 | 1553100675.1706 | 92 | 0.2058 | 0.003332 | 4.5e-05 | 108.043 | 97.4662 | 116.9013 | 20.9925 | 0.003027 | 0.002717 | 167.5071 | 0.26413 | 8.4043 | 0.003732 | 0.004288 | 0.004139 | 5.4e-05 | 0.001151 | 0.31581 | 1.347 | 2.5481 | 100.99 | 20.69 | 18.063 | 94.3965 | 111.4721 | 86.9115 | 0.004369 | 0.98051 | 0.023419 | 23.8841 | 0.9215 | 0.15713 | 0.018712 | 1.3175 | 1.3175 | 0.004902
2019-03-20 12:51 | 1553100676.1708 | 93 | 0.2058 | 0.003336 | 4.5e-05 | 108.0753 | 97.3507 | 116.756 | 20.9924 | 0.003017 | 0.002715 | 167.4507 | 0.26375 | 8.4019 | 0.003832 | 0.005108 | 0.005077 | 5.6e-05 | 0.001149 | 0.31492 | 1.3212 | 2.5479 | 100.99 | 20.69 | 18.068 | 94.1373 | 113.9063 | 86.8717 | 0.00437 | 0.98062 | 0.023405 | 23.868 | 0.93478 | 0.1569 | 0.018739 | 1.3186 | 1.3186 | 0.004891
2019-03-20 12:51 | 1553100677.1696 | 94 | 0.20578 | 0.003345 | 4.6e-05 | 107.9505 | 97.4187 | 116.9609 | 20.9914 | 0.00301 | 0.002749 | 167.4124 | 0.26392 | 8.4034 | 0.00377 | 0.00533 | 0.005197 | 5.2e-05 | 0.001148 | 0.30901 | 1.3565 | 2.5481 | 100.99 | 20.69 | 18.072 | 93.9839 | 112.3758 | 87.682 | 0.004371 | 0.97842 | 0.0234 | 23.9158 | 0.9182 | 0.15701 | 0.018829 | 1.3246 | 1.3246 | 0.004872
2019-03-20 12:51 | 1553100678.1697 | 95 | 0.20578 | 0.003346 | 4.5e-05 | 107.9999 | 97.2382 | 117.0719 | 20.9926 | 0.002979 | 0.002736 | 167.3301 | 0.26449 | 8.4022 | 0.003816 | 0.004983 | 0.004885 | 5.3e-05 | 0.001155 | 0.31526 | 1.345 | 2.5481 | 100.99 | 20.69 | 18.073 | 93.2306 | 113.5191 | 87.37 | 0.004371 | 0.97927 | 0.023421 | 23.9168 | 0.93888 | 0.15735 | 0.018831 | 1.326 | 1.326 | 0.004877
2019-03-20 12:51 | 1553100679.1699 | 96 | 0.20576 | 0.003355 | 4.6e-05 | 108.0973 | 97.5981 | 117.6496 | 20.9918 | 0.00302 | 0.002747 | 167.2806 | 0.26393 | 8.4016 | 0.003799 | 0.005579 | 0.005403 | 5.4e-05 | 0.001158 | 0.33804 | 1.3754 | 2.5482 | 100.99 | 20.69 | 18.072 | 94.226 | 113.0917 | 87.6504 | 0.004371 | 0.97851 | 0.023469 | 23.985 | 0.9139 | 0.15701 | 0.018928 | 1.3356 | 1.3356 | 0.004832
2019-03-20 12:51 | 1553100680.1708 | 97 | 0.20579 | 0.00336 | 4.6e-05 | 107.9957 | 97.8907 | 117.7439 | 20.9922 | 0.003011 | 0.002773 | 167.2052 | 0.26368 | 8.4048 | 0.003676 | 0.00525 | 0.005258 | 5.5e-05 | 0.001156 | 0.32549 | 1.3617 | 2.548 | 100.99 | 20.69 | 18.069 | 94.0128 | 110.1107 | 88.2669 | 0.00437 | 0.97684 | 0.023459 | 24.0151 | 0.87871 | 0.15686 | 0.01876 | 1.3231 | 1.3231 | 0.004873
2019-03-20 12:51 | 1553100681.1708 | 98 | 0.20574 | 0.003357 | 4.7e-05 | 108.0701 | 97.2247 | 118.1254 | 20.9925 | 0.003074 | 0.002706 | 167.1237 | 0.26462 | 8.3976 | 0.003668 | 0.004752 | 0.004616 | 5e-05 | 0.001156 | 0.30486 | 1.3702 | 2.548 | 100.99 | 20.69 | 18.066 | 95.5223 | 109.8995 | 86.6656 | 0.00437 | 0.98118 | 0.023549 | 24.0006 | 0.94593 | 0.15742 | 0.019034 | 1.3475 | 1.3475 | 0.004802
2019-03-20 12:51 | 1553100682.171 | 99 | 0.20576 | 0.003363 | 4.7e-05 | 108.0294 | 97.295 | 118.4094 | 20.9923 | 0.003037 | 0.002807 | 167.0482 | 0.26501 | 8.4013 | 0.003721 | 0.005366 | 0.005185 | 5.6e-05 | 0.00116 | 0.30287 | 1.3199 | 2.548 | 100.99 | 20.69 | 18.056 | 94.635 | 111.1948 | 89.0957 | 0.004367 | 0.9746 | 0.023498 | 24.1105 | 0.9361 | 0.15765 | 0.018908 | 1.3356 | 1.3356 | 0.004851
2019-03-20 12:51 | 1553100683.171 | 100 | 0.20573 | 0.003369 | 4.8e-05 | 108.0578 | 97.5177 | 118.2772 | 20.993 | 0.003012 | 0.002775 | 166.9916 | 0.26405 | 8.4017 | 0.003685 | 0.00538 | 0.005231 | 5.5e-05 | 0.001157 | 0.2907 | 1.383 | 2.5482 | 100.99 | 20.69 | 18.049 | 94.0362 | 110.3258 | 88.3199 | 0.004366 | 0.97669 | 0.02351 | 24.0712 | 0.918 | 0.15708 | 0.019091 | 1.3493 | 1.3493 | 0.004785
2019-03-20 12:51 | 1553100684.1711 | 101 | 0.20564 | 0.003384 | 4.8e-05 | 108.0695 | 97.4875 | 118.1798 | 20.9922 | 0.002982 | 0.002781 | 166.9464 | 0.26416 | 8.4021 | 0.00373 | 0.005263 | 0.005228 | 5.4e-05 | 0.001154 | 0.31924 | 1.4071 | 2.548 | 100.99 | 20.69 | 18.036 | 93.2961 | 111.4134 | 88.4504 | 0.004362 | 0.97634 | 0.023496 | 24.0656 | 0.92174 | 0.15715 | 0.019588 | 1.3836 | 1.3836 | 0.004668
2019-03-20 12:51 | 1553100685.1701 | 102 | 0.2057 | 0.003398 | 4.8e-05 | 108.0006 | 97.2318 | 117.8385 | 20.9917 | 0.00306 | 0.002765 | 166.8825 | 0.26472 | 8.4018 | 0.00375 | 0.00561 | 0.005437 | 5.4e-05 | 0.001149 | 0.3366 | 1.4169 | 2.5482 | 100.99 | 20.69 | 18.029 | 95.1769 | 111.8967 | 88.0775 | 0.004361 | 0.97735 | 0.023474 | 24.0185 | 0.93952 | 0.15748 | 0.019254 | 1.3588 | 1.3588 | 0.004763
2019-03-20 12:51 | 1553100686.1711 | 103 | 0.20569 | 0.003413 | 4.8e-05 | 108.0326 | 97.2445 | 117.8946 | 20.9924 | 0.003035 | 0.002792 | 166.8251 | 0.26429 | 8.4014 | 0.003686 | 0.005519 | 0.005339 | 5.8e-05 | 0.001143 | 0.29856 | 1.3818 | 2.5481 | 100.99 | 20.69 | 18.021 | 94.586 | 110.3443 | 88.7185 | 0.004359 | 0.97562 | 0.023459 | 24.0455 | 0.941 | 0.15723 | 0.019305 | 1.3614 | 1.3614 | 0.004746
2019-03-20 12:51 | 1553100687.1702 | 104 | 0.20566 | 0.003426 | 4.7e-05 | 108.0371 | 97.4236 | 117.8087 | 20.992 | 0.003046 | 0.002769 | 166.7545 | 0.2642 | 8.4013 | 0.003717 | 0.004454 | 0.004395 | 5.1e-05 | 0.001133 | 0.30021 | 1.3722 | 2.548 | 100.99 | 20.69 | 18.009 | 94.8444 | 111.1077 | 88.1769 | 0.004356 | 0.97708 | 0.023468 | 24.0187 | 0.92493 | 0.15717 | 0.019425 | 1.3704 | 1.3704 | 0.004714
2019-03-20 12:51 | 1553100688.1713 | 105 | 0.20563 | 0.003435 | 4.8e-05 | 107.9556 | 97.4699 | 117.9036 | 20.992 | 0.003029 | 0.002778 | 166.7138 | 0.26459 | 8.4043 | 0.003699 | 0.004853 | 0.004675 | 5.2e-05 | 0.001133 | 0.33931 | 1.3417 | 2.5481 | 100.99 | 20.69 | 18.003 | 94.4476 | 110.658 | 88.383 | 0.004354 | 0.97652 | 0.023471 | 24.0353 | 0.91392 | 0.15741 | 0.019616 | 1.3841 | 1.3841 | 0.004674
2019-03-20 12:51 | 1553100689.1703 | 106 | 0.20558 | 0.003448 | 4.7e-05 | 107.9831 | 97.8037 | 117.4235 | 20.9922 | 0.003081 | 0.002757 | 166.6294 | 0.26402 | 8.403 | 0.003725 | 0.005076 | 0.004943 | 5.4e-05 | 0.001127 | 0.32068 | 1.3442 | 2.5481 | 100.99 | 20.69 | 17.997 | 95.6863 | 111.2956 | 87.8741 | 0.004353 | 0.9779 | 0.02344 | 23.9694 | 0.88563 | 0.15707 | 0.019886 | 1.4013 | 1.4013 | 0.004607
2019-03-20 12:51 | 1553100690.1704 | 107 | 0.20558 | 0.003463 | 4.7e-05 | 108.0342 | 97.3682 | 117.3353 | 20.9921 | 0.003117 | 0.002828 | 166.5833 | 0.26392 | 8.4028 | 0.003634 | 0.005819 | 0.005683 | 5.3e-05 | 0.001125 | 0.3063 | 1.3332 | 2.5481 | 100.99 | 20.69 | 17.989 | 96.5705 | 109.0791 | 89.6009 | 0.004351 | 0.97325 | 0.023375 | 24.0176 | 0.92977 | 0.15701 | 0.01986 | 1.3956 | 1.3956 | 0.004624
2019-03-20 12:52 | 1553100691.1714 | 108 | 0.20556 | 0.003476 | 4.7e-05 | 108.0756 | 97.437 | 117.2912 | 20.992 | 0.003084 | 0.00278 | 166.5371 | 0.26325 | 8.4023 | 0.003695 | 0.004763 | 0.004608 | 5.5e-05 | 0.001128 | 0.31437 | 1.2874 | 2.5481 | 100.99 | 20.69 | 17.985 | 95.7747 | 110.5568 | 88.4299 | 0.00435 | 0.9764 | 0.023408 | 23.9743 | 0.92688 | 0.15661 | 0.019994 | 1.4071 | 1.4071 | 0.004574
2019-03-20 12:52 | 1553100692.1705 | 109 | 0.20557 | 0.003489 | 4.7e-05 | 108.0552 | 97.2161 | 117.3048 | 20.993 | 0.003096 | 0.002747 | 166.493 | 0.26428 | 8.4011 | 0.003704 | 0.004837 | 0.004728 | 5.5e-05 | 0.001126 | 0.31332 | 1.3134 | 2.548 | 100.99 | 20.69 | 17.976 | 96.0459 | 110.7891 | 87.6422 | 0.004348 | 0.97853 | 0.023435 | 23.9496 | 0.94549 | 0.15722 | 0.019913 | 1.403 | 1.403 | 0.004606
2019-03-20 12:52 | 1553100693.1705 | 110 | 0.20557 | 0.003505 | 4.7e-05 | 107.9537 | 97.0236 | 117.5198 | 20.9925 | 0.003041 | 0.002817 | 166.4085 | 0.26351 | 8.4037 | 0.003652 | 0.005352 | 0.005281 | 5.6e-05 | 0.001126 | 0.3115 | 1.315 | 2.5479 | 100.99 | 20.69 | 17.971 | 94.7206 | 109.5069 | 89.3224 | 0.004347 | 0.97399 | 0.023402 | 24.0273 | 0.95481 | 0.15676 | 0.019895 | 1.3997 | 1.3997 | 0.004603
2019-03-20 12:52 | 1553100694.1695 | 111 | 0.20556 | 0.00352 | 4.7e-05 | 107.9893 | 97.3709 | 117.4454 | 20.9935 | 0.003109 | 0.002748 | 166.332 | 0.26334 | 8.4036 | 0.003655 | 0.005133 | 0.005047 | 5.5e-05 | 0.001123 | 0.32084 | 1.3271 | 2.5481 | 100.99 | 20.69 | 17.966 | 96.3753 | 109.5796 | 87.6707 | 0.004346 | 0.97845 | 0.023448 | 23.9649 | 0.9258 | 0.15666 | 0.019943 | 1.4058 | 1.4058 | 0.00458
2019-03-20 12:52 | 1553100695.1697 | 112 | 0.20556 | 0.003527 | 4.8e-05 | 107.9741 | 97.4393 | 117.8477 | 20.9925 | 0.003069 | 0.002787 | 166.2706 | 0.26462 | 8.4013 | 0.003594 | 0.005072 | 0.004967 | 5.6e-05 | 0.001119 | 0.32388 | 1.363 | 2.5481 | 100.99 | 20.69 | 17.955 | 95.3995 | 108.1189 | 88.6086 | 0.004343 | 0.97591 | 0.023458 | 24.0371 | 0.91826 | 0.15742 | 0.019887 | 1.4024 | 1.4024 | 0.004614
2019-03-20 12:52 | 1553100696.1707 | 113 | 0.20556 | 0.003535 | 4.8e-05 | 107.9948 | 97.9625 | 117.5594 | 20.9927 | 0.003114 | 0.002806 | 166.2058 | 0.26355 | 8.4026 | 0.003591 | 0.004626 | 0.004434 | 5.6e-05 | 0.001123 | 0.33196 | 1.3212 | 2.5482 | 100.99 | 20.69 | 17.95 | 96.4866 | 108.0388 | 89.0679 | 0.004342 | 0.97468 | 0.023415 | 24.0229 | 0.87208 | 0.15678 | 0.019928 | 1.4027 | 1.4027 | 0.004594
2019-03-20 12:52 | 1553100697.1697 | 114 | 0.20555 | 0.003543 | 4.8e-05 | 108.0028 | 98.4525 | 117.1731 | 20.9936 | 0.003085 | 0.002814 | 166.1541 | 0.26388 | 8.4044 | 0.003595 | 0.004739 | 0.004711 | 5.8e-05 | 0.001121 | 0.34722 | 1.3298 | 2.5481 | 100.99 | 20.69 | 17.94 | 95.7952 | 108.1256 | 89.2651 | 0.004339 | 0.97415 | 0.02337 | 23.9899 | 0.82812 | 0.15699 | 0.019951 | 1.4017 | 1.4017 | 0.004603
2019-03-20 12:52 | 1553100698.1698 | 115 | 0.20555 | 0.003551 | 4.7e-05 | 107.9935 | 97.8321 | 116.974 | 20.9923 | 0.003109 | 0.002785 | 166.0776 | 0.2637 | 8.403 | 0.003575 | 0.004899 | 0.004741 | 5.2e-05 | 0.001115 | 0.35402 | 1.3294 | 2.548 | 100.99 | 20.69 | 17.933 | 96.3648 | 107.6477 | 88.5629 | 0.004338 | 0.97604 | 0.023372 | 23.9463 | 0.88389 | 0.15687 | 0.019932 | 1.4005 | 1.4005 | 0.004604
2019-03-20 12:52 | 1553100699.1709 | 116 | 0.20554 | 0.003558 | 4.6e-05 | 107.988 | 97.8037 | 116.8449 | 20.9923 | 0.003077 | 0.002798 | 166.0364 | 0.2631 | 8.4034 | 0.00357 | 0.004994 | 0.004892 | 5.6e-05 | 0.001121 | 0.33301 | 1.3291 | 2.548 | 100.99 | 20.69 | 17.927 | 95.595 | 107.5273 | 88.8677 | 0.004336 | 0.97522 | 0.02335 | 23.9431 | 0.88603 | 0.15652 | 0.019994 | 1.4036 | 1.4036 | 0.004583
2019-03-20 12:52 | 1553100700.1709 | 117 | 0.20558 | 0.00357 | 4.5e-05 | 108.0265 | 97.8115 | 116.4321 | 20.9921 | 0.003066 | 0.002759 | 165.9582 | 0.26391 | 8.4042 | 0.003612 | 0.004658 | 0.00458 | 5.9e-05 | 0.001127 | 0.31669 | 1.3653 | 2.548 | 100.99 | 20.69 | 17.919 | 95.3401 | 108.5338 | 87.9325 | 0.004334 | 0.97774 | 0.023339 | 23.8699 | 0.8885 | 0.157 | 0.01974 | 1.3852 | 1.3852 | 0.004658
2019-03-20 12:52 | 1553100701.17 | 118 | 0.20555 | 0.003567 | 4.4e-05 | 108.0721 | 98.0855 | 115.9471 | 20.9923 | 0.003107 | 0.00282 | 165.8812 | 0.26372 | 8.4029 | 0.003599 | 0.005067 | 0.004948 | 5.5e-05 | 0.001121 | 0.32621 | 1.3339 | 2.548 | 100.99 | 20.69 | 17.915 | 96.3145 | 108.2267 | 89.3933 | 0.004333 | 0.9738 | 0.023243 | 23.8683 | 0.86725 | 0.15689 | 0.019936 | 1.3932 | 1.3932 | 0.004628
2019-03-20 12:52 | 1553100702.1704 | 119 | 0.20551 | 0.003575 | 4.5e-05 | 107.966 | 98.9821 | 115.6596 | 20.9922 | 0.003127 | 0.00284 | 165.8094 | 0.26367 | 8.4014 | 0.00355 | 0.005297 | 0.005271 | 5.8e-05 | 0.001125 | 0.32173 | 1.3094 | 2.548 | 100.99 | 20.69 | 17.912 | 96.7932 | 107.0301 | 89.8761 | 0.004332 | 0.97251 | 0.023199 | 23.8545 | 0.77708 | 0.15686 | 0.020161 | 1.4062 | 1.4062 | 0.004584
2019-03-20 12:52 | 1553100703.17 | 120 | 0.20552 | 0.003581 | 4.5e-05 | 107.9964 | 98.8976 | 115.6845 | 20.9931 | 0.003158 | 0.002867 | 165.7401 | 0.26362 | 8.4044 | 0.003559 | 0.004346 | 0.004261 | 5.7e-05 | 0.001122 | 0.33107 | 1.3351 | 2.548 | 100.99 | 20.69 | 17.908 | 97.5541 | 107.2569 | 90.5436 | 0.004332 | 0.97072 | 0.02318 | 23.879 | 0.78723 | 0.15683 | 0.020071 | 1.3988 | 1.3988 | 0.004608
2019-03-20 12:52 | 1553100704.1701 | 121 | 0.20549 | 0.003584 | 4.6e-05 | 107.9899 | 98.2922 | 115.826 | 20.9919 | 0.003099 | 0.002788 | 165.6655 | 0.2637 | 8.4036 | 0.003559 | 0.004874 | 0.004731 | 5.4e-05 | 0.001131 | 0.36265 | 1.3016 | 2.5479 | 100.99 | 20.69 | 17.906 | 96.1333 | 107.2669 | 88.6344 | 0.004331 | 0.97584 | 0.023255 | 23.8308 | 0.84162 | 0.15688 | 0.020286 | 1.4184 | 1.4184 | 0.004546
2019-03-20 12:52 | 1553100705.1702 | 122 | 0.20545 | 0.0036 | 4.7e-05 | 108.0058 | 98.2959 | 115.7558 | 20.9918 | 0.003144 | 0.002802 | 165.5901 | 0.2639 | 8.4005 | 0.00357 | 0.005398 | 0.0053 | 5.5e-05 | 0.001127 | 0.31581 | 1.2997 | 2.5481 | 100.99 | 20.69 | 17.904 | 97.2027 | 107.5314 | 88.9657 | 0.004331 | 0.97495 | 0.023238 | 23.8345 | 0.8426 | 0.15699 | 0.020471 | 1.4301 | 1.4301 | 0.004512
2019-03-20 12:52 | 1553100706.1712 | 123 | 0.20542 | 0.003607 | 4.7e-05 | 108.0126 | 97.9591 | 116.1727 | 20.9917 | 0.003157 | 0.002826 | 165.5299 | 0.26315 | 8.4025 | 0.003529 | 0.004986 | 0.004908 | 5.4e-05 | 0.00113 | 0.32637 | 1.3026 | 2.548 | 100.99 | 20.69 | 17.905 | 97.5269 | 106.5367 | 89.5357 | 0.004331 | 0.97342 | 0.023261 | 23.8962 | 0.87386 | 0.15655 | 0.020677 | 1.446 | 1.446 | 0.00445
2019-03-20 12:52 | 1553100707.1707 | 124 | 0.2054 | 0.003618 | 4.7e-05 | 107.942 | 97.8978 | 116.5518 | 20.9914 | 0.0031 | 0.002822 | 165.4592 | 0.2644 | 8.4017 | 0.003496 | 0.005373 | 0.005195 | 5.9e-05 | 0.001135 | 0.32587 | 1.3214 | 2.5482 | 100.99 | 20.69 | 17.907 | 96.1374 | 105.7282 | 89.4481 | 0.004331 | 0.97366 | 0.023302 | 23.9322 | 0.87361 | 0.15729 | 0.020757 | 1.4541 | 1.4541 | 0.004446
2019-03-20 12:52 | 1553100708.1697 | 125 | 0.20541 | 0.003639 | 4.7e-05 | 108.0012 | 98.1367 | 116.4851 | 20.992 | 0.00306 | 0.002813 | 165.38 | 0.26344 | 8.4021 | 0.003584 | 0.005263 | 0.005187 | 5.2e-05 | 0.001131 | 0.31166 | 1.3067 | 2.548 | 100.99 | 20.69 | 17.909 | 95.1865 | 107.8741 | 89.236 | 0.004332 | 0.97423 | 0.023302 | 23.9184 | 0.85672 | 0.15672 | 0.020662 | 1.4475 | 1.4475 | 0.00445
2019-03-20 12:52 | 1553100709.1704 | 126 | 0.2054 | 0.003656 | 4.6e-05 | 107.9834 | 98.4032 | 116.8297 | 20.9915 | 0.003125 | 0.002875 | 165.3086 | 0.26358 | 8.4011 | 0.00365 | 0.004932 | 0.004808 | 5.6e-05 | 0.001135 | 0.36177 | 1.3262 | 2.5482 | 100.99 | 20.69 | 17.911 | 96.7532 | 109.4719 | 90.7326 | 0.004332 | 0.97022 | 0.023288 | 24.0032 | 0.83099 | 0.1568 | 0.02075 | 1.4528 | 1.4528 | 0.004436
2019-03-20 12:52 | 1553100710.1704 | 127 | 0.2054 | 0.003664 | 4.7e-05 | 107.9524 | 98.3359 | 117.1705 | 20.9925 | 0.003085 | 0.002814 | 165.2619 | 0.26273 | 8.403 | 0.003544 | 0.004972 | 0.004829 | 5.6e-05 | 0.00114 | 0.30254 | 1.33 | 2.5481 | 100.99 | 20.69 | 17.918 | 95.779 | 106.8879 | 89.2598 | 0.004334 | 0.97416 | 0.02337 | 23.9895 | 0.83454 | 0.1563 | 0.02072 | 1.4557 | 1.4557 | 0.004413
2019-03-20 12:52 | 1553100711.1694 | 128 | 0.20537 | 0.003677 | 4.7e-05 | 107.9724 | 98.5893 | 117.4031 | 20.9911 | 0.003174 | 0.002865 | 165.1965 | 0.26272 | 8.4013 | 0.003533 | 0.005032 | 0.004875 | 5.2e-05 | 0.001141 | 0.28744 | 1.3728 | 2.5482 | 100.99 | 20.69 | 17.92 | 97.9402 | 106.6172 | 90.4901 | 0.004334 | 0.97086 | 0.023353 | 24.054 | 0.81318 | 0.15629 | 0.020878 | 1.4658 | 1.4658 | 0.004382
2019-03-20 12:52 | 1553100712.1696 | 129 | 0.20539 | 0.003686 | 4.8e-05 | 108.0119 | 98.6356 | 117.4834 | 20.9911 | 0.003081 | 0.002859 | 165.108 | 0.26491 | 8.4012 | 0.003607 | 0.004974 | 0.004887 | 5.9e-05 | 0.001144 | 0.29103 | 1.3465 | 2.548 | 100.99 | 20.69 | 17.921 | 95.6783 | 108.4335 | 90.3307 | 0.004335 | 0.97129 | 0.023366 | 24.0569 | 0.81224 | 0.15759 | 0.020718 | 1.4553 | 1.4553 | 0.004451
2019-03-20 12:52 | 1553100713.1707 | 130 | 0.20539 | 0.003689 | 4.9e-05 | 107.9385 | 98.627 | 117.4097 | 20.9905 | 0.003182 | 0.002837 | 165.0428 | 0.26324 | 8.4 | 0.003592 | 0.005531 | 0.005354 | 5.3e-05 | 0.001146 | 0.33964 | 1.3512 | 2.5481 | 100.99 | 20.69 | 17.923 | 98.1363 | 108.0636 | 89.802 | 0.004335 | 0.97271 | 0.023376 | 24.0319 | 0.80694 | 0.1566 | 0.020725 | 1.4563 | 1.4563 | 0.00442
2019-03-20 12:52 | 1553100714.1707 | 131 | 0.20542 | 0.003702 | 4.9e-05 | 108.019 | 98.5691 | 117.2839 | 20.99 | 0.00313 | 0.002846 | 164.9936 | 0.26567 | 8.4027 | 0.003535 | 0.005001 | 0.004945 | 5.1e-05 | 0.001147 | 0.29319 | 1.3612 | 2.548 | 100.99 | 20.69 | 17.924 | 96.8781 | 106.6698 | 90.0194 | 0.004335 | 0.97212 | 0.023356 | 24.0262 | 0.81887 | 0.15805 | 0.020559 | 1.4435 | 1.4435 | 0.0045
2019-03-20 12:52 | 1553100715.17 | 132 | 0.20541 | 0.003696 | 5e-05 | 107.9631 | 98.2514 | 117.0585 | 20.9904 | 0.003118 | 0.002845 | 164.9142 | 0.26339 | 8.4013 | 0.0035 | 0.005353 | 0.005154 | 5.3e-05 | 0.001147 | 0.31349 | 1.3273 | 2.548 | 100.99 | 20.69 | 17.925 | 96.5783 | 105.8282 | 90.0035 | 0.004336 | 0.97217 | 0.023334 | 24.0026 | 0.84311 | 0.15669 | 0.020636 | 1.4475 | 1.4475 | 0.004449
2019-03-20 12:52 | 1553100716.1698 | 133 | 0.20539 | 0.003699 | 5e-05 | 108.0019 | 98.5074 | 116.9712 | 20.9897 | 0.00314 | 0.002844 | 164.854 | 0.26299 | 8.4033 | 0.003545 | 0.004715 | 0.004645 | 5.7e-05 | 0.001145 | 0.30596 | 1.3412 | 2.5479 | 100.99 | 20.69 | 17.928 | 97.112 | 106.9136 | 89.9777 | 0.004336 | 0.97223 | 0.023327 | 23.9928 | 0.82305 | 0.15646 | 0.020739 | 1.4542 | 1.4542 | 0.004422
2019-03-20 12:52 | 1553100717.1699 | 134 | 0.20541 | 0.003699 | 5e-05 | 107.9291 | 98.4671 | 116.7198 | 20.9904 | 0.003115 | 0.002871 | 164.7827 | 0.26365 | 8.4033 | 0.003531 | 0.005339 | 0.00524 | 5.9e-05 | 0.001148 | 0.33627 | 1.3326 | 2.548 | 100.99 | 20.69 | 17.93 | 96.5203 | 106.567 | 90.6232 | 0.004337 | 0.97051 | 0.023281 | 23.9883 | 0.82068 | 0.15685 | 0.020642 | 1.4446 | 1.4446 | 0.004462
2019-03-20 12:52 | 1553100718.1709 | 135 | 0.20544 | 0.003696 | 5.1e-05 | 107.9233 | 98.3904 | 116.6928 | 20.9903 | 0.003186 | 0.002861 | 164.7247 | 0.26445 | 8.4041 | 0.003587 | 0.005289 | 0.005193 | 5.4e-05 | 0.001144 | 0.31094 | 1.3086 | 2.548 | 100.99 | 20.68 | 17.935 | 98.2216 | 107.9293 | 90.3729 | 0.004335 | 0.97118 | 0.023286 | 23.9772 | 0.82716 | 0.15732 | 0.020441 | 1.4308 | 1.4308 | 0.004519
2019-03-20 12:52 | 1553100719.17 | 136 | 0.20544 | 0.003696 | 5.1e-05 | 107.9851 | 98.1995 | 116.3473 | 20.9894 | 0.003149 | 0.002846 | 164.6338 | 0.26471 | 8.4011 | 0.003656 | 0.00527 | 0.005096 | 5.4e-05 | 0.001142 | 0.29358 | 1.3191 | 2.548 | 100.99 | 20.69 | 17.938 | 97.3373 | 109.6231 | 90.029 | 0.004339 | 0.9721 | 0.023263 | 23.9304 | 0.84966 | 0.15747 | 0.020461 | 1.4307 | 1.4307 | 0.004524
2019-03-20 12:52 | 1553100720.171 | 137 | 0.20546 | 0.003701 | 5.2e-05 | 107.931 | 98.4962 | 116.1882 | 20.9908 | 0.003199 | 0.00287 | 164.5386 | 0.26412 | 8.4011 | 0.003636 | 0.005414 | 0.005257 | 5.4e-05 | 0.001142 | 0.32333 | 1.3031 | 2.5482 | 100.99 | 20.68 | 17.944 | 98.5389 | 109.1189 | 90.6086 | 0.004338 | 0.97055 | 0.023228 | 23.9331 | 0.8182 | 0.15712 | 0.020323 | 1.4189 | 1.4189 | 0.004551
2019-03-20 12:52 | 1553100721.17 | 138 | 0.2055 | 0.003696 | 5.2e-05 | 107.987 | 98.8625 | 116.0436 | 20.9897 | 0.003153 | 0.002915 | 164.4916 | 0.26513 | 8.4017 | 0.003679 | 0.004974 | 0.004835 | 5e-05 | 0.001135 | 0.31022 | 1.3429 | 2.5481 | 100.99 | 20.68 | 17.946 | 97.4168 | 110.169 | 91.6933 | 0.004338 | 0.96766 | 0.023179 | 23.9538 | 0.78963 | 0.15773 | 0.020084 | 1.3993 | 1.3993 | 0.004633
2019-03-20 12:52 | 1553100722.1705 | 139 | 0.20544 | 0.003695 | 5.3e-05 | 107.9414 | 98.8341 | 116.1361 | 20.9899 | 0.003181 | 0.002871 | 164.4506 | 0.26413 | 8.403 | 0.003616 | 0.00466 | 0.004587 | 5.4e-05 | 0.001142 | 0.3063 | 1.324 | 2.5479 | 100.99 | 20.68 | 17.949 | 98.1084 | 108.6419 | 90.6297 | 0.004339 | 0.97049 | 0.023222 | 23.9284 | 0.78843 | 0.15713 | 0.020433 | 1.4262 | 1.4262 | 0.004528
2019-03-20 12:52 | 1553100723.1692 | 140 | 0.20551 | 0.003697 | 5.3e-05 | 107.9071 | 99.3734 | 115.8281 | 20.9898 | 0.003178 | 0.002859 | 164.3882 | 0.264 | 8.4034 | 0.003687 | 0.005047 | 0.004967 | 5.5e-05 | 0.001136 | 0.31454 | 1.3342 | 2.5481 | 100.99 | 20.68 | 17.955 | 98.0385 | 110.3638 | 90.3287 | 0.00434 | 0.9713 | 0.023201 | 23.8868 | 0.73691 | 0.15706 | 0.020037 | 1.3972 | 1.3972 | 0.00462
2019-03-20 12:52 | 1553100724.1692 | 141 | 0.20544 | 0.003691 | 5.3e-05 | 107.9391 | 99.6167 | 115.6577 | 20.9897 | 0.003172 | 0.002909 | 164.2985 | 0.26412 | 8.401 | 0.003647 | 0.004158 | 0.004028 | 5.4e-05 | 0.001139 | 0.32964 | 1.3436 | 2.5481 | 100.99 | 20.68 | 17.958 | 97.8746 | 109.4024 | 91.5449 | 0.004341 | 0.96806 | 0.023145 | 23.9091 | 0.71769 | 0.15712 | 0.020412 | 1.42 | 1.42 | 0.004548
2019-03-20 12:52 | 1553100725.1692 | 142 | 0.20545 | 0.003692 | 5.2e-05 | 107.9621 | 99.9542 | 115.7805 | 20.9894 | 0.003162 | 0.002873 | 164.2514 | 0.26431 | 8.4022 | 0.003616 | 0.00579 | 0.005645 | 5.3e-05 | 0.001145 | 0.33284 | 1.3481 | 2.548 | 100.99 | 20.68 | 17.96 | 97.6382 | 108.6337 | 90.6612 | 0.004341 | 0.97041 | 0.023186 | 23.8928 | 0.68934 | 0.15724 | 0.020385 | 1.4207 | 1.4207 | 0.004549
2019-03-20 12:52 | 1553100726.1697 | 143 | 0.20545 | 0.003686 | 5.2e-05 | 108.0103 | 100.334 | 115.5026 | 20.9889 | 0.003221 | 0.00285 | 164.1539 | 0.26362 | 8.4042 | 0.003692 | 0.004846 | 0.004751 | 5.2e-05 | 0.001145 | 0.30093 | 1.3242 | 2.5481 | 100.99 | 20.68 | 17.966 | 99.0685 | 110.4797 | 90.1113 | 0.004343 | 0.97188 | 0.023175 | 23.8461 | 0.65941 | 0.15683 | 0.020369 | 1.4189 | 1.4189 | 0.004543
2019-03-20 12:52 | 1553100727.1694 | 144 | 0.20543 | 0.00368 | 5.3e-05 | 108.0028 | 100.199 | 115.6277 | 20.9891 | 0.003194 | 0.002877 | 164.0911 | 0.26364 | 8.4029 | 0.003636 | 0.005023 | 0.004992 | 4.9e-05 | 0.001143 | 0.31492 | 1.3763 | 2.548 | 100.99 | 20.68 | 17.97 | 98.4189 | 109.1219 | 90.7672 | 0.004344 | 0.97013 | 0.023167 | 23.8805 | 0.67083 | 0.15684 | 0.020536 | 1.4299 | 1.4299 | 0.004508
2019-03-20 12:52 | 1553100728.1704 | 145 | 0.20552 | 0.003675 | 5.4e-05 | 107.9178 | 99.949 | 115.6929 | 20.9889 | 0.003183 | 0.002919 | 164.014 | 0.26321 | 8.4032 | 0.00367 | 0.005064 | 0.004879 | 5.5e-05 | 0.001143 | 0.31421 | 1.3394 | 2.5481 | 100.99 | 20.68 | 17.974 | 98.149 | 109.9546 | 91.7848 | 0.004345 | 0.96742 | 0.023141 | 23.9206 | 0.68613 | 0.15659 | 0.019979 | 1.3896 | 1.3896 | 0.004631
2019-03-20 12:52 | 1553100729.1709 | 146 | 0.20545 | 0.00367 | 5.5e-05 | 107.932 | 100.3586 | 115.896 | 20.9886 | 0.003229 | 0.002867 | 163.947 | 0.263 | 8.4047 | 0.003539 | 0.004414 | 0.004196 | 5.4e-05 | 0.001146 | 0.31902 | 1.3358 | 2.548 | 100.99 | 20.68 | 17.976 | 99.2635 | 106.7613 | 90.5299 | 0.004345 | 0.97076 | 0.023202 | 23.9004 | 0.65073 | 0.15646 | 0.020385 | 1.4214 | 1.4214 | 0.004524
2019-03-20 12:52 | 1553100730.1706 | 147 | 0.20546 | 0.00367 | 5.5e-05 | 107.868 | 100.2499 | 115.7225 | 20.9884 | 0.00319 | 0.002891 | 163.8758 | 0.26437 | 8.4041 | 0.00359 | 0.005542 | 0.005456 | 5.6e-05 | 0.001146 | 0.3298 | 1.3426 | 2.548 | 100.99 | 20.68 | 17.975 | 98.32 | 108.0022 | 91.112 | 0.004345 | 0.96921 | 0.023166 | 23.9016 | 0.65512 | 0.15728 | 0.02034 | 1.4161 | 1.4161 | 0.004565
2019-03-20 12:52 | 1553100731.1697 | 148 | 0.20549 | 0.003676 | 5.5e-05 | 107.8874 | 99.8384 | 115.8983 | 20.988 | 0.003141 | 0.00289 | 163.7877 | 0.26346 | 8.4045 | 0.003524 | 0.004913 | 0.004815 | 5.7e-05 | 0.001148 | 0.33019 | 1.3302 | 2.5481 | 100.99 | 20.67 | 17.976 | 97.1267 | 106.3963 | 91.0872 | 0.004343 | 0.96927 | 0.023184 | 23.9189 | 0.69352 | 0.15673 | 0.020131 | 1.4027 | 1.4027 | 0.004592
2019-03-20 12:52 | 1553100732.1701 | 149 | 0.20552 | 0.00367 | 5.4e-05 | 107.9395 | 100.1945 | 115.8554 | 20.988 | 0.003188 | 0.002831 | 163.7234 | 0.26247 | 8.4037 | 0.003582 | 0.00548 | 0.005314 | 5.4e-05 | 0.001139 | 0.28854 | 1.3637 | 2.548 | 100.99 | 20.67 | 17.977 | 98.2602 | 107.8175 | 89.6564 | 0.004343 | 0.9731 | 0.023225 | 23.8675 | 0.66598 | 0.15615 | 0.019986 | 1.3951 | 1.3951 | 0.0046
2019-03-20 12:52 | 1553100733.1697 | 150 | 0.20557 | 0.003675 | 5.3e-05 | 107.9294 | 100.4162 | 116.0964 | 20.9879 | 0.003243 | 0.002931 | 163.6453 | 0.26419 | 8.4041 | 0.003604 | 0.005064 | 0.005052 | 5.4e-05 | 0.001139 | 0.33301 | 1.2894 | 2.5479 | 100.99 | 20.67 | 17.976 | 99.5938 | 108.3405 | 92.0576 | 0.004343 | 0.9667 | 0.023173 | 23.9712 | 0.64539 | 0.15716 | 0.019674 | 1.3702 | 1.3702 | 0.004714
2019-03-20 12:52 | 1553100734.1697 | 151 | 0.2056 | 0.003665 | 5.3e-05 | 107.9631 | 100.4648 | 116.2089 | 20.9885 | 0.003162 | 0.002864 | 163.5862 | 0.26446 | 8.402 | 0.003676 | 0.00494 | 0.004817 | 5.6e-05 | 0.001143 | 0.33157 | 1.3091 | 2.5478 | 100.99 | 20.67 | 17.975 | 97.6453 | 110.1007 | 90.4532 | 0.004342 | 0.97096 | 0.023235 | 23.9301 | 0.64385 | 0.15733 | 0.019496 | 1.3615 | 1.3615 | 0.004749
2019-03-20 12:52 | 1553100735.1703 | 152 | 0.20563 | 0.003656 | 5.3e-05 | 107.9026 | 100.5092 | 116.0739 | 20.9876 | 0.003195 | 0.002873 | 163.506 | 0.26353 | 8.4016 | 0.003601 | 0.005544 | 0.005372 | 5.5e-05 | 0.001145 | 0.31813 | 1.2818 | 2.548 | 100.99 | 20.67 | 17.973 | 98.4425 | 108.2703 | 90.6742 | 0.004342 | 0.97037 | 0.023215 | 23.9235 | 0.63488 | 0.15677 | 0.019339 | 1.3494 | 1.3494 | 0.004775
2019-03-20 12:52 | 1553100736.1713 | 153 | 0.20562 | 0.00364 | 5.5e-05 | 107.9443 | 100.8292 | 115.984 | 20.9893 | 0.003207 | 0.00293 | 163.4487 | 0.26318 | 8.4017 | 0.003608 | 0.005664 | 0.00558 | 5.4e-05 | 0.001145 | 0.33915 | 1.2719 | 2.548 | 100.99 | 20.67 | 17.97 | 98.7254 | 108.4555 | 92.0426 | 0.004341 | 0.96674 | 0.023162 | 23.9591 | 0.60991 | 0.15657 | 0.019432 | 1.3527 | 1.3527 | 0.004757
2019-03-20 12:52 | 1553100737.1699 | 154 | 0.20559 | 0.00362 | 5.6e-05 | 107.9747 | 101.3539 | 115.9178 | 20.988 | 0.003175 | 0.002906 | 163.3795 | 0.26358 | 8.4031 | 0.003684 | 0.004971 | 0.004863 | 5.6e-05 | 0.00114 | 0.30718 | 1.3098 | 2.548 | 100.99 | 20.67 | 17.969 | 97.9497 | 110.2949 | 91.462 | 0.004341 | 0.96828 | 0.023174 | 23.9333 | 0.566 | 0.1568 | 0.019616 | 1.3662 | 1.3662 | 0.004717
2019-03-20 12:52 | 1553100738.1703 | 155 | 0.20561 | 0.003605 | 5.7e-05 | 107.987 | 101.6903 | 115.7655 | 20.9877 | 0.003186 | 0.00287 | 163.3144 | 0.26335 | 8.4028 | 0.00368 | 0.005192 | 0.005041 | 5.8e-05 | 0.001135 | 0.32084 | 1.3242 | 2.5481 | 100.99 | 20.67 | 17.967 | 98.2205 | 110.1945 | 90.6008 | 0.00434 | 0.97057 | 0.023186 | 23.8893 | 0.53738 | 0.15667 | 0.019501 | 1.3588 | 1.3588 | 0.004739
2019-03-20 12:52 | 1553100739.17 | 156 | 0.20557 | 0.003586 | 5.8e-05 | 107.9508 | 100.9634 | 115.5622 | 20.988 | 0.003167 | 0.002871 | 163.2496 | 0.26335 | 8.402 | 0.003731 | 0.004938 | 0.004816 | 5.5e-05 | 0.001135 | 0.33323 | 1.3429 | 2.548 | 100.99 | 20.67 | 17.967 | 97.7685 | 111.4258 | 90.6341 | 0.00434 | 0.97048 | 0.023165 | 23.8694 | 0.59855 | 0.15667 | 0.019801 | 1.3783 | 1.3783 | 0.004672
2019-03-20 12:52 | 1553100740.1701 | 157 | 0.2056 | 0.003573 | 6e-05 | 107.9547 | 101.025 | 116.1668 | 20.988 | 0.003203 | 0.002867 | 163.1848 | 0.26377 | 8.4027 | 0.003667 | 0.005734 | 0.005609 | 5.4e-05 | 0.001131 | 0.29767 | 1.3379 | 2.5481 | 100.99 | 20.67 | 17.967 | 98.6333 | 109.8802 | 90.5213 | 0.00434 | 0.97078 | 0.023229 | 23.928 | 0.59341 | 0.15692 | 0.019616 | 1.3691 | 1.3691 | 0.004711
2019-03-20 12:52 | 1553100741.1695 | 158 | 0.20561 | 0.003569 | 6e-05 | 107.9327 | 101.0418 | 116.3091 | 20.9877 | 0.003165 | 0.002921 | 163.1037 | 0.26377 | 8.4028 | 0.003606 | 0.005138 | 0.004981 | 5.3e-05 | 0.001133 | 0.3413 | 1.3384 | 2.548 | 100.99 | 20.67 | 17.97 | 97.7119 | 108.4017 | 91.8166 | 0.004341 | 0.96734 | 0.023202 | 23.9853 | 0.5901 | 0.15692 | 0.019584 | 1.3653 | 1.3653 | 0.004724
2019-03-20 12:52 | 1553100742.1702 | 159 | 0.20561 | 0.003567 | 6.1e-05 | 107.9158 | 101.3461 | 116.2441 | 20.9882 | 0.003211 | 0.002918 | 163.0187 | 0.26409 | 8.4021 | 0.003589 | 0.005301 | 0.00515 | 5.7e-05 | 0.001131 | 0.32156 | 1.3206 | 2.5481 | 100.99 | 20.67 | 17.972 | 98.8365 | 107.9784 | 91.753 | 0.004342 | 0.9675 | 0.023197 | 23.9765 | 0.56181 | 0.15711 | 0.019546 | 1.3623 | 1.3623 | 0.00474
2019-03-20 12:52 | 1553100743.1702 | 160 | 0.20561 | 0.003559 | 6e-05 | 107.8984 | 100.6946 | 116.3325 | 20.9877 | 0.003259 | 0.002971 | 162.9645 | 0.2645 | 8.4045 | 0.003615 | 0.005568 | 0.005419 | 5.8e-05 | 0.001132 | 0.34164 | 1.2826 | 2.548 | 100.99 | 20.67 | 17.974 | 99.9788 | 108.6256 | 93.0335 | 0.004342 | 0.96412 | 0.023166 | 24.0276 | 0.61805 | 0.15735 | 0.019553 | 1.361 | 1.361 | 0.004752
2019-03-20 12:52 | 1553100744.1707 | 161 | 0.20565 | 0.003561 | 6e-05 | 107.9621 | 100.9865 | 116.4295 | 20.9875 | 0.00327 | 0.002884 | 162.8899 | 0.26404 | 8.4 | 0.003573 | 0.005087 | 0.004962 | 5.6e-05 | 0.001129 | 0.33229 | 1.2986 | 2.548 | 100.99 | 20.67 | 17.981 | 100.247 | 107.5973 | 90.936 | 0.004344 | 0.96968 | 0.023242 | 23.9687 | 0.59743 | 0.15707 | 0.019351 | 1.3514 | 1.3514 | 0.004777
2019-03-20 12:52 | 1553100745.1694 | 162 | 0.20569 | 0.003564 | 6.1e-05 | 107.9427 | 101.3464 | 116.6156 | 20.9878 | 0.003283 | 0.002901 | 162.8122 | 0.26317 | 8.4033 | 0.003627 | 0.004877 | 0.004786 | 5.5e-05 | 0.001135 | 0.3183 | 1.3193 | 2.5481 | 100.99 | 20.67 | 17.985 | 100.5784 | 108.8937 | 91.3396 | 0.004345 | 0.9686 | 0.023248 | 24.0012 | 0.564 | 0.15656 | 0.019117 | 1.3353 | 1.3353 | 0.004819
2019-03-20 12:52 | 1553100746.1694 | 163 | 0.20566 | 0.003557 | 6.3e-05 | 107.9378 | 101.339 | 116.8649 | 20.9878 | 0.003288 | 0.00296 | 162.7225 | 0.26408 | 8.4017 | 0.003586 | 0.005344 | 0.005189 | 5.6e-05 | 0.00113 | 0.31996 | 1.3262 | 2.548 | 100.99 | 20.68 | 17.995 | 100.6861 | 107.9023 | 92.7715 | 0.00435 | 0.96481 | 0.023227 | 24.0739 | 0.56426 | 0.1571 | 0.019253 | 1.3435 | 1.3435 | 0.004806
2019-03-20 12:52 | 1553100747.1705 | 164 | 0.20568 | 0.003548 | 6.4e-05 | 107.9935 | 101.3173 | 117.083 | 20.9878 | 0.003256 | 0.002898 | 162.6423 | 0.26365 | 8.4045 | 0.003578 | 0.005182 | 0.005095 | 5.7e-05 | 0.001132 | 0.30807 | 1.3392 | 2.548 | 100.99 | 20.68 | 18.001 | 99.903 | 107.7188 | 91.2623 | 0.004351 | 0.96881 | 0.023297 | 24.0467 | 0.57078 | 0.15684 | 0.019177 | 1.3422 | 1.3422 | 0.004803
2019-03-20 12:52 | 1553100748.1696 | 165 | 0.20569 | 0.003545 | 6.3e-05 | 107.943 | 101.0777 | 117.2719 | 20.9881 | 0.003275 | 0.002962 | 162.5588 | 0.26438 | 8.4024 | 0.003568 | 0.005755 | 0.005634 | 5.8e-05 | 0.001133 | 0.31957 | 1.3127 | 2.548 | 100.99 | 20.68 | 18.006 | 100.3794 | 107.4624 | 92.8156 | 0.004353 | 0.96469 | 0.023266 | 24.1173 | 0.58778 | 0.15728 | 0.01913 | 1.3371 | 1.3371 | 0.004834
2019-03-20 12:52 | 1553100749.1706 | 166 | 0.20576 | 0.003539 | 6.3e-05 | 107.9501 | 100.7384 | 117.4106 | 20.9882 | 0.003172 | 0.002894 | 162.4822 | 0.26402 | 8.4026 | 0.003635 | 0.005262 | 0.005185 | 5.4e-05 | 0.001128 | 0.33069 | 1.3684 | 2.5479 | 100.99 | 20.67 | 18.016 | 97.8902 | 109.1083 | 91.1713 | 0.004352 | 0.96905 | 0.023332 | 24.0773 | 0.61845 | 0.15706 | 0.018682 | 1.3095 | 1.3095 | 0.004929
2019-03-20 12:52 | 1553100750.1706 | 167 | 0.20569 | 0.003528 | 6.1e-05 | 107.9472 | 100.6184 | 117.2435 | 20.9878 | 0.003251 | 0.002929 | 162.4167 | 0.26419 | 8.4021 | 0.003625 | 0.00527 | 0.005113 | 5.9e-05 | 0.00113 | 0.30486 | 1.2991 | 2.548 | 100.99 | 20.67 | 18.02 | 99.7915 | 108.8614 | 92.0143 | 0.004353 | 0.96681 | 0.023289 | 24.088 | 0.62887 | 0.15717 | 0.019122 | 1.338 | 1.338 | 0.004828
2019-03-20 12:52 | 1553100751.1709 | 168 | 0.2057 | 0.003528 | 6.1e-05 | 107.9259 | 100.8355 | 117.2207 | 20.9875 | 0.003205 | 0.002959 | 162.3477 | 0.26359 | 8.4023 | 0.003641 | 0.004529 | 0.004499 | 5.4e-05 | 0.001128 | 0.34866 | 1.2783 | 2.548 | 100.99 | 20.67 | 18.027 | 98.6768 | 109.2425 | 92.736 | 0.004355 | 0.96491 | 0.023263 | 24.1094 | 0.60782 | 0.15681 | 0.019089 | 1.3342 | 1.3342 | 0.00483
2019-03-20 12:53 | 1553100752.1701 | 169 | 0.20573 | 0.003522 | 6e-05 | 107.8887 | 100.9779 | 117.2977 | 20.9883 | 0.003221 | 0.002947 | 162.2814 | 0.26383 | 8.4055 | 0.003619 | 0.00501 | 0.004843 | 5.3e-05 | 0.001126 | 0.29214 | 1.3208 | 2.548 | 100.99 | 20.67 | 18.03 | 99.0571 | 108.7232 | 92.463 | 0.004356 | 0.96563 | 0.02328 | 24.1083 | 0.59211 | 0.15695 | 0.018903 | 1.3222 | 1.3222 | 0.004879
2019-03-20 12:53 | 1553100753.1698 | 170 | 0.20573 | 0.00352 | 6e-05 | 107.8673 | 101.2052 | 117.4491 | 20.9873 | 0.003187 | 0.002917 | 162.2139 | 0.26413 | 8.4044 | 0.003627 | 0.005535 | 0.005462 | 5.7e-05 | 0.001125 | 0.32781 | 1.3381 | 2.5481 | 100.99 | 20.67 | 18.032 | 98.2344 | 108.916 | 91.7403 | 0.004356 | 0.96754 | 0.023318 | 24.1 | 0.57024 | 0.15713 | 0.018884 | 1.3231 | 1.3231 | 0.004881
2019-03-20 12:53 | 1553100754.1691 | 171 | 0.20573 | 0.003513 | 5.8e-05 | 107.9252 | 101.179 | 117.7531 | 20.9883 | 0.003232 | 0.002958 | 162.1357 | 0.2629 | 8.4065 | 0.00361 | 0.005652 | 0.005553 | 5.8e-05 | 0.001131 | 0.31045 | 1.3055 | 2.548 | 100.99 | 20.67 | 18.035 | 99.323 | 108.4954 | 92.7173 | 0.004357 | 0.96496 | 0.023317 | 24.1635 | 0.57735 | 0.1564 | 0.01894 | 1.327 | 1.327 | 0.004844
2019-03-20 12:53 | 1553100755.1701 | 172 | 0.20574 | 0.003499 | 5.7e-05 | 107.91 | 100.7421 | 117.7238 | 20.9877 | 0.003255 | 0.002916 | 162.0641 | 0.26349 | 8.4039 | 0.003693 | 0.00591 | 0.005811 | 6e-05 | 0.001128 | 0.3194 | 1.3435 | 2.5481 | 100.99 | 20.67 | 18.037 | 99.8934 | 110.5125 | 91.7062 | 0.004357 | 0.96763 | 0.023346 | 24.1271 | 0.61479 | 0.15675 | 0.018879 | 1.3245 | 1.3245 | 0.004864
2019-03-20 12:53 | 1553100756.1689 | 173 | 0.20574 | 0.003496 | 5.6e-05 | 107.9524 | 100.5208 | 118.0064 | 20.9878 | 0.003246 | 0.002938 | 162.0011 | 0.2646 | 8.404 | 0.003637 | 0.005205 | 0.005076 | 5.6e-05 | 0.001124 | 0.3413 | 1.2923 | 2.5479 | 100.99 | 20.67 | 18.039 | 99.6652 | 109.1483 | 92.2366 | 0.004358 | 0.96622 | 0.023357 | 24.1735 | 0.63797 | 0.15741 | 0.018895 | 1.3263 | 1.3263 | 0.004878
2019-03-20 12:53 | 1553100757.1701 | 174 | 0.20575 | 0.003486 | 5.4e-05 | 107.8563 | 100.5664 | 118.0042 | 20.9873 | 0.003229 | 0.002896 | 161.9213 | 0.26421 | 8.4025 | 0.003628 | 0.005237 | 0.005127 | 5.4e-05 | 0.001123 | 0.31598 | 1.274 | 2.5481 | 100.99 | 20.67 | 18.045 | 99.2577 | 108.9287 | 91.2203 | 0.004359 | 0.96892 | 0.023389 | 24.1397 | 0.62595 | 0.15718 | 0.01882 | 1.323 | 1.323 | 0.004883
2019-03-20 12:53 | 1553100758.1691 | 175 | 0.20576 | 0.00348 | 5.4e-05 | 107.9036 | 100.9241 | 117.9273 | 20.9878 | 0.003275 | 0.002935 | 161.8507 | 0.26289 | 8.405 | 0.003647 | 0.005064 | 0.005021 | 5.8e-05 | 0.001125 | 0.30773 | 1.3116 | 2.548 | 100.99 | 20.67 | 18.048 | 100.3728 | 109.391 | 92.1561 | 0.00436 | 0.96644 | 0.023352 | 24.1628 | 0.59811 | 0.15639 | 0.018802 | 1.3196 | 1.3196 | 0.004871
2019-03-20 12:53 | 1553100759.1701 | 176 | 0.2058 | 0.003478 | 5.4e-05 | 107.89 | 100.9189 | 117.7336 | 20.9876 | 0.00323 | 0.002893 | 161.7908 | 0.26304 | 8.4064 | 0.003602 | 0.004867 | 0.004743 | 5.6e-05 | 0.001122 | 0.33339 | 1.3364 | 2.5481 | 100.99 | 20.67 | 18.056 | 99.2909 | 108.3089 | 91.1573 | 0.004362 | 0.96909 | 0.023365 | 24.1099 | 0.59745 | 0.15648 | 0.018532 | 1.3013 | 1.3013 | 0.004942
2019-03-20 12:53 | 1553100760.1695 | 177 | 0.2058 | 0.003469 | 5.4e-05 | 107.9301 | 101.0848 | 117.4217 | 20.9871 | 0.003254 | 0.002984 | 161.7065 | 0.26405 | 8.4049 | 0.003615 | 0.004551 | 0.004357 | 5.6e-05 | 0.001123 | 0.32139 | 1.3077 | 2.5481 | 100.99 | 20.67 | 18.06 | 99.8582 | 108.6022 | 93.3497 | 0.004363 | 0.96329 | 0.023264 | 24.1503 | 0.58608 | 0.15708 | 0.018547 | 1.2967 | 1.2967 | 0.004979
2019-03-20 12:53 | 1553100761.1695 | 178 | 0.20583 | 0.003462 | 5.4e-05 | 107.9165 | 101.1652 | 117.205 | 20.9875 | 0.003226 | 0.002966 | 161.6263 | 0.26309 | 8.4011 | 0.003603 | 0.005302 | 0.00522 | 5.7e-05 | 0.001124 | 0.32477 | 1.3128 | 2.548 | 100.99 | 20.67 | 18.06 | 99.1982 | 108.3236 | 92.9019 | 0.004363 | 0.96447 | 0.023256 | 24.1132 | 0.57785 | 0.15651 | 0.018387 | 1.2852 | 1.2852 | 0.005005
2019-03-20 12:53 | 1553100762.1693 | 179 | 0.20584 | 0.003449 | 5.3e-05 | 107.9307 | 100.6857 | 117.1811 | 20.9879 | 0.003303 | 0.002919 | 161.5502 | 0.26309 | 8.4034 | 0.003588 | 0.00556 | 0.005477 | 5.8e-05 | 0.001129 | 0.35225 | 1.2977 | 2.5481 | 100.99 | 20.68 | 18.064 | 101.0551 | 107.968 | 91.7691 | 0.004367 | 0.96746 | 0.02329 | 24.0734 | 0.62152 | 0.15651 | 0.018344 | 1.284 | 1.284 | 0.00501
2019-03-20 12:53 | 1553100763.1703 | 180 | 0.20585 | 0.003432 | 5.3e-05 | 107.9812 | 101.1267 | 117.1966 | 20.9868 | 0.003238 | 0.002959 | 161.4589 | 0.2631 | 8.4023 | 0.003611 | 0.004667 | 0.004544 | 5.7e-05 | 0.001125 | 0.32621 | 1.3241 | 2.5481 | 100.99 | 20.68 | 18.08 | 99.476 | 108.524 | 92.7518 | 0.00437 | 0.96486 | 0.02326 | 24.1074 | 0.58661 | 0.15652 | 0.018295 | 1.279 | 1.279 | 0.00503
2019-03-20 12:53 | 1553100764.1697 | 181 | 0.20593 | 0.003417 | 5.2e-05 | 107.9779 | 101.0811 | 117.2987 | 20.9874 | 0.003175 | 0.002946 | 161.3921 | 0.26456 | 8.4034 | 0.003634 | 0.005114 | 0.00499 | 5.7e-05 | 0.001127 | 0.32405 | 1.2846 | 2.5481 | 100.99 | 20.68 | 18.086 | 97.9593 | 109.0722 | 92.4279 | 0.004372 | 0.96572 | 0.023281 | 24.1073 | 0.59037 | 0.15739 | 0.017859 | 1.2496 | 1.2496 | 0.005177
2019-03-20 12:53 | 1553100765.1695 | 182 | 0.20586 | 0.003406 | 5.1e-05 | 107.999 | 101.1106 | 117.3043 | 20.9866 | 0.003241 | 0.002888 | 161.3052 | 0.26365 | 8.4013 | 0.003631 | 0.004885 | 0.004852 | 5.5e-05 | 0.001127 | 0.30453 | 1.2859 | 2.5479 | 100.99 | 20.68 | 18.093 | 99.5558 | 108.9951 | 91.0312 | 0.004374 | 0.96942 | 0.023326 | 24.0617 | 0.5895 | 0.15684 | 0.018289 | 1.2823 | 1.2823 | 0.005027
2019-03-20 12:53 | 1553100766.1705 | 183 | 0.20628 | 0.00339 | 5e-05 | 107.9611 | 100.9387 | 116.9609 | 20.9875 | 0.003276 | 0.00294 | 161.2339 | 0.26295 | 8.4053 | 0.003643 | 0.005474 | 0.005358 | 5.8e-05 | 0.00113 | 0.30486 | 1.3377 | 2.5481 | 100.99 | 20.68 | 18.106 | 100.4079 | 109.3018 | 92.291 | 0.004377 | 0.96608 | 0.023252 | 24.068 | 0.60159 | 0.15643 | 0.01577 | 1.1021 | 1.1021 | 0.005834
2019-03-20 12:53 | 1553100767.1699 | 184 | 0.20595 | 0.003382 | 4.9e-05 | 108.0032 | 100.7443 | 116.7809 | 20.9865 | 0.003285 | 0.002928 | 161.1783 | 0.26331 | 8.4023 | 0.003623 | 0.004561 | 0.004422 | 5.6e-05 | 0.001131 | 0.32565 | 1.3778 | 2.5481 | 100.99 | 20.68 | 18.113 | 100.6125 | 108.8127 | 91.9842 | 0.004378 | 0.96689 | 0.023243 | 24.0394 | 0.62231 | 0.15665 | 0.017773 | 1.2418 | 1.2418 | 0.005185
2019-03-20 12:53 | 1553100768.1706 | 185 | 0.20604 | 0.003365 | 4.8e-05 | 108.0472 | 100.6804 | 116.8102 | 20.9871 | 0.003267 | 0.002925 | 161.1208 | 0.26279 | 8.4034 | 0.003649 | 0.005139 | 0.005011 | 5.7e-05 | 0.001134 | 0.3235 | 1.3367 | 2.548 | 100.99 | 20.68 | 18.125 | 100.1697 | 109.4496 | 91.9216 | 0.004381 | 0.96706 | 0.023248 | 24.0403 | 0.63163 | 0.15633 | 0.017262 | 1.2063 | 1.2063 | 0.005326
2019-03-20 12:53 | 1553100769.1697 | 186 | 0.20608 | 0.00335 | 4.8e-05 | 108.0624 | 100.2648 | 117.0332 | 20.9861 | 0.003193 | 0.002909 | 161.0371 | 0.26342 | 8.4043 | 0.003604 | 0.00549 | 0.005368 | 5.5e-05 | 0.001134 | 0.29662 | 1.3461 | 2.5482 | 100.99 | 20.68 | 18.129 | 98.3908 | 108.3569 | 91.5437 | 0.004382 | 0.96806 | 0.023283 | 24.0508 | 0.66989 | 0.15671 | 0.017009 | 1.1904 | 1.1904 | 0.005411
2019-03-20 12:53 | 1553100770.1691 | 187 | 0.20614 | 0.003327 | 4.8e-05 | 108.0679 | 100.5597 | 117.2726 | 20.9867 | 0.003235 | 0.002855 | 160.9727 | 0.26301 | 8.4052 | 0.003585 | 0.005236 | 0.00517 | 5.6e-05 | 0.001134 | 0.30326 | 1.3257 | 2.5481 | 100.99 | 20.68 | 18.133 | 99.4147 | 107.8794 | 90.2327 | 0.004383 | 0.97155 | 0.023348 | 24.0321 | 0.64408 | 0.15646 | 0.016713 | 1.1731 | 1.1731 | 0.005482
2019-03-20 12:53 | 1553100771.1688 | 188 | 0.2061 | 0.003304 | 4.7e-05 | 108.0245 | 101.222 | 117.3466 | 20.9864 | 0.003226 | 0.002928 | 160.9108 | 0.26322 | 8.4013 | 0.003546 | 0.00481 | 0.004683 | 5.5e-05 | 0.001132 | 0.32477 | 1.341 | 2.5481 | 100.99 | 20.68 | 18.138 | 99.1884 | 106.9377 | 92.0067 | 0.004384 | 0.96683 | 0.023299 | 24.0983 | 0.58177 | 0.15659 | 0.016997 | 1.1905 | 1.1905 | 0.005406
2019-03-20 12:53 | 1553100772.1688 | 189 | 0.20611 | 0.003279 | 4.8e-05 | 108.0734 | 101.9706 | 117.5996 | 20.987 | 0.00322 | 0.002891 | 160.8279 | 0.2628 | 8.403 | 0.003533 | 0.005143 | 0.005053 | 5.5e-05 | 0.00113 | 0.31006 | 1.3331 | 2.5482 | 100.99 | 20.68 | 18.14 | 99.0502 | 106.6149 | 91.0998 | 0.004385 | 0.96924 | 0.023353 | 24.0943 | 0.51991 | 0.15634 | 0.016935 | 1.1888 | 1.1888 | 0.005405
2019-03-20 12:53 | 1553100773.1692 | 190 | 0.20616 | 0.003248 | 4.8e-05 | 108.0556 | 102.2793 | 117.5111 | 20.9867 | 0.003219 | 0.002928 | 160.7691 | 0.26253 | 8.4021 | 0.003545 | 0.005339 | 0.005231 | 5.6e-05 | 0.001131 | 0.34092 | 1.2939 | 2.548 | 100.99 | 20.68 | 18.139 | 99.0258 | 106.9163 | 91.9948 | 0.004385 | 0.96686 | 0.023316 | 24.1148 | 0.4914 | 0.15618 | 0.016698 | 1.1703 | 1.1703 | 0.005485
2019-03-20 12:53 | 1553100774.1689 | 191 | 0.20612 | 0.003236 | 4.8e-05 | 108.1225 | 102.4744 | 117.2907 | 20.9859 | 0.003221 | 0.002945 | 160.6887 | 0.26271 | 8.4026 | 0.003581 | 0.005023 | 0.004871 | 5.8e-05 | 0.001132 | 0.32532 | 1.2882 | 2.5479 | 100.99 | 20.68 | 18.135 | 99.0687 | 107.7941 | 92.3989 | 0.004384 | 0.96579 | 0.023281 | 24.1055 | 0.47989 | 0.15629 | 0.016911 | 1.1835 | 1.1835 | 0.005428
2019-03-20 12:53 | 1553100775.1689 | 192 | 0.20616 | 0.00321 | 4.7e-05 | 108.2137 | 102.2827 | 117.0163 | 20.9852 | 0.003297 | 0.002903 | 160.6233 | 0.26401 | 8.4021 | 0.003618 | 0.005645 | 0.005555 | 5.9e-05 | 0.001131 | 0.30574 | 1.3088 | 2.5481 | 100.99 | 20.68 | 18.132 | 100.8945 | 108.6885 | 91.3847 | 0.004383 | 0.96848 | 0.023286 | 24.0438 | 0.50418 | 0.15706 | 0.016701 | 1.1691 | 1.1691 | 0.005521
2019-03-20 12:53 | 1553100776.17 | 193 | 0.20624 | 0.003189 | 4.6e-05 | 108.2195 | 102.1201 | 116.9067 | 20.9852 | 0.003303 | 0.002943 | 160.5529 | 0.26396 | 8.3992 | 0.003567 | 0.00566 | 0.005573 | 5.6e-05 | 0.00113 | 0.31852 | 1.3307 | 2.5482 | 100.99 | 20.69 | 18.121 | 101.047 | 107.4548 | 92.3491 | 0.004383 | 0.96593 | 0.023244 | 24.0643 | 0.51889 | 0.15703 | 0.016283 | 1.1378 | 1.1378 | 0.005672
2019-03-20 12:53 | 1553100777.1695 | 194 | 0.20633 | 0.003171 | 4.5e-05 | 108.2005 | 102.267 | 116.8802 | 20.9856 | 0.003248 | 0.002894 | 160.4972 | 0.26422 | 8.3996 | 0.003613 | 0.005191 | 0.004999 | 5.7e-05 | 0.001131 | 0.32659 | 1.3422 | 2.548 | 100.99 | 20.68 | 18.113 | 99.7282 | 108.5559 | 91.1859 | 0.004378 | 0.96901 | 0.023279 | 24.0233 | 0.50446 | 0.15718 | 0.015744 | 1.1018 | 1.1018 | 0.005863
2019-03-20 12:53 | 1553100778.1701 | 195 | 0.20633 | 0.00315 | 4.5e-05 | 108.2224 | 102.6336 | 117.0745 | 20.9845 | 0.003198 | 0.002835 | 160.4365 | 0.26357 | 8.4 | 0.003604 | 0.004982 | 0.004909 | 5.5e-05 | 0.001127 | 0.34418 | 1.281 | 2.548 | 100.99 | 20.68 | 18.093 | 98.4975 | 108.3491 | 89.7593 | 0.004374 | 0.97282 | 0.023344 | 23.9961 | 0.47427 | 0.1568 | 0.015809 | 1.1095 | 1.1095 | 0.005808
2019-03-20 12:53 | 1553100779.1701 | 196 | 0.20635 | 0.003127 | 4.5e-05 | 108.2765 | 103.0848 | 117.1707 | 20.9857 | 0.003235 | 0.002918 | 160.3859 | 0.26366 | 8.4022 | 0.003592 | 0.005281 | 0.005148 | 5.8e-05 | 0.001126 | 0.2943 | 1.2982 | 2.5481 | 100.99 | 20.68 | 18.082 | 99.4137 | 108.0627 | 91.7506 | 0.004371 | 0.96751 | 0.02329 | 24.0718 | 0.4395 | 0.15685 | 0.015727 | 1.1012 | 1.1012 | 0.005854
2019-03-20 12:53 | 1553100780.1702 | 197 | 0.20636 | 0.003099 | 4.6e-05 | 108.2415 | 103.3595 | 117.2797 | 20.9858 | 0.003199 | 0.00291 | 160.3248 | 0.26402 | 8.4031 | 0.003619 | 0.005215 | 0.005083 | 5.6e-05 | 0.001128 | 0.29839 | 1.3416 | 2.548 | 100.99 | 20.68 | 18.07 | 98.5354 | 108.7034 | 91.5612 | 0.004368 | 0.96801 | 0.023307 | 24.0767 | 0.41281 | 0.15706 | 0.015672 | 1.0981 | 1.0981 | 0.005879
2019-03-20 12:53 | 1553100781.1693 | 198 | 0.20637 | 0.003071 | 4.7e-05 | 108.3211 | 103.5632 | 117.3963 | 20.985 | 0.003174 | 0.002861 | 160.2777 | 0.26331 | 8.4025 | 0.003595 | 0.004961 | 0.004886 | 5.7e-05 | 0.001125 | 0.31365 | 1.3361 | 2.5481 | 100.99 | 20.68 | 18.042 | 97.9224 | 108.1354 | 90.3764 | 0.004361 | 0.97117 | 0.023356 | 24.0495 | 0.40177 | 0.15664 | 0.015655 | 1.0992 | 1.0992 | 0.005857
2019-03-20 12:53 | 1553100782.1695 | 199 | 0.20634 | 0.003046 | 4.8e-05 | 108.3635 | 103.5052 | 117.1449 | 20.986 | 0.003181 | 0.002906 | 160.2323 | 0.26398 | 8.4013 | 0.003662 | 0.004938 | 0.004768 | 5.5e-05 | 0.001126 | 0.32372 | 1.3469 | 2.5481 | 100.99 | 20.68 | 18.03 | 98.1036 | 109.7497 | 91.4748 | 0.004358 | 0.96824 | 0.023296 | 24.06 | 0.41028 | 0.15704 | 0.015837 | 1.109 | 1.109 | 0.00582
2019-03-20 12:53 | 1553100783.1699 | 200 | 0.20636 | 0.003032 | 4.8e-05 | 108.3812 | 103.4253 | 117.0372 | 20.9849 | 0.003174 | 0.002906 | 160.1791 | 0.26408 | 8.4021 | 0.003582 | 0.005239 | 0.005128 | 5.7e-05 | 0.001126 | 0.32731 | 1.3518 | 2.5481 | 100.99 | 20.68 | 18.018 | 97.9278 | 107.8191 | 91.4621 | 0.004355 | 0.96828 | 0.023286 | 24.0485 | 0.41866 | 0.1571 | 0.015765 | 1.1035 | 1.1035 | 0.005851
2019-03-20 12:53 | 1553100784.1704 | 201 | 0.20638 | 0.003019 | 4.9e-05 | 108.3693 | 103.198 | 117.0353 | 20.9845 | 0.003138 | 0.002914 | 160.1413 | 0.26405 | 8.404 | 0.003593 | 0.005084 | 0.004963 | 5.5e-05 | 0.001123 | 0.30414 | 1.2866 | 2.5481 | 100.99 | 20.68 | 17.996 | 97.0553 | 108.0844 | 91.6439 | 0.00435 | 0.96779 | 0.02328 | 24.0543 | 0.43734 | 0.15709 | 0.015676 | 1.0969 | 1.0969 | 0.005886
2019-03-20 12:53 | 1553100785.1685 | 202 | 0.20642 | 0.003007 | 5e-05 | 108.4492 | 103.1767 | 117.4949 | 20.9839 | 0.003208 | 0.002888 | 160.0993 | 0.26408 | 8.4044 | 0.003576 | 0.004948 | 0.004851 | 5.6e-05 | 0.001123 | 0.30973 | 1.3364 | 2.5482 | 100.99 | 20.69 | 17.985 | 98.7579 | 107.6654 | 91.0303 | 0.00435 | 0.96942 | 0.023345 | 24.0812 | 0.44578 | 0.1571 | 0.015425 | 1.0823 | 1.0823 | 0.005966
2019-03-20 12:53 | 1553100786.1685 | 203 | 0.20647 | 0.002998 | 5.1e-05 | 108.4424 | 103.192 | 117.7752 | 20.984 | 0.003224 | 0.002898 | 160.0522 | 0.26357 | 8.4039 | 0.003658 | 0.005384 | 0.00522 | 5.2e-05 | 0.001122 | 0.30381 | 1.3179 | 2.548 | 100.99 | 20.69 | 17.964 | 99.1454 | 109.6578 | 91.271 | 0.004345 | 0.96878 | 0.023365 | 24.1179 | 0.44389 | 0.1568 | 0.015148 | 1.0637 | 1.0637 | 0.006058
2019-03-20 12:53 | 1553100787.1695 | 204 | 0.20641 | 0.002983 | 5.3e-05 | 108.443 | 103.5613 | 117.7789 | 20.985 | 0.003225 | 0.002843 | 159.962 | 0.26334 | 8.4042 | 0.003658 | 0.005211 | 0.005131 | 5.1e-05 | 0.001121 | 0.3183 | 1.3644 | 2.5481 | 100.99 | 20.69 | 17.954 | 99.1723 | 109.651 | 89.9394 | 0.004343 | 0.97234 | 0.023408 | 24.0742 | 0.412 | 0.15666 | 0.015505 | 1.0907 | 1.0907 | 0.005903
2019-03-20 12:53 | 1553100788.1696 | 205 | 0.20642 | 0.00297 | 5.4e-05 | 108.5701 | 103.5467 | 117.6235 | 20.9846 | 0.003229 | 0.00289 | 159.9073 | 0.26317 | 8.404 | 0.003595 | 0.004647 | 0.004528 | 5.4e-05 | 0.001119 | 0.335 | 1.327 | 2.5481 | 100.99 | 20.69 | 17.945 | 99.2565 | 108.1262 | 91.0857 | 0.00434 | 0.96928 | 0.023356 | 24.0963 | 0.42373 | 0.15656 | 0.015486 | 1.0869 | 1.0869 | 0.00592
2019-03-20 12:53 | 1553100789.1687 | 206 | 0.20642 | 0.002963 | 5.5e-05 | 108.5158 | 103.4122 | 117.531 | 20.9844 | 0.003191 | 0.002922 | 159.8329 | 0.26279 | 8.4006 | 0.003581 | 0.005131 | 0.00497 | 5.6e-05 | 0.001118 | 0.33876 | 1.3341 | 2.5481 | 100.99 | 20.69 | 17.931 | 98.3326 | 107.7968 | 91.8532 | 0.004337 | 0.96724 | 0.023322 | 24.1121 | 0.43089 | 0.15633 | 0.015475 | 1.0845 | 1.0845 | 0.005925
2019-03-20 12:53 | 1553100790.1697 | 207 | 0.20643 | 0.002951 | 5.5e-05 | 108.5856 | 104.0113 | 117.52 | 20.984 | 0.003239 | 0.002846 | 159.7645 | 0.26316 | 8.4034 | 0.003545 | 0.004759 | 0.004651 | 5.5e-05 | 0.001108 | 0.30663 | 1.3313 | 2.5481 | 100.99 | 20.69 | 17.926 | 99.4967 | 106.9136 | 90.0244 | 0.004336 | 0.97211 | 0.02338 | 24.0505 | 0.38497 | 0.15655 | 0.015436 | 1.0844 | 1.0844 | 0.005934
2019-03-20 12:53 | 1553100791.1687 | 208 | 0.20634 | 0.002944 | 5.5e-05 | 108.6468 | 104.3951 | 117.7188 | 20.9846 | 0.003192 | 0.002892 | 159.6979 | 0.26255 | 8.4018 | 0.003546 | 0.005053 | 0.004858 | 5.6e-05 | 0.001101 | 0.28799 | 1.3364 | 2.5481 | 100.99 | 20.69 | 17.919 | 98.3603 | 106.9472 | 91.1326 | 0.004334 | 0.96915 | 0.023364 | 24.1076 | 0.35706 | 0.15619 | 0.015952 | 1.12 | 1.12 | 0.005732
2019-03-20 12:53 | 1553100792.1698 | 209 | 0.20646 | 0.002939 | 5.3e-05 | 108.6798 | 104.3495 | 117.9047 | 20.984 | 0.003201 | 0.002886 | 159.6319 | 0.26369 | 8.4048 | 0.003567 | 0.004983 | 0.004878 | 5.6e-05 | 0.001102 | 0.32676 | 1.3298 | 2.5482 | 100.99 | 20.69 | 17.917 | 98.569 | 107.4568 | 90.981 | 0.004334 | 0.96956 | 0.023387 | 24.1216 | 0.36368 | 0.15687 | 0.015251 | 1.0719 | 1.0719 | 0.006015
2019-03-20 12:53 | 1553100793.1688 | 210 | 0.20646 | 0.00293 | 5.1e-05 | 108.7341 | 104.5828 | 117.8592 | 20.9834 | 0.003215 | 0.002912 | 159.5706 | 0.26403 | 8.4 | 0.003562 | 0.005254 | 0.005105 | 5.5e-05 | 0.001101 | 0.33157 | 1.3021 | 2.5482 | 100.99 | 20.69 | 17.918 | 98.9292 | 107.3231 | 91.5958 | 0.004334 | 0.96792 | 0.023363 | 24.1373 | 0.34818 | 0.15707 | 0.015281 | 1.073 | 1.073 | 0.006017
2019-03-20 12:53 | 1553100794.169 | 211 | 0.20655 | 0.002924 | 4.8e-05 | 108.6633 | 104.7547 | 117.7709 | 20.9848 | 0.003209 | 0.002872 | 159.5029 | 0.26359 | 8.4029 | 0.003595 | 0.005708 | 0.005648 | 5.3e-05 | 0.001097 | 0.32621 | 1.302 | 2.548 | 100.99 | 20.69 | 17.913 | 98.7674 | 108.1145 | 90.6457 | 0.004333 | 0.97045 | 0.023385 | 24.0968 | 0.32766 | 0.15681 | 0.014723 | 1.0349 | 1.0349 | 0.006227
2019-03-20 12:53 | 1553100795.169 | 212 | 0.20658 | 0.002912 | 4.7e-05 | 108.7716 | 104.9259 | 117.7101 | 20.9848 | 0.00328 | 0.002871 | 159.4275 | 0.26337 | 8.4006 | 0.003496 | 0.00534 | 0.005145 | 5.3e-05 | 0.001102 | 0.31133 | 1.3309 | 2.548 | 100.99 | 20.69 | 17.911 | 100.4923 | 105.714 | 90.6223 | 0.004332 | 0.97051 | 0.023379 | 24.0898 | 0.32197 | 0.15668 | 0.014578 | 1.0246 | 1.0246 | 0.006285
2019-03-20 12:53 | 1553100796.1691 | 213 | 0.20661 | 0.002895 | 4.6e-05 | 108.7402 | 105.3942 | 117.6167 | 20.9827 | 0.00323 | 0.002897 | 159.3799 | 0.26344 | 8.402 | 0.003497 | 0.005064 | 0.004937 | 5.5e-05 | 0.001102 | 0.28871 | 1.3527 | 2.5481 | 100.99 | 20.69 | 17.909 | 99.286 | 105.7361 | 91.2457 | 0.004332 | 0.96885 | 0.02335 | 24.1009 | 0.27956 | 0.15672 | 0.014417 | 1.0119 | 1.0119 | 0.006365
2019-03-20 12:53 | 1553100797.1691 | 214 | 0.20663 | 0.002871 | 4.7e-05 | 108.7752 | 105.4734 | 117.6482 | 20.9835 | 0.003201 | 0.002919 | 159.3402 | 0.26323 | 8.4016 | 0.003477 | 0.004827 | 0.004727 | 5.5e-05 | 0.001107 | 0.31741 | 1.3176 | 2.5481 | 100.99 | 20.69 | 17.908 | 98.5765 | 105.2673 | 91.7639 | 0.004332 | 0.96748 | 0.023337 | 24.1212 | 0.27571 | 0.1566 | 0.014342 | 1.006 | 1.006 | 0.006398
2019-03-20 12:53 | 1553100798.1691 | 215 | 0.20662 | 0.002845 | 4.9e-05 | 108.8007 | 105.6162 | 117.7435 | 20.9839 | 0.003185 | 0.002854 | 159.269 | 0.26277 | 8.4049 | 0.003505 | 0.004816 | 0.00477 | 5.5e-05 | 0.001107 | 0.29806 | 1.3327 | 2.5481 | 100.99 | 20.69 | 17.906 | 98.1888 | 105.9477 | 90.2189 | 0.004331 | 0.97159 | 0.023396 | 24.0798 | 0.26571 | 0.15632 | 0.014439 | 1.0153 | 1.0153 | 0.006328
2019-03-20 12:53 | 1553100799.1692 | 216 | 0.20659 | 0.002827 | 5.1e-05 | 108.7939 | 105.1382 | 117.8151 | 20.9836 | 0.003257 | 0.002925 | 159.214 | 0.26295 | 8.4036 | 0.003501 | 0.004619 | 0.004506 | 5.5e-05 | 0.001106 | 0.30879 | 1.3077 | 2.548 | 100.99 | 20.69 | 17.903 | 99.9279 | 105.8317 | 91.9127 | 0.00433 | 0.96708 | 0.023348 | 24.1432 | 0.30573 | 0.15643 | 0.014608 | 1.025 | 1.025 | 0.006273
2019-03-20 12:53 | 1553100800.1692 | 217 | 0.20661 | 0.002814 | 5.2e-05 | 108.9052 | 104.9584 | 117.6876 | 20.9836 | 0.003105 | 0.002883 | 159.1728 | 0.26418 | 8.4031 | 0.003533 | 0.005057 | 0.004943 | 5.4e-05 | 0.001111 | 0.30613 | 1.3419 | 2.5481 | 100.99 | 20.69 | 17.9 | 96.2582 | 106.6312 | 90.9025 | 0.00433 | 0.96976 | 0.023368 | 24.0968 | 0.33018 | 0.15716 | 0.014515 | 1.0193 | 1.0193 | 0.006337
2019-03-20 12:53 | 1553100801.1693 | 218 | 0.20658 | 0.002803 | 5.3e-05 | 108.9427 | 104.563 | 117.4653 | 20.9838 | 0.003157 | 0.002877 | 159.1174 | 0.26294 | 8.4026 | 0.00354 | 0.005009 | 0.004905 | 5.5e-05 | 0.001107 | 0.31686 | 1.3246 | 2.5481 | 100.99 | 20.69 | 17.896 | 97.5286 | 106.7935 | 90.762 | 0.004329 | 0.97014 | 0.023351 | 24.0693 | 0.36702 | 0.15642 | 0.014716 | 1.0326 | 1.0326 | 0.006226
2019-03-20 12:53 | 1553100802.1683 | 219 | 0.20656 | 0.002793 | 5.4e-05 | 108.9744 | 104.449 | 117.5679 | 20.9828 | 0.003173 | 0.002874 | 159.0539 | 0.26258 | 8.4045 | 0.003608 | 0.004661 | 0.004666 | 5.5e-05 | 0.001107 | 0.3115 | 1.3389 | 2.5481 | 100.99 | 20.69 | 17.886 | 97.9024 | 108.4404 | 90.6908 | 0.004326 | 0.97033 | 0.023363 | 24.0775 | 0.37938 | 0.15621 | 0.014861 | 1.0433 | 1.0433 | 0.006154
2019-03-20 12:53 | 1553100803.1694 | 220 | 0.20661 | 0.002782 | 5.6e-05 | 108.9925 | 104.7394 | 117.6132 | 20.9828 | 0.00312 | 0.00292 | 158.9882 | 0.26419 | 8.4024 | 0.003564 | 0.004351 | 0.004191 | 5.7e-05 | 0.001112 | 0.32421 | 1.3491 | 2.5481 | 100.99 | 20.69 | 17.881 | 96.6148 | 107.3674 | 91.8 | 0.004325 | 0.96738 | 0.023332 | 24.1188 | 0.35603 | 0.15717 | 0.014529 | 1.0185 | 1.0185 | 0.006342
2019-03-20 12:53 | 1553100804.1684 | 221 | 0.2066 | 0.002768 | 5.8e-05 | 108.975 | 105.2137 | 117.558 | 20.9829 | 0.003145 | 0.0029 | 158.9448 | 0.26403 | 8.4021 | 0.00355 | 0.004932 | 0.004781 | 5.6e-05 | 0.001109 | 0.33251 | 1.3112 | 2.5481 | 100.99 | 20.69 | 17.868 | 97.23 | 107.0336 | 91.3076 | 0.004322 | 0.96869 | 0.023342 | 24.0969 | 0.31418 | 0.15707 | 0.014616 | 1.025 | 1.025 | 0.006298
2019-03-20 12:53 | 1553100805.1685 | 222 | 0.20662 | 0.002762 | 6.1e-05 | 109.0539 | 105.229 | 117.3015 | 20.9838 | 0.003212 | 0.002917 | 158.9047 | 0.26346 | 8.4019 | 0.003475 | 0.00462 | 0.004434 | 5.6e-05 | 0.001103 | 0.32836 | 1.3128 | 2.548 | 100.99 | 20.69 | 17.863 | 98.8479 | 105.2156 | 91.7393 | 0.004321 | 0.96754 | 0.023303 | 24.0848 | 0.31936 | 0.15673 | 0.014514 | 1.0158 | 1.0158 | 0.006341
2019-03-20 12:53 | 1553100806.1686 | 223 | 0.20662 | 0.002756 | 6.3e-05 | 109.086 | 105.2824 | 117.4048 | 20.9828 | 0.003208 | 0.002852 | 158.8297 | 0.26428 | 8.402 | 0.003466 | 0.004897 | 0.004816 | 5.5e-05 | 0.001103 | 0.35618 | 1.2927 | 2.5482 | 100.99 | 20.69 | 17.857 | 98.7607 | 105.0023 | 90.1687 | 0.004319 | 0.97172 | 0.023364 | 24.0435 | 0.31744 | 0.15722 | 0.014519 | 1.0187 | 1.0187 | 0.006343
2019-03-20 12:53 | 1553100807.1696 | 224 | 0.20664 | 0.002743 | 6.5e-05 | 109.0953 | 105.9477 | 117.7005 | 20.9827 | 0.003168 | 0.002861 | 158.7825 | 0.26311 | 8.404 | 0.003408 | 0.004947 | 0.004861 | 5.6e-05 | 0.001106 | 0.33644 | 1.3309 | 2.5481 | 100.99 | 20.68 | 17.847 | 97.7704 | 103.5981 | 90.3671 | 0.004314 | 0.97119 | 0.023387 | 24.0804 | 0.26187 | 0.15652 | 0.01441 | 1.0119 | 1.0119 | 0.006357
2019-03-20 12:53 | 1553100808.1696 | 225 | 0.20658 | 0.002745 | 6.7e-05 | 109.1898 | 106.1548 | 117.6885 | 20.9826 | 0.003172 | 0.002858 | 158.7452 | 0.2629 | 8.4029 | 0.003412 | 0.0048 | 0.004634 | 5.6e-05 | 0.001108 | 0.30757 | 1.3301 | 2.5481 | 100.99 | 20.68 | 17.842 | 97.8817 | 103.6814 | 90.3068 | 0.004313 | 0.97135 | 0.023387 | 24.0771 | 0.25214 | 0.1564 | 0.014786 | 1.0383 | 1.0383 | 0.006191
2019-03-20 12:53 | 1553100809.1687 | 226 | 0.20666 | 0.002729 | 6.9e-05 | 109.1021 | 106.2336 | 117.732 | 20.9827 | 0.003168 | 0.002895 | 158.6884 | 0.26233 | 8.4037 | 0.00348 | 0.005236 | 0.005149 | 5.8e-05 | 0.001097 | 0.32084 | 1.36 | 2.5481 | 100.99 | 20.68 | 17.837 | 97.7839 | 105.3257 | 91.205 | 0.004312 | 0.96896 | 0.023363 | 24.1113 | 0.23832 | 0.15606 | 0.01429 | 1.0023 | 1.0023 | 0.0064
2019-03-20 12:53 | 1553100810.1698 | 227 | 0.20667 | 0.002724 | 7.1e-05 | 109.1781 | 106.058 | 117.6148 | 20.9835 | 0.003085 | 0.002849 | 158.6294 | 0.2623 | 8.4056 | 0.003514 | 0.005378 | 0.005348 | 5.5e-05 | 0.001098 | 0.30807 | 1.316 | 2.548 | 100.99 | 20.69 | 17.83 | 95.7782 | 106.1517 | 90.0865 | 0.004313 | 0.97194 | 0.023387 | 24.0623 | 0.25935 | 0.15604 | 0.014242 | 0.99977 | 0.99977 | 0.006415
2019-03-20 12:54 | 1553100811.1688 | 228 | 0.20669 | 0.002713 | 7.4e-05 | 109.2981 | 105.9118 | 117.6876 | 20.9837 | 0.003169 | 0.002853 | 158.5852 | 0.26165 | 8.4014 | 0.003378 | 0.004759 | 0.004629 | 5.9e-05 | 0.001095 | 0.32029 | 1.3413 | 2.5481 | 100.99 | 20.69 | 17.828 | 97.8059 | 102.8556 | 90.1947 | 0.004312 | 0.97165 | 0.023391 | 24.0733 | 0.28151 | 0.15565 | 0.014161 | 0.99406 | 0.99406 | 0.006436
2019-03-20 12:54 | 1553100812.1699 | 229 | 0.20661 | 0.0027 | 7.8e-05 | 109.2726 | 105.824 | 117.7395 | 20.9831 | 0.003205 | 0.00288 | 158.5595 | 0.26279 | 8.4034 | 0.003417 | 0.005311 | 0.005164 | 5.5e-05 | 0.001098 | 0.33251 | 1.3323 | 2.548 | 100.99 | 20.69 | 17.827 | 98.6669 | 103.8177 | 90.8411 | 0.004312 | 0.96993 | 0.023375 | 24.1 | 0.28684 | 0.15633 | 0.014624 | 1.0258 | 1.0258 | 0.006264
2019-03-20 12:54 | 1553100813.1699 | 230 | 0.20666 | 0.002693 | 8.1e-05 | 109.2884 | 106.0419 | 117.9425 | 20.983 | 0.003188 | 0.002872 | 158.4873 | 0.26303 | 8.4015 | 0.00344 | 0.004713 | 0.00459 | 5.2e-05 | 0.001097 | 0.33411 | 1.3425 | 2.5481 | 100.99 | 20.69 | 17.828 | 98.2732 | 104.3578 | 90.6324 | 0.004312 | 0.97049 | 0.023402 | 24.1139 | 0.26973 | 0.15648 | 0.014339 | 1.0067 | 1.0067 | 0.006389
2019-03-20 12:54 | 1553100814.169 | 231 | 0.20664 | 0.002685 | 8.6e-05 | 109.3424 | 106.0288 | 118.3157 | 20.9832 | 0.003139 | 0.002876 | 158.4081 | 0.2632 | 8.4008 | 0.003467 | 0.005463 | 0.005337 | 5.9e-05 | 0.001091 | 0.33395 | 1.3102 | 2.5481 | 100.99 | 20.69 | 17.831 | 97.0774 | 105.009 | 90.7497 | 0.004313 | 0.97017 | 0.023435 | 24.1559 | 0.27526 | 0.15657 | 0.01447 | 1.0171 | 1.0171 | 0.006327
2019-03-20 12:54 | 1553100815.169 | 232 | 0.20667 | 0.002676 | 8.9e-05 | 109.3997 | 106.1697 | 118.7295 | 20.9831 | 0.00318 | 0.002877 | 158.3584 | 0.26278 | 8.4032 | 0.003513 | 0.005371 | 0.005246 | 5.4e-05 | 0.001092 | 0.33987 | 1.3078 | 2.548 | 100.99 | 20.69 | 17.833 | 98.0664 | 106.1397 | 90.769 | 0.004313 | 0.97012 | 0.023476 | 24.1988 | 0.26806 | 0.15633 | 0.014285 | 1.0055 | 1.0055 | 0.00639
2019-03-20 12:54 | 1553100816.1691 | 233 | 0.20663 | 0.002678 | 9.2e-05 | 109.4543 | 106.5603 | 118.9537 | 20.9829 | 0.003174 | 0.002863 | 158.3123 | 0.26223 | 8.4065 | 0.003535 | 0.005172 | 0.005139 | 5.5e-05 | 0.001089 | 0.2959 | 1.308 | 2.5482 | 100.99 | 20.69 | 17.836 | 97.918 | 106.6747 | 90.4181 | 0.004314 | 0.97106 | 0.023509 | 24.2099 | 0.23968 | 0.156 | 0.014546 | 1.0253 | 1.0253 | 0.006254
2019-03-20 12:54 | 1553100817.1692 | 234 | 0.20656 | 0.002676 | 9.5e-05 | 109.3842 | 106.2489 | 118.9553 | 20.9831 | 0.003113 | 0.002894 | 158.2471 | 0.26388 | 8.4012 | 0.003441 | 0.005046 | 0.004837 | 5.7e-05 | 0.00109 | 0.32549 | 1.3275 | 2.5481 | 100.99 | 20.69 | 17.842 | 96.4598 | 104.3869 | 91.1772 | 0.004316 | 0.96903 | 0.023485 | 24.2353 | 0.26012 | 0.15698 | 0.014939 | 1.0518 | 1.0518 | 0.006134
2019-03-20 12:54 | 1553100818.1692 | 235 | 0.20664 | 0.00267 | 9.9e-05 | 109.4634 | 106.8051 | 119.0347 | 20.9828 | 0.003122 | 0.002853 | 158.1792 | 0.26267 | 8.4011 | 0.003435 | 0.005475 | 0.005396 | 5.5e-05 | 0.00109 | 0.31404 | 1.2883 | 2.5481 | 100.99 | 20.69 | 17.846 | 96.6752 | 104.2508 | 90.1921 | 0.004317 | 0.97166 | 0.023525 | 24.2106 | 0.2199 | 0.15626 | 0.014473 | 1.0204 | 1.0204 | 0.006294
2019-03-20 12:54 | 1553100819.1692 | 236 | 0.20658 | 0.002671 | 0.000101 | 109.411 | 106.9082 | 119.187 | 20.9835 | 0.003213 | 0.002886 | 158.1346 | 0.26374 | 8.4037 | 0.003438 | 0.00492 | 0.004816 | 5.6e-05 | 0.001091 | 0.32068 | 1.3137 | 2.5481 | 100.99 | 20.69 | 17.85 | 98.8743 | 104.3212 | 90.9784 | 0.004317 | 0.96956 | 0.023514 | 24.2523 | 0.20698 | 0.1569 | 0.014837 | 1.0455 | 1.0455 | 0.006168
2019-03-20 12:54 | 1553100820.1693 | 237 | 0.20662 | 0.002676 | 0.000104 | 109.5271 | 107.3724 | 119.4356 | 20.9826 | 0.003103 | 0.002871 | 158.0759 | 0.2629 | 8.4009 | 0.003409 | 0.005526 | 0.005325 | 5.7e-05 | 0.001089 | 0.33588 | 1.3118 | 2.548 | 100.99 | 20.69 | 17.859 | 96.223 | 103.6106 | 90.6202 | 0.00432 | 0.97052 | 0.02355 | 24.2656 | 0.17771 | 0.1564 | 0.014555 | 1.027 | 1.027 | 0.006259
2019-03-20 12:54 | 1553100821.1699 | 238 | 0.20657 | 0.002679 | 0.000106 | 109.5168 | 107.2454 | 119.3936 | 20.9829 | 0.003087 | 0.002851 | 157.9925 | 0.2627 | 8.4032 | 0.003411 | 0.005041 | 0.00488 | 5.9e-05 | 0.001091 | 0.34772 | 1.3039 | 2.548 | 100.99 | 20.69 | 17.863 | 95.8364 | 103.6572 | 90.1399 | 0.004321 | 0.9718 | 0.023562 | 24.2453 | 0.18746 | 0.15628 | 0.01489 | 1.0511 | 1.0511 | 0.006111
2019-03-20 12:54 | 1553100822.1684 | 239 | 0.20659 | 0.002679 | 0.000108 | 109.508 | 107.0742 | 118.9363 | 20.9824 | 0.003132 | 0.002904 | 157.9095 | 0.26378 | 8.4037 | 0.003466 | 0.004406 | 0.004292 | 5.7e-05 | 0.001091 | 0.35579 | 1.322 | 2.548 | 100.99 | 20.69 | 17.872 | 96.901 | 104.9942 | 91.4245 | 0.004323 | 0.96838 | 0.023475 | 24.2416 | 0.20104 | 0.15692 | 0.01472 | 1.0351 | 1.0351 | 0.006231
2019-03-20 12:54 | 1553100823.1684 | 240 | 0.20659 | 0.00268 | 0.000111 | 109.5633 | 106.8611 | 118.7774 | 20.9825 | 0.003162 | 0.002883 | 157.8703 | 0.26332 | 8.4029 | 0.003365 | 0.005243 | 0.005164 | 5.8e-05 | 0.00109 | 0.33196 | 1.3165 | 2.5482 | 100.99 | 20.69 | 17.875 | 97.6308 | 102.5512 | 90.9 | 0.004324 | 0.96977 | 0.023476 | 24.208 | 0.22337 | 0.15665 | 0.014737 | 1.0362 | 1.0362 | 0.006213
2019-03-20 12:54 | 1553100824.1689 | 241 | 0.20664 | 0.002677 | 0.000113 | 109.6497 | 106.7857 | 118.296 | 20.9832 | 0.003196 | 0.002902 | 157.8153 | 0.2634 | 8.4018 | 0.003395 | 0.004856 | 0.004705 | 5.6e-05 | 0.001089 | 0.32925 | 1.3219 | 2.548 | 100.99 | 20.69 | 17.878 | 98.4641 | 103.2629 | 91.3539 | 0.004324 | 0.96856 | 0.023414 | 24.1739 | 0.23674 | 0.1567 | 0.014469 | 1.0145 | 1.0145 | 0.006348
2019-03-20 12:54 | 1553100825.1686 | 242 | 0.20661 | 0.002675 | 0.000116 | 109.6015 | 106.8735 | 118.1808 | 20.9836 | 0.003157 | 0.002868 | 157.7478 | 0.26245 | 8.401 | 0.003373 | 0.004882 | 0.004734 | 6e-05 | 0.001087 | 0.31868 | 1.3425 | 2.548 | 100.99 | 20.69 | 17.884 | 97.5222 | 102.7448 | 90.5554 | 0.004326 | 0.97069 | 0.023428 | 24.1357 | 0.22545 | 0.15613 | 0.014589 | 1.0233 | 1.0233 | 0.006271
2019-03-20 12:54 | 1553100826.1696 | 243 | 0.20659 | 0.002674 | 0.000119 | 109.5788 | 107.0357 | 118.2397 | 20.9825 | 0.003133 | 0.002817 | 157.6931 | 0.26318 | 8.4016 | 0.003423 | 0.00548 | 0.005369 | 5.7e-05 | 0.001089 | 0.323 | 1.3664 | 2.5481 | 100.99 | 20.69 | 17.887 | 96.9301 | 103.9563 | 89.3065 | 0.004326 | 0.97403 | 0.023474 | 24.1002 | 0.21004 | 0.15657 | 0.014759 | 1.0372 | 1.0372 | 0.006204
2019-03-20 12:54 | 1553100827.1689 | 244 | 0.20658 | 0.002663 | 0.000123 | 109.6558 | 107.2487 | 118.2531 | 20.9823 | 0.003144 | 0.002821 | 157.6273 | 0.26372 | 8.4031 | 0.003427 | 0.005445 | 0.005279 | 5.8e-05 | 0.001088 | 0.33107 | 1.3646 | 2.5481 | 100.99 | 20.69 | 17.895 | 97.2043 | 104.037 | 89.4111 | 0.004328 | 0.97375 | 0.023472 | 24.1051 | 0.19853 | 0.15689 | 0.014782 | 1.0385 | 1.0385 | 0.006209
2019-03-20 12:54 | 1553100828.1698 | 245 | 0.20658 | 0.002659 | 0.000127 | 109.6675 | 107.279 | 118.1977 | 20.9819 | 0.003158 | 0.002843 | 157.5574 | 0.2632 | 8.4023 | 0.003426 | 0.004718 | 0.004548 | 5.4e-05 | 0.001089 | 0.34291 | 1.3379 | 2.5479 | 100.99 | 20.69 | 17.9 | 97.5352 | 104.0351 | 89.9544 | 0.00433 | 0.9723 | 0.023449 | 24.1175 | 0.19696 | 0.15657 | 0.01481 | 1.0392 | 1.0392 | 0.006192
2019-03-20 12:54 | 1553100829.1699 | 246 | 0.20657 | 0.002659 | 0.000129 | 109.7587 | 107.1134 | 117.9599 | 20.9818 | 0.003133 | 0.002866 | 157.4962 | 0.26331 | 8.4022 | 0.003493 | 0.00507 | 0.004948 | 5.8e-05 | 0.001086 | 0.29247 | 1.3454 | 2.5481 | 100.99 | 20.7 | 17.905 | 96.9475 | 105.6385 | 90.5052 | 0.004333 | 0.97083 | 0.023408 | 24.1115 | 0.21821 | 0.15665 | 0.014866 | 1.0412 | 1.0412 | 0.006183
2019-03-20 12:54 | 1553100830.1689 | 247 | 0.20656 | 0.002655 | 0.00013 | 109.6439 | 107.1871 | 117.539 | 20.9825 | 0.003199 | 0.002874 | 157.4598 | 0.2634 | 8.4036 | 0.003413 | 0.005172 | 0.005089 | 5.8e-05 | 0.001089 | 0.31598 | 1.3327 | 2.5482 | 100.99 | 20.7 | 17.913 | 98.5339 | 103.7149 | 90.6919 | 0.004335 | 0.97033 | 0.02336 | 24.0746 | 0.2027 | 0.1567 | 0.014931 | 1.0436 | 1.0436 | 0.006171
2019-03-20 12:54 | 1553100831.169 | 248 | 0.20658 | 0.002656 | 0.000132 | 109.7454 | 107.0992 | 117.3599 | 20.9816 | 0.003145 | 0.002873 | 157.4075 | 0.26334 | 8.4023 | 0.003443 | 0.005294 | 0.005218 | 5.4e-05 | 0.001092 | 0.31614 | 1.3734 | 2.5482 | 100.99 | 20.7 | 17.916 | 97.2248 | 104.446 | 90.6621 | 0.004336 | 0.97041 | 0.023343 | 24.0552 | 0.21831 | 0.15666 | 0.014818 | 1.0348 | 1.0348 | 0.006222
2019-03-20 12:54 | 1553100832.168 | 249 | 0.20658 | 0.002648 | 0.000134 | 109.7095 | 106.8641 | 117.3994 | 20.9823 | 0.003132 | 0.002831 | 157.3275 | 0.26333 | 8.4018 | 0.003464 | 0.005138 | 0.004994 | 5.5e-05 | 0.001097 | 0.33157 | 1.3896 | 2.5482 | 100.99 | 20.7 | 17.919 | 96.9183 | 104.9531 | 89.6579 | 0.004337 | 0.97309 | 0.02338 | 24.0261 | 0.23504 | 0.15666 | 0.014822 | 1.0366 | 1.0366 | 0.006211
2019-03-20 12:54 | 1553100833.168 | 250 | 0.20656 | 0.00265 | 0.000137 | 109.8052 | 107.7234 | 117.3076 | 20.9824 | 0.003179 | 0.002862 | 157.2675 | 0.26299 | 8.3998 | 0.003382 | 0.005083 | 0.005033 | 5.5e-05 | 0.001099 | 0.33555 | 1.3882 | 2.548 | 100.99 | 20.7 | 17.925 | 98.0547 | 102.9572 | 90.3933 | 0.004338 | 0.97112 | 0.023347 | 24.041 | 0.17121 | 0.15646 | 0.014928 | 1.0423 | 1.0423 | 0.006169
2019-03-20 12:54 | 1553100834.1691 | 251 | 0.20655 | 0.002645 | 0.000137 | 109.7862 | 107.3332 | 117.0318 | 20.9817 | 0.003117 | 0.002922 | 157.186 | 0.26186 | 8.4024 | 0.003387 | 0.004694 | 0.004555 | 6e-05 | 0.001103 | 0.3465 | 1.3966 | 2.5482 | 100.99 | 20.7 | 17.928 | 96.5432 | 103.0674 | 91.8576 | 0.004339 | 0.96723 | 0.023272 | 24.061 | 0.20211 | 0.15578 | 0.015013 | 1.045 | 1.045 | 0.006127
2019-03-20 12:54 | 1553100835.1681 | 252 | 0.20633 | 0.002643 | 0.000138 | 109.8036 | 107.2584 | 116.5004 | 20.9817 | 0.00316 | 0.002889 | 157.1024 | 0.26254 | 8.4009 | 0.00343 | 0.005363 | 0.005251 | 5.7e-05 | 0.001097 | 0.29319 | 1.3466 | 2.5482 | 100.99 | 20.7 | 17.935 | 97.5892 | 104.1152 | 91.0626 | 0.004341 | 0.96934 | 0.023245 | 23.9802 | 0.20977 | 0.15619 | 0.016301 | 1.1338 | 1.1338 | 0.005662
2019-03-20 12:54 | 1553100836.1692 | 253 | 0.20603 | 0.002653 | 0.000136 | 109.768 | 107.4487 | 115.5164 | 20.982 | 0.003167 | 0.002946 | 156.981 | 0.26299 | 8.4041 | 0.003407 | 0.004861 | 0.004777 | 5.6e-05 | 0.001094 | 0.3079 | 1.3118 | 2.548 | 100.99 | 20.7 | 17.938 | 97.7533 | 103.5589 | 92.4264 | 0.004341 | 0.96572 | 0.023103 | 23.9234 | 0.19102 | 0.15646 | 0.018092 | 1.2515 | 1.2515 | 0.005138
2019-03-20 12:54 | 1553100837.1693 | 254 | 0.20555 | 0.002689 | 0.000134 | 109.7787 | 107.8359 | 114.7606 | 20.9807 | 0.003184 | 0.002939 | 156.8638 | 0.26295 | 8.4006 | 0.003341 | 0.004634 | 0.004436 | 5.6e-05 | 0.001093 | 0.33229 | 1.2972 | 2.5481 | 100.99 | 20.7 | 17.942 | 98.1549 | 101.9531 | 92.2506 | 0.004342 | 0.96619 | 0.023033 | 23.8392 | 0.15971 | 0.15643 | 0.020959 | 1.4464 | 1.4464 | 0.004445
2019-03-20 12:54 | 1553100838.1683 | 255 | 0.20479 | 0.002762 | 0.000132 | 109.8185 | 107.76 | 114.0142 | 20.9826 | 0.003157 | 0.002938 | 156.7462 | 0.26423 | 8.4007 | 0.003421 | 0.00463 | 0.004471 | 5.9e-05 | 0.001098 | 0.34219 | 1.3043 | 2.548 | 100.99 | 20.7 | 17.947 | 97.5266 | 103.8975 | 92.2393 | 0.004344 | 0.96622 | 0.022958 | 23.7612 | 0.16925 | 0.15719 | 0.025428 | 1.7499 | 1.7499 | 0.003692
2019-03-20 12:54 | 1553100839.1694 | 256 | 0.20405 | 0.002918 | 0.000132 | 109.8324 | 107.6789 | 113.1589 | 20.9801 | 0.003228 | 0.003018 | 156.6133 | 0.26231 | 8.4027 | 0.003372 | 0.004775 | 0.004689 | 5.5e-05 | 0.001095 | 0.30879 | 1.3094 | 2.5481 | 100.99 | 20.7 | 17.949 | 99.2332 | 102.7091 | 94.1622 | 0.004344 | 0.96116 | 0.022812 | 23.7341 | 0.17712 | 0.15605 | 0.02964 | 2.0272 | 2.0272 | 0.003164
2019-03-20 12:54 | 1553100840.1688 | 257 | 0.20339 | 0.003353 | 0.000134 | 109.8935 | 107.3837 | 112.7338 | 20.9815 | 0.00329 | 0.003032 | 156.4379 | 0.26258 | 8.4031 | 0.003354 | 0.005131 | 0.004995 | 5.4e-05 | 0.001092 | 0.33931 | 1.3173 | 2.5481 | 100.99 | 20.7 | 17.948 | 100.723 | 102.2856 | 94.5095 | 0.004344 | 0.96025 | 0.022759 | 23.7007 | 0.20666 | 0.15621 | 0.033071 | 2.2565 | 2.2565 | 0.002845
2019-03-20 12:54 | 1553100841.1685 | 258 | 0.20282 | 0.003852 | 0.000138 | 109.8243 | 106.3656 | 112.3162 | 20.9806 | 0.003342 | 0.003046 | 156.306 | 0.26313 | 8.4002 | 0.003376 | 0.004921 | 0.004827 | 5.7e-05 | 0.001084 | 0.34219 | 1.3577 | 2.548 | 100.99 | 20.7 | 17.949 | 101.9775 | 102.8094 | 94.8311 | 0.004344 | 0.95941 | 0.022707 | 23.6671 | 0.28623 | 0.15654 | 0.035872 | 2.4418 | 2.4418 | 0.002635
2019-03-20 12:54 | 1553100842.1695 | 259 | 0.20224 | 0.004334 | 0.000144 | 109.9158 | 105.8827 | 111.9036 | 20.9803 | 0.003294 | 0.003102 | 156.1974 | 0.2633 | 8.4021 | 0.003471 | 0.004931 | 0.004817 | 5.7e-05 | 0.001088 | 0.33229 | 1.3235 | 2.548 | 100.99 | 20.7 | 17.949 | 100.8326 | 105.1029 | 96.1751 | 0.004344 | 0.95592 | 0.022623 | 23.6667 | 0.33438 | 0.15664 | 0.038721 | 2.6255 | 2.6255 | 0.002452
2019-03-20 12:54 | 1553100843.1699 | 260 | 0.20189 | 0.004734 | 0.00015 | 109.8657 | 105.2742 | 111.1426 | 20.9812 | 0.003353 | 0.003149 | 156.059 | 0.26324 | 8.4012 | 0.003525 | 0.005231 | 0.00514 | 5.6e-05 | 0.00109 | 0.30342 | 1.3455 | 2.548 | 100.99 | 20.7 | 17.952 | 102.2553 | 106.4219 | 97.3104 | 0.004345 | 0.95299 | 0.022512 | 23.6223 | 0.38185 | 0.1566 | 0.040354 | 2.7223 | 2.7223 | 0.002364
2019-03-20 12:54 | 1553100844.17 | 261 | 0.20155 | 0.005165 | 0.000157 | 109.9453 | 104.8474 | 110.855 | 20.9801 | 0.003399 | 0.003137 | 155.9344 | 0.26295 | 8.4037 | 0.003524 | 0.004824 | 0.004678 | 5.7e-05 | 0.001091 | 0.33572 | 1.3175 | 2.548 | 100.99 | 20.7 | 17.953 | 103.3786 | 106.4051 | 97.0377 | 0.004345 | 0.95369 | 0.022491 | 23.5831 | 0.42466 | 0.15643 | 0.041855 | 2.8204 | 2.8204 | 0.00228
2019-03-20 12:54 | 1553100845.1687 | 262 | 0.20107 | 0.005534 | 0.000161 | 109.9498 | 103.9926 | 110.7867 | 20.9802 | 0.003441 | 0.003163 | 155.8058 | 0.26219 | 8.4036 | 0.003451 | 0.005015 | 0.004905 | 5.7e-05 | 0.001088 | 0.33339 | 1.3368 | 2.5481 | 100.99 | 20.7 | 17.954 | 104.3873 | 104.6287 | 97.6653 | 0.004345 | 0.95208 | 0.022465 | 23.5958 | 0.49825 | 0.15598 | 0.044224 | 2.9762 | 2.9762 | 0.002154
2019-03-20 12:54 | 1553100846.1687 | 263 | 0.20075 | 0.005868 | 0.000166 | 109.8867 | 103.5583 | 110.9499 | 20.9804 | 0.003404 | 0.003157 | 155.7029 | 0.26282 | 8.4035 | 0.003567 | 0.005148 | 0.005067 | 5.8e-05 | 0.001089 | 0.31581 | 1.3167 | 2.5481 | 100.99 | 20.7 | 17.959 | 103.4905 | 107.4312 | 97.5206 | 0.004347 | 0.95245 | 0.022486 | 23.6085 | 0.53055 | 0.15635 | 0.045724 | 3.0795 | 3.0795 | 0.002087
2019-03-20 12:54 | 1553100847.1688 | 264 | 0.20061 | 0.006209 | 0.000169 | 109.8897 | 103.2649 | 111.1825 | 20.9805 | 0.003444 | 0.003216 | 155.6005 | 0.26305 | 8.4017 | 0.00349 | 0.004616 | 0.004475 | 5.5e-05 | 0.001089 | 0.33771 | 1.3724 | 2.5482 | 100.99 | 20.7 | 17.961 | 104.4598 | 105.5626 | 98.9286 | 0.004347 | 0.94885 | 0.022467 | 23.6781 | 0.55616 | 0.15649 | 0.046175 | 3.107 | 3.107 | 0.00207
2019-03-20 12:54 | 1553100848.1679 | 265 | 0.20056 | 0.006424 | 0.000169 | 109.8997 | 103.1887 | 111.3416 | 20.9794 | 0.003474 | 0.003177 | 155.4908 | 0.262 | 8.4023 | 0.003581 | 0.005621 | 0.005499 | 5.8e-05 | 0.001092 | 0.28816 | 1.3546 | 2.5481 | 100.99 | 20.7 | 17.963 | 105.1934 | 107.7736 | 97.9851 | 0.004347 | 0.95126 | 0.022512 | 23.665 | 0.56359 | 0.15586 | 0.046179 | 3.1135 | 3.1135 | 0.002058
2019-03-20 12:54 | 1553100849.1679 | 266 | 0.20054 | 0.006629 | 0.000169 | 109.9446 | 102.7783 | 111.6047 | 20.9799 | 0.003487 | 0.003261 | 155.403 | 0.2624 | 8.4028 | 0.003666 | 0.004916 | 0.004787 | 5.6e-05 | 0.001093 | 0.32908 | 1.3775 | 2.5481 | 100.99 | 20.7 | 17.963 | 105.4905 | 109.8548 | 100.0215 | 0.004347 | 0.94607 | 0.022477 | 23.7578 | 0.60289 | 0.1561 | 0.046068 | 3.1012 | 3.1012 | 0.002069
2019-03-20 12:54 | 1553100850.1692 | 267 | 0.20056 | 0.006811 | 0.000168 | 109.912 | 102.8063 | 111.7784 | 20.9794 | 0.003503 | 0.003237 | 155.323 | 0.26367 | 8.4019 | 0.003629 | 0.005094 | 0.004955 | 5.8e-05 | 0.001096 | 0.3465 | 1.3299 | 2.5479 | 100.99 | 20.7 | 17.965 | 105.8809 | 108.9547 | 99.4498 | 0.004348 | 0.94752 | 0.022511 | 23.7581 | 0.59779 | 0.15686 | 0.045714 | 3.0822 | 3.0822 | 0.002092
2019-03-20 12:54 | 1553100851.169 | 268 | 0.2007 | 0.006948 | 0.000166 | 109.9207 | 102.8695 | 112.031 | 20.9807 | 0.003539 | 0.003203 | 155.214 | 0.26311 | 8.3997 | 0.003751 | 0.005675 | 0.005572 | 5.5e-05 | 0.001094 | 0.32908 | 1.341 | 2.5482 | 100.99 | 20.7 | 17.965 | 106.7726 | 111.9129 | 98.6161 | 0.004348 | 0.94964 | 0.022562 | 23.7583 | 0.59301 | 0.15652 | 0.044735 | 3.0232 | 3.0232 | 0.002128
2019-03-20 12:54 | 1553100852.1681 | 269 | 0.20088 | 0.007052 | 0.000164 | 109.9036 | 102.7858 | 112.2293 | 20.9804 | 0.003541 | 0.00321 | 155.0908 | 0.26351 | 8.4055 | 0.003771 | 0.005174 | 0.005053 | 5.4e-05 | 0.001095 | 0.33516 | 1.3356 | 2.548 | 100.99 | 20.7 | 17.967 | 106.8084 | 112.4107 | 98.7866 | 0.004348 | 0.94921 | 0.022577 | 23.7848 | 0.59889 | 0.15676 | 0.043509 | 2.9424 | 2.9424 | 0.00219
2019-03-20 12:54 | 1553100853.1696 | 270 | 0.201 | 0.007115 | 0.000164 | 109.9647 | 102.2188 | 112.6889 | 20.9804 | 0.00347 | 0.003218 | 154.9742 | 0.26277 | 8.402 | 0.003839 | 0.005039 | 0.004907 | 5.4e-05 | 0.001095 | 0.30917 | 1.3037 | 2.548 | 100.99 | 20.7 | 17.968 | 105.0882 | 114.055 | 98.9937 | 0.004349 | 0.94868 | 0.022617 | 23.8401 | 0.65334 | 0.15632 | 0.042731 | 2.895 | 2.895 | 0.002219
2019-03-20 12:54 | 1553100854.1692 | 271 | 0.20126 | 0.007144 | 0.000163 | 109.8861 | 101.9762 | 112.8619 | 20.9792 | 0.003554 | 0.003281 | 154.8517 | 0.26316 | 8.4031 | 0.003858 | 0.005514 | 0.00538 | 5.5e-05 | 0.001091 | 0.30359 | 1.3306 | 2.5481 | 100.99 | 20.7 | 17.968 | 107.1331 | 114.5107 | 100.5095 | 0.004349 | 0.94483 | 0.022588 | 23.9069 | 0.6682 | 0.15656 | 0.041161 | 2.7853 | 2.7853 | 0.00231
2019-03-20 12:54 | 1553100855.1684 | 272 | 0.20136 | 0.007141 | 0.000163 | 109.9097 | 102.4494 | 112.6479 | 20.9798 | 0.003542 | 0.003208 | 154.7569 | 0.2627 | 8.4037 | 0.003791 | 0.006284 | 0.006097 | 5.5e-05 | 0.001095 | 0.31116 | 1.3866 | 2.5481 | 100.99 | 20.7 | 17.97 | 106.844 | 112.8824 | 98.7389 | 0.004349 | 0.94933 | 0.02262 | 23.8275 | 0.62871 | 0.15628 | 0.040559 | 2.7485 | 2.7485 | 0.002337
2019-03-20 12:54 | 1553100856.1693 | 273 | 0.20148 | 0.007116 | 0.000164 | 109.9265 | 102.7091 | 112.703 | 20.9798 | 0.003574 | 0.003279 | 154.6595 | 0.26254 | 8.4023 | 0.003861 | 0.005039 | 0.004972 | 5.6e-05 | 0.0011 | 0.31885 | 1.3688 | 2.5482 | 100.99 | 20.7 | 17.972 | 107.6063 | 114.5991 | 100.4577 | 0.00435 | 0.94496 | 0.022574 | 23.8884 | 0.60743 | 0.15618 | 0.03986 | 2.6955 | 2.6955 | 0.002381
2019-03-20 12:54 | 1553100857.1683 | 274 | 0.20158 | 0.007067 | 0.000167 | 109.9579 | 103.0066 | 112.4434 | 20.9791 | 0.00349 | 0.003215 | 154.5619 | 0.26313 | 8.4025 | 0.003823 | 0.005397 | 0.005248 | 5.5e-05 | 0.001104 | 0.32444 | 1.3364 | 2.5481 | 100.99 | 20.7 | 17.973 | 105.5753 | 113.6569 | 98.9084 | 0.00435 | 0.9489 | 0.022595 | 23.8113 | 0.58411 | 0.15654 | 0.03935 | 2.6634 | 2.6634 | 0.002416
2019-03-20 12:54 | 1553100858.1684 | 275 | 0.20168 | 0.007011 | 0.000168 | 109.9644 | 103.3894 | 112.585 | 20.9798 | 0.003552 | 0.003239 | 154.4529 | 0.26286 | 8.4027 | 0.003862 | 0.005129 | 0.005003 | 5.5e-05 | 0.001103 | 0.34811 | 1.3479 | 2.5479 | 100.99 | 20.7 | 17.972 | 107.0888 | 114.6061 | 99.4943 | 0.00435 | 0.94741 | 0.022591 | 23.8451 | 0.55147 | 0.15638 | 0.038801 | 2.6258 | 2.6258 | 0.002448
2019-03-20 12:54 | 1553100859.1684 | 276 | 0.20185 | 0.006935 | 0.000171 | 109.9776 | 103.6218 | 112.9347 | 20.9798 | 0.003478 | 0.003314 | 154.3368 | 0.26256 | 8.4025 | 0.003855 | 0.00561 | 0.005456 | 5.3e-05 | 0.001104 | 0.32084 | 1.3502 | 2.548 | 100.99 | 20.7 | 17.973 | 105.2786 | 114.452 | 101.3085 | 0.00435 | 0.94282 | 0.022571 | 23.9401 | 0.53246 | 0.1562 | 0.037885 | 2.5614 | 2.5614 | 0.002506
2019-03-20 12:54 | 1553100860.1686 | 277 | 0.20189 | 0.006869 | 0.000173 | 109.9582 | 104.3241 | 112.9704 | 20.9784 | 0.003562 | 0.00325 | 154.2111 | 0.26297 | 8.4047 | 0.003884 | 0.006022 | 0.005902 | 5.6e-05 | 0.0011 | 0.31277 | 1.3138 | 2.5481 | 100.99 | 20.69 | 17.974 | 107.3245 | 115.141 | 99.7579 | 0.004347 | 0.94674 | 0.022622 | 23.8943 | 0.47046 | 0.15644 | 0.037734 | 2.5567 | 2.5567 | 0.002515
2019-03-20 12:54 | 1553100861.1695 | 278 | 0.20184 | 0.006803 | 0.000176 | 110.0358 | 105.192 | 112.7826 | 20.9795 | 0.003638 | 0.00326 | 154.1099 | 0.26299 | 8.4037 | 0.003942 | 0.006297 | 0.006229 | 5.6e-05 | 0.001104 | 0.3454 | 1.2794 | 2.548 | 100.99 | 20.7 | 17.976 | 109.1676 | 116.5569 | 99.9976 | 0.004351 | 0.94613 | 0.022596 | 23.8821 | 0.40268 | 0.15646 | 0.0381 | 2.5785 | 2.5785 | 0.002494
2019-03-20 12:54 | 1553100862.1686 | 279 | 0.20188 | 0.006739 | 0.000178 | 110.0054 | 105.5497 | 112.4636 | 20.98 | 0.003616 | 0.003283 | 153.9985 | 0.26284 | 8.4011 | 0.003942 | 0.005271 | 0.005148 | 5.6e-05 | 0.001099 | 0.32421 | 1.2961 | 2.5481 | 100.99 | 20.69 | 17.977 | 108.637 | 116.5732 | 100.5701 | 0.004348 | 0.94468 | 0.022546 | 23.8666 | 0.36984 | 0.15637 | 0.037929 | 2.5611 | 2.5611 | 0.002509
2019-03-20 12:54 | 1553100863.1686 | 280 | 0.20185 | 0.00668 | 0.00018 | 110.029 | 105.7007 | 112.1488 | 20.9793 | 0.003612 | 0.00331 | 153.8732 | 0.26273 | 8.4011 | 0.003915 | 0.004642 | 0.004548 | 5.5e-05 | 0.001101 | 0.33843 | 1.336 | 2.5481 | 100.99 | 20.69 | 17.98 | 108.5269 | 115.9142 | 101.2108 | 0.004349 | 0.94306 | 0.022495 | 23.8536 | 0.35897 | 0.1563 | 0.038195 | 2.5731 | 2.5731 | 0.002497
2019-03-20 12:54 | 1553100864.1677 | 281 | 0.2018 | 0.006641 | 0.000182 | 110.0261 | 106.0023 | 111.9902 | 20.979 | 0.00354 | 0.003302 | 153.7885 | 0.26236 | 8.4023 | 0.003972 | 0.005301 | 0.005147 | 5.7e-05 | 0.001098 | 0.34036 | 1.3958 | 2.5481 | 100.99 | 20.69 | 17.98 | 106.7825 | 117.3007 | 101.029 | 0.004349 | 0.94352 | 0.022485 | 23.8309 | 0.33325 | 0.15608 | 0.038522 | 2.5939 | 2.5939 | 0.002473
2019-03-20 12:54 | 1553100865.1681 | 282 | 0.20172 | 0.006607 | 0.000184 | 109.9964 | 106.2093 | 111.9594 | 20.9789 | 0.003642 | 0.00333 | 153.6491 | 0.26243 | 8.4032 | 0.00393 | 0.005541 | 0.005397 | 5.6e-05 | 0.001102 | 0.32444 | 1.3751 | 2.5481 | 100.99 | 20.69 | 17.981 | 109.265 | 116.2816 | 101.7065 | 0.004349 | 0.94181 | 0.022462 | 23.8492 | 0.31338 | 0.15612 | 0.039005 | 2.6235 | 2.6235 | 0.002446
2019-03-20 12:54 | 1553100866.1679 | 283 | 0.20167 | 0.006586 | 0.000187 | 110.0449 | 106.9886 | 111.8078 | 20.9785 | 0.003633 | 0.003302 | 153.532 | 0.26235 | 8.4039 | 0.003993 | 0.005441 | 0.005346 | 5.6e-05 | 0.001097 | 0.30558 | 1.3769 | 2.5481 | 100.99 | 20.69 | 17.98 | 109.0359 | 117.7993 | 101.0247 | 0.004349 | 0.94353 | 0.022467 | 23.8114 | 0.25193 | 0.15607 | 0.039352 | 2.6473 | 2.6473 | 0.002423
2019-03-20 12:54 | 1553100867.1678 | 284 | 0.20161 | 0.006574 | 0.000189 | 110.029 | 107.2293 | 111.6087 | 20.9785 | 0.003641 | 0.003308 | 153.3982 | 0.26298 | 8.4035 | 0.003999 | 0.005678 | 0.005608 | 5.7e-05 | 0.001097 | 0.32228 | 1.3102 | 2.5481 | 100.99 | 20.69 | 17.978 | 109.2329 | 117.968 | 101.1746 | 0.004348 | 0.94315 | 0.022442 | 23.7949 | 0.23054 | 0.15645 | 0.03974 | 2.6704 | 2.6704 | 0.002408
2019-03-20 12:54 | 1553100868.1696 | 285 | 0.20139 | 0.006585 | 0.000191 | 110.0468 | 107.1912 | 110.9393 | 20.9789 | 0.003606 | 0.003392 | 153.2957 | 0.26214 | 8.4034 | 0.004076 | 0.005306 | 0.005185 | 5.6e-05 | 0.001101 | 0.33035 | 1.3218 | 2.548 | 100.99 | 20.69 | 17.97 | 108.3972 | 119.8472 | 103.1857 | 0.004346 | 0.93811 | 0.022315 | 23.7871 | 0.23517 | 0.15595 | 0.041001 | 2.7394 | 2.7394 | 0.00234
2019-03-20 12:54 | 1553100869.168 | 286 | 0.20127 | 0.006608 | 0.000192 | 110.0831 | 107.5365 | 110.0828 | 20.9798 | 0.003682 | 0.00337 | 153.1795 | 0.26243 | 8.4033 | 0.004014 | 0.005837 | 0.005653 | 5.4e-05 | 0.001102 | 0.35977 | 1.2967 | 2.5482 | 100.99 | 20.69 | 17.965 | 110.2285 | 118.3195 | 102.6553 | 0.004345 | 0.93944 | 0.022244 | 23.6784 | 0.20934 | 0.15612 | 0.041693 | 2.7767 | 2.7767 | 0.002311
2019-03-20 12:54 | 1553100870.169 | 287 | 0.20114 | 0.006649 | 0.000194 | 110.0831 | 107.4278 | 109.6457 | 20.978 | 0.003671 | 0.003445 | 153.0569 | 0.26255 | 8.4018 | 0.003921 | 0.005284 | 0.005234 | 5.2e-05 | 0.001109 | 0.3366 | 1.3554 | 2.5481 | 100.99 | 20.69 | 17.955 | 109.9555 | 116.0419 | 104.4896 | 0.004343 | 0.93487 | 0.022146 | 23.689 | 0.2184 | 0.15619 | 0.042462 | 2.8154 | 2.8154 | 0.00228
2019-03-20 12:54 | 1553100871.1681 | 288 | 0.2011 | 0.006699 | 0.000195 | 110.0523 | 107.732 | 109.4242 | 20.9793 | 0.003676 | 0.003409 | 152.9372 | 0.26216 | 8.4014 | 0.003912 | 0.004799 | 0.004693 | 5.5e-05 | 0.001106 | 0.31492 | 1.3508 | 2.5481 | 100.99 | 20.69 | 17.948 | 110.0885 | 115.8254 | 103.597 | 0.004341 | 0.93709 | 0.02215 | 23.637 | 0.1906 | 0.15596 | 0.042613 | 2.8258 | 2.8258 | 0.002268
2019-03-20 12:55 | 1553100872.1685 | 289 | 0.20101 | 0.006759 | 0.000197 | 110.0701 | 107.6905 | 109.3768 | 20.9792 | 0.003762 | 0.003405 | 152.8193 | 0.26206 | 8.4011 | 0.003897 | 0.005143 | 0.005032 | 5.5e-05 | 0.001107 | 0.32731 | 1.3458 | 2.5481 | 100.99 | 20.69 | 17.943 | 112.1696 | 115.4717 | 103.5058 | 0.00434 | 0.93732 | 0.022148 | 23.629 | 0.19549 | 0.1559 | 0.043046 | 2.8541 | 2.8541 | 0.002245
2019-03-20 12:55 | 1553100873.1683 | 290 | 0.20087 | 0.006828 | 0.0002 | 110.116 | 107.9275 | 109.3878 | 20.9786 | 0.003722 | 0.003429 | 152.7156 | 0.26192 | 8.4029 | 0.003854 | 0.005629 | 0.005481 | 5.5e-05 | 0.00111 | 0.33411 | 1.3632 | 2.5479 | 100.99 | 20.69 | 17.932 | 111.1947 | 114.4114 | 104.0845 | 0.004337 | 0.93588 | 0.022132 | 23.6484 | 0.17957 | 0.15582 | 0.043841 | 2.9045 | 2.9045 | 0.002205
2019-03-20 12:55 | 1553100874.1699 | 291 | 0.20072 | 0.006903 | 0.000203 | 110.0426 | 107.6741 | 109.316 | 20.9788 | 0.00375 | 0.003479 | 152.573 | 0.26141 | 8.4031 | 0.00387 | 0.005374 | 0.00527 | 5.6e-05 | 0.001104 | 0.32388 | 1.352 | 2.5481 | 100.99 | 20.69 | 17.927 | 111.8832 | 114.7982 | 105.2981 | 0.004336 | 0.93288 | 0.022089 | 23.6786 | 0.19462 | 0.15551 | 0.044618 | 2.9501 | 2.9501 | 0.002167
2019-03-20 12:55 | 1553100875.1692 | 292 | 0.20061 | 0.006987 | 0.000205 | 110.0371 | 107.7578 | 109.2127 | 20.9785 | 0.003765 | 0.003441 | 152.462 | 0.26247 | 8.4026 | 0.003789 | 0.005523 | 0.00546 | 5.2e-05 | 0.001106 | 0.32012 | 1.3617 | 2.5481 | 100.99 | 20.69 | 17.922 | 112.2635 | 112.8244 | 104.3856 | 0.004335 | 0.93513 | 0.022105 | 23.6389 | 0.18723 | 0.15614 | 0.045176 | 2.9891 | 2.9891 | 0.002147
2019-03-20 12:55 | 1553100876.1686 | 293 | 0.20042 | 0.007078 | 0.000206 | 110.0704 | 107.6214 | 109.1394 | 20.9793 | 0.0038 | 0.003405 | 152.326 | 0.26228 | 8.404 | 0.003837 | 0.005044 | 0.00492 | 5e-05 | 0.001104 | 0.32587 | 1.3484 | 2.548 | 100.99 | 20.69 | 17.916 | 113.1055 | 114.0057 | 103.5137 | 0.004333 | 0.9373 | 0.022124 | 23.6036 | 0.20126 | 0.15603 | 0.046194 | 3.0588 | 3.0588 | 0.002097
2019-03-20 12:55 | 1553100877.1695 | 294 | 0.20041 | 0.007167 | 0.000204 | 110.0449 | 107.3459 | 108.9747 | 20.9788 | 0.003803 | 0.003467 | 152.1917 | 0.26238 | 8.4028 | 0.003797 | 0.005706 | 0.005579 | 5.7e-05 | 0.001103 | 0.28783 | 1.3324 | 2.5481 | 100.99 | 20.69 | 17.914 | 113.1707 | 113.0331 | 105.0236 | 0.004333 | 0.93355 | 0.022063 | 23.633 | 0.22211 | 0.15609 | 0.04613 | 3.0463 | 3.0463 | 0.002106
2019-03-20 12:55 | 1553100878.1685 | 295 | 0.20024 | 0.007254 | 0.000202 | 110.0433 | 107.3586 | 109.1732 | 20.9794 | 0.003776 | 0.003485 | 152.0383 | 0.26186 | 8.4021 | 0.003748 | 0.005398 | 0.005285 | 5.2e-05 | 0.001107 | 0.32477 | 1.3274 | 2.5482 | 100.99 | 20.69 | 17.913 | 112.5305 | 111.8459 | 105.456 | 0.004333 | 0.93249 | 0.02207 | 23.6681 | 0.22092 | 0.15578 | 0.047093 | 3.1109 | 3.1109 | 0.002058
2019-03-20 12:55 | 1553100879.1685 | 296 | 0.20017 | 0.007337 | 0.000199 | 110.1257 | 107.6793 | 109.2695 | 20.9778 | 0.003725 | 0.003449 | 151.9063 | 0.26122 | 8.4016 | 0.003664 | 0.005345 | 0.005192 | 5.7e-05 | 0.00111 | 0.32045 | 1.2946 | 2.5481 | 100.99 | 20.69 | 17.914 | 111.2847 | 109.7842 | 104.5783 | 0.004333 | 0.93465 | 0.022106 | 23.6511 | 0.20094 | 0.1554 | 0.047398 | 3.1362 | 3.1362 | 0.002037
2019-03-20 12:55 | 1553100880.1686 | 297 | 0.20008 | 0.007417 | 0.000197 | 110.0268 | 108.2751 | 109.4091 | 20.9775 | 0.003797 | 0.003525 | 151.7948 | 0.26249 | 8.404 | 0.003749 | 0.004993 | 0.00477 | 5.1e-05 | 0.00111 | 0.29247 | 1.321 | 2.5481 | 100.99 | 20.69 | 17.916 | 113.0421 | 111.8608 | 106.412 | 0.004333 | 0.93014 | 0.022066 | 23.7235 | 0.14355 | 0.15616 | 0.047862 | 3.1613 | 3.1613 | 0.00203
2019-03-20 12:55 | 1553100881.169 | 298 | 0.20006 | 0.007496 | 0.000195 | 110.1374 | 108.1962 | 109.2338 | 20.9779 | 0.003809 | 0.003567 | 151.6371 | 0.261 | 8.4019 | 0.003754 | 0.004885 | 0.004908 | 5.7e-05 | 0.001111 | 0.29927 | 1.333 | 2.548 | 100.99 | 20.69 | 17.921 | 113.3222 | 111.9896 | 107.4444 | 0.004335 | 0.92762 | 0.022019 | 23.7367 | 0.15906 | 0.15527 | 0.047832 | 3.1527 | 3.1527 | 0.002024
2019-03-20 12:55 | 1553100882.1677 | 299 | 0.20005 | 0.007576 | 0.000193 | 110.1277 | 108.5546 | 109.0495 | 20.9779 | 0.003896 | 0.003522 | 151.4949 | 0.26309 | 8.3994 | 0.003782 | 0.005119 | 0.004922 | 4.9e-05 | 0.001106 | 0.30182 | 1.3366 | 2.5481 | 100.99 | 20.69 | 17.926 | 115.434 | 112.6609 | 106.3541 | 0.004336 | 0.93028 | 0.022032 | 23.6827 | 0.12868 | 0.15651 | 0.047853 | 3.156 | 3.156 | 0.002038
2019-03-20 12:55 | 1553100883.1679 | 300 | 0.19992 | 0.007644 | 0.000192 | 110.0132 | 108.5277 | 109.1197 | 20.9765 | 0.00381 | 0.003537 | 151.366 | 0.26281 | 8.4014 | 0.003761 | 0.004893 | 0.004876 | 5.1e-05 | 0.001106 | 0.32709 | 1.3373 | 2.5481 | 100.99 | 20.69 | 17.936 | 113.3579 | 112.1567 | 106.7102 | 0.004338 | 0.92941 | 0.022028 | 23.7014 | 0.1216 | 0.15634 | 0.048548 | 3.2014 | 3.2014 | 0.002007
2019-03-20 12:55 | 1553100884.1688 | 301 | 0.19984 | 0.007708 | 0.000191 | 110.0539 | 108.4376 | 109.2338 | 20.9772 | 0.003904 | 0.003514 | 151.2547 | 0.26314 | 8.4013 | 0.003836 | 0.005078 | 0.004919 | 5.5e-05 | 0.001103 | 0.33157 | 1.3394 | 2.5482 | 100.99 | 20.69 | 17.942 | 115.6311 | 113.979 | 106.155 | 0.00434 | 0.93077 | 0.022056 | 23.6965 | 0.13234 | 0.15654 | 0.048946 | 3.2317 | 3.2317 | 0.001991
2019-03-20 12:55 | 1553100885.1682 | 302 | 0.19975 | 0.007769 | 0.00019 | 110.1982 | 108.2511 | 109.3836 | 20.9781 | 0.003924 | 0.003572 | 151.117 | 0.26242 | 8.402 | 0.00372 | 0.005748 | 0.005654 | 5.2e-05 | 0.001105 | 0.33699 | 1.3699 | 2.5482 | 100.99 | 20.69 | 17.947 | 116.1337 | 111.1611 | 107.5495 | 0.004341 | 0.92736 | 0.022031 | 23.7563 | 0.15945 | 0.15611 | 0.049413 | 3.2587 | 3.2587 | 0.001969
2019-03-20 12:55 | 1553100886.1689 | 303 | 0.19971 | 0.007824 | 0.000188 | 110.0627 | 108.1921 | 109.0871 | 20.978 | 0.003932 | 0.003527 | 150.9652 | 0.26233 | 8.4006 | 0.00377 | 0.005168 | 0.00497 | 5.2e-05 | 0.001111 | 0.32316 | 1.3469 | 2.548 | 100.99 | 20.69 | 17.96 | 116.3132 | 112.3799 | 106.4644 | 0.004344 | 0.93001 | 0.022032 | 23.6902 | 0.15333 | 0.15606 | 0.049552 | 3.2682 | 3.2682 | 0.001963
2019-03-20 12:55 | 1553100887.168 | 304 | 0.19965 | 0.007893 | 0.000186 | 110.0892 | 108.5677 | 109.3253 | 20.9772 | 0.003854 | 0.003585 | 150.8448 | 0.26244 | 8.4021 | 0.003842 | 0.005625 | 0.005534 | 5.7e-05 | 0.00111 | 0.31022 | 1.3514 | 2.5481 | 100.99 | 20.69 | 17.965 | 114.4218 | 114.1284 | 107.8754 | 0.004345 | 0.92657 | 0.022015 | 23.7601 | 0.12448 | 0.15613 | 0.049838 | 3.2846 | 3.2846 | 0.001954
2019-03-20 12:55 | 1553100888.169 | 305 | 0.19965 | 0.007952 | 0.000185 | 110.0413 | 108.3793 | 109.3167 | 20.9777 | 0.00393 | 0.003653 | 150.7276 | 0.26303 | 8.4015 | 0.003863 | 0.004894 | 0.004772 | 5.3e-05 | 0.001112 | 0.32045 | 1.3635 | 2.5481 | 100.99 | 20.69 | 17.976 | 116.268 | 114.6491 | 109.5192 | 0.004348 | 0.92259 | 0.021967 | 23.8104 | 0.13612 | 0.15648 | 0.049751 | 3.2718 | 3.2718 | 0.001966
2019-03-20 12:55 | 1553100889.1694 | 306 | 0.1996 | 0.008009 | 0.000184 | 110.0442 | 108.6641 | 109.6267 | 20.9772 | 0.003925 | 0.003593 | 150.5959 | 0.26214 | 8.4036 | 0.003879 | 0.005205 | 0.005064 | 5.7e-05 | 0.001114 | 0.30596 | 1.3228 | 2.5481 | 100.99 | 20.69 | 17.981 | 116.143 | 115.0164 | 108.0803 | 0.004349 | 0.92607 | 0.02204 | 23.7992 | 0.11289 | 0.15595 | 0.049997 | 3.2989 | 3.2989 | 0.001943
2019-03-20 12:55 | 1553100890.1697 | 307 | 0.1996 | 0.008073 | 0.000184 | 110.0417 | 108.7751 | 109.6103 | 20.977 | 0.00396 | 0.00358 | 150.459 | 0.26305 | 8.4039 | 0.003858 | 0.005364 | 0.005336 | 5.7e-05 | 0.001115 | 0.3314 | 1.3179 | 2.5481 | 100.99 | 20.69 | 17.986 | 116.9922 | 114.5276 | 107.7429 | 0.00435 | 0.92689 | 0.022048 | 23.7869 | 0.10354 | 0.15649 | 0.049945 | 3.2967 | 3.2967 | 0.001951
2019-03-20 12:55 | 1553100891.1695 | 308 | 0.19988 | 0.008129 | 0.000186 | 110.0507 | 108.6806 | 109.7474 | 20.9766 | 0.004002 | 0.003657 | 150.368 | 0.2631 | 8.4034 | 0.003818 | 0.005024 | 0.004882 | 5.8e-05 | 0.001117 | 0.31542 | 1.2892 | 2.5481 | 100.99 | 20.69 | 17.999 | 118.0233 | 113.5344 | 109.6144 | 0.004354 | 0.92236 | 0.022008 | 23.8602 | 0.11206 | 0.15652 | 0.048203 | 3.1759 | 3.1759 | 0.002026
2019-03-20 12:55 | 1553100892.1699 | 309 | 0.20014 | 0.008164 | 0.000189 | 110.0536 | 108.842 | 109.8446 | 20.977 | 0.003906 | 0.003645 | 150.2654 | 0.26243 | 8.4026 | 0.003839 | 0.005907 | 0.005759 | 5.4e-05 | 0.001111 | 0.30917 | 1.303 | 2.5479 | 100.99 | 20.69 | 18.004 | 115.6793 | 114.0644 | 109.3426 | 0.004355 | 0.92301 | 0.022025 | 23.8623 | 0.099017 | 0.15612 | 0.046561 | 3.0702 | 3.0702 | 0.00209
2019-03-20 12:55 | 1553100893.1683 | 310 | 0.19988 | 0.008179 | 0.000189 | 110.0876 | 108.7079 | 109.7002 | 20.9767 | 0.00399 | 0.003698 | 150.1421 | 0.26195 | 8.4049 | 0.003855 | 0.00557 | 0.005446 | 5.5e-05 | 0.001119 | 0.31022 | 1.327 | 2.5481 | 100.99 | 20.69 | 18.01 | 117.7238 | 114.4403 | 110.6325 | 0.004356 | 0.91991 | 0.021974 | 23.8868 | 0.11281 | 0.15584 | 0.048136 | 3.1665 | 3.1665 | 0.002023
2019-03-20 12:55 | 1553100894.1683 | 311 | 0.19987 | 0.008183 | 0.000188 | 110.0879 | 108.8495 | 109.5317 | 20.9754 | 0.004011 | 0.003661 | 149.9935 | 0.2626 | 8.4022 | 0.003827 | 0.005507 | 0.00536 | 5.3e-05 | 0.001118 | 0.32997 | 1.3305 | 2.5482 | 100.99 | 20.69 | 18.02 | 118.2392 | 113.7631 | 109.7211 | 0.004359 | 0.9221 | 0.021983 | 23.8401 | 0.10119 | 0.15622 | 0.048155 | 3.1691 | 3.1691 | 0.002026
2019-03-20 12:55 | 1553100895.1687 | 312 | 0.19986 | 0.00817 | 0.000185 | 110.0801 | 108.7183 | 109.5657 | 20.9765 | 0.004001 | 0.003697 | 149.8624 | 0.26224 | 8.4015 | 0.003769 | 0.005185 | 0.005137 | 5.4e-05 | 0.001118 | 0.36481 | 1.3278 | 2.548 | 100.99 | 20.69 | 18.023 | 118.0088 | 112.3366 | 110.5922 | 0.004359 | 0.92001 | 0.021961 | 23.8709 | 0.11134 | 0.15601 | 0.04828 | 3.1744 | 3.1744 | 0.00202
2019-03-20 12:55 | 1553100896.1688 | 313 | 0.19984 | 0.008143 | 0.000183 | 110.0012 | 108.3917 | 109.6432 | 20.9755 | 0.00401 | 0.00363 | 149.764 | 0.26237 | 8.4035 | 0.003796 | 0.005038 | 0.004866 | 5.4e-05 | 0.001113 | 0.32405 | 1.3082 | 2.5482 | 100.99 | 20.69 | 18.026 | 118.2227 | 113.0055 | 108.9645 | 0.00436 | 0.92393 | 0.022016 | 23.8286 | 0.13185 | 0.15608 | 0.048437 | 3.1927 | 3.1927 | 0.002009
2019-03-20 12:55 | 1553100897.1685 | 314 | 0.19969 | 0.008133 | 0.00018 | 110.0908 | 108.4399 | 109.1885 | 20.975 | 0.004034 | 0.003722 | 149.6411 | 0.26231 | 8.4016 | 0.003913 | 0.005595 | 0.005434 | 5.5e-05 | 0.001108 | 0.32621 | 1.3294 | 2.548 | 100.99 | 20.69 | 18.026 | 118.8109 | 115.8504 | 111.1936 | 0.00436 | 0.91857 | 0.021906 | 23.8484 | 0.13515 | 0.15605 | 0.049298 | 3.2333 | 3.2333 | 0.001984
2019-03-20 12:55 | 1553100898.1692 | 315 | 0.1996 | 0.008148 | 0.00018 | 110.0957 | 108.7635 | 108.8313 | 20.9745 | 0.004009 | 0.003678 | 149.5214 | 0.263 | 8.4041 | 0.003786 | 0.005151 | 0.005085 | 5.3e-05 | 0.001108 | 0.31404 | 1.3254 | 2.548 | 100.99 | 20.69 | 18.022 | 118.2016 | 112.7698 | 110.1389 | 0.004359 | 0.92109 | 0.021901 | 23.7767 | 0.10889 | 0.15646 | 0.049822 | 3.2668 | 3.2668 | 0.001968
2019-03-20 12:55 | 1553100899.1689 | 316 | 0.19959 | 0.008173 | 0.000179 | 110.0704 | 108.7221 | 108.239 | 20.9759 | 0.00407 | 0.003691 | 149.4051 | 0.26235 | 8.403 | 0.003886 | 0.005094 | 0.004875 | 5.1e-05 | 0.001109 | 0.32493 | 1.3445 | 2.548 | 100.99 | 20.69 | 18.018 | 119.6793 | 115.201 | 110.4616 | 0.004358 | 0.92032 | 0.021832 | 23.7219 | 0.11025 | 0.15607 | 0.049869 | 3.2596 | 3.2596 | 0.001968
2019-03-20 12:55 | 1553100900.1687 | 317 | 0.19955 | 0.008202 | 0.000178 | 110.0035 | 108.9444 | 108.2717 | 20.9757 | 0.004046 | 0.003736 | 149.273 | 0.26354 | 8.4018 | 0.003821 | 0.005409 | 0.005276 | 5.4e-05 | 0.00111 | 0.33085 | 1.3935 | 2.548 | 100.99 | 20.69 | 18.012 | 119.0923 | 113.607 | 111.5408 | 0.004357 | 0.91774 | 0.021804 | 23.7588 | 0.08653 | 0.15678 | 0.050072 | 3.2688 | 3.2688 | 0.001971
2019-03-20 12:55 | 1553100901.1677 | 318 | 0.19953 | 0.008233 | 0.000175 | 110.0536 | 108.7789 | 108.097 | 20.9752 | 0.004054 | 0.003717 | 149.1355 | 0.26176 | 8.4023 | 0.003745 | 0.005226 | 0.005157 | 5.3e-05 | 0.001114 | 0.33572 | 1.3563 | 2.5481 | 100.99 | 20.69 | 17.998 | 119.29 | 111.7731 | 111.0873 | 0.004353 | 0.91882 | 0.0218 | 23.7256 | 0.10421 | 0.15572 | 0.050177 | 3.2751 | 3.2751 | 0.001954
2019-03-20 12:55 | 1553100902.1677 | 319 | 0.19937 | 0.00826 | 0.000172 | 110.0365 | 108.5681 | 108.052 | 20.9752 | 0.004093 | 0.003716 | 148.9941 | 0.26251 | 8.4023 | 0.003794 | 0.005195 | 0.005075 | 5.2e-05 | 0.001116 | 0.32731 | 1.3555 | 2.5479 | 100.99 | 20.69 | 17.988 | 120.2434 | 112.9585 | 111.0645 | 0.004351 | 0.91888 | 0.021796 | 23.72 | 0.12017 | 0.15617 | 0.051073 | 3.3331 | 3.3331 | 0.001926
2019-03-20 12:55 | 1553100903.1678 | 320 | 0.19913 | 0.008304 | 0.000168 | 110.0297 | 108.9224 | 107.7445 | 20.9755 | 0.004135 | 0.003756 | 148.8538 | 0.26226 | 8.4053 | 0.00391 | 0.005536 | 0.005379 | 5.4e-05 | 0.001119 | 0.34075 | 1.3383 | 2.5481 | 100.99 | 20.69 | 17.978 | 121.2775 | 115.7813 | 112.0228 | 0.004348 | 0.91659 | 0.021738 | 23.7157 | 0.090471 | 0.15602 | 0.052473 | 3.4153 | 3.4153 | 0.001878
2019-03-20 12:55 | 1553100904.1679 | 321 | 0.19897 | 0.008353 | 0.000166 | 110.062 | 108.792 | 107.1971 | 20.9745 | 0.00412 | 0.003738 | 148.7311 | 0.2615 | 8.4042 | 0.003897 | 0.005011 | 0.004948 | 5.2e-05 | 0.001109 | 0.31996 | 1.3076 | 2.548 | 100.99 | 20.68 | 17.96 | 120.9171 | 115.4709 | 111.5965 | 0.004341 | 0.91761 | 0.021694 | 23.6423 | 0.10382 | 0.15557 | 0.053361 | 3.4661 | 3.4661 | 0.001845
2019-03-20 12:55 | 1553100905.1689 | 322 | 0.19894 | 0.008419 | 0.000165 | 110.0481 | 108.8757 | 107.0851 | 20.9753 | 0.004112 | 0.003752 | 148.6081 | 0.26252 | 8.4017 | 0.003838 | 0.005616 | 0.005473 | 5.2e-05 | 0.001108 | 0.31741 | 1.3271 | 2.548 | 100.99 | 20.68 | 17.951 | 120.7059 | 114.031 | 111.9255 | 0.004339 | 0.91682 | 0.021674 | 23.6401 | 0.095806 | 0.15617 | 0.053491 | 3.4713 | 3.4713 | 0.001849
2019-03-20 12:55 | 1553100906.1689 | 323 | 0.19882 | 0.008505 | 0.000164 | 109.9773 | 108.728 | 107.2185 | 20.9756 | 0.004123 | 0.003763 | 148.4749 | 0.26241 | 8.4034 | 0.003819 | 0.0056 | 0.00554 | 5.2e-05 | 0.00111 | 0.34202 | 1.2959 | 2.548 | 100.99 | 20.68 | 17.934 | 120.9746 | 113.5619 | 112.2055 | 0.004335 | 0.91616 | 0.021679 | 23.6634 | 0.10218 | 0.15611 | 0.054077 | 3.5102 | 3.5102 | 0.001828
2019-03-20 12:55 | 1553100907.169 | 324 | 0.19879 | 0.008596 | 0.000163 | 110.0672 | 108.7041 | 107.2969 | 20.9753 | 0.004096 | 0.00384 | 148.3274 | 0.2626 | 8.4036 | 0.003833 | 0.005504 | 0.005436 | 5.2e-05 | 0.001112 | 0.31708 | 1.3394 | 2.5481 | 100.99 | 20.68 | 17.929 | 120.3203 | 113.8966 | 114.0649 | 0.004334 | 0.91176 | 0.021635 | 23.7291 | 0.11146 | 0.15622 | 0.054161 | 3.5085 | 3.5085 | 0.00183
2019-03-20 12:55 | 1553100908.1684 | 325 | 0.19878 | 0.008672 | 0.000161 | 109.9983 | 108.8648 | 107.3762 | 20.9754 | 0.004187 | 0.003786 | 148.1778 | 0.26233 | 8.403 | 0.003907 | 0.005577 | 0.005421 | 5.1e-05 | 0.001107 | 0.33445 | 1.3543 | 2.5481 | 100.99 | 20.68 | 17.923 | 122.5489 | 115.7125 | 112.7629 | 0.004332 | 0.91483 | 0.02168 | 23.6979 | 0.092648 | 0.15606 | 0.054095 | 3.5115 | 3.5115 | 0.001827
2019-03-20 12:55 | 1553100909.1691 | 326 | 0.19887 | 0.008741 | 0.000161 | 110.105 | 108.894 | 107.4663 | 20.9747 | 0.004234 | 0.003809 | 148.0513 | 0.26321 | 8.403 | 0.003963 | 0.005732 | 0.005605 | 4.8e-05 | 0.001097 | 0.32997 | 1.3463 | 2.5482 | 100.99 | 20.68 | 17.915 | 123.6916 | 117.0771 | 113.3206 | 0.004331 | 0.91351 | 0.021673 | 23.725 | 0.098927 | 0.15659 | 0.0535 | 3.472 | 3.472 | 0.001854
2019-03-20 12:55 | 1553100910.1672 | 327 | 0.19889 | 0.008792 | 0.000161 | 110.0242 | 108.7856 | 107.0321 | 20.9758 | 0.004158 | 0.003783 | 147.9222 | 0.2624 | 8.4036 | 0.003928 | 0.005312 | 0.005223 | 4.7e-05 | 0.001101 | 0.31404 | 1.4067 | 2.5482 | 100.99 | 20.68 | 17.913 | 121.8399 | 116.2319 | 112.6786 | 0.00433 | 0.91503 | 0.021647 | 23.6573 | 0.10126 | 0.1561 | 0.053349 | 3.4581 | 3.4581 | 0.001855
2019-03-20 12:55 | 1553100911.1686 | 328 | 0.19886 | 0.008859 | 0.000163 | 110.0562 | 108.8256 | 107.0544 | 20.9744 | 0.004202 | 0.003862 | 147.7823 | 0.26311 | 8.4037 | 0.003855 | 0.00553 | 0.005416 | 4.8e-05 | 0.0011 | 0.32532 | 1.3571 | 2.5481 | 100.99 | 20.68 | 17.915 | 122.9094 | 114.4309 | 114.6039 | 0.004331 | 0.91049 | 0.021596 | 23.7187 | 0.10058 | 0.15652 | 0.053399 | 3.453 | 3.453 | 0.001863
2019-03-20 12:55 | 1553100912.1683 | 329 | 0.19877 | 0.0089 | 0.000163 | 110.0772 | 108.851 | 107.0617 | 20.9743 | 0.004179 | 0.003859 | 147.6347 | 0.26236 | 8.4044 | 0.003887 | 0.005141 | 0.005041 | 4.8e-05 | 0.001105 | 0.30038 | 1.3455 | 2.5482 | 100.99 | 20.68 | 17.92 | 122.3539 | 115.2137 | 114.5432 | 0.004332 | 0.91063 | 0.021598 | 23.7177 | 0.1002 | 0.15608 | 0.053908 | 3.4862 | 3.4862 | 0.00184
2019-03-20 12:55 | 1553100913.1683 | 330 | 0.19877 | 0.008927 | 0.000163 | 109.9993 | 108.5737 | 106.7807 | 20.974 | 0.00427 | 0.003882 | 147.4848 | 0.26278 | 8.4016 | 0.0039 | 0.005785 | 0.005691 | 5.2e-05 | 0.001109 | 0.32693 | 1.3735 | 2.5481 | 100.99 | 20.68 | 17.924 | 124.5878 | 115.5465 | 115.1091 | 0.004333 | 0.9093 | 0.021554 | 23.7038 | 0.11668 | 0.15633 | 0.05386 | 3.476 | 3.476 | 0.001848
2019-03-20 12:55 | 1553100914.1674 | 331 | 0.19878 | 0.008919 | 0.000162 | 110.0019 | 108.7908 | 107.0403 | 20.9741 | 0.004229 | 0.003867 | 147.3385 | 0.26188 | 8.4021 | 0.003869 | 0.005015 | 0.004843 | 5e-05 | 0.00111 | 0.33323 | 1.3289 | 2.5481 | 100.99 | 20.68 | 17.934 | 123.5712 | 114.7762 | 114.7294 | 0.004335 | 0.91019 | 0.021591 | 23.721 | 0.099019 | 0.15579 | 0.053819 | 3.4793 | 3.4793 | 0.00184
2019-03-20 12:55 | 1553100915.1674 | 332 | 0.19879 | 0.008914 | 0.00016 | 110.0569 | 108.5883 | 106.7363 | 20.9737 | 0.004341 | 0.003868 | 147.196 | 0.26177 | 8.4019 | 0.003831 | 0.005418 | 0.005256 | 4.8e-05 | 0.001112 | 0.29822 | 1.3266 | 2.5481 | 100.99 | 20.68 | 17.94 | 126.3138 | 113.8579 | 114.752 | 0.004337 | 0.91014 | 0.021559 | 23.688 | 0.12016 | 0.15572 | 0.053805 | 3.4734 | 3.4734 | 0.001843
2019-03-20 12:55 | 1553100916.1685 | 333 | 0.1987 | 0.008918 | 0.000158 | 110.0155 | 108.663 | 106.4969 | 20.9732 | 0.004343 | 0.003896 | 147.0432 | 0.26084 | 8.4001 | 0.003821 | 0.005146 | 0.005079 | 5.3e-05 | 0.001115 | 0.32278 | 1.2818 | 2.548 | 100.99 | 20.68 | 17.945 | 126.3665 | 113.5989 | 115.4484 | 0.004338 | 0.90851 | 0.021516 | 23.6827 | 0.11064 | 0.15517 | 0.054335 | 3.5006 | 3.5006 | 0.001822
2019-03-20 12:55 | 1553100917.1685 | 334 | 0.19863 | 0.008931 | 0.000156 | 110.0009 | 108.8648 | 106.3521 | 20.9727 | 0.004285 | 0.003914 | 146.9234 | 0.26253 | 8.4027 | 0.003755 | 0.00516 | 0.004976 | 5e-05 | 0.001117 | 0.34108 | 1.3402 | 2.5482 | 100.99 | 20.68 | 17.958 | 124.9462 | 111.9981 | 115.8657 | 0.004341 | 0.90753 | 0.02149 | 23.6792 | 0.092858 | 0.15618 | 0.054722 | 3.5213 | 3.5213 | 0.001823
2019-03-20 12:55 | 1553100918.169 | 335 | 0.1986 | 0.008949 | 0.000154 | 109.9792 | 108.919 | 106.0094 | 20.9724 | 0.004301 | 0.003899 | 146.7853 | 0.26191 | 8.4032 | 0.003813 | 0.005189 | 0.005092 | 5.1e-05 | 0.001114 | 0.34634 | 1.3415 | 2.548 | 100.99 | 20.68 | 17.963 | 125.3358 | 113.4101 | 115.5086 | 0.004342 | 0.90837 | 0.021465 | 23.6302 | 0.086645 | 0.15581 | 0.054895 | 3.5284 | 3.5284 | 0.001815
2019-03-20 12:55 | 1553100919.1686 | 336 | 0.19829 | 0.008986 | 0.000154 | 109.9744 | 108.9822 | 105.9827 | 20.9722 | 0.004268 | 0.003936 | 146.6253 | 0.26224 | 8.4045 | 0.003881 | 0.004633 | 0.004477 | 5.2e-05 | 0.001114 | 0.34971 | 1.324 | 2.5482 | 100.99 | 20.68 | 17.972 | 124.5432 | 115.07 | 116.4019 | 0.004344 | 0.90629 | 0.021438 | 23.6544 | 0.081064 | 0.15601 | 0.056649 | 3.6361 | 3.6361 | 0.001763
2019-03-20 12:55 | 1553100920.1691 | 337 | 0.19841 | 0.009041 | 0.000156 | 109.9686 | 109.0349 | 106.2627 | 20.9719 | 0.004334 | 0.003961 | 146.4556 | 0.26121 | 8.4031 | 0.003862 | 0.005841 | 0.005666 | 5.2e-05 | 0.001117 | 0.30486 | 1.3342 | 2.548 | 100.99 | 20.68 | 17.975 | 126.1623 | 114.5962 | 117.0207 | 0.004345 | 0.90485 | 0.021449 | 23.7044 | 0.076267 | 0.1554 | 0.05589 | 3.5894 | 3.5894 | 0.001779
2019-03-20 12:55 | 1553100921.1688 | 338 | 0.19819 | 0.009101 | 0.000158 | 109.9272 | 108.9377 | 106.0258 | 20.9704 | 0.004322 | 0.003937 | 146.2883 | 0.26231 | 8.4026 | 0.003865 | 0.005427 | 0.005353 | 5.1e-05 | 0.001118 | 0.36602 | 1.3622 | 2.5479 | 100.99 | 20.68 | 17.979 | 125.8487 | 114.6744 | 116.4262 | 0.004346 | 0.90623 | 0.021441 | 23.6599 | 0.080874 | 0.15605 | 0.057136 | 3.6676 | 3.6676 | 0.001749
2019-03-20 12:55 | 1553100922.1677 | 339 | 0.19813 | 0.009169 | 0.000162 | 109.9058 | 109.0046 | 106.0944 | 20.9714 | 0.004445 | 0.00398 | 146.1351 | 0.26191 | 8.4027 | 0.003796 | 0.005798 | 0.005611 | 5.3e-05 | 0.001117 | 0.33572 | 1.3343 | 2.5482 | 100.99 | 20.68 | 17.983 | 128.8751 | 113.01 | 117.4968 | 0.004347 | 0.90375 | 0.021419 | 23.7001 | 0.073646 | 0.15581 | 0.057399 | 3.6804 | 3.6804 | 0.00174
2019-03-20 12:55 | 1553100923.1682 | 340 | 0.19811 | 0.009219 | 0.000165 | 109.8887 | 108.4944 | 106.2716 | 20.9714 | 0.004342 | 0.004 | 145.9998 | 0.2627 | 8.4054 | 0.003919 | 0.005293 | 0.005302 | 5.6e-05 | 0.001117 | 0.3194 | 1.3337 | 2.5481 | 100.99 | 20.68 | 17.984 | 126.3377 | 115.9998 | 117.9808 | 0.004347 | 0.90263 | 0.021423 | 23.7345 | 0.11421 | 0.15628 | 0.057474 | 3.6858 | 3.6858 | 0.001743
2019-03-20 12:55 | 1553100924.168 | 341 | 0.19809 | 0.009262 | 0.000166 | 109.8926 | 108.4275 | 106.1045 | 20.9709 | 0.004382 | 0.00395 | 145.8686 | 0.26251 | 8.402 | 0.003872 | 0.005331 | 0.005115 | 5.4e-05 | 0.001124 | 0.34324 | 1.3886 | 2.5481 | 100.99 | 20.68 | 17.98 | 127.3365 | 114.8489 | 116.7622 | 0.004346 | 0.90545 | 0.02144 | 23.6789 | 0.12004 | 0.15617 | 0.057551 | 3.6935 | 3.6935 | 0.001738
2019-03-20 12:55 | 1553100925.168 | 342 | 0.19808 | 0.009323 | 0.000166 | 109.8751 | 108.4373 | 105.6637 | 20.9703 | 0.004429 | 0.003914 | 145.7088 | 0.2627 | 8.3994 | 0.003902 | 0.005718 | 0.005676 | 5.3e-05 | 0.001123 | 0.33716 | 1.3936 | 2.5481 | 100.99 | 20.67 | 17.977 | 128.4802 | 115.5856 | 115.8799 | 0.004343 | 0.9075 | 0.02142 | 23.6029 | 0.11782 | 0.15628 | 0.057517 | 3.6879 | 3.6879 | 0.001742
2019-03-20 12:55 | 1553100926.167 | 343 | 0.19809 | 0.009388 | 0.000166 | 109.7613 | 108.1028 | 105.4313 | 20.9715 | 0.004437 | 0.004004 | 145.5536 | 0.26215 | 8.4017 | 0.003834 | 0.005923 | 0.005754 | 5.3e-05 | 0.001124 | 0.33069 | 1.3405 | 2.548 | 100.99 | 20.67 | 17.975 | 128.6863 | 113.9364 | 118.0768 | 0.004342 | 0.90241 | 0.021336 | 23.6434 | 0.13618 | 0.15595 | 0.05738 | 3.6647 | 3.6647 | 0.001749
2019-03-20 12:55 | 1553100927.1685 | 344 | 0.19819 | 0.009437 | 0.000166 | 109.8609 | 108.2631 | 105.7073 | 20.9712 | 0.004438 | 0.003986 | 145.4016 | 0.26286 | 8.4001 | 0.003934 | 0.005214 | 0.005073 | 5.1e-05 | 0.001127 | 0.31813 | 1.3081 | 2.548 | 100.99 | 20.67 | 17.966 | 128.698 | 116.3667 | 117.636 | 0.00434 | 0.90342 | 0.021376 | 23.661 | 0.13104 | 0.15638 | 0.056721 | 3.6296 | 3.6296 | 0.001771
2019-03-20 12:55 | 1553100928.1681 | 345 | 0.19826 | 0.00947 | 0.000166 | 109.8861 | 108.506 | 106.1463 | 20.9705 | 0.00443 | 0.004017 | 145.261 | 0.26189 | 8.4005 | 0.003834 | 0.005306 | 0.005151 | 5.5e-05 | 0.001121 | 0.3126 | 1.2933 | 2.548 | 100.99 | 20.67 | 17.962 | 128.5069 | 113.9315 | 118.3919 | 0.004339 | 0.90168 | 0.0214 | 23.733 | 0.11305 | 0.1558 | 0.056289 | 3.606 | 3.606 | 0.001776
2019-03-20 12:55 | 1553100929.1682 | 346 | 0.19824 | 0.00948 | 0.000164 | 109.8638 | 108.6193 | 106.4162 | 20.9701 | 0.004399 | 0.003999 | 145.1085 | 0.26249 | 8.4023 | 0.003828 | 0.005774 | 0.005688 | 5.3e-05 | 0.001121 | 0.29927 | 1.331 | 2.5481 | 100.99 | 20.67 | 17.955 | 127.7349 | 113.7847 | 117.9541 | 0.004338 | 0.90269 | 0.021439 | 23.7499 | 0.1019 | 0.15615 | 0.05639 | 3.6192 | 3.6192 | 0.001773
2019-03-20 12:55 | 1553100930.1676 | 347 | 0.19824 | 0.009469 | 0.000163 | 109.7946 | 108.6436 | 106.4364 | 20.9705 | 0.004467 | 0.004005 | 144.9792 | 0.26178 | 8.4052 | 0.003818 | 0.005213 | 0.00506 | 5.8e-05 | 0.001126 | 0.30646 | 1.3022 | 2.548 | 100.99 | 20.67 | 17.953 | 129.4117 | 113.529 | 118.0971 | 0.004337 | 0.90236 | 0.021437 | 23.7565 | 0.094262 | 0.15573 | 0.056385 | 3.6186 | 3.6186 | 0.001769
2019-03-20 12:56 | 1553100931.1673 | 348 | 0.1983 | 0.009455 | 0.000163 | 109.7364 | 108.5841 | 106.5521 | 20.9702 | 0.004474 | 0.00402 | 144.8458 | 0.26292 | 8.4044 | 0.003875 | 0.005542 | 0.005498 | 5.3e-05 | 0.001124 | 0.33483 | 1.3306 | 2.548 | 100.99 | 20.67 | 17.952 | 129.598 | 114.9162 | 118.4646 | 0.004337 | 0.90151 | 0.021438 | 23.7805 | 0.094412 | 0.15641 | 0.056059 | 3.5981 | 3.5981 | 0.001787
2019-03-20 12:56 | 1553100932.1673 | 349 | 0.19822 | 0.009455 | 0.000164 | 109.7836 | 108.3468 | 106.1249 | 20.9703 | 0.004509 | 0.004011 | 144.7251 | 0.26245 | 8.4018 | 0.003839 | 0.004791 | 0.004611 | 5.5e-05 | 0.001122 | 0.31188 | 1.3213 | 2.5481 | 100.99 | 20.67 | 17.953 | 130.447 | 114.0514 | 118.2515 | 0.004337 | 0.902 | 0.021401 | 23.7264 | 0.11783 | 0.15613 | 0.056558 | 3.6236 | 3.6236 | 0.001771
2019-03-20 12:56 | 1553100933.1686 | 350 | 0.19828 | 0.009465 | 0.000165 | 109.7066 | 108.3554 | 106.2373 | 20.9698 | 0.004441 | 0.004091 | 144.595 | 0.26132 | 8.4016 | 0.003792 | 0.005395 | 0.005293 | 5.7e-05 | 0.001115 | 0.34307 | 1.3204 | 2.5481 | 100.99 | 20.66 | 17.954 | 128.7698 | 112.8986 | 120.1999 | 0.004335 | 0.89753 | 0.021359 | 23.7979 | 0.11085 | 0.15546 | 0.056137 | 3.5897 | 3.5897 | 0.00178
2019-03-20 12:56 | 1553100934.1674 | 351 | 0.19833 | 0.009474 | 0.000166 | 109.7043 | 108.3524 | 106.2502 | 20.9702 | 0.004436 | 0.004065 | 144.4347 | 0.26242 | 8.4026 | 0.003822 | 0.00536 | 0.005206 | 5.4e-05 | 0.001117 | 0.33124 | 1.3638 | 2.5482 | 100.99 | 20.66 | 17.956 | 128.6548 | 113.6299 | 119.5571 | 0.004335 | 0.899 | 0.021378 | 23.7799 | 0.11091 | 0.15611 | 0.055839 | 3.5737 | 3.5737 | 0.001795
2019-03-20 12:56 | 1553100935.1676 | 352 | 0.19837 | 0.00947 | 0.000169 | 109.6853 | 108.3307 | 106.2542 | 20.9695 | 0.004555 | 0.004038 | 144.3087 | 0.26237 | 8.4027 | 0.003847 | 0.005275 | 0.00519 | 5.6e-05 | 0.001116 | 0.32549 | 1.3725 | 2.5482 | 100.99 | 20.66 | 17.963 | 131.5758 | 114.2317 | 118.8983 | 0.004337 | 0.90051 | 0.021397 | 23.7604 | 0.11114 | 0.15608 | 0.055638 | 3.5638 | 3.5638 | 0.0018
2019-03-20 12:56 | 1553100936.1685 | 353 | 0.19836 | 0.009452 | 0.00017 | 109.6956 | 107.9656 | 106.296 | 20.9704 | 0.004555 | 0.004039 | 144.1718 | 0.26305 | 8.403 | 0.003847 | 0.005783 | 0.005638 | 5.5e-05 | 0.00112 | 0.28838 | 1.3465 | 2.5479 | 100.99 | 20.66 | 17.968 | 131.572 | 114.2467 | 118.9237 | 0.004338 | 0.90046 | 0.0214 | 23.7659 | 0.14219 | 0.15649 | 0.055674 | 3.5666 | 3.5666 | 0.001803
2019-03-20 12:56 | 1553100937.1687 | 354 | 0.19831 | 0.009424 | 0.00017 | 109.5003 | 108.0863 | 106.0134 | 20.9702 | 0.004519 | 0.004116 | 144.0214 | 0.26263 | 8.4072 | 0.003908 | 0.005514 | 0.005401 | 5.6e-05 | 0.001126 | 0.33716 | 1.3516 | 2.5479 | 100.99 | 20.66 | 17.979 | 130.7047 | 115.7215 | 120.8041 | 0.004341 | 0.89616 | 0.021321 | 23.7911 | 0.11625 | 0.15624 | 0.055994 | 3.5737 | 3.5737 | 0.001797
2019-03-20 12:56 | 1553100938.1676 | 355 | 0.19837 | 0.009385 | 0.00017 | 109.5248 | 108.1031 | 105.8559 | 20.9691 | 0.004518 | 0.00408 | 143.8674 | 0.26222 | 8.4008 | 0.003901 | 0.004885 | 0.004734 | 5.6e-05 | 0.001125 | 0.34866 | 1.3851 | 2.5482 | 100.99 | 20.65 | 17.986 | 130.6716 | 115.559 | 119.9319 | 0.00434 | 0.89815 | 0.021328 | 23.7471 | 0.11687 | 0.15599 | 0.055736 | 3.5586 | 3.5586 | 0.001802
2019-03-20 12:56 | 1553100939.1676 | 356 | 0.19835 | 0.009348 | 0.000168 | 109.4411 | 108.1521 | 105.4414 | 20.9709 | 0.004551 | 0.004071 | 143.723 | 0.26219 | 8.4021 | 0.003924 | 0.005139 | 0.004956 | 5.8e-05 | 0.001121 | 0.31797 | 1.3817 | 2.548 | 100.99 | 20.65 | 17.992 | 131.4934 | 116.1266 | 119.702 | 0.004341 | 0.89867 | 0.021293 | 23.6936 | 0.10597 | 0.15598 | 0.055901 | 3.5633 | 3.5633 | 0.001799
2019-03-20 12:56 | 1553100940.1678 | 357 | 0.19834 | 0.00933 | 0.000167 | 109.5155 | 108.1981 | 105.0013 | 20.9703 | 0.004539 | 0.004096 | 143.593 | 0.26275 | 8.4017 | 0.003981 | 0.005414 | 0.00528 | 5.8e-05 | 0.001115 | 0.33157 | 1.3902 | 2.5481 | 100.99 | 20.65 | 18.007 | 131.1973 | 117.5187 | 120.3131 | 0.004345 | 0.89728 | 0.021232 | 23.6625 | 0.10825 | 0.15631 | 0.055942 | 3.5557 | 3.5557 | 0.001807
2019-03-20 12:56 | 1553100941.1678 | 358 | 0.19834 | 0.009325 | 0.000166 | 109.4006 | 108.1158 | 104.5645 | 20.9695 | 0.004557 | 0.004116 | 143.459 | 0.26198 | 8.4018 | 0.003852 | 0.00511 | 0.004975 | 5.3e-05 | 0.001115 | 0.29695 | 1.3718 | 2.5482 | 100.99 | 20.65 | 18.007 | 131.6397 | 114.3652 | 120.8115 | 0.004345 | 0.89614 | 0.021174 | 23.6282 | 0.10567 | 0.15585 | 0.055978 | 3.5484 | 3.5484 | 0.001805
2019-03-20 12:56 | 1553100942.1679 | 359 | 0.19835 | 0.009336 | 0.000166 | 109.423 | 108.1996 | 104.5236 | 20.9692 | 0.004599 | 0.004085 | 143.3122 | 0.26228 | 8.4009 | 0.003906 | 0.005886 | 0.005729 | 5.4e-05 | 0.001115 | 0.34164 | 1.3405 | 2.5481 | 100.99 | 20.65 | 18.019 | 132.6722 | 115.6881 | 120.055 | 0.004348 | 0.89787 | 0.02119 | 23.6009 | 0.10057 | 0.15603 | 0.055918 | 3.5473 | 3.5473 | 0.001808
2019-03-20 12:56 | 1553100943.1669 | 360 | 0.19829 | 0.009346 | 0.000165 | 109.3434 | 108.4545 | 104.1678 | 20.9695 | 0.004586 | 0.004109 | 143.1575 | 0.2621 | 8.4013 | 0.003899 | 0.005792 | 0.005685 | 5.8e-05 | 0.00111 | 0.32781 | 1.3487 | 2.548 | 100.99 | 20.65 | 18.024 | 132.344 | 115.497 | 120.6415 | 0.004349 | 0.89653 | 0.021139 | 23.5782 | 0.073015 | 0.15593 | 0.056235 | 3.5587 | 3.5587 | 0.001801
2019-03-20 12:56 | 1553100944.168 | 361 | 0.19829 | 0.00935 | 0.000165 | 109.2574 | 108.302 | 104.2696 | 20.9694 | 0.00458 | 0.004134 | 143.0221 | 0.26246 | 8.4018 | 0.003891 | 0.005903 | 0.005815 | 5.8e-05 | 0.001105 | 0.32477 | 1.3437 | 2.548 | 100.99 | 20.65 | 18.034 | 132.1916 | 115.3197 | 121.2447 | 0.004351 | 0.89516 | 0.021133 | 23.6078 | 0.078561 | 0.15614 | 0.056224 | 3.557 | 3.557 | 0.001804
2019-03-20 12:56 | 1553100945.167 | 362 | 0.19814 | 0.009351 | 0.000164 | 109.2807 | 108.1723 | 104.5438 | 20.9702 | 0.004603 | 0.004076 | 142.8877 | 0.26217 | 8.4036 | 0.003907 | 0.005924 | 0.005828 | 5.4e-05 | 0.001103 | 0.3501 | 1.3498 | 2.5482 | 100.99 | 20.65 | 18.038 | 132.7508 | 115.6986 | 119.8196 | 0.004352 | 0.8984 | 0.021199 | 23.5961 | 0.091184 | 0.15596 | 0.057153 | 3.6269 | 3.6269 | 0.001767
2019-03-20 12:56 | 1553100946.167 | 363 | 0.19801 | 0.009363 | 0.000163 | 109.2603 | 108.2799 | 103.9943 | 20.9684 | 0.004597 | 0.004112 | 142.7668 | 0.26181 | 8.4001 | 0.003905 | 0.005332 | 0.005259 | 5.4e-05 | 0.001101 | 0.30845 | 1.3032 | 2.5481 | 100.99 | 20.65 | 18.044 | 132.6084 | 115.6471 | 120.7047 | 0.004354 | 0.89638 | 0.021119 | 23.5605 | 0.08062 | 0.15575 | 0.057916 | 3.6614 | 3.6614 | 0.001748
2019-03-20 12:56 | 1553100947.2388 | 364 | 0.19801 | 0.009385 | 0.000163 | 109.2548 | 108.1121 | 103.8363 | 20.9691 | 0.004617 | 0.004119 | 142.6167 | 0.2618 | 8.4024 | 0.003985 | 0.005306 | 0.005144 | 5.4e-05 | 0.001098 | 0.33572 | 1.3754 | 2.5481 | 100.99 | 20.64 | 18.054 | 133.1019 | 117.6149 | 120.8864 | 0.004353 | 0.89597 | 0.021098 | 23.548 | 0.094043 | 0.15574 | 0.057868 | 3.6548 | 3.6548 | 0.001751
2019-03-20 12:56 | 1553100948.1797 | 365 | 0.19796 | 0.009416 | 0.000163 | 109.1477 | 107.9809 | 103.9201 | 20.9699 | 0.004634 | 0.004178 | 142.4816 | 0.26208 | 8.4013 | 0.003927 | 0.005687 | 0.005546 | 5.1e-05 | 0.001102 | 0.3246 | 1.3808 | 2.5481 | 100.99 | 20.64 | 18.06 | 133.5176 | 116.1803 | 122.3132 | 0.004355 | 0.89274 | 0.021069 | 23.6001 | 0.096135 | 0.15591 | 0.058117 | 3.6653 | 3.6653 | 0.001748
2019-03-20 12:56 | 1553100949.1798 | 366 | 0.19795 | 0.00945 | 0.000162 | 109.084 | 107.9114 | 104.0772 | 20.9698 | 0.004632 | 0.004246 | 142.3415 | 0.26132 | 8.4016 | 0.003928 | 0.005476 | 0.005301 | 5.6e-05 | 0.001108 | 0.34971 | 1.3779 | 2.548 | 100.99 | 20.64 | 18.072 | 133.4847 | 116.2275 | 123.9946 | 0.004358 | 0.88896 | 0.02104 | 23.6681 | 0.096673 | 0.15546 | 0.058162 | 3.6631 | 3.6631 | 0.001744
2019-03-20 12:56 | 1553100950.181 | 367 | 0.19795 | 0.009479 | 0.00016 | 109.0148 | 107.7335 | 104.3635 | 20.9698 | 0.004654 | 0.004185 | 142.2011 | 0.26243 | 8.4003 | 0.004002 | 0.00506 | 0.004976 | 5.3e-05 | 0.001109 | 0.35314 | 1.3373 | 2.5481 | 100.99 | 20.64 | 18.077 | 134.0125 | 118.0315 | 122.4974 | 0.004359 | 0.89232 | 0.021109 | 23.6559 | 0.10575 | 0.15612 | 0.05812 | 3.6725 | 3.6725 | 0.001747
2019-03-20 12:56 | 1553100951.1799 | 368 | 0.19792 | 0.0095 | 0.000159 | 108.9459 | 107.6628 | 104.8771 | 20.9693 | 0.004702 | 0.004226 | 142.0475 | 0.26133 | 8.4026 | 0.004007 | 0.005675 | 0.005518 | 5.5e-05 | 0.001112 | 0.3298 | 1.329 | 2.5479 | 100.99 | 20.64 | 18.083 | 135.2052 | 118.1372 | 123.5067 | 0.00436 | 0.89005 | 0.021134 | 23.7443 | 0.10597 | 0.15546 | 0.058293 | 3.6878 | 3.6878 | 0.001733
2019-03-20 12:56 | 1553100952.1799 | 369 | 0.19763 | 0.009539 | 0.000157 | 108.9462 | 107.6572 | 104.6034 | 20.9687 | 0.004704 | 0.00415 | 141.9066 | 0.26158 | 8.404 | 0.003952 | 0.005252 | 0.005098 | 5.8e-05 | 0.001115 | 0.31885 | 1.3233 | 2.5478 | 100.99 | 20.65 | 18.094 | 135.2566 | 116.8005 | 121.6273 | 0.004366 | 0.89429 | 0.021156 | 23.657 | 0.10646 | 0.15562 | 0.059954 | 3.7966 | 3.7966 | 0.001685
2019-03-20 12:56 | 1553100953.18 | 370 | 0.1976 | 0.009606 | 0.000156 | 108.8175 | 107.4476 | 104.6112 | 20.9685 | 0.004678 | 0.004167 | 141.7659 | 0.26223 | 8.4024 | 0.003916 | 0.004855 | 0.004692 | 5.2e-05 | 0.001116 | 0.34683 | 1.368 | 2.5482 | 100.99 | 20.65 | 18.099 | 134.606 | 115.9233 | 122.0663 | 0.004367 | 0.8933 | 0.021145 | 23.6711 | 0.11332 | 0.156 | 0.060056 | 3.8012 | 3.8012 | 0.001687
2019-03-20 12:56 | 1553100954.1811 | 371 | 0.1976 | 0.009675 | 0.000155 | 108.7509 | 107.4737 | 104.659 | 20.968 | 0.004678 | 0.00419 | 141.6279 | 0.26239 | 8.4007 | 0.003938 | 0.00566 | 0.005503 | 5.5e-05 | 0.001117 | 0.31758 | 1.3755 | 2.5482 | 100.99 | 20.64 | 18.11 | 134.6168 | 116.4531 | 122.6203 | 0.004367 | 0.89205 | 0.021135 | 23.6931 | 0.10567 | 0.1561 | 0.059948 | 3.7926 | 3.7926 | 0.001692
2019-03-20 12:56 | 1553100955.1802 | 372 | 0.1976 | 0.00974 | 0.000155 | 108.6992 | 107.3298 | 104.7691 | 20.9674 | 0.004679 | 0.004297 | 141.461 | 0.26197 | 8.4046 | 0.003912 | 0.005616 | 0.005448 | 5.1e-05 | 0.001116 | 0.35026 | 1.3448 | 2.5481 | 100.99 | 20.65 | 18.115 | 134.6234 | 115.8344 | 125.2328 | 0.004371 | 0.8862 | 0.021077 | 23.7836 | 0.11339 | 0.15585 | 0.059903 | 3.7793 | 3.7793 | 0.001695
2019-03-20 12:56 | 1553100956.1802 | 373 | 0.19758 | 0.009784 | 0.000156 | 108.6293 | 107.544 | 105.275 | 20.9675 | 0.004712 | 0.004244 | 141.3017 | 0.26282 | 8.404 | 0.003966 | 0.006121 | 0.00598 | 5.4e-05 | 0.001113 | 0.33915 | 1.3342 | 2.5481 | 100.99 | 20.65 | 18.12 | 135.449 | 117.1444 | 123.9412 | 0.004372 | 0.88908 | 0.021162 | 23.8023 | 0.089814 | 0.15635 | 0.059955 | 3.7978 | 3.7978 | 0.001692
2019-03-20 12:56 | 1553100957.1792 | 374 | 0.19764 | 0.009808 | 0.000157 | 108.5866 | 107.5847 | 105.3569 | 20.9677 | 0.00473 | 0.004213 | 141.159 | 0.26191 | 8.4044 | 0.003991 | 0.005605 | 0.005484 | 5.5e-05 | 0.001108 | 0.34324 | 1.3136 | 2.5481 | 100.99 | 20.65 | 18.129 | 135.8888 | 117.7625 | 123.1683 | 0.004374 | 0.89081 | 0.021191 | 23.7884 | 0.08291 | 0.15581 | 0.059579 | 3.7792 | 3.7792 | 0.001695
2019-03-20 12:56 | 1553100958.1803 | 375 | 0.19764 | 0.009832 | 0.000158 | 108.5145 | 107.4076 | 105.2681 | 20.9667 | 0.004675 | 0.004205 | 140.9971 | 0.26204 | 8.4031 | 0.004085 | 0.005801 | 0.005677 | 5.3e-05 | 0.001103 | 0.36409 | 1.3442 | 2.548 | 100.99 | 20.65 | 18.132 | 134.5283 | 120.0445 | 122.99 | 0.004375 | 0.89121 | 0.021187 | 23.773 | 0.091708 | 0.15589 | 0.059536 | 3.7757 | 3.7757 | 0.001697
2019-03-20 12:56 | 1553100959.1794 | 376 | 0.19771 | 0.00985 | 0.000158 | 108.442 | 107.0432 | 105.2475 | 20.9673 | 0.004681 | 0.004256 | 140.8411 | 0.261 | 8.4031 | 0.004021 | 0.005845 | 0.005669 | 5.4e-05 | 0.0011 | 0.31061 | 1.3133 | 2.5481 | 100.99 | 20.65 | 18.139 | 134.6696 | 118.4872 | 124.2322 | 0.004377 | 0.88843 | 0.021152 | 23.8079 | 0.11613 | 0.15527 | 0.059092 | 3.7414 | 3.7414 | 0.001706
2019-03-20 12:56 | 1553100960.1793 | 377 | 0.19777 | 0.009864 | 0.000158 | 108.3906 | 106.9034 | 105.2214 | 20.967 | 0.004746 | 0.004251 | 140.6814 | 0.26279 | 8.3998 | 0.004067 | 0.005099 | 0.004946 | 5.1e-05 | 0.001101 | 0.32892 | 1.3145 | 2.5481 | 100.99 | 20.65 | 18.143 | 136.2737 | 119.6182 | 124.0991 | 0.004378 | 0.88873 | 0.021153 | 23.801 | 0.12358 | 0.15633 | 0.05873 | 3.7187 | 3.7187 | 0.001728
2019-03-20 12:56 | 1553100961.1798 | 378 | 0.19779 | 0.009858 | 0.000159 | 108.3098 | 107.1258 | 105.0092 | 20.9662 | 0.004745 | 0.004122 | 140.551 | 0.26157 | 8.4034 | 0.004027 | 0.005715 | 0.005627 | 5.4e-05 | 0.001104 | 0.30215 | 1.3537 | 2.5481 | 100.99 | 20.65 | 18.146 | 136.2667 | 118.6256 | 120.9437 | 0.004378 | 0.89584 | 0.021216 | 23.6824 | 0.098317 | 0.15561 | 0.058629 | 3.7234 | 3.7234 | 0.001718
2019-03-20 12:56 | 1553100962.1799 | 379 | 0.19785 | 0.00985 | 0.000158 | 108.2299 | 107.2375 | 105.4141 | 20.9659 | 0.004748 | 0.004258 | 140.3885 | 0.26216 | 8.4062 | 0.003964 | 0.005752 | 0.005622 | 5.4e-05 | 0.001099 | 0.29391 | 1.3211 | 2.548 | 100.99 | 20.65 | 18.152 | 136.3407 | 117.0892 | 124.284 | 0.00438 | 0.88831 | 0.021167 | 23.8283 | 0.082391 | 0.15596 | 0.05827 | 3.6922 | 3.6922 | 0.001736
2019-03-20 12:56 | 1553100963.1795 | 380 | 0.19798 | 0.009832 | 0.000158 | 108.1558 | 106.9991 | 105.2121 | 20.9661 | 0.004675 | 0.004215 | 140.2381 | 0.26182 | 8.4047 | 0.003969 | 0.005355 | 0.00518 | 5.4e-05 | 0.001107 | 0.30359 | 1.3043 | 2.548 | 100.99 | 20.64 | 18.155 | 134.519 | 117.2104 | 123.2333 | 0.004378 | 0.89067 | 0.021175 | 23.774 | 0.09618 | 0.15576 | 0.057521 | 3.6463 | 3.6463 | 0.001756
2019-03-20 12:56 | 1553100964.1796 | 381 | 0.19814 | 0.009803 | 0.000159 | 108.1374 | 106.9923 | 105.0231 | 20.965 | 0.004761 | 0.004209 | 140.0918 | 0.26138 | 8.4007 | 0.003859 | 0.005189 | 0.005028 | 5.5e-05 | 0.001104 | 0.31022 | 1.3141 | 2.5481 | 100.99 | 20.64 | 18.158 | 136.6441 | 114.5348 | 123.0899 | 0.004378 | 0.89099 | 0.021159 | 23.7483 | 0.095217 | 0.1555 | 0.056579 | 3.5841 | 3.5841 | 0.001783
2019-03-20 12:56 | 1553100965.18 | 382 | 0.19822 | 0.009763 | 0.000161 | 108.0345 | 106.7969 | 104.8144 | 20.9659 | 0.00473 | 0.004255 | 139.9484 | 0.26137 | 8.4031 | 0.003894 | 0.005786 | 0.005681 | 5.2e-05 | 0.001106 | 0.3115 | 1.3416 | 2.5481 | 100.99 | 20.64 | 18.166 | 135.8884 | 115.3859 | 124.2124 | 0.00438 | 0.88847 | 0.021109 | 23.7583 | 0.10306 | 0.15549 | 0.056188 | 3.5508 | 3.5508 | 0.0018
2019-03-20 12:56 | 1553100966.1801 | 383 | 0.19823 | 0.00972 | 0.000163 | 107.9873 | 107.0334 | 104.924 | 20.9657 | 0.004711 | 0.004311 | 139.7897 | 0.26214 | 8.4018 | 0.003821 | 0.00522 | 0.005171 | 5.4e-05 | 0.001105 | 0.30879 | 1.346 | 2.548 | 100.99 | 20.65 | 18.17 | 135.4065 | 113.5923 | 125.5767 | 0.004384 | 0.88543 | 0.021083 | 23.8115 | 0.07936 | 0.15595 | 0.056184 | 3.5462 | 3.5462 | 0.001807
2019-03-20 12:56 | 1553100967.1808 | 384 | 0.1982 | 0.009659 | 0.000166 | 107.921 | 106.8537 | 105.07 | 20.9657 | 0.004774 | 0.004309 | 139.6651 | 0.26191 | 8.4028 | 0.003848 | 0.005091 | 0.004992 | 5.3e-05 | 0.001103 | 0.33716 | 1.3162 | 2.5481 | 100.99 | 20.64 | 18.18 | 136.9748 | 114.2657 | 125.5352 | 0.004384 | 0.88552 | 0.021099 | 23.8268 | 0.088903 | 0.15581 | 0.056387 | 3.5615 | 3.5615 | 0.001798
2019-03-20 12:56 | 1553100968.1808 | 385 | 0.19819 | 0.009589 | 0.000167 | 107.803 | 106.7924 | 105.098 | 20.9647 | 0.004803 | 0.004265 | 139.5528 | 0.26129 | 8.4017 | 0.003869 | 0.005581 | 0.005433 | 5.3e-05 | 0.001109 | 0.31797 | 1.3319 | 2.548 | 100.99 | 20.64 | 18.186 | 137.6833 | 114.7653 | 124.4586 | 0.004385 | 0.88792 | 0.021131 | 23.7978 | 0.084246 | 0.15544 | 0.056529 | 3.5757 | 3.5757 | 0.001787
2019-03-20 12:56 | 1553100969.1799 | 386 | 0.19813 | 0.009537 | 0.000167 | 107.673 | 106.4938 | 104.5375 | 20.965 | 0.004822 | 0.004281 | 139.4292 | 0.26128 | 8.4034 | 0.00392 | 0.005229 | 0.005158 | 5.2e-05 | 0.001108 | 0.32139 | 1.357 | 2.5481 | 100.99 | 20.64 | 18.192 | 138.1468 | 116.0227 | 124.8327 | 0.004387 | 0.88709 | 0.021064 | 23.7454 | 0.098499 | 0.15544 | 0.056962 | 3.5916 | 3.5916 | 0.001779
2019-03-20 12:56 | 1553100970.1809 | 387 | 0.19779 | 0.009517 | 0.000166 | 107.6468 | 106.2164 | 104.2905 | 20.9651 | 0.00481 | 0.004381 | 139.2746 | 0.26124 | 8.4019 | 0.003809 | 0.004969 | 0.004704 | 5.3e-05 | 0.001108 | 0.30613 | 1.3 | 2.548 | 100.99 | 20.64 | 18.205 | 137.8681 | 113.3041 | 127.3094 | 0.00439 | 0.8816 | 0.020974 | 23.791 | 0.11965 | 0.15541 | 0.058997 | 3.7036 | 3.7036 | 0.001725
2019-03-20 12:56 | 1553100971.181 | 388 | 0.19779 | 0.009529 | 0.000166 | 107.5585 | 106.1862 | 104.106 | 20.9648 | 0.0048 | 0.004328 | 139.1254 | 0.26208 | 8.4023 | 0.003811 | 0.005837 | 0.005709 | 5.3e-05 | 0.001105 | 0.36226 | 1.2867 | 2.548 | 100.99 | 20.64 | 18.211 | 137.6162 | 113.3488 | 125.9918 | 0.004391 | 0.88451 | 0.02099 | 23.7308 | 0.11486 | 0.15591 | 0.05898 | 3.7054 | 3.7054 | 0.001729
2019-03-20 12:56 | 1553100972.1801 | 389 | 0.19771 | 0.009566 | 0.000164 | 107.4446 | 106.0987 | 103.8385 | 20.9639 | 0.004817 | 0.004326 | 139.0036 | 0.26259 | 8.4027 | 0.003865 | 0.005168 | 0.005053 | 5.5e-05 | 0.001109 | 0.3413 | 1.3118 | 2.5482 | 100.99 | 20.64 | 18.224 | 138.0351 | 114.6696 | 125.9426 | 0.004394 | 0.88462 | 0.020964 | 23.6988 | 0.11275 | 0.15621 | 0.059424 | 3.7286 | 3.7286 | 0.001722
2019-03-20 12:56 | 1553100973.1791 | 390 | 0.19763 | 0.009611 | 0.000163 | 107.303 | 105.8561 | 104.326 | 20.964 | 0.004816 | 0.004256 | 138.873 | 0.26244 | 8.405 | 0.003767 | 0.005249 | 0.005072 | 5.4e-05 | 0.001101 | 0.31421 | 1.3342 | 2.548 | 100.99 | 20.64 | 18.231 | 138.0056 | 112.2777 | 124.2417 | 0.004396 | 0.88841 | 0.021059 | 23.7037 | 0.12143 | 0.15613 | 0.059881 | 3.7741 | 3.7741 | 0.0017
2019-03-20 12:56 | 1553100974.1792 | 391 | 0.19752 | 0.00964 | 0.000161 | 107.1875 | 105.6154 | 104.5187 | 20.9643 | 0.004802 | 0.004252 | 138.7241 | 0.26193 | 8.4025 | 0.003762 | 0.005507 | 0.005447 | 4.9e-05 | 0.001093 | 0.34058 | 1.3033 | 2.548 | 100.99 | 20.64 | 18.237 | 137.6532 | 112.1761 | 124.1434 | 0.004397 | 0.88863 | 0.021081 | 23.7227 | 0.13216 | 0.15582 | 0.060478 | 3.8157 | 3.8157 | 0.001678
2019-03-20 12:56 | 1553100975.1793 | 392 | 0.19746 | 0.009678 | 0.000159 | 107.2124 | 105.7448 | 104.4706 | 20.9639 | 0.004794 | 0.004352 | 138.6098 | 0.26137 | 8.3987 | 0.003809 | 0.005435 | 0.005244 | 5.3e-05 | 0.001087 | 0.31061 | 1.3453 | 2.5481 | 100.99 | 20.64 | 18.248 | 137.4544 | 113.3124 | 126.5907 | 0.0044 | 0.88319 | 0.021011 | 23.7902 | 0.12329 | 0.15549 | 0.060805 | 3.8237 | 3.8237 | 0.001671
2019-03-20 12:56 | 1553100976.1796 | 393 | 0.19739 | 0.00972 | 0.000157 | 107.1015 | 105.6614 | 104.1938 | 20.965 | 0.004802 | 0.004359 | 138.4716 | 0.26167 | 8.4054 | 0.00382 | 0.005428 | 0.005376 | 5.2e-05 | 0.001082 | 0.30071 | 1.3472 | 2.5481 | 100.99 | 20.64 | 18.253 | 137.6749 | 113.5704 | 126.7596 | 0.004401 | 0.88281 | 0.020979 | 23.7637 | 0.12108 | 0.15567 | 0.061182 | 3.8415 | 3.8415 | 0.001666
2019-03-20 12:56 | 1553100977.1793 | 394 | 0.19724 | 0.009766 | 0.000155 | 106.9734 | 105.4211 | 104.3448 | 20.9654 | 0.004883 | 0.004321 | 138.3178 | 0.26145 | 8.4027 | 0.003823 | 0.00511 | 0.005006 | 5e-05 | 0.001083 | 0.30198 | 1.3251 | 2.5481 | 100.99 | 20.64 | 18.258 | 139.6542 | 113.6434 | 125.8303 | 0.004403 | 0.88487 | 0.021018 | 23.7532 | 0.13075 | 0.15553 | 0.061997 | 3.8999 | 3.8999 | 0.001639
2019-03-20 12:56 | 1553100978.1794 | 395 | 0.19713 | 0.009815 | 0.000154 | 106.8237 | 105.3004 | 104.2004 | 20.9648 | 0.004845 | 0.004358 | 138.1547 | 0.26206 | 8.399 | 0.003918 | 0.005113 | 0.005029 | 5.2e-05 | 0.00108 | 0.31996 | 1.3279 | 2.548 | 100.99 | 20.64 | 18.266 | 138.7316 | 115.9694 | 126.726 | 0.004404 | 0.88289 | 0.02098 | 23.7634 | 0.12847 | 0.1559 | 0.062616 | 3.9316 | 3.9316 | 0.00163
2019-03-20 12:56 | 1553100979.1798 | 396 | 0.1971 | 0.009879 | 0.000154 | 106.6325 | 105.226 | 104.8799 | 20.9655 | 0.004816 | 0.004333 | 137.9958 | 0.26214 | 8.3999 | 0.00404 | 0.004715 | 0.00452 | 5.1e-05 | 0.00108 | 0.31614 | 1.32 | 2.5481 | 100.99 | 20.64 | 18.269 | 138.011 | 118.9356 | 126.1096 | 0.004405 | 0.88425 | 0.021065 | 23.8224 | 0.11877 | 0.15594 | 0.062682 | 3.9516 | 3.9516 | 0.001622
2019-03-20 12:56 | 1553100980.1799 | 397 | 0.19715 | 0.009946 | 0.000154 | 106.6183 | 105.0574 | 105.0102 | 20.9657 | 0.004887 | 0.004339 | 137.8704 | 0.26225 | 8.4026 | 0.004046 | 0.00574 | 0.005629 | 5.2e-05 | 0.00108 | 0.30287 | 1.3194 | 2.548 | 100.99 | 20.64 | 18.274 | 139.7692 | 119.0864 | 126.2562 | 0.004406 | 0.88393 | 0.021074 | 23.8415 | 0.13192 | 0.15601 | 0.062345 | 3.9321 | 3.9321 | 0.001631
2019-03-20 12:56 | 1553100981.1805 | 398 | 0.19716 | 0.010015 | 0.000154 | 106.4602 | 105.0066 | 105.4411 | 20.9648 | 0.004813 | 0.004357 | 137.7249 | 0.26201 | 8.4055 | 0.004107 | 0.006082 | 0.005965 | 5.6e-05 | 0.001082 | 0.33754 | 1.2952 | 2.548 | 100.99 | 20.63 | 18.276 | 137.9464 | 120.5916 | 126.7147 | 0.004404 | 0.88291 | 0.021105 | 23.9041 | 0.12297 | 0.15587 | 0.062163 | 3.9265 | 3.9265 | 0.001632
2019-03-20 12:56 | 1553100982.1806 | 399 | 0.19728 | 0.010062 | 0.000156 | 106.2748 | 104.7737 | 105.6935 | 20.9642 | 0.004913 | 0.004301 | 137.5839 | 0.26171 | 8.4048 | 0.004102 | 0.005343 | 0.005258 | 5.2e-05 | 0.001083 | 0.30143 | 1.2848 | 2.5479 | 100.99 | 20.63 | 18.282 | 140.4006 | 120.4538 | 125.3245 | 0.004406 | 0.88599 | 0.021167 | 23.8911 | 0.12724 | 0.15569 | 0.061418 | 3.8909 | 3.8909 | 0.001645
2019-03-20 12:56 | 1553100983.1806 | 400 | 0.19745 | 0.010093 | 0.000158 | 106.1393 | 104.6216 | 105.7388 | 20.9646 | 0.004843 | 0.004329 | 137.4356 | 0.26224 | 8.4021 | 0.004009 | 0.005618 | 0.005477 | 5.4e-05 | 0.001088 | 0.34468 | 1.3412 | 2.548 | 100.99 | 20.63 | 18.283 | 138.6707 | 118.1929 | 126.0211 | 0.004406 | 0.88445 | 0.021153 | 23.9171 | 0.12882 | 0.15601 | 0.060366 | 3.8219 | 3.8219 | 0.001678
2019-03-20 12:56 | 1553100984.1807 | 401 | 0.19772 | 0.010099 | 0.000162 | 106.074 | 104.4647 | 105.6207 | 20.9643 | 0.004869 | 0.004358 | 137.3025 | 0.26104 | 8.4021 | 0.004037 | 0.005107 | 0.004994 | 5.3e-05 | 0.001091 | 0.35706 | 1.2849 | 2.5481 | 100.99 | 20.64 | 18.285 | 139.3252 | 118.8622 | 126.7256 | 0.004409 | 0.88289 | 0.021123 | 23.9248 | 0.13675 | 0.15529 | 0.058738 | 3.7137 | 3.7137 | 0.001719
2019-03-20 12:56 | 1553100985.18 | 402 | 0.19778 | 0.010072 | 0.000166 | 105.9973 | 104.061 | 105.7693 | 20.9647 | 0.004861 | 0.004304 | 137.1541 | 0.26158 | 8.4013 | 0.004067 | 0.005867 | 0.005709 | 5.2e-05 | 0.001096 | 0.36105 | 1.2843 | 2.548 | 100.99 | 20.64 | 18.29 | 139.1197 | 119.6086 | 125.4013 | 0.00441 | 0.88582 | 0.021173 | 23.9019 | 0.16491 | 0.15562 | 0.058412 | 3.7017 | 3.7017 | 0.001728
2019-03-20 12:56 | 1553100986.1808 | 403 | 0.19795 | 0.010036 | 0.000169 | 105.8702 | 103.9762 | 105.9791 | 20.9643 | 0.00489 | 0.004305 | 137.0066 | 0.26158 | 8.4019 | 0.004101 | 0.005451 | 0.005313 | 5.5e-05 | 0.001097 | 0.33323 | 1.2972 | 2.5482 | 100.99 | 20.64 | 18.293 | 139.8457 | 120.4278 | 125.4304 | 0.004411 | 0.88576 | 0.021193 | 23.9265 | 0.16147 | 0.15562 | 0.057447 | 3.6441 | 3.6441 | 0.001755
2019-03-20 12:56 | 1553100987.1789 | 404 | 0.19814 | 0.009976 | 0.000171 | 105.7577 | 103.7676 | 105.8261 | 20.9641 | 0.004804 | 0.004281 | 136.8719 | 0.26149 | 8.4022 | 0.004091 | 0.006122 | 0.005917 | 5.3e-05 | 0.001097 | 0.34987 | 1.3213 | 2.5481 | 100.99 | 20.64 | 18.294 | 137.718 | 120.1982 | 124.8311 | 0.004411 | 0.88709 | 0.021194 | 23.8913 | 0.16992 | 0.15556 | 0.05641 | 3.5785 | 3.5785 | 0.001787
2019-03-20 12:56 | 1553100988.1803 | 405 | 0.19834 | 0.009904 | 0.000173 | 105.575 | 103.7459 | 105.8219 | 20.9645 | 0.004871 | 0.004334 | 136.7423 | 0.26092 | 8.4004 | 0.004025 | 0.00611 | 0.005956 | 5.1e-05 | 0.001101 | 0.35441 | 1.3612 | 2.5481 | 100.99 | 20.64 | 18.296 | 139.3752 | 118.5819 | 126.1508 | 0.004412 | 0.88416 | 0.021158 | 23.9304 | 0.15632 | 0.15522 | 0.055296 | 3.5021 | 3.5021 | 0.001822
2019-03-20 12:56 | 1553100989.179 | 406 | 0.19842 | 0.009811 | 0.000174 | 105.5116 | 103.5452 | 105.535 | 20.9639 | 0.004782 | 0.004304 | 136.6346 | 0.26207 | 8.4036 | 0.004022 | 0.005635 | 0.005475 | 5.4e-05 | 0.001098 | 0.31133 | 1.3294 | 2.5481 | 100.99 | 20.64 | 18.297 | 137.1685 | 118.4924 | 125.4062 | 0.004412 | 0.88581 | 0.021149 | 23.8756 | 0.16827 | 0.15591 | 0.054948 | 3.4786 | 3.4786 | 0.001842
2019-03-20 12:56 | 1553100990.1789 | 407 | 0.19839 | 0.009724 | 0.000174 | 105.3528 | 103.3584 | 105.5885 | 20.9632 | 0.004842 | 0.004295 | 136.5057 | 0.26167 | 8.4012 | 0.004053 | 0.006075 | 0.005929 | 5.3e-05 | 0.001101 | 0.36658 | 1.3592 | 2.5481 | 100.99 | 20.64 | 18.3 | 138.6579 | 119.264 | 125.1818 | 0.004413 | 0.88631 | 0.021161 | 23.8749 | 0.17095 | 0.15567 | 0.055211 | 3.4971 | 3.4971 | 0.001829
2019-03-20 12:56 | 1553100991.1791 | 408 | 0.19876 | 0.009628 | 0.000175 | 105.2322 | 103.3404 | 105.4954 | 20.9632 | 0.004789 | 0.004365 | 136.3762 | 0.26104 | 8.4018 | 0.004051 | 0.005386 | 0.005221 | 5.4e-05 | 0.001098 | 0.33069 | 1.3553 | 2.5481 | 100.99 | 20.64 | 18.301 | 137.3509 | 119.2131 | 126.898 | 0.004413 | 0.88251 | 0.021106 | 23.9157 | 0.16226 | 0.15529 | 0.053084 | 3.354 | 3.354 | 0.001903
2019-03-20 12:57 | 1553100992.1792 | 409 | 0.19858 | 0.009544 | 0.000175 | 105.0339 | 103.2642 | 105.6057 | 20.9625 | 0.004863 | 0.004341 | 136.2433 | 0.26198 | 8.4021 | 0.004092 | 0.00549 | 0.005356 | 5.2e-05 | 0.001101 | 0.35115 | 1.358 | 2.5481 | 100.99 | 20.64 | 18.303 | 139.1704 | 120.215 | 126.3176 | 0.004413 | 0.88379 | 0.021132 | 23.9109 | 0.15199 | 0.15585 | 0.054262 | 3.4325 | 3.4325 | 0.001866
2019-03-20 12:57 | 1553100993.1802 | 410 | 0.19851 | 0.009468 | 0.000173 | 104.9213 | 102.7992 | 105.6787 | 20.9636 | 0.004859 | 0.004332 | 136.1072 | 0.26187 | 8.4 | 0.004016 | 0.005944 | 0.005809 | 5.1e-05 | 0.001103 | 0.33987 | 1.3438 | 2.5481 | 100.99 | 20.64 | 18.303 | 139.0647 | 118.3582 | 126.0942 | 0.004413 | 0.88428 | 0.021145 | 23.9125 | 0.18277 | 0.15579 | 0.054815 | 3.4696 | 3.4696 | 0.001845
2019-03-20 12:57 | 1553100994.1793 | 411 | 0.19851 | 0.00939 | 0.000171 | 104.8518 | 102.5163 | 106.1232 | 20.9628 | 0.004824 | 0.004331 | 135.9855 | 0.26149 | 8.4033 | 0.004051 | 0.005371 | 0.005214 | 5.5e-05 | 0.001103 | 0.32748 | 1.3675 | 2.5482 | 100.99 | 20.64 | 18.304 | 138.2071 | 119.2047 | 126.071 | 0.004414 | 0.88434 | 0.02119 | 23.962 | 0.20149 | 0.15556 | 0.054891 | 3.4819 | 3.4819 | 0.001836
2019-03-20 12:57 | 1553100995.1807 | 412 | 0.19851 | 0.009323 | 0.000169 | 104.7635 | 102.1336 | 106.2707 | 20.9638 | 0.004828 | 0.004297 | 135.8498 | 0.262 | 8.4001 | 0.00406 | 0.00599 | 0.005859 | 5.1e-05 | 0.001101 | 0.36785 | 1.3237 | 2.5482 | 100.99 | 20.64 | 18.305 | 138.3003 | 119.4374 | 125.2324 | 0.004414 | 0.8862 | 0.021227 | 23.9534 | 0.22741 | 0.15587 | 0.054965 | 3.4928 | 3.4928 | 0.001834
2019-03-20 12:57 | 1553100996.1793 | 413 | 0.19855 | 0.009271 | 0.00017 | 104.4948 | 102.1728 | 106.3601 | 20.962 | 0.004841 | 0.004342 | 135.7482 | 0.26171 | 8.4037 | 0.004063 | 0.00509 | 0.004999 | 5.8e-05 | 0.001102 | 0.32604 | 1.3031 | 2.5482 | 100.99 | 20.64 | 18.306 | 138.6328 | 119.5053 | 126.3267 | 0.004414 | 0.88377 | 0.021207 | 23.9964 | 0.201 | 0.15569 | 0.054779 | 3.4776 | 3.4776 | 0.00184
2019-03-20 12:57 | 1553100997.1804 | 414 | 0.1985 | 0.009235 | 0.000173 | 104.421 | 102.1885 | 106.6758 | 20.9621 | 0.004765 | 0.004267 | 135.638 | 0.26203 | 8.4035 | 0.004072 | 0.00572 | 0.005571 | 5.2e-05 | 0.0011 | 0.32803 | 1.3182 | 2.548 | 100.99 | 20.64 | 18.308 | 136.7532 | 119.7226 | 124.4967 | 0.004415 | 0.88784 | 0.021288 | 23.9769 | 0.19331 | 0.15588 | 0.055133 | 3.5131 | 3.5131 | 0.001824
2019-03-20 12:57 | 1553100998.1798 | 415 | 0.19878 | 0.009221 | 0.000176 | 104.236 | 101.9755 | 106.5655 | 20.9615 | 0.004891 | 0.004275 | 135.5272 | 0.26212 | 8.4054 | 0.004013 | 0.005678 | 0.005536 | 5.2e-05 | 0.001101 | 0.33411 | 1.3044 | 2.5479 | 100.99 | 20.64 | 18.31 | 139.871 | 118.2848 | 124.6941 | 0.004415 | 0.8874 | 0.021271 | 23.9704 | 0.19611 | 0.15593 | 0.053475 | 3.4051 | 3.4051 | 0.001882
2019-03-20 12:57 | 1553100999.1795 | 416 | 0.1991 | 0.009199 | 0.000178 | 104.0818 | 101.5894 | 106.7147 | 20.9611 | 0.004805 | 0.004248 | 135.3818 | 0.26182 | 8.4012 | 0.003944 | 0.005255 | 0.005152 | 5e-05 | 0.001101 | 0.31924 | 1.297 | 2.548 | 100.99 | 20.64 | 18.312 | 137.7279 | 116.5974 | 124.0352 | 0.004416 | 0.88887 | 0.021304 | 23.9674 | 0.2168 | 0.15576 | 0.051612 | 3.2916 | 3.2916 | 0.001945
2019-03-20 12:57 | 1553101000.1799 | 417 | 0.19932 | 0.009156 | 0.000179 | 103.9026 | 101.0803 | 106.3767 | 20.961 | 0.004829 | 0.00424 | 135.2633 | 0.26184 | 8.4057 | 0.00407 | 0.005356 | 0.005212 | 5.4e-05 | 0.001103 | 0.3418 | 1.3241 | 2.548 | 100.99 | 20.64 | 18.314 | 138.3271 | 119.6653 | 123.8299 | 0.004416 | 0.88933 | 0.021276 | 23.9232 | 0.24632 | 0.15577 | 0.05035 | 3.207 | 3.207 | 0.001996
2019-03-20 12:57 | 1553101001.1797 | 418 | 0.19941 | 0.009095 | 0.00018 | 103.7082 | 100.7335 | 106.2052 | 20.9606 | 0.004802 | 0.004274 | 135.1363 | 0.26133 | 8.4014 | 0.003998 | 0.006101 | 0.005975 | 5.4e-05 | 0.001107 | 0.30718 | 1.3359 | 2.5481 | 100.99 | 20.64 | 18.316 | 137.6713 | 117.9123 | 124.6798 | 0.004417 | 0.88743 | 0.021236 | 23.9295 | 0.26031 | 0.15546 | 0.049879 | 3.1711 | 3.1711 | 0.002015
2019-03-20 12:57 | 1553101002.1796 | 419 | 0.19953 | 0.009 | 0.00018 | 103.4922 | 100.6726 | 106.2676 | 20.9601 | 0.004744 | 0.004271 | 135.0179 | 0.2618 | 8.4022 | 0.003971 | 0.005329 | 0.005225 | 5.3e-05 | 0.001111 | 0.35131 | 1.3422 | 2.548 | 100.99 | 20.64 | 18.318 | 136.234 | 117.2473 | 124.5876 | 0.004417 | 0.88763 | 0.021244 | 23.9337 | 0.24707 | 0.15575 | 0.049279 | 3.1343 | 3.1343 | 0.002042
2019-03-20 12:57 | 1553101003.1797 | 420 | 0.19954 | 0.008893 | 0.000179 | 103.313 | 100.7114 | 106.2378 | 20.961 | 0.004807 | 0.004238 | 134.9011 | 0.26164 | 8.4031 | 0.003975 | 0.006058 | 0.005898 | 5.3e-05 | 0.001107 | 0.31813 | 1.3262 | 2.548 | 100.99 | 20.64 | 18.32 | 137.7734 | 117.3479 | 123.7736 | 0.004417 | 0.88945 | 0.021263 | 23.9059 | 0.22812 | 0.15565 | 0.049323 | 3.1399 | 3.1399 | 0.002037
2019-03-20 12:57 | 1553101004.1801 | 421 | 0.19951 | 0.008785 | 0.000177 | 103.0607 | 100.4061 | 106.0946 | 20.9613 | 0.004852 | 0.004212 | 134.7919 | 0.26207 | 8.4007 | 0.00401 | 0.004896 | 0.004744 | 5.1e-05 | 0.001108 | 0.32189 | 1.3264 | 2.548 | 100.99 | 20.64 | 18.321 | 138.8974 | 118.1967 | 123.1546 | 0.004418 | 0.89084 | 0.021265 | 23.8711 | 0.23341 | 0.15591 | 0.049625 | 3.1595 | 3.1595 | 0.002028
2019-03-20 12:57 | 1553101005.1792 | 422 | 0.19954 | 0.008682 | 0.000176 | 102.8673 | 100.3766 | 106.1653 | 20.9611 | 0.004785 | 0.00423 | 134.6843 | 0.26189 | 8.4032 | 0.004096 | 0.005821 | 0.005669 | 5.4e-05 | 0.001112 | 0.33699 | 1.3205 | 2.5481 | 100.99 | 20.64 | 18.324 | 137.2481 | 120.3062 | 123.5941 | 0.004418 | 0.88986 | 0.021261 | 23.8923 | 0.21924 | 0.1558 | 0.04959 | 3.1566 | 3.1566 | 0.002029
2019-03-20 12:57 | 1553101006.1799 | 423 | 0.19966 | 0.008591 | 0.000176 | 102.6358 | 100.0884 | 106.4779 | 20.9605 | 0.004732 | 0.004266 | 134.5896 | 0.26196 | 8.3997 | 0.003959 | 0.00561 | 0.005476 | 5.3e-05 | 0.001113 | 0.31365 | 1.3134 | 2.5481 | 100.99 | 20.64 | 18.326 | 135.9191 | 116.9626 | 124.4728 | 0.004419 | 0.88789 | 0.021268 | 23.9539 | 0.2248 | 0.15584 | 0.04901 | 3.121 | 3.121 | 0.002052
2019-03-20 12:57 | 1553101007.1792 | 424 | 0.19975 | 0.008524 | 0.000177 | 102.4168 | 99.8783 | 106.1836 | 20.9608 | 0.004801 | 0.004188 | 134.4691 | 0.26153 | 8.4023 | 0.003995 | 0.005677 | 0.005576 | 5.6e-05 | 0.001114 | 0.34396 | 1.3264 | 2.5481 | 100.99 | 20.64 | 18.328 | 137.6322 | 117.8348 | 122.5615 | 0.004419 | 0.89218 | 0.02129 | 23.8632 | 0.22449 | 0.15559 | 0.048557 | 3.0953 | 3.0953 | 0.002066
2019-03-20 12:57 | 1553101008.1793 | 425 | 0.19985 | 0.008454 | 0.000178 | 102.2115 | 99.5102 | 106.5697 | 20.96 | 0.00476 | 0.004208 | 134.3483 | 0.2617 | 8.4022 | 0.003967 | 0.005025 | 0.004874 | 5.3e-05 | 0.001109 | 0.33892 | 1.3347 | 2.5481 | 100.99 | 20.64 | 18.327 | 136.6223 | 117.1695 | 123.0477 | 0.004419 | 0.89108 | 0.021316 | 23.9213 | 0.23957 | 0.15568 | 0.048005 | 3.0638 | 3.0638 | 0.002088
2019-03-20 12:57 | 1553101009.179 | 426 | 0.20008 | 0.008391 | 0.000179 | 101.8897 | 98.8416 | 106.6427 | 20.9606 | 0.00475 | 0.004178 | 134.2473 | 0.26161 | 8.404 | 0.004051 | 0.005104 | 0.00502 | 5.1e-05 | 0.001107 | 0.31188 | 1.3284 | 2.5481 | 100.99 | 20.64 | 18.327 | 136.3682 | 119.2113 | 122.3103 | 0.004419 | 0.89275 | 0.021343 | 23.9072 | 0.27167 | 0.15563 | 0.046715 | 2.9854 | 2.9854 | 0.002143
2019-03-20 12:57 | 1553101010.1795 | 427 | 0.20028 | 0.008314 | 0.000179 | 101.6584 | 98.5908 | 107.1267 | 20.9603 | 0.00477 | 0.004155 | 134.104 | 0.26231 | 8.4035 | 0.003997 | 0.00486 | 0.004716 | 5.4e-05 | 0.001106 | 0.35081 | 1.3143 | 2.5482 | 100.99 | 20.64 | 18.326 | 136.8701 | 117.8851 | 121.7637 | 0.004419 | 0.89398 | 0.021406 | 23.9448 | 0.27407 | 0.15604 | 0.045649 | 2.926 | 2.926 | 0.002192
2019-03-20 12:57 | 1553101011.1791 | 428 | 0.2004 | 0.008227 | 0.000179 | 101.3366 | 98.1755 | 107.3548 | 20.9602 | 0.004771 | 0.004208 | 134.0058 | 0.26156 | 8.4047 | 0.003974 | 0.005314 | 0.005169 | 5.2e-05 | 0.001103 | 0.32084 | 1.343 | 2.5482 | 100.99 | 20.64 | 18.328 | 136.8997 | 117.3338 | 123.0384 | 0.004419 | 0.89111 | 0.021394 | 24.0089 | 0.28346 | 0.1556 | 0.044982 | 2.8817 | 2.8817 | 0.002219
2019-03-20 12:57 | 1553101012.1792 | 429 | 0.20053 | 0.008128 | 0.000178 | 101.2008 | 97.8881 | 107.3171 | 20.9609 | 0.004707 | 0.004155 | 133.8982 | 0.2609 | 8.4037 | 0.003957 | 0.005529 | 0.005437 | 5.5e-05 | 0.001101 | 0.34003 | 1.2832 | 2.548 | 100.99 | 20.64 | 18.329 | 135.3273 | 116.9195 | 121.7594 | 0.00442 | 0.89399 | 0.021425 | 23.9659 | 0.29769 | 0.15521 | 0.044335 | 2.8444 | 2.8444 | 0.002243
2019-03-20 12:57 | 1553101013.1806 | 430 | 0.20054 | 0.008022 | 0.000177 | 100.9369 | 97.3944 | 107.3351 | 20.9611 | 0.004734 | 0.004112 | 133.8195 | 0.26187 | 8.4001 | 0.003978 | 0.005381 | 0.005229 | 5.2e-05 | 0.001101 | 0.34108 | 1.3597 | 2.5482 | 100.99 | 20.64 | 18.329 | 135.9757 | 117.4313 | 120.6912 | 0.00442 | 0.89642 | 0.021456 | 23.9355 | 0.31956 | 0.15579 | 0.044442 | 2.8555 | 2.8555 | 0.002242
2019-03-20 12:57 | 1553101014.1793 | 431 | 0.20058 | 0.007921 | 0.000175 | 100.649 | 97.5271 | 107.453 | 20.9603 | 0.004644 | 0.004155 | 133.688 | 0.26105 | 8.4034 | 0.003981 | 0.005634 | 0.005456 | 5.6e-05 | 0.001102 | 0.34058 | 1.3381 | 2.548 | 100.99 | 20.64 | 18.33 | 133.7719 | 117.4884 | 121.7493 | 0.00442 | 0.89401 | 0.021439 | 23.9808 | 0.28184 | 0.1553 | 0.04433 | 2.8462 | 2.8462 | 0.002243
2019-03-20 12:57 | 1553101015.1792 | 432 | 0.20053 | 0.007826 | 0.000174 | 100.3195 | 97.1059 | 107.5346 | 20.9605 | 0.004673 | 0.004138 | 133.5717 | 0.26216 | 8.4033 | 0.003951 | 0.005781 | 0.005686 | 5.2e-05 | 0.001102 | 0.32532 | 1.2999 | 2.5481 | 100.99 | 20.64 | 18.329 | 134.4701 | 116.759 | 121.3462 | 0.00442 | 0.89493 | 0.021458 | 23.9777 | 0.29122 | 0.15596 | 0.044715 | 2.8734 | 2.8734 | 0.002231
2019-03-20 12:57 | 1553101016.1794 | 433 | 0.20051 | 0.007748 | 0.000176 | 100.095 | 96.7209 | 107.6213 | 20.9603 | 0.004643 | 0.004164 | 133.4447 | 0.26131 | 8.4021 | 0.003985 | 0.005424 | 0.005284 | 5.2e-05 | 0.001109 | 0.34346 | 1.3296 | 2.5481 | 100.99 | 20.64 | 18.327 | 133.731 | 117.589 | 121.9725 | 0.004419 | 0.89351 | 0.02145 | 24.0064 | 0.30671 | 0.15545 | 0.04492 | 2.8853 | 2.8853 | 0.002214
2019-03-20 12:57 | 1553101017.1804 | 434 | 0.20048 | 0.007691 | 0.00018 | 99.8156 | 96.4503 | 108.1078 | 20.9607 | 0.004718 | 0.004145 | 133.3305 | 0.26186 | 8.4024 | 0.003992 | 0.005637 | 0.005531 | 5.6e-05 | 0.001111 | 0.34418 | 1.353 | 2.5481 | 100.99 | 20.64 | 18.325 | 135.5798 | 117.7739 | 121.5091 | 0.004419 | 0.89456 | 0.021511 | 24.0464 | 0.30677 | 0.15578 | 0.045152 | 2.9083 | 2.9083 | 0.002201
2019-03-20 12:57 | 1553101018.1799 | 435 | 0.20051 | 0.007646 | 0.000184 | 99.606 | 96.3778 | 108.3477 | 20.9598 | 0.004647 | 0.004088 | 133.2346 | 0.26122 | 8.4 | 0.004008 | 0.004771 | 0.004584 | 5.3e-05 | 0.001107 | 0.30469 | 1.3334 | 2.548 | 100.99 | 20.64 | 18.322 | 133.8449 | 118.1611 | 120.125 | 0.004418 | 0.89771 | 0.021573 | 24.0309 | 0.29469 | 0.1554 | 0.045049 | 2.9098 | 2.9098 | 0.002195
2019-03-20 12:57 | 1553101019.18 | 436 | 0.20078 | 0.00761 | 0.000185 | 99.2147 | 95.3313 | 108.7759 | 20.9606 | 0.004649 | 0.004136 | 133.1206 | 0.26195 | 8.4054 | 0.003897 | 0.0062 | 0.00611 | 5.3e-05 | 0.001106 | 0.31526 | 1.332 | 2.548 | 100.99 | 20.64 | 18.317 | 133.8961 | 115.4501 | 121.2811 | 0.004417 | 0.89507 | 0.021584 | 24.1137 | 0.35714 | 0.15584 | 0.043491 | 2.8106 | 2.8106 | 0.002279
2019-03-20 12:57 | 1553101020.1786 | 437 | 0.20109 | 0.007564 | 0.000188 | 98.8745 | 94.4455 | 109.2059 | 20.9603 | 0.004599 | 0.004117 | 133.0449 | 0.26209 | 8.4025 | 0.003986 | 0.004989 | 0.004882 | 5.4e-05 | 0.001101 | 0.32748 | 1.305 | 2.5481 | 100.99 | 20.64 | 18.315 | 132.6449 | 117.63 | 120.8222 | 0.004416 | 0.89612 | 0.021639 | 24.1472 | 0.40991 | 0.15592 | 0.041699 | 2.7016 | 2.7016 | 0.002372
2019-03-20 12:57 | 1553101021.1786 | 438 | 0.20126 | 0.007511 | 0.000192 | 98.5177 | 94.1211 | 109.5918 | 20.9603 | 0.004684 | 0.004038 | 132.9291 | 0.26154 | 8.4042 | 0.003931 | 0.005797 | 0.005613 | 5.2e-05 | 0.001101 | 0.33019 | 1.3114 | 2.5481 | 100.99 | 20.64 | 18.309 | 134.7428 | 116.2911 | 118.8836 | 0.004415 | 0.90055 | 0.02173 | 24.1303 | 0.40836 | 0.15559 | 0.040741 | 2.6506 | 2.6506 | 0.002413
2019-03-20 12:57 | 1553101022.1801 | 439 | 0.20133 | 0.007438 | 0.000196 | 98.2616 | 93.8677 | 110.0013 | 20.9608 | 0.004645 | 0.004092 | 132.8418 | 0.26128 | 8.4038 | 0.003893 | 0.005561 | 0.005437 | 5.1e-05 | 0.001098 | 0.32045 | 1.3525 | 2.5483 | 100.99 | 20.64 | 18.306 | 133.7923 | 115.3475 | 120.2005 | 0.004414 | 0.89753 | 0.021735 | 24.2159 | 0.40919 | 0.15543 | 0.040382 | 2.6275 | 2.6275 | 0.002431
2019-03-20 12:57 | 1553101023.1797 | 440 | 0.20138 | 0.007352 | 0.000199 | 97.966 | 93.7234 | 110.6121 | 20.9599 | 0.004587 | 0.004084 | 132.7256 | 0.26132 | 8.4036 | 0.003975 | 0.005294 | 0.005082 | 5.1e-05 | 0.001102 | 0.31166 | 1.312 | 2.5481 | 100.99 | 20.64 | 18.301 | 132.3625 | 117.3641 | 120.0194 | 0.004413 | 0.89795 | 0.0218 | 24.2774 | 0.396 | 0.15546 | 0.040181 | 2.6221 | 2.6221 | 0.002437
2019-03-20 12:57 | 1553101024.1802 | 441 | 0.20135 | 0.007258 | 0.000201 | 97.781 | 93.5937 | 111.047 | 20.9607 | 0.004605 | 0.004056 | 132.6531 | 0.26166 | 8.4024 | 0.003955 | 0.005893 | 0.005803 | 5.3e-05 | 0.001102 | 0.3235 | 1.3424 | 2.5481 | 100.99 | 20.64 | 18.295 | 132.81 | 116.8685 | 119.3233 | 0.004411 | 0.89954 | 0.021862 | 24.3036 | 0.39148 | 0.15566 | 0.040455 | 2.6474 | 2.6474 | 0.002417
2019-03-20 12:57 | 1553101025.1788 | 442 | 0.20138 | 0.007166 | 0.000202 | 97.5653 | 92.9699 | 111.3674 | 20.9604 | 0.00452 | 0.004061 | 132.5579 | 0.26128 | 8.4032 | 0.003859 | 0.005471 | 0.005274 | 5.1e-05 | 0.001101 | 0.3553 | 1.3394 | 2.5483 | 100.99 | 20.64 | 18.292 | 130.7053 | 114.5215 | 119.4482 | 0.004411 | 0.89925 | 0.02189 | 24.3425 | 0.43153 | 0.15543 | 0.040386 | 2.6462 | 2.6462 | 0.002414
2019-03-20 12:57 | 1553101026.1799 | 443 | 0.20163 | 0.007072 | 0.000204 | 97.6665 | 92.6934 | 111.6024 | 20.9605 | 0.004568 | 0.004 | 132.4775 | 0.26135 | 8.4015 | 0.003865 | 0.005313 | 0.005235 | 5e-05 | 0.001103 | 0.33284 | 1.3555 | 2.5482 | 100.99 | 20.64 | 18.289 | 131.8974 | 114.679 | 117.9701 | 0.00441 | 0.90265 | 0.021954 | 24.3222 | 0.46746 | 0.15548 | 0.039012 | 2.5637 | 2.5637 | 0.002493
2019-03-20 12:57 | 1553101027.1802 | 444 | 0.20191 | 0.006978 | 0.000205 | 97.6924 | 92.3858 | 111.6731 | 20.9604 | 0.004527 | 0.00397 | 132.3966 | 0.26145 | 8.4039 | 0.003838 | 0.005648 | 0.005537 | 5e-05 | 0.001102 | 0.31636 | 1.3147 | 2.548 | 100.99 | 20.64 | 18.284 | 130.8898 | 114.0139 | 117.2429 | 0.004409 | 0.90433 | 0.021982 | 24.3073 | 0.49956 | 0.15553 | 0.037493 | 2.4668 | 2.4668 | 0.002591
2019-03-20 12:57 | 1553101028.18 | 445 | 0.20204 | 0.006876 | 0.000206 | 97.6843 | 92.7931 | 111.5329 | 20.9596 | 0.004526 | 0.003973 | 132.2992 | 0.26147 | 8.4026 | 0.003933 | 0.005557 | 0.005425 | 5.2e-05 | 0.0011 | 0.34363 | 1.3021 | 2.5479 | 100.99 | 20.64 | 18.281 | 130.8641 | 116.3316 | 117.3101 | 0.004408 | 0.90418 | 0.021966 | 24.2941 | 0.45946 | 0.15555 | 0.036819 | 2.4206 | 2.4206 | 0.002641
2019-03-20 12:57 | 1553101029.1791 | 446 | 0.20217 | 0.006761 | 0.000207 | 97.9391 | 93.174 | 111.4623 | 20.9594 | 0.004512 | 0.003973 | 132.2014 | 0.26046 | 8.3998 | 0.003912 | 0.00531 | 0.005175 | 5.4e-05 | 0.0011 | 0.29551 | 1.2936 | 2.5479 | 100.99 | 20.63 | 18.281 | 130.5072 | 115.817 | 117.31 | 0.004405 | 0.90418 | 0.021959 | 24.2864 | 0.44613 | 0.15495 | 0.036151 | 2.3758 | 2.3758 | 0.00268
2019-03-20 12:57 | 1553101030.1801 | 447 | 0.20249 | 0.006658 | 0.000207 | 98.2884 | 93.2656 | 111.5806 | 20.9599 | 0.004465 | 0.003948 | 132.1315 | 0.26097 | 8.4038 | 0.003855 | 0.005389 | 0.005282 | 5.2e-05 | 0.001101 | 0.34363 | 1.3564 | 2.548 | 100.99 | 20.63 | 18.281 | 129.3607 | 114.415 | 116.6921 | 0.004405 | 0.90561 | 0.021988 | 24.2801 | 0.46919 | 0.15525 | 0.034407 | 2.2641 | 2.2641 | 0.002818
2019-03-20 12:57 | 1553101031.1792 | 448 | 0.20267 | 0.006547 | 0.000207 | 98.5656 | 93.3026 | 111.6834 | 20.9592 | 0.004512 | 0.003962 | 132.0783 | 0.26098 | 8.4002 | 0.003904 | 0.005228 | 0.005061 | 5.1e-05 | 0.001101 | 0.33931 | 1.3025 | 2.5481 | 100.99 | 20.63 | 18.281 | 130.4992 | 115.6085 | 117.0272 | 0.004405 | 0.90483 | 0.021989 | 24.3017 | 0.49083 | 0.15526 | 0.033471 | 2.2025 | 2.2025 | 0.002897
2019-03-20 12:57 | 1553101032.1792 | 449 | 0.20277 | 0.006432 | 0.000206 | 98.6827 | 92.7318 | 112.0869 | 20.9592 | 0.004521 | 0.00394 | 131.9887 | 0.26072 | 8.4039 | 0.00387 | 0.005105 | 0.004962 | 5e-05 | 0.001099 | 0.32228 | 1.3321 | 2.548 | 100.99 | 20.63 | 18.284 | 130.7307 | 114.7821 | 116.5069 | 0.004406 | 0.90604 | 0.022043 | 24.3293 | 0.55633 | 0.1551 | 0.032992 | 2.1763 | 2.1763 | 0.002929
2019-03-20 12:57 | 1553101033.1804 | 450 | 0.20279 | 0.006312 | 0.000204 | 98.9534 | 93.2248 | 112.0554 | 20.9595 | 0.004524 | 0.003923 | 131.9127 | 0.26259 | 8.4029 | 0.003889 | 0.005693 | 0.005639 | 5e-05 | 0.001104 | 0.32637 | 1.3077 | 2.5481 | 100.99 | 20.63 | 18.286 | 130.7964 | 115.2655 | 116.0854 | 0.004407 | 0.90702 | 0.022052 | 24.3127 | 0.53341 | 0.15621 | 0.033017 | 2.1789 | 2.1789 | 0.002947
2019-03-20 12:57 | 1553101034.1784 | 451 | 0.20286 | 0.006185 | 0.000203 | 99.2115 | 93.9615 | 111.7268 | 20.959 | 0.004509 | 0.00391 | 131.8449 | 0.26154 | 8.4048 | 0.003837 | 0.005265 | 0.005137 | 5.4e-05 | 0.001102 | 0.3147 | 1.293 | 2.548 | 100.99 | 20.63 | 18.288 | 130.4304 | 113.9928 | 115.7707 | 0.004407 | 0.90776 | 0.022029 | 24.2672 | 0.4863 | 0.15559 | 0.03276 | 2.1598 | 2.1598 | 0.002961
2019-03-20 12:57 | 1553101035.1795 | 452 | 0.20306 | 0.006076 | 0.000201 | 99.3534 | 94.604 | 111.6578 | 20.9592 | 0.004428 | 0.003882 | 131.7715 | 0.26223 | 8.4052 | 0.00377 | 0.005109 | 0.004911 | 5.3e-05 | 0.001103 | 0.33373 | 1.3442 | 2.5481 | 100.99 | 20.63 | 18.294 | 128.4485 | 112.3456 | 115.0826 | 0.004408 | 0.90937 | 0.022041 | 24.2383 | 0.43814 | 0.156 | 0.031655 | 2.0881 | 2.0881 | 0.003071
2019-03-20 12:57 | 1553101036.1798 | 453 | 0.20315 | 0.005969 | 0.0002 | 99.5061 | 95.1026 | 111.743 | 20.9598 | 0.004427 | 0.003893 | 131.6984 | 0.26256 | 8.4042 | 0.003788 | 0.005504 | 0.005386 | 4.7e-05 | 0.001101 | 0.335 | 1.3327 | 2.5479 | 100.99 | 20.63 | 18.298 | 128.4132 | 112.7941 | 115.3419 | 0.004409 | 0.90876 | 0.022043 | 24.2556 | 0.40486 | 0.1562 | 0.031306 | 2.0652 | 2.0652 | 0.003109
2019-03-20 12:57 | 1553101037.1785 | 454 | 0.20318 | 0.00587 | 0.000199 | 99.6406 | 95.3115 | 112.219 | 20.9591 | 0.004369 | 0.003823 | 131.6265 | 0.26184 | 8.4013 | 0.003785 | 0.005222 | 0.005194 | 5.2e-05 | 0.001103 | 0.33124 | 1.3153 | 2.5482 | 100.99 | 20.64 | 18.302 | 126.9936 | 112.7185 | 113.6531 | 0.004413 | 0.91273 | 0.022138 | 24.2543 | 0.39731 | 0.15577 | 0.031197 | 2.0669 | 2.0669 | 0.003097
2019-03-20 12:57 | 1553101038.1795 | 455 | 0.20323 | 0.005771 | 0.000198 | 99.8266 | 95.3773 | 112.5279 | 20.9591 | 0.004361 | 0.003881 | 131.568 | 0.26199 | 8.4038 | 0.003799 | 0.005503 | 0.005336 | 5.2e-05 | 0.001105 | 0.31636 | 1.3754 | 2.5482 | 100.99 | 20.63 | 18.311 | 126.807 | 113.0581 | 115.0533 | 0.004413 | 0.90943 | 0.022128 | 24.3316 | 0.40782 | 0.15586 | 0.031067 | 2.0574 | 2.0574 | 0.003114
2019-03-20 12:57 | 1553101039.1796 | 456 | 0.20329 | 0.005683 | 0.000198 | 99.986 | 95.168 | 112.6993 | 20.9592 | 0.004314 | 0.003879 | 131.5265 | 0.26241 | 8.4036 | 0.003878 | 0.005072 | 0.005005 | 4.9e-05 | 0.00111 | 0.30182 | 1.3612 | 2.5482 | 100.99 | 20.63 | 18.317 | 125.6362 | 114.9956 | 115.0034 | 0.004414 | 0.90955 | 0.022146 | 24.3486 | 0.44175 | 0.15611 | 0.030803 | 2.0415 | 2.0415 | 0.003143
2019-03-20 12:57 | 1553101040.1786 | 457 | 0.20339 | 0.005592 | 0.000197 | 100.1953 | 94.8166 | 112.8528 | 20.9591 | 0.004388 | 0.003821 | 131.4551 | 0.26145 | 8.4021 | 0.003794 | 0.004938 | 0.004769 | 5.1e-05 | 0.001114 | 0.33804 | 1.3486 | 2.548 | 100.99 | 20.63 | 18.322 | 127.4546 | 112.9451 | 113.6024 | 0.004415 | 0.91285 | 0.022201 | 24.3211 | 0.49353 | 0.15554 | 0.030321 | 2.0146 | 2.0146 | 0.003173
2019-03-20 12:57 | 1553101041.1796 | 458 | 0.20341 | 0.00551 | 0.000195 | 100.3311 | 94.9385 | 112.8708 | 20.9589 | 0.004327 | 0.003846 | 131.3849 | 0.26169 | 8.4001 | 0.003718 | 0.005141 | 0.00494 | 5.1e-05 | 0.00111 | 0.30237 | 1.3588 | 2.5481 | 100.99 | 20.63 | 18.33 | 125.9725 | 111.0966 | 114.1951 | 0.004417 | 0.91145 | 0.022186 | 24.3417 | 0.49416 | 0.15568 | 0.030283 | 2.0108 | 2.0108 | 0.003182
2019-03-20 12:57 | 1553101042.1798 | 459 | 0.20352 | 0.005438 | 0.000195 | 100.5252 | 95.412 | 112.6347 | 20.9585 | 0.00431 | 0.0038 | 131.3116 | 0.26221 | 8.4037 | 0.003711 | 0.005322 | 0.005227 | 5e-05 | 0.001111 | 0.30879 | 1.3275 | 2.548 | 100.99 | 20.63 | 18.332 | 125.5504 | 110.9106 | 113.0955 | 0.004418 | 0.91404 | 0.022195 | 24.2817 | 0.46694 | 0.15599 | 0.029703 | 1.9729 | 1.9729 | 0.00325
2019-03-20 12:57 | 1553101043.1788 | 460 | 0.20361 | 0.005364 | 0.000193 | 100.6869 | 94.9987 | 112.63 | 20.9585 | 0.004333 | 0.003749 | 131.2278 | 0.26167 | 8.4054 | 0.003723 | 0.005572 | 0.005501 | 4.8e-05 | 0.001107 | 0.29894 | 1.308 | 2.5482 | 100.99 | 20.63 | 18.339 | 126.105 | 111.2105 | 111.8365 | 0.004419 | 0.91703 | 0.02223 | 24.2416 | 0.52015 | 0.15567 | 0.029259 | 1.9466 | 1.9466 | 0.003287
2019-03-20 12:57 | 1553101044.1802 | 461 | 0.20373 | 0.005299 | 0.000192 | 100.9275 | 95.5021 | 112.9274 | 20.9583 | 0.004275 | 0.003782 | 131.1439 | 0.26095 | 8.4023 | 0.003781 | 0.005144 | 0.004927 | 5.2e-05 | 0.001105 | 0.31006 | 1.2964 | 2.548 | 100.99 | 20.64 | 18.342 | 124.6829 | 112.6117 | 112.6497 | 0.004423 | 0.9151 | 0.022236 | 24.2992 | 0.49422 | 0.15524 | 0.028624 | 1.9048 | 1.9048 | 0.00335
2019-03-20 12:57 | 1553101045.1799 | 462 | 0.20381 | 0.005229 | 0.000193 | 101.0824 | 96.0429 | 113.2642 | 20.9583 | 0.004215 | 0.003716 | 131.0851 | 0.2616 | 8.4051 | 0.00383 | 0.00519 | 0.005007 | 4.8e-05 | 0.001109 | 0.29999 | 1.3178 | 2.548 | 100.99 | 20.63 | 18.345 | 123.2274 | 113.8143 | 111.0352 | 0.004421 | 0.91895 | 0.022316 | 24.2844 | 0.45743 | 0.15563 | 0.028247 | 1.8864 | 1.8864 | 0.003391
2019-03-20 12:57 | 1553101046.1794 | 463 | 0.2039 | 0.005162 | 0.000194 | 101.324 | 96.3501 | 113.1863 | 20.9578 | 0.004215 | 0.003773 | 131.0189 | 0.26153 | 8.4042 | 0.003813 | 0.00556 | 0.005454 | 5e-05 | 0.001111 | 0.32621 | 1.4005 | 2.5481 | 100.99 | 20.64 | 18.348 | 123.2099 | 113.4004 | 112.4335 | 0.004424 | 0.91561 | 0.022268 | 24.3202 | 0.45022 | 0.15559 | 0.027761 | 1.8497 | 1.8497 | 0.003457
2019-03-20 12:57 | 1553101047.18 | 464 | 0.20398 | 0.005092 | 0.000197 | 101.4317 | 96.6903 | 112.7194 | 20.9582 | 0.00417 | 0.003745 | 130.9544 | 0.26078 | 8.4018 | 0.003762 | 0.005079 | 0.004999 | 5e-05 | 0.001107 | 0.32621 | 1.3564 | 2.5482 | 100.99 | 20.64 | 18.346 | 122.1101 | 112.1528 | 111.7518 | 0.004424 | 0.91724 | 0.022242 | 24.2485 | 0.4282 | 0.15514 | 0.027383 | 1.8222 | 1.8222 | 0.003499
2019-03-20 12:57 | 1553101048.1804 | 465 | 0.20407 | 0.005022 | 0.000201 | 101.6477 | 96.3135 | 112.9725 | 20.9581 | 0.004208 | 0.003727 | 130.8993 | 0.26109 | 8.401 | 0.003701 | 0.004757 | 0.004614 | 5e-05 | 0.001102 | 0.34324 | 1.3462 | 2.5482 | 100.99 | 20.63 | 18.344 | 123.0378 | 110.679 | 111.2963 | 0.004421 | 0.91832 | 0.02228 | 24.2614 | 0.48215 | 0.15532 | 0.026915 | 1.7938 | 1.7938 | 0.003559
2019-03-20 12:57 | 1553101049.1798 | 466 | 0.20409 | 0.00495 | 0.000205 | 101.9314 | 96.1808 | 113.1999 | 20.958 | 0.004225 | 0.003777 | 130.8341 | 0.26001 | 8.4002 | 0.003684 | 0.005937 | 0.005775 | 5.2e-05 | 0.001101 | 0.33876 | 1.29 | 2.5481 | 100.99 | 20.63 | 18.339 | 123.4605 | 110.2616 | 112.5231 | 0.004419 | 0.9154 | 0.022267 | 24.3245 | 0.51941 | 0.15468 | 0.026911 | 1.7922 | 1.7922 | 0.003547
2019-03-20 12:57 | 1553101050.1782 | 467 | 0.20411 | 0.004884 | 0.000208 | 101.9779 | 95.6628 | 113.0424 | 20.9592 | 0.00423 | 0.00377 | 130.7587 | 0.26156 | 8.4032 | 0.003736 | 0.005462 | 0.005428 | 5e-05 | 0.001102 | 0.32388 | 1.3183 | 2.5481 | 100.99 | 20.63 | 18.335 | 123.5817 | 111.5245 | 112.3456 | 0.004418 | 0.91582 | 0.022256 | 24.302 | 0.5718 | 0.1556 | 0.026883 | 1.7894 | 1.7894 | 0.003574
2019-03-20 12:58 | 1553101051.1793 | 468 | 0.20415 | 0.004826 | 0.000209 | 102.1684 | 96.0997 | 112.8525 | 20.958 | 0.004326 | 0.003742 | 130.6801 | 0.26072 | 8.4057 | 0.003713 | 0.005196 | 0.004985 | 5.1e-05 | 0.001102 | 0.30237 | 1.3225 | 2.5481 | 100.99 | 20.63 | 18.326 | 125.9415 | 110.9632 | 111.6679 | 0.004416 | 0.91744 | 0.022257 | 24.2602 | 0.54773 | 0.1551 | 0.02669 | 1.7765 | 1.7765 | 0.003588
2019-03-20 12:58 | 1553101052.1783 | 469 | 0.20417 | 0.004775 | 0.00021 | 102.2936 | 96.4368 | 113.0617 | 20.9572 | 0.00421 | 0.003724 | 130.6346 | 0.26367 | 8.4019 | 0.003793 | 0.004986 | 0.004853 | 5.2e-05 | 0.001107 | 0.32228 | 1.2982 | 2.548 | 100.99 | 20.63 | 18.321 | 123.0988 | 112.9179 | 111.2423 | 0.004415 | 0.91845 | 0.02229 | 24.2692 | 0.52736 | 0.15686 | 0.026635 | 1.7754 | 1.7754 | 0.003631
2019-03-20 12:58 | 1553101053.18 | 470 | 0.20427 | 0.004728 | 0.00021 | 102.5171 | 96.9238 | 113.3314 | 20.9564 | 0.004204 | 0.003691 | 130.5882 | 0.26182 | 8.4014 | 0.003699 | 0.005173 | 0.005008 | 5.4e-05 | 0.001107 | 0.33085 | 1.2899 | 2.5479 | 100.99 | 20.63 | 18.313 | 122.937 | 110.6182 | 110.4424 | 0.004413 | 0.92037 | 0.02234 | 24.2729 | 0.50182 | 0.15576 | 0.026066 | 1.7412 | 1.7412 | 0.003677
2019-03-20 12:58 | 1553101054.1794 | 471 | 0.20437 | 0.004679 | 0.00021 | 102.6144 | 97.4718 | 113.4976 | 20.9569 | 0.004106 | 0.003652 | 130.5066 | 0.26245 | 8.4042 | 0.003696 | 0.00557 | 0.005492 | 5.3e-05 | 0.001106 | 0.33069 | 1.3109 | 2.548 | 100.99 | 20.63 | 18.308 | 120.5474 | 110.5557 | 109.491 | 0.004412 | 0.92265 | 0.022384 | 24.2605 | 0.45989 | 0.15613 | 0.025541 | 1.7094 | 1.7094 | 0.003754
2019-03-20 12:58 | 1553101055.1785 | 472 | 0.20451 | 0.004624 | 0.00021 | 102.846 | 97.4273 | 113.364 | 20.9579 | 0.004184 | 0.003689 | 130.4461 | 0.26146 | 8.4029 | 0.003689 | 0.00501 | 0.004851 | 5.3e-05 | 0.00111 | 0.33124 | 1.3333 | 2.5481 | 100.99 | 20.63 | 18.303 | 122.4634 | 110.3855 | 110.376 | 0.004411 | 0.92053 | 0.022345 | 24.2743 | 0.48413 | 0.15554 | 0.02476 | 1.654 | 1.654 | 0.003865
2019-03-20 12:58 | 1553101056.1786 | 473 | 0.20462 | 0.004568 | 0.00021 | 103.0633 | 97.3537 | 113.4565 | 20.9574 | 0.004101 | 0.003671 | 130.3848 | 0.26212 | 8.4025 | 0.003725 | 0.005229 | 0.005029 | 5.5e-05 | 0.001108 | 0.30702 | 1.3748 | 2.5481 | 100.99 | 20.63 | 18.292 | 120.4368 | 111.2601 | 109.936 | 0.004408 | 0.92158 | 0.022367 | 24.2702 | 0.50978 | 0.15593 | 0.02416 | 1.6152 | 1.6152 | 0.003968
2019-03-20 12:58 | 1553101057.1785 | 474 | 0.20472 | 0.00451 | 0.00021 | 103.2477 | 98.1255 | 113.8867 | 20.9563 | 0.004127 | 0.003639 | 130.3371 | 0.2622 | 8.4038 | 0.003678 | 0.004528 | 0.004488 | 5.3e-05 | 0.001112 | 0.31277 | 1.3411 | 2.5481 | 100.99 | 20.63 | 18.287 | 121.0721 | 110.1187 | 109.1669 | 0.004407 | 0.92344 | 0.022432 | 24.2918 | 0.45513 | 0.15598 | 0.023671 | 1.587 | 1.587 | 0.00404
2019-03-20 12:58 | 1553101058.179 | 475 | 0.20474 | 0.004449 | 0.000208 | 103.4919 | 98.6842 | 113.94 | 20.9561 | 0.004147 | 0.003643 | 130.2598 | 0.26143 | 8.3994 | 0.003721 | 0.005239 | 0.005045 | 5.2e-05 | 0.001116 | 0.30127 | 1.3283 | 2.5481 | 100.99 | 20.63 | 18.281 | 121.564 | 111.1585 | 109.2591 | 0.004405 | 0.92321 | 0.022434 | 24.3004 | 0.42547 | 0.15553 | 0.023618 | 1.5837 | 1.5837 | 0.004036
2019-03-20 12:58 | 1553101059.1797 | 476 | 0.20481 | 0.004377 | 0.000207 | 103.7357 | 98.648 | 113.8825 | 20.956 | 0.004137 | 0.003668 | 130.1873 | 0.26071 | 8.4006 | 0.003693 | 0.005519 | 0.00541 | 5.3e-05 | 0.001117 | 0.30502 | 1.3676 | 2.5481 | 100.99 | 20.63 | 18.27 | 121.3099 | 110.4712 | 109.8637 | 0.004403 | 0.92176 | 0.022411 | 24.3135 | 0.44981 | 0.1551 | 0.023253 | 1.5576 | 1.5576 | 0.004093
2019-03-20 12:58 | 1553101060.1797 | 477 | 0.20495 | 0.004311 | 0.000206 | 104.0083 | 99.1798 | 113.9374 | 20.9553 | 0.004053 | 0.003548 | 130.128 | 0.26122 | 8.4013 | 0.003683 | 0.005555 | 0.005343 | 5.2e-05 | 0.001118 | 0.30613 | 1.3295 | 2.548 | 100.99 | 20.63 | 18.264 | 119.2481 | 110.2273 | 106.9507 | 0.004401 | 0.92882 | 0.022502 | 24.2267 | 0.42519 | 0.1554 | 0.022504 | 1.5133 | 1.5133 | 0.00422
2019-03-20 12:58 | 1553101061.1788 | 478 | 0.20499 | 0.00424 | 0.000207 | 104.1422 | 100.4274 | 114.4296 | 20.9563 | 0.004099 | 0.003568 | 130.0842 | 0.26083 | 8.402 | 0.003662 | 0.005087 | 0.004965 | 5.6e-05 | 0.001118 | 0.33804 | 1.3406 | 2.5481 | 100.99 | 20.63 | 18.254 | 120.3697 | 109.7301 | 107.4423 | 0.004399 | 0.92762 | 0.022536 | 24.2946 | 0.32489 | 0.15517 | 0.022385 | 1.5075 | 1.5075 | 0.00423
2019-03-20 12:58 | 1553101062.1788 | 479 | 0.20519 | 0.004166 | 0.000207 | 104.3237 | 100.9301 | 114.5843 | 20.9552 | 0.004052 | 0.003562 | 130.0285 | 0.2614 | 8.4023 | 0.003594 | 0.004683 | 0.004588 | 5.2e-05 | 0.001119 | 0.2995 | 1.3211 | 2.548 | 100.99 | 20.63 | 18.249 | 119.229 | 108.0746 | 107.2907 | 0.004398 | 0.92799 | 0.022556 | 24.3062 | 0.2958 | 0.15551 | 0.021258 | 1.4325 | 1.4325 | 0.004462
2019-03-20 12:58 | 1553101063.1793 | 480 | 0.2053 | 0.004086 | 0.000208 | 104.6232 | 101.0717 | 114.5402 | 20.9537 | 0.004052 | 0.003547 | 130.0042 | 0.26104 | 8.402 | 0.003614 | 0.005183 | 0.005028 | 5.3e-05 | 0.001119 | 0.32693 | 1.3561 | 2.5481 | 100.99 | 20.63 | 18.245 | 119.2283 | 108.5569 | 106.9327 | 0.004397 | 0.92887 | 0.022562 | 24.2901 | 0.30889 | 0.15529 | 0.020709 | 1.3956 | 1.3956 | 0.004573
2019-03-20 12:58 | 1553101064.1796 | 481 | 0.20542 | 0.004009 | 0.000207 | 104.7716 | 101.2325 | 114.7646 | 20.9557 | 0.003994 | 0.00353 | 129.9658 | 0.26091 | 8.4022 | 0.003586 | 0.00538 | 0.005223 | 5.2e-05 | 0.001119 | 0.34683 | 1.346 | 2.548 | 100.99 | 20.63 | 18.239 | 117.8196 | 107.875 | 106.5212 | 0.004395 | 0.92987 | 0.022597 | 24.3007 | 0.30736 | 0.15522 | 0.020077 | 1.3548 | 1.3548 | 0.004709
2019-03-20 12:58 | 1553101065.179 | 482 | 0.20565 | 0.003926 | 0.000208 | 104.9928 | 101.4888 | 115.1545 | 20.9551 | 0.003967 | 0.00354 | 129.9275 | 0.26153 | 8.4028 | 0.003661 | 0.0054 | 0.005251 | 5.2e-05 | 0.001114 | 0.35651 | 1.3383 | 2.548 | 100.99 | 20.62 | 18.237 | 117.1625 | 109.7012 | 106.771 | 0.004392 | 0.92926 | 0.022627 | 24.35 | 0.30361 | 0.15559 | 0.018766 | 1.2675 | 1.2675 | 0.005045
2019-03-20 12:58 | 1553101066.178 | 483 | 0.20576 | 0.003837 | 0.000209 | 105.15 | 101.6802 | 115.8054 | 20.9545 | 0.003924 | 0.003524 | 129.8953 | 0.2606 | 8.404 | 0.003586 | 0.004828 | 0.004758 | 5.4e-05 | 0.001119 | 0.28926 | 1.3101 | 2.548 | 100.99 | 20.62 | 18.234 | 116.0985 | 107.8707 | 106.38 | 0.004391 | 0.93022 | 0.022703 | 24.4061 | 0.30014 | 0.15503 | 0.018251 | 1.2365 | 1.2365 | 0.005153
2019-03-20 12:58 | 1553101067.1784 | 484 | 0.20602 | 0.003736 | 0.000211 | 105.2997 | 102.4038 | 116.2725 | 20.9552 | 0.003923 | 0.003468 | 129.8648 | 0.26275 | 8.4028 | 0.003557 | 0.005288 | 0.005033 | 5e-05 | 0.001121 | 0.34683 | 1.3244 | 2.5481 | 100.99 | 20.62 | 18.234 | 116.0739 | 107.1844 | 105.0265 | 0.004391 | 0.93355 | 0.022789 | 24.4116 | 0.24944 | 0.15631 | 0.016811 | 1.1424 | 1.1424 | 0.005624
2019-03-20 12:58 | 1553101068.1782 | 485 | 0.20636 | 0.003623 | 0.000211 | 105.4362 | 103.4122 | 116.6417 | 20.9542 | 0.003951 | 0.003486 | 129.8536 | 0.2624 | 8.4019 | 0.003596 | 0.005043 | 0.004903 | 5.2e-05 | 0.001119 | 0.30663 | 1.306 | 2.5481 | 100.99 | 20.62 | 18.235 | 116.7719 | 108.1224 | 105.444 | 0.004392 | 0.93252 | 0.022813 | 24.4639 | 0.17338 | 0.1561 | 0.014896 | 1.0121 | 1.0121 | 0.006339
2019-03-20 12:58 | 1553101069.1791 | 486 | 0.20663 | 0.003485 | 0.00021 | 105.596 | 103.9765 | 117.1137 | 20.9543 | 0.003854 | 0.003433 | 129.84 | 0.26104 | 8.4028 | 0.00355 | 0.005213 | 0.005092 | 5.2e-05 | 0.00112 | 0.34634 | 1.3015 | 2.548 | 100.99 | 20.62 | 18.238 | 114.3978 | 106.9993 | 104.1732 | 0.004392 | 0.93566 | 0.022898 | 24.4721 | 0.13824 | 0.15529 | 0.013442 | 0.91572 | 0.91572 | 0.00697
2019-03-20 12:58 | 1553101070.1782 | 487 | 0.20692 | 0.00334 | 0.000209 | 105.7988 | 104.5207 | 117.3843 | 20.9538 | 0.003861 | 0.003435 | 129.8403 | 0.26176 | 8.4008 | 0.003583 | 0.005215 | 0.005159 | 5e-05 | 0.001123 | 0.29374 | 1.2959 | 2.548 | 100.99 | 20.62 | 18.242 | 114.5554 | 107.799 | 104.2143 | 0.004393 | 0.93556 | 0.022923 | 24.5017 | 0.10871 | 0.15572 | 0.011891 | 0.80973 | 0.80973 | 0.007904
2019-03-20 12:58 | 1553101071.1786 | 488 | 0.20718 | 0.00318 | 0.000207 | 105.8605 | 104.6889 | 117.6839 | 20.9536 | 0.003825 | 0.003432 | 129.8126 | 0.26217 | 8.4037 | 0.003548 | 0.005048 | 0.004909 | 5.2e-05 | 0.001121 | 0.30541 | 1.2994 | 2.548 | 100.99 | 20.62 | 18.247 | 113.7005 | 106.9585 | 104.1465 | 0.004394 | 0.93572 | 0.022954 | 24.5308 | 0.099548 | 0.15597 | 0.010513 | 0.71571 | 0.71571 | 0.008957
2019-03-20 12:58 | 1553101072.1784 | 489 | 0.20733 | 0.003026 | 0.000206 | 106.0404 | 105.1352 | 117.8611 | 20.9543 | 0.003855 | 0.00341 | 129.7873 | 0.26144 | 8.4033 | 0.003566 | 0.005435 | 0.005315 | 5.4e-05 | 0.001119 | 0.32676 | 1.3275 | 2.548 | 100.99 | 20.62 | 18.257 | 114.4221 | 107.3876 | 103.6025 | 0.004397 | 0.93707 | 0.022988 | 24.5315 | 0.076682 | 0.15553 | 0.009796 | 0.6671 | 0.6671 | 0.009582
2019-03-20 12:58 | 1553101073.1784 | 490 | 0.20742 | 0.002866 | 0.000205 | 106.29 | 105.7373 | 118.1547 | 20.9538 | 0.003858 | 0.003422 | 129.7495 | 0.26015 | 8.4022 | 0.003565 | 0.005167 | 0.005124 | 5e-05 | 0.001116 | 0.33395 | 1.3327 | 2.5481 | 100.99 | 20.62 | 18.263 | 114.4945 | 107.3563 | 103.8991 | 0.004398 | 0.93634 | 0.023007 | 24.5717 | 0.046638 | 0.15476 | 0.009437 | 0.64288 | 0.64288 | 0.009894
2019-03-20 12:58 | 1553101074.1795 | 491 | 0.20751 | 0.002705 | 0.000204 | 106.4259 | 105.9436 | 118.4596 | 20.9536 | 0.00388 | 0.003386 | 129.7576 | 0.26226 | 8.3997 | 0.003543 | 0.005254 | 0.005106 | 5.2e-05 | 0.001115 | 0.30702 | 1.3566 | 2.548 | 100.99 | 20.62 | 18.274 | 115.0218 | 106.8443 | 103.0376 | 0.004401 | 0.93848 | 0.023063 | 24.5753 | 0.040626 | 0.15602 | 0.009132 | 0.62329 | 0.62329 | 0.010288
2019-03-20 12:58 | 1553101075.1798 | 492 | 0.20766 | 0.002549 | 0.000202 | 106.6128 | 106.0524 | 118.8361 | 20.9536 | 0.003757 | 0.003334 | 129.7411 | 0.26126 | 8.4011 | 0.003598 | 0.004897 | 0.004792 | 5.1e-05 | 0.001109 | 0.32516 | 1.3368 | 2.548 | 100.99 | 20.62 | 18.278 | 112.0361 | 108.1682 | 101.7713 | 0.004402 | 0.94165 | 0.023139 | 24.5729 | 0.047145 | 0.15542 | 0.008421 | 0.57583 | 0.57583 | 0.011093
2019-03-20 12:58 | 1553101076.1796 | 493 | 0.20778 | 0.002404 | 0.000201 | 106.7748 | 106.4597 | 118.9065 | 20.9531 | 0.003794 | 0.003357 | 129.7414 | 0.26108 | 8.4037 | 0.003608 | 0.005495 | 0.005353 | 5e-05 | 0.001104 | 0.32139 | 1.3021 | 2.548 | 100.99 | 20.62 | 18.282 | 112.9249 | 108.4176 | 102.3293 | 0.004403 | 0.94025 | 0.023129 | 24.5984 | 0.026435 | 0.15531 | 0.007898 | 0.53925 | 0.53925 | 0.011838
2019-03-20 12:58 | 1553101077.1797 | 494 | 0.20796 | 0.002263 | 0.000198 | 106.8961 | 106.8062 | 118.9345 | 20.9533 | 0.003734 | 0.003382 | 129.7358 | 0.2616 | 8.4028 | 0.00355 | 0.005493 | 0.005379 | 5e-05 | 0.001102 | 0.32244 | 1.3315 | 2.5481 | 100.99 | 20.62 | 18.289 | 111.4801 | 107.0032 | 102.9304 | 0.004405 | 0.93875 | 0.023113 | 24.621 | 0.007527 | 0.15562 | 0.006986 | 0.47549 | 0.47549 | 0.013452
2019-03-20 12:58 | 1553101078.1787 | 495 | 0.208 | 0.002138 | 0.000195 | 107.0698 | 107.1157 | 119.2471 | 20.9534 | 0.003713 | 0.00331 | 129.7214 | 0.26175 | 8.4026 | 0.003491 | 0.005038 | 0.004852 | 5.1e-05 | 0.001103 | 0.33323 | 1.3702 | 2.548 | 100.99 | 20.62 | 18.291 | 110.9703 | 105.569 | 101.1782 | 0.004405 | 0.94314 | 0.023197 | 24.5958 | -0.003832 | 0.15572 | 0.006883 | 0.47023 | 0.47023 | 0.01361
2019-03-20 12:58 | 1553101079.1801 | 496 | 0.20807 | 0.002019 | 0.000191 | 107.1713 | 107.2398 | 119.48 | 20.9534 | 0.003712 | 0.003252 | 129.716 | 0.26157 | 8.4027 | 0.003543 | 0.005997 | 0.005881 | 5.2e-05 | 0.001103 | 0.35867 | 1.3419 | 2.5481 | 100.99 | 20.63 | 18.293 | 110.9278 | 106.83 | 99.7901 | 0.004408 | 0.94665 | 0.023263 | 24.5741 | -0.005708 | 0.15561 | 0.00665 | 0.45543 | 0.45543 | 0.014043
2019-03-20 12:58 | 1553101080.1788 | 497 | 0.20814 | 0.00191 | 0.000187 | 107.3405 | 107.51 | 119.8865 | 20.9544 | 0.003585 | 0.003266 | 129.6998 | 0.26135 | 8.4022 | 0.003533 | 0.00562 | 0.005566 | 5.4e-05 | 0.001104 | 0.2995 | 1.317 | 2.5481 | 100.99 | 20.62 | 18.294 | 107.8531 | 106.5813 | 100.1183 | 0.004406 | 0.94582 | 0.023293 | 24.6267 | -0.014112 | 0.15548 | 0.006353 | 0.43543 | 0.43543 | 0.014675
2019-03-20 12:58 | 1553101081.1788 | 498 | 0.2082 | 0.001812 | 0.000182 | 107.3864 | 107.6673 | 119.8565 | 20.9537 | 0.003657 | 0.003271 | 129.7078 | 0.26148 | 8.4021 | 0.00353 | 0.005202 | 0.005096 | 5.5e-05 | 0.001111 | 0.31974 | 1.3423 | 2.548 | 100.99 | 20.63 | 18.294 | 109.6038 | 106.5139 | 100.2462 | 0.004408 | 0.9455 | 0.023286 | 24.6279 | -0.023368 | 0.15555 | 0.0061 | 0.41783 | 0.41783 | 0.015301
2019-03-20 12:58 | 1553101082.1789 | 499 | 0.20825 | 0.001728 | 0.000178 | 107.4815 | 107.8628 | 119.9647 | 20.9525 | 0.003613 | 0.003177 | 129.7263 | 0.26085 | 8.4045 | 0.003672 | 0.005338 | 0.005235 | 4.8e-05 | 0.001112 | 0.32388 | 1.3704 | 2.5481 | 100.99 | 20.62 | 18.293 | 108.5423 | 109.9565 | 97.9817 | 0.004406 | 0.95127 | 0.023367 | 24.5642 | -0.031675 | 0.15518 | 0.005914 | 0.40643 | 0.40643 | 0.015692
2019-03-20 12:58 | 1553101083.1793 | 500 | 0.20833 | 0.001643 | 0.000174 | 107.5371 | 107.818 | 120.2182 | 20.953 | 0.003649 | 0.003289 | 129.7157 | 0.26119 | 8.4024 | 0.003575 | 0.005513 | 0.005452 | 4.8e-05 | 0.001114 | 0.32875 | 1.3536 | 2.5482 | 100.99 | 20.63 | 18.293 | 109.4169 | 107.6004 | 100.6864 | 0.004408 | 0.94439 | 0.023307 | 24.6795 | -0.023327 | 0.15538 | 0.005532 | 0.37874 | 0.37874 | 0.016862
2019-03-20 12:58 | 1553101084.1793 | 501 | 0.20838 | 0.001569 | 0.000169 | 107.6891 | 107.9525 | 119.8649 | 20.9533 | 0.003641 | 0.003253 | 129.6981 | 0.26125 | 8.4016 | 0.003577 | 0.00535 | 0.00512 | 5.1e-05 | 0.001113 | 0.32189 | 1.3131 | 2.5481 | 100.99 | 20.62 | 18.292 | 109.2051 | 107.6614 | 99.8175 | 0.004405 | 0.94659 | 0.0233 | 24.6146 | -0.021848 | 0.15542 | 0.005324 | 0.36432 | 0.36432 | 0.017533
2019-03-20 12:58 | 1553101085.1784 | 502 | 0.20837 | 0.001502 | 0.000165 | 107.7448 | 108.0224 | 119.5426 | 20.9529 | 0.003617 | 0.003184 | 129.7056 | 0.262 | 8.4028 | 0.003492 | 0.004987 | 0.004836 | 5.2e-05 | 0.001108 | 0.31581 | 1.392 | 2.5482 | 100.99 | 20.63 | 18.292 | 108.6278 | 105.5914 | 98.1467 | 0.004408 | 0.95084 | 0.023321 | 24.5264 | -0.023018 | 0.15586 | 0.005529 | 0.37929 | 0.37929 | 0.016889
2019-03-20 12:58 | 1553101086.1792 | 503 | 0.2084 | 0.001443 | 0.000159 | 107.9294 | 108.3752 | 119.2856 | 20.9537 | 0.00358 | 0.003141 | 129.6773 | 0.26136 | 8.4028 | 0.003462 | 0.004811 | 0.004724 | 5.1e-05 | 0.001113 | 0.30685 | 1.3799 | 2.5481 | 100.99 | 20.62 | 18.292 | 107.7421 | 104.8778 | 97.1059 | 0.004405 | 0.95352 | 0.023328 | 24.4656 | -0.036868 | 0.15548 | 0.005424 | 0.37239 | 0.37239 | 0.01716
2019-03-20 12:58 | 1553101087.1791 | 504 | 0.20845 | 0.001391 | 0.000154 | 107.9275 | 108.3251 | 119.0889 | 20.9528 | 0.003539 | 0.003162 | 129.6356 | 0.26031 | 8.4026 | 0.003525 | 0.005521 | 0.00535 | 4.9e-05 | 0.001112 | 0.34235 | 1.3561 | 2.5481 | 100.99 | 20.63 | 18.295 | 106.7333 | 106.3917 | 97.6089 | 0.004409 | 0.95222 | 0.023293 | 24.462 | -0.032894 | 0.15486 | 0.00515 | 0.3529 | 0.3529 | 0.018036
2019-03-20 12:58 | 1553101088.1793 | 505 | 0.20849 | 0.001349 | 0.00015 | 108.0539 | 108.5038 | 118.8281 | 20.9531 | 0.00357 | 0.00312 | 129.6292 | 0.26158 | 8.4022 | 0.003475 | 0.005415 | 0.005332 | 4.8e-05 | 0.001118 | 0.34396 | 1.3349 | 2.5481 | 100.99 | 20.63 | 18.296 | 107.4942 | 105.1887 | 96.589 | 0.004409 | 0.95485 | 0.0233 | 24.4016 | -0.03716 | 0.15562 | 0.005021 | 0.34415 | 0.34415 | 0.018584
2019-03-20 12:58 | 1553101089.1783 | 506 | 0.2085 | 0.001305 | 0.000146 | 108.1587 | 108.5367 | 118.773 | 20.9531 | 0.003531 | 0.003174 | 129.6364 | 0.26125 | 8.4013 | 0.003471 | 0.005313 | 0.005191 | 5e-05 | 0.001116 | 0.29767 | 1.3328 | 2.5481 | 100.99 | 20.62 | 18.298 | 106.5487 | 105.078 | 97.9102 | 0.004407 | 0.95145 | 0.023253 | 24.4395 | -0.031201 | 0.15542 | 0.005019 | 0.34351 | 0.34351 | 0.018596
2019-03-20 12:58 | 1553101090.1787 | 507 | 0.20852 | 0.001268 | 0.000143 | 108.2965 | 108.7837 | 118.5922 | 20.9531 | 0.003554 | 0.003196 | 129.6379 | 0.26103 | 8.4027 | 0.003471 | 0.005087 | 0.004957 | 5.2e-05 | 0.001119 | 0.32477 | 1.327 | 2.548 | 100.99 | 20.62 | 18.302 | 107.0959 | 105.0938 | 98.4292 | 0.004408 | 0.95012 | 0.023219 | 24.438 | -0.040151 | 0.15529 | 0.004923 | 0.3365 | 0.3365 | 0.018966
2019-03-20 12:58 | 1553101091.1784 | 508 | 0.20853 | 0.001234 | 0.000138 | 108.2926 | 108.9104 | 118.8103 | 20.9534 | 0.003539 | 0.003172 | 129.6261 | 0.26143 | 8.4001 | 0.003461 | 0.004948 | 0.004797 | 5.1e-05 | 0.001124 | 0.33804 | 1.3805 | 2.548 | 100.99 | 20.62 | 18.306 | 106.7314 | 104.8533 | 97.8587 | 0.004409 | 0.95158 | 0.023258 | 24.4416 | -0.050883 | 0.15553 | 0.004917 | 0.3369 | 0.3369 | 0.018973
2019-03-20 12:58 | 1553101092.1795 | 509 | 0.20857 | 0.001206 | 0.000134 | 108.4385 | 108.8656 | 119.0166 | 20.954 | 0.003468 | 0.003109 | 129.6045 | 0.26175 | 8.4019 | 0.003462 | 0.005566 | 0.005383 | 4.8e-05 | 0.001122 | 0.34899 | 1.323 | 2.5479 | 100.99 | 20.62 | 18.311 | 105.0253 | 104.8767 | 96.3289 | 0.00441 | 0.95552 | 0.023327 | 24.4124 | -0.03516 | 0.15572 | 0.004711 | 0.32368 | 0.32368 | 0.019772
2019-03-20 12:58 | 1553101093.1784 | 510 | 0.20853 | 0.001174 | 0.00013 | 108.4333 | 109.0528 | 118.93 | 20.9536 | 0.003444 | 0.003133 | 129.6097 | 0.26144 | 8.4025 | 0.003478 | 0.004872 | 0.004752 | 5.3e-05 | 0.001127 | 0.34634 | 1.317 | 2.548 | 100.99 | 20.62 | 18.314 | 104.425 | 105.2506 | 96.8981 | 0.004411 | 0.95405 | 0.0233 | 24.4223 | -0.050958 | 0.15553 | 0.004989 | 0.34311 | 0.34311 | 0.018631
2019-03-20 12:58 | 1553101094.1799 | 511 | 0.20864 | 0.001147 | 0.000126 | 108.4844 | 109.0611 | 119.545 | 20.9535 | 0.003477 | 0.003086 | 129.5948 | 0.26122 | 8.4025 | 0.003489 | 0.004944 | 0.004744 | 4.8e-05 | 0.001129 | 0.33301 | 1.3443 | 2.5482 | 100.99 | 20.62 | 18.317 | 105.2316 | 105.5322 | 95.7645 | 0.004411 | 0.95698 | 0.023396 | 24.4478 | -0.047418 | 0.1554 | 0.004377 | 0.3015 | 0.3015 | 0.021183
2019-03-20 12:58 | 1553101095.1795 | 512 | 0.20868 | 0.001127 | 0.000123 | 108.6008 | 109.2315 | 119.8489 | 20.9529 | 0.003437 | 0.00316 | 129.5556 | 0.26138 | 8.4009 | 0.003431 | 0.004674 | 0.00455 | 5e-05 | 0.001132 | 0.35513 | 1.3298 | 2.548 | 100.99 | 20.62 | 18.323 | 104.2653 | 104.1129 | 97.5611 | 0.004413 | 0.95235 | 0.023369 | 24.5384 | -0.05179 | 0.1555 | 0.004171 | 0.28683 | 0.28683 | 0.022281
2019-03-20 12:58 | 1553101096.1786 | 513 | 0.20869 | 0.001106 | 0.00012 | 108.6277 | 109.4453 | 120.1672 | 20.9534 | 0.003464 | 0.003099 | 129.5438 | 0.26168 | 8.4048 | 0.003469 | 0.005119 | 0.005055 | 5.2e-05 | 0.001129 | 0.31238 | 1.3096 | 2.548 | 100.99 | 20.62 | 18.325 | 104.9161 | 105.0366 | 96.0966 | 0.004413 | 0.95612 | 0.023446 | 24.5224 | -0.067069 | 0.15567 | 0.004118 | 0.28423 | 0.28423 | 0.02251
2019-03-20 12:58 | 1553101097.1786 | 514 | 0.20874 | 0.001083 | 0.000118 | 108.7267 | 109.4296 | 120.2832 | 20.9542 | 0.003391 | 0.003112 | 129.5542 | 0.26102 | 8.4036 | 0.003393 | 0.005015 | 0.004837 | 5.1e-05 | 0.001131 | 0.33019 | 1.3154 | 2.5481 | 100.99 | 20.62 | 18.33 | 103.156 | 103.2054 | 96.395 | 0.004414 | 0.95535 | 0.023448 | 24.5441 | -0.05764 | 0.15528 | 0.003879 | 0.26748 | 0.26748 | 0.023861
2019-03-20 12:58 | 1553101098.1791 | 515 | 0.20875 | 0.001058 | 0.000116 | 108.8224 | 109.4374 | 120.3799 | 20.9529 | 0.003376 | 0.00311 | 129.5404 | 0.26189 | 8.4037 | 0.003421 | 0.004923 | 0.004822 | 5.1e-05 | 0.001129 | 0.32892 | 1.3327 | 2.5482 | 100.99 | 20.62 | 18.332 | 102.7743 | 103.8842 | 96.3577 | 0.004415 | 0.95545 | 0.023459 | 24.5527 | -0.05041 | 0.1558 | 0.003802 | 0.26226 | 0.26226 | 0.024416
2019-03-20 12:58 | 1553101099.1778 | 516 | 0.20875 | 0.001038 | 0.000114 | 108.8997 | 109.7663 | 120.3419 | 20.952 | 0.00339 | 0.003041 | 129.5339 | 0.26126 | 8.405 | 0.003408 | 0.004937 | 0.004787 | 4.9e-05 | 0.001127 | 0.30287 | 1.3064 | 2.548 | 100.99 | 20.62 | 18.334 | 103.1168 | 103.5569 | 94.6945 | 0.004415 | 0.95977 | 0.023508 | 24.4936 | -0.070901 | 0.15543 | 0.003833 | 0.26516 | 0.26516 | 0.024091
2019-03-20 12:58 | 1553101100.1779 | 517 | 0.20877 | 0.001023 | 0.000112 | 108.9116 | 109.7218 | 120.4407 | 20.954 | 0.003422 | 0.00308 | 129.5118 | 0.26167 | 8.4024 | 0.003443 | 0.004667 | 0.004622 | 5.3e-05 | 0.001121 | 0.33373 | 1.2983 | 2.5481 | 100.99 | 20.62 | 18.336 | 103.8923 | 104.416 | 95.6303 | 0.004416 | 0.95733 | 0.023488 | 24.5348 | -0.066293 | 0.15567 | 0.003795 | 0.26238 | 0.26238 | 0.024385
2019-03-20 12:58 | 1553101101.179 | 518 | 0.2088 | 0.000999 | 0.00011 | 108.9734 | 109.7869 | 120.0503 | 20.953 | 0.003351 | 0.003059 | 129.492 | 0.26155 | 8.4029 | 0.00343 | 0.005498 | 0.005316 | 4.8e-05 | 0.001125 | 0.33124 | 1.3237 | 2.5481 | 100.99 | 20.62 | 18.337 | 102.1783 | 104.0902 | 95.129 | 0.004416 | 0.95864 | 0.023466 | 24.4783 | -0.066521 | 0.1556 | 0.003628 | 0.25038 | 0.25038 | 0.025541
2019-03-20 12:58 | 1553101102.1793 | 519 | 0.20883 | 0.000984 | 0.000109 | 109.0776 | 109.9326 | 119.6884 | 20.9523 | 0.003405 | 0.00306 | 129.4794 | 0.26208 | 8.4033 | 0.00339 | 0.005237 | 0.005133 | 4.9e-05 | 0.001124 | 0.30807 | 1.3107 | 2.5481 | 100.99 | 20.62 | 18.339 | 103.4841 | 103.1244 | 95.1503 | 0.004417 | 0.95858 | 0.02343 | 24.4421 | -0.069845 | 0.15591 | 0.003482 | 0.23969 | 0.23969 | 0.026735
2019-03-20 12:58 | 1553101103.1791 | 520 | 0.20882 | 0.000968 | 0.000109 | 109.1264 | 110.0769 | 119.498 | 20.9531 | 0.003377 | 0.003029 | 129.4579 | 0.26159 | 8.4046 | 0.003385 | 0.004706 | 0.004667 | 4.9e-05 | 0.001129 | 0.33892 | 1.3109 | 2.5481 | 100.99 | 20.62 | 18.341 | 102.8171 | 103.0041 | 94.388 | 0.004417 | 0.96057 | 0.023435 | 24.3974 | -0.077572 | 0.15562 | 0.003566 | 0.24573 | 0.24573 | 0.026028
2019-03-20 12:58 | 1553101104.1781 | 521 | 0.20887 | 0.000949 | 0.000108 | 109.0688 | 110.1023 | 119.0915 | 20.9522 | 0.003361 | 0.003012 | 129.4559 | 0.26082 | 8.4034 | 0.003358 | 0.005423 | 0.005269 | 5.3e-05 | 0.001131 | 0.34772 | 1.282 | 2.548 | 100.99 | 20.62 | 18.343 | 102.4093 | 102.3417 | 93.9824 | 0.004418 | 0.96163 | 0.023408 | 24.3424 | -0.084356 | 0.15516 | 0.00324 | 0.22246 | 0.22246 | 0.028665
2019-03-20 12:58 | 1553101105.1791 | 522 | 0.2088 | 0.000941 | 0.000107 | 109.172 | 110.1345 | 118.8577 | 20.9523 | 0.003308 | 0.003038 | 129.4691 | 0.26212 | 8.4029 | 0.0034 | 0.005078 | 0.004991 | 5.3e-05 | 0.001129 | 0.33268 | 1.3027 | 2.548 | 100.99 | 20.62 | 18.346 | 101.1389 | 103.3726 | 94.6219 | 0.004418 | 0.95996 | 0.023365 | 24.3396 | -0.078511 | 0.15594 | 0.003704 | 0.25485 | 0.25485 | 0.025148
2019-03-20 12:58 | 1553101106.1792 | 523 | 0.20886 | 0.000927 | 0.000107 | 109.238 | 110.2268 | 118.7136 | 20.9523 | 0.003313 | 0.002998 | 129.4578 | 0.26133 | 8.403 | 0.003336 | 0.00517 | 0.004981 | 5.1e-05 | 0.001126 | 0.32803 | 1.3224 | 2.5479 | 100.99 | 20.62 | 18.346 | 101.2589 | 101.8249 | 93.639 | 0.004418 | 0.96253 | 0.023382 | 24.2924 | -0.080601 | 0.15547 | 0.003341 | 0.22947 | 0.22947 | 0.027846
2019-03-20 12:58 | 1553101107.1783 | 524 | 0.20883 | 0.000916 | 0.000106 | 109.2858 | 110.173 | 118.5751 | 20.9515 | 0.003286 | 0.003007 | 129.4546 | 0.2612 | 8.4059 | 0.003389 | 0.004589 | 0.004439 | 5.1e-05 | 0.001123 | 0.33229 | 1.3122 | 2.5479 | 100.99 | 20.62 | 18.346 | 100.6011 | 103.089 | 93.8645 | 0.004418 | 0.96194 | 0.023361 | 24.2856 | -0.072314 | 0.15539 | 0.003528 | 0.24247 | 0.24247 | 0.026339
2019-03-20 12:58 | 1553101108.1793 | 525 | 0.20889 | 0.000909 | 0.000104 | 109.2635 | 110.3041 | 118.5601 | 20.9513 | 0.003228 | 0.003008 | 129.4554 | 0.26183 | 8.4025 | 0.00334 | 0.004521 | 0.004333 | 5e-05 | 0.001124 | 0.33804 | 1.3258 | 2.548 | 100.99 | 20.62 | 18.348 | 99.1988 | 101.9048 | 93.8985 | 0.004419 | 0.96185 | 0.023359 | 24.2852 | -0.084785 | 0.15576 | 0.003175 | 0.21773 | 0.21773 | 0.029403
2019-03-20 12:58 | 1553101109.1783 | 526 | 0.20892 | 0.000902 | 0.000104 | 109.3605 | 110.34 | 118.4674 | 20.952 | 0.003246 | 0.003012 | 129.4328 | 0.26188 | 8.4012 | 0.003332 | 0.005039 | 0.004925 | 4.8e-05 | 0.001119 | 0.31974 | 1.317 | 2.5482 | 100.99 | 20.62 | 18.349 | 99.6356 | 101.7099 | 93.9838 | 0.004419 | 0.96163 | 0.023347 | 24.2786 | -0.079755 | 0.15579 | 0.003042 | 0.20825 | 0.20825 | 0.030748
2019-03-20 12:58 | 1553101110.1787 | 527 | 0.20888 | 0.000891 | 0.000103 | 109.3402 | 110.4043 | 118.5512 | 20.9514 | 0.003254 | 0.003001 | 129.4304 | 0.26213 | 8.4026 | 0.003415 | 0.004044 | 0.003867 | 4.8e-05 | 0.001123 | 0.31598 | 1.3544 | 2.5479 | 100.99 | 20.62 | 18.35 | 99.8242 | 103.7301 | 93.7218 | 0.004419 | 0.96231 | 0.023363 | 24.2785 | -0.086631 | 0.15594 | 0.003302 | 0.22673 | 0.22673 | 0.028268
2019-03-20 12:58 | 1553101111.1785 | 528 | 0.20893 | 0.000879 | 0.000101 | 109.3867 | 110.6073 | 118.8582 | 20.9508 | 0.003243 | 0.002954 | 129.431 | 0.26067 | 8.4016 | 0.003431 | 0.004894 | 0.004746 | 5.3e-05 | 0.00112 | 0.34147 | 1.3358 | 2.548 | 100.99 | 20.62 | 18.351 | 99.5711 | 104.1249 | 92.594 | 0.004419 | 0.96528 | 0.02343 | 24.2725 | -0.099249 | 0.15507 | 0.003015 | 0.20721 | 0.20721 | 0.030759
2019-03-20 12:59 | 1553101112.1785 | 529 | 0.20891 | 0.000867 | 0.0001 | 109.4614 | 110.525 | 119.0276 | 20.9504 | 0.003242 | 0.00296 | 129.4219 | 0.26027 | 8.4025 | 0.003336 | 0.005203 | 0.005074 | 5.3e-05 | 0.001118 | 0.30287 | 1.329 | 2.5481 | 100.99 | 20.62 | 18.353 | 99.5463 | 101.8168 | 92.7257 | 0.00442 | 0.96493 | 0.023442 | 24.2942 | -0.08649 | 0.15483 | 0.003127 | 0.21532 | 0.21532 | 0.029555
2019-03-20 12:59 | 1553101113.1786 | 530 | 0.20891 | 0.000861 | 9.8e-05 | 109.5252 | 110.4634 | 119.3257 | 20.9509 | 0.003204 | 0.002935 | 129.4039 | 0.2611 | 8.4032 | 0.003389 | 0.004888 | 0.004811 | 5.2e-05 | 0.001113 | 0.35098 | 1.3459 | 2.5481 | 100.99 | 20.62 | 18.356 | 98.6143 | 103.1042 | 92.1233 | 0.004421 | 0.96652 | 0.023491 | 24.3045 | -0.076294 | 0.15532 | 0.003125 | 0.21568 | 0.21568 | 0.029599
2019-03-20 12:59 | 1553101114.1787 | 531 | 0.20894 | 0.000848 | 9.8e-05 | 109.5009 | 110.5516 | 119.5513 | 20.9493 | 0.003266 | 0.002971 | 129.4109 | 0.26238 | 8.4039 | 0.003442 | 0.00498 | 0.004831 | 5.1e-05 | 0.001113 | 0.31581 | 1.304 | 2.548 | 100.99 | 20.62 | 18.358 | 100.122 | 104.3743 | 92.9969 | 0.004421 | 0.96422 | 0.023485 | 24.3566 | -0.085414 | 0.15609 | 0.002957 | 0.20377 | 0.20377 | 0.031483
2019-03-20 12:59 | 1553101115.1787 | 532 | 0.20895 | 0.000837 | 9.6e-05 | 109.4721 | 110.3871 | 119.5889 | 20.9496 | 0.003195 | 0.002957 | 129.3873 | 0.26126 | 8.4035 | 0.003383 | 0.004492 | 0.004365 | 5e-05 | 0.001111 | 0.32477 | 1.3554 | 2.548 | 100.99 | 20.62 | 18.36 | 98.4083 | 102.9489 | 92.6553 | 0.004422 | 0.96512 | 0.0235 | 24.349 | -0.07445 | 0.15542 | 0.002949 | 0.20343 | 0.20343 | 0.031401
2019-03-20 12:59 | 1553101116.1793 | 533 | 0.20896 | 0.000829 | 9.4e-05 | 109.444 | 110.4533 | 119.656 | 20.9509 | 0.003256 | 0.002923 | 129.3945 | 0.26149 | 8.4031 | 0.003341 | 0.004785 | 0.004591 | 5.5e-05 | 0.001111 | 0.32709 | 1.3231 | 2.5481 | 100.99 | 20.62 | 18.361 | 99.8709 | 101.9367 | 91.8481 | 0.004422 | 0.96725 | 0.023532 | 24.329 | -0.082108 | 0.15556 | 0.002898 | 0.2002 | 0.2002 | 0.031935
2019-03-20 12:59 | 1553101117.1787 | 534 | 0.20896 | 0.000823 | 9.3e-05 | 109.5323 | 110.6738 | 119.8532 | 20.9493 | 0.003223 | 0.002941 | 129.3679 | 0.26059 | 8.4042 | 0.003396 | 0.004483 | 0.00443 | 5.3e-05 | 0.001111 | 0.33019 | 1.3007 | 2.5481 | 100.99 | 20.62 | 18.362 | 99.0739 | 103.2644 | 92.2797 | 0.004422 | 0.96611 | 0.023538 | 24.3634 | -0.092735 | 0.15502 | 0.002873 | 0.19855 | 0.19855 | 0.032091
2019-03-20 12:59 | 1553101118.1779 | 535 | 0.20899 | 0.000818 | 9.2e-05 | 109.4993 | 110.5602 | 119.7168 | 20.9489 | 0.003198 | 0.002974 | 129.3644 | 0.26203 | 8.4028 | 0.003408 | 0.005394 | 0.005207 | 5.2e-05 | 0.001111 | 0.33323 | 1.3032 | 2.5481 | 100.99 | 20.61 | 18.364 | 98.4862 | 103.5564 | 93.0645 | 0.00442 | 0.96404 | 0.023499 | 24.3757 | -0.086242 | 0.15588 | 0.002744 | 0.18916 | 0.18916 | 0.03387
2019-03-20 12:59 | 1553101119.1789 | 536 | 0.20898 | 0.000812 | 9.1e-05 | 109.4563 | 110.6809 | 119.7813 | 20.9496 | 0.003224 | 0.002935 | 129.3681 | 0.26156 | 8.4004 | 0.003339 | 0.004504 | 0.004373 | 5.6e-05 | 0.001107 | 0.31902 | 1.3086 | 2.5481 | 100.99 | 20.61 | 18.364 | 99.093 | 101.8835 | 92.1353 | 0.00442 | 0.96649 | 0.023535 | 24.3513 | -0.099517 | 0.1556 | 0.002801 | 0.19354 | 0.19354 | 0.033045
2019-03-20 12:59 | 1553101120.1781 | 537 | 0.20896 | 0.000811 | 8.9e-05 | 109.4805 | 110.6237 | 119.9257 | 20.9499 | 0.003155 | 0.002901 | 129.3525 | 0.26206 | 8.4019 | 0.003375 | 0.004085 | 0.004002 | 5.4e-05 | 0.001104 | 0.31542 | 1.3453 | 2.5481 | 100.99 | 20.62 | 18.364 | 97.4391 | 102.7472 | 91.3146 | 0.004423 | 0.96867 | 0.023576 | 24.3386 | -0.092913 | 0.1559 | 0.002897 | 0.20082 | 0.20082 | 0.031907
2019-03-20 12:59 | 1553101121.1783 | 538 | 0.20897 | 0.000804 | 8.5e-05 | 109.4657 | 110.4757 | 119.9846 | 20.95 | 0.003209 | 0.002908 | 129.3323 | 0.26117 | 8.4023 | 0.003382 | 0.005403 | 0.005223 | 5.5e-05 | 0.001103 | 0.33843 | 1.3245 | 2.548 | 100.99 | 20.62 | 18.364 | 98.7398 | 102.9202 | 91.4797 | 0.004423 | 0.96823 | 0.023576 | 24.3501 | -0.082153 | 0.15537 | 0.002888 | 0.20039 | 0.20039 | 0.031865
2019-03-20 12:59 | 1553101122.1785 | 539 | 0.20897 | 0.000803 | 8.4e-05 | 109.4896 | 110.6633 | 120.2632 | 20.9496 | 0.003202 | 0.002893 | 129.336 | 0.26188 | 8.4018 | 0.003377 | 0.004879 | 0.00478 | 5.2e-05 | 0.001101 | 0.33268 | 1.3285 | 2.5481 | 100.99 | 20.62 | 18.364 | 98.5683 | 102.8026 | 91.1111 | 0.004423 | 0.96921 | 0.023616 | 24.366 | -0.095376 | 0.15579 | 0.002861 | 0.19889 | 0.19889 | 0.032193
2019-03-20 12:59 | 1553101123.1791 | 540 | 0.209 | 0.000799 | 8.4e-05 | 109.4851 | 110.7216 | 120.1729 | 20.9503 | 0.003167 | 0.002901 | 129.3619 | 0.26087 | 8.4018 | 0.003365 | 0.00552 | 0.005373 | 5.3e-05 | 0.001103 | 0.32012 | 1.3512 | 2.548 | 100.99 | 20.62 | 18.364 | 97.7255 | 102.5197 | 91.3048 | 0.004423 | 0.96869 | 0.023601 | 24.3633 | -0.10046 | 0.15519 | 0.002714 | 0.18831 | 0.18831 | 0.033871
2019-03-20 12:59 | 1553101124.1792 | 541 | 0.209 | 0.000792 | 8.3e-05 | 109.4547 | 110.5538 | 119.9832 | 20.9498 | 0.003101 | 0.002878 | 129.3368 | 0.26097 | 8.4032 | 0.003422 | 0.005379 | 0.005261 | 5.4e-05 | 0.001103 | 0.32068 | 1.3392 | 2.5481 | 100.99 | 20.62 | 18.364 | 96.1244 | 103.8865 | 90.7516 | 0.004423 | 0.97017 | 0.0236 | 24.3256 | -0.089374 | 0.15525 | 0.002716 | 0.18852 | 0.18852 | 0.033847
2019-03-20 12:59 | 1553101125.1792 | 542 | 0.20899 | 0.000789 | 8.2e-05 | 109.4815 | 110.6155 | 120.1015 | 20.9499 | 0.003166 | 0.002855 | 129.3607 | 0.26163 | 8.4005 | 0.003413 | 0.005125 | 0.004979 | 5e-05 | 0.001101 | 0.30055 | 1.3334 | 2.5482 | 100.99 | 20.62 | 18.364 | 97.7111 | 103.6855 | 90.1914 | 0.004423 | 0.97166 | 0.02363 | 24.3188 | -0.092169 | 0.15564 | 0.002754 | 0.19149 | 0.19149 | 0.033406
2019-03-20 12:59 | 1553101126.1781 | 543 | 0.20903 | 0.00078 | 7.9e-05 | 109.5336 | 110.7639 | 120.0656 | 20.9499 | 0.003122 | 0.002877 | 129.3531 | 0.26029 | 8.4035 | 0.003375 | 0.00429 | 0.004132 | 5.3e-05 | 0.001098 | 0.33229 | 1.3478 | 2.5482 | 100.99 | 20.62 | 18.363 | 96.6369 | 102.7622 | 90.7267 | 0.004422 | 0.97023 | 0.023609 | 24.3331 | -0.099906 | 0.15485 | 0.002556 | 0.17739 | 0.17739 | 0.035877
2019-03-20 12:59 | 1553101127.1782 | 544 | 0.20901 | 0.000781 | 7.8e-05 | 109.4731 | 110.5998 | 120.2214 | 20.9487 | 0.00308 | 0.00285 | 129.3639 | 0.26063 | 8.4024 | 0.003394 | 0.004757 | 0.004613 | 5.4e-05 | 0.001097 | 0.35563 | 1.3171 | 2.5479 | 100.99 | 20.61 | 18.36 | 95.6244 | 103.2064 | 90.091 | 0.004419 | 0.97193 | 0.023645 | 24.3276 | -0.091587 | 0.15505 | 0.002652 | 0.18461 | 0.18461 | 0.034519
2019-03-20 12:59 | 1553101128.1777 | 545 | 0.209 | 0.000777 | 7.7e-05 | 109.4689 | 110.7381 | 120.5731 | 20.9495 | 0.003122 | 0.002844 | 129.3419 | 0.26078 | 8.4041 | 0.003375 | 0.004883 | 0.00482 | 5.1e-05 | 0.001095 | 0.31581 | 1.3117 | 2.5479 | 100.99 | 20.61 | 18.358 | 96.6497 | 102.7555 | 89.9354 | 0.004418 | 0.97235 | 0.023684 | 24.358 | -0.10311 | 0.15514 | 0.002725 | 0.19012 | 0.19012 | 0.033537
2019-03-20 12:59 | 1553101129.1774 | 546 | 0.20903 | 0.000774 | 7.7e-05 | 109.4621 | 110.7672 | 120.4545 | 20.9501 | 0.003039 | 0.002852 | 129.3418 | 0.26111 | 8.4025 | 0.003333 | 0.0045 | 0.004417 | 5.3e-05 | 0.00109 | 0.32532 | 1.2885 | 2.5478 | 100.99 | 20.61 | 18.356 | 94.6386 | 101.7357 | 90.1391 | 0.004418 | 0.9718 | 0.023666 | 24.3528 | -0.10602 | 0.15533 | 0.00258 | 0.17963 | 0.17963 | 0.035541
2019-03-20 12:59 | 1553101130.1785 | 547 | 0.20908 | 0.000766 | 7.7e-05 | 109.4453 | 110.6334 | 120.5083 | 20.9494 | 0.003133 | 0.002872 | 129.3661 | 0.26114 | 8.4004 | 0.003391 | 0.005165 | 0.005078 | 5.2e-05 | 0.001085 | 0.32101 | 1.3421 | 2.5481 | 100.99 | 20.62 | 18.351 | 96.8996 | 103.1434 | 90.605 | 0.004419 | 0.97056 | 0.023656 | 24.3739 | -0.096579 | 0.15535 | 0.002251 | 0.15611 | 0.15611 | 0.0409
2019-03-20 12:59 | 1553101131.1775 | 548 | 0.20902 | 0.000768 | 7.6e-05 | 109.4256 | 110.5385 | 120.6134 | 20.9497 | 0.003059 | 0.002884 | 129.3807 | 0.26063 | 8.4002 | 0.0034 | 0.005181 | 0.004969 | 5e-05 | 0.001083 | 0.32565 | 1.333 | 2.5482 | 100.99 | 20.62 | 18.347 | 95.1157 | 103.359 | 90.8898 | 0.004419 | 0.9698 | 0.023657 | 24.394 | -0.090513 | 0.15505 | 0.00262 | 0.18251 | 0.18251 | 0.034916
2019-03-20 12:59 | 1553101132.1775 | 549 | 0.20903 | 0.000766 | 7.5e-05 | 109.3576 | 110.6031 | 120.616 | 20.9494 | 0.003091 | 0.002851 | 129.3745 | 0.26094 | 8.4019 | 0.003429 | 0.005037 | 0.00493 | 5.1e-05 | 0.001083 | 0.34164 | 1.3174 | 2.5481 | 100.99 | 20.62 | 18.342 | 95.885 | 104.0583 | 90.097 | 0.004417 | 0.97192 | 0.023683 | 24.3677 | -0.1013 | 0.15524 | 0.002585 | 0.18029 | 0.18029 | 0.035389
2019-03-20 12:59 | 1553101133.179 | 550 | 0.209 | 0.00076 | 7.4e-05 | 109.3515 | 110.8255 | 120.7301 | 20.9484 | 0.003121 | 0.002844 | 129.3691 | 0.26056 | 8.4031 | 0.00338 | 0.004522 | 0.004404 | 5.1e-05 | 0.001083 | 0.30071 | 1.3133 | 2.548 | 100.99 | 20.62 | 18.334 | 96.6117 | 102.8699 | 89.9257 | 0.004415 | 0.97237 | 0.0237 | 24.3735 | -0.11977 | 0.15501 | 0.002738 | 0.1914 | 0.1914 | 0.033285
2019-03-20 12:59 | 1553101134.1786 | 551 | 0.20905 | 0.000756 | 7.3e-05 | 109.324 | 110.5411 | 121.0172 | 20.9498 | 0.003033 | 0.002817 | 129.3832 | 0.26148 | 8.4045 | 0.003383 | 0.004697 | 0.00456 | 5e-05 | 0.001083 | 0.27616 | 1.3233 | 2.5479 | 100.99 | 20.62 | 18.328 | 94.4988 | 102.9623 | 89.2833 | 0.004414 | 0.9741 | 0.023749 | 24.3808 | -0.09903 | 0.15555 | 0.002494 | 0.1744 | 0.1744 | 0.036658
2019-03-20 12:59 | 1553101135.179 | 552 | 0.20908 | 0.000755 | 7.2e-05 | 109.3256 | 110.4428 | 120.9479 | 20.9496 | 0.003028 | 0.002897 | 129.3842 | 0.26094 | 8.4018 | 0.003396 | 0.005053 | 0.00497 | 5.2e-05 | 0.001087 | 0.32587 | 1.2984 | 2.5479 | 100.99 | 20.61 | 18.316 | 94.3681 | 103.2734 | 91.2056 | 0.004408 | 0.96896 | 0.02368 | 24.4384 | -0.09094 | 0.15524 | 0.002315 | 0.16112 | 0.16112 | 0.039599
2019-03-20 12:59 | 1553101136.1778 | 553 | 0.20903 | 0.000753 | 7.2e-05 | 109.2419 | 110.3034 | 121.1179 | 20.9497 | 0.003022 | 0.002785 | 129.3824 | 0.26165 | 8.4027 | 0.003418 | 0.005169 | 0.005069 | 5.5e-05 | 0.001084 | 0.32781 | 1.3313 | 2.548 | 100.99 | 20.61 | 18.309 | 94.2326 | 103.7921 | 88.5145 | 0.004407 | 0.97617 | 0.023784 | 24.3651 | -0.086498 | 0.15565 | 0.002579 | 0.18082 | 0.18082 | 0.03538
2019-03-20 12:59 | 1553101137.1791 | 554 | 0.20903 | 0.000752 | 7.2e-05 | 109.28 | 110.6282 | 121.1329 | 20.9489 | 0.003043 | 0.002813 | 129.3851 | 0.26178 | 8.401 | 0.003397 | 0.004634 | 0.004514 | 5.3e-05 | 0.001082 | 0.32659 | 1.3397 | 2.5479 | 100.99 | 20.61 | 18.297 | 94.7257 | 103.2986 | 89.1973 | 0.004404 | 0.97433 | 0.023763 | 24.3896 | -0.10967 | 0.15574 | 0.002577 | 0.18051 | 0.18051 | 0.03546
2019-03-20 12:59 | 1553101138.1779 | 555 | 0.20906 | 0.000751 | 7.1e-05 | 109.2467 | 110.3225 | 121.2087 | 20.9497 | 0.003012 | 0.002789 | 129.4006 | 0.26065 | 8.4026 | 0.003487 | 0.004542 | 0.004438 | 5.4e-05 | 0.001083 | 0.31166 | 1.3782 | 2.5479 | 100.99 | 20.61 | 18.292 | 93.9886 | 105.4757 | 88.613 | 0.004403 | 0.9759 | 0.02379 | 24.3776 | -0.087647 | 0.15506 | 0.002444 | 0.17125 | 0.17125 | 0.037214
2019-03-20 12:59 | 1553101139.1793 | 556 | 0.20903 | 0.000754 | 6.9e-05 | 109.2017 | 110.3897 | 121.5216 | 20.9486 | 0.003034 | 0.002813 | 129.3898 | 0.26117 | 8.4028 | 0.003399 | 0.004807 | 0.004719 | 5.5e-05 | 0.001084 | 0.32117 | 1.3254 | 2.5481 | 100.99 | 20.61 | 18.286 | 94.5221 | 103.3453 | 89.1907 | 0.004401 | 0.97435 | 0.023802 | 24.4285 | -0.096781 | 0.15537 | 0.002606 | 0.18308 | 0.18308 | 0.034879
2019-03-20 12:59 | 1553101140.1794 | 557 | 0.20902 | 0.000756 | 6.8e-05 | 109.1756 | 110.1584 | 121.3671 | 20.9482 | 0.003001 | 0.002785 | 129.3625 | 0.26041 | 8.4041 | 0.003416 | 0.005054 | 0.004903 | 5.5e-05 | 0.001081 | 0.3235 | 1.3746 | 2.5481 | 100.99 | 20.62 | 18.276 | 93.718 | 103.749 | 88.5193 | 0.004401 | 0.97615 | 0.023809 | 24.3903 | -0.080162 | 0.15492 | 0.002625 | 0.18455 | 0.18455 | 0.034502
2019-03-20 12:59 | 1553101141.1792 | 558 | 0.20899 | 0.000748 | 6.7e-05 | 109.0928 | 110.3759 | 121.1758 | 20.9486 | 0.003034 | 0.0028 | 129.3316 | 0.26073 | 8.4051 | 0.003436 | 0.004501 | 0.004396 | 5.2e-05 | 0.001083 | 0.31885 | 1.3572 | 2.5479 | 100.99 | 20.62 | 18.272 | 94.5025 | 104.2432 | 88.8767 | 0.0044 | 0.97519 | 0.023778 | 24.3831 | -0.10459 | 0.15511 | 0.002821 | 0.19841 | 0.19841 | 0.03213
2019-03-20 12:59 | 1553101142.1784 | 559 | 0.20906 | 0.000745 | 6.7e-05 | 109.0184 | 110.2137 | 121.0627 | 20.9492 | 0.002992 | 0.002794 | 129.3158 | 0.26054 | 8.4042 | 0.003423 | 0.004686 | 0.004514 | 5.1e-05 | 0.001086 | 0.33085 | 1.287 | 2.5478 | 100.99 | 20.62 | 18.268 | 93.5091 | 103.9236 | 88.735 | 0.0044 | 0.97557 | 0.023772 | 24.367 | -0.097539 | 0.155 | 0.002438 | 0.1709 | 0.1709 | 0.037275
2019-03-20 12:59 | 1553101143.1781 | 560 | 0.20897 | 0.000749 | 6.6e-05 | 109.032 | 110.1625 | 121.0697 | 20.9495 | 0.002934 | 0.002809 | 129.2945 | 0.26033 | 8.4023 | 0.003391 | 0.005072 | 0.004909 | 5.5e-05 | 0.001085 | 0.32892 | 1.3279 | 2.548 | 100.99 | 20.62 | 18.262 | 92.114 | 103.1459 | 89.0916 | 0.004398 | 0.97461 | 0.023761 | 24.3797 | -0.092268 | 0.15487 | 0.002975 | 0.20934 | 0.20934 | 0.030406
2019-03-20 12:59 | 1553101144.1771 | 561 | 0.209 | 0.00075 | 6.5e-05 | 109.0708 | 110.263 | 121.0155 | 20.9474 | 0.002959 | 0.002786 | 129.2865 | 0.26062 | 8.4023 | 0.003428 | 0.004604 | 0.004466 | 5e-05 | 0.001089 | 0.33915 | 1.2827 | 2.548 | 100.99 | 20.62 | 18.26 | 92.7009 | 104.0438 | 88.5342 | 0.004398 | 0.97611 | 0.023774 | 24.3555 | -0.097244 | 0.15504 | 0.002776 | 0.19527 | 0.19527 | 0.032633
2019-03-20 12:59 | 1553101145.1774 | 562 | 0.20907 | 0.000746 | 6.5e-05 | 108.9317 | 110.2077 | 121.4275 | 20.9494 | 0.003004 | 0.002739 | 129.277 | 0.25983 | 8.4044 | 0.003438 | 0.005134 | 0.005069 | 5.6e-05 | 0.001086 | 0.32101 | 1.331 | 2.548 | 100.99 | 20.62 | 18.258 | 93.791 | 104.2864 | 87.413 | 0.004397 | 0.97915 | 0.023851 | 24.359 | -0.10417 | 0.15457 | 0.002359 | 0.16587 | 0.16587 | 0.038299
2019-03-20 12:59 | 1553101146.1773 | 563 | 0.20895 | 0.000747 | 6.4e-05 | 108.9802 | 110.1348 | 121.5275 | 20.9488 | 0.00295 | 0.002777 | 129.2726 | 0.26171 | 8.4005 | 0.003431 | 0.005027 | 0.004845 | 5.3e-05 | 0.001085 | 0.31542 | 1.3567 | 2.548 | 100.99 | 20.62 | 18.258 | 92.4973 | 104.1009 | 88.3153 | 0.004397 | 0.97671 | 0.023831 | 24.3995 | -0.094268 | 0.15569 | 0.003095 | 0.21868 | 0.21868 | 0.029261
2019-03-20 12:59 | 1553101147.1774 | 564 | 0.2091 | 0.000747 | 6.4e-05 | 108.8984 | 110.0276 | 121.5812 | 20.9484 | 0.002964 | 0.002813 | 129.269 | 0.26076 | 8.4032 | 0.003467 | 0.005407 | 0.005272 | 5.2e-05 | 0.00109 | 0.36033 | 1.3414 | 2.5479 | 100.99 | 20.62 | 18.258 | 92.8207 | 104.9956 | 89.185 | 0.004397 | 0.97436 | 0.023808 | 24.4343 | -0.09227 | 0.15512 | 0.002209 | 0.15486 | 0.15486 | 0.041171
2019-03-20 12:59 | 1553101148.1774 | 565 | 0.20892 | 0.000746 | 6.4e-05 | 108.8958 | 110.2477 | 121.6354 | 20.9488 | 0.002986 | 0.00282 | 129.2577 | 0.26232 | 8.4016 | 0.003531 | 0.005056 | 0.004939 | 5.1e-05 | 0.001087 | 0.32964 | 1.3144 | 2.548 | 100.99 | 20.62 | 18.26 | 93.35 | 106.5389 | 89.3589 | 0.004398 | 0.97389 | 0.023807 | 24.4456 | -0.11036 | 0.15606 | 0.003271 | 0.23114 | 0.23114 | 0.027749
2019-03-20 12:59 | 1553101149.1774 | 566 | 0.20903 | 0.000749 | 6.4e-05 | 108.8877 | 110.1046 | 121.7047 | 20.9481 | 0.002925 | 0.002809 | 129.2541 | 0.26091 | 8.4038 | 0.00351 | 0.004826 | 0.004659 | 5e-05 | 0.001085 | 0.32316 | 1.3294 | 2.5479 | 100.99 | 20.62 | 18.262 | 91.8907 | 106.029 | 89.0997 | 0.004398 | 0.97459 | 0.023823 | 24.4438 | -0.099404 | 0.15521 | 0.002621 | 0.18456 | 0.18456 | 0.034565
2019-03-20 12:59 | 1553101150.1775 | 567 | 0.20891 | 0.000748 | 6.4e-05 | 108.9459 | 110.062 | 121.841 | 20.9476 | 0.003025 | 0.002758 | 129.2553 | 0.26138 | 8.4016 | 0.003429 | 0.004967 | 0.004875 | 5.1e-05 | 0.001086 | 0.30431 | 1.3171 | 2.548 | 100.99 | 20.62 | 18.266 | 94.3003 | 104.0621 | 87.862 | 0.004399 | 0.97793 | 0.023877 | 24.4157 | -0.091161 | 0.1555 | 0.003333 | 0.23628 | 0.23628 | 0.027048
2019-03-20 12:59 | 1553101151.1785 | 568 | 0.20904 | 0.000752 | 6.4e-05 | 108.8903 | 110.0369 | 121.9385 | 20.9472 | 0.002979 | 0.002737 | 129.2889 | 0.26169 | 8.4011 | 0.003545 | 0.004587 | 0.004459 | 5.3e-05 | 0.00109 | 0.29999 | 1.3001 | 2.548 | 100.99 | 20.62 | 18.268 | 93.1746 | 106.8821 | 87.3548 | 0.0044 | 0.97931 | 0.023903 | 24.4083 | -0.093694 | 0.15568 | 0.002572 | 0.18173 | 0.18173 | 0.03521
2019-03-20 12:59 | 1553101152.1775 | 569 | 0.2097 | 0.000752 | 6.4e-05 | 108.8292 | 110.072 | 121.8415 | 20.9478 | 0.002942 | 0.002718 | 129.2931 | 0.26141 | 8.4015 | 0.003524 | 0.005086 | 0.004893 | 5e-05 | 0.001092 | 0.33716 | 1.3371 | 2.548 | 100.99 | 20.61 | 18.27 | 92.2994 | 106.3703 | 86.8924 | 0.004397 | 0.98057 | 0.023909 | 24.3829 | -0.10157 | 0.15551 | -0.001439 | -0.10774 | -0.10774 | -0.059322
2019-03-20 12:59 | 1553101153.1786 | 570 | 0.20926 | 0.000751 | 6.4e-05 | 108.7913 | 110.1943 | 121.8525 | 20.9475 | 0.00293 | 0.002764 | 129.2976 | 0.26156 | 8.4066 | 0.003638 | 0.005109 | 0.005007 | 5.2e-05 | 0.001094 | 0.29479 | 1.35 | 2.548 | 100.99 | 20.61 | 18.273 | 92.0027 | 109.1246 | 88.0103 | 0.004398 | 0.97753 | 0.023873 | 24.4218 | -0.11461 | 0.1556 | 0.00123 | 0.084795 | 0.084795 | 0.075419
2019-03-20 12:59 | 1553101154.1776 | 571 | 0.20889 | 0.000747 | 6.4e-05 | 108.6785 | 109.9517 | 121.8908 | 20.9485 | 0.002916 | 0.002719 | 129.2928 | 0.26007 | 8.4012 | 0.003548 | 0.004834 | 0.004753 | 5.1e-05 | 0.001095 | 0.30773 | 1.3599 | 2.548 | 100.99 | 20.61 | 18.276 | 91.6688 | 106.9424 | 86.9306 | 0.004399 | 0.98046 | 0.023913 | 24.3891 | -0.10418 | 0.15471 | 0.003478 | 0.24716 | 0.24716 | 0.025727
2019-03-20 12:59 | 1553101155.1787 | 572 | 0.20912 | 0.000747 | 6.4e-05 | 108.7984 | 110.0877 | 121.6378 | 20.9478 | 0.002942 | 0.002768 | 129.2987 | 0.26043 | 8.4015 | 0.003475 | 0.005192 | 0.005112 | 5.2e-05 | 0.001098 | 0.32764 | 1.3053 | 2.548 | 100.99 | 20.61 | 18.279 | 92.2907 | 105.1868 | 88.0996 | 0.004399 | 0.97729 | 0.023849 | 24.4033 | -0.10537 | 0.15493 | 0.002079 | 0.14577 | 0.14577 | 0.043682
2019-03-20 12:59 | 1553101156.1788 | 573 | 0.20905 | 0.000749 | 6.4e-05 | 108.7506 | 110.1079 | 121.3462 | 20.9488 | 0.002923 | 0.002776 | 129.3118 | 0.26094 | 8.4019 | 0.003483 | 0.004665 | 0.0045 | 5.1e-05 | 0.0011 | 0.31365 | 1.3068 | 2.548 | 100.99 | 20.61 | 18.282 | 91.8482 | 105.3665 | 88.3059 | 0.0044 | 0.97673 | 0.023814 | 24.381 | -0.11095 | 0.15524 | 0.00248 | 0.1744 | 0.1744 | 0.036584
2019-03-20 12:59 | 1553101157.1779 | 574 | 0.20912 | 0.00075 | 6.5e-05 | 108.769 | 109.9711 | 121.2157 | 20.9484 | 0.002963 | 0.002674 | 129.3083 | 0.26132 | 8.4008 | 0.003458 | 0.004885 | 0.004793 | 5.2e-05 | 0.001098 | 0.32731 | 1.3257 | 2.548 | 100.99 | 20.61 | 18.283 | 92.7959 | 104.7634 | 85.8384 | 0.0044 | 0.98345 | 0.023883 | 24.2846 | -0.098313 | 0.15546 | 0.002081 | 0.14612 | 0.14612 | 0.043726
2019-03-20 12:59 | 1553101158.1782 | 575 | 0.20902 | 0.000743 | 6.6e-05 | 108.7538 | 110.0078 | 121.0381 | 20.9478 | 0.002921 | 0.002723 | 129.3109 | 0.26092 | 8.4033 | 0.003469 | 0.004411 | 0.004283 | 5e-05 | 0.001098 | 0.31277 | 1.3103 | 2.548 | 100.99 | 20.61 | 18.285 | 91.7809 | 105.0427 | 87.0141 | 0.004401 | 0.98023 | 0.023826 | 24.3065 | -0.10254 | 0.15522 | 0.002687 | 0.18929 | 0.18929 | 0.033703
2019-03-20 12:59 | 1553101159.1771 | 576 | 0.20912 | 0.000743 | 6.6e-05 | 108.6655 | 109.9973 | 120.8726 | 20.9482 | 0.002955 | 0.002742 | 129.2993 | 0.26201 | 8.4024 | 0.003503 | 0.005492 | 0.005325 | 5.6e-05 | 0.001097 | 0.3126 | 1.3151 | 2.548 | 100.99 | 20.61 | 18.286 | 92.603 | 105.8686 | 87.4901 | 0.004401 | 0.97894 | 0.023794 | 24.3059 | -0.10895 | 0.15587 | 0.002097 | 0.14658 | 0.14658 | 0.043704
2019-03-20 12:59 | 1553101160.178 | 577 | 0.20906 | 0.000742 | 6.7e-05 | 108.5911 | 109.9158 | 120.9061 | 20.9478 | 0.002934 | 0.00266 | 129.2971 | 0.26175 | 8.4023 | 0.003405 | 0.004813 | 0.00467 | 5.2e-05 | 0.001095 | 0.28335 | 1.3238 | 2.5478 | 100.99 | 20.61 | 18.286 | 92.1068 | 103.4775 | 85.5107 | 0.004401 | 0.98434 | 0.023863 | 24.2425 | -0.10845 | 0.15572 | 0.002462 | 0.17326 | 0.17326 | 0.036938
2019-03-20 12:59 | 1553101161.1771 | 578 | 0.20903 | 0.000739 | 6.8e-05 | 108.6083 | 109.9364 | 120.939 | 20.9477 | 0.002892 | 0.002757 | 129.2751 | 0.26148 | 8.4032 | 0.003379 | 0.004743 | 0.004684 | 5.5e-05 | 0.001098 | 0.32117 | 1.3604 | 2.548 | 100.99 | 20.61 | 18.284 | 91.078 | 102.84 | 87.848 | 0.004401 | 0.97797 | 0.023789 | 24.3247 | -0.10871 | 0.15555 | 0.002631 | 0.18478 | 0.18478 | 0.0346
2019-03-20 12:59 | 1553101162.1781 | 579 | 0.20908 | 0.000736 | 6.9e-05 | 108.5824 | 109.8355 | 120.7897 | 20.9475 | 0.002892 | 0.002737 | 129.2812 | 0.26176 | 8.4027 | 0.003406 | 0.005294 | 0.005188 | 5.4e-05 | 0.001096 | 0.32333 | 1.3787 | 2.548 | 100.99 | 20.61 | 18.282 | 91.0879 | 103.5165 | 87.3481 | 0.0044 | 0.97933 | 0.023791 | 24.2928 | -0.10263 | 0.15572 | 0.002308 | 0.16156 | 0.16156 | 0.039614
2019-03-20 12:59 | 1553101163.1786 | 580 | 0.20908 | 0.000737 | 7e-05 | 108.584 | 109.7779 | 120.8071 | 20.9469 | 0.002966 | 0.002734 | 129.2717 | 0.26141 | 8.4058 | 0.003379 | 0.005027 | 0.004867 | 5.4e-05 | 0.001098 | 0.33428 | 1.3553 | 2.5479 | 100.99 | 20.61 | 18.274 | 92.8843 | 102.8465 | 87.2763 | 0.004398 | 0.97952 | 0.023795 | 24.2921 | -0.097809 | 0.15551 | 0.00234 | 0.16382 | 0.16382 | 0.039016
2019-03-20 12:59 | 1553101164.1772 | 581 | 0.20912 | 0.000734 | 7e-05 | 108.5892 | 109.9607 | 120.728 | 20.9476 | 0.002903 | 0.00267 | 129.271 | 0.26151 | 8.402 | 0.003478 | 0.005592 | 0.005465 | 5.4e-05 | 0.001098 | 0.31686 | 1.3357 | 2.5481 | 100.99 | 20.61 | 18.27 | 91.3562 | 105.2479 | 85.7486 | 0.004397 | 0.98369 | 0.023837 | 24.2327 | -0.11226 | 0.15557 | 0.002087 | 0.14595 | 0.14595 | 0.043811
2019-03-20 12:59 | 1553101165.1783 | 582 | 0.2091 | 0.000739 | 7.2e-05 | 108.575 | 109.9267 | 120.6125 | 20.9456 | 0.002907 | 0.002702 | 129.253 | 0.26072 | 8.4033 | 0.003471 | 0.005002 | 0.004932 | 5.5e-05 | 0.001095 | 0.32781 | 1.2968 | 2.548 | 100.99 | 20.61 | 18.266 | 91.4386 | 105.0879 | 86.5216 | 0.004396 | 0.98158 | 0.0238 | 24.2471 | -0.11067 | 0.1551 | 0.002187 | 0.15277 | 0.15277 | 0.041729
2019-03-20 12:59 | 1553101166.1773 | 583 | 0.20911 | 0.000729 | 7.2e-05 | 108.5934 | 109.9042 | 120.6327 | 20.9466 | 0.002835 | 0.002735 | 129.2132 | 0.26085 | 8.4029 | 0.003446 | 0.004822 | 0.004647 | 5.4e-05 | 0.001092 | 0.32045 | 1.3095 | 2.5479 | 100.99 | 20.61 | 18.255 | 89.7161 | 104.4731 | 87.2986 | 0.004394 | 0.97946 | 0.023777 | 24.2753 | -0.10732 | 0.15518 | 0.002181 | 0.15215 | 0.15215 | 0.04192
2019-03-20 12:59 | 1553101167.1783 | 584 | 0.20908 | 0.000729 | 7.3e-05 | 108.5168 | 109.7824 | 120.4975 | 20.9456 | 0.002888 | 0.002738 | 129.1987 | 0.26225 | 8.4022 | 0.003527 | 0.005026 | 0.004955 | 5.3e-05 | 0.001094 | 0.32261 | 1.3474 | 2.5479 | 100.99 | 20.61 | 18.251 | 90.9899 | 106.4404 | 87.3734 | 0.004393 | 0.97926 | 0.023761 | 24.2643 | -0.10372 | 0.15601 | 0.002303 | 0.16075 | 0.16075 | 0.039889
2019-03-20 12:59 | 1553101168.1779 | 585 | 0.20912 | 0.000728 | 7.3e-05 | 108.488 | 109.662 | 120.5482 | 20.9455 | 0.00289 | 0.002736 | 129.1841 | 0.26065 | 8.4026 | 0.003514 | 0.004657 | 0.00452 | 5.4e-05 | 0.001091 | 0.32493 | 1.3271 | 2.548 | 100.99 | 20.61 | 18.246 | 91.0288 | 106.1138 | 87.3399 | 0.004391 | 0.97935 | 0.023767 | 24.2682 | -0.096279 | 0.15506 | 0.002099 | 0.14615 | 0.14615 | 0.043607
2019-03-20 12:59 | 1553101169.1775 | 586 | 0.20912 | 0.00073 | 7.4e-05 | 108.5048 | 109.7559 | 120.6465 | 20.9466 | 0.002938 | 0.002729 | 129.1886 | 0.26164 | 8.4017 | 0.003528 | 0.004987 | 0.004766 | 5.4e-05 | 0.001094 | 0.3011 | 1.34 | 2.5478 | 100.99 | 20.61 | 18.235 | 92.1965 | 106.4722 | 87.1723 | 0.004389 | 0.97981 | 0.023782 | 24.2725 | -0.10254 | 0.15565 | 0.002064 | 0.14366 | 0.14366 | 0.044531
2019-03-20 12:59 | 1553101170.1786 | 587 | 0.20913 | 0.000725 | 7.4e-05 | 108.4372 | 109.6228 | 120.4426 | 20.9456 | 0.002835 | 0.002646 | 129.187 | 0.26053 | 8.4036 | 0.003549 | 0.004752 | 0.00464 | 5.4e-05 | 0.001101 | 0.34291 | 1.289 | 2.5479 | 100.99 | 20.61 | 18.23 | 89.7215 | 106.9655 | 85.1675 | 0.004388 | 0.98529 | 0.023829 | 24.1844 | -0.097265 | 0.15499 | 0.002012 | 0.1402 | 0.1402 | 0.045436
2019-03-20 12:59 | 1553101171.1777 | 588 | 0.20911 | 0.000719 | 7.6e-05 | 108.4498 | 109.6852 | 120.2203 | 20.9468 | 0.002927 | 0.002685 | 129.1761 | 0.26203 | 8.4039 | 0.003468 | 0.005139 | 0.005018 | 5.3e-05 | 0.001101 | 0.34147 | 1.2978 | 2.5479 | 100.99 | 20.61 | 18.22 | 91.926 | 105.0105 | 86.1022 | 0.004385 | 0.98272 | 0.023776 | 24.1936 | -0.10132 | 0.15588 | 0.002168 | 0.15102 | 0.15102 | 0.042421
2019-03-20 13:00 | 1553101172.1787 | 589 | 0.20911 | 0.000718 | 7.6e-05 | 108.4734 | 109.8257 | 120.6374 | 20.9456 | 0.002826 | 0.0027 | 129.1418 | 0.26089 | 8.4009 | 0.003504 | 0.00425 | 0.004175 | 5.4e-05 | 0.0011 | 0.3366 | 1.2657 | 2.5478 | 100.99 | 20.61 | 18.215 | 89.4863 | 105.8793 | 86.4624 | 0.004384 | 0.98174 | 0.023805 | 24.2476 | -0.11082 | 0.1552 | 0.002138 | 0.14902 | 0.14902 | 0.042806
2019-03-20 13:00 | 1553101173.1788 | 590 | 0.20914 | 0.000715 | 7.7e-05 | 108.4327 | 109.5484 | 120.4435 | 20.9468 | 0.002844 | 0.002677 | 129.142 | 0.26126 | 8.4063 | 0.003542 | 0.00523 | 0.005101 | 5.4e-05 | 0.001099 | 0.33826 | 1.3088 | 2.5481 | 100.99 | 20.61 | 18.213 | 89.9227 | 106.7951 | 85.9179 | 0.004384 | 0.98323 | 0.023804 | 24.2098 | -0.091568 | 0.15542 | 0.001984 | 0.13792 | 0.13792 | 0.046315
2019-03-20 13:00 | 1553101174.1778 | 591 | 0.20913 | 0.00072 | 7.7e-05 | 108.4252 | 109.5892 | 120.517 | 20.9461 | 0.002932 | 0.002713 | 129.1387 | 0.26083 | 8.4008 | 0.003541 | 0.004741 | 0.004599 | 5.4e-05 | 0.001096 | 0.33356 | 1.2684 | 2.548 | 100.99 | 20.6 | 18.212 | 92.0622 | 106.7723 | 86.7682 | 0.004381 | 0.9809 | 0.023783 | 24.2458 | -0.095507 | 0.15517 | 0.002065 | 0.14363 | 0.14363 | 0.044402
2019-03-20 13:00 | 1553101175.1769 | 592 | 0.20918 | 0.000713 | 7.8e-05 | 108.3971 | 109.7301 | 120.7456 | 20.9459 | 0.002895 | 0.002665 | 129.1208 | 0.26183 | 8.4026 | 0.003556 | 0.005441 | 0.005343 | 5.4e-05 | 0.001098 | 0.32532 | 1.275 | 2.548 | 100.99 | 20.6 | 18.215 | 91.1504 | 107.1454 | 85.6346 | 0.004381 | 0.984 | 0.023843 | 24.2306 | -0.10932 | 0.15576 | 0.001736 | 0.12026 | 0.12026 | 0.053235
2019-03-20 13:00 | 1553101176.1769 | 593 | 0.20916 | 0.000715 | 7.8e-05 | 108.3997 | 109.5013 | 120.8939 | 20.9447 | 0.002837 | 0.002707 | 129.1318 | 0.26156 | 8.4017 | 0.003565 | 0.005032 | 0.004937 | 5.2e-05 | 0.001089 | 0.30143 | 1.3211 | 2.2176 | 100.99 | 20.6 | 18.221 | 89.7582 | 107.3632 | 86.631 | 0.004383 | 0.98128 | 0.023825 | 24.2791 | -0.090444 | 0.1556 | 0.001862 | 0.12922 | 0.12922 | 0.049491
2019-03-20 13:00 | 1553101177.177 | 594 | 0.20913 | 0.000719 | 7.8e-05 | 108.2871 | 109.5081 | 121.1993 | 20.9453 | 0.002826 | 0.002705 | 129.1247 | 0.26191 | 8.4033 | 0.003549 | 0.005357 | 0.005218 | 5.2e-05 | 0.001088 | 0.30038 | 1.319 | 2.548 | 100.99 | 20.61 | 18.226 | 89.5046 | 106.9725 | 86.5761 | 0.004387 | 0.98143 | 0.023856 | 24.3079 | -0.10028 | 0.15581 | 0.002031 | 0.14151 | 0.14151 | 0.045255
2019-03-20 13:00 | 1553101178.1774 | 595 | 0.2091 | 0.000714 | 7.8e-05 | 108.2771 | 109.4774 | 121.1824 | 20.9455 | 0.002819 | 0.002653 | 129.1149 | 0.26123 | 8.4032 | 0.003507 | 0.005293 | 0.005155 | 5.6e-05 | 0.00109 | 0.33301 | 1.3423 | 2.5482 | 100.99 | 20.61 | 18.232 | 89.329 | 105.9581 | 85.3415 | 0.004388 | 0.98481 | 0.023896 | 24.2645 | -0.098609 | 0.15541 | 0.002242 | 0.15696 | 0.15696 | 0.040693
2019-03-20 13:00 | 1553101179.1781 | 596 | 0.20911 | 0.000715 | 7.8e-05 | 108.3214 | 109.4341 | 121 | 20.9459 | 0.002818 | 0.002648 | 129.092 | 0.26154 | 8.4007 | 0.003503 | 0.005156 | 0.005016 | 5.3e-05 | 0.001089 | 0.32117 | 1.4043 | 2.5481 | 100.99 | 20.6 | 18.244 | 89.3131 | 105.8545 | 85.2029 | 0.004388 | 0.98519 | 0.023882 | 24.2415 | -0.091407 | 0.15559 | 0.00219 | 0.15318 | 0.15318 | 0.041748
2019-03-20 13:00 | 1553101180.1781 | 597 | 0.20915 | 0.000713 | 7.8e-05 | 108.1898 | 109.3978 | 121.127 | 20.9457 | 0.002784 | 0.002696 | 129.0967 | 0.26136 | 8.4032 | 0.003546 | 0.005748 | 0.005647 | 5.5e-05 | 0.00109 | 0.32532 | 1.3494 | 2.548 | 100.99 | 20.6 | 18.252 | 88.4887 | 106.9073 | 86.3766 | 0.00439 | 0.98197 | 0.023856 | 24.2939 | -0.099319 | 0.15548 | 0.001941 | 0.13504 | 0.13504 | 0.047322
2019-03-20 13:00 | 1553101181.1782 | 598 | 0.20912 | 0.000706 | 7.8e-05 | 108.205 | 109.4677 | 121.277 | 20.9452 | 0.002875 | 0.002697 | 129.1178 | 0.26074 | 8.4006 | 0.003512 | 0.005222 | 0.005104 | 5.3e-05 | 0.001087 | 0.33948 | 1.3358 | 2.5481 | 100.99 | 20.61 | 18.268 | 90.6854 | 106.0764 | 86.3807 | 0.004397 | 0.98196 | 0.023871 | 24.3091 | -0.10378 | 0.15511 | 0.002122 | 0.14814 | 0.14814 | 0.043035
2019-03-20 13:00 | 1553101182.1776 | 599 | 0.20912 | 0.000709 | 7.9e-05 | 108.2212 | 109.4871 | 121.1561 | 20.9448 | 0.002824 | 0.002683 | 129.1409 | 0.26078 | 8.4039 | 0.003525 | 0.00521 | 0.005113 | 5.4e-05 | 0.001088 | 0.3501 | 1.3362 | 2.5481 | 100.99 | 20.61 | 18.277 | 89.4464 | 106.3807 | 86.0447 | 0.004399 | 0.98288 | 0.02387 | 24.2856 | -0.10403 | 0.15514 | 0.002129 | 0.14856 | 0.14856 | 0.04292
2019-03-20 13:00 | 1553101183.1783 | 600 | 0.20912 | 0.000707 | 7.9e-05 | 108.1519 | 109.2315 | 121.5103 | 20.9447 | 0.002808 | 0.002658 | 129.138 | 0.26127 | 8.4011 | 0.003444 | 0.005016 | 0.004849 | 5.3e-05 | 0.001091 | 0.31006 | 1.2924 | 2.548 | 100.99 | 20.61 | 18.285 | 89.0605 | 104.4271 | 85.4484 | 0.004401 | 0.98452 | 0.023925 | 24.3009 | -0.088839 | 0.15543 | 0.002126 | 0.14874 | 0.14874 | 0.042949
2019-03-20 13:00 | 1553101184.1784 | 601 | 0.20912 | 0.000708 | 8e-05 | 108.1736 | 109.3978 | 121.2913 | 20.945 | 0.002867 | 0.002631 | 129.1423 | 0.26083 | 8.3986 | 0.003417 | 0.004783 | 0.004606 | 5.4e-05 | 0.00109 | 0.33699 | 1.3379 | 2.548 | 100.99 | 20.61 | 18.302 | 90.4914 | 103.7609 | 84.8126 | 0.004405 | 0.98626 | 0.023924 | 24.2574 | -0.10066 | 0.15517 | 0.002091 | 0.14617 | 0.14617 | 0.043631
2019-03-20 13:00 | 1553101185.1774 | 602 | 0.20909 | 0.000703 | 7.9e-05 | 108.1128 | 109.2431 | 121.2244 | 20.9439 | 0.002882 | 0.002669 | 129.1656 | 0.26177 | 8.4048 | 0.003478 | 0.004711 | 0.004626 | 5.6e-05 | 0.001088 | 0.33107 | 1.3008 | 2.5479 | 100.99 | 20.61 | 18.31 | 90.85 | 105.2499 | 85.7145 | 0.004407 | 0.98379 | 0.023888 | 24.2813 | -0.093025 | 0.15573 | 0.002302 | 0.16117 | 0.16117 | 0.039712
2019-03-20 13:00 | 1553101186.1777 | 603 | 0.20911 | 0.000707 | 7.9e-05 | 108.0999 | 109.1679 | 121.2019 | 20.9447 | 0.002848 | 0.002678 | 129.1945 | 0.26045 | 8.4042 | 0.003582 | 0.005705 | 0.005436 | 5.5e-05 | 0.001084 | 0.29856 | 1.3154 | 2.548 | 100.99 | 20.61 | 18.318 | 90.0245 | 107.7618 | 85.9354 | 0.004409 | 0.98318 | 0.023878 | 24.2865 | -0.087942 | 0.15494 | 0.002176 | 0.152 | 0.152 | 0.041895
2019-03-20 13:00 | 1553101187.1775 | 604 | 0.20912 | 0.000705 | 8e-05 | 108.0649 | 109.289 | 120.8127 | 20.9442 | 0.00279 | 0.002671 | 129.2053 | 0.26156 | 8.4029 | 0.003488 | 0.005127 | 0.005008 | 5.7e-05 | 0.001082 | 0.32516 | 1.3157 | 2.5479 | 100.99 | 20.61 | 18.332 | 88.6226 | 105.4948 | 85.7648 | 0.004412 | 0.98365 | 0.023845 | 24.2417 | -0.10075 | 0.1556 | 0.002141 | 0.14926 | 0.14926 | 0.042845
2019-03-20 13:00 | 1553101188.1786 | 605 | 0.2091 | 0.000706 | 8e-05 | 107.9928 | 109.227 | 120.8399 | 20.9443 | 0.002829 | 0.002683 | 129.2167 | 0.26115 | 8.4034 | 0.003624 | 0.005182 | 0.005137 | 5.5e-05 | 0.001081 | 0.35297 | 1.2944 | 2.548 | 100.99 | 20.61 | 18.339 | 89.5656 | 108.8015 | 86.0529 | 0.004414 | 0.98286 | 0.023838 | 24.2542 | -0.10164 | 0.15536 | 0.002267 | 0.15832 | 0.15832 | 0.040331
2019-03-20 13:00 | 1553101189.1776 | 606 | 0.20911 | 0.000707 | 8e-05 | 108.0148 | 109.2236 | 121.1064 | 20.9441 | 0.002754 | 0.002676 | 129.2279 | 0.26224 | 8.4044 | 0.003589 | 0.00492 | 0.004724 | 5.5e-05 | 0.001078 | 0.3147 | 1.3171 | 2.548 | 100.99 | 20.6 | 18.352 | 87.7561 | 107.9447 | 85.891 | 0.004414 | 0.9833 | 0.02387 | 24.2754 | -0.099545 | 0.15601 | 0.002203 | 0.15391 | 0.15391 | 0.041661
2019-03-20 13:00 | 1553101190.1765 | 607 | 0.20909 | 0.000704 | 7.9e-05 | 108.0083 | 109.0349 | 120.8348 | 20.9443 | 0.002883 | 0.002711 | 129.2153 | 0.26176 | 8.4016 | 0.003575 | 0.004907 | 0.00479 | 5.6e-05 | 0.001078 | 0.30757 | 1.2848 | 2.548 | 100.99 | 20.6 | 18.359 | 90.8752 | 107.5897 | 86.7234 | 0.004416 | 0.98103 | 0.023816 | 24.2763 | -0.08461 | 0.15572 | 0.002275 | 0.15878 | 0.15878 | 0.040309
2019-03-20 13:00 | 1553101191.1769 | 608 | 0.20912 | 0.000703 | 8e-05 | 107.9317 | 109.2453 | 120.7925 | 20.9439 | 0.002811 | 0.002679 | 129.2133 | 0.26086 | 8.4025 | 0.003595 | 0.0056 | 0.005461 | 5.8e-05 | 0.001078 | 0.30469 | 1.3195 | 2.5481 | 100.99 | 20.6 | 18.365 | 89.1447 | 108.0863 | 85.9488 | 0.004417 | 0.98314 | 0.023837 | 24.2459 | -0.1082 | 0.15519 | 0.00211 | 0.14697 | 0.14697 | 0.043396
2019-03-20 13:00 | 1553101192.1778 | 609 | 0.20913 | 0.000699 | 7.8e-05 | 107.9207 | 109.2819 | 120.8256 | 20.9442 | 0.002795 | 0.00272 | 129.203 | 0.26038 | 8.4041 | 0.003599 | 0.005031 | 0.004919 | 5.8e-05 | 0.001077 | 0.31094 | 1.2993 | 2.548 | 100.99 | 20.61 | 18.379 | 88.749 | 108.1741 | 86.9358 | 0.004424 | 0.98045 | 0.023808 | 24.2825 | -0.11212 | 0.1549 | 0.002043 | 0.14209 | 0.14209 | 0.044804
2019-03-20 13:00 | 1553101193.1778 | 610 | 0.20912 | 0.000703 | 7.9e-05 | 108.008 | 109.1063 | 120.624 | 20.9428 | 0.002809 | 0.002672 | 129.2181 | 0.26158 | 8.4012 | 0.003602 | 0.005204 | 0.005112 | 5.8e-05 | 0.00108 | 0.32731 | 1.3232 | 2.548 | 100.99 | 20.61 | 18.384 | 89.0938 | 108.2506 | 85.7938 | 0.004425 | 0.98357 | 0.023826 | 24.2237 | -0.090491 | 0.15561 | 0.0021 | 0.1463 | 0.1463 | 0.043716
2019-03-20 13:00 | 1553101194.1768 | 611 | 0.20913 | 0.000697 | 7.9e-05 | 107.8952 | 109.03 | 120.7716 | 20.9437 | 0.002867 | 0.002655 | 129.2124 | 0.26154 | 8.4048 | 0.003609 | 0.004835 | 0.004636 | 5.8e-05 | 0.001071 | 0.31045 | 1.3577 | 2.548 | 100.99 | 20.6 | 18.395 | 90.4776 | 108.4162 | 85.3898 | 0.004425 | 0.98468 | 0.023854 | 24.2249 | -0.093591 | 0.15559 | 0.002094 | 0.14604 | 0.14604 | 0.043787
2019-03-20 13:00 | 1553101195.1772 | 612 | 0.20911 | 0.000701 | 7.9e-05 | 107.942 | 109.2685 | 120.789 | 20.9437 | 0.002763 | 0.002633 | 129.1935 | 0.26111 | 8.4026 | 0.003515 | 0.005211 | 0.005111 | 5.9e-05 | 0.001064 | 0.30198 | 1.3536 | 2.548 | 100.99 | 20.6 | 18.399 | 87.9752 | 106.1387 | 84.8606 | 0.004426 | 0.98613 | 0.023873 | 24.2088 | -0.10924 | 0.15533 | 0.002161 | 0.15094 | 0.15094 | 0.042296
2019-03-20 13:00 | 1553101196.177 | 613 | 0.20916 | 0.000701 | 7.9e-05 | 107.8644 | 108.8461 | 120.6846 | 20.9439 | 0.002789 | 0.002627 | 129.1743 | 0.26226 | 8.4026 | 0.003609 | 0.00543 | 0.005359 | 5.6e-05 | 0.001063 | 0.31598 | 1.3596 | 2.5482 | 100.99 | 20.61 | 18.403 | 88.6066 | 108.4243 | 84.7078 | 0.004429 | 0.98655 | 0.023868 | 24.1931 | -0.08104 | 0.15602 | 0.00188 | 0.13068 | 0.13068 | 0.049071
2019-03-20 13:00 | 1553101197.177 | 614 | 0.20915 | 0.0007 | 7.9e-05 | 107.8835 | 108.8686 | 120.773 | 20.9441 | 0.002799 | 0.002708 | 129.1833 | 0.26139 | 8.4036 | 0.003566 | 0.004763 | 0.004655 | 5.6e-05 | 0.001062 | 0.30685 | 1.3668 | 2.5481 | 100.99 | 20.61 | 18.406 | 88.8442 | 107.3878 | 86.6554 | 0.00443 | 0.98121 | 0.023812 | 24.2678 | -0.0813 | 0.1555 | 0.00197 | 0.13679 | 0.13679 | 0.046722
2019-03-20 13:00 | 1553101198.1781 | 615 | 0.20917 | 0.000701 | 7.9e-05 | 107.8822 | 109.2091 | 120.7073 | 20.9422 | 0.002765 | 0.00261 | 129.1708 | 0.26061 | 8.4021 | 0.003661 | 0.005417 | 0.005327 | 5.7e-05 | 0.001061 | 0.33931 | 1.3784 | 2.5481 | 100.99 | 20.6 | 18.405 | 88.019 | 109.698 | 84.3004 | 0.004427 | 0.98768 | 0.023884 | 24.1817 | -0.10934 | 0.15504 | 0.001813 | 0.12593 | 0.12593 | 0.050599
2019-03-20 13:00 | 1553101199.1772 | 616 | 0.20916 | 0.000697 | 7.9e-05 | 107.769 | 108.9803 | 120.8449 | 20.943 | 0.002791 | 0.002601 | 129.189 | 0.2605 | 8.403 | 0.003617 | 0.00502 | 0.004894 | 5.9e-05 | 0.00106 | 0.29717 | 1.3951 | 2.5481 | 100.99 | 20.61 | 18.399 | 88.6455 | 108.6093 | 84.0906 | 0.004428 | 0.98826 | 0.023904 | 24.1883 | -0.099974 | 0.15497 | 0.001895 | 0.13193 | 0.13193 | 0.048279
2019-03-20 13:00 | 1553101200.1782 | 617 | 0.20917 | 0.000691 | 7.9e-05 | 107.8625 | 108.9299 | 121.0991 | 20.9424 | 0.002731 | 0.002645 | 129.1937 | 0.26128 | 8.4002 | 0.00353 | 0.004491 | 0.004344 | 6.1e-05 | 0.001057 | 0.33052 | 1.3495 | 2.548 | 100.99 | 20.61 | 18.395 | 87.2095 | 106.4964 | 85.1441 | 0.004427 | 0.98535 | 0.023894 | 24.2494 | -0.088078 | 0.15544 | 0.001859 | 0.1293 | 0.1293 | 0.049408
2019-03-20 13:00 | 1553101201.1782 | 618 | 0.20923 | 0.000693 | 7.8e-05 | 107.8822 | 109.0696 | 121.2453 | 20.9438 | 0.00281 | 0.002603 | 129.189 | 0.26137 | 8.4048 | 0.003552 | 0.005044 | 0.004939 | 5.7e-05 | 0.001058 | 0.33588 | 1.3432 | 2.548 | 100.99 | 20.61 | 18.389 | 89.1212 | 107.0331 | 84.1334 | 0.004426 | 0.98814 | 0.023942 | 24.2298 | -0.097913 | 0.15549 | 0.001494 | 0.10322 | 0.10322 | 0.061911
2019-03-20 13:00 | 1553101202.1786 | 619 | 0.20915 | 0.000692 | 7.9e-05 | 107.7784 | 109.1025 | 121.2263 | 20.9427 | 0.002822 | 0.002652 | 129.182 | 0.26117 | 8.4041 | 0.00352 | 0.005063 | 0.004887 | 5.7e-05 | 0.001062 | 0.35297 | 1.3077 | 2.548 | 100.99 | 20.61 | 18.377 | 89.4039 | 106.2748 | 85.3092 | 0.004423 | 0.9849 | 0.023901 | 24.2678 | -0.10922 | 0.15537 | 0.001946 | 0.13559 | 0.13559 | 0.047098
2019-03-20 13:00 | 1553101203.1774 | 620 | 0.20916 | 0.000693 | 7.9e-05 | 107.8647 | 108.9545 | 121.408 | 20.9424 | 0.002768 | 0.002613 | 129.1693 | 0.261 | 8.4013 | 0.003464 | 0.005128 | 0.004949 | 5.6e-05 | 0.001064 | 0.32278 | 1.3073 | 2.5481 | 100.99 | 20.61 | 18.37 | 88.0943 | 104.9158 | 84.3779 | 0.004421 | 0.98746 | 0.02395 | 24.2544 | -0.089915 | 0.15527 | 0.001915 | 0.13364 | 0.13364 | 0.04775
2019-03-20 13:00 | 1553101204.1774 | 621 | 0.20917 | 0.00069 | 7.9e-05 | 107.8602 | 109.0618 | 121.5141 | 20.9427 | 0.002781 | 0.002643 | 129.1562 | 0.26035 | 8.4028 | 0.003476 | 0.005233 | 0.005208 | 5.6e-05 | 0.001065 | 0.30502 | 1.3061 | 2.5481 | 100.99 | 20.6 | 18.362 | 88.4008 | 105.1975 | 85.0818 | 0.004417 | 0.98552 | 0.023937 | 24.2888 | -0.099092 | 0.15488 | 0.001864 | 0.12991 | 0.12991 | 0.049001
2019-03-20 13:00 | 1553101205.1768 | 622 | 0.20918 | 0.00069 | 7.9e-05 | 107.8285 | 108.7994 | 121.3197 | 20.9421 | 0.002779 | 0.002597 | 129.1619 | 0.26057 | 8.4046 | 0.003544 | 0.004982 | 0.00482 | 5.7e-05 | 0.001069 | 0.32587 | 1.3416 | 2.5481 | 100.99 | 20.6 | 18.344 | 88.3673 | 106.8434 | 83.9743 | 0.004412 | 0.98858 | 0.023955 | 24.2319 | -0.080177 | 0.15501 | 0.001777 | 0.12368 | 0.12368 | 0.051513
2019-03-20 13:00 | 1553101206.1775 | 623 | 0.20912 | 0.000686 | 7.9e-05 | 107.8525 | 109.042 | 121.1364 | 20.9427 | 0.002781 | 0.002651 | 129.1671 | 0.2605 | 8.4051 | 0.003482 | 0.005069 | 0.004911 | 5.6e-05 | 0.001068 | 0.33107 | 1.3453 | 2.5481 | 100.99 | 20.6 | 18.335 | 88.407 | 105.3321 | 85.2803 | 0.00441 | 0.98498 | 0.023893 | 24.2578 | -0.098111 | 0.15497 | 0.002149 | 0.15025 | 0.15025 | 0.042392
2019-03-20 13:00 | 1553101207.1776 | 624 | 0.20913 | 0.000692 | 7.9e-05 | 107.7926 | 109.0782 | 121.2704 | 20.942 | 0.002789 | 0.002642 | 129.1663 | 0.26009 | 8.4005 | 0.003455 | 0.004572 | 0.004491 | 5.5e-05 | 0.001074 | 0.35225 | 1.3566 | 2.5482 | 100.99 | 20.61 | 18.316 | 88.603 | 104.6778 | 85.0602 | 0.004408 | 0.98558 | 0.023914 | 24.2637 | -0.10605 | 0.15473 | 0.002059 | 0.14388 | 0.14388 | 0.0442
2019-03-20 13:00 | 1553101208.177 | 625 | 0.20912 | 0.000689 | 7.9e-05 | 107.7926 | 108.9033 | 121.3258 | 20.9418 | 0.002773 | 0.002581 | 129.1677 | 0.26117 | 8.4022 | 0.003452 | 0.005229 | 0.005053 | 5.1e-05 | 0.001077 | 0.30165 | 1.3703 | 2.5482 | 100.99 | 20.61 | 18.308 | 88.2139 | 104.6217 | 83.5939 | 0.004406 | 0.98963 | 0.023968 | 24.2196 | -0.091693 | 0.15537 | 0.002131 | 0.14937 | 0.14937 | 0.042751
2019-03-20 13:00 | 1553101209.1768 | 626 | 0.20914 | 0.000685 | 7.8e-05 | 107.8353 | 108.9542 | 121.0902 | 20.9423 | 0.002771 | 0.002662 | 129.1388 | 0.26114 | 8.4023 | 0.003475 | 0.005422 | 0.005305 | 5.5e-05 | 0.001077 | 0.33356 | 1.3541 | 2.5481 | 100.99 | 20.61 | 18.3 | 88.17 | 105.1617 | 85.5491 | 0.004404 | 0.98424 | 0.02388 | 24.2622 | -0.092325 | 0.15535 | 0.002037 | 0.14205 | 0.14205 | 0.044949
2019-03-20 13:00 | 1553101210.1767 | 627 | 0.20916 | 0.000687 | 8e-05 | 107.9139 | 108.7834 | 120.7608 | 20.9416 | 0.002768 | 0.002615 | 129.1227 | 0.261 | 8.4029 | 0.003463 | 0.004793 | 0.004717 | 5.3e-05 | 0.00108 | 0.29446 | 1.3191 | 2.5481 | 100.99 | 20.61 | 18.288 | 88.099 | 104.8747 | 84.4153 | 0.004402 | 0.98736 | 0.023885 | 24.1909 | -0.071776 | 0.15527 | 0.001907 | 0.13265 | 0.13265 | 0.048107
2019-03-20 13:00 | 1553101211.1777 | 628 | 0.20917 | 0.000687 | 7.9e-05 | 107.8175 | 108.9568 | 120.6836 | 20.9416 | 0.002751 | 0.002608 | 129.1264 | 0.26168 | 8.4009 | 0.003424 | 0.004987 | 0.004863 | 5.3e-05 | 0.00108 | 0.34794 | 1.3494 | 2.5482 | 100.99 | 20.61 | 18.288 | 87.6928 | 103.9439 | 84.2453 | 0.004402 | 0.98783 | 0.023883 | 24.1774 | -0.094016 | 0.15567 | 0.001827 | 0.12692 | 0.12692 | 0.050411
2019-03-20 13:00 | 1553101212.1779 | 629 | 0.20917 | 0.000683 | 7.9e-05 | 107.8486 | 109.1526 | 120.7951 | 20.9415 | 0.00278 | 0.002607 | 129.1324 | 0.26123 | 8.4028 | 0.003434 | 0.004402 | 0.004306 | 5.4e-05 | 0.001083 | 0.3027 | 1.3463 | 2.548 | 100.99 | 20.61 | 18.278 | 88.3917 | 104.1677 | 84.2317 | 0.004399 | 0.98787 | 0.023895 | 24.1881 | -0.1075 | 0.15541 | 0.001853 | 0.12892 | 0.12892 | 0.049543
2019-03-20 13:00 | 1553101213.1778 | 630 | 0.20916 | 0.000688 | 7.9e-05 | 107.7923 | 108.9792 | 120.7029 | 20.9418 | 0.002783 | 0.002582 | 129.1232 | 0.26109 | 8.4057 | 0.003455 | 0.004975 | 0.004829 | 5.6e-05 | 0.001087 | 0.33804 | 1.3386 | 2.5481 | 100.99 | 20.61 | 18.273 | 88.4577 | 104.6902 | 83.6181 | 0.004398 | 0.98957 | 0.023906 | 24.1581 | -0.09795 | 0.15532 | 0.001929 | 0.13443 | 0.13443 | 0.047487
2019-03-20 13:00 | 1553101214.1769 | 631 | 0.20927 | 0.000683 | 7.9e-05 | 107.9213 | 109.159 | 120.5935 | 20.9414 | 0.002721 | 0.002578 | 129.1246 | 0.26123 | 8.4016 | 0.003483 | 0.004784 | 0.00464 | 5.6e-05 | 0.001087 | 0.32228 | 1.3762 | 2.548 | 100.99 | 20.61 | 18.264 | 86.9759 | 105.374 | 83.5287 | 0.004396 | 0.98981 | 0.023898 | 24.1441 | -0.10199 | 0.15541 | 0.001263 | 0.086327 | 0.086327 | 0.073989
2019-03-20 13:00 | 1553101215.1774 | 632 | 0.20916 | 0.000684 | 7.8e-05 | 107.8204 | 109.0039 | 120.8766 | 20.941 | 0.002771 | 0.00259 | 129.1338 | 0.26047 | 8.4035 | 0.003508 | 0.004668 | 0.0046 | 6e-05 | 0.001091 | 0.29374 | 1.3536 | 2.5481 | 100.99 | 20.61 | 18.259 | 88.1811 | 105.9765 | 83.8131 | 0.004395 | 0.98903 | 0.023917 | 24.1821 | -0.09764 | 0.15495 | 0.001882 | 0.13114 | 0.13114 | 0.048562
2019-03-20 13:00 | 1553101216.1772 | 633 | 0.20917 | 0.000685 | 7.9e-05 | 107.8376 | 109.187 | 121.0467 | 20.9417 | 0.002699 | 0.002571 | 129.1117 | 0.26113 | 8.4021 | 0.003527 | 0.0048 | 0.004682 | 5.6e-05 | 0.00109 | 0.33157 | 1.3477 | 2.5481 | 100.99 | 20.61 | 18.252 | 86.4344 | 106.4352 | 83.3584 | 0.004393 | 0.99029 | 0.023949 | 24.1837 | -0.11123 | 0.15535 | 0.001855 | 0.12926 | 0.12926 | 0.049393
2019-03-20 13:00 | 1553101217.1781 | 634 | 0.20917 | 0.000677 | 7.9e-05 | 107.8618 | 109.128 | 120.9897 | 20.9409 | 0.002745 | 0.002579 | 129.0873 | 0.26106 | 8.4038 | 0.003446 | 0.005261 | 0.005162 | 5.4e-05 | 0.001088 | 0.31542 | 1.3405 | 2.548 | 100.99 | 20.61 | 18.25 | 87.5446 | 104.4659 | 83.5579 | 0.004392 | 0.98973 | 0.023936 | 24.1848 | -0.10438 | 0.1553 | 0.001883 | 0.13128 | 0.13128 | 0.048621
2019-03-20 13:00 | 1553101218.1781 | 635 | 0.20917 | 0.000685 | 8e-05 | 107.9181 | 109.0872 | 120.8883 | 20.941 | 0.002772 | 0.002591 | 129.0945 | 0.26048 | 8.402 | 0.003491 | 0.005293 | 0.005116 | 5.4e-05 | 0.001091 | 0.31509 | 1.3556 | 2.548 | 100.99 | 20.61 | 18.248 | 88.1866 | 105.5688 | 83.8304 | 0.004392 | 0.98898 | 0.023917 | 24.1839 | -0.096377 | 0.15496 | 0.001861 | 0.12953 | 0.12953 | 0.049168
2019-03-20 13:00 | 1553101219.1772 | 636 | 0.20913 | 0.00068 | 7.9e-05 | 107.8926 | 109.2804 | 121.1711 | 20.9406 | 0.002747 | 0.002579 | 129.0785 | 0.26128 | 8.4022 | 0.003493 | 0.005199 | 0.005087 | 5.6e-05 | 0.001091 | 0.32388 | 1.327 | 2.5479 | 100.99 | 20.61 | 18.244 | 87.5857 | 105.6014 | 83.5516 | 0.004391 | 0.98975 | 0.023955 | 24.2027 | -0.11432 | 0.15543 | 0.002108 | 0.14758 | 0.14758 | 0.043288
2019-03-20 13:00 | 1553101220.1773 | 637 | 0.20913 | 0.000684 | 7.9e-05 | 107.8408 | 109.0446 | 121.5861 | 20.9413 | 0.002735 | 0.002619 | 129.075 | 0.26212 | 8.4017 | 0.00352 | 0.004798 | 0.004674 | 5.3e-05 | 0.00109 | 0.31758 | 1.3128 | 2.5481 | 100.99 | 20.61 | 18.244 | 87.2913 | 106.2695 | 84.5148 | 0.004391 | 0.98708 | 0.023963 | 24.2768 | -0.099292 | 0.15593 | 0.002104 | 0.14737 | 0.14737 | 0.043487
2019-03-20 13:00 | 1553101221.1764 | 638 | 0.20916 | 0.000677 | 8e-05 | 107.9152 | 109.1765 | 121.4387 | 20.9406 | 0.00269 | 0.002578 | 129.0584 | 0.26105 | 8.4013 | 0.003456 | 0.004845 | 0.004704 | 5.6e-05 | 0.001083 | 0.30845 | 1.3149 | 2.548 | 100.99 | 20.61 | 18.243 | 86.2172 | 104.7196 | 83.5366 | 0.004391 | 0.98979 | 0.023982 | 24.2289 | -0.10394 | 0.1553 | 0.001922 | 0.13429 | 0.13429 | 0.04753
2019-03-20 13:00 | 1553101222.1774 | 639 | 0.20913 | 0.000676 | 8e-05 | 107.8929 | 109.1881 | 121.0519 | 20.9408 | 0.002726 | 0.002627 | 129.0564 | 0.2612 | 8.4024 | 0.003413 | 0.005204 | 0.005081 | 5.4e-05 | 0.00108 | 0.33804 | 1.3612 | 2.548 | 100.99 | 20.61 | 18.242 | 87.0807 | 103.6632 | 84.7079 | 0.004391 | 0.98655 | 0.023904 | 24.2299 | -0.10674 | 0.15539 | 0.002112 | 0.14754 | 0.14754 | 0.043286
2019-03-20 13:00 | 1553101223.1778 | 640 | 0.20914 | 0.000681 | 8e-05 | 107.9113 | 109.1238 | 121.0392 | 20.9413 | 0.002735 | 0.002618 | 129.052 | 0.26123 | 8.4042 | 0.003472 | 0.005022 | 0.004945 | 5.3e-05 | 0.001077 | 0.34092 | 1.3793 | 2.548 | 100.99 | 20.61 | 18.243 | 87.2913 | 105.1042 | 84.4867 | 0.004391 | 0.98716 | 0.02391 | 24.2212 | -0.099943 | 0.15541 | 0.002054 | 0.14339 | 0.14339 | 0.044544
2019-03-20 13:00 | 1553101224.1775 | 641 | 0.20915 | 0.000677 | 8e-05 | 107.9647 | 109.0801 | 121.0965 | 20.9407 | 0.00276 | 0.002562 | 129.0576 | 0.26126 | 8.4035 | 0.003582 | 0.004807 | 0.004684 | 5.7e-05 | 0.001071 | 0.31022 | 1.339 | 2.5479 | 100.99 | 20.61 | 18.244 | 87.9013 | 107.7775 | 83.1381 | 0.004391 | 0.9909 | 0.023961 | 24.1812 | -0.091936 | 0.15542 | 0.001972 | 0.13774 | 0.13774 | 0.046375
2019-03-20 13:00 | 1553101225.1776 | 642 | 0.20914 | 0.000683 | 8e-05 | 107.9893 | 109.1586 | 121.1014 | 20.9403 | 0.002757 | 0.002572 | 129.0425 | 0.26103 | 8.402 | 0.003566 | 0.005062 | 0.004861 | 5.5e-05 | 0.001073 | 0.3115 | 1.3087 | 2.548 | 100.99 | 20.61 | 18.245 | 87.8331 | 107.3872 | 83.3849 | 0.004391 | 0.99021 | 0.023953 | 24.1901 | -0.096333 | 0.15529 | 0.002035 | 0.14232 | 0.14232 | 0.044845
2019-03-20 13:00 | 1553101226.1769 | 643 | 0.20911 | 0.000683 | 8e-05 | 107.8932 | 109.1489 | 120.947 | 20.9408 | 0.002673 | 0.002533 | 129.0615 | 0.26124 | 8.4048 | 0.003593 | 0.005056 | 0.004946 | 5.7e-05 | 0.001079 | 0.33395 | 1.29 | 2.548 | 100.99 | 20.61 | 18.246 | 85.8075 | 108.025 | 82.4494 | 0.004391 | 0.99282 | 0.023969 | 24.1429 | -0.1035 | 0.15541 | 0.00224 | 0.15724 | 0.15724 | 0.040624
2019-03-20 13:00 | 1553101227.1766 | 644 | 0.20918 | 0.000685 | 7.9e-05 | 107.9611 | 109.119 | 121.1456 | 20.9411 | 0.002765 | 0.002565 | 129.0555 | 0.26051 | 8.4035 | 0.003655 | 0.004987 | 0.004951 | 6e-05 | 0.001084 | 0.30845 | 1.3118 | 2.5478 | 100.99 | 20.61 | 18.25 | 88.0367 | 109.5472 | 83.2164 | 0.004392 | 0.99068 | 0.023963 | 24.1887 | -0.095417 | 0.15498 | 0.001782 | 0.12409 | 0.12409 | 0.051329
2019-03-20 13:00 | 1553101228.177 | 645 | 0.20916 | 0.000681 | 7.8e-05 | 107.9184 | 109.1317 | 121.2061 | 20.9417 | 0.002766 | 0.002622 | 129.0379 | 0.26092 | 8.4024 | 0.003602 | 0.005266 | 0.005159 | 5.7e-05 | 0.001083 | 0.31526 | 1.3097 | 2.5478 | 100.99 | 20.61 | 18.251 | 88.0537 | 108.2622 | 84.5878 | 0.004393 | 0.98688 | 0.023923 | 24.2413 | -0.099996 | 0.15522 | 0.00193 | 0.13466 | 0.13466 | 0.047376
2019-03-20 13:00 | 1553101229.1777 | 646 | 0.20917 | 0.000679 | 7.8e-05 | 107.9592 | 109.1728 | 120.9395 | 20.9401 | 0.002737 | 0.002607 | 129.0688 | 0.2617 | 8.4011 | 0.003644 | 0.005367 | 0.00527 | 5.5e-05 | 0.001084 | 0.31205 | 1.3183 | 2.548 | 100.99 | 20.61 | 18.254 | 87.3542 | 109.2796 | 84.2282 | 0.004393 | 0.98788 | 0.023909 | 24.2025 | -0.099989 | 0.15568 | 0.001857 | 0.12926 | 0.12926 | 0.049501
2019-03-20 13:00 | 1553101230.1768 | 647 | 0.2092 | 0.000682 | 7.7e-05 | 108.018 | 108.9852 | 120.5541 | 20.9404 | 0.002704 | 0.002615 | 129.0789 | 0.26093 | 8.403 | 0.003611 | 0.00431 | 0.004171 | 5.5e-05 | 0.001083 | 0.31725 | 1.3124 | 2.5479 | 100.99 | 20.61 | 18.256 | 86.5611 | 108.4793 | 84.4267 | 0.004394 | 0.98733 | 0.023864 | 24.1706 | -0.079728 | 0.15523 | 0.001669 | 0.11558 | 0.11558 | 0.0552
2019-03-20 13:00 | 1553101231.1772 | 648 | 0.20923 | 0.000679 | 7.7e-05 | 107.9676 | 109.2576 | 120.659 | 20.9414 | 0.002721 | 0.002574 | 129.0689 | 0.26096 | 8.4042 | 0.003601 | 0.004545 | 0.004314 | 5.6e-05 | 0.001088 | 0.37161 | 1.3069 | 2.5479 | 100.99 | 20.61 | 18.257 | 86.9657 | 108.2271 | 83.4398 | 0.004394 | 0.99006 | 0.023908 | 24.1477 | -0.10624 | 0.15524 | 0.001492 | 0.10303 | 0.10303 | 0.06193
2019-03-20 13:01 | 1553101232.178 | 649 | 0.20911 | 0.000679 | 7.6e-05 | 108.0672 | 109.3395 | 120.8674 | 20.9405 | 0.002712 | 0.002554 | 129.0702 | 0.26152 | 8.4031 | 0.003617 | 0.005348 | 0.005239 | 5.4e-05 | 0.001087 | 0.33483 | 1.2944 | 2.548 | 100.99 | 20.61 | 18.257 | 86.7416 | 108.6131 | 82.9412 | 0.004394 | 0.99145 | 0.023945 | 24.1516 | -0.10469 | 0.15558 | 0.002199 | 0.15427 | 0.15427 | 0.041449
2019-03-20 13:01 | 1553101233.178 | 650 | 0.20913 | 0.000678 | 7.7e-05 | 107.9634 | 109.2745 | 120.8871 | 20.9405 | 0.002753 | 0.002592 | 129.0744 | 0.26147 | 8.4021 | 0.003586 | 0.005258 | 0.005198 | 5.8e-05 | 0.001089 | 0.31564 | 1.3186 | 2.5481 | 100.99 | 20.61 | 18.258 | 87.7403 | 107.8584 | 83.858 | 0.004394 | 0.9889 | 0.023916 | 24.1847 | -0.10797 | 0.15555 | 0.002119 | 0.14832 | 0.14832 | 0.043102
2019-03-20 13:01 | 1553101234.178 | 651 | 0.20919 | 0.000677 | 7.8e-05 | 108.0918 | 109.1579 | 120.8883 | 20.9397 | 0.002703 | 0.002575 | 129.0609 | 0.26087 | 8.4017 | 0.003628 | 0.004997 | 0.004881 | 5.7e-05 | 0.00109 | 0.37056 | 1.365 | 2.548 | 100.99 | 20.61 | 18.258 | 86.52 | 108.8718 | 83.4601 | 0.004394 | 0.99 | 0.02393 | 24.1713 | -0.087784 | 0.15519 | 0.001754 | 0.12197 | 0.12197 | 0.052295
2019-03-20 13:01 | 1553101235.1781 | 652 | 0.20916 | 0.000671 | 7.7e-05 | 108.04 | 109.4374 | 120.654 | 20.9401 | 0.002789 | 0.00256 | 129.0398 | 0.26098 | 8.4035 | 0.003617 | 0.004431 | 0.004297 | 6.2e-05 | 0.001094 | 0.34562 | 1.3871 | 2.5481 | 100.99 | 20.61 | 18.257 | 88.5911 | 108.6183 | 83.093 | 0.004394 | 0.99102 | 0.023919 | 24.1354 | -0.11495 | 0.15526 | 0.001943 | 0.13559 | 0.13559 | 0.04706
2019-03-20 13:01 | 1553101236.1765 | 653 | 0.20917 | 0.00068 | 7.8e-05 | 108.0459 | 109.4449 | 120.9949 | 20.9407 | 0.002738 | 0.002628 | 129.0091 | 0.26091 | 8.4019 | 0.003606 | 0.005101 | 0.004992 | 5.6e-05 | 0.001091 | 0.34468 | 1.363 | 2.5481 | 100.99 | 20.61 | 18.256 | 87.3836 | 108.3493 | 84.7266 | 0.004394 | 0.9865 | 0.023898 | 24.2249 | -0.11507 | 0.15521 | 0.001887 | 0.13139 | 0.13139 | 0.048553
2019-03-20 13:01 | 1553101237.1762 | 654 | 0.20917 | 0.000675 | 7.8e-05 | 108.0785 | 109.2528 | 120.9653 | 20.9395 | 0.002744 | 0.002524 | 129.0197 | 0.26132 | 8.4012 | 0.003651 | 0.005012 | 0.004896 | 5.9e-05 | 0.001093 | 0.36713 | 1.3901 | 2.5482 | 100.99 | 20.61 | 18.253 | 87.53 | 109.4487 | 82.2236 | 0.004393 | 0.99345 | 0.023979 | 24.1371 | -0.096657 | 0.15546 | 0.001884 | 0.13163 | 0.13163 | 0.04854
2019-03-20 13:01 | 1553101238.1767 | 655 | 0.20913 | 0.000675 | 7.8e-05 | 108.0582 | 109.3847 | 121.0784 | 20.9398 | 0.002721 | 0.00254 | 129.0165 | 0.26078 | 8.4041 | 0.003583 | 0.005162 | 0.004995 | 6e-05 | 0.001092 | 0.35817 | 1.3661 | 2.5481 | 100.99 | 20.61 | 18.251 | 86.9666 | 107.7899 | 82.6176 | 0.004393 | 0.99235 | 0.023977 | 24.1617 | -0.10914 | 0.15514 | 0.00212 | 0.14869 | 0.14869 | 0.042884
2019-03-20 13:01 | 1553101239.1764 | 656 | 0.20916 | 0.000671 | 7.8e-05 | 108.1875 | 109.5712 | 121.0564 | 20.9398 | 0.002736 | 0.002546 | 129.0058 | 0.26113 | 8.4019 | 0.003649 | 0.00478 | 0.004587 | 5.9e-05 | 0.001087 | 0.33859 | 1.3384 | 2.548 | 100.99 | 20.61 | 18.249 | 87.3219 | 109.4001 | 82.7449 | 0.004392 | 0.99199 | 0.02397 | 24.1638 | -0.11367 | 0.15535 | 0.001937 | 0.13538 | 0.13538 | 0.047164
2019-03-20 13:01 | 1553101240.1763 | 657 | 0.20913 | 0.000673 | 7.7e-05 | 108.151 | 109.3761 | 121.2756 | 20.9389 | 0.002736 | 0.002553 | 128.9953 | 0.26098 | 8.404 | 0.003659 | 0.004498 | 0.004338 | 6e-05 | 0.001088 | 0.33699 | 1.309 | 2.548 | 100.99 | 20.61 | 18.245 | 87.3366 | 109.6372 | 82.914 | 0.004391 | 0.99152 | 0.023986 | 24.1915 | -0.10076 | 0.15526 | 0.002109 | 0.14804 | 0.14804 | 0.043103
2019-03-20 13:01 | 1553101241.1778 | 658 | 0.20917 | 0.000672 | 7.7e-05 | 108.1671 | 109.233 | 121.3042 | 20.9399 | 0.002715 | 0.002573 | 129.0117 | 0.26141 | 8.4046 | 0.003563 | 0.005056 | 0.00492 | 6.1e-05 | 0.001088 | 0.34307 | 1.3778 | 2.5481 | 100.99 | 20.61 | 18.245 | 86.8127 | 107.3089 | 83.3938 | 0.004391 | 0.99019 | 0.023973 | 24.2106 | -0.087704 | 0.15551 | 0.001896 | 0.13252 | 0.13252 | 0.048231
2019-03-20 13:01 | 1553101242.1764 | 659 | 0.20918 | 0.000677 | 7.6e-05 | 108.095 | 109.4606 | 121.4577 | 20.94 | 0.002684 | 0.002584 | 129.0144 | 0.26175 | 8.4031 | 0.003594 | 0.005035 | 0.004941 | 5.9e-05 | 0.001086 | 0.35579 | 1.3594 | 2.5482 | 100.99 | 20.61 | 18.242 | 86.0635 | 108.0618 | 83.6584 | 0.004391 | 0.98945 | 0.023979 | 24.2349 | -0.11229 | 0.15572 | 0.001802 | 0.12579 | 0.12579 | 0.05088
2019-03-20 13:01 | 1553101243.1775 | 660 | 0.20917 | 0.00067 | 7.6e-05 | 108.226 | 109.4808 | 121.4453 | 20.9385 | 0.002708 | 0.002576 | 129.0237 | 0.26038 | 8.4007 | 0.003676 | 0.00517 | 0.005085 | 5.7e-05 | 0.001091 | 0.32333 | 1.3346 | 2.548 | 100.99 | 20.61 | 18.24 | 86.6468 | 110.0439 | 83.4846 | 0.00439 | 0.98994 | 0.023984 | 24.2278 | -0.10311 | 0.1549 | 0.001879 | 0.13145 | 0.13145 | 0.048433
2019-03-20 13:01 | 1553101244.1766 | 661 | 0.20917 | 0.00067 | 7.6e-05 | 108.1584 | 109.4258 | 121.1676 | 20.939 | 0.002705 | 0.002562 | 129.0341 | 0.26058 | 8.4008 | 0.003627 | 0.004982 | 0.004765 | 5.9e-05 | 0.00109 | 0.32228 | 1.2997 | 2.5481 | 100.99 | 20.61 | 18.238 | 86.5732 | 108.8585 | 83.1347 | 0.00439 | 0.99091 | 0.023968 | 24.1882 | -0.10421 | 0.15502 | 0.001897 | 0.1327 | 0.1327 | 0.048015
2019-03-20 13:01 | 1553101245.1776 | 662 | 0.20917 | 0.000675 | 7.5e-05 | 108.2024 | 109.4957 | 121.231 | 20.9385 | 0.00271 | 0.002595 | 129.0456 | 0.26144 | 8.4043 | 0.003536 | 0.005549 | 0.005429 | 6.1e-05 | 0.001092 | 0.32997 | 1.2975 | 2.5481 | 100.99 | 20.61 | 18.232 | 86.6959 | 106.647 | 83.9326 | 0.004388 | 0.98869 | 0.023948 | 24.2216 | -0.10628 | 0.15553 | 0.001866 | 0.13033 | 0.13033 | 0.049047
2019-03-20 13:01 | 1553101246.1777 | 663 | 0.20919 | 0.000672 | 7.6e-05 | 108.1723 | 109.5555 | 121.2153 | 20.9375 | 0.002684 | 0.002597 | 129.0452 | 0.2607 | 8.4015 | 0.003594 | 0.004981 | 0.00491 | 6.1e-05 | 0.001092 | 0.3402 | 1.3371 | 2.548 | 100.99 | 20.61 | 18.23 | 86.0612 | 108.058 | 83.9908 | 0.004388 | 0.98853 | 0.023944 | 24.222 | -0.11365 | 0.15509 | 0.001755 | 0.12219 | 0.12219 | 0.052164
2019-03-20 13:01 | 1553101247.1766 | 664 | 0.20917 | 0.000672 | 7.7e-05 | 108.2017 | 109.4711 | 121.1221 | 20.9388 | 0.002718 | 0.002536 | 129.067 | 0.26122 | 8.4009 | 0.003619 | 0.005258 | 0.005146 | 6.1e-05 | 0.00109 | 0.32997 | 1.3238 | 2.5479 | 100.99 | 20.61 | 18.228 | 86.8843 | 108.6651 | 82.5263 | 0.004387 | 0.9926 | 0.023984 | 24.163 | -0.10432 | 0.1554 | 0.001895 | 0.13253 | 0.13253 | 0.048192
2019-03-20 13:01 | 1553101248.2277 | 665 | 0.20914 | 0.000673 | 7.7e-05 | 108.1875 | 109.4027 | 121.1601 | 20.9389 | 0.002717 | 0.00258 | 129.0854 | 0.26098 | 8.4036 | 0.003599 | 0.004695 | 0.00463 | 6e-05 | 0.001095 | 0.33483 | 1.3339 | 2.5479 | 100.99 | 20.61 | 18.228 | 86.8726 | 108.1829 | 83.5708 | 0.004387 | 0.9897 | 0.023953 | 24.2022 | -0.099904 | 0.15526 | 0.002027 | 0.14184 | 0.14184 | 0.044987
2019-03-20 13:01 | 1553101249.1812 | 666 | 0.20917 | 0.000675 | 7.7e-05 | 108.2632 | 109.4565 | 121.1585 | 20.9386 | 0.002669 | 0.002597 | 129.0698 | 0.2616 | 8.4012 | 0.003548 | 0.00486 | 0.004748 | 5.7e-05 | 0.001096 | 0.34578 | 1.3628 | 2.548 | 100.99 | 20.61 | 18.228 | 85.7055 | 106.9313 | 83.9832 | 0.004387 | 0.98855 | 0.023939 | 24.216 | -0.098048 | 0.15563 | 0.001874 | 0.13076 | 0.13076 | 0.048916
2019-03-20 13:01 | 1553101250.1811 | 667 | 0.20919 | 0.000675 | 7.7e-05 | 108.236 | 109.4841 | 121.0719 | 20.9378 | 0.002709 | 0.002568 | 129.0636 | 0.26095 | 8.404 | 0.003645 | 0.00501 | 0.00484 | 6e-05 | 0.001095 | 0.30757 | 1.3807 | 2.5482 | 100.99 | 20.61 | 18.231 | 86.6649 | 109.3002 | 83.2944 | 0.004388 | 0.99046 | 0.023953 | 24.184 | -0.10255 | 0.15524 | 0.001781 | 0.12411 | 0.12411 | 0.051411
2019-03-20 13:01 | 1553101251.1813 | 668 | 0.20917 | 0.000675 | 7.7e-05 | 108.2493 | 109.4359 | 121.1174 | 20.9376 | 0.002709 | 0.002519 | 129.0678 | 0.26102 | 8.4008 | 0.003625 | 0.004942 | 0.00482 | 5.6e-05 | 0.001092 | 0.33572 | 1.3479 | 2.548 | 100.99 | 20.61 | 18.233 | 86.6655 | 108.8027 | 82.1165 | 0.004388 | 0.99375 | 0.023998 | 24.1486 | -0.097517 | 0.15528 | 0.001848 | 0.12918 | 0.12918 | 0.049403
2019-03-20 13:01 | 1553101252.1805 | 669 | 0.20918 | 0.000672 | 7.7e-05 | 108.2001 | 109.4415 | 120.9399 | 20.938 | 0.002714 | 0.002547 | 129.0696 | 0.25873 | 8.402 | 0.003545 | 0.00481 | 0.004679 | 5.6e-05 | 0.001092 | 0.32853 | 1.3419 | 2.5481 | 100.99 | 20.61 | 18.234 | 86.7837 | 106.8707 | 82.7696 | 0.004389 | 0.99192 | 0.023958 | 24.1531 | -0.10204 | 0.15392 | 0.0018 | 0.12553 | 0.12553 | 0.050396
2019-03-20 13:01 | 1553101253.1807 | 670 | 0.20922 | 0.000676 | 7.7e-05 | 108.2674 | 109.6576 | 120.9397 | 20.9366 | 0.002696 | 0.002549 | 129.0442 | 0.26012 | 8.3994 | 0.003567 | 0.004903 | 0.004725 | 5.8e-05 | 0.001091 | 0.34451 | 1.3499 | 2.548 | 100.99 | 20.61 | 18.239 | 86.3553 | 107.3931 | 82.8175 | 0.00439 | 0.99179 | 0.023956 | 24.1547 | -0.11412 | 0.15475 | 0.001554 | 0.1077 | 0.1077 | 0.059053
2019-03-20 13:01 | 1553101254.1816 | 671 | 0.20922 | 0.000676 | 7.7e-05 | 108.2144 | 109.5869 | 121.2545 | 20.9371 | 0.002755 | 0.002573 | 129.0413 | 0.26066 | 8.4022 | 0.003639 | 0.004986 | 0.004872 | 5.5e-05 | 0.001091 | 0.33179 | 1.3602 | 2.548 | 100.99 | 20.61 | 18.244 | 87.79 | 109.1404 | 83.3953 | 0.004391 | 0.99018 | 0.023968 | 24.2057 | -0.11274 | 0.15507 | 0.001543 | 0.10697 | 0.10697 | 0.05958
2019-03-20 13:01 | 1553101255.1814 | 672 | 0.20919 | 0.000668 | 7.8e-05 | 108.2448 | 109.4957 | 121.6152 | 20.9369 | 0.002726 | 0.002553 | 129.0577 | 0.26052 | 8.3993 | 0.003554 | 0.005318 | 0.00519 | 6.3e-05 | 0.00109 | 0.3126 | 1.3193 | 2.548 | 100.99 | 20.61 | 18.256 | 87.0801 | 107.0759 | 82.9221 | 0.004394 | 0.9915 | 0.02402 | 24.2256 | -0.10278 | 0.15499 | 0.001777 | 0.12416 | 0.12416 | 0.051304
2019-03-20 13:01 | 1553101256.1804 | 673 | 0.20924 | 0.000674 | 7.8e-05 | 108.2774 | 109.4797 | 121.6504 | 20.9373 | 0.002692 | 0.002554 | 129.0646 | 0.25984 | 8.4034 | 0.00355 | 0.005004 | 0.004893 | 6.1e-05 | 0.001087 | 0.34468 | 1.2877 | 2.548 | 100.99 | 20.61 | 18.263 | 86.2599 | 106.9889 | 82.9562 | 0.004396 | 0.9914 | 0.024022 | 24.2302 | -0.098766 | 0.15458 | 0.001486 | 0.10301 | 0.10301 | 0.061675
2019-03-20 13:01 | 1553101257.1819 | 674 | 0.20922 | 0.000671 | 7.7e-05 | 108.206 | 109.4412 | 121.5382 | 20.9376 | 0.002704 | 0.00255 | 129.0303 | 0.25955 | 8.4029 | 0.003518 | 0.004834 | 0.00478 | 5.9e-05 | 0.001086 | 0.32532 | 1.2956 | 2.5481 | 100.99 | 20.61 | 18.271 | 86.5451 | 106.213 | 82.8543 | 0.004398 | 0.99169 | 0.024014 | 24.2156 | -0.10153 | 0.15441 | 0.001603 | 0.11156 | 0.11156 | 0.056887
2019-03-20 13:01 | 1553101258.1806 | 675 | 0.20925 | 0.000668 | 7.8e-05 | 108.268 | 109.5832 | 121.3249 | 20.937 | 0.002697 | 0.002498 | 129.0273 | 0.26128 | 8.4005 | 0.003576 | 0.005314 | 0.005058 | 5.5e-05 | 0.001086 | 0.31045 | 1.2835 | 2.5481 | 100.99 | 20.61 | 18.288 | 86.3947 | 107.618 | 81.6066 | 0.004402 | 0.99518 | 0.024035 | 24.1519 | -0.10799 | 0.15543 | 0.001425 | 0.098673 | 0.098673 | 0.064742
2019-03-20 13:01 | 1553101259.181 | 676 | 0.20919 | 0.000667 | 7.8e-05 | 108.2066 | 109.3963 | 121.3817 | 20.9369 | 0.002676 | 0.002538 | 129.0127 | 0.26137 | 8.4023 | 0.003556 | 0.005065 | 0.00491 | 5.8e-05 | 0.001085 | 0.31492 | 1.3034 | 2.5482 | 100.99 | 20.61 | 18.297 | 85.8903 | 107.1225 | 82.569 | 0.004404 | 0.99248 | 0.024008 | 24.1903 | -0.097807 | 0.15549 | 0.001742 | 0.12149 | 0.12149 | 0.0526
2019-03-20 13:01 | 1553101260.1817 | 677 | 0.20925 | 0.000669 | 7.8e-05 | 108.2312 | 109.3148 | 121.5171 | 20.9371 | 0.002648 | 0.002506 | 129.0264 | 0.26031 | 8.4014 | 0.003545 | 0.005389 | 0.00534 | 5.8e-05 | 0.00108 | 0.33826 | 1.2888 | 2.5481 | 100.99 | 20.61 | 18.315 | 85.1999 | 106.8762 | 81.7853 | 0.004408 | 0.99467 | 0.024048 | 24.1771 | -0.089111 | 0.15486 | 0.001414 | 0.097912 | 0.097912 | 0.065004
2019-03-20 13:01 | 1553101261.1807 | 678 | 0.20928 | 0.000669 | 7.7e-05 | 108.2072 | 109.3817 | 121.426 | 20.9366 | 0.002628 | 0.002536 | 128.991 | 0.26062 | 8.4026 | 0.003534 | 0.005401 | 0.005199 | 6.2e-05 | 0.001074 | 0.34523 | 1.2774 | 2.548 | 100.99 | 20.62 | 18.324 | 84.7133 | 106.5932 | 82.5235 | 0.004413 | 0.99261 | 0.024014 | 24.1932 | -0.096561 | 0.15504 | 0.001217 | 0.083509 | 0.083509 | 0.076306
2019-03-20 13:01 | 1553101262.1812 | 679 | 0.20925 | 0.000664 | 7.7e-05 | 108.1869 | 109.3754 | 121.5037 | 20.9364 | 0.002682 | 0.002583 | 128.958 | 0.26079 | 8.4049 | 0.003561 | 0.004871 | 0.004839 | 5.9e-05 | 0.00107 | 0.31437 | 1.3128 | 2.548 | 100.99 | 20.61 | 18.333 | 86.0249 | 107.262 | 83.6393 | 0.004412 | 0.98951 | 0.023985 | 24.2389 | -0.097726 | 0.15515 | 0.001388 | 0.095785 | 0.095785 | 0.066572
2019-03-20 13:01 | 1553101263.1808 | 680 | 0.20924 | 0.000666 | 7.7e-05 | 108.2179 | 109.4539 | 121.3646 | 20.9365 | 0.00268 | 0.002575 | 128.965 | 0.26106 | 8.4009 | 0.003623 | 0.004842 | 0.004747 | 5.8e-05 | 0.001072 | 0.33787 | 1.3241 | 2.5479 | 100.99 | 20.62 | 18.356 | 85.9798 | 108.7535 | 83.4468 | 0.004421 | 0.99004 | 0.023977 | 24.2184 | -0.10158 | 0.1553 | 0.001451 | 0.10038 | 0.10038 | 0.063591
2019-03-20 13:01 | 1553101264.1821 | 681 | 0.20924 | 0.000666 | 7.6e-05 | 108.1985 | 109.3731 | 121.4237 | 20.936 | 0.002666 | 0.002539 | 128.9413 | 0.26219 | 8.4028 | 0.003528 | 0.004493 | 0.004384 | 5.7e-05 | 0.001075 | 0.32925 | 1.3404 | 2.5479 | 100.99 | 20.62 | 18.366 | 85.6366 | 106.4594 | 82.5828 | 0.004423 | 0.99245 | 0.024012 | 24.195 | -0.09658 | 0.15598 | 0.001461 | 0.10127 | 0.10127 | 0.063304
2019-03-20 13:01 | 1553101265.181 | 682 | 0.20924 | 0.000663 | 7.6e-05 | 108.2021 | 109.4539 | 121.2547 | 20.9354 | 0.002706 | 0.002552 | 128.9151 | 0.26074 | 8.4033 | 0.003583 | 0.005407 | 0.005309 | 6.1e-05 | 0.001083 | 0.34844 | 1.3533 | 2.5481 | 100.99 | 20.62 | 18.387 | 86.5899 | 107.7801 | 82.8928 | 0.004428 | 0.99158 | 0.023985 | 24.1887 | -0.10289 | 0.15512 | 0.00147 | 0.10181 | 0.10181 | 0.062621
2019-03-20 13:01 | 1553101266.182 | 683 | 0.20926 | 0.000663 | 7.5e-05 | 108.2111 | 109.5417 | 121.3343 | 20.9352 | 0.002668 | 0.0025 | 128.8973 | 0.26081 | 8.3998 | 0.003614 | 0.005388 | 0.005203 | 6e-05 | 0.001087 | 0.35187 | 1.3243 | 2.5481 | 100.99 | 20.62 | 18.395 | 85.6741 | 108.5276 | 81.6563 | 0.00443 | 0.99504 | 0.024035 | 24.1545 | -0.10931 | 0.15515 | 0.001357 | 0.093862 | 0.093862 | 0.067939
2019-03-20 13:01 | 1553101267.1814 | 684 | 0.20926 | 0.000665 | 7.6e-05 | 108.2131 | 109.3642 | 121.3432 | 20.9357 | 0.002711 | 0.002514 | 128.8876 | 0.26072 | 8.4026 | 0.003506 | 0.004974 | 0.004881 | 6.1e-05 | 0.001088 | 0.33196 | 1.3409 | 2.5482 | 100.99 | 20.62 | 18.403 | 86.7105 | 105.9121 | 81.9733 | 0.004432 | 0.99415 | 0.024025 | 24.1662 | -0.094643 | 0.1551 | 0.001339 | 0.092548 | 0.092548 | 0.068879
2019-03-20 13:01 | 1553101268.1802 | 685 | 0.20925 | 0.000659 | 7.6e-05 | 108.2661 | 109.5006 | 121.2566 | 20.936 | 0.002712 | 0.002553 | 128.877 | 0.26015 | 8.4041 | 0.003568 | 0.005202 | 0.005039 | 5.6e-05 | 0.001088 | 0.32532 | 1.2969 | 2.5481 | 100.99 | 20.62 | 18.418 | 86.7528 | 107.4167 | 82.919 | 0.004436 | 0.99151 | 0.023984 | 24.1897 | -0.10141 | 0.15476 | 0.001404 | 0.097074 | 0.097074 | 0.065524
2019-03-20 13:01 | 1553101269.1803 | 686 | 0.20924 | 0.000659 | 7.5e-05 | 108.271 | 109.674 | 121.1615 | 20.9356 | 0.002702 | 0.002531 | 128.8868 | 0.26072 | 8.4031 | 0.003504 | 0.005093 | 0.004923 | 5.9e-05 | 0.001092 | 0.33483 | 1.2813 | 2.5481 | 100.99 | 20.63 | 18.425 | 86.508 | 105.8702 | 82.3929 | 0.00444 | 0.99298 | 0.023993 | 24.1624 | -0.11516 | 0.1551 | 0.001454 | 0.10073 | 0.10073 | 0.063282
2019-03-20 13:01 | 1553101270.1813 | 687 | 0.20929 | 0.000666 | 7.5e-05 | 108.2328 | 109.443 | 121.1824 | 20.9358 | 0.002692 | 0.002567 | 128.8592 | 0.26104 | 8.4048 | 0.003512 | 0.005663 | 0.005527 | 6e-05 | 0.001093 | 0.30901 | 1.2859 | 2.548 | 100.99 | 20.63 | 18.432 | 86.2616 | 106.0602 | 83.2581 | 0.004442 | 0.99057 | 0.023966 | 24.1938 | -0.09946 | 0.15529 | 0.001155 | 0.07897 | 0.07897 | 0.080821
2019-03-20 13:01 | 1553101271.1812 | 688 | 0.20918 | 0.000661 | 7.5e-05 | 108.215 | 109.3197 | 121.1798 | 20.9354 | 0.002683 | 0.002522 | 128.8635 | 0.26082 | 8.4019 | 0.003596 | 0.00489 | 0.004802 | 6.4e-05 | 0.001094 | 0.33428 | 1.3306 | 2.5481 | 100.99 | 20.63 | 18.444 | 86.0498 | 108.1023 | 82.1673 | 0.004445 | 0.99361 | 0.024002 | 24.1565 | -0.090845 | 0.15516 | 0.001826 | 0.12771 | 0.12771 | 0.049934
2019-03-20 13:01 | 1553101272.1808 | 689 | 0.20925 | 0.000654 | 7.4e-05 | 108.1924 | 109.4326 | 120.9827 | 20.9357 | 0.002697 | 0.002537 | 128.8694 | 0.26152 | 8.4024 | 0.003584 | 0.005348 | 0.005133 | 6e-05 | 0.001095 | 0.32748 | 1.368 | 2.5481 | 100.99 | 20.63 | 18.449 | 86.3809 | 107.8119 | 82.5411 | 0.004446 | 0.99256 | 0.02397 | 24.1496 | -0.10195 | 0.15558 | 0.001443 | 0.099903 | 0.099903 | 0.064004
2019-03-20 13:01 | 1553101273.1811 | 690 | 0.20925 | 0.000651 | 7.4e-05 | 108.1613 | 109.4692 | 120.9228 | 20.9354 | 0.002724 | 0.002509 | 128.8598 | 0.2616 | 8.4015 | 0.003565 | 0.00512 | 0.005006 | 6.3e-05 | 0.001099 | 0.32549 | 1.3561 | 2.5481 | 100.99 | 20.63 | 18.459 | 87.0365 | 107.3484 | 81.8708 | 0.004448 | 0.99444 | 0.023987 | 24.1208 | -0.10751 | 0.15563 | 0.001419 | 0.098265 | 0.098265 | 0.065092
2019-03-20 13:01 | 1553101274.1815 | 691 | 0.20927 | 0.000657 | 7.4e-05 | 108.1778 | 109.344 | 120.8587 | 20.9357 | 0.002675 | 0.002505 | 128.8629 | 0.26056 | 8.4049 | 0.003509 | 0.005182 | 0.005076 | 5.8e-05 | 0.001099 | 0.32421 | 1.3149 | 2.5479 | 100.99 | 20.63 | 18.462 | 85.8431 | 105.9924 | 81.7752 | 0.004449 | 0.9947 | 0.023983 | 24.1112 | -0.095907 | 0.15501 | 0.001302 | 0.089725 | 0.089725 | 0.071004
2019-03-20 13:01 | 1553101275.1814 | 692 | 0.20925 | 0.00065 | 7.4e-05 | 108.2098 | 109.335 | 120.7569 | 20.9352 | 0.002665 | 0.002539 | 128.8625 | 0.26061 | 8.4051 | 0.003529 | 0.00483 | 0.004736 | 5.7e-05 | 0.001099 | 0.33301 | 1.3414 | 2.5481 | 100.99 | 20.63 | 18.465 | 85.6031 | 106.487 | 82.575 | 0.00445 | 0.99247 | 0.023946 | 24.1282 | -0.092526 | 0.15504 | 0.001454 | 0.10059 | 0.10059 | 0.063345
2019-03-20 13:01 | 1553101276.1815 | 693 | 0.20922 | 0.000659 | 7.4e-05 | 108.1972 | 109.5682 | 120.5681 | 20.936 | 0.002609 | 0.002537 | 128.886 | 0.26101 | 8.4018 | 0.003561 | 0.005216 | 0.005068 | 6e-05 | 0.0011 | 0.31957 | 1.3642 | 2.5482 | 100.99 | 20.63 | 18.472 | 84.2788 | 107.2514 | 82.5305 | 0.004451 | 0.99259 | 0.023929 | 24.1078 | -0.11263 | 0.15527 | 0.001601 | 0.11117 | 0.11117 | 0.057403
2019-03-20 13:01 | 1553101277.1816 | 694 | 0.20923 | 0.000657 | 7.5e-05 | 108.1704 | 109.3563 | 120.7277 | 20.936 | 0.002752 | 0.002502 | 128.9014 | 0.26133 | 8.403 | 0.003599 | 0.004903 | 0.004831 | 6e-05 | 0.0011 | 0.3282 | 1.3363 | 2.548 | 100.99 | 20.63 | 18.474 | 87.7123 | 108.1737 | 81.6914 | 0.004452 | 0.99494 | 0.023973 | 24.0953 | -0.097531 | 0.15547 | 0.001554 | 0.10794 | 0.10794 | 0.059197
2019-03-20 13:01 | 1553101278.1806 | 695 | 0.20923 | 0.000656 | 7.4e-05 | 108.2571 | 109.6983 | 120.8496 | 20.9358 | 0.002656 | 0.002457 | 128.891 | 0.26086 | 8.4044 | 0.003524 | 0.005007 | 0.004912 | 5.6e-05 | 0.001099 | 0.33284 | 1.3331 | 2.5479 | 100.99 | 20.63 | 18.481 | 85.4006 | 106.3529 | 80.6048 | 0.004454 | 0.99799 | 0.024022 | 24.0705 | -0.11829 | 0.15519 | 0.001531 | 0.10653 | 0.10653 | 0.059871
2019-03-20 13:01 | 1553101279.1817 | 696 | 0.20923 | 0.000655 | 7.5e-05 | 108.1665 | 109.6056 | 120.9874 | 20.9355 | 0.002622 | 0.002498 | 128.8963 | 0.26144 | 8.4008 | 0.0036 | 0.00487 | 0.004722 | 5.9e-05 | 0.001097 | 0.31454 | 1.3265 | 2.5481 | 100.99 | 20.63 | 18.485 | 84.5812 | 108.1923 | 81.6075 | 0.004455 | 0.99517 | 0.024002 | 24.1183 | -0.11822 | 0.15553 | 0.001528 | 0.10617 | 0.10617 | 0.060207
2019-03-20 13:01 | 1553101280.1811 | 697 | 0.20924 | 0.000657 | 7.5e-05 | 108.1633 | 109.483 | 120.8629 | 20.9353 | 0.00261 | 0.002525 | 128.8928 | 0.26043 | 8.4047 | 0.003591 | 0.004747 | 0.004645 | 6e-05 | 0.001093 | 0.32316 | 1.3049 | 2.548 | 100.99 | 20.63 | 18.489 | 84.2921 | 107.9832 | 82.2523 | 0.004455 | 0.99337 | 0.023968 | 24.1278 | -0.10848 | 0.15493 | 0.001469 | 0.10177 | 0.10177 | 0.062568
2019-03-20 13:01 | 1553101281.1815 | 698 | 0.20921 | 0.000655 | 7.5e-05 | 108.1733 | 109.4064 | 120.9172 | 20.9347 | 0.002704 | 0.002484 | 128.8956 | 0.26058 | 8.4062 | 0.003537 | 0.004774 | 0.004642 | 6.1e-05 | 0.00109 | 0.32421 | 1.2956 | 2.5479 | 100.99 | 20.63 | 18.494 | 86.5596 | 106.6731 | 81.2489 | 0.004457 | 0.99618 | 0.024007 | 24.0991 | -0.10139 | 0.15502 | 0.001662 | 0.11587 | 0.11587 | 0.054987
2019-03-20 13:01 | 1553101282.1809 | 699 | 0.2092 | 0.000656 | 7.5e-05 | 108.184 | 109.3784 | 120.9827 | 20.9358 | 0.002675 | 0.002527 | 128.9148 | 0.26104 | 8.4053 | 0.003439 | 0.004885 | 0.00479 | 6e-05 | 0.001087 | 0.3246 | 1.2751 | 2.548 | 100.99 | 20.63 | 18.496 | 85.8488 | 104.3046 | 82.284 | 0.004457 | 0.99328 | 0.023979 | 24.1408 | -0.098211 | 0.15529 | 0.001735 | 0.121 | 0.121 | 0.052748
2019-03-20 13:01 | 1553101283.1804 | 700 | 0.2092 | 0.000658 | 7.5e-05 | 108.2205 | 109.4075 | 121.0226 | 20.9356 | 0.002613 | 0.002535 | 128.908 | 0.26032 | 8.4041 | 0.003489 | 0.004364 | 0.004198 | 6.1e-05 | 0.001092 | 0.335 | 1.2985 | 2.5481 | 100.99 | 20.63 | 18.496 | 84.3556 | 105.506 | 82.4888 | 0.004457 | 0.99271 | 0.023976 | 24.1518 | -0.097574 | 0.15487 | 0.001734 | 0.12099 | 0.12099 | 0.052608
2019-03-20 13:01 | 1553101284.18 | 701 | 0.20923 | 0.000659 | 7.5e-05 | 108.181 | 109.4991 | 121.0296 | 20.9351 | 0.002618 | 0.002536 | 128.8976 | 0.26094 | 8.4014 | 0.003481 | 0.005376 | 0.005242 | 6e-05 | 0.001087 | 0.34147 | 1.2835 | 2.548 | 100.99 | 20.63 | 18.491 | 84.4792 | 105.2994 | 82.5141 | 0.004456 | 0.99264 | 0.023975 | 24.1533 | -0.10832 | 0.15523 | 0.00154 | 0.10696 | 0.10696 | 0.05965
2019-03-20 13:01 | 1553101285.18 | 702 | 0.20924 | 0.000653 | 7.4e-05 | 108.1584 | 109.3896 | 121.0101 | 20.9352 | 0.002657 | 0.002526 | 128.9013 | 0.26127 | 8.402 | 0.003561 | 0.005056 | 0.004901 | 6.1e-05 | 0.001091 | 0.33339 | 1.3407 | 2.5481 | 100.99 | 20.63 | 18.485 | 85.4275 | 107.2514 | 82.2702 | 0.004455 | 0.99332 | 0.023982 | 24.1431 | -0.10124 | 0.15543 | 0.001503 | 0.10432 | 0.10432 | 0.061237
2019-03-20 13:01 | 1553101286.1813 | 703 | 0.20933 | 0.000655 | 7.4e-05 | 108.1367 | 109.3406 | 121.0094 | 20.9353 | 0.00271 | 0.002485 | 128.8874 | 0.26079 | 8.4014 | 0.003555 | 0.005528 | 0.005414 | 5.7e-05 | 0.001091 | 0.34971 | 1.3329 | 2.5481 | 100.99 | 20.63 | 18.469 | 86.6875 | 107.113 | 81.279 | 0.004451 | 0.9961 | 0.024015 | 24.1093 | -0.099029 | 0.15514 | 0.000966 | 0.065503 | 0.065503 | 0.097345
2019-03-20 13:01 | 1553101287.1811 | 704 | 0.20922 | 0.000656 | 7.4e-05 | 108.1481 | 109.52 | 120.7425 | 20.9353 | 0.002659 | 0.002519 | 128.88 | 0.26141 | 8.4008 | 0.003642 | 0.005135 | 0.005076 | 6e-05 | 0.001091 | 0.33085 | 1.3565 | 2.5483 | 100.99 | 20.63 | 18.458 | 85.4728 | 109.219 | 82.1071 | 0.004448 | 0.99377 | 0.023961 | 24.1109 | -0.11276 | 0.15551 | 0.001591 | 0.11063 | 0.11063 | 0.057774
2019-03-20 13:01 | 1553101288.1816 | 705 | 0.20919 | 0.000652 | 7.4e-05 | 108.15 | 109.3866 | 120.8808 | 20.9348 | 0.002692 | 0.002533 | 128.8902 | 0.26107 | 8.3993 | 0.003629 | 0.005395 | 0.005219 | 5.9e-05 | 0.001085 | 0.31222 | 1.3301 | 2.5482 | 100.99 | 20.63 | 18.446 | 86.2652 | 108.9081 | 82.4453 | 0.004445 | 0.99283 | 0.023963 | 24.1361 | -0.10169 | 0.15531 | 0.001817 | 0.12696 | 0.12696 | 0.050276
2019-03-20 13:01 | 1553101289.1803 | 706 | 0.20924 | 0.000654 | 7.4e-05 | 108.203 | 109.477 | 120.95 | 20.9343 | 0.002677 | 0.002545 | 128.902 | 0.26057 | 8.4014 | 0.003667 | 0.005224 | 0.005064 | 6e-05 | 0.001084 | 0.3402 | 1.342 | 2.5481 | 100.99 | 20.63 | 18.422 | 85.9114 | 109.8371 | 82.7208 | 0.004439 | 0.99206 | 0.023961 | 24.1524 | -0.1047 | 0.15501 | 0.001466 | 0.10154 | 0.10154 | 0.062745
2019-03-20 13:01 | 1553101290.1805 | 707 | 0.20927 | 0.000648 | 7.5e-05 | 108.0953 | 109.3429 | 121.1972 | 20.9352 | 0.002636 | 0.002481 | 128.9257 | 0.261 | 8.403 | 0.003687 | 0.00476 | 0.004678 | 6e-05 | 0.001085 | 0.34595 | 1.3228 | 2.5481 | 100.99 | 20.63 | 18.411 | 84.9241 | 110.3229 | 81.1907 | 0.004437 | 0.99634 | 0.024037 | 24.125 | -0.10264 | 0.15527 | 0.001353 | 0.093603 | 0.093603 | 0.068178
2019-03-20 13:02 | 1553101291.1809 | 708 | 0.20923 | 0.000653 | 7.5e-05 | 108.1403 | 109.4042 | 121.2024 | 20.9336 | 0.002662 | 0.002487 | 128.9205 | 0.26098 | 8.4021 | 0.00367 | 0.005839 | 0.005714 | 5.9e-05 | 0.001082 | 0.33682 | 1.2948 | 2.5481 | 100.99 | 20.63 | 18.39 | 85.55 | 109.8961 | 81.3201 | 0.004432 | 0.99598 | 0.024033 | 24.13 | -0.10393 | 0.15526 | 0.00155 | 0.10789 | 0.10789 | 0.059143
2019-03-20 13:02 | 1553101292.1814 | 709 | 0.20924 | 0.000649 | 7.5e-05 | 108.1164 | 109.421 | 121.2139 | 20.9344 | 0.002629 | 0.002502 | 128.9098 | 0.26115 | 8.4035 | 0.003712 | 0.005173 | 0.005107 | 5.8e-05 | 0.001081 | 0.32532 | 1.3168 | 2.5481 | 100.99 | 20.63 | 18.382 | 84.7462 | 110.9074 | 81.6836 | 0.00443 | 0.99496 | 0.024022 | 24.1435 | -0.10729 | 0.15536 | 0.00147 | 0.10205 | 0.10205 | 0.062573
2019-03-20 13:02 | 1553101293.1809 | 710 | 0.20924 | 0.000653 | 7.4e-05 | 108.0727 | 109.3548 | 121.4465 | 20.9343 | 0.002615 | 0.002504 | 128.9117 | 0.26149 | 8.4033 | 0.003634 | 0.005561 | 0.005418 | 6.1e-05 | 0.001084 | 0.34363 | 1.3492 | 2.548 | 100.99 | 20.63 | 18.373 | 84.4183 | 109.021 | 81.7404 | 0.004428 | 0.9948 | 0.024043 | 24.1686 | -0.10549 | 0.15556 | 0.001523 | 0.10602 | 0.10602 | 0.060305
2019-03-20 13:02 | 1553101294.1814 | 711 | 0.20923 | 0.000651 | 7.4e-05 | 108.1532 | 109.2961 | 121.55 | 20.9343 | 0.002589 | 0.00247 | 128.9289 | 0.26083 | 8.4009 | 0.003794 | 0.005398 | 0.005236 | 5.9e-05 | 0.001082 | 0.33483 | 1.3863 | 2.5479 | 100.99 | 20.63 | 18.361 | 83.7818 | 112.9144 | 80.9295 | 0.004425 | 0.99708 | 0.024081 | 24.1512 | -0.094025 | 0.15517 | 0.001542 | 0.10758 | 0.10758 | 0.059283
2019-03-20 13:02 | 1553101295.181 | 712 | 0.20924 | 0.000651 | 7.4e-05 | 108.0643 | 109.3197 | 121.8258 | 20.9342 | 0.002648 | 0.00247 | 128.9258 | 0.26065 | 8.4034 | 0.003701 | 0.005079 | 0.005004 | 5.9e-05 | 0.001081 | 0.3194 | 1.3512 | 2.5482 | 100.99 | 20.63 | 18.358 | 85.2032 | 110.6445 | 80.927 | 0.004424 | 0.99709 | 0.024108 | 24.1785 | -0.10331 | 0.15506 | 0.00147 | 0.10247 | 0.10247 | 0.062192
2019-03-20 13:02 | 1553101296.1806 | 713 | 0.20922 | 0.000652 | 7.3e-05 | 108.0714 | 109.3877 | 122.095 | 20.9331 | 0.002641 | 0.002506 | 128.9256 | 0.26137 | 8.4003 | 0.003679 | 0.00511 | 0.005009 | 5.8e-05 | 0.001078 | 0.31974 | 1.3386 | 2.5481 | 100.99 | 20.63 | 18.356 | 85.0375 | 110.122 | 81.7994 | 0.004423 | 0.99464 | 0.024105 | 24.235 | -0.10829 | 0.15549 | 0.001593 | 0.11146 | 0.11146 | 0.057334
2019-03-20 13:02 | 1553101297.1811 | 714 | 0.20924 | 0.000651 | 7.3e-05 | 108.1209 | 109.3795 | 122.4044 | 20.9332 | 0.002678 | 0.002545 | 125.1872 | 0.26118 | 8.4008 | 0.00366 | 0.004685 | 0.004615 | 6e-05 | 0.001081 | 0.35673 | 1.3055 | 2.5481 | 100.99 | 20.63 | 18.359 | 85.9353 | 109.6656 | 82.7298 | 0.004424 | 0.99204 | 0.024104 | 24.2975 | -0.10352 | 0.15538 | 0.001529 | 0.10685 | 0.10685 | 0.059764
2019-03-20 13:02 | 1553101298.1798 | 715 | 0.20926 | 0.000649 | 7.2e-05 | 108.0016 | 109.2846 | 122.4525 | 20.9332 | 0.002644 | 0.002512 | 97.156 | 0.26089 | 8.4056 | 0.003595 | 0.005566 | 0.005352 | 5.6e-05 | 0.001077 | 0.33931 | 1.3313 | 2.548 | 100.99 | 20.63 | 18.364 | 85.0945 | 108.0671 | 81.9367 | 0.004425 | 0.99425 | 0.024136 | 24.2752 | -0.10563 | 0.1552 | 0.00141 | 0.098313 | 0.098313 | 0.064882
2019-03-20 13:02 | 1553101299.1798 | 716 | 0.20923 | 0.000649 | 7.2e-05 | 107.9805 | 109.0502 | 122.6708 | 20.9332 | 0.002624 | 0.002472 | 69.0283 | 0.26151 | 8.3996 | 0.003686 | 0.004853 | 0.004708 | 6.1e-05 | 0.001075 | 0.3449 | 1.2998 | 2.548 | 100.99 | 20.63 | 18.376 | 84.6241 | 110.2808 | 80.9758 | 0.004428 | 0.99695 | 0.02419 | 24.2639 | -0.088171 | 0.15557 | 0.001578 | 0.11081 | 0.11081 | 0.057704
2019-03-20 13:02 | 1553101300.1799 | 717 | 0.20923 | 0.000649 | 7e-05 | 108.0083 | 109.2502 | 123.1069 | 20.933 | 0.002639 | 0.002477 | 40.9307 | 0.26115 | 8.4017 | 0.003657 | 0.006148 | 0.006035 | 5.8e-05 | 0.001072 | 0.35618 | 1.2811 | 2.548 | 100.99 | 20.63 | 18.384 | 84.9928 | 109.5927 | 81.1013 | 0.00443 | 0.9966 | 0.024229 | 24.3113 | -0.10225 | 0.15536 | 0.001574 | 0.11081 | 0.11081 | 0.057624
2019-03-20 13:02 | 1553101301.1805 | 718 | 0.20922 | 0.000652 | 6.8e-05 | 108.0636 | 109.3268 | 123.4597 | 20.9326 | 0.002639 | 0.0025 | 12.8554 | 0.26047 | 8.4036 | 0.003664 | 0.0045 | 0.004421 | 5.9e-05 | 0.001069 | 0.3095 | 1.3175 | 2.5481 | 100.99 | 20.63 | 18.384 | 84.9918 | 109.7469 | 81.6429 | 0.00443 | 0.99507 | 0.024245 | 24.3647 | -0.10394 | 0.15495 | 0.001623 | 0.11464 | 0.11464 | 0.055552
2019-03-20 13:02 | 1553101302.1803 | 719 | 0.20926 | 0.000646 | 6.3e-05 | 108.0025 | 109.2023 | 123.7317 | 20.9317 | 0.00258 | 0.002458 | -15.2517 | 0.2579 | 8.4041 | 0.00366 | 0.005073 | 0.004946 | 5.8e-05 | 0.001074 | 0.32549 | 1.2993 | 2.548 | 100.99 | 20.64 | 18.401 | 83.56 | 109.6532 | 80.6369 | 0.004437 | 0.9979 | 0.024306 | 24.3569 | -0.098817 | 0.15343 | 0.001395 | 0.098449 | 0.098449 | 0.064051
2019-03-20 13:02 | 1553101303.18 | 720 | 0.20929 | 0.000642 | 5.8e-05 | 108.0669 | 109.2846 | 124.023 | 20.9326 | 0.002578 | 0.002459 | -43.3552 | 0.23924 | 8.4012 | 0.003604 | 0.005427 | 0.005201 | 5.8e-05 | 0.001071 | 0.34036 | 1.322 | 2.548 | 100.99 | 20.64 | 18.416 | 83.5063 | 108.2887 | 80.6572 | 0.004441 | 0.99785 | 0.024334 | 24.3863 | -0.10022 | 0.14233 | 0.001227 | 0.086555 | 0.086555 | 0.067583
2019-03-20 13:02 | 1553101304.1811 | 721 | 0.20929 | 0.000643 | 5.2e-05 | 107.9779 | 109.2917 | 124.1876 | 20.9322 | 0.002576 | 0.002381 | -71.4406 | 0.23869 | 8.4022 | 0.003561 | 0.004567 | 0.004418 | 5.6e-05 | 0.001067 | 0.33572 | 1.3639 | 2.5481 | 100.99 | 20.64 | 18.422 | 83.4646 | 107.2572 | 78.7832 | 0.004442 | 1.0032 | 0.024415 | 24.3378 | -0.10817 | 0.142 | 0.001243 | 0.088371 | 0.088371 | 0.066041
2019-03-20 13:02 | 1553101305.1802 | 722 | 0.20935 | 0.000641 | 4.7e-05 | 108.0591 | 109.3758 | 124.3277 | 20.9324 | 0.002596 | 0.002413 | -99.543 | 0.23882 | 8.4013 | 0.003541 | 0.004347 | 0.004174 | 5.6e-05 | 0.00107 | 0.33445 | 1.3435 | 2.5481 | 100.99 | 20.64 | 18.428 | 83.9562 | 106.7701 | 79.5615 | 0.004444 | 1.0009 | 0.024401 | 24.3784 | -0.10833 | 0.14208 | 0.000912 | 0.064317 | 0.064317 | 0.090789
2019-03-20 13:02 | 1553101306.1812 | 723 | 0.20939 | 0.000631 | 4.3e-05 | 108.0365 | 109.1657 | 124.3469 | 20.9323 | 0.002546 | 0.002457 | -127.6438 | 0.23892 | 8.4034 | 0.00353 | 0.004615 | 0.004404 | 6e-05 | 0.001066 | 0.34003 | 1.3305 | 2.548 | 100.99 | 20.64 | 18.438 | 82.7515 | 106.5063 | 80.6024 | 0.004446 | 0.998 | 0.024367 | 24.4162 | -0.093005 | 0.14214 | 0.00065 | 0.04517 | 0.04517 | 0.12933
2019-03-20 13:02 | 1553101307.1802 | 724 | 0.20941 | 0.00062 | 3.9e-05 | 107.9602 | 109.2838 | 124.5192 | 20.9325 | 0.002519 | 0.002441 | -151.9823 | 0.23946 | 8.4014 | 0.003516 | 0.004726 | 0.004502 | 5.6e-05 | 0.001068 | 0.31238 | 1.2954 | 2.5482 | 100.99 | 20.64 | 18.441 | 82.1012 | 106.1452 | 80.2313 | 0.004447 | 0.99905 | 0.024397 | 24.4203 | -0.109 | 0.14245 | 0.000536 | 0.037082 | 0.037082 | 0.15789
2019-03-20 13:02 | 1553101308.1809 | 725 | 0.20943 | 0.000604 | 3.5e-05 | 108.0863 | 109.224 | 124.7772 | 20.9318 | 0.002546 | 0.002378 | -152.0558 | 0.23894 | 8.4032 | 0.003463 | 0.005385 | 0.005243 | 6e-05 | 0.001068 | 0.32173 | 1.3113 | 2.548 | 100.99 | 20.64 | 18.445 | 82.7488 | 104.8802 | 78.7157 | 0.004448 | 1.0034 | 0.024475 | 24.3931 | -0.093658 | 0.14214 | 0.000495 | 0.034445 | 0.034445 | 0.16961
2019-03-20 13:02 | 1553101309.1807 | 726 | 0.20941 | 0.000591 | 3.1e-05 | 108.0087 | 109.3051 | 124.803 | 20.9321 | 0.002603 | 0.002368 | -152.0329 | 0.23887 | 8.4011 | 0.003403 | 0.005111 | 0.004838 | 6e-05 | 0.001074 | 0.31957 | 1.3517 | 2.5481 | 100.99 | 20.64 | 18.444 | 84.1182 | 103.4103 | 78.476 | 0.004447 | 1.004 | 0.024486 | 24.3873 | -0.10672 | 0.14211 | 0.000627 | 0.044476 | 0.044476 | 0.13132
2019-03-20 13:02 | 1553101310.1814 | 727 | 0.20946 | 0.000582 | 2.7e-05 | 107.9828 | 109.301 | 124.9115 | 20.9317 | 0.002535 | 0.0023 | -152.0183 | 0.23926 | 8.4002 | 0.003459 | 0.004667 | 0.004405 | 5.7e-05 | 0.00107 | 0.2995 | 1.294 | 2.5481 | 100.99 | 20.64 | 18.443 | 82.4725 | 104.7833 | 76.8478 | 0.004447 | 1.0087 | 0.024553 | 24.3414 | -0.10853 | 0.14233 | 0.000355 | 0.024644 | 0.024644 | 0.23738
2019-03-20 13:02 | 1553101311.1804 | 728 | 0.20945 | 0.000565 | 2.4e-05 | 108.042 | 109.27 | 125.0652 | 20.9316 | 0.002503 | 0.002353 | -152.0209 | 0.23917 | 8.4005 | 0.003482 | 0.004546 | 0.004336 | 5.7e-05 | 0.001071 | 0.34219 | 1.3103 | 2.5481 | 100.99 | 20.64 | 18.44 | 81.7066 | 105.3226 | 78.1129 | 0.004446 | 1.0051 | 0.024524 | 24.4003 | -0.10109 | 0.14228 | 0.000409 | 0.028785 | 0.028785 | 0.20316
2019-03-20 13:02 | 1553101312.1818 | 729 | 0.2095 | 0.000548 | 2e-05 | 107.9841 | 109.227 | 125.3375 | 20.9316 | 0.002578 | 0.002325 | -152.0047 | 0.23851 | 8.4015 | 0.003511 | 0.004616 | 0.004439 | 5.8e-05 | 0.001074 | 0.32261 | 1.3024 | 2.5479 | 100.99 | 20.64 | 18.438 | 83.5161 | 106.0364 | 77.427 | 0.004446 | 1.007 | 0.024575 | 24.403 | -0.10236 | 0.14189 | 0.000116 | 0.00742 | 0.00742 | 0.78589
2019-03-20 13:02 | 1553101313.1806 | 730 | 0.20948 | 0.000545 | 1.6e-05 | 107.9983 | 109.3163 | 125.4032 | 20.9309 | 0.002534 | 0.002347 | -151.9986 | 0.23978 | 8.4007 | 0.003509 | 0.004658 | 0.004504 | 5.6e-05 | 0.001071 | 0.31758 | 1.2787 | 2.5481 | 100.99 | 20.64 | 18.431 | 82.4672 | 105.9761 | 77.9751 | 0.004444 | 1.0055 | 0.024562 | 24.4285 | -0.1085 | 0.14264 | 0.00029 | 0.020501 | 0.020501 | 0.28597
2019-03-20 13:02 | 1553101314.1805 | 731 | 0.20949 | 0.000531 | 1.4e-05 | 108.0297 | 109.2431 | 125.7684 | 20.9318 | 0.002476 | 0.002346 | -152.0231 | 0.23857 | 8.4016 | 0.003566 | 0.003743 | 0.003437 | 5.8e-05 | 0.00107 | 0.35745 | 1.3333 | 2.548 | 100.99 | 20.64 | 18.427 | 81.054 | 107.367 | 77.9495 | 0.004443 | 1.0055 | 0.024599 | 24.4631 | -0.099904 | 0.14193 | 0.000212 | 0.014922 | 0.014922 | 0.39093
2019-03-20 13:02 | 1553101315.1796 | 732 | 0.20952 | 0.000526 | 1.1e-05 | 108.0433 | 109.2808 | 125.7741 | 20.932 | 0.002432 | 0.002341 | -152.007 | 0.23874 | 8.4002 | 0.003459 | 0.004744 | 0.004552 | 5.9e-05 | 0.001074 | 0.33467 | 1.3108 | 2.5481 | 100.99 | 20.64 | 18.422 | 80.0039 | 104.7773 | 77.825 | 0.004442 | 1.0059 | 0.024604 | 24.4593 | -0.10187 | 0.14202 | 5.6e-05 | 0.00348 | 0.00348 | 1.6774
2019-03-20 13:02 | 1553101316.1797 | 733 | 0.20953 | 0.000516 | 9e-06 | 107.9482 | 109.2831 | 125.8532 | 20.9315 | 0.002503 | 0.002265 | -151.9982 | 0.2386 | 8.4042 | 0.003486 | 0.004648 | 0.004485 | 6e-05 | 0.001074 | 0.34346 | 1.3504 | 2.5481 | 100.99 | 20.64 | 18.413 | 81.7082 | 105.4329 | 75.9997 | 0.00444 | 1.0112 | 0.024676 | 24.4033 | -0.10993 | 0.14194 | 7e-06 | -7.9e-05 | -7.9e-05 | -73.4429
2019-03-20 13:02 | 1553101317.18 | 734 | 0.20951 | 0.000505 | 8e-06 | 107.9815 | 109.3481 | 125.9935 | 20.9317 | 0.002411 | 0.002271 | -152.0215 | 0.23851 | 8.4052 | 0.003549 | 0.004931 | 0.004711 | 5.7e-05 | 0.001076 | 0.29927 | 1.3383 | 2.5481 | 100.99 | 20.64 | 18.408 | 79.5054 | 106.9686 | 76.1424 | 0.004439 | 1.0107 | 0.024684 | 24.4219 | -0.11249 | 0.14189 | 0.000118 | 0.008306 | 0.008306 | 0.70211
2019-03-20 13:02 | 1553101318.1808 | 735 | 0.20954 | 0.000496 | 6e-06 | 108.0685 | 109.3776 | 126.2614 | 20.9308 | 0.002485 | 0.002253 | -152.0259 | 0.23844 | 8.403 | 0.003377 | 0.00439 | 0.004159 | 6e-05 | 0.001081 | 0.34274 | 1.3437 | 2.5481 | 100.99 | 20.64 | 18.4 | 81.2833 | 102.7968 | 75.7105 | 0.004437 | 1.012 | 0.024726 | 24.4328 | -0.1077 | 0.14185 | -3.4e-05 | -0.002905 | -0.002905 | -2.0067
2019-03-20 13:02 | 1553101319.1801 | 736 | 0.2095 | 0.000489 | 5e-06 | 108.0051 | 109.3784 | 126.355 | 20.9318 | 0.002429 | 0.0023 | -152.0265 | 0.23939 | 8.4022 | 0.003373 | 0.005031 | 0.004759 | 5.5e-05 | 0.00108 | 0.36624 | 1.3098 | 2.548 | 100.99 | 20.64 | 18.397 | 79.9477 | 102.6917 | 76.8284 | 0.004436 | 1.0088 | 0.024695 | 24.4809 | -0.11301 | 0.14242 | 0.00023 | 0.016842 | 0.016842 | 0.34753
2019-03-20 13:02 | 1553101320.1798 | 737 | 0.20953 | 0.000478 | 4e-06 | 107.9508 | 109.0125 | 125.9982 | 20.9317 | 0.002488 | 0.002267 | -152.0483 | 0.23904 | 8.4008 | 0.003433 | 0.005513 | 0.005353 | 6e-05 | 0.00108 | 0.33196 | 1.3123 | 2.548 | 100.99 | 20.64 | 18.395 | 81.3589 | 104.1424 | 76.0473 | 0.004436 | 1.011 | 0.024688 | 24.4191 | -0.087539 | 0.14221 | 3.3e-05 | 0.002235 | 0.002235 | 2.6151
2019-03-20 13:02 | 1553101321.1809 | 738 | 0.20953 | 0.000472 | 3e-06 | 107.9595 | 109.2434 | 125.887 | 20.9317 | 0.002426 | 0.002252 | -152.0613 | 0.24023 | 8.4026 | 0.003378 | 0.004673 | 0.004481 | 5.6e-05 | 0.001081 | 0.32211 | 1.3165 | 2.5482 | 100.99 | 20.65 | 18.391 | 79.8689 | 102.8236 | 75.6733 | 0.004437 | 1.0121 | 0.02469 | 24.3952 | -0.10575 | 0.14291 | 9.4e-05 | 0.006818 | 0.006818 | 0.86154
2019-03-20 13:02 | 1553101322.1799 | 739 | 0.20952 | 0.000473 | 2e-06 | 107.9349 | 109.1582 | 125.8536 | 20.9311 | 0.002424 | 0.002181 | -152.0806 | 0.23899 | 8.4033 | 0.003373 | 0.004598 | 0.004419 | 6e-05 | 0.001078 | 0.33229 | 1.3211 | 2.548 | 100.99 | 20.65 | 18.39 | 79.8137 | 102.7005 | 73.9834 | 0.004437 | 1.017 | 0.024747 | 24.3328 | -0.1008 | 0.14217 | 0.000124 | 0.009136 | 0.009136 | 0.6396
2019-03-20 13:02 | 1553101323.1814 | 740 | 0.20954 | 0.000463 | 2e-06 | 107.9084 | 109.3451 | 125.8496 | 20.9314 | 0.002382 | 0.002244 | -152.0573 | 0.23841 | 8.4015 | 0.003415 | 0.004951 | 0.00476 | 5.8e-05 | 0.001082 | 0.32173 | 1.3361 | 2.5481 | 100.99 | 20.65 | 18.389 | 78.8143 | 103.7003 | 75.4812 | 0.004437 | 1.0127 | 0.024694 | 24.3849 | -0.1183 | 0.14183 | 2.1e-05 | 0.00147 | 0.00147 | 3.9644
2019-03-20 13:02 | 1553101324.1811 | 741 | 0.20955 | 0.00046 | 1e-06 | 107.9734 | 109.2864 | 125.5494 | 20.9319 | 0.002352 | 0.002284 | -152.0379 | 0.23883 | 8.4016 | 0.003406 | 0.004937 | 0.004696 | 5.9e-05 | 0.001079 | 0.31094 | 1.3192 | 2.5482 | 100.99 | 20.65 | 18.387 | 78.0923 | 103.4847 | 76.452 | 0.004436 | 1.0098 | 0.02463 | 24.3897 | -0.10811 | 0.14208 | -2.9e-05 | -0.002212 | -0.002212 | -2.6398
2019-03-20 13:02 | 1553101325.1802 | 742 | 0.20955 | 0.000463 | 1e-06 | 108.0129 | 109.255 | 125.8403 | 20.9311 | 0.002419 | 0.002176 | -152.0544 | 0.23893 | 8.4015 | 0.003369 | 0.004778 | 0.004563 | 5.8e-05 | 0.001076 | 0.30541 | 1.3222 | 2.5481 | 100.99 | 20.64 | 18.386 | 79.6935 | 102.5968 | 73.8604 | 0.004433 | 1.0174 | 0.02475 | 24.3272 | -0.10228 | 0.14214 | -7e-06 | -0.000529 | -0.000529 | -11.0409
2019-03-20 13:02 | 1553101326.1802 | 743 | 0.20957 | 0.000455 | 0 | 107.9624 | 109.249 | 125.9966 | 20.9311 | 0.002371 | 0.002227 | -152.0592 | 0.23952 | 8.4018 | 0.003363 | 0.00501 | 0.00482 | 6.1e-05 | 0.001082 | 0.31797 | 1.3253 | 2.5482 | 100.99 | 20.64 | 18.385 | 78.5344 | 102.4562 | 75.0927 | 0.004433 | 1.0138 | 0.024722 | 24.3855 | -0.10597 | 0.14249 | -0.000157 | -0.011739 | -0.011739 | -0.49887
2019-03-20 13:02 | 1553101327.1803 | 744 | 0.20957 | 0.000451 | 0 | 107.9944 | 109.0663 | 125.8318 | 20.9303 | 0.002393 | 0.002184 | -152.0449 | 0.23908 | 8.4031 | 0.00343 | 0.004152 | 0.003935 | 6.3e-05 | 0.001082 | 0.36083 | 1.3424 | 2.5481 | 100.99 | 20.64 | 18.383 | 79.0605 | 104.0726 | 74.0576 | 0.004433 | 1.0168 | 0.024742 | 24.3333 | -0.088337 | 0.14223 | -0.000144 | -0.010798 | -0.010798 | -0.54133
2019-03-20 13:02 | 1553101328.1813 | 745 | 0.20952 | 0.000453 | 0 | 107.9349 | 109.2487 | 125.618 | 20.9305 | 0.002376 | 0.002193 | -152.0312 | 0.23902 | 8.4008 | 0.003378 | 0.004648 | 0.004509 | 5.8e-05 | 0.001082 | 0.35889 | 1.3005 | 2.548 | 100.99 | 20.64 | 18.381 | 78.6498 | 102.8178 | 74.2635 | 0.004432 | 1.0162 | 0.024714 | 24.3198 | -0.10821 | 0.14219 | 0.000163 | 0.012109 | 0.012109 | 0.48262
2019-03-20 13:02 | 1553101329.1794 | 746 | 0.20953 | 0.000453 | 0 | 107.9902 | 109.1149 | 125.7781 | 20.9319 | 0.002376 | 0.002236 | -152.0298 | 0.23885 | 8.4015 | 0.003342 | 0.004154 | 0.00389 | 6e-05 | 0.001078 | 0.32388 | 1.3247 | 2.548 | 100.99 | 20.64 | 18.374 | 78.663 | 101.937 | 75.2955 | 0.004431 | 1.0132 | 0.024693 | 24.3714 | -0.092669 | 0.14209 | 7.2e-05 | 0.005359 | 0.005359 | 1.0898
2019-03-20 13:02 | 1553101330.1807 | 747 | 0.20956 | 0.000454 | 0 | 107.9815 | 109.2752 | 126.0665 | 20.9311 | 0.002313 | 0.002224 | -152.0248 | 0.23875 | 8.4036 | 0.003313 | 0.004404 | 0.004144 | 6.3e-05 | 0.001084 | 0.34562 | 1.3117 | 2.5481 | 100.99 | 20.64 | 18.37 | 77.1583 | 101.2264 | 75.0194 | 0.00443 | 1.014 | 0.024731 | 24.3897 | -0.10653 | 0.14203 | -5.4e-05 | -0.004006 | -0.004006 | -1.4573
2019-03-20 13:02 | 1553101331.1795 | 748 | 0.20948 | 0.000449 | -1e-06 | 107.9492 | 109.1377 | 125.9738 | 20.9314 | 0.002322 | 0.002188 | -152.0235 | 0.23936 | 8.4028 | 0.003352 | 0.003599 | 0.003452 | 6.3e-05 | 0.001084 | 0.3095 | 1.3122 | 2.548 | 100.99 | 20.64 | 18.36 | 77.3708 | 102.1902 | 74.1529 | 0.004427 | 1.0165 | 0.024753 | 24.3504 | -0.09794 | 0.14239 | 0.000399 | 0.029873 | 0.029873 | 0.19591
2019-03-20 13:02 | 1553101332.1805 | 749 | 0.20953 | 0.00045 | 0 | 107.9828 | 109.4101 | 125.7863 | 20.9318 | 0.002274 | 0.002143 | -152.0206 | 0.23802 | 8.4013 | 0.003244 | 0.005137 | 0.004898 | 6e-05 | 0.001089 | 0.35795 | 1.3582 | 2.5479 | 100.99 | 20.64 | 18.354 | 76.2012 | 99.5761 | 73.0613 | 0.004426 | 1.0197 | 0.024774 | 24.294 | -0.11746 | 0.14159 | 0.000123 | 0.009234 | 0.009234 | 0.63023
2019-03-20 13:02 | 1553101333.1806 | 750 | 0.20954 | 0.000446 | -1e-06 | 107.9634 | 109.2576 | 125.7217 | 20.931 | 0.002352 | 0.002173 | -152.0461 | 0.23834 | 8.4042 | 0.003252 | 0.004701 | 0.004542 | 6.1e-05 | 0.001089 | 0.32045 | 1.3699 | 2.5482 | 100.99 | 20.64 | 18.348 | 78.0905 | 99.7557 | 73.7815 | 0.004424 | 1.0176 | 0.024741 | 24.313 | -0.10659 | 0.14179 | 5.2e-05 | 0.003921 | 0.003921 | 1.4862
2019-03-20 13:02 | 1553101334.1796 | 751 | 0.20955 | 0.000444 | -1e-06 | 107.9129 | 109.159 | 125.9562 | 20.9305 | 0.002318 | 0.002178 | -152.0666 | 0.2389 | 8.403 | 0.003296 | 0.004767 | 0.004465 | 5.9e-05 | 0.001092 | 0.34451 | 1.3287 | 2.548 | 100.99 | 20.65 | 18.334 | 77.2678 | 100.824 | 73.9026 | 0.004424 | 1.0173 | 0.02476 | 24.3399 | -0.10269 | 0.14212 | 1.1e-05 | 0.000863 | 0.000863 | 6.7646
2019-03-20 13:02 | 1553101335.1801 | 752 | 0.20954 | 0.000444 | 0 | 107.9369 | 109.0106 | 125.8217 | 20.9311 | 0.002326 | 0.002174 | -152.0555 | 0.23827 | 8.3998 | 0.003185 | 0.005008 | 0.004905 | 6.1e-05 | 0.001093 | 0.32012 | 1.3152 | 2.548 | 100.99 | 20.65 | 18.327 | 77.4577 | 98.1489 | 73.8003 | 0.004422 | 1.0176 | 0.024751 | 24.3233 | -0.088538 | 0.14175 | 8e-05 | 0.006008 | 0.006008 | 0.96969
2019-03-20 13:02 | 1553101336.1799 | 753 | 0.20951 | 0.000446 | 0 | 107.933 | 108.9366 | 125.9104 | 20.9306 | 0.002337 | 0.002162 | -152.0673 | 0.23831 | 8.405 | 0.003232 | 0.004386 | 0.004123 | 6e-05 | 0.001091 | 0.3402 | 1.3256 | 2.548 | 100.99 | 20.65 | 18.316 | 77.7309 | 99.2855 | 73.5186 | 0.004419 | 1.0184 | 0.024769 | 24.322 | -0.082785 | 0.14177 | 0.000251 | 0.018772 | 0.018772 | 0.31041
2019-03-20 13:02 | 1553101337.1799 | 754 | 0.20954 | 0.000444 | 0 | 107.9721 | 109.2147 | 126.2456 | 20.9309 | 0.002318 | 0.002097 | -152.0829 | 0.23958 | 8.4029 | 0.003188 | 0.004573 | 0.004368 | 6e-05 | 0.001091 | 0.34451 | 1.3267 | 2.548 | 100.99 | 20.65 | 18.31 | 77.2588 | 98.2161 | 71.9698 | 0.004418 | 1.023 | 0.024858 | 24.2999 | -0.10235 | 0.14253 | 4.3e-05 | 0.003212 | 0.003212 | 1.824
2019-03-20 13:02 | 1553101338.1809 | 755 | 0.20953 | 0.000446 | 0 | 107.9103 | 109.3193 | 126.3447 | 20.931 | 0.002313 | 0.002132 | -152.0854 | 0.23889 | 8.4024 | 0.003209 | 0.003724 | 0.003516 | 6.1e-05 | 0.001092 | 0.32477 | 1.3444 | 2.5481 | 100.99 | 20.65 | 18.305 | 77.1537 | 98.729 | 72.7923 | 0.004417 | 1.0205 | 0.024838 | 24.3384 | -0.11603 | 0.14211 | 0.000114 | 0.008537 | 0.008537 | 0.68422
2019-03-20 13:02 | 1553101339.1813 | 756 | 0.20953 | 0.000445 | -1e-06 | 107.9103 | 109.1605 | 126.1771 | 20.9305 | 0.002324 | 0.002127 | -152.0861 | 0.23891 | 8.4065 | 0.003121 | 0.004088 | 0.003914 | 6.4e-05 | 0.001092 | 0.34036 | 1.3458 | 2.548 | 100.99 | 20.65 | 18.298 | 77.4142 | 96.5953 | 72.6888 | 0.004415 | 1.0208 | 0.024825 | 24.3186 | -0.10303 | 0.14213 | 9.4e-05 | 0.007053 | 0.007053 | 0.82822
2019-03-20 13:02 | 1553101340.1799 | 757 | 0.20952 | 0.000444 | -1e-06 | 107.9408 | 108.9919 | 126.0142 | 20.9311 | 0.002235 | 0.002184 | -152.0965 | 0.23779 | 8.4028 | 0.003158 | 0.004637 | 0.004368 | 6.2e-05 | 0.001096 | 0.32964 | 1.3438 | 2.548 | 100.99 | 20.65 | 18.297 | 75.2782 | 97.4867 | 74.0421 | 0.004415 | 1.0169 | 0.024761 | 24.3504 | -0.086683 | 0.14146 | 0.000159 | 0.011924 | 0.011924 | 0.48757
2019-03-20 13:02 | 1553101341.1814 | 758 | 0.20953 | 0.000445 | 0 | 107.9188 | 109.1474 | 126.1346 | 20.9298 | 0.002292 | 0.002132 | -152.099 | 0.238 | 8.4022 | 0.003136 | 0.004517 | 0.004322 | 6.2e-05 | 0.001102 | 0.32156 | 1.3287 | 2.5482 | 100.99 | 20.65 | 18.297 | 76.6535 | 96.9479 | 72.8006 | 0.004415 | 1.0205 | 0.024817 | 24.3184 | -0.10126 | 0.14158 | 0.000141 | 0.010583 | 0.010583 | 0.54985
2019-03-20 13:02 | 1553101342.1811 | 759 | 0.20953 | 0.000442 | 0 | 107.91 | 109.1829 | 126.3214 | 20.9304 | 0.002247 | 0.002063 | -152.0946 | 0.23873 | 8.4018 | 0.003225 | 0.004566 | 0.004352 | 6.1e-05 | 0.001099 | 0.33356 | 1.3567 | 2.5481 | 101 | 20.66 | 18.3 | 75.5511 | 99.0994 | 71.1312 | 0.004418 | 1.0255 | 0.024896 | 24.2776 | -0.10489 | 0.14202 | 0.000105 | 0.007885 | 0.007885 | 0.74029
2019-03-20 13:02 | 1553101343.1807 | 760 | 0.20955 | 0.000443 | -1e-06 | 107.9511 | 109.2408 | 126.3914 | 20.9302 | 0.002339 | 0.00209 | -152.0918 | 0.23819 | 8.4018 | 0.003204 | 0.004189 | 0.003986 | 6.2e-05 | 0.001103 | 0.30901 | 1.3381 | 2.548 | 101 | 20.66 | 18.303 | 77.7625 | 98.6098 | 71.7889 | 0.004418 | 1.0235 | 0.024879 | 24.3075 | -0.10623 | 0.1417 | 2.5e-05 | 0.001917 | 0.001917 | 3.0374
2019-03-20 13:02 | 1553101344.1817 | 761 | 0.20954 | 0.000448 | -1e-06 | 107.847 | 109.2345 | 126.59 | 20.93 | 0.00221 | 0.002111 | -152.0844 | 0.23835 | 8.4045 | 0.003263 | 0.004285 | 0.004132 | 6e-05 | 0.001101 | 0.34634 | 1.3633 | 2.548 | 101 | 20.65 | 18.311 | 74.68 | 100.0301 | 72.301 | 0.004418 | 1.022 | 0.02488 | 24.3447 | -0.11434 | 0.1418 | 3.2e-05 | 0.002456 | 0.002456 | 2.3729
2019-03-20 13:02 | 1553101345.1811 | 762 | 0.20955 | 0.000447 | -1e-06 | 107.9152 | 109.156 | 126.6705 | 20.9297 | 0.002221 | 0.002069 | -152.0879 | 0.2396 | 8.4043 | 0.003171 | 0.004588 | 0.004333 | 5.8e-05 | 0.001103 | 0.30287 | 1.347 | 2.548 | 101 | 20.65 | 18.317 | 74.9246 | 97.8137 | 71.2834 | 0.004419 | 1.025 | 0.024925 | 24.3165 | -0.10225 | 0.14253 | -1.9e-05 | -0.001367 | -0.001367 | -4.2851
2019-03-20 13:02 | 1553101346.1793 | 763 | 0.20955 | 0.000443 | -1e-06 | 107.9042 | 108.931 | 126.7076 | 20.9297 | 0.002288 | 0.002106 | -152.0851 | 0.23889 | 8.4056 | 0.003175 | 0.005047 | 0.004867 | 5.9e-05 | 0.001111 | 0.347 | 1.3086 | 2.5481 | 100.99 | 20.65 | 18.33 | 76.5374 | 97.8999 | 72.169 | 0.004423 | 1.0224 | 0.024896 | 24.3513 | -0.084711 | 0.14211 | 3.2e-05 | 0.002419 | 0.002419 | 2.4143
2019-03-20 13:02 | 1553101347.1793 | 764 | 0.20953 | 0.000449 | -1e-06 | 107.8398 | 108.9456 | 126.6421 | 20.9297 | 0.002233 | 0.002105 | -152.0893 | 0.23977 | 8.4031 | 0.003133 | 0.004266 | 0.004054 | 5.8e-05 | 0.001108 | 0.32803 | 1.3098 | 2.5481 | 100.99 | 20.65 | 18.339 | 75.2271 | 96.8796 | 72.1534 | 0.004425 | 1.0224 | 0.02489 | 24.3445 | -0.091244 | 0.14264 | 8.8e-05 | 0.006707 | 0.006707 | 0.87412
2019-03-20 13:02 | 1553101348.1798 | 765 | 0.20951 | 0.000443 | -1e-06 | 107.9097 | 109.0655 | 126.5156 | 20.9291 | 0.00225 | 0.002136 | -152.0979 | 0.23838 | 8.4026 | 0.003233 | 0.004597 | 0.004374 | 5.6e-05 | 0.001112 | 0.31277 | 1.2936 | 2.548 | 100.99 | 20.65 | 18.347 | 75.6435 | 99.2973 | 72.8871 | 0.004427 | 1.0203 | 0.024851 | 24.3582 | -0.095296 | 0.14181 | 0.000255 | 0.019186 | 0.019186 | 0.30379
2019-03-20 13:02 | 1553101349.1794 | 766 | 0.20955 | 0.00044 | -1e-06 | 107.91 | 109.0902 | 126.6442 | 20.9297 | 0.002232 | 0.002076 | -152.1071 | 0.23933 | 8.4008 | 0.00321 | 0.004206 | 0.004 | 5.6e-05 | 0.00111 | 0.33284 | 1.3624 | 2.5482 | 100.99 | 20.65 | 18.365 | 75.1973 | 98.7568 | 71.4425 | 0.004431 | 1.0245 | 0.024916 | 24.3196 | -0.097292 | 0.14238 | 3.9e-05 | 0.002989 | 0.002989 | 1.9581
2019-03-20 13:02 | 1553101350.1809 | 767 | 0.20952 | 0.000437 | -1e-06 | 107.89 | 109.147 | 126.7674 | 20.9283 | 0.002248 | 0.002098 | -152.1081 | 0.23919 | 8.4023 | 0.003182 | 0.004117 | 0.003996 | 5.7e-05 | 0.001111 | 0.33588 | 1.303 | 2.5481 | 100.99 | 20.65 | 18.374 | 75.5842 | 98.064 | 71.9719 | 0.004433 | 1.023 | 0.024909 | 24.3501 | -0.10361 | 0.1423 | 0.000177 | 0.013363 | 0.013363 | 0.43766
2019-03-20 13:03 | 1553101351.1806 | 768 | 0.20955 | 0.000439 | -1e-06 | 107.8835 | 109.144 | 126.7799 | 20.9292 | 0.002171 | 0.002127 | -152.106 | 0.23863 | 8.405 | 0.003136 | 0.005075 | 0.004876 | 6e-05 | 0.001108 | 0.2975 | 1.3578 | 2.548 | 100.99 | 20.65 | 18.383 | 73.744 | 96.948 | 72.6791 | 0.004435 | 1.0209 | 0.024885 | 24.3763 | -0.1039 | 0.14196 | -2e-05 | -0.001456 | -0.001456 | -4.0063
2019-03-20 13:03 | 1553101352.1806 | 769 | 0.20955 | 0.00044 | -1e-06 | 107.9469 | 109.0472 | 126.9202 | 20.9294 | 0.002188 | 0.002108 | -152.1083 | 0.23964 | 8.4018 | 0.003175 | 0.004202 | 0.004015 | 6.1e-05 | 0.001108 | 0.35131 | 1.3709 | 2.5481 | 100.99 | 20.65 | 18.399 | 74.1331 | 97.8887 | 72.2108 | 0.004439 | 1.0223 | 0.024916 | 24.3732 | -0.090712 | 0.14256 | -1.4e-05 | -0.000958 | -0.000958 | -6.1156
2019-03-20 13:03 | 1553101353.1797 | 770 | 0.20954 | 0.000442 | -1e-06 | 107.8842 | 108.9415 | 127.0312 | 20.9291 | 0.00219 | 0.002073 | -152.1041 | 0.23946 | 8.403 | 0.003182 | 0.004375 | 0.00409 | 6.1e-05 | 0.001111 | 0.35867 | 1.3382 | 2.5481 | 100.99 | 20.65 | 18.406 | 74.1937 | 98.0694 | 71.3794 | 0.004441 | 1.0247 | 0.024957 | 24.3545 | -0.087232 | 0.14245 | 4.7e-05 | 0.003549 | 0.003549 | 1.6497
2019-03-20 13:03 | 1553101354.1797 | 771 | 0.20958 | 0.000443 | 0 | 107.8728 | 108.9807 | 127.035 | 20.9296 | 0.002247 | 0.002058 | -152.0768 | 0.23878 | 8.4031 | 0.003226 | 0.005116 | 0.004961 | 6e-05 | 0.001111 | 0.35563 | 1.3103 | 2.548 | 100.99 | 20.66 | 18.419 | 75.5653 | 99.1434 | 71.0165 | 0.004447 | 1.0258 | 0.02497 | 24.3421 | -0.091393 | 0.14205 | -0.000208 | -0.01566 | -0.01566 | -0.37281
2019-03-20 13:03 | 1553101355.1808 | 772 | 0.20955 | 0.00044 | -1e-06 | 107.8512 | 109.0966 | 127.0322 | 20.9291 | 0.002197 | 0.002109 | -152.0869 | 0.23909 | 8.4018 | 0.003222 | 0.004563 | 0.004388 | 6.2e-05 | 0.001103 | 0.29734 | 1.3089 | 2.548 | 100.99 | 20.66 | 18.424 | 74.3558 | 99.0447 | 72.2487 | 0.004448 | 1.0221 | 0.024925 | 24.3853 | -0.10269 | 0.14224 | 1e-05 | 0.000861 | 0.000861 | 6.7906
2019-03-20 13:03 | 1553101356.1809 | 773 | 0.20955 | 0.000439 | -1e-06 | 107.8906 | 108.9956 | 126.7745 | 20.9284 | 0.002181 | 0.002053 | -152.0839 | 0.23845 | 8.4031 | 0.003225 | 0.004785 | 0.004572 | 5.8e-05 | 0.001109 | 0.32084 | 1.3163 | 2.5478 | 100.99 | 20.66 | 18.428 | 73.9652 | 99.1107 | 70.9079 | 0.004449 | 1.0261 | 0.024948 | 24.3132 | -0.091145 | 0.14185 | 2.8e-05 | 0.002213 | 0.002213 | 2.6342
2019-03-20 13:03 | 1553101357.1811 | 774 | 0.20954 | 0.000441 | -1e-06 | 107.8075 | 109.0801 | 126.8226 | 20.9296 | 0.002188 | 0.002017 | -152.0731 | 0.23859 | 8.404 | 0.003214 | 0.004948 | 0.004685 | 6e-05 | 0.001114 | 0.35458 | 1.3356 | 2.5479 | 100.99 | 20.66 | 18.433 | 74.1316 | 98.8342 | 70.034 | 0.00445 | 1.0287 | 0.024985 | 24.287 | -0.10497 | 0.14194 | 6e-05 | 0.004572 | 0.004572 | 1.2758
2019-03-20 13:03 | 1553101358.1809 | 775 | 0.20955 | 0.000441 | -1e-06 | 107.7997 | 109.0895 | 126.893 | 20.9293 | 0.002209 | 0.002064 | -152.0645 | 0.23902 | 8.4048 | 0.003238 | 0.004531 | 0.004362 | 5.9e-05 | 0.001111 | 0.31542 | 1.3728 | 2.5481 | 100.99 | 20.66 | 18.435 | 74.6397 | 99.4222 | 71.1565 | 0.004451 | 1.0254 | 0.024951 | 24.3334 | -0.10638 | 0.14219 | -2e-05 | -0.001393 | -0.001393 | -4.195
2019-03-20 13:03 | 1553101359.1799 | 776 | 0.20956 | 0.000441 | -1e-06 | 107.9042 | 108.9878 | 126.7522 | 20.9292 | 0.002176 | 0.002054 | -152.0516 | 0.2389 | 8.4016 | 0.003205 | 0.004692 | 0.004457 | 6.1e-05 | 0.001108 | 0.31653 | 1.3455 | 2.5482 | 100.99 | 20.66 | 18.433 | 73.8545 | 98.6238 | 70.9304 | 0.00445 | 1.0261 | 0.024945 | 24.3119 | -0.089375 | 0.14212 | -4.5e-05 | -0.003333 | -0.003333 | -1.7524
2019-03-20 13:03 | 1553101360.18 | 777 | 0.20953 | 0.000441 | -2e-06 | 107.9029 | 109.033 | 126.5346 | 20.9295 | 0.00221 | 0.002045 | -152.0469 | 0.23874 | 8.4025 | 0.003168 | 0.004106 | 0.003929 | 5.8e-05 | 0.001108 | 0.33323 | 1.3118 | 2.5481 | 100.99 | 20.67 | 18.431 | 74.6688 | 97.7265 | 70.7117 | 0.004452 | 1.0267 | 0.024932 | 24.2833 | -0.093193 | 0.14202 | 0.000135 | 0.010274 | 0.010274 | 0.56817
2019-03-20 13:03 | 1553101361.1791 | 778 | 0.20957 | 0.000445 | -1e-06 | 107.8166 | 108.9699 | 126.6421 | 20.9282 | 0.002187 | 0.002087 | -152.0576 | 0.23952 | 8.4039 | 0.003172 | 0.004698 | 0.004454 | 6e-05 | 0.001108 | 0.30381 | 1.3492 | 2.5482 | 100.99 | 20.67 | 18.428 | 74.1085 | 97.8224 | 71.7075 | 0.004452 | 1.0237 | 0.024906 | 24.3288 | -0.09517 | 0.14249 | -0.000106 | -0.007928 | -0.007928 | -0.73866
2019-03-20 13:03 | 1553101362.1792 | 779 | 0.20952 | 0.000438 | -1e-06 | 107.8735 | 109.2535 | 126.3301 | 20.9289 | 0.002206 | 0.001994 | -152.0525 | 0.23895 | 8.4007 | 0.003225 | 0.004449 | 0.004191 | 6e-05 | 0.00111 | 0.35347 | 1.3174 | 2.5481 | 100.99 | 20.67 | 18.423 | 74.5723 | 99.1042 | 69.4907 | 0.004451 | 1.0304 | 0.024956 | 24.2206 | -0.1137 | 0.14215 | 0.000191 | 0.014497 | 0.014497 | 0.40301
2019-03-20 13:03 | 1553101363.1791 | 780 | 0.20955 | 0.000444 | -1e-06 | 107.835 | 108.9512 | 126.1372 | 20.9292 | 0.002198 | 0.002073 | -152.0512 | 0.23933 | 8.4021 | 0.00322 | 0.004645 | 0.004449 | 6.2e-05 | 0.00111 | 0.33931 | 1.333 | 2.548 | 100.99 | 20.67 | 18.421 | 74.3837 | 98.9885 | 71.3794 | 0.00445 | 1.0247 | 0.024869 | 24.2687 | -0.092107 | 0.14238 | -2.5e-05 | -0.001807 | -0.001807 | -3.238
2019-03-20 13:03 | 1553101364.1802 | 781 | 0.20954 | 0.000443 | -1e-06 | 107.8767 | 109.2431 | 125.6917 | 20.9286 | 0.002081 | 0.002089 | -152.0769 | 0.23995 | 8.4007 | 0.003198 | 0.005124 | 0.004855 | 6.1e-05 | 0.001111 | 0.33196 | 1.3047 | 2.5479 | 100.99 | 20.67 | 18.418 | 71.5723 | 98.4524 | 71.7571 | 0.004449 | 1.0236 | 0.024811 | 24.2391 | -0.11258 | 0.14275 | 4.7e-05 | 0.00353 | 0.00353 | 1.6621
2019-03-20 13:03 | 1553101365.1802 | 782 | 0.20953 | 0.00044 | -1e-06 | 107.8101 | 109.0487 | 125.5447 | 20.9288 | 0.002111 | 0.002019 | -152.0878 | 0.23899 | 8.4032 | 0.00322 | 0.005 | 0.004863 | 6.1e-05 | 0.001111 | 0.3449 | 1.31 | 2.5479 | 100.99 | 20.67 | 18.41 | 72.2984 | 98.991 | 70.0895 | 0.004447 | 1.0286 | 0.024857 | 24.1662 | -0.10218 | 0.14217 | 0.000108 | 0.008168 | 0.008168 | 0.71538
2019-03-20 13:03 | 1553101366.1793 | 783 | 0.20953 | 0.000443 | -1e-06 | 107.8308 | 109.1104 | 125.4727 | 20.9278 | 0.002181 | 0.002057 | -152.0825 | 0.23882 | 8.4033 | 0.003261 | 0.004169 | 0.003925 | 6.2e-05 | 0.001113 | 0.32244 | 1.3121 | 2.548 | 100.99 | 20.67 | 18.406 | 73.9732 | 99.9745 | 70.9969 | 0.004446 | 1.0259 | 0.024817 | 24.1912 | -0.10552 | 0.14208 | 0.00015 | 0.011315 | 0.011315 | 0.51608
2019-03-20 13:03 | 1553101367.1804 | 784 | 0.20949 | 0.000443 | -1e-06 | 107.8272 | 109.0966 | 125.0952 | 20.9288 | 0.002164 | 0.002019 | -152.0917 | 0.23953 | 8.4029 | 0.003175 | 0.004532 | 0.004359 | 6.1e-05 | 0.001113 | 0.34739 | 1.3396 | 2.5481 | 100.99 | 20.67 | 18.396 | 73.5689 | 97.9045 | 70.0724 | 0.004444 | 1.0286 | 0.024813 | 24.1223 | -0.10468 | 0.1425 | 0.000395 | 0.029605 | 0.029605 | 0.19782
2019-03-20 13:03 | 1553101368.1805 | 785 | 0.2096 | 0.000445 | -1e-06 | 107.8803 | 109.0439 | 124.9868 | 20.9293 | 0.002121 | 0.002032 | -152.1056 | 0.23868 | 8.4017 | 0.003202 | 0.004421 | 0.00422 | 5.9e-05 | 0.001115 | 0.34578 | 1.2886 | 2.5481 | 100.99 | 20.67 | 18.39 | 72.5327 | 98.5439 | 70.3919 | 0.004443 | 1.0277 | 0.02479 | 24.1231 | -0.095958 | 0.14199 | -0.000321 | -0.024014 | -0.024014 | -0.24301
2019-03-20 13:03 | 1553101369.1805 | 786 | 0.20958 | 0.000445 | -1e-06 | 107.8428 | 109.1044 | 124.9784 | 20.9279 | 0.002181 | 0.00205 | -152.1294 | 0.23991 | 8.4028 | 0.003271 | 0.004187 | 0.00403 | 6.2e-05 | 0.001113 | 0.30469 | 1.2932 | 2.5479 | 100.99 | 20.67 | 18.383 | 73.9734 | 100.2196 | 70.8217 | 0.004441 | 1.0264 | 0.024774 | 24.1374 | -0.10404 | 0.14272 | -0.000155 | -0.011522 | -0.011522 | -0.5091
2019-03-20 13:03 | 1553101370.1795 | 787 | 0.20956 | 0.000445 | 0 | 107.8347 | 109.0326 | 125.0769 | 20.9285 | 0.002066 | 0.002047 | -152.1185 | 0.2388 | 8.4005 | 0.003266 | 0.004283 | 0.00408 | 6.3e-05 | 0.001114 | 0.3298 | 1.3311 | 2.548 | 100.99 | 20.67 | 18.371 | 71.2159 | 100.1094 | 70.7559 | 0.004438 | 1.0266 | 0.024786 | 24.1446 | -0.09882 | 0.14206 | -8e-05 | -0.005962 | -0.005962 | -0.97925
2019-03-20 13:03 | 1553101371.1806 | 788 | 0.20958 | 0.000445 | -1e-06 | 107.8143 | 109.1138 | 124.9124 | 20.9275 | 0.002127 | 0.002047 | -152.1191 | 0.23921 | 8.402 | 0.003271 | 0.005285 | 0.005077 | 6.1e-05 | 0.001111 | 0.3131 | 1.3267 | 2.5479 | 100.99 | 20.67 | 18.365 | 72.6851 | 100.208 | 70.7525 | 0.004437 | 1.0266 | 0.02477 | 24.1286 | -0.10716 | 0.1423 | -0.000194 | -0.014451 | -0.014451 | -0.40472
2019-03-20 13:03 | 1553101372.1807 | 789 | 0.20959 | 0.000446 | -1e-06 | 107.8741 | 109.0061 | 125.1213 | 20.927 | 0.002197 | 0.002012 | -152.1278 | 0.23894 | 8.4031 | 0.003225 | 0.004934 | 0.004728 | 6e-05 | 0.001113 | 0.33892 | 1.3151 | 2.548 | 100.99 | 20.67 | 18.352 | 74.3607 | 99.0951 | 69.9116 | 0.004433 | 1.0291 | 0.024821 | 24.1192 | -0.093371 | 0.14215 | -0.000226 | -0.01689 | -0.01689 | -0.34591
2019-03-20 13:03 | 1553101373.1798 | 790 | 0.20956 | 0.000448 | -1e-06 | 107.8609 | 109.2106 | 125.1197 | 20.9273 | 0.002139 | 0.002014 | -152.1258 | 0.23898 | 8.3995 | 0.003138 | 0.004281 | 0.004123 | 6.2e-05 | 0.001107 | 0.31404 | 1.3236 | 2.5481 | 100.99 | 20.68 | 18.345 | 72.9625 | 96.9946 | 69.9501 | 0.004434 | 1.029 | 0.02482 | 24.1204 | -0.11123 | 0.14217 | -4e-05 | -0.002922 | -0.002922 | -1.9996
2019-03-20 13:03 | 1553101374.1804 | 791 | 0.20955 | 0.000446 | -1e-06 | 107.8334 | 108.9848 | 125.187 | 20.9275 | 0.002119 | 0.002018 | -152.1245 | 0.23898 | 8.4013 | 0.003162 | 0.004512 | 0.004223 | 6.2e-05 | 0.001103 | 0.33644 | 1.3764 | 2.5482 | 100.99 | 20.68 | 18.337 | 72.492 | 97.589 | 70.0671 | 0.004433 | 1.0286 | 0.024822 | 24.131 | -0.095002 | 0.14217 | 1.4e-05 | 0.001125 | 0.001125 | 5.1957
2019-03-20 13:03 | 1553101375.1799 | 792 | 0.20956 | 0.00045 | -2e-06 | 107.8538 | 108.9456 | 125.1537 | 20.9281 | 0.002167 | 0.002054 | -152.1077 | 0.23866 | 8.403 | 0.003266 | 0.00445 | 0.004265 | 5.8e-05 | 0.001105 | 0.33964 | 1.3782 | 2.548 | 100.99 | 20.67 | 18.318 | 73.6307 | 100.0901 | 70.9115 | 0.004425 | 1.0261 | 0.024788 | 24.1575 | -0.090091 | 0.14198 | -6.4e-05 | -0.004691 | -0.004691 | -1.244
2019-03-20 13:03 | 1553101376.1792 | 793 | 0.20956 | 0.000447 | -1e-06 | 107.8502 | 109.1541 | 125.3804 | 20.9274 | 0.002142 | 0.002025 | -152.1057 | 0.23855 | 8.404 | 0.003206 | 0.004885 | 0.00468 | 5.6e-05 | 0.001097 | 0.30397 | 1.3431 | 2.5481 | 100.99 | 20.67 | 18.31 | 73.0282 | 98.651 | 70.2355 | 0.004423 | 1.0281 | 0.024835 | 24.1556 | -0.10749 | 0.14192 | -7.5e-05 | -0.005569 | -0.005569 | -1.0474
2019-03-20 13:03 | 1553101377.1792 | 794 | 0.20957 | 0.000446 | 0 | 107.8877 | 109.0659 | 125.6973 | 20.9271 | 0.002089 | 0.001994 | -152.1011 | 0.23872 | 8.4033 | 0.003218 | 0.004329 | 0.00413 | 5.6e-05 | 0.001104 | 0.35225 | 1.3483 | 2.5482 | 100.99 | 20.67 | 18.302 | 71.7626 | 98.9362 | 69.4681 | 0.004421 | 1.0304 | 0.024894 | 24.159 | -0.09715 | 0.14202 | -8.6e-05 | -0.006467 | -0.006467 | -0.9026
2019-03-20 13:03 | 1553101378.1789 | 795 | 0.20956 | 0.000449 | -2e-06 | 107.8871 | 109.0027 | 125.8428 | 20.9275 | 0.002142 | 0.002071 | -152.0993 | 0.23855 | 8.4012 | 0.003216 | 0.004541 | 0.004389 | 5.6e-05 | 0.001102 | 0.32748 | 1.313 | 2.5481 | 100.99 | 20.67 | 18.285 | 73.0328 | 98.895 | 71.3249 | 0.004417 | 1.0249 | 0.024841 | 24.2384 | -0.092022 | 0.14191 | -5.1e-05 | -0.003658 | -0.003658 | -1.5943
2019-03-20 13:03 | 1553101379.179 | 796 | 0.20955 | 0.000447 | -1e-06 | 107.8172 | 109.0689 | 126.0285 | 20.9281 | 0.00218 | 0.002036 | -152.0893 | 0.23925 | 8.4039 | 0.003172 | 0.004897 | 0.004681 | 5.6e-05 | 0.001102 | 0.33859 | 1.3312 | 2.5479 | 100.99 | 20.67 | 18.276 | 73.961 | 97.8364 | 70.4873 | 0.004415 | 1.0274 | 0.02489 | 24.2268 | -0.10324 | 0.14233 | -1.1e-05 | -0.000774 | -0.000774 | -7.5609
2019-03-20 13:03 | 1553101380.1801 | 797 | 0.20957 | 0.000449 | -1e-06 | 107.8583 | 108.9471 | 126.3177 | 20.9273 | 0.002132 | 0.002043 | -152.1063 | 0.23921 | 8.4039 | 0.003181 | 0.004705 | 0.004549 | 5.4e-05 | 0.001099 | 0.31581 | 1.3414 | 2.5482 | 101 | 20.67 | 18.261 | 72.7919 | 98.0338 | 70.6456 | 0.004411 | 1.0269 | 0.024913 | 24.2602 | -0.089838 | 0.14231 | -9.6e-05 | -0.007158 | -0.007158 | -0.81707
2019-03-20 13:03 | 1553101381.1797 | 798 | 0.20956 | 0.000448 | -1e-06 | 107.8919 | 109.2846 | 126.5787 | 20.9271 | 0.002172 | 0.002045 | -152.0933 | 0.23909 | 8.4021 | 0.003247 | 0.004379 | 0.004224 | 5.4e-05 | 0.001106 | 0.30574 | 1.3255 | 2.548 | 101 | 20.67 | 18.255 | 73.7607 | 99.6467 | 70.7038 | 0.00441 | 1.0267 | 0.024936 | 24.2873 | -0.11471 | 0.14223 | -4.1e-05 | -0.002999 | -0.002999 | -1.9494
2019-03-20 13:03 | 1553101382.1802 | 799 | 0.20958 | 0.00044 | -1e-06 | 107.8819 | 109.1377 | 126.5923 | 20.9276 | 0.002115 | 0.00204 | -152.0822 | 0.23918 | 8.4038 | 0.003184 | 0.00451 | 0.004324 | 5.7e-05 | 0.001108 | 0.29336 | 1.2857 | 2.548 | 101 | 20.67 | 18.25 | 72.3845 | 98.1204 | 70.5828 | 0.004408 | 1.0271 | 0.024942 | 24.2843 | -0.10352 | 0.14229 | -0.000173 | -0.01292 | -0.01292 | -0.45265
2019-03-20 13:03 | 1553101383.1793 | 800 | 0.20955 | 0.000446 | -2e-06 | 107.8945 | 109.0992 | 126.8125 | 20.9277 | 0.002139 | 0.002006 | -152.0868 | 0.23914 | 8.4018 | 0.003237 | 0.004839 | 0.00462 | 5.6e-05 | 0.00111 | 0.32853 | 1.3071 | 2.5479 | 101 | 20.67 | 18.241 | 72.9691 | 99.4081 | 69.7664 | 0.004406 | 1.0295 | 0.024994 | 24.2765 | -0.099315 | 0.14226 | -1.4e-05 | -0.000939 | -0.000939 | -6.2242
2019-03-20 13:03 | 1553101384.1803 | 801 | 0.20956 | 0.00044 | -1e-06 | 107.9068 | 109.1978 | 127.1446 | 20.9273 | 0.00208 | 0.002016 | -152.0684 | 0.23923 | 8.4018 | 0.003253 | 0.004385 | 0.004124 | 5.4e-05 | 0.00111 | 0.31542 | 1.2851 | 2.548 | 101 | 20.67 | 18.239 | 71.5352 | 99.774 | 69.9977 | 0.004406 | 1.0288 | 0.025018 | 24.3165 | -0.10638 | 0.14232 | -3.3e-05 | -0.002391 | -0.002391 | -2.4466
2019-03-20 13:03 | 1553101385.1794 | 802 | 0.20955 | 0.000442 | -1e-06 | 107.8735 | 109.1728 | 127.2195 | 20.9265 | 0.002049 | 0.001947 | -152.0592 | 0.23943 | 8.4022 | 0.003268 | 0.00391 | 0.003828 | 5.8e-05 | 0.001112 | 0.30934 | 1.3186 | 2.5481 | 101 | 20.67 | 18.237 | 70.7999 | 100.1413 | 68.3552 | 0.004405 | 1.0338 | 0.025085 | 24.2654 | -0.10709 | 0.14244 | 1.4e-05 | 0.001124 | 0.001124 | 5.206
2019-03-20 13:03 | 1553101386.1804 | 803 | 0.20956 | 0.000448 | -1e-06 | 107.8395 | 108.9863 | 127.3279 | 20.9269 | 0.002107 | 0.002052 | -152.0866 | 0.23906 | 8.4009 | 0.003167 | 0.004936 | 0.004736 | 5.9e-05 | 0.001115 | 0.29574 | 1.3928 | 2.5482 | 101 | 20.67 | 18.239 | 72.1973 | 97.709 | 70.8728 | 0.004406 | 1.0262 | 0.025004 | 24.365 | -0.094615 | 0.14222 | -4.2e-05 | -0.003084 | -0.003084 | -1.8955
2019-03-20 13:03 | 1553101387.1804 | 804 | 0.20954 | 0.000437 | -1e-06 | 107.9084 | 109.1549 | 127.2545 | 20.9271 | 0.002091 | 0.002013 | -152.1036 | 0.23915 | 8.4042 | 0.003223 | 0.004408 | 0.004192 | 5.4e-05 | 0.001112 | 0.3178 | 1.3783 | 2.5479 | 100.99 | 20.68 | 18.241 | 71.8032 | 99.0616 | 69.942 | 0.004409 | 1.029 | 0.025031 | 24.325 | -0.10273 | 0.14227 | 6e-05 | 0.004611 | 0.004611 | 1.2682
2019-03-20 13:03 | 1553101388.1809 | 805 | 0.20955 | 0.00044 | -1e-06 | 107.8625 | 109.2576 | 127.346 | 20.9262 | 0.002105 | 0.001953 | -152.1092 | 0.23939 | 8.4027 | 0.003209 | 0.004734 | 0.004563 | 5.7e-05 | 0.001109 | 0.32875 | 1.3692 | 2.548 | 100.99 | 20.68 | 18.25 | 72.1445 | 98.7061 | 68.4905 | 0.004411 | 1.0334 | 0.025093 | 24.2822 | -0.11495 | 0.14241 | -1.4e-05 | -0.000976 | -0.000976 | -5.9959
2019-03-20 13:03 | 1553101389.18 | 806 | 0.20957 | 0.00044 | -1e-06 | 107.8654 | 109.3466 | 127.346 | 20.9259 | 0.002133 | 0.002052 | -152.1054 | 0.23908 | 8.4043 | 0.003248 | 0.004466 | 0.004162 | 5.9e-05 | 0.001104 | 0.30326 | 1.3822 | 2.5481 | 100.99 | 20.68 | 18.256 | 72.8162 | 99.666 | 70.8687 | 0.004413 | 1.0262 | 0.025006 | 24.3666 | -0.12199 | 0.14223 | -8.3e-05 | -0.006222 | -0.006222 | -0.9395
2019-03-20 13:03 | 1553101390.1796 | 807 | 0.20962 | 0.000441 | -1e-06 | 107.8495 | 109.0891 | 127.2143 | 20.9261 | 0.002033 | 0.001979 | -152.1085 | 0.23913 | 8.4042 | 0.003211 | 0.004064 | 0.003862 | 5.9e-05 | 0.00111 | 0.32372 | 1.3281 | 2.5481 | 100.99 | 20.68 | 18.27 | 70.4145 | 98.7633 | 69.1238 | 0.004416 | 1.0315 | 0.025057 | 24.2922 | -0.10222 | 0.14226 | -0.000427 | -0.032267 | -0.032267 | -0.1812
2019-03-20 13:03 | 1553101391.1807 | 808 | 0.20956 | 0.000447 | -2e-06 | 107.8932 | 109.2064 | 127.1103 | 20.9264 | 0.002107 | 0.001996 | -152.1256 | 0.23881 | 8.4046 | 0.00322 | 0.004605 | 0.004405 | 6e-05 | 0.001112 | 0.33843 | 1.3333 | 2.5481 | 100.99 | 20.68 | 18.27 | 72.1947 | 98.9754 | 69.5226 | 0.004416 | 1.0303 | 0.025032 | 24.2964 | -0.10821 | 0.14207 | -6e-05 | -0.00441 | -0.00441 | -1.3239
2019-03-20 13:03 | 1553101392.1786 | 809 | 0.20959 | 0.00044 | -1e-06 | 107.9113 | 109.1855 | 127.3254 | 20.9263 | 0.002058 | 0.001949 | -152.1476 | 0.23937 | 8.4016 | 0.003321 | 0.004562 | 0.004384 | 5.9e-05 | 0.001112 | 0.31421 | 1.2982 | 2.548 | 100.99 | 20.68 | 18.286 | 71.0126 | 101.4172 | 68.404 | 0.00442 | 1.0336 | 0.025094 | 24.2772 | -0.105 | 0.1424 | -0.000212 | -0.016011 | -0.016011 | -0.36555
2019-03-20 13:03 | 1553101393.1801 | 810 | 0.20951 | 0.000438 | -1e-06 | 107.9715 | 109.0614 | 127.3493 | 20.9264 | 0.001996 | 0.001985 | -152.1523 | 0.23758 | 8.4005 | 0.003278 | 0.003733 | 0.003423 | 6e-05 | 0.001111 | 0.3235 | 1.2818 | 2.548 | 100.99 | 20.68 | 18.292 | 69.5304 | 100.3821 | 69.2715 | 0.004422 | 1.031 | 0.025065 | 24.3103 | -0.089841 | 0.14134 | 0.00025 | 0.018963 | 0.018963 | 0.30634
2019-03-20 13:03 | 1553101394.1799 | 811 | 0.20953 | 0.000444 | -1e-06 | 107.9249 | 109.1298 | 127.242 | 20.9259 | 0.002117 | 0.002004 | -152.1583 | 0.23836 | 8.4028 | 0.003377 | 0.004571 | 0.004354 | 5.7e-05 | 0.001114 | 0.31797 | 1.324 | 2.548 | 100.99 | 20.68 | 18.306 | 72.4234 | 102.7888 | 69.7187 | 0.004425 | 1.0297 | 0.025038 | 24.3159 | -0.099307 | 0.1418 | 0.000163 | 0.012372 | 0.012372 | 0.47108
2019-03-20 13:03 | 1553101395.1799 | 812 | 0.20958 | 0.000443 | -1e-06 | 107.9841 | 109.0334 | 127.2916 | 20.9264 | 0.002065 | 0.001977 | -152.1691 | 0.23736 | 8.4043 | 0.003284 | 0.004645 | 0.004399 | 6.3e-05 | 0.001114 | 0.3366 | 1.3039 | 2.5481 | 100.99 | 20.67 | 18.312 | 71.1905 | 100.5202 | 69.0772 | 0.004424 | 1.0316 | 0.025066 | 24.2979 | -0.086496 | 0.14121 | -0.000194 | -0.014605 | -0.014605 | -0.39737
2019-03-20 13:03 | 1553101396.1793 | 813 | 0.20957 | 0.000441 | -1e-06 | 107.9576 | 109.2165 | 127.2732 | 20.9262 | 0.002116 | 0.001987 | -152.1495 | 0.23945 | 8.4031 | 0.003339 | 0.004039 | 0.003823 | 5.8e-05 | 0.001118 | 0.32493 | 1.3117 | 2.5481 | 100.99 | 20.68 | 18.319 | 72.4088 | 101.8576 | 69.3133 | 0.004428 | 1.0309 | 0.025056 | 24.3045 | -0.10371 | 0.14245 | -6.8e-05 | -0.005104 | -0.005104 | -1.147
2019-03-20 13:03 | 1553101397.1789 | 814 | 0.20957 | 0.000446 | -2e-06 | 107.9443 | 109.0924 | 127.5906 | 20.9267 | 0.002092 | 0.001938 | -152.1313 | 0.23878 | 8.4036 | 0.003303 | 0.004151 | 0.003972 | 6.1e-05 | 0.001117 | 0.3095 | 1.3219 | 2.5481 | 100.99 | 20.67 | 18.332 | 71.8393 | 100.9968 | 68.1316 | 0.004429 | 1.0345 | 0.02513 | 24.2928 | -0.094635 | 0.14205 | -0.000112 | -0.008366 | -0.008366 | -0.69787
2019-03-20 13:03 | 1553101398.1794 | 815 | 0.20954 | 0.000444 | -1e-06 | 107.966 | 109.107 | 127.4052 | 20.9264 | 0.002077 | 0.001995 | -152.1297 | 0.23932 | 8.4009 | 0.003343 | 0.004549 | 0.00437 | 6.2e-05 | 0.001111 | 0.29717 | 1.305 | 2.548 | 100.99 | 20.68 | 18.338 | 71.4836 | 101.9606 | 69.5118 | 0.004433 | 1.0303 | 0.025061 | 24.3241 | -0.094035 | 0.14237 | 7.7e-05 | 0.005892 | 0.005892 | 0.99304
2019-03-20 13:03 | 1553101399.1792 | 816 | 0.20954 | 0.000443 | -2e-06 | 107.9669 | 109.045 | 127.2232 | 20.9268 | 0.002065 | 0.002 | -152.1458 | 0.23848 | 8.402 | 0.003308 | 0.004152 | 0.003912 | 6.2e-05 | 0.001108 | 0.33892 | 1.3471 | 2.5481 | 100.99 | 20.67 | 18.343 | 71.187 | 101.1067 | 69.6165 | 0.004431 | 1.03 | 0.02504 | 24.3105 | -0.088867 | 0.14187 | 6.2e-05 | 0.00478 | 0.00478 | 1.2199
2019-03-20 13:03 | 1553101400.1808 | 817 | 0.20953 | 0.00044 | -2e-06 | 107.9589 | 109.0614 | 127.3218 | 20.9266 | 0.002073 | 0.001937 | -152.1487 | 0.23849 | 8.4029 | 0.003298 | 0.004444 | 0.004258 | 5.9e-05 | 0.001106 | 0.3131 | 1.3539 | 2.5482 | 100.99 | 20.68 | 18.353 | 71.3731 | 100.8813 | 68.1041 | 0.004436 | 1.0346 | 0.025105 | 24.2662 | -0.090886 | 0.14188 | 0.000127 | 0.009724 | 0.009724 | 0.59964
2019-03-20 13:03 | 1553101401.1806 | 818 | 0.20952 | 0.000444 | -2e-06 | 107.9197 | 109.0771 | 127.3718 | 20.9263 | 0.002071 | 0.001941 | -152.1376 | 0.23809 | 8.4016 | 0.00328 | 0.004652 | 0.004498 | 6.1e-05 | 0.001104 | 0.34235 | 1.3352 | 2.548 | 100.99 | 20.68 | 18.358 | 71.3238 | 100.4376 | 68.199 | 0.004438 | 1.0343 | 0.025106 | 24.2743 | -0.095416 | 0.14164 | 0.000214 | 0.016312 | 0.016312 | 0.35687
2019-03-20 13:03 | 1553101402.18 | 819 | 0.20955 | 0.000444 | -1e-06 | 107.8502 | 109.3275 | 127.3235 | 20.9273 | 0.001995 | 0.001966 | -152.1282 | 0.23875 | 8.4028 | 0.003355 | 0.004507 | 0.004258 | 6.3e-05 | 0.001108 | 0.32278 | 1.3165 | 2.5482 | 100.99 | 20.68 | 18.371 | 69.5156 | 102.2411 | 68.8022 | 0.004441 | 1.0324 | 0.025079 | 24.2912 | -0.12169 | 0.14203 | -3e-06 | -0.00016 | -0.00016 | -36.4224
2019-03-20 13:03 | 1553101403.1794 | 820 | 0.20955 | 0.000441 | -2e-06 | 107.9032 | 109.3302 | 127.499 | 20.9261 | 0.002077 | 0.001976 | -152.1097 | 0.23783 | 8.4011 | 0.003431 | 0.005092 | 0.004958 | 6.2e-05 | 0.001106 | 0.31205 | 1.316 | 2.5479 | 100.99 | 20.68 | 18.378 | 71.481 | 104.0993 | 69.045 | 0.004442 | 1.0317 | 0.025088 | 24.3165 | -0.11751 | 0.14149 | 2e-05 | 0.001637 | 0.001637 | 3.5531
2019-03-20 13:03 | 1553101404.1804 | 821 | 0.20959 | 0.000441 | -1e-06 | 107.855 | 109.1829 | 127.6242 | 20.9272 | 0.002062 | 0.001975 | -152.1143 | 0.239 | 8.4007 | 0.003343 | 0.004751 | 0.004554 | 6e-05 | 0.001108 | 0.34219 | 1.3117 | 2.5481 | 100.99 | 20.68 | 18.386 | 71.1093 | 101.9694 | 69.0251 | 0.004444 | 1.0318 | 0.025101 | 24.3278 | -0.10945 | 0.14218 | -0.000224 | -0.016866 | -0.016866 | -0.34647
2019-03-20 13:03 | 1553101405.1795 | 822 | 0.20957 | 0.000441 | -1e-06 | 108.031 | 109.1616 | 127.6237 | 20.9263 | 0.002098 | 0.002012 | -152.119 | 0.23868 | 8.4033 | 0.003432 | 0.004785 | 0.004613 | 6.3e-05 | 0.001108 | 0.33052 | 1.3131 | 2.5481 | 100.99 | 20.68 | 18.403 | 71.9731 | 104.1247 | 69.9078 | 0.004448 | 1.0291 | 0.025068 | 24.3591 | -0.093123 | 0.14199 | -8.1e-05 | -0.006008 | -0.006008 | -0.97128
2019-03-20 13:03 | 1553101406.1805 | 823 | 0.20954 | 0.000441 | -1e-06 | 107.9443 | 108.9822 | 127.5108 | 20.926 | 0.002044 | 0.001988 | -152.1148 | 0.23832 | 8.4039 | 0.003369 | 0.005172 | 0.005057 | 6.4e-05 | 0.001106 | 0.26792 | 1.312 | 2.5482 | 100.99 | 20.68 | 18.41 | 70.6725 | 102.5871 | 69.3242 | 0.00445 | 1.0309 | 0.025078 | 24.3276 | -0.085591 | 0.14178 | 9.8e-05 | 0.007499 | 0.007499 | 0.77706
2019-03-20 13:03 | 1553101407.1789 | 824 | 0.20953 | 0.00044 | -2e-06 | 107.8796 | 109.0618 | 127.3523 | 20.9275 | 0.002095 | 0.001931 | -152.1079 | 0.23834 | 8.4001 | 0.003311 | 0.004995 | 0.004727 | 6.2e-05 | 0.001106 | 0.32421 | 1.3328 | 2.5482 | 100.99 | 20.68 | 18.423 | 71.9026 | 101.1907 | 67.9694 | 0.004453 | 1.035 | 0.025113 | 24.2643 | -0.097483 | 0.14179 | 0.00016 | 0.012253 | 0.012253 | 0.47561
2019-03-20 13:03 | 1553101408.1796 | 825 | 0.20952 | 0.000437 | -2e-06 | 107.976 | 109.1878 | 127.5859 | 20.9261 | 0.002076 | 0.002007 | -152.1069 | 0.2382 | 8.4027 | 0.003348 | 0.004816 | 0.004569 | 6.5e-05 | 0.001108 | 0.30182 | 1.335 | 2.548 | 100.99 | 20.68 | 18.43 | 71.4584 | 102.0719 | 69.7922 | 0.004455 | 1.0295 | 0.025069 | 24.3514 | -0.09982 | 0.14171 | 0.000218 | 0.016563 | 0.016563 | 0.35163
2019-03-20 13:03 | 1553101409.1786 | 826 | 0.20953 | 0.00044 | -1e-06 | 107.9152 | 109.0741 | 127.9014 | 20.9261 | 0.002116 | 0.001955 | -152.0856 | 0.239 | 8.4034 | 0.003395 | 0.005606 | 0.005488 | 6.3e-05 | 0.001108 | 0.32764 | 1.3393 | 2.5481 | 100.99 | 20.68 | 18.435 | 72.4071 | 103.2277 | 68.5342 | 0.004456 | 1.0333 | 0.025146 | 24.3367 | -0.095546 | 0.14218 | 0.000129 | 0.009828 | 0.009828 | 0.59458
2019-03-20 13:03 | 1553101410.1787 | 827 | 0.20956 | 0.000439 | -1e-06 | 107.8842 | 109.2453 | 127.7293 | 20.9267 | 0.002069 | 0.001975 | -152.0809 | 0.23834 | 8.4015 | 0.003414 | 0.004685 | 0.004369 | 6.1e-05 | 0.001106 | 0.31902 | 1.3125 | 2.5481 | 100.99 | 20.68 | 18.446 | 71.2766 | 103.6895 | 69.023 | 0.004459 | 1.0318 | 0.025111 | 24.3377 | -0.11215 | 0.14179 | -4.9e-05 | -0.003644 | -0.003644 | -1.5991
2019-03-20 13:04 | 1553101411.1787 | 828 | 0.20949 | 0.000441 | -1e-06 | 107.9158 | 109.1119 | 127.7896 | 20.9266 | 0.002022 | 0.001934 | -152.0831 | 0.23911 | 8.4034 | 0.003436 | 0.004785 | 0.004603 | 6.3e-05 | 0.001107 | 0.34363 | 1.3008 | 2.5481 | 100.99 | 20.68 | 18.451 | 70.1482 | 104.2091 | 68.0339 | 0.00446 | 1.0348 | 0.025153 | 24.3082 | -0.098587 | 0.14225 | 0.000401 | 0.030499 | 0.030499 | 0.19169
2019-03-20 13:04 | 1553101412.1802 | 829 | 0.20955 | 0.000445 | -1e-06 | 107.8424 | 109.0797 | 127.7084 | 20.9262 | 0.002087 | 0.002007 | -152.078 | 0.23866 | 8.4016 | 0.003371 | 0.004979 | 0.004695 | 6.3e-05 | 0.00111 | 0.31708 | 1.2915 | 2.548 | 100.99 | 20.68 | 18.455 | 71.7048 | 102.6393 | 69.7938 | 0.004461 | 1.0295 | 0.025081 | 24.3631 | -0.10204 | 0.14198 | -5e-06 | -0.000352 | -0.000352 | -16.5799
2019-03-20 13:04 | 1553101413.1799 | 830 | 0.20953 | 0.000443 | -1e-06 | 107.9327 | 109.0771 | 127.3453 | 20.9258 | 0.00201 | 0.001978 | -152.1154 | 0.23837 | 8.4002 | 0.003396 | 0.005223 | 0.005063 | 6.3e-05 | 0.001108 | 0.31022 | 1.3213 | 2.5481 | 100.99 | 20.68 | 18.46 | 69.8691 | 103.2375 | 69.1048 | 0.004462 | 1.0315 | 0.02507 | 24.304 | -0.094344 | 0.14181 | 0.000127 | 0.009679 | 0.009679 | 0.60214
2019-03-20 13:04 | 1553101414.1793 | 831 | 0.20953 | 0.00044 | -1e-06 | 107.8955 | 109.2913 | 127.3063 | 20.9256 | 0.002049 | 0.001975 | -152.1317 | 0.23857 | 8.4013 | 0.003439 | 0.004231 | 0.004036 | 6.4e-05 | 0.001112 | 0.33467 | 1.306 | 2.5481 | 100.99 | 20.68 | 18.462 | 70.7951 | 104.2757 | 69.0188 | 0.004463 | 1.0318 | 0.02507 | 24.2972 | -0.11497 | 0.14192 | 0.000167 | 0.012697 | 0.012697 | 0.45942
2019-03-20 13:04 | 1553101415.18 | 832 | 0.20954 | 0.000438 | -1e-06 | 107.8842 | 109.1328 | 127.2925 | 20.925 | 0.002053 | 0.001952 | -152.1258 | 0.23881 | 8.4033 | 0.003352 | 0.004369 | 0.004104 | 6.2e-05 | 0.001113 | 0.31924 | 1.2993 | 2.548 | 100.99 | 20.68 | 18.462 | 70.8859 | 102.165 | 68.4716 | 0.004463 | 1.0334 | 0.025088 | 24.2765 | -0.10293 | 0.14207 | 8.2e-05 | 0.00626 | 0.00626 | 0.93273
2019-03-20 13:04 | 1553101416.1803 | 833 | 0.20952 | 0.000441 | -1e-06 | 107.9327 | 109.1642 | 127.3552 | 20.9249 | 0.00206 | 0.002006 | -152.1301 | 0.23894 | 8.4034 | 0.003477 | 0.004664 | 0.004473 | 6.6e-05 | 0.001112 | 0.33627 | 1.3388 | 2.548 | 100.99 | 20.69 | 18.461 | 71.0665 | 105.2046 | 69.7755 | 0.004465 | 1.0295 | 0.025047 | 24.3287 | -0.10148 | 0.14214 | 0.000195 | 0.014798 | 0.014798 | 0.39479
2019-03-20 13:04 | 1553101417.1802 | 834 | 0.20954 | 0.000442 | -1e-06 | 107.9876 | 109.1171 | 127.2402 | 20.9249 | 0.002086 | 0.001966 | -152.1348 | 0.23862 | 8.4002 | 0.003423 | 0.004345 | 0.004097 | 6.4e-05 | 0.001111 | 0.32278 | 1.3359 | 2.5481 | 100.99 | 20.69 | 18.457 | 71.6973 | 103.8975 | 68.7943 | 0.004464 | 1.0325 | 0.025071 | 24.2829 | -0.093067 | 0.14195 | 0.00011 | 0.008428 | 0.008428 | 0.69222
2019-03-20 13:04 | 1553101418.1805 | 835 | 0.20952 | 0.000439 | -1e-06 | 107.8861 | 109.0947 | 127.0327 | 20.9259 | 0.002095 | 0.001983 | -152.1463 | 0.23916 | 8.4033 | 0.003405 | 0.004645 | 0.004352 | 6.5e-05 | 0.00111 | 0.29839 | 1.3519 | 2.5479 | 100.99 | 20.69 | 18.455 | 71.8941 | 103.4614 | 69.2126 | 0.004464 | 1.0312 | 0.025036 | 24.278 | -0.099644 | 0.14228 | 0.000194 | 0.014696 | 0.014696 | 0.39789
2019-03-20 13:04 | 1553101419.1802 | 836 | 0.20955 | 0.00044 | -1e-06 | 107.965 | 109.1358 | 126.7231 | 20.9249 | 0.002084 | 0.001994 | -152.174 | 0.23827 | 8.4002 | 0.00342 | 0.004939 | 0.004692 | 6.4e-05 | 0.001111 | 0.32261 | 1.3168 | 2.548 | 100.99 | 20.69 | 18.451 | 71.6506 | 103.826 | 69.4774 | 0.004463 | 1.0304 | 0.024995 | 24.2577 | -0.096474 | 0.14175 | 5.1e-05 | 0.00395 | 0.00395 | 1.4748
2019-03-20 13:04 | 1553101420.1803 | 837 | 0.20955 | 0.000444 | -1e-06 | 107.9388 | 109.1601 | 126.9486 | 20.9247 | 0.002035 | 0.001942 | -152.173 | 0.23949 | 8.4014 | 0.003482 | 0.00507 | 0.004856 | 6.1e-05 | 0.001107 | 0.32892 | 1.3255 | 2.5482 | 100.99 | 20.69 | 18.444 | 70.4628 | 105.3169 | 68.2217 | 0.004461 | 1.0342 | 0.025064 | 24.2348 | -0.10064 | 0.14247 | 2.5e-05 | 0.001967 | 0.001967 | 2.9771
2019-03-20 13:04 | 1553101421.1796 | 838 | 0.20954 | 0.000443 | -1e-06 | 107.8466 | 109.0401 | 127.0496 | 20.9251 | 0.002106 | 0.001946 | -152.1746 | 0.2388 | 8.402 | 0.003422 | 0.004745 | 0.004538 | 6.4e-05 | 0.001108 | 0.323 | 1.3219 | 2.5481 | 100.99 | 20.69 | 18.439 | 72.1667 | 103.8703 | 68.332 | 0.00446 | 1.0339 | 0.025069 | 24.2483 | -0.098441 | 0.14206 | 8.6e-05 | 0.006592 | 0.006592 | 0.88567
2019-03-20 13:04 | 1553101422.1793 | 839 | 0.20954 | 0.000445 | -2e-06 | 107.9 | 109.0861 | 126.9503 | 20.9252 | 0.002042 | 0.001964 | -152.1683 | 0.23807 | 8.4032 | 0.003468 | 0.004694 | 0.00456 | 6.4e-05 | 0.001108 | 0.30879 | 1.318 | 2.548 | 100.99 | 20.69 | 18.435 | 70.6232 | 104.9853 | 68.7643 | 0.004459 | 1.0326 | 0.025044 | 24.2542 | -0.097787 | 0.14163 | 5.8e-05 | 0.004476 | 0.004476 | 1.3006
2019-03-20 13:04 | 1553101423.1794 | 840 | 0.20955 | 0.000438 | -1e-06 | 107.9437 | 109.0569 | 127.1449 | 20.9244 | 0.002083 | 0.001984 | -152.1422 | 0.23813 | 8.4045 | 0.003503 | 0.005001 | 0.0048 | 6.5e-05 | 0.001108 | 0.33356 | 1.3629 | 2.5481 | 100.99 | 20.69 | 18.425 | 71.6127 | 105.8319 | 69.2282 | 0.004457 | 1.0312 | 0.025046 | 24.2892 | -0.091777 | 0.14167 | 5.1e-05 | 0.003928 | 0.003928 | 1.4823
2019-03-20 13:04 | 1553101424.1784 | 841 | 0.20957 | 0.000438 | -1e-06 | 107.8602 | 109.1418 | 126.924 | 20.9247 | 0.002035 | 0.001967 | -152.118 | 0.23924 | 8.4026 | 0.003393 | 0.005266 | 0.005 | 6.3e-05 | 0.001108 | 0.3585 | 1.314 | 2.5482 | 100.99 | 20.69 | 18.421 | 70.4684 | 103.1688 | 68.816 | 0.004456 | 1.0324 | 0.025039 | 24.2535 | -0.10565 | 0.14232 | -8.5e-05 | -0.006339 | -0.006339 | -0.92279
2019-03-20 13:04 | 1553101425.1789 | 842 | 0.20954 | 0.000435 | -1e-06 | 107.9378 | 109.181 | 126.9881 | 20.925 | 0.002015 | 0.001971 | -152.112 | 0.23829 | 8.4023 | 0.00342 | 0.00524 | 0.005071 | 6.3e-05 | 0.001111 | 0.32565 | 1.3407 | 2.548 | 100.99 | 20.69 | 18.41 | 69.9869 | 103.8254 | 68.931 | 0.004453 | 1.0321 | 0.025041 | 24.2637 | -0.10243 | 0.14176 | 7.3e-05 | 0.005568 | 0.005568 | 1.0463
2019-03-20 13:04 | 1553101426.1785 | 843 | 0.20972 | 0.000443 | 0 | 107.9084 | 109.1261 | 127.097 | 20.925 | 0.002044 | 0.001937 | -152.1049 | 0.23906 | 8.402 | 0.003418 | 0.00472 | 0.004493 | 6.5e-05 | 0.001111 | 0.34844 | 1.3422 | 2.548 | 100.99 | 20.69 | 18.403 | 70.6744 | 103.7774 | 68.0963 | 0.004451 | 1.0346 | 0.025083 | 24.2445 | -0.10037 | 0.14222 | -0.000987 | -0.074752 | -0.074752 | -0.078194
2019-03-20 13:04 | 1553101427.18 | 844 | 0.20938 | 0.000437 | -1e-06 | 107.8392 | 108.95 | 126.9688 | 20.9247 | 0.002122 | 0.001939 | -152.1067 | 0.23844 | 8.4038 | 0.003449 | 0.004599 | 0.004405 | 6.4e-05 | 0.00111 | 0.30541 | 1.3626 | 2.5481 | 100.99 | 20.69 | 18.395 | 72.5532 | 104.5327 | 68.1467 | 0.004449 | 1.0344 | 0.025068 | 24.234 | -0.091666 | 0.14185 | 0.001036 | 0.078443 | 0.078443 | 0.074322
2019-03-20 13:04 | 1553101428.1796 | 845 | 0.20936 | 0.000441 | 0 | 107.8431 | 109.1354 | 126.8897 | 20.9258 | 0.002014 | 0.001949 | -152.0925 | 0.23893 | 8.4012 | 0.003424 | 0.004536 | 0.004405 | 6.3e-05 | 0.001108 | 0.37576 | 1.3655 | 2.5481 | 100.99 | 20.69 | 18.38 | 69.9587 | 103.9165 | 68.4028 | 0.004446 | 1.0336 | 0.025051 | 24.2356 | -0.10655 | 0.14214 | 0.001195 | 0.090389 | 0.090389 | 0.064632
2019-03-20 13:04 | 1553101429.1797 | 846 | 0.20951 | 0.000444 | -1e-06 | 107.8615 | 108.9736 | 126.9287 | 20.9243 | 0.002043 | 0.001927 | -152.0686 | 0.23908 | 8.4023 | 0.00339 | 0.004845 | 0.004655 | 6.4e-05 | 0.001101 | 0.32587 | 1.3288 | 2.548 | 100.99 | 20.69 | 18.372 | 70.6567 | 103.0883 | 67.8617 | 0.004444 | 1.0353 | 0.025075 | 24.2201 | -0.091748 | 0.14223 | 0.00025 | 0.01897 | 0.01897 | 0.30816
2019-03-20 13:04 | 1553101430.1792 | 847 | 0.20955 | 0.000444 | -1e-06 | 107.9582 | 109.1272 | 126.6308 | 20.9261 | 0.002033 | 0.001938 | -152.0459 | 0.23824 | 8.3998 | 0.003352 | 0.004906 | 0.004719 | 6.2e-05 | 0.001102 | 0.3282 | 1.329 | 2.548 | 100.99 | 20.69 | 18.357 | 70.4197 | 102.1835 | 68.1286 | 0.00444 | 1.0345 | 0.025036 | 24.2011 | -0.096332 | 0.14173 | 6e-06 | 0.000584 | 0.000584 | 9.9703
2019-03-20 13:04 | 1553101431.1788 | 848 | 0.20958 | 0.00044 | -1e-06 | 107.8457 | 109.1975 | 126.4015 | 20.9245 | 0.001992 | 0.001917 | -152.0439 | 0.23949 | 8.4039 | 0.003409 | 0.004406 | 0.004237 | 6.2e-05 | 0.001103 | 0.32731 | 1.3358 | 2.5481 | 100.99 | 20.69 | 18.35 | 69.4287 | 103.5444 | 67.6332 | 0.004438 | 1.036 | 0.025031 | 24.1616 | -0.11142 | 0.14247 | -0.00012 | -0.009005 | -0.009005 | -0.65027
2019-03-20 13:04 | 1553101432.1789 | 849 | 0.20957 | 0.000439 | -1e-06 | 107.9016 | 109.1672 | 126.252 | 20.9245 | 0.002026 | 0.001953 | -152.0443 | 0.23907 | 8.4041 | 0.003443 | 0.00461 | 0.00444 | 6.3e-05 | 0.001103 | 0.35961 | 1.3505 | 2.5481 | 100.99 | 20.69 | 18.341 | 70.256 | 104.3878 | 68.5007 | 0.004436 | 1.0334 | 0.024984 | 24.1781 | -0.1043 | 0.14222 | -0.000111 | -0.0083 | -0.0083 | -0.70429
2019-03-20 13:04 | 1553101433.18 | 850 | 0.20957 | 0.000436 | -1e-06 | 107.8845 | 109.2031 | 125.9719 | 20.9256 | 0.002029 | 0.002008 | -152.0578 | 0.2393 | 8.4029 | 0.003439 | 0.005224 | 0.005051 | 6.2e-05 | 0.001103 | 0.31277 | 1.3452 | 2.5479 | 100.99 | 20.7 | 18.323 | 70.3304 | 104.2726 | 69.8127 | 0.004435 | 1.0294 | 0.024909 | 24.1976 | -0.10866 | 0.14236 | -5.7e-05 | -0.004202 | -0.004202 | -1.3925
2019-03-20 13:04 | 1553101434.1789 | 851 | 0.20957 | 0.000441 | -1e-06 | 107.9343 | 109.2281 | 125.5652 | 20.9241 | 0.002054 | 0.00196 | -152.0746 | 0.23997 | 8.4011 | 0.003334 | 0.004869 | 0.004657 | 6.1e-05 | 0.001105 | 0.30237 | 1.3492 | 2.548 | 100.99 | 20.7 | 18.314 | 70.9264 | 101.7503 | 68.6539 | 0.004432 | 1.0329 | 0.024911 | 24.1176 | -0.10658 | 0.14276 | -7.1e-05 | -0.005286 | -0.005286 | -1.11
2019-03-20 13:04 | 1553101435.1793 | 852 | 0.20954 | 0.00044 | -1e-06 | 107.9469 | 109.0244 | 125.4588 | 20.9248 | 0.002065 | 0.001939 | -152.0784 | 0.23924 | 8.4006 | 0.003429 | 0.004479 | 0.004246 | 6e-05 | 0.001105 | 0.33356 | 1.3342 | 2.5479 | 100.99 | 20.7 | 18.305 | 71.1779 | 104.05 | 68.1586 | 0.00443 | 1.0344 | 0.024918 | 24.0899 | -0.088842 | 0.14232 | 7e-05 | 0.005376 | 0.005376 | 1.088
2019-03-20 13:04 | 1553101436.1792 | 853 | 0.20957 | 0.000445 | -1e-06 | 107.8939 | 109.255 | 125.0565 | 20.9251 | 0.002055 | 0.001937 | -152.0861 | 0.23922 | 8.4025 | 0.003316 | 0.005096 | 0.004903 | 6.5e-05 | 0.001103 | 0.32637 | 1.3704 | 2.5482 | 100.99 | 20.7 | 18.288 | 70.9381 | 101.303 | 68.0956 | 0.004426 | 1.0346 | 0.024881 | 24.0491 | -0.11214 | 0.14231 | -0.000106 | -0.007898 | -0.007898 | -0.74059
2019-03-20 13:04 | 1553101437.1801 | 854 | 0.20954 | 0.000438 | -1e-06 | 107.9443 | 109.0091 | 124.9654 | 20.9247 | 0.002021 | 0.001951 | -152.0973 | 0.23856 | 8.403 | 0.003358 | 0.004697 | 0.004555 | 6.2e-05 | 0.001093 | 0.3063 | 1.3489 | 2.5481 | 100.99 | 20.7 | 18.281 | 70.1199 | 102.3299 | 68.4432 | 0.004424 | 1.0335 | 0.024859 | 24.0525 | -0.087799 | 0.14192 | 9.3e-05 | 0.007007 | 0.007007 | 0.83246
2019-03-20 13:04 | 1553101438.1786 | 855 | 0.20958 | 0.00044 | -1e-06 | 107.8091 | 108.962 | 124.6136 | 20.9252 | 0.002022 | 0.001913 | -152.1115 | 0.23989 | 8.4006 | 0.003312 | 0.005145 | 0.004916 | 5.8e-05 | 0.001091 | 0.33892 | 1.3517 | 2.5481 | 100.99 | 20.7 | 18.271 | 70.1569 | 101.2048 | 67.5264 | 0.004422 | 1.0363 | 0.024857 | 23.9864 | -0.095143 | 0.14271 | -0.000134 | -0.009988 | -0.009988 | -0.58727
2019-03-20 13:04 | 1553101439.1782 | 856 | 0.20954 | 0.000441 | -1e-06 | 107.9246 | 109.1945 | 124.7199 | 20.925 | 0.002019 | 0.001959 | -152.1166 | 0.23932 | 8.403 | 0.003315 | 0.004521 | 0.004346 | 6e-05 | 0.001092 | 0.31492 | 1.3308 | 2.5481 | 100.99 | 20.7 | 18.266 | 70.071 | 101.2795 | 68.6408 | 0.004421 | 1.0329 | 0.024827 | 24.0358 | -0.10463 | 0.14237 | 8.2e-05 | 0.006207 | 0.006207 | 0.94277
2019-03-20 13:04 | 1553101440.1793 | 857 | 0.20955 | 0.00044 | 0 | 107.8392 | 109.0121 | 124.8394 | 20.9243 | 0.002013 | 0.001921 | -152.1313 | 0.23883 | 8.4043 | 0.003272 | 0.004963 | 0.004651 | 5.7e-05 | 0.001094 | 0.3449 | 1.363 | 2.5481 | 100.99 | 20.7 | 18.262 | 69.9348 | 100.249 | 67.7213 | 0.00442 | 1.0357 | 0.024873 | 24.015 | -0.096758 | 0.14208 | 2.4e-05 | 0.001815 | 0.001815 | 3.2172
2019-03-20 13:04 | 1553101441.1794 | 858 | 0.20954 | 0.000442 | -1e-06 | 107.9424 | 109.3017 | 124.8126 | 20.9248 | 0.002016 | 0.001973 | -152.1535 | 0.23804 | 8.4025 | 0.003235 | 0.004795 | 0.004594 | 6.1e-05 | 0.001092 | 0.34418 | 1.3633 | 2.5481 | 100.99 | 20.7 | 18.257 | 69.9952 | 99.3347 | 68.9627 | 0.004419 | 1.032 | 0.024825 | 24.0561 | -0.11194 | 0.14161 | 8.4e-05 | 0.006365 | 0.006365 | 0.91436
2019-03-20 13:04 | 1553101442.1795 | 859 | 0.20954 | 0.000443 | -1e-06 | 107.8838 | 109.1821 | 125.133 | 20.925 | 0.002037 | 0.001921 | -152.1618 | 0.23874 | 8.3993 | 0.003246 | 0.00481 | 0.004621 | 5.8e-05 | 0.001091 | 0.31006 | 1.3492 | 2.5482 | 100.99 | 20.7 | 18.255 | 70.5169 | 99.6187 | 67.7258 | 0.004418 | 1.0357 | 0.024902 | 24.0434 | -0.107 | 0.14203 | 7.7e-05 | 0.005819 | 0.005819 | 1.0031
2019-03-20 13:04 | 1553101443.1796 | 860 | 0.20954 | 0.000443 | -1e-06 | 107.9301 | 109.224 | 125.3311 | 20.9246 | 0.001967 | 0.001944 | -152.1604 | 0.23875 | 8.4049 | 0.003256 | 0.004631 | 0.004428 | 6.3e-05 | 0.001091 | 0.32803 | 1.3541 | 2.548 | 100.99 | 20.7 | 18.252 | 68.8339 | 99.8476 | 68.2651 | 0.004417 | 1.0341 | 0.024902 | 24.0814 | -0.10659 | 0.14203 | 7.4e-05 | 0.005673 | 0.005673 | 1.0291
2019-03-20 13:04 | 1553101444.1785 | 861 | 0.20954 | 0.000438 | -1e-06 | 107.8977 | 109.2786 | 125.6652 | 20.9245 | 0.002034 | 0.001951 | -152.1647 | 0.23861 | 8.4031 | 0.00326 | 0.004144 | 0.003959 | 6.1e-05 | 0.001085 | 0.3314 | 1.3792 | 2.548 | 100.99 | 20.7 | 18.251 | 70.4307 | 99.9493 | 68.4492 | 0.004417 | 1.0335 | 0.024928 | 24.12 | -0.11374 | 0.14195 | 0.00012 | 0.009043 | 0.009043 | 0.64515
2019-03-20 13:04 | 1553101445.1786 | 862 | 0.20955 | 0.000442 | 0 | 107.9142 | 109.1743 | 125.8379 | 20.924 | 0.00203 | 0.001915 | -152.1672 | 0.23972 | 8.4023 | 0.003244 | 0.00478 | 0.004594 | 5.8e-05 | 0.001093 | 0.33804 | 1.3767 | 2.548 | 100.99 | 20.7 | 18.251 | 70.3374 | 99.5527 | 67.5789 | 0.004417 | 1.0361 | 0.024977 | 24.1058 | -0.10384 | 0.14261 | 4.4e-05 | 0.003343 | 0.003343 | 1.7531
2019-03-20 13:04 | 1553101446.1796 | 863 | 0.20954 | 0.000438 | -1e-06 | 107.8822 | 109.2154 | 126.1595 | 20.924 | 0.001988 | 0.001944 | -152.1619 | 0.23921 | 8.4014 | 0.003279 | 0.003871 | 0.00363 | 6.3e-05 | 0.001091 | 0.3298 | 1.3404 | 2.5481 | 100.99 | 20.7 | 18.252 | 69.3425 | 100.4127 | 68.2857 | 0.004417 | 1.034 | 0.024983 | 24.1616 | -0.10986 | 0.14231 | 7.6e-05 | 0.005795 | 0.005795 | 1.0093
2019-03-20 13:04 | 1553101447.1799 | 864 | 0.20953 | 0.000439 | -1e-06 | 107.9013 | 109.0962 | 126.6224 | 20.9235 | 0.00201 | 0.001946 | -152.1467 | 0.23933 | 8.402 | 0.003319 | 0.004936 | 0.004735 | 5.9e-05 | 0.001098 | 0.3449 | 1.3448 | 2.548 | 100.99 | 20.7 | 18.253 | 69.8721 | 101.3767 | 68.3311 | 0.004418 | 1.0339 | 0.025027 | 24.2075 | -0.098507 | 0.14238 | 0.00014 | 0.010646 | 0.010646 | 0.54967
2019-03-20 13:04 | 1553101448.1791 | 865 | 0.20955 | 0.000443 | -1e-06 | 107.9201 | 108.9142 | 127.2362 | 20.9238 | 0.002055 | 0.001973 | -152.1328 | 0.23878 | 8.4034 | 0.003245 | 0.004402 | 0.004248 | 6.6e-05 | 0.0011 | 0.32764 | 1.3391 | 2.5479 | 100.99 | 20.7 | 18.252 | 70.9525 | 99.5879 | 68.9632 | 0.004417 | 1.032 | 0.025065 | 24.2885 | -0.082016 | 0.14205 | 5.9e-05 | 0.004525 | 0.004525 | 1.2903
2019-03-20 13:04 | 1553101449.1789 | 866 | 0.20953 | 0.000436 | -1e-06 | 107.8521 | 108.9986 | 127.3707 | 20.9234 | 0.002042 | 0.001921 | -152.1364 | 0.23956 | 8.4044 | 0.003161 | 0.004893 | 0.00464 | 6.4e-05 | 0.001103 | 0.32045 | 1.353 | 2.548 | 100.99 | 20.7 | 18.251 | 70.6183 | 97.548 | 67.7211 | 0.004417 | 1.0357 | 0.025124 | 24.2572 | -0.094581 | 0.14252 | 0.000178 | 0.013525 | 0.013525 | 0.43309
2019-03-20 13:04 | 1553101450.1797 | 867 | 0.20955 | 0.000439 | -2e-06 | 107.9058 | 109.0121 | 127.4284 | 20.9235 | 0.002046 | 0.002004 | -152.1378 | 0.23896 | 8.4008 | 0.003106 | 0.004794 | 0.004596 | 6.4e-05 | 0.001103 | 0.34523 | 1.359 | 2.548 | 100.99 | 20.7 | 18.25 | 70.7361 | 96.2366 | 69.7203 | 0.004417 | 1.0297 | 0.025056 | 24.3338 | -0.091234 | 0.14216 | 2.1e-05 | 0.001707 | 0.001707 | 3.4236
2019-03-20 13:04 | 1553101451.1799 | 868 | 0.20956 | 0.00044 | -1e-06 | 107.8997 | 109.1496 | 127.653 | 20.9242 | 0.001997 | 0.001947 | -152.1032 | 0.23898 | 8.4036 | 0.003131 | 0.004907 | 0.004738 | 6.3e-05 | 0.001098 | 0.35994 | 1.3897 | 2.5481 | 100.99 | 20.7 | 18.252 | 69.5423 | 96.8358 | 68.3498 | 0.004417 | 1.0338 | 0.025128 | 24.3065 | -0.10302 | 0.14217 | -4.2e-05 | -0.003088 | -0.003088 | -1.892
2019-03-20 13:04 | 1553101452.18 | 869 | 0.20956 | 0.000441 | -1e-06 | 107.8243 | 109.2128 | 127.5624 | 20.924 | 0.00205 | 0.001884 | -152.0987 | 0.23889 | 8.4033 | 0.003142 | 0.004166 | 0.003963 | 6.3e-05 | 0.001093 | 0.33428 | 1.357 | 2.5481 | 100.99 | 20.7 | 18.253 | 70.8162 | 97.0902 | 66.8378 | 0.004418 | 1.0384 | 0.025175 | 24.244 | -0.11445 | 0.14211 | -3.4e-05 | -0.00251 | -0.00251 | -2.3274
2019-03-20 13:04 | 1553101453.179 | 870 | 0.20954 | 0.000443 | -1e-06 | 107.8929 | 108.9033 | 127.4002 | 20.9239 | 0.001986 | 0.001933 | -152.1003 | 0.2394 | 8.4015 | 0.003196 | 0.004616 | 0.004354 | 6.1e-05 | 0.001084 | 0.31868 | 1.3441 | 2.5482 | 100.99 | 20.7 | 18.256 | 69.2878 | 98.3958 | 68.0048 | 0.004418 | 1.0349 | 0.025116 | 24.2701 | -0.083378 | 0.14242 | 7.6e-05 | 0.005855 | 0.005855 | 0.99977
2019-03-20 13:04 | 1553101454.1781 | 871 | 0.20955 | 0.000438 | -1e-06 | 107.8673 | 109.0831 | 127.2284 | 20.9232 | 0.002024 | 0.00191 | -152.1142 | 0.23902 | 8.4038 | 0.003088 | 0.005092 | 0.004825 | 6.3e-05 | 0.00109 | 0.32565 | 1.3922 | 2.5482 | 100.99 | 20.7 | 18.263 | 70.2042 | 95.7852 | 67.4581 | 0.00442 | 1.0365 | 0.025119 | 24.2343 | -0.10025 | 0.14219 | 5.6e-05 | 0.004309 | 0.004309 | 1.3562
2019-03-20 13:04 | 1553101455.1791 | 872 | 0.20957 | 0.000441 | -1e-06 | 107.8745 | 109.1844 | 126.9688 | 20.9235 | 0.002026 | 0.00191 | -152.1209 | 0.23869 | 8.4005 | 0.003064 | 0.004307 | 0.00416 | 6.1e-05 | 0.001091 | 0.3632 | 1.3604 | 2.5481 | 100.99 | 20.7 | 18.266 | 70.2411 | 95.221 | 67.463 | 0.004421 | 1.0365 | 0.025093 | 24.2098 | -0.10796 | 0.142 | -6e-05 | -0.004522 | -0.004522 | -1.2905
2019-03-20 13:04 | 1553101456.1782 | 873 | 0.20975 | 0.000443 | -1e-06 | 107.8732 | 109.0667 | 126.9855 | 20.9236 | 0.001942 | 0.001903 | -152.1227 | 0.23896 | 8.4023 | 0.003172 | 0.004351 | 0.004139 | 6.2e-05 | 0.001084 | 0.31996 | 1.3435 | 2.5481 | 100.99 | 20.7 | 18.275 | 68.2176 | 97.8328 | 67.2907 | 0.004423 | 1.037 | 0.025101 | 24.2052 | -0.098418 | 0.14216 | -0.001162 | -0.087999 | -0.087999 | -0.066394
2019-03-20 13:04 | 1553101457.1794 | 874 | 0.20973 | 0.00044 | -1e-06 | 107.846 | 109.1601 | 126.6921 | 20.924 | 0.002057 | 0.001927 | -152.1368 | 0.23827 | 8.4029 | 0.003173 | 0.004364 | 0.004107 | 6.3e-05 | 0.001088 | 0.335 | 1.3668 | 2.5481 | 100.99 | 20.7 | 18.28 | 71.0017 | 97.8472 | 67.8645 | 0.004424 | 1.0353 | 0.025051 | 24.1976 | -0.10833 | 0.14175 | -0.00103 | -0.07785 | -0.07785 | -0.074833
2019-03-20 13:04 | 1553101458.1793 | 875 | 0.20957 | 0.000441 | -1e-06 | 107.8709 | 109.0042 | 126.6043 | 20.9223 | 0.002006 | 0.001923 | -152.1461 | 0.23931 | 8.4016 | 0.003189 | 0.004729 | 0.004481 | 6.2e-05 | 0.001094 | 0.35795 | 1.3463 | 2.5479 | 100.99 | 20.7 | 18.284 | 69.7699 | 98.2336 | 67.7631 | 0.004425 | 1.0356 | 0.025046 | 24.1856 | -0.093486 | 0.14236 | -6.5e-05 | -0.004836 | -0.004836 | -1.2098
2019-03-20 13:04 | 1553101459.1793 | 876 | 0.20958 | 0.000446 | -2e-06 | 107.8987 | 109.0782 | 126.4766 | 20.924 | 0.001978 | 0.001934 | -152.1428 | 0.23936 | 8.4034 | 0.003253 | 0.00465 | 0.004566 | 5.7e-05 | 0.001092 | 0.32549 | 1.3333 | 2.548 | 100.99 | 20.7 | 18.293 | 69.0944 | 99.7754 | 68.0251 | 0.004427 | 1.0348 | 0.025024 | 24.1827 | -0.09725 | 0.1424 | -0.000176 | -0.013195 | -0.013195 | -0.44355
2019-03-20 13:04 | 1553101460.1787 | 877 | 0.20956 | 0.00044 | -1e-06 | 107.8745 | 109.1945 | 126.3421 | 20.9232 | 0.002017 | 0.001934 | -152.1532 | 0.23945 | 8.4003 | 0.003254 | 0.004357 | 0.004187 | 6.6e-05 | 0.001096 | 0.32549 | 1.3027 | 2.5481 | 100.99 | 20.7 | 18.296 | 70.0396 | 99.8147 | 68.0392 | 0.004428 | 1.0348 | 0.02501 | 24.1704 | -0.10879 | 0.14245 | 1e-06 | 0.000127 | 0.000127 | 46.2494
2019-03-20 13:04 | 1553101461.1798 | 878 | 0.20957 | 0.000442 | -2e-06 | 107.8767 | 109.1201 | 126.0156 | 20.9238 | 0.002046 | 0.001879 | -152.1662 | 0.23899 | 8.4008 | 0.003227 | 0.004643 | 0.004341 | 6.4e-05 | 0.0011 | 0.35347 | 1.3064 | 2.5482 | 100.99 | 20.7 | 18.305 | 70.7312 | 99.1524 | 66.7027 | 0.00443 | 1.0388 | 0.025027 | 24.0918 | -0.1025 | 0.14217 | -8.6e-05 | -0.006376 | -0.006376 | -0.91653
2019-03-20 13:04 | 1553101462.1795 | 879 | 0.2096 | 0.000445 | -1e-06 | 107.9061 | 109.1388 | 126.0057 | 20.9226 | 0.002098 | 0.001926 | -152.1732 | 0.23943 | 8.4023 | 0.00331 | 0.004566 | 0.004391 | 6.4e-05 | 0.001106 | 0.31437 | 1.3419 | 2.548 | 100.99 | 20.7 | 18.307 | 71.9858 | 101.153 | 67.8312 | 0.004431 | 1.0354 | 0.024985 | 24.1308 | -0.1016 | 0.14244 | -0.000253 | -0.018969 | -0.018969 | -0.30862
2019-03-20 13:04 | 1553101463.1786 | 880 | 0.20955 | 0.000442 | -1e-06 | 107.9136 | 109.1855 | 126.147 | 20.9233 | 0.00203 | 0.001925 | -152.1773 | 0.23958 | 8.4036 | 0.003262 | 0.005124 | 0.004925 | 6.6e-05 | 0.001105 | 0.32637 | 1.307 | 2.5481 | 100.99 | 20.7 | 18.31 | 70.3343 | 100.0078 | 67.8054 | 0.004431 | 1.0355 | 0.025 | 24.1434 | -0.10481 | 0.14252 | 9e-06 | 0.000796 | 0.000796 | 7.3569
2019-03-20 13:04 | 1553101464.1795 | 881 | 0.20954 | 0.000437 | -1e-06 | 107.878 | 109.2091 | 126.6977 | 20.9235 | 0.002031 | 0.001925 | -152.1625 | 0.23894 | 8.4014 | 0.003329 | 0.004876 | 0.004683 | 6.6e-05 | 0.001109 | 0.35153 | 1.3055 | 2.548 | 100.99 | 20.7 | 18.314 | 70.3651 | 101.6265 | 67.8265 | 0.004432 | 1.0354 | 0.025053 | 24.1968 | -0.10969 | 0.14215 | 9.2e-05 | 0.007021 | 0.007021 | 0.83214
2019-03-20 13:04 | 1553101465.179 | 882 | 0.20959 | 0.000438 | -1e-06 | 107.9252 | 109.0222 | 127.1765 | 20.9227 | 0.002018 | 0.0019 | -152.1524 | 0.23889 | 8.4048 | 0.00333 | 0.005397 | 0.005187 | 6.7e-05 | 0.001114 | 0.33915 | 1.3044 | 2.548 | 100.99 | 20.7 | 18.314 | 70.0562 | 101.639 | 67.2033 | 0.004432 | 1.0373 | 0.025124 | 24.2203 | -0.090453 | 0.14211 | -0.0002 | -0.015095 | -0.015095 | -0.38694
2019-03-20 13:04 | 1553101466.1797 | 883 | 0.20957 | 0.000435 | -1e-06 | 107.9113 | 109.2386 | 127.1817 | 20.9218 | 0.001981 | 0.001955 | -152.1372 | 0.23928 | 8.4026 | 0.003288 | 0.005216 | 0.005016 | 6.6e-05 | 0.001114 | 0.32637 | 1.3451 | 2.548 | 100.99 | 20.7 | 18.312 | 69.1648 | 100.6192 | 68.5425 | 0.004432 | 1.0332 | 0.025075 | 24.2684 | -0.10934 | 0.14235 | -4.3e-05 | -0.003188 | -0.003188 | -1.835
2019-03-20 13:04 | 1553101467.1791 | 884 | 0.20955 | 0.000437 | -2e-06 | 107.8715 | 108.9725 | 127.0395 | 20.9222 | 0.002031 | 0.001931 | -152.144 | 0.23932 | 8.4053 | 0.003269 | 0.004268 | 0.004086 | 6.5e-05 | 0.001113 | 0.33843 | 1.3195 | 2.5481 | 100.99 | 20.7 | 18.31 | 70.3544 | 100.1535 | 67.951 | 0.004431 | 1.035 | 0.025082 | 24.2338 | -0.090825 | 0.14237 | 5.4e-05 | 0.004177 | 0.004177 | 1.4009
2019-03-20 13:04 | 1553101468.1798 | 885 | 0.20957 | 0.000443 | -2e-06 | 107.9362 | 109.1679 | 126.9134 | 20.9223 | 0.001992 | 0.001965 | -152.1362 | 0.23936 | 8.4031 | 0.00329 | 0.00419 | 0.003935 | 6.5e-05 | 0.001116 | 0.32068 | 1.321 | 2.5482 | 100.99 | 20.7 | 18.306 | 69.4374 | 100.6763 | 68.7675 | 0.004431 | 1.0325 | 0.02504 | 24.2508 | -0.10149 | 0.1424 | -5e-05 | -0.003664 | -0.003664 | -1.5972
2019-03-20 13:04 | 1553101469.1789 | 886 | 0.20958 | 0.00044 | -2e-06 | 107.8654 | 109.2606 | 126.5754 | 20.9231 | 0.002004 | 0.001899 | -152.1451 | 0.23943 | 8.4043 | 0.003353 | 0.004815 | 0.004623 | 6.7e-05 | 0.001125 | 0.33268 | 1.3362 | 2.5481 | 100.99 | 20.71 | 18.296 | 69.7206 | 102.2016 | 67.2029 | 0.004431 | 1.0373 | 0.025064 | 24.163 | -0.11496 | 0.14244 | -0.000148 | -0.011131 | -0.011131 | -0.52592
2019-03-20 13:04 | 1553101470.1779 | 887 | 0.20957 | 0.000439 | -1e-06 | 107.8971 | 109.2128 | 126.8435 | 20.9228 | 0.002081 | 0.001913 | -152.1438 | 0.23941 | 8.4041 | 0.003338 | 0.004648 | 0.004428 | 6.7e-05 | 0.001124 | 0.32892 | 1.3591 | 2.548 | 100.99 | 20.71 | 18.291 | 71.565 | 101.8402 | 67.5169 | 0.00443 | 1.0363 | 0.025079 | 24.1997 | -0.10841 | 0.14242 | -6.1e-05 | -0.004537 | -0.004537 | -1.2901
2019-03-20 13:04 | 1553101471.1789 | 888 | 0.2095 | 0.000441 | -1e-06 | 107.9165 | 109.1534 | 127.2587 | 20.9236 | 0.002011 | 0.001893 | -152.1466 | 0.23949 | 8.4013 | 0.003319 | 0.004571 | 0.004422 | 6.8e-05 | 0.001123 | 0.33107 | 1.3342 | 2.548 | 100.99 | 20.71 | 18.288 | 69.891 | 101.3729 | 67.0372 | 0.004429 | 1.0378 | 0.025138 | 24.2222 | -0.10193 | 0.14247 | 0.000353 | 0.026864 | 0.026864 | 0.21797
2019-03-20 13:05 | 1553101472.1793 | 889 | 0.20957 | 0.000445 | -1e-06 | 107.9168 | 109.03 | 127.1268 | 20.923 | 0.001955 | 0.00195 | -152.1469 | 0.23874 | 8.4034 | 0.003231 | 0.004277 | 0.003997 | 6.6e-05 | 0.00113 | 0.34739 | 1.3254 | 2.5481 | 100.99 | 20.71 | 18.283 | 68.543 | 99.2498 | 68.4143 | 0.004428 | 1.0336 | 0.025074 | 24.2586 | -0.091794 | 0.14203 | -9.3e-05 | -0.006982 | -0.006982 | -0.83611
2019-03-20 13:05 | 1553101473.179 | 890 | 0.20953 | 0.000439 | -2e-06 | 107.9844 | 109.2797 | 127.1683 | 20.9236 | 0.001972 | 0.001929 | -152.1397 | 0.23894 | 8.4036 | 0.00334 | 0.004995 | 0.004829 | 7.1e-05 | 0.001129 | 0.335 | 1.2992 | 2.5481 | 100.99 | 20.71 | 18.279 | 68.949 | 101.8809 | 67.9194 | 0.004427 | 1.0351 | 0.025096 | 24.245 | -0.10665 | 0.14214 | 0.000143 | 0.010899 | 0.010899 | 0.536
2019-03-20 13:05 | 1553101474.1791 | 891 | 0.20949 | 0.000439 | -1e-06 | 107.8803 | 109.079 | 126.9641 | 20.9225 | 0.002011 | 0.001889 | -152.1376 | 0.23939 | 8.403 | 0.003342 | 0.005112 | 0.004923 | 6.7e-05 | 0.00113 | 0.34003 | 1.3654 | 2.548 | 100.99 | 20.71 | 18.274 | 69.8924 | 101.931 | 66.9573 | 0.004425 | 1.038 | 0.025112 | 24.1913 | -0.09884 | 0.14241 | 0.000436 | 0.033148 | 0.033148 | 0.17657
2019-03-20 13:05 | 1553101475.1792 | 892 | 0.20955 | 0.00044 | -2e-06 | 107.933 | 109.1978 | 126.7306 | 20.9237 | 0.002026 | 0.001899 | -152.1392 | 0.23902 | 8.4043 | 0.003313 | 0.004964 | 0.004735 | 6.8e-05 | 0.001128 | 0.33572 | 1.3066 | 2.5482 | 100.99 | 20.71 | 18.27 | 70.2482 | 101.2365 | 67.1979 | 0.004425 | 1.0373 | 0.02508 | 24.1776 | -0.10421 | 0.14219 | 6.6e-05 | 0.005107 | 0.005107 | 1.1444
2019-03-20 13:05 | 1553101476.1792 | 893 | 0.20956 | 0.000439 | -1e-06 | 107.9427 | 109.1844 | 126.7813 | 20.9234 | 0.002021 | 0.001953 | -152.1632 | 0.2395 | 8.4048 | 0.003277 | 0.004807 | 0.004623 | 7.2e-05 | 0.001126 | 0.33323 | 1.3567 | 2.548 | 100.99 | 20.71 | 18.265 | 70.1365 | 100.3635 | 68.4864 | 0.004423 | 1.0334 | 0.025037 | 24.2282 | -0.1023 | 0.14248 | -2.5e-05 | -0.001792 | -0.001792 | -3.267
2019-03-20 13:05 | 1553101477.1786 | 894 | 0.20956 | 0.000439 | -1e-06 | 107.964 | 109.292 | 126.6663 | 20.9233 | 0.001924 | 0.001949 | -152.1556 | 0.23881 | 8.4026 | 0.003398 | 0.004067 | 0.00395 | 7.1e-05 | 0.001127 | 0.33804 | 1.3456 | 2.5482 | 100.99 | 20.71 | 18.258 | 67.7881 | 103.2782 | 68.3949 | 0.004422 | 1.0337 | 0.025029 | 24.214 | -0.10935 | 0.14207 | 1.2e-05 | 0.001007 | 0.001007 | 5.7971
2019-03-20 13:05 | 1553101478.1783 | 895 | 0.20955 | 0.000439 | -1e-06 | 107.9537 | 109.2079 | 127.0289 | 20.9233 | 0.002028 | 0.001914 | -152.1606 | 0.23954 | 8.4044 | 0.003343 | 0.004728 | 0.004532 | 6.8e-05 | 0.001131 | 0.3413 | 1.3106 | 2.5482 | 100.99 | 20.71 | 18.253 | 70.3024 | 101.9467 | 67.5462 | 0.00442 | 1.0362 | 0.025096 | 24.2184 | -0.10332 | 0.1425 | 7.6e-05 | 0.005819 | 0.005819 | 1.0065
2019-03-20 13:05 | 1553101479.1788 | 896 | 0.20954 | 0.000445 | -1e-06 | 108.0235 | 109.2969 | 127.2744 | 20.9245 | 0.002009 | 0.001855 | -152.1465 | 0.23897 | 8.4027 | 0.003323 | 0.004002 | 0.003766 | 6.7e-05 | 0.001135 | 0.35513 | 1.3449 | 2.548 | 100.99 | 20.71 | 18.244 | 69.8436 | 101.4791 | 66.1353 | 0.004418 | 1.0406 | 0.025173 | 24.1916 | -0.10482 | 0.14216 | 0.000116 | 0.008859 | 0.008859 | 0.65953
2019-03-20 13:05 | 1553101480.1784 | 897 | 0.20953 | 0.000443 | -1e-06 | 108.0666 | 109.3608 | 127.0345 | 20.9235 | 0.001951 | 0.001924 | -152.1366 | 0.23933 | 8.4019 | 0.003479 | 0.004432 | 0.004229 | 6.8e-05 | 0.001137 | 0.36586 | 1.3479 | 2.5479 | 100.99 | 20.71 | 18.242 | 68.4489 | 105.2519 | 67.7984 | 0.004418 | 1.0355 | 0.025088 | 24.2279 | -0.10649 | 0.14238 | 0.000171 | 0.012986 | 0.012986 | 0.4506
2019-03-20 13:05 | 1553101481.1785 | 898 | 0.20955 | 0.000444 | -1e-06 | 108.0759 | 109.304 | 126.9897 | 20.9234 | 0.001947 | 0.001907 | -152.1393 | 0.23876 | 8.404 | 0.003458 | 0.00439 | 0.004243 | 6.4e-05 | 0.001135 | 0.33804 | 1.3417 | 2.5481 | 100.99 | 20.71 | 18.242 | 68.3352 | 104.7438 | 67.3757 | 0.004418 | 1.0368 | 0.025099 | 24.2086 | -0.10106 | 0.14204 | 6.1e-05 | 0.004673 | 0.004673 | 1.2491
2019-03-20 13:05 | 1553101482.1786 | 899 | 0.20958 | 0.000442 | -2e-06 | 108.105 | 109.2502 | 126.7989 | 20.9243 | 0.001975 | 0.001936 | -152.1339 | 0.23884 | 8.4029 | 0.003311 | 0.004745 | 0.004466 | 6.4e-05 | 0.001137 | 0.35762 | 1.3231 | 2.5482 | 100.99 | 20.71 | 18.237 | 69.0276 | 101.1836 | 68.0778 | 0.004417 | 1.0346 | 0.025054 | 24.2154 | -0.094248 | 0.14209 | -0.000152 | -0.011391 | -0.011391 | -0.51266
2019-03-20 13:05 | 1553101483.1789 | 900 | 0.20953 | 0.000437 | -1e-06 | 108.1371 | 109.3758 | 126.7414 | 20.9242 | 0.001938 | 0.001902 | -152.1306 | 0.23837 | 8.4008 | 0.003345 | 0.004717 | 0.004574 | 6.6e-05 | 0.001137 | 0.34451 | 1.3185 | 2.5481 | 100.99 | 20.71 | 18.234 | 68.1306 | 102.004 | 67.261 | 0.004416 | 1.0371 | 0.025078 | 24.1809 | -0.10187 | 0.1418 | 0.000193 | 0.014713 | 0.014713 | 0.39613
2019-03-20 13:05 | 1553101484.18 | 901 | 0.20954 | 0.000438 | -1e-06 | 108.1937 | 109.6112 | 126.7691 | 20.9245 | 0.002056 | 0.001904 | -152.1309 | 0.23933 | 8.4027 | 0.003365 | 0.004457 | 0.004103 | 6.2e-05 | 0.001134 | 0.35242 | 1.315 | 2.5481 | 100.99 | 20.71 | 18.232 | 70.9609 | 102.4914 | 67.3166 | 0.004415 | 1.0369 | 0.025079 | 24.1855 | -0.11643 | 0.14238 | 9.9e-05 | 0.007607 | 0.007607 | 0.76928
2019-03-20 13:05 | 1553101485.1778 | 902 | 0.20955 | 0.000443 | -2e-06 | 108.2047 | 109.44 | 126.6273 | 20.924 | 0.001953 | 0.001902 | -152.1411 | 0.23843 | 8.4021 | 0.003298 | 0.004274 | 0.004154 | 6.5e-05 | 0.001137 | 0.34274 | 1.3187 | 2.548 | 100.99 | 20.71 | 18.231 | 68.4807 | 100.8749 | 67.263 | 0.004415 | 1.0371 | 0.025067 | 24.1701 | -0.10154 | 0.14184 | 7.7e-05 | 0.005971 | 0.005971 | 0.97635
2019-03-20 13:05 | 1553101486.1788 | 903 | 0.20956 | 0.000438 | -1e-06 | 108.2771 | 109.5802 | 126.4599 | 20.924 | 0.001998 | 0.001885 | -152.1452 | 0.23917 | 8.4047 | 0.003287 | 0.004817 | 0.004561 | 6.3e-05 | 0.00113 | 0.33268 | 1.3153 | 2.548 | 100.99 | 20.71 | 18.229 | 69.5803 | 100.5923 | 66.8519 | 0.004415 | 1.0384 | 0.025066 | 24.1395 | -0.107 | 0.14228 | -2.5e-05 | -0.00178 | -0.00178 | -3.2856
2019-03-20 13:05 | 1553101487.1791 | 904 | 0.20953 | 0.000439 | -1e-06 | 108.215 | 109.4696 | 126.6684 | 20.9237 | 0.002029 | 0.001861 | -152.1473 | 0.2388 | 8.403 | 0.003171 | 0.004944 | 0.004688 | 6.3e-05 | 0.00113 | 0.323 | 1.3029 | 2.548 | 100.99 | 20.71 | 18.23 | 70.324 | 97.7858 | 66.2859 | 0.004415 | 1.0401 | 0.025107 | 24.1393 | -0.1031 | 0.14206 | 0.000173 | 0.013205 | 0.013205 | 0.44215
2019-03-20 13:05 | 1553101488.1779 | 905 | 0.20956 | 0.000442 | -1e-06 | 108.2396 | 109.5219 | 126.2822 | 20.9244 | 0.001994 | 0.001844 | -152.1396 | 0.23972 | 8.4044 | 0.003246 | 0.004362 | 0.004183 | 6.2e-05 | 0.001127 | 0.32388 | 1.3293 | 2.5481 | 100.99 | 20.71 | 18.231 | 69.488 | 99.5986 | 65.8612 | 0.004415 | 1.0414 | 0.025085 | 24.0874 | -0.10534 | 0.14261 | -3e-06 | -0.000147 | -0.000147 | -39.7898
2019-03-20 13:05 | 1553101489.1789 | 906 | 0.20958 | 0.000439 | -1e-06 | 108.2726 | 109.5828 | 126.4036 | 20.9241 | 0.00196 | 0.001913 | -152.1515 | 0.23911 | 8.4046 | 0.003326 | 0.004716 | 0.004469 | 6.2e-05 | 0.001124 | 0.34811 | 1.3124 | 2.5481 | 100.99 | 20.71 | 18.236 | 68.65 | 101.5525 | 67.5285 | 0.004416 | 1.0363 | 0.025035 | 24.1581 | -0.10759 | 0.14225 | -0.000123 | -0.009263 | -0.009263 | -0.63116
2019-03-20 13:05 | 1553101490.178 | 907 | 0.20952 | 0.000438 | 0 | 108.3373 | 109.5316 | 126.5207 | 20.9237 | 0.002005 | 0.001867 | -152.1564 | 0.23878 | 8.4021 | 0.003226 | 0.004539 | 0.00431 | 6.3e-05 | 0.001123 | 0.34075 | 1.2917 | 2.5481 | 100.99 | 20.71 | 18.241 | 69.7501 | 99.1226 | 66.4311 | 0.004417 | 1.0396 | 0.025087 | 24.1304 | -0.098068 | 0.14205 | 0.000261 | 0.019772 | 0.019772 | 0.29527
2019-03-20 13:05 | 1553101491.178 | 908 | 0.20957 | 0.000443 | -1e-06 | 108.269 | 109.5922 | 126.5073 | 20.9249 | 0.002029 | 0.001863 | -152.1595 | 0.23921 | 8.4033 | 0.003298 | 0.005207 | 0.004937 | 6.2e-05 | 0.001126 | 0.34739 | 1.3342 | 2.5481 | 100.99 | 20.72 | 18.25 | 70.329 | 100.8773 | 66.3277 | 0.004422 | 1.04 | 0.02509 | 24.1254 | -0.10865 | 0.14231 | -8.8e-05 | -0.006604 | -0.006604 | -0.88565
2019-03-20 13:05 | 1553101492.1795 | 909 | 0.20955 | 0.000443 | -1e-06 | 108.3424 | 109.705 | 127.1186 | 20.9231 | 0.001999 | 0.001953 | -152.1616 | 0.23904 | 8.4031 | 0.003277 | 0.004773 | 0.004618 | 6.1e-05 | 0.001124 | 0.31006 | 1.3101 | 2.548 | 100.99 | 20.72 | 18.254 | 69.5889 | 100.3671 | 68.4815 | 0.004423 | 1.0334 | 0.025071 | 24.2602 | -0.11179 | 0.1422 | 6e-05 | 0.004609 | 0.004609 | 1.268
2019-03-20 13:05 | 1553101493.1792 | 910 | 0.20957 | 0.000439 | -2e-06 | 108.3602 | 109.8097 | 127.4434 | 20.9239 | 0.001955 | 0.001872 | -152.1651 | 0.23861 | 8.3996 | 0.00327 | 0.005251 | 0.005025 | 6e-05 | 0.001121 | 0.33107 | 1.3125 | 2.5481 | 100.99 | 20.72 | 18.257 | 68.539 | 100.1818 | 66.5335 | 0.004424 | 1.0393 | 0.025175 | 24.2218 | -0.11885 | 0.14195 | -7.1e-05 | -0.005263 | -0.005263 | -1.1086
2019-03-20 13:05 | 1553101494.1782 | 911 | 0.20953 | 0.000441 | -2e-06 | 108.4146 | 109.6239 | 127.5519 | 20.9236 | 0.002005 | 0.001902 | -152.1917 | 0.23836 | 8.4014 | 0.003269 | 0.004622 | 0.004458 | 6.3e-05 | 0.00112 | 0.32101 | 1.3164 | 2.548 | 100.99 | 20.72 | 18.261 | 69.7406 | 100.1769 | 67.2542 | 0.004425 | 1.0371 | 0.025159 | 24.2578 | -0.099224 | 0.1418 | 0.000161 | 0.012363 | 0.012363 | 0.47141
2019-03-20 13:05 | 1553101495.1796 | 912 | 0.20954 | 0.00044 | -1e-06 | 108.4508 | 109.8209 | 127.5321 | 20.9243 | 0.001942 | 0.001905 | -152.1771 | 0.23831 | 8.4028 | 0.003314 | 0.004906 | 0.004689 | 6.5e-05 | 0.00112 | 0.3079 | 1.3437 | 2.5481 | 100.99 | 20.72 | 18.264 | 68.2158 | 101.2575 | 67.3406 | 0.004426 | 1.0369 | 0.025154 | 24.259 | -0.11229 | 0.14177 | 9.3e-05 | 0.007118 | 0.007118 | 0.81865
2019-03-20 13:05 | 1553101496.1792 | 913 | 0.20954 | 0.000441 | -2e-06 | 108.3812 | 109.7734 | 127.3641 | 20.9247 | 0.001955 | 0.001884 | -152.1654 | 0.23898 | 8.4037 | 0.003289 | 0.004888 | 0.00472 | 5.9e-05 | 0.001122 | 0.30718 | 1.3425 | 2.5481 | 100.99 | 20.72 | 18.267 | 68.5352 | 100.644 | 66.8315 | 0.004427 | 1.0384 | 0.025156 | 24.2249 | -0.11416 | 0.14217 | 0.000114 | 0.008789 | 0.008789 | 0.66481
2019-03-20 13:05 | 1553101497.1793 | 914 | 0.20951 | 0.000438 | -2e-06 | 108.4369 | 109.9345 | 127.5915 | 20.9247 | 0.001983 | 0.001882 | -152.1555 | 0.2394 | 8.4025 | 0.003306 | 0.004781 | 0.004506 | 6e-05 | 0.001126 | 0.33987 | 1.3463 | 2.5481 | 100.99 | 20.72 | 18.269 | 69.2097 | 101.0609 | 66.7855 | 0.004427 | 1.0386 | 0.02518 | 24.2449 | -0.12269 | 0.14242 | 0.000307 | 0.023436 | 0.023436 | 0.24977
2019-03-20 13:05 | 1553101498.1797 | 915 | 0.2096 | 0.000446 | -1e-06 | 108.4378 | 109.9573 | 127.5441 | 20.9254 | 0.001948 | 0.001911 | -152.1654 | 0.23902 | 8.4034 | 0.003314 | 0.004779 | 0.00466 | 6.2e-05 | 0.001126 | 0.3298 | 1.3353 | 2.5481 | 100.99 | 20.72 | 18.272 | 68.3775 | 101.2632 | 67.4711 | 0.004428 | 1.0365 | 0.02515 | 24.2648 | -0.12446 | 0.14219 | -0.000235 | -0.01778 | -0.01778 | -0.3287
2019-03-20 13:05 | 1553101499.1794 | 916 |  | 0.000439 |  | 108.5145 | 110.0119 | 127.4838 | 20.9245 | 0.001936 | 0.001904 | -152.167 | 0.23858 | 8.4023 | 0.003293 | 0.004469 | 0.004202 | 6.1e-05 | 0.001132 | 0.34396 | 1.3377 | 2.5482 | 100.99 | 20.72 | 18.278 | 68.0927 | 100.7593 | 67.3089 | 0.004429 | 1.037 | 0.02515 | 24.2533 | -0.12258 | 0.14193
2019-03-20 13:05 | 1553101500.1794 | 917 |  | 0.000441 |  | 108.5119 | 109.9495 | 127.1953 | 20.9234 | 0.001928 | 0.001906 | -152.1718 | 0.23825 | 8.4014 | 0.003348 | 0.004921 | 0.004735 | 6.1e-05 | 0.001129 | 0.31332 | 1.3158 | 2.5481 | 100.99 | 20.72 | 18.279 | 67.8943 | 102.0749 | 67.3615 | 0.004429 | 1.0368 | 0.02512 | 24.2277 | -0.11772 | 0.14174
2019-03-20 13:05 | 1553101501.1794 | 918 |  |  |  | 108.4184 | 109.7465 | 126.9646 | 20.9245 | 0.001984 | 0.001918 | -152.1772 | 0.23926 | 8.404 | 0.00334 | 0.004775 | 0.004565 | 6.1e-05 | 0.001131 | 0.30055 | 1.3079 | 2.548 | 100.99 | 20.73 | 18.28 | 69.2307 | 101.8891 | 67.659 | 0.004432 | 1.0359 | 0.025086 | 24.2163 | -0.1089 | 0.14234
2019-03-20 13:05 | 1553101502.1794 | 919 |  |  |  | 108.5488 | 110.022 | 126.7341 | 20.9233 | 0.001951 | 0.001884 | -152.1652 | 0.23855 | 8.402 | 0.003277 | 0.004547 | 0.004268 | 6.2e-05 | 0.001133 | 0.31598 | 1.2976 | 2.548 | 100.99 | 20.73 | 18.283 | 68.4411 | 100.3491 | 66.8419 | 0.004433 | 1.0384 | 0.025093 | 24.1653 | -0.12058 | 0.14191
2019-03-20 13:05 | 1553101503.1794 | 920 |  |  |  | 108.4737 | 109.7637 | 126.3778 | 20.924 | 0.001997 | 0.001902 | -152.1608 | 0.23985 | 8.4024 | 0.003323 | 0.005079 | 0.004846 | 6.4e-05 | 0.001139 | 0.31797 | 1.3269 | 2.548 | 100.99 | 20.73 | 18.285 | 69.5415 | 101.4716 | 67.2764 | 0.004434 | 1.0371 | 0.025042 | 24.1467 | -0.10574 | 0.14269
2019-03-20 13:05 | 1553101504.1794 | 921 |  |  |  | 108.5934 | 109.9117 |  | 20.9236 | 0.002 | 0.001924 | -152.1557 | 0.23879 | 8.4033 | 0.003312 | 0.005002 | 0.004775 | 6e-05 | 0.001139 | 0.35745 | 1.3218 | 2.548 | 100.99 | 20.73 | 18.293 | 69.6276 | 101.2057 | 67.7978 | 0.004436 | 1.0355 |  |  | -0.10793 | 0.14205
2019-03-20 13:05 | 1553101505.1794 | 922 |  |  |  | 108.5759 | 110.0552 |  | 20.9243 | 0.001998 |  | -152.167 | 0.23877 | 8.4034 | 0.003377 | 0.004705 | 0.004521 | 6.3e-05 | 0.00114 | 0.33428 | 1.3424 | 2.5481 | 100.99 | 20.73 | 18.298 | 69.5814 | 102.7764 |  | 0.004437 |  |  |  | -0.12104 | 0.14204
2019-03-20 13:05 | 1553101506.1794 | 923 |  |  |  | 108.5995 | 110.0459 |  | 20.9241 | 0.001964 |  |  | 0.23913 | 8.4015 | 0.00336 | 0.004824 | 0.004589 | 6.5e-05 | 0.001144 | 0.33682 | 1.3049 | 2.548 | 100.99 | 20.73 | 18.302 | 68.7553 | 102.374 |  | 0.004438 |  |  |  | -0.11834 | 0.14226
2019-03-20 13:05 | 1553101507.1794 | 924 |  |  |  | 108.6303 | 110.0246 |  | 20.9241 |  |  |  | 0.23862 | 8.4013 | 0.003314 | 0.004419 | 0.004182 | 5.9e-05 | 0.001146 | 0.3366 | 1.3159 | 2.5481 | 100.99 | 20.73 | 18.302 |  | 101.2587 |  | 0.004438 |  |  |  | -0.11408 | 0.14195
2019-03-20 13:05 | 1553101508.1794 | 925 |  |  |  | 108.6836 | 109.9977 |  | 20.9254 |  |  |  | 0.23896 | 8.4032 | 0.003287 | 0.004773 | 0.004632 | 6e-05 | 0.001148 | 0.31094 | 1.3797 | 2.548 | 100.99 | 20.73 | 18.312 |  | 100.592 |  | 0.00444 |  |  |  | -0.10749 | 0.14216
2019-03-20 13:05 | 1553101509.1794 | 926 |  |  |  | 108.6235 | 109.9857 |  | 20.9246 |  |  |  | 0.23947 | 8.4036 | 0.003273 | 0.004968 | 0.004749 | 6.2e-05 | 0.001146 | 0.30829 | 1.3196 | 2.548 | 100.99 | 20.73 | 18.317 |  | 100.2671 |  | 0.004441 |  |  |  | -0.11148 | 0.14246
2019-03-20 13:05 | 1553101510.1794 | 927 |  |  |  | 108.6859 | 109.9895 |  | 20.9243 |  |  |  | 0.23958 | 8.4015 | 0.0033 | 0.004643 | 0.004471 | 6.4e-05 | 0.001147 | 0.32333 | 1.303 | 2.548 | 100.99 | 20.73 | 18.319 |  | 100.9144 |  | 0.004442 |  |  |  | -0.10664 | 0.14252
2019-03-20 13:05 | 1553101511.1794

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-20 12:50 | 0 | 0 | 0 | 0 | 0 | 45.2 | -0.016754
2019-03-20 12:50 | 1 | 0 | 0 | 0 | 0 | 45.1574 | -0.010958
2019-03-20 12:50 | 2 | 0 | 0 | 0 | 0 | 45.3951 | -0.010001
2019-03-20 12:50 | 3 | 0 | 0 | 0 | 0 | 45.4555 | -0.025973
2019-03-20 12:50 | 4 | 0 | 0 | 0 | 0 | 45.4484 | 0.083699
2019-03-20 12:50 | 5 | 0 | 0 | 0 | 0 | 45.3204 | 0.010729
2019-03-20 12:50 | 6 | 0 | 0 | 0 | 0 | 45.2136 | 0.19228
2019-03-20 12:50 | 7 | 0 | 0 | 0 | 0 | 45.1552 | 0.56614
2019-03-20 12:50 | 8 | 0 | 0 | 0 | 0 | 45.0604 | 1.3801
2019-03-20 12:50 | 9 | 0 | 0 | 0 | 0 | 44.9613 | 1.9448
2019-03-20 12:50 | 10 | 0 | 0 | 0 | 0 | 44.8719 | 2.4649
2019-03-20 12:50 | 11 | 0 | 0 | 0 | 0 | 44.729 | 2.8068
2019-03-20 12:50 | 12 | 0 | 0 | 0 | 0 | 44.6087 | 3.087
2019-03-20 12:50 | 13 | 0 | 0 | 0 | 0 | 44.5049 | 3.3093
2019-03-20 12:50 | 14 | 0.12211 | 0 | 0 | 0 | 44.3741 | 3.4596
2019-03-20 12:50 | 15 | 0.1257 | 0 | 0 | 0 | 44.2608 | 3.6718
2019-03-20 12:50 | 16 | 0.14 | 0 | 0 | 0 | 44.1213 | 3.7395
2019-03-20 12:50 | 17 | 0.10958 | 0 | 0 | 0 | 43.9908 | 3.733
2019-03-20 12:50 | 18 | 0.1123 | 0 | 0 | 0 | 43.8943 | 3.7186
2019-03-20 12:50 | 19 | 0.14333 | 0 | 0 | 0 | 43.7598 | 3.6546
2019-03-20 12:50 | 20 | 0.11909 | 0 | 0 | 0 | 43.6205 | 3.5576
2019-03-20 12:50 | 21 | 0.10113 | 0 | 0 | 0 | 43.5206 | 3.4308
2019-03-20 12:50 | 22 | 0.11577 | 0 | 0 | 0 | 43.4104 | 3.1914
2019-03-20 12:50 | 23 | 0.11666 | 0 | 0 | 0 | 43.2921 | 2.9738
2019-03-20 12:50 | 24 | 0.099517 | 0 | 0 | 0 | 43.1817 | 2.937
2019-03-20 12:50 | 25 | 0.089062 | 0 | 0 | 0 | 43.0912 | 2.5936
2019-03-20 12:50 | 26 | 0.10477 | 0 | 0 | 0 | 42.9973 | 2.5146
2019-03-20 12:50 | 27 | 0.10473 | 0 | 0 | 0 | 42.8856 | 2.4031
2019-03-20 12:50 | 28 | 0.079808 | 0 | 0 | 0 | 42.7941 | 2.2899
2019-03-20 12:50 | 29 | 0.068567 | 0 | 0 | 0 | 42.7223 | 2.2538
2019-03-20 12:50 | 30 | 0.089716 | 0 | 0 | 0 | 42.6487 | 2.2232
2019-03-20 12:50 | 31 | 0.10029 | 0 | 0 | 0 | 42.5453 | 2.1927
2019-03-20 12:50 | 32 | 0.084393 | 0 | 0 | 0 | 42.4549 | 2.1668
2019-03-20 12:50 | 33 | 0.07295 | 0 | 0 | 0 | 42.3755 | 2.1417
2019-03-20 12:50 | 34 | 0.088308 | 0 | 0 | 0 | 42.303 | 2.0708
2019-03-20 12:50 | 35 | 0.074978 | 0 | 0 | 0 | 42.2054 | 1.9902
2019-03-20 12:50 | 36 | 0.057656 | 0 | 0 | 0 | 42.1534 | 1.8572
2019-03-20 12:50 | 37 | 0.088577 | 0 | 0 | 0 | 42.0788 | 1.7647
2019-03-20 12:50 | 38 | 0.093907 | 0 | 0 | 0 | 41.9821 | 1.6612
2019-03-20 12:50 | 39 | 0.075998 | 0 | 0 | 0 | 41.8975 | 1.6165
2019-03-20 12:50 | 40 | 0.06121 | 0 | 0 | 0 | 41.8297 | 1.5424
2019-03-20 12:50 | 41 | 0.056995 | 0 | 0 | 0 | 41.772 | 1.4509
2019-03-20 12:50 | 42 | 0.06152 | 0 | 0 | 0 | 41.7132 | 1.399
2019-03-20 12:50 | 43 | 0.064869 | 0 | 0 | 0 | 41.6488 | 1.3391
2019-03-20 12:50 | 44 | 0.073231 | 0 | 0 | 0 | 41.5822 | 1.2876
2019-03-20 12:50 | 45 | 0.083633 | 0 | 0 | 0 | 41.5026 | 1.2906
2019-03-20 12:50 | 46 | 0.064295 | 0 | 0 | 0 | 41.422 | 1.273
2019-03-20 12:50 | 47 | 0.060687 | 0 | 0 | 0 | 41.3708 | 1.2914
2019-03-20 12:50 | 48 | 0.05858 | 0 | 0 | 0 | 41.2997 | 1.2372
2019-03-20 12:50 | 49 | 0.062121 | 0 | 0 | 0 | 41.2524 | 1.2394
2019-03-20 12:51 | 50 | 0.058509 | 0 | 0 | 0 | 41.1773 | 1.232
2019-03-20 12:51 | 51 | 0.044265 | 0 | 0 | 0 | 41.1374 | 1.2334
2019-03-20 12:51 | 52 | 0.061626 | 0 | 0 | 0 | 41.0818 | 1.2391
2019-03-20 12:51 | 53 | 0.056557 | 0 | 0 | 0 | 41.0192 | 1.2293
2019-03-20 12:51 | 54 | 0.041836 | 0 | 0 | 0 | 40.9712 | 1.2259
2019-03-20 12:51 | 55 | 0.043046 | 0 | 0 | 0 | 40.9314 | 1.2174
2019-03-20 12:51 | 56 | 0.056913 | 0 | 0 | 0 | 40.8821 | 1.214
2019-03-20 12:51 | 57 | 0.046257 | 0 | 0 | 0 | 40.8232 | 1.2043
2019-03-20 12:51 | 58 | 0.045225 | 0.020624 | 2.6789 | 27.2792 | 40.7878 | 1.2337
2019-03-20 12:51 | 59 | 0.045814 | 0.020361 | 2.6159 | 26.6558 | 40.732 | 1.2212
2019-03-20 12:51 | 60 | 0.048802 | 0.018685 | 2.4459 | 25.2927 | 40.6952 | 1.2343
2019-03-20 12:51 | 61 | 0.061947 | 0.014744 | 1.9218 | 19.5787 | 40.6325 | 1.2128
2019-03-20 12:51 | 62 | 0.047527 | 0.018744 | 2.4908 | 26.1291 | 40.5775 | 1.2418
2019-03-20 12:51 | 63 | 0.052558 | 0.016917 | 2.2358 | 23.5841 | 40.5331 | 1.2395
2019-03-20 12:51 | 64 | 0.066742 | 0.013342 | 1.7579 | 18.4417 | 40.4703 | 1.2308
2019-03-20 12:51 | 65 | 0.053417 | 0.016646 | 2.1885 | 22.8792 | 40.4063 | 1.2221
2019-03-20 12:51 | 66 | 0.036117 | 0.024636 | 3.2322 | 34.0205 | 40.3643 | 1.2287
2019-03-20 12:51 | 67 | 0.05172 | 0.018092 | 2.2543 | 23.8628 | 40.326 | 1.2342
2019-03-20 12:51 | 68 | 0.064926 | 0.014419 | 1.7928 | 19.01 | 40.261 | 1.2342
2019-03-20 12:51 | 69 | 0.06339 | 0.014408 | 1.835 | 19.6534 | 40.2 | 1.2458
2019-03-20 12:51 | 70 | 0.057148 | 0.015976 | 2.0298 | 21.7096 | 40.1357 | 1.2407
2019-03-20 12:51 | 71 | 0.053801 | 0.016939 | 2.15 | 23.2307 | 40.0846 | 1.2498
2019-03-20 12:51 | 72 | 0.072593 | 0.012883 | 1.5944 | 17.1688 | 40.0235 | 1.2463
2019-03-20 12:51 | 73 | 0.054775 | 0.017092 | 2.1153 | 22.6648 | 39.948 | 1.2415
2019-03-20 12:51 | 74 | 0.035664 | 0.026826 | 3.2399 | 35.0413 | 39.9138 | 1.2497
2019-03-20 12:51 | 75 | 0.066672 | 0.014356 | 1.7333 | 18.4279 | 39.8645 | 1.2286
2019-03-20 12:51 | 76 | 0.083098 | 0.011803 | 1.3924 | 15.0919 | 39.7837 | 1.2541
2019-03-20 12:51 | 77 | 0.065842 | 0.014919 | 1.76 | 19.4639 | 39.7069 | 1.2815
2019-03-20 12:51 | 78 | 0.058752 | 0.016706 | 1.9713 | 23.0248 | 39.65 | 1.3528
2019-03-20 12:51 | 79 | 0.071602 | 0.013685 | 1.6224 | 16.9288 | 39.5848 | 1.2121
2019-03-20 12:51 | 80 | 0.061058 | 0.016045 | 1.9057 | 21.1878 | 39.5119 | 1.2937
2019-03-20 12:51 | 81 | 0.053755 | 0.01781 | 2.1704 | 23.7246 | 39.4621 | 1.2753
2019-03-20 12:51 | 82 | 0.057772 | 0.016564 | 2.0253 | 21.9311 | 39.401 | 1.267
2019-03-20 12:51 | 83 | 0.081586 | 0.011432 | 1.4324 | 15.7251 | 39.3422 | 1.2829
2019-03-20 12:51 | 84 | 0.072255 | 0.012905 | 1.6224 | 17.7088 | 39.2461 | 1.2796
2019-03-20 12:51 | 85 | 0.047265 | 0.019724 | 2.4789 | 27.4189 | 39.2018 | 1.2959
2019-03-20 12:51 | 86 | 0.05056 | 0.018393 | 2.3136 | 25.7005 | 39.1452 | 1.2994
2019-03-20 12:51 | 87 | 0.044071 | 0.021635 | 2.6639 | 29.9293 | 39.102 | 1.319
2019-03-20 12:51 | 88 | 0.049433 | 0.019224 | 2.3671 | 26.3817 | 39.0545 | 1.3041
2019-03-20 12:51 | 89 | 0.041446 | 0.0235 | 2.8357 | 31.6041 | 39.0055 | 1.3099
2019-03-20 12:51 | 90 | 0.056761 | 0.017523 | 2.0746 | 22.9705 | 38.9665 | 1.3038
2019-03-20 12:51 | 91 | 0.059472 | 0.017133 | 1.9846 | 22.1384 | 38.8949 | 1.3166
2019-03-20 12:51 | 92 | 0.049386 | 0.020625 | 2.3998 | 26.6774 | 38.8509 | 1.3175
2019-03-20 12:51 | 93 | 0.044695 | 0.022776 | 2.6533 | 29.5021 | 38.7946 | 1.3186
2019-03-20 12:51 | 94 | 0.061556 | 0.016901 | 1.9313 | 21.5193 | 38.7562 | 1.3246
2019-03-20 12:51 | 95 | 0.067387 | 0.015117 | 1.7663 | 19.677 | 38.6739 | 1.326
2019-03-20 12:51 | 96 | 0.05922 | 0.01762 | 2.0194 | 22.5524 | 38.6245 | 1.3356
2019-03-20 12:51 | 97 | 0.081102 | 0.01286 | 1.4761 | 16.3137 | 38.549 | 1.3231
2019-03-20 12:51 | 98 | 0.080541 | 0.013282 | 1.4908 | 16.7307 | 38.4676 | 1.3475
2019-03-20 12:51 | 99 | 0.06654 | 0.016042 | 1.8038 | 20.0727 | 38.3921 | 1.3356
2019-03-20 12:51 | 100 | 0.047784 | 0.022826 | 2.5176 | 28.2378 | 38.3354 | 1.3493
2019-03-20 12:51 | 101 | 0.054149 | 0.02013 | 2.2302 | 25.5509 | 38.2903 | 1.3836
2019-03-20 12:51 | 102 | 0.061133 | 0.017814 | 1.9817 | 22.2266 | 38.2263 | 1.3588
2019-03-20 12:51 | 103 | 0.065894 | 0.016516 | 1.8455 | 20.6608 | 38.1689 | 1.3614
2019-03-20 12:51 | 104 | 0.053127 | 0.020066 | 2.2986 | 25.7957 | 38.0984 | 1.3704
2019-03-20 12:51 | 105 | 0.063282 | 0.017207 | 1.935 | 21.8718 | 38.0576 | 1.3841
2019-03-20 12:51 | 106 | 0.068839 | 0.015468 | 1.7832 | 20.3565 | 37.9732 | 1.4013
2019-03-20 12:51 | 107 | 0.0431 | 0.024637 | 2.8525 | 32.3801 | 37.9272 | 1.3956
2019-03-20 12:51 | 108 | 0.041843 | 0.025412 | 2.9534 | 33.6274 | 37.881 | 1.4071
2019-03-20 12:52 | 109 | 0.064796 | 0.016429 | 1.9165 | 21.6521 | 37.8368 | 1.403
2019-03-20 12:52 | 110 | 0.085096 | 0.012493 | 1.464 | 16.4488 | 37.7524 | 1.3997
2019-03-20 12:52 | 111 | 0.068046 | 0.015653 | 1.8422 | 20.6595 | 37.6758 | 1.4058
2019-03-20 12:52 | 112 | 0.062928 | 0.017294 | 1.9969 | 22.286 | 37.6144 | 1.4024
2019-03-20 12:52 | 113 | 0.056451 | 0.019243 | 2.227 | 24.8487 | 37.5496 | 1.4027
2019-03-20 12:52 | 114 | 0.065975 | 0.016434 | 1.9061 | 21.2458 | 37.4979 | 1.4017
2019-03-20 12:52 | 115 | 0.057795 | 0.01837 | 2.181 | 24.2326 | 37.4214 | 1.4005
2019-03-20 12:52 | 116 | 0.056868 | 0.018255 | 2.2189 | 24.6815 | 37.3803 | 1.4036
2019-03-20 12:52 | 117 | 0.081141 | 0.01251 | 1.5596 | 17.0713 | 37.302 | 1.3852
2019-03-20 12:52 | 118 | 0.074485 | 0.01327 | 1.6906 | 18.7047 | 37.225 | 1.3932
2019-03-20 12:52 | 119 | 0.069713 | 0.014474 | 1.8069 | 20.1721 | 37.1532 | 1.4062
2019-03-20 12:52 | 120 | 0.071644 | 0.014072 | 1.7597 | 19.5248 | 37.0839 | 1.3988
2019-03-20 12:52 | 121 | 0.076661 | 0.013487 | 1.6513 | 18.5018 | 37.0094 | 1.4184
2019-03-20 12:52 | 122 | 0.067013 | 0.015752 | 1.896 | 21.3408 | 36.934 | 1.4301
2019-03-20 12:52 | 123 | 0.06349 | 0.016643 | 2.0071 | 22.7747 | 36.8737 | 1.446
2019-03-20 12:52 | 124 | 0.076503 | 0.013836 | 1.6737 | 19.0075 | 36.803 | 1.4541
2019-03-20 12:52 | 125 | 0.078066 | 0.013559 | 1.6497 | 18.5415 | 36.7238 | 1.4475
2019-03-20 12:52 | 126 | 0.056794 | 0.01823 | 2.2769 | 25.5807 | 36.6524 | 1.4528
2019-03-20 12:52 | 127 | 0.052085 | 0.020382 | 2.4969 | 27.9493 | 36.6058 | 1.4557
2019-03-20 12:52 | 128 | 0.080498 | 0.013178 | 1.6201 | 18.209 | 36.5403 | 1.4658
2019-03-20 12:52 | 129 | 0.080071 | 0.013538 | 1.6337 | 18.1752 | 36.4518 | 1.4553
2019-03-20 12:52 | 130 | 0.052696 | 0.021009 | 2.4855 | 27.6368 | 36.3867 | 1.4563
2019-03-20 12:52 | 131 | 0.064618 | 0.017118 | 2.0323 | 22.3392 | 36.3375 | 1.4435
2019-03-20 12:52 | 132 | 0.071397 | 0.015794 | 1.8346 | 20.2736 | 36.258 | 1.4475
2019-03-20 12:52 | 133 | 0.065233 | 0.017281 | 2.009 | 22.2924 | 36.1978 | 1.4542
2019-03-20 12:52 | 134 | 0.062852 | 0.0179 | 2.081 | 22.9841 | 36.1265 | 1.4446
2019-03-20 12:52 | 135 | 0.072963 | 0.015731 | 1.7915 | 19.6094 | 36.0685 | 1.4308
2019-03-20 12:52 | 136 | 0.099803 | 0.011489 | 1.3084 | 14.3353 | 35.9777 | 1.4307
2019-03-20 12:52 | 137 | 0.071988 | 0.016217 | 1.8137 | 19.7107 | 35.8824 | 1.4189
2019-03-20 12:52 | 138 | 0.038175 | 0.030516 | 3.4084 | 36.6545 | 35.8354 | 1.3993
2019-03-20 12:52 | 139 | 0.04888 | 0.024336 | 2.6662 | 29.1778 | 35.7945 | 1.4262
2019-03-20 12:52 | 140 | 0.081407 | 0.014599 | 1.6003 | 17.1635 | 35.7321 | 1.3972
2019-03-20 12:52 | 141 | 0.066484 | 0.017833 | 1.9516 | 21.3586 | 35.6423 | 1.42
2019-03-20 12:52 | 142 | 0.071626 | 0.016269 | 1.8152 | 19.8348 | 35.5953 | 1.4207
2019-03-20 12:52 | 143 | 0.08315 | 0.014008 | 1.5603 | 17.0642 | 35.4978 | 1.4189
2019-03-20 12:52 | 144 | 0.067949 | 0.017465 | 1.9056 | 21.0436 | 35.435 | 1.4299
2019-03-20 12:52 | 145 | 0.07291 | 0.016565 | 1.7716 | 19.0589 | 35.3578 | 1.3896
2019-03-20 12:52 | 146 | 0.066848 | 0.018451 | 1.9347 | 21.2632 | 35.2908 | 1.4214
2019-03-20 12:52 | 147 | 0.081964 | 0.015025 | 1.5754 | 17.277 | 35.2196 | 1.4161
2019-03-20 12:52 | 148 | 0.076403 | 0.016131 | 1.6942 | 18.3594 | 35.1316 | 1.4027
2019-03-20 12:52 | 149 | 0.070852 | 0.017108 | 1.8272 | 19.6904 | 35.0673 | 1.3951
2019-03-20 12:52 | 150 | 0.068004 | 0.017456 | 1.902 | 20.1494 | 34.9891 | 1.3702
2019-03-20 12:52 | 151 | 0.069923 | 0.017022 | 1.8497 | 19.4719 | 34.9301 | 1.3615
2019-03-20 12:52 | 152 | 0.069547 | 0.017099 | 1.8535 | 19.4021 | 34.8499 | 1.3494
2019-03-20 12:52 | 153 | 0.061726 | 0.019947 | 2.0745 | 21.9138 | 34.7925 | 1.3527
2019-03-20 12:52 | 154 | 0.068114 | 0.018415 | 1.8706 | 20.057 | 34.7233 | 1.3662
2019-03-20 12:52 | 155 | 0.064611 | 0.01977 | 1.9648 | 21.0308 | 34.6583 | 1.3588
2019-03-20 12:52 | 156 | 0.063421 | 0.020476 | 1.9894 | 21.7328 | 34.5934 | 1.3783
2019-03-20 12:52 | 157 | 0.072634 | 0.018546 | 1.7355 | 18.8494 | 34.5287 | 1.3691
2019-03-20 12:52 | 158 | 0.087006 | 0.015465 | 1.4455 | 15.692 | 34.4475 | 1.3653
2019-03-20 12:52 | 159 | 0.068175 | 0.020061 | 1.8434 | 19.9828 | 34.3625 | 1.3623
2019-03-20 12:52 | 160 | 0.061607 | 0.021806 | 2.0326 | 22.0914 | 34.3084 | 1.361
2019-03-20 12:52 | 161 | 0.076868 | 0.017534 | 1.6353 | 17.5804 | 34.2337 | 1.3514
2019-03-20 12:52 | 162 | 0.084688 | 0.016185 | 1.4859 | 15.7667 | 34.156 | 1.3353
2019-03-20 12:52 | 163 | 0.085676 | 0.016508 | 1.4646 | 15.6808 | 34.0664 | 1.3435
2019-03-20 12:52 | 164 | 0.081676 | 0.017644 | 1.5371 | 16.433 | 33.9861 | 1.3422
2019-03-20 12:52 | 165 | 0.081251 | 0.017436 | 1.5417 | 16.4569 | 33.9026 | 1.3371
2019-03-20 12:52 | 166 | 0.070149 | 0.020253 | 1.7878 | 18.6679 | 33.826 | 1.3095
2019-03-20 12:52 | 167 | 0.066548 | 0.020633 | 1.8752 | 20.105 | 33.7606 | 1.338
2019-03-20 12:52 | 168 | 0.067868 | 0.020209 | 1.8367 | 19.6589 | 33.6915 | 1.3342
2019-03-20 12:52 | 169 | 0.065788 | 0.020521 | 1.8929 | 20.0973 | 33.6252 | 1.3222
2019-03-20 12:53 | 170 | 0.073464 | 0.018407 | 1.6969 | 18.0096 | 33.5577 | 1.3231
2019-03-20 12:53 | 171 | 0.07653 | 0.01708 | 1.6256 | 17.3398 | 33.4796 | 1.327
2019-03-20 12:53 | 172 | 0.065373 | 0.019674 | 1.8979 | 20.2609 | 33.4079 | 1.3245
2019-03-20 12:53 | 173 | 0.071409 | 0.017704 | 1.7368 | 18.5727 | 33.3449 | 1.3263
2019-03-20 12:53 | 174 | 0.077468 | 0.015758 | 1.5985 | 17.0776 | 33.2652 | 1.323
2019-03-20 12:53 | 175 | 0.062509 | 0.019498 | 1.9745 | 21.1098 | 33.1946 | 1.3196
2019-03-20 12:53 | 176 | 0.071558 | 0.017042 | 1.7248 | 18.1858 | 33.1346 | 1.3013
2019-03-20 12:53 | 177 | 0.084579 | 0.014356 | 1.4492 | 15.3317 | 33.0504 | 1.2967
2019-03-20 12:53 | 178 | 0.076568 | 0.015852 | 1.5971 | 16.7846 | 32.9702 | 1.2852
2019-03-20 12:53 | 179 | 0.085372 | 0.013975 | 1.4291 | 15.0404 | 32.894 | 1.284
2019-03-20 12:53 | 180 | 0.078675 | 0.015145 | 1.5411 | 16.2567 | 32.8028 | 1.279
2019-03-20 12:53 | 181 | 0.076126 | 0.01537 | 1.5872 | 16.4148 | 32.7359 | 1.2496
2019-03-20 12:53 | 182 | 0.082016 | 0.014019 | 1.4713 | 15.6345 | 32.6491 | 1.2823
2019-03-20 12:53 | 183 | 0.062039 | 0.018112 | 1.9297 | 17.7641 | 32.5778 | 1.1021
2019-03-20 12:53 | 184 | 0.05307 | 0.020742 | 2.2497 | 23.3984 | 32.5221 | 1.2418
2019-03-20 12:53 | 185 | 0.07235 | 0.014907 | 1.6422 | 16.6734 | 32.4647 | 1.2063
2019-03-20 12:53 | 186 | 0.076489 | 0.014122 | 1.5488 | 15.5634 | 32.3809 | 1.1904
2019-03-20 12:53 | 187 | 0.05976 | 0.018126 | 1.9742 | 19.6294 | 32.3165 | 1.1731
2019-03-20 12:53 | 188 | 0.074169 | 0.01427 | 1.5764 | 16.0511 | 32.2547 | 1.1905
2019-03-20 12:53 | 189 | 0.070816 | 0.015299 | 1.6423 | 16.7875 | 32.1717 | 1.1888
2019-03-20 12:53 | 190 | 0.068863 | 0.015708 | 1.6703 | 16.9946 | 32.1129 | 1.1703
2019-03-20 12:53 | 191 | 0.074262 | 0.014544 | 1.5408 | 15.9369 | 32.0325 | 1.1835
2019-03-20 12:53 | 192 | 0.06785 | 0.01559 | 1.6732 | 17.2308 | 31.9672 | 1.1691
2019-03-20 12:53 | 193 | 0.063068 | 0.016386 | 1.7851 | 18.0414 | 31.8968 | 1.1378
2019-03-20 12:53 | 194 | 0.0578 | 0.017517 | 1.9397 | 19.062 | 31.8411 | 1.1018
2019-03-20 12:53 | 195 | 0.055222 | 0.018386 | 2.0224 | 20.0914 | 31.7804 | 1.1095
2019-03-20 12:53 | 196 | 0.056165 | 0.018036 | 1.9694 | 19.6067 | 31.7297 | 1.1012
2019-03-20 12:53 | 197 | 0.055099 | 0.018807 | 1.991 | 19.929 | 31.6687 | 1.0981
2019-03-20 12:53 | 198 | 0.044476 | 0.023855 | 2.4494 | 24.714 | 31.6215 | 1.0992
2019-03-20 12:53 | 199 | 0.050423 | 0.021434 | 2.1374 | 21.995 | 31.5761 | 1.109
2019-03-20 12:53 | 200 | 0.045791 | 0.023592 | 2.3418 | 24.099 | 31.523 | 1.1035
2019-03-20 12:53 | 201 | 0.038209 | 0.028855 | 2.7937 | 28.7075 | 31.4851 | 1.0969
2019-03-20 12:53 | 202 | 0.041341 | 0.02729 | 2.579 | 26.1807 | 31.4431 | 1.0823
2019-03-20 12:53 | 203 | 0.072007 | 0.015995 | 1.4775 | 14.7725 | 31.396 | 1.0637
2019-03-20 12:53 | 204 | 0.074354 | 0.016127 | 1.4263 | 14.6691 | 31.3059 | 1.0907
2019-03-20 12:53 | 205 | 0.062103 | 0.019629 | 1.6965 | 17.5018 | 31.2512 | 1.0869
2019-03-20 12:53 | 206 | 0.073196 | 0.016938 | 1.4339 | 14.8163 | 31.1768 | 1.0845
2019-03-20 12:53 | 207 | 0.067079 | 0.018528 | 1.5622 | 16.1655 | 31.1083 | 1.0844
2019-03-20 12:53 | 208 | 0.066577 | 0.018655 | 1.5691 | 16.8225 | 31.0417 | 1.12
2019-03-20 12:53 | 209 | 0.063049 | 0.019001 | 1.6558 | 17.0008 | 30.9757 | 1.0719
2019-03-20 12:53 | 210 | 0.063412 | 0.018161 | 1.6396 | 16.9206 | 30.9144 | 1.073
2019-03-20 12:53 | 211 | 0.074427 | 0.014577 | 1.3954 | 13.9047 | 30.8468 | 1.0349
2019-03-20 12:53 | 212 | 0.062849 | 0.016898 | 1.6452 | 16.302 | 30.7713 | 1.0246
2019-03-20 12:53 | 213 | 0.038671 | 0.026845 | 2.6549 | 26.1673 | 30.7237 | 1.0119
2019-03-20 12:53 | 214 | 0.056117 | 0.018891 | 1.8134 | 17.9273 | 30.684 | 1.006
2019-03-20 12:53 | 215 | 0.066912 | 0.016559 | 1.5108 | 15.1738 | 30.6128 | 1.0153
2019-03-20 12:53 | 216 | 0.045576 | 0.025252 | 2.1996 | 22.4898 | 30.5578 | 1.025
2019-03-20 12:53 | 217 | 0.046465 | 0.025276 | 2.1494 | 21.9364 | 30.5166 | 1.0193
2019-03-20 12:53 | 218 | 0.060434 | 0.019793 | 1.6449 | 17.0861 | 30.4612 | 1.0326
2019-03-20 12:53 | 219 | 0.067143 | 0.018161 | 1.4761 | 15.5387 | 30.3977 | 1.0433
2019-03-20 12:53 | 220 | 0.054981 | 0.022969 | 1.7931 | 18.5248 | 30.332 | 1.0185
2019-03-20 12:53 | 221 | 0.036994 | 0.035371 | 2.6526 | 27.7058 | 30.2887 | 1.025
2019-03-20 12:53 | 222 | 0.059591 | 0.023055 | 1.6404 | 17.0467 | 30.2485 | 1.0158
2019-03-20 12:53 | 223 | 0.06482 | 0.021948 | 1.5088 | 15.716 | 30.1736 | 1.0187
2019-03-20 12:53 | 224 | 0.038366 | 0.038295 | 2.5395 | 26.3749 | 30.1263 | 1.0119
2019-03-20 12:53 | 225 | 0.046033 | 0.0329 | 2.1181 | 22.5557 | 30.089 | 1.0383
2019-03-20 12:53 | 226 | 0.060728 | 0.025656 | 1.5946 | 16.5043 | 30.0322 | 1.0023
2019-03-20 12:53 | 227 | 0.05331 | 0.030104 | 1.815 | 18.7538 | 29.9732 | 0.99977
2019-03-20 12:53 | 228 | 0.029869 | 0.056011 | 3.2269 | 33.2809 | 29.929 | 0.99406
2019-03-20 12:54 | 229 | 0.046838 | 0.037623 | 2.0465 | 21.9001 | 29.9033 | 1.0258
2019-03-20 12:54 | 230 | 0.082032 | 0.022334 | 1.1668 | 12.2719 | 29.8311 | 1.0067
2019-03-20 12:54 | 231 | 0.065337 | 0.029814 | 1.4627 | 15.5667 | 29.7519 | 1.0171
2019-03-20 12:54 | 232 | 0.043826 | 0.046078 | 2.1771 | 22.9441 | 29.7022 | 1.0055
2019-03-20 12:54 | 233 | 0.055107 | 0.037934 | 1.7352 | 18.6049 | 29.6561 | 1.0253
2019-03-20 12:54 | 234 | 0.070069 | 0.030775 | 1.3622 | 15.0112 | 29.5909 | 1.0518
2019-03-20 12:54 | 235 | 0.055318 | 0.040692 | 1.7246 | 18.4458 | 29.5231 | 1.0204
2019-03-20 12:54 | 236 | 0.047683 | 0.048139 | 2.0005 | 21.9269 | 29.4784 | 1.0455
2019-03-20 12:54 | 237 | 0.07223 | 0.032773 | 1.3251 | 14.2184 | 29.4197 | 1.027
2019-03-20 12:54 | 238 | 0.088933 | 0.027143 | 1.078 | 11.8189 | 29.3363 | 1.0511
2019-03-20 12:54 | 239 | 0.059731 | 0.041024 | 1.5991 | 17.3299 | 29.2533 | 1.0351
2019-03-20 12:54 | 240 | 0.042385 | 0.059422 | 2.2545 | 24.4477 | 29.2142 | 1.0362
2019-03-20 12:54 | 241 | 0.063675 | 0.04016 | 1.4951 | 15.9324 | 29.1591 | 1.0145
2019-03-20 12:54 | 242 | 0.061195 | 0.042923 | 1.5554 | 16.7227 | 29.0916 | 1.0233
2019-03-20 12:54 | 243 | 0.058835 | 0.045889 | 1.6204 | 17.6286 | 29.037 | 1.0372
2019-03-20 12:54 | 244 | 0.069496 | 0.040152 | 1.3661 | 14.9434 | 28.9711 | 1.0385
2019-03-20 12:54 | 245 | 0.067922 | 0.042377 | 1.3942 | 15.3003 | 28.9013 | 1.0392
2019-03-20 12:54 | 246 | 0.046766 | 0.062407 | 2.0214 | 22.2646 | 28.8401 | 1.0412
2019-03-20 12:54 | 247 | 0.03998 | 0.073415 | 2.3561 | 26.1035 | 28.8036 | 1.0436
2019-03-20 12:54 | 248 | 0.069123 | 0.043084 | 1.3623 | 14.9701 | 28.7514 | 1.0348
2019-03-20 12:54 | 249 | 0.070513 | 0.042943 | 1.3335 | 14.7004 | 28.6714 | 1.0366
2019-03-20 12:54 | 250 | 0.068941 | 0.044842 | 1.363 | 15.1195 | 28.6114 | 1.0423
2019-03-20 12:54 | 251 | 0.081189 | 0.037955 | 1.1515 | 12.871 | 28.5298 | 1.045
2019-03-20 12:54 | 252 | 0.10304 | 0.030089 | 0.90556 | 11.0033 | 28.4463 | 1.1338
2019-03-20 12:54 | 253 | 0.12245 | 0.024801 | 0.76026 | 10.2204 | 28.3248 | 1.2515
2019-03-20 12:54 | 254 | 0.11577 | 0.025768 | 0.81258 | 12.494 | 28.2076 | 1.4464
2019-03-20 12:54 | 255 | 0.12177 | 0.024054 | 0.79092 | 14.3712 | 28.09 | 1.7499
2019-03-20 12:54 | 256 | 0.15904 | 0.0183 | 0.63569 | 12.7466 | 27.9571 | 2.0272
2019-03-20 12:54 | 257 | 0.15912 | 0.018525 | 0.72841 | 14.1817 | 27.7817 | 2.2565
2019-03-20 12:54 | 258 | 0.11412 | 0.026538 | 1.1641 | 21.3957 | 27.6498 | 2.4418
2019-03-20 12:54 | 259 | 0.12273 | 0.025655 | 1.2134 | 21.3926 | 27.5412 | 2.6255
2019-03-20 12:54 | 260 | 0.13364 | 0.024423 | 1.2112 | 20.3708 | 27.4028 | 2.7223
2019-03-20 12:54 | 261 | 0.12756 | 0.026756 | 1.3832 | 22.1113 | 27.2783 | 2.8204
2019-03-20 12:54 | 262 | 0.11612 | 0.030105 | 1.6261 | 25.6307 | 27.1497 | 2.9762
2019-03-20 12:54 | 263 | 0.0999 | 0.036113 | 2.0061 | 30.8258 | 27.0468 | 3.0795
2019-03-20 12:54 | 264 | 0.10785 | 0.034028 | 1.9646 | 28.81 | 26.9444 | 3.107
2019-03-20 12:54 | 265 | 0.10003 | 0.036762 | 2.1959 | 31.1266 | 26.8347 | 3.1135
2019-03-20 12:54 | 266 | 0.079703 | 0.046064 | 2.8393 | 38.9087 | 26.7468 | 3.1012
2019-03-20 12:54 | 267 | 0.092633 | 0.039459 | 2.5139 | 33.2727 | 26.6668 | 3.0822
2019-03-20 12:54 | 268 | 0.11902 | 0.030413 | 2.0004 | 25.4 | 26.5579 | 3.0232
2019-03-20 12:54 | 269 | 0.12062 | 0.029669 | 2.0048 | 24.3941 | 26.4347 | 2.9424
2019-03-20 12:54 | 270 | 0.12133 | 0.029549 | 2.0145 | 23.8618 | 26.318 | 2.895
2019-03-20 12:54 | 271 | 0.10893 | 0.03267 | 2.2501 | 25.5704 | 26.1955 | 2.7853
2019-03-20 12:54 | 272 | 0.093747 | 0.038013 | 2.617 | 29.3178 | 26.1007 | 2.7485
2019-03-20 12:54 | 273 | 0.096735 | 0.03699 | 2.5221 | 27.8652 | 26.0034 | 2.6955
2019-03-20 12:54 | 274 | 0.10274 | 0.035497 | 2.3605 | 25.9232 | 25.9058 | 2.6634
2019-03-20 12:54 | 275 | 0.1127 | 0.032547 | 2.1344 | 23.2979 | 25.7968 | 2.6258
2019-03-20 12:54 | 276 | 0.12358 | 0.030186 | 1.9238 | 20.7266 | 25.6807 | 2.5614
2019-03-20 12:54 | 277 | 0.1134 | 0.033357 | 2.0813 | 22.5469 | 25.5549 | 2.5567
2019-03-20 12:54 | 278 | 0.10309 | 0.037286 | 2.2648 | 25.012 | 25.4538 | 2.5785
2019-03-20 12:54 | 279 | 0.12259 | 0.031641 | 1.8824 | 20.8921 | 25.3423 | 2.5611
2019-03-20 12:54 | 280 | 0.10155 | 0.038539 | 2.2475 | 25.3393 | 25.2171 | 2.5731
2019-03-20 12:54 | 281 | 0.11052 | 0.035788 | 2.0521 | 23.4696 | 25.1324 | 2.5939
2019-03-20 12:54 | 282 | 0.13142 | 0.030396 | 1.7151 | 19.9626 | 24.993 | 2.6235
2019-03-20 12:54 | 283 | 0.12625 | 0.032164 | 1.7801 | 20.9691 | 24.8759 | 2.6473
2019-03-20 12:54 | 284 | 0.11839 | 0.034627 | 1.8927 | 22.5558 | 24.742 | 2.6704
2019-03-20 12:54 | 285 | 0.10619 | 0.038793 | 2.1017 | 25.797 | 24.6396 | 2.7394
2019-03-20 12:54 | 286 | 0.12082 | 0.034166 | 1.8478 | 22.9829 | 24.5233 | 2.7767
2019-03-20 12:54 | 287 | 0.1218 | 0.034093 | 1.8362 | 23.1148 | 24.4007 | 2.8154
2019-03-20 12:54 | 288 | 0.11974 | 0.034864 | 1.8821 | 23.5996 | 24.2811 | 2.8258
2019-03-20 12:54 | 289 | 0.10744 | 0.039249 | 2.1161 | 26.5635 | 24.1631 | 2.8541
2019-03-20 12:55 | 290 | 0.12458 | 0.03434 | 1.8423 | 23.3141 | 24.0594 | 2.9045
2019-03-20 12:55 | 291 | 0.12796 | 0.033869 | 1.8098 | 23.0544 | 23.9169 | 2.9501
2019-03-20 12:55 | 292 | 0.121 | 0.036196 | 1.9386 | 24.7029 | 23.8058 | 2.9891
2019-03-20 12:55 | 293 | 0.13564 | 0.032476 | 1.7535 | 22.5509 | 23.6699 | 3.0588
2019-03-20 12:55 | 294 | 0.1455 | 0.029898 | 1.6506 | 20.9362 | 23.5355 | 3.0463
2019-03-20 12:55 | 295 | 0.14602 | 0.029508 | 1.6652 | 21.304 | 23.3822 | 3.1109
2019-03-20 12:55 | 296 | 0.1161 | 0.036621 | 2.1218 | 27.0129 | 23.2501 | 3.1362
2019-03-20 12:55 | 297 | 0.13417 | 0.031314 | 1.8526 | 23.5615 | 23.1387 | 3.1613
2019-03-20 12:55 | 298 | 0.15492 | 0.026787 | 1.6181 | 20.3502 | 22.9809 | 3.1527
2019-03-20 12:55 | 299 | 0.13567 | 0.030292 | 1.8685 | 23.2621 | 22.8388 | 3.156
2019-03-20 12:55 | 300 | 0.1168 | 0.034999 | 2.1897 | 27.4101 | 22.7099 | 3.2014
2019-03-20 12:55 | 301 | 0.12189 | 0.033403 | 2.1183 | 26.5122 | 22.5986 | 3.2317
2019-03-20 12:55 | 302 | 0.14959 | 0.027045 | 1.7378 | 21.7842 | 22.4608 | 3.2587
2019-03-20 12:55 | 303 | 0.1375 | 0.029114 | 1.904 | 23.7679 | 22.309 | 3.2682
2019-03-20 12:55 | 304 | 0.115 | 0.034414 | 2.2949 | 28.5617 | 22.1887 | 3.2846
2019-03-20 12:55 | 305 | 0.12375 | 0.03174 | 2.1439 | 26.4385 | 22.0714 | 3.2718
2019-03-20 12:55 | 306 | 0.13932 | 0.028134 | 1.9243 | 23.6783 | 21.9398 | 3.2989
2019-03-20 12:55 | 307 | 0.11347 | 0.034554 | 2.3824 | 29.0519 | 21.8028 | 3.2967
2019-03-20 12:55 | 308 | 0.091191 | 0.043386 | 2.9797 | 34.827 | 21.7118 | 3.1759
2019-03-20 12:55 | 309 | 0.11176 | 0.035999 | 2.4436 | 27.4713 | 21.6093 | 3.0702
2019-03-20 12:55 | 310 | 0.13918 | 0.028841 | 1.9613 | 22.751 | 21.486 | 3.1665
2019-03-20 12:55 | 311 | 0.1447 | 0.027605 | 1.8881 | 21.9013 | 21.3373 | 3.1691
2019-03-20 12:55 | 312 | 0.11121 | 0.035308 | 2.4503 | 28.5428 | 21.2062 | 3.1744
2019-03-20 12:55 | 313 | 0.1082 | 0.03599 | 2.5166 | 29.508 | 21.1079 | 3.1927
2019-03-20 12:55 | 314 | 0.12363 | 0.030826 | 2.1887 | 26.1525 | 20.9849 | 3.2333
2019-03-20 12:55 | 315 | 0.11644 | 0.032721 | 2.3276 | 28.0545 | 20.8653 | 3.2668
2019-03-20 12:55 | 316 | 0.12347 | 0.030591 | 2.1949 | 26.4001 | 20.7489 | 3.2596
2019-03-20 12:55 | 317 | 0.13578 | 0.027627 | 2.0005 | 24.0744 | 20.6168 | 3.2688
2019-03-20 12:55 | 318 | 0.14002 | 0.026333 | 1.9468 | 23.3894 | 20.4793 | 3.2751
2019-03-20 12:55 | 319 | 0.1426 | 0.02541 | 1.9175 | 23.3735 | 20.3379 | 3.3331
2019-03-20 12:55 | 320 | 0.13139 | 0.026864 | 2.0866 | 25.9929 | 20.1977 | 3.4153
2019-03-20 12:55 | 321 | 0.12056 | 0.02887 | 2.2828 | 28.749 | 20.0749 | 3.4661
2019-03-20 12:55 | 322 | 0.12691 | 0.027236 | 2.1838 | 27.3524 | 19.9519 | 3.4713
2019-03-20 12:55 | 323 | 0.14103 | 0.024366 | 1.9857 | 24.8898 | 19.8187 | 3.5102
2019-03-20 12:55 | 324 | 0.15167 | 0.022472 | 1.8623 | 23.1321 | 19.6712 | 3.5085
2019-03-20 12:55 | 325 | 0.13801 | 0.024445 | 2.0691 | 25.4442 | 19.5216 | 3.5115
2019-03-20 12:55 | 326 | 0.12499 | 0.026983 | 2.3021 | 27.7788 | 19.3951 | 3.472
2019-03-20 12:55 | 327 | 0.13406 | 0.025127 | 2.1562 | 25.7951 | 19.266 | 3.4581
2019-03-20 12:55 | 328 | 0.14443 | 0.023557 | 2.0119 | 23.908 | 19.1261 | 3.453
2019-03-20 12:55 | 329 | 0.14969 | 0.022731 | 1.9504 | 23.2895 | 18.9786 | 3.4862
2019-03-20 12:55 | 330 | 0.14862 | 0.022847 | 1.9663 | 23.3877 | 18.8286 | 3.476
2019-03-20 12:55 | 331 | 0.14328 | 0.023595 | 2.0414 | 24.2838 | 18.6823 | 3.4793
2019-03-20 12:55 | 332 | 0.15005 | 0.022218 | 1.9452 | 23.1481 | 18.5398 | 3.4734
2019-03-20 12:55 | 333 | 0.13558 | 0.024234 | 2.1494 | 25.8188 | 18.3871 | 3.5006
2019-03-20 12:55 | 334 | 0.12438 | 0.026051 | 2.3437 | 28.3116 | 18.2673 | 3.5213
2019-03-20 12:55 | 335 | 0.14977 | 0.021333 | 1.948 | 23.5593 | 18.1292 | 3.5284
2019-03-20 12:55 | 336 | 0.16692 | 0.019117 | 1.7529 | 21.7837 | 17.9692 | 3.6361
2019-03-20 12:55 | 337 | 0.17052 | 0.018966 | 1.7272 | 21.0496 | 17.7994 | 3.5894
2019-03-20 12:55 | 338 | 0.1615 | 0.020274 | 1.8352 | 22.7102 | 17.6321 | 3.6676
2019-03-20 12:55 | 339 | 0.14341 | 0.023385 | 2.0799 | 25.6636 | 17.479 | 3.6804
2019-03-20 12:55 | 340 | 0.12939 | 0.026404 | 2.3182 | 28.4853 | 17.3436 | 3.6858
2019-03-20 12:55 | 341 | 0.14551 | 0.02364 | 2.0727 | 25.3835 | 17.2125 | 3.6935
2019-03-20 12:55 | 342 | 0.16018 | 0.021456 | 1.8936 | 23.0243 | 17.0527 | 3.6879
2019-03-20 12:55 | 343 | 0.15416 | 0.022205 | 1.9734 | 23.7722 | 16.8974 | 3.6647
2019-03-20 12:55 | 344 | 0.14503 | 0.023648 | 2.1126 | 25.0267 | 16.7455 | 3.6296
2019-03-20 12:55 | 345 | 0.14757 | 0.023266 | 2.0857 | 24.4356 | 16.6048 | 3.606
2019-03-20 12:55 | 346 | 0.14157 | 0.024004 | 2.1805 | 25.565 | 16.4523 | 3.6192
2019-03-20 12:55 | 347 | 0.13047 | 0.025885 | 2.363 | 27.7356 | 16.323 | 3.6186
2019-03-20 12:55 | 348 | 0.12656 | 0.026687 | 2.4326 | 28.4307 | 16.1896 | 3.5981
2019-03-20 12:56 | 349 | 0.12184 | 0.027842 | 2.5224 | 29.74 | 16.069 | 3.6236
2019-03-20 12:56 | 350 | 0.14888 | 0.022879 | 2.0624 | 24.1109 | 15.9388 | 3.5897
2019-03-20 12:56 | 351 | 0.14474 | 0.023697 | 2.1252 | 24.6902 | 15.7785 | 3.5737
2019-03-20 12:56 | 352 | 0.12746 | 0.027421 | 2.4146 | 27.961 | 15.6525 | 3.5638
2019-03-20 12:56 | 353 | 0.14423 | 0.024378 | 2.13 | 24.7279 | 15.5156 | 3.5666
2019-03-20 12:56 | 354 | 0.15411 | 0.022732 | 1.9803 | 23.1898 | 15.3653 | 3.5737
2019-03-20 12:56 | 355 | 0.15074 | 0.023248 | 2.0168 | 23.6081 | 15.2113 | 3.5586
2019-03-20 12:56 | 356 | 0.13606 | 0.025411 | 2.2219 | 26.189 | 15.0668 | 3.5633
2019-03-20 12:56 | 357 | 0.12972 | 0.026419 | 2.3194 | 27.4107 | 14.9369 | 3.5557
2019-03-20 12:56 | 358 | 0.14013 | 0.024244 | 2.1401 | 25.3231 | 14.8028 | 3.5484
2019-03-20 12:56 | 359 | 0.15343 | 0.022158 | 1.9583 | 23.1198 | 14.656 | 3.5473
2019-03-20 12:56 | 360 | 0.14579 | 0.023124 | 2.0582 | 24.4101 | 14.5013 | 3.5587
2019-03-20 12:56 | 361 | 0.13442 | 0.025072 | 2.2326 | 26.4615 | 14.3659 | 3.557
2019-03-20 12:56 | 362 | 0.12509 | 0.026862 | 2.4068 | 28.9936 | 14.2315 | 3.6269
2019-03-20 12:56 | 363 | 0.13558 | 0.02454 | 2.2151 | 27.0048 | 14.1107 | 3.6614
2019-03-20 12:56 | 364 | 0.14462 | 0.022983 | 2.0795 | 25.2718 | 13.9606 | 3.6548
2019-03-20 12:56 | 365 | 0.13639 | 0.024337 | 2.2093 | 26.8746 | 13.8255 | 3.6653
2019-03-20 12:56 | 366 | 0.1396 | 0.023599 | 2.1632 | 26.2406 | 13.6853 | 3.6631
2019-03-20 12:56 | 367 | 0.14804 | 0.022051 | 2.0529 | 24.8084 | 13.5449 | 3.6725
2019-03-20 12:56 | 368 | 0.14833 | 0.021896 | 2.0558 | 24.8621 | 13.3914 | 3.6878
2019-03-20 12:56 | 369 | 0.14001 | 0.022929 | 2.1892 | 27.1175 | 13.2505 | 3.7966
2019-03-20 12:56 | 370 | 0.13697 | 0.023277 | 2.2524 | 27.7525 | 13.1097 | 3.8012
2019-03-20 12:56 | 371 | 0.15284 | 0.020716 | 2.032 | 24.8144 | 12.9717 | 3.7926
2019-03-20 12:56 | 372 | 0.16685 | 0.018924 | 1.8687 | 22.6508 | 12.8048 | 3.7793
2019-03-20 12:56 | 373 | 0.14873 | 0.021453 | 2.1143 | 25.5344 | 12.6456 | 3.7978
2019-03-20 12:56 | 374 | 0.1514 | 0.02124 | 2.0851 | 24.9624 | 12.5029 | 3.7792
2019-03-20 12:56 | 375 | 0.1603 | 0.020184 | 1.9738 | 23.5546 | 12.341 | 3.7757
2019-03-20 12:56 | 376 | 0.15983 | 0.02021 | 1.9799 | 23.4091 | 12.1849 | 3.7414
2019-03-20 12:56 | 377 | 0.14264 | 0.022647 | 2.2218 | 26.071 | 12.0252 | 3.7187
2019-03-20 12:56 | 378 | 0.14503 | 0.022482 | 2.1903 | 25.674 | 11.8949 | 3.7234
2019-03-20 12:56 | 379 | 0.15951 | 0.020265 | 1.9852 | 23.1472 | 11.7323 | 3.6922
2019-03-20 12:56 | 380 | 0.14756 | 0.021913 | 2.1428 | 24.71 | 11.5819 | 3.6463
2019-03-20 12:56 | 381 | 0.14319 | 0.022709 | 2.2001 | 25.0305 | 11.4357 | 3.5841
2019-03-20 12:56 | 382 | 0.15367 | 0.021375 | 2.0368 | 23.1061 | 11.2923 | 3.5508
2019-03-20 12:56 | 383 | 0.14397 | 0.023071 | 2.1619 | 24.632 | 11.1335 | 3.5462
2019-03-20 12:56 | 384 | 0.11468 | 0.029519 | 2.6991 | 31.0565 | 11.0089 | 3.5615
2019-03-20 12:56 | 385 | 0.11437 | 0.029821 | 2.6908 | 31.2635 | 10.8966 | 3.5757
2019-03-20 12:56 | 386 | 0.14049 | 0.0242 | 2.1717 | 25.5649 | 10.773 | 3.5916
2019-03-20 12:56 | 387 | 0.1567 | 0.021474 | 1.9347 | 23.6344 | 10.6184 | 3.7036
2019-03-20 12:56 | 388 | 0.13432 | 0.025071 | 2.2616 | 27.5856 | 10.4693 | 3.7054
2019-03-20 12:56 | 389 | 0.12243 | 0.027142 | 2.4879 | 30.456 | 10.3474 | 3.7286
2019-03-20 12:56 | 390 | 0.14335 | 0.023144 | 2.1445 | 26.3283 | 10.2168 | 3.7741
2019-03-20 12:56 | 391 | 0.13113 | 0.025017 | 2.3539 | 29.0996 | 10.0679 | 3.8157
2019-03-20 12:56 | 392 | 0.12206 | 0.026454 | 2.5303 | 31.3267 | 9.9536 | 3.8237
2019-03-20 12:56 | 393 | 0.14724 | 0.02162 | 2.1034 | 26.0895 | 9.8154 | 3.8415
2019-03-20 12:56 | 394 | 0.16006 | 0.019672 | 1.9477 | 24.3647 | 9.6616 | 3.8999
2019-03-20 12:56 | 395 | 0.16459 | 0.018973 | 1.9002 | 23.8872 | 9.4986 | 3.9316
2019-03-20 12:56 | 396 | 0.14013 | 0.022375 | 2.2555 | 28.1993 | 9.3396 | 3.9516
2019-03-20 12:56 | 397 | 0.13299 | 0.023585 | 2.3937 | 29.5661 | 9.2143 | 3.9321
2019-03-20 12:56 | 398 | 0.14437 | 0.02176 | 2.2237 | 27.1984 | 9.0687 | 3.9265
2019-03-20 12:56 | 399 | 0.14554 | 0.021928 | 2.2226 | 26.7343 | 8.9277 | 3.8909
2019-03-20 12:56 | 400 | 0.14004 | 0.023067 | 2.3155 | 27.2915 | 8.7794 | 3.8219
2019-03-20 12:56 | 401 | 0.13957 | 0.023697 | 2.3215 | 26.6091 | 8.6463 | 3.7137
2019-03-20 12:56 | 402 | 0.15027 | 0.022606 | 2.1554 | 24.6337 | 8.4979 | 3.7017
2019-03-20 12:56 | 403 | 0.14143 | 0.024476 | 2.2841 | 25.7663 | 8.3504 | 3.6441
2019-03-20 12:56 | 404 | 0.13288 | 0.02636 | 2.4166 | 26.93 | 8.2158 | 3.5785
2019-03-20 12:56 | 405 | 0.11647 | 0.030376 | 2.7327 | 30.0697 | 8.0862 | 3.5021
2019-03-20 12:56 | 406 | 0.11644 | 0.030546 | 2.7065 | 29.8753 | 7.9785 | 3.4786
2019-03-20 12:56 | 407 | 0.13067 | 0.027234 | 2.3916 | 26.762 | 7.8496 | 3.4971
2019-03-20 12:56 | 408 | 0.13101 | 0.027249 | 2.3558 | 25.6011 | 7.7201 | 3.354
2019-03-20 12:56 | 409 | 0.13597 | 0.026286 | 2.2528 | 25.2434 | 7.5871 | 3.4325
2019-03-20 12:57 | 410 | 0.12799 | 0.027623 | 2.3756 | 27.1076 | 7.4511 | 3.4696
2019-03-20 12:57 | 411 | 0.13039 | 0.026859 | 2.3177 | 26.7038 | 7.3293 | 3.4819
2019-03-20 12:57 | 412 | 0.11909 | 0.029114 | 2.5238 | 29.3285 | 7.1936 | 3.4928
2019-03-20 12:57 | 413 | 0.10296 | 0.033844 | 2.9004 | 33.7772 | 7.092 | 3.4776
2019-03-20 12:57 | 414 | 0.10832 | 0.032861 | 2.7566 | 32.4331 | 6.9819 | 3.5131
2019-03-20 12:57 | 415 | 0.13045 | 0.027738 | 2.2837 | 26.103 | 6.871 | 3.4051
2019-03-20 12:57 | 416 | 0.13414 | 0.027323 | 2.2189 | 24.538 | 6.7256 | 3.2916
2019-03-20 12:57 | 417 | 0.12119 | 0.030372 | 2.4413 | 26.4616 | 6.6071 | 3.207
2019-03-20 12:57 | 418 | 0.12353 | 0.029907 | 2.3747 | 25.6704 | 6.4801 | 3.1711
2019-03-20 12:57 | 419 | 0.11752 | 0.031449 | 2.471 | 26.6706 | 6.3617 | 3.1343
2019-03-20 12:57 | 420 | 0.113 | 0.032555 | 2.5416 | 27.7875 | 6.2449 | 3.1399
2019-03-20 12:57 | 421 | 0.10886 | 0.033419 | 2.6065 | 29.0251 | 6.1357 | 3.1595
2019-03-20 12:57 | 422 | 0.098857 | 0.036585 | 2.836 | 31.9315 | 6.0281 | 3.1566
2019-03-20 12:57 | 423 | 0.10652 | 0.033964 | 2.6052 | 29.2995 | 5.9334 | 3.121
2019-03-20 12:57 | 424 | 0.12445 | 0.029267 | 2.2148 | 24.8724 | 5.8129 | 3.0953
2019-03-20 12:57 | 425 | 0.10738 | 0.034153 | 2.549 | 28.5332 | 5.6922 | 3.0638
2019-03-20 12:57 | 426 | 0.12427 | 0.029714 | 2.1889 | 24.0241 | 5.5911 | 2.9854
2019-03-20 12:57 | 427 | 0.12348 | 0.029991 | 2.189 | 23.6953 | 5.4478 | 2.926
2019-03-20 12:57 | 428 | 0.10155 | 0.036449 | 2.6325 | 28.3783 | 5.3496 | 2.8817
2019-03-20 12:57 | 429 | 0.089544 | 0 | 0 | 0 | 5.2421 | 2.8444
2019-03-20 12:57 | 430 | 0.104 | 0 | 0 | 0 | 5.1633 | 2.8555
2019-03-20 12:57 | 431 | 0.12739 | 0 | 0 | 0 | 5.0318 | 2.8462
2019-03-20 12:57 | 432 | 0.12142 | 0 | 0 | 0 | 4.9155 | 2.8734
2019-03-20 12:57 | 433 | 0.12304 | 0 | 0 | 0 | 4.7885 | 2.8853
2019-03-20 12:57 | 434 | 0.10251 | 0 | 0 | 0 | 4.6743 | 2.9083
2019-03-20 12:57 | 435 | 0.10657 | 0 | 0 | 0 | 4.5784 | 2.9098
2019-03-20 12:57 | 436 | 0.093015 | 0 | 0 | 0 | 4.4645 | 2.8106
2019-03-20 12:57 | 437 | 0.094975 | 0 | 0 | 0 | 4.3887 | 2.7016
2019-03-20 12:57 | 438 | 0.10245 | 0 | 0 | 0 | 4.2729 | 2.6506
2019-03-20 12:57 | 439 | 0.10301 | 0 | 0 | 0 | 4.1856 | 2.6275
2019-03-20 12:57 | 440 | 0.094873 | 0 | 0 | 0 | 4.0694 | 2.6221
2019-03-20 12:57 | 441 | 0.081411 | 0 | 0 | 0 | 3.9969 | 2.6474
2019-03-20 12:57 | 442 | 0.089661 | 0 | 0 | 0 | 3.9018 | 2.6462
2019-03-20 12:57 | 443 | 0.078035 | 0 | 0 | 0 | 3.8213 | 2.5637
2019-03-20 12:57 | 444 | 0.089164 | 0 | 0 | 0 | 3.7405 | 2.4668
2019-03-20 12:57 | 445 | 0.10133 | 0 | 0 | 0 | 3.643 | 2.4206
2019-03-20 12:57 | 446 | 0.085239 | 0 | 0 | 0 | 3.5452 | 2.3758
2019-03-20 12:57 | 447 | 0.056217 | 0 | 0 | 0 | 3.4754 | 2.2641
2019-03-20 12:57 | 448 | 0.07113 | 0 | 0 | 0 | 3.4221 | 2.2025
2019-03-20 12:57 | 449 | 0.086471 | 0 | 0 | 0 | 3.3326 | 2.1763
2019-03-20 12:57 | 450 | 0.070295 | 0 | 0 | 0 | 3.2566 | 2.1789
2019-03-20 12:57 | 451 | 0.069979 | 0 | 0 | 0 | 3.1888 | 2.1598
2019-03-20 12:57 | 452 | 0.073846 | 0 | 0 | 0 | 3.1153 | 2.0881
2019-03-20 12:57 | 453 | 0.073586 | 0 | 0 | 0 | 3.0422 | 2.0652
2019-03-20 12:57 | 454 | 0.066527 | 0 | 0 | 0 | 2.9703 | 2.0669
2019-03-20 12:57 | 455 | 0.046372 | 0 | 0 | 0 | 2.9118 | 2.0574
2019-03-20 12:57 | 456 | 0.055157 | 0 | 0 | 0 | 2.8703 | 2.0415
2019-03-20 12:57 | 457 | 0.073038 | 0 | 0 | 0 | 2.7989 | 2.0146
2019-03-20 12:57 | 458 | 0.070751 | 0 | 0 | 0 | 2.7287 | 2.0108
2019-03-20 12:57 | 459 | 0.078809 | 0 | 0 | 0 | 2.6554 | 1.9729
2019-03-20 12:57 | 460 | 0.086837 | 0 | 0 | 0 | 2.5716 | 1.9466
2019-03-20 12:57 | 461 | 0.070746 | 0 | 0 | 0 | 2.4877 | 1.9048
2019-03-20 12:57 | 462 | 0.06049 | 0 | 0 | 0 | 2.4289 | 1.8864
2019-03-20 12:57 | 463 | 0.06677 | 0 | 0 | 0 | 2.3628 | 1.8497
2019-03-20 12:57 | 464 | 0.058824 | 0 | 0 | 0 | 2.2982 | 1.8222
2019-03-20 12:57 | 465 | 0.058457 | 0 | 0 | 0 | 2.2432 | 1.7938
2019-03-20 12:57 | 466 | 0.070914 | 0 | 0 | 0 | 2.178 | 1.7922
2019-03-20 12:57 | 467 | 0.080612 | 0 | 0 | 0 | 2.1025 | 1.7894
2019-03-20 12:57 | 468 | 0.062217 | 0 | 0 | 0 | 2.024 | 1.7765
2019-03-20 12:58 | 469 | 0.040289 | 0 | 0 | 0 | 1.9785 | 1.7754
2019-03-20 12:58 | 470 | 0.065846 | 0 | 0 | 0 | 1.932 | 1.7412
2019-03-20 12:58 | 471 | 0.07388 | 0 | 0 | 0 | 1.8504 | 1.7094
2019-03-20 12:58 | 472 | 0.060279 | 0 | 0 | 0 | 1.79 | 1.654
2019-03-20 12:58 | 473 | 0.052141 | 0 | 0 | 0 | 1.7286 | 1.6152
2019-03-20 12:58 | 474 | 0.061758 | 0 | 0 | 0 | 1.6809 | 1.587
2019-03-20 12:58 | 475 | 0.078441 | 0 | 0 | 0 | 1.6036 | 1.5837
2019-03-20 12:58 | 476 | 0.066765 | 0 | 0 | 0 | 1.5312 | 1.5576
2019-03-20 12:58 | 477 | 0.049467 | 0 | 0 | 0 | 1.4719 | 1.5133
2019-03-20 12:58 | 478 | 0.051133 | 0 | 0 | 0 | 1.4281 | 1.5075
2019-03-20 12:58 | 479 | 0.039826 | 0 | 0 | 0 | 1.3723 | 1.4325
2019-03-20 12:58 | 480 | 0.028725 | 0 | 0 | 0 | 1.348 | 1.3956
2019-03-20 12:58 | 481 | 0.040029 | 0 | 0 | 0 | 1.3096 | 1.3548
2019-03-20 12:58 | 482 | 0.035368 | 0 | 0 | 0 | 1.2713 | 1.2675
2019-03-20 12:58 | 483 | 0.032461 | 0 | 0 | 0 | 1.2391 | 1.2365
2019-03-20 12:58 | 484 | 0.020511 | 0 | 0 | 0 | 1.2086 | 1.1424
2019-03-20 12:58 | 485 | 0.011964 | 0 | 0 | 0 | 1.1974 | 1.0121
2019-03-20 12:58 | 486 | 0.0045187 | 0 | 0 | 0 | 1.1838 | 0.91572
2019-03-20 12:58 | 487 | 0.012725 | 0 | 0 | 0 | 1.1841 | 0.80973
2019-03-20 12:58 | 488 | 0.027766 | 0 | 0 | 0 | 1.1564 | 0.71571
2019-03-20 12:58 | 489 | 0.035173 | 0 | 0 | 0 | 1.1312 | 0.6671
2019-03-20 12:58 | 490 | 0.013873 | 0 | 0 | 0 | 1.0933 | 0.64288
2019-03-20 12:58 | 491 | 0.0017861 | 0 | 0 | 0 | 1.1014 | 0.62329
2019-03-20 12:58 | 492 | 0.0096204 | 0 | 0 | 0 | 1.0849 | 0.57583
2019-03-20 12:58 | 493 | 0.00048733 | 0 | 0 | 0 | 1.0853 | 0.53925
2019-03-20 12:58 | 494 | 0.011263 | 0 | 0 | 0 | 1.0797 | 0.47549
2019-03-20 12:58 | 495 | 0.0097402 | 0 | 0 | 0 | 1.0653 | 0.47023
2019-03-20 12:58 | 496 | 0.01209 | 0 | 0 | 0 | 1.0598 | 0.45543
2019-03-20 12:58 | 497 | 0.0058918 | 0 | 0 | 0 | 1.0436 | 0.43543
2019-03-20 12:58 | 498 | -0.017739 | 0 | 0 | 0 | 1.0516 | 0.41783
2019-03-20 12:58 | 499 | -0.0053951 | 0 | 0 | 0 | 1.0702 | 0.40643
2019-03-20 12:58 | 500 | 0.018675 | 0 | 0 | 0 | 1.0595 | 0.37874
2019-03-20 12:58 | 501 | 0.0026433 | 0 | 0 | 0 | 1.0419 | 0.36432
2019-03-20 12:58 | 502 | 0.0071945 | 0 | 0 | 0 | 1.0494 | 0.37929
2019-03-20 12:58 | 503 | 0.040888 | 0 | 0 | 0 | 1.0211 | 0.37239
2019-03-20 12:58 | 504 | 0.026283 | 0 | 0 | 0 | 0.97946 | 0.3529
2019-03-20 12:58 | 505 | -0.0037937 | 0 | 0 | 0 | 0.97302 | 0.34415
2019-03-20 12:58 | 506 | -0.0065809 | 0 | 0 | 0 | 0.98023 | 0.34351
2019-03-20 12:58 | 507 | 0.0047953 | 0 | 0 | 0 | 0.9817 | 0.3365
2019-03-20 12:58 | 508 | 0.019995 | 0 | 0 | 0 | 0.96995 | 0.3369
2019-03-20 12:58 | 509 | 0.0073743 | 0 | 0 | 0 | 0.94837 | 0.32368
2019-03-20 12:58 | 510 | 0.00061825 | 0 | 0 | 0 | 0.95351 | 0.34311
2019-03-20 12:58 | 511 | 0.030958 | 0 | 0 | 0 | 0.93864 | 0.3015
2019-03-20 12:58 | 512 | 0.029391 | 0 | 0 | 0 | 0.89948 | 0.28683
2019-03-20 12:58 | 513 | -0.0035601 | 0 | 0 | 0 | 0.88761 | 0.28423
2019-03-20 12:58 | 514 | 0.00042467 | 0 | 0 | 0 | 0.89802 | 0.26748
2019-03-20 12:58 | 515 | 0.010857 | 0 | 0 | 0 | 0.88426 | 0.26226
2019-03-20 12:58 | 516 | 0.013881 | 0 | 0 | 0 | 0.87774 | 0.26516
2019-03-20 12:58 | 517 | 0.022876 | 0 | 0 | 0 | 0.85566 | 0.26238
2019-03-20 12:58 | 518 | 0.015263 | 0 | 0 | 0 | 0.8358 | 0.25038
2019-03-20 12:58 | 519 | 0.018036 | 0 | 0 | 0 | 0.82327 | 0.23969
2019-03-20 12:58 | 520 | 0.013761 | 0 | 0 | 0 | 0.80176 | 0.24573
2019-03-20 12:58 | 521 | -0.0092549 | 0 | 0 | 0 | 0.79977 | 0.22246
2019-03-20 12:58 | 522 | -0.0015046 | 0 | 0 | 0 | 0.81294 | 0.25485

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062444 | 0.062682 | 0 | 0.2095 | 2019-03-20 09:37 | Group A | 0.1 | 0.22 | Do Nothing | 0.20951 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025361 | 0.025478 | 0 | 0.0261 | 2019-03-20 09:37 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044884 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.00275 | 0.002762 | 0 | 0.0011 | 2019-03-20 09:37 | Group A | -0.01 | 0.1 | Do Nothing | 1.8009e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -20.1974 | 199.3757 | 0 | 100 | 2019-03-20 09:37 | None | 0 | 1000 | Do Nothing | 108.4626 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -12.5291 | 230.4066 | 0 | 100 | 2019-03-20 09:37 | None | 0 | 1000 | Do Nothing | 109.8026 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -15.3019 | 1447 | 0 |  | 2019-03-20 09:10 | None | -10 | 1000 | Do Nothing | 127.8918 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 20.9928 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001881 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001747 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -152.1334 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23895 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.403 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0034526 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0048009 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0045945 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 5.926e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0012069 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.32135 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.335 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5481 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 101 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.7284 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 17.9282 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 66.8271 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 104.6799 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 63.6029 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0043462 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0484 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025345 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.1756 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN