## Cone Summary
Date: | 2019-03-21
Time: | 12:29
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B4S-Test2
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0411 |  | Start Test | 0 | 2019-03-21 12:35
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 7 | 2019-03-21 12:35
Sample Area | 1 | m^2 | Wetting | 142 | 2019-03-21 12:38
Cone Heat Flux | 75 | kW/m^2 | Dripping | 145 | 2019-03-21 12:38
Cone Orientation |  |  | Bottom Ignition | 266 | 2019-03-21 12:40
Sample Grid |  |  | Flame Out | 536 | 2019-03-21 12:44
Sample Thickness | 106 | mm
Sample Initial Mass | 45.1 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 23.8387 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 3.6876 | kW/m2 | Average Heat Release Rate | 2.0339 | kW/m2
Total Heat Release | 1.0562 | MJ/m2 | Average Mass Loss Rate | 0.0888 | g/m2s
Peak Mass Loss Rate | 0.15829 | g/m2s | Effective Heat of Combustion | 22.9481 | MJ/kg
Total Mass Lost | 44.3157 | g | Average Specific Extinction Area | 144.7576 | m2/kg
Time to Ignition | 7 | s | Peak CO Yield | 0.061975 | kg/kg
Time to Peak Heat Release | 410 | s | Peak CO2 Yield | 4.3804 | kg/kg
Time to Flame Out | 536 | s | Average CO Yield | 0.024857 | kg/kg
Average Heat Release Rate | 1.9928 | kW/m2 | Average CO2 Yield | 1.9884 | kg/kg
Average Mass Loss Rate | 0.082952 | g/m2s
Effective Heat of Combustion | 23.8333 | MJ/kg
Dripping | 145 | s | Time to dripping
HRRD | 1.482 | kW | HRR at dripping time
THRD | 0.24387 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.7672 | kW | AHRR from Ignition to Dripping
Wetting | 142 | s | Time to wetting
HRRW | 1.5313 | kW | HRR at wetting time
THRW | 0.23938 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.7098 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-21 12:29 | 1553185747.3953 | -360 | 0.20955 | 0.000457 | -4e-06 | 100.5425 | 100.5526 | 124.4069 | 21.1981 | 0.001904 | 0.001793 | -152.1584 | 0.24023 | 8.403 | 0.003672 | 0.00472 | 0.004602 | 6e-05 | 0.001089 | 0.24585 | 1.3409 | 2.548 | 99.39 | 20.93 | 37.38 | 67.5697 | 110.2172 | 64.9142 | 0.009324 | 1.0443 | 0.024932 | 23.8744 | -0.000895 | 0.14291 | 1.3e-05 | 0.001214 | 0.001214 | 4.8375
2019-03-21 12:29 | 1553185748.3954 | -359 | 0.20955 | 0.000457 | -4e-06 | 100.5425 | 100.5526 | 124.4069 | 21.1981 | 0.001904 | 0.001793 | -152.1584 | 0.24023 | 8.403 | 0.003672 | 0.00472 | 0.004602 | 6e-05 | 0.001089 | 0.24585 | 1.3409 | 2.548 | 99.39 | 20.93 | 37.383 | 67.5697 | 110.2172 | 64.9142 | 0.009325 | 1.0443 | 0.024932 | 23.8744 | -0.000895 | 0.14291 | 1.3e-05 | 0.001214 | 0.001214 | 4.8378
2019-03-21 12:29 | 1553185749.3944 | -358 | 0.20961 | 0.000458 | -4e-06 | 100.5255 | 100.6251 | 124.4069 | 21.1982 | 0.001874 | 0.001773 | -152.1584 | 0.24019 | 8.4026 | 0.003586 | 0.004502 | 0.004294 | 5.5e-05 | 0.001092 | 0.24818 | 1.3318 | 2.5481 | 99.39 | 20.93 | 37.389 | 66.8638 | 108.1268 | 64.4326 | 0.009326 | 1.0458 | 0.02495 | 23.8574 | -0.008857 | 0.14289 | 1.2e-05 | 0.001156 | 0.001156 | 5.0793
2019-03-21 12:29 | 1553185750.3945 | -357 | 0.20959 | 0.000455 | -4e-06 | 100.516 | 100.7174 | 124.4069 | 21.1981 | 0.001912 | 0.001772 | -152.1584 | 0.23999 | 8.4039 | 0.003631 | 0.00437 | 0.004181 | 5.7e-05 | 0.001093 | 0.24657 | 1.364 | 2.5481 | 99.39 | 20.93 | 37.391 | 67.7827 | 109.2142 | 64.3921 | 0.009327 | 1.0459 | 0.024952 | 23.856 | -0.017906 | 0.14277 | 1.4e-05 | 0.00134 | 0.00134 | 4.3794
2019-03-21 12:29 | 1553185751.3949 | -356 | 0.20961 | 0.000456 | -4e-06 | 100.4862 | 100.6956 | 124.4069 | 21.1992 | 0.00189 | 0.00175 | -152.1584 | 0.24088 | 8.4012 | 0.003668 | 0.004742 | 0.00455 | 5.5e-05 | 0.001093 | 0.27256 | 1.3549 | 2.548 | 99.39 | 20.93 | 37.394 | 67.2509 | 110.1023 | 63.8706 | 0.009327 | 1.0475 | 0.024971 | 23.8376 | -0.018622 | 0.1433 | 1.3e-05 | 0.001253 | 0.001253 | 4.7016
2019-03-21 12:29 | 1553185752.3946 | -355 | 0.20955 | 0.000458 | -4e-06 | 100.437 | 100.4953 | 124.4069 | 21.198 | 0.001849 | 0.001758 | -152.1584 | 0.23936 | 8.4005 | 0.003659 | 0.004913 | 0.004747 | 5.7e-05 | 0.001097 | 0.27577 | 1.3764 | 2.548 | 99.39 | 20.93 | 37.398 | 66.2643 | 109.878 | 64.0741 | 0.009328 | 1.0469 | 0.024963 | 23.8448 | -0.005192 | 0.14239 | 1.1e-05 | 0.001073 | 0.001073 | 5.4547
2019-03-21 12:29 | 1553185753.3946 | -354 | 0.2096 | 0.000466 | -4e-06 | 100.4206 | 100.5117 | 124.4069 | 21.1994 | 0.001881 | 0.001785 | -152.1584 | 0.23976 | 8.4055 | 0.003673 | 0.005071 | 0.004905 | 5.6e-05 | 0.001096 | 0.25393 | 1.348 | 2.5478 | 99.39 | 20.93 | 37.401 | 67.0234 | 110.228 | 64.7283 | 0.009329 | 1.0449 | 0.024939 | 23.8679 | -0.008107 | 0.14264 | 1e-06 | 0.00029 | 0.00029 | 20.1951
2019-03-21 12:29 | 1553185754.3951 | -353 | 0.20958 | 0.000459 | -4e-06 | 100.5401 | 100.6592 | 124.4069 | 21.1991 | 0.001893 | 0.00174 | -152.1584 | 0.2396 | 8.4038 | 0.003627 | 0.004732 | 0.004544 | 5.4e-05 | 0.001099 | 0.26626 | 1.3409 | 2.548 | 99.39 | 20.93 | 37.403 | 67.31 | 109.101 | 63.6458 | 0.00933 | 1.0482 | 0.024979 | 23.8296 | -0.010585 | 0.14254 | 9e-06 | 0.000979 | 0.000979 | 5.9866
2019-03-21 12:29 | 1553185755.3962 | -352 | 0.20959 | 0.000459 | -4e-06 | 100.4123 | 100.4494 | 124.4069 | 21.1992 | 0.001918 | 0.001689 | -152.1584 | 0.23944 | 8.4033 | 0.003603 | 0.005207 | 0.005107 | 5.8e-05 | 0.001099 | 0.27632 | 1.3466 | 2.5479 | 99.39 | 20.93 | 37.407 | 67.9155 | 108.5304 | 62.3984 | 0.009331 | 1.0521 | 0.025026 | 23.7854 | -0.003301 | 0.14244 | 9e-06 | 0.000948 | 0.000948 | 6.1749
2019-03-21 12:29 | 1553185756.3948 | -351 | 0.20959 | 0.000457 | -4e-06 | 100.5371 | 100.5965 | 124.4069 | 21.1992 | 0.001837 | 0.001784 | -152.1584 | 0.24015 | 8.4 | 0.003537 | 0.005073 | 0.004837 | 5.4e-05 | 0.001101 | 0.26609 | 1.286 | 2.5478 | 99.39 | 20.93 | 37.41 | 65.9712 | 106.9203 | 64.6905 | 0.009331 | 1.045 | 0.024941 | 23.8665 | -0.00528 | 0.14286 | 1.2e-05 | 0.001105 | 0.001105 | 5.3155
2019-03-21 12:29 | 1553185757.3958 | -350 | 0.20957 | 0.00046 | -4e-06 | 100.5708 | 100.65 | 124.4069 | 21.1985 | 0.001875 | 0.00179 | -152.1584 | 0.23978 | 8.4027 | 0.003558 | 0.004779 | 0.004678 | 5.6e-05 | 0.001099 | 0.24945 | 1.2736 | 2.5479 | 99.39 | 20.93 | 37.415 | 66.8811 | 107.4378 | 64.8467 | 0.009333 | 1.0445 | 0.024935 | 23.8721 | -0.007038 | 0.14264 | 9e-06 | 0.000971 | 0.000971 | 6.0383
2019-03-21 12:29 | 1553185758.395 | -349 | 0.20957 | 0.000458 | -4e-06 | 100.5079 | 100.621 | 124.213 | 21.1997 | 0.001873 | 0.001725 | -152.1554 | 0.24026 | 8.4054 | 0.00362 | 0.005245 | 0.005098 | 5.7e-05 | 0.001096 | 0.2943 | 1.2952 | 2.5481 | 99.39 | 20.93 | 37.418 | 66.8395 | 108.9477 | 63.273 | 0.009333 | 1.0494 | 0.024974 | 23.7978 | -0.010058 | 0.14293 | 1.1e-05 | 0.001101 | 0.001101 | 5.3336
2019-03-21 12:29 | 1553185759.3961 | -348 | 0.20958 | 0.000457 | -4e-06 | 100.524 | 100.6922 | 124.2679 | 21.1994 | 0.001869 | 0.001769 | -152.1494 | 0.24015 | 8.4026 | 0.003551 | 0.004791 | 0.004608 | 5.5e-05 | 0.0011 | 0.26394 | 1.2805 | 2.5478 | 99.39 | 20.93 | 37.421 | 66.7447 | 107.2765 | 64.3237 | 0.009334 | 1.0461 | 0.02494 | 23.8403 | -0.014952 | 0.14286 | 1.2e-05 | 0.001179 | 0.001179 | 4.9809
2019-03-21 12:29 | 1553185760.3941 | -347 | 0.20961 | 0.000461 | -4e-06 | 100.5333 | 100.6939 | 124.1874 | 21.1998 | 0.001842 | 0.00173 | -152.1527 | 0.24025 | 8.3999 | 0.003485 | 0.005182 | 0.004982 | 5.3e-05 | 0.001099 | 0.2808 | 1.2875 | 2.548 | 99.39 | 20.93 | 37.422 | 66.0859 | 105.664 | 63.4036 | 0.009334 | 1.049 | 0.024966 | 23.8 | -0.014282 | 0.14292 | 7e-06 | 0.000758 | 0.000758 | 7.7482
2019-03-21 12:29 | 1553185761.3941 | -346 | 0.20961 | 0.000458 | -4e-06 | 100.502 | 100.5662 | 124.406 | 21.2003 | 0.001937 | 0.001717 | -152.1442 | 0.24022 | 8.4042 | 0.003524 | 0.00493 | 0.004766 | 5.6e-05 | 0.001102 | 0.28672 | 1.3107 | 2.548 | 99.39 | 20.93 | 37.422 | 68.3645 | 106.6191 | 63.0705 | 0.009334 | 1.05 | 0.025001 | 23.8092 | -0.005713 | 0.14291 | 1e-05 | 0.000968 | 0.000968 | 6.0667
2019-03-21 12:29 | 1553185762.3957 | -345 | 0.20958 | 0.000461 | -4e-06 | 100.4606 | 100.432 | 124.5468 | 21.1996 | 0.001873 | 0.001705 | -152.1515 | 0.24045 | 8.4016 | 0.003474 | 0.005044 | 0.004804 | 5.4e-05 | 0.001099 | 0.27328 | 1.3273 | 2.5481 | 99.39 | 20.93 | 37.423 | 66.8293 | 105.4048 | 62.801 | 0.009335 | 1.0509 | 0.025025 | 23.8131 | 0.002544 | 0.14304 | 7e-06 | 0.000802 | 0.000802 | 7.3319
2019-03-21 12:29 | 1553185763.3956 | -344 | 0.20959 | 0.000457 | -4e-06 | 100.485 | 100.8482 | 124.5919 | 21.2002 | 0.001842 | 0.001753 | -152.1562 | 0.24055 | 8.4027 | 0.00353 | 0.004832 | 0.004671 | 5.5e-05 | 0.001103 | 0.26017 | 1.3257 | 2.5479 | 99.39 | 20.93 | 37.424 | 66.079 | 106.7584 | 63.9481 | 0.009335 | 1.0473 | 0.024987 | 23.858 | -0.032272 | 0.1431 | 1.2e-05 | 0.001092 | 0.001092 | 5.3856
2019-03-21 12:29 | 1553185764.3942 | -343 | 0.20958 | 0.000458 | -3e-06 | 100.5201 | 100.5488 | 124.4914 | 21.2009 | 0.001887 | 0.001714 | -152.1674 | 0.23994 | 8.4042 | 0.003532 | 0.004747 | 0.004615 | 5.5e-05 | 0.001105 | 0.28838 | 1.3438 | 2.548 | 99.39 | 20.93 | 37.424 | 67.1612 | 106.8009 | 63.0195 | 0.009335 | 1.0502 | 0.025011 | 23.8155 | -0.00255 | 0.14274 | 1e-05 | 0.00095 | 0.00095 | 6.1731
2019-03-21 12:29 | 1553185765.3946 | -342 | 0.20955 | 0.000461 | -4e-06 | 100.6128 | 100.7957 | 124.6174 | 21.1991 | 0.001874 | 0.001748 | -152.1686 | 0.23963 | 8.4044 | 0.003543 | 0.004785 | 0.004667 | 5.2e-05 | 0.001103 | 0.22323 | 1.3649 | 2.5481 | 99.39 | 20.93 | 37.427 | 66.8514 | 107.0745 | 63.8323 | 0.009336 | 1.0477 | 0.024994 | 23.8564 | -0.016248 | 0.14256 | 7e-06 | 0.000742 | 0.000742 | 7.8935
2019-03-21 12:29 | 1553185766.3954 | -341 | 0.20953 | 0.000457 | -4e-06 | 100.5601 | 100.7065 | 124.5726 | 21.1997 | 0.001838 | 0.001776 | -152.163 | 0.23965 | 8.4018 | 0.003572 | 0.004574 | 0.004336 | 5.3e-05 | 0.001098 | 0.24386 | 1.296 | 2.548 | 99.39 | 20.93 | 37.428 | 65.9913 | 107.7755 | 64.4935 | 0.009336 | 1.0456 | 0.024965 | 23.8755 | -0.013015 | 0.14257 | 1.2e-05 | 0.001163 | 0.001163 | 5.0381
2019-03-21 12:29 | 1553185767.3954 | -340 | 0.20956 | 0.000457 | -4e-06 | 100.5154 | 100.5938 | 124.6327 | 21.1999 | 0.001872 | 0.001763 | -152.1465 | 0.24027 | 8.4 | 0.003524 | 0.005266 | 0.00512 | 5.5e-05 | 0.001098 | 0.25608 | 1.2579 | 2.548 | 99.39 | 20.93 | 37.431 | 66.8052 | 106.6021 | 64.1851 | 0.009337 | 1.0466 | 0.024982 | 23.8703 | -0.006973 | 0.14294 | 1.2e-05 | 0.001155 | 0.001155 | 5.0852
2019-03-21 12:29 | 1553185768.3955 | -339 | 0.20952 | 0.000455 | -4e-06 | 100.586 | 100.7123 | 124.7869 | 21.2003 | 0.001921 | 0.001699 | -152.1649 | 0.2404 | 8.4023 | 0.003566 | 0.004286 | 0.004084 | 5.7e-05 | 0.0011 | 0.26864 | 1.3036 | 2.5479 | 99.39 | 20.92 | 37.432 | 67.9875 | 107.635 | 62.6525 | 0.009331 | 1.0513 | 0.025054 | 23.8308 | -0.011224 | 0.14302 | 1.4e-05 | 0.001327 | 0.001327 | 4.4304
2019-03-21 12:29 | 1553185769.3946 | -338 | 0.20953 | 0.000459 | -4e-06 | 100.5732 | 100.635 | 125.1484 | 21.2002 | 0.00186 | 0.001793 | -152.1729 | 0.24028 | 8.4006 | 0.003535 | 0.005371 | 0.005237 | 5.5e-05 | 0.001098 | 0.23004 | 1.3214 | 2.5479 | 99.39 | 20.93 | 37.434 | 66.5255 | 106.8774 | 64.9016 | 0.009337 | 1.0444 | 0.025007 | 23.945 | -0.005495 | 0.14294 | 9e-06 | 0.000934 | 0.000934 | 6.2892
2019-03-21 12:29 | 1553185770.3946 | -337 | 0.20954 | 0.000455 | -4e-06 | 100.6497 | 100.7171 | 125.2364 | 21.2 | 0.001895 | 0.001725 | -152.161 | 0.24002 | 8.4013 | 0.003529 | 0.00516 | 0.005038 | 5.3e-05 | 0.001099 | 0.2338 | 1.2935 | 2.548 | 99.39 | 20.93 | 37.437 | 67.3608 | 106.725 | 63.2764 | 0.009338 | 1.0494 | 0.025076 | 23.8958 | -0.00598 | 0.14279 | 1.4e-05 | 0.001345 | 0.001345 | 4.3636
2019-03-21 12:30 | 1553185771.3956 | -336 | 0.20953 | 0.000463 | -4e-06 | 100.6015 | 100.8039 | 124.9139 | 21.1994 | 0.001856 | 0.001764 | -152.1698 | 0.24015 | 8.4001 | 0.003498 | 0.004745 | 0.004566 | 5.6e-05 | 0.001103 | 0.24513 | 1.282 | 2.548 | 99.39 | 20.93 | 37.437 | 66.4325 | 105.9823 | 64.2058 | 0.009338 | 1.0465 | 0.025009 | 23.898 | -0.01798 | 0.14286 | 4e-06 | 0.000512 | 0.000512 | 11.4689
2019-03-21 12:30 | 1553185772.3947 | -335 | 0.20953 | 0.000456 | -4e-06 | 100.6223 | 100.7903 | 124.7259 | 21.2005 | 0.00184 | 0.00176 | -152.1722 | 0.24042 | 8.4018 | 0.003502 | 0.00512 | 0.004897 | 5.1e-05 | 0.001101 | 0.25752 | 1.2755 | 2.548 | 99.39 | 20.93 | 37.438 | 66.0341 | 106.0778 | 64.1234 | 0.009338 | 1.0468 | 0.024994 | 23.877 | -0.014917 | 0.14303 | 1.3e-05 | 0.001194 | 0.001194 | 4.9228
2019-03-21 12:30 | 1553185773.3962 | -334 | 0.20952 | 0.000458 | -4e-06 | 100.6283 | 100.6724 | 124.8327 | 21.2006 | 0.001876 | 0.001755 | -152.1649 | 0.23878 | 8.4032 | 0.003484 | 0.004437 | 0.004355 | 5.3e-05 | 0.001102 | 0.28119 | 1.3022 | 2.5479 | 99.39 | 20.92 | 37.44 | 66.9106 | 105.6435 | 63.9962 | 0.009333 | 1.0472 | 0.025009 | 23.8828 | -0.003923 | 0.14205 | 1.1e-05 | 0.001076 | 0.001076 | 5.4243
2019-03-21 12:30 | 1553185774.3948 | -333 | 0.20953 | 0.000456 | -4e-06 | 100.6524 | 100.7085 | 124.7933 | 21.201 | 0.001857 | 0.001783 | -152.1867 | 0.23922 | 8.4025 | 0.003486 | 0.005138 | 0.004948 | 5.6e-05 | 0.001103 | 0.22307 | 1.32 | 2.5481 | 99.39 | 20.92 | 37.441 | 66.4514 | 105.6982 | 64.6742 | 0.009333 | 1.0451 | 0.02498 | 23.903 | -0.004985 | 0.14231 | 1.4e-05 | 0.001311 | 0.001311 | 4.4602
2019-03-21 12:30 | 1553185775.3962 | -332 | 0.20955 | 0.000459 | -4e-06 | 100.6879 | 100.9051 | 124.828 | 21.2001 | 0.001918 | 0.001733 | -152.1737 | 0.23975 | 8.401 | 0.003468 | 0.005189 | 0.005023 | 5.3e-05 | 0.001103 | 0.29192 | 1.3292 | 2.5479 | 99.39 | 20.92 | 37.441 | 67.9211 | 105.2574 | 63.4623 | 0.009333 | 1.0488 | 0.025028 | 23.8634 | -0.019272 | 0.14263 | 1e-05 | 0.001005 | 0.001005 | 5.8304
2019-03-21 12:30 | 1553185776.3943 | -331 | 0.20953 | 0.000457 | -4e-06 | 100.6459 | 100.7416 | 124.908 | 21.2011 | 0.001852 | 0.001782 | -152.1496 | 0.24004 | 8.4013 | 0.003491 | 0.004593 | 0.004405 | 5.5e-05 | 0.001107 | 0.29856 | 1.3358 | 2.5479 | 99.39 | 20.92 | 37.441 | 66.3383 | 105.8021 | 64.6455 | 0.009333 | 1.0451 | 0.024993 | 23.913 | -0.008502 | 0.1428 | 1.2e-05 | 0.001167 | 0.001167 | 5.0308
2019-03-21 12:30 | 1553185777.394 | -330 | 0.20955 | 0.000462 | -4e-06 | 100.6935 | 100.6878 | 124.5327 | 21.201 | 0.001889 | 0.00176 | -152.1424 | 0.24026 | 8.4038 | 0.003515 | 0.005261 | 0.005038 | 5.1e-05 | 0.001102 | 0.25105 | 1.3278 | 2.5481 | 99.39 | 20.92 | 37.44 | 67.2231 | 106.3885 | 64.1232 | 0.009333 | 1.0468 | 0.024974 | 23.8585 | 0.000513 | 0.14293 | 5e-06 | 0.000624 | 0.000624 | 9.4214
2019-03-21 12:30 | 1553185778.395 | -329 | 0.20955 | 0.000458 | -4e-06 | 100.6852 | 100.8308 | 124.3123 | 21.2014 | 0.001872 | 0.00179 | -152.1268 | 0.23995 | 8.4015 | 0.003491 | 0.005183 | 0.004996 | 5.6e-05 | 0.001104 | 0.26089 | 1.3463 | 2.548 | 99.39 | 20.92 | 37.439 | 66.8224 | 105.8048 | 64.844 | 0.009333 | 1.0445 | 0.024926 | 23.8629 | -0.012926 | 0.14275 | 1.1e-05 | 0.001083 | 0.001083 | 5.4161
2019-03-21 12:30 | 1553185779.3941 | -328 | 0.20955 | 0.000459 | -4e-06 | 100.7266 | 100.8642 | 123.8975 | 21.2019 | 0.001923 | 0.001774 | -152.1176 | 0.24041 | 8.4001 | 0.003537 | 0.004541 | 0.004383 | 5.5e-05 | 0.001106 | 0.26737 | 1.3197 | 2.5481 | 99.39 | 20.92 | 37.438 | 68.0491 | 106.9183 | 64.4635 | 0.009333 | 1.0457 | 0.024898 | 23.8096 | -0.012208 | 0.14302 | 1e-05 | 0.000982 | 0.000982 | 5.9853
2019-03-21 12:30 | 1553185780.3945 | -327 | 0.20955 | 0.000457 | -4e-06 | 100.7046 | 100.8846 | 123.991 | 21.2003 | 0.001938 | 0.001769 | -152.1321 | 0.24036 | 8.4054 | 0.003566 | 0.005057 | 0.004892 | 5.3e-05 | 0.001108 | 0.26825 | 1.3218 | 2.548 | 99.39 | 20.92 | 37.434 | 68.3902 | 107.6231 | 64.3348 | 0.009332 | 1.0461 | 0.024912 | 23.8141 | -0.015978 | 0.14299 | 1.2e-05 | 0.001098 | 0.001098 | 5.3545
2019-03-21 12:30 | 1553185781.3952 | -326 | 0.20958 | 0.00046 | -4e-06 | 100.7195 | 100.7389 | 124.1346 | 21.2017 | 0.001862 | 0.001783 | -152.1132 | 0.23967 | 8.4028 | 0.003536 | 0.004961 | 0.004783 | 5.3e-05 | 0.001108 | 0.27632 | 1.3125 | 2.548 | 99.39 | 20.93 | 37.432 | 66.5692 | 106.9055 | 64.6609 | 0.009337 | 1.0451 | 0.024914 | 23.8394 | -0.001722 | 0.14258 | 8e-06 | 0.000862 | 0.000862 | 6.7992
2019-03-21 12:30 | 1553185782.3953 | -325 | 0.20957 | 0.000459 | -4e-06 | 100.7549 | 101.0089 | 123.9583 | 21.2012 | 0.001859 | 0.001766 | -152.0937 | 0.23999 | 8.4049 | 0.003428 | 0.004844 | 0.004631 | 5.6e-05 | 0.001113 | 0.26737 | 1.3622 | 2.5481 | 99.39 | 20.93 | 37.431 | 66.493 | 104.2858 | 64.2733 | 0.009337 | 1.0463 | 0.024911 | 23.8088 | -0.022522 | 0.14277 | 1e-05 | 0.000977 | 0.000977 | 6.0089
2019-03-21 12:30 | 1553185783.3957 | -324 | 0.20958 | 0.000458 | -4e-06 | 100.8431 | 100.9664 | 124.0342 | 21.2015 | 0.001873 | 0.001796 | -152.0778 | 0.24066 | 8.4052 | 0.003472 | 0.004712 | 0.004596 | 5.3e-05 | 0.001116 | 0.28031 | 1.3224 | 2.5481 | 99.39 | 20.92 | 37.429 | 66.8365 | 105.3447 | 64.9933 | 0.00933 | 1.0441 | 0.024892 | 23.8414 | -0.010926 | 0.14317 | 1e-05 | 0.00098 | 0.00098 | 6.0053
2019-03-21 12:30 | 1553185784.3954 | -323 | 0.20953 | 0.000459 | -4e-06 | 100.7636 | 100.9688 | 124.1325 | 21.2019 | 0.001848 | 0.001826 | -152.0431 | 0.24002 | 8.403 | 0.003432 | 0.005136 | 0.004945 | 5.3e-05 | 0.001119 | 0.25642 | 1.3361 | 2.548 | 99.39 | 20.92 | 37.428 | 66.2305 | 104.3856 | 65.7093 | 0.00933 | 1.0419 | 0.024876 | 23.8761 | -0.018196 | 0.14279 | 9e-06 | 0.000914 | 0.000914 | 6.4217
2019-03-21 12:30 | 1553185785.3944 | -322 | 0.20955 | 0.000461 | -3e-06 | 100.8029 | 100.935 | 123.8539 | 21.2011 | 0.00188 | 0.001733 | -152.0478 | 0.24019 | 8.4015 | 0.003343 | 0.005202 | 0.005082 | 5.4e-05 | 0.001122 | 0.26305 | 1.2852 | 2.5481 | 99.39 | 20.93 | 37.426 | 67.0108 | 102.2371 | 63.4619 | 0.009335 | 1.0488 | 0.024931 | 23.7701 | -0.011718 | 0.14289 | 6e-06 | 0.000621 | 0.000621 | 9.4574
2019-03-21 12:30 | 1553185786.3955 | -321 | 0.20955 | 0.000456 | -4e-06 | 100.7847 | 100.9398 | 123.9454 | 21.2012 | 0.001919 | 0.001789 | -152.0693 | 0.23999 | 8.4039 | 0.003337 | 0.004792 | 0.004585 | 5.5e-05 | 0.001127 | 0.23651 | 1.3123 | 2.548 | 99.39 | 20.93 | 37.425 | 67.9448 | 102.0738 | 64.825 | 0.009335 | 1.0446 | 0.024889 | 23.827 | -0.013753 | 0.14277 | 1.2e-05 | 0.001137 | 0.001137 | 5.1594
2019-03-21 12:30 | 1553185787.3949 | -320 | 0.20955 | 0.000461 | -4e-06 | 100.827 | 100.8812 | 124.0015 | 21.2018 | 0.001897 | 0.001738 | -152.081 | 0.24079 | 8.4029 | 0.00338 | 0.005191 | 0.005001 | 5.3e-05 | 0.001129 | 0.25608 | 1.329 | 2.5478 | 99.38 | 20.93 | 37.424 | 67.4203 | 103.1209 | 63.5938 | 0.009336 | 1.0484 | 0.024941 | 23.7889 | -0.004807 | 0.14325 | 6e-06 | 0.000719 | 0.000719 | 8.1887
2019-03-21 12:30 | 1553185788.3956 | -319 | 0.20957 | 0.000459 | -4e-06 | 100.8625 | 101.1544 | 123.8806 | 21.2021 | 0.001845 | 0.001765 | -152.0535 | 0.24021 | 8.403 | 0.003379 | 0.005041 | 0.004832 | 5.2e-05 | 0.001131 | 0.2756 | 1.3532 | 2.548 | 99.38 | 20.93 | 37.419 | 66.159 | 103.1085 | 64.2443 | 0.009335 | 1.0464 | 0.024904 | 23.8003 | -0.025848 | 0.1429 | 9e-06 | 0.000907 | 0.000907 | 6.4735
2019-03-21 12:30 | 1553185789.3955 | -318 | 0.20957 | 0.000458 | -4e-06 | 100.8372 | 101.0273 | 124.0271 | 21.2014 | 0.001893 | 0.001752 | -152.0698 | 0.2396 | 8.404 | 0.003388 | 0.004631 | 0.004498 | 5.3e-05 | 0.00113 | 0.26521 | 1.3069 | 2.548 | 99.38 | 20.93 | 37.416 | 67.3077 | 103.3271 | 63.9181 | 0.009334 | 1.0474 | 0.024931 | 23.8028 | -0.016854 | 0.14254 | 1e-05 | 0.001006 | 0.001006 | 5.8224
2019-03-21 12:30 | 1553185790.396 | -317 | 0.20957 | 0.000461 | -3e-06 | 100.883 | 100.9991 | 124.1433 | 21.2027 | 0.001883 | 0.001765 | -152.0777 | 0.24053 | 8.4063 | 0.003455 | 0.005007 | 0.004881 | 5.3e-05 | 0.00113 | 0.24082 | 1.3427 | 2.548 | 99.38 | 20.93 | 37.413 | 67.0718 | 104.9446 | 64.2466 | 0.009333 | 1.0464 | 0.024931 | 23.8256 | -0.010281 | 0.14309 | 6e-06 | 0.00065 | 0.00065 | 9.0516
2019-03-21 12:30 | 1553185791.3947 | -316 | 0.20957 | 0.00046 | -4e-06 | 100.8348 | 100.8833 | 124.1755 | 21.2018 | 0.001835 | 0.001761 | -152.0849 | 0.2406 | 8.4013 | 0.003404 | 0.004039 | 0.003864 | 4.9e-05 | 0.001133 | 0.28368 | 1.3415 | 2.5479 | 99.38 | 20.93 | 37.408 | 65.93 | 103.7116 | 64.1316 | 0.009332 | 1.0467 | 0.024938 | 23.8246 | -0.004301 | 0.14314 | 8e-06 | 0.00083 | 0.00083 | 7.0885
2019-03-21 12:30 | 1553185792.3938 | -315 | 0.20958 | 0.000458 | -4e-06 | 100.8589 | 100.9667 | 123.9811 | 21.2017 | 0.001876 | 0.001735 | -152.1009 | 0.23995 | 8.4002 | 0.003358 | 0.004665 | 0.004503 | 5.2e-05 | 0.00113 | 0.26001 | 1.3611 | 2.5481 | 99.38 | 20.93 | 37.407 | 66.9092 | 102.5831 | 63.5248 | 0.009332 | 1.0486 | 0.024941 | 23.7845 | -0.009555 | 0.14274 | 1.1e-05 | 0.00105 | 0.00105 | 5.5854
2019-03-21 12:30 | 1553185793.3938 | -314 | 0.20958 | 0.000456 | -3e-06 | 100.8687 | 101.1428 | 123.8044 | 21.2022 | 0.001879 | 0.001791 | -152.1095 | 0.24101 | 8.4043 | 0.003457 | 0.004738 | 0.004589 | 5.4e-05 | 0.001132 | 0.26089 | 1.3098 | 2.5479 | 99.38 | 20.93 | 37.403 | 66.9799 | 104.9993 | 64.8652 | 0.009331 | 1.0445 | 0.024874 | 23.8148 | -0.024269 | 0.14338 | 1.2e-05 | 0.001111 | 0.001111 | 5.304
2019-03-21 12:30 | 1553185794.3942 | -313 | 0.20957 | 0.00046 | -4e-06 | 100.8234 | 100.9664 | 124.0182 | 21.2016 | 0.001888 | 0.001774 | -152.1413 | 0.24009 | 8.4039 | 0.003498 | 0.004989 | 0.004766 | 5.2e-05 | 0.00113 | 0.27057 | 1.288 | 2.548 | 99.38 | 20.93 | 37.399 | 67.1892 | 105.9939 | 64.4594 | 0.00933 | 1.0457 | 0.02491 | 23.8211 | -0.01267 | 0.14283 | 8e-06 | 0.000839 | 0.000839 | 6.9995
2019-03-21 12:30 | 1553185795.3943 | -312 | 0.2096 | 0.000456 | -3e-06 | 100.8562 | 100.9401 | 124.3426 | 21.2031 | 0.001889 | 0.001805 | -152.1454 | 0.24081 | 8.4014 | 0.003449 | 0.004988 | 0.004892 | 5.2e-05 | 0.001129 | 0.27273 | 1.2885 | 2.548 | 99.38 | 20.93 | 37.398 | 67.2251 | 104.8002 | 65.2125 | 0.009329 | 1.0434 | 0.024915 | 23.8788 | -0.00744 | 0.14326 | 1.2e-05 | 0.001096 | 0.001096 | 5.3699
2019-03-21 12:30 | 1553185796.394 | -311 | 0.20958 | 0.00046 | -4e-06 | 100.8696 | 100.8904 | 124.2675 | 21.203 | 0.001822 | 0.001774 | -152.131 | 0.24 | 8.4022 | 0.003439 | 0.004307 | 0.004075 | 5.3e-05 | 0.001127 | 0.26825 | 1.2728 | 2.5479 | 99.38 | 20.93 | 37.395 | 65.6012 | 104.5486 | 64.4592 | 0.009329 | 1.0457 | 0.024935 | 23.845 | -0.001843 | 0.14278 | 8e-06 | 0.000819 | 0.000819 | 7.1686
2019-03-21 12:30 | 1553185797.3954 | -310 | 0.20955 | 0.000463 | -4e-06 | 100.8944 | 100.9463 | 124.1297 | 21.2032 | 0.001897 | 0.001781 | -152.1201 | 0.24032 | 8.4027 | 0.003431 | 0.004951 | 0.004757 | 5.2e-05 | 0.001126 | 0.29264 | 1.2517 | 2.548 | 99.38 | 20.93 | 37.395 | 67.4272 | 104.3585 | 64.6287 | 0.009329 | 1.0452 | 0.024915 | 23.8377 | -0.004601 | 0.14297 | 4e-06 | 0.000511 | 0.000511 | 11.4996
2019-03-21 12:30 | 1553185798.3941 | -309 | 0.20958 | 0.000457 | -3e-06 | 100.9516 | 101.0154 | 124.6318 | 21.2029 | 0.001877 | 0.00178 | -152.1347 | 0.24008 | 8.4026 | 0.003515 | 0.004821 | 0.004713 | 5.1e-05 | 0.001121 | 0.23811 | 1.2804 | 2.5479 | 99.38 | 20.93 | 37.395 | 66.9267 | 106.3855 | 64.5971 | 0.009329 | 1.0453 | 0.024967 | 23.8848 | -0.005655 | 0.14282 | 1.1e-05 | 0.001069 | 0.001069 | 5.493
2019-03-21 12:30 | 1553185799.3952 | -308 | 0.20956 | 0.000456 | -4e-06 | 100.8789 | 100.885 | 124.6681 | 21.2028 | 0.001872 | 0.001758 | -152.125 | 0.24024 | 8.4044 | 0.003441 | 0.004768 | 0.004548 | 5.4e-05 | 0.001122 | 0.25409 | 1.311 | 2.548 | 99.38 | 20.93 | 37.396 | 66.8137 | 104.6097 | 64.0787 | 0.009329 | 1.0469 | 0.024989 | 23.8699 | -0.000541 | 0.14292 | 1.3e-05 | 0.00121 | 0.00121 | 4.8561
2019-03-21 12:30 | 1553185800.3942 | -307 | 0.20956 | 0.000457 | -4e-06 | 100.9203 | 100.9698 | 124.1759 | 21.203 | 0.001876 | 0.001781 | -152.1296 | 0.24038 | 8.4033 | 0.003537 | 0.004701 | 0.004565 | 5.1e-05 | 0.001125 | 0.2573 | 1.3245 | 2.548 | 99.38 | 20.93 | 37.397 | 66.9028 | 106.924 | 64.616 | 0.009329 | 1.0452 | 0.02492 | 23.8417 | -0.004386 | 0.143 | 1.1e-05 | 0.001073 | 0.001073 | 5.4777
2019-03-21 12:30 | 1553185801.3944 | -306 | 0.20957 | 0.000459 | -4e-06 | 100.8273 | 101.0491 | 123.9391 | 21.2033 | 0.001855 | 0.001738 | -152.132 | 0.24076 | 8.4057 | 0.003491 | 0.004616 | 0.004417 | 5.6e-05 | 0.001123 | 0.26089 | 1.3452 | 2.5479 | 99.38 | 20.93 | 37.398 | 66.4106 | 105.8225 | 63.5967 | 0.009329 | 1.0484 | 0.024934 | 23.783 | -0.019655 | 0.14323 | 9e-06 | 0.000933 | 0.000933 | 6.3127
2019-03-21 12:30 | 1553185802.3953 | -305 | 0.20957 | 0.000453 | -4e-06 | 100.9346 | 101.1632 | 123.9889 | 21.2026 | 0.00192 | 0.001788 | -152.1269 | 0.23938 | 8.4046 | 0.003531 | 0.005384 | 0.005227 | 5.5e-05 | 0.001123 | 0.27273 | 1.3406 | 2.5479 | 99.38 | 20.92 | 37.399 | 67.9609 | 106.7798 | 64.7898 | 0.009324 | 1.0447 | 0.024895 | 23.8299 | -0.020239 | 0.14241 | 1.6e-05 | 0.001457 | 0.001457 | 4.0176
2019-03-21 12:30 | 1553185803.3944 | -304 | 0.20957 | 0.000458 | -3e-06 | 100.8652 | 100.9371 | 124.0792 | 21.2039 | 0.001884 | 0.001779 | -152.1251 | 0.24084 | 8.4004 | 0.00353 | 0.004676 | 0.004501 | 5.7e-05 | 0.001123 | 0.24851 | 1.3509 | 2.548 | 99.38 | 20.93 | 37.401 | 67.1088 | 106.7534 | 64.588 | 0.00933 | 1.0453 | 0.024912 | 23.8315 | -0.006375 | 0.14327 | 1e-05 | 0.000935 | 0.000935 | 6.2992
2019-03-21 12:30 | 1553185804.3956 | -303 | 0.20957 | 0.000457 | -4e-06 | 100.9048 | 101.1493 | 123.9595 | 21.2031 | 0.001899 | 0.001769 | -152.1138 | 0.24019 | 8.402 | 0.003512 | 0.004531 | 0.004376 | 5.4e-05 | 0.001126 | 0.27489 | 1.337 | 2.548 | 99.38 | 20.92 | 37.402 | 67.4574 | 106.3314 | 64.3392 | 0.009325 | 1.0461 | 0.024909 | 23.8112 | -0.021645 | 0.14289 | 1.1e-05 | 0.001101 | 0.001101 | 5.3331
2019-03-21 12:30 | 1553185805.3955 | -302 | 0.20956 | 0.000458 | -4e-06 | 100.8321 | 100.9214 | 123.7917 | 21.2038 | 0.001892 | 0.001799 | -152.1055 | 0.24034 | 8.4015 | 0.003573 | 0.004438 | 0.00421 | 5.8e-05 | 0.001124 | 0.25912 | 1.3351 | 2.548 | 99.38 | 20.92 | 37.405 | 67.2929 | 107.8045 | 65.0578 | 0.009325 | 1.0439 | 0.024865 | 23.8204 | -0.00792 | 0.14298 | 1e-05 | 0.000953 | 0.000953 | 6.1635
2019-03-21 12:30 | 1553185806.3956 | -301 | 0.20955 | 0.000459 | -4e-06 | 100.8813 | 101.2344 | 123.7856 | 21.2038 | 0.001879 | 0.001775 | -152.128 | 0.23944 | 8.4022 | 0.003441 | 0.004953 | 0.004846 | 5.7e-05 | 0.001126 | 0.28031 | 1.3592 | 2.5481 | 99.38 | 20.92 | 37.408 | 66.9895 | 104.5942 | 64.47 | 0.009326 | 1.0457 | 0.024886 | 23.7991 | -0.031258 | 0.14244 | 9e-06 | 0.000905 | 0.000905 | 6.4673
2019-03-21 12:30 | 1553185807.3946 | -300 | 0.20957 | 0.000458 | -4e-06 | 100.847 | 101.0774 | 123.7961 | 21.204 | 0.00196 | 0.001765 | -152.1496 | 0.2412 | 8.404 | 0.00345 | 0.00421 | 0.004032 | 5.5e-05 | 0.001125 | 0.29551 | 1.3532 | 2.548 | 99.38 | 20.92 | 37.41 | 68.9327 | 104.8333 | 64.2524 | 0.009327 | 1.0464 | 0.024896 | 23.7924 | -0.020413 | 0.14349 | 1e-05 | 0.000952 | 0.000952 | 6.1943
2019-03-21 12:30 | 1553185808.395 | -299 | 0.20958 | 0.00046 | -3e-06 | 100.9096 | 101.1132 | 123.5445 | 21.2047 | 0.001908 | 0.001793 | -152.1603 | 0.24058 | 8.401 | 0.003495 | 0.004883 | 0.004686 | 5.9e-05 | 0.001128 | 0.23778 | 1.3353 | 2.5479 | 99.38 | 20.92 | 37.412 | 67.6775 | 105.9104 | 64.9059 | 0.009327 | 1.0443 | 0.024846 | 23.7913 | -0.018029 | 0.14312 | 7e-06 | 0.000759 | 0.000759 | 7.749
2019-03-21 12:30 | 1553185809.3937 | -298 | 0.20961 | 0.000462 | -4e-06 | 100.937 | 101.0587 | 123.5727 | 21.2047 | 0.001928 | 0.001836 | -152.1623 | 0.23943 | 8.3996 | 0.003505 | 0.004715 | 0.004515 | 5.9e-05 | 0.001132 | 0.29214 | 1.3126 | 2.548 | 99.38 | 20.92 | 37.415 | 68.165 | 106.1495 | 65.9415 | 0.009328 | 1.0412 | 0.024811 | 23.8304 | -0.010778 | 0.14244 | 5e-06 | 0.000629 | 0.000629 | 9.3014
2019-03-21 12:30 | 1553185810.3948 | -297 | 0.20957 | 0.000459 | -4e-06 | 100.881 | 101.1987 | 123.6443 | 21.2043 | 0.001918 | 0.001772 | -152.1384 | 0.24041 | 8.4016 | 0.003489 | 0.004549 | 0.00442 | 6.2e-05 | 0.001132 | 0.26305 | 1.3242 | 2.5481 | 99.38 | 20.92 | 37.416 | 67.9243 | 105.766 | 64.4065 | 0.009328 | 1.0459 | 0.024875 | 23.7833 | -0.028124 | 0.14302 | 9e-06 | 0.000918 | 0.000918 | 6.4058
2019-03-21 12:30 | 1553185811.3953 | -296 | 0.20953 | 0.000462 | -4e-06 | 100.8845 | 101.1121 | 123.6879 | 21.2041 | 0.001865 | 0.001804 | -152.1518 | 0.24002 | 8.4042 | 0.003511 | 0.004373 | 0.00414 | 5.7e-05 | 0.001135 | 0.28479 | 1.3382 | 2.548 | 99.38 | 20.92 | 37.42 | 66.6394 | 106.3081 | 65.1784 | 0.009329 | 1.0435 | 0.024851 | 23.8146 | -0.020158 | 0.14279 | 6e-06 | 0.000738 | 0.000738 | 7.9522
2019-03-21 12:30 | 1553185812.3938 | -295 | 0.20956 | 0.000457 | -4e-06 | 100.8524 | 101.1748 | 123.921 | 21.2044 | 0.001935 | 0.001817 | -152.1485 | 0.23882 | 8.4063 | 0.003491 | 0.004867 | 0.004746 | 5.6e-05 | 0.001137 | 0.26394 | 1.2795 | 2.5479 | 99.38 | 20.92 | 37.422 | 68.3239 | 105.8266 | 65.4902 | 0.00933 | 1.0425 | 0.024863 | 23.8481 | -0.028553 | 0.14208 | 1.1e-05 | 0.001042 | 0.001042 | 5.606
2019-03-21 12:30 | 1553185813.3939 | -294 | 0.2097 | 0.000457 | -3e-06 | 100.8711 | 101.0522 | 123.7668 | 21.2055 | 0.001886 | 0.00182 | -152.1485 | 0.23989 | 8.4015 | 0.003463 | 0.0049 | 0.004764 | 5.8e-05 | 0.001134 | 0.24585 | 1.3011 | 2.5479 | 99.38 | 20.92 | 37.422 | 67.1462 | 105.1306 | 65.5598 | 0.00933 | 1.0423 | 0.024844 | 23.8357 | -0.016042 | 0.14271 | 1.1e-05 | 0.001043 | 0.001043 | 5.6247
2019-03-21 12:30 | 1553185814.3939 | -293 | 0.20962 | 0.000458 | -4e-06 | 100.8711 | 101.122 | 123.9349 | 21.2053 | 0.001935 | 0.001794 | -152.1365 | 0.23935 | 8.4042 | 0.003456 | 0.004964 | 0.004793 | 5.6e-05 | 0.001131 | 0.25465 | 1.3242 | 2.5479 | 99.38 | 20.92 | 37.428 | 68.3388 | 104.9677 | 64.9423 | 0.009331 | 1.0442 | 0.024884 | 23.8301 | -0.02222 | 0.14239 | 1e-05 | 0.00101 | 0.00101 | 5.7963
2019-03-21 12:30 | 1553185815.3951 | -292 | 0.20952 | 0.000458 | -4e-06 | 100.7904 | 101.0621 | 123.9717 | 21.2044 | 0.001915 | 0.001799 | -152.1434 | 0.24003 | 8.4019 | 0.003483 | 0.005112 | 0.004898 | 5.7e-05 | 0.001135 | 0.28727 | 1.3285 | 2.5482 | 99.38 | 20.92 | 37.43 | 67.8448 | 105.6324 | 65.0702 | 0.009332 | 1.0438 | 0.024883 | 23.8381 | -0.024079 | 0.1428 | 1.1e-05 | 0.001055 | 0.001055 | 5.5611
2019-03-21 12:30 | 1553185816.3954 | -291 | 0.20962 | 0.000461 | 1e-06 | 100.9206 | 101.0679 | 123.7288 | 21.206 | 0.001931 | 0.001803 | -152.1085 | 0.2404 | 8.4051 | 0.003562 | 0.004721 | 0.00455 | 5.7e-05 | 0.001136 | 0.2808 | 1.3421 | 2.548 | 99.38 | 20.92 | 37.433 | 68.2412 | 107.5396 | 65.1499 | 0.009332 | 1.0436 | 0.024856 | 23.8176 | -0.013044 | 0.14302 | 0 | -9.2e-05 | -9.2e-05 | -63.7315
2019-03-21 12:30 | 1553185817.3942 | -290 | 0.20945 | 0.000461 | 7.3e-05 | 100.872 | 101.2307 | 123.9529 | 21.2049 | 0.001904 | 0.001786 | -152.0858 | 0.23982 | 8.4052 | 0.003464 | 0.00472 | 0.004553 | 5.9e-05 | 0.001137 | 0.27345 | 1.3653 | 2.548 | 99.38 | 20.92 | 37.435 | 67.5904 | 105.1646 | 64.742 | 0.009333 | 1.0448 | 0.024893 | 23.8248 | -0.031746 | 0.14267 | -9.1e-05 | -0.011349 | -0.011349 | -0.51666
2019-03-21 12:30 | 1553185818.3951 | -289 | 0.2097 | 0.00046 | 0.000207 | 100.8854 | 101.0931 | 123.8548 | 21.2056 | 0.001867 | 0.001759 | -152.0592 | 0.24103 | 8.4041 | 0.003489 | 0.004259 | 0.00408 | 5.9e-05 | 0.001135 | 0.2896 | 1.3351 | 2.5478 | 99.38 | 20.93 | 37.44 | 66.6874 | 105.7665 | 64.0971 | 0.00934 | 1.0468 | 0.024907 | 23.7926 | -0.018391 | 0.14339 | -0.000259 | -0.032315 | -0.032315 | -0.18237
2019-03-21 12:30 | 1553185819.3946 | -288 | 0.20976 | 0.000465 | 0.000329 | 100.8398 | 101.1673 | 123.9823 | 21.2064 | 0.001929 | 0.001784 | -152.0545 | 0.24142 | 8.3994 | 0.003541 | 0.00489 | 0.004639 | 5.8e-05 | 0.001137 | 0.26936 | 1.3115 | 2.548 | 99.38 | 20.93 | 37.442 | 68.1907 | 107.0377 | 64.7044 | 0.00934 | 1.045 | 0.024898 | 23.8263 | -0.029001 | 0.14362 | -0.00042 | -0.052032 | -0.052032 | -0.11344
2019-03-21 12:30 | 1553185820.3942 | -287 | 0.20976 | 0.000506 | 0.000375 | 100.8696 | 101.0338 | 124.0959 | 21.2058 | 0.001851 | 0.001801 | -152.0622 | 0.24013 | 8.4013 | 0.003501 | 0.005336 | 0.005176 | 5.7e-05 | 0.001133 | 0.25912 | 1.3247 | 2.5481 | 99.38 | 20.93 | 37.446 | 66.3077 | 106.0587 | 65.1115 | 0.009341 | 1.0437 | 0.024894 | 23.8515 | -0.014546 | 0.14286 | -0.000531 | -0.063183 | -0.063183 | -0.092927
2019-03-21 12:30 | 1553185821.3953 | -286 | 0.20962 | 0.00063 | 0.000394 | 100.8035 | 100.9357 | 123.8475 | 21.2061 | 0.001868 | 0.001774 | -152.0308 | 0.24008 | 8.4031 | 0.003458 | 0.004978 | 0.004864 | 6.3e-05 | 0.001135 | 0.29391 | 1.3031 | 2.548 | 99.38 | 20.93 | 37.448 | 66.723 | 105.0186 | 64.4691 | 0.009342 | 1.0457 | 0.024893 | 23.805 | -0.011725 | 0.14283 | -0.000712 | -0.07792 | -0.07792 | -0.075335
2019-03-21 12:30 | 1553185822.3945 | -285 | 0.20954 | 0.000777 | 0.000411 | 100.8294 | 101.1997 | 123.9154 | 21.2061 | 0.001949 | 0.001783 | -152.0356 | 0.23981 | 8.4033 | 0.00357 | 0.004492 | 0.004356 | 6.2e-05 | 0.001138 | 0.28711 | 1.2897 | 2.548 | 99.38 | 20.93 | 37.449 | 68.6703 | 107.7463 | 64.6746 | 0.009342 | 1.0451 | 0.024892 | 23.8188 | -0.032786 | 0.14266 | -0.00092 | -0.094496 | -0.094496 | -0.06205
2019-03-21 12:30 | 1553185823.3938 | -284 | 0.20943 | 0.000904 | 0.000425 | 100.9268 | 101.1217 | 124.1405 | 21.2053 | 0.001927 | 0.001774 | -152.037 | 0.23985 | 8.4034 | 0.00354 | 0.00428 | 0.004177 | 5.9e-05 | 0.001137 | 0.24779 | 1.3423 | 2.5479 | 99.38 | 20.93 | 37.451 | 68.1469 | 106.9949 | 64.4707 | 0.009343 | 1.0457 | 0.024922 | 23.8332 | -0.017251 | 0.14269 | -0.001098 | -0.10879 | -0.10879 | -0.053907
2019-03-21 12:30 | 1553185824.3944 | -283 | 0.20935 | 0.001014 | 0.000436 | 100.872 | 100.9643 | 124.0867 | 21.2064 | 0.001944 | 0.001772 | -152.0552 | 0.23937 | 8.4017 | 0.003546 | 0.005286 | 0.005082 | 5.8e-05 | 0.001139 | 0.27041 | 1.3137 | 2.5479 | 99.38 | 20.93 | 37.452 | 68.5396 | 107.1487 | 64.4116 | 0.009343 | 1.0459 | 0.024919 | 23.826 | -0.008182 | 0.1424 | -0.001251 | -0.12089 | -0.12089 | -0.048412
2019-03-21 12:30 | 1553185825.3935 | -282 | 0.20931 | 0.001105 | 0.000445 | 100.8327 | 100.822 | 124.0165 | 21.2066 | 0.001858 | 0.001773 | -152.0568 | 0.23993 | 8.4016 | 0.003577 | 0.004838 | 0.004654 | 6.1e-05 | 0.001141 | 0.26864 | 1.3082 | 2.548 | 99.38 | 20.93 | 37.454 | 66.4819 | 107.9177 | 64.433 | 0.009343 | 1.0458 | 0.024911 | 23.82 | 0.000952 | 0.14273 | -0.001379 | -0.13097 | -0.13097 | -0.044792
2019-03-21 12:30 | 1553185826.3935 | -281 | 0.20932 | 0.001181 | 0.000453 | 100.855 | 101.1203 | 123.9088 | 21.2061 | 0.001897 | 0.001816 | -152.0738 | 0.23938 | 8.4027 | 0.00358 | 0.004251 | 0.004189 | 6.2e-05 | 0.001139 | 0.25122 | 1.3291 | 2.5479 | 99.38 | 20.93 | 37.455 | 67.4098 | 107.9881 | 65.473 | 0.009344 | 1.0426 | 0.024862 | 23.8463 | -0.023496 | 0.14241 | -0.001485 | -0.13908 | -0.13908 | -0.042084
2019-03-21 12:30 | 1553185827.3946 | -280 | 0.20931 | 0.001249 | 0.000459 | 100.7838 | 101.0999 | 124.0278 | 21.2061 | 0.001899 | 0.001766 | -152.079 | 0.24054 | 8.4005 | 0.003537 | 0.004719 | 0.004552 | 5.9e-05 | 0.001143 | 0.27079 | 1.3426 | 2.548 | 99.38 | 20.93 | 37.455 | 67.4759 | 106.9428 | 64.2788 | 0.009344 | 1.0463 | 0.024918 | 23.8156 | -0.028006 | 0.1431 | -0.001579 | -0.14686 | -0.14686 | -0.040046
2019-03-21 12:30 | 1553185828.3947 | -279 | 0.20931 | 0.001302 | 0.000464 | 100.827 | 101.0662 | 124.21 | 21.2062 | 0.001906 | 0.001744 | -152.1046 | 0.23896 | 8.3999 | 0.003493 | 0.004886 | 0.004713 | 5.9e-05 | 0.001141 | 0.25481 | 1.3106 | 2.5481 | 99.38 | 20.93 | 37.454 | 67.629 | 105.8568 | 63.7292 | 0.009343 | 1.048 | 0.024956 | 23.8137 | -0.021189 | 0.14216 | -0.001652 | -0.15284 | -0.15284 | -0.038226
2019-03-21 12:30 | 1553185829.3943 | -278 | 0.20929 | 0.001341 | 0.000468 | 100.8169 | 101.013 | 123.7337 | 21.2063 | 0.001911 | 0.001808 | -152.1058 | 0.24032 | 8.4013 | 0.003501 | 0.004683 | 0.00458 | 5.6e-05 | 0.001139 | 0.24834 | 1.2963 | 2.5478 | 99.38 | 20.93 | 37.455 | 67.7656 | 106.0543 | 65.289 | 0.009344 | 1.0432 | 0.024851 | 23.8229 | -0.017384 | 0.14297 | -0.001708 | -0.1566 | -0.1566 | -0.037522
2019-03-21 12:30 | 1553185830.3944 | -277 | 0.2093 | 0.001381 | 0.000472 | 100.8202 | 101.0369 | 123.607 | 21.2059 | 0.001928 | 0.001768 | -152.1115 | 0.24072 | 8.4023 | 0.003509 | 0.005404 | 0.005176 | 5.9e-05 | 0.001139 | 0.27079 | 1.3069 | 2.548 | 99.38 | 20.93 | 37.455 | 68.1538 | 106.2587 | 64.3043 | 0.009344 | 1.0462 | 0.024875 | 23.7761 | -0.019204 | 0.1432 | -0.001763 | -0.16111 | -0.16111 | -0.036533
2019-03-21 12:31 | 1553185831.3942 | -276 | 0.20932 | 0.00141 | 0.000474 | 100.794 | 101.1336 | 123.702 | 21.2065 | 0.00194 | 0.001812 | -152.1286 | 0.24004 | 8.4023 | 0.003447 | 0.0046 | 0.004502 | 5.8e-05 | 0.00114 | 0.27544 | 1.3227 | 2.548 | 99.38 | 20.93 | 37.455 | 68.4538 | 104.7495 | 65.3806 | 0.009344 | 1.0429 | 0.024845 | 23.8231 | -0.03009 | 0.1428 | -0.001803 | -0.1641 | -0.1641 | -0.035765
2019-03-21 12:31 | 1553185832.3949 | -275 | 0.20932 | 0.001435 | 0.000477 | 100.7755 | 100.8281 | 123.7844 | 21.2069 | 0.001944 | 0.001795 | -152.135 | 0.24005 | 8.4017 | 0.003412 | 0.004337 | 0.004125 | 5.7e-05 | 0.001139 | 0.27688 | 1.3125 | 2.548 | 99.38 | 20.93 | 37.455 | 68.5495 | 103.9022 | 64.9696 | 0.009344 | 1.0441 | 0.024868 | 23.8166 | -0.004668 | 0.14281 | -0.001837 | -0.16698 | -0.16698 | -0.03515
2019-03-21 12:31 | 1553185833.395 | -274 | 0.20929 | 0.001457 | 0.000479 | 100.8589 | 100.8434 | 123.3466 | 21.2066 | 0.001896 | 0.001784 | -152.1451 | 0.24043 | 8.4019 | 0.003429 | 0.004225 | 0.004046 | 5.7e-05 | 0.00114 | 0.2636 | 1.3009 | 2.5479 | 99.38 | 20.93 | 37.455 | 67.4009 | 104.3145 | 64.6897 | 0.009344 | 1.045 | 0.024834 | 23.7646 | 0.001374 | 0.14303 | -0.001868 | -0.16912 | -0.16912 | -0.03476
2019-03-21 12:31 | 1553185834.3954 | -273 | 0.20929 | 0.001474 | 0.00048 | 100.8366 | 100.9858 | 123.246 | 21.2069 | 0.001877 | 0.001752 | -152.1377 | 0.24027 | 8.4015 | 0.003434 | 0.004707 | 0.00452 | 5.6e-05 | 0.001142 | 0.27328 | 1.3001 | 2.548 | 99.38 | 20.93 | 37.455 | 66.9494 | 104.4391 | 63.9296 | 0.009344 | 1.0474 | 0.024852 | 23.7281 | -0.013227 | 0.14293 | -0.001891 | -0.17104 | -0.17104 | -0.034346
2019-03-21 12:31 | 1553185835.395 | -272 | 0.20947 | 0.001489 | 0.000482 | 100.8232 | 100.9742 | 123.1711 | 21.2072 | 0.00193 | 0.001818 | -152.1458 | 0.24103 | 8.4029 | 0.003377 | 0.004692 | 0.004531 | 5.8e-05 | 0.001142 | 0.28119 | 1.2775 | 2.5479 | 99.38 | 20.93 | 37.455 | 68.2034 | 103.0482 | 65.5108 | 0.009344 | 1.0425 | 0.024786 | 23.7765 | -0.013391 | 0.14339 | -0.001913 | -0.17244 | -0.17244 | -0.034175
2019-03-21 12:31 | 1553185836.3945 | -271 | 0.20961 | 0.0015 | 0.000482 | 100.8991 | 101.0032 | 123.0924 | 21.2073 | 0.001891 | 0.001771 | -152.1412 | 0.2399 | 8.4029 | 0.003468 | 0.004903 | 0.004673 | 5.8e-05 | 0.001138 | 0.26864 | 1.2533 | 2.5479 | 99.38 | 20.93 | 37.455 | 67.2768 | 105.2669 | 64.3906 | 0.009344 | 1.0459 | 0.02482 | 23.7296 | -0.009217 | 0.14272 | -0.001928 | -0.17395 | -0.17395 | -0.03372
2019-03-21 12:31 | 1553185837.3944 | -270 | 0.20958 | 0.001512 | 0.000483 | 100.7946 | 101.0706 | 123.0469 | 21.2078 | 0.001974 | 0.001769 | -152.141 | 0.24086 | 8.4027 | 0.003448 | 0.005065 | 0.004892 | 5.7e-05 | 0.001141 | 0.28119 | 1.2681 | 2.5481 | 99.38 | 20.93 | 37.456 | 69.2623 | 104.7893 | 64.3295 | 0.009344 | 1.0461 | 0.024817 | 23.723 | -0.024462 | 0.14329 | -0.001943 | -0.17508 | -0.17508 | -0.033637
2019-03-21 12:31 | 1553185838.3935 | -269 | 0.20959 | 0.001517 | 0.000483 | 100.8518 | 100.9545 | 123.1495 | 21.2078 | 0.001939 | 0.001809 | -152.144 | 0.24014 | 8.4002 | 0.003516 | 0.004719 | 0.00454 | 5.9e-05 | 0.001137 | 0.27057 | 1.3156 | 2.548 | 99.38 | 20.93 | 37.455 | 68.4402 | 106.4281 | 65.3005 | 0.009344 | 1.0431 | 0.024792 | 23.767 | -0.009103 | 0.14286 | -0.001951 | -0.17548 | -0.17548 | -0.033459
2019-03-21 12:31 | 1553185839.3933 | -268 | 0.20957 | 0.001527 | 0.000484 | 100.8413 | 100.9027 | 123.7492 | 21.2083 | 0.001963 | 0.0018 | -152.1433 | 0.23973 | 8.402 | 0.003497 | 0.004515 | 0.00433 | 6e-05 | 0.001137 | 0.2505 | 1.3233 | 2.548 | 99.38 | 20.93 | 37.455 | 69.003 | 105.9612 | 65.0763 | 0.009344 | 1.0438 | 0.02486 | 23.817 | -0.00544 | 0.14262 | -0.001964 | -0.17702 | -0.17702 | -0.033112
2019-03-21 12:31 | 1553185840.3943 | -267 | 0.20958 | 0.001532 | 0.000485 | 100.8389 | 100.9681 | 124.1149 | 21.207 | 0.001925 | 0.001785 | -152.1404 | 0.23941 | 8.4027 | 0.003539 | 0.005421 | 0.005313 | 6e-05 | 0.001138 | 0.27129 | 1.3149 | 2.548 | 99.38 | 20.93 | 37.453 | 68.0923 | 106.9753 | 64.7364 | 0.009343 | 1.0449 | 0.02491 | 23.8401 | -0.011447 | 0.14243 | -0.001971 | -0.17792 | -0.17792 | -0.032901
2019-03-21 12:31 | 1553185841.3947 | -266 | 0.20964 | 0.001533 | 0.000485 | 100.8482 | 100.8516 | 124.2165 | 21.2078 | 0.001954 | 0.001773 | -152.1566 | 0.23884 | 8.4019 | 0.00357 | 0.004722 | 0.004561 | 5.8e-05 | 0.001138 | 0.25249 | 1.2964 | 2.548 | 99.38 | 20.93 | 37.452 | 68.796 | 107.7392 | 64.4373 | 0.009343 | 1.0458 | 0.024931 | 23.8393 | -0.000303 | 0.14209 | -0.001973 | -0.17825 | -0.17825 | -0.032761
2019-03-21 12:31 | 1553185842.3944 | -265 | 0.20962 | 0.001537 | 0.000485 | 100.8929 | 100.9466 | 124.1433 | 21.2084 | 0.00196 | 0.001807 | -152.16 | 0.24004 | 8.4023 | 0.003528 | 0.004971 | 0.004781 | 6.3e-05 | 0.001138 | 0.3011 | 1.2958 | 2.548 | 99.38 | 20.93 | 37.451 | 68.9252 | 106.7231 | 65.2591 | 0.009343 | 1.0433 | 0.024893 | 23.8613 | -0.004764 | 0.1428 | -0.001979 | -0.17847 | -0.17847 | -0.032886
2019-03-21 12:31 | 1553185843.3946 | -264 | 0.20962 | 0.00154 | 0.000486 | 100.8369 | 100.9136 | 124.2574 | 21.2081 | 0.001881 | 0.001804 | -152.163 | 0.24019 | 8.4016 | 0.003544 | 0.005265 | 0.005088 | 5.9e-05 | 0.001138 | 0.27007 | 1.3196 | 2.548 | 99.38 | 20.93 | 37.447 | 67.0263 | 107.1031 | 65.1861 | 0.009342 | 1.0435 | 0.024907 | 23.8697 | -0.006803 | 0.14289 | -0.001983 | -0.17892 | -0.17892 | -0.032822
2019-03-21 12:31 | 1553185844.394 | -263 | 0.20958 | 0.001542 | 0.000486 | 100.8148 | 100.9398 | 124.0022 | 21.2081 | 0.001907 | 0.001834 | -152.1586 | 0.24058 | 8.4005 | 0.003473 | 0.004862 | 0.004671 | 5.9e-05 | 0.001138 | 0.26808 | 1.2973 | 2.548 | 99.38 | 20.93 | 37.446 | 67.6558 | 105.3896 | 65.9073 | 0.009341 | 1.0413 | 0.024855 | 23.8706 | -0.011083 | 0.14312 | -0.001985 | -0.17869 | -0.17869 | -0.03292
2019-03-21 12:31 | 1553185845.3946 | -262 | 0.20958 | 0.001547 | 0.000485 | 100.8401 | 101.0106 | 123.9619 | 21.2086 | 0.001952 | 0.001871 | -152.1433 | 0.24045 | 8.4031 | 0.0036 | 0.005272 | 0.005116 | 5.7e-05 | 0.001142 | 0.27295 | 1.3347 | 2.5482 | 99.38 | 20.93 | 37.445 | 68.7358 | 108.467 | 66.7873 | 0.009341 | 1.0386 | 0.024819 | 23.8976 | -0.015112 | 0.14304 | -0.001991 | -0.17878 | -0.17878 | -0.032885
2019-03-21 12:31 | 1553185846.3946 | -261 | 0.20959 | 0.001548 | 0.000485 | 100.8208 | 100.9017 | 124.3224 | 21.2083 | 0.00193 | 0.001799 | -152.1589 | 0.23986 | 8.3997 | 0.003559 | 0.005039 | 0.004782 | 6e-05 | 0.001141 | 0.27455 | 1.3169 | 2.5481 | 99.38 | 20.93 | 37.442 | 68.2028 | 107.4695 | 65.0538 | 0.00934 | 1.0439 | 0.024919 | 23.8713 | -0.007173 | 0.14269 | -0.001993 | -0.17969 | -0.17969 | -0.032638
2019-03-21 12:31 | 1553185847.3953 | -260 | 0.20956 | 0.001548 | 0.000486 | 100.8059 | 100.9565 | 124.1086 | 21.2086 | 0.001927 | 0.00176 | -152.1668 | 0.24056 | 8.4018 | 0.003564 | 0.004905 | 0.004773 | 5.6e-05 | 0.00114 | 0.26488 | 1.3204 | 2.5481 | 99.38 | 20.93 | 37.44 | 68.1526 | 107.5888 | 64.1171 | 0.00934 | 1.0468 | 0.024932 | 23.8177 | -0.013355 | 0.14311 | -0.001993 | -0.1798 | -0.1798 | -0.032712
2019-03-21 12:31 | 1553185848.3939 | -259 | 0.20958 | 0.001553 | 0.000485 | 100.8461 | 101.0508 | 124.0687 | 21.2085 | 0.001923 | 0.001798 | -152.1511 | 0.23963 | 8.4027 | 0.003599 | 0.004531 | 0.00441 | 5.9e-05 | 0.001143 | 0.27815 | 1.385 | 2.5481 | 99.38 | 20.93 | 37.439 | 68.0556 | 108.43 | 65.0398 | 0.00934 | 1.0439 | 0.024894 | 23.8464 | -0.018141 | 0.14255 | -0.001999 | -0.17995 | -0.17995 | -0.032559
2019-03-21 12:31 | 1553185849.3949 | -258 | 0.2096 | 0.001556 | 0.000486 | 100.8488 | 101.0978 | 123.7311 | 21.2101 | 0.001938 | 0.001796 | -152.1446 | 0.24031 | 8.4024 | 0.003551 | 0.00484 | 0.004594 | 6.2e-05 | 0.001142 | 0.29463 | 1.3568 | 2.5479 | 99.38 | 20.93 | 37.435 | 68.4105 | 107.2664 | 64.9961 | 0.009339 | 1.0441 | 0.024862 | 23.8124 | -0.022063 | 0.14296 | -0.002003 | -0.18008 | -0.18008 | -0.032628
2019-03-21 12:31 | 1553185850.3949 | -257 | 0.20961 | 0.001551 | 0.000485 | 100.8184 | 101.0934 | 123.6415 | 21.2093 | 0.001938 | 0.00185 | -152.1417 | 0.24035 | 8.4019 | 0.003565 | 0.005212 | 0.005013 | 5.8e-05 | 0.001141 | 0.27328 | 1.3309 | 2.5479 | 99.38 | 20.94 | 37.435 | 68.4159 | 107.6129 | 66.2901 | 0.009344 | 1.0401 | 0.024805 | 23.8493 | -0.024367 | 0.14298 | -0.001997 | -0.17917 | -0.17917 | -0.0328
2019-03-21 12:31 | 1553185851.3952 | -256 | 0.20955 | 0.001555 | 0.000485 | 100.8503 | 101.0263 | 124.0161 | 21.2103 | 0.001963 | 0.001831 | -152.1262 | 0.23984 | 8.4014 | 0.003562 | 0.004555 | 0.004423 | 5.8e-05 | 0.001138 | 0.25912 | 1.2997 | 2.5478 | 99.38 | 20.93 | 37.434 | 69.0119 | 107.5542 | 65.8283 | 0.009338 | 1.0415 | 0.02486 | 23.8691 | -0.0156 | 0.14268 | -0.002 | -0.17972 | -0.17972 | -0.032628
2019-03-21 12:31 | 1553185852.395 | -255 | 0.20957 | 0.001551 | 0.000485 | 100.8208 | 100.9827 | 124.0478 | 21.2091 | 0.001945 | 0.001823 | -152.107 | 0.2396 | 8.4035 | 0.003493 | 0.00482 | 0.004678 | 6.1e-05 | 0.00114 | 0.25841 | 1.2769 | 2.5478 | 99.38 | 20.93 | 37.432 | 68.5742 | 105.8808 | 65.6391 | 0.009338 | 1.0421 | 0.02487 | 23.8655 | -0.014356 | 0.14253 | -0.001997 | -0.17957 | -0.17957 | -0.032623
2019-03-21 12:31 | 1553185853.395 | -254 | 0.20958 | 0.001551 | 0.000485 | 100.8184 | 101.0818 | 123.8767 | 21.2101 | 0.001971 | 0.001865 | -152.1123 | 0.24017 | 8.399 | 0.003523 | 0.004746 | 0.004575 | 6.1e-05 | 0.00114 | 0.25442 | 1.2831 | 2.5478 | 99.38 | 20.93 | 37.43 | 69.1912 | 106.6068 | 66.6577 | 0.009337 | 1.039 | 0.024815 | 23.8849 | -0.023342 | 0.14288 | -0.001996 | -0.17913 | -0.17913 | -0.032782
2019-03-21 12:31 | 1553185854.3935 | -253 | 0.20959 | 0.001553 | 0.000485 | 100.8661 | 101.0178 | 123.9107 | 21.2094 | 0.001923 | 0.001807 | -152.1136 | 0.23979 | 8.4018 | 0.00345 | 0.004346 | 0.004191 | 6e-05 | 0.001139 | 0.24204 | 1.3357 | 2.548 | 99.38 | 20.93 | 37.426 | 68.0543 | 104.8232 | 65.267 | 0.009336 | 1.0432 | 0.02487 | 23.8392 | -0.013446 | 0.14265 | -0.001999 | -0.17971 | -0.17971 | -0.032624
2019-03-21 12:31 | 1553185855.3931 | -252 | 0.20949 | 0.001552 | 0.000485 | 100.799 | 100.9098 | 123.9445 | 21.2101 | 0.001907 | 0.001826 | -152.1106 | 0.24009 | 8.3986 | 0.003526 | 0.004918 | 0.004695 | 5.8e-05 | 0.001139 | 0.25874 | 1.3281 | 2.548 | 99.38 | 20.93 | 37.425 | 67.6546 | 106.6728 | 65.7103 | 0.009336 | 1.0419 | 0.024857 | 23.8581 | -0.009828 | 0.14283 | -0.001997 | -0.1794 | -0.1794 | -0.032722
2019-03-21 12:31 | 1553185856.3942 | -251 | 0.20948 | 0.001552 | 0.000484 | 100.8464 | 100.9401 | 123.909 | 21.2102 | 0.001969 | 0.001821 | -152.1077 | 0.23904 | 8.4016 | 0.003599 | 0.004168 | 0.004049 | 6.1e-05 | 0.001145 | 0.27433 | 1.3434 | 2.548 | 99.38 | 20.93 | 37.421 | 69.1493 | 108.4317 | 65.594 | 0.009335 | 1.0422 | 0.024858 | 23.8506 | -0.008312 | 0.14221 | -0.001996 | -0.1793 | -0.1793 | -0.032597
2019-03-21 12:31 | 1553185857.3942 | -250 | 0.20947 | 0.001551 | 0.000485 | 100.8032 | 101.0488 | 123.9764 | 21.2099 | 0.001919 | 0.001806 | -152.0954 | 0.23975 | 8.4024 | 0.003574 | 0.00522 | 0.005027 | 6.2e-05 | 0.001141 | 0.25785 | 1.3535 | 2.5481 | 99.38 | 20.93 | 37.42 | 67.9533 | 107.8459 | 65.2309 | 0.009335 | 1.0433 | 0.024878 | 23.8443 | -0.021766 | 0.14263 | -0.001995 | -0.1794 | -0.1794 | -0.032676
2019-03-21 12:31 | 1553185858.3947 | -249 | 0.20947 | 0.001552 | 0.000484 | 100.8327 | 100.8635 | 123.7001 | 21.2095 | 0.001934 | 0.00184 | -152.1065 | 0.24042 | 8.3998 | 0.003558 | 0.004609 | 0.004436 | 6.1e-05 | 0.001145 | 0.29391 | 1.3529 | 2.548 | 99.38 | 20.93 | 37.419 | 68.3046 | 107.4503 | 66.0578 | 0.009335 | 1.0408 | 0.02482 | 23.8468 | -0.002734 | 0.14303 | -0.001996 | -0.17902 | -0.17902 | -0.032836
2019-03-21 12:31 | 1553185859.3949 | -248 | 0.20948 | 0.001554 | 0.000484 | 100.816 | 100.9221 | 123.5912 | 21.2105 | 0.001964 | 0.001815 | -152.1243 | 0.24024 | 8.4038 | 0.003576 | 0.005044 | 0.004797 | 6.1e-05 | 0.001146 | 0.28727 | 1.315 | 2.548 | 99.38 | 20.94 | 37.415 | 69.0255 | 107.8924 | 65.4516 | 0.009339 | 1.0427 | 0.024831 | 23.8149 | -0.009408 | 0.14292 | -0.001998 | -0.17925 | -0.17925 | -0.03277
2019-03-21 12:31 | 1553185860.3944 | -247 | 0.20944 | 0.001555 | 0.000484 | 100.8658 | 101.1568 | 123.4934 | 21.2108 | 0.001962 | 0.001812 | -152.1296 | 0.23951 | 8.4022 | 0.003602 | 0.005038 | 0.00488 | 6.4e-05 | 0.001145 | 0.26985 | 1.3543 | 2.5481 | 99.38 | 20.94 | 37.413 | 68.9751 | 108.526 | 65.3769 | 0.009339 | 1.0429 | 0.024824 | 23.8029 | -0.025768 | 0.14249 | -0.002 | -0.17932 | -0.17932 | -0.032658
2019-03-21 12:31 | 1553185861.3945 | -246 | 0.20945 | 0.001549 | 0.000484 | 100.8282 | 100.9398 | 123.2849 | 21.2113 | 0.001925 | 0.001778 | -152.1339 | 0.24061 | 8.4023 | 0.003564 | 0.005677 | 0.005548 | 6.2e-05 | 0.001145 | 0.25442 | 1.3048 | 2.548 | 99.38 | 20.94 | 37.412 | 68.1003 | 107.6064 | 64.5586 | 0.009339 | 1.0454 | 0.024833 | 23.754 | -0.009893 | 0.14314 | -0.001991 | -0.17874 | -0.17874 | -0.032913
2019-03-21 12:31 | 1553185862.3938 | -245 | 0.20945 | 0.001547 | 0.000484 | 100.8461 | 101.0267 | 123.2122 | 21.2102 | 0.001942 | 0.001813 | -152.1218 | 0.24026 | 8.4041 | 0.003576 | 0.004678 | 0.004528 | 6.4e-05 | 0.001143 | 0.26864 | 1.3258 | 2.548 | 99.38 | 20.94 | 37.408 | 68.5082 | 107.8928 | 65.4061 | 0.009338 | 1.0428 | 0.024794 | 23.7768 | -0.016 | 0.14293 | -0.001989 | -0.17834 | -0.17834 | -0.03294
2019-03-21 12:31 | 1553185863.3946 | -244 | 0.20946 | 0.001548 | 0.000487 | 100.8735 | 101.2286 | 123.7119 | 21.2104 | 0.00198 | 0.001803 | -152.101 | 0.23955 | 8.4018 | 0.00359 | 0.00525 | 0.00506 | 6.7e-05 | 0.001143 | 0.27527 | 1.2977 | 2.5479 | 99.38 | 20.94 | 37.406 | 69.4261 | 108.2363 | 65.158 | 0.009337 | 1.0436 | 0.024854 | 23.8162 | -0.031433 | 0.14251 | -0.001995 | -0.1794 | -0.1794 | -0.032648
2019-03-21 12:31 | 1553185864.3936 | -243 | 0.20943 | 0.001552 | 0.00054 | 100.8425 | 100.9708 | 123.9799 | 21.2112 | 0.001962 | 0.001796 | -152.0908 | 0.23939 | 8.3997 | 0.003574 | 0.004962 | 0.004805 | 6.5e-05 | 0.001142 | 0.27831 | 1.2824 | 2.548 | 99.38 | 20.94 | 37.405 | 68.9957 | 107.8496 | 65.001 | 0.009337 | 1.044 | 0.024886 | 23.8365 | -0.011371 | 0.14241 | -0.002067 | -0.18826 | -0.18826 | -0.03109
2019-03-21 12:31 | 1553185865.3946 | -242 | 0.20944 | 0.001728 | 0.000659 | 100.8133 | 101.104 | 124.3799 | 21.2107 | 0.001928 | 0.00184 | -152.0969 | 0.23982 | 8.4008 | 0.003526 | 0.004875 | 0.004619 | 6.5e-05 | 0.001143 | 0.24978 | 1.2977 | 2.5478 | 99.38 | 20.94 | 37.404 | 68.1647 | 106.6625 | 66.0558 | 0.009337 | 1.0408 | 0.024888 | 23.9121 | -0.02575 | 0.14267 | -0.002441 | -0.22369 | -0.22369 | -0.026213
2019-03-21 12:31 | 1553185866.3941 | -241 | 0.20945 | 0.003441 | 0.000751 | 100.8607 | 100.998 | 124.5475 | 21.2113 | 0.001933 | 0.00181 | -152.0859 | 0.2406 | 8.403 | 0.003537 | 0.005017 | 0.004883 | 5.9e-05 | 0.001141 | 0.2995 | 1.2998 | 2.548 | 99.38 | 20.94 | 37.403 | 68.2939 | 106.9424 | 65.3259 | 0.009336 | 1.043 | 0.024931 | 23.9025 | -0.012172 | 0.14314 | -0.004743 | -0.40243 | -0.40243 | -0.014618
2019-03-21 12:31 | 1553185867.3938 | -240 | 0.20939 | 0.006185 | 0.000762 | 100.8404 | 101.1418 | 124.6557 | 21.2116 | 0.001933 | 0.001816 | -152.0942 | 0.23988 | 8.4029 | 0.003596 | 0.004794 | 0.004584 | 6.3e-05 | 0.001144 | 0.29192 | 1.3038 | 2.5481 | 99.38 | 20.95 | 37.401 | 68.2903 | 108.3621 | 65.4659 | 0.009342 | 1.0426 | 0.024937 | 23.9178 | -0.026689 | 0.14271 | -0.008274 | -0.66802 | -0.66802 | -0.00878
2019-03-21 12:31 | 1553185868.3942 | -239 | 0.20933 | 0.009183 | 0.000681 | 100.8869 | 100.873 | 124.8012 | 21.2123 | 0.001965 | 0.001812 | -152.0903 | 0.24025 | 8.402 | 0.003573 | 0.004483 | 0.004399 | 6.1e-05 | 0.001141 | 0.26106 | 1.3048 | 2.548 | 99.38 | 20.95 | 37.4 | 69.0662 | 107.812 | 65.3901 | 0.009341 | 1.0428 | 0.024954 | 23.9291 | 0.001229 | 0.14292 | -0.012041 | -0.94604 | -0.94604 | -0.006209
2019-03-21 12:31 | 1553185869.3939 | -238 | 0.20933 | 0.012291 | 0.000518 | 100.7502 | 100.9132 | 124.8836 | 21.2121 | 0.001917 | 0.001823 | -152.0833 | 0.24057 | 8.4041 | 0.003502 | 0.00494 | 0.004749 | 6.1e-05 | 0.001142 | 0.27671 | 1.3106 | 2.5481 | 99.38 | 20.95 | 37.397 | 67.8953 | 106.0813 | 65.6482 | 0.009341 | 1.0421 | 0.024953 | 23.9461 | -0.014467 | 0.14311 | -0.015872 | -1.2235 | -1.2235 | -0.004808
2019-03-21 12:31 | 1553185870.393 | -237 | 0.20933 | 0.014858 | 0.000358 | 100.7725 | 100.8996 | 124.7627 | 21.2131 | 0.001925 | 0.001814 | -152.0803 | 0.24082 | 8.4009 | 0.003447 | 0.004653 | 0.004548 | 6.1e-05 | 0.001138 | 0.27472 | 1.3371 | 2.548 | 99.38 | 20.95 | 37.395 | 68.1069 | 104.7559 | 65.426 | 0.00934 | 1.0427 | 0.024949 | 23.9266 | -0.011275 | 0.14326 | -0.019026 | -1.4502 | -1.4502 | -0.00406
2019-03-21 12:31 | 1553185871.3934 | -236 | 0.20947 | 0.015665 | 0.000232 | 100.7722 | 100.8547 | 124.5449 | 21.2125 | 0.001961 | 0.001836 | -152.0724 | 0.239 | 8.4 | 0.003412 | 0.005032 | 0.004776 | 5.8e-05 | 0.001139 | 0.25426 | 1.3738 | 2.548 | 99.38 | 20.95 | 37.393 | 68.9737 | 103.9217 | 65.9467 | 0.00934 | 1.0411 | 0.024908 | 23.9241 | -0.007316 | 0.14218 | -0.019924 | -1.5081 | -1.5081 | -0.003875
2019-03-21 12:31 | 1553185872.394 | -235 | 0.20955 | 0.01463 | 0.000166 | 100.833 | 101.0089 | 124.498 | 21.2123 | 0.001916 | 0.00181 | -152.0858 | 0.23982 | 8.4029 | 0.003486 | 0.004739 | 0.004644 | 6e-05 | 0.001135 | 0.27616 | 1.3094 | 2.5479 | 99.38 | 20.95 | 37.39 | 67.8907 | 105.7085 | 65.3379 | 0.009339 | 1.043 | 0.024926 | 23.8981 | -0.015595 | 0.14267 | -0.01848 | -1.3969 | -1.3969 | -0.004198
2019-03-21 12:31 | 1553185873.3945 | -234 | 0.20954 | 0.012086 | 0.000142 | 100.7946 | 100.8911 | 124.1748 | 21.2117 | 0.001887 | 0.001831 | -152.1022 | 0.24022 | 8.4022 | 0.003461 | 0.00477 | 0.004581 | 6e-05 | 0.001137 | 0.26034 | 1.2942 | 2.5478 | 99.38 | 20.95 | 37.389 | 67.1866 | 105.1005 | 65.8328 | 0.009339 | 1.0415 | 0.024875 | 23.8845 | -0.008564 | 0.14291 | -0.015118 | -1.1404 | -1.1404 | -0.00515
2019-03-21 12:31 | 1553185874.3931 | -233 | 0.20957 | 0.009181 | 0.000119 | 100.7219 | 100.8676 | 123.9302 | 21.2127 | 0.001935 | 0.001766 | -152.1034 | 0.23995 | 8.4038 | 0.003474 | 0.005351 | 0.005135 | 5.9e-05 | 0.001137 | 0.26089 | 1.2501 | 2.548 | 99.38 | 20.95 | 37.386 | 68.3366 | 105.4114 | 64.2664 | 0.009338 | 1.0463 | 0.024908 | 23.8058 | -0.012934 | 0.14275 | -0.011313 | -0.85511 | -0.85511 | -0.006861
2019-03-21 12:31 | 1553185875.3942 | -232 | 0.20957 | 0.006551 | 9.9e-05 | 100.8398 | 101.121 | 123.6882 | 21.2129 | 0.001904 | 0.001824 | -152.092 | 0.23861 | 8.4037 | 0.003496 | 0.00504 | 0.004886 | 5.9e-05 | 0.001137 | 0.24563 | 1.2875 | 2.5479 | 99.38 | 20.95 | 37.383 | 67.5973 | 105.9357 | 65.663 | 0.009337 | 1.042 | 0.024833 | 23.8317 | -0.024904 | 0.14195 | -0.007894 | -0.59563 | -0.59563 | -0.009795
2019-03-21 12:31 | 1553185876.3946 | -231 | 0.20957 | 0.004501 | 8.2e-05 | 100.8041 | 100.9688 | 123.2042 | 21.213 | 0.001969 | 0.001839 | -152.0967 | 0.23994 | 8.3998 | 0.00346 | 0.005256 | 0.005085 | 5.8e-05 | 0.001139 | 0.26648 | 1.3287 | 2.5479 | 99.38 | 20.95 | 37.38 | 69.1667 | 105.0807 | 66.033 | 0.009336 | 1.0409 | 0.024771 | 23.798 | -0.0146 | 0.14274 | -0.005243 | -0.3955 | -0.3955 | -0.014833
2019-03-21 12:31 | 1553185877.3943 | -230 | 0.20955 | 0.003041 | 6.7e-05 | 100.8044 | 101.0587 | 122.9561 | 21.2135 | 0.001934 | 0.001845 | -152.1108 | 0.23968 | 8.4007 | 0.003504 | 0.00479 | 0.004628 | 5.6e-05 | 0.001139 | 0.27113 | 1.2966 | 2.5479 | 99.38 | 20.95 | 37.375 | 68.31 | 106.1315 | 66.1733 | 0.009335 | 1.0404 | 0.024741 | 23.779 | -0.022535 | 0.14259 | -0.00336 | -0.25396 | -0.25396 | -0.023076
2019-03-21 12:31 | 1553185878.3943 | -229 | 0.20956 | 0.002071 | 5.4e-05 | 100.7805 | 101.0273 | 122.8209 | 21.2125 | 0.001974 | 0.001802 | -152.1128 | 0.2388 | 8.4032 | 0.003504 | 0.004936 | 0.004759 | 5.3e-05 | 0.001136 | 0.24729 | 1.2917 | 2.548 | 99.38 | 20.95 | 37.372 | 69.2808 | 106.147 | 65.1475 | 0.009334 | 1.0436 | 0.024764 | 23.7299 | -0.021876 | 0.14206 | -0.00211 | -0.16036 | -0.16036 | -0.03641
2019-03-21 12:31 | 1553185879.3937 | -228 | 0.20958 | 0.001628 | 4.4e-05 | 100.8047 | 100.7763 | 122.7838 | 21.2141 | 0.001949 | 0.0018 | -152.1053 | 0.24072 | 8.3999 | 0.003405 | 0.004789 | 0.00453 | 5.7e-05 | 0.001136 | 0.26576 | 1.3226 | 2.5481 | 99.38 | 20.95 | 37.365 | 68.674 | 103.73 | 65.0844 | 0.009333 | 1.0438 | 0.024763 | 23.7242 | 0.002517 | 0.1432 | -0.001534 | -0.11689 | -0.11689 | -0.050355
2019-03-21 12:31 | 1553185880.3934 | -227 | 0.20957 | 0.001422 | 3.6e-05 | 100.7817 | 100.9255 | 122.7047 | 21.2141 | 0.001946 | 0.001801 | -152.1099 | 0.24071 | 8.4027 | 0.003428 | 0.005045 | 0.004958 | 5.6e-05 | 0.001136 | 0.20742 | 1.3143 | 2.5478 | 99.38 | 20.95 | 37.362 | 68.6142 | 104.297 | 65.107 | 0.009332 | 1.0437 | 0.024754 | 23.7173 | -0.01275 | 0.1432 | -0.001262 | -0.096111 | -0.096111 | -0.061235
2019-03-21 12:31 | 1553185881.3945 | -226 | 0.20957 | 0.001274 | 2.9e-05 | 100.7114 | 100.8591 | 122.7199 | 21.2131 | 0.001981 | 0.001821 | -152.1198 | 0.23982 | 8.4029 | 0.003442 | 0.00448 | 0.004299 | 5.4e-05 | 0.001138 | 0.25371 | 1.3382 | 2.548 | 99.38 | 20.95 | 37.359 | 69.4367 | 104.6341 | 65.6031 | 0.009331 | 1.0422 | 0.024738 | 23.7362 | -0.013105 | 0.14267 | -0.001065 | -0.081005 | -0.081005 | -0.072388
2019-03-21 12:31 | 1553185882.3935 | -225 | 0.2096 | 0.001146 | 2.3e-05 | 100.7144 | 100.9994 | 122.3632 | 21.2145 | 0.001964 | 0.001824 | -152.1226 | 0.24016 | 8.4039 | 0.003402 | 0.004564 | 0.00441 | 5.7e-05 | 0.001142 | 0.24779 | 1.304 | 2.5481 | 99.38 | 20.95 | 37.353 | 69.0419 | 103.6613 | 65.6701 | 0.00933 | 1.042 | 0.024699 | 23.704 | -0.025276 | 0.14287 | -0.000895 | -0.067881 | -0.067881 | -0.086505
2019-03-21 12:31 | 1553185883.395 | -224 | 0.20957 | 0.001031 | 1.9e-05 | 100.6685 | 100.8748 | 122.3988 | 21.2144 | 0.001934 | 0.001827 | -152.1148 | 0.2409 | 8.4004 | 0.003442 | 0.004311 | 0.004106 | 5.8e-05 | 0.001142 | 0.29192 | 1.3369 | 2.5481 | 99.38 | 20.96 | 37.35 | 68.325 | 104.6368 | 65.7503 | 0.009335 | 1.0417 | 0.0247 | 23.7102 | -0.018305 | 0.14331 | -0.000745 | -0.056435 | -0.056435 | -0.10437
2019-03-21 12:31 | 1553185884.3931 | -223 | 0.20948 | 0.000939 | 1.4e-05 | 100.7254 | 100.8342 | 122.424 | 21.2144 | 0.001949 | 0.001825 | -152.1155 | 0.23992 | 8.403 | 0.003429 | 0.004536 | 0.004344 | 5.4e-05 | 0.001141 | 0.27776 | 1.3172 | 2.5481 | 99.38 | 20.96 | 37.347 | 68.6784 | 104.3285 | 65.6943 | 0.009334 | 1.0419 | 0.024704 | 23.7107 | -0.009656 | 0.14273 | -0.000622 | -0.047019 | -0.047019 | -0.12476
2019-03-21 12:31 | 1553185885.3938 | -222 | 0.20948 | 0.000863 | 1.2e-05 | 100.6578 | 100.7681 | 122.1406 | 21.2138 | 0.001944 | 0.001851 | -152.1155 | 0.24012 | 8.4028 | 0.003418 | 0.004678 | 0.004556 | 5.6e-05 | 0.001147 | 0.2896 | 1.3206 | 2.548 | 99.38 | 20.96 | 37.339 | 68.5639 | 104.0522 | 66.3179 | 0.009332 | 1.04 | 0.024653 | 23.705 | -0.009801 | 0.14284 | -0.000522 | -0.039341 | -0.039341 | -0.14923
2019-03-21 12:31 | 1553185886.3928 | -221 | 0.20946 | 0.00079 | 9e-06 | 100.6301 | 100.712 | 122.323 | 21.2156 | 0.002064 | 0.001839 | -152.1243 | 0.24045 | 8.4024 | 0.003484 | 0.005362 | 0.005152 | 5.5e-05 | 0.001145 | 0.26737 | 1.2974 | 2.5481 | 99.38 | 20.96 | 37.335 | 71.4506 | 105.6567 | 66.0418 | 0.009331 | 1.0408 | 0.024682 | 23.7131 | -0.007274 | 0.14305 | -0.000427 | -0.032218 | -0.032218 | -0.18248
2019-03-21 12:31 | 1553185887.3928 | -220 | 0.20946 | 0.000736 | 7e-06 | 100.6521 | 100.8267 | 122.3411 | 21.2157 | 0.001964 | 0.001893 | -152.121 | 0.24019 | 8.4038 | 0.003437 | 0.005132 | 0.005038 | 5.7e-05 | 0.001148 | 0.31741 | 1.3563 | 2.5481 | 99.38 | 20.96 | 37.329 | 69.0431 | 104.5084 | 67.3364 | 0.009329 | 1.0369 | 0.024636 | 23.7601 | -0.015502 | 0.14289 | -0.000356 | -0.026764 | -0.026764 | -0.21942
2019-03-21 12:31 | 1553185888.3942 | -219 | 0.20945 | 0.000693 | 5e-06 | 100.6581 | 100.7068 | 122.2369 | 21.2155 | 0.002005 | 0.001828 | -152.1183 | 0.24034 | 8.4059 | 0.003551 | 0.005033 | 0.004871 | 5.5e-05 | 0.001144 | 0.28047 | 1.2801 | 2.5478 | 99.38 | 20.96 | 37.326 | 70.0183 | 107.2916 | 65.7764 | 0.009329 | 1.0417 | 0.024683 | 23.6955 | -0.004331 | 0.14298 | -0.000299 | -0.022484 | -0.022484 | -0.26137
2019-03-21 12:31 | 1553185889.3939 | -218 | 0.20944 | 0.000654 | 4e-06 | 100.6063 | 100.7096 | 122.1141 | 21.2152 | 0.00191 | 0.001838 | -152.116 | 0.24028 | 8.4013 | 0.003433 | 0.004884 | 0.004737 | 5.4e-05 | 0.001139 | 0.24961 | 1.3106 | 2.548 | 99.38 | 20.96 | 37.323 | 67.7499 | 104.4206 | 66.018 | 0.009328 | 1.0409 | 0.024661 | 23.692 | -0.009181 | 0.14295 | -0.000247 | -0.018551 | -0.018551 | -0.3167
2019-03-21 12:31 | 1553185890.394 | -217 | 0.20945 | 0.000622 | 2e-06 | 100.6638 | 100.9272 | 122.1111 | 21.2155 | 0.001957 | 0.001883 | -152.1235 | 0.24019 | 8.4065 | 0.003474 | 0.005057 | 0.004856 | 5.4e-05 | 0.001138 | 0.24441 | 1.3174 | 2.5479 | 99.38 | 20.96 | 37.318 | 68.8657 | 105.4267 | 67.0936 | 0.009327 | 1.0376 | 0.024622 | 23.7292 | -0.023378 | 0.14289 | -0.000205 | -0.015311 | -0.015311 | -0.38357
2019-03-21 12:32 | 1553185891.3934 | -216 | 0.20947 | 0.000596 | 1e-06 | 100.561 | 100.7177 | 122.308 | 21.2167 | 0.001998 | 0.001828 | -152.1176 | 0.24111 | 8.4079 | 0.003409 | 0.004678 | 0.00455 | 5.7e-05 | 0.001142 | 0.25321 | 1.3015 | 2.5479 | 99.38 | 20.96 | 37.317 | 69.8533 | 103.8529 | 65.7643 | 0.009326 | 1.0417 | 0.02469 | 23.7019 | -0.013934 | 0.14344 | -0.000171 | -0.012766 | -0.012766 | -0.4618
2019-03-21 12:32 | 1553185892.3942 | -215 | 0.20949 | 0.000576 | 1e-06 | 100.5604 | 100.7457 | 122.5775 | 21.2167 | 0.001946 | 0.001865 | -152.095 | 0.24099 | 8.4007 | 0.003378 | 0.004992 | 0.004783 | 5.6e-05 | 0.001142 | 0.28998 | 1.2845 | 2.5479 | 99.38 | 20.96 | 37.313 | 68.5986 | 103.0913 | 66.6663 | 0.009325 | 1.0389 | 0.024685 | 23.7596 | -0.016467 | 0.14336 | -0.000145 | -0.01082 | -0.01082 | -0.54457
2019-03-21 12:32 | 1553185893.3932 | -214 | 0.20947 | 0.000559 | 0 | 100.5473 | 100.5386 | 122.6897 | 21.2158 | 0.001968 | 0.001849 | -152.1089 | 0.24051 | 8.4018 | 0.00341 | 0.004614 | 0.00452 | 5.5e-05 | 0.00114 | 0.28224 | 1.2817 | 2.5479 | 99.38 | 20.96 | 37.312 | 69.1359 | 103.8709 | 66.262 | 0.009325 | 1.0402 | 0.024711 | 23.7563 | 0.000771 | 0.14308 | -0.000121 | -0.008991 | -0.008991 | -0.65402
2019-03-21 12:32 | 1553185894.3932 | -213 | 0.20946 | 0.000549 | -2e-06 | 100.4856 | 100.6374 | 122.9511 | 21.2168 | 0.001984 | 0.001842 | -152.1207 | 0.24019 | 8.4003 | 0.003382 | 0.00461 | 0.004477 | 5.5e-05 | 0.001129 | 0.28351 | 1.2797 | 2.5479 | 99.38 | 20.96 | 37.312 | 69.5119 | 103.1904 | 66.0982 | 0.009325 | 1.0407 | 0.024743 | 23.7759 | -0.0135 | 0.14289 | -0.000108 | -0.007954 | -0.007954 | -0.73831
2019-03-21 12:32 | 1553185895.3946 | -212 | 0.20947 | 0.000531 | -1e-06 | 100.4609 | 100.6561 | 123.2068 | 21.2161 | 0.001991 | 0.001842 | -152.1292 | 0.24043 | 8.4066 | 0.003421 | 0.004772 | 0.004617 | 5.7e-05 | 0.001124 | 0.24762 | 1.3294 | 2.548 | 99.38 | 20.96 | 37.31 | 69.679 | 104.1273 | 66.1086 | 0.009325 | 1.0406 | 0.024768 | 23.8009 | -0.017366 | 0.14303 | -8.6e-05 | -0.00633 | -0.00633 | -0.92876
2019-03-21 12:32 | 1553185896.3943 | -211 | 0.20946 | 0.000517 | -2e-06 | 100.4877 | 100.5989 | 123.522 | 21.2169 | 0.001955 | 0.001845 | -152.1249 | 0.23918 | 8.4028 | 0.003413 | 0.004483 | 0.004338 | 5.8e-05 | 0.001124 | 0.26985 | 1.3133 | 2.5477 | 99.38 | 20.96 | 37.31 | 68.8275 | 103.9396 | 66.1891 | 0.009325 | 1.0404 | 0.024797 | 23.8342 | -0.009893 | 0.14229 | -6.7e-05 | -0.004889 | -0.004889 | -1.1961
2019-03-21 12:32 | 1553185897.3934 | -210 | 0.20947 | 0.000511 | -2e-06 | 100.4638 | 100.7181 | 123.9217 | 21.2168 | 0.00194 | 0.001794 | -152.1025 | 0.23981 | 8.4039 | 0.003394 | 0.004598 | 0.004508 | 5.9e-05 | 0.001122 | 0.26178 | 1.335 | 2.548 | 99.38 | 20.96 | 37.309 | 68.474 | 103.4829 | 64.9417 | 0.009324 | 1.0442 | 0.024883 | 23.8288 | -0.022608 | 0.14266 | -6e-05 | -0.004342 | -0.004342 | -1.3504
2019-03-21 12:32 | 1553185898.3938 | -209 | 0.20946 | 0.000504 | -2e-06 | 100.4147 | 100.4742 | 124.1036 | 21.2164 | 0.001958 | 0.001843 | -152.1091 | 0.23901 | 8.4044 | 0.003444 | 0.004295 | 0.004133 | 5.6e-05 | 0.001124 | 0.25011 | 1.298 | 2.548 | 99.38 | 20.96 | 37.309 | 68.8979 | 104.6979 | 66.1246 | 0.009324 | 1.0406 | 0.024858 | 23.888 | -0.005303 | 0.14219 | -5e-05 | -0.003612 | -0.003612 | -1.6179
2019-03-21 12:32 | 1553185899.3938 | -208 | 0.20947 | 0.000501 | -3e-06 | 100.3438 | 100.4293 | 124.3142 | 21.2173 | 0.001954 | 0.001801 | -152.139 | 0.24056 | 8.4051 | 0.003472 | 0.00588 | 0.005672 | 5.7e-05 | 0.00112 | 0.28191 | 1.3389 | 2.548 | 99.38 | 20.96 | 37.309 | 68.8098 | 105.3672 | 65.1123 | 0.009324 | 1.0437 | 0.024916 | 23.8725 | -0.007618 | 0.14311 | -4.6e-05 | -0.003291 | -0.003291 | -1.7874
2019-03-21 12:32 | 1553185900.3935 | -207 | 0.20948 | 0.000495 | -3e-06 | 100.3652 | 100.4613 | 124.7151 | 21.217 | 0.001942 | 0.001826 | -152.1318 | 0.23987 | 8.4016 | 0.003519 | 0.004749 | 0.004581 | 5.8e-05 | 0.001119 | 0.3063 | 1.338 | 2.5481 | 99.38 | 20.96 | 37.31 | 68.5179 | 106.5063 | 65.717 | 0.009325 | 1.0418 | 0.024934 | 23.9324 | -0.008557 | 0.1427 | -3.7e-05 | -0.002634 | -0.002634 | -2.2265
2019-03-21 12:32 | 1553185901.393 | -206 | 0.20947 | 0.000493 | -3e-06 | 100.3798 | 100.59 | 124.6855 | 21.2167 | 0.001973 | 0.00181 | -152.1457 | 0.24007 | 8.4012 | 0.00349 | 0.004689 | 0.004553 | 5.8e-05 | 0.001119 | 0.23308 | 1.3127 | 2.5479 | 99.38 | 20.96 | 37.311 | 69.2578 | 105.8128 | 65.3405 | 0.009325 | 1.043 | 0.024945 | 23.9162 | -0.018711 | 0.14282 | -3.6e-05 | -0.002522 | -0.002522 | -2.3278
2019-03-21 12:32 | 1553185902.3926 | -205 | 0.20944 | 0.000484 | -4e-06 | 100.3223 | 100.5417 | 124.6757 | 21.2157 | 0.001974 | 0.001851 | -152.1811 | 0.24041 | 8.404 | 0.003531 | 0.005057 | 0.004906 | 5.6e-05 | 0.001121 | 0.22052 | 1.2976 | 2.548 | 99.38 | 20.96 | 37.312 | 69.2832 | 106.7947 | 66.3266 | 0.009325 | 1.04 | 0.024907 | 23.9501 | -0.019535 | 0.14302 | -2.2e-05 | -0.001456 | -0.001456 | -4.0362
2019-03-21 12:32 | 1553185903.3927 | -204 | 0.20946 | 0.000483 | -4e-06 | 100.387 | 100.2797 | 124.7383 | 21.2169 | 0.001939 | 0.001771 | -152.1742 | 0.23954 | 8.4009 | 0.003449 | 0.004692 | 0.004463 | 5.8e-05 | 0.001124 | 0.25249 | 1.2814 | 2.548 | 99.38 | 20.96 | 37.312 | 68.4428 | 104.8099 | 64.4024 | 0.009325 | 1.0459 | 0.024985 | 23.8881 | 0.009558 | 0.1425 | -2.2e-05 | -0.001406 | -0.001406 | -4.1648
2019-03-21 12:32 | 1553185904.3929 | -203 | 0.20945 | 0.000475 | -3e-06 | 100.2743 | 100.4936 | 124.5677 | 21.2168 | 0.00197 | 0.001796 | -152.1723 | 0.24016 | 8.403 | 0.003506 | 0.004031 | 0.003798 | 6.1e-05 | 0.001123 | 0.28136 | 1.3092 | 2.548 | 99.38 | 20.96 | 37.312 | 69.1752 | 106.1884 | 65.0043 | 0.009325 | 1.044 | 0.024945 | 23.893 | -0.01954 | 0.14287 | -1.3e-05 | -0.00074 | -0.00074 | -7.9324
2019-03-21 12:32 | 1553185905.3942 | -202 | 0.20944 | 0.000478 | -3e-06 | 100.3622 | 100.6506 | 124.3571 | 21.2178 | 0.002008 | 0.001798 | -152.1724 | 0.24033 | 8.4022 | 0.003491 | 0.005299 | 0.005117 | 6.3e-05 | 0.001126 | 0.25697 | 1.3175 | 2.5481 | 99.38 | 20.96 | 37.311 | 70.101 | 105.8359 | 65.0442 | 0.009325 | 1.0439 | 0.024923 | 23.8742 | -0.025667 | 0.14297 | -1.6e-05 | -0.00103 | -0.00103 | -5.707
2019-03-21 12:32 | 1553185906.3942 | -201 | 0.20946 | 0.00048 | -3e-06 | 100.3146 | 100.4058 | 124.2402 | 21.2173 | 0.0019 | 0.001816 | -152.1623 | 0.23995 | 8.3987 | 0.003563 | 0.004297 | 0.00418 | 6.1e-05 | 0.001128 | 0.25929 | 1.2713 | 2.5478 | 99.38 | 20.96 | 37.311 | 67.5002 | 107.5727 | 65.4714 | 0.009325 | 1.0426 | 0.024895 | 23.8781 | -0.008128 | 0.14274 | -1.8e-05 | -0.001191 | -0.001191 | -4.926
2019-03-21 12:32 | 1553185907.3929 | -200 | 0.20944 | 0.000478 | -4e-06 | 100.3131 | 100.4572 | 124.0473 | 21.2177 | 0.001951 | 0.001805 | -152.1724 | 0.24049 | 8.4033 | 0.003523 | 0.004993 | 0.004788 | 5.9e-05 | 0.001132 | 0.25802 | 1.3193 | 2.5481 | 99.38 | 20.96 | 37.311 | 68.7251 | 106.6111 | 65.219 | 0.009325 | 1.0434 | 0.024885 | 23.8507 | -0.012841 | 0.14306 | -1.4e-05 | -0.000814 | -0.000814 | -7.2261
2019-03-21 12:32 | 1553185908.3939 | -199 | 0.20942 | 0.000475 | -4e-06 | 100.3414 | 100.3608 | 123.9079 | 21.2187 | 0.001928 | 0.00184 | -152.1498 | 0.24013 | 8.4022 | 0.003466 | 0.004731 | 0.004585 | 5.9e-05 | 0.00113 | 0.24994 | 1.3245 | 2.548 | 99.38 | 20.96 | 37.31 | 68.1739 | 105.2293 | 66.0503 | 0.009325 | 1.0408 | 0.024841 | 23.8665 | -0.001732 | 0.14286 | -1.1e-05 | -0.000561 | -0.000561 | -10.4652
2019-03-21 12:32 | 1553185909.394 | -198 | 0.20943 | 0.000471 | -4e-06 | 100.2848 | 100.3615 | 123.8762 | 21.2172 | 0.001901 | 0.001826 | -152.1083 | 0.24092 | 8.4045 | 0.003438 | 0.003718 | 0.003542 | 5.8e-05 | 0.001128 | 0.28694 | 1.3954 | 2.548 | 99.38 | 20.96 | 37.308 | 67.5322 | 104.5372 | 65.7124 | 0.009324 | 1.0419 | 0.02485 | 23.8516 | -0.00684 | 0.14333 | -7e-06 | -0.000304 | -0.000304 | -19.3767
2019-03-21 12:32 | 1553185910.394 | -197 | 0.20943 | 0.000469 | -4e-06 | 100.3184 | 100.3445 | 123.8055 | 21.2181 | 0.001948 | 0.001809 | -152.1058 | 0.23935 | 8.4025 | 0.003438 | 0.004854 | 0.004693 | 5.9e-05 | 0.001129 | 0.27383 | 1.3503 | 2.5481 | 99.38 | 20.96 | 37.307 | 68.6656 | 104.5406 | 65.3074 | 0.009324 | 1.0431 | 0.024858 | 23.8305 | -0.00232 | 0.14239 | -4e-06 | -7.8e-05 | -7.8e-05 | -75.2941
2019-03-21 12:32 | 1553185911.3946 | -196 | 0.20942 | 0.000474 | -4e-06 | 100.3393 | 100.5175 | 123.9213 | 21.2179 | 0.00199 | 0.001832 | -152.0852 | 0.239 | 8.4013 | 0.003479 | 0.004534 | 0.004336 | 5.9e-05 | 0.001132 | 0.23684 | 1.3182 | 2.5477 | 99.38 | 20.96 | 37.306 | 69.6682 | 105.5384 | 65.8552 | 0.009324 | 1.0414 | 0.024849 | 23.8609 | -0.01587 | 0.14218 | -9e-06 | -0.000475 | -0.000475 | -12.3023
2019-03-21 12:32 | 1553185912.3936 | -195 | 0.20944 | 0.000467 | -3e-06 | 100.3006 | 100.2784 | 123.9198 | 21.2173 | 0.001914 | 0.001797 | -152.0805 | 0.23955 | 8.4013 | 0.003474 | 0.004595 | 0.004432 | 5.6e-05 | 0.001129 | 0.27007 | 1.3152 | 2.548 | 99.38 | 20.96 | 37.304 | 67.8271 | 105.4098 | 65.0156 | 0.009323 | 1.044 | 0.02488 | 23.8312 | 0.001978 | 0.14251 | -2e-06 | 9e-05 | 9e-05 | 65.2211
2019-03-21 12:32 | 1553185913.3932 | -194 | 0.20945 | 0.00047 | -4e-06 | 100.2875 | 100.4807 | 123.7473 | 21.2177 | 0.00198 | 0.001794 | -152.0711 | 0.23978 | 8.4023 | 0.003476 | 0.004822 | 0.004583 | 5.8e-05 | 0.001132 | 0.28982 | 1.3282 | 2.548 | 99.38 | 20.95 | 37.303 | 69.4294 | 105.4747 | 64.9496 | 0.009317 | 1.0442 | 0.024865 | 23.8123 | -0.017218 | 0.14265 | -5e-06 | -0.000155 | -0.000155 | -37.757
2019-03-21 12:32 | 1553185914.3933 | -193 | 0.20942 | 0.000471 | -4e-06 | 100.291 | 100.3751 | 123.8907 | 21.2186 | 0.001941 | 0.001866 | -152.0698 | 0.23923 | 8.4021 | 0.003453 | 0.004924 | 0.004793 | 5.6e-05 | 0.00113 | 0.25929 | 1.294 | 2.548 | 99.38 | 20.95 | 37.302 | 68.4983 | 104.9005 | 66.69 | 0.009317 | 1.0389 | 0.024816 | 23.8873 | -0.007496 | 0.14232 | -6e-06 | -0.000183 | -0.000183 | -31.9195
2019-03-21 12:32 | 1553185915.3938 | -192 | 0.20944 | 0.000467 | -3e-06 | 100.2311 | 100.2685 | 124.3266 | 21.2189 | 0.001942 | 0.001846 | -152.0751 | 0.23887 | 8.4044 | 0.003364 | 0.005333 | 0.005145 | 5.8e-05 | 0.001132 | 0.27007 | 1.3052 | 2.548 | 99.38 | 20.95 | 37.3 | 68.5204 | 102.7493 | 66.2003 | 0.009316 | 1.0404 | 0.024877 | 23.9121 | -0.003334 | 0.1421 | -1e-06 | 0.000109 | 0.000109 | 53.4753
2019-03-21 12:32 | 1553185916.3934 | -191 | 0.20943 | 0.000464 | -4e-06 | 100.2431 | 100.2944 | 124.3107 | 21.2192 | 0.001926 | 0.00182 | -152.0929 | 0.23917 | 8.3998 | 0.003395 | 0.004359 | 0.004171 | 5.9e-05 | 0.001136 | 0.26913 | 1.2838 | 2.5479 | 99.38 | 20.95 | 37.3 | 68.1322 | 103.5069 | 65.5707 | 0.009316 | 1.0423 | 0.024899 | 23.8884 | -0.004579 | 0.14228 | 3e-06 | 0.000417 | 0.000417 | 14.0272
2019-03-21 12:32 | 1553185917.3925 | -190 | 0.20943 | 0.000469 | -4e-06 | 100.1814 | 100.4068 | 124.2651 | 21.2178 | 0.002011 | 0.001849 | -152.0957 | 0.23984 | 8.4007 | 0.003458 | 0.004669 | 0.004532 | 6e-05 | 0.001132 | 0.28495 | 1.2866 | 2.5479 | 99.38 | 20.96 | 37.3 | 70.1757 | 105.0209 | 66.2866 | 0.009322 | 1.0401 | 0.024868 | 23.9092 | -0.020102 | 0.14268 | -3e-06 | 1.7e-05 | 1.7e-05 | 355.0072
2019-03-21 12:32 | 1553185918.3936 | -189 | 0.20942 | 0.000467 | -4e-06 | 100.2243 | 100.3261 | 124.5482 | 21.2191 | 0.001948 | 0.001808 | -152.1108 | 0.24 | 8.4015 | 0.00346 | 0.005263 | 0.005018 | 5.6e-05 | 0.001135 | 0.24674 | 1.2958 | 2.548 | 99.38 | 20.96 | 37.3 | 68.6486 | 105.0787 | 65.2851 | 0.009322 | 1.0432 | 0.024933 | 23.9011 | -0.009079 | 0.14278 | -1e-06 | 0.000213 | 0.000213 | 27.5184
2019-03-21 12:32 | 1553185919.394 | -188 | 0.20945 | 0.000466 | -4e-06 | 100.2279 | 100.3966 | 124.4689 | 21.219 | 0.00197 | 0.001735 | -152.1246 | 0.23967 | 8.4036 | 0.003493 | 0.00494 | 0.004794 | 5.7e-05 | 0.001133 | 0.28119 | 1.3076 | 2.5481 | 99.38 | 20.95 | 37.3 | 69.1936 | 105.8771 | 63.526 | 0.009316 | 1.0486 | 0.02499 | 23.8313 | -0.015043 | 0.14258 | 1e-06 | 0.00031 | 0.00031 | 18.8948
2019-03-21 12:32 | 1553185920.3928 | -187 | 0.20942 | 0.000466 | -4e-06 | 100.1906 | 100.2266 | 124.1442 | 21.2184 | 0.001951 | 0.001789 | -152.118 | 0.23972 | 8.4019 | 0.003483 | 0.004954 | 0.004784 | 5.9e-05 | 0.001131 | 0.29927 | 1.3292 | 2.548 | 99.38 | 20.95 | 37.303 | 68.7243 | 105.6429 | 64.8333 | 0.009317 | 1.0446 | 0.024909 | 23.8464 | -0.003213 | 0.14261 | 0 | 0.000269 | 0.000269 | 21.7625
2019-03-21 12:32 | 1553185921.3937 | -186 | 0.20943 | 0.000465 | -4e-06 | 100.2204 | 100.3121 | 123.9522 | 21.219 | 0.001923 | 0.001873 | -152.1172 | 0.23978 | 8.404 | 0.003521 | 0.004954 | 0.004782 | 6.1e-05 | 0.001129 | 0.22876 | 1.3103 | 2.548 | 99.38 | 20.95 | 37.305 | 68.0518 | 106.5696 | 66.862 | 0.009318 | 1.0383 | 0.024815 | 23.8993 | -0.008179 | 0.14265 | 2e-06 | 0.000399 | 0.000399 | 14.6831
2019-03-21 12:32 | 1553185922.3942 | -185 | 0.20944 | 0.000466 | -4e-06 | 100.2341 | 100.3233 | 123.9278 | 21.2195 | 0.001915 | 0.001822 | -152.097 | 0.24 | 8.403 | 0.003551 | 0.004901 | 0.004735 | 6e-05 | 0.00113 | 0.25658 | 1.3328 | 2.548 | 99.38 | 20.95 | 37.308 | 67.8549 | 107.2931 | 65.6155 | 0.009318 | 1.0422 | 0.024859 | 23.8531 | -0.007958 | 0.14278 | 1e-06 | 0.000302 | 0.000302 | 19.447
2019-03-21 12:32 | 1553185923.3929 | -184 | 0.20943 | 0.000463 | -4e-06 | 100.2711 | 100.3877 | 123.8238 | 21.2186 | 0.001941 | 0.001783 | -152.0954 | 0.23995 | 8.402 | 0.003525 | 0.004845 | 0.004654 | 5.7e-05 | 0.001134 | 0.29983 | 1.3322 | 2.5479 | 99.38 | 20.95 | 37.309 | 68.4959 | 106.6537 | 64.7 | 0.009319 | 1.045 | 0.024882 | 23.8109 | -0.010399 | 0.14275 | 5.8e-05 | 0.004589 | 0.004589 | 1.2784
2019-03-21 12:32 | 1553185924.3929 | -183 | 0.20944 | 0.000466 | -4e-06 | 100.2669 | 100.0897 | 123.6246 | 21.2191 | 0.001952 | 0.001783 | -152.0829 | 0.23963 | 8.4037 | 0.003542 | 0.004957 | 0.004784 | 6e-05 | 0.001136 | 0.26056 | 1.3526 | 2.5479 | 99.38 | 20.95 | 37.309 | 68.7421 | 107.0706 | 64.6925 | 0.009319 | 1.045 | 0.024862 | 23.7914 | 0.015821 | 0.14255 | -4e-06 | -5.5e-05 | -5.5e-05 | -105.9905
2019-03-21 12:32 | 1553185925.3943 | -182 | 0.20944 | 0.000464 | -4e-06 | 100.2475 | 100.2276 | 123.3805 | 21.2191 | 0.001897 | 0.001838 | -152.0714 | 0.24013 | 8.4016 | 0.003489 | 0.004918 | 0.004817 | 6.2e-05 | 0.001137 | 0.28998 | 1.353 | 2.5479 | 99.38 | 20.95 | 37.308 | 67.4301 | 105.7853 | 66.0233 | 0.009318 | 1.0409 | 0.024789 | 23.8147 | 0.001775 | 0.14285 | -5.4e-05 | -0.003805 | -0.003805 | -1.543
2019-03-21 12:32 | 1553185926.393 | -181 | 0.20943 | 0.000464 | -4e-06 | 100.2863 | 100.3502 | 123.1741 | 21.2193 | 0.00193 | 0.001826 | -152.067 | 0.23962 | 8.4023 | 0.003475 | 0.004992 | 0.004785 | 6.2e-05 | 0.001135 | 0.2891 | 1.3193 | 2.5479 | 99.38 | 20.95 | 37.309 | 68.2134 | 105.4549 | 65.7154 | 0.009319 | 1.0418 | 0.024779 | 23.784 | -0.005705 | 0.14255 | 4.8e-05 | 0.00384 | 0.00384 | 1.5258
2019-03-21 12:32 | 1553185927.3945 | -180 | 0.20942 | 0.000463 | -3e-06 | 100.2097 | 100.3049 | 123.145 | 21.2201 | 0.001966 | 0.001814 | -152.0792 | 0.23979 | 8.4043 | 0.003511 | 0.005492 | 0.005316 | 6e-05 | 0.001137 | 0.23938 | 1.3279 | 2.548 | 99.38 | 20.95 | 37.308 | 69.0947 | 106.3259 | 65.4248 | 0.009318 | 1.0427 | 0.024787 | 23.771 | -0.008499 | 0.14265 | 6.5e-05 | 0.005083 | 0.005083 | 1.1534
2019-03-21 12:32 | 1553185928.3932 | -179 | 0.20942 | 0.000462 | -3e-06 | 100.2809 | 100.3332 | 122.7953 | 21.2193 | 0.001929 | 0.001813 | -152.0725 | 0.24023 | 8.4013 | 0.003454 | 0.004527 | 0.00439 | 6.1e-05 | 0.001135 | 0.25608 | 1.2706 | 2.5477 | 99.38 | 20.95 | 37.307 | 68.2001 | 104.9428 | 65.4219 | 0.009318 | 1.0427 | 0.024752 | 23.7371 | -0.004666 | 0.14291 | 6.7e-05 | 0.005211 | 0.005211 | 1.1273
2019-03-21 12:32 | 1553185929.3936 | -178 | 0.20943 | 0.000462 | -4e-06 | 100.3134 | 100.3373 | 122.981 | 21.2199 | 0.001936 | 0.001844 | -152.0788 | 0.23959 | 8.4022 | 0.003487 | 0.004508 | 0.004309 | 5.5e-05 | 0.001133 | 0.2625 | 1.2999 | 2.5481 | 99.38 | 20.95 | 37.304 | 68.3584 | 105.7387 | 66.1467 | 0.009317 | 1.0405 | 0.024744 | 23.7804 | -0.002134 | 0.14253 | 1.8e-05 | 0.001583 | 0.001583 | 3.7012
2019-03-21 12:32 | 1553185930.3947 | -177 | 0.20945 | 0.000457 | -4e-06 | 100.266 | 100.4139 | 122.9734 | 21.2195 | 0.001937 | 0.001824 | -152.0769 | 0.23997 | 8.4015 | 0.003453 | 0.0051 | 0.004913 | 5.6e-05 | 0.001133 | 0.29966 | 1.2511 | 2.548 | 99.38 | 20.95 | 37.299 | 68.3915 | 104.9038 | 65.6858 | 0.009316 | 1.0419 | 0.02476 | 23.7636 | -0.013188 | 0.14276 | -5.8e-05 | -0.004046 | -0.004046 | -1.4503
2019-03-21 12:32 | 1553185931.3923 | -176 | 0.20948 | 0.000461 | -4e-06 | 100.2559 | 100.4133 | 122.904 | 21.2199 | 0.001993 | 0.001878 | -152.082 | 0.24054 | 8.4025 | 0.003402 | 0.005274 | 0.005088 | 6.2e-05 | 0.001135 | 0.28368 | 1.2912 | 2.5479 | 99.38 | 20.95 | 37.297 | 69.7355 | 103.6743 | 66.9822 | 0.009316 | 1.038 | 0.024706 | 23.8023 | -0.014031 | 0.1431 | -0.000262 | -0.019181 | -0.019181 | -0.30662
2019-03-21 12:32 | 1553185932.3933 | -175 | 0.20944 | 0.000461 | -4e-06 | 100.2678 | 100.3097 | 122.7807 | 21.2203 | 0.001938 | 0.001828 | -152.1069 | 0.24019 | 8.4019 | 0.003405 | 0.005114 | 0.004937 | 5.7e-05 | 0.00113 | 0.24419 | 1.3179 | 2.5481 | 99.38 | 20.95 | 37.295 | 68.4281 | 103.7433 | 65.7693 | 0.009315 | 1.0417 | 0.024738 | 23.7479 | -0.00374 | 0.14289 | -1.6e-05 | -0.000947 | -0.000947 | -6.1985
2019-03-21 12:32 | 1553185933.3934 | -174 | 0.20946 | 0.000465 | -4e-06 | 100.2246 | 100.5372 | 122.6357 | 21.2205 | 0.001947 | 0.001852 | -152.1026 | 0.24022 | 8.402 | 0.003374 | 0.004397 | 0.004273 | 5.6e-05 | 0.001132 | 0.26344 | 1.2897 | 2.548 | 99.38 | 20.95 | 37.288 | 68.6388 | 102.996 | 66.3486 | 0.009313 | 1.0399 | 0.024702 | 23.7541 | -0.027859 | 0.14291 | -0.000131 | -0.009445 | -0.009445 | -0.62189
2019-03-21 12:32 | 1553185934.3929 | -173 | 0.20943 | 0.000462 | -3e-06 | 100.2818 | 100.4306 | 122.7087 | 21.2198 | 0.001961 | 0.00179 | -152.1067 | 0.24016 | 8.4042 | 0.003444 | 0.005541 | 0.005286 | 5.5e-05 | 0.001135 | 0.28766 | 1.324 | 2.5479 | 99.38 | 20.95 | 37.286 | 68.9783 | 104.6956 | 64.8503 | 0.009313 | 1.0445 | 0.024764 | 23.7087 | -0.013266 | 0.14287 | 4.8e-05 | 0.003811 | 0.003811 | 1.541
2019-03-21 12:32 | 1553185935.3936 | -172 | 0.20943 | 0.00046 | -3e-06 | 100.2597 | 100.3942 | 122.8744 | 21.2212 | 0.001898 | 0.001772 | -152.1013 | 0.2403 | 8.4036 | 0.003426 | 0.004164 | 0.004006 | 5.7e-05 | 0.001135 | 0.28102 | 1.3303 | 2.548 | 99.38 | 20.95 | 37.283 | 67.4555 | 104.2669 | 64.4239 | 0.009312 | 1.0458 | 0.024796 | 23.7097 | -0.011986 | 0.14296 | 1.5e-05 | 0.001326 | 0.001326 | 4.4327
2019-03-21 12:32 | 1553185936.3926 | -171 | 0.20942 | 0.000464 | -4e-06 | 100.2767 | 100.3206 | 123.1178 | 21.2221 | 0.001959 | 0.001832 | -152.0896 | 0.23969 | 8.4033 | 0.00343 | 0.004945 | 0.004729 | 6.1e-05 | 0.001137 | 0.28655 | 1.3387 | 2.5481 | 99.38 | 20.95 | 37.283 | 68.9135 | 104.3433 | 65.8725 | 0.009312 | 1.0414 | 0.024768 | 23.7841 | -0.003914 | 0.14259 | 0.000115 | 0.008809 | 0.008809 | 0.66531
2019-03-21 12:32 | 1553185937.394 | -170 | 0.20944 | 0.000465 | -4e-06 | 100.221 | 100.3598 | 123.1185 | 21.2211 | 0.001908 | 0.001829 | -152.0867 | 0.23988 | 8.4046 | 0.003453 | 0.005343 | 0.005216 | 5.9e-05 | 0.001137 | 0.25465 | 1.3422 | 2.5479 | 99.38 | 20.95 | 37.281 | 67.7059 | 104.91 | 65.8013 | 0.009312 | 1.0416 | 0.024771 | 23.7816 | -0.012377 | 0.14271 | -2.4e-05 | -0.001541 | -0.001541 | -3.8067
2019-03-21 12:32 | 1553185938.3926 | -169 | 0.20944 | 0.000461 | -4e-06 | 100.2946 | 100.228 | 123.3593 | 21.2217 | 0.001938 | 0.001805 | -152.0906 | 0.24007 | 8.402 | 0.003365 | 0.004479 | 0.004317 | 6e-05 | 0.001142 | 0.29247 | 1.3066 | 2.5479 | 99.38 | 20.95 | 37.28 | 68.4273 | 102.7693 | 65.2205 | 0.009311 | 1.0434 | 0.024816 | 23.7845 | 0.005944 | 0.14282 | -2.4e-05 | -0.001556 | -0.001556 | -3.7722
2019-03-21 12:32 | 1553185939.3927 | -168 | 0.20944 | 0.000461 | -3e-06 | 100.2317 | 100.3557 | 123.2927 | 21.2214 | 0.001906 | 0.001829 | -152.0876 | 0.23953 | 8.4038 | 0.003394 | 0.004841 | 0.004663 | 6e-05 | 0.001139 | 0.28512 | 1.2893 | 2.5479 | 99.38 | 20.95 | 37.28 | 67.6415 | 103.4912 | 65.8023 | 0.009311 | 1.0416 | 0.024788 | 23.7985 | -0.011055 | 0.1425 | -4.2e-05 | -0.002891 | -0.002891 | -2.0261
2019-03-21 12:32 | 1553185940.3927 | -167 | 0.20944 | 0.000464 | -4e-06 | 100.266 | 100.4047 | 123.5933 | 21.2214 | 0.00197 | 0.001873 | -152.0726 | 0.23897 | 8.4024 | 0.003368 | 0.004872 | 0.004716 | 6e-05 | 0.00114 | 0.27096 | 1.3031 | 2.548 | 99.38 | 20.95 | 37.281 | 69.1859 | 102.8499 | 66.8644 | 0.009312 | 1.0383 | 0.024779 | 23.8648 | -0.012368 | 0.14216 | -4.9e-05 | -0.003433 | -0.003433 | -1.7019
2019-03-21 12:32 | 1553185941.3927 | -166 | 0.20946 | 0.000465 | -4e-06 | 100.2627 | 100.325 | 123.8818 | 21.2213 | 0.001997 | 0.001851 | -152.0799 | 0.23913 | 8.4012 | 0.003383 | 0.005145 | 0.004893 | 6e-05 | 0.00114 | 0.30397 | 1.3171 | 2.548 | 99.38 | 20.95 | 37.282 | 69.8321 | 103.209 | 66.3228 | 0.009312 | 1.04 | 0.024828 | 23.8736 | -0.005558 | 0.14226 | -0.000126 | -0.009179 | -0.009179 | -0.63695
2019-03-21 12:32 | 1553185942.3939 | -165 | 0.20946 | 0.000463 | -3e-06 | 100.2457 | 100.3611 | 124.2529 | 21.2217 | 0.00193 | 0.001791 | -152.0606 | 0.23956 | 8.4013 | 0.003482 | 0.004444 | 0.004344 | 6.3e-05 | 0.001141 | 0.31045 | 1.2724 | 2.5479 | 99.38 | 20.95 | 37.283 | 68.2244 | 105.6108 | 64.8779 | 0.009312 | 1.0444 | 0.024918 | 23.8584 | -0.010291 | 0.14252 | -0.000137 | -0.010054 | -0.010054 | -0.58258
2019-03-21 12:32 | 1553185943.3943 | -164 | 0.20945 | 0.00046 | -3e-06 | 100.179 | 100.3513 | 124.5738 | 21.2224 | 0.001957 | 0.001822 | -152.0543 | 0.23907 | 8.4028 | 0.003438 | 0.004692 | 0.004571 | 6e-05 | 0.001142 | 0.29999 | 1.2905 | 2.5479 | 99.38 | 20.95 | 37.285 | 68.8779 | 104.5401 | 65.6356 | 0.009313 | 1.0421 | 0.024923 | 23.9159 | -0.015368 | 0.14222 | -6.9e-05 | -0.004932 | -0.004932 | -1.1852
2019-03-21 12:32 | 1553185944.3944 | -163 | 0.20944 | 0.000463 | -3e-06 | 100.2982 | 100.4848 | 124.7867 | 21.2223 | 0.001889 | 0.001821 | -152.0686 | 0.23904 | 8.4015 | 0.003474 | 0.004462 | 0.004234 | 5.8e-05 | 0.001142 | 0.30182 | 1.283 | 2.5481 | 99.38 | 20.95 | 37.285 | 67.2313 | 105.4186 | 65.6161 | 0.009313 | 1.0422 | 0.024945 | 23.9357 | -0.016626 | 0.14221 | -4.7e-05 | -0.003285 | -0.003285 | -1.7791
2019-03-21 12:32 | 1553185945.3931 | -162 | 0.20943 | 0.00046 | -4e-06 | 100.291 | 100.5345 | 124.6808 | 21.2227 | 0.002003 | 0.001835 | -152.0966 | 0.23928 | 8.4042 | 0.003472 | 0.004235 | 0.004073 | 5.9e-05 | 0.001142 | 0.24746 | 1.3026 | 2.5479 | 99.38 | 20.95 | 37.285 | 69.9762 | 105.3654 | 65.9433 | 0.009313 | 1.0411 | 0.024922 | 23.9371 | -0.021689 | 0.14235 | 6.7e-05 | 0.005219 | 0.005219 | 1.1209
2019-03-21 12:32 | 1553185946.3944 | -161 | 0.20942 | 0.000458 | -3e-06 | 100.2529 | 100.4838 | 124.8719 | 21.2223 | 0.001966 | 0.001816 | -152.106 | 0.23886 | 8.4016 | 0.003528 | 0.004861 | 0.004698 | 6e-05 | 0.001142 | 0.28672 | 1.2681 | 2.5481 | 99.38 | 20.96 | 37.285 | 69.0954 | 106.7437 | 65.485 | 0.009318 | 1.0426 | 0.024958 | 23.9392 | -0.020574 | 0.1421 | 0.00012 | 0.009196 | 0.009196 | 0.6351
2019-03-21 12:32 | 1553185947.3931 | -160 | 0.20942 | 0.000458 | -4e-06 | 100.2061 | 100.4405 | 124.7768 | 21.2224 | 0.001983 | 0.001852 | -152.0897 | 0.24002 | 8.4067 | 0.003533 | 0.004411 | 0.004171 | 6.1e-05 | 0.001144 | 0.26504 | 1.3094 | 2.548 | 99.38 | 20.96 | 37.284 | 69.5127 | 106.8557 | 66.3526 | 0.009318 | 1.0399 | 0.024917 | 23.9607 | -0.020898 | 0.14279 | 7.7e-05 | 0.006003 | 0.006003 | 0.97763
2019-03-21 12:32 | 1553185948.3932 | -159 | 0.20944 | 0.000459 | -4e-06 | 100.2121 | 100.3806 | 124.7219 | 21.222 | 0.002001 | 0.001874 | -152.0986 | 0.23984 | 8.4009 | 0.003495 | 0.004841 | 0.004831 | 6e-05 | 0.001142 | 0.28174 | 1.3557 | 2.5481 | 99.38 | 20.96 | 37.283 | 69.9359 | 105.9329 | 66.8758 | 0.009318 | 1.0383 | 0.024892 | 23.9739 | -0.015026 | 0.14268 | -3.1e-05 | -0.002066 | -0.002066 | -2.8385
2019-03-21 12:32 | 1553185949.3937 | -158 | 0.20941 | 0.000459 | -4e-06 | 100.2821 | 100.2433 | 124.7707 | 21.2234 | 0.00192 | 0.001827 | -152.099 | 0.239 | 8.3986 | 0.003471 | 0.004949 | 0.004666 | 6.1e-05 | 0.001141 | 0.25536 | 1.3167 | 2.548 | 99.38 | 20.96 | 37.282 | 67.9921 | 105.3457 | 65.7583 | 0.009318 | 1.0417 | 0.024938 | 23.9392 | 0.00346 | 0.14218 | 0.000146 | 0.01114 | 0.01114 | 0.52453
2019-03-21 12:32 | 1553185950.3933 | -157 | 0.20946 | 0.000461 | -3e-06 | 100.193 | 100.2201 | 124.636 | 21.2236 | 0.001996 | 0.001877 | -152.1089 | 0.23891 | 8.4036 | 0.003492 | 0.005094 | 0.004947 | 6.2e-05 | 0.00114 | 0.24635 | 1.2804 | 2.548 | 99.38 | 20.96 | 37.28 | 69.8187 | 105.8564 | 66.9586 | 0.009317 | 1.038 | 0.02488 | 23.9686 | -0.002423 | 0.14213 | -0.000129 | -0.009429 | -0.009429 | -0.61952
2019-03-21 12:32 | 1553185951.3924 | -156 | 0.20944 | 0.000462 | -3e-06 | 100.2883 | 100.5042 | 124.2658 | 21.2227 | 0.001971 | 0.001859 | -152.1036 | 0.23939 | 8.404 | 0.003477 | 0.004513 | 0.004372 | 6.2e-05 | 0.001144 | 0.29048 | 1.2839 | 2.5479 | 99.38 | 20.96 | 37.277 | 69.2167 | 105.5007 | 66.5194 | 0.009316 | 1.0394 | 0.024859 | 23.9175 | -0.019231 | 0.14241 | -4.3e-05 | -0.003008 | -0.003008 | -1.9459
2019-03-21 12:33 | 1553185952.3923 | -155 | 0.2094 | 0.00046 | -3e-06 | 100.3152 | 100.3186 | 123.8933 | 21.2235 | 0.002031 | 0.001866 | -152.1182 | 0.23967 | 8.4056 | 0.003531 | 0.004348 | 0.004125 | 6.3e-05 | 0.001139 | 0.26416 | 1.2823 | 2.5479 | 99.38 | 20.96 | 37.275 | 70.6585 | 106.8016 | 66.6784 | 0.009316 | 1.0389 | 0.024816 | 23.8872 | -0.000304 | 0.14258 | 0.000202 | 0.015269 | 0.015269 | 0.38378
2019-03-21 12:33 | 1553185953.3929 | -154 | 0.20943 | 0.000459 | -4e-06 | 100.2398 | 100.3799 | 123.9299 | 21.2234 | 0.001944 | 0.001878 | -152.1361 | 0.23969 | 8.4036 | 0.003443 | 0.004518 | 0.004469 | 6.6e-05 | 0.001139 | 0.29048 | 1.2876 | 2.548 | 99.38 | 20.96 | 37.271 | 68.5529 | 104.67 | 66.9793 | 0.009315 | 1.038 | 0.024809 | 23.9013 | -0.012493 | 0.14259 | 5e-05 | 0.003955 | 0.003955 | 1.4817
2019-03-21 12:33 | 1553185954.3939 | -153 | 0.20945 | 0.000462 | -4e-06 | 100.2285 | 100.2695 | 123.7605 | 21.2244 | 0.001972 | 0.001829 | -152.1282 | 0.23911 | 8.4022 | 0.003511 | 0.004839 | 0.004601 | 6.3e-05 | 0.001136 | 0.25608 | 1.2933 | 2.548 | 99.38 | 20.96 | 37.268 | 69.2433 | 106.3185 | 65.8124 | 0.009314 | 1.0415 | 0.024835 | 23.8439 | -0.003665 | 0.14225 | -7.7e-05 | -0.005487 | -0.005487 | -1.0656
2019-03-21 12:33 | 1553185955.3926 | -152 | 0.20944 | 0.000456 | -4e-06 | 100.2404 | 100.2883 | 123.7558 | 21.2236 | 0.001938 | 0.001834 | -152.1236 | 0.2385 | 8.401 | 0.003512 | 0.004491 | 0.004331 | 6.4e-05 | 0.001135 | 0.28672 | 1.2903 | 2.548 | 99.38 | 20.96 | 37.266 | 68.4229 | 106.335 | 65.9185 | 0.009314 | 1.0412 | 0.02483 | 23.8472 | -0.004272 | 0.14188 | -1.8e-05 | -0.001031 | -0.001031 | -5.6539
2019-03-21 12:33 | 1553185956.3941 | -151 | 0.20946 | 0.000465 | -4e-06 | 100.2612 | 100.4763 | 123.2462 | 21.2241 | 0.001983 | 0.001836 | -152.1155 | 0.24017 | 8.403 | 0.003515 | 0.004824 | 0.004646 | 6.4e-05 | 0.001138 | 0.28838 | 1.3473 | 2.548 | 99.38 | 20.96 | 37.263 | 69.5137 | 106.4196 | 65.9704 | 0.009313 | 1.0411 | 0.024777 | 23.7999 | -0.019163 | 0.14288 | -0.000135 | -0.009827 | -0.009827 | -0.59755
2019-03-21 12:33 | 1553185957.393 | -150 | 0.20944 | 0.000458 | -4e-06 | 100.2949 | 100.4524 | 123.2859 | 21.2241 | 0.001963 | 0.001816 | -152.1145 | 0.23997 | 8.4012 | 0.003485 | 0.004793 | 0.00453 | 6.5e-05 | 0.001138 | 0.29502 | 1.3327 | 2.5479 | 99.38 | 20.96 | 37.261 | 69.02 | 105.6934 | 65.4763 | 0.009312 | 1.0426 | 0.024799 | 23.7864 | -0.014037 | 0.14276 | -1.1e-05 | -0.0006 | -0.0006 | -9.7742
2019-03-21 12:33 | 1553185958.3927 | -149 | 0.20943 | 0.000461 | -4e-06 | 100.3047 | 100.3039 | 123.0922 | 21.2238 | 0.001944 | 0.001825 | -152.0914 | 0.23945 | 8.4024 | 0.003554 | 0.005596 | 0.005512 | 6.5e-05 | 0.001134 | 0.31045 | 1.3615 | 2.548 | 99.38 | 20.96 | 37.259 | 68.5603 | 107.365 | 65.6979 | 0.009312 | 1.0419 | 0.024772 | 23.7755 | 7.2e-05 | 0.14245 | 5.7e-05 | 0.004464 | 0.004464 | 1.3116
2019-03-21 12:33 | 1553185959.3941 | -148 | 0.20946 | 0.00046 | -4e-06 | 100.387 | 100.4126 | 122.876 | 21.2246 | 0.001937 | 0.001816 | -152.0847 | 0.23911 | 8.4016 | 0.003523 | 0.004582 | 0.004397 | 6.5e-05 | 0.001141 | 0.27328 | 1.3465 | 2.5481 | 99.38 | 20.96 | 37.255 | 68.4065 | 106.6186 | 65.4893 | 0.009311 | 1.0425 | 0.024757 | 23.7473 | -0.002282 | 0.14225 | -0.000132 | -0.009605 | -0.009605 | -0.60867
2019-03-21 12:33 | 1553185960.3938 | -147 | 0.20945 | 0.000455 | -4e-06 | 100.345 | 100.3312 | 122.8805 | 21.2237 | 0.001968 | 0.001826 | -152.0909 | 0.24007 | 8.4045 | 0.003515 | 0.00548 | 0.005278 | 6.6e-05 | 0.001136 | 0.28672 | 1.299 | 2.548 | 99.38 | 20.96 | 37.253 | 69.1427 | 106.4287 | 65.7384 | 0.00931 | 1.0418 | 0.024749 | 23.7564 | 0.001228 | 0.14282 | -6.4e-05 | -0.004528 | -0.004528 | -1.2964
2019-03-21 12:33 | 1553185961.3943 | -146 | 0.20939 | 0.00046 | -4e-06 | 100.3986 | 100.4688 | 122.999 | 21.2237 | 0.001949 | 0.001871 | -152.0851 | 0.23985 | 8.4026 | 0.003543 | 0.004249 | 0.004107 | 6.3e-05 | 0.001133 | 0.27577 | 1.3492 | 2.548 | 99.38 | 20.96 | 37.251 | 68.6763 | 107.0955 | 66.8012 | 0.00931 | 1.0385 | 0.024722 | 23.8051 | -0.006251 | 0.14269 | 0.000253 | 0.019034 | 0.019034 | 0.3081
2019-03-21 12:33 | 1553185962.3946 | -145 | 0.20944 | 0.00046 | -4e-06 | 100.3739 | 100.4878 | 123.2098 | 21.2247 | 0.001993 | 0.001837 | -152.0656 | 0.24006 | 8.4027 | 0.003453 | 0.004753 | 0.004548 | 6.4e-05 | 0.001135 | 0.28766 | 1.3752 | 2.548 | 99.38 | 20.96 | 37.25 | 69.7457 | 104.9174 | 65.9821 | 0.00931 | 1.041 | 0.024773 | 23.7968 | -0.010152 | 0.14281 | -4e-06 | -5.7e-05 | -5.7e-05 | -102.3812
2019-03-21 12:33 | 1553185963.3918 | -144 | 0.20947 | 0.000463 | -4e-06 | 100.3813 | 100.4654 | 123.0934 | 21.2238 | 0.001915 | 0.001827 | -152.0617 | 0.23908 | 8.4022 | 0.003471 | 0.004997 | 0.004835 | 6.2e-05 | 0.001132 | 0.26792 | 1.3376 | 2.5481 | 99.38 | 20.96 | 37.248 | 67.8646 | 105.3546 | 65.7438 | 0.009309 | 1.0418 | 0.02477 | 23.7772 | -0.007487 | 0.14223 | -0.000197 | -0.014426 | -0.014426 | -0.40521
2019-03-21 12:33 | 1553185964.3929 | -143 | 0.20944 | 0.000457 | -4e-06 | 100.4266 | 100.5297 | 123.2952 | 21.2246 | 0.001986 | 0.001825 | -152.0798 | 0.23955 | 8.4013 | 0.003421 | 0.004292 | 0.004133 | 6.3e-05 | 0.001137 | 0.26576 | 1.3146 | 2.548 | 99.38 | 20.96 | 37.248 | 69.5858 | 104.1468 | 65.7111 | 0.009309 | 1.0419 | 0.024792 | 23.7955 | -0.009182 | 0.14251 | -2.4e-05 | -0.00155 | -0.00155 | -3.7799
2019-03-21 12:33 | 1553185965.3919 | -142 | 0.20944 | 0.00046 | -4e-06 | 100.3509 | 100.6782 | 123.295 | 21.2251 | 0.001942 | 0.001835 | -152.0785 | 0.23996 | 8.4024 | 0.003466 | 0.005364 | 0.00513 | 6.1e-05 | 0.001135 | 0.27721 | 1.2961 | 2.5479 | 99.38 | 20.96 | 37.248 | 68.5289 | 105.2237 | 65.9536 | 0.009309 | 1.0411 | 0.024783 | 23.804 | -0.029128 | 0.14275 | -3.3e-05 | -0.002229 | -0.002229 | -2.6318
2019-03-21 12:33 | 1553185966.3923 | -141 | 0.20947 | 0.000455 | -4e-06 | 100.3682 | 100.4936 | 123.5389 | 21.2254 | 0.002014 | 0.001793 | -152.0954 | 0.23935 | 8.4007 | 0.003433 | 0.005083 | 0.004973 | 6.1e-05 | 0.001134 | 0.29286 | 1.3265 | 2.5481 | 99.38 | 20.96 | 37.247 | 70.261 | 104.4195 | 64.9361 | 0.009309 | 1.0442 | 0.024844 | 23.7918 | -0.011172 | 0.14239 | -0.000189 | -0.013837 | -0.013837 | -0.42292
2019-03-21 12:33 | 1553185967.3934 | -140 | 0.20944 | 0.000458 | -4e-06 | 100.4144 | 100.5989 | 123.9827 | 21.2246 | 0.002035 | 0.00178 | -152.1138 | 0.2398 | 8.4022 | 0.003501 | 0.00476 | 0.004597 | 6e-05 | 0.00113 | 0.3095 | 1.306 | 2.548 | 99.38 | 20.96 | 37.247 | 70.7649 | 106.0771 | 64.6206 | 0.009309 | 1.0452 | 0.024901 | 23.8234 | -0.01642 | 0.14266 | -9e-06 | -0.000433 | -0.000433 | -13.5258
2019-03-21 12:33 | 1553185968.3932 | -139 | 0.20944 | 0.000461 | -4e-06 | 100.4144 | 100.3077 | 124.2785 | 21.2256 | 0.001917 | 0.001858 | -152.1242 | 0.23993 | 8.4005 | 0.003508 | 0.004688 | 0.004537 | 6.3e-05 | 0.001131 | 0.27577 | 1.3414 | 2.5481 | 99.38 | 20.96 | 37.247 | 67.9233 | 106.2457 | 66.5056 | 0.009309 | 1.0394 | 0.024861 | 23.9182 | 0.00951 | 0.14273 | -4.5e-05 | -0.003142 | -0.003142 | -1.8668
2019-03-21 12:33 | 1553185969.3932 | -138 | 0.20943 | 0.00046 | -4e-06 | 100.4573 | 100.5951 | 124.3557 | 21.2256 | 0.001975 | 0.001852 | -152.144 | 0.23949 | 8.4044 | 0.003478 | 0.004678 | 0.004507 | 5.9e-05 | 0.001129 | 0.23922 | 1.3094 | 2.548 | 99.38 | 20.96 | 37.25 | 69.3077 | 105.5295 | 66.3488 | 0.00931 | 1.0399 | 0.024875 | 23.9201 | -0.012266 | 0.14247 | 3.5e-05 | 0.002823 | 0.002823 | 2.0745
2019-03-21 12:33 | 1553185970.3935 | -137 | 0.20945 | 0.000462 | -4e-06 | 100.4531 | 100.6401 | 124.5463 | 21.2254 | 0.002023 | 0.001892 | -152.1307 | 0.24024 | 8.4023 | 0.003512 | 0.005071 | 0.004938 | 6e-05 | 0.00113 | 0.26034 | 1.3542 | 2.548 | 99.38 | 20.96 | 37.251 | 70.4639 | 106.3508 | 67.3216 | 0.00931 | 1.0369 | 0.024858 | 23.9727 | -0.016634 | 0.14292 | -9.2e-05 | -0.006581 | -0.006581 | -0.89252
2019-03-21 12:33 | 1553185971.3932 | -136 | 0.20944 | 0.000459 | -4e-06 | 100.4504 | 100.4422 | 124.5592 | 21.2262 | 0.001999 | 0.001823 | -152.1726 | 0.24013 | 8.4055 | 0.003584 | 0.004723 | 0.004547 | 5.8e-05 | 0.001132 | 0.26305 | 1.3403 | 2.548 | 99.38 | 20.96 | 37.252 | 69.8875 | 108.0865 | 65.6665 | 0.00931 | 1.042 | 0.02492 | 23.9156 | 0.000732 | 0.14285 | -4.6e-05 | -0.003194 | -0.003194 | -1.8383
2019-03-21 12:33 | 1553185972.3934 | -135 | 0.20942 | 0.000461 | -4e-06 | 100.476 | 100.6125 | 124.5221 | 21.2257 | 0.001996 | 0.001872 | -152.2119 | 0.23945 | 8.4039 | 0.003557 | 0.004934 | 0.004778 | 5.8e-05 | 0.001136 | 0.28766 | 1.3378 | 2.548 | 99.38 | 20.96 | 37.253 | 69.8217 | 107.4402 | 66.8358 | 0.00931 | 1.0384 | 0.024873 | 23.9533 | -0.012139 | 0.14245 | 0.000106 | 0.008196 | 0.008196 | 0.7143
2019-03-21 12:33 | 1553185973.3928 | -134 | 0.20947 | 0.000459 | -4e-06 | 100.4087 | 100.5202 | 124.5841 | 21.2257 | 0.001989 | 0.001857 | -152.2051 | 0.24058 | 8.4026 | 0.003506 | 0.004725 | 0.004535 | 5.7e-05 | 0.001132 | 0.24923 | 1.3359 | 2.548 | 99.38 | 20.96 | 37.254 | 69.6453 | 106.1929 | 66.4872 | 0.009311 | 1.0395 | 0.024892 | 23.947 | -0.009926 | 0.14312 | -0.000165 | -0.012132 | -0.012132 | -0.48486
2019-03-21 12:33 | 1553185974.3927 | -133 | 0.20946 | 0.000457 | -4e-06 | 100.5207 | 100.652 | 124.1919 | 21.2265 | 0.00197 | 0.001834 | -152.1785 | 0.23955 | 8.4021 | 0.003514 | 0.004794 | 0.004664 | 5.8e-05 | 0.001135 | 0.29607 | 1.3247 | 2.548 | 99.38 | 20.96 | 37.257 | 69.1853 | 106.3967 | 65.9275 | 0.009311 | 1.0412 | 0.024874 | 23.8895 | -0.011673 | 0.14251 | -0.000106 | -0.007639 | -0.007639 | -0.76674
2019-03-21 12:33 | 1553185975.3936 | -132 | 0.20943 | 0.000462 | -4e-06 | 100.5192 | 100.8002 | 124.0281 | 21.2264 | 0.001958 | 0.00178 | -152.1738 | 0.23984 | 8.4067 | 0.003539 | 0.004725 | 0.004483 | 5.7e-05 | 0.001136 | 0.26593 | 1.3353 | 2.5479 | 99.38 | 20.96 | 37.26 | 68.9045 | 107.0045 | 64.614 | 0.009312 | 1.0452 | 0.024905 | 23.8275 | -0.024962 | 0.14268 | 5.5e-05 | 0.00437 | 0.00437 | 1.3418
2019-03-21 12:33 | 1553185976.3929 | -131 | 0.20947 | 0.000462 | -4e-06 | 100.5458 | 100.6496 | 124.1499 | 21.2269 | 0.001925 | 0.00186 | -152.1765 | 0.23958 | 8.4015 | 0.003522 | 0.004693 | 0.004515 | 5.8e-05 | 0.001135 | 0.27992 | 1.3183 | 2.5481 | 99.38 | 20.95 | 37.261 | 68.1186 | 106.5941 | 66.5491 | 0.009307 | 1.0393 | 0.024847 | 23.9074 | -0.009234 | 0.14252 | -0.000171 | -0.012481 | -0.012481 | -0.46932
2019-03-21 12:33 | 1553185977.3926 | -130 | 0.20941 | 0.000458 | -4e-06 | 100.4808 | 100.5754 | 123.576 | 21.2269 | 0.001937 | 0.001793 | -152.1691 | 0.23982 | 8.402 | 0.003561 | 0.004793 | 0.004593 | 5.9e-05 | 0.001136 | 0.2844 | 1.285 | 2.5478 | 99.38 | 20.95 | 37.263 | 68.3979 | 107.5358 | 64.925 | 0.009307 | 1.0443 | 0.024849 | 23.795 | -0.008414 | 0.14267 | 0.000196 | 0.014838 | 0.014838 | 0.39518
2019-03-21 12:33 | 1553185978.3928 | -129 | 0.2095 | 0.00046 | -4e-06 | 100.5386 | 100.6592 | 123.5248 | 21.2266 | 0.001985 | 0.001805 | -152.179 | 0.23989 | 8.4034 | 0.003571 | 0.004898 | 0.004745 | 5.8e-05 | 0.001134 | 0.27577 | 1.3158 | 2.548 | 99.38 | 20.95 | 37.266 | 69.5472 | 107.771 | 65.2253 | 0.009308 | 1.0434 | 0.024832 | 23.8006 | -0.010718 | 0.14271 | -0.000375 | -0.027677 | -0.027677 | -0.21193
2019-03-21 12:33 | 1553185979.3917 | -128 | 0.20944 | 0.00046 | -4e-06 | 100.6116 | 100.7498 | 123.7466 | 21.2271 | 0.001957 | 0.00183 | -152.1678 | 0.2399 | 8.4032 | 0.003544 | 0.005212 | 0.005087 | 5.5e-05 | 0.001134 | 0.2792 | 1.2883 | 2.5479 | 99.38 | 20.95 | 37.267 | 68.8895 | 107.1185 | 65.8228 | 0.009308 | 1.0415 | 0.024833 | 23.843 | -0.012273 | 0.14272 | 5e-06 | 0.000604 | 0.000604 | 9.7038
2019-03-21 12:33 | 1553185980.3932 | -127 | 0.20952 | 0.000459 | -4e-06 | 100.6086 | 100.7549 | 123.5933 | 21.228 | 0.001921 | 0.001852 | -152.1872 | 0.23955 | 8.4003 | 0.003555 | 0.004809 | 0.004617 | 6e-05 | 0.001141 | 0.30469 | 1.3057 | 2.548 | 99.38 | 20.96 | 37.268 | 68.0185 | 107.4022 | 66.3591 | 0.009314 | 1.0399 | 0.024798 | 23.847 | -0.012991 | 0.14251 | -0.000475 | -0.035107 | -0.035107 | -0.16684
2019-03-21 12:33 | 1553185981.4284 | -126 | 0.20943 | 0.000461 | -4e-06 | 100.6 | 100.5696 | 123.7755 | 21.2275 | 0.001965 | 0.001852 | -152.1496 | 0.23974 | 8.4007 | 0.003562 | 0.004773 | 0.00459 | 5.7e-05 | 0.00113 | 0.25769 | 1.3005 | 2.548 | 99.38 | 20.96 | 37.269 | 69.0653 | 107.5585 | 66.3554 | 0.009314 | 1.0399 | 0.024816 | 23.8645 | 0.002704 | 0.14262 | 3.9e-05 | 0.003152 | 0.003152 | 1.8596
2019-03-21 12:33 | 1553185982.3759 | -125 | 0.20945 | 0.000462 | -4e-06 | 100.5931 | 100.7787 | 123.7546 | 21.2271 | 0.002003 | 0.001829 | -152.112 | 0.23965 | 8.4022 | 0.003557 | 0.005041 | 0.00485 | 5.5e-05 | 0.001128 | 0.29535 | 1.2906 | 2.548 | 99.38 | 20.96 | 37.274 | 69.9797 | 107.4447 | 65.8064 | 0.009316 | 1.0416 | 0.024834 | 23.8432 | -0.016485 | 0.14257 | -8.7e-05 | -0.006234 | -0.006234 | -0.93993
2019-03-21 12:33 | 1553185983.3773 | -124 | 0.20934 | 0.000459 | -4e-06 | 100.6751 | 100.8233 | 123.726 | 21.2271 | 0.002014 | 0.001848 | -152.1189 | 0.23958 | 8.4012 | 0.003534 | 0.004529 | 0.00444 | 6e-05 | 0.001123 | 0.26394 | 1.3062 | 2.5479 | 99.38 | 20.95 | 37.274 | 70.2441 | 106.8823 | 66.2591 | 0.00931 | 1.0402 | 0.024815 | 23.8563 | -0.013162 | 0.14252 | 0.000585 | 0.043775 | 0.043775 | 0.13381
2019-03-21 12:33 | 1553185984.376 | -123 | 0.2097 | 0.000458 | -4e-06 | 100.6909 | 100.7372 | 123.7736 | 21.2278 | 0.001944 | 0.001846 | -152.1064 | 0.23992 | 8.4 | 0.003488 | 0.00561 | 0.005339 | 5.7e-05 | 0.001117 | 0.26897 | 1.2864 | 2.5482 | 99.38 | 20.96 | 37.275 | 68.5736 | 105.7584 | 66.2059 | 0.009316 | 1.0403 | 0.024822 | 23.859 | -0.004111 | 0.14273 | -0.001573 | -0.11703 | -0.11703 | -0.050127
2019-03-21 12:33 | 1553185985.3772 | -122 | 0.20963 | 0.000458 | -4e-06 | 100.6781 | 100.8363 | 123.8088 | 21.2278 | 0.00193 | 0.001824 | -152.1091 | 0.23978 | 8.4047 | 0.003576 | 0.004634 | 0.00455 | 6e-05 | 0.001125 | 0.29822 | 1.3088 | 2.5481 | 99.38 | 20.95 | 37.275 | 68.231 | 107.895 | 65.6938 | 0.00931 | 1.0419 | 0.024844 | 23.8444 | -0.014045 | 0.14265 | -0.001176 | -0.087445 | -0.087445 | -0.067046
2019-03-21 12:33 | 1553185986.3775 | -121 | 0.20946 | 0.000458 | -4e-06 | 100.6646 | 100.8206 | 123.5124 | 21.2276 | 0.00198 | 0.001834 | -152.0922 | 0.23952 | 8.4029 | 0.003523 | 0.004652 | 0.004415 | 5.9e-05 | 0.001123 | 0.25785 | 1.3202 | 2.5481 | 99.38 | 20.95 | 37.276 | 69.432 | 106.6167 | 65.9269 | 0.00931 | 1.0412 | 0.024805 | 23.8241 | -0.013847 | 0.14249 | -0.000149 | -0.010885 | -0.010885 | -0.53803
2019-03-21 12:33 | 1553185987.3772 | -120 | 0.20962 | 0.000462 | -4e-06 | 100.7442 | 100.7699 | 123.5152 | 21.2279 | 0.001907 | 0.001827 | -152.0849 | 0.24043 | 8.4023 | 0.003541 | 0.005061 | 0.004908 | 5.8e-05 | 0.001127 | 0.25299 | 1.3047 | 2.5478 | 99.38 | 20.95 | 37.277 | 67.6829 | 107.0525 | 65.7477 | 0.009311 | 1.0417 | 0.024812 | 23.818 | -0.002277 | 0.14303 | -0.001078 | -0.080091 | -0.080091 | -0.073399
2019-03-21 12:33 | 1553185988.3772 | -119 | 0.2093 | 0.000461 | -4e-06 | 100.7585 | 100.6438 | 123.479 | 21.2275 | 0.001976 | 0.001821 | -152.075 | 0.23979 | 8.4007 | 0.003568 | 0.004369 | 0.004222 | 5.7e-05 | 0.00113 | 0.25426 | 1.3294 | 2.548 | 99.38 | 20.95 | 37.276 | 69.3298 | 107.7034 | 65.6226 | 0.00931 | 1.0421 | 0.024813 | 23.8101 | 0.010185 | 0.14265 | 0.000852 | 0.063707 | 0.063707 | 0.09203
2019-03-21 12:33 | 1553185989.3776 | -118 | 0.20944 | 0.000459 | -4e-06 | 100.7084 | 100.8666 | 123.0495 | 21.228 | 0.001938 | 0.001858 | -152.0722 | 0.23978 | 8.401 | 0.003451 | 0.004877 | 0.004644 | 5.7e-05 | 0.001131 | 0.26017 | 1.3079 | 2.548 | 99.38 | 20.95 | 37.277 | 68.4161 | 104.8726 | 66.5079 | 0.009311 | 1.0394 | 0.024738 | 23.7997 | -0.014034 | 0.14265 | -9e-06 | -0.000437 | -0.000437 | -13.4176
2019-03-21 12:33 | 1553185990.3763 | -117 | 0.20944 | 0.000461 | -4e-06 | 100.6751 | 101.0123 | 123.1131 | 21.2281 | 0.001951 | 0.001826 | -152.047 | 0.24024 | 8.4029 | 0.003434 | 0.004523 | 0.004352 | 6e-05 | 0.001135 | 0.30702 | 1.3011 | 2.5478 | 99.38 | 20.95 | 37.279 | 68.7469 | 104.4512 | 65.7347 | 0.009311 | 1.0418 | 0.024772 | 23.7788 | -0.029915 | 0.14292 | -1.7e-05 | -0.00101 | -0.00101 | -5.8183
2019-03-21 12:33 | 1553185991.3778 | -116 | 0.20945 | 0.000457 | -4e-06 | 100.7689 | 100.5321 | 123.034 | 21.2278 | 0.002014 | 0.00184 | -152.0553 | 0.24047 | 8.4013 | 0.003404 | 0.004715 | 0.004562 | 5.7e-05 | 0.001136 | 0.26017 | 1.3296 | 2.5481 | 99.38 | 20.95 | 37.28 | 70.2501 | 103.7206 | 66.0616 | 0.009311 | 1.0408 | 0.024752 | 23.7826 | 0.021044 | 0.14305 | -6.6e-05 | -0.004674 | -0.004674 | -1.2579
2019-03-21 12:33 | 1553185992.3773 | -115 | 0.20947 | 0.000459 | -4e-06 | 100.7367 | 100.8713 | 123.3114 | 21.2279 | 0.001949 | 0.001819 | -152.0691 | 0.24001 | 8.4037 | 0.003457 | 0.004844 | 0.004607 | 5.8e-05 | 0.001138 | 0.30702 | 1.2915 | 2.548 | 99.38 | 20.94 | 37.282 | 68.6999 | 105.0267 | 65.559 | 0.009306 | 1.0423 | 0.024799 | 23.7917 | -0.011943 | 0.14278 | -0.000225 | -0.016521 | -0.016521 | -0.3552
2019-03-21 12:33 | 1553185993.3769 | -114 | 0.20942 | 0.000459 | -4e-06 | 100.6804 | 100.8318 | 123.7417 | 21.2291 | 0.00193 | 0.001812 | -152.0632 | 0.2396 | 8.4019 | 0.003474 | 0.004946 | 0.004777 | 5.8e-05 | 0.001138 | 0.28672 | 1.2862 | 2.5481 | 99.38 | 20.94 | 37.282 | 68.2301 | 105.4369 | 65.4002 | 0.009306 | 1.0428 | 0.024848 | 23.8276 | -0.01344 | 0.14253 | 9.3e-05 | 0.007204 | 0.007204 | 0.81316
2019-03-21 12:33 | 1553185994.3769 | -113 | 0.20943 | 0.000458 | -4e-06 | 100.7555 | 101.0154 | 124.0398 | 21.2276 | 0.001964 | 0.001835 | -152.075 | 0.24039 | 8.4022 | 0.003483 | 0.004889 | 0.004597 | 5.7e-05 | 0.001141 | 0.31293 | 1.3005 | 2.548 | 99.38 | 20.94 | 37.284 | 69.0412 | 105.6377 | 65.9543 | 0.009307 | 1.0411 | 0.024857 | 23.8758 | -0.023042 | 0.14301 | 7.1e-05 | 0.005484 | 0.005484 | 1.0718
2019-03-21 12:33 | 1553185995.3756 | -112 | 0.20942 | 0.000464 | -4e-06 | 100.7254 | 100.9269 | 124.3071 | 21.2284 | 0.001941 | 0.001818 | -152.0589 | 0.24031 | 8.4053 | 0.003526 | 0.005014 | 0.004884 | 5.8e-05 | 0.001142 | 0.28423 | 1.2983 | 2.5481 | 99.38 | 20.94 | 37.285 | 68.5008 | 106.6883 | 65.5493 | 0.009307 | 1.0424 | 0.024899 | 23.8873 | -0.01787 | 0.14296 | 0.000103 | 0.007933 | 0.007933 | 0.7407
2019-03-21 12:33 | 1553185996.377 | -111 | 0.20945 | 0.000458 | -4e-06 | 100.7186 | 100.935 | 124.3531 | 21.2287 | 0.001943 | 0.001805 | -152.0512 | 0.23965 | 8.406 | 0.003545 | 0.004574 | 0.004381 | 6.1e-05 | 0.001141 | 0.28064 | 1.339 | 2.548 | 99.38 | 20.94 | 37.288 | 68.5533 | 107.1572 | 65.2164 | 0.009308 | 1.0434 | 0.024916 | 23.8799 | -0.019203 | 0.14257 | -4.8e-05 | -0.003298 | -0.003298 | -1.7765
2019-03-21 12:33 | 1553185997.3761 | -110 | 0.20944 | 0.00046 | -4e-06 | 100.7701 | 100.7746 | 124.3928 | 21.2284 | 0.001985 | 0.001824 | -152.0549 | 0.23957 | 8.4005 | 0.003513 | 0.004974 | 0.004766 | 5.5e-05 | 0.001143 | 0.26936 | 1.3428 | 2.5481 | 99.38 | 20.94 | 37.29 | 69.5669 | 106.3853 | 65.686 | 0.009308 | 1.0419 | 0.024903 | 23.9003 | -0.0004 | 0.14252 | -3e-05 | -0.001986 | -0.001986 | -2.9501
2019-03-21 12:33 | 1553185998.3757 | -109 | 0.20947 | 0.000456 | -4e-06 | 100.7713 | 100.9051 | 124.271 | 21.2285 | 0.001954 | 0.001878 | -152.0558 | 0.23914 | 8.4019 | 0.003522 | 0.004661 | 0.004507 | 5.6e-05 | 0.001147 | 0.25354 | 1.2988 | 2.5478 | 99.38 | 20.94 | 37.292 | 68.8169 | 106.5822 | 66.9735 | 0.009309 | 1.038 | 0.024843 | 23.934 | -0.011864 | 0.14226 | -0.000158 | -0.011521 | -0.011521 | -0.50753
2019-03-21 12:33 | 1553185999.3758 | -108 | 0.20945 | 0.000459 | -4e-06 | 100.8047 | 100.9013 | 124.4102 | 21.2288 | 0.001949 | 0.001827 | -152.0696 | 0.24006 | 8.4014 | 0.003534 | 0.00425 | 0.004086 | 5.8e-05 | 0.001154 | 0.28711 | 1.3072 | 2.548 | 99.38 | 20.94 | 37.296 | 68.6993 | 106.8731 | 65.7573 | 0.00931 | 1.0417 | 0.024902 | 23.9045 | -0.008571 | 0.14281 | -6.2e-05 | -0.004397 | -0.004397 | -1.335
2019-03-21 12:33 | 1553186000.3763 | -107 | 0.20946 | 0.000458 | -4e-06 | 100.7597 | 100.8826 | 124.1187 | 21.2288 | 0.002013 | 0.00185 | -152.0956 | 0.23975 | 8.4023 | 0.003476 | 0.004583 | 0.004381 | 5.8e-05 | 0.001152 | 0.28943 | 1.3499 | 2.548 | 99.38 | 20.94 | 37.298 | 70.2301 | 105.4737 | 66.299 | 0.00931 | 1.0401 | 0.024853 | 23.8956 | -0.010903 | 0.14263 | -0.000131 | -0.009495 | -0.009495 | -0.6174
2019-03-21 12:33 | 1553186001.3759 | -106 | 0.20944 | 0.00046 | -4e-06 | 100.7549 | 100.8935 | 123.9999 | 21.2296 | 0.002005 | 0.001787 | -152.1052 | 0.24019 | 8.4002 | 0.003464 | 0.004505 | 0.004339 | 5.7e-05 | 0.001152 | 0.28528 | 1.3099 | 2.548 | 99.38 | 20.94 | 37.3 | 70.0368 | 105.177 | 64.7918 | 0.009311 | 1.0447 | 0.024896 | 23.831 | -0.012292 | 0.14289 | 1.2e-05 | 0.001143 | 0.001143 | 5.1393
2019-03-21 12:33 | 1553186002.3759 | -105 | 0.20944 | 0.000459 | -4e-06 | 100.7654 | 100.7521 | 123.5673 | 21.2298 | 0.001991 | 0.001843 | -152.0994 | 0.23937 | 8.4021 | 0.003528 | 0.004877 | 0.004631 | 5.6e-05 | 0.00115 | 0.24851 | 1.3286 | 2.5481 | 99.38 | 20.94 | 37.302 | 69.6954 | 106.7345 | 66.1327 | 0.009311 | 1.0406 | 0.024803 | 23.8366 | 0.001173 | 0.1424 | -2.6e-05 | -0.001708 | -0.001708 | -3.4272
2019-03-21 12:33 | 1553186003.377 | -104 | 0.20946 | 0.00046 | -4e-06 | 100.8086 | 101.0154 | 123.1157 | 21.2288 | 0.001984 | 0.001833 | -152.0959 | 0.24031 | 8.4028 | 0.003484 | 0.004675 | 0.004574 | 6.1e-05 | 0.001152 | 0.24746 | 1.3621 | 2.548 | 99.38 | 20.94 | 37.303 | 69.5337 | 105.6791 | 65.8947 | 0.009311 | 1.0413 | 0.024767 | 23.7846 | -0.018336 | 0.14296 | -0.000125 | -0.009044 | -0.009044 | -0.64969
2019-03-21 12:33 | 1553186004.3765 | -103 | 0.20951 | 0.000463 | -4e-06 | 100.8422 | 101.0035 | 123.0415 | 21.2278 | 0.002005 | 0.001831 | -152.1048 | 0.23873 | 8.4032 | 0.003463 | 0.005613 | 0.005443 | 6e-05 | 0.00115 | 0.2557 | 1.3243 | 2.548 | 99.38 | 20.95 | 37.305 | 70.0458 | 105.1723 | 65.8516 | 0.009318 | 1.0414 | 0.024761 | 23.776 | -0.014293 | 0.14202 | -0.000435 | -0.032097 | -0.032097 | -0.18185
2019-03-21 12:33 | 1553186005.3765 | -102 | 0.20945 | 0.000461 | -4e-06 | 100.8264 | 100.9817 | 122.7664 | 21.2295 | 0.001968 | 0.001873 | -152.1111 | 0.23996 | 8.4045 | 0.00355 | 0.004562 | 0.004358 | 5.7e-05 | 0.001153 | 0.27759 | 1.3198 | 2.548 | 99.38 | 20.95 | 37.306 | 69.1531 | 107.2759 | 66.8542 | 0.009318 | 1.0384 | 0.024697 | 23.7845 | -0.013763 | 0.14275 | -7.8e-05 | -0.005536 | -0.005536 | -1.0599
2019-03-21 12:33 | 1553186006.3762 | -101 | 0.20947 | 0.000459 | -4e-06 | 100.7665 | 100.871 | 123.0486 | 21.2281 | 0.00198 | 0.001884 | -152.0965 | 0.24018 | 8.4015 | 0.003537 | 0.005004 | 0.004847 | 6.3e-05 | 0.001154 | 0.27057 | 1.3078 | 2.5479 | 99.38 | 20.94 | 37.306 | 69.4263 | 106.9543 | 67.1331 | 0.009312 | 1.0375 | 0.024715 | 23.8215 | -0.009268 | 0.14288 | -0.000209 | -0.015246 | -0.015246 | -0.38518
2019-03-21 12:33 | 1553186007.3772 | -100 | 0.20944 | 0.000458 | -4e-06 | 100.7677 | 100.9228 | 123.352 | 21.2292 | 0.001958 | 0.001863 | -152.1023 | 0.23958 | 8.4044 | 0.003587 | 0.004911 | 0.004673 | 6.3e-05 | 0.001156 | 0.2808 | 1.2754 | 2.5479 | 99.38 | 20.94 | 37.308 | 68.8993 | 108.1635 | 66.6123 | 0.009313 | 1.0391 | 0.024764 | 23.8326 | -0.013752 | 0.14252 | 1e-05 | 0.001027 | 0.001027 | 5.7014
2019-03-21 12:33 | 1553186008.3763 | -99 | 0.20946 | 0.000463 | -4e-06 | 100.85 | 101.0324 | 123.3121 | 21.2291 | 0.001985 | 0.001845 | -152.1208 | 0.24016 | 8.4054 | 0.003578 | 0.00507 | 0.004924 | 6e-05 | 0.001159 | 0.29319 | 1.2816 | 2.5479 | 99.38 | 20.95 | 37.306 | 69.5568 | 107.9483 | 66.1971 | 0.009318 | 1.0404 | 0.024776 | 23.8142 | -0.01617 | 0.14287 | -0.000108 | -0.007795 | -0.007795 | -0.75327
2019-03-21 12:33 | 1553186009.3754 | -98 | 0.20947 | 0.000462 | -4e-06 | 100.7698 | 100.8543 | 123.0647 | 21.2297 | 0.00198 | 0.001833 | -152.1031 | 0.23993 | 8.4014 | 0.003518 | 0.00467 | 0.004453 | 6.2e-05 | 0.001156 | 0.28567 | 1.2957 | 2.548 | 99.38 | 20.95 | 37.303 | 69.4427 | 106.4907 | 65.9057 | 0.009317 | 1.0413 | 0.024761 | 23.7801 | -0.007497 | 0.14274 | -0.000198 | -0.014439 | -0.014439 | -0.40629
2019-03-21 12:33 | 1553186010.3764 | -97 | 0.20945 | 0.000459 | -4e-06 | 100.7963 | 101.0668 | 123.1607 | 21.2295 | 0.001974 | 0.001831 | -152.0854 | 0.23962 | 8.4014 | 0.00362 | 0.004754 | 0.004604 | 6e-05 | 0.001157 | 0.26576 | 1.2718 | 2.548 | 99.38 | 20.95 | 37.302 | 69.2896 | 108.9699 | 65.8626 | 0.009317 | 1.0414 | 0.024773 | 23.7879 | -0.023972 | 0.14255 | -5.4e-05 | -0.003795 | -0.003795 | -1.544
2019-03-21 12:34 | 1553186011.377 | -96 | 0.20949 | 0.000461 | -4e-06 | 100.8967 | 101.1745 | 123.2518 | 21.2293 | 0.00196 | 0.001816 | -152.0698 | 0.23909 | 8.4009 | 0.003525 | 0.005058 | 0.004945 | 6.2e-05 | 0.001155 | 0.27721 | 1.2913 | 2.548 | 99.38 | 20.95 | 37.301 | 68.9447 | 106.6649 | 65.497 | 0.009317 | 1.0425 | 0.024795 | 23.7838 | -0.024587 | 0.14224 | -0.000279 | -0.020507 | -0.020507 | -0.28507
2019-03-21 12:34 | 1553186012.3759 | -95 | 0.20945 | 0.000459 | -5e-06 | 100.8652 | 101.0733 | 123.5297 | 21.2304 | 0.00196 | 0.001849 | -152.0779 | 0.23932 | 8.4025 | 0.003615 | 0.004816 | 0.004551 | 6.1e-05 | 0.00115 | 0.27367 | 1.3411 | 2.548 | 99.38 | 20.95 | 37.3 | 68.9667 | 108.8535 | 66.2926 | 0.009316 | 1.0401 | 0.024794 | 23.8386 | -0.01844 | 0.14237 | -4.5e-05 | -0.003061 | -0.003061 | -1.9117
2019-03-21 12:34 | 1553186013.3756 | -94 | 0.20948 | 0.000462 | -4e-06 | 100.8428 | 101.0536 | 123.6811 | 21.2307 | 0.001935 | 0.001811 | -152.0955 | 0.23976 | 8.4044 | 0.003593 | 0.005138 | 0.005064 | 6e-05 | 0.001146 | 0.25481 | 1.3086 | 2.5479 | 99.38 | 20.95 | 37.3 | 68.3634 | 108.324 | 65.3743 | 0.009316 | 1.0429 | 0.024843 | 23.8209 | -0.018673 | 0.14263 | -0.000248 | -0.018224 | -0.018224 | -0.32168
2019-03-21 12:34 | 1553186014.3766 | -93 | 0.20945 | 0.000456 | -4e-06 | 100.8601 | 101.2208 | 123.6823 | 21.2294 | 0.001981 | 0.001824 | -152.0853 | 0.24026 | 8.4019 | 0.003612 | 0.005295 | 0.005109 | 6e-05 | 0.001143 | 0.286 | 1.3134 | 2.548 | 99.38 | 20.95 | 37.297 | 69.4696 | 108.7767 | 65.6901 | 0.009316 | 1.0419 | 0.024831 | 23.8321 | -0.03193 | 0.14293 | -7.8e-05 | -0.005586 | -0.005586 | -1.0516
2019-03-21 12:34 | 1553186015.376 | -92 | 0.20948 | 0.000456 | -4e-06 | 100.8455 | 101.1128 | 123.9062 | 21.2297 | 0.001993 | 0.001856 | -152.0847 | 0.23968 | 8.4048 | 0.003604 | 0.004573 | 0.004446 | 6.3e-05 | 0.001145 | 0.27721 | 1.2646 | 2.5479 | 99.38 | 20.95 | 37.296 | 69.7443 | 108.5856 | 66.4519 | 0.009315 | 1.0396 | 0.024826 | 23.8805 | -0.023679 | 0.14259 | -0.000221 | -0.016203 | -0.016203 | -0.36168
2019-03-21 12:34 | 1553186016.3768 | -91 | 0.20946 | 0.000461 | -4e-06 | 100.8592 | 100.9851 | 123.5955 | 21.2298 | 0.001934 | 0.001848 | -152.1 | 0.23993 | 8.4001 | 0.003592 | 0.004846 | 0.004726 | 6.2e-05 | 0.001144 | 0.28263 | 1.29 | 2.5479 | 99.38 | 20.95 | 37.297 | 68.3219 | 108.2933 | 66.2617 | 0.009316 | 1.0402 | 0.024802 | 23.8438 | -0.011158 | 0.14274 | -0.000107 | -0.007713 | -0.007713 | -0.76059
2019-03-21 12:34 | 1553186017.3768 | -90 | 0.20946 | 0.000454 | -4e-06 | 100.9331 | 100.9017 | 123.4335 | 21.2303 | 0.002015 | 0.001895 | -152.0886 | 0.23908 | 8.4011 | 0.003605 | 0.004828 | 0.004606 | 6.1e-05 | 0.001137 | 0.25642 | 1.308 | 2.5479 | 99.38 | 20.95 | 37.295 | 70.2768 | 108.601 | 67.3927 | 0.009315 | 1.0367 | 0.024744 | 23.8679 | 0.002786 | 0.14223 | -0.000142 | -0.010317 | -0.010317 | -0.56658
2019-03-21 12:34 | 1553186018.3768 | -89 | 0.20946 | 0.000458 | -4e-06 | 100.8744 | 101.0338 | 123.3164 | 21.2307 | 0.001989 | 0.001839 | -152.0736 | 0.24028 | 8.4029 | 0.003565 | 0.004951 | 0.00474 | 5.7e-05 | 0.001137 | 0.262 | 1.3299 | 2.548 | 99.38 | 20.95 | 37.295 | 69.6574 | 107.6427 | 66.0353 | 0.009315 | 1.0409 | 0.024782 | 23.8089 | -0.014123 | 0.14295 | -9.6e-05 | -0.006892 | -0.006892 | -0.85242
2019-03-21 12:34 | 1553186019.3763 | -88 | 0.20948 | 0.000455 | -4e-06 | 100.8512 | 101.0846 | 123.7621 | 21.2307 | 0.001975 | 0.001841 | -152.0871 | 0.24049 | 8.403 | 0.003556 | 0.00524 | 0.005067 | 6e-05 | 0.001137 | 0.24065 | 1.2969 | 2.5479 | 99.38 | 20.95 | 37.296 | 69.3055 | 107.4277 | 66.1065 | 0.009315 | 1.0406 | 0.024824 | 23.8544 | -0.020676 | 0.14306 | -0.000218 | -0.015995 | -0.015995 | -0.36762
2019-03-21 12:34 | 1553186020.377 | -87 | 0.20947 | 0.000457 | -4e-06 | 100.8297 | 101.0076 | 123.9518 | 21.2303 | 0.001991 | 0.001827 | -152.0936 | 0.24014 | 8.4022 | 0.003518 | 0.005061 | 0.004931 | 6.3e-05 | 0.001137 | 0.25066 | 1.335 | 2.548 | 99.38 | 20.95 | 37.296 | 69.6896 | 106.4902 | 65.7549 | 0.009315 | 1.0417 | 0.024856 | 23.8603 | -0.015765 | 0.14286 | -0.000161 | -0.011713 | -0.011713 | -0.50127
2019-03-21 12:34 | 1553186021.376 | -86 | 0.20946 | 0.000459 | -4e-06 | 100.9429 | 100.8666 | 124.2125 | 21.231 | 0.001973 | 0.001886 | -152.1046 | 0.23967 | 8.4009 | 0.003608 | 0.005088 | 0.004889 | 5.9e-05 | 0.001137 | 0.2792 | 1.2696 | 2.5478 | 99.38 | 20.95 | 37.298 | 69.2675 | 108.6901 | 67.1827 | 0.009316 | 1.0374 | 0.02483 | 23.9357 | 0.006768 | 0.14258 | -0.000146 | -0.010639 | -0.010639 | -0.55079
2019-03-21 12:34 | 1553186022.3765 | -85 | 0.20947 | 0.000461 | -4e-06 | 100.8747 | 100.9671 | 124.0996 | 21.231 | 0.001914 | 0.001829 | -152.0999 | 0.23982 | 8.4024 | 0.003574 | 0.00483 | 0.004629 | 6.3e-05 | 0.001137 | 0.30143 | 1.2637 | 2.5479 | 99.38 | 20.95 | 37.303 | 67.8622 | 107.8564 | 65.8032 | 0.009317 | 1.0416 | 0.024869 | 23.8763 | -0.008185 | 0.14267 | -0.000174 | -0.012713 | -0.012713 | -0.46124
2019-03-21 12:34 | 1553186023.3761 | -84 | 0.20944 | 0.000457 | -4e-06 | 100.8437 | 100.9582 | 124.0055 | 21.2311 | 0.002002 | 0.001857 | -152.0878 | 0.23971 | 8.4026 | 0.003514 | 0.005049 | 0.004874 | 6.2e-05 | 0.001137 | 0.27273 | 1.2741 | 2.5478 | 99.38 | 20.95 | 37.306 | 69.9579 | 106.399 | 66.4736 | 0.009318 | 1.0395 | 0.024835 | 23.8908 | -0.010149 | 0.14261 | -1.9e-05 | -0.001194 | -0.001194 | -4.9093
2019-03-21 12:34 | 1553186024.3772 | -83 | 0.20946 | 0.000459 | -5e-06 | 100.8631 | 100.8022 | 123.8555 | 21.232 | 0.002033 | 0.001854 | -152.1115 | 0.24046 | 8.4051 | 0.003512 | 0.004891 | 0.004703 | 5.7e-05 | 0.001134 | 0.2912 | 1.3086 | 2.5479 | 99.38 | 20.95 | 37.308 | 70.7206 | 106.3573 | 66.4026 | 0.009318 | 1.0397 | 0.024823 | 23.8739 | 0.0054 | 0.14305 | -0.000116 | -0.008383 | -0.008383 | -0.70132
2019-03-21 12:34 | 1553186025.3756 | -82 | 0.20947 | 0.000461 | -4e-06 | 100.8839 | 101.1009 | 123.6243 | 21.2306 | 0.002016 | 0.001839 | -152.1261 | 0.24003 | 8.4033 | 0.003536 | 0.005088 | 0.004954 | 6e-05 | 0.001136 | 0.27096 | 1.3094 | 2.548 | 99.38 | 20.95 | 37.317 | 70.299 | 106.9346 | 66.0588 | 0.009321 | 1.0408 | 0.024812 | 23.8395 | -0.019216 | 0.1428 | -0.000168 | -0.012227 | -0.012227 | -0.47998
2019-03-21 12:34 | 1553186026.3763 | -81 | 0.20945 | 0.000457 | -3e-06 | 100.8291 | 100.8945 | 123.944 | 21.2307 | 0.002018 | 0.001858 | -152.1284 | 0.24076 | 8.3997 | 0.003571 | 0.005151 | 0.004974 | 6.1e-05 | 0.001135 | 0.24906 | 1.2858 | 2.5479 | 99.38 | 20.95 | 37.32 | 70.3591 | 107.7954 | 66.4958 | 0.009321 | 1.0395 | 0.024828 | 23.8857 | -0.005799 | 0.14323 | -3.6e-05 | -0.002493 | -0.002493 | -2.3608
2019-03-21 12:34 | 1553186027.3763 | -80 | 0.20944 | 0.000461 | -4e-06 | 100.903 | 101.0485 | 124.0351 | 21.2316 | 0.001989 | 0.001832 | -152.1286 | 0.23861 | 8.405 | 0.003578 | 0.004704 | 0.004572 | 6.4e-05 | 0.001133 | 0.27489 | 1.337 | 2.548 | 99.38 | 20.95 | 37.324 | 69.651 | 107.9587 | 65.8803 | 0.009322 | 1.0413 | 0.02486 | 23.8728 | -0.012884 | 0.14195 | -2.6e-05 | -0.001719 | -0.001719 | -3.3939
2019-03-21 12:34 | 1553186028.3753 | -79 | 0.20944 | 0.000459 | -4e-06 | 100.8807 | 101.0781 | 124.2956 | 21.2312 | 0.001995 | 0.001877 | -152.1308 | 0.23939 | 8.4041 | 0.003599 | 0.004534 | 0.004415 | 6.3e-05 | 0.001133 | 0.26808 | 1.3832 | 2.5481 | 99.38 | 20.95 | 37.327 | 69.7972 | 108.4578 | 66.9636 | 0.009323 | 1.038 | 0.024846 | 23.936 | -0.017488 | 0.14242 | -4e-06 | -2.7e-05 | -2.7e-05 | -213.2841
2019-03-21 12:34 | 1553186029.3754 | -78 | 0.20946 | 0.000458 | -4e-06 | 100.8419 | 101.0171 | 124.5109 | 21.2315 | 0.001973 | 0.001843 | -152.1325 | 0.23918 | 8.4021 | 0.003544 | 0.005494 | 0.005226 | 6.3e-05 | 0.001133 | 0.28064 | 1.3113 | 2.548 | 99.38 | 20.95 | 37.334 | 69.2596 | 107.1175 | 66.1455 | 0.009325 | 1.0405 | 0.024898 | 23.9279 | -0.015526 | 0.14229 | -0.00012 | -0.008743 | -0.008743 | -0.66891
2019-03-21 12:34 | 1553186030.3755 | -77 | 0.20944 | 0.00046 | -4e-06 | 100.9009 | 101.0662 | 124.3874 | 21.2317 | 0.001919 | 0.001864 | -152.1203 | 0.24021 | 8.4043 | 0.003586 | 0.005185 | 0.00506 | 6.5e-05 | 0.00113 | 0.24906 | 1.3537 | 2.5479 | 99.38 | 20.95 | 37.339 | 67.9745 | 108.155 | 66.6575 | 0.009326 | 1.039 | 0.024866 | 23.934 | -0.014637 | 0.1429 | -1.4e-05 | -0.000758 | -0.000758 | -7.7525
2019-03-21 12:34 | 1553186031.3759 | -76 | 0.20942 | 0.000459 | -5e-06 | 100.8789 | 101.029 | 124.2614 | 21.2315 | 0.001962 | 0.00181 | -152.1086 | 0.24017 | 8.4007 | 0.003664 | 0.005273 | 0.005181 | 6.4e-05 | 0.00113 | 0.26864 | 1.3202 | 2.548 | 99.38 | 20.95 | 37.343 | 69.0021 | 110.0394 | 65.3493 | 0.009327 | 1.043 | 0.024902 | 23.8758 | -0.013305 | 0.14288 | 9.5e-05 | 0.007377 | 0.007377 | 0.79601
2019-03-21 12:34 | 1553186032.3765 | -75 | 0.20946 | 0.000454 | -5e-06 | 100.8664 | 101.0066 | 124.5651 | 21.2313 | 0.001987 | 0.001861 | -152.114 | 0.24114 | 8.4015 | 0.003653 | 0.004585 | 0.004314 | 6.4e-05 | 0.001132 | 0.25752 | 1.3024 | 2.548 | 99.38 | 20.95 | 37.351 | 69.6118 | 109.7726 | 66.5751 | 0.009329 | 1.0392 | 0.024887 | 23.9482 | -0.012424 | 0.14346 | -0.00013 | -0.009441 | -0.009441 | -0.62449
2019-03-21 12:34 | 1553186033.3756 | -74 | 0.20945 | 0.000457 | -4e-06 | 100.8172 | 101.0144 | 124.4468 | 21.2307 | 0.001987 | 0.001836 | -152.1377 | 0.23998 | 8.4017 | 0.003691 | 0.004229 | 0.004107 | 6.8e-05 | 0.001131 | 0.26394 | 1.3179 | 2.548 | 99.38 | 20.95 | 37.356 | 69.6144 | 110.7107 | 65.987 | 0.00933 | 1.041 | 0.024897 | 23.9161 | -0.017478 | 0.14276 | -5.1e-05 | -0.003579 | -0.003579 | -1.6395
2019-03-21 12:34 | 1553186034.3771 | -73 | 0.20942 | 0.000455 | -5e-06 | 100.8318 | 100.9408 | 124.4944 | 21.2313 | 0.001982 | 0.001857 | -152.1219 | 0.2397 | 8.4029 | 0.003664 | 0.005164 | 0.004987 | 6.4e-05 | 0.001132 | 0.24779 | 1.3094 | 2.5479 | 99.38 | 20.95 | 37.359 | 69.477 | 110.0519 | 66.4885 | 0.009331 | 1.0395 | 0.024883 | 23.9384 | -0.009667 | 0.1426 | 0.000125 | 0.009668 | 0.009668 | 0.60621
2019-03-21 12:34 | 1553186035.3758 | -72 | 0.20946 | 0.000459 | -4e-06 | 100.8461 | 101.1023 | 124.4125 | 21.2314 | 0.001974 | 0.001833 | -152.1061 | 0.24017 | 8.4013 | 0.003687 | 0.00504 | 0.004882 | 6.2e-05 | 0.00113 | 0.26377 | 1.2794 | 2.548 | 99.38 | 20.95 | 37.366 | 69.2835 | 110.6034 | 65.8916 | 0.009333 | 1.0413 | 0.024897 | 23.9095 | -0.022692 | 0.14288 | -9.4e-05 | -0.006762 | -0.006762 | -0.86845
2019-03-21 12:34 | 1553186036.3758 | -71 | 0.20944 | 0.000461 | -4e-06 | 100.897 | 101.0175 | 124.0332 | 21.2313 | 0.002021 | 0.00179 | -152.1157 | 0.2399 | 8.401 | 0.00365 | 0.004681 | 0.004482 | 6.3e-05 | 0.001129 | 0.27577 | 1.2505 | 2.5478 | 99.38 | 20.96 | 37.368 | 70.4282 | 109.7041 | 64.8615 | 0.009339 | 1.0445 | 0.024897 | 23.8367 | -0.010669 | 0.14272 | -2.6e-05 | -0.001724 | -0.001724 | -3.4023
2019-03-21 12:34 | 1553186037.3768 | -70 | 0.20943 | 0.000456 | -4e-06 | 100.7713 | 100.9132 | 123.9248 | 21.2316 | 0.002001 | 0.001839 | -152.1235 | 0.24028 | 8.4031 | 0.003696 | 0.004399 | 0.004253 | 6.8e-05 | 0.001127 | 0.26217 | 1.2826 | 2.5479 | 99.38 | 20.96 | 37.373 | 69.9334 | 110.8237 | 66.0403 | 0.00934 | 1.0408 | 0.024843 | 23.8678 | -0.012589 | 0.14295 | 4.2e-05 | 0.003374 | 0.003374 | 1.7415
2019-03-21 12:34 | 1553186038.376 | -69 | 0.20948 | 0.000457 | -4e-06 | 100.7558 | 101.0147 | 123.7346 | 21.2314 | 0.00194 | 0.001799 | -152.1142 | 0.23961 | 8.4019 | 0.003644 | 0.004924 | 0.004761 | 6.6e-05 | 0.00112 | 0.26377 | 1.3262 | 2.5481 | 99.38 | 20.96 | 37.376 | 68.4882 | 109.5672 | 65.0769 | 0.009341 | 1.0438 | 0.024859 | 23.8156 | -0.022956 | 0.14255 | -0.000251 | -0.018491 | -0.018491 | -0.31684
2019-03-21 12:34 | 1553186039.3763 | -68 | 0.20949 | 0.00046 | -4e-06 | 100.8416 | 100.8264 | 123.5903 | 21.2315 | 0.00195 | 0.001807 | -152.0889 | 0.23948 | 8.402 | 0.00366 | 0.004988 | 0.004823 | 6.4e-05 | 0.001117 | 0.25138 | 1.3868 | 2.5479 | 99.38 | 20.96 | 37.379 | 68.7245 | 109.9393 | 65.28 | 0.009342 | 1.0432 | 0.024837 | 23.8088 | 0.001352 | 0.14247 | -0.000279 | -0.020597 | -0.020597 | -0.28428
2019-03-21 12:34 | 1553186040.377 | -67 | 0.20943 | 0.000462 | -4e-06 | 100.8249 | 101.0093 | 123.4638 | 21.2319 | 0.002014 | 0.001808 | -152.1154 | 0.23919 | 8.4018 | 0.003612 | 0.004768 | 0.004583 | 6.5e-05 | 0.001118 | 0.29031 | 1.3656 | 2.5479 | 99.38 | 20.96 | 37.383 | 70.2541 | 108.7823 | 65.3115 | 0.009343 | 1.0431 | 0.024823 | 23.7977 | -0.016339 | 0.14229 | 6.7e-05 | 0.005278 | 0.005278 | 1.108
2019-03-21 12:34 | 1553186041.3765 | -66 | 0.20946 | 0.000458 | -5e-06 | 100.794 | 100.8928 | 123.5192 | 21.2315 | 0.001968 | 0.001811 | -152.1088 | 0.23931 | 8.4012 | 0.003642 | 0.00412 | 0.004006 | 6.9e-05 | 0.00112 | 0.30431 | 1.3555 | 2.548 | 99.38 | 20.96 | 37.384 | 69.1575 | 109.516 | 65.3628 | 0.009343 | 1.0429 | 0.024827 | 23.8049 | -0.008768 | 0.14236 | -9.8e-05 | -0.00704 | -0.00704 | -0.83119
2019-03-21 12:34 | 1553186042.3755 | -65 | 0.2095 | 0.000458 | -5e-06 | 100.7877 | 100.8441 | 123.3037 | 21.2327 | 0.002022 | 0.001841 | -152.1101 | 0.23986 | 8.4014 | 0.003601 | 0.005174 | 0.005013 | 6.8e-05 | 0.001121 | 0.28064 | 1.3582 | 2.5481 | 99.38 | 20.96 | 37.385 | 70.4528 | 108.5074 | 66.099 | 0.009343 | 1.0407 | 0.024778 | 23.81 | -0.005004 | 0.14269 | -0.000366 | -0.026965 | -0.026965 | -0.21749
2019-03-21 12:34 | 1553186043.3761 | -64 | 0.20949 | 0.000458 | -4e-06 | 100.8577 | 100.9684 | 123.3671 | 21.2309 | 0.001958 | 0.001804 | -152.0928 | 0.23947 | 8.4031 | 0.003686 | 0.00493 | 0.004748 | 6.9e-05 | 0.001118 | 0.25802 | 1.3454 | 2.548 | 99.38 | 20.96 | 37.387 | 68.9151 | 110.5706 | 65.2131 | 0.009344 | 1.0434 | 0.024817 | 23.785 | -0.009812 | 0.14246 | -0.000305 | -0.022481 | -0.022481 | -0.26044
2019-03-21 12:34 | 1553186044.3752 | -63 | 0.20944 | 0.000458 | -4e-06 | 100.8288 | 101.0341 | 123.5236 | 21.232 | 0.001968 | 0.001807 | -152.0921 | 0.23901 | 8.4006 | 0.003661 | 0.004315 | 0.004168 | 6.8e-05 | 0.001122 | 0.26217 | 1.3241 | 2.548 | 99.38 | 20.96 | 37.386 | 69.1548 | 109.9658 | 65.2695 | 0.009344 | 1.0432 | 0.024831 | 23.802 | -0.018197 | 0.14219 | 3.3e-05 | 0.00265 | 0.00265 | 2.2054
2019-03-21 12:34 | 1553186045.3753 | -62 | 0.20943 | 0.00046 | -4e-06 | 100.8395 | 101.0222 | 123.5955 | 21.2315 | 0.001919 | 0.00184 | -152.0914 | 0.23943 | 8.4021 | 0.003659 | 0.004865 | 0.004686 | 6.6e-05 | 0.001117 | 0.27256 | 1.3391 | 2.5481 | 99.38 | 20.96 | 37.387 | 67.9648 | 109.9291 | 66.0721 | 0.009344 | 1.0408 | 0.024809 | 23.8372 | -0.01619 | 0.14244 | 6.8e-05 | 0.005297 | 0.005297 | 1.1052
2019-03-21 12:34 | 1553186046.3753 | -61 | 0.20944 | 0.000459 | -5e-06 | 100.8482 | 101.0863 | 124.0086 | 21.232 | 0.001989 | 0.00189 | -152.0941 | 0.23889 | 8.4013 | 0.00367 | 0.004992 | 0.004775 | 7.1e-05 | 0.001119 | 0.25984 | 1.3213 | 2.548 | 99.38 | 20.96 | 37.388 | 69.6451 | 110.1914 | 67.2758 | 0.009344 | 1.0371 | 0.024806 | 23.9193 | -0.021091 | 0.14212 | -1.3e-05 | -0.00071 | -0.00071 | -8.2323
2019-03-21 12:34 | 1553186047.3764 | -60 | 0.20944 | 0.000455 | -5e-06 | 100.7722 | 100.9749 | 124.068 | 21.2333 | 0.00202 | 0.0018 | -152.1004 | 0.2393 | 8.4035 | 0.003735 | 0.004997 | 0.004825 | 6.2e-05 | 0.001115 | 0.25608 | 1.2954 | 2.5482 | 99.38 | 20.96 | 37.389 | 70.3967 | 111.7663 | 65.103 | 0.009344 | 1.0437 | 0.024891 | 23.8485 | -0.017972 | 0.14236 | -1.7e-05 | -0.001014 | -0.001014 | -5.7678
2019-03-21 12:34 | 1553186048.3765 | -59 | 0.20946 | 0.000461 | -4e-06 | 100.7221 | 100.807 | 124.1832 | 21.2322 | 0.001928 | 0.001868 | -152.095 | 0.23979 | 8.4006 | 0.003785 | 0.004924 | 0.00475 | 6.4e-05 | 0.001118 | 0.28368 | 1.3009 | 2.548 | 99.38 | 20.96 | 37.39 | 68.1831 | 112.9922 | 66.7477 | 0.009345 | 1.0387 | 0.024843 | 23.9176 | -0.00753 | 0.14265 | -0.000109 | -0.007836 | -0.007836 | -0.74823
2019-03-21 12:34 | 1553186049.3765 | -58 | 0.20944 | 0.000458 | -5e-06 | 100.7832 | 100.8703 | 123.8454 | 21.2321 | 0.001938 | 0.00183 | -152.1141 | 0.24067 | 8.403 | 0.003805 | 0.004864 | 0.004679 | 6.4e-05 | 0.00112 | 0.21588 | 1.2856 | 2.5479 | 99.38 | 20.96 | 37.391 | 68.4193 | 113.4757 | 65.8377 | 0.009345 | 1.0415 | 0.024842 | 23.853 | -0.007726 | 0.14317 | 8e-06 | 0.00091 | 0.00091 | 6.4632
2019-03-21 12:34 | 1553186050.3755 | -57 | 0.20945 | 0.000461 | -5e-06 | 100.8038 | 101.092 | 123.929 | 21.2329 | 0.001968 | 0.001881 | -152.1224 | 0.23976 | 8.4026 | 0.003748 | 0.004942 | 0.004771 | 6.8e-05 | 0.001123 | 0.27936 | 1.2888 | 2.548 | 99.38 | 20.96 | 37.393 | 69.1394 | 112.084 | 67.0706 | 0.009345 | 1.0377 | 0.024805 | 23.9044 | -0.025542 | 0.14264 | -7.7e-05 | -0.005435 | -0.005435 | -1.0786
2019-03-21 12:34 | 1553186051.377 | -56 | 0.20948 | 0.000457 | -5e-06 | 100.8056 | 100.9616 | 123.572 | 21.2324 | 0.001921 | 0.001821 | -152.1339 | 0.23862 | 8.4024 | 0.00369 | 0.004893 | 0.004679 | 6.9e-05 | 0.001116 | 0.24458 | 1.3173 | 2.5479 | 99.38 | 20.96 | 37.394 | 68.0321 | 110.6825 | 65.618 | 0.009346 | 1.0421 | 0.024823 | 23.8189 | -0.013834 | 0.14195 | -0.000259 | -0.019021 | -0.019021 | -0.30673
2019-03-21 12:34 | 1553186052.3769 | -55 | 0.20949 | 0.000457 | -4e-06 | 100.7797 | 100.8094 | 123.3981 | 21.2331 | 0.001962 | 0.001812 | -152.1098 | 0.24013 | 8.4013 | 0.003771 | 0.004419 | 0.004224 | 6.7e-05 | 0.001115 | 0.28152 | 1.3178 | 2.5479 | 99.38 | 20.96 | 37.395 | 69.0155 | 112.6386 | 65.4117 | 0.009346 | 1.0428 | 0.024813 | 23.7949 | -0.002636 | 0.14285 | -0.000295 | -0.021737 | -0.021737 | -0.2701
2019-03-21 12:34 | 1553186053.3767 | -54 | 0.2095 | 0.00046 | -5e-06 | 100.7904 | 100.8788 | 123.641 | 21.2335 | 0.001967 | 0.001882 | -152.1239 | 0.23928 | 8.4005 | 0.003693 | 0.005291 | 0.005163 | 6.7e-05 | 0.001114 | 0.27903 | 1.333 | 2.5481 | 99.38 | 20.96 | 37.396 | 69.1312 | 110.7616 | 67.0805 | 0.009346 | 1.0377 | 0.024776 | 23.877 | -0.007847 | 0.14235 | -0.000364 | -0.026787 | -0.026787 | -0.21841
2019-03-21 12:34 | 1553186054.3768 | -53 | 0.20945 | 0.000459 | -4e-06 | 100.7505 | 101.0045 | 123.5025 | 21.233 | 0.001946 | 0.001842 | -152.1167 | 0.23984 | 8.4047 | 0.003674 | 0.004418 | 0.004223 | 6.6e-05 | 0.001118 | 0.28512 | 1.3404 | 2.5479 | 99.38 | 20.96 | 37.395 | 68.6208 | 110.2989 | 66.1102 | 0.009346 | 1.0406 | 0.024798 | 23.8295 | -0.022527 | 0.14268 | -4.5e-05 | -0.003089 | -0.003089 | -1.8984
2019-03-21 12:34 | 1553186055.3768 | -52 | 0.20947 | 0.000455 | -4e-06 | 100.7716 | 100.7835 | 123.603 | 21.2335 | 0.001994 | 0.00189 | -152.1188 | 0.23941 | 8.3988 | 0.003648 | 0.005036 | 0.004828 | 6.4e-05 | 0.001116 | 0.29535 | 1.2857 | 2.5479 | 99.38 | 20.96 | 37.393 | 69.7848 | 109.6581 | 67.2803 | 0.009345 | 1.0371 | 0.024765 | 23.8803 | -0.001053 | 0.14242 | -0.000171 | -0.012448 | -0.012448 | -0.47024
2019-03-21 12:34 | 1553186056.3753 | -51 | 0.20951 | 0.00046 | -5e-06 | 100.6602 | 100.9626 | 123.1652 | 21.2334 | 0.002016 | 0.001821 | -152.1114 | 0.23968 | 8.4007 | 0.003631 | 0.00445 | 0.004264 | 6.7e-05 | 0.001122 | 0.28584 | 1.2885 | 2.5479 | 99.38 | 20.96 | 37.392 | 70.3016 | 109.2393 | 65.6056 | 0.009345 | 1.0422 | 0.024782 | 23.7793 | -0.026835 | 0.14259 | -0.000394 | -0.029024 | -0.029024 | -0.20191
2019-03-21 12:34 | 1553186057.3766 | -50 | 0.20943 | 0.000457 | -4e-06 | 100.7686 | 101.0318 | 123.1854 | 21.2337 | 0.002036 | 0.001842 | -152.1103 | 0.24009 | 8.4019 | 0.003665 | 0.00487 | 0.004807 | 6.4e-05 | 0.001122 | 0.29175 | 1.3285 | 2.5481 | 99.38 | 20.96 | 37.39 | 70.7807 | 110.0669 | 66.1289 | 0.009345 | 1.0406 | 0.024765 | 23.7996 | -0.023326 | 0.14283 | 5.8e-05 | 0.004549 | 0.004549 | 1.2904
2019-03-21 12:34 | 1553186058.3749 | -49 | 0.20952 | 0.000458 | -4e-06 | 100.7332 | 100.9255 | 122.9537 | 21.2338 | 0.001977 | 0.001883 | -152.1011 | 0.23951 | 8.4021 | 0.003555 | 0.005142 | 0.004941 | 6.4e-05 | 0.001129 | 0.30541 | 1.3008 | 2.548 | 99.38 | 20.96 | 37.385 | 69.3686 | 107.3882 | 67.0972 | 0.009343 | 1.0376 | 0.024707 | 23.8111 | -0.017061 | 0.14249 | -0.000489 | -0.036013 | -0.036013 | -0.16261
2019-03-21 12:34 | 1553186059.3763 | -48 | 0.20955 | 0.000455 | -4e-06 | 100.751 | 101.0318 | 123.3401 | 21.2338 | 0.001952 | 0.001821 | -152.1007 | 0.23945 | 8.4052 | 0.003636 | 0.004774 | 0.004734 | 6.6e-05 | 0.001126 | 0.28622 | 1.3214 | 2.548 | 99.38 | 20.96 | 37.384 | 68.7556 | 109.3692 | 65.606 | 0.009343 | 1.0422 | 0.0248 | 23.7962 | -0.024887 | 0.14245 | -0.000643 | -0.047616 | -0.047616 | -0.12296
2019-03-21 12:34 | 1553186060.3751 | -47 | 0.20943 | 0.000461 | -5e-06 | 100.7454 | 100.8291 | 123.3802 | 21.2337 | 0.00198 | 0.001885 | -152.1001 | 0.24019 | 8.4036 | 0.003594 | 0.005174 | 0.004995 | 6.8e-05 | 0.001126 | 0.26576 | 1.3465 | 2.548 | 99.38 | 20.96 | 37.381 | 69.4403 | 108.3483 | 67.1485 | 0.009342 | 1.0375 | 0.024748 | 23.8542 | -0.00743 | 0.14289 | 6.8e-05 | 0.005389 | 0.005389 | 1.0899
2019-03-21 12:34 | 1553186061.3752 | -46 | 0.20944 | 0.000456 | -4e-06 | 100.7433 | 100.9497 | 123.5703 | 21.2341 | 0.001955 | 0.001789 | -152.083 | 0.24023 | 8.402 | 0.003551 | 0.004813 | 0.004649 | 6.5e-05 | 0.001126 | 0.27417 | 1.2836 | 2.5481 | 99.38 | 20.96 | 37.373 | 68.8285 | 107.2916 | 64.8406 | 0.00934 | 1.0445 | 0.024851 | 23.7914 | -0.018305 | 0.14291 | 6e-06 | 0.000735 | 0.000735 | 7.9901
2019-03-21 12:34 | 1553186062.3762 | -45 | 0.20952 | 0.000455 | -4e-06 | 100.7317 | 100.6016 | 123.8133 | 21.2347 | 0.001929 | 0.001817 | -152.1028 | 0.24039 | 8.4021 | 0.003589 | 0.005361 | 0.00522 | 6.6e-05 | 0.001132 | 0.31293 | 1.2599 | 2.548 | 99.38 | 20.96 | 37.369 | 68.2151 | 108.2177 | 65.5266 | 0.009339 | 1.0424 | 0.02485 | 23.839 | 0.011558 | 0.14301 | -0.000503 | -0.037305 | -0.037305 | -0.15756
2019-03-21 12:34 | 1553186063.3766 | -44 | 0.20949 | 0.000457 | -5e-06 | 100.6941 | 100.8989 | 123.487 | 21.2337 | 0.00191 | 0.001888 | -152.0826 | 0.2392 | 8.3982 | 0.003538 | 0.005076 | 0.00493 | 6.4e-05 | 0.001126 | 0.26913 | 1.2569 | 2.548 | 99.38 | 20.96 | 37.359 | 67.7564 | 106.9786 | 67.2383 | 0.009337 | 1.0372 | 0.024755 | 23.8676 | -0.018173 | 0.1423 | -0.000277 | -0.020272 | -0.020272 | -0.28851
2019-03-21 12:34 | 1553186064.3753 | -43 | 0.20949 | 0.000463 | -5e-06 | 100.667 | 100.8257 | 123.5185 | 21.2337 | 0.001915 | 0.001813 | -152.0964 | 0.23914 | 8.4023 | 0.003615 | 0.005198 | 0.004961 | 6.4e-05 | 0.001132 | 0.25857 | 1.2982 | 2.5481 | 99.38 | 20.96 | 37.354 | 67.8788 | 108.8449 | 65.4352 | 0.009336 | 1.0427 | 0.024824 | 23.8074 | -0.014088 | 0.14226 | -0.000318 | -0.023377 | -0.023377 | -0.25012
2019-03-21 12:34 | 1553186065.3764 | -42 | 0.20945 | 0.000461 | -4e-06 | 100.7528 | 100.9926 | 123.7084 | 21.2349 | 0.001927 | 0.001786 | -152.0976 | 0.23946 | 8.3999 | 0.003576 | 0.004815 | 0.00463 | 6.6e-05 | 0.001133 | 0.30215 | 1.296 | 2.5481 | 99.38 | 20.96 | 37.345 | 68.1632 | 107.9122 | 64.7659 | 0.009333 | 1.0448 | 0.024868 | 23.8021 | -0.02126 | 0.14245 | -3.7e-05 | -0.002516 | -0.002516 | -2.3271
2019-03-21 12:34 | 1553186066.3768 | -41 | 0.2094 | 0.000455 | -5e-06 | 100.7329 | 100.9017 | 124.1464 | 21.234 | 0.00195 | 0.001852 | -152.0993 | 0.23844 | 8.4023 | 0.003583 | 0.004775 | 0.004664 | 6.4e-05 | 0.00113 | 0.26969 | 1.302 | 2.548 | 99.38 | 20.96 | 37.34 | 68.7246 | 108.0752 | 66.367 | 0.009332 | 1.0398 | 0.024853 | 23.9006 | -0.014974 | 0.14185 | 0.000248 | 0.018826 | 0.018826 | 0.30968
2019-03-21 12:34 | 1553186067.3764 | -40 | 0.20954 | 0.000456 | -4e-06 | 100.7251 | 100.9054 | 124.1762 | 21.234 | 0.00194 | 0.001782 | -152.0959 | 0.23825 | 8.4014 | 0.003559 | 0.004779 | 0.004563 | 6.7e-05 | 0.00113 | 0.27129 | 1.3242 | 2.5479 | 99.38 | 20.96 | 37.334 | 68.4733 | 107.4915 | 64.6836 | 0.009331 | 1.045 | 0.024918 | 23.8441 | -0.015994 | 0.14173 | -0.000617 | -0.04593 | -0.04593 | -0.12683
2019-03-21 12:34 | 1553186068.3758 | -39 | 0.20945 | 0.000456 | -4e-06 | 100.732 | 100.6544 | 124.5336 | 21.2336 | 0.001919 | 0.001835 | -152.1221 | 0.23887 | 8.4022 | 0.003579 | 0.004767 | 0.004622 | 6.2e-05 | 0.001132 | 0.30071 | 1.283 | 2.5479 | 99.38 | 20.97 | 37.323 | 67.9715 | 107.9835 | 65.9484 | 0.009334 | 1.0411 | 0.024907 | 23.9231 | 0.006892 | 0.1421 | -3.4e-05 | -0.002237 | -0.002237 | -2.6104
2019-03-21 12:34 | 1553186069.3765 | -38 | 0.20947 | 0.00046 | -5e-06 | 100.7868 | 101.0525 | 124.4682 | 21.2346 | 0.001937 | 0.001866 | -152.1055 | 0.23877 | 8.4035 | 0.003565 | 0.005493 | 0.005377 | 6.5e-05 | 0.001137 | 0.27256 | 1.3037 | 2.548 | 99.38 | 20.97 | 37.318 | 68.4069 | 107.6465 | 66.6985 | 0.009332 | 1.0388 | 0.024873 | 23.9433 | -0.023552 | 0.14204 | -0.000201 | -0.014713 | -0.014713 | -0.39678
2019-03-21 12:34 | 1553186070.377 | -37 | 0.20949 | 0.000461 | -4e-06 | 100.7496 | 101.0485 | 124.578 | 21.2348 | 0.002009 | 0.001856 | -152.0812 | 0.23958 | 8.4014 | 0.003645 | 0.004342 | 0.00418 | 6.1e-05 | 0.001138 | 0.25962 | 1.3286 | 2.5479 | 99.38 | 20.97 | 37.313 | 70.1353 | 109.5768 | 66.4662 | 0.009331 | 1.0395 | 0.024892 | 23.9456 | -0.026496 | 0.14252 | -0.000319 | -0.023603 | -0.023603 | -0.24817
2019-03-21 12:34 | 1553186071.3766 | -36 | 0.20941 | 0.000456 | -4e-06 | 100.8017 | 100.7961 | 124.5775 | 21.2343 | 0.001993 | 0.001856 | -152.0688 | 0.24 | 8.3987 | 0.003556 | 0.004595 | 0.004532 | 6e-05 | 0.001138 | 0.26521 | 1.2874 | 2.548 | 99.38 | 20.97 | 37.303 | 69.7606 | 107.4329 | 66.4595 | 0.009329 | 1.0396 | 0.024893 | 23.9453 | 0.000499 | 0.14278 | 0.00017 | 0.012955 | 0.012955 | 0.45296
2019-03-21 12:35 | 1553186072.3761 | -35 | 0.20948 | 0.000459 | -4e-06 | 100.7269 | 100.6724 | 124.4998 | 21.2343 | 0.001965 | 0.001829 | -152.0517 | 0.23967 | 8.4001 | 0.003575 | 0.004825 | 0.004509 | 6.3e-05 | 0.00114 | 0.27721 | 1.3277 | 2.5481 | 99.38 | 20.97 | 37.298 | 69.0743 | 107.8919 | 65.8016 | 0.009327 | 1.0416 | 0.024909 | 23.9147 | 0.004838 | 0.14258 | -0.000214 | -0.015738 | -0.015738 | -0.37235
2019-03-21 12:35 | 1553186073.3762 | -34 | 0.20953 | 0.00046 | -5e-06 | 100.7621 | 100.8744 | 124.6757 | 21.235 | 0.001961 | 0.001845 | -152.0589 | 0.23971 | 8.4046 | 0.003614 | 0.004273 | 0.004158 | 6.4e-05 | 0.001146 | 0.27704 | 1.326 | 2.5479 | 99.38 | 20.97 | 37.291 | 68.9931 | 108.8272 | 66.191 | 0.009326 | 1.0404 | 0.024912 | 23.9453 | -0.009967 | 0.14261 | -0.000556 | -0.041268 | -0.041268 | -0.14203
2019-03-21 12:35 | 1553186074.3748 | -33 | 0.20946 | 0.000458 | -4e-06 | 100.7189 | 100.9418 | 124.4825 | 21.235 | 0.001948 | 0.001834 | -152.058 | 0.23969 | 8.4022 | 0.003556 | 0.004836 | 0.004624 | 6.5e-05 | 0.00115 | 0.26017 | 1.3351 | 2.548 | 99.38 | 20.97 | 37.288 | 68.6623 | 107.4318 | 65.9233 | 0.009325 | 1.0412 | 0.024903 | 23.9173 | -0.01978 | 0.14259 | -0.000121 | -0.008759 | -0.008759 | -0.66911
2019-03-21 12:35 | 1553186075.3751 | -32 | 0.20943 | 0.000459 | -4e-06 | 100.7999 | 100.9224 | 124.3738 | 21.2351 | 0.001953 | 0.001908 | -152.0541 | 0.23879 | 8.4048 | 0.003595 | 0.00488 | 0.004709 | 6.3e-05 | 0.001149 | 0.27041 | 1.2857 | 2.548 | 99.38 | 20.97 | 37.285 | 68.8007 | 108.3718 | 67.7039 | 0.009324 | 1.0358 | 0.024827 | 23.9696 | -0.010865 | 0.14205 | 4.3e-05 | 0.003473 | 0.003473 | 1.6809
2019-03-21 12:35 | 1553186076.375 | -31 | 0.20946 | 0.000454 | -5e-06 | 100.7707 | 100.8397 | 124.3785 | 21.2348 | 0.001962 | 0.001801 | -152.057 | 0.23937 | 8.4046 | 0.003556 | 0.004431 | 0.004224 | 6.2e-05 | 0.001148 | 0.26322 | 1.2593 | 2.548 | 99.38 | 20.97 | 37.279 | 69.0131 | 107.4159 | 65.1442 | 0.009323 | 1.0436 | 0.024921 | 23.8798 | -0.006119 | 0.1424 | -9.2e-05 | -0.00661 | -0.00661 | -0.88546
2019-03-21 12:35 | 1553186077.3753 | -30 | 0.20944 | 0.00046 | -4e-06 | 100.7356 | 100.9728 | 124.2729 | 21.2347 | 0.001958 | 0.001778 | -152.0372 | 0.23873 | 8.4017 | 0.003514 | 0.005232 | 0.00505 | 6.4e-05 | 0.001156 | 0.29142 | 1.2814 | 2.548 | 99.38 | 20.97 | 37.278 | 68.9055 | 106.4128 | 64.5828 | 0.009322 | 1.0453 | 0.024931 | 23.8499 | -0.021044 | 0.14202 | 3.1e-05 | 0.002586 | 0.002586 | 2.2573
2019-03-21 12:35 | 1553186078.376 | -29 | 0.20946 | 0.000456 | -5e-06 | 100.782 | 100.7498 | 124.1349 | 21.2353 | 0.001992 | 0.001863 | -152.0359 | 0.23969 | 8.4029 | 0.00359 | 0.005254 | 0.00506 | 6.2e-05 | 0.001152 | 0.26272 | 1.3111 | 2.5481 | 99.38 | 20.97 | 37.275 | 69.7404 | 108.2522 | 66.6359 | 0.009322 | 1.039 | 0.024842 | 23.909 | 0.002865 | 0.14259 | -0.000127 | -0.0092 | -0.0092 | -0.63704
2019-03-21 12:35 | 1553186079.3761 | -28 | 0.20942 | 0.000459 | -5e-06 | 100.7889 | 100.7389 | 123.9877 | 21.2342 | 0.001959 | 0.001862 | -152.0383 | 0.24016 | 8.4036 | 0.003608 | 0.004985 | 0.004806 | 6.1e-05 | 0.001151 | 0.23883 | 1.3622 | 2.5482 | 99.38 | 20.97 | 37.274 | 68.9377 | 108.675 | 66.6154 | 0.009321 | 1.0391 | 0.024828 | 23.8941 | 0.004441 | 0.14287 | 0.000148 | 0.011305 | 0.011305 | 0.51942
2019-03-21 12:35 | 1553186080.3755 | -27 | 0.20944 | 0.000457 | -5e-06 | 100.7522 | 101.0917 | 123.8515 | 21.2355 | 0.001948 | 0.001818 | -152.0354 | 0.23854 | 8.402 | 0.003606 | 0.004874 | 0.00465 | 6.5e-05 | 0.001155 | 0.25713 | 1.3132 | 2.548 | 99.38 | 20.97 | 37.272 | 68.67 | 108.6282 | 65.5563 | 0.009321 | 1.0423 | 0.024853 | 23.8437 | -0.030086 | 0.14191 | 4.4e-05 | 0.003581 | 0.003581 | 1.6285
2019-03-21 12:35 | 1553186081.3751 | -26 | 0.20945 | 0.000456 | -5e-06 | 100.8309 | 100.9841 | 123.8065 | 21.2357 | 0.002019 | 0.001858 | -152.0564 | 0.23926 | 8.4009 | 0.003592 | 0.004642 | 0.004501 | 6.5e-05 | 0.001156 | 0.27024 | 1.3205 | 2.5482 | 99.38 | 20.97 | 37.269 | 70.3779 | 108.2872 | 66.5171 | 0.00932 | 1.0394 | 0.024813 | 23.8731 | -0.013578 | 0.14233 | -2.4e-05 | -0.001485 | -0.001485 | -3.9391
2019-03-21 12:35 | 1553186082.3766 | -25 | 0.20943 | 0.000459 | -5e-06 | 100.7722 | 100.8782 | 123.8473 | 21.2356 | 0.00189 | 0.001821 | -152.0647 | 0.23992 | 8.4052 | 0.003565 | 0.004947 | 0.004791 | 6.4e-05 | 0.001154 | 0.24331 | 1.3458 | 2.548 | 99.38 | 20.97 | 37.268 | 67.2691 | 107.6489 | 65.6249 | 0.00932 | 1.0421 | 0.02485 | 23.8457 | -0.0094 | 0.14273 | 0.000101 | 0.00785 | 0.00785 | 0.74724
2019-03-21 12:35 | 1553186083.3767 | -24 | 0.20945 | 0.000459 | -4e-06 | 100.8017 | 100.9064 | 123.7419 | 21.2344 | 0.00195 | 0.001869 | -152.0493 | 0.23984 | 8.3993 | 0.003644 | 0.005549 | 0.005357 | 6.3e-05 | 0.001153 | 0.27903 | 1.3318 | 2.548 | 99.38 | 20.97 | 37.268 | 68.7103 | 109.5588 | 66.7724 | 0.00932 | 1.0386 | 0.024798 | 23.8759 | -0.009288 | 0.14268 | -2.9e-05 | -0.001894 | -0.001894 | -3.0966
2019-03-21 12:35 | 1553186084.3753 | -23 | 0.20945 | 0.000458 | -4e-06 | 100.774 | 100.8478 | 123.8537 | 21.235 | 0.001927 | 0.001875 | -152.0341 | 0.23845 | 8.3999 | 0.003619 | 0.005797 | 0.005601 | 6.1e-05 | 0.001154 | 0.27831 | 1.3396 | 2.5481 | 99.38 | 20.96 | 37.265 | 68.1734 | 108.9462 | 66.9089 | 0.009313 | 1.0382 | 0.024804 | 23.8915 | -0.006553 | 0.14185 | -7.4e-05 | -0.005254 | -0.005254 | -1.1097
2019-03-21 12:35 | 1553186085.3764 | -22 | 0.20945 | 0.000457 | -5e-06 | 100.8178 | 101.0287 | 123.7185 | 21.2357 | 0.00199 | 0.001811 | -152.0354 | 0.23961 | 8.4024 | 0.003586 | 0.004647 | 0.004531 | 6.9e-05 | 0.001157 | 0.2844 | 1.3319 | 2.5481 | 99.38 | 20.96 | 37.264 | 69.6761 | 108.1494 | 65.3815 | 0.009313 | 1.0429 | 0.024846 | 23.8247 | -0.018692 | 0.14255 | -2.7e-05 | -0.001742 | -0.001742 | -3.3634
2019-03-21 12:35 | 1553186086.3754 | -21 | 0.20945 | 0.000458 | -5e-06 | 100.8747 | 101.0696 | 123.7393 | 21.235 | 0.001972 | 0.001817 | -152.0209 | 0.23883 | 8.399 | 0.003535 | 0.004712 | 0.004556 | 6.7e-05 | 0.001157 | 0.31509 | 1.3177 | 2.5481 | 99.38 | 20.96 | 37.264 | 69.2558 | 106.9066 | 65.5137 | 0.009313 | 1.0425 | 0.024843 | 23.8314 | -0.017263 | 0.14208 | -7.6e-05 | -0.005379 | -0.005379 | -1.0857
2019-03-21 12:35 | 1553186087.3749 | -20 | 0.20945 | 0.000457 | -4e-06 | 100.7913 | 101.0955 | 123.8299 | 21.2351 | 0.001972 | 0.001876 | -152.0267 | 0.23925 | 8.4037 | 0.003549 | 0.004771 | 0.004569 | 6.3e-05 | 0.001158 | 0.29717 | 1.3197 | 2.5482 | 99.38 | 20.96 | 37.264 | 69.2576 | 107.2635 | 66.9378 | 0.009313 | 1.0381 | 0.0248 | 23.8902 | -0.026953 | 0.14233 | -6.3e-05 | -0.004393 | -0.004393 | -1.3315
2019-03-21 12:35 | 1553186088.3756 | -19 | 0.20953 | 0.000455 | -4e-06 | 100.8038 | 101.0832 | 123.9694 | 21.2354 | 0.001923 | 0.001792 | -152.0297 | 0.23846 | 8.4062 | 0.003573 | 0.004811 | 0.004644 | 6.2e-05 | 0.001158 | 0.23938 | 1.3513 | 2.5481 | 99.38 | 20.96 | 37.264 | 68.0795 | 107.8396 | 64.9121 | 0.009313 | 1.0443 | 0.024889 | 23.8323 | -0.024758 | 0.14186 | -0.00052 | -0.038628 | -0.038628 | -0.15094
2019-03-21 12:35 | 1553186089.3757 | -18 | 0.20946 | 0.00046 | -4e-06 | 100.8404 | 101.1217 | 124.3412 | 21.2354 | 0.001962 | 0.00181 | -152.0627 | 0.23908 | 8.4045 | 0.003554 | 0.004832 | 0.00466 | 6.6e-05 | 0.001159 | 0.24923 | 1.2784 | 2.5479 | 99.38 | 20.96 | 37.263 | 68.9973 | 107.3684 | 65.3625 | 0.009313 | 1.0429 | 0.024909 | 23.8839 | -0.024912 | 0.14223 | -8e-05 | -0.005779 | -0.005779 | -1.0115
2019-03-21 12:35 | 1553186090.3747 | -17 | 0.20944 | 0.000461 | -4e-06 | 100.8264 | 101.0454 | 124.7752 | 21.2356 | 0.001956 | 0.001813 | -152.0855 | 0.23871 | 8.4025 | 0.003453 | 0.005004 | 0.004898 | 6.6e-05 | 0.001155 | 0.24298 | 1.304 | 2.5479 | 99.38 | 20.96 | 37.263 | 68.865 | 104.9351 | 65.4353 | 0.009313 | 1.0427 | 0.02495 | 23.9282 | -0.019403 | 0.14201 | 3e-05 | 0.002503 | 0.002503 | 2.3322
2019-03-21 12:35 | 1553186091.3757 | -16 | 0.20943 | 0.000456 | -5e-06 | 100.8461 | 100.9275 | 125.0587 | 21.236 | 0.001938 | 0.001798 | -152.1011 | 0.23881 | 8.4015 | 0.003512 | 0.004747 | 0.004638 | 6.2e-05 | 0.001149 | 0.25697 | 1.3129 | 2.548 | 99.38 | 20.96 | 37.263 | 68.4286 | 106.3495 | 65.0611 | 0.009313 | 1.0439 | 0.024992 | 23.9421 | -0.007221 | 0.14207 | 8.8e-05 | 0.006904 | 0.006904 | 0.84571
2019-03-21 12:35 | 1553186092.376 | -15 | 0.20944 | 0.000456 | -5e-06 | 100.8017 | 101.042 | 125.3038 | 21.2348 | 0.001939 | 0.001821 | -152.1144 | 0.23791 | 8.4034 | 0.003523 | 0.00439 | 0.004196 | 6.4e-05 | 0.001143 | 0.25841 | 1.3186 | 2.548 | 99.38 | 20.96 | 37.263 | 68.4626 | 106.6111 | 65.6109 | 0.009313 | 1.0422 | 0.024996 | 23.985 | -0.021296 | 0.14153 | -2e-06 | 0.000169 | 0.000169 | 34.4508
2019-03-21 12:35 | 1553186093.3747 | -14 | 0.20943 | 0.000459 | -4e-06 | 100.796 | 100.8931 | 125.7181 | 21.2357 | 0.001933 | 0.001809 | -152.1174 | 0.23844 | 8.4025 | 0.003619 | 0.005043 | 0.00481 | 6.2e-05 | 0.001143 | 0.2724 | 1.2808 | 2.548 | 99.38 | 20.96 | 37.264 | 68.3081 | 108.957 | 65.3308 | 0.009313 | 1.043 | 0.025048 | 24.0147 | -0.008613 | 0.14185 | 5.9e-05 | 0.004689 | 0.004689 | 1.2434
2019-03-21 12:35 | 1553186094.3759 | -13 | 0.20943 | 0.000459 | -5e-06 | 100.8372 | 100.7695 | 125.9479 | 21.2359 | 0.001926 | 0.001798 | -152.1306 | 0.23908 | 8.4002 | 0.003632 | 0.004899 | 0.004789 | 6.3e-05 | 0.00114 | 0.25625 | 1.3161 | 2.548 | 99.38 | 20.96 | 37.265 | 68.1453 | 109.2789 | 65.0548 | 0.009313 | 1.0439 | 0.025081 | 24.0268 | 0.006002 | 0.14223 | 7.5e-05 | 0.005953 | 0.005953 | 0.98194
2019-03-21 12:35 | 1553186095.376 | -12 | 0.20946 | 0.000459 | -5e-06 | 100.8094 | 101.0273 | 126.2762 | 21.2363 | 0.00195 | 0.001765 | -152.1449 | 0.23993 | 8.4033 | 0.003676 | 0.005091 | 0.004845 | 6e-05 | 0.001139 | 0.2656 | 1.3455 | 2.5482 | 99.38 | 20.96 | 37.27 | 68.7112 | 110.3352 | 64.2744 | 0.009315 | 1.0463 | 0.025143 | 24.0304 | -0.019312 | 0.14274 | -0.000109 | -0.007968 | -0.007968 | -0.73624
2019-03-21 12:35 | 1553186096.3749 | -11 | 0.20944 | 0.000462 | -4e-06 | 100.858 | 100.9766 | 126.2612 | 21.2365 | 0.001944 | 0.001706 | -140.6112 | 0.23972 | 8.3997 | 0.003697 | 0.005566 | 0.005337 | 6.3e-05 | 0.001139 | 0.30702 | 1.3163 | 2.5481 | 99.38 | 20.96 | 37.272 | 68.5804 | 110.8577 | 62.8613 | 0.009315 | 1.0507 | 0.025194 | 23.9786 | -0.01051 | 0.14261 | -2.4e-05 | -0.001583 | -0.001583 | -3.703
2019-03-21 12:35 | 1553186097.3751 | -10 | 0.20947 | 0.000461 | -5e-06 | 100.8074 | 100.9524 | 126.3077 | 21.2365 | 0.001926 | 0.00179 | -99.0948 | 0.23916 | 8.4015 | 0.003715 | 0.004832 | 0.004684 | 6.3e-05 | 0.00114 | 0.30973 | 1.2952 | 2.548 | 99.38 | 20.96 | 37.274 | 68.1414 | 111.2925 | 64.8756 | 0.009316 | 1.0444 | 0.025124 | 24.0548 | -0.01286 | 0.14228 | -0.000144 | -0.010541 | -0.010541 | -0.55474
2019-03-21 12:35 | 1553186098.3754 | -9 | 0.20944 | 0.00047 | -4e-06 | 100.8556 | 101.0859 | 126.5145 | 21.2363 | 0.001864 | 0.001756 | -57.4972 | 0.24021 | 8.4017 | 0.003735 | 0.00483 | 0.0047 | 6.5e-05 | 0.001139 | 0.29662 | 1.3044 | 2.548 | 99.38 | 20.96 | 37.278 | 66.6533 | 111.7668 | 64.0511 | 0.009317 | 1.047 | 0.025175 | 24.0451 | -0.0204 | 0.1429 | -1.7e-05 | -0.001073 | -0.001073 | -5.4735
2019-03-21 12:35 | 1553186099.3761 | -8 | 0.20946 | 0.000468 | -4e-06 | 100.8047 | 100.9017 | 126.2807 | 21.2353 | 0.001936 | 0.001798 | -15.8748 | 0.23995 | 8.4039 | 0.003753 | 0.004545 | 0.004307 | 6e-05 | 0.001138 | 0.26985 | 1.323 | 2.548 | 99.38 | 20.96 | 37.28 | 68.3818 | 112.2133 | 65.0556 | 0.009317 | 1.0439 | 0.025114 | 24.0586 | -0.008601 | 0.14275 | -9e-05 | -0.006485 | -0.006485 | -0.90469
2019-03-21 12:35 | 1553186100.3756 | -7 | 0.20946 | 0.000463 | -5e-06 | 100.8985 | 101.138 | 126.0504 | 21.2367 | 0.001899 | 0.001816 | 25.7678 | 0.23855 | 8.4036 | 0.003658 | 0.005576 | 0.005449 | 6.3e-05 | 0.001138 | 0.31277 | 1.3313 | 2.5481 | 99.38 | 20.96 | 37.284 | 67.5074 | 109.8917 | 65.5052 | 0.009318 | 1.0425 | 0.025075 | 24.0526 | -0.021206 | 0.14192 | -0.000129 | -0.009403 | -0.009403 | -0.62029
2019-03-21 12:35 | 1553186101.3752 | -6 | 0.20946 | 0.000468 | -4e-06 | 100.9137 | 101.0277 | 125.8629 | 21.2367 | 0.001902 | 0.001827 | 67.3761 | 0.2483 | 8.4031 | 0.003742 | 0.005031 | 0.004822 | 6.2e-05 | 0.001138 | 0.26128 | 1.2929 | 2.548 | 99.38 | 20.96 | 37.285 | 67.5696 | 111.9533 | 65.7633 | 0.009318 | 1.0417 | 0.025046 | 24.0439 | -0.010094 | 0.14771 | -0.000112 | -0.008163 | -0.008163 | -0.74377
2019-03-21 12:35 | 1553186102.3762 | -5 | 0.20943 | 0.00046 | -4e-06 | 100.7922 | 101.0314 | 125.535 | 21.2361 | 0.001912 | 0.001747 | 263.1572 | 0.2716 | 8.4018 | 0.003708 | 0.005713 | 0.005659 | 6e-05 | 0.001136 | 0.30326 | 1.3141 | 2.5481 | 99.39 | 20.96 | 37.288 | 67.8095 | 111.1083 | 63.8307 | 0.009318 | 1.0477 | 0.025085 | 23.944 | -0.021207 | 0.16157 | 9.4e-05 | 0.007327 | 0.007327 | 0.90634
2019-03-21 12:35 | 1553186103.3757 | -4 | 0.20946 | 0.000464 | -4e-06 | 100.8401 | 101.0154 | 125.1918 | 21.2363 | 0.001896 | 0.001804 | 263.8077 | 0.27143 | 8.3989 | 0.003729 | 0.004663 | 0.004615 | 6.3e-05 | 0.001134 | 0.2943 | 1.2919 | 2.5479 | 99.39 | 20.96 | 37.293 | 67.4343 | 111.6293 | 65.2151 | 0.009319 | 1.0434 | 0.025 | 23.9603 | -0.015534 | 0.16147 | -8.8e-05 | -0.006369 | -0.006369 | -1.042
2019-03-21 12:35 | 1553186104.3753 | -3 | 0.20939 | 0.000459 | -4e-06 | 100.788 | 100.9412 | 124.8721 | 21.2362 | 0.001864 | 0.001758 | 263.9024 | 0.27182 | 8.4013 | 0.003736 | 0.005331 | 0.005225 | 6.3e-05 | 0.001128 | 0.25442 | 1.2966 | 2.548 | 99.39 | 20.96 | 37.296 | 66.6424 | 111.8037 | 64.1004 | 0.00932 | 1.0468 | 0.025009 | 23.8902 | -0.013583 | 0.16171 | 0.000295 | 0.022431 | 0.022431 | 0.29629
2019-03-21 12:35 | 1553186105.3745 | -2 | 0.20945 | 0.000465 | -4e-06 | 100.8696 | 101.0788 | 124.7083 | 21.2368 | 0.001902 | 0.001829 | 264.2173 | 0.27154 | 8.4036 | 0.003803 | 0.005059 | 0.004973 | 6.3e-05 | 0.001126 | 0.24978 | 1.2718 | 2.548 | 99.39 | 20.96 | 37.297 | 67.5634 | 113.4275 | 65.8096 | 0.00932 | 1.0416 | 0.02493 | 23.935 | -0.018525 | 0.16154 | -4.7e-05 | -0.003214 | -0.003214 | -2.0659
2019-03-21 12:35 | 1553186106.3755 | -1 | 0.20944 | 0.000462 | -4e-06 | 100.8625 | 101.0382 | 124.7172 | 21.2362 | 0.00189 | 0.001804 | 263.8756 | 0.27228 | 8.4023 | 0.003766 | 0.004945 | 0.004919 | 6.3e-05 | 0.001125 | 0.27455 | 1.3303 | 2.5479 | 99.39 | 20.96 | 37.299 | 67.2806 | 112.5414 | 65.2158 | 0.009321 | 1.0434 | 0.024952 | 23.9149 | -0.015572 | 0.16198 | 3.3e-05 | 0.002701 | 0.002701 | 2.4648
2019-03-21 12:35 | 1553186107.3745 | 0 | 0.20944 | 0.000458 | -4e-06 | 100.8521 | 100.8601 | 124.3804 | 21.236 | 0.001926 | 0.001846 | 264.1583 | 0.27094 | 8.4043 | 0.00377 | 0.005276 | 0.005131 | 6e-05 | 0.001127 | 0.29358 | 1.3161 | 2.706 | 99.39 | 20.96 | 37.3 | 68.1503 | 112.6154 | 66.2317 | 0.009321 | 1.0403 | 0.024881 | 23.9184 | -0.000714 | 0.16118 | 0 | 0.000285 | 0.000285 | 23.2176
2019-03-21 12:35 | 1553186108.3747 | 1 | 0.20947 | 0.00046 | -4e-06 | 99.2345 | 93.185 | 123.5368 | 21.2112 | 0.001917 | 0.001714 | 264.2306 | 0.27314 | 8.404 | 0.003763 | 0.005331 | 0.007464 | 6.1e-05 | 0.001124 | 0.30165 | 1.3461 | 2.5947 | 99.39 | 20.96 | 37.304 | 67.8968 | 112.4321 | 63.0177 | 0.009322 | 1.0502 | 0.024915 | 23.724 | 0.56261 | 0.16249 | -0.000169 | -0.01244 | -0.01244 | -0.53685
2019-03-21 12:35 | 1553186109.3746 | 2 | 0.20939 | 0.000465 | -4e-06 | 91.5455 | 73.2684 | 122.6171 | 21.2089 | 0.001919 | 0.001768 | 264.2491 | 0.27215 | 8.4029 | 0.00374 | 0.006921 | 0.010442 | 6.5e-05 | 0.001151 | 0.27848 | 1.2842 | 2.5481 | 99.39 | 20.96 | 37.305 | 67.9391 | 111.8757 | 64.3243 | 0.009322 | 1.0461 | 0.024774 | 23.6814 | 1.992 | 0.1619 | 0.000307 | 0.023052 | 0.023052 | 0.28867
2019-03-21 12:35 | 1553186110.3756 | 3 | 0.20942 | 0.000461 | -3e-06 | 92.9527 | 68.9137 | 121.1961 | 21.2125 | 0.00187 | 0.001808 | 264.408 | 0.27128 | 8.4022 | 0.00367 | 0.01175 | 0.013848 | 6.1e-05 | 0.00118 | 0.30182 | 1.2773 | 2.5481 | 99.39 | 20.96 | 37.306 | 66.7862 | 110.1596 | 65.2748 | 0.009323 | 1.0432 | 0.024595 | 23.5769 | 2.6765 | 0.16139 | 0.000136 | 0.01021 | 0.01021 | 0.64967
2019-03-21 12:35 | 1553186111.3758 | 4 | 0.20944 | 0.000466 | -1e-06 | 96.2205 | 74.6825 | 119.8134 | 21.2114 | 0.001865 | 0.001828 | 264.187 | 0.27242 | 8.403 | 0.003657 | 0.009599 | 0.012105 | 6e-05 | 0.001123 | 0.50312 | 1.3304 | 2.5484 | 99.38 | 20.96 | 37.309 | 66.6517 | 109.8595 | 65.7738 | 0.009324 | 1.0417 | 0.024437 | 23.4593 | 2.2665 | 0.16206 | -1.4e-05 | -0.000952 | -0.000952 | -6.9977
2019-03-21 12:35 | 1553186112.3748 | 5 | 0.2093 | 0.000467 | 2e-06 | 98.3827 | 84.2745 | 117.9586 | 21.2177 | 0.00195 | 0.001866 | 264.5943 | 0.27014 | 8.4036 | 0.003675 | 0.005667 | 0.007853 | 6.5e-05 | 0.001116 | 0.3533 | 1.329 | 2.5482 | 99.38 | 20.96 | 37.311 | 68.7011 | 110.3053 | 66.6776 | 0.009325 | 1.0389 | 0.024215 | 23.308 | 1.3845 | 0.1607 | 0.000875 | 0.063487 | 0.063487 | 0.10404
2019-03-21 12:35 | 1553186113.3749 | 6 | 0.20889 | 0.000466 | 8e-06 | 99.8286 | 91.9922 | 116.101 | 21.2192 | 0.001981 | 0.00192 | 264.1403 | 0.27179 | 8.4049 | 0.003667 | 0.00602 | 0.007735 | 7.4e-05 | 0.001112 | 0.37482 | 1.3316 | 2.5483 | 99.38 | 20.96 | 37.314 | 69.4542 | 110.1072 | 67.9958 | 0.009326 | 1.0349 | 0.023977 | 23.1686 | 0.73122 | 0.16169 | 0.003307 | 0.23744 | 0.23744 | 0.027987
2019-03-21 12:35 | 1553186114.3749 | 7 | 0.20815 | 0.000476 | 1.8e-05 | 100.2639 | 94.6965 | 114.0943 | 21.232 | 0.002084 | 0.002054 | 264.0781 | 0.27102 | 8.4017 | 0.003655 | 0.004749 | 0.005014 | 5.7e-05 | 0.001112 | 0.29087 | 1.3214 | 1.8755 | 99.38 | 20.96 | 37.316 | 71.9414 | 109.8309 | 71.2251 | 0.009326 | 1.0252 | 0.023657 | 23.0759 | 0.51099 | 0.16123 | 0.007736 | 0.54799 | 0.54799 | 0.012092
2019-03-21 12:35 | 1553186115.3749 | 8 | 0.20698 | 0.00051 | 2.9e-05 | 100.5553 | 94.128 | 111.552 | 21.236 | 0.00209 | 0.002092 | 263.9908 | 0.27224 | 8.4019 | 0.003729 | 0.005504 | 0.00542 | 5.8e-05 | 0.00111 | 0.25155 | 1.3308 | 2.5479 | 99.38 | 20.96 | 37.318 | 72.081 | 111.6251 | 72.1317 | 0.009327 | 1.0225 | 0.023361 | 22.8474 | 0.5908 | 0.16195 | 0.014741 | 1.0306 | 1.0306 | 0.006459
2019-03-21 12:35 | 1553186116.376 | 9 | 0.20528 | 0.00059 | 4.3e-05 | 100.7719 | 89.5117 | 109.1137 | 21.2359 | 0.002161 | 0.002229 | 263.9137 | 0.27148 | 8.4009 | 0.003782 | 0.005289 | 0.005256 | 6e-05 | 0.001111 | 0.25929 | 1.3042 | 2.548 | 99.38 | 20.96 | 37.321 | 73.788 | 112.9106 | 75.4296 | 0.009327 | 1.0128 | 0.022995 | 22.704 | 1.0598 | 0.16151 | 0.024795 | 1.7049 | 1.7049 | 0.003893
2019-03-21 12:35 | 1553186117.3751 | 10 | 0.2037 | 0.000867 | 5.8e-05 | 100.7969 | 85.5514 | 107.4629 | 21.2372 | 0.002252 | 0.002298 | 263.7419 | 0.27205 | 8.4027 | 0.003754 | 0.004865 | 0.004799 | 6.2e-05 | 0.001106 | 0.22987 | 1.3336 | 2.548 | 99.38 | 20.96 | 37.323 | 75.9866 | 112.2327 | 77.0833 | 0.009328 | 1.008 | 0.022766 | 22.585 | 1.4668 | 0.16184 | 0.033857 | 2.3027 | 2.3027 | 0.002889
2019-03-21 12:35 | 1553186118.3742 | 11 | 0.20268 | 0.001386 | 7.4e-05 | 100.8288 | 81.5464 | 105.9687 | 21.2371 | 0.00233 | 0.002351 | 263.5973 | 0.2715 | 8.4026 | 0.003809 | 0.005263 | 0.005158 | 5.8e-05 | 0.001096 | 0.23042 | 1.3107 | 2.5481 | 99.38 | 20.96 | 37.328 | 77.8528 | 113.5739 | 78.3477 | 0.009329 | 1.0044 | 0.022567 | 22.4678 | 1.8985 | 0.16152 | 0.039313 | 2.6485 | 2.6485 | 0.002506
2019-03-21 12:35 | 1553186119.3741 | 12 | 0.20173 | 0.002032 | 9e-05 | 100.7946 | 78.287 | 104.7305 | 21.2371 | 0.002375 | 0.002406 | 263.4457 | 0.27239 | 8.4027 | 0.003835 | 0.00568 | 0.005663 | 5.6e-05 | 0.001098 | 0.28241 | 1.3398 | 2.548 | 99.38 | 20.96 | 37.332 | 78.9341 | 114.2072 | 79.6886 | 0.00933 | 1.0006 | 0.022392 | 22.3788 | 2.2603 | 0.16205 | 0.044144 | 2.9491 | 2.9491 | 0.002258
2019-03-21 12:35 | 1553186120.3752 | 13 | 0.20122 | 0.002725 | 0.000102 | 100.8604 | 76.6671 | 104.3639 | 21.2365 | 0.002504 | 0.002449 | 263.3296 | 0.2711 | 8.404 | 0.003878 | 0.005147 | 0.005045 | 5.3e-05 | 0.001097 | 0.27743 | 1.327 | 2.5481 | 99.38 | 20.96 | 37.336 | 82.0306 | 115.2677 | 80.7024 | 0.009331 | 0.99772 | 0.022321 | 22.3716 | 2.4532 | 0.16128 | 0.046298 | 3.0819 | 3.0819 | 0.002151
2019-03-21 12:35 | 1553186121.3743 | 14 | 0.20049 | 0.003368 | 0.000114 | 100.8398 | 74.4646 | 104.1402 | 21.2369 | 0.002522 | 0.002532 | 263.2075 | 0.27143 | 8.4044 | 0.003962 | 0.005528 | 0.005474 | 5.6e-05 | 0.001097 | 0.23523 | 1.3064 | 2.5479 | 99.38 | 20.96 | 37.345 | 82.4646 | 117.3174 | 82.7142 | 0.009333 | 0.99208 | 0.022234 | 22.4111 | 2.7121 | 0.16148 | 0.049877 | 3.3059 | 3.3059 | 0.002008
2019-03-21 12:35 | 1553186122.3753 | 15 | 0.19999 | 0.00397 | 0.000122 | 100.8112 | 72.0021 | 104.4141 | 21.2376 | 0.002591 | 0.002569 | 263.0746 | 0.27115 | 8.4014 | 0.003988 | 0.005678 | 0.005559 | 5.6e-05 | 0.001092 | 0.23452 | 1.3597 | 2.5481 | 99.38 | 20.97 | 37.349 | 84.1176 | 117.9343 | 83.6095 | 0.00934 | 0.98959 | 0.022235 | 22.4687 | 3.0103 | 0.16131 | 0.052121 | 3.4537 | 3.4537 | 0.00192
2019-03-21 12:35 | 1553186123.3744 | 16 | 0.19989 | 0.004472 | 0.00013 | 100.7734 | 71.123 | 104.8603 | 21.2366 | 0.002613 | 0.002583 | 262.9596 | 0.27225 | 8.4028 | 0.004104 | 0.005677 | 0.005674 | 5.6e-05 | 0.001096 | 0.25249 | 1.3315 | 2.548 | 99.38 | 20.97 | 37.352 | 84.6556 | 120.7822 | 83.9269 | 0.009341 | 0.98871 | 0.022272 | 22.5267 | 3.1168 | 0.16196 | 0.052104 | 3.458 | 3.458 | 0.001925
2019-03-21 12:35 | 1553186124.3754 | 17 | 0.19983 | 0.004842 | 0.000136 | 100.8973 | 71.3277 | 105.6044 | 21.2369 | 0.002733 | 0.002573 | 262.8268 | 0.27109 | 8.4043 | 0.004119 | 0.006014 | 0.005947 | 5.6e-05 | 0.001097 | 0.25802 | 1.3091 | 2.5481 | 99.38 | 20.97 | 37.359 | 87.534 | 121.1421 | 83.6894 | 0.009343 | 0.98937 | 0.022359 | 22.599 | 3.1021 | 0.16127 | 0.052 | 3.4642 | 3.4642 | 0.001913
2019-03-21 12:35 | 1553186125.3751 | 18 | 0.19985 | 0.005123 | 0.000139 | 100.8544 | 70.7716 | 106.7562 | 21.2377 | 0.002719 | 0.002662 | 262.705 | 0.27239 | 8.4026 | 0.004138 | 0.005585 | 0.005491 | 5.8e-05 | 0.001091 | 0.26001 | 1.3243 | 2.548 | 99.38 | 20.97 | 37.36 | 87.2067 | 121.62 | 85.8312 | 0.009343 | 0.98347 | 0.022413 | 22.79 | 3.1683 | 0.16205 | 0.051554 | 3.4427 | 3.4427 | 0.001935
2019-03-21 12:35 | 1553186126.3749 | 19 | 0.19993 | 0.005332 | 0.000141 | 100.8842 | 70.8489 | 107.5474 | 21.2381 | 0.00275 | 0.002643 | 262.5903 | 0.272 | 8.402 | 0.004202 | 0.00573 | 0.005576 | 5.8e-05 | 0.001081 | 0.24961 | 1.337 | 2.5479 | 99.38 | 20.97 | 37.364 | 87.9638 | 123.1669 | 85.3805 | 0.009344 | 0.9847 | 0.02251 | 22.8599 | 3.1612 | 0.16181 | 0.050814 | 3.4081 | 3.4081 | 0.001951
2019-03-21 12:35 | 1553186127.3756 | 20 | 0.20015 | 0.005517 | 0.000142 | 100.8383 | 70.6622 | 108.1837 | 21.2373 | 0.00276 | 0.002684 | 262.4935 | 0.27256 | 8.4017 | 0.004251 | 0.005344 | 0.005337 | 5.5e-05 | 0.001073 | 0.26322 | 1.3067 | 2.548 | 99.38 | 20.97 | 37.366 | 88.1989 | 124.3871 | 86.3608 | 0.009344 | 0.98202 | 0.022546 | 22.9588 | 3.1807 | 0.16214 | 0.049309 | 3.3126 | 3.3126 | 0.002012
2019-03-21 12:35 | 1553186128.3746 | 21 | 0.20038 | 0.005679 | 0.000141 | 100.8574 | 70.2402 | 109.1161 | 21.2387 | 0.00288 | 0.002681 | 262.4032 | 0.27251 | 8.4025 | 0.004212 | 0.005559 | 0.005533 | 5.4e-05 | 0.001073 | 0.24027 | 1.3055 | 2.5479 | 99.38 | 20.97 | 37.368 | 91.0811 | 123.413 | 86.2936 | 0.009345 | 0.9822 | 0.022645 | 23.0553 | 3.236 | 0.16212 | 0.047729 | 3.2208 | 3.2208 | 0.002069
2019-03-21 12:35 | 1553186129.3747 | 22 | 0.20077 | 0.005816 | 0.00014 | 100.8708 | 70.6016 | 109.7315 | 21.2385 | 0.002845 | 0.002707 | 262.2966 | 0.27068 | 8.4026 | 0.004159 | 0.00583 | 0.005747 | 5.7e-05 | 0.001075 | 0.24065 | 1.3257 | 2.5481 | 99.38 | 20.97 | 37.37 | 90.2517 | 122.1314 | 86.9195 | 0.009345 | 0.98049 | 0.022689 | 23.1404 | 3.1913 | 0.16103 | 0.045235 | 3.0589 | 3.0589 | 0.002164
2019-03-21 12:35 | 1553186130.3748 | 23 | 0.20102 | 0.005914 | 0.000138 | 100.8541 | 72.1428 | 109.8233 | 21.238 | 0.00285 | 0.002685 | 262.2165 | 0.27127 | 8.4039 | 0.004202 | 0.0054 | 0.005395 | 6e-05 | 0.001078 | 0.25625 | 1.2701 | 2.5478 | 99.38 | 20.97 | 37.369 | 90.3574 | 123.1739 | 86.3833 | 0.009345 | 0.98196 | 0.022715 | 23.1328 | 2.9967 | 0.16138 | 0.043633 | 2.9544 | 2.9544 | 0.002245
2019-03-21 12:36 | 1553186131.3749 | 24 | 0.2012 | 0.005981 | 0.000136 | 100.763 | 73.1162 | 110.0045 | 21.2379 | 0.002857 | 0.002737 | 262.1126 | 0.27068 | 8.4035 | 0.004132 | 0.005882 | 0.005756 | 5.6e-05 | 0.001081 | 0.26377 | 1.2973 | 2.548 | 99.38 | 20.97 | 37.365 | 90.5286 | 121.4575 | 87.645 | 0.009344 | 0.97852 | 0.022694 | 23.1925 | 2.8687 | 0.16103 | 0.042478 | 2.8737 | 2.8737 | 0.002303
2019-03-21 12:36 | 1553186132.3759 | 25 | 0.20141 | 0.006021 | 0.000133 | 100.8446 | 72.6775 | 110.1392 | 21.239 | 0.002883 | 0.002816 | 262.0221 | 0.27132 | 8.4013 | 0.004192 | 0.006042 | 0.006028 | 5.6e-05 | 0.00108 | 0.26504 | 1.3163 | 2.548 | 99.38 | 20.97 | 37.365 | 91.1586 | 122.9366 | 89.5315 | 0.009344 | 0.97343 | 0.022649 | 23.2672 | 2.9298 | 0.16141 | 0.041188 | 2.7812 | 2.7812 | 0.002385
2019-03-21 12:36 | 1553186133.3749 | 26 | 0.20152 | 0.006056 | 0.00013 | 100.8509 | 71.5951 | 110.0895 | 21.2392 | 0.002917 | 0.002772 | 261.9343 | 0.27146 | 8.4018 | 0.004194 | 0.005399 | 0.005327 | 5.7e-05 | 0.001077 | 0.26161 | 1.3082 | 2.5482 | 99.38 | 20.97 | 37.363 | 91.9722 | 122.9766 | 88.4906 | 0.009344 | 0.97623 | 0.022677 | 23.2286 | 3.0645 | 0.16149 | 0.040486 | 2.7373 | 2.7373 | 0.002425
2019-03-21 12:36 | 1553186134.3741 | 27 | 0.20153 | 0.006059 | 0.000128 | 100.8744 | 72.2708 | 109.9022 | 21.2387 | 0.002952 | 0.002781 | 261.8418 | 0.27134 | 8.402 | 0.004205 | 0.005932 | 0.005884 | 5.2e-05 | 0.001079 | 0.27079 | 1.3189 | 2.548 | 99.38 | 20.97 | 37.361 | 92.8084 | 123.2638 | 88.6924 | 0.009343 | 0.97569 | 0.022651 | 23.2153 | 2.9826 | 0.16142 | 0.040445 | 2.7316 | 2.7316 | 0.002429
2019-03-21 12:36 | 1553186135.3741 | 28 | 0.20157 | 0.006058 | 0.000125 | 100.968 | 74.2347 | 109.6778 | 21.2391 | 0.002948 | 0.002767 | 261.7353 | 0.27137 | 8.4016 | 0.004173 | 0.005968 | 0.005841 | 5.6e-05 | 0.001076 | 0.24779 | 1.337 | 2.5481 | 99.38 | 20.97 | 37.362 | 92.7205 | 122.4712 | 88.3513 | 0.009343 | 0.97661 | 0.022638 | 23.1806 | 2.7511 | 0.16144 | 0.040218 | 2.715 | 2.715 | 0.002444
2019-03-21 12:36 | 1553186136.374 | 29 | 0.20158 | 0.006045 | 0.000123 | 100.8446 | 74.2275 | 109.7641 | 21.2385 | 0.003054 | 0.002857 | 261.6116 | 0.27225 | 8.4014 | 0.004216 | 0.005606 | 0.005569 | 5.3e-05 | 0.001078 | 0.26913 | 1.3601 | 2.548 | 99.38 | 20.97 | 37.364 | 95.2711 | 123.514 | 90.5205 | 0.009344 | 0.97078 | 0.02258 | 23.2592 | 2.741 | 0.16196 | 0.040171 | 2.7049 | 2.7049 | 0.002461
2019-03-21 12:36 | 1553186137.3741 | 30 | 0.20159 | 0.006041 | 0.000121 | 100.8544 | 74.6148 | 109.5076 | 21.2387 | 0.003005 | 0.002911 | 261.5494 | 0.27245 | 8.4031 | 0.004284 | 0.005617 | 0.005516 | 5.5e-05 | 0.001088 | 0.29717 | 1.3345 | 2.548 | 99.38 | 20.97 | 37.368 | 94.0953 | 125.176 | 91.8261 | 0.009345 | 0.96731 | 0.022513 | 23.2737 | 2.6953 | 0.16208 | 0.040119 | 2.6935 | 2.6935 | 0.002473
2019-03-21 12:36 | 1553186138.3742 | 31 | 0.20182 | 0.006031 | 0.000118 | 100.8696 | 74.159 | 108.8213 | 21.2394 | 0.003087 | 0.002855 | 261.4254 | 0.2715 | 8.404 | 0.004331 | 0.005166 | 0.005105 | 5.6e-05 | 0.001085 | 0.25608 | 1.3216 | 2.5479 | 99.38 | 20.97 | 37.371 | 96.0846 | 126.3475 | 90.4938 | 0.009346 | 0.97086 | 0.022483 | 23.1583 | 2.7515 | 0.16151 | 0.038732 | 2.5973 | 2.5973 | 0.002556
2019-03-21 12:36 | 1553186139.3752 | 32 | 0.20168 | 0.006028 | 0.000117 | 100.8011 | 74.1768 | 108.9288 | 21.2396 | 0.003099 | 0.002923 | 261.3399 | 0.27115 | 8.4012 | 0.004235 | 0.00558 | 0.005513 | 5.6e-05 | 0.001089 | 0.25962 | 1.3205 | 2.5476 | 99.38 | 20.97 | 37.376 | 96.3559 | 123.9759 | 92.1116 | 0.009347 | 0.96656 | 0.022445 | 23.2212 | 2.7433 | 0.1613 | 0.0396 | 2.6509 | 2.6509 | 0.002501
2019-03-21 12:36 | 1553186140.3743 | 33 | 0.20161 | 0.006021 | 0.000115 | 100.8696 | 75.2898 | 109.1644 | 21.239 | 0.003113 | 0.002937 | 261.2347 | 0.27217 | 8.4032 | 0.004275 | 0.005673 | 0.005567 | 5.7e-05 | 0.001088 | 0.27848 | 1.3382 | 2.5479 | 99.38 | 20.97 | 37.379 | 96.7117 | 124.9696 | 92.449 | 0.009348 | 0.96566 | 0.022458 | 23.257 | 2.6161 | 0.16192 | 0.039987 | 2.6785 | 2.6785 | 0.002484
2019-03-21 12:36 | 1553186141.3743 | 34 | 0.20177 | 0.00602 | 0.000113 | 100.8172 | 76.6944 | 109.4747 | 21.2395 | 0.003131 | 0.002941 | 261.1587 | 0.27135 | 8.4023 | 0.004278 | 0.00618 | 0.006149 | 5.3e-05 | 0.001085 | 0.27433 | 1.3536 | 2.548 | 99.38 | 20.97 | 37.383 | 97.1447 | 125.04 | 92.5512 | 0.009349 | 0.96539 | 0.022487 | 23.2933 | 2.4462 | 0.16143 | 0.03905 | 2.6193 | 2.6193 | 0.002533
2019-03-21 12:36 | 1553186142.3744 | 35 | 0.20195 | 0.006018 | 0.000111 | 100.8801 | 77.6691 | 109.8451 | 21.2396 | 0.003175 | 0.002933 | 261.0685 | 0.27124 | 8.4039 | 0.004361 | 0.005313 | 0.005303 | 5.4e-05 | 0.001086 | 0.24945 | 1.3285 | 2.548 | 99.38 | 20.97 | 37.391 | 98.2036 | 127.0882 | 92.3618 | 0.009351 | 0.96589 | 0.022531 | 23.3267 | 2.3388 | 0.16136 | 0.038009 | 2.5547 | 2.5547 | 0.002596
2019-03-21 12:36 | 1553186143.3745 | 36 | 0.20209 | 0.006006 | 0.00011 | 100.8595 | 78.7927 | 109.9639 | 21.2401 | 0.003149 | 0.002938 | 260.9352 | 0.2707 | 8.4035 | 0.004359 | 0.005869 | 0.005847 | 5.9e-05 | 0.001083 | 0.25033 | 1.3075 | 2.548 | 99.38 | 20.97 | 37.395 | 97.5716 | 127.0411 | 92.4829 | 0.009352 | 0.96557 | 0.022539 | 23.3431 | 2.2085 | 0.16104 | 0.037199 | 2.5013 | 2.5013 | 0.002646
2019-03-21 12:36 | 1553186144.3755 | 37 | 0.20213 | 0.006003 | 0.000107 | 100.8431 | 78.1368 | 110.4955 | 21.2406 | 0.003149 | 0.002969 | 260.8345 | 0.27189 | 8.4019 | 0.004385 | 0.006356 | 0.006264 | 5.9e-05 | 0.001089 | 0.25785 | 1.2881 | 2.548 | 99.38 | 20.97 | 37.399 | 97.5728 | 127.6755 | 93.2421 | 0.009353 | 0.96357 | 0.02257 | 23.4238 | 2.2818 | 0.16174 | 0.036965 | 2.4892 | 2.4892 | 0.002671
2019-03-21 12:36 | 1553186145.3746 | 38 | 0.20227 | 0.005985 | 0.000105 | 100.8678 | 78.2567 | 110.8518 | 21.2402 | 0.003129 | 0.002988 | 260.7452 | 0.27118 | 8.4015 | 0.00443 | 0.005821 | 0.005778 | 5.8e-05 | 0.001085 | 0.26217 | 1.2858 | 2.548 | 99.38 | 20.97 | 37.402 | 97.0891 | 128.7616 | 93.6791 | 0.009353 | 0.96242 | 0.022593 | 23.4755 | 2.2703 | 0.16132 | 0.036107 | 2.4341 | 2.4341 | 0.002724
2019-03-21 12:36 | 1553186146.3757 | 39 | 0.2025 | 0.005961 | 0.000102 | 100.8619 | 78.4484 | 111.1657 | 21.24 | 0.003117 | 0.00298 | 260.7021 | 0.27089 | 8.4041 | 0.004436 | 0.006691 | 0.006616 | 5.7e-05 | 0.001087 | 0.26145 | 1.3019 | 2.5479 | 99.38 | 20.97 | 37.404 | 96.8114 | 128.9167 | 93.4872 | 0.009354 | 0.96293 | 0.022631 | 23.5025 | 2.2479 | 0.16115 | 0.034817 | 2.3513 | 2.3513 | 0.002817
2019-03-21 12:36 | 1553186147.3747 | 40 | 0.20283 | 0.00593 | 9.9e-05 | 100.8342 | 78.2802 | 111.7767 | 21.2397 | 0.003202 | 0.002919 | 260.6053 | 0.27203 | 8.4021 | 0.004506 | 0.006205 | 0.00614 | 5.9e-05 | 0.001089 | 0.27433 | 1.2998 | 2.5481 | 99.38 | 20.97 | 37.405 | 98.8663 | 130.6341 | 92.0298 | 0.009354 | 0.96677 | 0.022739 | 23.5202 | 2.2646 | 0.16183 | 0.032849 | 2.2292 | 2.2292 | 0.002984
2019-03-21 12:36 | 1553186148.3758 | 41 | 0.20326 | 0.005879 | 9.5e-05 | 100.8538 | 79.3908 | 112.6059 | 21.241 | 0.003227 | 0.00291 | 260.5249 | 0.27128 | 8.3996 | 0.004448 | 0.006115 | 0.006051 | 6.1e-05 | 0.001086 | 0.2568 | 1.3079 | 2.548 | 99.38 | 20.97 | 37.406 | 99.4661 | 129.2097 | 91.8018 | 0.009354 | 0.96737 | 0.02283 | 23.5999 | 2.1403 | 0.16139 | 0.030347 | 2.068 | 2.068 | 0.003207
2019-03-21 12:36 | 1553186149.3739 | 42 | 0.2035 | 0.005817 | 9.2e-05 | 100.8372 | 79.8557 | 113.1747 | 21.2404 | 0.003175 | 0.002983 | 260.4592 | 0.27184 | 8.3995 | 0.004418 | 0.005415 | 0.005426 | 5.4e-05 | 0.001083 | 0.26017 | 1.2977 | 2.5481 | 99.38 | 20.97 | 37.405 | 98.1935 | 128.4749 | 93.573 | 0.009354 | 0.9627 | 0.022832 | 23.7167 | 2.0866 | 0.16172 | 0.028963 | 1.9741 | 1.9741 | 0.003367
2019-03-21 12:36 | 1553186150.3739 | 43 | 0.20369 | 0.005745 | 8.9e-05 | 100.8559 | 79.8622 | 113.8112 | 21.2404 | 0.00314 | 0.002916 | 260.3891 | 0.27189 | 8.4011 | 0.004403 | 0.005931 | 0.005847 | 5.9e-05 | 0.001083 | 0.2385 | 1.321 | 2.548 | 99.38 | 20.96 | 37.403 | 97.3565 | 128.1057 | 91.9592 | 0.009348 | 0.96696 | 0.022947 | 23.731 | 2.0876 | 0.16175 | 0.027919 | 1.9128 | 1.9128 | 0.003476
2019-03-21 12:36 | 1553186151.3742 | 44 | 0.20385 | 0.005666 | 8.5e-05 | 100.8947 | 79.9347 | 113.9532 | 21.2403 | 0.00316 | 0.002926 | 260.2885 | 0.27152 | 8.4017 | 0.004415 | 0.006267 | 0.006219 | 5.7e-05 | 0.001085 | 0.27295 | 1.3203 | 2.5478 | 99.38 | 20.96 | 37.402 | 97.8319 | 128.4109 | 92.1819 | 0.009348 | 0.96637 | 0.022954 | 23.753 | 2.0829 | 0.16153 | 0.027075 | 1.8557 | 1.8557 | 0.003577
2019-03-21 12:36 | 1553186152.3749 | 45 | 0.20403 | 0.005575 | 8.2e-05 | 100.9107 | 80.5566 | 114.003 | 21.24 | 0.00316 | 0.002901 | 260.1852 | 0.27125 | 8.4051 | 0.004375 | 0.005716 | 0.0056 | 5.9e-05 | 0.001084 | 0.27096 | 1.3476 | 2.5478 | 99.38 | 20.96 | 37.398 | 97.8506 | 127.4113 | 91.5804 | 0.009347 | 0.96796 | 0.022978 | 23.7386 | 2.015 | 0.16137 | 0.026098 | 1.7908 | 1.7908 | 0.003703
2019-03-21 12:36 | 1553186153.3751 | 46 | 0.20429 | 0.005482 | 7.8e-05 | 100.8154 | 80.668 | 114.6112 | 21.241 | 0.003186 | 0.002853 | 260.1352 | 0.27258 | 8.4026 | 0.004396 | 0.006001 | 0.005933 | 5.8e-05 | 0.00108 | 0.2505 | 1.3134 | 2.548 | 99.38 | 20.96 | 37.396 | 98.4802 | 127.9471 | 90.4456 | 0.009346 | 0.97098 | 0.023075 | 23.7648 | 1.9942 | 0.16216 | 0.024659 | 1.6994 | 1.6994 | 0.003922
2019-03-21 12:36 | 1553186154.3741 | 47 | 0.20455 | 0.005383 | 7.4e-05 | 100.8118 | 82.2248 | 114.8879 | 21.2404 | 0.00321 | 0.002849 | 260.0786 | 0.27105 | 8.4036 | 0.004384 | 0.006383 | 0.00634 | 5.8e-05 | 0.001078 | 0.2338 | 1.306 | 2.548 | 99.38 | 20.96 | 37.394 | 99.0427 | 127.6515 | 90.3395 | 0.009346 | 0.97127 | 0.023106 | 23.79 | 1.8229 | 0.16125 | 0.023243 | 1.6043 | 1.6043 | 0.004131
2019-03-21 12:36 | 1553186155.3755 | 48 | 0.2046 | 0.005279 | 7e-05 | 100.7585 | 83.0474 | 114.8682 | 21.2417 | 0.003185 | 0.002883 | 260.0253 | 0.27256 | 8.4046 | 0.004337 | 0.006018 | 0.005986 | 5.7e-05 | 0.001083 | 0.28191 | 1.3224 | 2.5479 | 99.38 | 20.96 | 37.391 | 98.4347 | 126.486 | 91.1668 | 0.009345 | 0.96906 | 0.023078 | 23.815 | 1.7291 | 0.16214 | 0.023064 | 1.5901 | 1.5901 | 0.004191
2019-03-21 12:36 | 1553186156.3752 | 49 | 0.20479 | 0.005169 | 6.6e-05 | 100.872 | 83.7309 | 114.9745 | 21.2411 | 0.003172 | 0.002901 | 259.932 | 0.27037 | 8.4023 | 0.00429 | 0.005819 | 0.005747 | 5.3e-05 | 0.00108 | 0.2589 | 1.3221 | 2.5479 | 99.38 | 20.96 | 37.39 | 98.1331 | 125.3495 | 91.6002 | 0.009345 | 0.96791 | 0.023075 | 23.8402 | 1.6659 | 0.16084 | 0.022045 | 1.5199 | 1.5199 | 0.004349
2019-03-21 12:36 | 1553186157.3753 | 50 | 0.20494 | 0.005059 | 6.3e-05 | 100.8342 | 84.6229 | 115.0133 | 21.242 | 0.003147 | 0.002855 | 259.8516 | 0.27208 | 8.4027 | 0.004387 | 0.006106 | 0.005965 | 5.6e-05 | 0.00108 | 0.27903 | 1.3278 | 2.5479 | 99.38 | 20.96 | 37.388 | 97.5375 | 127.7297 | 90.4933 | 0.009344 | 0.97086 | 0.023114 | 23.808 | 1.5677 | 0.16186 | 0.021291 | 1.4706 | 1.4706 | 0.004524
2019-03-21 12:36 | 1553186158.3757 | 51 | 0.20503 | 0.004948 | 5.9e-05 | 100.8771 | 85.3211 | 115.4163 | 21.2419 | 0.003164 | 0.002887 | 259.8114 | 0.272 | 8.4021 | 0.004345 | 0.006327 | 0.006358 | 5.2e-05 | 0.001076 | 0.28744 | 1.3745 | 2.5481 | 99.38 | 20.96 | 37.387 | 97.9456 | 126.6828 | 91.2579 | 0.009344 | 0.96882 | 0.02313 | 23.8747 | 1.498 | 0.16181 | 0.020894 | 1.4444 | 1.4444 | 0.004604
2019-03-21 12:36 | 1553186159.3754 | 52 | 0.2052 | 0.004837 | 5.7e-05 | 100.8434 | 85.7111 | 115.1952 | 21.2428 | 0.003131 | 0.00291 | 259.7284 | 0.27137 | 8.4043 | 0.004307 | 0.005988 | 0.005932 | 5.5e-05 | 0.001078 | 0.27704 | 1.3642 | 2.5479 | 99.38 | 20.96 | 37.386 | 97.1409 | 125.7451 | 91.8154 | 0.009344 | 0.96734 | 0.02309 | 23.8701 | 1.4543 | 0.16144 | 0.020046 | 1.3835 | 1.3835 | 0.004796
2019-03-21 12:36 | 1553186160.3743 | 53 | 0.20533 | 0.004733 | 5.5e-05 | 100.8196 | 86.4805 | 115.106 | 21.2415 | 0.003183 | 0.002845 | 259.6558 | 0.27152 | 8.4016 | 0.004298 | 0.005729 | 0.005618 | 5.7e-05 | 0.001082 | 0.2625 | 1.3613 | 2.548 | 99.38 | 20.96 | 37.382 | 98.4069 | 125.5233 | 90.2396 | 0.009343 | 0.97153 | 0.023132 | 23.8093 | 1.3722 | 0.16153 | 0.019338 | 1.3371 | 1.3371 | 0.004965
2019-03-21 12:36 | 1553186161.3755 | 54 | 0.20538 | 0.00463 | 5.2e-05 | 100.9256 | 87.2931 | 115.1762 | 21.2422 | 0.003177 | 0.002854 | 259.5949 | 0.27213 | 8.4036 | 0.00424 | 0.005388 | 0.005305 | 5.6e-05 | 0.00108 | 0.27936 | 1.3464 | 2.5479 | 99.38 | 20.96 | 37.382 | 98.2603 | 124.1043 | 90.4538 | 0.009343 | 0.97096 | 0.023132 | 23.8236 | 1.298 | 0.16189 | 0.019225 | 1.3294 | 1.3294 | 0.005005
2019-03-21 12:36 | 1553186162.3755 | 55 | 0.2054 | 0.004533 | 5e-05 | 100.8422 | 87.3807 | 115.4733 | 21.2418 | 0.003178 | 0.002838 | 259.5517 | 0.27066 | 8.4028 | 0.00419 | 0.005734 | 0.005641 | 5.2e-05 | 0.001083 | 0.2891 | 1.3091 | 2.548 | 99.38 | 20.96 | 37.382 | 98.271 | 122.8829 | 90.0711 | 0.009343 | 0.97198 | 0.023174 | 23.8417 | 1.2816 | 0.16101 | 0.01921 | 1.3309 | 1.3309 | 0.004972
2019-03-21 12:36 | 1553186163.3746 | 56 | 0.20547 | 0.004438 | 4.9e-05 | 100.8422 | 86.9921 | 115.5646 | 21.2403 | 0.003156 | 0.002878 | 259.5053 | 0.27063 | 8.4002 | 0.00415 | 0.00552 | 0.005452 | 5.3e-05 | 0.001079 | 0.26792 | 1.2876 | 2.5479 | 99.38 | 20.96 | 37.381 | 97.753 | 121.9053 | 91.0279 | 0.009342 | 0.96943 | 0.023152 | 23.8825 | 1.3215 | 0.161 | 0.018897 | 1.3081 | 1.3081 | 0.005058
2019-03-21 12:36 | 1553186164.3746 | 57 | 0.20553 | 0.004346 | 4.7e-05 | 100.8258 | 87.4495 | 115.8456 | 21.2419 | 0.003176 | 0.002862 | 259.447 | 0.27109 | 8.4029 | 0.004153 | 0.005949 | 0.005839 | 5.5e-05 | 0.001085 | 0.26609 | 1.2863 | 2.5479 | 99.38 | 20.96 | 37.382 | 98.2348 | 121.9815 | 90.6652 | 0.009343 | 0.9704 | 0.023192 | 23.8996 | 1.2731 | 0.16127 | 0.018686 | 1.2959 | 1.2959 | 0.005115
2019-03-21 12:36 | 1553186165.374 | 58 | 0.20552 | 0.00427 | 4.6e-05 | 100.8145 | 87.8473 | 116.1716 | 21.2428 | 0.003133 | 0.002835 | 259.369 | 0.27161 | 8.4031 | 0.004127 | 0.005581 | 0.005543 | 5.4e-05 | 0.001081 | 0.26288 | 1.2814 | 2.548 | 99.38 | 20.96 | 37.384 | 97.1887 | 121.3343 | 90.0048 | 0.009343 | 0.97216 | 0.023246 | 23.9115 | 1.2315 | 0.16158 | 0.018804 | 1.3071 | 1.3071 | 0.005081
2019-03-21 12:36 | 1553186166.3737 | 59 | 0.20555 | 0.0042 | 4.3e-05 | 100.864 | 87.6109 | 116.2413 | 21.2427 | 0.003195 | 0.002815 | 259.317 | 0.27111 | 8.4035 | 0.004064 | 0.005345 | 0.005285 | 5.4e-05 | 0.001087 | 0.26001 | 1.3163 | 2.5481 | 99.38 | 20.96 | 37.385 | 98.6803 | 119.7935 | 89.5282 | 0.009343 | 0.97344 | 0.023268 | 23.903 | 1.26 | 0.16128 | 0.018751 | 1.3048 | 1.3048 | 0.00508
2019-03-21 12:36 | 1553186167.3738 | 60 | 0.20557 | 0.004133 | 4.2e-05 | 100.8199 | 87.4396 | 116.1812 | 21.2425 | 0.003148 | 0.0028 | 259.2874 | 0.27196 | 8.4022 | 0.004065 | 0.005758 | 0.005688 | 5.5e-05 | 0.001088 | 0.2568 | 1.3275 | 2.5479 | 99.38 | 20.96 | 37.386 | 97.5621 | 119.8407 | 89.1545 | 0.009344 | 0.97444 | 0.023274 | 23.8845 | 1.2736 | 0.16179 | 0.018711 | 1.3024 | 1.3024 | 0.005105
2019-03-21 12:36 | 1553186168.3752 | 61 | 0.20561 | 0.004065 | 4e-05 | 100.8121 | 88.3735 | 116.1768 | 21.2419 | 0.003133 | 0.002884 | 259.2214 | 0.27101 | 8.4043 | 0.004115 | 0.00507 | 0.005018 | 5.5e-05 | 0.001086 | 0.274 | 1.2874 | 2.5481 | 99.38 | 20.96 | 37.388 | 97.1987 | 121.054 | 91.1771 | 0.009344 | 0.96903 | 0.023209 | 23.9506 | 1.1779 | 0.16122 | 0.018526 | 1.2861 | 1.2861 | 0.005152
2019-03-21 12:36 | 1553186169.3739 | 62 | 0.2057 | 0.004008 | 3.9e-05 | 100.7353 | 88.7035 | 116.3507 | 21.2432 | 0.00311 | 0.002811 | 259.1734 | 0.27157 | 8.4033 | 0.003968 | 0.005894 | 0.005796 | 5.3e-05 | 0.001077 | 0.25227 | 1.2894 | 2.5481 | 99.38 | 20.96 | 37.389 | 96.6321 | 117.4559 | 89.4193 | 0.009344 | 0.97373 | 0.023283 | 23.9106 | 1.1377 | 0.16156 | 0.018094 | 1.2601 | 1.2601 | 0.005269
2019-03-21 12:36 | 1553186170.3743 | 63 | 0.20578 | 0.003949 | 3.8e-05 | 100.8047 | 88.4883 | 116.457 | 21.2425 | 0.003153 | 0.00284 | 259.1382 | 0.27248 | 8.4019 | 0.00404 | 0.004932 | 0.004826 | 5.4e-05 | 0.001081 | 0.30326 | 1.3156 | 2.5479 | 99.38 | 20.96 | 37.394 | 97.6836 | 119.2256 | 90.1344 | 0.009346 | 0.97181 | 0.02327 | 23.9452 | 1.1656 | 0.1621 | 0.017684 | 1.231 | 1.231 | 0.005412
2019-03-21 12:36 | 1553186171.3744 | 64 | 0.2058 | 0.003897 | 3.7e-05 | 100.8309 | 88.8309 | 116.4528 | 21.2427 | 0.003133 | 0.002824 | 259.1048 | 0.27182 | 8.4019 | 0.004047 | 0.005327 | 0.005272 | 5.9e-05 | 0.001081 | 0.24585 | 1.3399 | 2.5481 | 99.38 | 20.96 | 37.398 | 97.1981 | 119.3961 | 89.7502 | 0.009347 | 0.97284 | 0.023282 | 23.9321 | 1.1334 | 0.16171 | 0.017597 | 1.2257 | 1.2257 | 0.005423
2019-03-21 12:36 | 1553186172.3751 | 65 | 0.20582 | 0.003845 | 3.6e-05 | 100.833 | 89.1745 | 116.377 | 21.2429 | 0.003099 | 0.002798 | 259.0547 | 0.27174 | 8.4012 | 0.004009 | 0.005491 | 0.005425 | 5.4e-05 | 0.001079 | 0.2724 | 1.3319 | 2.548 | 99.38 | 20.96 | 37.402 | 96.3741 | 118.4673 | 89.1134 | 0.009348 | 0.97455 | 0.023295 | 23.9033 | 1.099 | 0.16166 | 0.01754 | 1.2224 | 1.2224 | 0.005435
2019-03-21 12:36 | 1553186173.3751 | 66 | 0.20585 | 0.003797 | 3.6e-05 | 100.8512 | 89.6323 | 116.2545 | 21.2428 | 0.003158 | 0.002803 | 258.9804 | 0.27093 | 8.4047 | 0.003999 | 0.005775 | 0.005725 | 5.6e-05 | 0.001086 | 0.28368 | 1.3284 | 2.5479 | 99.38 | 20.96 | 37.402 | 97.8039 | 118.2215 | 89.2431 | 0.009348 | 0.97421 | 0.023279 | 23.8949 | 1.0548 | 0.16118 | 0.017417 | 1.213 | 1.213 | 0.005461
2019-03-21 12:36 | 1553186174.3752 | 67 | 0.20586 | 0.003746 | 3.6e-05 | 100.8148 | 90.0393 | 115.9813 | 21.2425 | 0.003117 | 0.002819 | 258.9301 | 0.27237 | 8.4023 | 0.004033 | 0.005439 | 0.005407 | 5.3e-05 | 0.001085 | 0.274 | 1.3052 | 2.5481 | 99.38 | 20.96 | 37.411 | 96.81 | 119.0439 | 89.6076 | 0.00935 | 0.97323 | 0.02324 | 23.8789 | 1.0111 | 0.16204 | 0.017429 | 1.2118 | 1.2118 | 0.005496
2019-03-21 12:36 | 1553186175.3756 | 68 | 0.20587 | 0.0037 | 3.6e-05 | 100.7728 | 89.9167 | 115.8599 | 21.2423 | 0.003101 | 0.002822 | 258.866 | 0.27157 | 8.4019 | 0.003937 | 0.005715 | 0.005584 | 5.6e-05 | 0.00108 | 0.27704 | 1.326 | 2.5482 | 99.38 | 20.96 | 37.415 | 96.4242 | 116.7011 | 89.6848 | 0.009351 | 0.97302 | 0.023225 | 23.8689 | 1.0195 | 0.16156 | 0.017404 | 1.2092 | 1.2092 | 0.005491
2019-03-21 12:36 | 1553186176.3753 | 69 | 0.20589 | 0.00366 | 3.7e-05 | 100.8005 | 89.9773 | 115.8289 | 21.2431 | 0.0031 | 0.002805 | 258.811 | 0.27135 | 8.4041 | 0.003949 | 0.005022 | 0.005003 | 6e-05 | 0.001084 | 0.24834 | 1.3209 | 2.5479 | 99.38 | 20.96 | 37.419 | 96.3891 | 117.0011 | 89.2894 | 0.009352 | 0.97408 | 0.023234 | 23.8527 | 1.016 | 0.16142 | 0.017349 | 1.2059 | 1.2059 | 0.005502
2019-03-21 12:36 | 1553186177.3753 | 70 | 0.20588 | 0.003624 | 3.7e-05 | 100.7901 | 89.7549 | 115.7888 | 21.2429 | 0.003096 | 0.002805 | 258.7406 | 0.27141 | 8.4034 | 0.003856 | 0.005239 | 0.005198 | 6e-05 | 0.001078 | 0.24546 | 1.2836 | 2.548 | 99.38 | 20.96 | 37.426 | 96.2949 | 114.7236 | 89.2745 | 0.009354 | 0.97412 | 0.023231 | 23.8481 | 1.0372 | 0.16146 | 0.017444 | 1.2123 | 1.2123 | 0.005474
2019-03-21 12:36 | 1553186178.3748 | 71 | 0.20585 | 0.003591 | 3.7e-05 | 100.794 | 89.6544 | 115.7266 | 21.2437 | 0.003073 | 0.002853 | 258.6909 | 0.27121 | 8.4027 | 0.003802 | 0.005326 | 0.005272 | 5.7e-05 | 0.001077 | 0.23203 | 1.2956 | 2.548 | 99.38 | 20.96 | 37.429 | 95.7311 | 113.4185 | 90.4388 | 0.009354 | 0.971 | 0.023187 | 23.8799 | 1.0475 | 0.16134 | 0.017669 | 1.2256 | 1.2256 | 0.005411
2019-03-21 12:36 | 1553186179.3748 | 72 | 0.20586 | 0.003561 | 3.8e-05 | 100.7722 | 89.9947 | 115.5904 | 21.2439 | 0.003089 | 0.002825 | 258.6298 | 0.27123 | 8.3987 | 0.003785 | 0.00492 | 0.004892 | 5.8e-05 | 0.001075 | 0.25033 | 1.3063 | 2.548 | 99.38 | 20.96 | 37.433 | 96.1173 | 112.9928 | 89.7692 | 0.009355 | 0.97279 | 0.023195 | 23.8439 | 1.0117 | 0.16135 | 0.017658 | 1.2252 | 1.2252 | 0.005413
2019-03-21 12:36 | 1553186180.3754 | 73 | 0.20587 | 0.003535 | 3.9e-05 | 100.7895 | 90.3635 | 115.4207 | 21.2432 | 0.003097 | 0.002787 | 258.5521 | 0.27141 | 8.4031 | 0.003757 | 0.005259 | 0.005193 | 6.1e-05 | 0.001074 | 0.25642 | 1.2891 | 2.5479 | 99.38 | 20.95 | 37.44 | 96.3203 | 112.3262 | 88.8467 | 0.009351 | 0.97527 | 0.023208 | 23.7961 | 0.97668 | 0.16146 | 0.017637 | 1.2244 | 1.2244 | 0.00542
2019-03-21 12:36 | 1553186181.3738 | 74 | 0.20581 | 0.003507 | 3.9e-05 | 100.8592 | 90.0362 | 115.1231 | 21.2436 | 0.003126 | 0.002795 | 258.5102 | 0.27104 | 8.4021 | 0.003804 | 0.005101 | 0.005026 | 5.6e-05 | 0.001075 | 0.29374 | 1.2351 | 2.548 | 99.38 | 20.95 | 37.443 | 97.0289 | 113.4624 | 89.0322 | 0.009352 | 0.97477 | 0.023172 | 23.7715 | 1.0153 | 0.16124 | 0.018039 | 1.2503 | 1.2503 | 0.0053
2019-03-21 12:36 | 1553186182.375 | 75 | 0.20585 | 0.003482 | 4e-05 | 100.8234 | 89.9184 | 115.0365 | 21.2436 | 0.003102 | 0.002783 | 258.4321 | 0.27113 | 8.4032 | 0.003713 | 0.005335 | 0.005337 | 5.7e-05 | 0.001068 | 0.25536 | 1.3194 | 2.548 | 99.38 | 20.95 | 37.448 | 96.4374 | 111.2475 | 88.7487 | 0.009353 | 0.97554 | 0.023172 | 23.7532 | 1.0239 | 0.1613 | 0.017839 | 1.2364 | 1.2364 | 0.005362
2019-03-21 12:36 | 1553186183.3736 | 76 | 0.20585 | 0.003466 | 4e-05 | 100.7236 | 89.7481 | 115.0586 | 21.2429 | 0.003135 | 0.002735 | 258.3888 | 0.27209 | 8.4031 | 0.003743 | 0.005335 | 0.005199 | 6e-05 | 0.001069 | 0.28982 | 1.3091 | 2.548 | 99.38 | 20.95 | 37.45 | 97.2342 | 111.9847 | 87.5937 | 0.009354 | 0.97866 | 0.023211 | 23.7176 | 1.032 | 0.16187 | 0.017836 | 1.2383 | 1.2383 | 0.005373
2019-03-21 12:36 | 1553186184.3747 | 77 | 0.20582 | 0.003453 | 4.1e-05 | 100.7564 | 89.6585 | 115.052 | 21.2435 | 0.003113 | 0.002766 | 258.3044 | 0.27211 | 8.4023 | 0.003705 | 0.005248 | 0.005131 | 5.9e-05 | 0.001068 | 0.29015 | 1.3228 | 2.5479 | 99.38 | 20.95 | 37.453 | 96.7028 | 111.046 | 88.3354 | 0.009355 | 0.97665 | 0.023187 | 23.7413 | 1.0438 | 0.16188 | 0.018021 | 1.2498 | 1.2498 | 0.005323
2019-03-21 12:36 | 1553186185.3737 | 78 | 0.20581 | 0.003441 | 4.2e-05 | 100.7588 | 89.9249 | 115.1677 | 21.2433 | 0.003017 | 0.002792 | 258.2655 | 0.27173 | 8.4039 | 0.003744 | 0.004908 | 0.004972 | 5.9e-05 | 0.001062 | 0.23971 | 1.3174 | 2.5481 | 99.38 | 20.95 | 37.459 | 94.3958 | 112.0097 | 88.9763 | 0.009356 | 0.97492 | 0.023178 | 23.7742 | 1.0175 | 0.16165 | 0.018121 | 1.2561 | 1.2561 | 0.005289
2019-03-21 12:36 | 1553186186.3752 | 79 | 0.20583 | 0.003426 | 4.3e-05 | 100.7648 | 89.948 | 115.2989 | 21.2428 | 0.003087 | 0.002852 | 258.1952 | 0.27115 | 8.4044 | 0.003814 | 0.004576 | 0.004424 | 5.8e-05 | 0.001064 | 0.25321 | 1.3097 | 2.548 | 99.38 | 20.95 | 37.461 | 96.0786 | 113.7036 | 90.4035 | 0.009357 | 0.9711 | 0.023146 | 23.8346 | 1.0157 | 0.1613 | 0.01804 | 1.2488 | 1.2488 | 0.005309
2019-03-21 12:36 | 1553186187.3738 | 80 | 0.20582 | 0.003419 | 4.3e-05 | 100.7719 | 89.5863 | 115.4027 | 21.2427 | 0.003103 | 0.00276 | 258.1108 | 0.27172 | 8.4035 | 0.003698 | 0.005225 | 0.005218 | 6.1e-05 | 0.001064 | 0.27865 | 1.3535 | 2.5481 | 99.38 | 20.94 | 37.465 | 96.4702 | 110.882 | 88.208 | 0.009352 | 0.977 | 0.023226 | 23.7732 | 1.0524 | 0.16165 | 0.018089 | 1.2565 | 1.2565 | 0.005287
2019-03-21 12:36 | 1553186188.3739 | 81 | 0.20579 | 0.003411 | 4.4e-05 | 100.794 | 89.5362 | 114.8346 | 21.2431 | 0.003089 | 0.0028 | 258.0732 | 0.27176 | 8.4017 | 0.003765 | 0.005395 | 0.005331 | 5.7e-05 | 0.001062 | 0.29679 | 1.3143 | 2.5481 | 99.38 | 20.94 | 37.466 | 96.1228 | 112.5053 | 89.1536 | 0.009352 | 0.97445 | 0.023139 | 23.7457 | 1.0593 | 0.16167 | 0.018266 | 1.264 | 1.264 | 0.005257
2019-03-21 12:36 | 1553186189.3754 | 82 | 0.20572 | 0.003409 | 4.4e-05 | 100.7987 | 89.5179 | 114.5769 | 21.2425 | 0.003081 | 0.002787 | 257.9924 | 0.27251 | 8.3991 | 0.003596 | 0.005197 | 0.005148 | 5.9e-05 | 0.001061 | 0.29966 | 1.2952 | 2.548 | 99.38 | 20.94 | 37.468 | 95.9358 | 108.3939 | 88.8569 | 0.009353 | 0.97525 | 0.023122 | 23.7093 | 1.0616 | 0.16212 | 0.018691 | 1.2925 | 1.2925 | 0.005155
2019-03-21 12:36 | 1553186190.375 | 83 | 0.20571 | 0.003411 | 4.4e-05 | 100.7472 | 89.5131 | 114.5002 | 21.2434 | 0.003129 | 0.002791 | 257.9322 | 0.27177 | 8.4 | 0.003659 | 0.005031 | 0.004962 | 6.1e-05 | 0.001068 | 0.24978 | 1.3321 | 2.5481 | 99.38 | 20.94 | 37.474 | 97.0953 | 109.9337 | 88.9516 | 0.009354 | 0.97499 | 0.023112 | 23.7044 | 1.0575 | 0.16168 | 0.018742 | 1.2953 | 1.2953 | 0.00513
2019-03-21 12:36 | 1553186191.3741 | 84 | 0.20574 | 0.003402 | 4.5e-05 | 100.7189 | 89.0608 | 114.852 | 21.2443 | 0.00304 | 0.002807 | 257.8419 | 0.27236 | 8.4052 | 0.003667 | 0.005485 | 0.005418 | 5.9e-05 | 0.001074 | 0.28152 | 1.312 | 2.5478 | 99.38 | 20.94 | 37.476 | 94.9542 | 110.1199 | 89.3376 | 0.009355 | 0.97395 | 0.023135 | 23.7535 | 1.1003 | 0.16203 | 0.01858 | 1.2854 | 1.2854 | 0.005181
2019-03-21 12:37 | 1553186192.3753 | 85 | 0.20567 | 0.003405 | 4.5e-05 | 100.782 | 89.3258 | 114.8323 | 21.2442 | 0.003086 | 0.002831 | 257.8062 | 0.27126 | 8.4057 | 0.003656 | 0.005605 | 0.005482 | 6.1e-05 | 0.001078 | 0.29175 | 1.3279 | 2.5479 | 99.38 | 20.94 | 37.48 | 96.0635 | 109.8626 | 89.9008 | 0.009356 | 0.97244 | 0.023115 | 23.7699 | 1.0793 | 0.16137 | 0.01897 | 1.3113 | 1.3113 | 0.005058
2019-03-21 12:37 | 1553186193.3742 | 86 | 0.20561 | 0.003408 | 4.6e-05 | 100.6903 | 89.3912 | 114.6145 | 21.2437 | 0.003172 | 0.002822 | 257.731 | 0.27116 | 8.4025 | 0.003652 | 0.005148 | 0.00506 | 5.9e-05 | 0.001079 | 0.2912 | 1.2819 | 2.548 | 99.38 | 20.94 | 37.483 | 98.1412 | 109.7746 | 89.6893 | 0.009356 | 0.97301 | 0.0231 | 23.7404 | 1.0646 | 0.16132 | 0.019364 | 1.3376 | 1.3376 | 0.004957
2019-03-21 12:37 | 1553186194.3753 | 87 | 0.20557 | 0.003408 | 4.6e-05 | 100.7281 | 89.1207 | 114.3319 | 21.2434 | 0.003103 | 0.00286 | 257.6679 | 0.2724 | 8.4003 | 0.003711 | 0.005255 | 0.005231 | 6e-05 | 0.00108 | 0.25033 | 1.2973 | 2.548 | 99.38 | 20.94 | 37.484 | 96.4576 | 111.1991 | 90.6074 | 0.009357 | 0.97055 | 0.023042 | 23.7411 | 1.0951 | 0.16205 | 0.0196 | 1.3505 | 1.3505 | 0.004932
2019-03-21 12:37 | 1553186195.3752 | 88 | 0.20553 | 0.003413 | 4.6e-05 | 100.7576 | 88.9634 | 114.1161 | 21.243 | 0.003112 | 0.002776 | 257.6227 | 0.27078 | 8.401 | 0.003736 | 0.005188 | 0.005129 | 5.8e-05 | 0.001082 | 0.28335 | 1.3114 | 2.548 | 99.38 | 20.94 | 37.487 | 96.6867 | 111.8128 | 88.5894 | 0.009357 | 0.97597 | 0.023084 | 23.6528 | 1.1135 | 0.16109 | 0.019841 | 1.3696 | 1.3696 | 0.004834
2019-03-21 12:37 | 1553186196.3744 | 89 | 0.20554 | 0.003429 | 4.7e-05 | 100.7755 | 89.0475 | 114.0093 | 21.2436 | 0.003191 | 0.002818 | 257.5565 | 0.27172 | 8.4009 | 0.003674 | 0.004315 | 0.00423 | 5.6e-05 | 0.001078 | 0.26128 | 1.3254 | 2.5481 | 99.38 | 20.94 | 37.487 | 98.5988 | 110.3037 | 89.5869 | 0.009357 | 0.97328 | 0.023042 | 23.6743 | 1.1067 | 0.16165 | 0.019744 | 1.3603 | 1.3603 | 0.004884
2019-03-21 12:37 | 1553186197.3737 | 90 | 0.2055 | 0.003433 | 4.7e-05 | 100.7799 | 89.4535 | 113.8434 | 21.244 | 0.003085 | 0.002842 | 257.5026 | 0.27224 | 8.404 | 0.003711 | 0.005515 | 0.005403 | 5.7e-05 | 0.001078 | 0.27544 | 1.3285 | 2.5481 | 99.38 | 20.94 | 37.486 | 96.021 | 111.2043 | 90.1662 | 0.009357 | 0.97173 | 0.023007 | 23.676 | 1.0664 | 0.16195 | 0.01995 | 1.3724 | 1.3724 | 0.00485
2019-03-21 12:37 | 1553186198.3734 | 91 | 0.20547 | 0.003443 | 4.7e-05 | 100.7117 | 89.0931 | 114.1997 | 21.2443 | 0.003132 | 0.002794 | 257.4167 | 0.27183 | 8.4025 | 0.003633 | 0.004931 | 0.004907 | 5.4e-05 | 0.001079 | 0.2724 | 1.3276 | 2.548 | 99.38 | 20.94 | 37.489 | 97.1752 | 109.3086 | 89.025 | 0.009358 | 0.97479 | 0.023079 | 23.6757 | 1.0964 | 0.16171 | 0.020121 | 1.3885 | 1.3885 | 0.004787
2019-03-21 12:37 | 1553186199.3735 | 92 | 0.20547 | 0.003454 | 4.8e-05 | 100.6989 | 88.6599 | 114.2607 | 21.2439 | 0.003105 | 0.00281 | 257.3576 | 0.27166 | 8.4022 | 0.003631 | 0.005274 | 0.00514 | 5.3e-05 | 0.001078 | 0.28191 | 1.2827 | 2.5481 | 99.38 | 20.94 | 37.49 | 96.5098 | 109.2545 | 89.3935 | 0.009358 | 0.9738 | 0.023073 | 23.6941 | 1.1389 | 0.16161 | 0.020141 | 1.3896 | 1.3896 | 0.00478
2019-03-21 12:37 | 1553186200.3746 | 93 | 0.20544 | 0.003464 | 4.9e-05 | 100.7847 | 89.1017 | 114.452 | 21.2451 | 0.003081 | 0.002798 | 257.2676 | 0.27136 | 8.4022 | 0.003646 | 0.004857 | 0.004793 | 5.6e-05 | 0.001077 | 0.25426 | 1.3341 | 2.5481 | 99.38 | 20.94 | 37.491 | 95.9313 | 109.608 | 89.1071 | 0.009358 | 0.97457 | 0.023102 | 23.7045 | 1.102 | 0.16143 | 0.020291 | 1.4015 | 1.4015 | 0.004734
2019-03-21 12:37 | 1553186201.3735 | 94 | 0.20536 | 0.003476 | 4.9e-05 | 100.8252 | 89.0727 | 114.6072 | 21.2433 | 0.003105 | 0.002848 | 257.2038 | 0.27047 | 8.404 | 0.003686 | 0.005043 | 0.004998 | 5.7e-05 | 0.00108 | 0.26681 | 1.2999 | 2.5478 | 99.38 | 20.94 | 37.49 | 96.5075 | 110.5922 | 90.3134 | 0.009358 | 0.97134 | 0.023079 | 23.7601 | 1.1085 | 0.1609 | 0.020761 | 1.4326 | 1.4326 | 0.004616
2019-03-21 12:37 | 1553186202.3747 | 95 | 0.20537 | 0.003488 | 4.8e-05 | 100.7174 | 88.7233 | 114.441 | 21.2439 | 0.003164 | 0.002822 | 257.1306 | 0.27086 | 8.4023 | 0.003631 | 0.005063 | 0.00505 | 5.5e-05 | 0.00108 | 0.25426 | 1.3119 | 2.5481 | 99.38 | 20.94 | 37.49 | 97.95 | 109.2458 | 89.6895 | 0.009358 | 0.97301 | 0.023082 | 23.7224 | 1.1341 | 0.16113 | 0.020691 | 1.428 | 1.428 | 0.004638
2019-03-21 12:37 | 1553186203.3737 | 96 | 0.20534 | 0.003507 | 4.9e-05 | 100.732 | 88.9423 | 114.4746 | 21.2439 | 0.00317 | 0.002839 | 257.051 | 0.27097 | 8.4026 | 0.003667 | 0.005327 | 0.00521 | 5.3e-05 | 0.00108 | 0.27776 | 1.2768 | 2.548 | 99.38 | 20.95 | 37.489 | 98.0761 | 110.1313 | 90.0897 | 0.009364 | 0.97193 | 0.023073 | 23.739 | 1.1134 | 0.1612 | 0.020857 | 1.4388 | 1.4388 | 0.004605
2019-03-21 12:37 | 1553186204.3747 | 97 | 0.20535 | 0.003523 | 4.8e-05 | 100.7204 | 89.2863 | 114.4964 | 21.2442 | 0.003198 | 0.002788 | 257.008 | 0.27132 | 8.4005 | 0.003586 | 0.005635 | 0.005501 | 5.6e-05 | 0.001079 | 0.29231 | 1.306 | 2.5479 | 99.38 | 20.95 | 37.489 | 98.7687 | 108.1694 | 88.8727 | 0.009364 | 0.9752 | 0.023114 | 23.7015 | 1.0778 | 0.16141 | 0.020754 | 1.4343 | 1.4343 | 0.004625
2019-03-21 12:37 | 1553186205.3741 | 98 | 0.20535 | 0.003534 | 4.9e-05 | 100.7704 | 89.6755 | 114.4664 | 21.2448 | 0.003166 | 0.002824 | 256.907 | 0.27183 | 8.4046 | 0.003634 | 0.004991 | 0.004924 | 5.9e-05 | 0.001076 | 0.30414 | 1.3433 | 2.5481 | 99.38 | 20.94 | 37.488 | 97.997 | 109.3185 | 89.7366 | 0.009358 | 0.97288 | 0.023083 | 23.7266 | 1.0434 | 0.16171 | 0.020746 | 1.4318 | 1.4318 | 0.004642
2019-03-21 12:37 | 1553186206.3739 | 99 | 0.20525 | 0.003545 | 4.9e-05 | 100.7552 | 89.4773 | 114.4793 | 21.2438 | 0.003177 | 0.002859 | 256.8717 | 0.2716 | 8.4047 | 0.003697 | 0.005605 | 0.005556 | 6.1e-05 | 0.001075 | 0.25984 | 1.3255 | 2.548 | 99.38 | 20.94 | 37.488 | 98.2502 | 110.8673 | 90.5741 | 0.009358 | 0.97064 | 0.023058 | 23.7553 | 1.0618 | 0.16157 | 0.021313 | 1.4693 | 1.4693 | 0.00452
2019-03-21 12:37 | 1553186207.3749 | 100 | 0.20521 | 0.003558 | 5e-05 | 100.785 | 89.2641 | 114.4218 | 21.2437 | 0.003132 | 0.002865 | 256.8121 | 0.27026 | 8.4006 | 0.003606 | 0.005477 | 0.00542 | 5.9e-05 | 0.001074 | 0.27743 | 1.2972 | 2.5479 | 99.38 | 20.94 | 37.487 | 97.1653 | 108.6355 | 90.7265 | 0.009357 | 0.97023 | 0.023047 | 23.7543 | 1.0858 | 0.16078 | 0.021559 | 1.4855 | 1.4855 | 0.004448
2019-03-21 12:37 | 1553186208.374 | 101 | 0.20514 | 0.003572 | 5e-05 | 100.7496 | 88.6262 | 114.2774 | 21.2442 | 0.003107 | 0.002851 | 256.7305 | 0.27235 | 8.4003 | 0.003691 | 0.005118 | 0.005089 | 5.7e-05 | 0.001076 | 0.28241 | 1.2885 | 2.548 | 99.38 | 20.94 | 37.487 | 96.5743 | 110.72 | 90.3836 | 0.009357 | 0.97115 | 0.023044 | 23.7281 | 1.1468 | 0.16202 | 0.021944 | 1.5118 | 1.5118 | 0.004405
2019-03-21 12:37 | 1553186209.3744 | 102 | 0.20514 | 0.003589 | 5.2e-05 | 100.7248 | 88.2192 | 114.253 | 21.2442 | 0.003174 | 0.002876 | 256.6944 | 0.27055 | 8.4012 | 0.003614 | 0.004923 | 0.004827 | 5.6e-05 | 0.001076 | 0.2975 | 1.3334 | 2.5481 | 99.38 | 20.94 | 37.486 | 98.1878 | 108.8481 | 90.9872 | 0.009357 | 0.96954 | 0.023022 | 23.7453 | 1.1858 | 0.16095 | 0.021955 | 1.5111 | 1.5111 | 0.004378
2019-03-21 12:37 | 1553186210.3744 | 103 | 0.20507 | 0.003606 | 5.2e-05 | 100.8011 | 87.8187 | 114.1178 | 21.2438 | 0.003164 | 0.002872 | 256.6475 | 0.27043 | 8.4 | 0.003636 | 0.005155 | 0.005096 | 5.9e-05 | 0.001078 | 0.28423 | 1.336 | 2.5481 | 99.38 | 20.94 | 37.486 | 97.9467 | 109.3671 | 90.8954 | 0.009357 | 0.96978 | 0.023011 | 23.7283 | 1.2332 | 0.16088 | 0.022345 | 1.5371 | 1.5371 | 0.004302
2019-03-21 12:37 | 1553186211.3732 | 104 | 0.2051 | 0.003625 | 5.2e-05 | 100.8097 | 88.3874 | 113.7908 | 21.2443 | 0.003207 | 0.002895 | 256.6005 | 0.27109 | 8.4021 | 0.003625 | 0.005679 | 0.005586 | 5.6e-05 | 0.00107 | 0.30127 | 1.3366 | 2.548 | 99.38 | 20.94 | 37.485 | 98.9907 | 109.1042 | 91.4484 | 0.009357 | 0.96831 | 0.022961 | 23.7122 | 1.1763 | 0.16127 | 0.022132 | 1.5191 | 1.5191 | 0.004363
2019-03-21 12:37 | 1553186212.3735 | 105 | 0.20502 | 0.003647 | 5.3e-05 | 100.8217 | 88.6729 | 113.7277 | 21.2442 | 0.003144 | 0.002861 | 256.5253 | 0.27093 | 8.403 | 0.003639 | 0.005152 | 0.005151 | 5.7e-05 | 0.001066 | 0.25984 | 1.3206 | 2.5479 | 99.38 | 20.94 | 37.485 | 97.4541 | 109.4598 | 90.6238 | 0.009357 | 0.97051 | 0.02298 | 23.6788 | 1.1485 | 0.16118 | 0.022594 | 1.5521 | 1.5521 | 0.004268
2019-03-21 12:37 | 1553186213.3742 | 106 | 0.20509 | 0.003663 | 5.3e-05 | 100.8047 | 88.678 | 113.6258 | 21.2442 | 0.003235 | 0.002884 | 256.4895 | 0.27099 | 8.4057 | 0.003709 | 0.004771 | 0.004712 | 5.6e-05 | 0.001064 | 0.2808 | 1.3292 | 2.5481 | 99.38 | 20.94 | 37.483 | 99.651 | 111.151 | 91.1905 | 0.009356 | 0.969 | 0.022952 | 23.6866 | 1.1465 | 0.16121 | 0.022129 | 1.5184 | 1.5184 | 0.004364
2019-03-21 12:37 | 1553186214.3739 | 107 | 0.2052 | 0.003686 | 5.4e-05 | 100.8553 | 88.9726 | 113.6634 | 21.2442 | 0.003211 | 0.002885 | 256.3969 | 0.27158 | 8.4016 | 0.003714 | 0.004966 | 0.00499 | 6e-05 | 0.001065 | 0.28241 | 1.3014 | 2.548 | 99.38 | 20.94 | 37.483 | 99.0793 | 111.2702 | 91.2132 | 0.009356 | 0.96894 | 0.022955 | 23.6913 | 1.1213 | 0.16156 | 0.021458 | 1.4725 | 1.4725 | 0.004509
2019-03-21 12:37 | 1553186215.3733 | 108 | 0.20503 | 0.003705 | 5.4e-05 | 100.8077 | 89.261 | 113.872 | 21.2449 | 0.003165 | 0.002911 | 256.3466 | 0.27128 | 8.4016 | 0.003695 | 0.005386 | 0.005352 | 5.9e-05 | 0.00106 | 0.2557 | 1.3122 | 2.5481 | 99.38 | 20.94 | 37.483 | 97.9761 | 110.8146 | 91.8478 | 0.009356 | 0.96725 | 0.022956 | 23.7337 | 1.0881 | 0.16139 | 0.022466 | 1.5417 | 1.5417 | 0.004302
2019-03-21 12:37 | 1553186216.3734 | 109 | 0.20507 | 0.003719 | 5.5e-05 | 100.8291 | 89.2774 | 113.5427 | 21.2438 | 0.003149 | 0.002857 | 256.2661 | 0.27169 | 8.4028 | 0.003689 | 0.004611 | 0.004543 | 5.9e-05 | 0.001054 | 0.25083 | 1.2866 | 2.5479 | 99.38 | 20.94 | 37.483 | 97.5706 | 110.6613 | 90.5256 | 0.009356 | 0.97077 | 0.022965 | 23.6564 | 1.0884 | 0.16162 | 0.0222 | 1.5239 | 1.5239 | 0.004359
2019-03-21 12:37 | 1553186217.3748 | 110 | 0.205 | 0.003735 | 5.6e-05 | 100.8223 | 88.9797 | 113.5277 | 21.2456 | 0.003169 | 0.002867 | 256.1698 | 0.27083 | 8.4045 | 0.003703 | 0.004501 | 0.004424 | 5.8e-05 | 0.001053 | 0.25337 | 1.2933 | 2.5479 | 99.38 | 20.94 | 37.482 | 98.0662 | 110.9945 | 90.7823 | 0.009356 | 0.97008 | 0.022955 | 23.6631 | 1.1176 | 0.16112 | 0.022627 | 1.5525 | 1.5525 | 0.004265
2019-03-21 12:37 | 1553186218.3749 | 111 | 0.20497 | 0.003753 | 5.7e-05 | 100.8348 | 88.9937 | 113.0841 | 21.2446 | 0.003205 | 0.002927 | 256.124 | 0.27061 | 8.4057 | 0.003641 | 0.004995 | 0.004885 | 6.1e-05 | 0.001054 | 0.28136 | 1.2736 | 2.548 | 99.38 | 20.94 | 37.483 | 98.933 | 109.4972 | 92.2206 | 0.009356 | 0.96627 | 0.022865 | 23.6635 | 1.1173 | 0.16099 | 0.022756 | 1.5552 | 1.5552 | 0.004254
2019-03-21 12:37 | 1553186219.3735 | 112 | 0.20493 | 0.003775 | 5.6e-05 | 100.7987 | 89.3486 | 113.2413 | 21.2442 | 0.003199 | 0.002937 | 256.0186 | 0.27157 | 8.4014 | 0.003613 | 0.005521 | 0.005488 | 5.9e-05 | 0.001053 | 0.27865 | 1.3328 | 2.5481 | 99.38 | 20.94 | 37.483 | 98.7782 | 108.8197 | 92.4688 | 0.009356 | 0.96561 | 0.022873 | 23.688 | 1.0785 | 0.16156 | 0.022966 | 1.5702 | 1.5702 | 0.004229
2019-03-21 12:37 | 1553186220.3735 | 113 | 0.20497 | 0.003787 | 5.6e-05 | 100.8276 | 89.3591 | 112.8383 | 21.2453 | 0.003216 | 0.002913 | 255.9563 | 0.27143 | 8.4049 | 0.00366 | 0.004844 | 0.004716 | 5.9e-05 | 0.001047 | 0.25841 | 1.3047 | 2.548 | 99.38 | 20.94 | 37.482 | 99.187 | 109.9711 | 91.8772 | 0.009356 | 0.96718 | 0.022851 | 23.6267 | 1.08 | 0.16148 | 0.022703 | 1.5507 | 1.5507 | 0.00428
2019-03-21 12:37 | 1553186221.3739 | 114 | 0.20491 | 0.003807 | 5.6e-05 | 100.8366 | 89.3792 | 112.7186 | 21.2459 | 0.003232 | 0.002895 | 255.867 | 0.27157 | 8.4025 | 0.003607 | 0.005302 | 0.005287 | 6.4e-05 | 0.001049 | 0.2907 | 1.3445 | 2.548 | 99.38 | 20.94 | 37.483 | 99.5758 | 108.6748 | 91.4548 | 0.009356 | 0.9683 | 0.022852 | 23.6004 | 1.0788 | 0.16156 | 0.023034 | 1.5733 | 1.5733 | 0.00422
2019-03-21 12:37 | 1553186222.3747 | 115 | 0.20491 | 0.003817 | 5.6e-05 | 100.8634 | 89.6037 | 112.8871 | 21.2445 | 0.003244 | 0.002851 | 255.7986 | 0.2706 | 8.4018 | 0.003609 | 0.005367 | 0.005414 | 6.3e-05 | 0.001049 | 0.26394 | 1.3065 | 2.5479 | 99.38 | 20.94 | 37.483 | 99.8701 | 108.7089 | 90.3974 | 0.009356 | 0.97111 | 0.022903 | 23.5838 | 1.0588 | 0.16098 | 0.023028 | 1.5764 | 1.5764 | 0.004197
2019-03-21 12:37 | 1553186223.3751 | 116 | 0.20492 | 0.003834 | 5.6e-05 | 100.7981 | 89.338 | 113.0533 | 21.2451 | 0.003215 | 0.002965 | 255.7094 | 0.27202 | 8.4019 | 0.003662 | 0.005041 | 0.004819 | 6.2e-05 | 0.001045 | 0.28838 | 1.3267 | 2.5478 | 99.38 | 20.94 | 37.483 | 99.1703 | 110.011 | 93.1377 | 0.009356 | 0.96385 | 0.022833 | 23.6899 | 1.0795 | 0.16182 | 0.022971 | 1.5677 | 1.5677 | 0.004242
2019-03-21 12:37 | 1553186224.3748 | 117 | 0.20493 | 0.003847 | 5.6e-05 | 100.9066 | 89.2062 | 113.1735 | 21.2444 | 0.003192 | 0.002917 | 255.6669 | 0.27108 | 8.4033 | 0.003633 | 0.005049 | 0.004989 | 6e-05 | 0.001047 | 0.24674 | 1.2936 | 2.548 | 99.38 | 20.94 | 37.483 | 98.6131 | 109.2912 | 91.9881 | 0.009356 | 0.96688 | 0.022882 | 23.6653 | 1.1024 | 0.16127 | 0.022862 | 1.5636 | 1.5636 | 0.004239
2019-03-21 12:37 | 1553186225.3739 | 118 | 0.20495 | 0.003859 | 5.6e-05 | 100.844 | 89.6606 | 112.9531 | 21.2459 | 0.0033 | 0.002953 | 255.6216 | 0.27171 | 8.3994 | 0.003632 | 0.004697 | 0.004684 | 6.2e-05 | 0.001045 | 0.26128 | 1.2957 | 2.5481 | 99.38 | 20.94 | 37.483 | 101.2411 | 109.2912 | 92.86 | 0.009356 | 0.96458 | 0.022832 | 23.6705 | 1.0514 | 0.16164 | 0.022722 | 1.5507 | 1.5507 | 0.004284
2019-03-21 12:37 | 1553186226.3753 | 119 | 0.20495 | 0.003869 | 5.6e-05 | 100.8595 | 89.7447 | 112.8658 | 21.2451 | 0.003227 | 0.00294 | 255.5245 | 0.27107 | 8.4012 | 0.003607 | 0.004609 | 0.00455 | 5.8e-05 | 0.001043 | 0.28279 | 1.3325 | 2.5481 | 99.38 | 20.94 | 37.482 | 99.4598 | 108.6731 | 92.5331 | 0.009356 | 0.96544 | 0.022833 | 23.6507 | 1.0444 | 0.16126 | 0.02276 | 1.5534 | 1.5534 | 0.004267
2019-03-21 12:37 | 1553186227.373 | 120 | 0.20496 | 0.003876 | 5.6e-05 | 100.8875 | 89.5884 | 112.9791 | 21.2448 | 0.003257 | 0.002987 | 255.5068 | 0.27189 | 8.4016 | 0.003528 | 0.00487 | 0.004808 | 6.2e-05 | 0.001047 | 0.26648 | 1.3695 | 2.5482 | 99.38 | 20.94 | 37.482 | 100.185 | 106.7485 | 93.6616 | 0.009356 | 0.96247 | 0.02281 | 23.6991 | 1.0624 | 0.16174 | 0.022653 | 1.5445 | 1.5445 | 0.004304
2019-03-21 12:37 | 1553186228.3731 | 121 | 0.20492 | 0.003885 | 5.5e-05 | 100.8461 | 89.3551 | 113.4308 | 21.2453 | 0.003265 | 0.002964 | 255.4212 | 0.27155 | 8.4023 | 0.003555 | 0.005083 | 0.005058 | 5.8e-05 | 0.001047 | 0.26128 | 1.3412 | 2.548 | 99.38 | 20.94 | 37.482 | 100.3757 | 107.4128 | 93.1184 | 0.009356 | 0.9639 | 0.022872 | 23.7288 | 1.0821 | 0.16155 | 0.022885 | 1.5646 | 1.5646 | 0.004244
2019-03-21 12:37 | 1553186229.3734 | 122 | 0.20493 | 0.003888 | 5.5e-05 | 100.8783 | 89.554 | 113.3 | 21.2461 | 0.003273 | 0.002894 | 255.3705 | 0.27155 | 8.4036 | 0.003594 | 0.00512 | 0.004952 | 6e-05 | 0.001045 | 0.27505 | 1.3116 | 2.548 | 99.38 | 20.94 | 37.482 | 100.5779 | 108.3605 | 91.4222 | 0.009356 | 0.96838 | 0.022912 | 23.6602 | 1.0651 | 0.16154 | 0.02283 | 1.5636 | 1.5636 | 0.004246
2019-03-21 12:37 | 1553186230.3746 | 123 | 0.20495 | 0.003895 | 5.5e-05 | 100.8604 | 89.8091 | 113.4474 | 21.2457 | 0.003312 | 0.002945 | 255.3025 | 0.27136 | 8.4011 | 0.003607 | 0.005732 | 0.005692 | 6.2e-05 | 0.001049 | 0.25946 | 1.3417 | 2.5482 | 99.38 | 20.94 | 37.481 | 101.5165 | 108.6635 | 92.6691 | 0.009356 | 0.96508 | 0.022888 | 23.716 | 1.038 | 0.16143 | 0.022702 | 1.5532 | 1.5532 | 0.004272
2019-03-21 12:37 | 1553186231.3731 | 124 | 0.20495 | 0.003903 | 5.6e-05 | 100.911 | 90.4136 | 113.5002 | 21.2452 | 0.003232 | 0.002931 | 255.2241 | 0.27074 | 8.4035 | 0.00353 | 0.005254 | 0.005229 | 5.9e-05 | 0.001046 | 0.28783 | 1.3325 | 2.548 | 99.38 | 20.94 | 37.485 | 99.5885 | 106.8009 | 92.3126 | 0.009357 | 0.96602 | 0.022904 | 23.71 | 0.98251 | 0.16106 | 0.022705 | 1.5544 | 1.5544 | 0.004259
2019-03-21 12:37 | 1553186232.3742 | 125 | 0.20492 | 0.003899 | 5.6e-05 | 100.8765 | 90.3547 | 113.3899 | 21.2456 | 0.003264 | 0.002999 | 255.1603 | 0.27174 | 8.4029 | 0.00356 | 0.005253 | 0.005172 | 6.4e-05 | 0.001047 | 0.25658 | 1.3392 | 2.548 | 99.38 | 20.94 | 37.485 | 100.3655 | 107.5407 | 93.966 | 0.009357 | 0.96167 | 0.022842 | 23.752 | 0.98528 | 0.16166 | 0.022869 | 1.5613 | 1.5613 | 0.004255
2019-03-21 12:37 | 1553186233.3749 | 126 | 0.20492 | 0.003907 | 5.6e-05 | 100.8434 | 89.8741 | 112.8184 | 21.2453 | 0.003274 | 0.002963 | 255.1022 | 0.27106 | 8.4038 | 0.003642 | 0.004902 | 0.004824 | 6.3e-05 | 0.001045 | 0.24563 | 1.2899 | 2.548 | 99.38 | 20.94 | 37.485 | 100.6093 | 109.527 | 93.08 | 0.009357 | 0.964 | 0.022812 | 23.6635 | 1.03 | 0.16125 | 0.022841 | 1.5574 | 1.5574 | 0.004255
2019-03-21 12:37 | 1553186234.3748 | 127 | 0.20494 | 0.003912 | 5.5e-05 | 100.8276 | 89.6711 | 112.6576 | 21.2445 | 0.003255 | 0.002978 | 255.0238 | 0.2708 | 8.4066 | 0.003607 | 0.005487 | 0.005447 | 6e-05 | 0.001046 | 0.27616 | 1.3222 | 2.5481 | 99.38 | 20.94 | 37.486 | 100.1394 | 108.682 | 93.4487 | 0.009357 | 0.96303 | 0.022784 | 23.6585 | 1.0489 | 0.1611 | 0.02276 | 1.5501 | 1.5501 | 0.004272
2019-03-21 12:37 | 1553186235.3748 | 128 | 0.20494 | 0.003918 | 5.5e-05 | 100.8544 | 89.6483 | 112.6893 | 21.2452 | 0.003294 | 0.00295 | 254.9152 | 0.27032 | 8.4036 | 0.003614 | 0.00516 | 0.00513 | 5.8e-05 | 0.001047 | 0.29358 | 1.2932 | 2.5481 | 99.38 | 20.94 | 37.488 | 101.0732 | 108.8536 | 92.7669 | 0.009358 | 0.96482 | 0.022808 | 23.6398 | 1.0535 | 0.16082 | 0.022747 | 1.5508 | 1.5508 | 0.004262
2019-03-21 12:37 | 1553186236.3744 | 129 | 0.20495 | 0.003921 | 5.6e-05 | 100.925 | 89.8336 | 112.7512 | 21.2449 | 0.003277 | 0.002948 | 254.858 | 0.27078 | 8.4028 | 0.003576 | 0.004861 | 0.004788 | 5.9e-05 | 0.001048 | 0.27417 | 1.3078 | 2.548 | 99.38 | 20.94 | 37.492 | 100.6679 | 107.9219 | 92.7167 | 0.009359 | 0.96496 | 0.022816 | 23.6447 | 1.0413 | 0.16109 | 0.022649 | 1.5447 | 1.5447 | 0.004286
2019-03-21 12:37 | 1553186237.3736 | 130 | 0.20497 | 0.003922 | 5.5e-05 | 100.8801 | 89.7873 | 112.6153 | 21.2459 | 0.003289 | 0.002974 | 254.7623 | 0.27145 | 8.4002 | 0.003564 | 0.005092 | 0.005046 | 6.3e-05 | 0.001049 | 0.25553 | 1.2772 | 2.5479 | 99.38 | 20.94 | 37.494 | 100.9686 | 107.6253 | 93.3482 | 0.009359 | 0.96329 | 0.022783 | 23.6508 | 1.0419 | 0.16149 | 0.022586 | 1.5381 | 1.5381 | 0.004315
2019-03-21 12:37 | 1553186238.3746 | 131 | 0.20501 | 0.003925 | 5.5e-05 | 100.9042 | 89.9044 | 112.5214 | 21.2461 | 0.00329 | 0.002962 | 254.6748 | 0.27163 | 8.4018 | 0.003579 | 0.004627 | 0.004596 | 6.3e-05 | 0.001048 | 0.26842 | 1.3176 | 2.548 | 99.38 | 20.94 | 37.495 | 100.9821 | 107.9996 | 93.0682 | 0.009359 | 0.96403 | 0.022782 | 23.6319 | 1.0324 | 0.16159 | 0.022333 | 1.5208 | 1.5208 | 0.004367
2019-03-21 12:37 | 1553186239.3736 | 132 | 0.20499 | 0.00393 | 5.5e-05 | 100.9373 | 89.977 | 112.6494 | 21.2456 | 0.003303 | 0.002979 | 254.6214 | 0.27145 | 8.4044 | 0.003585 | 0.005402 | 0.005308 | 6.1e-05 | 0.001053 | 0.28208 | 1.3196 | 2.548 | 99.38 | 20.94 | 37.495 | 101.3139 | 108.1385 | 93.4707 | 0.009359 | 0.96297 | 0.022782 | 23.6583 | 1.0281 | 0.16149 | 0.022425 | 1.5271 | 1.5271 | 0.004346
2019-03-21 12:37 | 1553186240.3737 | 133 | 0.20501 | 0.003929 | 5.4e-05 | 100.9197 | 90.0713 | 112.7937 | 21.2444 | 0.003298 | 0.002939 | 254.5438 | 0.27072 | 8.4044 | 0.003569 | 0.005528 | 0.005526 | 6.2e-05 | 0.001052 | 0.27959 | 1.3375 | 2.548 | 99.38 | 20.94 | 37.497 | 101.1872 | 107.7537 | 92.5129 | 0.00936 | 0.96549 | 0.022827 | 23.6425 | 1.0172 | 0.16105 | 0.022333 | 1.5239 | 1.5239 | 0.004344
2019-03-21 12:37 | 1553186241.3748 | 134 | 0.20502 | 0.003928 | 5.5e-05 | 100.8771 | 90.1663 | 112.805 | 21.2463 | 0.003296 | 0.00296 | 254.5003 | 0.27027 | 8.4022 | 0.003627 | 0.005436 | 0.005326 | 6e-05 | 0.001054 | 0.2557 | 1.3643 | 2.5481 | 99.38 | 20.94 | 37.501 | 101.1442 | 109.1543 | 93.0137 | 0.009361 | 0.96417 | 0.022812 | 23.6599 | 1.004 | 0.16079 | 0.02225 | 1.5172 | 1.5172 | 0.004356
2019-03-21 12:37 | 1553186242.3731 | 135 | 0.20504 | 0.003926 | 5.5e-05 | 100.8774 | 90.4112 | 112.8324 | 21.2455 | 0.003262 | 0.003035 | 254.3984 | 0.27118 | 8.4038 | 0.003702 | 0.005206 | 0.005116 | 6.3e-05 | 0.001054 | 0.24834 | 1.3322 | 2.548 | 99.38 | 20.94 | 37.502 | 100.3053 | 110.9716 | 94.8183 | 0.009361 | 0.95944 | 0.022759 | 23.721 | 0.97976 | 0.16133 | 0.022104 | 1.5037 | 1.5037 | 0.004409
2019-03-21 12:37 | 1553186243.3738 | 136 | 0.20503 | 0.003929 | 5.5e-05 | 100.9492 | 90.8438 | 113.2794 | 21.2458 | 0.00324 | 0.002931 | 254.3319 | 0.27084 | 8.4048 | 0.003627 | 0.005326 | 0.005291 | 5.9e-05 | 0.001054 | 0.25642 | 1.3273 | 2.5478 | 99.38 | 20.94 | 37.503 | 99.7759 | 109.165 | 92.3204 | 0.009361 | 0.966 | 0.022882 | 23.6871 | 0.94344 | 0.16112 | 0.022166 | 1.5161 | 1.5161 | 0.004368
2019-03-21 12:37 | 1553186244.3739 | 137 | 0.20504 | 0.003922 | 5.5e-05 | 100.8985 | 90.7457 | 113.2207 | 21.2465 | 0.003285 | 0.002953 | 254.2582 | 0.27057 | 8.4045 | 0.003755 | 0.005165 | 0.005105 | 5.9e-05 | 0.001059 | 0.25227 | 1.3634 | 2.5479 | 99.38 | 20.93 | 37.508 | 100.8703 | 112.2711 | 92.8575 | 0.009357 | 0.96459 | 0.022859 | 23.6984 | 0.94861 | 0.16096 | 0.022142 | 1.5129 | 1.5129 | 0.004373
2019-03-21 12:37 | 1553186245.3732 | 138 | 0.20504 | 0.003925 | 5.5e-05 | 100.8729 | 90.6936 | 113.1709 | 21.2463 | 0.003249 | 0.00294 | 254.196 | 0.27127 | 8.4006 | 0.003604 | 0.005124 | 0.005094 | 5.9e-05 | 0.001056 | 0.2385 | 1.3366 | 2.5481 | 99.38 | 20.94 | 37.509 | 99.9952 | 108.598 | 92.5366 | 0.009363 | 0.96543 | 0.022864 | 23.6828 | 0.95148 | 0.16138 | 0.022153 | 1.514 | 1.514 | 0.004381
2019-03-21 12:37 | 1553186246.373 | 139 | 0.205 | 0.003924 | 5.5e-05 | 100.9408 | 90.3642 | 113.4427 | 21.2458 | 0.00325 | 0.002979 | 254.1176 | 0.27036 | 8.4025 | 0.003616 | 0.005852 | 0.005814 | 6.2e-05 | 0.001061 | 0.29015 | 1.3365 | 2.5481 | 99.38 | 20.93 | 37.51 | 100.0276 | 108.8883 | 93.481 | 0.009357 | 0.96295 | 0.022862 | 23.7418 | 0.99004 | 0.16084 | 0.022351 | 1.5274 | 1.5274 | 0.004328
2019-03-21 12:37 | 1553186247.3734 | 140 | 0.20497 | 0.003918 | 5.5e-05 | 100.8443 | 90.1289 | 113.4843 | 21.2453 | 0.003261 | 0.003004 | 254.0269 | 0.27163 | 8.4006 | 0.00361 | 0.004341 | 0.004296 | 6.2e-05 | 0.001059 | 0.27168 | 1.313 | 2.548 | 99.38 | 20.93 | 37.513 | 100.2894 | 108.7404 | 94.0785 | 0.009358 | 0.96138 | 0.022848 | 23.7655 | 1.0048 | 0.16159 | 0.022539 | 1.5393 | 1.5393 | 0.004315
2019-03-21 12:37 | 1553186248.3731 | 141 | 0.20501 | 0.003916 | 5.4e-05 | 100.8482 | 91.1789 | 113.3441 | 21.246 | 0.0033 | 0.003018 | 253.9958 | 0.27177 | 8.4028 | 0.003657 | 0.005528 | 0.005396 | 6e-05 | 0.001057 | 0.28119 | 1.2771 | 2.5481 | 99.38 | 20.93 | 37.515 | 101.2398 | 109.8843 | 94.4257 | 0.009359 | 0.96047 | 0.022823 | 23.7621 | 0.90155 | 0.16168 | 0.022304 | 1.5217 | 1.5217 | 0.004367
2019-03-21 12:37 | 1553186249.3741 | 142 | 0.20499 | 0.003918 | 5.3e-05 | 100.8735 | 91.0944 | 113.1122 | 21.2457 | 0.003299 | 0.002969 | 253.8964 | 0.27014 | 8.4013 | 0.003596 | 0.004487 | 0.004474 | 6.1e-05 | 0.001059 | 0.24834 | 1.3057 | 2.548 | 99.38 | 20.93 | 37.518 | 101.201 | 108.4177 | 93.2323 | 0.009359 | 0.9636 | 0.022836 | 23.6992 | 0.91208 | 0.16071 | 0.022431 | 1.5313 | 1.5313 | 0.004313
2019-03-21 12:37 | 1553186250.3731 | 143 | 0.20505 | 0.003916 | 5.4e-05 | 100.8354 | 90.9889 | 113.2338 | 21.2466 | 0.003354 | 0.002986 | 253.8388 | 0.27078 | 8.4019 | 0.0037 | 0.005731 | 0.005715 | 6.2e-05 | 0.001061 | 0.29645 | 1.3148 | 2.5481 | 99.38 | 20.93 | 37.521 | 102.5273 | 110.9404 | 93.647 | 0.00936 | 0.96251 | 0.022836 | 23.7253 | 0.91907 | 0.16109 | 0.022099 | 1.5085 | 1.5085 | 0.004389
2019-03-21 12:38 | 1553186251.3737 | 144 | 0.20507 | 0.003921 | 5.5e-05 | 100.8208 | 91.0897 | 113.3303 | 21.2454 | 0.003325 | 0.002977 | 253.7616 | 0.27133 | 8.4005 | 0.003692 | 0.004892 | 0.004852 | 6.2e-05 | 0.001057 | 0.25769 | 1.3197 | 2.548 | 99.38 | 20.93 | 37.524 | 101.8294 | 110.7434 | 93.4266 | 0.009361 | 0.96309 | 0.022852 | 23.7283 | 0.90787 | 0.16142 | 0.021952 | 1.4996 | 1.4996 | 0.004424
2019-03-21 12:38 | 1553186252.3733 | 145 | 0.20511 | 0.003916 | 5.4e-05 | 100.9113 | 91.4956 | 113.3521 | 21.2446 | 0.003384 | 0.003001 | 253.7047 | 0.27107 | 8.4014 | 0.003639 | 0.00519 | 0.005018 | 6.4e-05 | 0.001058 | 0.25802 | 1.3116 | 2.548 | 99.38 | 20.93 | 37.527 | 103.2748 | 109.4444 | 94.0143 | 0.009362 | 0.96155 | 0.022836 | 23.7496 | 0.87613 | 0.16126 | 0.02171 | 1.482 | 1.482 | 0.004472
2019-03-21 12:38 | 1553186253.3733 | 146 | 0.20515 | 0.003915 | 5.4e-05 | 100.892 | 91.4633 | 113.2296 | 21.2461 | 0.003337 | 0.002977 | 253.607 | 0.27064 | 8.405 | 0.003647 | 0.005455 | 0.005417 | 6.4e-05 | 0.00106 | 0.262 | 1.2923 | 2.5478 | 99.38 | 20.93 | 37.528 | 102.1201 | 109.6399 | 93.4328 | 0.009362 | 0.96307 | 0.022842 | 23.718 | 0.87757 | 0.161 | 0.021512 | 1.4689 | 1.4689 | 0.004505
2019-03-21 12:38 | 1553186254.3737 | 147 | 0.20512 | 0.003912 | 5.4e-05 | 100.8476 | 90.684 | 113.3707 | 21.246 | 0.003347 | 0.002955 | 253.5335 | 0.27177 | 8.4027 | 0.003714 | 0.005231 | 0.005178 | 6.9e-05 | 0.001062 | 0.26272 | 1.3041 | 2.5479 | 99.38 | 20.93 | 37.529 | 102.3751 | 111.2719 | 92.9054 | 0.009362 | 0.96446 | 0.022873 | 23.7156 | 0.95017 | 0.16168 | 0.021682 | 1.4825 | 1.4825 | 0.004482
2019-03-21 12:38 | 1553186255.3749 | 148 | 0.20515 | 0.003908 | 5.4e-05 | 100.8616 | 90.9667 | 113.404 | 21.2462 | 0.003269 | 0.003013 | 253.4821 | 0.27139 | 8.4017 | 0.003694 | 0.005147 | 0.005057 | 6.7e-05 | 0.00106 | 0.27975 | 1.311 | 2.5479 | 99.38 | 20.93 | 37.531 | 100.4707 | 110.7814 | 94.3056 | 0.009363 | 0.96078 | 0.022833 | 23.7645 | 0.92357 | 0.16145 | 0.021478 | 1.4659 | 1.4659 | 0.004526
2019-03-21 12:38 | 1553186256.3735 | 149 | 0.20517 | 0.0039 | 5.4e-05 | 100.923 | 91.3284 | 113.251 | 21.2465 | 0.003316 | 0.002989 | 253.3941 | 0.27018 | 8.4059 | 0.003679 | 0.005536 | 0.005512 | 6.7e-05 | 0.001056 | 0.24602 | 1.3146 | 2.5479 | 99.38 | 20.93 | 37.532 | 101.624 | 110.4217 | 93.7141 | 0.009363 | 0.96233 | 0.022835 | 23.7293 | 0.89352 | 0.16073 | 0.02138 | 1.4594 | 1.4594 | 0.004526
2019-03-21 12:38 | 1553186257.3748 | 150 | 0.20516 | 0.003896 | 5.4e-05 | 100.8664 | 91.1258 | 113.2735 | 21.2467 | 0.003352 | 0.003009 | 253.339 | 0.27124 | 8.403 | 0.003684 | 0.005455 | 0.005355 | 6.5e-05 | 0.001061 | 0.25752 | 1.3217 | 2.548 | 99.38 | 20.93 | 37.534 | 102.5015 | 110.5569 | 94.2025 | 0.009363 | 0.96105 | 0.022823 | 23.7474 | 0.90837 | 0.16136 | 0.021464 | 1.4644 | 1.4644 | 0.004529
2019-03-21 12:38 | 1553186258.3728 | 151 | 0.20515 | 0.003891 | 5.5e-05 | 100.838 | 91.401 | 113.7089 | 21.2465 | 0.003302 | 0.00298 | 253.2296 | 0.27072 | 8.403 | 0.003629 | 0.005873 | 0.005777 | 6.2e-05 | 0.001063 | 0.28368 | 1.2819 | 2.5479 | 99.38 | 20.93 | 37.534 | 101.2879 | 109.2051 | 93.503 | 0.009363 | 0.96289 | 0.022888 | 23.7704 | 0.87889 | 0.16105 | 0.021518 | 1.4722 | 1.4722 | 0.004496
2019-03-21 12:38 | 1553186259.373 | 152 | 0.2052 | 0.003885 | 5.5e-05 | 100.8956 | 91.3615 | 113.7758 | 21.246 | 0.003365 | 0.002994 | 253.1611 | 0.27068 | 8.4043 | 0.003589 | 0.005451 | 0.005446 | 6.7e-05 | 0.001066 | 0.26017 | 1.3011 | 2.5479 | 99.38 | 20.93 | 37.535 | 102.7993 | 108.2438 | 93.827 | 0.009364 | 0.96204 | 0.022885 | 23.7879 | 0.88785 | 0.16103 | 0.02125 | 1.4537 | 1.4537 | 0.004553
2019-03-21 12:38 | 1553186260.3737 | 153 | 0.20518 | 0.00388 | 5.6e-05 | 100.8681 | 91.2453 | 113.5995 | 21.2465 | 0.003312 | 0.003021 | 253.081 | 0.27093 | 8.3995 | 0.003651 | 0.005555 | 0.005442 | 6.5e-05 | 0.001063 | 0.28672 | 1.3557 | 2.548 | 99.38 | 20.93 | 37.536 | 101.5156 | 109.746 | 94.4812 | 0.009364 | 0.96032 | 0.022847 | 23.7906 | 0.8968 | 0.16118 | 0.021337 | 1.4571 | 1.4571 | 0.004546
2019-03-21 12:38 | 1553186261.3728 | 154 | 0.20518 | 0.003872 | 5.7e-05 | 100.9239 | 91.4326 | 113.5059 | 21.2477 | 0.003287 | 0.002995 | 252.9892 | 0.27098 | 8.4022 | 0.003688 | 0.004545 | 0.004505 | 6.3e-05 | 0.00106 | 0.29911 | 1.3459 | 2.5481 | 99.38 | 20.93 | 37.536 | 100.9109 | 110.6542 | 93.8711 | 0.009364 | 0.96192 | 0.022856 | 23.7611 | 0.8834 | 0.1612 | 0.021372 | 1.46 | 1.46 | 0.004538
2019-03-21 12:38 | 1553186262.3732 | 155 | 0.20517 | 0.003864 | 6e-05 | 100.7975 | 91.3083 | 113.5073 | 21.2467 | 0.003312 | 0.003017 | 252.9277 | 0.27009 | 8.4022 | 0.003662 | 0.004988 | 0.00491 | 6.1e-05 | 0.001055 | 0.27184 | 1.3108 | 2.548 | 99.38 | 20.94 | 37.536 | 101.5315 | 110.008 | 94.4023 | 0.00937 | 0.96053 | 0.02284 | 23.7784 | 0.88436 | 0.16068 | 0.021394 | 1.4603 | 1.4603 | 0.004522
2019-03-21 12:38 | 1553186263.3729 | 156 | 0.20517 | 0.003854 | 6.1e-05 | 100.8232 | 91.3148 | 113.4624 | 21.2469 | 0.003267 | 0.002999 | 252.8466 | 0.27101 | 8.401 | 0.003646 | 0.004875 | 0.00491 | 6.2e-05 | 0.001052 | 0.2557 | 1.3572 | 2.548 | 99.38 | 20.94 | 37.536 | 100.4273 | 109.6273 | 93.9517 | 0.00937 | 0.96171 | 0.022849 | 23.7591 | 0.886 | 0.16122 | 0.021435 | 1.4637 | 1.4637 | 0.004527
2019-03-21 12:38 | 1553186264.373 | 157 | 0.20518 | 0.003856 | 6.1e-05 | 100.8234 | 91.3342 | 113.5347 | 21.2468 | 0.003346 | 0.003036 | 252.768 | 0.2706 | 8.403 | 0.003732 | 0.005235 | 0.004999 | 6.4e-05 | 0.001053 | 0.26576 | 1.321 | 2.5479 | 99.38 | 20.94 | 37.536 | 102.35 | 111.7061 | 94.8642 | 0.00937 | 0.95932 | 0.022828 | 23.7962 | 0.88413 | 0.16098 | 0.021397 | 1.4597 | 1.4597 | 0.004533
2019-03-21 12:38 | 1553186265.3741 | 158 | 0.20518 | 0.00385 | 6.1e-05 | 100.8604 | 91.2055 | 113.2735 | 21.2462 | 0.003332 | 0.003044 | 252.6818 | 0.27107 | 8.4024 | 0.003633 | 0.005601 | 0.005616 | 6e-05 | 0.001051 | 0.24906 | 1.2991 | 2.5481 | 99.38 | 20.94 | 37.536 | 102.0132 | 109.293 | 95.0492 | 0.00937 | 0.95884 | 0.022796 | 23.7748 | 0.90002 | 0.16126 | 0.02139 | 1.4572 | 1.4572 | 0.004548
2019-03-21 12:38 | 1553186266.3731 | 159 | 0.20518 | 0.003842 | 6.2e-05 | 100.8112 | 91.5144 | 113.1275 | 21.2468 | 0.00331 | 0.003047 | 252.6258 | 0.27107 | 8.4028 | 0.003693 | 0.005033 | 0.005029 | 6.5e-05 | 0.001052 | 0.2756 | 1.3527 | 2.548 | 99.38 | 20.94 | 37.537 | 101.4689 | 110.7715 | 95.1257 | 0.00937 | 0.95864 | 0.022779 | 23.7619 | 0.86542 | 0.16126 | 0.02139 | 1.4561 | 1.4561 | 0.004552
2019-03-21 12:38 | 1553186267.3732 | 160 | 0.20518 | 0.003835 | 6.1e-05 | 100.8735 | 91.5464 | 113.1636 | 21.2472 | 0.003343 | 0.003047 | 252.5532 | 0.27044 | 8.4021 | 0.003652 | 0.005041 | 0.004935 | 6.3e-05 | 0.001053 | 0.26089 | 1.3307 | 2.5479 | 99.38 | 20.94 | 37.537 | 102.2688 | 109.779 | 95.123 | 0.00937 | 0.95865 | 0.022783 | 23.7656 | 0.86781 | 0.16089 | 0.021421 | 1.4585 | 1.4585 | 0.004534
2019-03-21 12:38 | 1553186268.3732 | 161 | 0.20523 | 0.003835 | 6e-05 | 100.8708 | 91.794 | 112.544 | 21.2475 | 0.00332 | 0.003002 | 252.4985 | 0.27181 | 8.406 | 0.003758 | 0.005242 | 0.005156 | 6.5e-05 | 0.001052 | 0.29103 | 1.3466 | 2.5481 | 99.38 | 20.94 | 37.537 | 101.7264 | 112.3432 | 94.0415 | 0.00937 | 0.96147 | 0.022754 | 23.6657 | 0.84342 | 0.1617 | 0.021122 | 1.4364 | 1.4364 | 0.004627
2019-03-21 12:38 | 1553186269.3742 | 162 | 0.20529 | 0.003826 | 5.9e-05 | 100.8422 | 92.0682 | 112.7076 | 21.2476 | 0.00335 | 0.003038 | 252.4419 | 0.27085 | 8.4022 | 0.003732 | 0.005374 | 0.005317 | 6.2e-05 | 0.001054 | 0.27472 | 1.3302 | 2.548 | 99.38 | 20.94 | 37.537 | 102.4513 | 111.7061 | 94.9089 | 0.00937 | 0.95921 | 0.022744 | 23.7108 | 0.8142 | 0.16113 | 0.020736 | 1.4095 | 1.4095 | 0.004698
2019-03-21 12:38 | 1553186270.3737 | 163 | 0.20529 | 0.003822 | 5.8e-05 | 100.8491 | 92.2269 | 112.7008 | 21.2476 | 0.003328 | 0.002972 | 252.3624 | 0.27139 | 8.4026 | 0.003771 | 0.005513 | 0.005464 | 6e-05 | 0.001048 | 0.25155 | 1.3462 | 2.5481 | 99.38 | 20.94 | 37.536 | 101.9054 | 112.6718 | 93.3053 | 0.00937 | 0.96341 | 0.022793 | 23.6584 | 0.7994 | 0.16145 | 0.020761 | 1.4143 | 1.4143 | 0.004692
2019-03-21 12:38 | 1553186271.3732 | 164 | 0.20533 | 0.003808 | 5.7e-05 | 100.8491 | 92.2238 | 112.5907 | 21.2474 | 0.003351 | 0.003018 | 252.299 | 0.27051 | 8.4029 | 0.003808 | 0.005545 | 0.005464 | 6.2e-05 | 0.001047 | 0.24818 | 1.3216 | 2.5479 | 99.38 | 20.95 | 37.535 | 102.4748 | 113.5636 | 94.4312 | 0.009375 | 0.96046 | 0.022747 | 23.6831 | 0.7997 | 0.16093 | 0.020526 | 1.3955 | 1.3955 | 0.00474
2019-03-21 12:38 | 1553186272.3744 | 165 | 0.20533 | 0.003803 | 5.7e-05 | 100.8985 | 92.358 | 112.6677 | 21.2469 | 0.003316 | 0.002991 | 252.2158 | 0.27164 | 8.4015 | 0.003798 | 0.005721 | 0.005714 | 6.1e-05 | 0.001049 | 0.28711 | 1.268 | 2.548 | 99.38 | 20.95 | 37.534 | 101.622 | 113.3196 | 93.7623 | 0.009375 | 0.96221 | 0.022775 | 23.6697 | 0.79108 | 0.1616 | 0.020514 | 1.3965 | 1.3965 | 0.004756
2019-03-21 12:38 | 1553186273.3748 | 166 | 0.20535 | 0.00379 | 5.7e-05 | 100.8509 | 92.4081 | 112.8827 | 21.248 | 0.003344 | 0.003037 | 252.166 | 0.26965 | 8.4025 | 0.00379 | 0.005714 | 0.005566 | 6.1e-05 | 0.001047 | 0.30685 | 1.3069 | 2.548 | 99.38 | 20.95 | 37.533 | 102.2994 | 113.1249 | 94.8762 | 0.009375 | 0.95929 | 0.022762 | 23.7282 | 0.78201 | 0.16042 | 0.020441 | 1.3907 | 1.3907 | 0.004741
2019-03-21 12:38 | 1553186274.3725 | 167 | 0.20541 | 0.003778 | 5.6e-05 | 100.8637 | 92.481 | 112.8062 | 21.2479 | 0.003351 | 0.00298 | 252.1028 | 0.27081 | 8.4036 | 0.003817 | 0.005707 | 0.005649 | 6e-05 | 0.001046 | 0.26897 | 1.2882 | 2.548 | 99.38 | 20.95 | 37.531 | 102.4743 | 113.7836 | 93.5137 | 0.009374 | 0.96286 | 0.022797 | 23.6762 | 0.77609 | 0.1611 | 0.020096 | 1.3693 | 1.3693 | 0.004836
2019-03-21 12:38 | 1553186275.3726 | 168 | 0.20545 | 0.00377 | 5.6e-05 | 100.7999 | 92.3079 | 113.2622 | 21.2476 | 0.003283 | 0.002997 | 252.0157 | 0.27055 | 8.4037 | 0.00379 | 0.005561 | 0.005525 | 5.9e-05 | 0.001044 | 0.25337 | 1.2846 | 2.548 | 99.38 | 20.95 | 37.53 | 100.8221 | 113.126 | 93.9133 | 0.009374 | 0.96181 | 0.02283 | 23.7369 | 0.78719 | 0.16095 | 0.019891 | 1.3574 | 1.3574 | 0.004873
2019-03-21 12:38 | 1553186276.3735 | 169 | 0.20541 | 0.003757 | 5.4e-05 | 100.8926 | 92.3032 | 113.4296 | 21.2484 | 0.003292 | 0.003031 | 251.9556 | 0.27092 | 8.4026 | 0.003802 | 0.00557 | 0.005521 | 6e-05 | 0.001044 | 0.24458 | 1.2962 | 2.548 | 99.38 | 20.95 | 37.528 | 101.0449 | 113.4231 | 94.7286 | 0.009373 | 0.95968 | 0.022822 | 23.7808 | 0.79587 | 0.16117 | 0.020103 | 1.3714 | 1.3714 | 0.00483
2019-03-21 12:38 | 1553186277.373 | 170 | 0.20542 | 0.003743 | 5.3e-05 | 100.9164 | 92.169 | 113.3803 | 21.2477 | 0.003334 | 0.003022 | 251.8697 | 0.27079 | 8.4039 | 0.003807 | 0.005694 | 0.00565 | 6e-05 | 0.001039 | 0.30613 | 1.3096 | 2.5481 | 99.38 | 20.95 | 37.527 | 102.0541 | 113.5503 | 94.5104 | 0.009373 | 0.96025 | 0.022824 | 23.7686 | 0.81099 | 0.16109 | 0.020053 | 1.3682 | 1.3682 | 0.004839
2019-03-21 12:38 | 1553186278.3737 | 171 | 0.20544 | 0.003734 | 5.2e-05 | 100.8923 | 92.4071 | 113.773 | 21.2484 | 0.003356 | 0.003021 | 251.816 | 0.27109 | 8.4021 | 0.003771 | 0.005714 | 0.005666 | 5.7e-05 | 0.001041 | 0.27527 | 1.3194 | 2.5481 | 99.38 | 20.95 | 37.525 | 102.5818 | 112.6726 | 94.5033 | 0.009373 | 0.96027 | 0.022863 | 23.8095 | 0.78578 | 0.16127 | 0.019985 | 1.3659 | 1.3659 | 0.004853
2019-03-21 12:38 | 1553186279.3728 | 172 | 0.20548 | 0.003717 | 5.1e-05 | 100.8854 | 92.0345 | 113.9464 | 21.249 | 0.003309 | 0.003004 | 251.7168 | 0.2706 | 8.4006 | 0.003844 | 0.00526 | 0.005209 | 5.8e-05 | 0.00104 | 0.29286 | 1.3282 | 2.5481 | 99.38 | 20.95 | 37.522 | 101.4522 | 114.4472 | 94.0774 | 0.009372 | 0.96138 | 0.022894 | 23.8138 | 0.82131 | 0.16098 | 0.019742 | 1.3512 | 1.3512 | 0.004897
2019-03-21 12:38 | 1553186280.3742 | 173 | 0.20547 | 0.003713 | 5e-05 | 100.8649 | 92.0886 | 113.8288 | 21.2486 | 0.00332 | 0.002975 | 251.6462 | 0.26964 | 8.4017 | 0.003809 | 0.00481 | 0.004696 | 6e-05 | 0.001039 | 0.28584 | 1.3347 | 2.5481 | 99.38 | 20.95 | 37.521 | 101.7258 | 113.5909 | 93.3718 | 0.009372 | 0.96323 | 0.022904 | 23.7787 | 0.81422 | 0.16041 | 0.019826 | 1.3577 | 1.3577 | 0.004856
2019-03-21 12:38 | 1553186281.3729 | 174 | 0.20549 | 0.003702 | 5e-05 | 100.7701 | 92.3182 | 113.9626 | 21.2479 | 0.003319 | 0.003009 | 251.5667 | 0.27089 | 8.4013 | 0.003771 | 0.005083 | 0.004965 | 5.9e-05 | 0.001038 | 0.26465 | 1.3316 | 2.548 | 99.38 | 20.95 | 37.519 | 101.6874 | 112.6602 | 94.1991 | 0.009371 | 0.96106 | 0.022892 | 23.8195 | 0.78355 | 0.16115 | 0.019699 | 1.3483 | 1.3483 | 0.004913
2019-03-21 12:38 | 1553186282.397 | 175 | 0.20553 | 0.003687 | 4.8e-05 | 100.827 | 92.3107 | 113.9065 | 21.2479 | 0.003355 | 0.003021 | 251.5059 | 0.27072 | 8.4018 | 0.003717 | 0.005135 | 0.005102 | 5.7e-05 | 0.001038 | 0.29662 | 1.3634 | 2.548 | 99.38 | 20.95 | 37.516 | 102.5522 | 111.3417 | 94.4886 | 0.00937 | 0.96031 | 0.022877 | 23.823 | 0.78933 | 0.16105 | 0.019485 | 1.3329 | 1.3329 | 0.004966
2019-03-21 12:38 | 1553186283.3836 | 176 | 0.20554 | 0.003679 | 4.7e-05 | 100.8485 | 92.2531 | 113.8636 | 21.2486 | 0.003321 | 0.003013 | 251.4324 | 0.27099 | 8.4021 | 0.003764 | 0.005037 | 0.00501 | 6e-05 | 0.00104 | 0.31708 | 1.3134 | 2.548 | 99.38 | 20.95 | 37.514 | 101.7485 | 112.5029 | 94.31 | 0.00937 | 0.96077 | 0.022879 | 23.8127 | 0.79681 | 0.16121 | 0.019468 | 1.3318 | 1.3318 | 0.004975
2019-03-21 12:38 | 1553186284.3847 | 177 | 0.20557 | 0.003665 | 4.6e-05 | 100.8801 | 92.5692 | 113.6345 | 21.2493 | 0.003381 | 0.003026 | 251.329 | 0.27057 | 8.4023 | 0.003688 | 0.005852 | 0.005801 | 5.8e-05 | 0.001038 | 0.29894 | 1.2999 | 2.5479 | 99.38 | 20.95 | 37.511 | 103.1995 | 110.64 | 94.6061 | 0.009369 | 0.96 | 0.022846 | 23.7983 | 0.76902 | 0.16096 | 0.019274 | 1.3168 | 1.3168 | 0.005024
2019-03-21 12:38 | 1553186285.3841 | 178 | 0.20558 | 0.003654 | 4.5e-05 | 100.8342 | 92.4415 | 113.4286 | 21.2483 | 0.003293 | 0.002987 | 251.2673 | 0.26998 | 8.404 | 0.003674 | 0.005455 | 0.005474 | 5.9e-05 | 0.001039 | 0.27721 | 1.3347 | 2.5479 | 99.38 | 20.96 | 37.509 | 101.0675 | 110.3114 | 93.6612 | 0.009374 | 0.96247 | 0.022855 | 23.7462 | 0.7773 | 0.16061 | 0.019274 | 1.3173 | 1.3173 | 0.005011
2019-03-21 12:38 | 1553186286.3838 | 179 | 0.20565 | 0.003643 | 4.5e-05 | 100.8059 | 92.3478 | 113.6681 | 21.2491 | 0.003314 | 0.003062 | 251.1664 | 0.27083 | 8.4021 | 0.003714 | 0.005014 | 0.004917 | 6.1e-05 | 0.001042 | 0.26842 | 1.3078 | 2.5482 | 99.38 | 20.96 | 37.507 | 101.5779 | 111.2796 | 95.4817 | 0.009374 | 0.95772 | 0.022823 | 23.8302 | 0.78385 | 0.16112 | 0.018855 | 1.2868 | 1.2868 | 0.005146
2019-03-21 12:38 | 1553186287.3853 | 180 | 0.20562 | 0.003627 | 4.5e-05 | 100.7993 | 92.2947 | 113.9495 | 21.2493 | 0.003308 | 0.003017 | 251.1092 | 0.27074 | 8.401 | 0.003742 | 0.005357 | 0.005231 | 6.2e-05 | 0.001041 | 0.24082 | 1.3356 | 2.548 | 99.38 | 20.96 | 37.506 | 101.4165 | 111.9517 | 94.4041 | 0.009374 | 0.96053 | 0.022884 | 23.8247 | 0.78842 | 0.16106 | 0.019041 | 1.3031 | 1.3031 | 0.00508
2019-03-21 12:38 | 1553186288.3853 | 181 | 0.20565 | 0.003615 | 4.5e-05 | 100.7695 | 92.4479 | 113.7643 | 21.2495 | 0.003353 | 0.002979 | 251.0194 | 0.27183 | 8.4049 | 0.003713 | 0.005215 | 0.005092 | 5.6e-05 | 0.00104 | 0.28279 | 1.3386 | 2.5481 | 99.38 | 20.96 | 37.505 | 102.5238 | 111.254 | 93.4776 | 0.009373 | 0.96295 | 0.022895 | 23.7753 | 0.77093 | 0.16171 | 0.018888 | 1.2931 | 1.2931 | 0.00514
2019-03-21 12:38 | 1553186289.3849 | 182 | 0.20565 | 0.003599 | 4.4e-05 | 100.8026 | 92.1441 | 113.5246 | 21.2487 | 0.003354 | 0.002986 | 250.9555 | 0.27088 | 8.4032 | 0.003719 | 0.004818 | 0.004724 | 5.9e-05 | 0.001046 | 0.29789 | 1.2888 | 2.5479 | 99.38 | 20.96 | 37.503 | 102.5301 | 111.408 | 93.6599 | 0.009373 | 0.96247 | 0.022865 | 23.7562 | 0.80331 | 0.16115 | 0.018875 | 1.2906 | 1.2906 | 0.005132
2019-03-21 12:38 | 1553186290.383 | 183 | 0.20567 | 0.003584 | 4.5e-05 | 100.8434 | 91.9881 | 113.4423 | 21.2485 | 0.003315 | 0.002991 | 250.8991 | 0.27072 | 8.4024 | 0.003737 | 0.0052 | 0.005099 | 6e-05 | 0.001042 | 0.28567 | 1.322 | 2.548 | 99.38 | 20.96 | 37.502 | 101.5945 | 111.8481 | 93.781 | 0.009373 | 0.96216 | 0.022853 | 23.7515 | 0.82209 | 0.16105 | 0.018778 | 1.2833 | 1.2833 | 0.005158
2019-03-21 12:38 | 1553186291.383 | 184 | 0.20573 | 0.003575 | 4.5e-05 | 100.8157 | 91.9745 | 113.393 | 21.2481 | 0.003358 | 0.003013 | 250.8318 | 0.27137 | 8.3997 | 0.003787 | 0.005123 | 0.00507 | 5.7e-05 | 0.001042 | 0.2589 | 1.3206 | 2.5482 | 99.38 | 20.96 | 37.501 | 102.6316 | 113.0577 | 94.2963 | 0.009372 | 0.96081 | 0.022832 | 23.763 | 0.82095 | 0.16144 | 0.018414 | 1.2572 | 1.2572 | 0.005278
2019-03-21 12:38 | 1553186292.3845 | 185 | 0.20571 | 0.003559 | 4.5e-05 | 100.7937 | 92.1073 | 113.546 | 21.2491 | 0.003322 | 0.002985 | 250.7641 | 0.27063 | 8.4035 | 0.003905 | 0.00544 | 0.005333 | 5.7e-05 | 0.00104 | 0.26665 | 1.3423 | 2.548 | 99.38 | 20.96 | 37.499 | 101.7577 | 115.9242 | 93.6345 | 0.009372 | 0.96254 | 0.022868 | 23.7576 | 0.80609 | 0.161 | 0.018558 | 1.2691 | 1.2691 | 0.005214
2019-03-21 12:38 | 1553186293.3842 | 186 | 0.20573 | 0.003544 | 4.6e-05 | 100.7153 | 92.1567 | 113.6375 | 21.2493 | 0.003317 | 0.003022 | 250.6937 | 0.27105 | 8.4032 | 0.003883 | 0.005459 | 0.005419 | 6.2e-05 | 0.001042 | 0.26106 | 1.382 | 2.548 | 99.38 | 20.96 | 37.498 | 101.6566 | 115.4052 | 94.5245 | 0.009372 | 0.96021 | 0.022849 | 23.796 | 0.79434 | 0.16125 | 0.018472 | 1.2621 | 1.2621 | 0.005251
2019-03-21 12:38 | 1553186294.3833 | 187 | 0.20575 | 0.003529 | 4.6e-05 | 100.7975 | 92.5964 | 114.045 | 21.2486 | 0.003296 | 0.003044 | 250.6619 | 0.27213 | 8.3988 | 0.003861 | 0.00527 | 0.005183 | 5.9e-05 | 0.001043 | 0.27743 | 1.3369 | 2.548 | 99.38 | 20.96 | 37.498 | 101.1276 | 114.8635 | 95.0396 | 0.009372 | 0.95887 | 0.022874 | 23.8553 | 0.75906 | 0.16189 | 0.018406 | 1.2589 | 1.2589 | 0.005285
2019-03-21 12:38 | 1553186295.3837 | 188 | 0.20576 | 0.003517 | 4.7e-05 | 100.7716 | 93.0872 | 114.0539 | 21.2493 | 0.003309 | 0.003008 | 250.6127 | 0.26988 | 8.4035 | 0.003847 | 0.005401 | 0.00535 | 5.9e-05 | 0.001041 | 0.25011 | 1.3057 | 2.548 | 99.38 | 20.97 | 37.498 | 101.462 | 114.514 | 94.173 | 0.009377 | 0.96113 | 0.022902 | 23.8282 | 0.70948 | 0.16055 | 0.018325 | 1.2549 | 1.2549 | 0.005258
2019-03-21 12:38 | 1553186296.3833 | 189 | 0.20579 | 0.003506 | 4.7e-05 | 100.768 | 92.9067 | 113.6845 | 21.2492 | 0.003327 | 0.003033 | 250.5489 | 0.27027 | 8.4021 | 0.003823 | 0.005652 | 0.005585 | 6.3e-05 | 0.001038 | 0.26017 | 1.2805 | 2.548 | 99.38 | 20.97 | 37.497 | 101.8827 | 113.932 | 94.7945 | 0.009377 | 0.95951 | 0.022846 | 23.8097 | 0.72652 | 0.16079 | 0.018165 | 1.2409 | 1.2409 | 0.005325
2019-03-21 12:38 | 1553186297.3844 | 190 | 0.20581 | 0.003493 | 4.7e-05 | 100.7555 | 92.4071 | 113.3056 | 21.2489 | 0.003315 | 0.003032 | 250.4723 | 0.27009 | 8.4001 | 0.003878 | 0.005669 | 0.005668 | 5.9e-05 | 0.00104 | 0.27887 | 1.2838 | 2.548 | 99.38 | 20.97 | 37.495 | 101.5887 | 115.272 | 94.7654 | 0.009377 | 0.95958 | 0.022808 | 23.769 | 0.77365 | 0.16068 | 0.01808 | 1.233 | 1.233 | 0.005356
2019-03-21 12:38 | 1553186298.3847 | 191 | 0.2058 | 0.003482 | 4.8e-05 | 100.7659 | 92.6935 | 113.5181 | 21.2487 | 0.003336 | 0.003004 | 250.3942 | 0.27018 | 8.4027 | 0.003818 | 0.004858 | 0.00476 | 5.9e-05 | 0.001036 | 0.25962 | 1.3229 | 2.5481 | 99.38 | 20.97 | 37.494 | 102.103 | 113.7963 | 94.0904 | 0.009376 | 0.96135 | 0.022851 | 23.7694 | 0.74689 | 0.16073 | 0.018106 | 1.2371 | 1.2371 | 0.00534
2019-03-21 12:38 | 1553186299.3835 | 192 | 0.2058 | 0.003472 | 4.8e-05 | 100.7505 | 92.6499 | 113.7338 | 21.2487 | 0.003325 | 0.003048 | 250.3507 | 0.27077 | 8.4017 | 0.003798 | 0.004981 | 0.004906 | 6.4e-05 | 0.001038 | 0.27201 | 1.298 | 2.5481 | 99.38 | 20.97 | 37.492 | 101.847 | 113.3112 | 95.1455 | 0.009376 | 0.95859 | 0.02284 | 23.8262 | 0.74972 | 0.16108 | 0.018172 | 1.2409 | 1.2409 | 0.005335
2019-03-21 12:38 | 1553186300.3835 | 193 | 0.20587 | 0.003463 | 4.9e-05 | 100.7272 | 93.1342 | 114.2159 | 21.2491 | 0.003306 | 0.002984 | 250.2843 | 0.27011 | 8.4045 | 0.003844 | 0.005223 | 0.005214 | 6.5e-05 | 0.001039 | 0.28191 | 1.2707 | 2.548 | 99.38 | 20.97 | 37.489 | 101.3754 | 114.4433 | 93.6041 | 0.009375 | 0.96262 | 0.022936 | 23.8266 | 0.70102 | 0.16069 | 0.017738 | 1.2164 | 1.2164 | 0.005429
2019-03-21 12:38 | 1553186301.3846 | 194 | 0.20582 | 0.003446 | 4.9e-05 | 100.6709 | 93.2476 | 114.3816 | 21.2481 | 0.003336 | 0.003001 | 250.2358 | 0.27031 | 8.4006 | 0.003846 | 0.005379 | 0.00533 | 5.8e-05 | 0.001035 | 0.26913 | 1.2847 | 2.5478 | 99.38 | 20.97 | 37.487 | 102.0976 | 114.4796 | 94.0201 | 0.009375 | 0.96153 | 0.02294 | 23.8574 | 0.68514 | 0.16081 | 0.018069 | 1.2393 | 1.2393 | 0.005333
2019-03-21 12:38 | 1553186302.384 | 195 | 0.2059 | 0.003436 | 4.9e-05 | 100.7513 | 93.5916 | 114.4598 | 21.2492 | 0.003295 | 0.003009 | 250.1573 | 0.27116 | 8.4029 | 0.003782 | 0.005459 | 0.00542 | 5.4e-05 | 0.001033 | 0.28064 | 1.3178 | 2.5479 | 99.38 | 20.98 | 37.483 | 101.1234 | 112.9372 | 94.205 | 0.009379 | 0.96105 | 0.022942 | 23.8716 | 0.65934 | 0.16131 | 0.017605 | 1.2076 | 1.2076 | 0.00549
2019-03-21 12:38 | 1553186303.3847 | 196 | 0.20597 | 0.003425 | 4.9e-05 | 100.7195 | 93.5494 | 114.2551 | 21.2486 | 0.00333 | 0.003002 | 250.0876 | 0.27102 | 8.4003 | 0.003815 | 0.005306 | 0.005211 | 5.6e-05 | 0.001033 | 0.2959 | 1.2775 | 2.548 | 99.38 | 20.98 | 37.481 | 101.9543 | 113.7307 | 94.0315 | 0.009379 | 0.9615 | 0.022927 | 23.8446 | 0.66055 | 0.16123 | 0.017208 | 1.1796 | 1.1796 | 0.005618
2019-03-21 12:38 | 1553186304.3842 | 197 | 0.20597 | 0.003408 | 4.9e-05 | 100.7865 | 93.5273 | 114.0147 | 21.2491 | 0.003314 | 0.002968 | 250.0257 | 0.27021 | 8.4006 | 0.00388 | 0.005473 | 0.005422 | 5.8e-05 | 0.001034 | 0.31116 | 1.318 | 2.5481 | 99.38 | 20.98 | 37.479 | 101.5779 | 115.3126 | 93.2206 | 0.009378 | 0.96363 | 0.022928 | 23.7932 | 0.66862 | 0.16075 | 0.017211 | 1.1798 | 1.1798 | 0.0056
2019-03-21 12:38 | 1553186305.3832 | 198 | 0.20604 | 0.003397 | 4.9e-05 | 100.7162 | 93.9319 | 114.0183 | 21.2484 | 0.00335 | 0.002956 | 249.9127 | 0.2702 | 8.4055 | 0.003834 | 0.005512 | 0.005493 | 5.7e-05 | 0.001033 | 0.26936 | 1.2997 | 2.5478 | 99.38 | 20.98 | 37.475 | 102.4419 | 114.2007 | 92.9336 | 0.009377 | 0.96439 | 0.022937 | 23.7842 | 0.62376 | 0.16074 | 0.016802 | 1.1522 | 1.1522 | 0.005734
2019-03-21 12:38 | 1553186306.3828 | 199 | 0.2061 | 0.003376 | 4.9e-05 | 100.6757 | 94.0875 | 114.2537 | 21.2492 | 0.003303 | 0.002979 | 249.8831 | 0.27067 | 8.4016 | 0.003826 | 0.005698 | 0.005615 | 5.8e-05 | 0.00103 | 0.22357 | 1.2885 | 2.548 | 99.38 | 20.98 | 37.472 | 101.3048 | 114.0158 | 93.4812 | 0.009377 | 0.96294 | 0.022944 | 23.8265 | 0.60536 | 0.16102 | 0.016464 | 1.1293 | 1.1293 | 0.00586
2019-03-21 12:38 | 1553186307.3829 | 200 | 0.20613 | 0.003354 | 4.8e-05 | 100.6733 | 94.2336 | 114.4849 | 21.2486 | 0.003275 | 0.002988 | 249.8309 | 0.2701 | 8.4028 | 0.003822 | 0.005515 | 0.005484 | 5.5e-05 | 0.001032 | 0.26233 | 1.2794 | 2.548 | 99.38 | 20.98 | 37.468 | 100.6387 | 113.9163 | 93.7076 | 0.009376 | 0.96235 | 0.02296 | 23.858 | 0.59126 | 0.16068 | 0.016302 | 1.119 | 1.119 | 0.005902
2019-03-21 12:38 | 1553186308.3843 | 201 | 0.20615 | 0.003338 | 4.8e-05 | 100.6614 | 94.3181 | 114.5405 | 21.249 | 0.003281 | 0.002962 | 249.7836 | 0.27025 | 8.4002 | 0.003788 | 0.005045 | 0.004938 | 5.6e-05 | 0.001027 | 0.26145 | 1.2976 | 2.5479 | 99.38 | 20.98 | 37.467 | 100.7677 | 113.0823 | 93.0595 | 0.009375 | 0.96405 | 0.022986 | 23.8427 | 0.58219 | 0.16077 | 0.016195 | 1.113 | 1.113 | 0.005937
2019-03-21 12:38 | 1553186309.3844 | 202 | 0.2062 | 0.003314 | 4.8e-05 | 100.6274 | 94.5255 | 114.5912 | 21.2486 | 0.003279 | 0.003014 | 249.696 | 0.27116 | 8.4003 | 0.003708 | 0.005407 | 0.005246 | 5.4e-05 | 0.001028 | 0.24585 | 1.3071 | 2.5481 | 99.38 | 20.98 | 37.465 | 100.7147 | 111.1357 | 94.333 | 0.009375 | 0.96071 | 0.022951 | 23.8894 | 0.55952 | 0.16131 | 0.015927 | 1.0929 | 1.0929 | 0.006067
2019-03-21 12:38 | 1553186310.3831 | 203 | 0.20625 | 0.003288 | 4.8e-05 | 100.6325 | 94.4438 | 114.7375 | 21.2497 | 0.003313 | 0.003019 | 249.6074 | 0.27155 | 8.403 | 0.003765 | 0.005071 | 0.005103 | 5.8e-05 | 0.001025 | 0.24441 | 1.2927 | 2.5478 | 99.38 | 20.98 | 37.462 | 101.5384 | 112.5142 | 94.4568 | 0.009374 | 0.96039 | 0.022962 | 23.9087 | 0.56771 | 0.16155 | 0.015662 | 1.0752 | 1.0752 | 0.006175
2019-03-21 12:38 | 1553186311.3841 | 204 | 0.20625 | 0.003262 | 4.7e-05 | 100.6107 | 94.9495 | 114.8762 | 21.2493 | 0.003266 | 0.002956 | 249.5235 | 0.27094 | 8.4024 | 0.003755 | 0.004745 | 0.004666 | 5.6e-05 | 0.001028 | 0.29247 | 1.2805 | 2.5479 | 99.38 | 20.98 | 37.462 | 100.4069 | 112.2668 | 92.9354 | 0.009374 | 0.96438 | 0.023023 | 23.8736 | 0.51801 | 0.16118 | 0.015701 | 1.0808 | 1.0808 | 0.006129
2019-03-21 12:39 | 1553186312.3845 | 205 | 0.20628 | 0.003238 | 4.7e-05 | 100.5529 | 94.7128 | 114.9816 | 21.2502 | 0.00331 | 0.002932 | 249.446 | 0.27061 | 8.4005 | 0.003722 | 0.005505 | 0.00541 | 5.6e-05 | 0.001028 | 0.26288 | 1.2683 | 2.548 | 99.38 | 20.98 | 37.461 | 101.4807 | 111.4836 | 92.3463 | 0.009374 | 0.96593 | 0.023052 | 23.8653 | 0.53519 | 0.16099 | 0.015567 | 1.0729 | 1.0729 | 0.006167
2019-03-21 12:39 | 1553186313.3832 | 206 | 0.20627 | 0.003203 | 4.7e-05 | 100.5911 | 95.2574 | 115.2445 | 21.2484 | 0.003277 | 0.002932 | 249.4085 | 0.27013 | 8.4019 | 0.003719 | 0.005581 | 0.005503 | 5.4e-05 | 0.001028 | 0.28279 | 1.313 | 2.548 | 99.38 | 20.98 | 37.461 | 100.6801 | 111.4028 | 92.3446 | 0.009374 | 0.96594 | 0.023079 | 23.8925 | 0.4873 | 0.1607 | 0.015656 | 1.0803 | 1.0803 | 0.006114
2019-03-21 12:39 | 1553186314.3846 | 207 | 0.20628 | 0.003179 | 4.7e-05 | 100.5291 | 95.2901 | 115.2426 | 21.2508 | 0.003276 | 0.002934 | 249.319 | 0.27139 | 8.4039 | 0.003717 | 0.00547 | 0.005509 | 5.5e-05 | 0.001027 | 0.28727 | 1.3124 | 2.548 | 99.38 | 20.99 | 37.461 | 100.6556 | 111.3531 | 92.3947 | 0.00938 | 0.96581 | 0.023077 | 23.894 | 0.47872 | 0.16145 | 0.015624 | 1.078 | 1.078 | 0.006155
2019-03-21 12:39 | 1553186315.3835 | 208 | 0.20639 | 0.003153 | 4.6e-05 | 100.5321 | 95.7162 | 115.5008 | 21.2497 | 0.003243 | 0.002933 | 249.2908 | 0.27018 | 8.4024 | 0.003792 | 0.005083 | 0.004997 | 5.4e-05 | 0.001027 | 0.26001 | 1.3269 | 2.5479 | 99.38 | 20.99 | 37.461 | 99.855 | 113.1787 | 92.3789 | 0.00938 | 0.96585 | 0.023103 | 23.9202 | 0.43908 | 0.16073 | 0.015022 | 1.0376 | 1.0376 | 0.006366
2019-03-21 12:39 | 1553186316.3834 | 209 | 0.20642 | 0.003125 | 4.7e-05 | 100.5371 | 95.5268 | 115.398 | 21.2496 | 0.003261 | 0.002992 | 249.2132 | 0.27146 | 8.401 | 0.003686 | 0.0053 | 0.005161 | 5.4e-05 | 0.001029 | 0.30574 | 1.301 | 2.5477 | 99.38 | 20.99 | 37.461 | 100.2802 | 110.5977 | 93.7844 | 0.00938 | 0.96215 | 0.023049 | 23.9555 | 0.45724 | 0.16149 | 0.01484 | 1.0226 | 1.0226 | 0.006491
2019-03-21 12:39 | 1553186317.3846 | 210 | 0.20643 | 0.003101 | 4.7e-05 | 100.563 | 95.7574 | 115.2388 | 21.2495 | 0.00331 | 0.003009 | 249.1641 | 0.27038 | 8.4035 | 0.003701 | 0.005097 | 0.005079 | 5.8e-05 | 0.001028 | 0.25227 | 1.2779 | 2.5479 | 99.38 | 20.99 | 37.461 | 101.4675 | 110.9675 | 94.2094 | 0.00938 | 0.96103 | 0.023019 | 23.9528 | 0.43799 | 0.16085 | 0.014837 | 1.0211 | 1.0211 | 0.006475
2019-03-21 12:39 | 1553186318.3845 | 211 | 0.20645 | 0.003071 | 4.8e-05 | 100.5377 | 95.4819 | 115.183 | 21.2497 | 0.003302 | 0.002904 | 249.1206 | 0.27146 | 8.4012 | 0.003662 | 0.004823 | 0.004756 | 5.9e-05 | 0.00103 | 0.28296 | 1.3219 | 2.5479 | 99.38 | 20.99 | 37.46 | 101.2763 | 110.0124 | 91.6742 | 0.009379 | 0.96771 | 0.023094 | 23.8642 | 0.46151 | 0.16149 | 0.014747 | 1.0181 | 1.0181 | 0.006519
2019-03-21 12:39 | 1553186319.3849 | 212 | 0.20646 | 0.003048 | 4.9e-05 | 100.6125 | 95.7986 | 115.1557 | 21.2504 | 0.003271 | 0.002929 | 249.0752 | 0.2702 | 8.4033 | 0.003711 | 0.005411 | 0.005363 | 5.5e-05 | 0.001032 | 0.26792 | 1.3043 | 2.548 | 99.38 | 20.99 | 37.46 | 100.5241 | 111.2092 | 92.2699 | 0.009379 | 0.96614 | 0.023072 | 23.8809 | 0.43854 | 0.16074 | 0.014698 | 1.0137 | 1.0137 | 0.006517
2019-03-21 12:39 | 1553186320.3836 | 213 | 0.20647 | 0.00302 | 5e-05 | 100.5321 | 95.8708 | 114.8494 | 21.2495 | 0.003268 | 0.00293 | 249.0418 | 0.27064 | 8.4013 | 0.003687 | 0.005188 | 0.005182 | 5.7e-05 | 0.001028 | 0.25625 | 1.3064 | 2.5479 | 99.38 | 20.99 | 37.458 | 100.462 | 110.6117 | 92.3011 | 0.009379 | 0.96605 | 0.02304 | 23.8501 | 0.42464 | 0.16101 | 0.014707 | 1.0129 | 1.0129 | 0.006533
2019-03-21 12:39 | 1553186321.3828 | 214 | 0.20651 | 0.002997 | 5.1e-05 | 100.5356 | 95.9737 | 114.7976 | 21.2502 | 0.003235 | 0.002921 | 248.9815 | 0.27059 | 8.4029 | 0.003717 | 0.00522 | 0.005132 | 5.6e-05 | 0.001031 | 0.27057 | 1.3169 | 2.5479 | 99.38 | 20.99 | 37.456 | 99.6524 | 111.3557 | 92.0854 | 0.009378 | 0.96662 | 0.023042 | 23.8377 | 0.41537 | 0.16097 | 0.014461 | 0.99596 | 0.99596 | 0.006643
2019-03-21 12:39 | 1553186322.3841 | 215 | 0.20647 | 0.002973 | 5.3e-05 | 100.5219 | 96.097 | 114.8353 | 21.2503 | 0.003254 | 0.00294 | 248.9459 | 0.27096 | 8.405 | 0.003717 | 0.005429 | 0.005385 | 5.6e-05 | 0.00103 | 0.21909 | 1.3134 | 2.5479 | 99.38 | 20.99 | 37.456 | 100.1198 | 111.3511 | 92.5359 | 0.009378 | 0.96543 | 0.023032 | 23.8563 | 0.40266 | 0.1612 | 0.014716 | 1.0129 | 1.0129 | 0.006541
2019-03-21 12:39 | 1553186323.3828 | 216 | 0.20651 | 0.002954 | 5.5e-05 | 100.5347 | 95.722 | 114.9769 | 21.2498 | 0.003201 | 0.002939 | 248.8764 | 0.2706 | 8.4059 | 0.003684 | 0.005486 | 0.005439 | 5.5e-05 | 0.001028 | 0.29192 | 1.3426 | 2.5479 | 99.38 | 21 | 37.453 | 98.8378 | 110.5488 | 92.5072 | 0.009383 | 0.96551 | 0.023047 | 23.8701 | 0.43878 | 0.16098 | 0.014513 | 0.99948 | 0.99948 | 0.00662
2019-03-21 12:39 | 1553186324.3828 | 217 | 0.20647 | 0.002935 | 5.7e-05 | 100.471 | 95.5725 | 115.1287 | 21.2502 | 0.003226 | 0.002924 | 248.847 | 0.27113 | 8.4011 | 0.003648 | 0.005525 | 0.005508 | 5.5e-05 | 0.001025 | 0.2855 | 1.3214 | 2.5479 | 99.38 | 21 | 37.452 | 99.4507 | 109.6786 | 92.1588 | 0.009383 | 0.96643 | 0.023073 | 23.8744 | 0.44708 | 0.16129 | 0.014752 | 1.0171 | 1.0171 | 0.006518
2019-03-21 12:39 | 1553186325.3838 | 218 | 0.2065 | 0.002916 | 5.7e-05 | 100.4591 | 95.4802 | 115.3299 | 21.2503 | 0.00323 | 0.002934 | 248.7713 | 0.27103 | 8.4041 | 0.003571 | 0.005074 | 0.005014 | 5.7e-05 | 0.001017 | 0.27113 | 1.2637 | 2.548 | 99.38 | 21 | 37.449 | 99.5448 | 107.7942 | 92.4029 | 0.009382 | 0.96579 | 0.023085 | 23.9033 | 0.45467 | 0.16123 | 0.014643 | 1.0101 | 1.0101 | 0.00656
2019-03-21 12:39 | 1553186326.3839 | 219 | 0.20649 | 0.0029 | 5.7e-05 | 100.5088 | 95.3518 | 115.6287 | 21.2498 | 0.003235 | 0.00295 | 248.7264 | 0.27106 | 8.4025 | 0.003635 | 0.004452 | 0.004435 | 5.9e-05 | 0.001018 | 0.26681 | 1.3023 | 2.5481 | 99.38 | 21 | 37.446 | 99.6659 | 109.3526 | 92.7738 | 0.009382 | 0.96481 | 0.023104 | 23.9464 | 0.47113 | 0.16125 | 0.014699 | 1.0147 | 1.0147 | 0.006531
2019-03-21 12:39 | 1553186327.3836 | 220 | 0.2065 | 0.002887 | 5.6e-05 | 100.493 | 95.2244 | 115.8425 | 21.2499 | 0.003278 | 0.002856 | 248.6489 | 0.26953 | 8.4017 | 0.003629 | 0.005192 | 0.005065 | 6.4e-05 | 0.001017 | 0.26969 | 1.3116 | 2.548 | 99.38 | 21 | 37.444 | 100.697 | 109.2147 | 90.5166 | 0.009381 | 0.97079 | 0.023197 | 23.8945 | 0.48168 | 0.16034 | 0.01465 | 1.0154 | 1.0154 | 0.00649
2019-03-21 12:39 | 1553186328.384 | 221 | 0.20654 | 0.002879 | 5.7e-05 | 100.4364 | 95.7891 | 115.9385 | 21.2504 | 0.003291 | 0.002919 | 248.5514 | 0.26984 | 8.4035 | 0.003665 | 0.005016 | 0.004969 | 6.1e-05 | 0.001019 | 0.24945 | 1.2818 | 2.5479 | 99.38 | 21 | 37.438 | 101.0171 | 110.0833 | 92.0434 | 0.00938 | 0.96674 | 0.023158 | 23.9545 | 0.42376 | 0.16053 | 0.014439 | 0.99903 | 0.99903 | 0.006604
2019-03-21 12:39 | 1553186329.3831 | 222 | 0.20651 | 0.002861 | 5.8e-05 | 100.4302 | 96.0275 | 116.0557 | 21.2503 | 0.003204 | 0.002879 | 248.5067 | 0.26988 | 8.4031 | 0.003719 | 0.005113 | 0.005044 | 5.7e-05 | 0.001021 | 0.23794 | 1.2923 | 2.5479 | 99.38 | 21 | 37.436 | 98.9088 | 111.4063 | 91.0682 | 0.009379 | 0.96932 | 0.0232 | 23.9346 | 0.40097 | 0.16055 | 0.014654 | 1.0158 | 1.0158 | 0.006496
2019-03-21 12:39 | 1553186330.3841 | 223 | 0.2065 | 0.002858 | 6e-05 | 100.4853 | 96.1736 | 115.9914 | 21.2499 | 0.003271 | 0.002913 | 248.4217 | 0.27011 | 8.4037 | 0.003709 | 0.006072 | 0.005946 | 5.7e-05 | 0.001018 | 0.2505 | 1.3107 | 2.5479 | 99.38 | 21 | 37.435 | 100.533 | 111.1525 | 91.8872 | 0.009379 | 0.96715 | 0.023168 | 23.9548 | 0.39227 | 0.16069 | 0.014666 | 1.0151 | 1.0151 | 0.006506
2019-03-21 12:39 | 1553186331.3842 | 224 | 0.20652 | 0.002849 | 6.3e-05 | 100.4352 | 96.3534 | 116.0282 | 21.2499 | 0.003237 | 0.002958 | 248.3863 | 0.26982 | 8.4039 | 0.003738 | 0.00475 | 0.004742 | 5.9e-05 | 0.00102 | 0.23794 | 1.3081 | 2.5478 | 99.38 | 21 | 37.431 | 99.7188 | 111.8576 | 92.9734 | 0.009378 | 0.96428 | 0.023137 | 23.9942 | 0.37111 | 0.16052 | 0.014595 | 1.0086 | 1.0086 | 0.006541
2019-03-21 12:39 | 1553186332.3846 | 225 | 0.20647 | 0.00283 | 6.7e-05 | 100.4251 | 95.5977 | 115.9331 | 21.2501 | 0.00326 | 0.00289 | 248.335 | 0.27017 | 8.4014 | 0.003694 | 0.004782 | 0.004673 | 5.5e-05 | 0.001021 | 0.25442 | 1.3158 | 2.5478 | 99.38 | 21 | 37.43 | 100.2595 | 110.7842 | 91.3401 | 0.009378 | 0.9686 | 0.023179 | 23.9308 | 0.44064 | 0.16072 | 0.014921 | 1.0329 | 1.0329 | 0.006396
2019-03-21 12:39 | 1553186333.3843 | 226 | 0.2064 | 0.002822 | 7e-05 | 100.322 | 95.6815 | 115.9 | 21.2508 | 0.003219 | 0.002898 | 248.2736 | 0.27046 | 8.4013 | 0.003631 | 0.005046 | 0.004959 | 5.7e-05 | 0.001018 | 0.26952 | 1.3437 | 2.548 | 99.38 | 21.01 | 37.428 | 99.2731 | 109.2715 | 91.5345 | 0.009383 | 0.96808 | 0.02317 | 23.9338 | 0.42362 | 0.1609 | 0.015343 | 1.0615 | 1.0615 | 0.00623
2019-03-21 12:39 | 1553186334.3834 | 227 | 0.20641 | 0.002813 | 7.4e-05 | 100.2812 | 95.943 | 116.0892 | 21.2503 | 0.00326 | 0.002907 | 248.2073 | 0.27039 | 8.4042 | 0.003616 | 0.005135 | 0.005116 | 5.5e-05 | 0.001018 | 0.23739 | 1.3421 | 2.5481 | 99.38 | 21 | 37.426 | 100.2611 | 108.899 | 91.7507 | 0.009377 | 0.96751 | 0.023182 | 23.9604 | 0.39556 | 0.16085 | 0.015264 | 1.0564 | 1.0564 | 0.006258
2019-03-21 12:39 | 1553186335.3825 | 228 | 0.20644 | 0.002805 | 7.7e-05 | 100.3607 | 96.1276 | 115.9437 | 21.2503 | 0.00325 | 0.002899 | 248.1558 | 0.26992 | 8.4031 | 0.003676 | 0.005243 | 0.0052 | 5.5e-05 | 0.001015 | 0.24834 | 1.3391 | 2.548 | 99.38 | 21.01 | 37.424 | 100.0195 | 110.364 | 91.5598 | 0.009382 | 0.96802 | 0.023174 | 23.9391 | 0.38546 | 0.16057 | 0.015091 | 1.0438 | 1.0438 | 0.006323
2019-03-21 12:39 | 1553186336.3824 | 229 | 0.20647 | 0.002796 | 8.1e-05 | 100.3098 | 96.0527 | 115.9547 | 21.2513 | 0.003235 | 0.002934 | 248.1091 | 0.27057 | 8.4036 | 0.003694 | 0.005233 | 0.005177 | 5.6e-05 | 0.001014 | 0.27743 | 1.3262 | 2.5479 | 99.38 | 21 | 37.421 | 99.6754 | 110.7854 | 92.3961 | 0.009375 | 0.9658 | 0.023148 | 23.9677 | 0.38789 | 0.16096 | 0.01495 | 1.0327 | 1.0327 | 0.006406
2019-03-21 12:39 | 1553186337.3839 | 230 | 0.20647 | 0.00279 | 8.4e-05 | 100.2952 | 96.1045 | 116.0597 | 21.25 | 0.00321 | 0.002936 | 248.0495 | 0.27057 | 8.4036 | 0.003651 | 0.005165 | 0.005095 | 5.7e-05 | 0.001016 | 0.2568 | 1.2842 | 2.548 | 99.38 | 21 | 37.42 | 99.0598 | 109.7484 | 92.4486 | 0.009375 | 0.96566 | 0.023157 | 23.9803 | 0.38177 | 0.16096 | 0.014953 | 1.0331 | 1.0331 | 0.006404
2019-03-21 12:39 | 1553186338.3826 | 231 | 0.20646 | 0.002793 | 8.7e-05 | 100.2699 | 95.845 | 115.9782 | 21.25 | 0.003279 | 0.002919 | 248.0155 | 0.27023 | 8.4031 | 0.00368 | 0.005137 | 0.00506 | 5.6e-05 | 0.001017 | 0.28528 | 1.2721 | 2.548 | 99.38 | 21 | 37.419 | 100.738 | 110.4393 | 92.0437 | 0.009375 | 0.96674 | 0.023162 | 23.9586 | 0.4037 | 0.16076 | 0.015005 | 1.0368 | 1.0368 | 0.006373
2019-03-21 12:39 | 1553186339.3836 | 232 | 0.20643 | 0.002789 | 9e-05 | 100.2731 | 96.0054 | 115.8855 | 21.2509 | 0.003197 | 0.002979 | 247.9608 | 0.27017 | 8.4023 | 0.003642 | 0.00594 | 0.005913 | 5.8e-05 | 0.001015 | 0.26128 | 1.3388 | 2.5479 | 99.38 | 21.01 | 37.419 | 98.7456 | 109.5175 | 93.4735 | 0.009381 | 0.96296 | 0.023107 | 23.9958 | 0.38903 | 0.16072 | 0.015166 | 1.0452 | 1.0452 | 0.00632
2019-03-21 12:39 | 1553186340.3826 | 233 | 0.20644 | 0.002783 | 9.3e-05 | 100.2142 | 96.3259 | 115.9822 | 21.2512 | 0.003209 | 0.002934 | 247.9271 | 0.2707 | 8.4016 | 0.003643 | 0.005446 | 0.005357 | 6.3e-05 | 0.001016 | 0.28385 | 1.2874 | 2.548 | 99.38 | 21.01 | 37.418 | 99.0432 | 109.5574 | 92.3929 | 0.00938 | 0.96581 | 0.023151 | 23.9704 | 0.35396 | 0.16104 | 0.015119 | 1.0438 | 1.0438 | 0.006341
2019-03-21 12:39 | 1553186341.3826 | 234 | 0.2064 | 0.00278 | 9.7e-05 | 100.1468 | 96.4774 | 115.9329 | 21.2502 | 0.003256 | 0.00293 | 247.8577 | 0.27046 | 8.4029 | 0.00363 | 0.005069 | 0.005045 | 6e-05 | 0.00102 | 0.27903 | 1.2798 | 2.5479 | 99.38 | 21.01 | 37.416 | 100.1693 | 109.2291 | 92.2979 | 0.00938 | 0.96606 | 0.023149 | 23.9622 | 0.33388 | 0.16089 | 0.015362 | 1.0604 | 1.0604 | 0.006236
2019-03-21 12:39 | 1553186342.3827 | 235 | 0.20643 | 0.00278 | 0.0001 | 100.2973 | 96.5411 | 115.8883 | 21.2506 | 0.003154 | 0.002954 | 247.7821 | 0.27098 | 8.402 | 0.003619 | 0.004513 | 0.004396 | 5.9e-05 | 0.001016 | 0.23435 | 1.2801 | 2.5479 | 99.38 | 21.01 | 37.415 | 97.7174 | 108.9678 | 92.8771 | 0.00938 | 0.96453 | 0.023126 | 23.9766 | 0.34141 | 0.1612 | 0.01516 | 1.0452 | 1.0452 | 0.006339
2019-03-21 12:39 | 1553186343.3828 | 236 | 0.20638 | 0.002775 | 0.000102 | 100.1403 | 96.3851 | 115.9688 | 21.251 | 0.003187 | 0.002905 | 247.7293 | 0.27069 | 8.402 | 0.003568 | 0.005272 | 0.005256 | 6.1e-05 | 0.001017 | 0.27671 | 1.3095 | 2.548 | 99.38 | 21 | 37.415 | 98.5064 | 107.7202 | 91.7068 | 0.009374 | 0.96763 | 0.023171 | 23.9466 | 0.34186 | 0.16103 | 0.015443 | 1.0667 | 1.0667 | 0.006205
2019-03-21 12:39 | 1553186344.3829 | 237 | 0.20641 | 0.002776 | 0.000104 | 100.201 | 96.6773 | 116.0268 | 21.2498 | 0.00326 | 0.002943 | 247.6915 | 0.27109 | 8.4018 | 0.003657 | 0.005002 | 0.004936 | 6.4e-05 | 0.001017 | 0.24707 | 1.3219 | 2.548 | 99.38 | 21.01 | 37.413 | 100.2661 | 109.9028 | 92.6198 | 0.009379 | 0.96521 | 0.023148 | 23.9825 | 0.32021 | 0.16127 | 0.015312 | 1.0564 | 1.0564 | 0.006274
2019-03-21 12:39 | 1553186345.3839 | 238 | 0.20637 | 0.002773 | 0.000106 | 100.1993 | 96.522 | 115.9198 | 21.25 | 0.003194 | 0.002909 | 247.646 | 0.27018 | 8.4025 | 0.003629 | 0.00486 | 0.004749 | 6.2e-05 | 0.001017 | 0.29048 | 1.334 | 2.5481 | 99.38 | 21.01 | 37.412 | 98.6765 | 109.2109 | 91.7919 | 0.009379 | 0.9674 | 0.023164 | 23.9443 | 0.33443 | 0.16073 | 0.015504 | 1.0703 | 1.0703 | 0.006172
2019-03-21 12:39 | 1553186346.383 | 239 | 0.2064 | 0.002776 | 0.000108 | 100.1105 | 96.4205 | 115.8029 | 21.25 | 0.00323 | 0.002923 | 247.5657 | 0.27059 | 8.4014 | 0.003572 | 0.005516 | 0.005461 | 6.3e-05 | 0.001012 | 0.27759 | 1.329 | 2.548 | 99.38 | 21.01 | 37.413 | 99.543 | 107.8371 | 92.1374 | 0.009379 | 0.96649 | 0.023141 | 23.9435 | 0.33591 | 0.16097 | 0.015326 | 1.0568 | 1.0568 | 0.00626
2019-03-21 12:39 | 1553186347.383 | 240 | 0.2064 | 0.002779 | 0.000111 | 100.0878 | 96.2448 | 115.7116 | 21.2511 | 0.003212 | 0.002963 | 247.51 | 0.27001 | 8.4017 | 0.003584 | 0.004751 | 0.004711 | 6.4e-05 | 0.001009 | 0.25841 | 1.2952 | 2.5478 | 99.38 | 21.01 | 37.412 | 99.1088 | 108.1094 | 93.0982 | 0.009379 | 0.96395 | 0.023102 | 23.9656 | 0.35021 | 0.16063 | 0.015324 | 1.0547 | 1.0547 | 0.00626
2019-03-21 12:39 | 1553186348.3832 | 241 | 0.2064 | 0.002783 | 0.000113 | 100.1003 | 96.1978 | 115.7972 | 21.2508 | 0.003189 | 0.002966 | 247.4601 | 0.27083 | 8.4016 | 0.003626 | 0.005173 | 0.005126 | 6.2e-05 | 0.001011 | 0.286 | 1.319 | 2.548 | 99.38 | 21.01 | 37.41 | 98.5565 | 109.1359 | 93.177 | 0.009378 | 0.96374 | 0.023108 | 23.977 | 0.35569 | 0.16112 | 0.015323 | 1.0548 | 1.0548 | 0.006278
2019-03-21 12:39 | 1553186349.3842 | 242 | 0.20638 | 0.002785 | 0.000115 | 100.098 | 96.0694 | 115.6031 | 21.2512 | 0.003159 | 0.002927 | 247.3912 | 0.27039 | 8.404 | 0.003666 | 0.005111 | 0.00508 | 6.3e-05 | 0.001012 | 0.2672 | 1.3232 | 2.5481 | 99.38 | 21.01 | 37.408 | 97.8247 | 110.1062 | 92.2364 | 0.009378 | 0.96623 | 0.023118 | 23.9261 | 0.36743 | 0.16085 | 0.015422 | 1.062 | 1.062 | 0.006225
2019-03-21 12:39 | 1553186350.3834 | 243 | 0.20639 | 0.002786 | 0.000117 | 100.137 | 96.3395 | 115.5463 | 21.2503 | 0.00321 | 0.002923 | 247.3333 | 0.27124 | 8.4024 | 0.003634 | 0.005994 | 0.005882 | 6.3e-05 | 0.001011 | 0.28136 | 1.3253 | 2.548 | 99.38 | 21.01 | 37.407 | 99.0507 | 109.3268 | 92.129 | 0.009378 | 0.96651 | 0.023116 | 23.9167 | 0.3458 | 0.16136 | 0.015361 | 1.0575 | 1.0575 | 0.006271
2019-03-21 12:39 | 1553186351.3832 | 244 | 0.20628 | 0.002782 | 0.000119 | 100.109 | 96.4005 | 115.6465 | 21.2503 | 0.003251 | 0.002888 | 247.2588 | 0.26989 | 8.4023 | 0.003649 | 0.005237 | 0.005225 | 6.3e-05 | 0.001008 | 0.2625 | 1.3273 | 2.548 | 99.38 | 21.01 | 37.404 | 100.0596 | 109.7006 | 91.28 | 0.009377 | 0.96876 | 0.023153 | 23.8993 | 0.33764 | 0.16056 | 0.016058 | 1.1075 | 1.1075 | 0.005959
2019-03-21 12:39 | 1553186352.3833 | 245 | 0.20632 | 0.002784 | 0.000121 | 100.0559 | 96.4754 | 115.3834 | 21.2498 | 0.003222 | 0.002963 | 247.2282 | 0.2699 | 8.4022 | 0.003628 | 0.005308 | 0.005154 | 6.2e-05 | 0.001007 | 0.23866 | 1.2992 | 2.548 | 99.37 | 21.01 | 37.403 | 99.3413 | 109.1838 | 93.0861 | 0.009378 | 0.96398 | 0.023069 | 23.9311 | 0.32595 | 0.16056 | 0.015791 | 1.0849 | 1.0849 | 0.006083
2019-03-21 12:39 | 1553186353.3834 | 246 | 0.20627 | 0.002785 | 0.000123 | 100.0551 | 96.6654 | 115.0769 | 21.2504 | 0.003227 | 0.002913 | 247.1512 | 0.27053 | 8.4003 | 0.003663 | 0.004824 | 0.004858 | 6.3e-05 | 0.001009 | 0.25608 | 1.2901 | 2.548 | 99.37 | 21.01 | 37.403 | 99.4628 | 110.0439 | 91.8877 | 0.009378 | 0.96715 | 0.023076 | 23.8602 | 0.30827 | 0.16094 | 0.016074 | 1.1047 | 1.1047 | 0.005988
2019-03-21 12:39 | 1553186354.3835 | 247 | 0.20625 | 0.002786 | 0.000124 | 100.0181 | 96.6593 | 114.8403 | 21.2502 | 0.003215 | 0.002951 | 247.0735 | 0.27081 | 8.4019 | 0.003751 | 0.004867 | 0.00482 | 6.3e-05 | 0.001004 | 0.25658 | 1.3523 | 2.5481 | 99.37 | 21.02 | 37.399 | 99.1896 | 112.1791 | 92.8174 | 0.009382 | 0.96469 | 0.023023 | 23.866 | 0.30554 | 0.1611 | 0.016224 | 1.1124 | 1.1124 | 0.005952
2019-03-21 12:39 | 1553186355.3824 | 248 | 0.20625 | 0.002796 | 0.000125 | 100.0241 | 97.0091 | 114.7898 | 21.2509 | 0.003264 | 0.002948 | 247.0175 | 0.27107 | 8.402 | 0.003745 | 0.004933 | 0.004816 | 6.6e-05 | 0.001007 | 0.26377 | 1.3858 | 2.5481 | 99.37 | 21.02 | 37.399 | 100.3763 | 112.0415 | 92.7248 | 0.009382 | 0.96493 | 0.023021 | 23.8577 | 0.27376 | 0.16126 | 0.016169 | 1.1085 | 1.1085 | 0.005979
2019-03-21 12:39 | 1553186356.3835 | 249 | 0.20626 | 0.0028 | 0.000127 | 99.8983 | 96.9246 | 114.7297 | 21.251 | 0.003282 | 0.00288 | 246.9474 | 0.27024 | 8.4024 | 0.003786 | 0.005155 | 0.005046 | 6.2e-05 | 0.001004 | 0.26576 | 1.3321 | 2.548 | 99.37 | 21.02 | 37.397 | 100.8034 | 113.0417 | 91.0923 | 0.009382 | 0.96926 | 0.023067 | 23.7982 | 0.2703 | 0.16077 | 0.016109 | 1.1064 | 1.1064 | 0.005972
2019-03-21 12:39 | 1553186357.3826 | 250 | 0.20626 | 0.002801 | 0.000128 | 99.9418 | 96.8197 | 114.4762 | 21.2512 | 0.003284 | 0.002968 | 246.9037 | 0.27067 | 8.4026 | 0.003718 | 0.005024 | 0.005015 | 6.4e-05 | 0.001007 | 0.27721 | 1.3118 | 2.548 | 99.37 | 21.02 | 37.395 | 100.8563 | 111.3654 | 93.2133 | 0.009381 | 0.96365 | 0.022974 | 23.841 | 0.28388 | 0.16102 | 0.016141 | 1.1041 | 1.1041 | 0.005994
2019-03-21 12:39 | 1553186358.3836 | 251 | 0.20627 | 0.002806 | 0.000131 | 99.8977 | 96.4417 | 114.3758 | 21.252 | 0.003207 | 0.002893 | 246.8453 | 0.27096 | 8.4 | 0.003696 | 0.005462 | 0.005318 | 6.7e-05 | 0.001006 | 0.27328 | 1.3036 | 2.548 | 99.37 | 21.02 | 37.39 | 98.9951 | 110.8413 | 91.4145 | 0.00938 | 0.9684 | 0.023021 | 23.772 | 0.31493 | 0.1612 | 0.016039 | 1.0992 | 1.0992 | 0.006027
2019-03-21 12:39 | 1553186359.3837 | 252 | 0.20625 | 0.00281 | 0.000132 | 99.9138 | 96.2915 | 114.3821 | 21.2514 | 0.003227 | 0.002977 | 246.7927 | 0.27079 | 8.4009 | 0.003737 | 0.005065 | 0.005038 | 6.6e-05 | 0.001004 | 0.25011 | 1.3447 | 2.548 | 99.38 | 21.02 | 37.388 | 99.4714 | 111.8351 | 93.4442 | 0.009379 | 0.96304 | 0.022958 | 23.8387 | 0.33031 | 0.16109 | 0.01614 | 1.103 | 1.103 | 0.006003
2019-03-21 12:39 | 1553186360.3841 | 253 | 0.20624 | 0.002818 | 0.000134 | 99.8379 | 96.3426 | 114.3283 | 21.2507 | 0.00321 | 0.002931 | 246.7109 | 0.27044 | 8.4081 | 0.00371 | 0.005576 | 0.005502 | 6.6e-05 | 0.000998 | 0.27865 | 1.2949 | 2.548 | 99.38 | 21.02 | 37.385 | 99.0581 | 111.1747 | 92.3165 | 0.009378 | 0.96601 | 0.022988 | 23.7964 | 0.31876 | 0.16088 | 0.016194 | 1.1081 | 1.1081 | 0.005967
2019-03-21 12:39 | 1553186361.3838 | 254 | 0.20622 | 0.002821 | 0.000135 | 99.8489 | 96.5244 | 114.1626 | 21.2508 | 0.003266 | 0.002884 | 246.6524 | 0.26953 | 8.4026 | 0.003641 | 0.005241 | 0.005209 | 6.5e-05 | 0.000992 | 0.25802 | 1.3116 | 2.548 | 99.38 | 21.02 | 37.383 | 100.4052 | 109.497 | 91.1952 | 0.009377 | 0.96899 | 0.023006 | 23.7427 | 0.30288 | 0.16034 | 0.016363 | 1.1205 | 1.1205 | 0.005881
2019-03-21 12:39 | 1553186362.3828 | 255 | 0.20624 | 0.002822 | 0.000136 | 99.7938 | 96.3327 | 114.0777 | 21.2513 | 0.003208 | 0.002952 | 246.5613 | 0.2706 | 8.4023 | 0.003656 | 0.005214 | 0.005131 | 6.3e-05 | 0.000985 | 0.26842 | 1.3491 | 2.5479 | 99.38 | 21.02 | 37.382 | 99.0047 | 109.8807 | 92.8331 | 0.009377 | 0.96465 | 0.022946 | 23.7871 | 0.31572 | 0.16098 | 0.01623 | 1.1084 | 1.1084 | 0.005969
2019-03-21 12:39 | 1553186363.3831 | 256 | 0.2062 | 0.002833 | 0.000136 | 99.8155 | 96.3964 | 114.0875 | 21.2516 | 0.003258 | 0.002974 | 246.4903 | 0.2707 | 8.4023 | 0.003725 | 0.005398 | 0.005361 | 6.5e-05 | 0.00099 | 0.25426 | 1.3136 | 2.548 | 99.38 | 21.02 | 37.38 | 100.2084 | 111.5575 | 93.3678 | 0.009377 | 0.96324 | 0.02293 | 23.8055 | 0.31176 | 0.16104 | 0.01644 | 1.122 | 1.122 | 0.005899
2019-03-21 12:39 | 1553186364.3843 | 257 | 0.20618 | 0.002843 | 0.000135 | 99.7398 | 96.2789 | 113.9453 | 21.2511 | 0.003216 | 0.002978 | 246.4223 | 0.27124 | 8.4018 | 0.003724 | 0.005064 | 0.004946 | 6.8e-05 | 0.00099 | 0.30845 | 1.3082 | 2.5479 | 99.37 | 21.03 | 37.379 | 99.2006 | 111.5334 | 93.4652 | 0.009383 | 0.96299 | 0.022913 | 23.7938 | 0.31589 | 0.16136 | 0.016546 | 1.1285 | 1.1285 | 0.005877
2019-03-21 12:39 | 1553186365.384 | 258 | 0.20618 | 0.002845 | 0.000134 | 99.75 | 96.3606 | 113.4376 | 21.251 | 0.003196 | 0.002978 | 246.3353 | 0.27042 | 8.4024 | 0.003683 | 0.005369 | 0.005279 | 6.6e-05 | 0.00099 | 0.24994 | 1.3426 | 2.548 | 99.37 | 21.03 | 37.377 | 98.7243 | 110.5249 | 93.4557 | 0.009383 | 0.96301 | 0.022862 | 23.7405 | 0.30921 | 0.16087 | 0.016572 | 1.1278 | 1.1278 | 0.005862
2019-03-21 12:39 | 1553186366.384 | 259 | 0.2062 | 0.002848 | 0.000133 | 99.6552 | 96.5425 | 113.247 | 21.2517 | 0.003238 | 0.002896 | 246.2933 | 0.27035 | 8.4062 | 0.003673 | 0.004957 | 0.004888 | 6.7e-05 | 0.000991 | 0.28385 | 1.3538 | 2.5481 | 99.37 | 21.03 | 37.375 | 99.7455 | 110.2818 | 91.4918 | 0.009382 | 0.9682 | 0.022905 | 23.6569 | 0.28384 | 0.16083 | 0.01644 | 1.121 | 1.121 | 0.005897
2019-03-21 12:39 | 1553186367.3831 | 260 | 0.2062 | 0.002856 | 0.000133 | 99.6984 | 96.5748 | 113.1784 | 21.2508 | 0.003249 | 0.002947 | 246.2225 | 0.27085 | 8.4007 | 0.003665 | 0.005897 | 0.005853 | 6.3e-05 | 0.000986 | 0.28672 | 1.3056 | 2.5481 | 99.37 | 21.03 | 37.373 | 100.0043 | 110.0971 | 92.717 | 0.009382 | 0.96496 | 0.022859 | 23.6894 | 0.28472 | 0.16113 | 0.016394 | 1.1156 | 1.1156 | 0.005936
2019-03-21 12:39 | 1553186368.3821 | 261 | 0.20619 | 0.002862 | 0.000133 | 99.6746 | 96.3947 | 112.9036 | 21.2507 | 0.003234 | 0.002961 | 246.1451 | 0.27066 | 8.4029 | 0.003697 | 0.005133 | 0.005031 | 6.4e-05 | 0.000981 | 0.27649 | 1.3295 | 2.5481 | 99.37 | 21.03 | 37.371 | 99.6323 | 110.8639 | 93.0455 | 0.009381 | 0.96409 | 0.022821 | 23.6713 | 0.29928 | 0.16101 | 0.016454 | 1.1178 | 1.1178 | 0.00592
2019-03-21 12:39 | 1553186369.3825 | 262 | 0.2062 | 0.002864 | 0.000135 | 99.6469 | 96.3 | 112.5963 | 21.251 | 0.003243 | 0.002942 | 246.0715 | 0.27074 | 8.4043 | 0.003655 | 0.005426 | 0.005415 | 6.3e-05 | 0.000988 | 0.2891 | 1.3191 | 2.548 | 99.37 | 21.03 | 37.37 | 99.8492 | 109.8464 | 92.5795 | 0.009381 | 0.96532 | 0.022805 | 23.624 | 0.30559 | 0.16106 | 0.016369 | 1.1111 | 1.1111 | 0.005958
2019-03-21 12:39 | 1553186370.3837 | 263 | 0.20615 | 0.002868 | 0.000138 | 99.5468 | 95.9975 | 112.5811 | 21.2508 | 0.003213 | 0.002924 | 246.0051 | 0.27092 | 8.4037 | 0.003666 | 0.004875 | 0.00483 | 6.5e-05 | 0.000988 | 0.29927 | 1.2593 | 2.5479 | 99.37 | 21.03 | 37.365 | 99.1413 | 110.1143 | 92.1633 | 0.00938 | 0.96642 | 0.022816 | 23.609 | 0.32473 | 0.16117 | 0.016688 | 1.1333 | 1.1333 | 0.005845
2019-03-21 12:40 | 1553186371.3823 | 264 | 0.20605 | 0.002875 | 0.000139 | 99.5298 | 96.158 | 112.2381 | 21.2508 | 0.003231 | 0.002959 | 245.9473 | 0.27002 | 8.4029 | 0.003615 | 0.005162 | 0.005065 | 6.3e-05 | 0.000983 | 0.26217 | 1.3171 | 2.548 | 99.37 | 21.03 | 37.364 | 99.5721 | 108.8672 | 92.9971 | 0.009379 | 0.96422 | 0.022755 | 23.5999 | 0.30827 | 0.16064 | 0.017262 | 1.1693 | 1.1693 | 0.005646
2019-03-21 12:40 | 1553186372.3827 | 265 | 0.206 | 0.002882 | 0.000139 | 99.5894 | 96.1184 | 112.1715 | 21.251 | 0.003271 | 0.003015 | 245.8596 | 0.27035 | 8.4018 | 0.003601 | 0.004804 | 0.004735 | 6.4e-05 | 0.000985 | 0.25155 | 1.2804 | 2.548 | 99.37 | 21.03 | 37.361 | 100.5264 | 108.5363 | 94.3602 | 0.009379 | 0.96064 | 0.022706 | 23.6367 | 0.3173 | 0.16083 | 0.017589 | 1.1889 | 1.1889 | 0.00556
2019-03-21 12:40 | 1553186373.3838 | 266 | 0.20577 | 0.002886 | 0.000137 | 99.5521 | 95.8773 | 111.822 | 21.2515 | 0.003218 | 0.002984 | 245.7963 | 0.27073 | 8.4023 | 0.003605 | 0.005326 | 0.005242 | 6.4e-05 | 0.000987 | 0.24779 | 1.3068 | 2.5481 | 99.37 | 21.03 | 37.358 | 99.2598 | 108.6211 | 93.6064 | 0.009378 | 0.96261 | 0.022694 | 23.5757 | 0.33642 | 0.16106 | 0.018945 | 1.2805 | 1.2805 | 0.00517
2019-03-21 12:40 | 1553186374.3825 | 267 | 0.20539 | 0.00291 | 0.000135 | 99.4618 | 95.9311 | 111.5088 | 21.2514 | 0.003258 | 0.003062 | 245.7199 | 0.27052 | 8.4031 | 0.003627 | 0.004892 | 0.004891 | 6.5e-05 | 0.000986 | 0.27129 | 1.3315 | 2.548 | 99.37 | 21.03 | 37.358 | 100.2121 | 109.1617 | 95.4847 | 0.009378 | 0.95771 | 0.022605 | 23.6028 | 0.32329 | 0.16093 | 0.021205 | 1.4283 | 1.4283 | 0.004631
2019-03-21 12:40 | 1553186375.3825 | 268 | 0.20488 | 0.002933 | 0.000131 | 99.542 | 95.8715 | 111.3239 | 21.2516 | 0.003304 | 0.003031 | 245.6293 | 0.27052 | 8.4037 | 0.003635 | 0.005708 | 0.005577 | 6.5e-05 | 0.000986 | 0.26792 | 1.3528 | 2.548 | 99.37 | 21.04 | 37.357 | 101.3432 | 109.3677 | 94.7279 | 0.009383 | 0.95968 | 0.022609 | 23.559 | 0.33605 | 0.16093 | 0.024265 | 1.6355 | 1.6355 | 0.004044
2019-03-21 12:40 | 1553186376.384 | 269 | 0.20452 | 0.002982 | 0.000129 | 99.4291 | 96.0687 | 111.1227 | 21.2519 | 0.003364 | 0.003078 | 245.5263 | 0.26988 | 8.4044 | 0.003698 | 0.004884 | 0.004812 | 6.7e-05 | 0.000987 | 0.26521 | 1.308 | 2.5481 | 99.37 | 21.04 | 37.355 | 102.7894 | 110.8886 | 95.8791 | 0.009383 | 0.95669 | 0.022553 | 23.5745 | 0.30752 | 0.16055 | 0.026351 | 1.7721 | 1.7721 | 0.003724
2019-03-21 12:40 | 1553186377.3839 | 270 | 0.20416 | 0.003052 | 0.000127 | 99.4416 | 95.7615 | 111.0077 | 21.2518 | 0.00331 | 0.003094 | 245.3995 | 0.27079 | 8.4003 | 0.003705 | 0.004947 | 0.004863 | 6.7e-05 | 0.000984 | 0.25409 | 1.3056 | 2.548 | 99.37 | 21.04 | 37.353 | 101.4808 | 111.0629 | 96.2676 | 0.009382 | 0.95568 | 0.02253 | 23.5747 | 0.3373 | 0.16109 | 0.0284 | 1.9082 | 1.9082 | 0.00347
2019-03-21 12:40 | 1553186378.3826 | 271 | 0.20387 | 0.003161 | 0.000125 | 99.4934 | 95.3545 | 110.649 | 21.2519 | 0.003397 | 0.003104 | 245.2998 | 0.27038 | 8.4027 | 0.003708 | 0.005562 | 0.005522 | 6.7e-05 | 0.000989 | 0.28567 | 1.3064 | 2.548 | 99.37 | 21.04 | 37.349 | 103.5901 | 111.1302 | 96.5082 | 0.009381 | 0.95506 | 0.022486 | 23.5443 | 0.38005 | 0.16085 | 0.03 | 2.0119 | 2.0119 | 0.003286
2019-03-21 12:40 | 1553186379.3827 | 272 | 0.20355 | 0.00329 | 0.000125 | 99.4475 | 95.0286 | 110.3704 | 21.2521 | 0.00333 | 0.003082 | 245.1685 | 0.27025 | 8.3998 | 0.003684 | 0.004507 | 0.004451 | 6.6e-05 | 0.000991 | 0.2677 | 1.3107 | 2.548 | 99.37 | 21.04 | 37.347 | 101.9602 | 110.5535 | 95.971 | 0.009381 | 0.95645 | 0.022474 | 23.4975 | 0.40655 | 0.16077 | 0.031786 | 2.1307 | 2.1307 | 0.003101
2019-03-21 12:40 | 1553186380.3842 | 273 | 0.20334 | 0.003446 | 0.000124 | 99.4517 | 94.9652 | 109.7754 | 21.2525 | 0.003411 | 0.003132 | 245.0596 | 0.2698 | 8.4062 | 0.003736 | 0.005356 | 0.005255 | 6.3e-05 | 0.000986 | 0.24226 | 1.3182 | 2.548 | 99.37 | 21.04 | 37.345 | 103.9255 | 111.8151 | 97.1669 | 0.00938 | 0.95336 | 0.022377 | 23.472 | 0.4129 | 0.1605 | 0.032847 | 2.1923 | 2.1923 | 0.003009
2019-03-21 12:40 | 1553186381.3841 | 274 | 0.20282 | 0.003615 | 0.000125 | 99.4032 | 94.8552 | 109.2562 | 21.2514 | 0.003363 | 0.003134 | 244.9353 | 0.27013 | 8.4044 | 0.003683 | 0.005058 | 0.005113 | 6.3e-05 | 0.000992 | 0.29015 | 1.3432 | 2.548 | 99.37 | 21.04 | 37.342 | 102.7566 | 110.5299 | 97.2222 | 0.00938 | 0.95322 | 0.022323 | 23.4182 | 0.41889 | 0.1607 | 0.035735 | 2.3791 | 2.3791 | 0.002776
2019-03-21 12:40 | 1553186382.3835 | 275 | 0.2026 | 0.003798 | 0.000127 | 99.3853 | 94.7173 | 108.8461 | 21.2511 | 0.003399 | 0.003163 | 244.8497 | 0.27067 | 8.4043 | 0.003666 | 0.005225 | 0.005183 | 6.8e-05 | 0.000993 | 0.28838 | 1.2837 | 2.548 | 99.38 | 21.04 | 37.342 | 103.6232 | 110.1029 | 97.9147 | 0.009379 | 0.95144 | 0.02226 | 23.3961 | 0.4303 | 0.16102 | 0.036853 | 2.4464 | 2.4464 | 0.002705
2019-03-21 12:40 | 1553186383.3823 | 276 | 0.20243 | 0.003983 | 0.000129 | 99.4621 | 94.3314 | 108.6947 | 21.2516 | 0.003439 | 0.003206 | 244.7042 | 0.26953 | 8.4024 | 0.003643 | 0.004916 | 0.004818 | 6.4e-05 | 0.000992 | 0.26753 | 1.3511 | 2.5481 | 99.38 | 21.04 | 37.341 | 104.5992 | 109.5521 | 98.9548 | 0.009378 | 0.94878 | 0.022213 | 23.4125 | 0.47373 | 0.16034 | 0.037632 | 2.4927 | 2.4927 | 0.002644
2019-03-21 12:40 | 1553186384.3834 | 277 | 0.20201 | 0.004171 | 0.000132 | 99.4458 | 94.219 | 108.5926 | 21.2522 | 0.003475 | 0.003181 | 244.5684 | 0.2705 | 8.4031 | 0.003666 | 0.005486 | 0.005375 | 6.3e-05 | 0.000987 | 0.25122 | 1.3059 | 2.548 | 99.38 | 21.04 | 37.339 | 105.4864 | 110.1047 | 98.3531 | 0.009378 | 0.95032 | 0.022221 | 23.3826 | 0.48292 | 0.16092 | 0.039879 | 2.6421 | 2.6421 | 0.002503
2019-03-21 12:40 | 1553186385.3821 | 278 | 0.20189 | 0.004354 | 0.000134 | 99.4505 | 93.2667 | 108.6215 | 21.2521 | 0.003386 | 0.003232 | 244.4633 | 0.27008 | 8.4026 | 0.003636 | 0.004988 | 0.004978 | 6e-05 | 0.000989 | 0.26897 | 1.3097 | 2.5479 | 99.38 | 21.04 | 37.337 | 103.3073 | 109.3903 | 99.5961 | 0.009377 | 0.94715 | 0.022187 | 23.4248 | 0.57421 | 0.16067 | 0.040365 | 2.6701 | 2.6701 | 0.002473
2019-03-21 12:40 | 1553186386.3831 | 279 | 0.20172 | 0.004545 | 0.000136 | 99.3567 | 92.9918 | 108.5701 | 21.2513 | 0.0035 | 0.003168 | 244.3324 | 0.27005 | 8.4044 | 0.003669 | 0.005455 | 0.005388 | 6.1e-05 | 0.000984 | 0.27273 | 1.3295 | 2.5479 | 99.38 | 21.04 | 37.332 | 106.0795 | 110.1889 | 98.0427 | 0.009376 | 0.95111 | 0.022228 | 23.3704 | 0.59217 | 0.16065 | 0.041155 | 2.7272 | 2.7272 | 0.002421
2019-03-21 12:40 | 1553186387.3835 | 280 | 0.20173 | 0.004726 | 0.000136 | 99.4073 | 92.6346 | 108.5689 | 21.2512 | 0.003455 | 0.003225 | 244.2501 | 0.2709 | 8.4009 | 0.003653 | 0.005297 | 0.00525 | 6.3e-05 | 0.000975 | 0.26753 | 1.3201 | 2.5482 | 99.38 | 21.04 | 37.329 | 104.9797 | 109.8048 | 99.4101 | 0.009375 | 0.94762 | 0.022187 | 23.4133 | 0.63116 | 0.16116 | 0.040859 | 2.7026 | 2.7026 | 0.002451
2019-03-21 12:40 | 1553186388.3832 | 281 | 0.20173 | 0.004913 | 0.000136 | 99.3644 | 92.0682 | 108.7905 | 21.2512 | 0.003556 | 0.003222 | 244.1116 | 0.27092 | 8.403 | 0.003602 | 0.005148 | 0.005042 | 6.1e-05 | 0.000974 | 0.26537 | 1.3107 | 2.5481 | 99.38 | 21.04 | 37.327 | 107.4509 | 108.546 | 99.3521 | 0.009375 | 0.94777 | 0.022211 | 23.4353 | 0.68215 | 0.16117 | 0.040649 | 2.6917 | 2.6917 | 0.002461
2019-03-21 12:40 | 1553186389.3826 | 282 | 0.20171 | 0.005092 | 0.000135 | 99.3704 | 92.2524 | 108.8123 | 21.2512 | 0.003476 | 0.003173 | 244.0254 | 0.27094 | 8.3995 | 0.003685 | 0.005358 | 0.00527 | 6.4e-05 | 0.000979 | 0.29645 | 1.3015 | 2.548 | 99.38 | 21.04 | 37.325 | 105.4925 | 110.5786 | 98.1726 | 0.009374 | 0.95078 | 0.022249 | 23.4005 | 0.66481 | 0.16118 | 0.040565 | 2.6907 | 2.6907 | 0.002462
2019-03-21 12:40 | 1553186390.3833 | 283 | 0.20185 | 0.005241 | 0.000131 | 99.4228 | 92.1233 | 109.124 | 21.2509 | 0.00351 | 0.003253 | 243.9565 | 0.26942 | 8.4035 | 0.003602 | 0.005376 | 0.005245 | 6e-05 | 0.00098 | 0.27455 | 1.2944 | 2.548 | 99.38 | 21.04 | 37.324 | 106.3137 | 108.5646 | 100.1032 | 0.009374 | 0.94586 | 0.022223 | 23.4949 | 0.68205 | 0.16028 | 0.039539 | 2.6199 | 2.6199 | 0.002514
2019-03-21 12:40 | 1553186391.3836 | 284 | 0.20206 | 0.005376 | 0.000128 | 99.3674 | 91.8808 | 109.4125 | 21.2509 | 0.003497 | 0.003245 | 243.8652 | 0.27108 | 8.4019 | 0.003707 | 0.005592 | 0.00551 | 5.9e-05 | 0.000977 | 0.2557 | 1.284 | 2.5479 | 99.38 | 21.04 | 37.323 | 106.0114 | 111.118 | 99.9117 | 0.009374 | 0.94635 | 0.022258 | 23.5199 | 0.70064 | 0.16127 | 0.03812 | 2.5301 | 2.5301 | 0.00262
2019-03-21 12:40 | 1553186392.3833 | 285 | 0.20224 | 0.005487 | 0.000125 | 99.4183 | 92.1625 | 109.4954 | 21.2499 | 0.003536 | 0.003206 | 243.773 | 0.26979 | 8.4002 | 0.003671 | 0.005095 | 0.005085 | 5.9e-05 | 0.000981 | 0.27632 | 1.2975 | 2.5479 | 99.38 | 21.04 | 37.322 | 106.9591 | 110.2413 | 98.9609 | 0.009374 | 0.94876 | 0.022295 | 23.4988 | 0.67785 | 0.1605 | 0.036956 | 2.4572 | 2.4572 | 0.002685
2019-03-21 12:40 | 1553186393.3839 | 286 | 0.2026 | 0.005569 | 0.000122 | 99.438 | 92.6438 | 109.6245 | 21.2503 | 0.003518 | 0.003226 | 243.7066 | 0.27006 | 8.4016 | 0.003744 | 0.004861 | 0.004741 | 5.9e-05 | 0.000982 | 0.28241 | 1.2858 | 2.5478 | 99.38 | 21.04 | 37.321 | 106.5084 | 112.0103 | 99.443 | 0.009373 | 0.94754 | 0.022293 | 23.5279 | 0.63303 | 0.16066 | 0.034715 | 2.3084 | 2.3084 | 0.002861
2019-03-21 12:40 | 1553186394.3826 | 287 | 0.20268 | 0.00562 | 0.000119 | 99.3933 | 93.5964 | 109.6942 | 21.2507 | 0.003523 | 0.003202 | 243.6419 | 0.27025 | 8.4032 | 0.003693 | 0.005094 | 0.005082 | 5.7e-05 | 0.000983 | 0.29192 | 1.2889 | 2.5479 | 99.38 | 21.04 | 37.317 | 106.6465 | 110.7587 | 98.8676 | 0.009372 | 0.949 | 0.022318 | 23.5172 | 0.53751 | 0.16077 | 0.034134 | 2.2724 | 2.2724 | 0.002908
2019-03-21 12:40 | 1553186395.3826 | 288 | 0.20276 | 0.005646 | 0.000117 | 99.3436 | 94.4233 | 110.1528 | 21.2504 | 0.003549 | 0.003305 | 243.5465 | 0.2702 | 8.4024 | 0.003666 | 0.004805 | 0.004683 | 6.2e-05 | 0.000984 | 0.28495 | 1.2904 | 2.5479 | 99.38 | 21.04 | 37.316 | 107.2722 | 110.0999 | 101.3602 | 0.009372 | 0.94269 | 0.02229 | 23.6451 | 0.45435 | 0.16074 | 0.033655 | 2.2378 | 2.2378 | 0.002952
2019-03-21 12:40 | 1553186396.3826 | 289 | 0.20297 | 0.005652 | 0.000114 | 99.3614 | 94.7694 | 110.3632 | 21.2508 | 0.003499 | 0.003226 | 243.481 | 0.26991 | 8.4055 | 0.003683 | 0.005463 | 0.005427 | 5.9e-05 | 0.000983 | 0.25282 | 1.3427 | 2.548 | 99.38 | 21.04 | 37.318 | 106.0537 | 110.5278 | 99.452 | 0.009373 | 0.94751 | 0.022368 | 23.6073 | 0.42324 | 0.16057 | 0.032357 | 2.1593 | 2.1593 | 0.003056
2019-03-21 12:40 | 1553186397.383 | 290 | 0.20307 | 0.005641 | 0.000113 | 99.3561 | 95.4274 | 110.5474 | 21.2511 | 0.003548 | 0.003276 | 243.4031 | 0.27006 | 8.4002 | 0.003663 | 0.004807 | 0.004751 | 6e-05 | 0.000984 | 0.25857 | 1.3699 | 2.5481 | 99.38 | 21.04 | 37.319 | 107.2547 | 110.0296 | 100.645 | 0.009373 | 0.94449 | 0.022351 | 23.6648 | 0.36086 | 0.16066 | 0.031783 | 2.1194 | 2.1194 | 0.003116
2019-03-21 12:40 | 1553186398.3828 | 291 | 0.20315 | 0.005617 | 0.000111 | 99.3409 | 95.9921 | 110.8878 | 21.2522 | 0.003535 | 0.003225 | 243.3277 | 0.26988 | 8.4034 | 0.003664 | 0.005652 | 0.005571 | 5.9e-05 | 0.000987 | 0.28888 | 1.3617 | 2.548 | 99.38 | 21.04 | 37.32 | 106.9227 | 110.0583 | 99.4345 | 0.009373 | 0.94756 | 0.022422 | 23.6628 | 0.30672 | 0.16055 | 0.031362 | 2.0981 | 2.0981 | 0.003145
2019-03-21 12:40 | 1553186399.3828 | 292 | 0.20319 | 0.005585 | 0.00011 | 99.3641 | 96.0091 | 111.1636 | 21.2512 | 0.003562 | 0.003203 | 243.2572 | 0.27029 | 8.402 | 0.003708 | 0.004423 | 0.004384 | 5.7e-05 | 0.000988 | 0.29374 | 1.3873 | 2.548 | 99.38 | 21.04 | 37.321 | 107.5895 | 111.1309 | 98.902 | 0.009373 | 0.94891 | 0.022466 | 23.6752 | 0.30723 | 0.1608 | 0.031145 | 2.0877 | 2.0877 | 0.003166
2019-03-21 12:40 | 1553186400.3832 | 293 | 0.20326 | 0.005547 | 0.000108 | 99.3427 | 96.096 | 111.2443 | 21.2509 | 0.003517 | 0.003265 | 243.148 | 0.27122 | 8.4018 | 0.003646 | 0.005346 | 0.005287 | 6e-05 | 0.000992 | 0.28528 | 1.3714 | 2.548 | 99.38 | 21.04 | 37.321 | 106.5019 | 109.62 | 100.377 | 0.009373 | 0.94517 | 0.02243 | 23.7307 | 0.29721 | 0.16135 | 0.030757 | 2.0585 | 2.0585 | 0.003221
2019-03-21 12:40 | 1553186401.3819 | 294 | 0.20337 | 0.005516 | 0.000107 | 99.4011 | 96.207 | 111.4389 | 21.2517 | 0.003525 | 0.003244 | 243.0856 | 0.27032 | 8.4043 | 0.0037 | 0.005042 | 0.005003 | 6e-05 | 0.000992 | 0.27057 | 1.3574 | 2.5481 | 99.38 | 21.04 | 37.322 | 106.6845 | 110.9398 | 99.8835 | 0.009374 | 0.94642 | 0.022464 | 23.7358 | 0.29214 | 0.16082 | 0.03015 | 2.0211 | 2.0211 | 0.00327
2019-03-21 12:40 | 1553186402.382 | 295 | 0.20355 | 0.005471 | 0.000106 | 99.3555 | 96.3041 | 111.6302 | 21.2513 | 0.003562 | 0.003212 | 242.9944 | 0.27105 | 8.4021 | 0.003654 | 0.005446 | 0.005392 | 5.5e-05 | 0.000996 | 0.30143 | 1.3786 | 2.5481 | 99.38 | 21.04 | 37.325 | 107.5755 | 109.827 | 99.1184 | 0.009374 | 0.94836 | 0.022506 | 23.7318 | 0.27901 | 0.16125 | 0.029159 | 1.9584 | 1.9584 | 0.003384
2019-03-21 12:40 | 1553186403.382 | 296 | 0.2036 | 0.005424 | 0.000106 | 99.3924 | 96.6099 | 111.6065 | 21.2518 | 0.003557 | 0.003231 | 242.9476 | 0.26999 | 8.4039 | 0.003627 | 0.005024 | 0.004962 | 5.6e-05 | 0.000993 | 0.25802 | 1.3716 | 2.5481 | 99.38 | 21.04 | 37.326 | 107.4617 | 109.1667 | 99.5589 | 0.009375 | 0.94724 | 0.022491 | 23.7433 | 0.25398 | 0.16062 | 0.028894 | 1.9392 | 1.9392 | 0.003404
2019-03-21 12:40 | 1553186404.3825 | 297 | 0.20362 | 0.005373 | 0.000106 | 99.4231 | 96.7945 | 111.7701 | 21.2515 | 0.003519 | 0.003232 | 242.8396 | 0.27012 | 8.4054 | 0.003711 | 0.00523 | 0.005147 | 5.6e-05 | 0.000993 | 0.24635 | 1.3839 | 2.548 | 99.37 | 21.04 | 37.327 | 106.5386 | 111.2046 | 99.5955 | 0.009376 | 0.94715 | 0.022506 | 23.7619 | 0.23966 | 0.16069 | 0.028828 | 1.9361 | 1.9361 | 0.003411
2019-03-21 12:40 | 1553186405.3832 | 298 | 0.20361 | 0.005331 | 0.000107 | 99.4309 | 97.3994 | 111.645 | 21.2519 | 0.00357 | 0.00323 | 242.7762 | 0.27079 | 8.4016 | 0.003644 | 0.005159 | 0.005166 | 5.6e-05 | 0.000997 | 0.29103 | 1.3403 | 2.5481 | 99.37 | 21.04 | 37.328 | 107.7902 | 109.5778 | 99.5387 | 0.009376 | 0.94729 | 0.022495 | 23.7467 | 0.18464 | 0.16109 | 0.028971 | 1.9447 | 1.9447 | 0.003405
2019-03-21 12:40 | 1553186406.3832 | 299 | 0.20357 | 0.005279 | 0.000109 | 99.3135 | 97.1014 | 111.4814 | 21.2514 | 0.003521 | 0.003233 | 242.7042 | 0.27042 | 8.4046 | 0.003587 | 0.005438 | 0.005297 | 5.6e-05 | 0.000995 | 0.30596 | 1.3502 | 2.548 | 99.37 | 21.03 | 37.329 | 106.5849 | 108.1884 | 99.6235 | 0.009371 | 0.94708 | 0.022476 | 23.732 | 0.20148 | 0.16087 | 0.029239 | 1.9609 | 1.9609 | 0.003372
2019-03-21 12:40 | 1553186407.3827 | 300 | 0.20364 | 0.005237 | 0.000111 | 99.3394 | 97.1794 | 111.5868 | 21.2534 | 0.003542 | 0.003267 | 242.6258 | 0.27094 | 8.404 | 0.003566 | 0.005572 | 0.005518 | 5.8e-05 | 0.000998 | 0.26985 | 1.3412 | 2.548 | 99.37 | 21.03 | 37.328 | 107.1018 | 107.6768 | 100.4416 | 0.00937 | 0.945 | 0.022462 | 23.7693 | 0.19663 | 0.16118 | 0.028882 | 1.9356 | 1.9356 | 0.003422
2019-03-21 12:40 | 1553186408.3836 | 301 | 0.20363 | 0.005202 | 0.000113 | 99.3796 | 97.2305 | 111.4527 | 21.2527 | 0.003579 | 0.00326 | 242.5805 | 0.27014 | 8.4019 | 0.003523 | 0.004593 | 0.004528 | 5.9e-05 | 0.001002 | 0.26449 | 1.338 | 2.5481 | 99.37 | 21.03 | 37.328 | 108.0022 | 106.631 | 100.2801 | 0.00937 | 0.94541 | 0.022453 | 23.7499 | 0.19555 | 0.1607 | 0.028996 | 1.9425 | 1.9425 | 0.0034
2019-03-21 12:40 | 1553186409.3824 | 302 | 0.20363 | 0.005175 | 0.000114 | 99.3689 | 97.4239 | 111.4807 | 21.2529 | 0.00349 | 0.003242 | 242.4598 | 0.26951 | 8.402 | 0.003596 | 0.005607 | 0.005505 | 5.6e-05 | 0.001002 | 0.274 | 1.3716 | 2.5478 | 99.37 | 21.03 | 37.326 | 105.8325 | 108.4019 | 99.8245 | 0.00937 | 0.94657 | 0.02247 | 23.7384 | 0.17681 | 0.16033 | 0.02901 | 1.9448 | 1.9448 | 0.003388
2019-03-21 12:40 | 1553186410.3841 | 303 | 0.20363 | 0.005146 | 0.000114 | 99.4645 | 97.522 | 111.4391 | 21.2518 | 0.003568 | 0.003225 | 242.3892 | 0.27044 | 8.4009 | 0.003587 | 0.005614 | 0.00555 | 5.9e-05 | 0.001 | 0.26952 | 1.3237 | 2.5479 | 99.37 | 21.03 | 37.326 | 107.7335 | 108.188 | 99.4183 | 0.00937 | 0.9476 | 0.022478 | 23.721 | 0.17641 | 0.16088 | 0.029063 | 1.949 | 1.949 | 0.003393
2019-03-21 12:40 | 1553186411.3829 | 304 | 0.20348 | 0.005131 | 0.000114 | 99.357 | 97.1276 | 111.5518 | 21.2527 | 0.003522 | 0.00327 | 242.2886 | 0.27029 | 8.4006 | 0.003607 | 0.005017 | 0.004955 | 5.5e-05 | 0.000999 | 0.28263 | 1.3042 | 2.5478 | 99.37 | 21.03 | 37.326 | 106.6076 | 108.6862 | 100.5218 | 0.00937 | 0.9448 | 0.022456 | 23.7681 | 0.20298 | 0.1608 | 0.02994 | 2.0059 | 2.0059 | 0.003295
2019-03-21 12:40 | 1553186412.3835 | 305 | 0.20342 | 0.005118 | 0.000114 | 99.4139 | 97.1065 | 111.679 | 21.2526 | 0.003516 | 0.003271 | 242.2242 | 0.27008 | 8.404 | 0.003661 | 0.005268 | 0.005216 | 5.8e-05 | 0.000998 | 0.28622 | 1.2989 | 2.548 | 99.37 | 21.03 | 37.326 | 106.4745 | 109.9861 | 100.5413 | 0.00937 | 0.94475 | 0.022468 | 23.7823 | 0.21005 | 0.16067 | 0.030341 | 2.0339 | 2.0339 | 0.003247
2019-03-21 12:40 | 1553186413.3836 | 306 | 0.20336 | 0.005112 | 0.000116 | 99.3483 | 97.3704 | 111.7257 | 21.2528 | 0.003509 | 0.003245 | 242.1035 | 0.2699 | 8.4011 | 0.003615 | 0.005176 | 0.005118 | 5.6e-05 | 0.000999 | 0.27472 | 1.3019 | 2.5481 | 99.37 | 21.03 | 37.325 | 106.2944 | 108.8674 | 99.9002 | 0.00937 | 0.94637 | 0.022492 | 23.7668 | 0.17987 | 0.16056 | 0.030725 | 2.0617 | 2.0617 | 0.003201
2019-03-21 12:40 | 1553186414.3829 | 307 | 0.20329 | 0.005114 | 0.000119 | 99.3543 | 97.3108 | 111.6607 | 21.2534 | 0.003542 | 0.003231 | 242.0375 | 0.26947 | 8.4046 | 0.003682 | 0.005572 | 0.005553 | 5.8e-05 | 0.001001 | 0.29662 | 1.328 | 2.5479 | 99.37 | 21.03 | 37.324 | 107.0986 | 110.5027 | 99.58 | 0.009369 | 0.94719 | 0.022495 | 23.7497 | 0.18588 | 0.16031 | 0.031146 | 2.0901 | 2.0901 | 0.003152
2019-03-21 12:40 | 1553186415.3816 | 308 | 0.20321 | 0.00512 | 0.000122 | 99.3272 | 97.0656 | 111.887 | 21.2523 | 0.003581 | 0.003262 | 241.9525 | 0.26956 | 8.4022 | 0.003695 | 0.004512 | 0.00444 | 5.6e-05 | 0.001 | 0.28351 | 1.356 | 2.5481 | 99.37 | 21.03 | 37.324 | 108.0499 | 110.8117 | 100.3126 | 0.009369 | 0.94533 | 0.022496 | 23.7971 | 0.20601 | 0.16036 | 0.031583 | 2.1193 | 2.1193 | 0.00311
2019-03-21 12:40 | 1553186416.3817 | 309 | 0.20315 | 0.005137 | 0.000126 | 99.3114 | 96.6518 | 112.159 | 21.2529 | 0.003604 | 0.003228 | 241.8469 | 0.26997 | 8.3999 | 0.003746 | 0.005289 | 0.005135 | 5.7e-05 | 0.000996 | 0.26465 | 1.3252 | 2.548 | 99.37 | 21.03 | 37.326 | 108.6186 | 112.0674 | 99.4995 | 0.00937 | 0.94739 | 0.022548 | 23.8001 | 0.2428 | 0.16061 | 0.031906 | 2.1456 | 2.1456 | 0.003076
2019-03-21 12:40 | 1553186417.3819 | 310 | 0.2031 | 0.005156 | 0.000132 | 99.3808 | 96.6126 | 111.6882 | 21.2519 | 0.003569 | 0.003295 | 241.7597 | 0.26969 | 8.4026 | 0.003671 | 0.005272 | 0.005245 | 5.8e-05 | 0.000996 | 0.25337 | 1.3331 | 2.548 | 99.37 | 21.03 | 37.33 | 107.7618 | 110.2401 | 101.1228 | 0.009371 | 0.94328 | 0.022452 | 23.8018 | 0.25268 | 0.16044 | 0.032169 | 2.1538 | 2.1538 | 0.003062
2019-03-21 12:40 | 1553186418.3819 | 311 | 0.20309 | 0.005176 | 0.000138 | 99.3829 | 96.4168 | 111.3466 | 21.2525 | 0.003562 | 0.003219 | 241.6603 | 0.26971 | 8.4054 | 0.003725 | 0.00523 | 0.005175 | 5.8e-05 | 0.000994 | 0.30071 | 1.3378 | 2.548 | 99.37 | 21.03 | 37.33 | 107.5977 | 111.5513 | 99.282 | 0.009371 | 0.94795 | 0.022473 | 23.7068 | 0.27102 | 0.16045 | 0.032198 | 2.1574 | 2.1574 | 0.003057
2019-03-21 12:40 | 1553186419.3828 | 312 | 0.203 | 0.005201 | 0.000143 | 99.3698 | 96.6593 | 111.1833 | 21.2533 | 0.003589 | 0.003309 | 241.5842 | 0.27028 | 8.4012 | 0.003833 | 0.004928 | 0.004824 | 5.7e-05 | 0.000993 | 0.28816 | 1.3258 | 2.5481 | 99.37 | 21.03 | 37.332 | 108.2321 | 114.1883 | 101.4628 | 0.009371 | 0.94243 | 0.022391 | 23.7587 | 0.24737 | 0.16079 | 0.032725 | 2.1844 | 2.1844 | 0.003025
2019-03-21 12:40 | 1553186420.3829 | 313 | 0.20296 | 0.005228 | 0.000148 | 99.3611 | 97.0125 | 110.8511 | 21.2527 | 0.003539 | 0.003245 | 241.4804 | 0.26976 | 8.4043 | 0.003835 | 0.005807 | 0.005761 | 5.7e-05 | 0.000992 | 0.27903 | 1.3104 | 2.5479 | 99.37 | 21.03 | 37.335 | 107.0382 | 114.2271 | 99.9002 | 0.009372 | 0.94637 | 0.022404 | 23.6736 | 0.21397 | 0.16048 | 0.032946 | 2.2002 | 2.2002 | 0.002998
2019-03-21 12:40 | 1553186421.382 | 314 | 0.20289 | 0.005254 | 0.000154 | 99.3725 | 97.1058 | 110.7378 | 21.2528 | 0.00368 | 0.003297 | 241.4058 | 0.26928 | 8.4021 | 0.003825 | 0.00513 | 0.005081 | 5.7e-05 | 0.000991 | 0.28152 | 1.3382 | 2.5479 | 99.37 | 21.03 | 37.337 | 110.4452 | 113.9757 | 101.1716 | 0.009373 | 0.94316 | 0.022355 | 23.7018 | 0.20639 | 0.16019 | 0.033314 | 2.2195 | 2.2195 | 0.002966
2019-03-21 12:40 | 1553186422.383 | 315 | 0.2028 | 0.005288 | 0.000158 | 99.3832 | 97.287 | 110.7216 | 21.253 | 0.00356 | 0.003271 | 241.3002 | 0.26986 | 8.4021 | 0.003895 | 0.004916 | 0.004805 | 5.8e-05 | 0.000987 | 0.2557 | 1.295 | 2.548 | 99.37 | 21.03 | 37.34 | 107.5495 | 115.6827 | 100.5353 | 0.009373 | 0.94477 | 0.022372 | 23.6799 | 0.19068 | 0.16054 | 0.033784 | 2.2524 | 2.2524 | 0.002929
2019-03-21 12:40 | 1553186423.3824 | 316 | 0.20267 | 0.005324 | 0.000163 | 99.396 | 97.1698 | 110.5373 | 21.2533 | 0.003634 | 0.003336 | 241.1845 | 0.27037 | 8.4046 | 0.003959 | 0.005225 | 0.005197 | 6e-05 | 0.000985 | 0.25536 | 1.2873 | 2.548 | 99.37 | 21.03 | 37.342 | 109.3354 | 117.2559 | 102.1051 | 0.009374 | 0.94081 | 0.022307 | 23.7099 | 0.20261 | 0.16084 | 0.034542 | 2.2959 | 2.2959 | 0.002879
2019-03-21 12:40 | 1553186424.3831 | 317 | 0.20256 | 0.005362 | 0.000168 | 99.4163 | 97.131 | 110.1876 | 21.2517 | 0.003639 | 0.003322 | 241.082 | 0.27006 | 8.4004 | 0.003915 | 0.005615 | 0.005477 | 6e-05 | 0.000987 | 0.24458 | 1.3475 | 2.548 | 99.37 | 21.03 | 37.343 | 109.4536 | 116.1856 | 101.7617 | 0.009374 | 0.94168 | 0.022281 | 23.6615 | 0.208 | 0.16066 | 0.035132 | 2.3323 | 2.3323 | 0.002831
2019-03-21 12:40 | 1553186425.3821 | 318 | 0.20244 | 0.005398 | 0.000173 | 99.2947 | 96.9883 | 109.6418 | 21.2529 | 0.003656 | 0.003283 | 240.9713 | 0.2708 | 8.402 | 0.003953 | 0.005331 | 0.005312 | 6e-05 | 0.000987 | 0.28871 | 1.2912 | 2.548 | 99.37 | 21.03 | 37.347 | 109.8753 | 117.0947 | 100.8164 | 0.009375 | 0.94406 | 0.022254 | 23.573 | 0.21021 | 0.1611 | 0.035796 | 2.3732 | 2.3732 | 0.00279
2019-03-21 12:40 | 1553186426.3833 | 319 | 0.20231 | 0.005445 | 0.000175 | 99.3987 | 96.7117 | 109.0224 | 21.253 | 0.003691 | 0.00334 | 240.8737 | 0.26954 | 8.4022 | 0.003909 | 0.005842 | 0.005711 | 5.8e-05 | 0.000986 | 0.2636 | 1.2931 | 2.5478 | 99.37 | 21.03 | 37.348 | 110.7192 | 116.0278 | 102.2028 | 0.009375 | 0.94057 | 0.02215 | 23.5499 | 0.24512 | 0.16035 | 0.036484 | 2.4074 | 2.4074 | 0.002738
2019-03-21 12:40 | 1553186427.3833 | 320 | 0.20225 | 0.005489 | 0.000178 | 99.3719 | 96.692 | 108.809 | 21.2539 | 0.003617 | 0.003333 | 240.7577 | 0.26996 | 8.4014 | 0.00394 | 0.005373 | 0.005307 | 5.7e-05 | 0.000987 | 0.25033 | 1.2853 | 2.548 | 99.37 | 21.03 | 37.349 | 108.93 | 116.7799 | 102.0371 | 0.009376 | 0.94098 | 0.022134 | 23.5216 | 0.24453 | 0.1606 | 0.036803 | 2.4265 | 2.4265 | 0.00272
2019-03-21 12:40 | 1553186428.3834 | 321 | 0.20221 | 0.005539 | 0.000179 | 99.407 | 96.4532 | 109.071 | 21.2521 | 0.003657 | 0.003327 | 240.645 | 0.2707 | 8.4023 | 0.003983 | 0.005708 | 0.005595 | 5.6e-05 | 0.000987 | 0.23562 | 1.2953 | 2.5481 | 99.37 | 21.03 | 37.349 | 109.9012 | 117.8378 | 101.8817 | 0.009376 | 0.94137 | 0.022165 | 23.5451 | 0.26981 | 0.16104 | 0.037018 | 2.444 | 2.444 | 0.002708
2019-03-21 12:40 | 1553186429.3818 | 322 | 0.20214 | 0.005591 | 0.000179 | 99.3281 | 96.5663 | 108.9604 | 21.2528 | 0.003663 | 0.003408 | 240.5554 | 0.26894 | 8.4047 | 0.00393 | 0.005183 | 0.005156 | 5.8e-05 | 0.000989 | 0.28263 | 1.3026 | 2.548 | 99.37 | 21.03 | 37.35 | 110.0438 | 116.5348 | 103.8497 | 0.009376 | 0.93646 | 0.022096 | 23.5948 | 0.25222 | 0.15999 | 0.037327 | 2.4567 | 2.4567 | 0.002677
2019-03-21 12:40 | 1553186430.3815 | 323 | 0.20207 | 0.005646 | 0.000178 | 99.2899 | 96.8354 | 109.0194 | 21.2535 | 0.003647 | 0.003359 | 240.4458 | 0.26966 | 8.4045 | 0.003923 | 0.005212 | 0.005076 | 5.8e-05 | 0.000987 | 0.26017 | 1.2984 | 2.5481 | 99.37 | 21.03 | 37.353 | 109.6566 | 116.3788 | 102.6708 | 0.009377 | 0.9394 | 0.022136 | 23.5642 | 0.2239 | 0.16042 | 0.037695 | 2.4855 | 2.4855 | 0.002653
2019-03-21 12:40 | 1553186431.3825 | 324 | 0.20204 | 0.0057 | 0.000176 | 99.3918 | 96.8432 | 108.579 | 21.2528 | 0.003706 | 0.003364 | 240.3035 | 0.27061 | 8.4023 | 0.003939 | 0.004673 | 0.004652 | 5.9e-05 | 0.000987 | 0.27151 | 1.297 | 2.5479 | 99.37 | 21.03 | 37.354 | 111.0938 | 116.7588 | 102.7878 | 0.009377 | 0.93911 | 0.022088 | 23.5203 | 0.23235 | 0.16099 | 0.037804 | 2.4874 | 2.4874 | 0.00266
2019-03-21 12:41 | 1553186432.3825 | 325 | 0.20198 | 0.005751 | 0.000176 | 99.3346 | 96.7077 | 108.1051 | 21.2528 | 0.003664 | 0.003385 | 240.1793 | 0.271 | 8.4029 | 0.003928 | 0.004801 | 0.004729 | 6e-05 | 0.000988 | 0.25713 | 1.3037 | 2.548 | 99.37 | 21.03 | 37.357 | 110.0589 | 116.4833 | 103.3017 | 0.009378 | 0.93782 | 0.022025 | 23.4849 | 0.23973 | 0.16122 | 0.038094 | 2.4993 | 2.4993 | 0.002651
2019-03-21 12:41 | 1553186433.3816 | 326 | 0.20185 | 0.005808 | 0.000175 | 99.3602 | 97.1409 | 107.7319 | 21.2531 | 0.003684 | 0.003356 | 240.1041 | 0.2697 | 8.4045 | 0.003883 | 0.004886 | 0.004798 | 5.8e-05 | 0.00099 | 0.27151 | 1.3437 | 2.5482 | 99.37 | 21.03 | 37.357 | 110.5449 | 115.3948 | 102.5917 | 0.009378 | 0.9396 | 0.022007 | 23.4222 | 0.20206 | 0.16045 | 0.038841 | 2.5464 | 2.5464 | 0.00259
2019-03-21 12:41 | 1553186434.3827 | 327 | 0.20174 | 0.005867 | 0.000174 | 99.2935 | 97.2196 | 107.5791 | 21.2525 | 0.003703 | 0.003433 | 239.9836 | 0.27018 | 8.4012 | 0.003861 | 0.005513 | 0.005398 | 5.8e-05 | 0.000993 | 0.29574 | 1.3088 | 2.548 | 99.37 | 21.03 | 37.357 | 111.0076 | 114.8674 | 104.4547 | 0.009378 | 0.93496 | 0.021937 | 23.4636 | 0.1888 | 0.16073 | 0.039401 | 2.5749 | 2.5749 | 0.002566
2019-03-21 12:41 | 1553186435.3817 | 328 | 0.20153 | 0.005927 | 0.000172 | 99.3632 | 96.9713 | 107.4068 | 21.2524 | 0.00371 | 0.003417 | 239.8783 | 0.26965 | 8.4042 | 0.003913 | 0.005542 | 0.005487 | 5.6e-05 | 0.000985 | 0.30309 | 1.2887 | 2.5479 | 99.37 | 21.03 | 37.356 | 111.1859 | 116.1165 | 104.0655 | 0.009377 | 0.93593 | 0.021931 | 23.4327 | 0.21795 | 0.16041 | 0.040602 | 2.6527 | 2.6527 | 0.002485
2019-03-21 12:41 | 1553186436.3818 | 329 | 0.20151 | 0.005991 | 0.00017 | 99.329 | 96.9083 | 107.3427 | 21.2527 | 0.003729 | 0.003427 | 239.7449 | 0.27025 | 8.4 | 0.003914 | 0.005216 | 0.005207 | 5.8e-05 | 0.000985 | 0.25409 | 1.3178 | 2.548 | 99.37 | 21.03 | 37.356 | 111.6535 | 116.1548 | 104.3015 | 0.009377 | 0.93534 | 0.021918 | 23.433 | 0.22068 | 0.16077 | 0.040638 | 2.6535 | 2.6535 | 0.00249
2019-03-21 12:41 | 1553186437.3819 | 330 | 0.20145 | 0.006058 | 0.000168 | 99.3057 | 96.9052 | 107.4075 | 21.2526 | 0.003701 | 0.003459 | 239.6224 | 0.27046 | 8.4034 | 0.003829 | 0.004594 | 0.004516 | 6e-05 | 0.000988 | 0.26504 | 1.2926 | 2.5481 | 99.37 | 21.03 | 37.355 | 110.973 | 114.0823 | 105.0996 | 0.009377 | 0.93337 | 0.021901 | 23.4649 | 0.21887 | 0.1609 | 0.040912 | 2.6695 | 2.6695 | 0.002477
2019-03-21 12:41 | 1553186438.383 | 331 | 0.20128 | 0.006127 | 0.000166 | 99.3156 | 97.2352 | 107.2153 | 21.2525 | 0.003747 | 0.003452 | 239.5041 | 0.26895 | 8.4 | 0.003878 | 0.005256 | 0.005158 | 5.9e-05 | 0.000988 | 0.28119 | 1.2773 | 2.548 | 99.37 | 21.03 | 37.355 | 112.0796 | 115.2818 | 104.9084 | 0.009377 | 0.93384 | 0.021887 | 23.4379 | 0.18935 | 0.16 | 0.041887 | 2.7315 | 2.7315 | 0.002407
2019-03-21 12:41 | 1553186439.3823 | 332 | 0.20112 | 0.006194 | 0.000163 | 99.3269 | 97.052 | 107.0075 | 21.2526 | 0.003787 | 0.003502 | 239.3638 | 0.27055 | 8.4016 | 0.003886 | 0.005408 | 0.005313 | 5.8e-05 | 0.000987 | 0.24889 | 1.2972 | 2.548 | 99.37 | 21.03 | 37.356 | 113.0605 | 115.4814 | 106.1295 | 0.009377 | 0.93083 | 0.021831 | 23.453 | 0.20724 | 0.16095 | 0.042715 | 2.7783 | 2.7783 | 0.002381
2019-03-21 12:41 | 1553186440.3821 | 333 | 0.20093 | 0.006265 | 0.000161 | 99.3391 | 97.0578 | 107.1744 | 21.2525 | 0.003767 | 0.003437 | 239.2561 | 0.27002 | 8.4026 | 0.003994 | 0.005982 | 0.005876 | 5.9e-05 | 0.000991 | 0.25874 | 1.3087 | 2.5479 | 99.37 | 21.03 | 37.356 | 112.5779 | 118.1012 | 104.5655 | 0.009377 | 0.93469 | 0.021893 | 23.4228 | 0.20781 | 0.16064 | 0.043757 | 2.8542 | 2.8542 | 0.002313
2019-03-21 12:41 | 1553186441.383 | 334 | 0.20082 | 0.006339 | 0.000161 | 99.3153 | 96.9202 | 107.2089 | 21.2529 | 0.003848 | 0.003454 | 239.1375 | 0.27023 | 8.4019 | 0.004049 | 0.005104 | 0.005111 | 5.7e-05 | 0.000989 | 0.31868 | 1.2919 | 2.548 | 99.37 | 21.03 | 37.356 | 114.538 | 119.4499 | 104.9663 | 0.009377 | 0.9337 | 0.021885 | 23.439 | 0.21835 | 0.16076 | 0.04437 | 2.8931 | 2.8931 | 0.002284
2019-03-21 12:41 | 1553186442.382 | 335 | 0.20084 | 0.006416 | 0.000162 | 99.3022 | 97.3207 | 107.4249 | 21.2528 | 0.003878 | 0.003475 | 239.0293 | 0.26996 | 8.4032 | 0.003997 | 0.005295 | 0.005224 | 5.5e-05 | 0.000994 | 0.29479 | 1.2722 | 2.5481 | 99.37 | 21.03 | 37.355 | 115.2838 | 118.1869 | 105.4775 | 0.009377 | 0.93244 | 0.021892 | 23.4785 | 0.18028 | 0.1606 | 0.044109 | 2.877 | 2.877 | 0.002294
2019-03-21 12:41 | 1553186443.3825 | 336 | 0.20084 | 0.006497 | 0.000164 | 99.3543 | 97.2659 | 108.1157 | 21.2527 | 0.003838 | 0.003503 | 238.9047 | 0.26934 | 8.4039 | 0.003976 | 0.005458 | 0.005397 | 6e-05 | 0.000993 | 0.27168 | 1.2875 | 2.5478 | 99.37 | 21.03 | 37.355 | 114.3063 | 117.6664 | 106.1558 | 0.009377 | 0.93077 | 0.021943 | 23.5749 | 0.19002 | 0.16023 | 0.044068 | 2.8809 | 2.8809 | 0.002286
2019-03-21 12:41 | 1553186444.3832 | 337 | 0.2009 | 0.006584 | 0.000167 | 99.3361 | 97.1116 | 108.5624 | 21.2526 | 0.003827 | 0.003499 | 238.7837 | 0.27025 | 8.402 | 0.003912 | 0.005747 | 0.005612 | 5.6e-05 | 0.000989 | 0.28655 | 1.3021 | 2.548 | 99.37 | 21.04 | 37.352 | 114.0382 | 116.0955 | 106.0517 | 0.009382 | 0.93102 | 0.021991 | 23.6203 | 0.20258 | 0.16077 | 0.043566 | 2.8542 | 2.8542 | 0.002315
2019-03-21 12:41 | 1553186445.3822 | 338 | 0.20104 | 0.006661 | 0.00017 | 99.2896 | 96.4723 | 108.8708 | 21.2533 | 0.003809 | 0.003477 | 238.6743 | 0.27118 | 8.4052 | 0.003848 | 0.005706 | 0.005642 | 6e-05 | 0.000991 | 0.28799 | 1.2801 | 2.5479 | 99.37 | 21.04 | 37.351 | 113.5868 | 114.537 | 105.5212 | 0.009382 | 0.93233 | 0.022038 | 23.6373 | 0.25747 | 0.16132 | 0.04266 | 2.8007 | 2.8007 | 0.002367
2019-03-21 12:41 | 1553186446.3826 | 339 | 0.20113 | 0.006727 | 0.000173 | 99.3483 | 96.4904 | 109.0687 | 21.2523 | 0.003855 | 0.003417 | 238.5651 | 0.26991 | 8.4035 | 0.003812 | 0.005264 | 0.005274 | 6.3e-05 | 0.000992 | 0.27848 | 1.305 | 2.548 | 99.37 | 21.04 | 37.352 | 114.7044 | 113.6647 | 104.063 | 0.009382 | 0.93593 | 0.0221 | 23.6132 | 0.26108 | 0.16057 | 0.042015 | 2.766 | 2.766 | 0.002386
2019-03-21 12:41 | 1553186447.3814 | 340 | 0.20118 | 0.006775 | 0.000175 | 99.3471 | 96.6447 | 109.1877 | 21.2527 | 0.003855 | 0.003436 | 238.4762 | 0.26986 | 8.4016 | 0.003755 | 0.005637 | 0.005538 | 5.7e-05 | 0.000992 | 0.28031 | 1.3268 | 2.548 | 99.37 | 21.04 | 37.35 | 114.7206 | 112.274 | 104.5264 | 0.009382 | 0.93478 | 0.022099 | 23.6406 | 0.24668 | 0.16054 | 0.041674 | 2.7433 | 2.7433 | 0.002405
2019-03-21 12:41 | 1553186448.3815 | 341 | 0.20122 | 0.006803 | 0.000176 | 99.281 | 96.5959 | 108.9607 | 21.2532 | 0.003791 | 0.003504 | 238.356 | 0.26932 | 8.4015 | 0.003789 | 0.00528 | 0.005259 | 6.2e-05 | 0.000994 | 0.26145 | 1.2808 | 2.5479 | 99.37 | 21.04 | 37.349 | 113.156 | 113.111 | 106.1907 | 0.009381 | 0.93068 | 0.022027 | 23.668 | 0.24524 | 0.16022 | 0.041406 | 2.7168 | 2.7168 | 0.002424
2019-03-21 12:41 | 1553186449.3824 | 342 | 0.2012 | 0.006824 | 0.000175 | 99.3284 | 96.7897 | 108.9215 | 21.253 | 0.003866 | 0.003544 | 238.2648 | 0.27036 | 8.4019 | 0.003798 | 0.005466 | 0.005349 | 6e-05 | 0.000996 | 0.29966 | 1.3097 | 2.548 | 99.37 | 21.04 | 37.347 | 114.9771 | 113.3344 | 107.1534 | 0.009381 | 0.92833 | 0.021995 | 23.6937 | 0.23158 | 0.16084 | 0.041473 | 2.7174 | 2.7174 | 0.002433
2019-03-21 12:41 | 1553186450.3824 | 343 | 0.2013 | 0.006832 | 0.000173 | 99.3269 | 96.8006 | 108.6544 | 21.2533 | 0.003844 | 0.003528 | 238.1463 | 0.27016 | 8.4041 | 0.003861 | 0.005258 | 0.005217 | 5.7e-05 | 0.000992 | 0.27527 | 1.3272 | 2.5479 | 99.37 | 21.05 | 37.346 | 114.4348 | 114.8642 | 106.7522 | 0.009386 | 0.92931 | 0.02198 | 23.6522 | 0.23044 | 0.16072 | 0.040862 | 2.6756 | 2.6756 | 0.002469
2019-03-21 12:41 | 1553186451.3826 | 344 | 0.20122 | 0.006837 | 0.000171 | 99.3415 | 96.8599 | 108.752 | 21.2528 | 0.00393 | 0.003578 | 238.0479 | 0.27007 | 8.402 | 0.003868 | 0.004998 | 0.004942 | 6.1e-05 | 0.000995 | 0.29822 | 1.3103 | 2.548 | 99.37 | 21.05 | 37.343 | 116.5326 | 115.0196 | 107.9827 | 0.009386 | 0.92631 | 0.021954 | 23.7011 | 0.22628 | 0.16067 | 0.041383 | 2.7066 | 2.7066 | 0.00244
2019-03-21 12:41 | 1553186452.3816 | 345 | 0.20113 | 0.006833 | 0.000167 | 99.3203 | 97.0663 | 108.5858 | 21.2535 | 0.003896 | 0.003547 | 237.9311 | 0.2691 | 8.4002 | 0.003917 | 0.005941 | 0.005796 | 5.9e-05 | 0.000991 | 0.262 | 1.3222 | 2.5481 | 99.37 | 21.05 | 37.338 | 115.7258 | 116.2173 | 107.2138 | 0.009384 | 0.92818 | 0.02196 | 23.6591 | 0.20533 | 0.16009 | 0.041924 | 2.7428 | 2.7428 | 0.002399
2019-03-21 12:41 | 1553186453.3818 | 346 | 0.20096 | 0.006835 | 0.000164 | 99.3522 | 96.9863 | 107.8366 | 21.2526 | 0.003899 | 0.003495 | 237.8097 | 0.26976 | 8.4046 | 0.003884 | 0.005699 | 0.005646 | 5.7e-05 | 0.000993 | 0.27256 | 1.3298 | 2.5481 | 99.37 | 21.05 | 37.335 | 115.7955 | 115.4134 | 105.9731 | 0.009384 | 0.93122 | 0.02192 | 23.5388 | 0.21558 | 0.16048 | 0.042891 | 2.8011 | 2.8011 | 0.002355
2019-03-21 12:41 | 1553186454.3817 | 347 | 0.20093 | 0.006841 | 0.000161 | 99.3179 | 97.0077 | 107.4998 | 21.2529 | 0.003889 | 0.003576 | 237.7121 | 0.2698 | 8.4029 | 0.003845 | 0.005608 | 0.005587 | 5.4e-05 | 0.00099 | 0.2589 | 1.3541 | 2.548 | 99.37 | 21.05 | 37.332 | 115.5394 | 114.4653 | 107.9157 | 0.009383 | 0.92647 | 0.02183 | 23.5622 | 0.21052 | 0.1605 | 0.043093 | 2.8029 | 2.8029 | 0.002354
2019-03-21 12:41 | 1553186455.3827 | 348 | 0.20073 | 0.00686 | 0.000159 | 99.3433 | 97.2087 | 107.1758 | 21.253 | 0.003857 | 0.003605 | 237.5935 | 0.26961 | 8.4017 | 0.00392 | 0.005657 | 0.005523 | 6e-05 | 0.000991 | 0.24978 | 1.3309 | 2.5479 | 99.37 | 21.05 | 37.327 | 114.773 | 116.2959 | 108.6313 | 0.009382 | 0.92473 | 0.021776 | 23.5487 | 0.19429 | 0.16039 | 0.044237 | 2.8702 | 2.8702 | 0.002297
2019-03-21 12:41 | 1553186456.3818 | 349 | 0.20051 | 0.006894 | 0.000158 | 99.2563 | 97.2993 | 106.9674 | 21.2528 | 0.003938 | 0.003589 | 237.4937 | 0.26967 | 8.4008 | 0.003917 | 0.006052 | 0.006029 | 5.9e-05 | 0.000989 | 0.27151 | 1.2999 | 2.5479 | 99.37 | 21.05 | 37.324 | 116.7352 | 116.2255 | 108.2442 | 0.009381 | 0.92567 | 0.021766 | 23.5139 | 0.17812 | 0.16043 | 0.045554 | 2.9542 | 2.9542 | 0.002232
2019-03-21 12:41 | 1553186457.3829 | 350 | 0.20043 | 0.006935 | 0.000157 | 99.2801 | 97.2833 | 106.582 | 21.2534 | 0.003926 | 0.003554 | 237.3511 | 0.2694 | 8.4051 | 0.003937 | 0.005827 | 0.005763 | 5.8e-05 | 0.000992 | 0.2521 | 1.3035 | 2.548 | 99.37 | 21.05 | 37.318 | 116.4461 | 116.7177 | 107.3924 | 0.009379 | 0.92774 | 0.021751 | 23.4452 | 0.18174 | 0.16027 | 0.045939 | 2.9772 | 2.9772 | 0.002212
2019-03-21 12:41 | 1553186458.3829 | 351 | 0.20032 | 0.006988 | 0.000157 | 99.2965 | 97.2764 | 106.4057 | 21.2526 | 0.003936 | 0.003612 | 237.2388 | 0.27036 | 8.4008 | 0.003936 | 0.00562 | 0.00549 | 5.9e-05 | 0.000993 | 0.29015 | 1.2887 | 2.548 | 99.37 | 21.05 | 37.316 | 116.6883 | 116.6978 | 108.8033 | 0.009379 | 0.92432 | 0.021693 | 23.4692 | 0.18384 | 0.16084 | 0.046569 | 3.0098 | 3.0098 | 0.002196
2019-03-21 12:41 | 1553186459.383 | 352 | 0.20015 | 0.007058 | 0.000156 | 99.2744 | 97.1637 | 106.3595 | 21.2528 | 0.003937 | 0.003634 | 237.117 | 0.27033 | 8.4018 | 0.003906 | 0.005819 | 0.005712 | 5.6e-05 | 0.000993 | 0.2672 | 1.3133 | 2.5479 | 99.37 | 21.05 | 37.315 | 116.7121 | 115.9585 | 109.3239 | 0.009379 | 0.92306 | 0.021673 | 23.4801 | 0.19223 | 0.16082 | 0.047496 | 3.0669 | 3.0669 | 0.002155
2019-03-21 12:41 | 1553186460.383 | 353 | 0.20009 | 0.007129 | 0.000155 | 99.3591 | 97.2737 | 106.2093 | 21.2532 | 0.00392 | 0.003578 | 236.9909 | 0.26918 | 8.4061 | 0.00409 | 0.005597 | 0.00567 | 5.7e-05 | 0.000992 | 0.2589 | 1.2542 | 2.5477 | 99.37 | 21.05 | 37.31 | 116.2939 | 120.4408 | 107.9853 | 0.009377 | 0.9263 | 0.021696 | 23.4224 | 0.18973 | 0.16014 | 0.047737 | 3.0857 | 3.0857 | 0.002133
2019-03-21 12:41 | 1553186461.3815 | 354 | 0.20007 | 0.007202 | 0.000155 | 99.3394 | 97.1211 | 105.879 | 21.2534 | 0.003934 | 0.003635 | 236.8732 | 0.26981 | 8.4006 | 0.004145 | 0.005914 | 0.005795 | 5.6e-05 | 0.000996 | 0.28943 | 1.311 | 2.548 | 99.37 | 21.05 | 37.308 | 116.64 | 121.7887 | 109.3551 | 0.009377 | 0.92298 | 0.021624 | 23.428 | 0.202 | 0.16051 | 0.047784 | 3.0784 | 3.0784 | 0.002143
2019-03-21 12:41 | 1553186462.3825 | 355 | 0.20008 | 0.007269 | 0.000155 | 99.3156 | 97.0183 | 105.4701 | 21.2533 | 0.003988 | 0.003635 | 236.7323 | 0.26929 | 8.404 | 0.004013 | 0.005509 | 0.005442 | 6e-05 | 0.000993 | 0.26034 | 1.2964 | 2.5479 | 99.37 | 21.05 | 37.305 | 117.9681 | 118.5718 | 109.3575 | 0.009376 | 0.92298 | 0.021582 | 23.3828 | 0.20933 | 0.1602 | 0.047677 | 3.0656 | 3.0656 | 0.002148
2019-03-21 12:41 | 1553186463.3822 | 356 | 0.20009 | 0.007335 | 0.000156 | 99.3185 | 96.9461 | 105.4138 | 21.2525 | 0.003999 | 0.003705 | 236.6234 | 0.27009 | 8.4009 | 0.004046 | 0.005428 | 0.005409 | 5.6e-05 | 0.000989 | 0.274 | 1.3054 | 2.5481 | 99.37 | 21.05 | 37.304 | 118.2311 | 119.3686 | 111.0517 | 0.009376 | 0.91891 | 0.021528 | 23.4282 | 0.21626 | 0.16068 | 0.047538 | 3.0491 | 3.0491 | 0.002166
2019-03-21 12:41 | 1553186464.3812 | 357 | 0.20005 | 0.007392 | 0.000156 | 99.3483 | 97.1841 | 105.2253 | 21.2524 | 0.00401 | 0.003696 | 236.4859 | 0.26916 | 8.4021 | 0.004076 | 0.005704 | 0.005593 | 5.5e-05 | 0.000994 | 0.28152 | 1.3249 | 2.548 | 99.37 | 21.05 | 37.302 | 118.4933 | 120.1124 | 110.8395 | 0.009375 | 0.91941 | 0.021515 | 23.4008 | 0.197 | 0.16012 | 0.047663 | 3.0552 | 3.0552 | 0.002154
2019-03-21 12:41 | 1553186465.3814 | 358 | 0.20001 | 0.007437 | 0.000154 | 99.3352 | 97.1089 | 105.0931 | 21.2533 | 0.004037 | 0.003689 | 236.3568 | 0.26955 | 8.4031 | 0.004167 | 0.005808 | 0.00575 | 5.9e-05 | 0.000994 | 0.27688 | 1.3085 | 2.548 | 99.37 | 21.05 | 37.302 | 119.1624 | 122.3423 | 110.6692 | 0.009375 | 0.91982 | 0.021506 | 23.3809 | 0.20275 | 0.16036 | 0.04786 | 3.0666 | 3.0666 | 0.002149
2019-03-21 12:41 | 1553186466.3823 | 359 | 0.19987 | 0.007478 | 0.000151 | 99.3138 | 97.0912 | 105.0809 | 21.2532 | 0.004048 | 0.003705 | 236.2379 | 0.27077 | 8.4031 | 0.00404 | 0.005647 | 0.005609 | 5.3e-05 | 0.00099 | 0.24923 | 1.3202 | 2.548 | 99.37 | 21.05 | 37.302 | 119.4145 | 119.24 | 111.0555 | 0.009375 | 0.9189 | 0.021494 | 23.3913 | 0.20245 | 0.16108 | 0.04867 | 3.1168 | 3.1168 | 0.002124
2019-03-21 12:41 | 1553186467.3819 | 360 | 0.19948 | 0.00752 | 0.000146 | 99.3078 | 97.0397 | 105.2384 | 21.2534 | 0.00403 | 0.003731 | 236.1185 | 0.26962 | 8.4011 | 0.00411 | 0.005513 | 0.005393 | 5.5e-05 | 0.000991 | 0.2625 | 1.3163 | 2.548 | 99.37 | 21.05 | 37.302 | 118.9968 | 120.9306 | 111.696 | 0.009375 | 0.91737 | 0.021492 | 23.4284 | 0.20665 | 0.1604 | 0.050969 | 3.2636 | 3.2636 | 0.00202
2019-03-21 12:41 | 1553186468.3819 | 361 | 0.19919 | 0.007558 | 0.000142 | 99.3242 | 97.2005 | 105.2091 | 21.2531 | 0.004061 | 0.003745 | 236.0261 | 0.27049 | 8.4021 | 0.004102 | 0.005513 | 0.005487 | 5.9e-05 | 0.000995 | 0.30486 | 1.3407 | 2.5481 | 99.37 | 21.05 | 37.301 | 119.7395 | 120.7367 | 112.024 | 0.009375 | 0.91659 | 0.02148 | 23.4351 | 0.19332 | 0.16092 | 0.052627 | 3.3678 | 3.3678 | 0.001964
2019-03-21 12:41 | 1553186469.3816 | 362 | 0.19919 | 0.007612 | 0.000138 | 99.3546 | 97.4862 | 105.4213 | 21.253 | 0.004094 | 0.003796 | 235.8943 | 0.26988 | 8.4069 | 0.004101 | 0.005319 | 0.005244 | 5.4e-05 | 0.000993 | 0.26432 | 1.3334 | 2.548 | 99.37 | 21.06 | 37.301 | 120.5572 | 120.7192 | 113.2759 | 0.009381 | 0.91362 | 0.021467 | 23.4968 | 0.1698 | 0.16055 | 0.052557 | 3.3615 | 3.3615 | 0.001963
2019-03-21 12:41 | 1553186470.3816 | 363 | 0.1992 | 0.007683 | 0.000134 | 99.3397 | 97.3037 | 105.4351 | 21.2525 | 0.004104 | 0.003779 | 235.7775 | 0.26986 | 8.3991 | 0.004009 | 0.005392 | 0.005301 | 5.8e-05 | 0.000997 | 0.29662 | 1.3145 | 2.548 | 99.37 | 21.06 | 37.299 | 120.7979 | 118.4831 | 112.8668 | 0.00938 | 0.91459 | 0.02148 | 23.4859 | 0.18523 | 0.16054 | 0.052421 | 3.355 | 3.355 | 0.001967
2019-03-21 12:41 | 1553186471.382 | 364 | 0.19928 | 0.007763 | 0.000132 | 99.3513 | 97.3275 | 105.5685 | 21.2538 | 0.004104 | 0.00374 | 235.6535 | 0.27016 | 8.4026 | 0.004013 | 0.006077 | 0.006005 | 5.3e-05 | 0.000996 | 0.29192 | 1.3316 | 2.5478 | 99.37 | 21.05 | 37.298 | 120.8041 | 118.5611 | 111.9198 | 0.009374 | 0.91683 | 0.02152 | 23.4719 | 0.18408 | 0.16072 | 0.051833 | 3.3238 | 3.3238 | 0.001987
2019-03-21 12:41 | 1553186472.3817 | 365 | 0.1993 | 0.007841 | 0.000132 | 99.3197 | 97.3636 | 105.891 | 21.254 | 0.004127 | 0.003797 | 235.5632 | 0.26909 | 8.403 | 0.003948 | 0.005653 | 0.005604 | 5.7e-05 | 0.000992 | 0.26808 | 1.3603 | 2.548 | 99.37 | 21.06 | 37.298 | 121.3693 | 116.9825 | 113.3085 | 0.00938 | 0.91354 | 0.021514 | 23.5501 | 0.17792 | 0.16008 | 0.051638 | 3.3104 | 3.3104 | 0.001987
2019-03-21 12:41 | 1553186473.3817 | 366 | 0.19932 | 0.007923 | 0.000132 | 99.3433 | 97.629 | 106.1999 | 21.2538 | 0.004098 | 0.003761 | 235.4077 | 0.26932 | 8.4061 | 0.003931 | 0.004864 | 0.004824 | 5.7e-05 | 0.000997 | 0.29518 | 1.3297 | 2.5478 | 99.37 | 21.05 | 37.298 | 120.6548 | 116.5572 | 112.4345 | 0.009374 | 0.91561 | 0.02157 | 23.5577 | 0.1557 | 0.16022 | 0.051455 | 3.3073 | 3.3073 | 0.001991
2019-03-21 12:41 | 1553186474.3818 | 367 | 0.19942 | 0.007988 | 0.000132 | 99.309 | 97.5012 | 106.3111 | 21.254 | 0.004149 | 0.003769 | 235.3023 | 0.27042 | 8.4023 | 0.003995 | 0.005469 | 0.005491 | 5.7e-05 | 0.000995 | 0.23059 | 1.3235 | 2.548 | 99.37 | 21.05 | 37.298 | 121.8845 | 118.1351 | 112.6085 | 0.009374 | 0.9152 | 0.021576 | 23.5754 | 0.16432 | 0.16087 | 0.050746 | 3.2628 | 3.2628 | 0.002026
2019-03-21 12:41 | 1553186475.3822 | 368 | 0.19952 | 0.008039 | 0.000134 | 99.3713 | 97.6027 | 106.3996 | 21.2537 | 0.004153 | 0.003862 | 235.1825 | 0.26931 | 8.402 | 0.003829 | 0.00582 | 0.005748 | 5.5e-05 | 0.000994 | 0.30773 | 1.3034 | 2.5481 | 99.37 | 21.05 | 37.298 | 121.9995 | 114.0907 | 114.8812 | 0.009374 | 0.90984 | 0.021522 | 23.6545 | 0.16062 | 0.16021 | 0.050079 | 3.2118 | 3.2118 | 0.00205
2019-03-21 12:41 | 1553186476.3819 | 369 | 0.19961 | 0.008079 | 0.000136 | 99.3814 | 97.3865 | 106.455 | 21.2539 | 0.004146 | 0.003813 | 235.0367 | 0.26954 | 8.4053 | 0.003874 | 0.005692 | 0.005657 | 5.8e-05 | 0.000996 | 0.28241 | 1.2874 | 2.548 | 99.37 | 21.06 | 37.297 | 121.8267 | 115.1828 | 113.6828 | 0.00938 | 0.91266 | 0.021561 | 23.6241 | 0.18138 | 0.16035 | 0.049489 | 3.1797 | 3.1797 | 0.002073
2019-03-21 12:41 | 1553186477.3809 | 370 | 0.19967 | 0.008105 | 0.000139 | 99.3713 | 97.4859 | 106.0813 | 21.2541 | 0.004197 | 0.003778 | 234.9237 | 0.26932 | 8.4021 | 0.003865 | 0.005143 | 0.005047 | 5.6e-05 | 0.000992 | 0.25481 | 1.3259 | 2.548 | 99.37 | 21.06 | 37.296 | 123.0691 | 114.9662 | 112.8282 | 0.00938 | 0.91468 | 0.021547 | 23.5566 | 0.17133 | 0.16022 | 0.049094 | 3.1522 | 3.1522 | 0.002089
2019-03-21 12:41 | 1553186478.3821 | 371 | 0.19971 | 0.008123 | 0.000141 | 99.3656 | 97.7676 | 106.2182 | 21.2541 | 0.004161 | 0.003784 | 234.7623 | 0.27028 | 8.4037 | 0.003919 | 0.005223 | 0.005166 | 5.7e-05 | 0.000993 | 0.26305 | 1.3694 | 2.548 | 99.37 | 21.06 | 37.295 | 122.1906 | 116.2718 | 112.9957 | 0.009379 | 0.91428 | 0.021556 | 23.5769 | 0.14502 | 0.16079 | 0.048853 | 3.138 | 3.138 | 0.002106
2019-03-21 12:41 | 1553186479.3821 | 372 | 0.1998 | 0.008124 | 0.000143 | 99.4035 | 97.2066 | 106.0607 | 21.2537 | 0.004153 | 0.003814 | 234.6599 | 0.27001 | 8.4028 | 0.003921 | 0.006037 | 0.006018 | 5.7e-05 | 0.000989 | 0.2855 | 1.3035 | 2.5481 | 99.37 | 21.06 | 37.295 | 121.9899 | 116.3219 | 113.7096 | 0.009379 | 0.91259 | 0.02152 | 23.5812 | 0.19989 | 0.16063 | 0.048292 | 3.0968 | 3.0968 | 0.002132
2019-03-21 12:41 | 1553186480.3816 | 373 | 0.19991 | 0.008118 | 0.000146 | 99.4246 | 97.3074 | 106.0816 | 21.2548 | 0.004174 | 0.003821 | 234.5214 | 0.26903 | 8.4041 | 0.003941 | 0.005251 | 0.005141 | 5.5e-05 | 0.000993 | 0.28512 | 1.289 | 2.5479 | 99.37 | 21.06 | 37.294 | 122.5161 | 116.8122 | 113.8773 | 0.009379 | 0.9122 | 0.021517 | 23.5886 | 0.19252 | 0.16005 | 0.047679 | 3.0571 | 3.0571 | 0.002152
2019-03-21 12:41 | 1553186481.3815 | 374 | 0.19978 | 0.0081 | 0.000148 | 99.4565 | 97.0697 | 105.9288 | 21.255 | 0.004209 | 0.003829 | 234.3902 | 0.26897 | 8.406 | 0.003997 | 0.005772 | 0.005697 | 5.6e-05 | 0.000994 | 0.26952 | 1.3184 | 2.548 | 99.37 | 21.06 | 37.294 | 123.3617 | 118.171 | 114.0831 | 0.009379 | 0.91171 | 0.021496 | 23.5779 | 0.21727 | 0.16001 | 0.048476 | 3.1049 | 3.1049 | 0.002118
2019-03-21 12:41 | 1553186482.3812 | 375 | 0.19978 | 0.008081 | 0.00015 | 99.4812 | 97.3912 | 105.4356 | 21.2547 | 0.004244 | 0.00378 | 234.2312 | 0.26887 | 8.4025 | 0.003966 | 0.004538 | 0.004527 | 5.6e-05 | 0.000995 | 0.28191 | 1.3107 | 2.548 | 99.37 | 21.06 | 37.293 | 124.2286 | 117.4189 | 112.8834 | 0.009379 | 0.91455 | 0.021479 | 23.4864 | 0.18992 | 0.15995 | 0.048482 | 3.1027 | 3.1027 | 0.002119
2019-03-21 12:41 | 1553186483.3812 | 376 | 0.19971 | 0.008065 | 0.000152 | 99.4455 | 97.193 | 105.1694 | 21.2536 | 0.004245 | 0.003847 | 234.1043 | 0.26885 | 8.4 | 0.004071 | 0.005987 | 0.005882 | 5.5e-05 | 0.000998 | 0.2724 | 1.2917 | 2.5481 | 99.37 | 21.06 | 37.291 | 124.2356 | 119.9962 | 114.5166 | 0.009378 | 0.91069 | 0.021407 | 23.5064 | 0.20493 | 0.15994 | 0.048912 | 3.1195 | 3.1195 | 0.002107
2019-03-21 12:41 | 1553186484.3817 | 377 | 0.19961 | 0.008058 | 0.000153 | 99.4475 | 97.1726 | 104.8054 | 21.2541 | 0.004242 | 0.00381 | 233.9885 | 0.26869 | 8.4017 | 0.004091 | 0.005817 | 0.005713 | 5.9e-05 | 0.000994 | 0.28368 | 1.27 | 2.548 | 99.37 | 21.06 | 37.29 | 124.1773 | 120.4876 | 113.6199 | 0.009378 | 0.9128 | 0.021395 | 23.4385 | 0.20699 | 0.15984 | 0.049513 | 3.1559 | 3.1559 | 0.002082
2019-03-21 12:41 | 1553186485.3814 | 378 | 0.19956 | 0.008064 | 0.000153 | 99.4869 | 97.0036 | 104.4786 | 21.2542 | 0.00422 | 0.003847 | 233.8486 | 0.26971 | 8.3995 | 0.004205 | 0.005589 | 0.005519 | 5.8e-05 | 0.000993 | 0.25658 | 1.3357 | 2.5481 | 99.37 | 21.06 | 37.29 | 123.6359 | 123.2775 | 114.5188 | 0.009378 | 0.91069 | 0.021337 | 23.4291 | 0.22609 | 0.16045 | 0.049776 | 3.164 | 3.164 | 0.002084
2019-03-21 12:41 | 1553186486.3818 | 379 | 0.19953 | 0.00807 | 0.000153 | 99.5101 | 97.1314 | 104.238 | 21.2543 | 0.004255 | 0.00386 | 233.7387 | 0.26858 | 8.4061 | 0.004191 | 0.006157 | 0.006181 | 5.5e-05 | 0.000994 | 0.27992 | 1.3203 | 2.5479 | 99.37 | 21.06 | 37.289 | 124.5002 | 122.9167 | 114.8304 | 0.009378 | 0.90996 | 0.021303 | 23.4115 | 0.21641 | 0.15978 | 0.049949 | 3.17 | 3.17 | 0.002072
2019-03-21 12:41 | 1553186487.3819 | 380 | 0.19945 | 0.008084 | 0.000151 | 99.4919 | 97.2707 | 104.3742 | 21.2543 | 0.004249 | 0.003897 | 233.6025 | 0.27038 | 8.406 | 0.004158 | 0.005406 | 0.005334 | 5.4e-05 | 0.000994 | 0.28888 | 1.2923 | 2.5478 | 99.37 | 21.06 | 37.29 | 124.3443 | 122.1207 | 115.7354 | 0.009378 | 0.90784 | 0.021293 | 23.4541 | 0.20196 | 0.16085 | 0.050413 | 3.1979 | 3.1979 | 0.002067
2019-03-21 12:41 | 1553186488.3826 | 381 | 0.19924 | 0.008093 | 0.000149 | 99.4925 | 97.2369 | 104.3512 | 21.2541 | 0.004246 | 0.003875 | 233.4753 | 0.26891 | 8.4035 | 0.004089 | 0.005402 | 0.005382 | 5.2e-05 | 0.000994 | 0.26449 | 1.3208 | 2.548 | 99.37 | 21.06 | 37.289 | 124.2687 | 120.417 | 115.1981 | 0.009378 | 0.90909 | 0.021305 | 23.4353 | 0.20512 | 0.15998 | 0.051689 | 3.2805 | 3.2805 | 0.002004
2019-03-21 12:41 | 1553186489.3816 | 382 | 0.19911 | 0.008105 | 0.000147 | 99.6233 | 97.5717 | 104.3155 | 21.2535 | 0.0043 | 0.003815 | 233.317 | 0.26923 | 8.405 | 0.004096 | 0.005234 | 0.005122 | 5.4e-05 | 0.000995 | 0.27096 | 1.2937 | 2.548 | 99.37 | 21.06 | 37.287 | 125.6005 | 120.5927 | 113.7307 | 0.009377 | 0.91254 | 0.021342 | 23.387 | 0.18612 | 0.16016 | 0.052454 | 3.3348 | 3.3348 | 0.001974
2019-03-21 12:41 | 1553186490.3821 | 383 | 0.1989 | 0.008137 | 0.000146 | 99.5828 | 97.3871 | 104.3946 | 21.2541 | 0.004364 | 0.003902 | 233.1874 | 0.26891 | 8.4023 | 0.004076 | 0.005865 | 0.005851 | 5.7e-05 | 0.000995 | 0.29645 | 1.3159 | 2.5479 | 99.37 | 21.06 | 37.287 | 127.1639 | 120.1203 | 115.8562 | 0.009377 | 0.90756 | 0.021291 | 23.46 | 0.19942 | 0.15998 | 0.053638 | 3.4019 | 3.4019 | 0.001933
2019-03-21 12:42 | 1553186491.3827 | 384 | 0.19891 | 0.008172 | 0.000144 | 99.5637 | 97.3582 | 104.3308 | 21.2545 | 0.004295 | 0.003887 | 233.0573 | 0.26931 | 8.4043 | 0.003989 | 0.005261 | 0.005162 | 5.4e-05 | 0.000993 | 0.28423 | 1.3435 | 2.548 | 99.37 | 21.06 | 37.286 | 125.4678 | 117.993 | 115.496 | 0.009377 | 0.9084 | 0.021295 | 23.442 | 0.20037 | 0.16021 | 0.053546 | 3.3967 | 3.3967 | 0.001939
2019-03-21 12:42 | 1553186492.3818 | 385 | 0.19893 | 0.008222 | 0.000143 | 99.5712 | 97.1382 | 104.8314 | 21.2537 | 0.004279 | 0.003859 | 232.9539 | 0.26938 | 8.4004 | 0.004034 | 0.005425 | 0.005268 | 5.9e-05 | 0.000993 | 0.2469 | 1.3028 | 2.548 | 99.37 | 21.06 | 37.285 | 125.0883 | 119.0938 | 114.8125 | 0.009377 | 0.91 | 0.021364 | 23.4775 | 0.22127 | 0.16026 | 0.0534 | 3.3986 | 3.3986 | 0.001938
2019-03-21 12:42 | 1553186493.3818 | 386 | 0.1989 | 0.008272 | 0.000142 | 99.6659 | 97.5935 | 104.5204 | 21.2544 | 0.004332 | 0.003947 | 232.8335 | 0.26982 | 8.4011 | 0.004082 | 0.005823 | 0.005747 | 5.2e-05 | 0.000993 | 0.26609 | 1.2964 | 2.5481 | 99.37 | 21.06 | 37.285 | 126.3808 | 120.2674 | 116.9695 | 0.009377 | 0.90497 | 0.021274 | 23.5077 | 0.18795 | 0.16051 | 0.053503 | 3.3908 | 3.3908 | 0.001946
2019-03-21 12:42 | 1553186494.3812 | 387 | 0.19889 | 0.008326 | 0.000141 | 99.578 | 97.771 | 104.7575 | 21.2541 | 0.0043 | 0.00395 | 232.6684 | 0.26925 | 8.4062 | 0.004073 | 0.005719 | 0.005723 | 5.5e-05 | 0.000994 | 0.27831 | 1.3044 | 2.5479 | 99.37 | 21.06 | 37.285 | 125.6003 | 120.0426 | 117.0317 | 0.009377 | 0.90482 | 0.021296 | 23.5362 | 0.16381 | 0.16018 | 0.053529 | 3.396 | 3.396 | 0.001939
2019-03-21 12:42 | 1553186495.3823 | 388 | 0.19889 | 0.008372 | 0.000142 | 99.6942 | 97.7676 | 104.93 | 21.254 | 0.0043 | 0.003942 | 232.5233 | 0.26926 | 8.403 | 0.004062 | 0.005242 | 0.005139 | 5.8e-05 | 0.000992 | 0.23418 | 1.3253 | 2.5481 | 99.37 | 21.06 | 37.286 | 125.6073 | 119.7697 | 116.8305 | 0.009377 | 0.90529 | 0.021319 | 23.5495 | 0.17455 | 0.16018 | 0.053419 | 3.3927 | 3.3927 | 0.00194
2019-03-21 12:42 | 1553186496.3809 | 389 | 0.19887 | 0.008415 | 0.000144 | 99.7014 | 97.5465 | 104.6216 | 21.2537 | 0.004369 | 0.003964 | 232.3714 | 0.26958 | 8.4051 | 0.004101 | 0.004848 | 0.004853 | 5.8e-05 | 0.000989 | 0.28943 | 1.3629 | 2.548 | 99.37 | 21.06 | 37.287 | 127.2792 | 120.7182 | 117.364 | 0.009377 | 0.90405 | 0.021273 | 23.531 | 0.19544 | 0.16037 | 0.053503 | 3.3906 | 3.3906 | 0.001944
2019-03-21 12:42 | 1553186497.3815 | 390 | 0.19875 | 0.008451 | 0.000146 | 99.7136 | 97.6828 | 104.6023 | 21.2542 | 0.004368 | 0.003948 | 232.2304 | 0.26983 | 8.4011 | 0.00414 | 0.005724 | 0.005604 | 5.7e-05 | 0.000988 | 0.2844 | 1.3452 | 2.5481 | 99.37 | 21.06 | 37.288 | 127.2637 | 121.6823 | 116.9771 | 0.009378 | 0.90495 | 0.021282 | 23.5172 | 0.18405 | 0.16052 | 0.05415 | 3.4327 | 3.4327 | 0.001922
2019-03-21 12:42 | 1553186498.3825 | 391 | 0.19877 | 0.008487 | 0.000148 | 99.7741 | 97.6729 | 104.4608 | 21.2542 | 0.004385 | 0.00396 | 232.1059 | 0.26967 | 8.4008 | 0.004094 | 0.005522 | 0.00541 | 5.2e-05 | 0.000985 | 0.2995 | 1.3796 | 2.5482 | 99.37 | 21.06 | 37.289 | 127.6915 | 120.5425 | 117.2773 | 0.009378 | 0.90425 | 0.021259 | 23.5103 | 0.19038 | 0.16043 | 0.053991 | 3.419 | 3.419 | 0.001929
2019-03-21 12:42 | 1553186499.3821 | 392 | 0.19873 | 0.008519 | 0.000149 | 99.7923 | 97.5053 | 104.2157 | 21.254 | 0.004394 | 0.003934 | 231.9764 | 0.27009 | 8.4 | 0.004139 | 0.005205 | 0.005123 | 5.7e-05 | 0.000985 | 0.27616 | 1.2911 | 2.5481 | 99.37 | 21.06 | 37.288 | 127.9064 | 121.6495 | 116.653 | 0.009378 | 0.9057 | 0.021251 | 23.4639 | 0.20737 | 0.16068 | 0.054185 | 3.4299 | 3.4299 | 0.001925
2019-03-21 12:42 | 1553186500.3812 | 393 | 0.19875 | 0.008551 | 0.000149 | 99.8373 | 97.3759 | 104.1484 | 21.2542 | 0.004398 | 0.003964 | 231.8384 | 0.27011 | 8.4039 | 0.004171 | 0.005033 | 0.005021 | 5.6e-05 | 0.000987 | 0.31564 | 1.3013 | 2.5479 | 99.37 | 21.06 | 37.287 | 127.995 | 122.4289 | 117.3691 | 0.009377 | 0.90404 | 0.021225 | 23.4779 | 0.22328 | 0.16069 | 0.054033 | 3.416 | 3.416 | 0.001933
2019-03-21 12:42 | 1553186501.3813 | 394 | 0.19875 | 0.008583 | 0.000149 | 99.8965 | 97.6613 | 104.2326 | 21.2547 | 0.004369 | 0.003912 | 231.6813 | 0.26997 | 8.4018 | 0.004079 | 0.005456 | 0.005295 | 5.8e-05 | 0.000988 | 0.28727 | 1.3177 | 2.548 | 99.37 | 21.06 | 37.286 | 127.2887 | 120.175 | 116.1014 | 0.009377 | 0.90699 | 0.021268 | 23.4492 | 0.20241 | 0.1606 | 0.053994 | 3.4205 | 3.4205 | 0.00193
2019-03-21 12:42 | 1553186502.3814 | 395 | 0.19877 | 0.008615 | 0.00015 | 99.9234 | 97.8009 | 103.9299 | 21.2546 | 0.004376 | 0.003995 | 231.5344 | 0.26946 | 8.4024 | 0.004127 | 0.005512 | 0.005464 | 5.6e-05 | 0.000989 | 0.27759 | 1.3182 | 2.548 | 99.37 | 21.06 | 37.286 | 127.4608 | 121.3605 | 118.1396 | 0.009377 | 0.90226 | 0.021182 | 23.4764 | 0.19203 | 0.1603 | 0.053864 | 3.3984 | 3.3984 | 0.001939
2019-03-21 12:42 | 1553186503.3828 | 396 | 0.19876 | 0.008635 | 0.000152 | 99.9847 | 98.0612 | 104.4141 | 21.2542 | 0.004396 | 0.003978 | 231.3893 | 0.26911 | 8.4015 | 0.004212 | 0.005903 | 0.00591 | 5.4e-05 | 0.000991 | 0.28838 | 1.3369 | 2.5478 | 99.37 | 21.06 | 37.294 | 127.9577 | 123.4325 | 117.7068 | 0.009379 | 0.90326 | 0.021243 | 23.518 | 0.17376 | 0.16009 | 0.05387 | 3.4085 | 3.4085 | 0.00193
2019-03-21 12:42 | 1553186504.3818 | 397 | 0.1988 | 0.008662 | 0.000152 | 99.9955 | 98.0921 | 104.2953 | 21.2542 | 0.004465 | 0.004021 | 231.2657 | 0.26967 | 8.4037 | 0.004198 | 0.005819 | 0.005724 | 5.6e-05 | 0.00099 | 0.23612 | 1.2962 | 2.5479 | 99.37 | 21.06 | 37.302 | 129.6432 | 123.0876 | 118.7582 | 0.009381 | 0.90084 | 0.021202 | 23.5362 | 0.17189 | 0.16043 | 0.053642 | 3.3876 | 3.3876 | 0.001946
2019-03-21 12:42 | 1553186505.3815 | 398 | 0.19888 | 0.008681 | 0.000153 | 100.0116 | 97.977 | 104.5413 | 21.2543 | 0.004394 | 0.003961 | 231.1495 | 0.26967 | 8.3994 | 0.004178 | 0.005372 | 0.005284 | 5.8e-05 | 0.00099 | 0.31492 | 1.3191 | 2.548 | 99.37 | 21.06 | 37.311 | 127.8996 | 122.6092 | 117.313 | 0.009383 | 0.90417 | 0.021266 | 23.5204 | 0.18383 | 0.16043 | 0.053103 | 3.3637 | 3.3637 | 0.00196
2019-03-21 12:42 | 1553186506.3816 | 399 | 0.19907 | 0.008695 | 0.000153 | 100.106 | 97.7206 | 104.6476 | 21.2538 | 0.004431 | 0.004008 | 231.0242 | 0.26939 | 8.3993 | 0.004121 | 0.0053 | 0.005205 | 5.5e-05 | 0.000989 | 0.30541 | 1.3609 | 2.548 | 99.37 | 21.06 | 37.332 | 128.8053 | 121.2072 | 118.4446 | 0.009389 | 0.90156 | 0.021247 | 23.5665 | 0.21572 | 0.16026 | 0.051961 | 3.2886 | 3.2886 | 0.002003
2019-03-21 12:42 | 1553186507.3818 | 400 | 0.19909 | 0.008698 | 0.000153 | 100.0866 | 97.7849 | 104.404 | 21.2543 | 0.00441 | 0.003991 | 230.8984 | 0.26969 | 8.4011 | 0.004068 | 0.005813 | 0.005657 | 5.8e-05 | 0.000993 | 0.28479 | 1.33 | 2.548 | 99.37 | 21.06 | 37.342 | 128.3012 | 119.9215 | 118.0291 | 0.009391 | 0.90252 | 0.021233 | 23.5265 | 0.2081 | 0.16044 | 0.051863 | 3.2803 | 3.2803 | 0.00201
2019-03-21 12:42 | 1553186508.3816 | 401 | 0.19911 | 0.008684 | 0.000154 | 100.1179 | 98.134 | 104.477 | 21.2539 | 0.004446 | 0.004013 | 230.7537 | 0.26911 | 8.403 | 0.004104 | 0.005629 | 0.005597 | 5.3e-05 | 0.000992 | 0.30541 | 1.3149 | 2.548 | 99.37 | 21.06 | 37.352 | 129.1714 | 120.7883 | 118.5749 | 0.009394 | 0.90126 | 0.021226 | 23.5512 | 0.17902 | 0.16009 | 0.051751 | 3.272 | 3.272 | 0.002011
2019-03-21 12:42 | 1553186509.3807 | 402 | 0.19892 | 0.008667 | 0.000154 | 100.232 | 98.0714 | 104.3498 | 21.2543 | 0.004415 | 0.004041 | 230.6372 | 0.26977 | 8.4025 | 0.004048 | 0.005725 | 0.005579 | 5.3e-05 | 0.000991 | 0.24818 | 1.3215 | 2.548 | 99.37 | 21.05 | 37.371 | 128.4265 | 119.4297 | 119.2457 | 0.009393 | 0.89972 | 0.021195 | 23.557 | 0.19492 | 0.16049 | 0.052872 | 3.3378 | 3.3378 | 0.001976
2019-03-21 12:42 | 1553186510.3807 | 403 | 0.19886 | 0.008653 | 0.000153 | 100.2356 | 98.2267 | 103.7937 | 21.2539 | 0.004471 | 0.004014 | 230.4942 | 0.26843 | 8.4043 | 0.004222 | 0.006454 | 0.006385 | 5.2e-05 | 0.000992 | 0.27007 | 1.3374 | 2.548 | 99.37 | 21.05 | 37.38 | 129.8 | 123.6816 | 118.5954 | 0.009395 | 0.90121 | 0.021156 | 23.4746 | 0.18109 | 0.15969 | 0.053258 | 3.3559 | 3.3559 | 0.001956
2019-03-21 12:42 | 1553186511.3807 | 404 | 0.19867 | 0.008647 | 0.000152 | 100.2311 | 98.1933 | 103.6548 | 21.2541 | 0.004463 | 0.004111 | 230.3981 | 0.26942 | 8.404 | 0.004239 | 0.005497 | 0.005439 | 5.4e-05 | 0.000993 | 0.25122 | 1.3342 | 2.5479 | 99.37 | 21.05 | 37.397 | 129.6062 | 124.1018 | 120.9781 | 0.009399 | 0.89576 | 0.021077 | 23.5301 | 0.18373 | 0.16028 | 0.054417 | 3.4161 | 3.4161 | 0.001928
2019-03-21 12:42 | 1553186512.381 | 405 | 0.19852 | 0.008647 | 0.000151 | 100.3199 | 98.0465 | 103.7045 | 21.2545 | 0.004443 | 0.00404 | 230.2422 | 0.26951 | 8.4034 | 0.004108 | 0.005733 | 0.005677 | 5.2e-05 | 0.000988 | 0.25962 | 1.3232 | 2.5479 | 99.37 | 21.05 | 37.406 | 129.1059 | 120.8889 | 119.2381 | 0.009401 | 0.89973 | 0.021129 | 23.4838 | 0.20503 | 0.16033 | 0.055286 | 3.479 | 3.479 | 0.001894
2019-03-21 12:42 | 1553186513.3813 | 406 | 0.19845 | 0.008657 | 0.000149 | 100.3747 | 98.2471 | 103.3416 | 21.254 | 0.004497 | 0.004105 | 230.107 | 0.26951 | 8.4016 | 0.004272 | 0.005979 | 0.005888 | 5.2e-05 | 0.000993 | 0.27544 | 1.3594 | 2.5481 | 99.37 | 21.05 | 37.415 | 130.4311 | 124.9033 | 120.817 | 0.009404 | 0.89613 | 0.02105 | 23.4898 | 0.19163 | 0.16033 | 0.055702 | 3.492 | 3.492 | 0.001887
2019-03-21 12:42 | 1553186514.3814 | 407 | 0.19832 | 0.008685 | 0.000147 | 100.3643 | 98.4252 | 103.256 | 21.2533 | 0.004491 | 0.004026 | 229.9756 | 0.2698 | 8.4038 | 0.004221 | 0.005874 | 0.005787 | 5.6e-05 | 0.000986 | 0.28622 | 1.3173 | 2.5481 | 99.37 | 21.05 | 37.429 | 130.2931 | 123.654 | 118.8865 | 0.009407 | 0.90054 | 0.021093 | 23.4224 | 0.1745 | 0.1605 | 0.056452 | 3.5462 | 3.5462 | 0.00186
2019-03-21 12:42 | 1553186515.382 | 408 | 0.19815 | 0.008722 | 0.000146 | 100.3998 | 98.1061 | 103.0982 | 21.2537 | 0.004496 | 0.004051 | 229.8162 | 0.2699 | 8.3997 | 0.004237 | 0.005986 | 0.005901 | 5.4e-05 | 0.000985 | 0.28871 | 1.3036 | 2.5481 | 99.37 | 21.06 | 37.432 | 130.4006 | 124.0561 | 119.4974 | 0.009414 | 0.89914 | 0.02106 | 23.4228 | 0.20671 | 0.16057 | 0.057445 | 3.6028 | 3.6028 | 0.001832
2019-03-21 12:42 | 1553186516.3815 | 409 | 0.19806 | 0.008764 | 0.000146 | 100.4841 | 98.1344 | 103.3442 | 21.2542 | 0.004493 | 0.004098 | 229.6873 | 0.26944 | 8.3997 | 0.004277 | 0.005902 | 0.00579 | 5.5e-05 | 0.000985 | 0.26792 | 1.3149 | 2.5482 | 99.37 | 21.06 | 37.433 | 130.3356 | 125.0402 | 120.6408 | 0.009414 | 0.89653 | 0.021055 | 23.4848 | 0.21164 | 0.16029 | 0.057925 | 3.6318 | 3.6318 | 0.001814
2019-03-21 12:42 | 1553186517.3821 | 410 | 0.1979 | 0.008824 | 0.000147 | 100.4838 | 98.1817 | 103.5444 | 21.2546 | 0.004516 | 0.004101 | 229.5402 | 0.26936 | 8.4056 | 0.004334 | 0.005708 | 0.005659 | 5.7e-05 | 0.000987 | 0.2896 | 1.315 | 2.5479 | 99.37 | 21.06 | 37.431 | 130.9035 | 126.4254 | 120.731 | 0.009414 | 0.89633 | 0.021073 | 23.5102 | 0.2073 | 0.16024 | 0.058768 | 3.6876 | 3.6876 | 0.001786
2019-03-21 12:42 | 1553186518.3822 | 411 | 0.19789 | 0.008884 | 0.000148 | 100.5866 | 98.321 | 103.2914 | 21.254 | 0.004579 | 0.004097 | 229.3993 | 0.26984 | 8.4045 | 0.004303 | 0.005984 | 0.005881 | 5.7e-05 | 0.000987 | 0.28208 | 1.3296 | 2.548 | 99.37 | 21.06 | 37.428 | 132.4567 | 125.6782 | 120.6218 | 0.009413 | 0.89657 | 0.02105 | 23.4782 | 0.20377 | 0.16053 | 0.058786 | 3.6847 | 3.6847 | 0.001791
2019-03-21 12:42 | 1553186519.3817 | 412 | 0.19788 | 0.008954 | 0.00015 | 100.5928 | 98.0727 | 103.6001 | 21.2545 | 0.004552 | 0.004112 | 229.261 | 0.26928 | 8.3995 | 0.00434 | 0.005855 | 0.005834 | 5.2e-05 | 0.000984 | 0.27887 | 1.3788 | 2.5481 | 99.37 | 21.06 | 37.424 | 131.792 | 126.5735 | 120.9929 | 0.009412 | 0.89573 | 0.021071 | 23.5244 | 0.22694 | 0.16019 | 0.058741 | 3.6855 | 3.6855 | 0.001786
2019-03-21 12:42 | 1553186520.3817 | 413 | 0.19789 | 0.009021 | 0.000152 | 100.5982 | 97.9324 | 103.6141 | 21.2545 | 0.004553 | 0.004141 | 229.1332 | 0.26971 | 8.4032 | 0.004276 | 0.005428 | 0.005372 | 5.5e-05 | 0.000982 | 0.27295 | 1.3685 | 2.548 | 99.37 | 21.06 | 37.422 | 131.807 | 125.0079 | 121.6938 | 0.009411 | 0.89414 | 0.021054 | 23.5469 | 0.24022 | 0.16045 | 0.058615 | 3.6745 | 3.6745 | 0.001795
2019-03-21 12:42 | 1553186521.3815 | 414 | 0.19792 | 0.009084 | 0.000154 | 100.6858 | 98.0234 | 103.2423 | 21.2544 | 0.004549 | 0.004127 | 228.9799 | 0.26882 | 8.4019 | 0.004314 | 0.005272 | 0.0052 | 5.7e-05 | 0.000985 | 0.2959 | 1.2941 | 2.5481 | 99.37 | 21.06 | 37.42 | 131.7055 | 125.9445 | 121.3515 | 0.009411 | 0.89492 | 0.021026 | 23.4944 | 0.23971 | 0.15992 | 0.058379 | 3.6547 | 3.6547 | 0.001798
2019-03-21 12:42 | 1553186522.3804 | 415 | 0.1981 | 0.009141 | 0.000156 | 100.7788 | 98.3214 | 103.1712 | 21.2539 | 0.004573 | 0.004069 | 228.8658 | 0.2699 | 8.4026 | 0.004245 | 0.005744 | 0.005671 | 5.6e-05 | 0.000986 | 0.24707 | 1.3495 | 2.5481 | 99.37 | 21.06 | 37.416 | 132.3017 | 124.2572 | 119.9495 | 0.00941 | 0.89811 | 0.021056 | 23.4445 | 0.22081 | 0.16056 | 0.057194 | 3.5858 | 3.5858 | 0.00184
2019-03-21 12:42 | 1553186523.3805 | 416 | 0.19814 | 0.009183 | 0.000157 | 100.7481 | 98.365 | 103.4834 | 21.254 | 0.004498 | 0.00414 | 228.755 | 0.26904 | 8.4035 | 0.004328 | 0.006101 | 0.006098 | 5.6e-05 | 0.000987 | 0.25841 | 1.3306 | 2.5479 | 99.37 | 21.06 | 37.415 | 130.4511 | 126.2847 | 121.6726 | 0.00941 | 0.89419 | 0.021041 | 23.5314 | 0.21412 | 0.16005 | 0.05692 | 3.5662 | 3.5662 | 0.001845
2019-03-21 12:42 | 1553186524.3808 | 417 | 0.19814 | 0.009215 | 0.000157 | 100.7525 | 98.3295 | 103.2463 | 21.2541 | 0.004545 | 0.004183 | 228.6062 | 0.2693 | 8.406 | 0.004273 | 0.005871 | 0.005784 | 5.6e-05 | 0.000987 | 0.2844 | 1.32 | 2.5481 | 99.37 | 21.06 | 37.414 | 131.6254 | 124.9416 | 122.7172 | 0.009409 | 0.89183 | 0.02099 | 23.5355 | 0.21774 | 0.16021 | 0.056872 | 3.5545 | 3.5545 | 0.001852
2019-03-21 12:42 | 1553186525.3805 | 418 | 0.19814 | 0.00923 | 0.000157 | 100.8497 | 98.4549 | 103.2456 | 21.2541 | 0.004598 | 0.004141 | 228.5005 | 0.26833 | 8.4018 | 0.00429 | 0.005468 | 0.005401 | 5.7e-05 | 0.000988 | 0.2891 | 1.2965 | 2.548 | 99.37 | 21.06 | 37.412 | 132.9188 | 125.3418 | 121.7006 | 0.009409 | 0.89412 | 0.021017 | 23.5052 | 0.21496 | 0.15963 | 0.056902 | 3.5609 | 3.5609 | 0.001842
2019-03-21 12:42 | 1553186526.3808 | 419 | 0.19814 | 0.009234 | 0.000157 | 100.8735 | 98.5097 | 103.121 | 21.2545 | 0.004659 | 0.004173 | 228.347 | 0.26941 | 8.4029 | 0.004301 | 0.006187 | 0.006145 | 5.5e-05 | 0.000987 | 0.29358 | 1.2785 | 2.548 | 99.37 | 21.06 | 37.411 | 134.4173 | 125.6191 | 122.4919 | 0.009409 | 0.89234 | 0.020983 | 23.5145 | 0.21209 | 0.16027 | 0.05684 | 3.5513 | 3.5513 | 0.001855
2019-03-21 12:42 | 1553186527.381 | 420 | 0.19809 | 0.009232 | 0.000156 | 100.9432 | 98.3796 | 103.1884 | 21.2538 | 0.004644 | 0.004167 | 228.2069 | 0.26989 | 8.3982 | 0.004257 | 0.005775 | 0.005661 | 5.4e-05 | 0.000988 | 0.2756 | 1.273 | 2.5481 | 99.37 | 21.05 | 37.405 | 134.0567 | 124.5388 | 122.3347 | 0.009401 | 0.89269 | 0.020994 | 23.5175 | 0.2301 | 0.16056 | 0.057151 | 3.5727 | 3.5727 | 0.001847
2019-03-21 12:42 | 1553186528.382 | 421 | 0.19802 | 0.009232 | 0.000156 | 100.8851 | 98.3251 | 103.494 | 21.2543 | 0.004584 | 0.004171 | 228.0702 | 0.26967 | 8.3989 | 0.004197 | 0.006559 | 0.006451 | 5.8e-05 | 0.000988 | 0.26056 | 1.3042 | 2.548 | 99.37 | 21.05 | 37.402 | 132.5641 | 123.0789 | 122.4321 | 0.0094 | 0.89247 | 0.021022 | 23.5552 | 0.2299 | 0.16043 | 0.057601 | 3.6056 | 3.6056 | 0.001829
2019-03-21 12:42 | 1553186529.3807 | 422 | 0.19805 | 0.009243 | 0.000155 | 100.9072 | 98.2318 | 103.451 | 21.2547 | 0.004602 | 0.004161 | 227.9086 | 0.26941 | 8.4006 | 0.004163 | 0.00621 | 0.00618 | 5.6e-05 | 0.000986 | 0.21101 | 1.3171 | 2.5481 | 99.37 | 21.06 | 37.396 | 133.016 | 122.2308 | 122.1958 | 0.009405 | 0.893 | 0.021024 | 23.5433 | 0.24035 | 0.16027 | 0.057422 | 3.5948 | 3.5948 | 0.001832
2019-03-21 12:42 | 1553186530.3812 | 423 | 0.19805 | 0.009257 | 0.000156 | 100.9927 | 98.4716 | 103.6693 | 21.2545 | 0.004637 | 0.004138 | 227.7829 | 0.26932 | 8.4047 | 0.004218 | 0.005437 | 0.005376 | 5.6e-05 | 0.000985 | 0.27151 | 1.2832 | 2.548 | 99.37 | 21.06 | 37.394 | 133.8768 | 123.5957 | 121.6246 | 0.009404 | 0.8943 | 0.021062 | 23.5511 | 0.22612 | 0.16022 | 0.057355 | 3.5969 | 3.5969 | 0.001831
2019-03-21 12:42 | 1553186531.3819 | 424 | 0.19812 | 0.009266 | 0.000157 | 101.0353 | 98.6197 | 103.9714 | 21.2549 | 0.00462 | 0.004213 | 227.6533 | 0.26863 | 8.4012 | 0.00411 | 0.005895 | 0.005843 | 5.8e-05 | 0.000985 | 0.30021 | 1.3229 | 2.5481 | 99.37 | 21.06 | 37.39 | 133.4726 | 120.9304 | 123.4721 | 0.009403 | 0.89013 | 0.021043 | 23.6405 | 0.21644 | 0.15981 | 0.056941 | 3.5679 | 3.5679 | 0.001841
2019-03-21 12:42 | 1553186532.3809 | 425 | 0.19826 | 0.009275 | 0.000158 | 101.0207 | 98.5053 | 104.0737 | 21.2547 | 0.004614 | 0.004164 | 227.5313 | 0.26957 | 8.403 | 0.004204 | 0.005806 | 0.005717 | 5.5e-05 | 0.000987 | 0.27865 | 1.3209 | 2.5479 | 99.37 | 21.05 | 37.386 | 133.3253 | 123.2478 | 122.2666 | 0.009396 | 0.89284 | 0.021086 | 23.6162 | 0.22554 | 0.16037 | 0.056102 | 3.5225 | 3.5225 | 0.001871
2019-03-21 12:42 | 1553186533.3824 | 426 | 0.1983 | 0.009278 | 0.000158 | 101.0335 | 98.4552 | 103.7005 | 21.2543 | 0.004673 | 0.004199 | 227.398 | 0.26944 | 8.4006 | 0.004271 | 0.005466 | 0.005392 | 5.6e-05 | 0.000989 | 0.27815 | 1.3024 | 2.5479 | 99.37 | 21.05 | 37.386 | 134.7778 | 124.883 | 123.1154 | 0.009396 | 0.89093 | 0.021025 | 23.5991 | 0.23122 | 0.16029 | 0.055868 | 3.4979 | 3.4979 | 0.001883
2019-03-21 12:42 | 1553186534.3821 | 427 | 0.19844 | 0.009282 | 0.000158 | 101.0624 | 98.5656 | 104.1296 | 21.2548 | 0.00459 | 0.004197 | 227.2694 | 0.27005 | 8.4039 | 0.004286 | 0.00545 | 0.005348 | 5.8e-05 | 0.000987 | 0.30127 | 1.2954 | 2.5479 | 99.37 | 21.06 | 37.385 | 132.7252 | 125.2582 | 123.0738 | 0.009402 | 0.89102 | 0.02107 | 23.6466 | 0.22376 | 0.16065 | 0.055056 | 3.4545 | 3.4545 | 0.001911
2019-03-21 12:42 | 1553186535.3821 | 428 | 0.19845 | 0.009275 | 0.000158 | 101.1184 | 98.7348 | 104.1028 | 21.2546 | 0.004684 | 0.004171 | 227.1099 | 0.26865 | 8.4018 | 0.004215 | 0.006059 | 0.005942 | 5.7e-05 | 0.000988 | 0.26344 | 1.291 | 2.548 | 99.37 | 21.05 | 37.383 | 135.0308 | 123.5218 | 122.4388 | 0.009396 | 0.89245 | 0.021084 | 23.6246 | 0.21337 | 0.15982 | 0.054963 | 3.451 | 3.451 | 0.001903
2019-03-21 12:42 | 1553186536.3811 | 429 | 0.19848 | 0.009257 | 0.000158 | 101.1172 | 98.5169 | 103.939 | 21.2545 | 0.004666 | 0.004146 | 226.962 | 0.26955 | 8.3999 | 0.004254 | 0.005463 | 0.005401 | 5.6e-05 | 0.000987 | 0.27793 | 1.2948 | 2.5479 | 99.37 | 21.05 | 37.38 | 134.6069 | 124.4711 | 121.8248 | 0.009395 | 0.89384 | 0.021084 | 23.5877 | 0.23303 | 0.16036 | 0.054798 | 3.4406 | 3.4406 | 0.001916
2019-03-21 12:42 | 1553186537.3816 | 430 | 0.1985 | 0.009241 | 0.00016 | 101.2441 | 98.774 | 103.9845 | 21.2545 | 0.004576 | 0.00419 | 226.8201 | 0.2696 | 8.4047 | 0.004234 | 0.006084 | 0.006056 | 6.3e-05 | 0.00099 | 0.29015 | 1.324 | 2.548 | 99.37 | 21.05 | 37.379 | 132.371 | 123.9804 | 122.9106 | 0.009395 | 0.89139 | 0.021059 | 23.6253 | 0.22093 | 0.16039 | 0.054741 | 3.433 | 3.433 | 0.00192
2019-03-21 12:42 | 1553186538.3812 | 431 | 0.19853 | 0.009216 | 0.00016 | 101.2087 | 98.7897 | 103.8956 | 21.2546 | 0.004629 | 0.004158 | 226.6828 | 0.26901 | 8.4037 | 0.004164 | 0.005915 | 0.005872 | 5.5e-05 | 0.000987 | 0.26488 | 1.2917 | 2.548 | 99.37 | 21.05 | 37.377 | 133.6849 | 122.2672 | 122.1274 | 0.009394 | 0.89316 | 0.021071 | 23.5918 | 0.21638 | 0.16004 | 0.054559 | 3.4235 | 3.4235 | 0.001921
2019-03-21 12:42 | 1553186539.3813 | 432 | 0.19871 | 0.009195 | 0.000161 | 101.1509 | 99.0046 | 103.8594 | 21.2545 | 0.004673 | 0.004128 | 226.583 | 0.27004 | 8.4018 | 0.004257 | 0.005876 | 0.005714 | 5.5e-05 | 0.000989 | 0.28119 | 1.3151 | 2.5481 | 99.37 | 21.06 | 37.375 | 134.7644 | 124.5338 | 121.3786 | 0.009399 | 0.89485 | 0.021088 | 23.5653 | 0.19184 | 0.16065 | 0.053546 | 3.3627 | 3.3627 | 0.001963
2019-03-21 12:42 | 1553186540.3803 | 433 | 0.19905 | 0.009171 | 0.000162 | 101.237 | 99.0615 | 104.0479 | 21.2549 | 0.00462 | 0.004182 | 226.4682 | 0.26908 | 8.4011 | 0.004197 | 0.005865 | 0.005776 | 5.8e-05 | 0.000988 | 0.28584 | 1.2982 | 2.548 | 99.36 | 21.05 | 37.375 | 133.4613 | 123.0827 | 122.6997 | 0.009395 | 0.89187 | 0.021071 | 23.6262 | 0.19431 | 0.16007 | 0.051521 | 3.2334 | 3.2334 | 0.002035
2019-03-21 12:42 | 1553186541.3818 | 434 | 0.19909 | 0.00914 | 0.000164 | 101.2394 | 98.9443 | 104.0305 | 21.2539 | 0.004579 | 0.004194 | 226.3459 | 0.26821 | 8.404 | 0.0042 | 0.006327 | 0.00625 | 5.6e-05 | 0.000987 | 0.24906 | 1.3182 | 2.548 | 99.36 | 21.05 | 37.373 | 132.464 | 123.1532 | 123.0052 | 0.009394 | 0.89118 | 0.021062 | 23.6333 | 0.20511 | 0.15956 | 0.051286 | 3.2171 | 3.2171 | 0.002038
2019-03-21 12:42 | 1553186542.3805 | 435 | 0.19918 | 0.009097 | 0.000166 | 101.2686 | 98.9681 | 103.9108 | 21.2547 | 0.004582 | 0.004202 | 226.2056 | 0.26971 | 8.4005 | 0.004222 | 0.005299 | 0.005313 | 5.7e-05 | 0.00099 | 0.28279 | 1.3473 | 2.548 | 99.36 | 21.05 | 37.371 | 132.5207 | 123.6763 | 123.2008 | 0.009394 | 0.89074 | 0.021044 | 23.6255 | 0.20553 | 0.16045 | 0.050846 | 3.1868 | 3.1868 | 0.002069
2019-03-21 12:42 | 1553186543.3811 | 436 | 0.19923 | 0.009045 | 0.000168 | 101.2847 | 99.1303 | 103.9315 | 21.2556 | 0.004656 | 0.004146 | 226.0876 | 0.26967 | 8.4027 | 0.004162 | 0.005603 | 0.005492 | 5.7e-05 | 0.000989 | 0.25321 | 1.2981 | 2.548 | 99.36 | 21.05 | 37.371 | 134.3408 | 122.2246 | 121.8139 | 0.009394 | 0.89387 | 0.021083 | 23.5865 | 0.19231 | 0.16043 | 0.050581 | 3.176 | 3.176 | 0.002076
2019-03-21 12:42 | 1553186544.3809 | 437 | 0.19927 | 0.008993 | 0.000169 | 101.2155 | 99.1953 | 103.9662 | 21.2551 | 0.004658 | 0.004135 | 225.9527 | 0.26939 | 8.4015 | 0.004113 | 0.005125 | 0.005003 | 5.7e-05 | 0.00099 | 0.29679 | 1.3028 | 2.5481 | 99.36 | 21.05 | 37.368 | 134.4019 | 121.0172 | 121.5452 | 0.009393 | 0.89448 | 0.021094 | 23.5824 | 0.18034 | 0.16026 | 0.050425 | 3.1678 | 3.1678 | 0.002079
2019-03-21 12:42 | 1553186545.3806 | 438 | 0.19926 | 0.008929 | 0.00017 | 101.3046 | 99.1027 | 103.7794 | 21.255 | 0.004683 | 0.004192 | 225.8402 | 0.26883 | 8.4038 | 0.004164 | 0.006058 | 0.005952 | 5.8e-05 | 0.00099 | 0.24635 | 1.3363 | 2.5482 | 99.36 | 21.05 | 37.367 | 135.0265 | 122.2651 | 122.9603 | 0.009393 | 0.89128 | 0.021037 | 23.6034 | 0.19656 | 0.15993 | 0.050535 | 3.1661 | 3.1661 | 0.002076
2019-03-21 12:42 | 1553186546.3807 | 439 | 0.19921 | 0.008865 | 0.000171 | 101.2373 | 98.9079 | 104.0512 | 21.2549 | 0.004627 | 0.004163 | 225.7391 | 0.26883 | 8.4013 | 0.004084 | 0.005891 | 0.005848 | 5.7e-05 | 0.000993 | 0.26017 | 1.3446 | 2.5481 | 99.36 | 21.06 | 37.366 | 133.6365 | 120.3053 | 122.2382 | 0.009398 | 0.89291 | 0.021084 | 23.6128 | 0.20822 | 0.15993 | 0.050924 | 3.1975 | 3.1975 | 0.002056
2019-03-21 12:42 | 1553186547.3816 | 440 | 0.19925 | 0.008814 | 0.000171 | 101.2096 | 98.5615 | 104.0848 | 21.2545 | 0.004665 | 0.004182 | 225.6345 | 0.26895 | 8.4061 | 0.00418 | 0.005216 | 0.005176 | 6e-05 | 0.000992 | 0.29463 | 1.3595 | 2.5479 | 99.36 | 21.06 | 37.366 | 134.5677 | 122.6628 | 122.6993 | 0.009398 | 0.89187 | 0.021075 | 23.6303 | 0.23715 | 0.16 | 0.050746 | 3.185 | 3.185 | 0.002065
2019-03-21 12:42 | 1553186548.3818 | 441 | 0.19917 | 0.008755 | 0.00017 | 101.2576 | 98.742 | 104.1 | 21.2553 | 0.004681 | 0.004192 | 225.5227 | 0.26859 | 8.4032 | 0.004178 | 0.005598 | 0.005466 | 6e-05 | 0.000991 | 0.28279 | 1.2907 | 2.548 | 99.36 | 21.06 | 37.366 | 134.9612 | 122.603 | 122.9471 | 0.009398 | 0.89131 | 0.02107 | 23.6395 | 0.22502 | 0.15978 | 0.051298 | 3.2188 | 3.2188 | 0.00204
2019-03-21 12:42 | 1553186549.3807 | 442 | 0.19916 | 0.008719 | 0.00017 | 101.2945 | 98.7996 | 103.9939 | 21.2551 | 0.004647 | 0.004168 | 225.3824 | 0.26895 | 8.4002 | 0.004126 | 0.005913 | 0.005876 | 6e-05 | 0.000989 | 0.2907 | 1.3099 | 2.5481 | 99.36 | 21.06 | 37.365 | 134.1315 | 121.3252 | 122.3607 | 0.009398 | 0.89263 | 0.021075 | 23.6099 | 0.22307 | 0.16 | 0.051374 | 3.2243 | 3.2243 | 0.00204
2019-03-21 12:42 | 1553186550.3818 | 443 | 0.1992 | 0.008689 | 0.000171 | 101.2629 | 99.0482 | 104.0545 | 21.2546 | 0.004617 | 0.004131 | 225.2568 | 0.26903 | 8.4041 | 0.004184 | 0.005687 | 0.005568 | 5.6e-05 | 0.000987 | 0.2417 | 1.2878 | 2.5479 | 99.36 | 21.06 | 37.364 | 133.3772 | 122.7471 | 121.4453 | 0.009398 | 0.8947 | 0.021106 | 23.5894 | 0.1978 | 0.16005 | 0.051222 | 3.2194 | 3.2194 | 0.002043
2019-03-21 12:42 | 1553186551.3812 | 444 | 0.19921 | 0.008663 | 0.000172 | 101.2754 | 99.088 | 104.3296 | 21.2555 | 0.004648 | 0.004139 | 225.1266 | 0.26884 | 8.4046 | 0.004105 | 0.005873 | 0.005827 | 5.9e-05 | 0.000983 | 0.26178 | 1.2604 | 2.5479 | 99.36 | 21.06 | 37.363 | 134.1433 | 120.8249 | 121.6631 | 0.009397 | 0.89421 | 0.021128 | 23.6271 | 0.19531 | 0.15993 | 0.05114 | 3.2176 | 3.2176 | 0.002043
2019-03-21 12:43 | 1553186552.382 | 445 | 0.19921 | 0.008647 | 0.000173 | 101.2671 | 98.9031 | 104.4514 | 21.2552 | 0.004674 | 0.004173 | 225.0078 | 0.26897 | 8.4032 | 0.004144 | 0.005575 | 0.005564 | 5.8e-05 | 0.00098 | 0.27417 | 1.327 | 2.5479 | 99.36 | 21.06 | 37.361 | 134.7915 | 121.7841 | 122.4886 | 0.009397 | 0.89234 | 0.021118 | 23.6656 | 0.21128 | 0.16001 | 0.051211 | 3.2205 | 3.2205 | 0.002042
2019-03-21 12:43 | 1553186553.381 | 446 | 0.19928 | 0.008628 | 0.000175 | 101.282 | 98.7219 | 104.7105 | 21.2549 | 0.004662 | 0.004197 | 224.885 | 0.26919 | 8.4034 | 0.004065 | 0.005013 | 0.004955 | 5.9e-05 | 0.000978 | 0.27848 | 1.3243 | 2.548 | 99.36 | 21.06 | 37.359 | 134.4968 | 119.8466 | 123.0819 | 0.009396 | 0.89101 | 0.021128 | 23.7127 | 0.229 | 0.16014 | 0.050781 | 3.195 | 3.195 | 0.00206
2019-03-21 12:43 | 1553186554.3802 | 447 | 0.1994 | 0.008606 | 0.000174 | 101.1944 | 98.8159 | 104.7655 | 21.2559 | 0.004735 | 0.004175 | 224.7761 | 0.26942 | 8.4041 | 0.004172 | 0.005734 | 0.00564 | 5.9e-05 | 0.000979 | 0.25155 | 1.3195 | 2.548 | 99.36 | 21.06 | 37.355 | 136.3038 | 122.456 | 122.5361 | 0.009395 | 0.89224 | 0.021148 | 23.7026 | 0.21274 | 0.16028 | 0.050071 | 3.1534 | 3.1534 | 0.002089
2019-03-21 12:43 | 1553186555.3812 | 448 | 0.1995 | 0.008584 | 0.000174 | 101.2221 | 98.8581 | 104.8321 | 21.2556 | 0.004707 | 0.004171 | 224.6758 | 0.26962 | 8.4015 | 0.004179 | 0.005408 | 0.005294 | 5.9e-05 | 0.00098 | 0.26465 | 1.3304 | 2.548 | 99.36 | 21.06 | 37.353 | 135.62 | 122.6345 | 122.4278 | 0.009395 | 0.89248 | 0.021158 | 23.7069 | 0.21137 | 0.1604 | 0.0495 | 3.119 | 3.119 | 0.002114
2019-03-21 12:43 | 1553186556.3812 | 449 | 0.19958 | 0.008552 | 0.000174 | 101.2868 | 98.4774 | 104.7889 | 21.2563 | 0.004616 | 0.004147 | 224.5446 | 0.26905 | 8.4026 | 0.004136 | 0.005515 | 0.005478 | 5.6e-05 | 0.00098 | 0.30989 | 1.3469 | 2.5481 | 99.36 | 21.06 | 37.352 | 133.3758 | 121.5836 | 121.8521 | 0.009395 | 0.89378 | 0.021169 | 23.6847 | 0.2516 | 0.16006 | 0.04909 | 3.0949 | 3.0949 | 0.002126
2019-03-21 12:43 | 1553186557.3802 | 450 | 0.19964 | 0.008514 | 0.000175 | 101.2987 | 98.8309 | 104.8847 | 21.2556 | 0.004627 | 0.004201 | 224.4209 | 0.26858 | 8.4029 | 0.0041 | 0.005625 | 0.005547 | 5.7e-05 | 0.000982 | 0.25354 | 1.3348 | 2.548 | 99.36 | 21.06 | 37.345 | 133.6322 | 120.6994 | 123.1747 | 0.009393 | 0.8908 | 0.021143 | 23.7352 | 0.2206 | 0.15978 | 0.048767 | 3.0707 | 3.0707 | 0.002139
2019-03-21 12:43 | 1553186558.3807 | 451 | 0.19958 | 0.008474 | 0.000175 | 101.2278 | 98.9968 | 104.8502 | 21.2554 | 0.004654 | 0.004144 | 224.3066 | 0.26845 | 8.4033 | 0.004095 | 0.005191 | 0.005155 | 5.7e-05 | 0.000983 | 0.26233 | 1.3279 | 2.548 | 99.36 | 21.06 | 37.342 | 134.3085 | 120.5692 | 121.7702 | 0.009392 | 0.89397 | 0.021177 | 23.6892 | 0.19934 | 0.1597 | 0.049207 | 3.1034 | 3.1034 | 0.002115
2019-03-21 12:43 | 1553186559.3804 | 452 | 0.19963 | 0.008431 | 0.000175 | 101.2051 | 98.932 | 104.9342 | 21.2557 | 0.004615 | 0.004169 | 224.1965 | 0.26905 | 8.4055 | 0.004158 | 0.005733 | 0.00554 | 5.6e-05 | 0.000983 | 0.29407 | 1.3553 | 2.5481 | 99.36 | 21.06 | 37.337 | 133.3334 | 122.127 | 122.3978 | 0.009391 | 0.89255 | 0.021169 | 23.7175 | 0.20319 | 0.16006 | 0.04893 | 3.0848 | 3.0848 | 0.002133
2019-03-21 12:43 | 1553186560.3808 | 453 | 0.19941 | 0.008388 | 0.000174 | 101.2126 | 98.6994 | 104.7704 | 21.2555 | 0.004659 | 0.004189 | 224.0677 | 0.26939 | 8.4005 | 0.004135 | 0.00529 | 0.005277 | 5.3e-05 | 0.00098 | 0.26161 | 1.3488 | 2.5482 | 99.36 | 21.06 | 37.335 | 134.4198 | 121.559 | 122.8687 | 0.00939 | 0.89149 | 0.02114 | 23.7131 | 0.2249 | 0.16026 | 0.050295 | 3.1663 | 3.1663 | 0.00208
2019-03-21 12:43 | 1553186561.3805 | 454 | 0.19948 | 0.00836 | 0.000173 | 101.2114 | 98.665 | 104.6307 | 21.2562 | 0.004702 | 0.004154 | 223.9468 | 0.26929 | 8.4035 | 0.004132 | 0.005341 | 0.005244 | 6e-05 | 0.000977 | 0.2855 | 1.3164 | 2.5479 | 99.36 | 21.06 | 37.333 | 135.4844 | 121.4914 | 122.0235 | 0.00939 | 0.89339 | 0.021148 | 23.672 | 0.22791 | 0.1602 | 0.04995 | 3.1459 | 3.1459 | 0.002093
2019-03-21 12:43 | 1553186562.3815 | 455 | 0.19942 | 0.008338 | 0.000173 | 101.1983 | 98.7413 | 104.749 | 21.256 | 0.004681 | 0.004206 | 223.8285 | 0.26856 | 8.4033 | 0.004108 | 0.006116 | 0.006047 | 5.8e-05 | 0.000982 | 0.26322 | 1.2943 | 2.5479 | 99.36 | 21.06 | 37.33 | 134.9769 | 120.8926 | 123.2902 | 0.009389 | 0.89054 | 0.021127 | 23.7233 | 0.21984 | 0.15977 | 0.050319 | 3.1658 | 3.1658 | 0.002074
2019-03-21 12:43 | 1553186563.3819 | 456 | 0.19947 | 0.008334 | 0.000173 | 101.1271 | 98.585 | 104.6352 | 21.2559 | 0.004659 | 0.004169 | 223.6832 | 0.26941 | 8.4035 | 0.004136 | 0.004945 | 0.004871 | 5.8e-05 | 0.000981 | 0.28047 | 1.2933 | 2.548 | 99.36 | 21.06 | 37.329 | 134.4283 | 121.5768 | 122.3917 | 0.009389 | 0.89256 | 0.021139 | 23.6835 | 0.22772 | 0.16027 | 0.050002 | 3.1478 | 3.1478 | 0.002093
2019-03-21 12:43 | 1553186564.3817 | 457 | 0.19952 | 0.008329 | 0.000175 | 101.1837 | 98.8615 | 104.6824 | 21.2552 | 0.004659 | 0.004156 | 223.5756 | 0.26916 | 8.4023 | 0.004131 | 0.005726 | 0.00567 | 5.1e-05 | 0.000977 | 0.29048 | 1.3135 | 2.548 | 99.36 | 21.06 | 37.328 | 134.4291 | 121.4523 | 122.0644 | 0.009389 | 0.8933 | 0.021153 | 23.6791 | 0.20767 | 0.16012 | 0.049714 | 3.1316 | 3.1316 | 0.002101
2019-03-21 12:43 | 1553186565.3807 | 458 | 0.19955 | 0.008324 | 0.000177 | 101.1065 | 98.8139 | 104.8364 | 21.2559 | 0.004702 | 0.004138 | 223.4309 | 0.26969 | 8.4006 | 0.004149 | 0.005515 | 0.005463 | 5.8e-05 | 0.000979 | 0.30613 | 1.2914 | 2.548 | 99.36 | 21.06 | 37.328 | 135.4755 | 121.8896 | 121.6214 | 0.009389 | 0.8943 | 0.02118 | 23.6832 | 0.20516 | 0.16044 | 0.049545 | 3.125 | 3.125 | 0.00211
2019-03-21 12:43 | 1553186566.3817 | 459 | 0.19958 | 0.008323 | 0.000179 | 101.1163 | 98.2488 | 104.9612 | 21.2552 | 0.004654 | 0.004141 | 223.3245 | 0.26923 | 8.4042 | 0.004167 | 0.005454 | 0.005385 | 5.6e-05 | 0.000982 | 0.28031 | 1.295 | 2.548 | 99.36 | 21.07 | 37.327 | 134.2987 | 122.3481 | 121.712 | 0.009394 | 0.8941 | 0.02119 | 23.7 | 0.25732 | 0.16017 | 0.049338 | 3.1133 | 3.1133 | 0.002114
2019-03-21 12:43 | 1553186567.3811 | 460 | 0.19953 | 0.008309 | 0.000182 | 101.1431 | 98.4961 | 105.0098 | 21.2548 | 0.004647 | 0.004147 | 223.2106 | 0.26876 | 8.4018 | 0.00413 | 0.00563 | 0.005607 | 5.8e-05 | 0.00098 | 0.2896 | 1.3082 | 2.548 | 99.36 | 21.06 | 37.325 | 134.1258 | 121.4399 | 121.8489 | 0.009388 | 0.89379 | 0.021191 | 23.7096 | 0.23721 | 0.15988 | 0.049675 | 3.1346 | 3.1346 | 0.002096
2019-03-21 12:43 | 1553186568.3809 | 461 | 0.19994 | 0.008293 | 0.000183 | 101.0326 | 98.2127 | 105.3833 | 21.2556 | 0.004593 | 0.00415 | 223.1146 | 0.26887 | 8.4004 | 0.004135 | 0.005223 | 0.00513 | 5.7e-05 | 0.000981 | 0.27223 | 1.2919 | 2.548 | 99.36 | 21.07 | 37.325 | 132.8018 | 121.567 | 121.926 | 0.009394 | 0.89361 | 0.021227 | 23.754 | 0.2532 | 0.15995 | 0.047231 | 2.9856 | 2.9856 | 0.002202
2019-03-21 12:43 | 1553186569.3809 | 462 | 0.20005 | 0.008265 | 0.000181 | 101.0424 | 97.9396 | 105.5065 | 21.2556 | 0.004633 | 0.004161 | 223.0042 | 0.26845 | 8.4025 | 0.004135 | 0.005517 | 0.005361 | 5.5e-05 | 0.000975 | 0.25194 | 1.2924 | 2.5481 | 99.36 | 21.07 | 37.324 | 133.7804 | 121.556 | 122.2007 | 0.009393 | 0.89299 | 0.021232 | 23.7762 | 0.27898 | 0.1597 | 0.046632 | 2.9486 | 2.9486 | 0.002226
2019-03-21 12:43 | 1553186570.3803 | 463 | 0.20028 | 0.008233 | 0.00018 | 101.0198 | 97.9883 | 105.86 | 21.2554 | 0.004687 | 0.004135 | 222.8884 | 0.26902 | 8.4018 | 0.00408 | 0.005396 | 0.005351 | 5.7e-05 | 0.000981 | 0.29645 | 1.3096 | 2.548 | 99.36 | 21.07 | 37.323 | 135.1132 | 120.2057 | 121.5557 | 0.009393 | 0.89445 | 0.021285 | 23.7965 | 0.27253 | 0.16004 | 0.045296 | 2.8714 | 2.8714 | 0.002291
2019-03-21 12:43 | 1553186571.38 | 464 | 0.20029 | 0.008188 | 0.000178 | 100.9644 | 98.0053 | 106.2449 | 21.2557 | 0.004625 | 0.004101 | 222.8139 | 0.26993 | 8.4019 | 0.004167 | 0.005491 | 0.005361 | 5.8e-05 | 0.000982 | 0.30309 | 1.267 | 2.548 | 99.36 | 21.07 | 37.322 | 133.5917 | 122.3406 | 120.7314 | 0.009393 | 0.89632 | 0.021346 | 23.8149 | 0.26607 | 0.16058 | 0.0453 | 2.88 | 2.88 | 0.002292
2019-03-21 12:43 | 1553186572.38 | 465 | 0.20054 | 0.008128 | 0.000177 | 100.9432 | 98.1133 | 106.4653 | 21.2553 | 0.004595 | 0.004109 | 222.7125 | 0.26932 | 8.4042 | 0.004084 | 0.00558 | 0.005569 | 5.7e-05 | 0.00098 | 0.26128 | 1.3329 | 2.548 | 99.36 | 21.07 | 37.32 | 132.8407 | 120.3152 | 120.9071 | 0.009392 | 0.89592 | 0.021363 | 23.8449 | 0.25434 | 0.16022 | 0.043891 | 2.7928 | 2.7928 | 0.002358
2019-03-21 12:43 | 1553186573.3815 | 466 | 0.20072 | 0.008067 | 0.000176 | 100.8813 | 98.3418 | 106.6463 | 21.2556 | 0.004599 | 0.004105 | 222.6422 | 0.26843 | 8.4038 | 0.004131 | 0.0053 | 0.005225 | 5.7e-05 | 0.000984 | 0.27793 | 1.3807 | 2.5482 | 99.36 | 21.07 | 37.319 | 132.9506 | 121.452 | 120.8098 | 0.009392 | 0.89615 | 0.021384 | 23.8622 | 0.22804 | 0.15969 | 0.042888 | 2.7318 | 2.7318 | 0.002402
2019-03-21 12:43 | 1553186574.3811 | 467 | 0.20072 | 0.007991 | 0.000177 | 100.8488 | 98.1167 | 106.7242 | 21.2545 | 0.004606 | 0.004056 | 222.5253 | 0.26911 | 8.405 | 0.004159 | 0.005672 | 0.005582 | 5.7e-05 | 0.000983 | 0.24674 | 1.3299 | 2.5479 | 99.36 | 21.07 | 37.317 | 133.1165 | 122.133 | 119.6211 | 0.009392 | 0.89886 | 0.021424 | 23.8349 | 0.24566 | 0.16009 | 0.042971 | 2.7422 | 2.7422 | 0.002399
2019-03-21 12:43 | 1553186575.3805 | 468 | 0.20065 | 0.007921 | 0.000177 | 100.8342 | 98.1947 | 107.2092 | 21.2554 | 0.004564 | 0.00404 | 222.4444 | 0.26856 | 8.4023 | 0.004135 | 0.005851 | 0.005796 | 5.7e-05 | 0.00098 | 0.28639 | 1.332 | 2.548 | 99.36 | 21.06 | 37.315 | 132.0748 | 121.566 | 119.2351 | 0.009385 | 0.89974 | 0.021483 | 23.8772 | 0.23726 | 0.15976 | 0.043466 | 2.7814 | 2.7814 | 0.002361
2019-03-21 12:43 | 1553186576.3802 | 469 | 0.20129 | 0.007843 | 0.000179 | 100.7925 | 97.9382 | 107.5575 | 21.2562 | 0.004575 | 0.004035 | 222.3554 | 0.26881 | 8.4019 | 0.004194 | 0.005078 | 0.005004 | 5.7e-05 | 0.000982 | 0.28816 | 1.2857 | 2.548 | 99.36 | 21.06 | 37.314 | 132.3671 | 122.9997 | 119.1112 | 0.009385 | 0.90003 | 0.021522 | 23.9122 | 0.25695 | 0.15992 | 0.039707 | 2.5455 | 2.5455 | 0.002582
2019-03-21 12:43 | 1553186577.3813 | 470 | 0.20102 | 0.007761 | 0.000179 | 100.7254 | 97.7076 | 107.8213 | 21.2562 | 0.004532 | 0.004048 | 222.2649 | 0.26925 | 8.4029 | 0.004166 | 0.005587 | 0.005546 | 6e-05 | 0.000981 | 0.2557 | 1.3055 | 2.548 | 99.36 | 21.06 | 37.311 | 131.2858 | 122.3261 | 119.4356 | 0.009384 | 0.89928 | 0.021539 | 23.9514 | 0.27208 | 0.16017 | 0.04144 | 2.6587 | 2.6587 | 0.002476
2019-03-21 12:43 | 1553186578.3807 | 471 | 0.20114 | 0.007679 | 0.00018 | 100.7403 | 97.4209 | 107.765 | 21.2561 | 0.004535 | 0.004055 | 222.1619 | 0.26845 | 8.4017 | 0.004162 | 0.00568 | 0.005598 | 5.8e-05 | 0.000976 | 0.27865 | 1.3083 | 2.5479 | 99.36 | 21.06 | 37.311 | 131.3695 | 122.2099 | 119.5979 | 0.009384 | 0.89891 | 0.021529 | 23.9501 | 0.29969 | 0.1597 | 0.040833 | 2.6185 | 2.6185 | 0.002507
2019-03-21 12:43 | 1553186579.3818 | 472 | 0.20123 | 0.007595 | 0.000182 | 100.6697 | 97.1337 | 108.0213 | 21.2566 | 0.004541 | 0.004009 | 222.0599 | 0.26976 | 8.4016 | 0.004109 | 0.005825 | 0.00574 | 6.1e-05 | 0.000974 | 0.26217 | 1.2911 | 2.548 | 99.36 | 21.06 | 37.31 | 131.5113 | 120.9115 | 118.4759 | 0.009384 | 0.90149 | 0.021585 | 23.9443 | 0.31982 | 0.16048 | 0.040383 | 2.5963 | 2.5963 | 0.00254
2019-03-21 12:43 | 1553186580.3805 | 473 | 0.20175 | 0.007509 | 0.000184 | 100.6241 | 97.3701 | 108.234 | 21.2559 | 0.004471 | 0.004 | 222 | 0.26908 | 8.4037 | 0.004101 | 0.005688 | 0.005641 | 5.9e-05 | 0.000974 | 0.25658 | 1.3236 | 2.548 | 99.36 | 21.06 | 37.308 | 129.7963 | 120.7164 | 118.2652 | 0.009384 | 0.90197 | 0.021612 | 23.9614 | 0.29403 | 0.16007 | 0.037355 | 2.4045 | 2.4045 | 0.002736
2019-03-21 12:43 | 1553186581.3816 | 474 | 0.2016 | 0.007426 | 0.000186 | 100.58 | 97.3531 | 108.5699 | 21.2565 | 0.004521 | 0.003986 | 221.8975 | 0.26864 | 8.401 | 0.004086 | 0.00598 | 0.005895 | 5.9e-05 | 0.000975 | 0.30055 | 1.3026 | 2.5481 | 99.36 | 21.06 | 37.307 | 131.0314 | 120.3585 | 117.9093 | 0.009383 | 0.90279 | 0.021656 | 23.9876 | 0.29168 | 0.15981 | 0.038379 | 2.4754 | 2.4754 | 0.002653
2019-03-21 12:43 | 1553186582.3818 | 475 | 0.20175 | 0.007336 | 0.000186 | 100.5055 | 97.225 | 108.7314 | 21.256 | 0.004451 | 0.004024 | 221.819 | 0.26997 | 8.4015 | 0.004083 | 0.005594 | 0.005505 | 6.1e-05 | 0.00097 | 0.29103 | 1.2833 | 2.5478 | 99.36 | 21.06 | 37.306 | 129.3146 | 120.2817 | 118.8518 | 0.009383 | 0.90062 | 0.021646 | 24.0344 | 0.29682 | 0.16061 | 0.037553 | 2.4209 | 2.4209 | 0.002727
2019-03-21 12:43 | 1553186583.4327 | 476 | 0.2019 | 0.007242 | 0.000187 | 100.4677 | 97.0966 | 109.0264 | 21.2562 | 0.004527 | 0.003954 | 221.7222 | 0.26824 | 8.4032 | 0.004017 | 0.005105 | 0.005038 | 6.4e-05 | 0.000969 | 0.28479 | 1.2858 | 2.548 | 99.36 | 21.07 | 37.305 | 131.1818 | 118.68 | 117.1208 | 0.009389 | 0.90462 | 0.021723 | 24.0138 | 0.30528 | 0.15957 | 0.036783 | 2.3796 | 2.3796 | 0.002756
2019-03-21 12:43 | 1553186584.3856 | 477 | 0.20229 | 0.007145 | 0.000188 | 100.4978 | 96.8105 | 109.0996 | 21.2567 | 0.004486 | 0.003913 | 221.6432 | 0.2688 | 8.4004 | 0.004016 | 0.005351 | 0.005281 | 6.3e-05 | 0.00096 | 0.28943 | 1.2859 | 2.5479 | 99.36 | 21.07 | 37.306 | 130.1761 | 118.6473 | 116.1352 | 0.009389 | 0.90691 | 0.021758 | 23.9915 | 0.33435 | 0.15991 | 0.034584 | 2.2409 | 2.2409 | 0.002933
2019-03-21 12:43 | 1553186585.3846 | 478 | 0.20233 | 0.007048 | 0.000189 | 100.4087 | 96.8596 | 109.1799 | 21.2567 | 0.004421 | 0.003925 | 221.5992 | 0.2694 | 8.4021 | 0.004036 | 0.005058 | 0.005018 | 5.9e-05 | 0.000964 | 0.2568 | 1.2869 | 2.548 | 99.36 | 21.07 | 37.305 | 128.5604 | 119.1404 | 116.4159 | 0.009389 | 0.90625 | 0.021758 | 24.0089 | 0.32189 | 0.16027 | 0.034453 | 2.2323 | 2.2323 | 0.002951
2019-03-21 12:43 | 1553186586.3847 | 479 | 0.20247 | 0.006944 | 0.000188 | 100.3727 | 96.723 | 109.5508 | 21.2568 | 0.004459 | 0.003939 | 221.5348 | 0.26857 | 8.4018 | 0.004053 | 0.005712 | 0.005644 | 5.8e-05 | 0.000963 | 0.2891 | 1.3526 | 2.548 | 99.36 | 21.07 | 37.304 | 129.5094 | 119.5464 | 116.7723 | 0.009388 | 0.90543 | 0.021785 | 24.0607 | 0.3313 | 0.15977 | 0.033728 | 2.188 | 2.188 | 0.003001
2019-03-21 12:43 | 1553186587.3853 | 480 | 0.20261 | 0.006839 | 0.000187 | 100.2875 | 96.5994 | 109.8928 | 21.2566 | 0.004461 | 0.00388 | 221.4824 | 0.26819 | 8.4028 | 0.004062 | 0.005201 | 0.005116 | 5.9e-05 | 0.000965 | 0.26449 | 1.3322 | 2.5481 | 99.36 | 21.07 | 37.303 | 129.56 | 119.7801 | 115.3161 | 0.009388 | 0.90882 | 0.02186 | 24.0532 | 0.33514 | 0.15955 | 0.033027 | 2.15 | 2.15 | 0.00305
2019-03-21 12:43 | 1553186588.3858 | 481 | 0.20277 | 0.006728 | 0.000184 | 100.2064 | 96.3 | 110.3031 | 21.2568 | 0.0044 | 0.003891 | 221.416 | 0.26867 | 8.399 | 0.004012 | 0.005655 | 0.005582 | 5.7e-05 | 0.000965 | 0.28136 | 1.3083 | 2.548 | 99.36 | 21.07 | 37.3 | 128.0423 | 118.5404 | 115.6064 | 0.009387 | 0.90814 | 0.021893 | 24.107 | 0.35567 | 0.15983 | 0.032229 | 2.1012 | 2.1012 | 0.003126
2019-03-21 12:43 | 1553186589.3859 | 482 | 0.2031 | 0.006619 | 0.00018 | 100.1558 | 96.3514 | 110.3892 | 21.2569 | 0.004351 | 0.003879 | 221.3555 | 0.26835 | 8.4016 | 0.003991 | 0.005621 | 0.005564 | 5.7e-05 | 0.000967 | 0.27743 | 1.3256 | 2.5478 | 99.36 | 21.07 | 37.299 | 126.8594 | 118.0405 | 115.2942 | 0.009387 | 0.90887 | 0.02191 | 24.1068 | 0.34637 | 0.15964 | 0.030378 | 1.9821 | 1.9821 | 0.00331
2019-03-21 12:43 | 1553186590.3845 | 483 | 0.20333 | 0.006512 | 0.000176 | 100.1808 | 96.4018 | 110.5401 | 21.2569 | 0.004375 | 0.003808 | 221.2779 | 0.26948 | 8.403 | 0.003959 | 0.005662 | 0.005593 | 5.6e-05 | 0.000971 | 0.26377 | 1.3225 | 2.5479 | 99.36 | 21.07 | 37.296 | 127.4329 | 117.2539 | 113.5791 | 0.009386 | 0.9129 | 0.021973 | 24.0699 | 0.34393 | 0.16031 | 0.029125 | 1.9059 | 1.9059 | 0.003457
2019-03-21 12:43 | 1553186591.3859 | 484 | 0.20344 | 0.006398 | 0.000172 | 100.1203 | 96.5966 | 110.3927 | 21.257 | 0.004376 | 0.003856 | 221.1876 | 0.26859 | 8.4045 | 0.003976 | 0.005554 | 0.005522 | 6e-05 | 0.000969 | 0.30574 | 1.2887 | 2.5478 | 99.36 | 21.07 | 37.295 | 127.4592 | 117.6809 | 114.7453 | 0.009386 | 0.91016 | 0.021926 | 24.0901 | 0.32047 | 0.15979 | 0.028602 | 1.8678 | 1.8678 | 0.003516
2019-03-21 12:43 | 1553186592.385 | 485 | 0.20372 | 0.006282 | 0.000166 | 100.0595 | 96.3334 | 110.6845 | 21.2566 | 0.004265 | 0.003797 | 221.1314 | 0.26919 | 8.4007 | 0.003917 | 0.005542 | 0.005418 | 5.6e-05 | 0.00097 | 0.24298 | 1.2974 | 2.548 | 99.36 | 21.07 | 37.294 | 124.7492 | 116.2275 | 113.3063 | 0.009386 | 0.91355 | 0.021996 | 24.0771 | 0.33945 | 0.16014 | 0.027099 | 1.7754 | 1.7754 | 0.003707
2019-03-21 12:43 | 1553186593.3856 | 486 | 0.20386 | 0.006148 | 0.000161 | 99.9272 | 95.9883 | 110.9558 | 21.2561 | 0.004305 | 0.003804 | 221.0667 | 0.26941 | 8.4001 | 0.004012 | 0.005937 | 0.005883 | 5.7e-05 | 0.000969 | 0.25033 | 1.2981 | 2.5481 | 99.36 | 21.07 | 37.291 | 125.7289 | 118.5428 | 113.4716 | 0.009385 | 0.91316 | 0.022018 | 24.1118 | 0.35971 | 0.16027 | 0.026414 | 1.7324 | 1.7324 | 0.003802
2019-03-21 12:43 | 1553186594.3851 | 487 | 0.20401 | 0.006023 | 0.000154 | 99.89 | 96.3177 | 111.4347 | 21.2564 | 0.004293 | 0.003781 | 221.0325 | 0.26907 | 8.4008 | 0.004 | 0.005808 | 0.005677 | 5.8e-05 | 0.000971 | 0.27887 | 1.3149 | 2.548 | 99.36 | 21.07 | 37.29 | 125.4179 | 118.2476 | 112.9135 | 0.009385 | 0.91448 | 0.022081 | 24.1463 | 0.32574 | 0.16007 | 0.025694 | 1.6903 | 1.6903 | 0.003892
2019-03-21 12:43 | 1553186595.3851 | 488 | 0.20416 | 0.005886 | 0.000149 | 99.8933 | 96.6293 | 111.4729 | 21.257 | 0.004228 | 0.00374 | 220.9424 | 0.26858 | 8.405 | 0.004041 | 0.00588 | 0.005746 | 5.9e-05 | 0.000971 | 0.28119 | 1.3393 | 2.548 | 99.36 | 21.06 | 37.287 | 123.8295 | 119.2579 | 111.9261 | 0.009378 | 0.91682 | 0.022113 | 24.1196 | 0.29714 | 0.15978 | 0.024957 | 1.6444 | 1.6444 | 0.003993
2019-03-21 12:43 | 1553186596.3858 | 489 | 0.20427 | 0.005753 | 0.000144 | 99.8533 | 96.7577 | 111.2183 | 21.2569 | 0.004301 | 0.00371 | 220.8935 | 0.2685 | 8.4013 | 0.004052 | 0.00564 | 0.005563 | 5.6e-05 | 0.000972 | 0.27688 | 1.2965 | 2.5479 | 99.36 | 21.06 | 37.286 | 125.6278 | 119.5315 | 111.1948 | 0.009378 | 0.91856 | 0.022109 | 24.0691 | 0.28169 | 0.15973 | 0.024462 | 1.6117 | 1.6117 | 0.004073
2019-03-21 12:43 | 1553186597.3853 | 490 | 0.20445 | 0.005618 | 0.00014 | 99.7068 | 97.4171 | 111.1471 | 21.2567 | 0.004273 | 0.003756 | 220.8357 | 0.26906 | 8.4033 | 0.003963 | 0.005171 | 0.005039 | 5.8e-05 | 0.000972 | 0.26698 | 1.2972 | 2.548 | 99.36 | 21.06 | 37.286 | 124.9417 | 117.3613 | 112.3142 | 0.009378 | 0.9159 | 0.02207 | 24.0964 | 0.2078 | 0.16007 | 0.023565 | 1.5499 | 1.5499 | 0.004245
2019-03-21 12:43 | 1553186598.386 | 491 | 0.20458 | 0.005487 | 0.000137 | 99.7199 | 97.7008 | 111.2124 | 21.2564 | 0.004184 | 0.003643 | 220.7573 | 0.26914 | 8.4011 | 0.00402 | 0.005357 | 0.005359 | 5.8e-05 | 0.000976 | 0.28407 | 1.2972 | 2.5481 | 99.36 | 21.07 | 37.285 | 122.744 | 118.7476 | 109.5671 | 0.009384 | 0.92247 | 0.022155 | 24.0174 | 0.18296 | 0.16011 | 0.022932 | 1.5142 | 1.5142 | 0.004346
2019-03-21 12:43 | 1553186599.386 | 492 | 0.20472 | 0.005346 | 0.000136 | 99.6942 | 98.0002 | 111.4617 | 21.2567 | 0.004221 | 0.003689 | 220.7325 | 0.26824 | 8.402 | 0.003995 | 0.005112 | 0.00497 | 5.9e-05 | 0.000976 | 0.27831 | 1.3253 | 2.548 | 99.36 | 21.06 | 37.286 | 123.6623 | 118.1346 | 110.6873 | 0.009378 | 0.91978 | 0.022148 | 24.0795 | 0.1533 | 0.15957 | 0.022276 | 1.4703 | 1.4703 | 0.004461
2019-03-21 12:43 | 1553186600.3864 | 493 | 0.20486 | 0.005217 | 0.000134 | 99.4985 | 98.0414 | 111.571 | 21.2568 | 0.004123 | 0.003702 | 220.6666 | 0.26895 | 8.4026 | 0.003979 | 0.00483 | 0.004779 | 5.6e-05 | 0.000978 | 0.30453 | 1.3579 | 2.5481 | 99.36 | 21.06 | 37.287 | 121.2522 | 117.7457 | 110.9858 | 0.009378 | 0.91906 | 0.02215 | 24.1007 | 0.13196 | 0.16 | 0.021574 | 1.4241 | 1.4241 | 0.004618
2019-03-21 12:43 | 1553186601.3841 | 494 | 0.20503 | 0.00509 | 0.000134 | 99.4726 | 98.1831 | 111.8283 | 21.2569 | 0.004215 | 0.003668 | 220.6624 | 0.26948 | 8.4033 | 0.003922 | 0.00511 | 0.004999 | 5.9e-05 | 0.000982 | 0.26842 | 1.3387 | 2.548 | 99.36 | 21.06 | 37.289 | 123.5036 | 116.3481 | 110.178 | 0.009379 | 0.921 | 0.022199 | 24.1031 | 0.11671 | 0.16031 | 0.020691 | 1.3686 | 1.3686 | 0.004814
2019-03-21 12:43 | 1553186602.3842 | 495 | 0.20516 | 0.00496 | 0.000135 | 99.4341 | 98.2509 | 112.1116 | 21.2566 | 0.004131 | 0.003643 | 220.6033 | 0.26843 | 8.4033 | 0.003961 | 0.005316 | 0.005273 | 5.7e-05 | 0.000984 | 0.2557 | 1.2851 | 2.548 | 99.36 | 21.06 | 37.289 | 121.4551 | 117.2941 | 109.5578 | 0.009379 | 0.92249 | 0.022245 | 24.1141 | 0.10708 | 0.15969 | 0.020108 | 1.3327 | 1.3327 | 0.004925
2019-03-21 12:43 | 1553186603.3856 | 496 | 0.20532 | 0.004839 | 0.000135 | 99.3656 | 97.8592 | 112.2196 | 21.2571 | 0.004103 | 0.003651 | 220.5733 | 0.26902 | 8.4018 | 0.003935 | 0.005419 | 0.005303 | 5.6e-05 | 0.000982 | 0.30431 | 1.2641 | 2.548 | 99.36 | 21.06 | 37.288 | 120.7677 | 116.673 | 109.7554 | 0.009379 | 0.92202 | 0.02225 | 24.1319 | 0.13664 | 0.16004 | 0.019293 | 1.2787 | 1.2787 | 0.005144
2019-03-21 12:43 | 1553186604.3846 | 497 | 0.20545 | 0.004712 | 0.000135 | 99.3719 | 98.1991 | 112.2576 | 21.2566 | 0.00412 | 0.003592 | 220.5211 | 0.26863 | 8.4006 | 0.003896 | 0.005378 | 0.005248 | 5.9e-05 | 0.000987 | 0.2656 | 1.3133 | 2.5479 | 99.36 | 21.06 | 37.288 | 121.1877 | 115.7123 | 108.3266 | 0.009379 | 0.92547 | 0.022295 | 24.0909 | 0.10619 | 0.15981 | 0.01868 | 1.2405 | 1.2405 | 0.005295
2019-03-21 12:43 | 1553186605.3843 | 498 | 0.20554 | 0.0046 | 0.000136 | 99.2715 | 98.0363 | 112.4789 | 21.2572 | 0.00414 | 0.003585 | 220.4846 | 0.26858 | 8.4053 | 0.003903 | 0.005046 | 0.005024 | 5.9e-05 | 0.000989 | 0.27649 | 1.2771 | 2.548 | 99.36 | 21.06 | 37.289 | 121.6782 | 115.8917 | 108.1421 | 0.009379 | 0.92592 | 0.022323 | 24.1088 | 0.11199 | 0.15978 | 0.018244 | 1.2128 | 1.2128 | 0.005414
2019-03-21 12:43 | 1553186606.3844 | 499 | 0.20564 | 0.004489 | 0.000138 | 99.2193 | 98.1715 | 112.8653 | 21.257 | 0.004014 | 0.003567 | 220.436 | 0.26936 | 8.3998 | 0.003895 | 0.004892 | 0.004809 | 6e-05 | 0.000985 | 0.25249 | 1.335 | 2.5481 | 99.36 | 21.06 | 37.289 | 118.5921 | 115.6898 | 107.7126 | 0.009379 | 0.92696 | 0.022374 | 24.1366 | 0.094962 | 0.16024 | 0.017769 | 1.1837 | 1.1837 | 0.005564
2019-03-21 12:43 | 1553186607.3848 | 500 | 0.20576 | 0.004384 | 0.00014 | 99.1445 | 98.0148 | 112.9381 | 21.2574 | 0.004037 | 0.003524 | 220.3811 | 0.269 | 8.4018 | 0.003829 | 0.005232 | 0.005248 | 5.8e-05 | 0.000988 | 0.2521 | 1.2991 | 2.5479 | 99.36 | 21.06 | 37.289 | 119.1703 | 114.076 | 106.6624 | 0.009379 | 0.92953 | 0.022412 | 24.111 | 0.1025 | 0.16003 | 0.01718 | 1.1461 | 1.1461 | 0.005739
2019-03-21 12:43 | 1553186608.3854 | 501 | 0.20587 | 0.004277 | 0.000142 | 99.0483 | 98.036 | 112.7965 | 21.2575 | 0.004005 | 0.003513 | 220.387 | 0.2688 | 8.4035 | 0.003849 | 0.004931 | 0.00482 | 5.8e-05 | 0.000988 | 0.24961 | 1.305 | 2.548 | 99.36 | 21.06 | 37.289 | 118.3843 | 114.5845 | 106.3917 | 0.009379 | 0.93019 | 0.022406 | 24.0873 | 0.09189 | 0.15991 | 0.016646 | 1.1099 | 1.1099 | 0.005922
2019-03-21 12:43 | 1553186609.3855 | 502 | 0.20598 | 0.004176 | 0.000144 | 99.0221 | 98.0727 | 113.2716 | 21.2567 | 0.003984 | 0.003569 | 220.3013 | 0.26863 | 8.4017 | 0.003858 | 0.005742 | 0.005702 | 5.9e-05 | 0.000986 | 0.26952 | 1.3261 | 2.548 | 99.36 | 21.06 | 37.289 | 117.8699 | 114.783 | 107.7658 | 0.009379 | 0.92683 | 0.022412 | 24.1817 | 0.086167 | 0.15981 | 0.016126 | 1.0752 | 1.0752 | 0.006109
2019-03-21 12:43 | 1553186610.3849 | 503 | 0.2061 | 0.004078 | 0.000147 | 98.9652 | 97.9069 | 113.6695 | 21.2571 | 0.003972 | 0.003543 | 220.2819 | 0.26899 | 8.4031 | 0.00387 | 0.005443 | 0.005335 | 6e-05 | 0.000987 | 0.27815 | 1.3028 | 2.5481 | 99.36 | 21.06 | 37.288 | 117.5677 | 115.0935 | 107.1384 | 0.009379 | 0.92836 | 0.02247 | 24.2041 | 0.096166 | 0.16002 | 0.015555 | 1.0395 | 1.0395 | 0.006327
2019-03-21 12:43 | 1553186611.3846 | 504 | 0.2062 | 0.00398 | 0.000151 | 98.9041 | 97.6732 | 113.7641 | 21.2569 | 0.003955 | 0.003547 | 220.2367 | 0.26919 | 8.4017 | 0.003952 | 0.004773 | 0.004703 | 5.9e-05 | 0.000984 | 0.28512 | 1.2745 | 2.5479 | 99.36 | 21.06 | 37.285 | 117.1525 | 117.0835 | 107.2375 | 0.009378 | 0.92812 | 0.022477 | 24.2174 | 0.11201 | 0.16014 | 0.015024 | 1.0038 | 1.0038 | 0.006557
2019-03-21 12:44 | 1553186612.3851 | 505 | 0.20627 | 0.003886 | 0.000153 | 98.7745 | 97.8279 | 113.8408 | 21.2574 | 0.003979 | 0.003512 | 220.2084 | 0.26907 | 8.3997 | 0.003913 | 0.005592 | 0.00547 | 5.7e-05 | 0.000987 | 0.28119 | 1.3068 | 2.5479 | 99.36 | 21.06 | 37.283 | 117.7541 | 116.1429 | 106.3853 | 0.009377 | 0.9302 | 0.022509 | 24.1984 | 0.086137 | 0.16007 | 0.014748 | 0.98651 | 0.98651 | 0.006669
2019-03-21 12:44 | 1553186613.3858 | 506 | 0.20639 | 0.003796 | 0.000156 | 98.7331 | 97.7362 | 114.1873 | 21.2572 | 0.00394 | 0.00349 | 220.1904 | 0.26892 | 8.4018 | 0.00402 | 0.005279 | 0.005236 | 6.3e-05 | 0.000991 | 0.28456 | 1.3409 | 2.548 | 99.36 | 21.06 | 37.283 | 116.7895 | 118.735 | 105.8444 | 0.009377 | 0.93153 | 0.02256 | 24.2179 | 0.090766 | 0.15998 | 0.014118 | 0.94605 | 0.94605 | 0.00695
2019-03-21 12:44 | 1553186614.3862 | 507 | 0.20651 | 0.003702 | 0.00016 | 98.6303 | 97.7046 | 114.375 | 21.2575 | 0.00388 | 0.003515 | 220.181 | 0.26946 | 8.4009 | 0.003967 | 0.00589 | 0.005811 | 6.2e-05 | 0.000989 | 0.29607 | 1.3016 | 2.5481 | 99.36 | 21.06 | 37.283 | 115.3224 | 117.4493 | 106.45 | 0.009377 | 0.93005 | 0.02256 | 24.2572 | 0.084348 | 0.1603 | 0.01349 | 0.90342 | 0.90342 | 0.007293
2019-03-21 12:44 | 1553186615.3849 | 508 | 0.20656 | 0.003613 | 0.000163 | 98.6097 | 97.6 | 114.3473 | 21.2575 | 0.003903 | 0.003473 | 220.1557 | 0.26967 | 8.3997 | 0.004005 | 0.005514 | 0.005419 | 6.3e-05 | 0.000989 | 0.30718 | 1.3055 | 2.548 | 99.36 | 21.06 | 37.282 | 115.8856 | 118.3825 | 105.4348 | 0.009377 | 0.93254 | 0.022588 | 24.2218 | 0.092061 | 0.16043 | 0.013319 | 0.89276 | 0.89276 | 0.007386
2019-03-21 12:44 | 1553186616.3838 | 509 | 0.20664 | 0.00352 | 0.000166 | 98.5665 | 97.6613 | 114.4117 | 21.2566 | 0.003868 | 0.003464 | 220.1038 | 0.26768 | 8.4032 | 0.003988 | 0.005377 | 0.005302 | 6e-05 | 0.000988 | 0.29374 | 1.3044 | 2.548 | 99.36 | 21.06 | 37.28 | 115.0257 | 117.957 | 105.2034 | 0.009377 | 0.93311 | 0.022601 | 24.2212 | 0.082525 | 0.15924 | 0.012958 | 0.86868 | 0.86868 | 0.007534
2019-03-21 12:44 | 1553186617.3853 | 510 | 0.20678 | 0.003435 | 0.00017 | 98.5442 | 97.1487 | 114.7027 | 21.257 | 0.003856 | 0.003434 | 220.1053 | 0.2688 | 8.4041 | 0.003991 | 0.00573 | 0.005718 | 6.2e-05 | 0.000986 | 0.27599 | 1.3249 | 2.548 | 99.36 | 21.06 | 37.28 | 114.7333 | 118.0335 | 104.4945 | 0.009377 | 0.93486 | 0.022651 | 24.2292 | 0.12757 | 0.15991 | 0.012229 | 0.82095 | 0.82095 | 0.008006
2019-03-21 12:44 | 1553186618.384 | 511 | 0.2069 | 0.003344 | 0.000174 | 98.4581 | 96.8756 | 115.1334 | 21.2575 | 0.003903 | 0.003436 | 220.0678 | 0.26888 | 8.4013 | 0.003953 | 0.005352 | 0.005271 | 5.9e-05 | 0.000989 | 0.27505 | 1.3307 | 2.5479 | 99.36 | 21.06 | 37.277 | 115.8842 | 117.1105 | 104.5343 | 0.009376 | 0.93476 | 0.022692 | 24.276 | 0.14493 | 0.15996 | 0.011577 | 0.77789 | 0.77789 | 0.008451
2019-03-21 12:44 | 1553186619.3844 | 512 | 0.20705 | 0.00326 | 0.000178 | 98.487 | 97.0295 | 115.1625 | 21.2571 | 0.003806 | 0.003412 | 220.0609 | 0.26921 | 8.4036 | 0.003969 | 0.005578 | 0.005521 | 6.1e-05 | 0.000989 | 0.28655 | 1.3111 | 2.548 | 99.36 | 21.06 | 37.277 | 113.5327 | 117.4877 | 103.9591 | 0.009376 | 0.93619 | 0.022712 | 24.2605 | 0.13336 | 0.16015 | 0.010779 | 0.7241 | 0.7241 | 0.00909
2019-03-21 12:44 | 1553186620.384 | 513 | 0.20711 | 0.003165 | 0.000182 | 98.49 | 97.0278 | 115.0755 | 21.2571 | 0.003874 | 0.003404 | 220.0256 | 0.26938 | 8.4021 | 0.003986 | 0.005763 | 0.005668 | 5.9e-05 | 0.000989 | 0.26449 | 1.3054 | 2.5479 | 99.36 | 21.06 | 37.275 | 115.1764 | 117.9084 | 103.75 | 0.009375 | 0.93671 | 0.02271 | 24.2446 | 0.13378 | 0.16025 | 0.010556 | 0.70861 | 0.70861 | 0.009295
2019-03-21 12:44 | 1553186621.3842 | 514 | 0.20716 | 0.003081 | 0.000188 | 98.5138 | 97.1872 | 115.4327 | 21.2573 | 0.003796 | 0.003389 | 219.9975 | 0.26922 | 8.4048 | 0.00392 | 0.005218 | 0.005135 | 5.8e-05 | 0.00099 | 0.25321 | 1.3191 | 2.5481 | 99.36 | 21.06 | 37.275 | 113.2872 | 116.2962 | 103.3872 | 0.009375 | 0.93761 | 0.022756 | 24.2706 | 0.12126 | 0.16016 | 0.010354 | 0.69595 | 0.69595 | 0.009458
2019-03-21 12:44 | 1553186622.3846 | 515 | 0.20725 | 0.002991 | 0.000193 | 98.4316 | 97.3565 | 115.3595 | 21.2576 | 0.00374 | 0.003349 | 219.9903 | 0.26897 | 8.4018 | 0.003904 | 0.005791 | 0.005732 | 5.7e-05 | 0.00099 | 0.274 | 1.3051 | 2.5481 | 99.36 | 21.06 | 37.274 | 111.9235 | 115.9199 | 102.4249 | 0.009375 | 0.94001 | 0.022778 | 24.2318 | 0.098231 | 0.16001 | 0.009871 | 0.66337 | 0.66337 | 0.009914
2019-03-21 12:44 | 1553186623.3843 | 516 | 0.20738 | 0.002899 | 0.000198 | 98.391 | 97.382 | 115.2172 | 21.2577 | 0.003771 | 0.003314 | 219.9724 | 0.2683 | 8.4027 | 0.003922 | 0.005809 | 0.005695 | 5.9e-05 | 0.00099 | 0.2349 | 1.2945 | 2.548 | 99.36 | 21.06 | 37.274 | 112.6706 | 116.341 | 101.5745 | 0.009375 | 0.94215 | 0.02279 | 24.1895 | 0.092201 | 0.15961 | 0.009215 | 0.61863 | 0.61863 | 0.010604
2019-03-21 12:44 | 1553186624.3853 | 517 | 0.20751 | 0.002807 | 0.000202 | 98.3568 | 97.4396 | 115.1222 | 21.2573 | 0.003713 | 0.0033 | 219.9524 | 0.26838 | 8.403 | 0.003876 | 0.005415 | 0.005351 | 6e-05 | 0.000991 | 0.25962 | 1.2899 | 2.548 | 99.36 | 21.06 | 37.273 | 111.2585 | 115.2351 | 101.2309 | 0.009375 | 0.94301 | 0.022791 | 24.1684 | 0.0838 | 0.15966 | 0.008545 | 0.57262 | 0.57262 | 0.01146
2019-03-21 12:44 | 1553186625.3847 | 518 | 0.20767 | 0.002715 | 0.000204 | 98.324 | 97.4781 | 115.1175 | 21.258 | 0.003757 | 0.003295 | 219.9461 | 0.26936 | 8.4037 | 0.003837 | 0.005026 | 0.005022 | 6.1e-05 | 0.00099 | 0.3011 | 1.2895 | 2.5481 | 99.36 | 21.06 | 37.271 | 112.3376 | 114.2918 | 101.1179 | 0.009374 | 0.9433 | 0.022794 | 24.1642 | 0.077287 | 0.16024 | 0.007702 | 0.515 | 0.515 | 0.012788
2019-03-21 12:44 | 1553186626.3844 | 519 | 0.20779 | 0.002628 | 0.000204 | 98.3189 | 97.5288 | 115.1794 | 21.2574 | 0.003709 | 0.003247 | 219.9546 | 0.26837 | 8.4045 | 0.003839 | 0.005835 | 0.005783 | 6.4e-05 | 0.000989 | 0.27489 | 1.274 | 2.5478 | 99.36 | 21.06 | 37.271 | 111.156 | 114.3385 | 99.9565 | 0.009374 | 0.94623 | 0.022836 | 24.1332 | 0.072171 | 0.15965 | 0.00707 | 0.47274 | 0.47274 | 0.01388
2019-03-21 12:44 | 1553186627.3856 | 520 | 0.20787 | 0.002537 | 0.000205 | 98.3383 | 97.255 | 115.297 | 21.258 | 0.003683 | 0.003295 | 219.9228 | 0.26861 | 8.4013 | 0.003848 | 0.005595 | 0.00549 | 6e-05 | 0.000989 | 0.28208 | 1.2896 | 2.5479 | 99.36 | 21.06 | 37.271 | 110.5206 | 114.5391 | 101.1179 | 0.009374 | 0.9433 | 0.022812 | 24.1831 | 0.099082 | 0.1598 | 0.006709 | 0.44747 | 0.44747 | 0.014677
2019-03-21 12:44 | 1553186628.3861 | 521 | 0.20794 | 0.002445 | 0.000205 | 98.4429 | 97.4386 | 115.5233 | 21.2582 | 0.00368 | 0.003265 | 219.8973 | 0.26875 | 8.4023 | 0.003839 | 0.005862 | 0.005853 | 5.6e-05 | 0.000985 | 0.3147 | 1.3013 | 2.5479 | 99.36 | 21.06 | 37.271 | 110.4483 | 114.329 | 100.4022 | 0.009374 | 0.9451 | 0.022856 | 24.1836 | 0.091721 | 0.15988 | 0.006385 | 0.42611 | 0.42611 | 0.015421
2019-03-21 12:44 | 1553186629.3846 | 522 | 0.20803 | 0.002358 | 0.000206 | 98.3973 | 97.4481 | 115.4304 | 21.2577 | 0.003649 | 0.003222 | 219.8933 | 0.269 | 8.4026 | 0.003771 | 0.005811 | 0.00567 | 5.8e-05 | 0.000982 | 0.25514 | 1.3026 | 2.5481 | 99.36 | 21.05 | 37.271 | 109.6938 | 112.6699 | 99.3553 | 0.009369 | 0.94776 | 0.022879 | 24.14 | 0.086703 | 0.16003 | 0.005971 | 0.39804 | 0.39804 | 0.016524
2019-03-21 12:44 | 1553186630.3847 | 523 | 0.20806 | 0.002276 | 0.000206 | 98.4208 | 97.6099 | 115.6761 | 21.2573 | 0.003638 | 0.003226 | 219.8964 | 0.26904 | 8.4032 | 0.003865 | 0.005274 | 0.00526 | 6e-05 | 0.000983 | 0.26985 | 1.276 | 2.548 | 99.36 | 21.05 | 37.271 | 109.4309 | 114.9503 | 99.459 | 0.009369 | 0.9475 | 0.0229 | 24.1691 | 0.074005 | 0.16005 | 0.00591 | 0.39427 | 0.39427 | 0.016684
2019-03-21 12:44 | 1553186631.386 | 524 | 0.20808 | 0.002193 | 0.000205 | 98.4426 | 97.4931 | 115.5069 | 21.2575 | 0.003601 | 0.003195 | 219.8982 | 0.26856 | 8.3987 | 0.003813 | 0.005706 | 0.005604 | 5.7e-05 | 0.000982 | 0.27041 | 1.2931 | 2.5478 | 99.36 | 21.05 | 37.27 | 108.5324 | 113.6863 | 98.6947 | 0.009368 | 0.94944 | 0.022907 | 24.1266 | 0.086693 | 0.15977 | 0.005912 | 0.39462 | 0.39462 | 0.01664
2019-03-21 12:44 | 1553186632.3847 | 525 | 0.20813 | 0.002116 | 0.000202 | 98.5436 | 97.6708 | 115.4442 | 21.2577 | 0.003657 | 0.003216 | 219.8702 | 0.26841 | 8.4025 | 0.003876 | 0.005712 | 0.005663 | 6.1e-05 | 0.000987 | 0.25658 | 1.3301 | 2.548 | 99.36 | 21.05 | 37.269 | 109.905 | 115.2309 | 99.2134 | 0.009368 | 0.94812 | 0.022885 | 24.1369 | 0.079569 | 0.15968 | 0.00568 | 0.37845 | 0.37845 | 0.017341
2019-03-21 12:44 | 1553186633.3838 | 526 | 0.20818 | 0.002044 | 0.000198 | 98.5871 | 97.6514 | 115.697 | 21.2579 | 0.00357 | 0.003149 | 219.8521 | 0.26818 | 8.4023 | 0.00389 | 0.004699 | 0.004574 | 5.8e-05 | 0.000985 | 0.26897 | 1.3392 | 2.5479 | 99.36 | 21.05 | 37.267 | 107.7844 | 115.5609 | 97.5991 | 0.009368 | 0.95225 | 0.02296 | 24.1108 | 0.085295 | 0.15954 | 0.005469 | 0.36536 | 0.36536 | 0.017947
2019-03-21 12:44 | 1553186634.3839 | 527 | 0.20829 | 0.001979 | 0.000194 | 98.7346 | 97.8936 | 115.9266 | 21.2576 | 0.003544 | 0.003201 | 219.8279 | 0.26862 | 8.4025 | 0.00381 | 0.005016 | 0.004943 | 6.4e-05 | 0.000981 | 0.25227 | 1.3241 | 2.5479 | 99.36 | 21.05 | 37.266 | 107.1525 | 113.628 | 98.8563 | 0.009367 | 0.94903 | 0.022943 | 24.1756 | 0.076512 | 0.1598 | 0.004912 | 0.32702 | 0.32702 | 0.020084
2019-03-21 12:44 | 1553186635.3853 | 528 | 0.20833 | 0.001913 | 0.000189 | 98.6219 | 97.8766 | 116.2082 | 21.2571 | 0.00352 | 0.00314 | 219.8189 | 0.26912 | 8.4012 | 0.003838 | 0.005517 | 0.005483 | 6.2e-05 | 0.000982 | 0.25409 | 1.3212 | 2.5482 | 99.36 | 21.05 | 37.262 | 106.5701 | 114.3075 | 97.379 | 0.009366 | 0.95281 | 0.023017 | 24.1569 | 0.06786 | 0.1601 | 0.004742 | 0.31661 | 0.31661 | 0.020783
2019-03-21 12:44 | 1553186636.385 | 529 | 0.20837 | 0.00185 | 0.000183 | 98.6059 | 97.7713 | 116.3587 | 21.2576 | 0.00353 | 0.003156 | 219.8136 | 0.26938 | 8.4051 | 0.003778 | 0.005625 | 0.005465 | 5.9e-05 | 0.000984 | 0.27975 | 1.3065 | 2.5479 | 99.36 | 21.05 | 37.261 | 106.8232 | 112.8514 | 97.7541 | 0.009366 | 0.95185 | 0.02302 | 24.1847 | 0.076023 | 0.16026 | 0.004569 | 0.30508 | 0.30508 | 0.02159
2019-03-21 12:44 | 1553186637.384 | 530 | 0.20843 | 0.001792 | 0.000176 | 98.6386 | 97.8776 | 116.1334 | 21.258 | 0.00352 | 0.003126 | 219.8173 | 0.26876 | 8.4018 | 0.00377 | 0.006138 | 0.006035 | 6.1e-05 | 0.00098 | 0.28639 | 1.2832 | 2.5481 | 99.36 | 21.06 | 37.261 | 106.5748 | 112.6416 | 97.0272 | 0.009372 | 0.95372 | 0.023021 | 24.1376 | 0.069279 | 0.15989 | 0.004332 | 0.28911 | 0.28911 | 0.02273
2019-03-21 12:44 | 1553186638.3855 | 531 | 0.20843 | 0.00174 | 0.000169 | 98.6562 | 97.9246 | 115.6787 | 21.2576 | 0.003514 | 0.003106 | 219.834 | 0.26835 | 8.4033 | 0.003858 | 0.005989 | 0.005959 | 6e-05 | 0.00098 | 0.262 | 1.2892 | 2.5479 | 99.36 | 21.06 | 37.257 | 106.4241 | 114.7942 | 96.563 | 0.009371 | 0.95492 | 0.02299 | 24.0752 | 0.066579 | 0.15964 | 0.004399 | 0.29374 | 0.29374 | 0.022337
2019-03-21 12:44 | 1553186639.3852 | 532 | 0.20846 | 0.001685 | 0.000162 | 98.6237 | 97.9477 | 115.7599 | 21.2575 | 0.00351 | 0.003086 | 219.8079 | 0.26861 | 8.4032 | 0.003788 | 0.005312 | 0.005265 | 6.1e-05 | 0.000981 | 0.2891 | 1.2702 | 2.548 | 99.36 | 21.06 | 37.255 | 106.3376 | 113.0867 | 96.0726 | 0.00937 | 0.95619 | 0.023013 | 24.0677 | 0.061519 | 0.1598 | 0.004275 | 0.28587 | 0.28587 | 0.022975
2019-03-21 12:44 | 1553186640.3841 | 533 | 0.2085 | 0.00163 | 0.000156 | 98.6553 | 97.9358 | 115.763 | 21.2571 | 0.003487 | 0.003122 | 219.8384 | 0.26863 | 8.4044 | 0.003835 | 0.005281 | 0.005216 | 5.9e-05 | 0.000982 | 0.28528 | 1.337 | 2.548 | 99.36 | 21.06 | 37.254 | 105.7674 | 114.2339 | 96.9426 | 0.00937 | 0.95394 | 0.022986 | 24.0964 | 0.065472 | 0.15981 | 0.004141 | 0.27667 | 0.27667 | 0.02374
2019-03-21 12:44 | 1553186641.3853 | 534 | 0.20853 | 0.001585 | 0.00015 | 98.5632 | 97.9069 | 115.901 | 21.2577 | 0.003438 | 0.003148 | 219.8279 | 0.26763 | 8.4004 | 0.00377 | 0.005476 | 0.005374 | 5.6e-05 | 0.000984 | 0.2912 | 1.3186 | 2.548 | 99.36 | 21.06 | 37.251 | 104.5824 | 112.6603 | 97.5605 | 0.009369 | 0.95235 | 0.022981 | 24.1308 | 0.059765 | 0.15921 | 0.003974 | 0.26542 | 0.26542 | 0.024653
2019-03-21 12:44 | 1553186642.3842 | 535 | 0.20857 | 0.00154 | 0.000146 | 98.5829 | 97.8507 | 116.0935 | 21.2579 | 0.00348 | 0.003087 | 219.8549 | 0.26854 | 8.4025 | 0.003768 | 0.005315 | 0.005385 | 5.9e-05 | 0.000981 | 0.26537 | 1.3161 | 2.5479 | 99.36 | 21.06 | 37.25 | 105.606 | 112.5934 | 96.095 | 0.009369 | 0.95613 | 0.023046 | 24.1031 | 0.066685 | 0.15975 | 0.003796 | 0.25412 | 0.25412 | 0.025837
2019-03-21 12:44 | 1553186643.3843 | 536 | 0.20849 | 0.001495 | 0.00014 | 98.6234 | 97.9024 | 116.0115 | 21.2572 | 0.003457 | 0.003098 | 219.8426 | 0.26875 | 8.4024 | 0.003708 | 0.005052 | 0.004966 | 6.1e-05 | 0.000986 | 0.23308 | 1.3215 | 2.5479 | 99.36 | 21.06 | 37.247 | 105.0396 | 111.1322 | 96.3542 | 0.009368 | 0.95546 | 0.023029 | 24.103 | 0.06563 | 0.15988 | 0.004394 | 0.29557 | 0.29557 | 0.022232
2019-03-21 12:44 | 1553186644.3848 | 537 | 0.20855 | 0.001456 | 0.000135 | 98.5543 | 97.7757 | 115.9397 | 21.2575 | 0.003411 | 0.003092 | 219.8359 | 0.26877 | 8.4034 | 0.003753 | 0.005166 | 0.00503 | 5.7e-05 | 0.000983 | 0.2912 | 1.3807 | 2.548 | 99.36 | 21.06 | 37.247 | 103.9195 | 112.2257 | 96.2081 | 0.009368 | 0.95583 | 0.023027 | 24.0908 | 0.070941 | 0.15989 | 0.004042 | 0.27149 | 0.27149 | 0.024205
2019-03-21 12:44 | 1553186645.3848 | 538 | 0.20852 | 0.001418 | 0.00013 | 98.5332 | 97.7696 | 116.0449 | 21.2578 | 0.003382 | 0.003089 | 219.8201 | 0.26783 | 8.4037 | 0.003771 | 0.004553 | 0.004522 | 5.8e-05 | 0.000979 | 0.25752 | 1.34 | 2.5478 | 99.36 | 21.06 | 37.246 | 103.217 | 112.6682 | 96.1535 | 0.009368 | 0.95598 | 0.023039 | 24.0999 | 0.069582 | 0.15933 | 0.004325 | 0.29155 | 0.29155 | 0.022461
2019-03-21 12:44 | 1553186646.3839 | 539 | 0.20855 | 0.001386 | 0.000125 | 98.4438 | 97.6327 | 115.8796 | 21.2577 | 0.003372 | 0.003069 | 219.7968 | 0.26813 | 8.4022 | 0.003728 | 0.005458 | 0.005395 | 5.8e-05 | 0.000981 | 0.25874 | 1.302 | 2.5481 | 99.36 | 21.06 | 37.245 | 102.9956 | 111.6307 | 95.6593 | 0.009368 | 0.95726 | 0.023038 | 24.0667 | 0.073999 | 0.15951 | 0.004193 | 0.2827 | 0.2827 | 0.02319
2019-03-21 12:44 | 1553186647.3856 | 540 | 0.20858 | 0.001356 | 0.000121 | 98.4354 | 97.8163 | 116.1254 | 21.258 | 0.003375 | 0.003029 | 219.7965 | 0.26839 | 8.4044 | 0.003759 | 0.005426 | 0.005389 | 5.8e-05 | 0.000981 | 0.24226 | 1.3348 | 2.5479 | 99.36 | 21.06 | 37.246 | 103.0499 | 112.3929 | 94.6958 | 0.009368 | 0.95976 | 0.023093 | 24.0607 | 0.056439 | 0.15966 | 0.004034 | 0.27259 | 0.27259 | 0.024073
2019-03-21 12:44 | 1553186648.3849 | 541 | 0.20865 | 0.001329 | 0.000117 | 98.2701 | 97.7938 | 116.5131 | 21.2579 | 0.003353 | 0.003067 | 219.7571 | 0.26847 | 8.4035 | 0.00366 | 0.005017 | 0.004967 | 5.8e-05 | 0.00098 | 0.26936 | 1.3047 | 2.5479 | 99.36 | 21.05 | 37.245 | 102.5275 | 109.9803 | 95.6016 | 0.009362 | 0.95741 | 0.023103 | 24.1305 | 0.043456 | 0.15971 | 0.003667 | 0.24745 | 0.24745 | 0.026526
2019-03-21 12:44 | 1553186649.385 | 542 | 0.20861 | 0.001304 | 0.000114 | 98.2927 | 97.6596 | 116.7671 | 21.2576 | 0.003415 | 0.003009 | 219.7817 | 0.26895 | 8.403 | 0.003722 | 0.005496 | 0.005374 | 5.9e-05 | 0.000981 | 0.27168 | 1.3252 | 2.5479 | 99.36 | 21.05 | 37.245 | 104.0313 | 111.4758 | 94.2211 | 0.009362 | 0.961 | 0.023171 | 24.1115 | 0.057799 | 0.16 | 0.003896 | 0.26431 | 0.26431 | 0.024879
2019-03-21 12:44 | 1553186650.3838 | 543 | 0.20861 | 0.001282 | 0.000111 | 98.1846 | 97.5462 | 116.8131 | 21.2576 | 0.003378 | 0.003062 | 219.746 | 0.26871 | 8.4008 | 0.003648 | 0.004472 | 0.004419 | 6.1e-05 | 0.000984 | 0.28672 | 1.3664 | 2.5481 | 99.36 | 21.05 | 37.246 | 103.1327 | 109.6747 | 95.5001 | 0.009362 | 0.95767 | 0.023136 | 24.1582 | 0.058345 | 0.15986 | 0.003969 | 0.26911 | 0.26911 | 0.024414
2019-03-21 12:44 | 1553186651.3837 | 544 | 0.20863 | 0.001253 | 0.000108 | 98.06 | 97.4052 | 116.9387 | 21.2574 | 0.003329 | 0.003109 | 219.7416 | 0.26832 | 8.4026 | 0.003705 | 0.005387 | 0.005344 | 6.1e-05 | 0.000983 | 0.28335 | 1.3204 | 2.548 | 99.36 | 21.06 | 37.246 | 101.9346 | 111.0779 | 96.6319 | 0.009368 | 0.95474 | 0.023113 | 24.2082 | 0.059931 | 0.15962 | 0.003877 | 0.26266 | 0.26266 | 0.024978
2019-03-21 12:44 | 1553186652.3838 | 545 | 0.20865 | 0.00123 | 0.000105 | 97.976 | 97.4266 | 117.0155 | 21.2578 | 0.003336 | 0.003027 | 219.7178 | 0.26835 | 8.3994 | 0.003721 | 0.005212 | 0.005126 | 5.9e-05 | 0.000978 | 0.28296 | 1.3098 | 2.5479 | 99.36 | 21.06 | 37.247 | 102.1087 | 111.458 | 94.6541 | 0.009368 | 0.95987 | 0.023182 | 24.1513 | 0.050294 | 0.15964 | 0.003779 | 0.25681 | 0.25681 | 0.02555
2019-03-21 12:44 | 1553186653.3851 | 546 | 0.20866 | 0.001209 | 0.000102 | 97.9108 | 97.1767 | 117.0948 | 21.2571 | 0.003324 | 0.003027 | 219.7206 | 0.26804 | 8.3995 | 0.003732 | 0.005314 | 0.005265 | 6.4e-05 | 0.000975 | 0.30038 | 1.2717 | 2.548 | 99.36 | 21.06 | 37.25 | 101.8174 | 111.7225 | 94.6465 | 0.009369 | 0.95989 | 0.02319 | 24.1593 | 0.067316 | 0.15946 | 0.003765 | 0.25608 | 0.25608 | 0.025592
2019-03-21 12:44 | 1553186654.384 | 547 | 0.20868 | 0.001193 | 0.0001 | 97.805 | 97.3813 | 116.9809 | 21.2575 | 0.003334 | 0.003042 | 219.7415 | 0.26861 | 8.4025 | 0.003697 | 0.005115 | 0.005103 | 6.3e-05 | 0.000976 | 0.26753 | 1.3376 | 2.548 | 99.36 | 21.06 | 37.252 | 102.0738 | 110.8839 | 95.0131 | 0.00937 | 0.95894 | 0.023168 | 24.1596 | 0.038827 | 0.15979 | 0.003631 | 0.24668 | 0.24668 | 0.026624
2019-03-21 12:44 | 1553186655.385 | 548 | 0.20871 | 0.001174 | 9.7e-05 | 97.7791 | 96.9883 | 116.7622 | 21.2566 | 0.003315 | 0.002996 | 219.7167 | 0.26802 | 8.4028 | 0.003658 | 0.005509 | 0.005318 | 6.2e-05 | 0.000976 | 0.27113 | 1.3392 | 2.548 | 99.36 | 21.06 | 37.254 | 101.6129 | 109.9222 | 93.8868 | 0.00937 | 0.96188 | 0.023181 | 24.1 | 0.07263 | 0.15944 | 0.003485 | 0.23681 | 0.23681 | 0.027673
2019-03-21 12:44 | 1553186656.3844 | 549 | 0.2087 | 0.001161 | 9.5e-05 | 97.6533 | 97.1085 | 116.7389 | 21.257 | 0.003343 | 0.002935 | 219.7131 | 0.26856 | 8.4027 | 0.003675 | 0.005374 | 0.005288 | 6.3e-05 | 0.000972 | 0.27433 | 1.3351 | 2.5481 | 99.36 | 21.06 | 37.259 | 102.2774 | 110.3351 | 92.4345 | 0.009371 | 0.9657 | 0.023225 | 24.0499 | 0.05004 | 0.15977 | 0.003571 | 0.24338 | 0.24338 | 0.02698
2019-03-21 12:44 | 1553186657.3851 | 550 | 0.2087 | 0.001146 | 9.3e-05 | 97.5714 | 96.8466 | 116.6519 | 21.2581 | 0.003254 | 0.002999 | 219.6876 | 0.26917 | 8.4064 | 0.003698 | 0.005201 | 0.005226 | 6.3e-05 | 0.000973 | 0.2756 | 1.3522 | 2.548 | 99.36 | 21.06 | 37.263 | 100.1299 | 110.9025 | 93.9758 | 0.009372 | 0.96165 | 0.023168 | 24.0915 | 0.066689 | 0.16013 | 0.003611 | 0.24569 | 0.24569 | 0.026788
2019-03-21 12:44 | 1553186658.3851 | 551 | 0.20874 | 0.001125 | 9.1e-05 | 97.4796 | 96.7666 | 116.4373 | 21.2576 | 0.003269 | 0.002956 | 219.6976 | 0.26808 | 8.4027 | 0.003671 | 0.005442 | 0.005334 | 6.1e-05 | 0.000972 | 0.27472 | 1.3203 | 2.5479 | 99.36 | 21.06 | 37.273 | 100.4919 | 110.2323 | 92.9246 | 0.009375 | 0.96441 | 0.023179 | 24.0349 | 0.065668 | 0.15948 | 0.0034 | 0.23125 | 0.23125 | 0.028344
2019-03-21 12:44 | 1553186659.3856 | 552 | 0.20868 | 0.001115 | 9e-05 | 97.4174 | 96.7373 | 116.2751 | 21.257 | 0.003238 | 0.002965 | 219.6909 | 0.26753 | 8.4005 | 0.003624 | 0.004793 | 0.00474 | 5.9e-05 | 0.000972 | 0.26626 | 1.376 | 2.548 | 99.36 | 21.06 | 37.28 | 99.746 | 109.0914 | 93.1563 | 0.009377 | 0.9638 | 0.023156 | 24.0257 | 0.06266 | 0.15915 | 0.00378 | 0.25745 | 0.25745 | 0.025407
2019-03-21 12:44 | 1553186660.3842 | 553 | 0.20875 | 0.001101 | 9e-05 | 97.345 | 96.5694 | 116.246 | 21.2579 | 0.003246 | 0.002923 | 219.6684 | 0.26807 | 8.4046 | 0.003699 | 0.005217 | 0.005199 | 5.9e-05 | 0.000974 | 0.26145 | 1.322 | 2.5478 | 99.36 | 21.06 | 37.288 | 99.9359 | 110.9181 | 92.1451 | 0.009379 | 0.96647 | 0.023185 | 23.9896 | 0.071549 | 0.15947 | 0.003351 | 0.22795 | 0.22795 | 0.028753
2019-03-21 12:44 | 1553186661.3843 | 554 | 0.20875 | 0.00109 | 8.9e-05 | 97.1724 | 96.3715 | 116.3223 | 21.2574 | 0.003231 | 0.002947 | 219.6554 | 0.2691 | 8.3992 | 0.00364 | 0.005055 | 0.004932 | 6e-05 | 0.000971 | 0.28031 | 1.3408 | 2.548 | 99.36 | 21.06 | 37.303 | 99.565 | 109.4757 | 92.7248 | 0.009382 | 0.96493 | 0.023174 | 24.0165 | 0.074031 | 0.16009 | 0.003407 | 0.23182 | 0.23182 | 0.028383
2019-03-21 12:44 | 1553186662.3847 | 555 | 0.20873 | 0.001078 | 8.9e-05 | 97.2022 | 96.4488 | 116.4871 | 21.2577 | 0.00326 | 0.002978 | 219.6398 | 0.26762 | 8.4017 | 0.003613 | 0.005397 | 0.005332 | 5.9e-05 | 0.000971 | 0.25066 | 1.332 | 2.548 | 99.36 | 21.06 | 37.312 | 100.285 | 108.8407 | 93.4591 | 0.009385 | 0.963 | 0.023167 | 24.0576 | 0.069601 | 0.15921 | 0.003488 | 0.2374 | 0.2374 | 0.027562
2019-03-21 12:44 | 1553186663.3844 | 556 | 0.20874 | 0.00106 | 9e-05 | 97.137 | 96.6303 | 116.6728 | 21.2573 | 0.003197 | 0.00298 | 219.6249 | 0.26869 | 8.4043 | 0.003632 | 0.005363 | 0.005347 | 6.2e-05 | 0.000975 | 0.29806 | 1.3519 | 2.5479 | 99.36 | 21.06 | 37.321 | 98.7525 | 109.2859 | 93.5087 | 0.009387 | 0.96287 | 0.023184 | 24.0784 | 0.046775 | 0.15984 | 0.003455 | 0.23519 | 0.23519 | 0.027933
2019-03-21 12:44 | 1553186664.3835 | 557 | 0.20871 | 0.001054 | 8.9e-05 | 97.1155 | 96.5997 | 116.9446 | 21.2574 | 0.003179 | 0.002964 | 219.5998 | 0.26855 | 8.4019 | 0.003633 | 0.004691 | 0.004604 | 6e-05 | 0.000973 | 0.24635 | 1.3242 | 2.548 | 99.36 | 21.05 | 37.337 | 98.3183 | 109.3268 | 93.12 | 0.009385 | 0.96389 | 0.023224 | 24.0936 | 0.047637 | 0.15976 | 0.003648 | 0.24903 | 0.24903 | 0.026367
2019-03-21 12:44 | 1553186665.3848 | 558 | 0.20872 | 0.001047 | 9e-05 | 96.9999 | 96.161 | 117.4114 | 21.258 | 0.003224 | 0.002903 | 219.6242 | 0.26781 | 8.4025 | 0.00354 | 0.004613 | 0.004553 | 6.4e-05 | 0.000975 | 0.28368 | 1.3341 | 2.5481 | 99.36 | 21.05 | 37.345 | 99.4058 | 107.0488 | 91.647 | 0.009387 | 0.96779 | 0.023317 | 24.0931 | 0.077695 | 0.15932 | 0.003589 | 0.24591 | 0.24591 | 0.026629
2019-03-21 12:44 | 1553186666.3845 | 559 | 0.20877 | 0.001032 | 9e-05 | 96.907 | 95.9849 | 117.8598 | 21.2585 | 0.003196 | 0.002905 | 219.6223 | 0.26933 | 8.401 | 0.003629 | 0.004832 | 0.004763 | 6.3e-05 | 0.000974 | 0.27887 | 1.3707 | 2.5481 | 99.36 | 21.05 | 37.36 | 98.72 | 109.2182 | 91.6952 | 0.009391 | 0.96766 | 0.02336 | 24.1406 | 0.085513 | 0.16023 | 0.003333 | 0.22842 | 0.22842 | 0.028829
2019-03-21 12:44 | 1553186667.3847 | 560 | 0.20873 | 0.001023 | 8.9e-05 | 96.8388 | 96.2005 | 118.0607 | 21.2574 | 0.003235 | 0.002957 | 219.5987 | 0.26824 | 8.4041 | 0.003607 | 0.004749 | 0.004631 | 6.1e-05 | 0.000975 | 0.31581 | 1.3585 | 2.548 | 99.36 | 21.05 | 37.368 | 99.6632 | 108.6922 | 92.9464 | 0.009393 | 0.96435 | 0.02334 | 24.2026 | 0.059145 | 0.15957 | 0.003587 | 0.24607 | 0.24607 | 0.026653
2019-03-21 12:44 | 1553186668.3846 | 561 | 0.20875 | 0.001016 | 8.8e-05 | 96.8057 | 96.1852 | 118.3679 | 21.2583 | 0.003218 | 0.002877 | 219.5866 | 0.26819 | 8.4014 | 0.003521 | 0.004553 | 0.004563 | 6.1e-05 | 0.000975 | 0.30309 | 1.387 | 2.5481 | 99.36 | 21.05 | 37.374 | 99.2506 | 106.6067 | 91.0422 | 0.009394 | 0.96939 | 0.023431 | 24.1709 | 0.057515 | 0.15955 | 0.003441 | 0.23683 | 0.23683 | 0.027689
2019-03-21 12:44 | 1553186669.3842 | 562 | 0.20875 | 0.00101 | 8.7e-05 | 96.7425 | 96.0881 | 118.3848 | 21.2588 | 0.003156 | 0.0029 | 219.5643 | 0.26723 | 8.404 | 0.003596 | 0.0049 | 0.004787 | 6.2e-05 | 0.000975 | 0.26432 | 1.3077 | 2.5479 | 99.36 | 21.05 | 37.384 | 97.7694 | 108.4186 | 91.5886 | 0.009397 | 0.96794 | 0.023415 | 24.1908 | 0.060708 | 0.15898 | 0.003496 | 0.24056 | 0.24056 | 0.027161
2019-03-21 12:44 | 1553186670.3848 | 563 | 0.20882 | 0.001 | 8.5e-05 | 96.6463 | 96.0333 | 118.2442 | 21.2586 | 0.0032 | 0.002921 | 219.5775 | 0.26914 | 8.4027 | 0.00357 | 0.004945 | 0.004864 | 6.3e-05 | 0.000977 | 0.28191 | 1.353 | 2.548 | 99.36 | 21.05 | 37.389 | 98.8173 | 107.7851 | 92.0949 | 0.009398 | 0.9666 | 0.023385 | 24.1932 | 0.056912 | 0.16011 | 0.003081 | 0.21125 | 0.21125 | 0.031151
2019-03-21 12:44 | 1553186671.3848 | 564 | 0.20879 | 0.000995 | 8.4e-05 | 96.5608 | 95.7632 | 118.079 | 21.2579 | 0.00314 | 0.002904 | 219.5556 | 0.26833 | 8.4036 | 0.003571 | 0.004448 | 0.004462 | 6.5e-05 | 0.000977 | 0.25083 | 1.3225 | 2.5479 | 99.36 | 21.05 | 37.393 | 97.3704 | 107.8056 | 91.6744 | 0.009399 | 0.96771 | 0.023382 | 24.1624 | 0.074186 | 0.15963 | 0.003262 | 0.22401 | 0.22401 | 0.029288
2019-03-21 12:45 | 1553186672.3842 | 565 | 0.20881 | 0.000983 | 8.2e-05 | 96.4583 | 95.9025 | 118.0867 | 21.2584 | 0.003151 | 0.002853 | 219.5509 | 0.26811 | 8.4054 | 0.003569 | 0.004589 | 0.004515 | 6.5e-05 | 0.000979 | 0.25608 | 1.3167 | 2.5479 | 99.36 | 21.05 | 37.395 | 97.6334 | 107.7699 | 90.4618 | 0.0094 | 0.97094 | 0.023422 | 24.123 | 0.051685 | 0.1595 | 0.003165 | 0.21768 | 0.21768 | 0.030114
2019-03-21 12:45 | 1553186673.3839 | 566 | 0.20882 | 0.000984 | 8e-05 | 96.4875 | 95.9148 | 118.1682 | 21.2579 | 0.00309 | 0.00287 | 219.533 | 0.26821 | 8.4039 | 0.003608 | 0.005115 | 0.005042 | 6.8e-05 | 0.00098 | 0.25177 | 1.2862 | 2.548 | 99.36 | 21.05 | 37.396 | 96.1615 | 108.7165 | 90.8723 | 0.0094 | 0.96985 | 0.023417 | 24.1449 | 0.053249 | 0.15956 | 0.00309 | 0.21247 | 0.21247 | 0.030865
2019-03-21 12:45 | 1553186674.3853 | 567 | 0.20885 | 0.000973 | 7.9e-05 | 96.4738 | 95.6794 | 118.402 | 21.2584 | 0.003097 | 0.002896 | 219.5138 | 0.26814 | 8.4019 | 0.00355 | 0.005368 | 0.005202 | 6.4e-05 | 0.000982 | 0.2636 | 1.3195 | 2.5479 | 99.36 | 21.05 | 37.397 | 96.3453 | 107.3075 | 91.4998 | 0.0094 | 0.96818 | 0.02342 | 24.1896 | 0.073953 | 0.15952 | 0.002942 | 0.20216 | 0.20216 | 0.03243
2019-03-21 12:45 | 1553186675.3851 | 568 | 0.20884 | 0.000971 | 7.7e-05 | 96.4264 | 95.6971 | 118.2008 | 21.2586 | 0.003172 | 0.002842 | 219.5137 | 0.2695 | 8.4019 | 0.003588 | 0.004853 | 0.004797 | 6.8e-05 | 0.000982 | 0.29302 | 1.3202 | 2.548 | 99.36 | 21.05 | 37.398 | 98.1485 | 108.2232 | 90.1884 | 0.0094 | 0.97167 | 0.023442 | 24.1255 | 0.067904 | 0.16032 | 0.002995 | 0.2062 | 0.2062 | 0.031956
2019-03-21 12:45 | 1553186676.3851 | 569 | 0.20884 | 0.00097 | 7.6e-05 | 96.3081 | 95.6614 | 118.0238 | 21.2585 | 0.003143 | 0.0029 | 219.5152 | 0.26922 | 8.4048 | 0.003541 | 0.005211 | 0.005178 | 6.5e-05 | 0.000986 | 0.28694 | 1.3418 | 2.5479 | 99.36 | 21.05 | 37.398 | 97.4384 | 107.0918 | 91.5835 | 0.0094 | 0.96795 | 0.02338 | 24.1537 | 0.060268 | 0.16016 | 0.00303 | 0.20818 | 0.20818 | 0.031619
2019-03-21 12:45 | 1553186677.3842 | 570 | 0.20884 | 0.000962 | 7.4e-05 | 96.2497 | 95.4478 | 118.1679 | 21.2576 | 0.003118 | 0.002897 | 219.5399 | 0.26908 | 8.402 | 0.003526 | 0.004606 | 0.004551 | 6.6e-05 | 0.000985 | 0.29927 | 1.3115 | 2.5479 | 99.36 | 21.05 | 37.4 | 96.845 | 106.7264 | 91.5021 | 0.009401 | 0.96817 | 0.023397 | 24.1658 | 0.074832 | 0.16007 | 0.003044 | 0.20945 | 0.20945 | 0.031411
2019-03-21 12:45 | 1553186678.3853 | 571 | 0.20889 | 0.000956 | 7.4e-05 | 96.1997 | 95.3903 | 118.1496 | 21.2582 | 0.003123 | 0.002821 | 219.5362 | 0.26939 | 8.4008 | 0.003565 | 0.004829 | 0.004691 | 6.7e-05 | 0.000984 | 0.26161 | 1.3131 | 2.548 | 99.36 | 21.05 | 37.4 | 96.971 | 107.6539 | 89.6786 | 0.009401 | 0.97304 | 0.023453 | 24.1034 | 0.075575 | 0.16026 | 0.002696 | 0.18544 | 0.18544 | 0.03552
2019-03-21 12:45 | 1553186679.3836 | 572 | 0.20886 | 0.000948 | 7.4e-05 | 96.2184 | 95.5371 | 118.2825 | 21.2582 | 0.003119 | 0.002925 | 219.5482 | 0.26862 | 8.4027 | 0.003557 | 0.0052 | 0.005036 | 6.5e-05 | 0.000983 | 0.25536 | 1.3287 | 2.5481 | 99.36 | 21.05 | 37.402 | 96.8778 | 107.4594 | 92.1931 | 0.009401 | 0.96634 | 0.023386 | 24.2004 | 0.063566 | 0.1598 | 0.002937 | 0.20181 | 0.20181 | 0.032545
2019-03-21 12:45 | 1553186680.3843 | 573 | 0.2089 | 0.000944 | 7.3e-05 | 96.0623 | 95.329 | 118.4555 | 21.2583 | 0.003059 | 0.00286 | 219.5483 | 0.26977 | 8.4052 | 0.003585 | 0.005419 | 0.005439 | 7e-05 | 0.000987 | 0.2656 | 1.2908 | 2.548 | 99.36 | 21.05 | 37.402 | 95.4076 | 108.1407 | 90.617 | 0.009401 | 0.97053 | 0.023454 | 24.1658 | 0.068546 | 0.16049 | 0.002695 | 0.18539 | 0.18539 | 0.03558
2019-03-21 12:45 | 1553186681.3844 | 574 | 0.20889 | 0.000936 | 7.2e-05 | 96.0653 | 95.3409 | 118.5191 | 21.2581 | 0.003089 | 0.002867 | 219.538 | 0.26911 | 8.4032 | 0.003509 | 0.005194 | 0.005126 | 6.7e-05 | 0.000991 | 0.26792 | 1.3078 | 2.548 | 99.36 | 21.05 | 37.403 | 96.1327 | 106.3144 | 90.7912 | 0.009402 | 0.97006 | 0.023454 | 24.178 | 0.067704 | 0.1601 | 0.002733 | 0.18817 | 0.18817 | 0.034969
2019-03-21 12:45 | 1553186682.3847 | 575 | 0.20887 | 0.000932 | 7.2e-05 | 96.0155 | 95.4206 | 118.6419 | 21.2582 | 0.003112 | 0.002837 | 219.5379 | 0.26796 | 8.4054 | 0.003593 | 0.005219 | 0.005146 | 6.6e-05 | 0.000996 | 0.29789 | 1.269 | 2.5481 | 99.36 | 21.05 | 37.403 | 96.7039 | 108.3471 | 90.0716 | 0.009402 | 0.97198 | 0.02349 | 24.1666 | 0.055597 | 0.15941 | 0.002889 | 0.19948 | 0.19948 | 0.032844
2019-03-21 12:45 | 1553186683.3847 | 576 | 0.20893 | 0.000924 | 7.1e-05 | 96.0355 | 95.3944 | 118.6463 | 21.2584 | 0.003033 | 0.002877 | 219.5242 | 0.26932 | 8.4024 | 0.003487 | 0.004771 | 0.004669 | 6.6e-05 | 0.000994 | 0.25514 | 1.2567 | 2.5481 | 99.36 | 21.05 | 37.404 | 94.7987 | 105.7842 | 91.0284 | 0.009402 | 0.96943 | 0.023459 | 24.1989 | 0.059915 | 0.16022 | 0.002515 | 0.17293 | 0.17293 | 0.038079
2019-03-21 12:45 | 1553186684.3852 | 577 | 0.2089 | 0.000913 | 7.1e-05 | 96.0537 | 95.3889 | 118.3672 | 21.2581 | 0.003084 | 0.002869 | 219.5474 | 0.26871 | 8.4034 | 0.00351 | 0.004967 | 0.00486 | 6.6e-05 | 0.000993 | 0.28639 | 1.2991 | 2.548 | 99.36 | 21.05 | 37.403 | 96.0341 | 106.3413 | 90.8296 | 0.009402 | 0.96996 | 0.023438 | 24.1638 | 0.062118 | 0.15986 | 0.002739 | 0.18849 | 0.18849 | 0.034858
2019-03-21 12:45 | 1553186685.3839 | 578 | 0.20888 | 0.000907 | 7.1e-05 | 95.9881 | 95.2053 | 118.5597 | 21.2583 | 0.003105 | 0.002822 | 219.5171 | 0.26881 | 8.4025 | 0.003559 | 0.00526 | 0.005281 | 6.7e-05 | 0.000994 | 0.27936 | 1.3112 | 2.5481 | 99.36 | 21.05 | 37.403 | 96.5229 | 107.5256 | 89.7129 | 0.009402 | 0.97294 | 0.023493 | 24.1463 | 0.073243 | 0.15992 | 0.002883 | 0.19908 | 0.19908 | 0.033014
2019-03-21 12:45 | 1553186686.3836 | 579 | 0.20891 | 0.000901 | 7.1e-05 | 95.908 | 95.2506 | 118.6642 | 21.2581 | 0.003066 | 0.002835 | 219.5016 | 0.26926 | 8.4038 | 0.003538 | 0.004995 | 0.004842 | 6.8e-05 | 0.000994 | 0.2724 | 1.2685 | 2.5478 | 99.36 | 21.05 | 37.402 | 95.5892 | 107.0198 | 90.026 | 0.009401 | 0.9721 | 0.023493 | 24.1674 | 0.061517 | 0.16019 | 0.002648 | 0.18253 | 0.18253 | 0.036068
2019-03-21 12:45 | 1553186687.3847 | 580 | 0.20892 | 0.000896 | 7.1e-05 | 95.793 | 94.9778 | 118.5811 | 21.2584 | 0.003006 | 0.002841 | 219.4773 | 0.26832 | 8.4037 | 0.003518 | 0.005096 | 0.005063 | 6.6e-05 | 0.00099 | 0.28854 | 1.3037 | 2.5481 | 99.36 | 21.05 | 37.398 | 94.1421 | 106.5178 | 90.1694 | 0.0094 | 0.97172 | 0.02348 | 24.1637 | 0.07644 | 0.15962 | 0.002651 | 0.18267 | 0.18267 | 0.035916
2019-03-21 12:45 | 1553186688.3847 | 581 | 0.20889 | 0.000893 | 7.1e-05 | 95.7438 | 94.8337 | 118.5961 | 21.259 | 0.003076 | 0.002809 | 219.49 | 0.26889 | 8.4031 | 0.003533 | 0.005186 | 0.005052 | 6.5e-05 | 0.000993 | 0.2672 | 1.3134 | 2.548 | 99.36 | 21.05 | 37.396 | 95.8413 | 106.8908 | 89.405 | 0.0094 | 0.97377 | 0.023507 | 24.1398 | 0.085426 | 0.15996 | 0.002778 | 0.19182 | 0.19182 | 0.034275
2019-03-21 12:45 | 1553186689.3853 | 582 | 0.2089 | 0.000885 | 7e-05 | 95.7965 | 94.9431 | 118.5632 | 21.2587 | 0.00304 | 0.002827 | 219.4519 | 0.2689 | 8.4018 | 0.00357 | 0.004906 | 0.004801 | 6.2e-05 | 0.000992 | 0.28335 | 1.3413 | 2.548 | 99.36 | 21.05 | 37.39 | 94.9549 | 107.7922 | 89.8259 | 0.009398 | 0.97264 | 0.02349 | 24.1504 | 0.080046 | 0.15997 | 0.00277 | 0.19119 | 0.19119 | 0.034388
2019-03-21 12:45 | 1553186690.3849 | 583 | 0.20891 | 0.000884 | 6.9e-05 | 95.6479 | 94.8685 | 118.6184 | 21.2587 | 0.003055 | 0.002817 | 219.4156 | 0.26876 | 8.402 | 0.003478 | 0.004563 | 0.004546 | 6.5e-05 | 0.000993 | 0.29264 | 1.3244 | 2.5478 | 99.36 | 21.05 | 37.386 | 95.3345 | 105.5627 | 89.589 | 0.009397 | 0.97328 | 0.023503 | 24.1482 | 0.073183 | 0.15988 | 0.002728 | 0.18837 | 0.18837 | 0.034884
2019-03-21 12:45 | 1553186691.3849 | 584 | 0.20889 | 0.000875 | 6.8e-05 | 95.5314 | 94.7421 | 118.8226 | 21.2585 | 0.00301 | 0.002822 | 219.4448 | 0.26858 | 8.4028 | 0.003543 | 0.005249 | 0.005181 | 6.5e-05 | 0.000993 | 0.29358 | 1.3145 | 2.548 | 99.36 | 21.05 | 37.384 | 94.2432 | 107.1408 | 89.6974 | 0.009397 | 0.97299 | 0.02352 | 24.1726 | 0.074204 | 0.15978 | 0.002856 | 0.1976 | 0.1976 | 0.033234
2019-03-21 12:45 | 1553186692.3843 | 585 | 0.2089 | 0.000873 | 6.9e-05 | 95.481 | 94.8767 | 118.5602 | 21.2583 | 0.003003 | 0.002828 | 219.4227 | 0.26749 | 8.4029 | 0.003468 | 0.004704 | 0.0047 | 6.6e-05 | 0.000993 | 0.28694 | 1.2926 | 2.548 | 99.36 | 21.05 | 37.379 | 94.0614 | 105.3176 | 89.859 | 0.009396 | 0.97255 | 0.023488 | 24.1512 | 0.056796 | 0.15913 | 0.002767 | 0.191 | 0.191 | 0.034241
2019-03-21 12:45 | 1553186693.384 | 586 | 0.20892 | 0.000868 | 7e-05 | 95.4328 | 94.6509 | 118.4177 | 21.2585 | 0.00304 | 0.002854 | 219.4588 | 0.26872 | 8.4023 | 0.003495 | 0.005255 | 0.005164 | 6.7e-05 | 0.000991 | 0.28495 | 1.3136 | 2.548 | 99.36 | 21.05 | 37.377 | 94.953 | 105.9701 | 90.4759 | 0.009395 | 0.9709 | 0.023454 | 24.1572 | 0.073587 | 0.15987 | 0.00264 | 0.18175 | 0.18175 | 0.03615
2019-03-21 12:45 | 1553186694.383 | 587 | 0.20887 | 0.000865 | 6.9e-05 | 95.3589 | 94.6154 | 118.528 | 21.2586 | 0.003044 | 0.002764 | 219.4377 | 0.26823 | 8.4014 | 0.003493 | 0.004915 | 0.004809 | 6.5e-05 | 0.000989 | 0.26985 | 1.2695 | 2.548 | 99.35 | 21.05 | 37.374 | 95.0525 | 105.919 | 88.314 | 0.009395 | 0.97671 | 0.023535 | 24.0965 | 0.070007 | 0.15957 | 0.002971 | 0.20579 | 0.20579 | 0.031869
2019-03-21 12:45 | 1553186695.3845 | 588 | 0.20893 | 0.000868 | 6.9e-05 | 95.2704 | 94.3917 | 118.328 | 21.2591 | 0.003006 | 0.002759 | 219.4528 | 0.26852 | 8.4007 | 0.003436 | 0.005012 | 0.004903 | 6.9e-05 | 0.000991 | 0.27024 | 1.2831 | 2.548 | 99.35 | 21.05 | 37.368 | 94.1452 | 104.5421 | 88.1982 | 0.009394 | 0.97702 | 0.023519 | 24.0723 | 0.082882 | 0.15974 | 0.002626 | 0.18135 | 0.18135 | 0.036203
2019-03-21 12:45 | 1553186696.3834 | 589 | 0.20891 | 0.000862 | 6.8e-05 | 95.0654 | 94.3123 | 118.2776 | 21.2587 | 0.003013 | 0.002807 | 219.4601 | 0.26867 | 8.4022 | 0.003572 | 0.005231 | 0.005165 | 6.3e-05 | 0.000993 | 0.28816 | 1.3485 | 2.5481 | 99.35 | 21.05 | 37.365 | 94.3035 | 107.8354 | 89.339 | 0.009393 | 0.97395 | 0.023477 | 24.1052 | 0.071139 | 0.15983 | 0.002761 | 0.19055 | 0.19055 | 0.034474
2019-03-21 12:45 | 1553186697.3843 | 590 | 0.20892 | 0.000857 | 6.7e-05 | 95.0338 | 94.0535 | 118.3339 | 21.2588 | 0.003046 | 0.002804 | 219.4955 | 0.26964 | 8.4036 | 0.003544 | 0.00459 | 0.004592 | 6.3e-05 | 0.000993 | 0.27041 | 1.3548 | 2.5479 | 99.35 | 21.05 | 37.358 | 95.1029 | 107.1481 | 89.2649 | 0.009391 | 0.97415 | 0.023485 | 24.1084 | 0.092749 | 0.16041 | 0.002676 | 0.18471 | 0.18471 | 0.035694
2019-03-21 12:45 | 1553186698.3832 | 591 | 0.20887 | 0.000858 | 6.7e-05 | 95.0466 | 94.2796 | 118.3768 | 21.2592 | 0.00301 | 0.002737 | 219.4838 | 0.26788 | 8.4066 | 0.003526 | 0.004699 | 0.004651 | 6.5e-05 | 0.000993 | 0.27831 | 1.3103 | 2.5479 | 99.35 | 21.05 | 37.356 | 94.2352 | 106.7114 | 87.6656 | 0.009391 | 0.97847 | 0.023541 | 24.0595 | 0.072474 | 0.15936 | 0.002972 | 0.20606 | 0.20606 | 0.031786
2019-03-21 12:45 | 1553186699.3837 | 592 | 0.20893 | 0.000858 | 6.7e-05 | 94.9921 | 94.0674 | 118.4503 | 21.2589 | 0.002937 | 0.002758 | 219.4884 | 0.26851 | 8.4023 | 0.003507 | 0.004642 | 0.004546 | 6.6e-05 | 0.000986 | 0.2792 | 1.3778 | 2.5481 | 99.35 | 21.05 | 37.354 | 92.4683 | 106.2524 | 88.1542 | 0.00939 | 0.97714 | 0.023533 | 24.0833 | 0.087495 | 0.15974 | 0.002613 | 0.18061 | 0.18061 | 0.03635
2019-03-21 12:45 | 1553186700.3844 | 593 | 0.20893 | 0.000855 | 6.7e-05 | 94.8476 | 94.1131 | 118.4487 | 21.2594 | 0.002991 | 0.002776 | 219.5008 | 0.26858 | 8.4023 | 0.003478 | 0.00502 | 0.004952 | 6.5e-05 | 0.000988 | 0.27671 | 1.4017 | 2.5481 | 99.35 | 21.05 | 37.35 | 93.7769 | 105.5665 | 88.5921 | 0.009389 | 0.97596 | 0.023518 | 24.0977 | 0.069539 | 0.15978 | 0.002638 | 0.18227 | 0.18227 | 0.036028
2019-03-21 12:45 | 1553186701.3843 | 594 | 0.20896 | 0.000854 | 6.6e-05 | 94.8258 | 94.0163 | 118.1806 | 21.2586 | 0.002979 | 0.00275 | 219.439 | 0.26815 | 8.4025 | 0.003463 | 0.005135 | 0.005047 | 6.4e-05 | 0.000989 | 0.26681 | 1.3541 | 2.548 | 99.35 | 21.05 | 37.347 | 93.485 | 105.2 | 87.9729 | 0.009389 | 0.97763 | 0.023512 | 24.0498 | 0.076686 | 0.15952 | 0.002458 | 0.16958 | 0.16958 | 0.038661
2019-03-21 12:45 | 1553186702.3839 | 595 | 0.20893 | 0.000849 | 6.5e-05 | 94.6492 | 93.7847 | 118.355 | 21.2585 | 0.002958 | 0.00282 | 219.4701 | 0.26923 | 8.4023 | 0.003461 | 0.004528 | 0.004512 | 6.4e-05 | 0.000989 | 0.27256 | 1.3075 | 2.5482 | 99.35 | 21.05 | 37.341 | 92.989 | 105.1369 | 89.6483 | 0.009387 | 0.97312 | 0.023475 | 24.1233 | 0.082066 | 0.16017 | 0.00262 | 0.18077 | 0.18077 | 0.036416
2019-03-21 12:45 | 1553186703.3837 | 596 | 0.20889 | 0.000848 | 6.5e-05 | 94.5771 | 93.5576 | 118.3482 | 21.2589 | 0.002953 | 0.002729 | 219.4294 | 0.26923 | 8.4006 | 0.003464 | 0.005141 | 0.005015 | 6e-05 | 0.000987 | 0.24906 | 1.3072 | 2.5479 | 99.35 | 21.05 | 37.338 | 92.8642 | 105.2168 | 87.4764 | 0.009386 | 0.97898 | 0.023545 | 24.0503 | 0.096942 | 0.16016 | 0.002852 | 0.19776 | 0.19776 | 0.033287
2019-03-21 12:45 | 1553186704.3849 | 597 | 0.20891 | 0.000844 | 6.5e-05 | 94.4853 | 93.4407 | 118.3975 | 21.259 | 0.002989 | 0.00275 | 219.4369 | 0.26912 | 8.4006 | 0.003342 | 0.005333 | 0.005318 | 6.2e-05 | 0.000985 | 0.26416 | 1.3219 | 2.5479 | 99.35 | 21.06 | 37.335 | 93.7271 | 102.2544 | 87.9709 | 0.009391 | 0.97764 | 0.023534 | 24.0718 | 0.099434 | 0.1601 | 0.002746 | 0.19011 | 0.19011 | 0.034612
2019-03-21 12:45 | 1553186705.3851 | 598 | 0.20896 | 0.000844 | 6.5e-05 | 94.3908 | 93.4564 | 118.3747 | 21.259 | 0.00291 | 0.002748 | 219.4168 | 0.2687 | 8.4042 | 0.003463 | 0.005022 | 0.004922 | 6.4e-05 | 0.000988 | 0.25841 | 1.3565 | 2.548 | 99.35 | 21.06 | 37.329 | 91.8351 | 105.191 | 87.9247 | 0.00939 | 0.97776 | 0.023533 | 24.068 | 0.088989 | 0.15985 | 0.002444 | 0.16882 | 0.16882 | 0.038917
2019-03-21 12:45 | 1553186706.3837 | 599 | 0.20891 | 0.000839 | 6.4e-05 | 94.3616 | 93.3467 | 118.021 | 21.2589 | 0.002951 | 0.002721 | 219.4023 | 0.26943 | 8.3994 | 0.003388 | 0.005429 | 0.005377 | 6.5e-05 | 0.000991 | 0.24961 | 1.3614 | 2.5481 | 99.35 | 21.06 | 37.325 | 92.8071 | 103.3637 | 87.2673 | 0.009389 | 0.97955 | 0.023519 | 24.0101 | 0.096724 | 0.16028 | 0.002804 | 0.19415 | 0.19415 | 0.03393
2019-03-21 12:45 | 1553186707.3838 | 600 | 0.20895 | 0.000837 | 6.3e-05 | 94.3181 | 93.4292 | 117.9098 | 21.2586 | 0.002965 | 0.002761 | 219.3789 | 0.26864 | 8.4034 | 0.003441 | 0.0051 | 0.004952 | 6.2e-05 | 0.000991 | 0.25713 | 1.3075 | 2.5479 | 99.35 | 21.06 | 37.322 | 93.1533 | 104.6643 | 88.2406 | 0.009388 | 0.97691 | 0.023476 | 24.0312 | 0.084706 | 0.15981 | 0.00254 | 0.17524 | 0.17524 | 0.037482
2019-03-21 12:45 | 1553186708.3852 | 601 | 0.20889 | 0.000834 | 6.2e-05 | 94.2535 | 93.7142 | 117.6947 | 21.2591 | 0.002883 | 0.002751 | 219.3661 | 0.26882 | 8.4038 | 0.003415 | 0.004927 | 0.004909 | 6.3e-05 | 0.000992 | 0.25857 | 1.3123 | 2.5479 | 99.35 | 21.06 | 37.321 | 91.1642 | 104.0194 | 87.9967 | 0.009388 | 0.97757 | 0.023463 | 24.0011 | 0.051322 | 0.15992 | 0.002905 | 0.20091 | 0.20091 | 0.032715
2019-03-21 12:45 | 1553186709.3839 | 602 | 0.20897 | 0.000832 | 6.2e-05 | 94.1859 | 93.1632 | 117.5708 | 21.2581 | 0.002917 | 0.002767 | 219.3878 | 0.26863 | 8.4004 | 0.003473 | 0.005245 | 0.005207 | 6.1e-05 | 0.000989 | 0.23308 | 1.3031 | 2.548 | 99.35 | 21.06 | 37.319 | 91.9865 | 105.4435 | 88.3761 | 0.009387 | 0.97654 | 0.023438 | 24.0011 | 0.097654 | 0.15981 | 0.002452 | 0.1689 | 0.1689 | 0.038887
2019-03-21 12:45 | 1553186710.3829 | 603 | 0.20893 | 0.000828 | 6.1e-05 | 94.0765 | 93.1509 | 117.4361 | 21.2589 | 0.002948 | 0.002796 | 219.3576 | 0.26805 | 8.3995 | 0.003453 | 0.005315 | 0.005242 | 6.5e-05 | 0.000988 | 0.27024 | 1.3493 | 2.548 | 99.35 | 21.06 | 37.315 | 92.7531 | 104.9463 | 89.0816 | 0.009386 | 0.97464 | 0.023402 | 24.0107 | 0.088441 | 0.15946 | 0.002667 | 0.18375 | 0.18375 | 0.035668
2019-03-21 12:45 | 1553186711.384 | 604 | 0.20889 | 0.000828 | 6e-05 | 93.9451 | 92.9568 | 117.6046 | 21.2586 | 0.002914 | 0.002747 | 219.4006 | 0.26921 | 8.4003 | 0.00349 | 0.00513 | 0.00511 | 6.3e-05 | 0.000988 | 0.25282 | 1.3277 | 2.5479 | 99.35 | 21.06 | 37.312 | 91.9158 | 105.8428 | 87.8932 | 0.009386 | 0.97785 | 0.023457 | 23.9885 | 0.0946 | 0.16015 | 0.002901 | 0.20076 | 0.20076 | 0.032788
2019-03-21 12:45 | 1553186712.384 | 605 | 0.20892 | 0.000828 | 5.9e-05 | 93.7887 | 92.8699 | 117.6178 | 21.2592 | 0.002961 | 0.002753 | 219.384 | 0.26908 | 8.401 | 0.003462 | 0.005141 | 0.005108 | 6.2e-05 | 0.000988 | 0.24115 | 1.3277 | 2.5482 | 99.35 | 21.06 | 37.31 | 93.0677 | 105.1785 | 88.0358 | 0.009385 | 0.97746 | 0.023454 | 23.9946 | 0.088056 | 0.16007 | 0.002751 | 0.19024 | 0.19024 | 0.034584
2019-03-21 12:45 | 1553186713.3842 | 606 | 0.20893 | 0.000825 | 5.8e-05 | 93.7065 | 92.6751 | 117.6471 | 21.2594 | 0.002915 | 0.002745 | 219.4199 | 0.26791 | 8.4031 | 0.003471 | 0.00485 | 0.004774 | 6.6e-05 | 0.000992 | 0.25299 | 1.2756 | 2.548 | 99.35 | 21.06 | 37.31 | 91.9487 | 105.3816 | 87.8516 | 0.009385 | 0.97796 | 0.023463 | 23.9914 | 0.098993 | 0.15938 | 0.002663 | 0.18413 | 0.18413 | 0.035576
2019-03-21 12:45 | 1553186714.3841 | 607 | 0.20891 | 0.000819 | 5.8e-05 | 93.635 | 92.7037 | 117.7476 | 21.2591 | 0.002913 | 0.002736 | 219.4177 | 0.26824 | 8.4011 | 0.003461 | 0.0045 | 0.00442 | 6.5e-05 | 0.000994 | 0.28655 | 1.305 | 2.548 | 99.35 | 21.05 | 37.306 | 91.9079 | 105.1329 | 87.6405 | 0.009378 | 0.97853 | 0.02348 | 23.9947 | 0.089404 | 0.15958 | 0.002811 | 0.19468 | 0.19468 | 0.033689
2019-03-21 12:45 | 1553186715.3837 | 608 | 0.2089 | 0.000823 | 5.7e-05 | 93.5527 | 92.5395 | 118.0388 | 21.259 | 0.002976 | 0.002763 | 219.4133 | 0.26907 | 8.4014 | 0.003427 | 0.00497 | 0.00488 | 6.5e-05 | 0.000994 | 0.28456 | 1.3475 | 2.5481 | 99.35 | 21.05 | 37.303 | 93.4213 | 104.3126 | 88.2802 | 0.009377 | 0.9768 | 0.023488 | 24.0456 | 0.097399 | 0.16007 | 0.002849 | 0.19744 | 0.19744 | 0.03332
2019-03-21 12:45 | 1553186716.3842 | 609 | 0.20895 | 0.000819 | 5.8e-05 | 93.416 | 92.3839 | 118.4175 | 21.2595 | 0.002959 | 0.002716 | 219.4221 | 0.26859 | 8.4024 | 0.003513 | 0.00489 | 0.004762 | 6.6e-05 | 0.000996 | 0.27505 | 1.343 | 2.548 | 99.35 | 21.05 | 37.302 | 93.0008 | 106.4072 | 87.1656 | 0.009377 | 0.97982 | 0.023562 | 24.047 | 0.099371 | 0.15979 | 0.002566 | 0.17806 | 0.17806 | 0.036882
2019-03-21 12:45 | 1553186717.3833 | 610 | 0.20896 | 0.000821 | 5.8e-05 | 93.4818 | 92.4915 | 118.2618 | 21.2589 | 0.002946 | 0.002732 | 219.4076 | 0.26893 | 8.4024 | 0.003458 | 0.004803 | 0.004663 | 6.3e-05 | 0.000996 | 0.28296 | 1.3012 | 2.548 | 99.35 | 21.05 | 37.299 | 92.6921 | 105.0665 | 87.5328 | 0.009376 | 0.97882 | 0.023534 | 24.0434 | 0.095259 | 0.15998 | 0.002479 | 0.17172 | 0.17172 | 0.038292
2019-03-21 12:45 | 1553186718.3836 | 611 | 0.20898 | 0.000823 | 5.8e-05 | 93.3755 | 92.5968 | 118.4625 | 21.2591 | 0.002907 | 0.002702 | 219.4301 | 0.26862 | 8.4037 | 0.003525 | 0.005295 | 0.005312 | 6.5e-05 | 0.000999 | 0.29247 | 1.317 | 2.5479 | 99.35 | 21.06 | 37.297 | 91.7553 | 106.6878 | 86.8059 | 0.009382 | 0.9808 | 0.023578 | 24.0396 | 0.074904 | 0.1598 | 0.002406 | 0.16685 | 0.16685 | 0.039365
2019-03-21 12:45 | 1553186719.3843 | 612 | 0.20892 | 0.00082 | 5.8e-05 | 93.273 | 92.4452 | 118.4346 | 21.259 | 0.002901 | 0.00275 | 219.3982 | 0.26822 | 8.4008 | 0.003448 | 0.005145 | 0.00506 | 6.4e-05 | 0.000998 | 0.26488 | 1.3185 | 2.548 | 99.35 | 21.06 | 37.294 | 91.6128 | 104.829 | 87.9754 | 0.009381 | 0.97762 | 0.023537 | 24.0757 | 0.079733 | 0.15957 | 0.002747 | 0.1906 | 0.1906 | 0.034409
2019-03-21 12:45 | 1553186720.3834 | 613 | 0.20901 | 0.000815 | 5.8e-05 | 93.1082 | 92.0535 | 118.4576 | 21.2591 | 0.002919 | 0.002721 | 219.4331 | 0.26835 | 8.4055 | 0.003433 | 0.005791 | 0.005646 | 6.8e-05 | 0.000997 | 0.2385 | 1.2989 | 2.5478 | 99.35 | 21.06 | 37.292 | 92.0352 | 104.4745 | 87.286 | 0.009381 | 0.9795 | 0.023562 | 24.0551 | 0.1019 | 0.15964 | 0.002203 | 0.15235 | 0.15235 | 0.043065
2019-03-21 12:45 | 1553186721.3845 | 614 | 0.20891 | 0.000815 | 5.9e-05 | 93.012 | 92.0651 | 118.3097 | 21.2595 | 0.002928 | 0.002713 | 219.4038 | 0.26794 | 8.4045 | 0.003451 | 0.005105 | 0.005125 | 7e-05 | 0.000998 | 0.28208 | 1.3083 | 2.5478 | 99.35 | 21.06 | 37.291 | 92.2694 | 104.9059 | 87.0912 | 0.00938 | 0.98003 | 0.023553 | 24.0336 | 0.091521 | 0.1594 | 0.002792 | 0.19389 | 0.19389 | 0.033788
2019-03-21 12:45 | 1553186722.3846 | 615 | 0.20901 | 0.000809 | 5.9e-05 | 92.8791 | 91.8853 | 118.0405 | 21.2594 | 0.002904 | 0.002767 | 219.3954 | 0.26947 | 8.4015 | 0.003417 | 0.004782 | 0.004684 | 7.1e-05 | 0.000998 | 0.29999 | 1.3145 | 2.5481 | 99.36 | 21.06 | 37.29 | 91.6853 | 104.0864 | 88.3906 | 0.009379 | 0.9765 | 0.023484 | 24.0495 | 0.096222 | 0.16031 | 0.002195 | 0.15125 | 0.15125 | 0.04356
2019-03-21 12:45 | 1553186723.3846 | 616 | 0.20889 | 0.000811 | 5.9e-05 | 92.723 | 91.7143 | 117.9874 | 21.2593 | 0.0029 | 0.002709 | 219.3459 | 0.26902 | 8.4013 | 0.003408 | 0.005076 | 0.005068 | 6.9e-05 | 0.000999 | 0.27743 | 1.322 | 2.548 | 99.36 | 21.06 | 37.29 | 91.5922 | 103.8646 | 86.9811 | 0.009379 | 0.98033 | 0.023525 | 23.9971 | 0.097833 | 0.16004 | 0.002932 | 0.20357 | 0.20357 | 0.032311
2019-03-21 12:45 | 1553186724.3846 | 617 | 0.20904 | 0.000809 | 5.9e-05 | 92.6309 | 91.6349 | 117.8407 | 21.2596 | 0.002945 | 0.002728 | 219.3294 | 0.26966 | 8.4034 | 0.003421 | 0.004626 | 0.004519 | 7e-05 | 0.001001 | 0.27455 | 1.3319 | 2.548 | 99.36 | 21.06 | 37.289 | 92.6722 | 104.1842 | 87.4419 | 0.009379 | 0.97907 | 0.023495 | 23.9976 | 0.09669 | 0.16042 | 0.002028 | 0.13954 | 0.13954 | 0.047249
2019-03-21 12:45 | 1553186725.3837 | 618 | 0.20899 | 0.000811 | 5.9e-05 | 92.4694 | 91.4132 | 117.6065 | 21.2593 | 0.002899 | 0.002726 | 219.3129 | 0.26856 | 8.4009 | 0.003376 | 0.004366 | 0.004378 | 6.6e-05 | 0.001 | 0.31277 | 1.3228 | 2.548 | 99.36 | 21.06 | 37.289 | 91.5514 | 103.0731 | 87.4009 | 0.009379 | 0.97918 | 0.023473 | 23.9723 | 0.10275 | 0.15976 | 0.002308 | 0.15914 | 0.15914 | 0.04126
2019-03-21 12:45 | 1553186726.3828 | 619 | 0.20899 | 0.000813 | 5.9e-05 | 92.4393 | 91.2712 | 117.6201 | 21.2594 | 0.002913 | 0.002628 | 219.3063 | 0.26796 | 8.4014 | 0.003357 | 0.004814 | 0.004724 | 6.9e-05 | 0.000998 | 0.28854 | 1.3124 | 2.548 | 99.36 | 21.06 | 37.29 | 91.8901 | 102.6263 | 85.0351 | 0.009379 | 0.98565 | 0.023552 | 23.8949 | 0.11375 | 0.15941 | 0.002307 | 0.15961 | 0.15961 | 0.04105
2019-03-21 12:45 | 1553186727.385 | 620 | 0.209 | 0.000806 | 5.9e-05 | 92.3002 | 91.091 | 117.9276 | 21.259 | 0.002923 | 0.002713 | 219.2978 | 0.26898 | 8.4039 | 0.00343 | 0.005405 | 0.005278 | 7e-05 | 0.000994 | 0.25227 | 1.3124 | 2.5481 | 99.35 | 21.06 | 37.292 | 92.1469 | 104.3942 | 87.0939 | 0.009381 | 0.98002 | 0.023515 | 23.9948 | 0.11795 | 0.16002 | 0.002299 | 0.15877 | 0.15877 | 0.041422
2019-03-21 12:45 | 1553186728.3828 | 621 | 0.209 | 0.000802 | 5.9e-05 | 92.077 | 90.9783 | 118.0635 | 21.2592 | 0.002866 | 0.002714 | 219.2674 | 0.26949 | 8.4024 | 0.003472 | 0.004902 | 0.00486 | 6.7e-05 | 0.000986 | 0.25265 | 1.3531 | 2.548 | 99.35 | 21.06 | 37.295 | 90.7634 | 105.4215 | 87.1027 | 0.009381 | 0.97999 | 0.023529 | 24.0089 | 0.10737 | 0.16032 | 0.002256 | 0.1558 | 0.1558 | 0.042291
2019-03-21 12:45 | 1553186729.3846 | 622 | 0.20896 | 0.000805 | 6e-05 | 91.979 | 90.6759 | 118.2355 | 21.2595 | 0.002894 | 0.002685 | 219.2836 | 0.26918 | 8.4033 | 0.003571 | 0.004386 | 0.004413 | 6.5e-05 | 0.000985 | 0.27527 | 1.3203 | 2.548 | 99.35 | 21.06 | 37.295 | 91.4304 | 107.8095 | 86.4028 | 0.009381 | 0.9819 | 0.023569 | 24.0031 | 0.12763 | 0.16014 | 0.002495 | 0.17296 | 0.17296 | 0.038053
2019-03-21 12:45 | 1553186730.3829 | 623 | 0.20902 | 0.0008 | 6.1e-05 | 91.7829 | 90.4998 | 118.3597 | 21.2593 | 0.002874 | 0.002654 | 219.2594 | 0.26822 | 8.4042 | 0.003514 | 0.005357 | 0.005248 | 6.6e-05 | 0.000985 | 0.27505 | 1.2792 | 2.5479 | 99.35 | 21.06 | 37.294 | 90.9541 | 106.4369 | 85.664 | 0.009381 | 0.98392 | 0.023605 | 23.991 | 0.12593 | 0.15956 | 0.002182 | 0.151 | 0.151 | 0.043431
2019-03-21 12:46 | 1553186731.383 | 624 | 0.20903 | 0.0008 | 6.2e-05 | 91.6861 | 90.6299 | 118.5388 | 21.2595 | 0.002847 | 0.002713 | 219.2804 | 0.26926 | 8.4027 | 0.003532 | 0.004624 | 0.004556 | 6.5e-05 | 0.000985 | 0.2975 | 1.2808 | 2.5479 | 99.35 | 21.06 | 37.293 | 90.2971 | 106.8636 | 87.0874 | 0.009381 | 0.98003 | 0.023576 | 24.0567 | 0.10364 | 0.16019 | 0.002085 | 0.14388 | 0.14388 | 0.045757
2019-03-21 12:46 | 1553186732.3841 | 625 | 0.20905 | 0.000798 | 6.3e-05 | 91.5705 | 90.324 | 118.3564 | 21.2592 | 0.002868 | 0.002739 | 219.271 | 0.26879 | 8.4014 | 0.003474 | 0.005117 | 0.005072 | 6.7e-05 | 0.000984 | 0.26936 | 1.3073 | 2.548 | 99.35 | 21.06 | 37.293 | 90.8234 | 105.4663 | 87.6986 | 0.009381 | 0.97837 | 0.023538 | 24.0586 | 0.12259 | 0.1599 | 0.00201 | 0.1383 | 0.1383 | 0.047519
2019-03-21 12:46 | 1553186733.3835 | 626 | 0.20902 | 0.0008 | 6.4e-05 | 91.3715 | 89.9879 | 118.4219 | 21.2595 | 0.002853 | 0.00269 | 219.2867 | 0.26893 | 8.399 | 0.003562 | 0.004752 | 0.004696 | 6.4e-05 | 0.000983 | 0.26864 | 1.3223 | 2.548 | 99.35 | 21.06 | 37.291 | 90.4424 | 107.5835 | 86.5301 | 0.00938 | 0.98155 | 0.023583 | 24.0262 | 0.13648 | 0.15998 | 0.002135 | 0.14735 | 0.14735 | 0.044626
2019-03-21 12:46 | 1553186734.3832 | 627 | 0.20894 | 0.00079 | 6.5e-05 | 91.3441 | 90.2913 | 118.4595 | 21.2592 | 0.002882 | 0.002724 | 219.2601 | 0.26794 | 8.4029 | 0.003574 | 0.005216 | 0.005146 | 6.4e-05 | 0.000981 | 0.26593 | 1.3066 | 2.5479 | 99.35 | 21.06 | 37.289 | 91.1463 | 107.8921 | 87.3574 | 0.00938 | 0.9793 | 0.02356 | 24.0577 | 0.10368 | 0.1594 | 0.002686 | 0.18609 | 0.18609 | 0.035205
2019-03-21 12:46 | 1553186735.3842 | 628 | 0.20911 | 0.000792 | 6.7e-05 | 91.1415 | 89.8465 | 118.39 | 21.2597 | 0.002871 | 0.002663 | 219.2911 | 0.26774 | 8.3997 | 0.003476 | 0.004978 | 0.004888 | 6.6e-05 | 0.000977 | 0.25785 | 1.3136 | 2.5481 | 99.35 | 21.07 | 37.288 | 90.8911 | 105.5117 | 85.8838 | 0.009385 | 0.98332 | 0.023601 | 24.0014 | 0.128 | 0.15928 | 0.001631 | 0.1116 | 0.1116 | 0.058657
2019-03-21 12:46 | 1553186736.3846 | 629 | 0.20902 | 0.000787 | 6.7e-05 | 91.0321 | 89.6561 | 118.3299 | 21.2595 | 0.002911 | 0.002663 | 219.289 | 0.26885 | 8.4039 | 0.003599 | 0.004967 | 0.004925 | 6.3e-05 | 0.000982 | 0.274 | 1.3118 | 2.548 | 99.35 | 21.07 | 37.287 | 91.8418 | 108.4839 | 85.8711 | 0.009385 | 0.98335 | 0.023595 | 23.9949 | 0.13623 | 0.15994 | 0.002201 | 0.1519 | 0.1519 | 0.043274
2019-03-21 12:46 | 1553186737.3842 | 630 | 0.20893 | 0.000786 | 6.8e-05 | 90.9284 | 89.4644 | 118.1823 | 21.2595 | 0.002881 | 0.002684 | 219.2931 | 0.26907 | 8.4028 | 0.003532 | 0.004925 | 0.004874 | 6.6e-05 | 0.000979 | 0.28208 | 1.3356 | 2.548 | 99.35 | 21.07 | 37.283 | 91.1319 | 106.8661 | 86.3966 | 0.009384 | 0.98192 | 0.023564 | 23.9974 | 0.14519 | 0.16007 | 0.00274 | 0.18978 | 0.18978 | 0.034666
2019-03-21 12:46 | 1553186738.385 | 631 | 0.20911 | 0.000783 | 6.9e-05 | 90.7702 | 89.6439 | 117.9621 | 21.2596 | 0.002858 | 0.002702 | 219.2837 | 0.26894 | 8.4026 | 0.003567 | 0.00546 | 0.005316 | 6.6e-05 | 0.000981 | 0.28982 | 1.386 | 2.548 | 99.35 | 21.07 | 37.28 | 90.5738 | 107.7121 | 86.8206 | 0.009383 | 0.98076 | 0.023528 | 23.9892 | 0.11169 | 0.15999 | 0.001623 | 0.11056 | 0.11056 | 0.059479
2019-03-21 12:46 | 1553186739.3839 | 632 | 0.20898 | 0.000779 | 6.9e-05 | 90.5298 | 89.4024 | 117.8506 | 21.2596 | 0.00284 | 0.002683 | 219.2902 | 0.26995 | 8.4044 | 0.003551 | 0.005238 | 0.005218 | 7.1e-05 | 0.000978 | 0.28368 | 1.3784 | 2.548 | 99.35 | 21.07 | 37.277 | 90.1291 | 107.3307 | 86.3622 | 0.009383 | 0.98201 | 0.023532 | 23.9626 | 0.11209 | 0.16059 | 0.002434 | 0.16783 | 0.16783 | 0.039327
2019-03-21 12:46 | 1553186740.3841 | 633 | 0.20892 | 0.00077 | 7e-05 | 90.3582 | 89.1735 | 117.7274 | 21.2599 | 0.002848 | 0.002712 | 219.2828 | 0.26774 | 8.4053 | 0.003506 | 0.005071 | 0.004973 | 6.7e-05 | 0.000979 | 0.25912 | 1.3459 | 2.5478 | 99.35 | 21.07 | 37.27 | 90.3428 | 106.2383 | 87.0557 | 0.009381 | 0.98012 | 0.023497 | 23.9732 | 0.11804 | 0.15928 | 0.002778 | 0.1918 | 0.1918 | 0.034132
2019-03-21 12:46 | 1553186741.3835 | 634 | 0.20902 | 0.000772 | 7e-05 | 90.2476 | 88.8418 | 117.7586 | 21.2594 | 0.002856 | 0.002682 | 219.2646 | 0.2693 | 8.4029 | 0.003468 | 0.005241 | 0.005199 | 6.7e-05 | 0.000981 | 0.30558 | 1.3148 | 2.5479 | 99.35 | 21.07 | 37.267 | 90.522 | 105.3053 | 86.337 | 0.00938 | 0.98208 | 0.023523 | 23.9524 | 0.14043 | 0.1602 | 0.002218 | 0.15246 | 0.15246 | 0.043187
2019-03-21 12:46 | 1553186742.383 | 635 | 0.20921 | 0.000771 | 7e-05 | 90.0194 | 88.7767 | 118.271 | 21.2593 | 0.002862 | 0.002644 | 219.2731 | 0.26854 | 8.4047 | 0.003477 | 0.00439 | 0.004327 | 6.6e-05 | 0.000981 | 0.27721 | 1.3247 | 2.5481 | 99.35 | 21.07 | 37.264 | 90.6727 | 105.5428 | 85.4307 | 0.009379 | 0.98456 | 0.023604 | 23.9742 | 0.12434 | 0.15975 | 0.001019 | 0.068058 | 0.068058 | 0.096474
2019-03-21 12:46 | 1553186743.3827 | 636 | 0.20886 | 0.00077 | 7.1e-05 | 89.9008 | 88.4814 | 118.3029 | 21.2598 | 0.002856 | 0.002623 | 219.2588 | 0.26822 | 8.4034 | 0.0035 | 0.004866 | 0.004844 | 6.8e-05 | 0.000983 | 0.27079 | 1.2839 | 2.548 | 99.35 | 21.07 | 37.258 | 90.5147 | 106.0871 | 84.928 | 0.009378 | 0.98594 | 0.023624 | 23.9606 | 0.14235 | 0.15957 | 0.003173 | 0.22077 | 0.22077 | 0.029707
2019-03-21 12:46 | 1553186744.3841 | 637 | 0.20905 | 0.000768 | 7.1e-05 | 89.8657 | 88.5353 | 118.344 | 21.2595 | 0.002851 | 0.002632 | 219.2708 | 0.26809 | 8.4067 | 0.003431 | 0.005048 | 0.004972 | 6.9e-05 | 0.000981 | 0.24441 | 1.3186 | 2.5479 | 99.35 | 21.06 | 37.255 | 90.4016 | 104.4136 | 85.1304 | 0.009371 | 0.98539 | 0.023621 | 23.9715 | 0.13341 | 0.15949 | 0.002043 | 0.14063 | 0.14063 | 0.046611
2019-03-21 12:46 | 1553186745.3828 | 638 | 0.20916 | 0.000767 | 7.2e-05 | 89.7697 | 88.4814 | 118.5005 | 21.2596 | 0.002814 | 0.00268 | 219.236 | 0.26961 | 8.4008 | 0.003435 | 0.00478 | 0.004726 | 6.8e-05 | 0.000978 | 0.30718 | 1.296 | 2.5481 | 99.35 | 21.06 | 37.249 | 89.5173 | 104.5205 | 86.2826 | 0.00937 | 0.98223 | 0.023599 | 24.0259 | 0.12929 | 0.16039 | 0.00133 | 0.089956 | 0.089956 | 0.073281
2019-03-21 12:46 | 1553186746.3828 | 639 | 0.20926 | 0.000759 | 7.3e-05 | 89.8856 | 88.4685 | 118.6186 | 21.2595 | 0.002801 | 0.002705 | 219.2387 | 0.26925 | 8.4018 | 0.003456 | 0.005378 | 0.005239 | 6.5e-05 | 0.000978 | 0.25608 | 1.2613 | 2.548 | 99.35 | 21.06 | 37.246 | 89.1992 | 105.0286 | 86.8992 | 0.009369 | 0.98055 | 0.023591 | 24.0585 | 0.14214 | 0.16017 | 0.000738 | 0.047992 | 0.047992 | 0.13717
2019-03-21 12:46 | 1553186747.3832 | 640 | 0.20897 | 0.00076 | 7.3e-05 | 89.9384 | 88.6773 | 118.6003 | 21.2597 | 0.002869 | 0.002673 | 219.2233 | 0.26867 | 8.402 | 0.003439 | 0.005158 | 0.005052 | 6.6e-05 | 0.000974 | 0.30342 | 1.2768 | 2.5482 | 99.35 | 21.06 | 37.242 | 90.8444 | 104.6166 | 86.1188 | 0.009368 | 0.98268 | 0.023614 | 24.0306 | 0.12631 | 0.15983 | 0.002492 | 0.17228 | 0.17228 | 0.038129
2019-03-21 12:46 | 1553186748.384 | 641 | 0.20841 | 0.000759 | 7.4e-05 | 90.085 | 88.5223 | 118.4193 | 21.2593 | 0.00293 | 0.002704 | 219.2211 | 0.26869 | 8.4035 | 0.003452 | 0.004315 | 0.004289 | 6.6e-05 | 0.000972 | 0.29031 | 1.2882 | 2.5481 | 99.35 | 21.06 | 37.237 | 92.3116 | 104.9144 | 86.8756 | 0.009367 | 0.98061 | 0.023571 | 24.0375 | 0.15652 | 0.15984 | 0.005903 | 0.41314 | 0.41314 | 0.015902
2019-03-21 12:46 | 1553186749.3829 | 642 | 0.20898 | 0.000756 | 7.4e-05 | 90.3972 | 89.0771 | 118.1419 | 21.259 | 0.002876 | 0.002664 | 219.2116 | 0.26858 | 8.4017 | 0.003425 | 0.004892 | 0.004823 | 7.2e-05 | 0.00097 | 0.31509 | 1.3079 | 2.5481 | 99.35 | 21.06 | 37.234 | 91.0023 | 104.2653 | 85.912 | 0.009366 | 0.98324 | 0.023575 | 23.9772 | 0.13158 | 0.15978 | 0.002437 | 0.16809 | 0.16809 | 0.039068
2019-03-21 12:46 | 1553186750.384 | 643 | 0.20898 | 0.000754 | 7.4e-05 | 90.5378 | 89.1711 | 118.1417 | 21.2598 | 0.00283 | 0.002646 | 219.2079 | 0.26942 | 8.4014 | 0.003464 | 0.004627 | 0.004514 | 6.9e-05 | 0.000968 | 0.26394 | 1.3611 | 2.5481 | 99.35 | 21.06 | 37.23 | 89.8932 | 105.2184 | 85.4766 | 0.009365 | 0.98444 | 0.02359 | 23.9626 | 0.13605 | 0.16028 | 0.002475 | 0.17089 | 0.17089 | 0.038547
2019-03-21 12:46 | 1553186751.3844 | 644 | 0.2084 | 0.000752 | 7.4e-05 | 90.7753 | 89.3278 | 118.1245 | 21.2592 | 0.002824 | 0.002641 | 219.2027 | 0.2691 | 8.4045 | 0.003581 | 0.004975 | 0.004905 | 7e-05 | 0.000967 | 0.29358 | 1.3575 | 2.548 | 99.35 | 21.06 | 37.228 | 89.7653 | 108.0511 | 85.3519 | 0.009364 | 0.98478 | 0.023592 | 23.9567 | 0.14378 | 0.16009 | 0.005926 | 0.41513 | 0.41513 | 0.015849
2019-03-21 12:46 | 1553186752.3844 | 645 | 0.20902 | 0.000755 | 7.3e-05 | 90.9675 | 89.5805 | 117.8574 | 21.2592 | 0.002832 | 0.002608 | 219.1629 | 0.26793 | 8.4055 | 0.003485 | 0.0043 | 0.004361 | 6.8e-05 | 0.00097 | 0.27024 | 1.319 | 2.5479 | 99.35 | 21.06 | 37.227 | 89.942 | 105.7142 | 84.555 | 0.009364 | 0.98697 | 0.023592 | 23.903 | 0.13742 | 0.15939 | 0.002212 | 0.15229 | 0.15229 | 0.043017
2019-03-21 12:46 | 1553186753.3833 | 646 | 0.20964 | 0.000749 | 7.3e-05 | 91.1707 | 89.8649 | 117.6936 | 21.2593 | 0.002786 | 0.002648 | 219.1805 | 0.26889 | 8.4031 | 0.003501 | 0.004951 | 0.004889 | 6.5e-05 | 0.000972 | 0.30989 | 1.3186 | 2.5481 | 99.35 | 21.06 | 37.225 | 88.8381 | 106.1143 | 85.5166 | 0.009364 | 0.98433 | 0.023544 | 23.9185 | 0.12903 | 0.15996 | -0.001522 | -0.1119 | -0.1119 | -0.058751
2019-03-21 12:46 | 1553186754.3832 | 647 | 0.20934 | 0.00075 | 7.3e-05 | 91.3575 | 90.1687 | 117.6534 | 21.2595 | 0.002789 | 0.00266 | 219.1611 | 0.26819 | 8.4022 | 0.003471 | 0.005199 | 0.005129 | 6.5e-05 | 0.000968 | 0.27815 | 1.3067 | 2.548 | 99.35 | 21.06 | 37.226 | 88.9227 | 105.3856 | 85.8095 | 0.009364 | 0.98352 | 0.02353 | 23.9241 | 0.11715 | 0.15955 | 0.000299 | 0.016766 | 0.016766 | 0.39112
2019-03-21 12:46 | 1553186755.383 | 648 | 0.20914 | 0.000746 | 7.3e-05 | 91.6507 | 90.5314 | 117.707 | 21.2598 | 0.002834 | 0.00265 | 219.1674 | 0.26754 | 8.4028 | 0.003427 | 0.004816 | 0.004759 | 6.7e-05 | 0.000966 | 0.28854 | 1.3225 | 2.5481 | 99.35 | 21.06 | 37.225 | 90.0043 | 104.3203 | 85.5772 | 0.009364 | 0.98416 | 0.023543 | 23.9218 | 0.1099 | 0.15916 | 0.001504 | 0.10197 | 0.10197 | 0.064149
2019-03-21 12:46 | 1553186756.3824 | 649 | 0.20863 | 0.000749 | 7.3e-05 | 91.9394 | 90.7593 | 117.6175 | 21.2595 | 0.002833 | 0.002681 | 219.1381 | 0.26793 | 8.4013 | 0.003566 | 0.00505 | 0.004925 | 6.4e-05 | 0.000958 | 0.2724 | 1.2882 | 2.5481 | 99.35 | 21.06 | 37.224 | 89.9766 | 107.6906 | 86.3044 | 0.009363 | 0.98217 | 0.02351 | 23.937 | 0.11555 | 0.15939 | 0.004585 | 0.31917 | 0.31917 | 0.020524
2019-03-21 12:46 | 1553186757.3833 | 650 | 0.20923 | 0.000745 | 7.2e-05 | 92.1715 | 90.9112 | 117.8081 | 21.2593 | 0.002789 | 0.002644 | 219.1435 | 0.26822 | 8.4014 | 0.003497 | 0.005087 | 0.005099 | 6.5e-05 | 0.000958 | 0.27057 | 1.3617 | 2.5479 | 99.35 | 21.07 | 37.225 | 88.9226 | 106.0277 | 85.423 | 0.009369 | 0.98458 | 0.023558 | 23.927 | 0.12314 | 0.15956 | 0.000972 | 0.064489 | 0.064489 | 0.10169
2019-03-21 12:46 | 1553186758.3824 | 651 | 0.20906 | 0.000741 | 7.2e-05 | 92.2409 | 91.2256 | 117.8494 | 21.2594 | 0.002817 | 0.002666 | 219.1351 | 0.26876 | 8.4027 | 0.003578 | 0.004524 | 0.004461 | 6.1e-05 | 0.000955 | 0.28136 | 1.326 | 2.5482 | 99.35 | 21.07 | 37.225 | 89.5827 | 107.9702 | 85.96 | 0.009369 | 0.98311 | 0.023545 | 23.9491 | 0.099001 | 0.15988 | 0.001984 | 0.13596 | 0.13596 | 0.048334
2019-03-21 12:46 | 1553186759.3826 | 652 | 0.20908 | 0.00074 | 7.2e-05 | 92.42 | 91.4984 | 118.1125 | 21.2598 | 0.002803 | 0.002643 | 219.1271 | 0.26808 | 8.4022 | 0.003533 | 0.005276 | 0.005248 | 6.2e-05 | 0.000958 | 0.28031 | 1.3381 | 2.5481 | 99.35 | 21.07 | 37.225 | 89.2377 | 106.894 | 85.397 | 0.009369 | 0.98466 | 0.023589 | 23.957 | 0.08964 | 0.15948 | 0.001859 | 0.12739 | 0.12739 | 0.051454
2019-03-21 12:46 | 1553186760.3827 | 653 | 0.20893 | 0.00074 | 7.2e-05 | 92.6452 | 91.6976 | 118.1412 | 21.2592 | 0.002813 | 0.002653 | 219.0973 | 0.26884 | 8.4051 | 0.003525 | 0.005343 | 0.005278 | 6.2e-05 | 0.000961 | 0.30879 | 1.3166 | 2.548 | 99.35 | 21.07 | 37.224 | 89.4958 | 106.6862 | 85.6493 | 0.009369 | 0.98396 | 0.023584 | 23.9683 | 0.091957 | 0.15993 | 0.0028 | 0.19392 | 0.19392 | 0.033897
2019-03-21 12:46 | 1553186761.3832 | 654 | 0.2091 | 0.000737 | 7.2e-05 | 92.8663 | 91.8168 | 118.0424 | 21.2596 | 0.002784 | 0.002628 | 219.1136 | 0.26842 | 8.4012 | 0.003508 | 0.005329 | 0.005293 | 5.8e-05 | 0.00095 | 0.29551 | 1.318 | 2.5479 | 99.35 | 21.07 | 37.222 | 88.7955 | 106.2742 | 85.0315 | 0.009369 | 0.98566 | 0.023594 | 23.9377 | 0.10165 | 0.15968 | 0.001776 | 0.12151 | 0.12151 | 0.05401
2019-03-21 12:46 | 1553186762.3836 | 655 | 0.20909 | 0.000733 | 7.3e-05 | 93.0522 | 92.0535 | 118.3083 | 21.2594 | 0.002802 | 0.002622 | 219.1376 | 0.26747 | 8.4036 | 0.003515 | 0.005368 | 0.005226 | 6.2e-05 | 0.00095 | 0.26449 | 1.3102 | 2.5479 | 99.35 | 21.07 | 37.22 | 89.2265 | 106.4503 | 84.8925 | 0.009368 | 0.98604 | 0.023626 | 23.96 | 0.096514 | 0.15912 | 0.001829 | 0.12536 | 0.12536 | 0.052168
2019-03-21 12:46 | 1553186763.3836 | 656 | 0.20916 | 0.000735 | 7.3e-05 | 93.1782 | 92.0563 | 118.2447 | 21.2598 | 0.00278 | 0.002661 | 219.1098 | 0.26823 | 8.4072 | 0.003537 | 0.004592 | 0.004574 | 6.1e-05 | 0.000955 | 0.27256 | 1.312 | 2.5479 | 99.35 | 21.07 | 37.214 | 88.6895 | 106.9916 | 85.8247 | 0.009367 | 0.98348 | 0.023589 | 23.9847 | 0.10836 | 0.15957 | 0.001391 | 0.094165 | 0.094165 | 0.069647
2019-03-21 12:46 | 1553186764.3842 | 657 | 0.20895 | 0.000736 | 7.4e-05 | 93.3343 | 92.3161 | 118.3346 | 21.26 | 0.002824 | 0.002638 | 219.1296 | 0.2687 | 8.4025 | 0.003495 | 0.004595 | 0.004534 | 6e-05 | 0.000959 | 0.27704 | 1.2879 | 2.548 | 99.35 | 21.07 | 37.212 | 89.7451 | 105.9638 | 85.287 | 0.009366 | 0.98496 | 0.023615 | 23.9758 | 0.098115 | 0.15985 | 0.002643 | 0.18302 | 0.18302 | 0.035897
2019-03-21 12:46 | 1553186765.3833 | 658 | 0.20902 | 0.000728 | 7.3e-05 | 93.517 | 92.4006 | 118.3193 | 21.2595 | 0.002773 | 0.002669 | 219.1272 | 0.26972 | 8.4038 | 0.003517 | 0.004899 | 0.00476 | 6e-05 | 0.00096 | 0.30702 | 1.3024 | 2.5478 | 99.35 | 21.07 | 37.21 | 88.5351 | 106.4936 | 86.0179 | 0.009366 | 0.98295 | 0.02359 | 23.9987 | 0.10742 | 0.16046 | 0.002275 | 0.15676 | 0.15676 | 0.042069
2019-03-21 12:46 | 1553186766.3829 | 659 | 0.20905 | 0.000732 | 7.4e-05 | 93.6445 | 92.7858 | 118.3663 | 21.2597 | 0.00279 | 0.002647 | 219.1404 | 0.26848 | 8.4024 | 0.003486 | 0.004752 | 0.004684 | 6e-05 | 0.00096 | 0.29031 | 1.3109 | 2.5477 | 99.35 | 21.07 | 37.206 | 88.9323 | 105.7455 | 85.4874 | 0.009365 | 0.98441 | 0.023612 | 23.9857 | 0.082399 | 0.15972 | 0.002039 | 0.14013 | 0.14013 | 0.046846
2019-03-21 12:46 | 1553186767.383 | 660 | 0.20911 | 0.000733 | 7.4e-05 | 93.7804 | 92.9282 | 118.1421 | 21.2595 | 0.002786 | 0.002602 | 219.1187 | 0.26844 | 8.4006 | 0.003448 | 0.004674 | 0.004577 | 6e-05 | 0.000956 | 0.28694 | 1.2646 | 2.5477 | 99.35 | 21.07 | 37.204 | 88.8452 | 104.839 | 84.4188 | 0.009364 | 0.98735 | 0.023625 | 23.9273 | 0.081652 | 0.15969 | 0.001723 | 0.1178 | 0.1178 | 0.055717
2019-03-21 12:46 | 1553186768.3834 | 661 | 0.20904 | 0.000731 | 7.4e-05 | 93.9988 | 93.2415 | 118.1839 | 21.2601 | 0.002786 | 0.002633 | 219.1387 | 0.26884 | 8.3993 | 0.003412 | 0.00486 | 0.004778 | 6.1e-05 | 0.00096 | 0.27273 | 1.3057 | 2.5479 | 99.35 | 21.07 | 37.2 | 88.8287 | 103.9598 | 85.1678 | 0.009363 | 0.98529 | 0.023604 | 23.9566 | 0.072349 | 0.15993 | 0.002102 | 0.14459 | 0.14459 | 0.045459
2019-03-21 12:46 | 1553186769.3841 | 662 | 0.20901 | 0.000734 | 7.5e-05 | 94.2133 | 93.4077 | 118.1985 | 21.2599 | 0.00276 | 0.002636 | 219.1198 | 0.26926 | 8.4001 | 0.003419 | 0.004659 | 0.004634 | 5.7e-05 | 0.000958 | 0.25033 | 1.2976 | 2.5481 | 99.35 | 21.07 | 37.199 | 88.2247 | 104.127 | 85.2215 | 0.009363 | 0.98514 | 0.023604 | 23.9598 | 0.076809 | 0.16018 | 0.002307 | 0.15902 | 0.15902 | 0.041401
2019-03-21 12:46 | 1553186770.3841 | 663 | 0.20897 | 0.000724 | 7.4e-05 | 94.4579 | 93.4748 | 118.3653 | 21.2596 | 0.002802 | 0.002637 | 219.1486 | 0.26825 | 8.4011 | 0.003455 | 0.005015 | 0.004955 | 5.9e-05 | 0.000963 | 0.29231 | 1.2891 | 2.5479 | 99.35 | 21.07 | 37.197 | 89.2289 | 104.9911 | 85.2547 | 0.009362 | 0.98505 | 0.023619 | 23.9779 | 0.093579 | 0.15958 | 0.002565 | 0.17751 | 0.17751 | 0.036948
2019-03-21 12:46 | 1553186771.3826 | 664 | 0.20902 | 0.000725 | 7.4e-05 | 94.6283 | 94.0085 | 118.2273 | 21.2597 | 0.002799 | 0.00264 | 219.1465 | 0.26921 | 8.4012 | 0.003545 | 0.004258 | 0.004176 | 5.8e-05 | 0.000958 | 0.26432 | 1.2876 | 2.548 | 99.35 | 21.07 | 37.195 | 89.1541 | 107.1868 | 85.3173 | 0.009362 | 0.98487 | 0.023603 | 23.966 | 0.058779 | 0.16015 | 0.002284 | 0.15746 | 0.15746 | 0.041801
2019-03-21 12:46 | 1553186772.3833 | 665 | 0.20908 | 0.000728 | 7.4e-05 | 94.869 | 94.1318 | 118.0065 | 21.2597 | 0.002819 | 0.002583 | 219.1468 | 0.26845 | 8.4013 | 0.003493 | 0.005039 | 0.004889 | 5.7e-05 | 0.000953 | 0.28694 | 1.2833 | 2.5479 | 99.35 | 21.07 | 37.196 | 89.6299 | 105.9116 | 83.9517 | 0.009362 | 0.98864 | 0.023626 | 23.8979 | 0.069782 | 0.1597 | 0.001876 | 0.12864 | 0.12864 | 0.051023
2019-03-21 12:46 | 1553186773.3823 | 666 | 0.20901 | 0.000726 | 7.4e-05 | 95.0657 | 94.44 | 117.9783 | 21.2596 | 0.002832 | 0.002644 | 219.1575 | 0.26802 | 8.4031 | 0.003516 | 0.005199 | 0.005124 | 5.6e-05 | 0.000946 | 0.26936 | 1.3219 | 2.548 | 99.35 | 21.07 | 37.197 | 89.9565 | 106.4821 | 85.4115 | 0.009362 | 0.98462 | 0.023576 | 23.9439 | 0.059062 | 0.15944 | 0.002298 | 0.15826 | 0.15826 | 0.041409
2019-03-21 12:46 | 1553186774.3837 | 667 | 0.20908 | 0.000727 | 7.4e-05 | 95.2513 | 94.6287 | 117.8858 | 21.2597 | 0.002793 | 0.002628 | 219.1685 | 0.26987 | 8.4008 | 0.003534 | 0.004808 | 0.004725 | 5.8e-05 | 0.000948 | 0.26017 | 1.2895 | 2.5479 | 99.35 | 21.07 | 37.2 | 89.0058 | 106.9041 | 85.0496 | 0.009363 | 0.98561 | 0.023578 | 23.9224 | 0.058656 | 0.16055 | 0.001856 | 0.12697 | 0.12697 | 0.051967
2019-03-21 12:46 | 1553186775.3824 | 668 | 0.20903 | 0.000727 | 7.4e-05 | 95.4917 | 94.7115 | 117.8443 | 21.2597 | 0.002827 | 0.002615 | 219.1419 | 0.26906 | 8.4037 | 0.003522 | 0.004775 | 0.004711 | 5.3e-05 | 0.000951 | 0.25265 | 1.3263 | 2.5479 | 99.35 | 21.07 | 37.201 | 89.8214 | 106.6341 | 84.7343 | 0.009363 | 0.98648 | 0.023584 | 23.9076 | 0.073387 | 0.16007 | 0.002173 | 0.14943 | 0.14943 | 0.044025
2019-03-21 12:46 | 1553186776.3835 | 669 | 0.20901 | 0.000722 | 7.5e-05 | 95.6437 | 95.0194 | 117.6807 | 21.2597 | 0.002803 | 0.00263 | 219.1305 | 0.26926 | 8.4038 | 0.003507 | 0.004574 | 0.004479 | 5.7e-05 | 0.000957 | 0.27295 | 1.3265 | 2.5478 | 99.35 | 21.06 | 37.203 | 89.2376 | 106.2531 | 85.0965 | 0.009358 | 0.98548 | 0.023556 | 23.9031 | 0.058579 | 0.16018 | 0.002295 | 0.15782 | 0.15782 | 0.041716
2019-03-21 12:46 | 1553186777.3828 | 670 | 0.20906 | 0.000723 | 7.5e-05 | 95.756 | 94.9843 | 117.7238 | 21.2596 | 0.002713 | 0.002631 | 219.159 | 0.26868 | 8.401 | 0.00349 | 0.005014 | 0.005017 | 5.7e-05 | 0.000962 | 0.29679 | 1.3093 | 2.548 | 99.35 | 21.06 | 37.205 | 87.093 | 105.8531 | 85.1056 | 0.009359 | 0.98546 | 0.02356 | 23.9078 | 0.07238 | 0.15984 | 0.002008 | 0.13755 | 0.13755 | 0.047759
2019-03-21 12:46 | 1553186778.3839 | 671 | 0.20905 | 0.000723 | 7.6e-05 | 95.9044 | 95.2585 | 117.8072 | 21.2602 | 0.002734 | 0.002623 | 219.1416 | 0.26817 | 8.4039 | 0.003519 | 0.004706 | 0.00464 | 5.5e-05 | 0.000959 | 0.29336 | 1.3111 | 2.5479 | 99.35 | 21.06 | 37.206 | 87.5855 | 106.5424 | 84.9089 | 0.009359 | 0.986 | 0.023575 | 23.9097 | 0.060448 | 0.15954 | 0.002044 | 0.14021 | 0.14021 | 0.046765
2019-03-21 12:46 | 1553186779.3835 | 672 | 0.20901 | 0.000722 | 7.6e-05 | 96.1422 | 95.6198 | 117.7274 | 21.26 | 0.002778 | 0.002628 | 219.1724 | 0.26748 | 8.4065 | 0.003533 | 0.005092 | 0.005014 | 5.8e-05 | 0.000959 | 0.28711 | 1.3105 | 2.5481 | 99.35 | 21.06 | 37.209 | 88.6568 | 106.8943 | 85.0492 | 0.00936 | 0.98561 | 0.023562 | 23.9063 | 0.048728 | 0.15913 | 0.002313 | 0.15911 | 0.15911 | 0.041105
2019-03-21 12:46 | 1553186780.3837 | 673 | 0.20904 | 0.000723 | 7.5e-05 | 96.2884 | 95.7959 | 117.5004 | 21.26 | 0.002771 | 0.002668 | 219.1702 | 0.26837 | 8.4049 | 0.003507 | 0.004676 | 0.004655 | 5.3e-05 | 0.000957 | 0.2943 | 1.3521 | 2.548 | 99.35 | 21.06 | 37.211 | 88.4684 | 106.2486 | 86.0012 | 0.00936 | 0.983 | 0.023508 | 23.915 | 0.045871 | 0.15965 | 0.002153 | 0.1475 | 0.1475 | 0.044485
2019-03-21 12:46 | 1553186781.3841 | 674 | 0.20911 | 0.00072 | 7.6e-05 | 96.6055 | 95.9148 | 117.5565 | 21.2597 | 0.002742 | 0.002562 | 219.1751 | 0.26857 | 8.3991 | 0.003464 | 0.005567 | 0.005477 | 5.6e-05 | 0.000958 | 0.30254 | 1.3518 | 2.5479 | 99.35 | 21.06 | 37.216 | 87.7725 | 105.2201 | 83.4547 | 0.009361 | 0.99002 | 0.023598 | 23.8357 | 0.06418 | 0.15977 | 0.00171 | 0.11666 | 0.11666 | 0.056288
2019-03-21 12:46 | 1553186782.3827 | 675 | 0.20901 | 0.000723 | 7.5e-05 | 96.6823 | 96.1181 | 117.479 | 21.2597 | 0.00283 | 0.002601 | 219.1672 | 0.26828 | 8.4009 | 0.003494 | 0.004711 | 0.004624 | 5.2e-05 | 0.000957 | 0.25227 | 1.328 | 2.548 | 99.35 | 21.06 | 37.217 | 89.905 | 105.9548 | 84.394 | 0.009362 | 0.98742 | 0.023559 | 23.8592 | 0.052352 | 0.1596 | 0.002321 | 0.15967 | 0.15967 | 0.041083
2019-03-21 12:46 | 1553186783.3827 | 676 | 0.20908 | 0.000716 | 7.5e-05 | 96.8653 | 96.2118 | 117.6229 | 21.2601 | 0.002746 | 0.002586 | 219.1556 | 0.26816 | 8.401 | 0.003508 | 0.004533 | 0.004533 | 5.3e-05 | 0.000958 | 0.28799 | 1.307 | 2.5479 | 99.35 | 21.06 | 37.217 | 87.8791 | 106.2771 | 84.031 | 0.009362 | 0.98842 | 0.023585 | 23.8617 | 0.060549 | 0.15953 | 0.001896 | 0.12978 | 0.12978 | 0.050519
2019-03-21 12:46 | 1553186784.383 | 677 | 0.20905 | 0.000721 | 7.5e-05 | 96.9362 | 96.6314 | 117.5351 | 21.26 | 0.002714 | 0.002593 | 219.138 | 0.26986 | 8.4051 | 0.003529 | 0.005283 | 0.005249 | 5.3e-05 | 0.000956 | 0.23131 | 1.3123 | 2.5479 | 99.35 | 21.05 | 37.219 | 87.1192 | 106.7902 | 84.1978 | 0.009356 | 0.98796 | 0.023571 | 23.8584 | 0.02817 | 0.16054 | 0.002095 | 0.14384 | 0.14384 | 0.045872
2019-03-21 12:46 | 1553186785.3829 | 678 | 0.20902 | 0.000722 | 7.4e-05 | 97.1197 | 96.8214 | 117.3506 | 21.2599 | 0.002744 | 0.002666 | 219.1532 | 0.26886 | 8.4041 | 0.003481 | 0.004953 | 0.004855 | 4.7e-05 | 0.000958 | 0.26305 | 1.3315 | 2.548 | 99.35 | 21.05 | 37.221 | 87.8336 | 105.6213 | 85.9434 | 0.009357 | 0.98316 | 0.023495 | 23.8978 | 0.027515 | 0.15994 | 0.002237 | 0.15343 | 0.15343 | 0.042846
2019-03-21 12:46 | 1553186786.383 | 679 | 0.20907 | 0.000721 | 7.5e-05 | 97.1719 | 96.7114 | 117.2718 | 21.26 | 0.002799 | 0.002604 | 219.1637 | 0.26829 | 8.4047 | 0.003452 | 0.004591 | 0.00449 | 5.4e-05 | 0.00096 | 0.26936 | 1.3164 | 2.548 | 99.35 | 21.05 | 37.223 | 89.1506 | 104.9175 | 84.4499 | 0.009357 | 0.98726 | 0.023536 | 23.84 | 0.042484 | 0.1596 | 0.001959 | 0.13399 | 0.13399 | 0.048957
2019-03-21 12:46 | 1553186787.383 | 680 | 0.2091 | 0.000713 | 7.5e-05 | 97.3012 | 96.8453 | 117.288 | 21.2599 | 0.002737 | 0.002617 | 219.1245 | 0.26854 | 8.4015 | 0.00342 | 0.004763 | 0.004692 | 5.4e-05 | 0.000961 | 0.29286 | 1.3007 | 2.5479 | 99.35 | 21.05 | 37.224 | 87.6711 | 104.138 | 84.7775 | 0.009358 | 0.98636 | 0.023527 | 23.8526 | 0.042008 | 0.15975 | 0.001792 | 0.12214 | 0.12214 | 0.053755
2019-03-21 12:46 | 1553186788.3831 | 681 | 0.20906 | 0.000714 | 7.5e-05 | 97.3852 | 97.099 | 117.425 | 21.2606 | 0.002723 | 0.002586 | 219.1252 | 0.26932 | 8.4033 | 0.003409 | 0.005785 | 0.005677 | 5.1e-05 | 0.000962 | 0.27671 | 1.3042 | 2.548 | 99.35 | 21.05 | 37.227 | 87.3218 | 103.8825 | 84.0211 | 0.009358 | 0.98845 | 0.023566 | 23.8413 | 0.026324 | 0.16022 | 0.002013 | 0.13797 | 0.13797 | 0.047729
2019-03-21 12:46 | 1553186789.3821 | 682 | 0.20908 | 0.000711 | 7.5e-05 | 97.5273 | 97.1365 | 117.6922 | 21.26 | 0.002779 | 0.002607 | 219.0835 | 0.26838 | 8.4048 | 0.00346 | 0.005975 | 0.005973 | 4.9e-05 | 0.00096 | 0.27505 | 1.3267 | 2.5479 | 99.35 | 21.05 | 37.228 | 88.671 | 105.1195 | 84.5427 | 0.009359 | 0.98701 | 0.023575 | 23.8858 | 0.035917 | 0.15966 | 0.001878 | 0.12842 | 0.12842 | 0.051097
2019-03-21 12:46 | 1553186790.3832 | 683 | 0.20904 | 0.000716 | 7.5e-05 | 97.6027 | 97.304 | 117.7513 | 21.2602 | 0.002807 | 0.002609 | 219.0785 | 0.26941 | 8.4016 | 0.003459 | 0.004656 | 0.00461 | 5e-05 | 0.000965 | 0.31404 | 1.3372 | 2.5479 | 99.35 | 21.05 | 37.229 | 89.3412 | 105.0835 | 84.5873 | 0.009359 | 0.98688 | 0.02358 | 23.8933 | 0.027411 | 0.16027 | 0.002163 | 0.14863 | 0.14863 | 0.044318
2019-03-21 12:47 | 1553186791.3828 | 684 | 0.20905 | 0.000713 | 7.5e-05 | 97.6453 | 97.4008 | 117.8818 | 21.26 | 0.002791 | 0.002629 | 219.0517 | 0.26864 | 8.4039 | 0.003491 | 0.004879 | 0.00484 | 5.3e-05 | 0.000964 | 0.28296 | 1.3087 | 2.5479 | 99.35 | 21.05 | 37.229 | 88.9512 | 105.8765 | 85.0617 | 0.009359 | 0.98558 | 0.023577 | 23.9224 | 0.022428 | 0.15981 | 0.002062 | 0.14149 | 0.14149 | 0.046423
2019-03-21 12:47 | 1553186792.3833 | 685 | 0.20904 | 0.000713 | 7.5e-05 | 97.715 | 97.2005 | 118.1853 | 21.2603 | 0.002742 | 0.00264 | 219.0544 | 0.26929 | 8.4017 | 0.003464 | 0.004913 | 0.00484 | 5.5e-05 | 0.000964 | 0.27671 | 1.3492 | 2.5481 | 99.35 | 21.05 | 37.228 | 87.777 | 105.2143 | 85.3179 | 0.009359 | 0.98487 | 0.023599 | 23.9617 | 0.047221 | 0.1602 | 0.002136 | 0.14688 | 0.14688 | 0.044828
2019-03-21 12:47 | 1553186793.3833 | 686 | 0.2091 | 0.000712 | 7.4e-05 | 97.7788 | 97.583 | 118.6217 | 21.2601 | 0.002793 | 0.002597 | 219.0641 | 0.26873 | 8.4008 | 0.00348 | 0.00504 | 0.004968 | 5.4e-05 | 0.000966 | 0.27079 | 1.2903 | 2.548 | 99.35 | 21.05 | 37.228 | 89.0069 | 105.598 | 84.2891 | 0.009359 | 0.98771 | 0.023677 | 23.9715 | 0.017928 | 0.15987 | 0.001807 | 0.124 | 0.124 | 0.05299
2019-03-21 12:47 | 1553186794.3838 | 687 | 0.20909 | 0.000717 | 7.4e-05 | 97.8431 | 97.5472 | 118.8388 | 21.2603 | 0.002828 | 0.002599 | 219.0418 | 0.2696 | 8.4013 | 0.003485 | 0.005267 | 0.00518 | 5.4e-05 | 0.000961 | 0.26698 | 1.348 | 2.548 | 99.35 | 21.05 | 37.228 | 89.8571 | 105.7225 | 84.3337 | 0.009359 | 0.98759 | 0.023697 | 23.9949 | 0.027092 | 0.16038 | 0.001852 | 0.12737 | 0.12737 | 0.051754
2019-03-21 12:47 | 1553186795.3824 | 688 | 0.20905 | 0.000713 | 7.5e-05 | 97.9957 | 97.7587 | 118.9247 | 21.2606 | 0.002752 | 0.002606 | 219.0312 | 0.26826 | 8.4041 | 0.003457 | 0.00521 | 0.005189 | 5.5e-05 | 0.000959 | 0.2959 | 1.3461 | 2.5481 | 99.35 | 21.05 | 37.228 | 88.0208 | 105.0477 | 84.5189 | 0.009359 | 0.98707 | 0.023699 | 24.0098 | 0.021654 | 0.15959 | 0.002087 | 0.14403 | 0.14403 | 0.04554
2019-03-21 12:47 | 1553186796.3825 | 689 | 0.20908 | 0.000707 | 7.5e-05 | 98.0585 | 97.7468 | 119.4322 | 21.2599 | 0.002749 | 0.002587 | 219.0292 | 0.26873 | 8.4032 | 0.00343 | 0.005169 | 0.00502 | 5.1e-05 | 0.00096 | 0.25177 | 1.3039 | 2.5481 | 99.35 | 21.05 | 37.229 | 87.9616 | 104.3849 | 84.0587 | 0.009359 | 0.98835 | 0.023765 | 24.0455 | 0.028481 | 0.15987 | 0.001919 | 0.13242 | 0.13242 | 0.04962
2019-03-21 12:47 | 1553186797.384 | 690 | 0.2091 | 0.000709 | 7.5e-05 | 98.2599 | 97.6838 | 119.5554 | 21.2602 | 0.002746 | 0.002605 | 219.0594 | 0.26812 | 8.4048 | 0.003449 | 0.004849 | 0.004872 | 5.4e-05 | 0.000957 | 0.24386 | 1.3643 | 2.5481 | 99.35 | 21.05 | 37.228 | 87.8659 | 104.8484 | 84.4789 | 0.009359 | 0.98718 | 0.023763 | 24.072 | 0.052602 | 0.15951 | 0.001824 | 0.12562 | 0.12562 | 0.052188
2019-03-21 12:47 | 1553186798.3836 | 691 | 0.2091 | 0.000709 | 7.6e-05 | 98.383 | 98.146 | 119.6563 | 21.2605 | 0.002751 | 0.002563 | 219.0528 | 0.26865 | 8.4019 | 0.003413 | 0.004149 | 0.004035 | 5.6e-05 | 0.000953 | 0.26017 | 1.3279 | 2.5479 | 99.35 | 21.05 | 37.229 | 88.0089 | 103.9718 | 83.4773 | 0.009359 | 0.98996 | 0.023807 | 24.0484 | 0.021576 | 0.15982 | 0.001769 | 0.12189 | 0.12189 | 0.053889
2019-03-21 12:47 | 1553186799.383 | 692 | 0.20908 | 0.000709 | 7.5e-05 | 98.506 | 98.1385 | 119.4913 | 21.2604 | 0.00276 | 0.002601 | 219.0779 | 0.2688 | 8.4017 | 0.003423 | 0.004744 | 0.004645 | 5.3e-05 | 0.00095 | 0.27417 | 1.3138 | 2.548 | 99.35 | 21.04 | 37.23 | 88.2134 | 104.2249 | 84.3886 | 0.009353 | 0.98743 | 0.02376 | 24.0625 | 0.033439 | 0.15991 | 0.001889 | 0.13025 | 0.13025 | 0.050459
2019-03-21 12:47 | 1553186800.383 | 693 | 0.20909 | 0.000707 | 7.5e-05 | 98.6225 | 98.4654 | 119.4876 | 21.2603 | 0.002683 | 0.002577 | 219.0628 | 0.26855 | 8.4007 | 0.00338 | 0.004543 | 0.004488 | 5.4e-05 | 0.00095 | 0.26488 | 1.3441 | 2.548 | 99.35 | 21.04 | 37.23 | 86.359 | 103.1802 | 83.8102 | 0.009353 | 0.98903 | 0.023779 | 24.0427 | 0.01426 | 0.15976 | 0.001834 | 0.12647 | 0.12647 | 0.051919
2019-03-21 12:47 | 1553186801.3828 | 694 | 0.20912 | 0.000705 | 7.5e-05 | 98.6228 | 98.5809 | 119.5765 | 21.2605 | 0.002744 | 0.002617 | 219.0602 | 0.26864 | 8.4048 | 0.003443 | 0.004696 | 0.004597 | 5e-05 | 0.000954 | 0.2844 | 1.3397 | 2.548 | 99.35 | 21.04 | 37.231 | 87.8349 | 104.7127 | 84.7781 | 0.009354 | 0.98636 | 0.023756 | 24.0842 | 0.003805 | 0.15981 | 0.001685 | 0.11574 | 0.11574 | 0.056752
2019-03-21 12:47 | 1553186802.3842 | 695 | 0.20906 | 0.00071 | 7.4e-05 | 98.7313 | 98.617 | 119.3714 | 21.2603 | 0.002753 | 0.002545 | 219.0432 | 0.2691 | 8.4063 | 0.003436 | 0.00525 | 0.005208 | 5e-05 | 0.000956 | 0.30021 | 1.338 | 2.5479 | 99.35 | 21.04 | 37.236 | 88.0425 | 104.5496 | 83.0428 | 0.009355 | 0.99116 | 0.023793 | 24.0051 | 0.01036 | 0.16009 | 0.002023 | 0.14004 | 0.14004 | 0.046984
2019-03-21 12:47 | 1553186803.3823 | 696 | 0.20911 | 0.000709 | 7.5e-05 | 98.8037 | 98.681 | 118.9299 | 21.2605 | 0.002767 | 0.002542 | 219.0516 | 0.26939 | 8.4018 | 0.003392 | 0.00497 | 0.00493 | 5.4e-05 | 0.000958 | 0.27721 | 1.327 | 2.5479 | 99.35 | 21.04 | 37.236 | 88.3909 | 103.472 | 82.972 | 0.009355 | 0.99136 | 0.023751 | 23.9583 | 0.011112 | 0.16026 | 0.001726 | 0.11862 | 0.11862 | 0.055527
2019-03-21 12:47 | 1553186804.3819 | 697 | 0.20911 | 0.0007 | 7.5e-05 | 98.8308 | 98.4143 | 118.7203 | 21.2603 | 0.002762 | 0.002573 | 219.0675 | 0.2686 | 8.4021 | 0.003362 | 0.004445 | 0.004322 | 5.2e-05 | 0.00096 | 0.25481 | 1.3381 | 2.548 | 99.35 | 21.03 | 37.241 | 88.2627 | 102.7545 | 83.709 | 0.00935 | 0.98931 | 0.023706 | 23.9619 | 0.03777 | 0.15979 | 0.001745 | 0.11971 | 0.11971 | 0.054862
2019-03-21 12:47 | 1553186805.3823 | 698 | 0.20911 | 0.000711 | 7.4e-05 | 98.8585 | 98.7692 | 118.8742 | 21.2603 | 0.002673 | 0.002643 | 219.0796 | 0.26859 | 8.4004 | 0.003357 | 0.004561 | 0.004538 | 5.2e-05 | 0.00096 | 0.26864 | 1.3596 | 2.548 | 99.35 | 21.03 | 37.245 | 86.1115 | 102.6131 | 85.399 | 0.009351 | 0.98465 | 0.023665 | 24.0342 | 0.008081 | 0.15979 | 0.001718 | 0.11767 | 0.11767 | 0.05581
2019-03-21 12:47 | 1553186806.3834 | 699 | 0.20909 | 0.000709 | 7.4e-05 | 98.9732 | 98.7178 | 118.5597 | 21.2607 | 0.002764 | 0.002612 | 219.0695 | 0.26808 | 8.4047 | 0.003373 | 0.004198 | 0.004163 | 5.1e-05 | 0.000962 | 0.24099 | 1.322 | 2.5479 | 99.35 | 21.03 | 37.252 | 88.3014 | 103.0213 | 84.6501 | 0.009353 | 0.98671 | 0.023659 | 23.9773 | 0.023112 | 0.15948 | 0.001836 | 0.12606 | 0.12606 | 0.051997
2019-03-21 12:47 | 1553186807.3821 | 700 | 0.20909 | 0.00071 | 7.4e-05 | 99.0054 | 98.7287 | 118.3891 | 21.2606 | 0.002786 | 0.002612 | 219.0532 | 0.26811 | 8.4032 | 0.003418 | 0.004721 | 0.00469 | 5.3e-05 | 0.000963 | 0.29535 | 1.3005 | 2.5479 | 99.35 | 21.03 | 37.255 | 88.8441 | 104.0938 | 84.6517 | 0.009354 | 0.98671 | 0.023642 | 23.9601 | 0.025032 | 0.1595 | 0.001845 | 0.12658 | 0.12658 | 0.05179
2019-03-21 12:47 | 1553186808.3821 | 701 | 0.2091 | 0.000701 | 7.4e-05 | 98.9839 | 98.8043 | 118.2752 | 21.2607 | 0.002746 | 0.002562 | 219.0595 | 0.2684 | 8.4 | 0.00337 | 0.004307 | 0.004191 | 5.3e-05 | 0.000965 | 0.29479 | 1.3155 | 2.5479 | 99.35 | 21.03 | 37.264 | 87.8834 | 102.9425 | 83.4525 | 0.009356 | 0.99003 | 0.02367 | 23.9084 | 0.016246 | 0.15967 | 0.001787 | 0.12261 | 0.12261 | 0.053522
2019-03-21 12:47 | 1553186809.3826 | 702 | 0.20907 | 0.000702 | 7.4e-05 | 99.0641 | 98.9395 | 118.1999 | 21.2607 | 0.002703 | 0.002578 | 219.0302 | 0.26913 | 8.4038 | 0.003436 | 0.005236 | 0.005161 | 5.4e-05 | 0.000965 | 0.26753 | 1.3271 | 2.5479 | 99.35 | 21.03 | 37.267 | 86.8363 | 104.5391 | 83.8474 | 0.009357 | 0.98893 | 0.023649 | 23.914 | 0.011254 | 0.1601 | 0.001967 | 0.13528 | 0.13528 | 0.048641
2019-03-21 12:47 | 1553186810.3822 | 703 | 0.20911 | 0.000706 | 7.4e-05 | 99.1302 | 99.0638 | 118.225 | 21.2606 | 0.002684 | 0.002569 | 219.0393 | 0.26819 | 8.4038 | 0.003381 | 0.004138 | 0.004116 | 5.1e-05 | 0.000968 | 0.27151 | 1.3278 | 2.5479 | 99.35 | 21.03 | 37.272 | 86.3883 | 103.2176 | 83.628 | 0.009358 | 0.98954 | 0.023659 | 23.9092 | 0.005992 | 0.15955 | 0.001725 | 0.11817 | 0.11817 | 0.055493
2019-03-21 12:47 | 1553186811.3823 | 704 | 0.20907 | 0.0007 | 7.3e-05 | 99.1177 | 99.0046 | 118.0982 | 21.2604 | 0.00275 | 0.002619 | 219.0065 | 0.26884 | 8.4017 | 0.003383 | 0.004711 | 0.004616 | 5.3e-05 | 0.000968 | 0.26681 | 1.3009 | 2.5479 | 99.35 | 21.02 | 37.279 | 87.9761 | 103.2503 | 84.8271 | 0.009354 | 0.98622 | 0.023607 | 23.9365 | 0.010216 | 0.15993 | 0.002016 | 0.13851 | 0.13851 | 0.047458
2019-03-21 12:47 | 1553186812.3837 | 705 | 0.20907 | 0.0007 | 7.3e-05 | 99.1946 | 99.0298 | 118.0924 | 21.2608 | 0.002713 | 0.00261 | 219.0146 | 0.26936 | 8.4017 | 0.003386 | 0.004767 | 0.00469 | 5.1e-05 | 0.000971 | 0.28998 | 1.3311 | 2.548 | 99.35 | 21.02 | 37.283 | 87.0729 | 103.3207 | 84.6028 | 0.009355 | 0.98684 | 0.023614 | 23.9284 | 0.014873 | 0.16024 | 0.002012 | 0.13831 | 0.13831 | 0.047615
2019-03-21 12:47 | 1553186813.3834 | 706 | 0.20907 | 0.000699 | 7.3e-05 | 99.2849 | 99.1456 | 118.1064 | 21.2606 | 0.002709 | 0.002568 | 218.999 | 0.26951 | 8.4001 | 0.003334 | 0.004411 | 0.004308 | 5.2e-05 | 0.000965 | 0.24873 | 1.3335 | 2.5481 | 99.35 | 21.02 | 37.289 | 86.983 | 102.0594 | 83.6008 | 0.009357 | 0.98961 | 0.023648 | 23.8963 | 0.012557 | 0.16033 | 0.001979 | 0.13616 | 0.13616 | 0.048395
2019-03-21 12:47 | 1553186814.3825 | 707 | 0.20903 | 0.000697 | 7.3e-05 | 99.3311 | 99.07 | 118.0069 | 21.2607 | 0.002752 | 0.002571 | 219.0145 | 0.26793 | 8.3994 | 0.003367 | 0.004834 | 0.004765 | 5.1e-05 | 0.000971 | 0.26665 | 1.3393 | 2.5479 | 99.35 | 21.02 | 37.292 | 88.0139 | 102.8662 | 83.6781 | 0.009357 | 0.9894 | 0.023636 | 23.8888 | 0.02354 | 0.15939 | 0.002217 | 0.15301 | 0.15301 | 0.042813
2019-03-21 12:47 | 1553186815.3839 | 708 | 0.20909 | 0.000694 | 7.3e-05 | 99.3966 | 99.3506 | 117.9245 | 21.2606 | 0.002708 | 0.002537 | 219.0039 | 0.26866 | 8.4021 | 0.00337 | 0.004783 | 0.004607 | 5.3e-05 | 0.000969 | 0.25658 | 1.3111 | 2.5479 | 99.35 | 21.02 | 37.295 | 86.9705 | 102.948 | 82.8401 | 0.009358 | 0.99173 | 0.023655 | 23.8524 | 0.004138 | 0.15983 | 0.001861 | 0.12781 | 0.12781 | 0.051394
2019-03-21 12:47 | 1553186816.3825 | 709 | 0.20912 | 0.000698 | 7.3e-05 | 99.4326 | 99.2685 | 117.8403 | 21.2608 | 0.002745 | 0.002551 | 219.0036 | 0.26923 | 8.4023 | 0.003341 | 0.004891 | 0.00479 | 5.3e-05 | 0.000972 | 0.27887 | 1.3301 | 2.548 | 99.35 | 21.02 | 37.301 | 87.8501 | 102.2286 | 83.1923 | 0.00936 | 0.99075 | 0.023635 | 23.8557 | 0.014775 | 0.16017 | 0.001703 | 0.1165 | 0.1165 | 0.056506
2019-03-21 12:47 | 1553186817.3836 | 710 | 0.20913 | 0.000698 | 7.2e-05 | 99.3814 | 99.2662 | 118.0025 | 21.2609 | 0.002746 | 0.002585 | 219.0041 | 0.26922 | 8.4042 | 0.003321 | 0.004873 | 0.004808 | 4.9e-05 | 0.000967 | 0.27577 | 1.3658 | 2.548 | 99.35 | 21.02 | 37.304 | 87.8806 | 101.7592 | 83.9957 | 0.00936 | 0.98852 | 0.023625 | 23.899 | 0.010378 | 0.16016 | 0.001632 | 0.11144 | 0.11144 | 0.059069
2019-03-21 12:47 | 1553186818.3817 | 711 | 0.20915 | 0.000696 | 7.2e-05 | 99.4952 | 99.3796 | 117.9327 | 21.2608 | 0.00274 | 0.002596 | 218.9987 | 0.26958 | 8.4035 | 0.003354 | 0.005256 | 0.005141 | 5e-05 | 0.000967 | 0.24419 | 1.2991 | 2.548 | 99.35 | 21.01 | 37.307 | 87.7384 | 102.5404 | 84.2782 | 0.009355 | 0.98774 | 0.023608 | 23.9014 | 0.010402 | 0.16037 | 0.001514 | 0.10305 | 0.10305 | 0.063966
2019-03-21 12:47 | 1553186819.3821 | 712 | 0.2091 | 0.000699 | 7.1e-05 | 99.5423 | 99.4715 | 117.8731 | 21.2608 | 0.002686 | 0.002558 | 219.0384 | 0.26826 | 8.4044 | 0.003378 | 0.004727 | 0.004634 | 5.1e-05 | 0.000968 | 0.2724 | 1.3028 | 2.5479 | 99.35 | 21.01 | 37.311 | 86.443 | 103.1296 | 83.3668 | 0.009356 | 0.99026 | 0.023632 | 23.8649 | 0.006361 | 0.15959 | 0.001823 | 0.12512 | 0.12512 | 0.052424
2019-03-21 12:47 | 1553186820.3828 | 713 | 0.20912 | 0.000698 | 7.1e-05 | 99.4043 | 99.3765 | 117.8982 | 21.2608 | 0.002706 | 0.002549 | 219.0175 | 0.26997 | 8.4018 | 0.003385 | 0.005394 | 0.005339 | 5e-05 | 0.000969 | 0.23562 | 1.3021 | 2.548 | 99.35 | 21.01 | 37.313 | 86.92 | 103.2994 | 83.1322 | 0.009357 | 0.99091 | 0.023643 | 23.8595 | 0.002505 | 0.16061 | 0.001692 | 0.11585 | 0.11585 | 0.05698
2019-03-21 12:47 | 1553186821.3818 | 714 | 0.2091 | 0.000693 | 7.1e-05 | 99.4803 | 99.4194 | 118.0255 | 21.2609 | 0.002694 | 0.002602 | 219.0364 | 0.26867 | 8.4035 | 0.003294 | 0.00509 | 0.005101 | 5e-05 | 0.000973 | 0.28495 | 1.3076 | 2.5481 | 99.35 | 21.01 | 37.316 | 86.6281 | 101.1004 | 84.4131 | 0.009358 | 0.98737 | 0.023613 | 23.9153 | 0.005477 | 0.15983 | 0.001823 | 0.12501 | 0.12501 | 0.052549
2019-03-21 12:47 | 1553186822.3823 | 715 | 0.20913 | 0.000693 | 7.1e-05 | 99.5515 | 99.6197 | 118.1161 | 21.2609 | 0.002763 | 0.002578 | 219.0073 | 0.26897 | 8.4008 | 0.003325 | 0.005489 | 0.005377 | 5.1e-05 | 0.000968 | 0.2568 | 1.2952 | 2.5478 | 99.35 | 21.01 | 37.318 | 88.2984 | 101.8549 | 83.8459 | 0.009358 | 0.98893 | 0.023641 | 23.9055 | -0.006122 | 0.16001 | 0.001663 | 0.11383 | 0.11383 | 0.057773
2019-03-21 12:47 | 1553186823.3828 | 716 | 0.2091 | 0.000693 | 7.1e-05 | 99.6394 | 99.7188 | 118.1081 | 21.2607 | 0.002704 | 0.002564 | 218.9958 | 0.26852 | 8.4022 | 0.003399 | 0.004744 | 0.004731 | 5.3e-05 | 0.000969 | 0.25929 | 1.2976 | 2.548 | 99.35 | 21.01 | 37.32 | 86.8669 | 103.6428 | 83.5112 | 0.009359 | 0.98986 | 0.023651 | 23.8935 | -0.007123 | 0.15974 | 0.001799 | 0.1235 | 0.1235 | 0.053161
2019-03-21 12:47 | 1553186824.383 | 717 | 0.20911 | 0.000691 | 7.1e-05 | 99.6269 | 99.6037 | 118.4268 | 21.2609 | 0.002718 | 0.0026 | 218.9481 | 0.26874 | 8.4029 | 0.003335 | 0.00445 | 0.004355 | 5.2e-05 | 0.000969 | 0.27273 | 1.3176 | 2.5479 | 99.35 | 21 | 37.325 | 87.198 | 102.0982 | 84.3773 | 0.009354 | 0.98746 | 0.023654 | 23.9547 | 0.002086 | 0.15987 | 0.001778 | 0.12205 | 0.12205 | 0.053837
2019-03-21 12:47 | 1553186825.3831 | 718 | 0.20912 | 0.000692 | 7.1e-05 | 99.6111 | 99.5536 | 118.4853 | 21.2608 | 0.002745 | 0.002577 | 218.9315 | 0.26884 | 8.4019 | 0.003337 | 0.005283 | 0.00516 | 5.1e-05 | 0.000967 | 0.23562 | 1.3151 | 2.5481 | 99.35 | 21 | 37.327 | 87.8491 | 102.1351 | 83.803 | 0.009355 | 0.98905 | 0.023679 | 23.9414 | 0.005165 | 0.15993 | 0.001713 | 0.11755 | 0.11755 | 0.055918
2019-03-21 12:47 | 1553186826.382 | 719 | 0.20912 | 0.000688 | 7e-05 | 99.6257 | 99.5809 | 118.6062 | 21.2608 | 0.002735 | 0.002567 | 218.9245 | 0.2684 | 8.4009 | 0.003297 | 0.004116 | 0.004048 | 5.2e-05 | 0.00097 | 0.22285 | 1.33 | 2.5481 | 99.35 | 21 | 37.333 | 87.6226 | 101.1787 | 83.5743 | 0.009356 | 0.98969 | 0.023699 | 23.9459 | 0.004028 | 0.15967 | 0.001712 | 0.11764 | 0.11764 | 0.055783
2019-03-21 12:47 | 1553186827.3835 | 720 | 0.20913 | 0.000692 | 7e-05 | 99.6522 | 99.6384 | 118.4423 | 21.2608 | 0.002721 | 0.002549 | 218.9284 | 0.26821 | 8.4045 | 0.003266 | 0.004977 | 0.004913 | 5.2e-05 | 0.000966 | 0.25658 | 1.3063 | 2.5481 | 99.35 | 21 | 37.335 | 87.2668 | 100.4182 | 83.1375 | 0.009357 | 0.9909 | 0.023697 | 23.9147 | 0.00124 | 0.15956 | 0.001646 | 0.11294 | 0.11294 | 0.058064
2019-03-21 12:47 | 1553186828.3832 | 721 | 0.20908 | 0.00069 | 7e-05 | 99.5867 | 99.5424 | 118.4067 | 21.261 | 0.00272 | 0.002527 | 218.9489 | 0.26898 | 8.4021 | 0.003376 | 0.004673 | 0.004566 | 5.4e-05 | 0.000969 | 0.28816 | 1.35 | 2.548 | 99.35 | 21 | 37.337 | 87.2418 | 103.093 | 82.6183 | 0.009357 | 0.99235 | 0.023711 | 23.8937 | 0.003981 | 0.16002 | 0.001924 | 0.13278 | 0.13278 | 0.049531
2019-03-21 12:47 | 1553186829.3832 | 722 | 0.20911 | 0.000685 | 7.1e-05 | 99.6159 | 99.6636 | 118.7005 | 21.261 | 0.002718 | 0.002592 | 218.9166 | 0.2692 | 8.4005 | 0.003309 | 0.004892 | 0.004872 | 5.3e-05 | 0.000971 | 0.27024 | 1.3655 | 2.548 | 99.35 | 21 | 37.341 | 87.1974 | 101.4584 | 84.1731 | 0.009358 | 0.98803 | 0.023689 | 23.9755 | -0.004285 | 0.16015 | 0.001744 | 0.11979 | 0.11979 | 0.054949
2019-03-21 12:47 | 1553186830.3837 | 723 | 0.20914 | 0.000688 | 7.1e-05 | 99.6642 | 99.4773 | 118.6536 | 21.2608 | 0.002737 | 0.002542 | 218.9639 | 0.2696 | 8.4006 | 0.003312 | 0.00485 | 0.004754 | 5.4e-05 | 0.000972 | 0.2912 | 1.3392 | 2.5479 | 99.35 | 21 | 37.343 | 87.6525 | 101.5243 | 82.9598 | 0.009359 | 0.9914 | 0.023724 | 23.9301 | 0.016783 | 0.16038 | 0.001597 | 0.10949 | 0.10949 | 0.060204
2019-03-21 12:47 | 1553186831.3834 | 724 | 0.20913 | 0.000688 | 7.1e-05 | 99.6358 | 99.6674 | 118.3132 | 21.261 | 0.00268 | 0.002571 | 218.9628 | 0.26871 | 8.4045 | 0.003393 | 0.004581 | 0.004573 | 5.5e-05 | 0.000974 | 0.23042 | 1.303 | 2.548 | 99.35 | 21 | 37.346 | 86.2801 | 103.489 | 83.6776 | 0.009359 | 0.9894 | 0.023666 | 23.9198 | -0.002829 | 0.15986 | 0.001624 | 0.11118 | 0.11118 | 0.059094
2019-03-21 12:47 | 1553186832.3825 | 725 | 0.20911 | 0.000683 | 7.1e-05 | 99.713 | 99.8138 | 118.3628 | 21.261 | 0.002712 | 0.00256 | 218.9962 | 0.26809 | 8.4023 | 0.00326 | 0.004777 | 0.004723 | 5.3e-05 | 0.000969 | 0.24275 | 1.363 | 2.5481 | 99.35 | 21 | 37.352 | 87.0624 | 100.2827 | 83.4053 | 0.009361 | 0.99016 | 0.02368 | 23.9157 | -0.009038 | 0.15948 | 0.001789 | 0.123 | 0.123 | 0.053293
2019-03-21 12:47 | 1553186833.3835 | 726 | 0.20912 | 0.000687 | 7.1e-05 | 99.7529 | 99.6445 | 118.5008 | 21.261 | 0.002676 | 0.002588 | 218.9891 | 0.26836 | 8.4026 | 0.003379 | 0.005053 | 0.004964 | 5.3e-05 | 0.000965 | 0.23634 | 1.3372 | 2.548 | 99.35 | 21 | 37.354 | 86.1988 | 103.1606 | 84.0848 | 0.009361 | 0.98827 | 0.023671 | 23.9524 | 0.009724 | 0.15965 | 0.001732 | 0.11893 | 0.11893 | 0.055174
2019-03-21 12:47 | 1553186834.3815 | 727 | 0.20909 | 0.000683 | 7.1e-05 | 99.8158 | 99.7958 | 118.1111 | 21.261 | 0.00272 | 0.002522 | 219.0025 | 0.26934 | 8.4014 | 0.003304 | 0.004856 | 0.004711 | 5.7e-05 | 0.000966 | 0.29607 | 1.3328 | 2.5481 | 99.35 | 21 | 37.359 | 87.2498 | 101.3517 | 82.4801 | 0.009363 | 0.99273 | 0.023686 | 23.8592 | 0.001796 | 0.16023 | 0.001925 | 0.13268 | 0.13268 | 0.049633
2019-03-21 12:47 | 1553186835.3826 | 728 | 0.20911 | 0.000684 | 7.1e-05 | 99.7115 | 99.9773 | 118.0994 | 21.261 | 0.00269 | 0.002511 | 219.0192 | 0.26791 | 8.4011 | 0.003361 | 0.004625 | 0.004524 | 5.4e-05 | 0.000963 | 0.27312 | 1.3389 | 2.548 | 99.35 | 21 | 37.362 | 86.5284 | 102.7104 | 82.2198 | 0.009363 | 0.99346 | 0.023693 | 23.8493 | -0.02381 | 0.15938 | 0.00179 | 0.12313 | 0.12313 | 0.053202
2019-03-21 12:47 | 1553186836.382 | 729 | 0.20913 | 0.00069 | 7.1e-05 | 99.7759 | 99.7035 | 117.937 | 21.2611 | 0.002732 | 0.002558 | 219.015 | 0.2689 | 8.4015 | 0.003281 | 0.005001 | 0.004883 | 5.8e-05 | 0.000966 | 0.26465 | 1.3535 | 2.548 | 99.35 | 21 | 37.364 | 87.5474 | 100.7873 | 83.35 | 0.009364 | 0.99031 | 0.023639 | 23.8707 | 0.006494 | 0.15997 | 0.001628 | 0.11131 | 0.11131 | 0.059067
2019-03-21 12:47 | 1553186837.383 | 730 | 0.2091 | 0.000688 | 7.1e-05 | 99.7818 | 99.7277 | 118.0104 | 21.2611 | 0.002671 | 0.002526 | 219.0247 | 0.26894 | 8.4015 | 0.003327 | 0.005091 | 0.00504 | 5.1e-05 | 0.00096 | 0.26521 | 1.304 | 2.548 | 99.35 | 20.99 | 37.368 | 86.0733 | 101.9045 | 82.5871 | 0.009359 | 0.99243 | 0.023672 | 23.8526 | 0.004859 | 0.15999 | 0.001847 | 0.12702 | 0.12702 | 0.051769
2019-03-21 12:47 | 1553186838.3817 | 731 | 0.20911 | 0.000689 | 7.2e-05 | 99.8107 | 99.7317 | 118.1818 | 21.261 | 0.002667 | 0.002557 | 218.9774 | 0.26886 | 8.4004 | 0.003312 | 0.005026 | 0.004891 | 5.7e-05 | 0.000953 | 0.2756 | 1.2932 | 2.548 | 99.35 | 20.99 | 37.371 | 85.989 | 101.5259 | 83.3288 | 0.00936 | 0.99037 | 0.023665 | 23.8948 | 0.007082 | 0.15994 | 0.001757 | 0.12058 | 0.12058 | 0.054515
2019-03-21 12:47 | 1553186839.3818 | 732 | 0.20911 | 0.000684 | 7.2e-05 | 99.8796 | 99.8142 | 118.347 | 21.261 | 0.002756 | 0.002558 | 219.002 | 0.26837 | 8.4022 | 0.003358 | 0.004565 | 0.004553 | 5.4e-05 | 0.00095 | 0.2417 | 1.3178 | 2.5479 | 99.35 | 20.99 | 37.378 | 88.1235 | 102.6421 | 83.3461 | 0.009362 | 0.99032 | 0.023681 | 23.9121 | 0.005859 | 0.15965 | 0.00177 | 0.12159 | 0.12159 | 0.053967
2019-03-21 12:47 | 1553186840.3822 | 733 | 0.20913 | 0.000686 | 7.2e-05 | 99.8721 | 99.8445 | 118.5055 | 21.261 | 0.002704 | 0.002553 | 218.9505 | 0.26802 | 8.4025 | 0.003377 | 0.004883 | 0.004787 | 5.9e-05 | 0.00095 | 0.27865 | 1.359 | 2.5479 | 99.34 | 20.99 | 37.383 | 86.8648 | 103.0976 | 83.2335 | 0.009364 | 0.99063 | 0.0237 | 23.9243 | 0.002476 | 0.15944 | 0.001631 | 0.11176 | 0.11176 | 0.058636
2019-03-21 12:47 | 1553186841.3819 | 734 | 0.20911 | 0.000691 | 7.2e-05 | 99.8951 | 99.6827 | 118.8996 | 21.2609 | 0.002703 | 0.002516 | 218.9725 | 0.26858 | 8.4008 | 0.003366 | 0.00507 | 0.004989 | 5.5e-05 | 0.000952 | 0.27793 | 1.3269 | 2.548 | 99.34 | 20.99 | 37.386 | 86.8539 | 102.8549 | 82.3553 | 0.009365 | 0.99308 | 0.023769 | 23.9345 | 0.019035 | 0.15978 | 0.001758 | 0.12114 | 0.12114 | 0.054209
2019-03-21 12:47 | 1553186842.3819 | 735 | 0.20913 | 0.000688 | 7.1e-05 | 99.9544 | 99.9834 | 118.7871 | 21.2611 | 0.002746 | 0.002588 | 218.9329 | 0.26895 | 8.4013 | 0.003415 | 0.005343 | 0.005299 | 5.7e-05 | 0.000953 | 0.23418 | 1.3545 | 2.5479 | 99.34 | 20.99 | 37.396 | 87.8776 | 104.0204 | 84.0747 | 0.009367 | 0.9883 | 0.0237 | 23.981 | -0.002602 | 0.16 | 0.00166 | 0.11391 | 0.11391 | 0.057728
2019-03-21 12:47 | 1553186843.3834 | 736 | 0.20914 | 0.000692 | 7.1e-05 | 99.966 | 99.8223 | 118.643 | 21.2609 | 0.002661 | 0.00252 | 218.9492 | 0.26892 | 8.4048 | 0.003381 | 0.004579 | 0.004514 | 5.7e-05 | 0.000957 | 0.2724 | 1.3191 | 2.548 | 99.34 | 20.99 | 37.4 | 85.8248 | 103.213 | 82.4529 | 0.009368 | 0.99281 | 0.02374 | 23.912 | 0.012861 | 0.15998 | 0.001565 | 0.10732 | 0.10732 | 0.061266
2019-03-21 12:47 | 1553186844.3821 | 737 | 0.20915 | 0.000689 | 7.1e-05 | 99.9981 | 100.1565 | 118.867 | 21.2608 | 0.002668 | 0.002534 | 203.7348 | 0.26818 | 8.3995 | 0.003371 | 0.004868 | 0.004808 | 5.5e-05 | 0.000954 | 0.29911 | 1.3125 | 2.548 | 99.34 | 20.99 | 37.41 | 85.9941 | 102.964 | 82.7788 | 0.009371 | 0.9919 | 0.023751 | 23.9455 | -0.014149 | 0.15954 | 0.001536 | 0.10526 | 0.10526 | 0.062292
2019-03-21 12:47 | 1553186845.3821 | 738 | 0.20916 | 0.000688 | 7e-05 | 100.013 | 99.8819 | 119.1193 | 21.261 | 0.002669 | 0.002526 | 166.6409 | 0.26839 | 8.4017 | 0.003516 | 0.005317 | 0.005306 | 5.9e-05 | 0.00096 | 0.27959 | 1.3142 | 2.5479 | 99.34 | 20.99 | 37.414 | 86.0155 | 106.474 | 82.5895 | 0.009372 | 0.99243 | 0.023783 | 23.9645 | 0.011733 | 0.15966 | 0.001493 | 0.10239 | 0.10239 | 0.064093
2019-03-21 12:47 | 1553186846.3835 | 739 | 0.20912 | 0.000687 | 7e-05 | 100.1161 | 100.229 | 119.6211 | 21.2609 | 0.002635 | 0.002463 | 129.5362 | 0.26765 | 8.4033 | 0.003462 | 0.00524 | 0.005181 | 5.6e-05 | 0.000958 | 0.26913 | 1.2979 | 2.5479 | 99.34 | 20.99 | 37.419 | 85.2091 | 105.1638 | 81.0794 | 0.009373 | 0.99666 | 0.023884 | 23.9639 | -0.010078 | 0.15923 | 0.001695 | 0.11733 | 0.11733 | 0.055777
2019-03-21 12:47 | 1553186847.3822 | 740 | 0.20916 | 0.000689 | 6.8e-05 | 100.1224 | 100.0318 | 120.065 | 21.261 | 0.00267 | 0.002546 | 92.4057 | 0.26737 | 8.4047 | 0.003571 | 0.004987 | 0.004965 | 5.6e-05 | 0.000958 | 0.25155 | 1.2974 | 2.5479 | 99.35 | 20.99 | 37.428 | 86.0499 | 107.8118 | 83.0581 | 0.009374 | 0.99112 | 0.023862 | 24.0753 | 0.008097 | 0.15906 | 0.001463 | 0.10073 | 0.10073 | 0.064898
2019-03-21 12:47 | 1553186848.3833 | 741 | 0.20914 | 0.000685 | 6.6e-05 | 100.1856 | 100.0682 | 120.3011 | 21.261 | 0.0027 | 0.00244 | 55.3132 | 0.26848 | 8.4032 | 0.003442 | 0.00453 | 0.004494 | 5.5e-05 | 0.000957 | 0.28622 | 1.2864 | 2.5478 | 99.35 | 20.99 | 37.432 | 86.7663 | 104.6736 | 80.5242 | 0.009375 | 0.99822 | 0.02397 | 24.0131 | 0.01048 | 0.15972 | 0.001593 | 0.11065 | 0.11065 | 0.059326
2019-03-21 12:47 | 1553186849.3823 | 742 | 0.2092 | 0.000685 | 6.3e-05 | 100.1787 | 100.1261 | 120.0967 | 21.261 | 0.00267 | 0.002535 | 18.1889 | 0.25556 | 8.4042 | 0.003464 | 0.004853 | 0.004686 | 5.4e-05 | 0.000959 | 0.24331 | 1.3085 | 2.5479 | 99.35 | 20.99 | 37.44 | 86.0551 | 105.2062 | 82.8085 | 0.009377 | 0.99182 | 0.023873 | 24.07 | 0.004694 | 0.15204 | 0.001239 | 0.085035 | 0.085035 | 0.073483
2019-03-21 12:47 | 1553186850.3827 | 743 | 0.20917 | 0.000685 | 5.9e-05 | 100.2094 | 100.2058 | 120.1422 | 21.2608 | 0.002651 | 0.002497 | -18.9124 | 0.23915 | 8.4022 | 0.003489 | 0.00466 | 0.004542 | 5.8e-05 | 0.000962 | 0.29463 | 1.3536 | 2.5481 | 99.35 | 20.99 | 37.442 | 85.5924 | 105.8293 | 81.8982 | 0.009378 | 0.99436 | 0.023908 | 24.0438 | 0.000317 | 0.14227 | 0.001456 | 0.10091 | 0.10091 | 0.057945
2019-03-21 12:48 | 1553186851.3818 | 744 | 0.20923 | 0.000678 | 5.7e-05 | 100.1605 | 100.1905 | 120.4241 | 21.2611 | 0.002644 | 0.002495 | -56.0413 | 0.23908 | 8.4036 | 0.003428 | 0.005072 | 0.004861 | 5.8e-05 | 0.000959 | 0.28528 | 1.3537 | 2.548 | 99.35 | 20.99 | 37.444 | 85.4254 | 104.3388 | 81.8407 | 0.009378 | 0.99452 | 0.023938 | 24.07 | -0.002677 | 0.14223 | 0.001106 | 0.076053 | 0.076053 | 0.076863
2019-03-21 12:48 | 1553186852.3815 | 745 | 0.20924 | 0.000679 | 5.4e-05 | 100.2112 | 100.291 | 120.826 | 21.261 | 0.002678 | 0.002464 | -93.1287 | 0.2405 | 8.3992 | 0.003405 | 0.004157 | 0.004006 | 5.5e-05 | 0.000961 | 0.26648 | 1.3535 | 2.5481 | 99.35 | 20.99 | 37.446 | 86.2339 | 103.7782 | 81.0943 | 0.009379 | 0.99662 | 0.024003 | 24.0848 | -0.007121 | 0.14307 | 0.00099 | 0.068088 | 0.068088 | 0.086364
2019-03-21 12:48 | 1553186853.3829 | 746 | 0.20924 | 0.000677 | 5.1e-05 | 100.2159 | 100.3751 | 121.6825 | 21.2609 | 0.002654 | 0.002405 | -130.2598 | 0.23951 | 8.4018 | 0.003378 | 0.004433 | 0.004289 | 5.5e-05 | 0.000964 | 0.26394 | 1.3072 | 2.5479 | 99.35 | 20.99 | 37.448 | 85.6586 | 103.144 | 79.6807 | 0.009379 | 1.0006 | 0.024136 | 24.1217 | -0.014195 | 0.14249 | 0.000998 | 0.069199 | 0.069199 | 0.084627
2019-03-21 12:48 | 1553186854.3826 | 747 | 0.20929 | 0.000669 | 4.8e-05 | 100.2305 | 100.2811 | 122.2784 | 21.261 | 0.002615 | 0.002433 | -152.1407 | 0.23993 | 8.4021 | 0.003442 | 0.004261 | 0.004056 | 5.8e-05 | 0.000967 | 0.26825 | 1.329 | 2.5479 | 99.35 | 20.99 | 37.448 | 84.7341 | 104.6958 | 80.3551 | 0.009379 | 0.9987 | 0.024172 | 24.2038 | -0.004511 | 0.14274 | 0.00076 | 0.052248 | 0.052248 | 0.11228
2019-03-21 12:48 | 1553186855.3816 | 748 | 0.20933 | 0.000656 | 4.4e-05 | 100.3071 | 100.2218 | 122.55 | 21.2609 | 0.002592 | 0.00237 | -152.1634 | 0.23921 | 8.4045 | 0.003417 | 0.004764 | 0.004624 | 5.6e-05 | 0.000972 | 0.25929 | 1.3153 | 2.5479 | 99.35 | 20.99 | 37.448 | 84.1667 | 104.0832 | 78.8346 | 0.009379 | 1.003 | 0.024251 | 24.1785 | 0.007607 | 0.1423 | 0.000515 | 0.034847 | 0.034847 | 0.16784
2019-03-21 12:48 | 1553186856.3831 | 749 | 0.20934 | 0.000657 | 3.9e-05 | 100.2681 | 100.3421 | 122.627 | 21.2609 | 0.002622 | 0.002421 | -152.1753 | 0.24025 | 8.4028 | 0.00345 | 0.004434 | 0.004305 | 6.1e-05 | 0.000973 | 0.28567 | 1.3205 | 2.548 | 99.35 | 20.98 | 37.447 | 84.8884 | 104.8862 | 80.0681 | 0.009373 | 0.99951 | 0.024217 | 24.2285 | -0.006598 | 0.14293 | 0.000458 | 0.030946 | 0.030946 | 0.18982
2019-03-21 12:48 | 1553186857.3818 | 750 | 0.20934 | 0.000648 | 3.4e-05 | 100.3512 | 100.481 | 122.5716 | 21.2611 | 0.002567 | 0.002433 | -152.1825 | 0.23887 | 8.4012 | 0.003416 | 0.00502 | 0.004845 | 6e-05 | 0.000974 | 0.26665 | 1.3238 | 2.5479 | 99.35 | 20.99 | 37.447 | 83.5802 | 104.0536 | 80.3538 | 0.009379 | 0.9987 | 0.024201 | 24.2328 | -0.011564 | 0.1421 | 0.000507 | 0.034763 | 0.034763 | 0.168
2019-03-21 12:48 | 1553186858.3828 | 751 | 0.20936 | 0.000638 | 2.9e-05 | 100.2812 | 100.3203 | 122.6009 | 21.261 | 0.002596 | 0.002441 | -152.1869 | 0.24002 | 8.4034 | 0.00344 | 0.004222 | 0.004052 | 5.9e-05 | 0.000979 | 0.26609 | 1.2893 | 2.5479 | 99.35 | 20.99 | 37.447 | 84.2753 | 104.6415 | 80.5556 | 0.009379 | 0.99813 | 0.024197 | 24.2426 | -0.003485 | 0.14279 | 0.000364 | 0.024683 | 0.024683 | 0.23775
2019-03-21 12:48 | 1553186859.3828 | 752 | 0.20937 | 0.000625 | 2.5e-05 | 100.3455 | 100.4681 | 122.7901 | 21.261 | 0.002534 | 0.002387 | -152.1858 | 0.23929 | 8.4013 | 0.003412 | 0.004917 | 0.00476 | 6.2e-05 | 0.000978 | 0.2625 | 1.286 | 2.548 | 99.35 | 20.98 | 37.447 | 82.7866 | 103.9558 | 79.2518 | 0.009373 | 1.0018 | 0.024261 | 24.2165 | -0.010916 | 0.14235 | 0.000353 | 0.024201 | 0.024201 | 0.24175
2019-03-21 12:48 | 1553186860.3823 | 753 | 0.2094 | 0.000615 | 2e-05 | 100.3262 | 100.4003 | 122.9472 | 21.261 | 0.002586 | 0.002345 | -152.1761 | 0.23969 | 8.4044 | 0.003439 | 0.004836 | 0.004632 | 5.6e-05 | 0.000978 | 0.26969 | 1.3421 | 2.5479 | 99.35 | 20.98 | 37.447 | 84.0396 | 104.6054 | 78.2283 | 0.009373 | 1.0047 | 0.024312 | 24.1968 | -0.006607 | 0.14259 | 0.000166 | 0.010913 | 0.010913 | 0.53704
2019-03-21 12:48 | 1553186861.382 | 754 | 0.20942 | 0.000615 | 1.6e-05 | 100.3178 | 100.2147 | 122.8594 | 21.2611 | 0.002541 | 0.002358 | -152.164 | 0.23993 | 8.4021 | 0.003466 | 0.005267 | 0.005066 | 6.4e-05 | 0.000983 | 0.24945 | 1.3068 | 2.5479 | 99.35 | 20.98 | 37.447 | 82.9396 | 105.2662 | 78.5402 | 0.009373 | 1.0039 | 0.024292 | 24.1989 | 0.009202 | 0.14274 | 5.8e-05 | 0.003243 | 0.003243 | 1.809
2019-03-21 12:48 | 1553186862.383 | 755 | 0.20944 | 0.000597 | 1.3e-05 | 100.2857 | 100.2695 | 123.1047 | 21.2611 | 0.002513 | 0.002312 | -152.1748 | 0.24155 | 8.4044 | 0.003513 | 0.004464 | 0.004337 | 6.6e-05 | 0.000985 | 0.27688 | 1.3056 | 2.5479 | 99.35 | 20.98 | 37.448 | 82.2643 | 106.4151 | 77.4527 | 0.009373 | 1.007 | 0.024354 | 24.1856 | 0.001439 | 0.1437 | -3.7e-05 | -0.003511 | -0.003511 | -1.6821
2019-03-21 12:48 | 1553186863.3835 | 756 | 0.20946 | 0.000586 | 1e-05 | 100.378 | 100.4453 | 122.9974 | 21.2612 | 0.002512 | 0.00236 | -152.1644 | 0.24039 | 8.4025 | 0.003479 | 0.005336 | 0.005181 | 6.3e-05 | 0.000984 | 0.25265 | 1.3025 | 2.548 | 99.35 | 20.98 | 37.448 | 82.2565 | 105.5723 | 78.6035 | 0.009373 | 1.0037 | 0.024304 | 24.2147 | -0.00599 | 0.14301 | -0.000137 | -0.010629 | -0.010629 | -0.553
2019-03-21 12:48 | 1553186864.3821 | 757 | 0.20947 | 0.000577 | 8e-06 | 100.3843 | 100.4398 | 123.1084 | 21.261 | 0.002476 | 0.002275 | -152.1885 | 0.23891 | 8.4045 | 0.00348 | 0.00484 | 0.004699 | 6.5e-05 | 0.00098 | 0.26648 | 1.3435 | 2.548 | 99.35 | 20.98 | 37.448 | 81.3751 | 105.6051 | 76.5524 | 0.009373 | 1.0096 | 0.024386 | 24.1549 | -0.004947 | 0.14213 | -0.000177 | -0.013458 | -0.013458 | -0.43404
2019-03-21 12:48 | 1553186865.3831 | 758 | 0.20948 | 0.000561 | 6e-06 | 100.3893 | 100.214 | 123.069 | 21.2611 | 0.002466 | 0.002311 | -152.1891 | 0.24013 | 8.3996 | 0.003394 | 0.004947 | 0.0048 | 6e-05 | 0.000982 | 0.23828 | 1.3256 | 2.548 | 99.35 | 20.98 | 37.449 | 81.1334 | 103.5308 | 77.4309 | 0.009374 | 1.007 | 0.024351 | 24.1813 | 0.015636 | 0.14286 | -0.000214 | -0.01598 | -0.01598 | -0.36742
2019-03-21 12:48 | 1553186866.3826 | 759 | 0.20949 | 0.000559 | 5e-06 | 100.3509 | 100.4572 | 122.9953 | 21.261 | 0.002472 | 0.002297 | -152.1902 | 0.23968 | 8.4044 | 0.003402 | 0.004731 | 0.004574 | 6.2e-05 | 0.000981 | 0.27544 | 1.3637 | 2.548 | 99.35 | 20.98 | 37.45 | 81.2768 | 103.7146 | 77.0848 | 0.009374 | 1.008 | 0.024356 | 24.1621 | -0.009468 | 0.14259 | -0.000261 | -0.019412 | -0.019412 | -0.30189
2019-03-21 12:48 | 1553186867.3813 | 760 | 0.20947 | 0.000547 | 4e-06 | 100.4209 | 100.4037 | 122.9784 | 21.2611 | 0.00245 | 0.002291 | -152.1795 | 0.24004 | 8.4017 | 0.003348 | 0.004938 | 0.004642 | 6.2e-05 | 0.000981 | 0.29518 | 1.4064 | 2.5479 | 99.34 | 20.98 | 37.451 | 80.7544 | 102.3962 | 76.9484 | 0.009375 | 1.0084 | 0.024359 | 24.1558 | 0.001533 | 0.1428 | -0.000142 | -0.010602 | -0.010602 | -0.55357
2019-03-21 12:48 | 1553186868.3813 | 761 | 0.20955 | 0.000542 | 2e-06 | 100.3858 | 100.4919 | 122.8493 | 21.2611 | 0.002501 | 0.002267 | -152.1799 | 0.2398 | 8.4016 | 0.003427 | 0.004001 | 0.003907 | 6.3e-05 | 0.00098 | 0.26913 | 1.3572 | 2.5481 | 99.34 | 20.98 | 37.453 | 81.9857 | 104.3311 | 76.3668 | 0.009376 | 1.0101 | 0.024367 | 24.123 | -0.009454 | 0.14266 | -0.000579 | -0.042486 | -0.042486 | -0.138
2019-03-21 12:48 | 1553186869.3814 | 762 | 0.20948 | 0.000536 | 2e-06 | 100.3697 | 100.5474 | 122.7573 | 21.2611 | 0.00243 | 0.002299 | -152.1832 | 0.23941 | 8.4048 | 0.003385 | 0.005112 | 0.004906 | 6.5e-05 | 0.000978 | 0.27057 | 1.3391 | 2.5481 | 99.34 | 20.98 | 37.454 | 80.2804 | 103.3045 | 77.1416 | 0.009376 | 1.0079 | 0.02433 | 24.1407 | -0.015828 | 0.14242 | -0.000183 | -0.013475 | -0.013475 | -0.4344
2019-03-21 12:48 | 1553186870.3824 | 763 | 0.20951 | 0.000522 | 1e-06 | 100.3819 | 100.6105 | 122.7467 | 21.2611 | 0.002413 | 0.002247 | -152.1852 | 0.24008 | 8.4023 | 0.003396 | 0.004529 | 0.004351 | 6e-05 | 0.000976 | 0.25553 | 1.3231 | 2.548 | 99.34 | 20.98 | 37.454 | 79.8788 | 103.57 | 75.8882 | 0.009376 | 1.0115 | 0.024373 | 24.0964 | -0.020343 | 0.14283 | -0.000345 | -0.025337 | -0.025337 | -0.23168
2019-03-21 12:48 | 1553186871.3825 | 764 | 0.20948 | 0.000523 | 0 | 100.4057 | 100.369 | 122.7082 | 21.2612 | 0.002419 | 0.002252 | -152.2003 | 0.24004 | 8.4021 | 0.003475 | 0.004524 | 0.00436 | 5.9e-05 | 0.000978 | 0.28008 | 1.3464 | 2.5481 | 99.34 | 20.98 | 37.455 | 80.0123 | 105.4819 | 75.9966 | 0.009376 | 1.0112 | 0.024365 | 24.0964 | 0.003275 | 0.1428 | -0.000164 | -0.011988 | -0.011988 | -0.48956
2019-03-21 12:48 | 1553186872.3826 | 765 | 0.2095 | 0.000519 | -1e-06 | 100.3986 | 100.4259 | 122.6237 | 21.2612 | 0.002426 | 0.002301 | -152.1975 | 0.23979 | 8.4022 | 0.003442 | 0.004769 | 0.004582 | 6e-05 | 0.000977 | 0.26416 | 1.2928 | 2.5479 | 99.34 | 20.98 | 37.455 | 80.186 | 104.6951 | 77.1822 | 0.009376 | 1.0077 | 0.024316 | 24.129 | -0.00243 | 0.14265 | -0.000261 | -0.019024 | -0.019024 | -0.30819
2019-03-21 12:48 | 1553186873.3816 | 766 | 0.20949 | 0.000507 | -1e-06 | 100.4066 | 100.419 | 122.4831 | 21.2611 | 0.002474 | 0.002262 | -152.1951 | 0.23974 | 8.4021 | 0.003455 | 0.004707 | 0.004543 | 6e-05 | 0.000978 | 0.24065 | 1.3387 | 2.548 | 99.34 | 20.98 | 37.456 | 81.3439 | 105.0034 | 76.2389 | 0.009376 | 1.0105 | 0.024335 | 24.0826 | -0.001106 | 0.14262 | -0.000222 | -0.016138 | -0.016138 | -0.36322
2019-03-21 12:48 | 1553186874.3816 | 767 | 0.20949 | 0.000506 | -1e-06 | 100.392 | 100.3867 | 122.3549 | 21.2612 | 0.002416 | 0.002243 | -152.1772 | 0.24004 | 8.4006 | 0.003478 | 0.004701 | 0.004523 | 5.6e-05 | 0.00098 | 0.29031 | 1.3176 | 2.5481 | 99.34 | 20.98 | 37.457 | 79.9328 | 105.559 | 75.7924 | 0.009377 | 1.0118 | 0.024337 | 24.0546 | 0.000476 | 0.1428 | -0.000204 | -0.014876 | -0.014876 | -0.39453
2019-03-21 12:48 | 1553186875.3826 | 768 | 0.20952 | 0.000504 | -2e-06 | 100.3873 | 100.3537 | 122.3049 | 21.2611 | 0.002412 | 0.002229 | -152.1909 | 0.23955 | 8.4008 | 0.003386 | 0.004768 | 0.004571 | 5.8e-05 | 0.000981 | 0.27024 | 1.3262 | 2.5481 | 99.34 | 20.98 | 37.458 | 79.857 | 103.3367 | 75.4572 | 0.009377 | 1.0127 | 0.024344 | 24.0382 | 0.002995 | 0.14251 | -0.000362 | -0.026374 | -0.026374 | -0.22208
2019-03-21 12:48 | 1553186876.3831 | 769 | 0.20952 | 0.000502 | -2e-06 | 100.3244 | 100.433 | 122.5453 | 21.2612 | 0.002344 | 0.002227 | -152.1825 | 0.24 | 8.4037 | 0.003343 | 0.005391 | 0.005256 | 5.9e-05 | 0.000982 | 0.26969 | 1.344 | 2.5478 | 99.34 | 20.98 | 37.459 | 78.2082 | 102.2793 | 75.4135 | 0.009377 | 1.0129 | 0.02437 | 24.0603 | -0.009678 | 0.14278 | -0.000341 | -0.024867 | -0.024867 | -0.23598
2019-03-21 12:48 | 1553186877.3828 | 770 | 0.2095 | 0.000496 | -2e-06 | 100.3557 | 100.4783 | 122.8812 | 21.2612 | 0.0024 | 0.00224 | -152.2034 | 0.23944 | 8.4011 | 0.003325 | 0.005155 | 0.004987 | 5.7e-05 | 0.000983 | 0.27632 | 1.3817 | 2.5479 | 99.34 | 20.97 | 37.459 | 79.5521 | 101.8511 | 75.7177 | 0.009371 | 1.012 | 0.024392 | 24.1037 | -0.010923 | 0.14244 | -0.000233 | -0.016912 | -0.016912 | -0.34617
2019-03-21 12:48 | 1553186878.3819 | 771 | 0.20954 | 0.000496 | -3e-06 | 100.297 | 100.37 | 123.1187 | 21.2611 | 0.002418 | 0.00226 | -152.1981 | 0.23931 | 8.4047 | 0.00333 | 0.004183 | 0.003987 | 5.9e-05 | 0.000978 | 0.24154 | 1.3711 | 2.5481 | 99.34 | 20.97 | 37.461 | 79.9957 | 101.9714 | 76.1907 | 0.009372 | 1.0106 | 0.024399 | 24.1434 | -0.006509 | 0.14236 | -0.000467 | -0.034093 | -0.034093 | -0.17163
2019-03-21 12:48 | 1553186879.3832 | 772 | 0.20951 | 0.000491 | -2e-06 | 100.325 | 100.572 | 123.2413 | 21.2611 | 0.002391 | 0.00221 | -152.2148 | 0.24036 | 8.4009 | 0.003384 | 0.004655 | 0.004448 | 6.3e-05 | 0.000983 | 0.2912 | 1.3118 | 2.548 | 99.34 | 20.97 | 37.461 | 79.3533 | 103.273 | 75.005 | 0.009372 | 1.014 | 0.024453 | 24.1144 | -0.021992 | 0.14299 | -0.000313 | -0.022848 | -0.022848 | -0.25722
2019-03-21 12:48 | 1553186880.3823 | 773 | 0.20946 | 0.00049 | -3e-06 | 100.3053 | 100.3162 | 123.138 | 21.2612 | 0.002414 | 0.002184 | -152.2103 | 0.24033 | 8.4043 | 0.003418 | 0.004456 | 0.004289 | 6.1e-05 | 0.000984 | 0.30381 | 1.3386 | 2.5479 | 99.34 | 20.97 | 37.461 | 79.896 | 104.0922 | 74.3685 | 0.009372 | 1.0159 | 0.024465 | 24.0822 | -0.000968 | 0.14297 | 1.1e-05 | 0.001008 | 0.001008 | 5.8304
2019-03-21 12:48 | 1553186881.381 | 774 | 0.20948 | 0.000489 | -3e-06 | 100.3557 | 100.513 | 123.1882 | 21.2612 | 0.002349 | 0.002248 | -152.2065 | 0.24004 | 8.4002 | 0.003372 | 0.004827 | 0.004737 | 6.1e-05 | 0.000978 | 0.24082 | 1.3369 | 2.5478 | 99.34 | 20.98 | 37.461 | 78.3259 | 102.9902 | 75.9046 | 0.009378 | 1.0114 | 0.024416 | 24.1403 | -0.014015 | 0.1428 | -0.000128 | -0.009245 | -0.009245 | -0.63483
2019-03-21 12:48 | 1553186882.3825 | 775 | 0.20957 | 0.000485 | -3e-06 | 100.2994 | 100.3966 | 122.9328 | 21.2612 | 0.002318 | 0.002213 | -152.1993 | 0.24025 | 8.4032 | 0.00342 | 0.00443 | 0.004214 | 6.4e-05 | 0.000984 | 0.26432 | 1.3488 | 2.5481 | 99.34 | 20.98 | 37.462 | 77.5885 | 104.1479 | 75.0683 | 0.009378 | 1.0139 | 0.02442 | 24.0864 | -0.008664 | 0.14293 | -0.000643 | -0.046983 | -0.046983 | -0.12503
2019-03-21 12:48 | 1553186883.3811 | 776 | 0.20947 | 0.000485 | -3e-06 | 100.3312 | 100.5655 | 122.6685 | 21.2612 | 0.002374 | 0.002181 | -152.199 | 0.23975 | 8.4005 | 0.003313 | 0.004829 | 0.004661 | 6.5e-05 | 0.000982 | 0.2557 | 1.3136 | 2.548 | 99.34 | 20.98 | 37.463 | 78.9456 | 101.5568 | 74.2872 | 0.009378 | 1.0161 | 0.024421 | 24.0334 | -0.020859 | 0.14263 | -2.5e-05 | -0.001676 | -0.001676 | -3.498
2019-03-21 12:48 | 1553186884.4418 | 777 | 0.2095 | 0.000483 | -3e-06 | 100.3315 | 100.3649 | 122.535 | 21.2611 | 0.002371 | 0.00221 | -152.2042 | 0.23844 | 8.4011 | 0.003404 | 0.004588 | 0.004421 | 6.1e-05 | 0.000986 | 0.27295 | 1.2863 | 2.548 | 99.34 | 20.98 | 37.463 | 78.8545 | 103.7643 | 74.9825 | 0.009378 | 1.0141 | 0.024384 | 24.0444 | -0.002972 | 0.14185 | -0.000207 | -0.014953 | -0.014953 | -0.38988
2019-03-21 12:48 | 1553186885.3881 | 778 | 0.20946 | 0.000479 | -3e-06 | 100.3223 | 100.3632 | 122.7061 | 21.2612 | 0.002299 | 0.002185 | -152.1861 | 0.2393 | 8.4041 | 0.003449 | 0.004442 | 0.00417 | 6.6e-05 | 0.000987 | 0.26504 | 1.2843 | 2.5478 | 99.34 | 20.98 | 37.462 | 77.1434 | 104.8556 | 74.4011 | 0.009378 | 1.0158 | 0.024421 | 24.0411 | -0.003644 | 0.14236 | -1e-06 | 0.000109 | 0.000109 | 53.891
2019-03-21 12:48 | 1553186886.3888 | 779 | 0.20949 | 0.00048 | -3e-06 | 100.3128 | 100.5434 | 122.7199 | 21.2612 | 0.002377 | 0.002193 | -152.1791 | 0.23991 | 8.402 | 0.003454 | 0.004623 | 0.004494 | 6.8e-05 | 0.000989 | 0.29999 | 1.3437 | 2.548 | 99.34 | 20.98 | 37.463 | 78.9987 | 104.9664 | 74.5842 | 0.009378 | 1.0153 | 0.024416 | 24.0488 | -0.020537 | 0.14272 | -0.000151 | -0.010845 | -0.010845 | -0.54089
2019-03-21 12:48 | 1553186887.3888 | 780 | 0.20947 | 0.000479 | -3e-06 | 100.2952 | 100.3734 | 122.6913 | 21.2612 | 0.002321 | 0.002171 | -152.1674 | 0.23922 | 8.4032 | 0.003483 | 0.004927 | 0.004657 | 6.4e-05 | 0.000993 | 0.25818 | 1.3453 | 2.548 | 99.34 | 20.98 | 37.463 | 77.6587 | 105.6839 | 74.0629 | 0.009378 | 1.0168 | 0.024432 | 24.0279 | -0.006972 | 0.14231 | -5.5e-05 | -0.003871 | -0.003871 | -1.511
2019-03-21 12:48 | 1553186888.3883 | 781 | 0.20949 | 0.000478 | -3e-06 | 100.3083 | 100.3707 | 123.0042 | 21.2611 | 0.002324 | 0.002152 | -152.1834 | 0.23986 | 8.4035 | 0.003544 | 0.004658 | 0.004479 | 6.7e-05 | 0.000994 | 0.28385 | 1.3784 | 2.548 | 99.34 | 20.98 | 37.463 | 77.7341 | 107.1582 | 73.603 | 0.009378 | 1.0181 | 0.024479 | 24.0426 | -0.00556 | 0.14269 | -0.000122 | -0.008815 | -0.008815 | -0.66533
2019-03-21 12:48 | 1553186889.3889 | 782 | 0.20949 | 0.000476 | -3e-06 | 100.3104 | 100.3176 | 123.3781 | 21.2612 | 0.002281 | 0.002149 | -152.1709 | 0.23998 | 8.4024 | 0.003413 | 0.004491 | 0.004396 | 6.9e-05 | 0.000998 | 0.27544 | 1.3694 | 2.5481 | 99.34 | 20.98 | 37.463 | 76.7023 | 103.982 | 73.5163 | 0.009378 | 1.0184 | 0.024519 | 24.0761 | -0.000638 | 0.14276 | -0.000165 | -0.011989 | -0.011989 | -0.48942
2019-03-21 12:48 | 1553186890.389 | 783 | 0.20949 | 0.000475 | -3e-06 | 100.2752 | 100.3898 | 123.7358 | 21.2612 | 0.002274 | 0.002179 | -152.1655 | 0.2395 | 8.4039 | 0.003484 | 0.004892 | 0.004715 | 6.9e-05 | 0.000999 | 0.27616 | 1.3798 | 2.5481 | 99.34 | 20.98 | 37.462 | 76.536 | 105.7028 | 74.2553 | 0.009378 | 1.0162 | 0.024529 | 24.1367 | -0.010209 | 0.14248 | -0.000146 | -0.010591 | -0.010591 | -0.55292
2019-03-21 12:48 | 1553186891.388 | 784 | 0.2095 | 0.000476 | -3e-06 | 100.2395 | 100.5209 | 123.8741 | 21.2612 | 0.002298 | 0.002116 | -152.161 | 0.23951 | 8.4037 | 0.003454 | 0.005146 | 0.00494 | 6.9e-05 | 0.000996 | 0.23075 | 1.3135 | 2.5478 | 99.34 | 20.98 | 37.463 | 77.117 | 104.98 | 72.7439 | 0.009378 | 1.0207 | 0.024596 | 24.0976 | -0.025075 | 0.14249 | -0.000211 | -0.01538 | -0.01538 | -0.38077
2019-03-21 12:48 | 1553186892.3895 | 785 | 0.20949 | 0.000469 | -4e-06 | 100.3563 | 100.5093 | 124.1118 | 21.2612 | 0.002294 | 0.002148 | -152.1651 | 0.24019 | 8.4042 | 0.003473 | 0.004897 | 0.004734 | 6.5e-05 | 0.000998 | 0.27649 | 1.2771 | 2.5478 | 99.34 | 20.98 | 37.463 | 77.0139 | 105.4346 | 73.4947 | 0.009378 | 1.0185 | 0.024593 | 24.1468 | -0.013629 | 0.14289 | -0.000149 | -0.010795 | -0.010795 | -0.54404
2019-03-21 12:48 | 1553186893.3892 | 786 | 0.20947 | 0.000472 | -4e-06 | 100.3301 | 100.369 | 124.1123 | 21.2612 | 0.002213 | 0.002139 | -152.1699 | 0.23946 | 8.4023 | 0.003394 | 0.004418 | 0.004243 | 6.7e-05 | 0.000996 | 0.25874 | 1.317 | 2.548 | 99.34 | 20.98 | 37.463 | 75.0638 | 103.5231 | 73.2881 | 0.009378 | 1.0191 | 0.0246 | 24.1397 | -0.003469 | 0.14245 | -4.4e-05 | -0.003023 | -0.003023 | -1.937
2019-03-21 12:48 | 1553186894.3883 | 787 | 0.2095 | 0.000471 | -3e-06 | 100.2434 | 100.3118 | 124.109 | 21.2612 | 0.002288 | 0.002141 | -152.1754 | 0.23942 | 8.4016 | 0.003325 | 0.004933 | 0.004774 | 6.6e-05 | 0.000993 | 0.27472 | 1.3437 | 2.5481 | 99.34 | 20.98 | 37.461 | 76.8611 | 101.8606 | 73.3246 | 0.009378 | 1.019 | 0.024598 | 24.1406 | -0.006102 | 0.14243 | -0.00019 | -0.013866 | -0.013866 | -0.42218
2019-03-21 12:48 | 1553186895.3892 | 788 | 0.20949 | 0.000466 | -4e-06 | 100.3301 | 100.3339 | 123.8694 | 21.2612 | 0.002312 | 0.002199 | -152.193 | 0.23885 | 8.4032 | 0.003384 | 0.004681 | 0.004586 | 6.6e-05 | 0.000993 | 0.26178 | 1.2978 | 2.5479 | 99.34 | 20.98 | 37.462 | 77.4451 | 103.2715 | 74.7332 | 0.009378 | 1.0148 | 0.024525 | 24.1663 | -0.000343 | 0.14209 | -0.000136 | -0.009754 | -0.009754 | -0.59876
2019-03-21 12:48 | 1553186896.3885 | 789 | 0.2095 | 0.000469 | -3e-06 | 100.4358 | 100.6479 | 123.8637 | 21.2612 | 0.002287 | 0.002105 | -152.2091 | 0.23968 | 8.4034 | 0.003385 | 0.004454 | 0.004242 | 6.4e-05 | 0.000991 | 0.26017 | 1.2984 | 2.548 | 99.34 | 20.98 | 37.463 | 76.8503 | 103.3015 | 72.4796 | 0.009378 | 1.0215 | 0.024604 | 24.0873 | -0.018869 | 0.14259 | -0.000179 | -0.012987 | -0.012987 | -0.45124
2019-03-21 12:48 | 1553186897.3884 | 790 | 0.20948 | 0.000471 | -3e-06 | 100.3667 | 100.4201 | 123.6952 | 21.2612 | 0.002299 | 0.002106 | -152.2023 | 0.24023 | 8.4031 | 0.003368 | 0.004521 | 0.004388 | 6.4e-05 | 0.00099 | 0.26609 | 1.2872 | 2.548 | 99.34 | 20.98 | 37.463 | 77.1396 | 102.8975 | 72.4865 | 0.009378 | 1.0214 | 0.024587 | 24.0712 | -0.004754 | 0.14291 | -7.8e-05 | -0.005592 | -0.005592 | -1.0504
2019-03-21 12:48 | 1553186898.3885 | 791 | 0.20949 | 0.000465 | -3e-06 | 100.4197 | 100.4221 | 123.4739 | 21.2612 | 0.002264 | 0.002103 | -152.2011 | 0.23895 | 8.4015 | 0.003335 | 0.004332 | 0.004171 | 6.5e-05 | 0.000994 | 0.22428 | 1.2856 | 2.548 | 99.34 | 20.99 | 37.463 | 76.3023 | 102.094 | 72.4149 | 0.009384 | 1.0216 | 0.024568 | 24.0472 | -0.000211 | 0.14215 | -0.000161 | -0.01169 | -0.01169 | -0.49979
2019-03-21 12:48 | 1553186899.3875 | 792 | 0.2095 | 0.000472 | -3e-06 | 100.3613 | 100.3148 | 123.4323 | 21.2612 | 0.002237 | 0.002153 | -152.2147 | 0.23956 | 8.401 | 0.003406 | 0.004811 | 0.004636 | 6.2e-05 | 0.000991 | 0.27024 | 1.3284 | 2.5479 | 99.34 | 20.99 | 37.464 | 75.632 | 103.802 | 73.6146 | 0.009384 | 1.0181 | 0.024521 | 24.0848 | 0.004146 | 0.14251 | -0.000207 | -0.015017 | -0.015017 | -0.39005
2019-03-21 12:48 | 1553186900.3876 | 793 | 0.20948 | 0.000466 | -3e-06 | 100.415 | 100.4855 | 123.4178 | 21.2612 | 0.00218 | 0.002082 | -152.2578 | 0.23923 | 8.4024 | 0.00337 | 0.00448 | 0.004374 | 6.3e-05 | 0.000994 | 0.25155 | 1.3637 | 2.548 | 99.34 | 20.99 | 37.463 | 74.2748 | 102.9343 | 71.9106 | 0.009384 | 1.0231 | 0.02458 | 24.0241 | -0.006277 | 0.14232 | -9.6e-05 | -0.006897 | -0.006897 | -0.84812
2019-03-21 12:48 | 1553186901.389 | 794 | 0.20949 | 0.000468 | -3e-06 | 100.381 | 100.417 | 123.5166 | 21.2612 | 0.002224 | 0.002096 | -152.257 | 0.2403 | 8.4028 | 0.003419 | 0.004608 | 0.004381 | 6e-05 | 0.000991 | 0.27096 | 1.291 | 2.5479 | 99.34 | 20.99 | 37.462 | 75.324 | 104.1252 | 72.2434 | 0.009384 | 1.0222 | 0.024578 | 24.0453 | -0.003207 | 0.14296 | -0.000149 | -0.010801 | -0.010801 | -0.54398
2019-03-21 12:48 | 1553186902.3886 | 795 | 0.20951 | 0.000469 | -4e-06 | 100.3935 | 100.3489 | 123.4323 | 21.2612 | 0.002267 | 0.002104 | -152.2653 | 0.24019 | 8.4027 | 0.003342 | 0.004341 | 0.004046 | 6.1e-05 | 0.000995 | 0.29319 | 1.2953 | 2.5479 | 99.34 | 20.99 | 37.46 | 76.3655 | 102.2683 | 72.4351 | 0.009383 | 1.0216 | 0.024563 | 24.0438 | 0.003978 | 0.14289 | -0.00028 | -0.020385 | -0.020385 | -0.28809
2019-03-21 12:48 | 1553186903.3887 | 796 | 0.20949 | 0.000462 | -4e-06 | 100.4647 | 100.5822 | 123.6403 | 21.2612 | 0.002197 | 0.002116 | -152.2637 | 0.23988 | 8.4014 | 0.003332 | 0.004749 | 0.004606 | 6.5e-05 | 0.000997 | 0.29142 | 1.349 | 2.548 | 99.34 | 20.99 | 37.46 | 74.6832 | 102.0303 | 72.7338 | 0.009383 | 1.0207 | 0.024573 | 24.0744 | -0.010452 | 0.1427 | -0.000114 | -0.008183 | -0.008183 | -0.71671
2019-03-21 12:48 | 1553186904.3882 | 797 | 0.20949 | 0.000468 | -3e-06 | 100.4883 | 100.635 | 123.6969 | 21.2612 | 0.0023 | 0.002108 | -152.2555 | 0.23999 | 8.4005 | 0.003331 | 0.004733 | 0.004527 | 6.2e-05 | 0.000996 | 0.29574 | 1.3352 | 2.5481 | 99.34 | 20.99 | 37.459 | 77.1566 | 101.9981 | 72.5444 | 0.009383 | 1.0213 | 0.024585 | 24.0734 | -0.01305 | 0.14277 | -0.000157 | -0.011429 | -0.011429 | -0.5134
2019-03-21 12:48 | 1553186905.3888 | 798 | 0.20948 | 0.000464 | -4e-06 | 100.482 | 100.6796 | 123.8093 | 21.2612 | 0.002247 | 0.002093 | -152.2416 | 0.23995 | 8.4046 | 0.003266 | 0.004946 | 0.004723 | 6.2e-05 | 0.000997 | 0.27848 | 1.3216 | 2.5481 | 99.34 | 20.99 | 37.457 | 75.8893 | 100.4171 | 72.1708 | 0.009382 | 1.0224 | 0.02461 | 24.0713 | -0.017571 | 0.14275 | -6.7e-05 | -0.004702 | -0.004702 | -1.2479
2019-03-21 12:48 | 1553186906.3889 | 799 | 0.20947 | 0.000463 | -3e-06 | 100.5771 | 100.761 | 123.5499 | 21.2612 | 0.002196 | 0.002109 | -152.2215 | 0.24047 | 8.4015 | 0.003339 | 0.004938 | 0.004753 | 6.6e-05 | 0.000998 | 0.27417 | 1.2745 | 2.548 | 99.34 | 20.99 | 37.458 | 74.6554 | 102.1951 | 72.556 | 0.009383 | 1.0212 | 0.02457 | 24.0595 | -0.016344 | 0.14306 | -5e-06 | -0.000202 | -0.000202 | -29.0366
2019-03-21 12:48 | 1553186907.3883 | 800 | 0.20945 | 0.000458 | -4e-06 | 100.5327 | 100.7089 | 123.4422 | 21.2612 | 0.00217 | 0.002084 | -152.2086 | 0.23969 | 8.4022 | 0.003322 | 0.004853 | 0.0046 | 6.3e-05 | 0.000996 | 0.26504 | 1.269 | 2.5478 | 99.34 | 20.99 | 37.457 | 74.0331 | 101.7805 | 71.9714 | 0.009382 | 1.023 | 0.02458 | 24.0286 | -0.015665 | 0.14259 | 0.000112 | 0.008465 | 0.008465 | 0.69233
2019-03-21 12:48 | 1553186908.3884 | 801 | 0.20945 | 0.000461 | -4e-06 | 100.5401 | 100.6408 | 123.4586 | 21.2612 | 0.002213 | 0.002047 | -152.2117 | 0.24017 | 8.4031 | 0.003346 | 0.00375 | 0.003665 | 6.4e-05 | 0.000994 | 0.25608 | 1.2954 | 2.5479 | 99.34 | 20.99 | 37.456 | 75.0559 | 102.3532 | 71.0811 | 0.009382 | 1.0256 | 0.024614 | 23.9992 | -0.008951 | 0.14288 | 0.000133 | 0.010086 | 0.010086 | 0.58224
2019-03-21 12:48 | 1553186909.389 | 802 | 0.20944 | 0.000463 | -4e-06 | 100.4555 | 100.6442 | 123.361 | 21.2612 | 0.00227 | 0.002102 | -152.2125 | 0.2396 | 8.4022 | 0.003374 | 0.004358 | 0.00419 | 6.1e-05 | 0.000995 | 0.26145 | 1.3308 | 2.548 | 99.34 | 20.99 | 37.456 | 76.4387 | 103.0445 | 72.39 | 0.009382 | 1.0217 | 0.024557 | 24.0353 | -0.016785 | 0.14253 | 0.00014 | 0.010559 | 0.010559 | 0.5548
2019-03-21 12:48 | 1553186910.3895 | 803 | 0.20948 | 0.000461 | -3e-06 | 100.5851 | 100.6156 | 123.4185 | 21.2612 | 0.002189 | 0.00209 | -152.1893 | 0.2394 | 8.4043 | 0.003314 | 0.004757 | 0.004571 | 6.1e-05 | 0.000993 | 0.22412 | 1.2924 | 2.548 | 99.34 | 20.99 | 37.454 | 74.4915 | 101.5828 | 72.1062 | 0.009382 | 1.0226 | 0.024573 | 24.031 | -0.002709 | 0.14242 | -8.6e-05 | -0.006133 | -0.006133 | -0.95436
2019-03-21 12:48 | 1553186911.3885 | 804 | 0.20945 | 0.000462 | -3e-06 | 100.5866 | 100.6074 | 123.598 | 21.2612 | 0.00222 | 0.002066 | -152.1951 | 0.24012 | 8.4041 | 0.003329 | 0.004003 | 0.003841 | 6e-05 | 0.000995 | 0.25962 | 1.3384 | 2.5481 | 99.34 | 20.99 | 37.455 | 75.2312 | 101.9383 | 71.524 | 0.009382 | 1.0243 | 0.024612 | 24.0282 | -0.00185 | 0.14284 | 9.4e-05 | 0.007149 | 0.007149 | 0.82121
2019-03-21 12:49 | 1553186912.3882 | 805 | 0.20945 | 0.000468 | -3e-06 | 100.5377 | 100.5931 | 123.8386 | 21.2612 | 0.002216 | 0.002064 | -152.2083 | 0.23931 | 8.4062 | 0.003373 | 0.004935 | 0.004756 | 6e-05 | 0.000992 | 0.27367 | 1.3148 | 2.548 | 99.34 | 20.99 | 37.455 | 75.1495 | 103.006 | 71.4817 | 0.009382 | 1.0244 | 0.024637 | 24.0501 | -0.004924 | 0.14236 | 0.000123 | 0.009339 | 0.009339 | 0.62656
2019-03-21 12:49 | 1553186913.3886 | 806 | 0.20946 | 0.00046 | -3e-06 | 100.5562 | 100.854 | 123.7537 | 21.2612 | 0.002231 | 0.002108 | -152.1928 | 0.2393 | 8.4051 | 0.003388 | 0.005034 | 0.004942 | 6.2e-05 | 0.000992 | 0.26609 | 1.3463 | 2.5481 | 99.34 | 20.99 | 37.456 | 75.4924 | 103.373 | 72.5536 | 0.009382 | 1.0212 | 0.024591 | 24.0792 | -0.026449 | 0.14236 | 4.9e-05 | 0.0038 | 0.0038 | 1.5399
2019-03-21 12:49 | 1553186914.3872 | 807 | 0.20947 | 0.000463 | -3e-06 | 100.6089 | 100.6493 | 123.7001 | 21.2612 | 0.002215 | 0.002065 | -152.197 | 0.23946 | 8.4018 | 0.003383 | 0.005755 | 0.005603 | 6e-05 | 0.000992 | 0.26969 | 1.3333 | 2.548 | 99.34 | 20.99 | 37.456 | 75.1089 | 103.2448 | 71.516 | 0.009382 | 1.0243 | 0.024622 | 24.0379 | -0.003587 | 0.14245 | -1.9e-05 | -0.001206 | -0.001206 | -4.855
2019-03-21 12:49 | 1553186915.3883 | 808 | 0.20946 | 0.000465 | -4e-06 | 100.5729 | 100.5168 | 123.7084 | 21.2612 | 0.002189 | 0.002054 | -152.1897 | 0.2402 | 8.402 | 0.003369 | 0.004621 | 0.004491 | 6.1e-05 | 0.000992 | 0.25946 | 1.386 | 2.5481 | 99.34 | 20.99 | 37.455 | 74.4869 | 102.9227 | 71.2455 | 0.009382 | 1.0251 | 0.024633 | 24.0292 | 0.004989 | 0.14289 | 3.6e-05 | 0.002873 | 0.002873 | 2.0439
2019-03-21 12:49 | 1553186916.3874 | 809 | 0.20947 | 0.000465 | -3e-06 | 100.5374 | 100.6731 | 124.159 | 21.2612 | 0.00216 | 0.002128 | -152.2094 | 0.23991 | 8.4026 | 0.003413 | 0.004921 | 0.004669 | 5.9e-05 | 0.000993 | 0.29734 | 1.3064 | 2.5479 | 99.34 | 20.99 | 37.456 | 73.8032 | 103.9915 | 73.0324 | 0.009382 | 1.0198 | 0.024614 | 24.1353 | -0.012065 | 0.14272 | -3.8e-05 | -0.002625 | -0.002625 | -2.2346
2019-03-21 12:49 | 1553186917.3874 | 810 | 0.20953 | 0.000459 | -4e-06 | 100.67 | 100.4313 | 124.47 | 21.2612 | 0.00221 | 0.002076 | -152.216 | 0.23961 | 8.4023 | 0.003485 | 0.004582 | 0.004388 | 6e-05 | 0.000993 | 0.29015 | 1.2795 | 2.548 | 99.34 | 20.99 | 37.457 | 75.002 | 105.7273 | 71.7672 | 0.009382 | 1.0236 | 0.02469 | 24.1213 | 0.021234 | 0.14255 | -0.000339 | -0.024914 | -0.024914 | -0.23516
2019-03-21 12:49 | 1553186918.3885 | 811 | 0.20946 | 0.000462 | -4e-06 | 100.6283 | 100.8305 | 124.5266 | 21.2612 | 0.002191 | 0.00209 | -152.2272 | 0.23862 | 8.4018 | 0.003433 | 0.005429 | 0.005149 | 6.1e-05 | 0.000994 | 0.31365 | 1.2966 | 2.5479 | 99.34 | 20.99 | 37.457 | 74.5439 | 104.4555 | 72.1186 | 0.009382 | 1.0225 | 0.024683 | 24.1391 | -0.017953 | 0.14195 | 6.1e-05 | 0.004754 | 0.004754 | 1.2274
2019-03-21 12:49 | 1553186919.389 | 812 | 0.20946 | 0.000464 | -4e-06 | 100.5991 | 100.7365 | 124.6857 | 21.2612 | 0.00217 | 0.002096 | -152.2203 | 0.23972 | 8.4033 | 0.003492 | 0.004225 | 0.004177 | 6e-05 | 0.000993 | 0.28152 | 1.3434 | 2.5481 | 99.34 | 20.99 | 37.457 | 74.0341 | 105.8894 | 72.2607 | 0.009382 | 1.0221 | 0.024693 | 24.1595 | -0.012206 | 0.14261 | 6.9e-05 | 0.005374 | 0.005374 | 1.0907
2019-03-21 12:49 | 1553186920.3876 | 813 | 0.20948 | 0.000459 | -4e-06 | 100.5994 | 100.5621 | 124.6681 | 21.2612 | 0.002197 | 0.002073 | -152.196 | 0.2391 | 8.4021 | 0.003429 | 0.004925 | 0.004645 | 6e-05 | 0.000994 | 0.27168 | 1.2698 | 2.548 | 99.34 | 20.99 | 37.457 | 74.6748 | 104.3722 | 71.7032 | 0.009382 | 1.0238 | 0.024712 | 24.1383 | 0.003318 | 0.14224 | -5e-05 | -0.003498 | -0.003498 | -1.6714
2019-03-21 12:49 | 1553186921.3886 | 814 | 0.20946 | 0.000457 | -4e-06 | 100.6042 | 100.7092 | 124.1525 | 21.2612 | 0.002167 | 0.002062 | -152.2182 | 0.23878 | 8.4035 | 0.003426 | 0.004639 | 0.004517 | 6.3e-05 | 0.000996 | 0.30845 | 1.2997 | 2.548 | 99.34 | 20.99 | 37.459 | 73.9692 | 104.3021 | 71.449 | 0.009383 | 1.0245 | 0.02467 | 24.0794 | -0.009336 | 0.14205 | 3.8e-05 | 0.003043 | 0.003043 | 1.9188
2019-03-21 12:49 | 1553186922.3888 | 815 | 0.20946 | 0.000458 | -4e-06 | 100.5315 | 100.6057 | 123.8532 | 21.2612 | 0.002201 | 0.002067 | -152.2153 | 0.23902 | 8.4041 | 0.003418 | 0.0046 | 0.004479 | 6.2e-05 | 0.000996 | 0.28119 | 1.3169 | 2.5481 | 99.34 | 20.99 | 37.459 | 74.7729 | 104.0976 | 71.5504 | 0.009383 | 1.0242 | 0.024636 | 24.0539 | -0.006601 | 0.14219 | 7.5e-05 | 0.005753 | 0.005753 | 1.0158
2019-03-21 12:49 | 1553186923.3888 | 816 | 0.20946 | 0.000458 | -4e-06 | 100.6602 | 100.6636 | 123.7736 | 21.2612 | 0.002188 | 0.002047 | -152.2362 | 0.24049 | 8.4036 | 0.003373 | 0.005181 | 0.004971 | 6.3e-05 | 0.000997 | 0.29502 | 1.3158 | 2.5479 | 99.34 | 20.99 | 37.458 | 74.4737 | 103.0029 | 71.0811 | 0.009383 | 1.0256 | 0.024645 | 24.0298 | -0.000303 | 0.14306 | 5e-05 | 0.003917 | 0.003917 | 1.501
2019-03-21 12:49 | 1553186924.3878 | 817 | 0.20947 | 0.000465 | -3e-06 | 100.5842 | 100.7477 | 123.8835 | 21.2612 | 0.002198 | 0.00207 | -152.2436 | 0.23912 | 8.4045 | 0.003412 | 0.004048 | 0.003935 | 6.4e-05 | 0.001 | 0.29374 | 1.3423 | 2.5481 | 99.34 | 20.99 | 37.459 | 74.7089 | 103.968 | 71.6376 | 0.009383 | 1.0239 | 0.024636 | 24.0599 | -0.014529 | 0.14225 | 2.4e-05 | 0.001973 | 0.001973 | 2.9635
2019-03-21 12:49 | 1553186925.3879 | 818 | 0.20946 | 0.000462 | -3e-06 | 100.488 | 100.6411 | 124.0609 | 21.2612 | 0.002183 | 0.002084 | -152.2569 | 0.23959 | 8.4055 | 0.003333 | 0.004843 | 0.004638 | 6.5e-05 | 0.001 | 0.26288 | 1.3038 | 2.5479 | 99.34 | 20.99 | 37.459 | 74.3411 | 102.0531 | 71.9776 | 0.009383 | 1.0229 | 0.024642 | 24.089 | -0.013621 | 0.14253 | 8.9e-05 | 0.006798 | 0.006798 | 0.86175
2019-03-21 12:49 | 1553186926.3889 | 819 | 0.20947 | 0.000462 | -3e-06 | 100.5666 | 100.7068 | 124.0464 | 21.2612 | 0.002152 | 0.002077 | -152.2468 | 0.2401 | 8.4008 | 0.003343 | 0.00424 | 0.004128 | 6.1e-05 | 0.000999 | 0.29789 | 1.2937 | 2.548 | 99.34 | 20.99 | 37.459 | 73.6038 | 102.2801 | 71.7905 | 0.009383 | 1.0235 | 0.024647 | 24.081 | -0.012463 | 0.14283 | 2.9e-05 | 0.002366 | 0.002366 | 2.4811
2019-03-21 12:49 | 1553186927.387 | 820 | 0.20949 | 0.000466 | -3e-06 | 100.5359 | 100.5256 | 123.9105 | 21.2612 | 0.002142 | 0.002087 | -152.2596 | 0.23985 | 8.4046 | 0.00339 | 0.003782 | 0.003529 | 6.2e-05 | 0.000997 | 0.25946 | 1.3013 | 2.5479 | 99.34 | 20.99 | 37.459 | 73.3698 | 103.4129 | 72.0408 | 0.009383 | 1.0228 | 0.024624 | 24.0766 | 0.000915 | 0.14269 | -0.00012 | -0.008714 | -0.008714 | -0.673
2019-03-21 12:49 | 1553186928.388 | 821 | 0.20944 | 0.000462 | -3e-06 | 100.5616 | 100.3581 | 123.9189 | 21.2612 | 0.002138 | 0.002051 | -152.2429 | 0.23922 | 8.4024 | 0.003334 | 0.004085 | 0.003862 | 6.3e-05 | 0.001 | 0.26952 | 1.2768 | 2.5479 | 99.34 | 20.99 | 37.459 | 73.2708 | 102.0569 | 71.1683 | 0.009383 | 1.0253 | 0.024656 | 24.047 | 0.018117 | 0.14231 | 0.000156 | 0.011726 | 0.011726 | 0.49882
2019-03-21 12:49 | 1553186929.3884 | 822 | 0.20947 | 0.000463 | -3e-06 | 100.4832 | 100.573 | 123.4835 | 21.2612 | 0.00212 | 0.00199 | -152.2526 | 0.24019 | 8.4009 | 0.003325 | 0.004515 | 0.004368 | 6.2e-05 | 0.000999 | 0.23114 | 1.335 | 2.548 | 99.34 | 20.99 | 37.458 | 72.8313 | 101.8497 | 69.718 | 0.009383 | 1.0297 | 0.024665 | 23.9541 | -0.007989 | 0.14289 | 7e-06 | 0.000723 | 0.000723 | 8.1172
2019-03-21 12:49 | 1553186930.3876 | 823 | 0.20945 | 0.000468 | -3e-06 | 100.5184 | 100.5832 | 123.5079 | 21.2612 | 0.002138 | 0.002014 | -152.272 | 0.23965 | 8.4018 | 0.003343 | 0.005167 | 0.004994 | 6e-05 | 0.001003 | 0.27345 | 1.3461 | 2.5479 | 99.34 | 20.99 | 37.457 | 73.27 | 102.2927 | 70.2804 | 0.009382 | 1.028 | 0.024647 | 23.9761 | -0.005769 | 0.14257 | 0.000119 | 0.008996 | 0.008996 | 0.65135
2019-03-21 12:49 | 1553186931.3881 | 824 | 0.20947 | 0.000464 | -3e-06 | 100.4895 | 100.6353 | 123.741 | 21.2612 | 0.002118 | 0.002042 | -152.246 | 0.23951 | 8.4046 | 0.003377 | 0.005128 | 0.004921 | 5.8e-05 | 0.001003 | 0.26842 | 1.3328 | 2.5479 | 99.34 | 20.99 | 37.455 | 72.7744 | 103.0978 | 70.9617 | 0.009382 | 1.026 | 0.024646 | 24.0225 | -0.012974 | 0.14249 | 8e-06 | 0.000802 | 0.000802 | 7.299
2019-03-21 12:49 | 1553186932.3886 | 825 | 0.20948 | 0.000461 | -3e-06 | 100.5339 | 100.6718 | 124.0597 | 21.2612 | 0.00215 | 0.002043 | -152.206 | 0.23943 | 8.4044 | 0.003322 | 0.004778 | 0.00458 | 6.1e-05 | 0.001004 | 0.24729 | 1.3452 | 2.5479 | 99.34 | 21 | 37.455 | 73.5601 | 101.7735 | 70.9758 | 0.009388 | 1.0259 | 0.024677 | 24.0539 | -0.012262 | 0.14244 | -3.7e-05 | -0.002565 | -0.002565 | -2.2823
2019-03-21 12:49 | 1553186933.3886 | 826 | 0.20947 | 0.000458 | -4e-06 | 100.4346 | 100.433 | 124.402 | 21.2612 | 0.002116 | 0.002033 | -152.2026 | 0.23997 | 8.4025 | 0.003386 | 0.004455 | 0.004246 | 6.4e-05 | 0.001005 | 0.30845 | 1.3045 | 2.548 | 99.34 | 21 | 37.455 | 72.7261 | 103.3369 | 70.7435 | 0.009388 | 1.0266 | 0.02472 | 24.0789 | 0.000145 | 0.14276 | 1e-06 | 0.000316 | 0.000316 | 18.5449
2019-03-21 12:49 | 1553186934.3873 | 827 | 0.20944 | 0.000457 | -4e-06 | 100.4662 | 100.6425 | 124.4691 | 21.2612 | 0.002171 | 0.002052 | -152.1982 | 0.23943 | 8.4004 | 0.003379 | 0.005009 | 0.004852 | 6.3e-05 | 0.001002 | 0.2808 | 1.332 | 2.548 | 99.34 | 21 | 37.454 | 74.0605 | 103.1545 | 71.1957 | 0.009387 | 1.0253 | 0.02471 | 24.1012 | -0.015678 | 0.14244 | 0.000157 | 0.011846 | 0.011846 | 0.4942
2019-03-21 12:49 | 1553186935.3888 | 828 | 0.2095 | 0.000462 | -4e-06 | 100.4662 | 100.6857 | 124.2907 | 21.2612 | 0.002153 | 0.001992 | -152.179 | 0.24002 | 8.4012 | 0.00341 | 0.004588 | 0.004374 | 6.3e-05 | 0.001003 | 0.27295 | 1.2959 | 2.548 | 99.34 | 21 | 37.453 | 73.6256 | 103.9113 | 69.7572 | 0.009387 | 1.0296 | 0.024744 | 24.0336 | -0.019522 | 0.14279 | -0.000165 | -0.012019 | -0.012019 | -0.48828
2019-03-21 12:49 | 1553186936.3885 | 829 | 0.20945 | 0.00046 | -4e-06 | 100.3944 | 100.5134 | 124.0323 | 21.2612 | 0.002105 | 0.002073 | -152.1773 | 0.23982 | 8.4023 | 0.003375 | 0.004463 | 0.004314 | 6.3e-05 | 0.001006 | 0.22395 | 1.3542 | 2.548 | 99.34 | 21 | 37.45 | 72.4688 | 103.0642 | 71.6915 | 0.009386 | 1.0238 | 0.024649 | 24.0762 | -0.010594 | 0.14267 | 0.000116 | 0.008772 | 0.008772 | 0.66847
2019-03-21 12:49 | 1553186937.3885 | 830 | 0.20946 | 0.000462 | -3e-06 | 100.3575 | 100.4235 | 123.9224 | 21.2612 | 0.002129 | 0.001974 | -152.1996 | 0.23995 | 8.4003 | 0.003365 | 0.00485 | 0.004732 | 6.4e-05 | 0.001007 | 0.30558 | 1.3231 | 2.548 | 99.34 | 21 | 37.449 | 73.0533 | 102.8096 | 69.3164 | 0.009386 | 1.0309 | 0.024723 | 23.9825 | -0.005881 | 0.14275 | 5.4e-05 | 0.004213 | 0.004213 | 1.3927
2019-03-21 12:49 | 1553186938.3879 | 831 | 0.20947 | 0.000463 | -3e-06 | 100.3599 | 100.3499 | 123.7412 | 21.2612 | 0.002171 | 0.00206 | -152.2029 | 0.23984 | 8.4023 | 0.003462 | 0.0048 | 0.004626 | 6e-05 | 0.001008 | 0.30397 | 1.312 | 2.548 | 99.34 | 21 | 37.449 | 74.0567 | 105.1624 | 71.3937 | 0.009386 | 1.0247 | 0.024631 | 24.0376 | 0.000886 | 0.14268 | 1e-05 | 0.000936 | 0.000936 | 6.268
2019-03-21 12:49 | 1553186939.3877 | 832 | 0.20949 | 0.00046 | -3e-06 | 100.367 | 100.261 | 123.8037 | 21.2612 | 0.002137 | 0.001999 | -152.206 | 0.23887 | 8.4021 | 0.003399 | 0.005112 | 0.004888 | 6.2e-05 | 0.001008 | 0.27721 | 1.3152 | 2.548 | 99.34 | 21 | 37.446 | 73.228 | 103.6503 | 69.9326 | 0.009385 | 1.029 | 0.024689 | 23.9926 | 0.009451 | 0.1421 | -0.000129 | -0.009401 | -0.009401 | -0.62128
2019-03-21 12:49 | 1553186940.3887 | 833 | 0.20945 | 0.000465 | -3e-06 | 100.3342 | 100.4838 | 123.5764 | 21.2612 | 0.002075 | 0.00203 | -152.2069 | 0.23959 | 8.4034 | 0.00344 | 0.004442 | 0.004234 | 6.2e-05 | 0.001012 | 0.29231 | 1.3127 | 2.548 | 99.34 | 21 | 37.444 | 71.753 | 104.6261 | 70.6771 | 0.009385 | 1.0268 | 0.02464 | 23.9966 | -0.01332 | 0.14253 | 0.000123 | 0.009287 | 0.009287 | 0.63079
2019-03-21 12:49 | 1553186941.3877 | 834 | 0.20949 | 0.000465 | -3e-06 | 100.2898 | 100.3169 | 123.6452 | 21.2612 | 0.002147 | 0.001996 | -152.2194 | 0.2395 | 8.4024 | 0.003411 | 0.004406 | 0.004233 | 6e-05 | 0.001008 | 0.23828 | 1.3115 | 2.5479 | 99.34 | 21 | 37.442 | 73.4813 | 103.9357 | 69.8437 | 0.009384 | 1.0293 | 0.024677 | 23.9741 | -0.002411 | 0.14248 | -0.000111 | -0.007994 | -0.007994 | -0.73255
2019-03-21 12:49 | 1553186942.3887 | 835 | 0.20947 | 0.000458 | -3e-06 | 100.2547 | 100.2883 | 123.438 | 21.2612 | 0.002121 | 0.001986 | -152.2491 | 0.23931 | 8.4045 | 0.003403 | 0.004772 | 0.004653 | 6.1e-05 | 0.001009 | 0.25713 | 1.3305 | 2.5479 | 99.34 | 21 | 37.442 | 72.8508 | 103.747 | 69.6113 | 0.009384 | 1.03 | 0.024664 | 23.9459 | -0.002996 | 0.14236 | -2.5e-05 | -0.00168 | -0.00168 | -3.4823
2019-03-21 12:49 | 1553186943.3868 | 836 | 0.20948 | 0.000461 | -4e-06 | 100.2952 | 100.3203 | 123.356 | 21.2612 | 0.002067 | 0.002006 | -152.251 | 0.2402 | 8.4017 | 0.00333 | 0.005434 | 0.005215 | 6.4e-05 | 0.001007 | 0.29806 | 1.2938 | 2.548 | 99.34 | 21 | 37.441 | 71.5605 | 101.9792 | 70.0939 | 0.009384 | 1.0286 | 0.024639 | 23.9548 | -0.002236 | 0.14289 | -6.1e-05 | -0.004275 | -0.004275 | -1.3737
2019-03-21 12:49 | 1553186944.3879 | 837 | 0.20947 | 0.000463 | -3e-06 | 100.241 | 100.23 | 123.3945 | 21.2612 | 0.002107 | 0.001966 | -152.2458 | 0.23967 | 8.402 | 0.003377 | 0.004655 | 0.00449 | 6e-05 | 0.001007 | 0.26034 | 1.2761 | 2.548 | 99.34 | 21 | 37.436 | 72.5185 | 103.1041 | 69.1332 | 0.009383 | 1.0314 | 0.024677 | 23.925 | 0.000977 | 0.14258 | -3.4e-05 | -0.002321 | -0.002321 | -2.5248
2019-03-21 12:49 | 1553186945.3873 | 838 | 0.20945 | 0.000459 | -3e-06 | 100.2657 | 100.3312 | 123.4441 | 21.2612 | 0.00208 | 0.002019 | -152.2644 | 0.23953 | 8.4019 | 0.003377 | 0.004611 | 0.004503 | 6.2e-05 | 0.001003 | 0.30431 | 1.334 | 2.5481 | 99.34 | 21 | 37.434 | 71.8662 | 103.1143 | 70.395 | 0.009382 | 1.0277 | 0.024637 | 23.9739 | -0.005839 | 0.1425 | 9.8e-05 | 0.007479 | 0.007479 | 0.78307
2019-03-21 12:49 | 1553186946.388 | 839 | 0.20945 | 0.000463 | -3e-06 | 100.3137 | 100.3986 | 123.591 | 21.2612 | 0.002091 | 0.001969 | -152.2666 | 0.23936 | 8.4031 | 0.003386 | 0.005088 | 0.004901 | 6.2e-05 | 0.001 | 0.29358 | 1.3059 | 2.5481 | 99.34 | 21 | 37.43 | 72.1451 | 103.318 | 69.1957 | 0.009381 | 1.0313 | 0.024695 | 23.9462 | -0.007571 | 0.14239 | 0.000116 | 0.008809 | 0.008809 | 0.66438
2019-03-21 12:49 | 1553186947.388 | 840 | 0.20946 | 0.000461 | -4e-06 | 100.2973 | 100.3829 | 123.6816 | 21.2612 | 0.002094 | 0.002043 | -152.2539 | 0.23881 | 8.4032 | 0.003448 | 0.004602 | 0.004397 | 6.3e-05 | 0.001 | 0.26609 | 1.3217 | 2.5482 | 99.34 | 21 | 37.424 | 72.1997 | 104.8273 | 70.9884 | 0.00938 | 1.0259 | 0.024639 | 24.0176 | -0.007636 | 0.14207 | 6.1e-05 | 0.004756 | 0.004756 | 1.2277
2019-03-21 12:49 | 1553186948.3871 | 841 | 0.2095 | 0.000461 | -4e-06 | 100.2866 | 100.2797 | 123.5591 | 21.2612 | 0.002078 | 0.002017 | -152.2719 | 0.23852 | 8.4007 | 0.00335 | 0.004633 | 0.004473 | 6.3e-05 | 0.000999 | 0.31045 | 1.3035 | 2.5481 | 99.34 | 21 | 37.42 | 71.8231 | 102.4506 | 70.3465 | 0.009379 | 1.0278 | 0.02465 | 23.9833 | 0.000607 | 0.14189 | -0.000173 | -0.012557 | -0.012557 | -0.46443
2019-03-21 12:49 | 1553186949.3874 | 842 | 0.20946 | 0.000462 | -3e-06 | 100.2422 | 100.387 | 123.5272 | 21.2612 | 0.002163 | 0.002009 | -152.2488 | 0.23939 | 8.4015 | 0.003495 | 0.005118 | 0.004938 | 6.4e-05 | 0.001 | 0.28783 | 1.3392 | 2.548 | 99.34 | 21 | 37.414 | 73.857 | 105.9721 | 70.1525 | 0.009377 | 1.0284 | 0.024654 | 23.9735 | -0.012917 | 0.14242 | 7.5e-05 | 0.005697 | 0.005697 | 1.0274
2019-03-21 12:49 | 1553186950.3872 | 843 | 0.20949 | 0.000457 | -3e-06 | 100.2544 | 100.2985 | 123.3952 | 21.2612 | 0.002083 | 0.002013 | -152.2444 | 0.23911 | 8.4042 | 0.003516 | 0.005049 | 0.004864 | 6.1e-05 | 0.001002 | 0.27936 | 1.3268 | 2.5478 | 99.34 | 21.01 | 37.41 | 71.9433 | 106.483 | 70.2638 | 0.009382 | 1.028 | 0.024637 | 23.9646 | -0.003934 | 0.14225 | -0.000142 | -0.010306 | -0.010306 | -0.56726
2019-03-21 12:49 | 1553186951.3873 | 844 | 0.2095 | 0.000464 | -3e-06 | 100.1864 | 100.4017 | 123.2007 | 21.2612 | 0.002119 | 0.002009 | -152.2358 | 0.239 | 8.4008 | 0.003469 | 0.004961 | 0.004772 | 6e-05 | 0.001002 | 0.27417 | 1.3212 | 2.5481 | 99.34 | 21.01 | 37.407 | 72.8023 | 105.3495 | 70.1734 | 0.009381 | 1.0283 | 0.02462 | 23.9425 | -0.019195 | 0.14218 | -0.000188 | -0.013683 | -0.013683 | -0.42708
2019-03-21 12:49 | 1553186952.3883 | 845 | 0.20952 | 0.000461 | -3e-06 | 100.1486 | 100.2222 | 123.2058 | 21.2612 | 0.00208 | 0.002016 | -152.2134 | 0.2395 | 8.4022 | 0.003427 | 0.004758 | 0.004605 | 6.2e-05 | 0.001003 | 0.26864 | 1.3275 | 2.548 | 99.34 | 21.01 | 37.399 | 71.8731 | 104.313 | 70.3357 | 0.009379 | 1.0278 | 0.024615 | 23.9487 | -0.006569 | 0.14248 | -0.000316 | -0.023186 | -0.023186 | -0.25256
2019-03-21 12:49 | 1553186953.3877 | 846 | 0.2095 | 0.000465 | -3e-06 | 100.1516 | 100.1609 | 123.1697 | 21.2612 | 0.002098 | 0.001984 | -152.2164 | 0.23841 | 8.4028 | 0.003442 | 0.00483 | 0.004597 | 6.2e-05 | 0.001001 | 0.23092 | 1.3559 | 2.5482 | 99.34 | 21.01 | 37.396 | 72.3136 | 104.6843 | 69.5714 | 0.009379 | 1.0301 | 0.024639 | 23.9185 | -0.00083 | 0.14183 | -0.000162 | -0.011819 | -0.011819 | -0.49322
2019-03-21 12:49 | 1553186954.3874 | 847 | 0.2095 | 0.00046 | -3e-06 | 100.1617 | 100.2958 | 123.3197 | 21.2612 | 0.00208 | 0.002072 | -152.2317 | 0.23908 | 8.4029 | 0.003516 | 0.004871 | 0.004736 | 6.1e-05 | 0.001 | 0.25371 | 1.3644 | 2.548 | 99.34 | 21.01 | 37.389 | 71.8669 | 106.4883 | 71.6689 | 0.009377 | 1.0239 | 0.024579 | 24.0062 | -0.011962 | 0.14223 | -0.000191 | -0.013884 | -0.013884 | -0.42103
2019-03-21 12:49 | 1553186955.3874 | 848 | 0.20947 | 0.000464 | -4e-06 | 100.1754 | 100.1554 | 123.4105 | 21.2612 | 0.00208 | 0.00202 | -152.2221 | 0.23904 | 8.4019 | 0.003511 | 0.004786 | 0.004631 | 6.2e-05 | 0.001006 | 0.27831 | 1.367 | 2.5481 | 99.34 | 21.01 | 37.385 | 71.8786 | 106.3663 | 70.4371 | 0.009376 | 1.0275 | 0.024632 | 23.9721 | 0.001785 | 0.14221 | 1.6e-05 | 0.001434 | 0.001434 | 4.0763
2019-03-21 12:49 | 1553186956.3874 | 849 | 0.20948 | 0.000461 | -3e-06 | 100.193 | 100.0267 | 123.2673 | 21.2612 | 0.00212 | 0.002049 | -152.2117 | 0.23936 | 8.4038 | 0.003479 | 0.005176 | 0.00506 | 6.4e-05 | 0.001009 | 0.2756 | 1.3141 | 2.548 | 99.34 | 21.01 | 37.383 | 72.822 | 105.5714 | 71.1208 | 0.009375 | 1.0255 | 0.024593 | 23.982 | 0.01486 | 0.14239 | -6.4e-05 | -0.004497 | -0.004497 | -1.3014
2019-03-21 12:49 | 1553186957.3876 | 850 | 0.20944 | 0.000458 | -4e-06 | 100.168 | 100.275 | 123.2079 | 21.2612 | 0.002072 | 0.002053 | -152.2101 | 0.23935 | 8.4042 | 0.003497 | 0.005223 | 0.004995 | 6.3e-05 | 0.001011 | 0.26233 | 1.3008 | 2.5479 | 99.34 | 21.01 | 37.378 | 71.6821 | 106.0177 | 71.2323 | 0.009374 | 1.0252 | 0.024583 | 23.9801 | -0.00955 | 0.14239 | 0.00016 | 0.012022 | 0.012022 | 0.48678
2019-03-21 12:49 | 1553186958.3876 | 851 | 0.20954 | 0.000462 | -3e-06 | 100.1105 | 100.2481 | 123.38 | 21.2612 | 0.002131 | 0.002001 | -152.1866 | 0.2401 | 8.4041 | 0.003529 | 0.004832 | 0.004721 | 6.1e-05 | 0.001014 | 0.25912 | 1.2795 | 2.5479 | 99.34 | 21.01 | 37.376 | 73.1009 | 106.804 | 69.9824 | 0.009374 | 1.0289 | 0.024645 | 23.9533 | -0.012286 | 0.14283 | -0.000442 | -0.032503 | -0.032503 | -0.18061
2019-03-21 12:49 | 1553186959.387 | 852 | 0.20953 | 0.000464 | -3e-06 | 100.168 | 100.3036 | 123.33 | 21.2612 | 0.002109 | 0.002049 | -152.2128 | 0.23916 | 8.4013 | 0.003438 | 0.004865 | 0.004633 | 6e-05 | 0.001019 | 0.27632 | 1.2736 | 2.548 | 99.34 | 21.01 | 37.375 | 72.5567 | 104.5835 | 71.1357 | 0.009373 | 1.0254 | 0.024599 | 23.9886 | -0.012101 | 0.14228 | -0.000374 | -0.02746 | -0.02746 | -0.21295
2019-03-21 12:49 | 1553186960.3868 | 853 | 0.20953 | 0.000464 | -4e-06 | 100.0807 | 100.2671 | 123.7673 | 21.2612 | 0.002088 | 0.00196 | -152.2119 | 0.23882 | 8.4024 | 0.003426 | 0.005102 | 0.004928 | 6.1e-05 | 0.001013 | 0.25066 | 1.3142 | 2.5479 | 99.34 | 21.02 | 37.372 | 72.0572 | 104.2902 | 68.9748 | 0.009378 | 1.0319 | 0.02472 | 23.9556 | -0.01665 | 0.14208 | -0.000366 | -0.026917 | -0.026917 | -0.21694
2019-03-21 12:49 | 1553186961.3868 | 854 | 0.20951 | 0.000463 | -4e-06 | 100.0819 | 100.3032 | 123.8156 | 21.2612 | 0.002113 | 0.001944 | -152.2202 | 0.23918 | 8.4039 | 0.003409 | 0.004719 | 0.004619 | 5.9e-05 | 0.001014 | 0.29502 | 1.2902 | 2.548 | 99.34 | 21.02 | 37.37 | 72.6644 | 103.8902 | 68.5981 | 0.009378 | 1.0331 | 0.024739 | 23.9471 | -0.019763 | 0.14229 | -0.000224 | -0.016433 | -0.016433 | -0.35588
2019-03-21 12:49 | 1553186962.3868 | 855 | 0.20946 | 0.000464 | -4e-06 | 100.0789 | 100.2229 | 124.0386 | 21.2612 | 0.00206 | 0.001976 | -152.2198 | 0.23855 | 8.4023 | 0.003452 | 0.004341 | 0.004199 | 6e-05 | 0.001015 | 0.26626 | 1.3118 | 2.5479 | 99.34 | 21.02 | 37.366 | 71.3858 | 104.9367 | 69.3695 | 0.009377 | 1.0307 | 0.024733 | 23.9956 | -0.012859 | 0.14192 | 8.7e-05 | 0.006686 | 0.006686 | 0.87239
2019-03-21 12:49 | 1553186963.3868 | 856 | 0.20947 | 0.000462 | -3e-06 | 99.9731 | 99.9463 | 123.9691 | 21.2612 | 0.002047 | 0.00193 | -152.2141 | 0.24013 | 8.4031 | 0.003357 | 0.004363 | 0.004117 | 5.8e-05 | 0.001009 | 0.27544 | 1.3233 | 2.548 | 99.34 | 21.02 | 37.365 | 71.0739 | 102.6329 | 68.2551 | 0.009377 | 1.0341 | 0.024766 | 23.9499 | 0.002399 | 0.14286 | 1.2e-05 | 0.001075 | 0.001075 | 5.4594
2019-03-21 12:49 | 1553186964.3879 | 857 | 0.20948 | 0.000459 | -3e-06 | 100.1003 | 100.0481 | 123.6933 | 21.2612 | 0.002059 | 0.001998 | -152.2002 | 0.23903 | 8.4022 | 0.003371 | 0.004378 | 0.00415 | 6.3e-05 | 0.001019 | 0.2557 | 1.3472 | 2.5479 | 99.34 | 21.02 | 37.363 | 71.3546 | 102.9574 | 69.8882 | 0.009376 | 1.0292 | 0.02468 | 23.9804 | 0.004665 | 0.1422 | -7.8e-05 | -0.00556 | -0.00556 | -1.0511
2019-03-21 12:49 | 1553186965.3874 | 858 | 0.20944 | 0.000462 | -3e-06 | 100.0607 | 100.0757 | 123.4396 | 21.2612 | 0.002054 | 0.001993 | -152.2117 | 0.239 | 8.4036 | 0.003399 | 0.005 | 0.004828 | 6.2e-05 | 0.001019 | 0.26897 | 1.3283 | 2.548 | 99.34 | 21.02 | 37.359 | 71.2565 | 103.6535 | 69.7711 | 0.009375 | 1.0295 | 0.024659 | 23.9517 | -0.001342 | 0.14218 | 0.0002 | 0.015043 | 0.015043 | 0.38846
2019-03-21 12:49 | 1553186966.388 | 859 | 0.20951 | 0.000462 | -3e-06 | 100.0494 | 100.0703 | 123.4044 | 21.2612 | 0.002031 | 0.002021 | -152.2171 | 0.2395 | 8.4044 | 0.003368 | 0.004236 | 0.004051 | 6.1e-05 | 0.001017 | 0.24475 | 1.3779 | 2.548 | 99.34 | 21.02 | 37.358 | 70.7045 | 102.8881 | 70.4611 | 0.009375 | 1.0275 | 0.024631 | 23.9723 | -0.001867 | 0.14248 | -0.000218 | -0.01591 | -0.01591 | -0.36806
2019-03-21 12:49 | 1553186967.3881 | 860 | 0.20946 | 0.00046 | -3e-06 | 100.0229 | 100.199 | 123.7389 | 21.2612 | 0.002102 | 0.001935 | -152.2031 | 0.24082 | 8.401 | 0.003393 | 0.004864 | 0.004624 | 6.2e-05 | 0.001017 | 0.286 | 1.3409 | 2.548 | 99.34 | 21.02 | 37.355 | 72.3885 | 103.4861 | 68.383 | 0.009374 | 1.0337 | 0.024739 | 23.9321 | -0.015738 | 0.14326 | 5.7e-05 | 0.004405 | 0.004405 | 1.3366
2019-03-21 12:49 | 1553186968.3871 | 861 | 0.20944 | 0.000461 | -3e-06 | 100.03 | 99.9593 | 123.8978 | 21.2612 | 0.002127 | 0.00193 | -152.2083 | 0.24021 | 8.4016 | 0.003379 | 0.004992 | 0.004861 | 5.8e-05 | 0.001016 | 0.27256 | 1.3465 | 2.5479 | 99.34 | 21.02 | 37.354 | 72.9918 | 103.1552 | 68.2628 | 0.009374 | 1.0341 | 0.024759 | 23.9432 | 0.006331 | 0.1429 | 0.000218 | 0.016424 | 0.016424 | 0.35761
2019-03-21 12:49 | 1553186969.3881 | 862 | 0.20946 | 0.000462 | -3e-06 | 100.0262 | 99.8853 | 123.9877 | 21.2612 | 0.002074 | 0.001968 | -152.2007 | 0.2392 | 8.402 | 0.003338 | 0.00508 | 0.004933 | 6.2e-05 | 0.001016 | 0.26178 | 1.311 | 2.548 | 99.34 | 21.02 | 37.353 | 71.7179 | 102.1608 | 69.176 | 0.009374 | 1.0313 | 0.024735 | 23.9839 | 0.0126 | 0.1423 | 9.3e-05 | 0.007154 | 0.007154 | 0.81756
2019-03-21 12:49 | 1553186970.3883 | 863 | 0.20947 | 0.000462 | -4e-06 | 99.9663 | 100.0321 | 123.722 | 21.2612 | 0.002023 | 0.001975 | -152.1869 | 0.24058 | 8.4026 | 0.003329 | 0.004385 | 0.004208 | 6.4e-05 | 0.001015 | 0.27455 | 1.384 | 2.548 | 99.34 | 21.02 | 37.351 | 70.508 | 101.9478 | 69.3361 | 0.009373 | 1.0308 | 0.024703 | 23.9638 | -0.005892 | 0.14312 | 2.1e-05 | 0.001806 | 0.001806 | 3.2573
2019-03-21 12:50 | 1553186971.3873 | 864 | 0.20945 | 0.000456 | -3e-06 | 100.0157 | 100.0478 | 123.8076 | 21.2612 | 0.002026 | 0.00194 | -152.175 | 0.23956 | 8.4033 | 0.003289 | 0.004301 | 0.004163 | 5.9e-05 | 0.001011 | 0.26985 | 1.3218 | 2.548 | 99.34 | 21.02 | 37.349 | 70.5736 | 100.9853 | 68.5164 | 0.009373 | 1.0333 | 0.024741 | 23.9434 | -0.002868 | 0.14251 | 0.000155 | 0.011692 | 0.011692 | 0.50096
2019-03-21 12:50 | 1553186972.3884 | 865 | 0.20944 | 0.000461 | -3e-06 | 99.9174 | 99.8625 | 123.5642 | 21.2612 | 0.002047 | 0.001947 | -152.2026 | 0.24048 | 8.4027 | 0.00332 | 0.004555 | 0.004354 | 6.1e-05 | 0.001015 | 0.27721 | 1.3401 | 2.5482 | 99.34 | 21.02 | 37.348 | 71.0717 | 101.728 | 68.6801 | 0.009372 | 1.0328 | 0.024711 | 23.9256 | 0.004914 | 0.14306 | 0.000172 | 0.012942 | 0.012942 | 0.45432
2019-03-21 12:50 | 1553186973.3884 | 866 | 0.20944 | 0.00046 | -3e-06 | 99.9704 | 99.9909 | 124.044 | 21.2612 | 0.002078 | 0.002015 | -152.2088 | 0.23933 | 8.4018 | 0.003327 | 0.00469 | 0.004489 | 6.2e-05 | 0.001012 | 0.24275 | 1.3427 | 2.5479 | 99.34 | 21.02 | 37.348 | 71.8311 | 101.8906 | 70.2967 | 0.009372 | 1.0279 | 0.0247 | 24.0286 | -0.001833 | 0.14238 | 0.000218 | 0.016342 | 0.016342 | 0.35807
2019-03-21 12:50 | 1553186974.3869 | 867 | 0.20944 | 0.000459 | -3e-06 | 99.9853 | 99.9763 | 124.5203 | 21.2612 | 0.002032 | 0.002028 | -152.2155 | 0.23936 | 8.4033 | 0.003317 | 0.004988 | 0.004815 | 6.2e-05 | 0.001006 | 0.24945 | 1.3152 | 2.5481 | 99.34 | 21.02 | 37.348 | 70.7077 | 101.6629 | 70.6139 | 0.009372 | 1.027 | 0.024736 | 24.0858 | 0.00081 | 0.1424 | 0.000211 | 0.015895 | 0.015895 | 0.3682
2019-03-21 12:50 | 1553186975.3865 | 868 | 0.20946 | 0.000464 | -3e-06 | 99.9415 | 99.8138 | 124.643 | 21.2612 | 0.002064 | 0.001968 | -152.2196 | 0.23947 | 8.4024 | 0.003321 | 0.004538 | 0.004468 | 6e-05 | 0.001003 | 0.25066 | 1.3738 | 2.5479 | 99.34 | 21.02 | 37.347 | 71.4841 | 101.7515 | 69.169 | 0.009372 | 1.0313 | 0.0248 | 24.047 | 0.011438 | 0.14246 | 7.4e-05 | 0.005667 | 0.005667 | 1.0332
2019-03-21 12:50 | 1553186976.387 | 869 | 0.20945 | 0.000459 | -4e-06 | 99.9621 | 99.979 | 124.7717 | 21.2612 | 0.002066 | 0.001973 | -152.2344 | 0.24015 | 8.4037 | 0.003307 | 0.003886 | 0.003735 | 5.9e-05 | 0.001004 | 0.27345 | 1.3667 | 2.548 | 99.34 | 21.02 | 37.347 | 71.5324 | 101.417 | 69.2899 | 0.009372 | 1.031 | 0.024809 | 24.0636 | -0.001513 | 0.14286 | 0.000137 | 0.010402 | 0.010402 | 0.56449
2019-03-21 12:50 | 1553186977.3867 | 870 | 0.20944 | 0.000461 | -3e-06 | 99.9812 | 100.1163 | 124.5339 | 21.2612 | 0.002037 | 0.001947 | -152.2237 | 0.23953 | 8.4017 | 0.003301 | 0.00428 | 0.004061 | 6e-05 | 0.001004 | 0.23811 | 1.3329 | 2.5479 | 99.34 | 21.03 | 37.346 | 70.8307 | 101.2631 | 68.6753 | 0.009378 | 1.0328 | 0.024808 | 24.0191 | -0.012078 | 0.1425 | 0.000198 | 0.014978 | 0.014978 | 0.391
2019-03-21 12:50 | 1553186978.3877 | 871 | 0.20945 | 0.000457 | -4e-06 | 99.9386 | 99.9146 | 124.6003 | 21.2612 | 0.002021 | 0.001969 | -152.2045 | 0.23961 | 8.4041 | 0.003441 | 0.00459 | 0.004438 | 5.8e-05 | 0.001009 | 0.25105 | 1.331 | 2.5481 | 99.34 | 21.03 | 37.345 | 70.4541 | 104.6661 | 69.1988 | 0.009377 | 1.0312 | 0.024795 | 24.0439 | 0.002141 | 0.14254 | 0.000122 | 0.009305 | 0.009305 | 0.62962
2019-03-21 12:50 | 1553186979.3881 | 872 | 0.20944 | 0.000462 | -3e-06 | 99.898 | 99.6851 | 124.7564 | 21.2612 | 0.002044 | 0.001973 | -152.2164 | 0.2392 | 8.4041 | 0.00341 | 0.004533 | 0.004379 | 6.3e-05 | 0.001006 | 0.27975 | 1.3329 | 2.548 | 99.34 | 21.03 | 37.344 | 71.0037 | 103.9193 | 69.3024 | 0.009377 | 1.0309 | 0.024807 | 24.0626 | 0.019088 | 0.1423 | 0.000167 | 0.012619 | 0.012619 | 0.46346
2019-03-21 12:50 | 1553186980.3871 | 873 | 0.20944 | 0.000458 | -3e-06 | 99.9475 | 99.9174 | 124.8634 | 21.2612 | 0.00202 | 0.002002 | -152.2247 | 0.23987 | 8.4002 | 0.003386 | 0.004229 | 0.004186 | 5.9e-05 | 0.001 | 0.27848 | 1.322 | 2.548 | 99.34 | 21.03 | 37.341 | 70.4371 | 103.3282 | 69.9879 | 0.009376 | 1.0289 | 0.024793 | 24.097 | 0.002697 | 0.1427 | 0.000206 | 0.015584 | 0.015584 | 0.37634
2019-03-21 12:50 | 1553186981.3867 | 874 | 0.20946 | 0.000463 | -4e-06 | 99.9165 | 99.9926 | 124.9691 | 21.2612 | 0.002061 | 0.001988 | -152.2427 | 0.24055 | 8.4039 | 0.00336 | 0.005083 | 0.004915 | 6.2e-05 | 0.001009 | 0.27936 | 1.299 | 2.548 | 99.34 | 21.03 | 37.34 | 71.4187 | 102.7026 | 69.6478 | 0.009376 | 1.0299 | 0.024816 | 24.0953 | -0.006812 | 0.1431 | 9.3e-05 | 0.007137 | 0.007137 | 0.82412
2019-03-21 12:50 | 1553186982.3884 | 875 | 0.20944 | 0.000463 | -3e-06 | 99.9558 | 99.981 | 125.3209 | 21.2612 | 0.002072 | 0.001954 | -152.2182 | 0.24039 | 8.4039 | 0.003319 | 0.004768 | 0.004603 | 5.9e-05 | 0.001008 | 0.25122 | 1.3146 | 2.5478 | 99.34 | 21.03 | 37.339 | 71.6821 | 101.6951 | 68.8462 | 0.009376 | 1.0323 | 0.02488 | 24.1009 | -0.002256 | 0.14301 | 0.000195 | 0.014761 | 0.014761 | 0.39818
2019-03-21 12:50 | 1553186983.3873 | 876 | 0.20945 | 0.000462 | -3e-06 | 99.9928 | 99.963 | 125.1733 | 21.2612 | 0.002011 | 0.001957 | -152.2325 | 0.23907 | 8.4059 | 0.003239 | 0.004772 | 0.00459 | 6.3e-05 | 0.001007 | 0.22898 | 1.3119 | 2.548 | 99.34 | 21.03 | 37.339 | 70.2241 | 99.7721 | 68.9063 | 0.009376 | 1.0321 | 0.024863 | 24.0888 | 0.002665 | 0.14222 | 0.000123 | 0.009377 | 0.009377 | 0.62337
2019-03-21 12:50 | 1553186984.388 | 877 | 0.20962 | 0.000465 | -3e-06 | 100.0413 | 100.0437 | 124.7266 | 21.2612 | 0.002057 | 0.001963 | -152.2346 | 0.23888 | 8.4013 | 0.003269 | 0.005027 | 0.004856 | 6.1e-05 | 0.001009 | 0.26233 | 1.3474 | 2.548 | 99.34 | 21.03 | 37.337 | 71.3056 | 100.5035 | 69.0669 | 0.009375 | 1.0316 | 0.024813 | 24.0515 | -0.000212 | 0.14211 | -0.000892 | -0.066241 | -0.066241 | -0.088173
2019-03-21 12:50 | 1553186985.3881 | 878 | 0.20953 | 0.000464 | -3e-06 | 99.9514 | 100.0226 | 124.7188 | 21.2612 | 0.002076 | 0.001944 | -152.2416 | 0.23972 | 8.4055 | 0.003263 | 0.004474 | 0.004305 | 6.1e-05 | 0.001016 | 0.26394 | 1.3269 | 2.5479 | 99.34 | 21.03 | 37.336 | 71.7703 | 100.3455 | 68.5936 | 0.009375 | 1.0331 | 0.024829 | 24.0341 | -0.006372 | 0.14261 | -0.000383 | -0.028324 | -0.028324 | -0.20694
2019-03-21 12:50 | 1553186986.3871 | 879 | 0.20936 | 0.000468 | -4e-06 | 99.8784 | 99.7801 | 124.4257 | 21.2612 | 0.002001 | 0.001957 | -152.2258 | 0.23984 | 8.4054 | 0.003223 | 0.004633 | 0.00458 | 6.4e-05 | 0.001014 | 0.23866 | 1.3401 | 2.548 | 99.34 | 21.03 | 37.335 | 69.9792 | 99.386 | 68.9039 | 0.009375 | 1.0321 | 0.024788 | 24.0167 | 0.008805 | 0.14268 | 0.000695 | 0.052008 | 0.052008 | 0.11276
2019-03-21 12:50 | 1553186987.3872 | 880 | 0.20942 | 0.000463 | -3e-06 | 99.9216 | 100.0689 | 124.2757 | 21.2612 | 0.002028 | 0.001937 | -152.2541 | 0.23939 | 8.4073 | 0.003274 | 0.005169 | 0.00508 | 6.3e-05 | 0.001017 | 0.27743 | 1.3224 | 2.548 | 99.34 | 21.04 | 37.333 | 70.6188 | 100.6053 | 68.4359 | 0.00938 | 1.0335 | 0.02479 | 23.9858 | -0.01318 | 0.14241 | 0.000291 | 0.021875 | 0.021875 | 0.26757
2019-03-21 12:50 | 1553186988.3872 | 881 | 0.2094 | 0.000462 | -3e-06 | 99.9076 | 99.8979 | 123.9478 | 21.2612 | 0.001981 | 0.001947 | -152.2746 | 0.23851 | 8.4049 | 0.00323 | 0.004518 | 0.004282 | 6.4e-05 | 0.00102 | 0.25393 | 1.3066 | 2.548 | 99.34 | 21.04 | 37.331 | 69.4935 | 99.5596 | 68.6839 | 0.00938 | 1.0328 | 0.024749 | 23.9628 | 0.000862 | 0.14189 | 0.000415 | 0.031025 | 0.031025 | 0.18797
2019-03-21 12:50 | 1553186989.3867 | 882 | 0.20948 | 0.00046 | -4e-06 | 99.8995 | 99.8969 | 124.0419 | 21.2612 | 0.001994 | 0.001943 | -152.254 | 0.23835 | 8.403 | 0.003282 | 0.004045 | 0.003895 | 6.4e-05 | 0.001024 | 0.22827 | 1.3248 | 2.5481 | 99.34 | 21.04 | 37.329 | 69.8086 | 100.8039 | 68.5671 | 0.009379 | 1.0332 | 0.024762 | 23.9678 | 0.000233 | 0.14179 | -6e-05 | -0.004246 | -0.004246 | -1.3725
2019-03-21 12:50 | 1553186990.3874 | 883 | 0.2094 | 0.000465 | -3e-06 | 99.9222 | 99.9446 | 124.0895 | 21.2612 | 0.00201 | 0.001943 | -152.2667 | 0.23863 | 8.4042 | 0.003129 | 0.004655 | 0.004494 | 6.3e-05 | 0.001025 | 0.24978 | 1.3517 | 2.5481 | 99.34 | 21.04 | 37.328 | 70.1852 | 97.1055 | 68.5686 | 0.009379 | 1.0331 | 0.024767 | 23.9725 | -0.002008 | 0.14196 | 0.000405 | 0.030317 | 0.030317 | 0.19245
2019-03-21 12:50 | 1553186991.3867 | 884 | 0.20943 | 0.00046 | -4e-06 | 99.9552 | 99.8779 | 123.9424 | 21.2612 | 0.00205 | 0.001931 | -152.2607 | 0.23932 | 8.4048 | 0.003226 | 0.004737 | 0.004559 | 6.2e-05 | 0.00103 | 0.25498 | 1.3832 | 2.5481 | 99.34 | 21.04 | 37.327 | 71.1566 | 99.4443 | 68.2909 | 0.009379 | 1.034 | 0.024762 | 23.9485 | 0.006928 | 0.14237 | 0.000249 | 0.018739 | 0.018739 | 0.31227
2019-03-21 12:50 | 1553186992.3864 | 885 | 0.20945 | 0.000466 | -3e-06 | 99.9445 | 99.7552 | 123.6159 | 21.2612 | 0.002001 | 0.00186 | -152.2822 | 0.2388 | 8.4041 | 0.00325 | 0.004492 | 0.004441 | 6.4e-05 | 0.001031 | 0.26073 | 1.3777 | 2.5479 | 99.34 | 21.04 | 37.327 | 69.974 | 100.0287 | 66.5919 | 0.009379 | 1.0392 | 0.024792 | 23.8574 | 0.016955 | 0.14206 | 0.000126 | 0.009572 | 0.009572 | 0.60999
2019-03-21 12:50 | 1553186993.3865 | 886 | 0.20944 | 0.000452 | -3e-06 | 99.9079 | 100.0921 | 123.4427 | 21.2612 | 0.001961 | 0.001879 | -152.2576 | 0.23931 | 8.4041 | 0.003179 | 0.004301 | 0.004158 | 6.2e-05 | 0.001035 | 0.29839 | 1.3307 | 2.548 | 99.34 | 21.04 | 37.327 | 68.9984 | 98.3127 | 67.0397 | 0.009379 | 1.0378 | 0.024758 | 23.8564 | -0.016477 | 0.14236 | 0.00022 | 0.016575 | 0.016575 | 0.35302
2019-03-21 12:50 | 1553186994.3874 | 887 | 0.20945 | 0.000458 | -3e-06 | 99.91 | 99.9306 | 123.5307 | 21.2612 | 0.001963 | 0.001922 | -152.2534 | 0.23878 | 8.4016 | 0.003152 | 0.004039 | 0.003826 | 6.2e-05 | 0.00104 | 0.24834 | 1.3379 | 2.548 | 99.34 | 21.04 | 37.326 | 69.0536 | 97.677 | 68.081 | 0.009378 | 1.0346 | 0.024729 | 23.9014 | -0.001852 | 0.14205 | 0.00014 | 0.010555 | 0.010555 | 0.5531
2019-03-21 12:50 | 1553186995.3876 | 888 | 0.20944 | 0.000457 | -3e-06 | 99.8805 | 99.9371 | 123.6196 | 21.2612 | 0.002037 | 0.00189 | -152.2518 | 0.23839 | 8.4051 | 0.003135 | 0.004643 | 0.004478 | 6.2e-05 | 0.001039 | 0.25929 | 1.363 | 2.548 | 99.34 | 21.03 | 37.324 | 70.8346 | 97.2678 | 67.3049 | 0.009372 | 1.037 | 0.024766 | 23.8828 | -0.005072 | 0.14182 | 0.000214 | 0.016064 | 0.016064 | 0.36283
2019-03-21 12:50 | 1553186996.3867 | 889 | 0.20944 | 0.000458 | -3e-06 | 99.9207 | 99.9797 | 123.8959 | 21.2612 | 0.002045 | 0.001976 | -152.2729 | 0.2396 | 8.3994 | 0.003088 | 0.004761 | 0.004674 | 6.5e-05 | 0.001042 | 0.23883 | 1.3466 | 2.548 | 99.34 | 21.03 | 37.323 | 71.0328 | 96.1334 | 69.3743 | 0.009372 | 1.0307 | 0.024719 | 23.982 | -0.005281 | 0.14253 | 0.000219 | 0.016442 | 0.016442 | 0.3563
2019-03-21 12:50 | 1553186997.3881 | 890 | 0.20944 | 0.000455 | -3e-06 | 99.9219 | 99.9105 | 123.9198 | 21.2612 | 0.001956 | 0.001929 | -152.2694 | 0.23905 | 8.4017 | 0.003105 | 0.004436 | 0.004186 | 6.4e-05 | 0.001039 | 0.29264 | 1.3642 | 2.5481 | 99.34 | 21.03 | 37.324 | 68.8993 | 96.5431 | 68.245 | 0.009372 | 1.0341 | 0.024762 | 23.9447 | 0.001014 | 0.14221 | 0.00023 | 0.017306 | 0.017306 | 0.33774
2019-03-21 12:50 | 1553186998.3867 | 891 | 0.20941 | 0.00046 | -4e-06 | 100.0056 | 99.8935 | 124.2844 | 21.2612 | 0.001996 | 0.001886 | -152.2906 | 0.23887 | 8.4028 | 0.003132 | 0.004394 | 0.004305 | 6.2e-05 | 0.001041 | 0.24818 | 1.3348 | 2.548 | 99.34 | 21.03 | 37.325 | 69.8538 | 97.1918 | 67.2176 | 0.009372 | 1.0372 | 0.024836 | 23.9438 | 0.01003 | 0.1421 | 0.000396 | 0.029745 | 0.029745 | 0.19635
2019-03-21 12:50 | 1553186999.3869 | 892 | 0.20943 | 0.00046 | -3e-06 | 99.898 | 99.9269 | 124.1165 | 21.2612 | 0.001965 | 0.00191 | -152.3148 | 0.23937 | 8.4023 | 0.003106 | 0.00494 | 0.004761 | 6.1e-05 | 0.001037 | 0.23075 | 1.3645 | 2.5481 | 99.34 | 21.03 | 37.325 | 69.112 | 96.5477 | 67.7887 | 0.009372 | 1.0355 | 0.024798 | 23.9477 | -0.002584 | 0.1424 | 0.000258 | 0.019394 | 0.019394 | 0.30178
2019-03-21 12:50 | 1553187000.3869 | 893 | 0.20944 | 0.000454 | -3e-06 | 99.9451 | 99.9909 | 124.0555 | 21.2612 | 0.001952 | 0.001874 | -152.3152 | 0.23975 | 8.4041 | 0.003097 | 0.004906 | 0.004637 | 6.7e-05 | 0.001044 | 0.262 | 1.3046 | 2.548 | 99.34 | 21.03 | 37.327 | 68.7955 | 96.3285 | 66.918 | 0.009373 | 1.0382 | 0.024824 | 23.9112 | -0.004099 | 0.14263 | 0.000202 | 0.015231 | 0.015231 | 0.38487
2019-03-21 12:50 | 1553187001.387 | 894 | 0.20947 | 0.000457 | -4e-06 | 99.9603 | 100.2283 | 124.0928 | 21.2612 | 0.001999 | 0.00194 | -152.3084 | 0.23967 | 8.4024 | 0.003048 | 0.004368 | 0.004204 | 6.3e-05 | 0.001045 | 0.28152 | 1.3369 | 2.548 | 99.34 | 21.03 | 37.327 | 69.9333 | 95.1625 | 68.4938 | 0.009373 | 1.0334 | 0.02477 | 23.9702 | -0.02395 | 0.14258 | 2.4e-05 | 0.002034 | 0.002034 | 2.8817
2019-03-21 12:50 | 1553187002.387 | 895 | 0.20945 | 0.000459 | -4e-06 | 99.9946 | 100.0045 | 124.1292 | 21.2612 | 0.00197 | 0.0019 | -152.3097 | 0.23939 | 8.4045 | 0.003153 | 0.004335 | 0.004109 | 6.2e-05 | 0.001043 | 0.25393 | 1.3279 | 2.5479 | 99.34 | 21.04 | 37.326 | 69.2354 | 97.6973 | 67.5451 | 0.009378 | 1.0362 | 0.024808 | 23.9404 | -0.000892 | 0.14242 | 0.000123 | 0.009381 | 0.009381 | 0.62397
2019-03-21 12:50 | 1553187003.3881 | 896 | 0.20944 | 0.000458 | -4e-06 | 100.0554 | 100.0727 | 124.1644 | 21.2612 | 0.001973 | 0.001876 | -152.3187 | 0.23863 | 8.4016 | 0.003089 | 0.004667 | 0.004557 | 6.2e-05 | 0.001046 | 0.26178 | 1.3517 | 2.548 | 99.34 | 21.04 | 37.326 | 69.3048 | 96.1429 | 66.9733 | 0.009378 | 1.038 | 0.024833 | 23.9237 | -0.001548 | 0.14196 | 0.00022 | 0.0166 | 0.0166 | 0.35147
2019-03-21 12:50 | 1553187004.3876 | 897 | 0.20945 | 0.000458 | -3e-06 | 100.0562 | 100.1725 | 123.9083 | 21.2612 | 0.001972 | 0.001842 | -152.3378 | 0.24013 | 8.3976 | 0.003043 | 0.00462 | 0.004384 | 6.3e-05 | 0.001047 | 0.26288 | 1.3246 | 2.5481 | 99.34 | 21.04 | 37.326 | 69.2729 | 95.0442 | 66.1525 | 0.009378 | 1.0405 | 0.024837 | 23.8701 | -0.010383 | 0.14285 | 0.000152 | 0.011512 | 0.011512 | 0.51001
2019-03-21 12:50 | 1553187005.3862 | 898 | 0.20945 | 0.000454 | -4e-06 | 100.0649 | 100.0533 | 123.8311 | 21.2612 | 0.001988 | 0.001885 | -152.3237 | 0.24012 | 8.4015 | 0.003173 | 0.004064 | 0.00391 | 6.2e-05 | 0.001043 | 0.30342 | 1.3523 | 2.5481 | 99.34 | 21.04 | 37.328 | 69.6557 | 98.1629 | 67.1714 | 0.009379 | 1.0374 | 0.024792 | 23.8985 | 0.00104 | 0.14285 | 0.000158 | 0.011949 | 0.011949 | 0.49133
2019-03-21 12:50 | 1553187006.3862 | 899 | 0.20945 | 0.000461 | -3e-06 | 100.0238 | 100.2416 | 123.5793 | 21.2612 | 0.001957 | 0.001884 | -152.3218 | 0.23979 | 8.4017 | 0.003148 | 0.004424 | 0.004297 | 6.3e-05 | 0.001045 | 0.24834 | 1.3141 | 2.5481 | 99.34 | 21.04 | 37.329 | 68.9094 | 97.5804 | 67.1481 | 0.009379 | 1.0375 | 0.024768 | 23.8734 | -0.019458 | 0.14265 | 0.000128 | 0.009672 | 0.009672 | 0.60619
2019-03-21 12:50 | 1553187007.3862 | 900 | 0.20946 | 0.000454 | -3e-06 | 100.0556 | 99.9773 | 123.8123 | 21.2612 | 0.002014 | 0.001895 | -152.3238 | 0.23909 | 8.4035 | 0.003161 | 0.004663 | 0.004508 | 6.4e-05 | 0.00104 | 0.24137 | 1.3173 | 2.548 | 99.34 | 21.04 | 37.329 | 70.2866 | 97.8947 | 67.4283 | 0.009379 | 1.0366 | 0.024781 | 23.9057 | 0.007007 | 0.14223 | 5.7e-05 | 0.004423 | 0.004423 | 1.3216
2019-03-21 12:50 | 1553187008.3873 | 901 | 0.20944 | 0.000456 | -3e-06 | 100.0235 | 100.1885 | 123.952 | 21.2612 | 0.001943 | 0.001858 | -152.2988 | 0.24061 | 8.401 | 0.003141 | 0.00421 | 0.00403 | 6.3e-05 | 0.001045 | 0.25194 | 1.3082 | 2.5478 | 99.34 | 21.04 | 37.329 | 68.5866 | 97.3956 | 66.545 | 0.009379 | 1.0393 | 0.024827 | 23.8882 | -0.014743 | 0.14314 | 0.000207 | 0.015606 | 0.015606 | 0.37697
2019-03-21 12:50 | 1553187009.3874 | 902 | 0.20945 | 0.000459 | -3e-06 | 99.997 | 100.1575 | 124.1795 | 21.2612 | 0.001999 | 0.001892 | -152.2862 | 0.23981 | 8.4021 | 0.003117 | 0.004837 | 0.004567 | 6.4e-05 | 0.001046 | 0.25769 | 1.3067 | 2.5479 | 99.34 | 21.04 | 37.328 | 69.9348 | 96.8213 | 67.3612 | 0.009379 | 1.0368 | 0.02482 | 23.9388 | -0.014347 | 0.14266 | 0.00014 | 0.010637 | 0.010637 | 0.55122
2019-03-21 12:50 | 1553187010.3864 | 903 | 0.20944 | 0.000459 | -3e-06 | 100.0771 | 99.8557 | 124.3269 | 21.2612 | 0.001967 | 0.001895 | -152.2934 | 0.23942 | 8.4008 | 0.003135 | 0.004421 | 0.004327 | 6.3e-05 | 0.001042 | 0.2995 | 1.3247 | 2.5479 | 99.34 | 21.04 | 37.328 | 69.1464 | 97.2459 | 67.4352 | 0.009379 | 1.0366 | 0.024832 | 23.9556 | 0.019809 | 0.14243 | 0.000205 | 0.015458 | 0.015458 | 0.3787
2019-03-21 12:50 | 1553187011.3875 | 904 | 0.20945 | 0.000459 | -3e-06 | 100.114 | 100.0573 | 124.7045 | 21.2612 | 0.00198 | 0.001917 | -152.2948 | 0.23953 | 8.4013 | 0.003138 | 0.004578 | 0.004403 | 6.4e-05 | 0.001042 | 0.2756 | 1.3558 | 2.5481 | 99.34 | 21.04 | 37.328 | 69.456 | 97.3367 | 67.9531 | 0.009379 | 1.035 | 0.024851 | 24.0102 | 0.005068 | 0.1425 | 0.000163 | 0.012368 | 0.012368 | 0.47353
2019-03-21 12:50 | 1553187012.3865 | 905 | 0.20947 | 0.000461 | -3e-06 | 100.1012 | 100.1364 | 125.0963 | 21.2612 | 0.001949 | 0.001875 | -152.2941 | 0.23987 | 8.4012 | 0.003076 | 0.004039 | 0.00385 | 6.4e-05 | 0.001042 | 0.27743 | 1.3306 | 2.5482 | 99.34 | 21.04 | 37.328 | 68.7176 | 95.8327 | 66.9413 | 0.009379 | 1.0381 | 0.024927 | 24.0122 | -0.003138 | 0.1427 | 4.6e-05 | 0.003656 | 0.003656 | 1.604
2019-03-21 12:50 | 1553187013.3865 | 906 | 0.20946 | 0.00046 | -3e-06 | 100.1382 | 100.1691 | 125.1702 | 21.2612 | 0.001999 | 0.001944 | -152.2929 | 0.23913 | 8.402 | 0.003124 | 0.004012 | 0.003835 | 6.3e-05 | 0.001037 | 0.24923 | 1.3635 | 2.5481 | 99.34 | 21.04 | 37.329 | 69.9341 | 96.9882 | 68.5919 | 0.009379 | 1.0331 | 0.024874 | 24.0775 | -0.002757 | 0.14226 | 5.5e-05 | 0.004349 | 0.004349 | 1.3444
2019-03-21 12:50 | 1553187014.3865 | 907 | 0.20945 | 0.000457 | -4e-06 | 100.0732 | 100.0825 | 125.6378 | 21.2612 | 0.00195 | 0.001843 | -152.2793 | 0.23989 | 8.4027 | 0.00309 | 0.004371 | 0.004183 | 6.3e-05 | 0.001038 | 0.26089 | 1.3601 | 2.548 | 99.34 | 21.04 | 37.33 | 68.7395 | 96.1617 | 66.1657 | 0.009379 | 1.0405 | 0.025009 | 24.0366 | -0.000833 | 0.14271 | 0.000131 | 0.01004 | 0.01004 | 0.58423
2019-03-21 12:50 | 1553187015.3874 | 908 | 0.20944 | 0.000458 | -4e-06 | 100.1453 | 100.1408 | 125.9535 | 21.2612 | 0.00205 | 0.001845 | -152.2948 | 0.24013 | 8.4015 | 0.003105 | 0.004255 | 0.00405 | 6.1e-05 | 0.001032 | 0.23595 | 1.3989 | 2.548 | 99.34 | 21.04 | 37.331 | 71.1395 | 96.5414 | 66.2227 | 0.00938 | 1.0403 | 0.025039 | 24.0688 | 0.000406 | 0.14286 | 0.000201 | 0.015329 | 0.015329 | 0.38302
2019-03-21 12:50 | 1553187016.387 | 909 | 0.20942 | 0.000457 | -4e-06 | 100.1328 | 100.245 | 126.6173 | 21.2612 | 0.001992 | 0.001928 | -152.2906 | 0.23965 | 8.4033 | 0.003173 | 0.004318 | 0.004233 | 6.3e-05 | 0.001032 | 0.23131 | 1.374 | 2.548 | 99.34 | 21.04 | 37.33 | 69.7447 | 98.181 | 68.2137 | 0.009379 | 1.0342 | 0.025031 | 24.2028 | -0.010016 | 0.14257 | 0.000301 | 0.022815 | 0.022815 | 0.25683
2019-03-21 12:50 | 1553187017.3877 | 910 | 0.20944 | 0.000462 | -4e-06 | 100.1623 | 100.2385 | 126.5534 | 21.2612 | 0.001972 | 0.001833 | -152.2859 | 0.23898 | 8.4034 | 0.003089 | 0.004646 | 0.004418 | 6.1e-05 | 0.001032 | 0.29856 | 1.3783 | 2.5479 | 99.34 | 21.04 | 37.33 | 69.2798 | 96.1412 | 65.9245 | 0.009379 | 1.0412 | 0.025109 | 24.1155 | -0.006804 | 0.14217 | 0.000189 | 0.014487 | 0.014487 | 0.40334
2019-03-21 12:50 | 1553187018.3872 | 911 | 0.20945 | 0.000457 | -3e-06 | 100.1331 | 100.1479 | 126.5382 | 21.2612 | 0.00197 | 0.001869 | -152.29 | 0.23915 | 8.403 | 0.003135 | 0.004461 | 0.004319 | 5.9e-05 | 0.001031 | 0.2453 | 1.3958 | 2.548 | 99.34 | 21.04 | 37.329 | 69.2166 | 97.2475 | 66.8079 | 0.009379 | 1.0385 | 0.025075 | 24.1454 | -0.001324 | 0.14227 | 0.000139 | 0.010706 | 0.010706 | 0.54619
2019-03-21 12:50 | 1553187019.3878 | 912 | 0.20945 | 0.000458 | -4e-06 | 100.1379 | 100.2688 | 126.5203 | 21.2612 | 0.001909 | 0.00184 | -152.2861 | 0.2391 | 8.4026 | 0.003096 | 0.004256 | 0.004009 | 6.2e-05 | 0.001031 | 0.24187 | 1.3523 | 2.5479 | 99.34 | 21.04 | 37.33 | 67.7526 | 96.3159 | 66.1042 | 0.009379 | 1.0407 | 0.025099 | 24.1187 | -0.01169 | 0.14224 | 0.000149 | 0.011464 | 0.011464 | 0.50997
2019-03-21 12:50 | 1553187020.3868 | 913 | 0.20945 | 0.000454 | -3e-06 | 100.2055 | 100.2849 | 126.5807 | 21.2612 | 0.001989 | 0.001858 | -152.2689 | 0.23975 | 8.4028 | 0.003013 | 0.004202 | 0.004068 | 6e-05 | 0.001034 | 0.28102 | 1.3609 | 2.5481 | 99.34 | 21.04 | 37.331 | 69.6791 | 94.3185 | 66.5339 | 0.00938 | 1.0393 | 0.025089 | 24.1397 | -0.007079 | 0.14263 | 0.000156 | 0.011959 | 0.011959 | 0.49017
2019-03-21 12:50 | 1553187021.3869 | 914 |  | 0.000459 |  | 100.1724 | 100.2532 | 126.3746 | 21.2612 | 0.001964 | 0.001857 | -152.2844 | 0.24003 | 8.4033 | 0.003061 | 0.003834 | 0.003645 | 6.2e-05 | 0.001028 | 0.27256 | 1.3302 | 2.5481 | 99.34 | 21.04 | 37.331 | 69.0842 | 95.4639 | 66.5099 | 0.00938 | 1.0394 | 0.02507 | 24.1192 | -0.007206 | 0.1428
2019-03-21 12:50 | 1553187022.3869 | 915 |  | 0.000457 |  | 100.1763 | 100.3884 | 126.0671 | 21.2612 | 0.002005 | 0.001837 | -152.2833 | 0.24064 | 8.4048 | 0.003102 | 0.004312 | 0.004194 | 6.4e-05 | 0.001027 | 0.28385 | 1.3208 | 2.548 | 99.34 | 21.04 | 37.33 | 70.0689 | 96.4688 | 66.033 | 0.009379 | 1.0409 | 0.025057 | 24.0729 | -0.018916 | 0.14316
2019-03-21 12:50 | 1553187023.3869 | 916 |  |  |  | 100.291 | 100.3574 | 126.123 | 21.2612 | 0.001958 | 0.001922 | -152.2856 | 0.23982 | 8.4033 | 0.002986 | 0.004416 | 0.004188 | 6.2e-05 | 0.001027 | 0.25713 | 1.2999 | 2.5479 | 99.34 | 21.04 | 37.328 | 68.932 | 93.6574 | 68.0779 | 0.009379 | 1.0346 | 0.024987 | 24.1507 | -0.005918 | 0.14267
2019-03-21 12:50 | 1553187024.3869 | 917 |  |  |  | 100.2153 | 100.2661 | 126.2377 | 21.2612 | 0.001934 | 0.001899 | -152.2915 | 0.23861 | 8.4028 | 0.002962 | 0.005151 | 0.005105 | 6.2e-05 | 0.001025 | 0.2417 | 1.3073 | 2.5479 | 99.34 | 21.04 | 37.328 | 68.3604 | 93.0854 | 67.5179 | 0.009379 | 1.0363 | 0.025019 | 24.1419 | -0.004531 | 0.14195
2019-03-21 12:50 | 1553187025.3869 | 918 |  |  |  | 100.277 | 100.3942 | 126.1483 | 21.2612 | 0.00203 | 0.00184 | -152.2781 | 0.23939 | 8.4012 | 0.002986 | 0.004241 | 0.004053 | 5.8e-05 | 0.00102 | 0.26128 | 1.2921 | 2.548 | 99.34 | 21.04 | 37.327 | 70.6639 | 93.6637 | 66.0955 | 0.009379 | 1.0407 | 0.025063 | 24.0829 | -0.010445 | 0.14242
2019-03-21 12:50 | 1553187026.3869 | 919 |  |  |  | 100.2782 | 100.3996 |  | 21.2612 | 0.00195 | 0.001894 | -152.2786 | 0.23931 | 8.4036 | 0.003051 | 0.003418 | 0.003268 | 6.3e-05 | 0.001019 | 0.28672 | 1.3237 | 2.548 | 99.34 | 21.04 | 37.328 | 68.7517 | 95.2187 | 67.3931 | 0.009379 | 1.0367 |  |  | -0.010824 | 0.14236
2019-03-21 12:50 | 1553187027.3869 | 920 |  |  |  | 100.3289 | 100.4276 |  | 21.2612 | 0.001967 |  | -152.2817 | 0.23966 | 8.4021 | 0.003065 | 0.004714 | 0.004482 | 6.2e-05 | 0.001016 | 0.26449 | 1.3198 | 2.548 | 99.34 | 21.04 | 37.327 | 69.1464 | 95.5759 |  | 0.009379 |  |  |  | -0.008795 | 0.14257
2019-03-21 12:50 | 1553187028.3869 | 921 |  |  |  | 100.297 | 100.4081 |  | 21.2612 | 0.001941 |  |  | 0.23976 | 8.4013 | 0.002973 | 0.004219 | 0.004029 | 6.6e-05 | 0.001016 | 0.29551 | 1.3472 | 2.5479 | 99.34 | 21.04 | 37.327 | 68.5234 | 93.3584 |  | 0.009379 |  |  |  | -0.009908 | 0.14264
2019-03-21 12:50 | 1553187029.3869 | 922 |  |  |  | 100.3664 | 100.5277 |  | 21.2612 |  |  |  | 0.24009 | 8.4035 | 0.00294 | 0.004809 | 0.004598 | 6.5e-05 | 0.001016 | 0.26145 | 1.314 | 2.548 | 99.34 | 21.04 | 37.327 |  | 92.5497 |  | 0.009379 |  |  |  | -0.014362 | 0.14283
2019-03-21 12:50 | 1553187030.3869 | 923 |  |  |  | 100.3453 | 100.244 |  | 21.2612 |  |  |  | 0.23992 | 8.4012 | 0.002982 | 0.00445 | 0.004272 | 6.2e-05 | 0.001013 | 0.25946 | 1.3048 | 2.5479 | 99.34 | 21.04 | 37.326 |  | 93.5612 |  | 0.009378 |  |  |  | 0.009031 | 0.14273
2019-03-21 12:50 | 1553187031.3869 | 924 |  |  |  | 100.3548 | 100.4323 |  | 21.2612 |  |  |  | 0.24126 | 8.4042 | 0.003053 | 0.004084 | 0.004021 | 6.5e-05 | 0.001009 | 0.26808 | 1.3197 | 2.5478 | 99.34 | 21.04 | 37.324 |  | 95.2776 |  | 0.009378 |  |  |  | -0.006909 | 0.14353
2019-03-21 12:51 | 1553187032.3869 | 925 |  |  |  | 100.333 | 100.5624 |  | 21.2612 |  |  |  | 0.23955 | 8.4057 | 0.002996 | 0.003907 | 0.003752 | 6.2e-05 | 0.001005 | 0.27151 | 1.3159 | 2.5479 | 99.33 | 21.04 | 37.325 |  | 93.8998 |  | 0.009379 |  |  |  | -0.020427 | 0.14251
2019-03-21 12:51 | 1553187033.3869

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-21 12:35 | 0 | 0 | 0 | 0 | 0 | 45.1 | 0.000285
2019-03-21 12:35 | 1 | 0 | 0 | 0 | 0 | 45.1723 | -0.01244
2019-03-21 12:35 | 2 | 0 | 0 | 0 | 0 | 45.1909 | 0.023052
2019-03-21 12:35 | 3 | 0 | 0 | 0 | 0 | 45.3497 | 0.01021
2019-03-21 12:35 | 4 | 0 | 0 | 0 | 0 | 45.1288 | -0.000952
2019-03-21 12:35 | 5 | 0 | 0 | 0 | 0 | 45.5361 | 0.063487
2019-03-21 12:35 | 6 | 0 | 0 | 0 | 0 | 45.082 | 0.23744
2019-03-21 12:35 | 7 | 0 | 0 | 0 | 0 | 45.0198 | 0.54799
2019-03-21 12:35 | 8 | 0 | 0 | 0 | 0 | 44.9325 | 1.0306
2019-03-21 12:35 | 9 | 0 | 0 | 0 | 0 | 44.8554 | 1.7049
2019-03-21 12:35 | 10 | 0 | 0 | 0 | 0 | 44.6836 | 2.3027
2019-03-21 12:35 | 11 | 0 | 0 | 0 | 0 | 44.539 | 2.6485
2019-03-21 12:35 | 12 | 0 | 0 | 0 | 0 | 44.3874 | 2.9491
2019-03-21 12:35 | 13 | 0 | 0 | 0 | 0 | 44.2713 | 3.0819
2019-03-21 12:35 | 14 | 0 | 0 | 0 | 0 | 44.1492 | 3.3059
2019-03-21 12:35 | 15 | 0.12338 | 0 | 0 | 0 | 44.0163 | 3.4537
2019-03-21 12:35 | 16 | 0.12327 | 0 | 0 | 0 | 43.9013 | 3.458
2019-03-21 12:35 | 17 | 0.12941 | 0 | 0 | 0 | 43.7686 | 3.4642
2019-03-21 12:35 | 18 | 0.11883 | 0 | 0 | 0 | 43.6467 | 3.4427
2019-03-21 12:35 | 19 | 0.10564 | 0 | 0 | 0 | 43.532 | 3.4081
2019-03-21 12:35 | 20 | 0.090699 | 0 | 0 | 0 | 43.4353 | 3.3126
2019-03-21 12:35 | 21 | 0.10013 | 0 | 0 | 0 | 43.345 | 3.2208
2019-03-21 12:35 | 22 | 0.092774 | 0 | 0 | 0 | 43.2383 | 3.0589
2019-03-21 12:35 | 23 | 0.09095 | 0 | 0 | 0 | 43.1582 | 2.9544
2019-03-21 12:35 | 24 | 0.099348 | 0 | 0 | 0 | 43.0543 | 2.8737
2019-03-21 12:36 | 25 | 0.087642 | 0 | 0 | 0 | 42.9639 | 2.7812
2019-03-21 12:36 | 26 | 0.088776 | 0 | 0 | 0 | 42.876 | 2.7373
2019-03-21 12:36 | 27 | 0.098449 | 0 | 0 | 0 | 42.7835 | 2.7316
2019-03-21 12:36 | 28 | 0.12142 | 0 | 0 | 0 | 42.677 | 2.715
2019-03-21 12:36 | 29 | 0.089193 | 0 | 0 | 0 | 42.5533 | 2.7049
2019-03-21 12:36 | 30 | 0.091189 | 0 | 0 | 0 | 42.4911 | 2.6935
2019-03-21 12:36 | 31 | 0.10829 | 0 | 0 | 0 | 42.3671 | 2.5973
2019-03-21 12:36 | 32 | 0.094567 | 0 | 0 | 0 | 42.2816 | 2.6509
2019-03-21 12:36 | 33 | 0.091067 | 0 | 0 | 0 | 42.1764 | 2.6785
2019-03-21 12:36 | 34 | 0.077053 | 0 | 0 | 0 | 42.1004 | 2.6193
2019-03-21 12:36 | 35 | 0.1156 | 0 | 0 | 0 | 42.0102 | 2.5547
2019-03-21 12:36 | 36 | 0.12158 | 0 | 0 | 0 | 41.8769 | 2.5013
2019-03-21 12:36 | 37 | 0.096144 | 0 | 0 | 0 | 41.7762 | 2.4892
2019-03-21 12:36 | 38 | 0.060781 | 0 | 0 | 0 | 41.6869 | 2.4341
2019-03-21 12:36 | 39 | 0.067448 | 0 | 0 | 0 | 41.6438 | 2.3513
2019-03-21 12:36 | 40 | 0.094247 | 0 | 0 | 0 | 41.5471 | 2.2292
2019-03-21 12:36 | 41 | 0.071367 | 0 | 0 | 0 | 41.4667 | 2.068
2019-03-21 12:36 | 42 | 0.064142 | 0 | 0 | 0 | 41.4009 | 1.9741
2019-03-21 12:36 | 43 | 0.08548 | 0 | 0 | 0 | 41.3308 | 1.9128
2019-03-21 12:36 | 44 | 0.10892 | 0 | 0 | 0 | 41.2302 | 1.8557
2019-03-21 12:36 | 45 | 0.076345 | 0 | 0 | 0 | 41.127 | 1.7908
2019-03-21 12:36 | 46 | 0.049189 | 0 | 0 | 0 | 41.0769 | 1.6994
2019-03-21 12:36 | 47 | 0.052139 | 0 | 0 | 0 | 41.0203 | 1.6043
2019-03-21 12:36 | 48 | 0.074088 | 0 | 0 | 0 | 40.967 | 1.5901
2019-03-21 12:36 | 49 | 0.093507 | 0 | 0 | 0 | 40.8737 | 1.5199
2019-03-21 12:36 | 50 | 0.055647 | 0 | 0 | 0 | 40.7933 | 1.4706
2019-03-21 12:36 | 51 | 0.059135 | 0 | 0 | 0 | 40.7531 | 1.4444
2019-03-21 12:36 | 52 | 0.082336 | 0.01545 | 2.0602 | 16.8033 | 40.6701 | 1.3835
2019-03-21 12:36 | 53 | 0.067379 | 0.01825 | 2.4679 | 19.8447 | 40.5975 | 1.3371
2019-03-21 12:36 | 54 | 0.050844 | 0.022866 | 3.1993 | 26.1472 | 40.5366 | 1.3294
2019-03-21 12:36 | 55 | 0.042313 | 0.026467 | 3.7707 | 31.4542 | 40.4934 | 1.3309
2019-03-21 12:36 | 56 | 0.050936 | 0.021526 | 3.0637 | 25.682 | 40.447 | 1.3081
2019-03-21 12:36 | 57 | 0.071288 | 0.014778 | 2.1474 | 18.178 | 40.3887 | 1.2959
2019-03-21 12:36 | 58 | 0.068502 | 0.015087 | 2.2008 | 19.081 | 40.3108 | 1.3071
2019-03-21 12:36 | 59 | 0.035636 | 0.027136 | 4.165 | 36.6154 | 40.2587 | 1.3048
2019-03-21 12:36 | 60 | 0.047423 | 0.019922 | 3.0807 | 27.4641 | 40.2291 | 1.3024
2019-03-21 12:36 | 61 | 0.061073 | 0.014692 | 2.3462 | 21.0584 | 40.1631 | 1.2861
2019-03-21 12:36 | 62 | 0.040304 | 0.021775 | 3.5166 | 31.2651 | 40.1151 | 1.2601
2019-03-21 12:36 | 63 | 0.031862 | 0.026823 | 4.3804 | 38.6346 | 40.0799 | 1.231
2019-03-21 12:36 | 64 | 0.039561 | 0.021046 | 3.4833 | 30.9813 | 40.0465 | 1.2257
2019-03-21 12:36 | 65 | 0.065604 | 0.012355 | 2.0736 | 18.6326 | 39.9964 | 1.2224
2019-03-21 12:36 | 66 | 0.063159 | 0.012825 | 2.1256 | 19.205 | 39.9221 | 1.213
2019-03-21 12:36 | 67 | 0.055938 | 0.014456 | 2.3638 | 21.6632 | 39.8718 | 1.2118
2019-03-21 12:36 | 68 | 0.059434 | 0.013597 | 2.196 | 20.346 | 39.8077 | 1.2092
2019-03-21 12:36 | 69 | 0.063677 | 0.013048 | 2.0283 | 18.9373 | 39.7527 | 1.2059
2019-03-21 12:36 | 70 | 0.060343 | 0.013767 | 2.119 | 20.0903 | 39.6823 | 1.2123
2019-03-21 12:36 | 71 | 0.052309 | 0.015852 | 2.4176 | 23.43 | 39.6326 | 1.2256
2019-03-21 12:36 | 72 | 0.073362 | 0.011612 | 1.71 | 16.7013 | 39.5715 | 1.2252
2019-03-21 12:36 | 73 | 0.058164 | 0.01504 | 2.1423 | 21.0504 | 39.4938 | 1.2244
2019-03-21 12:36 | 74 | 0.05992 | 0.014577 | 2.0598 | 20.8663 | 39.4519 | 1.2503
2019-03-21 12:36 | 75 | 0.060294 | 0.014858 | 2.0325 | 20.5066 | 39.3738 | 1.2364
2019-03-21 12:36 | 76 | 0.06475 | 0.013859 | 1.8871 | 19.1242 | 39.3305 | 1.2383
2019-03-21 12:36 | 77 | 0.062447 | 0.014714 | 1.9473 | 20.0136 | 39.2461 | 1.2498
2019-03-21 12:36 | 78 | 0.049636 | 0.018955 | 2.4404 | 25.3067 | 39.2072 | 1.2561
2019-03-21 12:36 | 79 | 0.083834 | 0.011474 | 1.4366 | 14.8961 | 39.1369 | 1.2488
2019-03-21 12:36 | 80 | 0.058526 | 0.016493 | 2.0608 | 21.4688 | 39.0525 | 1.2565
2019-03-21 12:36 | 81 | 0.05704 | 0.017252 | 2.1016 | 22.1601 | 39.015 | 1.264
2019-03-21 12:36 | 82 | 0.071634 | 0.013727 | 1.6712 | 18.0427 | 38.9341 | 1.2925
2019-03-21 12:36 | 83 | 0.078089 | 0.012587 | 1.5333 | 16.588 | 38.8739 | 1.2953
2019-03-21 12:36 | 84 | 0.062236 | 0.016168 | 1.9207 | 20.6541 | 38.7836 | 1.2854
2019-03-21 12:36 | 85 | 0.05187 | 0.019382 | 2.3046 | 25.2803 | 38.7479 | 1.3113
2019-03-21 12:37 | 86 | 0.073928 | 0.013892 | 1.6174 | 18.0928 | 38.6727 | 1.3376
2019-03-21 12:37 | 87 | 0.051433 | 0.019918 | 2.3189 | 26.2575 | 38.6096 | 1.3505
2019-03-21 12:37 | 88 | 0.055182 | 0.018599 | 2.1685 | 24.8192 | 38.5644 | 1.3696
2019-03-21 12:37 | 89 | 0.059099 | 0.017711 | 2.0305 | 23.0181 | 38.4982 | 1.3603
2019-03-21 12:37 | 90 | 0.071138 | 0.014691 | 1.6863 | 19.2919 | 38.4444 | 1.3724
2019-03-21 12:37 | 91 | 0.072617 | 0.014437 | 1.662 | 19.121 | 38.3584 | 1.3885
2019-03-21 12:37 | 92 | 0.074474 | 0.014373 | 1.6253 | 18.6583 | 38.2993 | 1.3896
2019-03-21 12:37 | 93 | 0.078676 | 0.013906 | 1.5449 | 17.8139 | 38.2093 | 1.4015
2019-03-21 12:37 | 94 | 0.065816 | 0.016607 | 1.8513 | 21.7669 | 38.1455 | 1.4326
2019-03-21 12:37 | 95 | 0.080232 | 0.013347 | 1.5241 | 17.7984 | 38.0723 | 1.428
2019-03-21 12:37 | 96 | 0.056981 | 0.019177 | 2.1568 | 25.2513 | 37.9927 | 1.4388
2019-03-21 12:37 | 97 | 0.074437 | 0.014406 | 1.6615 | 19.2687 | 37.9497 | 1.4343
2019-03-21 12:37 | 98 | 0.070997 | 0.015398 | 1.7451 | 20.1676 | 37.8487 | 1.4318
2019-03-21 12:37 | 99 | 0.04011 | 0.027225 | 3.0952 | 36.6323 | 37.8134 | 1.4693
2019-03-21 12:37 | 100 | 0.076405 | 0.014577 | 1.6301 | 19.4428 | 37.7539 | 1.4855
2019-03-21 12:37 | 101 | 0.059829 | 0.018614 | 2.0896 | 25.2685 | 37.6722 | 1.5118
2019-03-21 12:37 | 102 | 0.037681 | 0.030707 | 3.3304 | 40.1009 | 37.6361 | 1.5111
2019-03-21 12:37 | 103 | 0.045477 | 0.02543 | 2.7712 | 33.7994 | 37.5892 | 1.5371
2019-03-21 12:37 | 104 | 0.064384 | 0.017924 | 1.9635 | 23.5947 | 37.5422 | 1.5191
2019-03-21 12:37 | 105 | 0.053156 | 0.022146 | 2.3946 | 29.1996 | 37.467 | 1.5521
2019-03-21 12:37 | 106 | 0.064424 | 0.01825 | 1.982 | 23.5684 | 37.4312 | 1.5184
2019-03-21 12:37 | 107 | 0.073677 | 0.016261 | 1.7442 | 19.986 | 37.3386 | 1.4725
2019-03-21 12:37 | 108 | 0.060547 | 0.019788 | 2.1335 | 25.4625 | 37.2883 | 1.5417
2019-03-21 12:37 | 109 | 0.095086 | 0.012839 | 1.3642 | 16.027 | 37.2079 | 1.5239
2019-03-21 12:37 | 110 | 0.067422 | 0.018428 | 1.9314 | 23.0262 | 37.1115 | 1.5525
2019-03-21 12:37 | 111 | 0.074977 | 0.016801 | 1.7383 | 20.7427 | 37.0657 | 1.5552
2019-03-21 12:37 | 112 | 0.086546 | 0.014305 | 1.5153 | 18.1426 | 36.9603 | 1.5702
2019-03-21 12:37 | 113 | 0.073981 | 0.016718 | 1.7766 | 20.9605 | 36.8981 | 1.5507
2019-03-21 12:37 | 114 | 0.079408 | 0.015576 | 1.664 | 19.8132 | 36.8087 | 1.5733
2019-03-21 12:37 | 115 | 0.080932 | 0.015317 | 1.6406 | 19.4781 | 36.7403 | 1.5764
2019-03-21 12:37 | 116 | 0.067314 | 0.018359 | 1.9752 | 23.2901 | 36.6511 | 1.5677
2019-03-21 12:37 | 117 | 0.035682 | 0.03471 | 3.7469 | 43.8207 | 36.6087 | 1.5636
2019-03-21 12:37 | 118 | 0.078095 | 0.015824 | 1.7136 | 19.8566 | 36.5633 | 1.5507
2019-03-21 12:37 | 119 | 0.056076 | 0.022038 | 2.3927 | 27.7011 | 36.4662 | 1.5534
2019-03-21 12:37 | 120 | 0.047907 | 0.025771 | 2.8029 | 32.2386 | 36.4485 | 1.5445
2019-03-21 12:37 | 121 | 0.072382 | 0.016798 | 1.8645 | 21.616 | 36.3629 | 1.5646
2019-03-21 12:37 | 122 | 0.055575 | 0.021916 | 2.4345 | 28.1349 | 36.3122 | 1.5636
2019-03-21 12:37 | 123 | 0.075854 | 0.01604 | 1.785 | 20.476 | 36.2442 | 1.5532
2019-03-21 12:37 | 124 | 0.072473 | 0.017105 | 1.8734 | 21.4486 | 36.1658 | 1.5544
2019-03-21 12:37 | 125 | 0.058006 | 0.021314 | 2.3319 | 26.9167 | 36.102 | 1.5613
2019-03-21 12:37 | 126 | 0.065227 | 0.018929 | 2.0753 | 23.8771 | 36.0439 | 1.5574
2019-03-21 12:37 | 127 | 0.0995 | 0.012173 | 1.3605 | 15.5786 | 35.9655 | 1.5501
2019-03-21 12:37 | 128 | 0.082216 | 0.014747 | 1.6508 | 18.8629 | 35.8569 | 1.5508
2019-03-21 12:37 | 129 | 0.072848 | 0.016952 | 1.8652 | 21.2038 | 35.7997 | 1.5447
2019-03-21 12:37 | 130 | 0.097638 | 0.012404 | 1.39 | 15.7531 | 35.704 | 1.5381
2019-03-21 12:37 | 131 | 0.067751 | 0.017875 | 2.0046 | 22.4472 | 35.6165 | 1.5208
2019-03-21 12:37 | 132 | 0.065491 | 0.018492 | 2.0764 | 23.3185 | 35.5631 | 1.5271
2019-03-21 12:37 | 133 | 0.057709 | 0.020645 | 2.3604 | 26.4064 | 35.4856 | 1.5239
2019-03-21 12:37 | 134 | 0.072836 | 0.016649 | 1.8685 | 20.8305 | 35.442 | 1.5172
2019-03-21 12:37 | 135 | 0.08845 | 0.013678 | 1.5343 | 17.0008 | 35.3401 | 1.5037
2019-03-21 12:37 | 136 | 0.068125 | 0.017855 | 2.0044 | 22.2542 | 35.2736 | 1.5161
2019-03-21 12:37 | 137 | 0.067224 | 0.018076 | 2.0255 | 22.5057 | 35.1999 | 1.5129
2019-03-21 12:37 | 138 | 0.068283 | 0.0178 | 1.9961 | 22.1724 | 35.1377 | 1.514
2019-03-21 12:37 | 139 | 0.090863 | 0.013375 | 1.4995 | 16.8097 | 35.0593 | 1.5274
2019-03-21 12:37 | 140 | 0.05625 | 0.021592 | 2.4171 | 27.366 | 34.9686 | 1.5393
2019-03-21 12:37 | 141 | 0.06377 | 0.018679 | 2.1287 | 23.8621 | 34.9375 | 1.5217
2019-03-21 12:37 | 142 | 0.082577 | 0.014166 | 1.6456 | 18.5434 | 34.8381 | 1.5313
2019-03-21 12:37 | 143 | 0.065544 | 0.018184 | 2.0722 | 23.0156 | 34.7805 | 1.5085
2019-03-21 12:37 | 144 | 0.0653 | 0.018603 | 2.0841 | 22.9643 | 34.7034 | 1.4996
2019-03-21 12:38 | 145 | 0.07763 | 0.015353 | 1.7496 | 19.0904 | 34.6464 | 1.482
2019-03-21 12:38 | 146 | 0.090798 | 0.01313 | 1.4959 | 16.1773 | 34.5488 | 1.4689
2019-03-21 12:38 | 147 | 0.057433 | 0.020786 | 2.3663 | 25.8122 | 34.4752 | 1.4825
2019-03-21 12:38 | 148 | 0.07061 | 0.016877 | 1.9193 | 20.7609 | 34.4238 | 1.4659
2019-03-21 12:38 | 149 | 0.070062 | 0.017011 | 1.9306 | 20.8307 | 34.3358 | 1.4594
2019-03-21 12:38 | 150 | 0.082916 | 0.014366 | 1.6288 | 17.6608 | 34.2807 | 1.4644
2019-03-21 12:38 | 151 | 0.09249 | 0.013155 | 1.4624 | 15.9178 | 34.1713 | 1.4722
2019-03-21 12:38 | 152 | 0.069919 | 0.017399 | 1.9313 | 20.7911 | 34.1028 | 1.4537
2019-03-21 12:38 | 153 | 0.089461 | 0.013823 | 1.505 | 16.2879 | 34.0227 | 1.4571
2019-03-21 12:38 | 154 | 0.075994 | 0.016569 | 1.7687 | 19.2124 | 33.9309 | 1.46
2019-03-21 12:38 | 155 | 0.068947 | 0.019211 | 1.9441 | 21.1805 | 33.8694 | 1.4603
2019-03-21 12:38 | 156 | 0.080847 | 0.016663 | 1.6543 | 18.1041 | 33.7884 | 1.4637
2019-03-21 12:38 | 157 | 0.084769 | 0.015877 | 1.5771 | 17.2197 | 33.7097 | 1.4597
2019-03-21 12:38 | 158 | 0.070372 | 0.019098 | 1.8942 | 20.7065 | 33.6235 | 1.4572
2019-03-21 12:38 | 159 | 0.063243 | 0.021583 | 2.1017 | 23.0237 | 33.5675 | 1.4561
2019-03-21 12:38 | 160 | 0.064855 | 0.020711 | 2.0461 | 22.4878 | 33.4949 | 1.4585
2019-03-21 12:38 | 161 | 0.05228 | 0.02524 | 2.5351 | 27.4747 | 33.4402 | 1.4364
2019-03-21 12:38 | 162 | 0.069554 | 0.018647 | 1.9002 | 20.2646 | 33.3836 | 1.4095
2019-03-21 12:38 | 163 | 0.071676 | 0.017827 | 1.846 | 19.7316 | 33.3041 | 1.4143
2019-03-21 12:38 | 164 | 0.074754 | 0.016764 | 1.7599 | 18.6679 | 33.2407 | 1.3955
2019-03-21 12:38 | 165 | 0.067024 | 0.01872 | 1.9627 | 20.835 | 33.1575 | 1.3965
2019-03-21 12:38 | 166 | 0.051695 | 0.024257 | 2.5346 | 26.9021 | 33.1077 | 1.3907
2019-03-21 12:38 | 167 | 0.078515 | 0.015715 | 1.6661 | 17.4398 | 33.0445 | 1.3693
2019-03-21 12:38 | 168 | 0.073426 | 0.016829 | 1.7803 | 18.4865 | 32.9574 | 1.3574
2019-03-21 12:38 | 169 | 0.073477 | 0.016211 | 1.7723 | 18.6645 | 32.8974 | 1.3714
2019-03-21 12:38 | 170 | 0.068203 | 0.017142 | 1.9024 | 20.0604 | 32.8114 | 1.3682
2019-03-21 12:38 | 171 | 0.076122 | 0.015095 | 1.7033 | 17.9441 | 32.7577 | 1.3659
2019-03-21 12:38 | 172 | 0.087905 | 0.012838 | 1.4703 | 15.3713 | 32.6585 | 1.3512
2019-03-21 12:38 | 173 | 0.074217 | 0.014914 | 1.7404 | 18.293 | 32.588 | 1.3577
2019-03-21 12:38 | 174 | 0.069833 | 0.015842 | 1.8432 | 19.3069 | 32.5084 | 1.3483
2019-03-21 12:38 | 175 | 0.063142 | 0.016809 | 2.0289 | 21.1094 | 32.4477 | 1.3329
2019-03-21 12:38 | 176 | 0.09298 | 0.011178 | 1.3749 | 14.324 | 32.3741 | 1.3318
2019-03-21 12:38 | 177 | 0.081722 | 0.012429 | 1.5561 | 16.1127 | 32.2708 | 1.3168
2019-03-21 12:38 | 178 | 0.081521 | 0.012194 | 1.5559 | 16.1589 | 32.209 | 1.3173
2019-03-21 12:38 | 179 | 0.079592 | 0.012472 | 1.5866 | 16.1678 | 32.1081 | 1.2868
2019-03-21 12:38 | 180 | 0.071987 | 0.013826 | 1.7512 | 18.1011 | 32.0509 | 1.3031
2019-03-21 12:38 | 181 | 0.080194 | 0.012417 | 1.5675 | 16.1251 | 31.9611 | 1.2931
2019-03-21 12:38 | 182 | 0.0571 | 0.017029 | 2.1889 | 22.6028 | 31.8972 | 1.2906
2019-03-21 12:38 | 183 | 0.06122 | 0.016236 | 2.032 | 20.9616 | 31.8408 | 1.2833
2019-03-21 12:38 | 184 | 0.06816 | 0.014569 | 1.8188 | 18.4451 | 31.7735 | 1.2572
2019-03-21 12:38 | 185 | 0.072323 | 0.013752 | 1.7092 | 17.5476 | 31.7058 | 1.2691
2019-03-21 12:38 | 186 | 0.049913 | 0.020353 | 2.4641 | 25.2868 | 31.6354 | 1.2621
2019-03-21 12:38 | 187 | 0.036037 | 0.02822 | 3.4021 | 34.9345 | 31.6036 | 1.2589
2019-03-21 12:38 | 188 | 0.05688 | 0.01829 | 2.1508 | 22.0619 | 31.5544 | 1.2549
2019-03-21 12:38 | 189 | 0.071297 | 0.014556 | 1.7063 | 17.4048 | 31.4906 | 1.2409
2019-03-21 12:38 | 190 | 0.081261 | 0.01275 | 1.489 | 15.1739 | 31.414 | 1.233
2019-03-21 12:38 | 191 | 0.059014 | 0.017964 | 2.0478 | 20.9623 | 31.3359 | 1.2371
2019-03-21 12:38 | 192 | 0.053604 | 0.019767 | 2.2469 | 23.1501 | 31.2924 | 1.2409
2019-03-21 12:38 | 193 | 0.056821 | 0.019117 | 2.1231 | 21.4075 | 31.226 | 1.2164
2019-03-21 12:38 | 194 | 0.062731 | 0.017319 | 1.9139 | 19.7563 | 31.1776 | 1.2393
2019-03-21 12:38 | 195 | 0.077276 | 0.01406 | 1.5493 | 15.6269 | 31.099 | 1.2076
2019-03-21 12:38 | 196 | 0.060816 | 0.017854 | 1.9611 | 19.3956 | 31.0293 | 1.1796
2019-03-21 12:38 | 197 | 0.093773 | 0.01158 | 1.2656 | 12.5815 | 30.9674 | 1.1798
2019-03-21 12:38 | 198 | 0.07368 | 0.014743 | 1.6061 | 15.6382 | 30.8544 | 1.1522
2019-03-21 12:38 | 199 | 0.034313 | 0.031668 | 3.4286 | 32.9129 | 30.8248 | 1.1293
2019-03-21 12:38 | 200 | 0.048263 | 0.022071 | 2.4234 | 23.1861 | 30.7727 | 1.119
2019-03-21 12:38 | 201 | 0.066985 | 0.01592 | 1.7397 | 16.6156 | 30.7253 | 1.113
2019-03-21 12:38 | 202 | 0.091852 | 0.011592 | 1.2577 | 11.8979 | 30.6377 | 1.0929
2019-03-21 12:38 | 203 | 0.08685 | 0.012266 | 1.3203 | 12.3796 | 30.5491 | 1.0752
2019-03-21 12:38 | 204 | 0.083598 | 0.012511 | 1.3644 | 12.9289 | 30.4652 | 1.0808
2019-03-21 12:38 | 205 | 0.052635 | 0.019895 | 2.1539 | 20.3845 | 30.3878 | 1.0729
2019-03-21 12:39 | 206 | 0.065327 | 0.016048 | 1.7186 | 16.5372 | 30.3502 | 1.0803
2019-03-21 12:39 | 207 | 0.059065 | 0.017748 | 1.8864 | 18.2516 | 30.2607 | 1.078
2019-03-21 12:39 | 208 | 0.050154 | 0.02048 | 2.2059 | 20.6893 | 30.2325 | 1.0376
2019-03-21 12:39 | 209 | 0.067949 | 0.015409 | 1.61 | 15.0493 | 30.1549 | 1.0226
2019-03-21 12:39 | 210 | 0.043727 | 0.023913 | 2.4794 | 23.3505 | 30.1058 | 1.0211
2019-03-21 12:39 | 211 | 0.04497 | 0.023825 | 2.3953 | 22.6393 | 30.0624 | 1.0181
2019-03-21 12:39 | 212 | 0.037355 | 0.029251 | 2.8593 | 27.1376 | 30.0169 | 1.0137
2019-03-21 12:39 | 213 | 0.047899 | 0.023245 | 2.2063 | 21.1466 | 29.9835 | 1.0129
2019-03-21 12:39 | 214 | 0.04735 | 0.023987 | 2.2151 | 21.0341 | 29.9233 | 0.99596
2019-03-21 12:39 | 215 | 0.053843 | 0.021912 | 1.9315 | 18.8122 | 29.8876 | 1.0129
2019-03-21 12:39 | 216 | 0.048417 | 0.025304 | 2.1356 | 20.6429 | 29.8181 | 0.99948
2019-03-21 12:39 | 217 | 0.051772 | 0.024552 | 1.9866 | 19.6448 | 29.7887 | 1.0171
2019-03-21 12:39 | 218 | 0.061425 | 0.020705 | 1.6645 | 16.4445 | 29.7131 | 1.0101
2019-03-21 12:39 | 219 | 0.057008 | 0.022327 | 1.7851 | 17.7994 | 29.6681 | 1.0147
2019-03-21 12:39 | 220 | 0.094611 | 0.013271 | 1.0751 | 10.7325 | 29.5906 | 1.0154
2019-03-21 12:39 | 221 | 0.069395 | 0.018385 | 1.4592 | 14.3964 | 29.4931 | 0.99903
2019-03-21 12:39 | 222 | 0.0646 | 0.020132 | 1.5605 | 15.7242 | 29.4484 | 1.0158
2019-03-21 12:39 | 223 | 0.062239 | 0.021587 | 1.6158 | 16.3097 | 29.3634 | 1.0151
2019-03-21 12:39 | 224 | 0.038357 | 0.03673 | 2.6101 | 26.2962 | 29.328 | 1.0086
2019-03-21 12:39 | 225 | 0.057271 | 0.026209 | 1.7396 | 18.0348 | 29.2767 | 1.0329
2019-03-21 12:39 | 226 | 0.065934 | 0.023775 | 1.5062 | 16.0997 | 29.2153 | 1.0615
2019-03-21 12:39 | 227 | 0.059696 | 0.027775 | 1.6591 | 17.6961 | 29.149 | 1.0564
2019-03-21 12:39 | 228 | 0.04679 | 0.036859 | 2.11 | 22.3073 | 29.0975 | 1.0438
2019-03-21 12:39 | 229 | 0.054871 | 0.033027 | 1.7915 | 18.8206 | 29.0508 | 1.0327
2019-03-21 12:39 | 230 | 0.046158 | 0.040731 | 2.1259 | 22.3817 | 28.9912 | 1.0331
2019-03-21 12:39 | 231 | 0.043978 | 0.044287 | 2.2342 | 23.575 | 28.9572 | 1.0368
2019-03-21 12:39 | 232 | 0.042987 | 0.046758 | 2.277 | 24.3154 | 28.9025 | 1.0452
2019-03-21 12:39 | 233 | 0.049268 | 0.042238 | 1.9862 | 21.1855 | 28.8688 | 1.0438
2019-03-21 12:39 | 234 | 0.077364 | 0.028053 | 1.2634 | 13.7064 | 28.7994 | 1.0604
2019-03-21 12:39 | 235 | 0.065967 | 0.033883 | 1.4802 | 15.8436 | 28.7238 | 1.0452
2019-03-21 12:39 | 236 | 0.042768 | 0.053411 | 2.2835 | 24.9407 | 28.671 | 1.0667
2019-03-21 12:39 | 237 | 0.037544 | 0.061975 | 2.5996 | 28.1373 | 28.6332 | 1.0564
2019-03-21 12:39 | 238 | 0.065537 | 0.036211 | 1.4886 | 16.3314 | 28.5877 | 1.0703
2019-03-21 12:39 | 239 | 0.071338 | 0.03386 | 1.3677 | 14.8141 | 28.5075 | 1.0568
2019-03-21 12:39 | 240 | 0.049228 | 0.050347 | 1.9808 | 21.4247 | 28.4517 | 1.0547
2019-03-21 12:39 | 241 | 0.059855 | 0.042165 | 1.6319 | 17.6228 | 28.4018 | 1.0548
2019-03-21 12:39 | 242 | 0.063567 | 0.040423 | 1.5383 | 16.7077 | 28.3329 | 1.062
2019-03-21 12:39 | 243 | 0.068935 | 0.03792 | 1.4189 | 15.341 | 28.275 | 1.0575
2019-03-21 12:39 | 244 | 0.050087 | 0.053166 | 1.9532 | 22.1107 | 28.2005 | 1.1075
2019-03-21 12:39 | 245 | 0.050101 | 0.053849 | 1.947 | 21.6547 | 28.1699 | 1.0849
2019-03-21 12:39 | 246 | 0.082989 | 0.033056 | 1.1762 | 13.3113 | 28.0929 | 1.1047
2019-03-21 12:39 | 247 | 0.065689 | 0.042005 | 1.483 | 16.9348 | 28.0152 | 1.1124
2019-03-21 12:39 | 248 | 0.063485 | 0.04381 | 1.5399 | 17.4603 | 27.9593 | 1.1085
2019-03-21 12:39 | 249 | 0.056855 | 0.049801 | 1.7254 | 19.4603 | 27.8891 | 1.1064
2019-03-21 12:39 | 250 | 0.04933 | 0.057616 | 1.9813 | 22.3808 | 27.8454 | 1.1041
2019-03-21 12:39 | 251 | 0.054334 | 0.053646 | 1.8057 | 20.2305 | 27.787 | 1.0992
2019-03-21 12:39 | 252 | 0.068607 | 0.042692 | 1.4282 | 16.0768 | 27.7344 | 1.103
2019-03-21 12:39 | 253 | 0.069837 | 0.042631 | 1.4088 | 15.8668 | 27.6527 | 1.1081
2019-03-21 12:39 | 254 | 0.074535 | 0.040274 | 1.3225 | 15.0335 | 27.5941 | 1.1205
2019-03-21 12:39 | 255 | 0.084039 | 0.03589 | 1.1703 | 13.1891 | 27.5031 | 1.1084
2019-03-21 12:39 | 256 | 0.066256 | 0.045491 | 1.4891 | 16.935 | 27.432 | 1.122
2019-03-21 12:39 | 257 | 0.080964 | 0.036926 | 1.222 | 13.9389 | 27.364 | 1.1285
2019-03-21 12:39 | 258 | 0.063683 | 0.046495 | 1.5512 | 17.7103 | 27.2771 | 1.1278
2019-03-21 12:39 | 259 | 0.052108 | 0.056505 | 1.9014 | 21.5128 | 27.235 | 1.121
2019-03-21 12:39 | 260 | 0.076848 | 0.038237 | 1.2903 | 14.5174 | 27.1642 | 1.1156
2019-03-21 12:39 | 261 | 0.076697 | 0.038249 | 1.2934 | 14.5743 | 27.0868 | 1.1178
2019-03-21 12:39 | 262 | 0.07038 | 0.042279 | 1.4095 | 15.7865 | 27.0132 | 1.1111
2019-03-21 12:39 | 263 | 0.059002 | 0.051578 | 1.6844 | 19.2071 | 26.9468 | 1.1333
2019-03-21 12:39 | 264 | 0.074054 | 0.041281 | 1.3417 | 15.7895 | 26.889 | 1.1693
2019-03-21 12:40 | 265 | 0.0769 | 0.039668 | 1.2924 | 15.4608 | 26.8013 | 1.1889
2019-03-21 12:40 | 266 | 0.066643 | 0.045091 | 1.4926 | 19.2139 | 26.738 | 1.2805
2019-03-21 12:40 | 267 | 0.08353 | 0.03531 | 1.1961 | 17.0992 | 26.6616 | 1.4283
2019-03-21 12:40 | 268 | 0.096032 | 0.029809 | 1.0488 | 17.0308 | 26.571 | 1.6355
2019-03-21 12:40 | 269 | 0.1182 | 0.02379 | 0.86417 | 14.9928 | 26.468 | 1.7721
2019-03-21 12:40 | 270 | 0.11259 | 0.024563 | 0.92758 | 16.9483 | 26.3412 | 1.9082
2019-03-21 12:40 | 271 | 0.11507 | 0.023608 | 0.93815 | 17.4842 | 26.2415 | 2.0119
2019-03-21 12:40 | 272 | 0.12146 | 0.022354 | 0.92455 | 17.5414 | 26.1103 | 2.1307
2019-03-21 12:40 | 273 | 0.118 | 0.022727 | 0.99248 | 18.578 | 26.0013 | 2.1923
2019-03-21 12:40 | 274 | 0.1012 | 0.02665 | 1.2111 | 23.509 | 25.877 | 2.3791
2019-03-21 12:40 | 275 | 0.11312 | 0.024155 | 1.1351 | 21.6268 | 25.7914 | 2.4464
2019-03-21 12:40 | 276 | 0.14824 | 0.018683 | 0.9065 | 16.8159 | 25.6459 | 2.4927
2019-03-21 12:40 | 277 | 0.11751 | 0.024125 | 1.1979 | 22.4846 | 25.5101 | 2.6421
2019-03-21 12:40 | 278 | 0.11949 | 0.024049 | 1.2279 | 22.3464 | 25.405 | 2.6701
2019-03-21 12:40 | 279 | 0.10402 | 0.02809 | 1.4752 | 26.219 | 25.2741 | 2.7272
2019-03-21 12:40 | 280 | 0.1107 | 0.026345 | 1.4386 | 24.4142 | 25.1919 | 2.7026
2019-03-21 12:40 | 281 | 0.11854 | 0.02463 | 1.3982 | 22.7076 | 25.0533 | 2.6917
2019-03-21 12:40 | 282 | 0.071331 | 0.040698 | 2.4123 | 37.7217 | 24.9671 | 2.6907
2019-03-21 12:40 | 283 | 0.07858 | 0.035807 | 2.2512 | 33.3404 | 24.8982 | 2.6199
2019-03-21 12:40 | 284 | 0.095775 | 0.028751 | 1.8976 | 26.4176 | 24.8069 | 2.5301
2019-03-21 12:40 | 285 | 0.079488 | 0.033886 | 2.3375 | 30.9134 | 24.7147 | 2.4572
2019-03-21 12:40 | 286 | 0.06081 | 0.043228 | 3.1008 | 37.9605 | 24.6483 | 2.3084
2019-03-21 12:40 | 287 | 0.082409 | 0.031148 | 2.3116 | 27.5742 | 24.5836 | 2.2724
2019-03-21 12:40 | 288 | 0.081956 | 0.030756 | 2.3323 | 27.3052 | 24.4882 | 2.2378
2019-03-21 12:40 | 289 | 0.069406 | 0.035509 | 2.7665 | 31.111 | 24.4228 | 2.1593
2019-03-21 12:40 | 290 | 0.078143 | 0.031239 | 2.4506 | 27.1226 | 24.3448 | 2.1194
2019-03-21 12:40 | 291 | 0.069515 | 0.034604 | 2.7517 | 30.1825 | 24.2694 | 2.0981
2019-03-21 12:40 | 292 | 0.093309 | 0.025598 | 2.0423 | 22.374 | 24.1989 | 2.0877
2019-03-21 12:40 | 293 | 0.086643 | 0.027023 | 2.181 | 23.7583 | 24.0897 | 2.0585
2019-03-21 12:40 | 294 | 0.076603 | 0.030327 | 2.4568 | 26.3836 | 24.0273 | 2.0211
2019-03-21 12:40 | 295 | 0.066285 | 0.034785 | 2.8213 | 29.5448 | 23.9361 | 1.9584
2019-03-21 12:40 | 296 | 0.077448 | 0.029752 | 2.3923 | 25.0386 | 23.8893 | 1.9392
2019-03-21 12:40 | 297 | 0.090114 | 0.025587 | 2.0381 | 21.4847 | 23.7813 | 1.9361
2019-03-21 12:40 | 298 | 0.063439 | 0.036671 | 2.8711 | 30.6551 | 23.7179 | 1.9447
2019-03-21 12:40 | 299 | 0.078621 | 0.030117 | 2.2921 | 24.9415 | 23.6459 | 1.9609
2019-03-21 12:40 | 300 | 0.056072 | 0.042977 | 3.1863 | 34.5203 | 23.5676 | 1.9356
2019-03-21 12:40 | 301 | 0.084418 | 0.029049 | 2.1014 | 23.01 | 23.5222 | 1.9425
2019-03-21 12:40 | 302 | 0.099485 | 0.024886 | 1.7752 | 19.5483 | 23.4015 | 1.9448
2019-03-21 12:40 | 303 | 0.084462 | 0.029323 | 2.08 | 23.0758 | 23.3309 | 1.949
2019-03-21 12:40 | 304 | 0.080302 | 0.030812 | 2.1793 | 24.9798 | 23.2303 | 2.0059
2019-03-21 12:40 | 305 | 0.094066 | 0.026318 | 1.8567 | 21.6218 | 23.1659 | 2.0339
2019-03-21 12:40 | 306 | 0.096451 | 0.026145 | 1.8106 | 21.3761 | 23.0453 | 2.0617
2019-03-21 12:40 | 307 | 0.069257 | 0.037357 | 2.5228 | 30.1787 | 22.9792 | 2.0901
2019-03-21 12:40 | 308 | 0.09836 | 0.026969 | 1.7785 | 21.5461 | 22.8942 | 2.1193
2019-03-21 12:40 | 309 | 0.097115 | 0.028275 | 1.8115 | 22.0935 | 22.7887 | 2.1456
2019-03-21 12:40 | 310 | 0.093758 | 0.030551 | 1.8753 | 22.9716 | 22.7014 | 2.1538
2019-03-21 12:40 | 311 | 0.086412 | 0.034688 | 2.0445 | 24.9665 | 22.602 | 2.1574
2019-03-21 12:40 | 312 | 0.090435 | 0.03422 | 1.9558 | 24.1545 | 22.526 | 2.1844
2019-03-21 12:40 | 313 | 0.088936 | 0.036034 | 2.0003 | 24.7389 | 22.4221 | 2.2002
2019-03-21 12:40 | 314 | 0.086787 | 0.03834 | 2.0555 | 25.5743 | 22.3475 | 2.2195
2019-03-21 12:40 | 315 | 0.11437 | 0.02987 | 1.571 | 19.6928 | 22.242 | 2.2524
2019-03-21 12:40 | 316 | 0.10928 | 0.032158 | 1.6505 | 21.0088 | 22.1262 | 2.2959
2019-03-21 12:40 | 317 | 0.10655 | 0.033955 | 1.703 | 21.8893 | 22.0237 | 2.3323
2019-03-21 12:40 | 318 | 0.1033 | 0.036022 | 1.7662 | 22.9742 | 21.913 | 2.3732
2019-03-21 12:40 | 319 | 0.10598 | 0.035351 | 1.7285 | 22.7158 | 21.8154 | 2.4074
2019-03-21 12:40 | 320 | 0.11784 | 0.032314 | 1.5659 | 20.5907 | 21.6994 | 2.4265
2019-03-21 12:40 | 321 | 0.099254 | 0.038635 | 1.8787 | 24.6235 | 21.5867 | 2.444
2019-03-21 12:40 | 322 | 0.094897 | 0.040283 | 1.9772 | 25.8884 | 21.4971 | 2.4567
2019-03-21 12:40 | 323 | 0.12913 | 0.029492 | 1.47 | 19.2485 | 21.3875 | 2.4855
2019-03-21 12:40 | 324 | 0.14005 | 0.026829 | 1.3654 | 17.7611 | 21.2452 | 2.4874
2019-03-21 12:40 | 325 | 0.0944 | 0.039689 | 2.0379 | 26.476 | 21.1211 | 2.4993
2019-03-21 12:41 | 326 | 0.095059 | 0.039157 | 2.0422 | 26.7871 | 21.0458 | 2.5464
2019-03-21 12:41 | 327 | 0.11432 | 0.032271 | 1.7099 | 22.5238 | 20.9253 | 2.5749
2019-03-21 12:41 | 328 | 0.11903 | 0.03063 | 1.6586 | 22.287 | 20.82 | 2.6527
2019-03-21 12:41 | 329 | 0.13062 | 0.027571 | 1.5269 | 20.3154 | 20.6866 | 2.6535
2019-03-21 12:41 | 330 | 0.11762 | 0.030235 | 1.7132 | 22.6964 | 20.5642 | 2.6695
2019-03-21 12:41 | 331 | 0.13169 | 0.026666 | 1.5466 | 20.742 | 20.4458 | 2.7315
2019-03-21 12:41 | 332 | 0.12491 | 0.027533 | 1.6441 | 22.2419 | 20.3055 | 2.7783
2019-03-21 12:41 | 333 | 0.11129 | 0.030611 | 1.8718 | 25.6465 | 20.1979 | 2.8542
2019-03-21 12:41 | 334 | 0.11296 | 0.030149 | 1.8653 | 25.6125 | 20.0792 | 2.8931
2019-03-21 12:41 | 335 | 0.11585 | 0.029588 | 1.8415 | 24.8343 | 19.971 | 2.877
2019-03-21 12:41 | 336 | 0.12511 | 0.0278 | 1.7307 | 23.0262 | 19.8464 | 2.8809
2019-03-21 12:41 | 337 | 0.1149 | 0.030894 | 1.914 | 24.8419 | 19.7255 | 2.8542
2019-03-21 12:41 | 338 | 0.11006 | 0.0329 | 2.0257 | 25.4467 | 19.6161 | 2.8007
2019-03-21 12:41 | 339 | 0.096467 | 0.038306 | 2.3407 | 28.6737 | 19.5068 | 2.766
2019-03-21 12:41 | 340 | 0.10527 | 0.035508 | 2.1602 | 26.0609 | 19.4179 | 2.7433
2019-03-21 12:41 | 341 | 0.10603 | 0.035338 | 2.1465 | 25.6229 | 19.2977 | 2.7168
2019-03-21 12:41 | 342 | 0.10411 | 0.035736 | 2.1898 | 26.1023 | 19.2065 | 2.7174
2019-03-21 12:41 | 343 | 0.10916 | 0.033668 | 2.0893 | 24.5102 | 19.088 | 2.6756
2019-03-21 12:41 | 344 | 0.10556 | 0.034372 | 2.1596 | 25.6396 | 18.9896 | 2.7066
2019-03-21 12:41 | 345 | 0.12261 | 0.028909 | 1.8587 | 22.3702 | 18.8728 | 2.7428
2019-03-21 12:41 | 346 | 0.10811 | 0.032138 | 2.1048 | 25.909 | 18.7514 | 2.8011
2019-03-21 12:41 | 347 | 0.10769 | 0.031543 | 2.1062 | 26.0269 | 18.6538 | 2.8029
2019-03-21 12:41 | 348 | 0.10736 | 0.031169 | 2.1132 | 26.7339 | 18.5352 | 2.8702
2019-03-21 12:41 | 349 | 0.12216 | 0.027209 | 1.8656 | 24.1831 | 18.4354 | 2.9542
2019-03-21 12:41 | 350 | 0.13027 | 0.025337 | 1.7587 | 22.8542 | 18.2928 | 2.9772
2019-03-21 12:41 | 351 | 0.11414 | 0.028838 | 2.0171 | 26.3683 | 18.1805 | 3.0098
2019-03-21 12:41 | 352 | 0.12541 | 0.026057 | 1.8526 | 24.4553 | 18.0588 | 3.0669
2019-03-21 12:41 | 353 | 0.12035 | 0.027006 | 1.9519 | 25.6386 | 17.9326 | 3.0857
2019-03-21 12:41 | 354 | 0.1313 | 0.024672 | 1.8015 | 23.4457 | 17.8149 | 3.0784
2019-03-21 12:41 | 355 | 0.12447 | 0.025975 | 1.9142 | 24.6287 | 17.674 | 3.0656
2019-03-21 12:41 | 356 | 0.12124 | 0.026772 | 1.9781 | 25.1484 | 17.5651 | 3.0491
2019-03-21 12:41 | 357 | 0.1365 | 0.023765 | 1.7696 | 22.3828 | 17.4276 | 3.0552
2019-03-21 12:41 | 358 | 0.1232 | 0.025982 | 1.9717 | 24.8909 | 17.2985 | 3.0666
2019-03-21 12:41 | 359 | 0.12054 | 0.026023 | 2.0252 | 25.8568 | 17.1797 | 3.1168
2019-03-21 12:41 | 360 | 0.10271 | 0.029529 | 2.39 | 31.7764 | 17.0602 | 3.2636
2019-03-21 12:41 | 361 | 0.11112 | 0.02653 | 2.2189 | 30.3072 | 16.9678 | 3.3678
2019-03-21 12:41 | 362 | 0.12696 | 0.022553 | 1.9549 | 26.4779 | 16.836 | 3.3615
2019-03-21 12:41 | 363 | 0.12195 | 0.022813 | 2.0554 | 27.5116 | 16.7192 | 3.355
2019-03-21 12:41 | 364 | 0.10231 | 0.026836 | 2.4801 | 32.4882 | 16.5952 | 3.3238
2019-03-21 12:41 | 365 | 0.12429 | 0.022083 | 2.0613 | 26.6341 | 16.5049 | 3.3104
2019-03-21 12:41 | 366 | 0.13469 | 0.020432 | 1.9272 | 24.5554 | 16.3494 | 3.3073
2019-03-21 12:41 | 367 | 0.10623 | 0.025912 | 2.4641 | 30.714 | 16.244 | 3.2628
2019-03-21 12:41 | 368 | 0.13672 | 0.020388 | 1.922 | 23.4923 | 16.1242 | 3.2118
2019-03-21 12:41 | 369 | 0.12758 | 0.022214 | 2.0737 | 24.9231 | 15.9784 | 3.1797
2019-03-21 12:41 | 370 | 0.1394 | 0.020766 | 1.9028 | 22.6134 | 15.8654 | 3.1522
2019-03-21 12:41 | 371 | 0.13286 | 0.022111 | 2.0017 | 23.6191 | 15.704 | 3.138
2019-03-21 12:41 | 372 | 0.11614 | 0.02561 | 2.2864 | 26.6651 | 15.6017 | 3.0968
2019-03-21 12:41 | 373 | 0.1356 | 0.022391 | 1.9565 | 22.5446 | 15.4631 | 3.0571
2019-03-21 12:41 | 374 | 0.1472 | 0.020889 | 1.7965 | 21.0923 | 15.3319 | 3.1049
2019-03-21 12:41 | 375 | 0.14619 | 0.021301 | 1.8033 | 21.2243 | 15.1729 | 3.1027
2019-03-21 12:41 | 376 | 0.11665 | 0.02696 | 2.2479 | 26.7418 | 15.046 | 3.1195
2019-03-21 12:41 | 377 | 0.12941 | 0.024448 | 2.0234 | 24.3866 | 14.9302 | 3.1559
2019-03-21 12:41 | 378 | 0.1247 | 0.025303 | 2.0957 | 25.3727 | 14.7903 | 3.164
2019-03-21 12:41 | 379 | 0.12128 | 0.025975 | 2.1529 | 26.1379 | 14.6804 | 3.17
2019-03-21 12:41 | 380 | 0.13131 | 0.023666 | 1.991 | 24.3535 | 14.5442 | 3.1979
2019-03-21 12:41 | 381 | 0.1444 | 0.021247 | 1.8135 | 22.7179 | 14.417 | 3.2805
2019-03-21 12:41 | 382 | 0.14647 | 0.020702 | 1.7937 | 22.768 | 14.2587 | 3.3348
2019-03-21 12:41 | 383 | 0.12965 | 0.023173 | 2.0295 | 26.2385 | 14.1292 | 3.4019
2019-03-21 12:41 | 384 | 0.11543 | 0.025677 | 2.2899 | 29.4277 | 13.9991 | 3.3967
2019-03-21 12:42 | 385 | 0.10595 | 0.027869 | 2.518 | 32.0769 | 13.8956 | 3.3986
2019-03-21 12:42 | 386 | 0.14584 | 0.020021 | 1.8327 | 23.2503 | 13.7752 | 3.3908
2019-03-21 12:42 | 387 | 0.15829 | 0.018335 | 1.7013 | 21.4544 | 13.6101 | 3.396
2019-03-21 12:42 | 388 | 0.1477 | 0.01981 | 1.8354 | 22.9702 | 13.465 | 3.3927
2019-03-21 12:42 | 389 | 0.14838 | 0.019954 | 1.8324 | 22.8512 | 13.3131 | 3.3906
2019-03-21 12:42 | 390 | 0.13147 | 0.022843 | 2.0778 | 26.1106 | 13.1721 | 3.4327
2019-03-21 12:42 | 391 | 0.12492 | 0.024344 | 2.1937 | 27.3696 | 13.0476 | 3.419
2019-03-21 12:42 | 392 | 0.13254 | 0.023091 | 2.0746 | 25.8787 | 12.9181 | 3.4299
2019-03-21 12:42 | 393 | 0.1491 | 0.020501 | 1.8488 | 22.9113 | 12.7801 | 3.416
2019-03-21 12:42 | 394 | 0.15376 | 0.01992 | 1.8032 | 22.2459 | 12.623 | 3.4205
2019-03-21 12:42 | 395 | 0.14699 | 0.020892 | 1.8856 | 23.1199 | 12.4761 | 3.3984
2019-03-21 12:42 | 396 | 0.13478 | 0.023155 | 2.0671 | 25.2894 | 12.331 | 3.4085
2019-03-21 12:42 | 397 | 0.11732 | 0.026549 | 2.3775 | 28.8735 | 12.2075 | 3.3876
2019-03-21 12:42 | 398 | 0.12014 | 0.026177 | 2.3339 | 27.9997 | 12.0912 | 3.3637
2019-03-21 12:42 | 399 | 0.12475 | 0.025186 | 2.2492 | 26.3611 | 11.9659 | 3.2886
2019-03-21 12:42 | 400 | 0.13762 | 0.022815 | 2.0382 | 23.835 | 11.8401 | 3.2803
2019-03-21 12:42 | 401 | 0.12992 | 0.024317 | 2.1548 | 25.1844 | 11.6954 | 3.272
2019-03-21 12:42 | 402 | 0.13134 | 0.02402 | 2.1243 | 25.4131 | 11.5789 | 3.3378
2019-03-21 12:42 | 403 | 0.1168 | 0.026786 | 2.3806 | 28.7332 | 11.4359 | 3.3559
2019-03-21 12:42 | 404 | 0.12379 | 0.025015 | 2.2362 | 27.5968 | 11.3398 | 3.4161
2019-03-21 12:42 | 405 | 0.15084 | 0.020444 | 1.8397 | 23.065 | 11.1839 | 3.479
2019-03-21 12:42 | 406 | 0.12924 | 0.023456 | 2.1416 | 27.0199 | 11.0487 | 3.492
2019-03-21 12:42 | 407 | 0.14764 | 0.020299 | 1.8846 | 24.0195 | 10.9173 | 3.5462
2019-03-21 12:42 | 408 | 0.145 | 0.020496 | 1.9241 | 24.8474 | 10.7579 | 3.6028
2019-03-21 12:42 | 409 | 0.13599 | 0.021847 | 2.0608 | 26.7056 | 10.629 | 3.6318
2019-03-21 12:42 | 410 | 0.14568 | 0.020552 | 1.9386 | 25.3122 | 10.4819 | 3.6876
2019-03-21 12:42 | 411 | 0.13997 | 0.021513 | 2.0293 | 26.3254 | 10.3411 | 3.6847
2019-03-21 12:42 | 412 | 0.13073 | 0.023367 | 2.1919 | 28.1915 | 10.2027 | 3.6855
2019-03-21 12:42 | 413 | 0.14294 | 0.021639 | 2.0181 | 25.707 | 10.0749 | 3.6745
2019-03-21 12:42 | 414 | 0.13613 | 0.02299 | 2.1311 | 26.8478 | 9.9216 | 3.6547
2019-03-21 12:42 | 415 | 0.10601 | 0.029949 | 2.7577 | 33.8265 | 9.8075 | 3.5858
2019-03-21 12:42 | 416 | 0.13312 | 0.023984 | 2.2045 | 26.7888 | 9.6967 | 3.5662
2019-03-21 12:42 | 417 | 0.1264 | 0.025198 | 2.3241 | 28.1201 | 9.5479 | 3.5545
2019-03-21 12:42 | 418 | 0.12707 | 0.025097 | 2.3186 | 28.0226 | 9.4422 | 3.5609
2019-03-21 12:42 | 419 | 0.1511 | 0.021072 | 1.9476 | 23.5034 | 9.2888 | 3.5513
2019-03-21 12:42 | 420 | 0.13522 | 0.02341 | 2.177 | 26.4218 | 9.1486 | 3.5727
2019-03-21 12:42 | 421 | 0.15182 | 0.020877 | 1.9415 | 23.7481 | 9.0119 | 3.6056
2019-03-21 12:42 | 422 | 0.14545 | 0.021655 | 2.0292 | 24.7156 | 8.8503 | 3.5948
2019-03-21 12:42 | 423 | 0.12535 | 0.025335 | 2.3624 | 28.6952 | 8.7246 | 3.5969
2019-03-21 12:42 | 424 | 0.12512 | 0.02552 | 2.3669 | 28.5152 | 8.595 | 3.5679
2019-03-21 12:42 | 425 | 0.12738 | 0.025279 | 2.3319 | 27.6537 | 8.4731 | 3.5225
2019-03-21 12:42 | 426 | 0.12934 | 0.024825 | 2.2907 | 27.045 | 8.3397 | 3.4979
2019-03-21 12:42 | 427 | 0.14461 | 0.02225 | 2.0541 | 23.8884 | 8.2111 | 3.4545
2019-03-21 12:42 | 428 | 0.15679 | 0.020536 | 1.8944 | 22.0108 | 8.0516 | 3.451
2019-03-21 12:42 | 429 | 0.14435 | 0.022305 | 2.0536 | 23.8355 | 7.9037 | 3.4406
2019-03-21 12:42 | 430 | 0.1422 | 0.022901 | 2.0785 | 24.1413 | 7.7618 | 3.433
2019-03-21 12:42 | 431 | 0.11689 | 0.027877 | 2.5233 | 29.2896 | 7.6245 | 3.4235
2019-03-21 12:42 | 432 | 0.1036 | 0.031675 | 2.8427 | 32.4583 | 7.5247 | 3.3627
2019-03-21 12:42 | 433 | 0.11831 | 0.027886 | 2.4808 | 27.33 | 7.4099 | 3.2334
2019-03-21 12:42 | 434 | 0.13375 | 0.024961 | 2.186 | 24.0527 | 7.2876 | 3.2171
2019-03-21 12:42 | 435 | 0.12924 | 0.026125 | 2.2498 | 24.6584 | 7.1473 | 3.1868
2019-03-21 12:42 | 436 | 0.1265 | 0.027063 | 2.2897 | 25.108 | 7.0293 | 3.176
2019-03-21 12:42 | 437 | 0.12603 | 0.02734 | 2.2862 | 25.1359 | 6.8944 | 3.1678
2019-03-21 12:42 | 438 | 0.10462 | 0.03304 | 2.727 | 30.2642 | 6.782 | 3.1661
2019-03-21 12:42 | 439 | 0.10133 | 0.03439 | 2.8016 | 31.5557 | 6.6808 | 3.1975
2019-03-21 12:42 | 440 | 0.10612 | 0.032823 | 2.6586 | 30.0131 | 6.5762 | 3.185
2019-03-21 12:42 | 441 | 0.1279 | 0.027067 | 2.1905 | 25.1662 | 6.4644 | 3.2188
2019-03-21 12:42 | 442 | 0.13495 | 0.02566 | 2.0681 | 23.8931 | 6.3241 | 3.2243
2019-03-21 12:42 | 443 | 0.12759 | 0.027339 | 2.183 | 25.2319 | 6.1985 | 3.2194
2019-03-21 12:42 | 444 | 0.12453 | 0.028205 | 2.2323 | 25.8379 | 6.0683 | 3.2176
2019-03-21 12:42 | 445 | 0.12101 | 0.02918 | 2.2919 | 26.6132 | 5.9495 | 3.2205
2019-03-21 12:43 | 446 | 0.11692 | 0.030565 | 2.368 | 27.3264 | 5.8267 | 3.195
2019-03-21 12:43 | 447 | 0.10087 | 0.035259 | 2.7404 | 31.2626 | 5.7178 | 3.1534
2019-03-21 12:43 | 448 | 0.11567 | 0.030762 | 2.3848 | 26.9651 | 5.6175 | 3.119
2019-03-21 12:43 | 449 | 0.13084 | 0.02721 | 2.1016 | 23.6544 | 5.4863 | 3.0949
2019-03-21 12:43 | 450 | 0.11871 | 0.030126 | 2.3032 | 25.8684 | 5.3626 | 3.0707
2019-03-21 12:43 | 451 | 0.10987 | 0.032601 | 2.4807 | 28.2457 | 5.2483 | 3.1034
2019-03-21 12:43 | 452 | 0.11972 | 0 | 0 | 0 | 5.1382 | 3.0848
2019-03-21 12:43 | 453 | 0.12662 | 0 | 0 | 0 | 5.0095 | 3.1663
2019-03-21 12:43 | 454 | 0.11675 | 0 | 0 | 0 | 4.8885 | 3.1459
2019-03-21 12:43 | 455 | 0.13471 | 0 | 0 | 0 | 4.7702 | 3.1658
2019-03-21 12:43 | 456 | 0.12557 | 0 | 0 | 0 | 4.6249 | 3.1478
2019-03-21 12:43 | 457 | 0.12619 | 0 | 0 | 0 | 4.5173 | 3.1316
2019-03-21 12:43 | 458 | 0.12802 | 0 | 0 | 0 | 4.3726 | 3.125
2019-03-21 12:43 | 459 | 0.10846 | 0 | 0 | 0 | 4.2662 | 3.1133
2019-03-21 12:43 | 460 | 0.1044 | 0 | 0 | 0 | 4.1523 | 3.1346
2019-03-21 12:43 | 461 | 0.10124 | 0 | 0 | 0 | 4.0563 | 2.9856
2019-03-21 12:43 | 462 | 0.11771 | 0 | 0 | 0 | 3.9459 | 2.9486
2019-03-21 12:43 | 463 | 0.093352 | 0 | 0 | 0 | 3.8301 | 2.8714
2019-03-21 12:43 | 464 | 0.087136 | 0 | 0 | 0 | 3.7556 | 2.88
2019-03-21 12:43 | 465 | 0.084221 | 0 | 0 | 0 | 3.6542 | 2.7928
2019-03-21 12:43 | 466 | 0.093972 | 0 | 0 | 0 | 3.5839 | 2.7318
2019-03-21 12:43 | 467 | 0.10212 | 0 | 0 | 0 | 3.467 | 2.7422
2019-03-21 12:43 | 468 | 0.081811 | 0 | 0 | 0 | 3.3861 | 2.7814
2019-03-21 12:43 | 469 | 0.08938 | 0 | 0 | 0 | 3.2972 | 2.5455
2019-03-21 12:43 | 470 | 0.097006 | 0 | 0 | 0 | 3.2066 | 2.6587
2019-03-21 12:43 | 471 | 0.10702 | 0 | 0 | 0 | 3.1036 | 2.6185
2019-03-21 12:43 | 472 | 0.077334 | 0 | 0 | 0 | 3.0016 | 2.5963
2019-03-21 12:43 | 473 | 0.07971 | 0 | 0 | 0 | 2.9417 | 2.4045
2019-03-21 12:43 | 474 | 0.09248 | 0 | 0 | 0 | 2.8392 | 2.4754
2019-03-21 12:43 | 475 | 0.087153 | 0 | 0 | 0 | 2.7607 | 2.4209
2019-03-21 12:43 | 476 | 0.092344 | 0 | 0 | 0 | 2.6639 | 2.3796
2019-03-21 12:43 | 477 | 0.058284 | 0 | 0 | 0 | 2.5849 | 2.2409
2019-03-21 12:43 | 478 | 0.052322 | 0 | 0 | 0 | 2.5409 | 2.2323
2019-03-21 12:43 | 479 | 0.058909 | 0 | 0 | 0 | 2.4765 | 2.188
2019-03-21 12:43 | 480 | 0.058874 | 0 | 0 | 0 | 2.4241 | 2.15
2019-03-21 12:43 | 481 | 0.06325 | 0 | 0 | 0 | 2.3577 | 2.1012
2019-03-21 12:43 | 482 | 0.067482 | 0 | 0 | 0 | 2.2972 | 1.9821
2019-03-21 12:43 | 483 | 0.088209 | 0 | 0 | 0 | 2.2196 | 1.9059
2019-03-21 12:43 | 484 | 0.073603 | 0 | 0 | 0 | 2.1293 | 1.8678
2019-03-21 12:43 | 485 | 0.0601 | 0 | 0 | 0 | 2.0731 | 1.7754
2019-03-21 12:43 | 486 | 0.045533 | 0 | 0 | 0 | 2.0084 | 1.7324
2019-03-21 12:43 | 487 | 0.063041 | 0 | 0 | 0 | 1.9742 | 1.6903
2019-03-21 12:43 | 488 | 0.073385 | 0 | 0 | 0 | 1.8842 | 1.6444
2019-03-21 12:43 | 489 | 0.04826 | 0 | 0 | 0 | 1.8352 | 1.6117
2019-03-21 12:43 | 490 | 0.07328 | 0 | 0 | 0 | 1.7774 | 1.5499
2019-03-21 12:43 | 491 | 0.049876 | 0 | 0 | 0 | 1.6991 | 1.5142
2019-03-21 12:43 | 492 | 0.046055 | 0 | 0 | 0 | 1.6742 | 1.4703
2019-03-21 12:43 | 493 | 0.033868 | 0 | 0 | 0 | 1.6083 | 1.4241
2019-03-21 12:43 | 494 | 0.028931 | 0 | 0 | 0 | 1.6041 | 1.3686
2019-03-21 12:43 | 495 | 0.047269 | 0 | 0 | 0 | 1.545 | 1.3327
2019-03-21 12:43 | 496 | 0.039972 | 0 | 0 | 0 | 1.515 | 1.2787
2019-03-21 12:43 | 497 | 0.045236 | 0 | 0 | 0 | 1.4628 | 1.2405
2019-03-21 12:43 | 498 | 0.0407 | 0 | 0 | 0 | 1.4263 | 1.2128
2019-03-21 12:43 | 499 | 0.057821 | 0 | 0 | 0 | 1.3778 | 1.1837
2019-03-21 12:43 | 500 | 0.017452 | 0 | 0 | 0 | 1.3228 | 1.1461
2019-03-21 12:43 | 501 | 0.040334 | 0 | 0 | 0 | 1.3287 | 1.1099
2019-03-21 12:43 | 502 | 0.058035 | 0 | 0 | 0 | 1.243 | 1.0752
2019-03-21 12:43 | 503 | 0.028153 | 0 | 0 | 0 | 1.2236 | 1.0395
2019-03-21 12:43 | 504 | 0.039764 | 0 | 0 | 0 | 1.1785 | 1.0038
2019-03-21 12:43 | 505 | 0.022477 | 0 | 0 | 0 | 1.1501 | 0.98651
2019-03-21 12:44 | 506 | 0.011505 | 0 | 0 | 0 | 1.1321 | 0.94605
2019-03-21 12:44 | 507 | 0.014439 | 0 | 0 | 0 | 1.1227 | 0.90342
2019-03-21 12:44 | 508 | 0.044366 | 0 | 0 | 0 | 1.0974 | 0.89276
2019-03-21 12:44 | 509 | 0.024147 | 0 | 0 | 0 | 1.0455 | 0.86868
2019-03-21 12:44 | 510 | 0.01611 | 0 | 0 | 0 | 1.047 | 0.82095
2019-03-21 12:44 | 511 | 0.023102 | 0 | 0 | 0 | 1.0095 | 0.77789
2019-03-21 12:44 | 512 | 0.01914 | 0 | 0 | 0 | 1.0026 | 0.7241
2019-03-21 12:44 | 513 | 0.035801 | 0 | 0 | 0 | 0.96729 | 0.70861
2019-03-21 12:44 | 514 | 0.01615 | 0 | 0 | 0 | 0.93921 | 0.69595
2019-03-21 12:44 | 515 | 0.010641 | 0 | 0 | 0 | 0.932 | 0.66337
2019-03-21 12:44 | 516 | 0.02098 | 0 | 0 | 0 | 0.9141 | 0.61863
2019-03-21 12:44 | 517 | 0.014545 | 0 | 0 | 0 | 0.89411 | 0.57262
2019-03-21 12:44 | 518 | -0.0056047 | 0 | 0 | 0 | 0.88782 | 0.515
2019-03-21 12:44 | 519 | 0.010973 | 0 | 0 | 0 | 0.89631 | 0.47274
2019-03-21 12:44 | 520 | 0.03379 | 0 | 0 | 0 | 0.86448 | 0.44747
2019-03-21 12:44 | 521 | 0.014814 | 0 | 0 | 0 | 0.83902 | 0.42611
2019-03-21 12:44 | 522 | -0.0014376 | 0 | 0 | 0 | 0.83498 | 0.39804
2019-03-21 12:44 | 523 | -0.0055133 | 0 | 0 | 0 | 0.8381 | 0.39427
2019-03-21 12:44 | 524 | 0.013997 | 0 | 0 | 0 | 0.83987 | 0.39462
2019-03-21 12:44 | 525 | 0.025013 | 0 | 0 | 0 | 0.81196 | 0.37845
2019-03-21 12:44 | 526 | 0.021609 | 0 | 0 | 0 | 0.79379 | 0.36536
2019-03-21 12:44 | 527 | 0.017402 | 0 | 0 | 0 | 0.76964 | 0.32702
2019-03-21 12:44 | 528 | 0.0066257 | 0 | 0 | 0 | 0.76061 | 0.31661
2019-03-21 12:44 | 529 | 0.0015384 | 0 | 0 | 0 | 0.75536 | 0.30508
2019-03-21 12:44 | 530 | -0.014463 | 0 | 0 | 0 | 0.75906 | 0.28911
2019-03-21 12:44 | 531 | 0.0083757 | 0 | 0 | 0 | 0.77567 | 0.29374
2019-03-21 12:44 | 532 | -0.0020878 | 0 | 0 | 0 | 0.74959 | 0.28587
2019-03-21 12:44 | 533 | -0.011586 | 0 | 0 | 0 | 0.78012 | 0.27667
2019-03-21 12:44 | 534 | -0.0080682 | 0 | 0 | 0 | 0.76958 | 0.26542
2019-03-21 12:44 | 535 | -0.010052 | 0 | 0 | 0 | 0.79657 | 0.25412
2019-03-21 12:44 | 536 | 0.012017 | 0 | 0 | 0 | 0.78434 | 0.29557

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062478 | 0.062678 | 0 | 0.2095 | 2019-03-21 09:44 | Group A | 0.1 | 0.22 | Do Nothing | 0.20944 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025146 | 0.025271 | 0 | 0.0261 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | 0.00046247 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002726 | 0.002739 | 0 | 0.0011 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | -3.729e-06 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.6721 | 183.6734 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 100.3003 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4138 | 209.9622 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 100.4016 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -16.237 | 1447 | 0 |  | 2019-03-21 09:27 | None | -10 | 1000 | Do Nothing | 123.5973 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2225 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0019542 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018286 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -152.0968 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23966 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4026 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0034751 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0048204 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0046504 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 6.0774e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011366 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.2778 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.315 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.548 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 99.38 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.9554 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 37.2756 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 68.8089 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 105.4506 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 65.7878 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0093134 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0416 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.024819 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.8273 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN