## Cone Summary
Date: | 2019-03-28
Time: | 15:32
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B4S-Test3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0413 |  | Start Test | 0 | `
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 7 | 2019-03-28 15:38
Sample Area | 1 | m^2 | Wetting | 145 | 2019-03-28 15:41
Cone Heat Flux | 75 | kW/m^2 | Dripping | 157 | 2019-03-28 15:41
Cone Orientation |  |  | Flame Out | 430 | 2019-03-28 15:51
Sample Grid |  |  | NO BOTTOM IGNITION
Sample Thickness | 106 | mm
Sample Initial Mass | 45.9 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 24.1882 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 3.5517 | kW/m2 | Average Heat Release Rate | 1.3924 | kW/m2
Total Heat Release | 0.52942 | MJ/m2 | Average Mass Loss Rate | 0.065697 | g/m2s
Peak Mass Loss Rate | 0.15239 | g/m2s | Effective Heat of Combustion | 21.3235 | MJ/kg
Total Mass Lost | 24.5692 | g | Average Specific Extinction Area | 308.0332 | m2/kg
Time to Ignition | 7 | s | Peak CO Yield | 0.19221 | kg/kg
Time to Peak Heat Release | 16 | s | Peak CO2 Yield | 8.8019 | kg/kg
Time to Flame Out | 430 | s | Average CO Yield | 0.031349 | kg/kg
Average Heat Release Rate | 1.2486 | kW/m2 | Average CO2 Yield | 2.1219 | kg/kg
Average Mass Loss Rate | 0.056807 | g/m2s
Effective Heat of Combustion | 21.5483 | MJ/kg
Dripping | 157 | s | Time to dripping
HRRD | 1.567 | kW | HRR at dripping time
THRD | 0.25311 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.6874 | kW | AHRR from Ignition to Dripping
Wetting | 145 | s | Time to wetting
HRRW | 1.5838 | kW | HRR at wetting time
THRW | 0.23408 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.6284 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-28 15:32 | 1553801550.2796 | -351 | 0.20944 | 0.000453 | 0 | 96.4803 | 95.8591 | 126.8809 | 21.2612 | 0.00188 | 0.001762 | -152.0184 | 0.23366 | 8.402 | 0.002818 | 0.004388 | 0.004271 | 5.2e-05 | 0.001251 | 0.29983 | 1.3334 | 2.5472 | 100.64 | 21.15 | 26.076 | 67.0553 | 89.604 | 64.2144 | 0.006511 | 1.0465 | 0.025328 | 24.2029 | 0.05778 | 0.13901 | -9e-06 | -0.000662 | -0.000662 | -8.671
2019-03-28 15:32 | 1553801551.2783 | -350 | 0.20944 | 0.000453 | 0 | 96.4803 | 95.8591 | 126.8809 | 21.2612 | 0.00188 | 0.001762 | -152.0184 | 0.23366 | 8.402 | 0.002818 | 0.004388 | 0.004271 | 5.2e-05 | 0.001251 | 0.29983 | 1.3334 | 2.5472 | 100.64 | 21.15 | 26.074 | 67.0553 | 89.604 | 64.2144 | 0.006511 | 1.0465 | 0.025328 | 24.2029 | 0.05778 | 0.13901 | -9e-06 | -0.000662 | -0.000662 | -8.672
2019-03-28 15:33 | 1553801552.2787 | -349 | 0.20943 | 0.000452 | 0 | 96.4827 | 95.8784 | 126.8809 | 21.2612 | 0.001848 | 0.00175 | -152.0184 | 0.23429 | 8.4048 | 0.0028 | 0.003899 | 0.003748 | 4.8e-05 | 0.001257 | 0.3115 | 1.3208 | 2.5474 | 100.64 | 21.15 | 26.069 | 66.285 | 89.1845 | 63.9317 | 0.006509 | 1.0474 | 0.025338 | 24.1928 | 0.056195 | 0.13938 | -8e-06 | -0.000611 | -0.000611 | -9.4261
2019-03-28 15:33 | 1553801553.2794 | -348 | 0.20943 | 0.000451 | 0 | 96.4264 | 96.1261 | 126.8809 | 21.2612 | 0.001935 | 0.001775 | -152.0184 | 0.234 | 8.404 | 0.002868 | 0.003886 | 0.003601 | 5.1e-05 | 0.001261 | 0.32156 | 1.3099 | 2.5473 | 100.64 | 21.15 | 26.066 | 68.3822 | 90.8153 | 64.5352 | 0.006509 | 1.0455 | 0.025316 | 24.2144 | 0.0279 | 0.13921 | -7e-06 | -0.000556 | -0.000556 | -10.338
2019-03-28 15:33 | 1553801554.2796 | -347 | 0.20944 | 0.000456 | 0 | 96.4741 | 95.9583 | 126.8809 | 21.2612 | 0.00185 | 0.001787 | -152.0184 | 0.23389 | 8.4052 | 0.002787 | 0.004077 | 0.003969 | 5.3e-05 | 0.001265 | 0.31924 | 1.297 | 2.5473 | 100.64 | 21.15 | 26.063 | 66.342 | 88.8594 | 64.835 | 0.006508 | 1.0446 | 0.025305 | 24.2251 | 0.047949 | 0.13914 | -1.3e-05 | -0.001007 | -0.001007 | -5.7044
2019-03-28 15:33 | 1553801555.2783 | -346 | 0.20944 | 0.000453 | 0 | 96.5432 | 95.9583 | 126.8809 | 21.2612 | 0.00189 | 0.00177 | -152.0184 | 0.2334 | 8.4034 | 0.00284 | 0.004157 | 0.003996 | 5.3e-05 | 0.001265 | 0.2844 | 1.278 | 2.5473 | 100.64 | 21.15 | 26.06 | 67.308 | 90.1377 | 64.419 | 0.006507 | 1.0458 | 0.02532 | 24.2102 | 0.054353 | 0.13885 | -8e-06 | -0.000612 | -0.000612 | -9.3672
2019-03-28 15:33 | 1553801556.2795 | -345 | 0.20943 | 0.000455 | 0 | 96.5277 | 96.1013 | 126.8809 | 21.2612 | 0.001905 | 0.001764 | -152.0184 | 0.23372 | 8.4053 | 0.002812 | 0.003976 | 0.003775 | 5e-05 | 0.001266 | 0.30038 | 1.2948 | 2.5474 | 100.64 | 21.15 | 26.058 | 67.6615 | 89.4797 | 64.2674 | 0.006507 | 1.0463 | 0.025326 | 24.2048 | 0.039596 | 0.13904 | -1.2e-05 | -0.000974 | -0.000974 | -5.8936
2019-03-28 15:33 | 1553801557.2789 | -344 | 0.20943 | 0.000453 | 0 | 96.4824 | 95.995 | 126.8809 | 21.2612 | 0.001891 | 0.001702 | -152.0184 | 0.23464 | 8.402 | 0.002811 | 0.004571 | 0.004383 | 5.4e-05 | 0.001265 | 0.32836 | 1.2718 | 2.5474 | 100.64 | 21.15 | 26.057 | 67.3338 | 89.4455 | 62.79 | 0.006506 | 1.0509 | 0.025382 | 24.1517 | 0.045303 | 0.13959 | -9e-06 | -0.000674 | -0.000674 | -8.5552
2019-03-28 15:33 | 1553801558.2789 | -343 | 0.20944 | 0.000457 | 0 | 96.5625 | 96.0289 | 126.8809 | 21.2612 | 0.001858 | 0.00177 | -152.0184 | 0.23476 | 8.4002 | 0.002806 | 0.004146 | 0.004029 | 5.2e-05 | 0.001266 | 0.32278 | 1.3273 | 2.5474 | 100.64 | 21.15 | 26.054 | 66.5402 | 89.3258 | 64.4127 | 0.006506 | 1.0459 | 0.02532 | 24.21 | 0.049563 | 0.13966 | -1.4e-05 | -0.001034 | -0.001034 | -5.5808
2019-03-28 15:33 | 1553801559.2797 | -342 | 0.20946 | 0.000456 | 0 | 96.561 | 96.0252 | 126.8809 | 21.2612 | 0.001838 | 0.001767 | -152.0184 | 0.23407 | 8.4012 | 0.002881 | 0.004154 | 0.003997 | 5.1e-05 | 0.001265 | 0.28423 | 1.2904 | 2.5474 | 100.64 | 21.15 | 26.053 | 66.0527 | 91.135 | 64.3487 | 0.006505 | 1.0461 | 0.025323 | 24.2077 | 0.049773 | 0.13925 | -1.3e-05 | -0.001037 | -0.001037 | -5.5461
2019-03-28 15:33 | 1553801560.2786 | -341 | 0.20949 | 0.000456 | 0 | 96.5589 | 95.8954 | 126.8809 | 21.2612 | 0.001899 | 0.001766 | -152.0184 | 0.23366 | 8.4065 | 0.002875 | 0.003836 | 0.003683 | 5.4e-05 | 0.001261 | 0.33539 | 1.3049 | 2.5472 | 100.64 | 21.15 | 26.054 | 67.5218 | 90.9912 | 64.3254 | 0.006506 | 1.0461 | 0.025324 | 24.2069 | 0.061676 | 0.13901 | -1.3e-05 | -0.001031 | -0.001031 | -5.5669
2019-03-28 15:33 | 1553801561.2787 | -340 | 0.20948 | 0.000458 | -1e-06 | 96.6158 | 96.2358 | 126.8131 | 21.2612 | 0.001871 | 0.001825 | -152.0019 | 0.2337 | 8.3999 | 0.00291 | 0.004134 | 0.003911 | 5.5e-05 | 0.001269 | 0.28888 | 1.2799 | 2.5472 | 100.64 | 21.15 | 26.054 | 66.8492 | 91.8372 | 65.7429 | 0.006506 | 1.0418 | 0.025264 | 24.2512 | 0.03525 | 0.13903 | -1.5e-05 | -0.001105 | -0.001105 | -5.1949
2019-03-28 15:33 | 1553801562.2777 | -339 | 0.20947 | 0.000453 | -1e-06 | 96.5411 | 95.9443 | 126.6511 | 21.2612 | 0.001892 | 0.001797 | -151.9962 | 0.23317 | 8.4 | 0.002884 | 0.003911 | 0.003779 | 5.7e-05 | 0.001271 | 0.32156 | 1.2928 | 2.5474 | 100.64 | 21.15 | 26.054 | 67.347 | 91.2015 | 65.0614 | 0.006506 | 1.0439 | 0.025273 | 24.2113 | 0.055459 | 0.13871 | -8e-06 | -0.000556 | -0.000556 | -10.3088
2019-03-28 15:33 | 1553801563.2791 | -338 | 0.20946 | 0.000454 | -1e-06 | 96.5119 | 96.0272 | 126.2728 | 21.2612 | 0.001934 | 0.001769 | -152.0214 | 0.2343 | 8.4005 | 0.002855 | 0.004488 | 0.004264 | 5.7e-05 | 0.001273 | 0.33035 | 1.3375 | 2.5474 | 100.64 | 21.15 | 26.054 | 68.3673 | 90.5113 | 64.3831 | 0.006506 | 1.046 | 0.025261 | 24.1509 | 0.045029 | 0.13939 | -9e-06 | -0.00065 | -0.00065 | -8.8539
2019-03-28 15:33 | 1553801564.2787 | -337 | 0.20948 | 0.000453 | 0 | 96.4773 | 96.1472 | 125.9075 | 21.2612 | 0.001897 | 0.001772 | -152.0289 | 0.23438 | 8.4036 | 0.002851 | 0.004546 | 0.004367 | 5.8e-05 | 0.001274 | 0.29839 | 1.4042 | 2.5475 | 100.64 | 21.15 | 26.054 | 67.4655 | 90.4183 | 64.4596 | 0.006506 | 1.0457 | 0.025221 | 24.1186 | 0.030663 | 0.13943 | -9e-06 | -0.000645 | -0.000645 | -8.9339
2019-03-28 15:33 | 1553801565.2791 | -336 | 0.20949 | 0.000451 | 0 | 96.5059 | 96.172 | 125.4735 | 21.2612 | 0.001844 | 0.001771 | -152.0217 | 0.23438 | 8.4037 | 0.002799 | 0.004288 | 0.004142 | 5.5e-05 | 0.001273 | 0.33445 | 1.3694 | 2.5474 | 100.64 | 21.15 | 26.056 | 66.2085 | 89.1643 | 64.4502 | 0.006506 | 1.0457 | 0.025178 | 24.0767 | 0.031007 | 0.13943 | -6e-06 | -0.000464 | -0.000464 | -12.4025
2019-03-28 15:33 | 1553801566.2792 | -335 | 0.20948 | 0.000451 | 0 | 96.5491 | 96.0238 | 125.6904 | 21.2612 | 0.001908 | 0.001765 | -152.0238 | 0.23338 | 8.405 | 0.002822 | 0.003718 | 0.003592 | 5.7e-05 | 0.001271 | 0.2943 | 1.36 | 2.5473 | 100.64 | 21.15 | 26.057 | 67.737 | 89.7049 | 64.3056 | 0.006506 | 1.0462 | 0.025205 | 24.0923 | 0.048796 | 0.13884 | -7e-06 | -0.00051 | -0.00051 | -11.2529
2019-03-28 15:33 | 1553801567.2788 | -334 | 0.20949 | 0.000457 | 0 | 96.5238 | 96.1241 | 125.7928 | 21.2612 | 0.001898 | 0.001749 | -152.0186 | 0.23381 | 8.401 | 0.002752 | 0.004167 | 0.003971 | 5.8e-05 | 0.001263 | 0.29142 | 1.3094 | 2.5474 | 100.64 | 21.15 | 26.059 | 67.5023 | 88.0145 | 63.9037 | 0.006507 | 1.0474 | 0.025231 | 24.0878 | 0.037119 | 0.13909 | -1.3e-05 | -0.000997 | -0.000997 | -5.762
2019-03-28 15:33 | 1553801568.2792 | -333 | 0.2095 | 0.000454 | 0 | 96.5679 | 95.9617 | 126.1723 | 21.2612 | 0.001867 | 0.001804 | -152.0012 | 0.23406 | 8.403 | 0.002737 | 0.004042 | 0.003835 | 5.5e-05 | 0.001261 | 0.31581 | 1.3235 | 2.5474 | 100.64 | 21.15 | 26.065 | 66.754 | 87.6513 | 65.223 | 0.006508 | 1.0434 | 0.025219 | 24.1713 | 0.056327 | 0.13924 | -1e-05 | -0.000785 | -0.000785 | -7.3233
2019-03-28 15:33 | 1553801569.2778 | -332 | 0.20948 | 0.000451 | 0 | 96.6063 | 96.0999 | 126.6718 | 21.2612 | 0.001904 | 0.001754 | -151.9817 | 0.2347 | 8.4052 | 0.002842 | 0.003646 | 0.003525 | 5.5e-05 | 0.001258 | 0.3147 | 1.2874 | 2.5473 | 100.64 | 21.15 | 26.069 | 67.6379 | 90.201 | 64.0246 | 0.006509 | 1.0471 | 0.025314 | 24.1761 | 0.047006 | 0.13962 | -8e-06 | -0.000605 | -0.000605 | -9.5304
2019-03-28 15:33 | 1553801570.2789 | -331 | 0.2095 | 0.000452 | 0 | 96.6412 | 96.0853 | 126.6112 | 21.2612 | 0.001935 | 0.001767 | -151.9643 | 0.23419 | 8.4021 | 0.002748 | 0.003997 | 0.003816 | 5.6e-05 | 0.001257 | 0.30127 | 1.2582 | 2.5472 | 100.64 | 21.15 | 26.074 | 68.3813 | 87.9153 | 64.3533 | 0.006511 | 1.046 | 0.025296 | 24.1821 | 0.051592 | 0.13932 | -8e-06 | -0.000614 | -0.000614 | -9.3652
2019-03-28 15:33 | 1553801571.2793 | -330 | 0.20949 | 0.000453 | 0 | 96.7162 | 96.2192 | 126.8112 | 21.2612 | 0.001857 | 0.001747 | -151.9785 | 0.2337 | 8.4058 | 0.002745 | 0.003565 | 0.003334 | 5.6e-05 | 0.001259 | 0.2844 | 1.2918 | 2.5472 | 100.64 | 21.15 | 26.079 | 66.5113 | 87.8489 | 63.8551 | 0.006512 | 1.0476 | 0.025334 | 24.1834 | 0.046085 | 0.13903 | -9e-06 | -0.000739 | -0.000739 | -7.7689
2019-03-28 15:33 | 1553801572.2783 | -329 | 0.20948 | 0.000452 | 0 | 96.6811 | 96.033 | 127.0713 | 21.2612 | 0.001902 | 0.001761 | -152.0012 | 0.23381 | 8.4039 | 0.002791 | 0.004547 | 0.00439 | 5.2e-05 | 0.001259 | 0.31614 | 1.2864 | 2.5474 | 100.64 | 21.15 | 26.081 | 67.5951 | 88.9524 | 64.2065 | 0.006512 | 1.0465 | 0.025347 | 24.2208 | 0.060157 | 0.13909 | -9e-06 | -0.000684 | -0.000684 | -8.3988
2019-03-28 15:33 | 1553801573.279 | -328 | 0.20951 | 0.000457 | 0 | 96.7523 | 96.0802 | 127.2011 | 21.2612 | 0.001931 | 0.001727 | -152.0099 | 0.23334 | 8.401 | 0.002826 | 0.004477 | 0.004313 | 5.5e-05 | 0.001262 | 0.30287 | 1.3291 | 2.5475 | 100.64 | 21.15 | 26.08 | 68.294 | 89.808 | 63.3798 | 0.006512 | 1.0491 | 0.025391 | 24.2034 | 0.062343 | 0.13881 | -1.4e-05 | -0.001075 | -0.001075 | -5.3332
2019-03-28 15:33 | 1553801574.278 | -327 | 0.20951 | 0.000455 | 0 | 96.698 | 96.1506 | 127.2844 | 21.2612 | 0.001896 | 0.001789 | -151.9862 | 0.23451 | 8.4018 | 0.002809 | 0.00475 | 0.004589 | 5.6e-05 | 0.001264 | 0.30663 | 1.3133 | 2.5474 | 100.64 | 21.15 | 26.078 | 67.446 | 89.4002 | 64.8858 | 0.006512 | 1.0444 | 0.025343 | 24.2655 | 0.050786 | 0.13951 | -1.2e-05 | -0.00088 | -0.00088 | -6.546
2019-03-28 15:33 | 1553801575.2771 | -326 | 0.20948 | 0.000452 | 0 | 96.7543 | 96.2953 | 127.3142 | 21.2612 | 0.001892 | 0.001784 | -151.992 | 0.23466 | 8.4055 | 0.0028 | 0.00443 | 0.004187 | 5.6e-05 | 0.001264 | 0.31293 | 1.3092 | 2.5474 | 100.64 | 21.15 | 26.076 | 67.3532 | 89.1838 | 64.7484 | 0.006511 | 1.0448 | 0.025351 | 24.2634 | 0.042539 | 0.1396 | -8e-06 | -0.000591 | -0.000591 | -9.7598
2019-03-28 15:33 | 1553801576.2785 | -325 | 0.20946 | 0.000455 | 0 | 96.6861 | 96.2372 | 126.9143 | 21.2612 | 0.001865 | 0.001802 | -151.9775 | 0.23415 | 8.4033 | 0.002786 | 0.004176 | 0.004084 | 5.8e-05 | 0.001267 | 0.32045 | 1.3343 | 2.5474 | 100.64 | 21.15 | 26.066 | 66.7088 | 88.8532 | 65.1879 | 0.006509 | 1.0435 | 0.025295 | 24.241 | 0.041629 | 0.1393 | -1.2e-05 | -0.00094 | -0.00094 | -6.1226
2019-03-28 15:33 | 1553801577.2781 | -324 | 0.20948 | 0.000454 | 0 | 96.7588 | 96.4424 | 126.8124 | 21.2612 | 0.001829 | 0.001788 | -151.9771 | 0.23405 | 8.3999 | 0.002696 | 0.0045 | 0.004272 | 5.3e-05 | 0.001261 | 0.29551 | 1.3172 | 2.5473 | 100.64 | 21.16 | 26.06 | 65.8482 | 86.6785 | 64.8554 | 0.006511 | 1.0445 | 0.025297 | 24.2193 | 0.029299 | 0.13923 | -1.1e-05 | -0.000835 | -0.000835 | -6.8894
2019-03-28 15:33 | 1553801578.2782 | -323 | 0.20947 | 0.000461 | 0 | 96.8428 | 96.4145 | 126.6725 | 21.2612 | 0.001864 | 0.001773 | -151.9814 | 0.23381 | 8.4024 | 0.00277 | 0.004133 | 0.004025 | 5.6e-05 | 0.001256 | 0.31614 | 1.3354 | 2.5474 | 100.64 | 21.16 | 26.045 | 66.6869 | 88.4449 | 64.4822 | 0.006507 | 1.0457 | 0.025297 | 24.1926 | 0.039644 | 0.13909 | -1.8e-05 | -0.001394 | -0.001394 | -4.1204
2019-03-28 15:33 | 1553801579.2785 | -322 | 0.20947 | 0.000458 | 0 | 96.9378 | 96.3568 | 126.5657 | 21.2612 | 0.001919 | 0.001754 | -151.9486 | 0.23433 | 8.4059 | 0.002767 | 0.004028 | 0.003946 | 5.6e-05 | 0.001255 | 0.28711 | 1.2934 | 2.5474 | 100.64 | 21.16 | 26.037 | 67.9961 | 88.39 | 64.0433 | 0.006505 | 1.047 | 0.025303 | 24.1667 | 0.053775 | 0.1394 | -1.5e-05 | -0.001129 | -0.001129 | -5.0995
2019-03-28 15:33 | 1553801580.2782 | -321 | 0.20947 | 0.000463 | 0 | 96.9983 | 96.5008 | 126.9105 | 21.2612 | 0.001914 | 0.001745 | -151.9468 | 0.2348 | 8.4025 | 0.002766 | 0.004433 | 0.004109 | 5.9e-05 | 0.001251 | 0.30702 | 1.2937 | 2.5473 | 100.64 | 21.16 | 26.028 | 67.8892 | 88.3517 | 63.813 | 0.006503 | 1.0477 | 0.025346 | 24.1913 | 0.045991 | 0.13968 | -2.2e-05 | -0.001692 | -0.001692 | -3.4099
2019-03-28 15:33 | 1553801581.2777 | -320 | 0.20946 | 0.00046 | 0 | 97.0352 | 96.5185 | 126.8645 | 21.2612 | 0.001912 | 0.001754 | -151.9601 | 0.2336 | 8.4026 | 0.002755 | 0.003965 | 0.003705 | 5.5e-05 | 0.001251 | 0.32549 | 1.2812 | 2.5474 | 100.64 | 21.16 | 26.009 | 67.8339 | 88.0855 | 64.0214 | 0.006498 | 1.0471 | 0.025333 | 24.1944 | 0.047759 | 0.13897 | -1.8e-05 | -0.001385 | -0.001385 | -4.1428
2019-03-28 15:33 | 1553801582.2774 | -319 | 0.20947 | 0.000462 | 0 | 97.012 | 96.6099 | 126.5307 | 21.2612 | 0.001898 | 0.001785 | -151.9467 | 0.23439 | 8.4015 | 0.002795 | 0.003944 | 0.003812 | 5.2e-05 | 0.001253 | 0.32117 | 1.2783 | 2.5474 | 100.64 | 21.15 | 26 | 67.485 | 89.0641 | 64.7873 | 0.006492 | 1.0447 | 0.025271 | 24.19 | 0.037152 | 0.13944 | -2.1e-05 | -0.001581 | -0.001581 | -3.6425
2019-03-28 15:33 | 1553801583.2787 | -318 | 0.20946 | 0.000463 | 0 | 97.1049 | 96.562 | 126.8316 | 21.2612 | 0.001845 | 0.001726 | -151.9305 | 0.23328 | 8.4039 | 0.002749 | 0.003943 | 0.003731 | 5.2e-05 | 0.001258 | 0.29717 | 1.2948 | 2.5474 | 100.64 | 21.16 | 25.984 | 66.2297 | 87.9512 | 63.3461 | 0.006492 | 1.0492 | 0.025356 | 24.167 | 0.050152 | 0.13878 | -2.1e-05 | -0.001623 | -0.001623 | -3.5304
2019-03-28 15:33 | 1553801584.2787 | -317 | 0.20947 | 0.000462 | 0 | 97.1749 | 96.6812 | 126.9849 | 21.2612 | 0.001858 | 0.001768 | -151.9342 | 0.23368 | 8.4046 | 0.002754 | 0.004586 | 0.004447 | 5.4e-05 | 0.001258 | 0.33395 | 1.2826 | 2.5474 | 100.64 | 21.16 | 25.98 | 66.5301 | 88.0754 | 64.38 | 0.006491 | 1.046 | 0.025332 | 24.2188 | 0.045558 | 0.13901 | -2e-05 | -0.001546 | -0.001546 | -3.7138
2019-03-28 15:33 | 1553801585.2783 | -316 | 0.2095 | 0.000463 | 0 | 97.0635 | 96.725 | 127.3544 | 21.2612 | 0.001872 | 0.001752 | -151.9226 | 0.2342 | 8.4008 | 0.002702 | 0.004331 | 0.004159 | 5.8e-05 | 0.001258 | 0.29518 | 1.3055 | 2.5475 | 100.64 | 21.16 | 25.976 | 66.8649 | 86.826 | 63.9747 | 0.00649 | 1.0472 | 0.025384 | 24.2394 | 0.031245 | 0.13932 | -2.3e-05 | -0.001743 | -0.001743 | -3.302
2019-03-28 15:33 | 1553801586.2786 | -315 | 0.2095 | 0.000461 | 0 | 97.1168 | 96.6534 | 127.7886 | 21.2612 | 0.001867 | 0.00179 | -151.9398 | 0.23459 | 8.4033 | 0.002709 | 0.004608 | 0.004375 | 5.5e-05 | 0.001253 | 0.28296 | 1.2944 | 2.5472 | 100.64 | 21.16 | 25.969 | 66.755 | 86.9885 | 64.8889 | 0.006488 | 1.0444 | 0.025393 | 24.3136 | 0.042788 | 0.13956 | -2e-05 | -0.001499 | -0.001499 | -3.8447
2019-03-28 15:33 | 1553801587.2775 | -314 | 0.20947 | 0.000461 | 0 | 97.1853 | 96.5881 | 127.8189 | 21.2612 | 0.001891 | 0.00176 | -151.9445 | 0.23368 | 8.4035 | 0.002656 | 0.004269 | 0.004055 | 5.5e-05 | 0.001247 | 0.31078 | 1.3504 | 2.5474 | 100.64 | 21.16 | 25.965 | 67.3237 | 85.7229 | 64.1723 | 0.006487 | 1.0466 | 0.025423 | 24.2907 | 0.055134 | 0.13902 | -2e-05 | -0.00154 | -0.00154 | -3.7278
2019-03-28 15:33 | 1553801588.2778 | -313 | 0.20948 | 0.00046 | 0 | 97.2935 | 96.8759 | 127.6363 | 21.2612 | 0.001876 | 0.001759 | -151.9384 | 0.23392 | 8.4028 | 0.002699 | 0.0046 | 0.004406 | 5.5e-05 | 0.001242 | 0.29662 | 1.2854 | 2.5474 | 100.64 | 21.16 | 25.962 | 66.9639 | 86.744 | 64.1473 | 0.006487 | 1.0467 | 0.025406 | 24.2724 | 0.038472 | 0.13916 | -1.8e-05 | -0.001397 | -0.001397 | -4.1146
2019-03-28 15:33 | 1553801589.2775 | -312 | 0.20947 | 0.000464 | -1e-06 | 97.3045 | 96.8623 | 127.3865 | 21.2612 | 0.001867 | 0.001755 | -151.972 | 0.23423 | 8.4033 | 0.002611 | 0.004097 | 0.003916 | 5.5e-05 | 0.001239 | 0.32278 | 1.3225 | 2.5474 | 100.64 | 21.16 | 25.955 | 66.7463 | 84.6227 | 64.0628 | 0.006485 | 1.047 | 0.025384 | 24.2456 | 0.04074 | 0.13935 | -2.2e-05 | -0.001624 | -0.001624 | -3.5438
2019-03-28 15:33 | 1553801590.2779 | -311 | 0.20948 | 0.000457 | -1e-06 | 97.3503 | 97.005 | 127.0755 | 21.2612 | 0.001888 | 0.001722 | -151.9602 | 0.23408 | 8.4034 | 0.002671 | 0.003741 | 0.003587 | 5.8e-05 | 0.001241 | 0.29192 | 1.3196 | 2.5475 | 100.64 | 21.16 | 25.951 | 67.251 | 86.0828 | 63.257 | 0.006484 | 1.0495 | 0.025383 | 24.1871 | 0.031789 | 0.13925 | -1.3e-05 | -0.000973 | -0.000973 | -5.9124
2019-03-28 15:33 | 1553801591.277 | -310 | 0.20944 | 0.000457 | 0 | 97.4647 | 96.9764 | 126.8697 | 21.2612 | 0.001845 | 0.001759 | -151.9446 | 0.23401 | 8.4006 | 0.002637 | 0.003663 | 0.003465 | 5.3e-05 | 0.001243 | 0.31868 | 1.3384 | 2.5474 | 100.64 | 21.16 | 25.943 | 66.2134 | 85.25 | 64.1449 | 0.006482 | 1.0467 | 0.025329 | 24.1993 | 0.044924 | 0.13922 | -1.4e-05 | -0.001026 | -0.001026 | -5.6044
2019-03-28 15:33 | 1553801592.2766 | -309 | 0.20946 | 0.000459 | 0 | 97.4793 | 97.1738 | 127.1286 | 21.2612 | 0.001892 | 0.001794 | -151.9494 | 0.23373 | 8.3997 | 0.002633 | 0.003931 | 0.003762 | 5.5e-05 | 0.001245 | 0.31293 | 1.3601 | 2.5474 | 100.64 | 21.16 | 25.939 | 67.3626 | 85.1588 | 64.9967 | 0.006481 | 1.0441 | 0.025323 | 24.2546 | 0.028076 | 0.13905 | -1.6e-05 | -0.001217 | -0.001217 | -4.7197
2019-03-28 15:33 | 1553801593.277 | -308 | 0.20949 | 0.00046 | 0 | 97.5696 | 97.1884 | 126.8737 | 21.2612 | 0.001888 | 0.001833 | -151.9569 | 0.2336 | 8.4046 | 0.002556 | 0.004811 | 0.004607 | 5.5e-05 | 0.001251 | 0.30773 | 1.3324 | 2.5475 | 100.64 | 21.16 | 25.936 | 67.2621 | 83.3088 | 65.9255 | 0.00648 | 1.0412 | 0.025263 | 24.2635 | 0.035009 | 0.13897 | -1.7e-05 | -0.001311 | -0.001311 | -4.3787
2019-03-28 15:33 | 1553801594.2767 | -307 | 0.20948 | 0.000461 | 0 | 97.6002 | 97.2445 | 127.0215 | 21.2612 | 0.001928 | 0.001764 | -151.976 | 0.23381 | 8.4029 | 0.002585 | 0.004011 | 0.003822 | 5.6e-05 | 0.001242 | 0.31852 | 1.364 | 2.5475 | 100.64 | 21.16 | 25.928 | 68.2051 | 83.9939 | 64.2729 | 0.006478 | 1.0463 | 0.02534 | 24.2184 | 0.032663 | 0.13909 | -1.9e-05 | -0.001382 | -0.001382 | -4.1563
2019-03-28 15:33 | 1553801595.2766 | -306 | 0.2094 | 0.000459 | 0 | 97.7003 | 97.2659 | 126.9182 | 21.2612 | 0.001938 | 0.001746 | -151.9868 | 0.23437 | 8.4052 | 0.002701 | 0.003944 | 0.003805 | 5.6e-05 | 0.00124 | 0.33467 | 1.3129 | 2.5474 | 100.64 | 21.16 | 25.924 | 68.4601 | 86.794 | 63.8481 | 0.006477 | 1.0476 | 0.025345 | 24.1933 | 0.039862 | 0.13943 | -1.7e-05 | -0.001238 | -0.001238 | -4.65
2019-03-28 15:33 | 1553801596.2767 | -305 | 0.20944 | 0.000459 | -1e-06 | 97.7673 | 97.4157 | 126.8044 | 21.2612 | 0.001884 | 0.001744 | -152.0053 | 0.23364 | 8.4007 | 0.002619 | 0.003771 | 0.003502 | 5.6e-05 | 0.001238 | 0.30596 | 1.3331 | 2.5473 | 100.64 | 21.16 | 25.914 | 67.1474 | 84.8161 | 63.7966 | 0.006475 | 1.0478 | 0.025336 | 24.1806 | 0.032228 | 0.13899 | -1.6e-05 | -0.001153 | -0.001153 | -4.9791
2019-03-28 15:33 | 1553801597.2768 | -304 | 0.20943 | 0.000458 | 0 | 97.921 | 97.6803 | 126.8549 | 21.2612 | 0.001904 | 0.001793 | -152.0141 | 0.23425 | 8.4034 | 0.002624 | 0.004191 | 0.003991 | 5.4e-05 | 0.001247 | 0.33052 | 1.3079 | 2.5475 | 100.64 | 21.16 | 25.91 | 67.6504 | 84.9402 | 64.981 | 0.006474 | 1.0441 | 0.025297 | 24.2279 | 0.022012 | 0.13936 | -1.5e-05 | -0.001116 | -0.001116 | -5.1551
2019-03-28 15:33 | 1553801598.2782 | -303 | 0.20945 | 0.000454 | -1e-06 | 98.0003 | 97.6056 | 127.2732 | 21.2612 | 0.001855 | 0.001775 | -152.002 | 0.23461 | 8.3994 | 0.002625 | 0.003732 | 0.003506 | 5.5e-05 | 0.001251 | 0.33826 | 1.3583 | 2.5476 | 100.64 | 21.16 | 25.907 | 66.4668 | 84.959 | 64.5407 | 0.006473 | 1.0455 | 0.025355 | 24.252 | 0.036092 | 0.13957 | -1e-05 | -0.000732 | -0.000732 | -7.8786
2019-03-28 15:33 | 1553801599.2768 | -302 | 0.20943 | 0.000455 | 0 | 98.0702 | 97.715 | 127.3285 | 21.2612 | 0.00188 | 0.001756 | -151.9913 | 0.2342 | 8.4059 | 0.002652 | 0.003465 | 0.003351 | 5.5e-05 | 0.001246 | 0.30182 | 1.3174 | 2.5474 | 100.64 | 21.16 | 25.902 | 67.0637 | 85.613 | 64.0871 | 0.006472 | 1.0469 | 0.025377 | 24.241 | 0.03246 | 0.13933 | -1.2e-05 | -0.00089 | -0.00089 | -6.4672
2019-03-28 15:33 | 1553801600.2779 | -301 | 0.20941 | 0.000456 | 0 | 98.1977 | 97.8139 | 127.4492 | 21.2612 | 0.001847 | 0.001756 | -152.0041 | 0.23366 | 8.4044 | 0.002667 | 0.003825 | 0.003556 | 5.4e-05 | 0.001245 | 0.30934 | 1.2896 | 2.5475 | 100.64 | 21.16 | 25.902 | 66.268 | 85.9844 | 64.0864 | 0.006472 | 1.0469 | 0.025389 | 24.2524 | 0.035034 | 0.13901 | -1.3e-05 | -0.001016 | -0.001016 | -5.6493
2019-03-28 15:33 | 1553801601.2769 | -300 | 0.20945 | 0.000456 | 0 | 98.2936 | 97.9919 | 127.6114 | 21.2612 | 0.001818 | 0.001704 | -151.9981 | 0.23393 | 8.4041 | 0.002667 | 0.003769 | 0.003619 | 5.5e-05 | 0.001245 | 0.31653 | 1.3046 | 2.5474 | 100.64 | 21.16 | 25.904 | 65.5774 | 85.9837 | 62.8391 | 0.006472 | 1.0508 | 0.025453 | 24.2229 | 0.027502 | 0.13917 | -1.3e-05 | -0.000946 | -0.000946 | -6.0778
2019-03-28 15:33 | 1553801602.277 | -299 | 0.20947 | 0.000459 | 0 | 98.3446 | 98.106 | 127.5823 | 21.2612 | 0.00189 | 0.001715 | -151.9929 | 0.2341 | 8.3996 | 0.002526 | 0.004015 | 0.003879 | 5.4e-05 | 0.001238 | 0.27671 | 1.3281 | 2.5475 | 100.64 | 21.16 | 25.905 | 67.315 | 82.5779 | 63.0899 | 0.006472 | 1.05 | 0.02544 | 24.2292 | 0.021723 | 0.13926 | -1.7e-05 | -0.00127 | -0.00127 | -4.5281
2019-03-28 15:33 | 1553801603.2773 | -298 | 0.20947 | 0.000449 | 0 | 98.4018 | 98.1295 | 127.4666 | 21.2612 | 0.001891 | 0.001761 | -151.9924 | 0.23359 | 8.4026 | 0.002538 | 0.003815 | 0.003514 | 5.6e-05 | 0.001236 | 0.3115 | 1.356 | 2.5475 | 100.64 | 21.16 | 25.906 | 67.3261 | 82.8797 | 64.2096 | 0.006473 | 1.0465 | 0.025386 | 24.2585 | 0.024786 | 0.13896 | -3e-06 | -0.000229 | -0.000229 | -25.0425
2019-03-28 15:33 | 1553801604.277 | -297 | 0.20947 | 0.000454 | 0 | 98.4572 | 98.3479 | 127.23 | 21.2612 | 0.00184 | 0.001778 | -151.9773 | 0.23388 | 8.4028 | 0.002573 | 0.004353 | 0.004175 | 5.2e-05 | 0.001233 | 0.31868 | 1.3335 | 2.5475 | 100.64 | 21.16 | 25.905 | 66.0893 | 83.7075 | 64.6158 | 0.006472 | 1.0452 | 0.025348 | 24.2506 | 0.009931 | 0.13913 | -1e-05 | -0.000762 | -0.000762 | -7.5394
2019-03-28 15:33 | 1553801605.2771 | -296 | 0.20939 | 0.000456 | 0 | 98.5766 | 98.407 | 127.3523 | 21.2612 | 0.001874 | 0.001756 | -151.9672 | 0.23379 | 8.4027 | 0.002573 | 0.003815 | 0.003669 | 5.5e-05 | 0.001237 | 0.28816 | 1.3172 | 2.5475 | 100.64 | 21.16 | 25.905 | 66.9093 | 83.7089 | 64.0794 | 0.006472 | 1.0469 | 0.02538 | 24.243 | 0.015401 | 0.13908 | -1.2e-05 | -0.000921 | -0.000921 | -6.2362
2019-03-28 15:33 | 1553801606.2764 | -295 | 0.20951 | 0.000456 | 1e-06 | 98.7237 | 98.5249 | 127.4138 | 21.2612 | 0.001878 | 0.001715 | -151.9417 | 0.23325 | 8.4046 | 0.002612 | 0.004017 | 0.0038 | 5.3e-05 | 0.00124 | 0.29374 | 1.326 | 2.5474 | 100.64 | 21.16 | 25.903 | 67.0038 | 84.6617 | 63.083 | 0.006472 | 1.05 | 0.025424 | 24.213 | 0.018033 | 0.13876 | -1.3e-05 | -0.001053 | -0.001053 | -5.4417
2019-03-28 15:33 | 1553801607.2771 | -294 | 0.20952 | 0.000454 | 0 | 98.7282 | 98.3822 | 127.198 | 21.2612 | 0.001942 | 0.00178 | -151.91 | 0.23379 | 8.4061 | 0.002565 | 0.004074 | 0.003927 | 5.3e-05 | 0.001236 | 0.2896 | 1.3126 | 2.5475 | 100.64 | 21.16 | 25.897 | 68.5647 | 83.5163 | 64.6579 | 0.00647 | 1.0451 | 0.025343 | 24.249 | 0.031401 | 0.13908 | -1e-05 | -0.00074 | -0.00074 | -7.765
2019-03-28 15:33 | 1553801608.2771 | -293 | 0.20951 | 0.000455 | 0 | 98.8712 | 98.6316 | 126.7164 | 21.2612 | 0.001871 | 0.001717 | -151.9095 | 0.23309 | 8.4014 | 0.002567 | 0.003254 | 0.003127 | 5.2e-05 | 0.001236 | 0.29767 | 1.3053 | 2.5474 | 100.64 | 21.16 | 25.892 | 66.8555 | 83.5834 | 63.146 | 0.006469 | 1.0498 | 0.025352 | 24.1489 | 0.021702 | 0.13867 | -1.2e-05 | -0.000889 | -0.000889 | -6.4389
2019-03-28 15:33 | 1553801609.2775 | -292 | 0.20949 | 0.000453 | 0 | 98.9298 | 98.7372 | 126.5948 | 21.2612 | 0.001819 | 0.001707 | -151.9141 | 0.23297 | 8.4035 | 0.002596 | 0.003968 | 0.003795 | 5e-05 | 0.001231 | 0.29391 | 1.3209 | 2.5474 | 100.64 | 21.16 | 25.887 | 65.6053 | 84.2622 | 62.9125 | 0.006468 | 1.0505 | 0.025348 | 24.1289 | 0.017432 | 0.1386 | -9e-06 | -0.000687 | -0.000687 | -8.3289
2019-03-28 15:33 | 1553801610.2773 | -291 | 0.20952 | 0.000457 | 0 | 98.9915 | 98.8745 | 126.4882 | 21.2612 | 0.001823 | 0.001752 | -151.9049 | 0.2333 | 8.4035 | 0.00266 | 0.003293 | 0.003142 | 5.2e-05 | 0.001226 | 0.31509 | 1.3097 | 2.5474 | 100.64 | 21.16 | 25.882 | 65.6953 | 85.8199 | 63.9755 | 0.006467 | 1.0472 | 0.025298 | 24.1569 | 0.01058 | 0.13879 | -1.5e-05 | -0.001162 | -0.001162 | -4.9312
2019-03-28 15:34 | 1553801611.2766 | -290 | 0.20953 | 0.000453 | 0 | 99.0678 | 98.99 | 126.4866 | 21.2612 | 0.00186 | 0.001777 | -151.9139 | 0.23275 | 8.4027 | 0.002622 | 0.003992 | 0.003846 | 5.4e-05 | 0.001226 | 0.32068 | 1.3227 | 2.5473 | 100.64 | 21.16 | 25.872 | 66.5784 | 84.8919 | 64.597 | 0.006464 | 1.0453 | 0.025274 | 24.179 | 0.007024 | 0.13846 | -9e-06 | -0.000666 | -0.000666 | -8.585
2019-03-28 15:34 | 1553801612.2762 | -289 | 0.20949 | 0.000454 | 0 | 99.0469 | 98.9326 | 126.3136 | 21.2612 | 0.001897 | 0.001736 | -151.9034 | 0.23348 | 8.401 | 0.002701 | 0.003649 | 0.003461 | 5.2e-05 | 0.001223 | 0.32587 | 1.3603 | 2.5475 | 100.64 | 21.16 | 25.868 | 67.4742 | 86.7986 | 63.6039 | 0.006463 | 1.0484 | 0.025294 | 24.1269 | 0.010331 | 0.1389 | -1.1e-05 | -0.000806 | -0.000806 | -7.1193
2019-03-28 15:34 | 1553801613.2773 | -288 | 0.2095 | 0.000455 | 0 | 99.1166 | 98.8045 | 126.4509 | 21.2612 | 0.0019 | 0.001779 | -151.883 | 0.23428 | 8.4027 | 0.002778 | 0.00422 | 0.004045 | 5.4e-05 | 0.001226 | 0.29662 | 1.3351 | 2.5475 | 100.64 | 21.16 | 25.864 | 67.5367 | 88.6377 | 64.6454 | 0.006462 | 1.0451 | 0.025269 | 24.1773 | 0.02821 | 0.13937 | -1.1e-05 | -0.000871 | -0.000871 | -6.606
2019-03-28 15:34 | 1553801614.2777 | -287 | 0.20947 | 0.000458 | 0 | 99.2643 | 99.1001 | 126.526 | 21.2612 | 0.001894 | 0.001727 | -151.8929 | 0.23429 | 8.4025 | 0.002701 | 0.00437 | 0.004207 | 5.3e-05 | 0.001228 | 0.34252 | 1.3166 | 2.5475 | 100.64 | 21.16 | 25.863 | 67.4071 | 86.7993 | 63.3805 | 0.006462 | 1.0491 | 0.025324 | 24.1391 | 0.014815 | 0.13938 | -1.6e-05 | -0.001221 | -0.001221 | -4.7144
2019-03-28 15:34 | 1553801615.2777 | -286 | 0.20956 | 0.000455 | 0 | 99.3799 | 99.3229 | 126.6216 | 21.2612 | 0.00187 | 0.001781 | -151.9184 | 0.23392 | 8.4039 | 0.002677 | 0.003766 | 0.003538 | 5.8e-05 | 0.001229 | 0.31509 | 1.3333 | 2.5475 | 100.64 | 21.17 | 25.865 | 66.8148 | 86.2147 | 64.6735 | 0.006466 | 1.0451 | 0.025285 | 24.1946 | 0.005131 | 0.13916 | -1.2e-05 | -0.000904 | -0.000904 | -6.3599
2019-03-28 15:34 | 1553801616.2767 | -285 | 0.20944 | 0.000453 | 0 | 99.379 | 99.2451 | 126.4594 | 21.2612 | 0.001859 | 0.001779 | -151.9266 | 0.23368 | 8.4029 | 0.002753 | 0.004524 | 0.004394 | 5.6e-05 | 0.001225 | 0.32029 | 1.3525 | 2.5474 | 100.64 | 21.16 | 25.87 | 66.552 | 88.0354 | 64.6374 | 0.006464 | 1.0452 | 0.02527 | 24.1778 | 0.012059 | 0.13902 | -9e-06 | -0.000672 | -0.000672 | -8.5423
2019-03-28 15:34 | 1553801617.2774 | -284 | 0.2096 | 0.000454 | 0 | 99.4809 | 99.4061 | 126.6216 | 21.2612 | 0.001899 | 0.001749 | -151.948 | 0.23429 | 8.4033 | 0.002714 | 0.004004 | 0.003804 | 5.8e-05 | 0.001228 | 0.32516 | 1.3242 | 2.5473 | 100.64 | 21.16 | 25.873 | 67.5304 | 87.1165 | 63.9138 | 0.006464 | 1.0474 | 0.025313 | 24.1674 | 0.006722 | 0.13938 | -1.1e-05 | -0.000824 | -0.000824 | -6.9865
2019-03-28 15:34 | 1553801618.2765 | -283 | 0.20944 | 0.000452 | 0 | 99.463 | 99.4388 | 126.7319 | 21.2612 | 0.001911 | 0.001785 | -151.9548 | 0.23358 | 8.4025 | 0.002687 | 0.003814 | 0.003634 | 5.4e-05 | 0.001228 | 0.3314 | 1.338 | 2.5475 | 100.64 | 21.16 | 25.877 | 67.7995 | 86.4678 | 64.7703 | 0.006465 | 1.0448 | 0.025292 | 24.2086 | 0.00218 | 0.13896 | -9e-06 | -0.000685 | -0.000685 | -8.3787
2019-03-28 15:34 | 1553801619.2768 | -282 | 0.20942 | 0.000456 | 0 | 99.6194 | 99.5529 | 126.8823 | 21.2612 | 0.001843 | 0.001758 | -151.9562 | 0.23367 | 8.3994 | 0.00268 | 0.003623 | 0.003482 | 5.4e-05 | 0.001227 | 0.30359 | 1.3289 | 2.5475 | 100.64 | 21.17 | 25.883 | 66.1657 | 86.2796 | 64.1362 | 0.006471 | 1.0467 | 0.025331 | 24.2002 | 0.00597 | 0.13901 | -1.2e-05 | -0.000939 | -0.000939 | -6.1125
2019-03-28 15:34 | 1553801620.2766 | -281 | 0.20952 | 0.000454 | 1e-06 | 99.6447 | 99.7007 | 126.8654 | 21.2612 | 0.001863 | 0.001735 | -151.9461 | 0.23335 | 8.4012 | 0.002616 | 0.004135 | 0.003996 | 5.7e-05 | 0.001223 | 0.32659 | 1.3414 | 2.5475 | 100.63 | 21.17 | 25.888 | 66.6573 | 84.7466 | 63.567 | 0.006473 | 1.0485 | 0.025351 | 24.1782 | -0.005025 | 0.13882 | -1.1e-05 | -0.000851 | -0.000851 | -6.739
2019-03-28 15:34 | 1553801621.2775 | -280 | 0.20942 | 0.000457 | 0 | 99.7379 | 99.6861 | 126.8166 | 21.2612 | 0.001841 | 0.00174 | -151.9559 | 0.23432 | 8.4022 | 0.002718 | 0.004181 | 0.003936 | 5.3e-05 | 0.00123 | 0.30359 | 1.297 | 2.5474 | 100.63 | 21.16 | 25.898 | 66.1282 | 87.2077 | 63.7045 | 0.006471 | 1.0481 | 0.025341 | 24.1785 | 0.00465 | 0.1394 | -1.4e-05 | -0.001083 | -0.001083 | -5.3141
2019-03-28 15:34 | 1553801622.2766 | -279 | 0.20949 | 0.000454 | -1e-06 | 99.81 | 99.6416 | 126.9382 | 21.2612 | 0.001882 | 0.001743 | -151.9709 | 0.23324 | 8.4033 | 0.00276 | 0.003898 | 0.003714 | 5.7e-05 | 0.001237 | 0.32964 | 1.2948 | 2.5472 | 100.63 | 21.16 | 25.901 | 67.1185 | 88.2205 | 63.7577 | 0.006472 | 1.0479 | 0.025351 | 24.192 | 0.015106 | 0.13875 | -1e-05 | -0.000689 | -0.000689 | -8.3121
2019-03-28 15:34 | 1553801623.2766 | -278 | 0.20966 | 0.000454 | 0 | 99.8079 | 99.733 | 126.8532 | 21.2612 | 0.001866 | 0.001724 | -151.9942 | 0.23444 | 8.4044 | 0.002718 | 0.003698 | 0.003517 | 5.8e-05 | 0.00124 | 0.30829 | 1.3023 | 2.5474 | 100.63 | 21.16 | 25.903 | 66.7362 | 87.2087 | 63.3165 | 0.006473 | 1.0493 | 0.025359 | 24.168 | 0.00672 | 0.13947 | -1.1e-05 | -0.000829 | -0.000829 | -6.9523
2019-03-28 15:34 | 1553801624.2766 | -277 | 0.20959 | 0.000454 | 0 | 99.7647 | 99.6708 | 126.7169 | 21.2612 | 0.001863 | 0.001744 | -151.9993 | 0.23371 | 8.4025 | 0.002682 | 0.004115 | 0.003897 | 5.5e-05 | 0.00124 | 0.32421 | 1.2775 | 2.5473 | 100.63 | 21.16 | 25.904 | 66.6612 | 86.3451 | 63.7998 | 0.006473 | 1.0478 | 0.025327 | 24.1724 | 0.008426 | 0.13903 | -1.1e-05 | -0.000854 | -0.000854 | -6.7252
2019-03-28 15:34 | 1553801625.2757 | -276 | 0.20932 | 0.000453 | 0 | 99.8633 | 99.843 | 126.3882 | 21.2612 | 0.001878 | 0.00174 | -151.9757 | 0.23403 | 8.403 | 0.002667 | 0.004213 | 0.004079 | 5.4e-05 | 0.001235 | 0.33196 | 1.3039 | 2.5475 | 100.63 | 21.16 | 25.902 | 67.007 | 85.9837 | 63.6898 | 0.006472 | 1.0481 | 0.025298 | 24.1371 | 0.001817 | 0.13922 | -9e-06 | -0.000689 | -0.000689 | -8.3406
2019-03-28 15:34 | 1553801626.2757 | -275 | 0.20951 | 0.000453 | 0 | 99.8931 | 99.825 | 126.565 | 21.2612 | 0.001843 | 0.001784 | -151.9874 | 0.23358 | 8.4059 | 0.002712 | 0.004454 | 0.004292 | 5.6e-05 | 0.001232 | 0.3298 | 1.2775 | 2.5475 | 100.63 | 21.16 | 25.894 | 66.1643 | 87.065 | 64.7428 | 0.00647 | 1.0448 | 0.025277 | 24.1917 | 0.006097 | 0.13896 | -9e-06 | -0.000741 | -0.000741 | -7.7425
2019-03-28 15:34 | 1553801627.2761 | -274 | 0.20948 | 0.000455 | 0 | 100.023 | 99.839 | 126.7955 | 21.2612 | 0.00189 | 0.001736 | -151.9654 | 0.2338 | 8.4026 | 0.002682 | 0.003687 | 0.003554 | 5.4e-05 | 0.001233 | 0.2943 | 1.3001 | 2.5474 | 100.63 | 21.16 | 25.89 | 67.3056 | 86.3388 | 63.5897 | 0.006469 | 1.0484 | 0.025343 | 24.1723 | 0.01647 | 0.13909 | -1.3e-05 | -0.000991 | -0.000991 | -5.7994
2019-03-28 15:34 | 1553801628.2767 | -273 | 0.20944 | 0.000457 | 0 | 100.0501 | 99.9623 | 126.8823 | 21.2612 | 0.001871 | 0.001731 | -151.9681 | 0.23383 | 8.405 | 0.002651 | 0.0043 | 0.004051 | 5e-05 | 0.00124 | 0.31614 | 1.3358 | 2.5475 | 100.63 | 21.16 | 25.886 | 66.8346 | 85.5957 | 63.4821 | 0.006468 | 1.0488 | 0.025356 | 24.1767 | 0.007852 | 0.1391 | -1.4e-05 | -0.001091 | -0.001091 | -5.2634
2019-03-28 15:34 | 1553801629.2772 | -272 | 0.20947 | 0.000455 | 1e-06 | 100.0924 | 100.071 | 126.5772 | 21.2612 | 0.001833 | 0.001726 | -151.9702 | 0.2341 | 8.4017 | 0.002649 | 0.003673 | 0.003618 | 5.7e-05 | 0.00125 | 0.31094 | 1.2969 | 2.5474 | 100.63 | 21.16 | 25.88 | 65.9262 | 85.5348 | 63.3468 | 0.006467 | 1.0492 | 0.02533 | 24.1428 | 0.001911 | 0.13926 | -1.3e-05 | -0.001016 | -0.001016 | -5.6596
2019-03-28 15:34 | 1553801630.2762 | -271 | 0.20952 | 0.000455 | 0 | 100.1784 | 100.176 | 126.6009 | 21.2612 | 0.0018 | 0.001693 | -151.9726 | 0.23353 | 8.401 | 0.002646 | 0.00369 | 0.003443 | 5.5e-05 | 0.001253 | 0.31526 | 1.3223 | 2.5474 | 100.63 | 21.16 | 25.877 | 65.1386 | 85.4701 | 62.5721 | 0.006466 | 1.0516 | 0.025362 | 24.1173 | 0.000221 | 0.13893 | -1.2e-05 | -0.000888 | -0.000888 | -6.4635
2019-03-28 15:34 | 1553801631.2759 | -270 | 0.20937 | 0.000459 | 0 | 100.0578 | 100.3954 | 126.6185 | 21.2612 | 0.001852 | 0.001781 | -151.9642 | 0.23338 | 8.4032 | 0.002669 | 0.004105 | 0.003955 | 5.7e-05 | 0.001255 | 0.31581 | 1.3424 | 2.5474 | 100.63 | 21.16 | 25.875 | 66.3795 | 86.0251 | 64.6743 | 0.006466 | 1.0451 | 0.025284 | 24.1943 | -0.03013 | 0.13884 | -1.7e-05 | -0.001321 | -0.001321 | -4.3422
2019-03-28 15:34 | 1553801632.2773 | -269 | 0.20952 | 0.000452 | 0 | 100.2213 | 100.1413 | 126.472 | 21.2612 | 0.001858 | 0.001736 | -151.9507 | 0.23305 | 8.4041 | 0.002611 | 0.003941 | 0.003715 | 5.8e-05 | 0.001259 | 0.32493 | 1.3077 | 2.5473 | 100.63 | 21.16 | 25.875 | 66.5402 | 84.6336 | 63.5976 | 0.006466 | 1.0484 | 0.02531 | 24.1418 | 0.007145 | 0.13864 | -8e-06 | -0.000642 | -0.000642 | -8.9137
2019-03-28 15:34 | 1553801633.2763 | -268 | 0.20958 | 0.000456 | 0 | 100.241 | 100.2789 | 126.2359 | 21.2612 | 0.001877 | 0.001747 | -151.9432 | 0.23353 | 8.4017 | 0.002626 | 0.004432 | 0.004277 | 5.5e-05 | 0.00126 | 0.28744 | 1.3476 | 2.5474 | 100.63 | 21.16 | 25.877 | 66.9928 | 84.9864 | 63.8669 | 0.006466 | 1.0476 | 0.025276 | 24.1289 | -0.003382 | 0.13893 | -1.3e-05 | -0.000959 | -0.000959 | -5.9827
2019-03-28 15:34 | 1553801634.2759 | -267 | 0.20952 | 0.000462 | 0 | 100.3101 | 100.3051 | 126.1737 | 21.2612 | 0.00187 | 0.001824 | -151.9295 | 0.2339 | 8.4026 | 0.002681 | 0.004005 | 0.003753 | 5.4e-05 | 0.001258 | 0.31045 | 1.3293 | 2.5474 | 100.63 | 21.16 | 25.883 | 66.8283 | 86.3061 | 65.7147 | 0.006468 | 1.0418 | 0.025201 | 24.189 | 0.000449 | 0.13914 | -2.1e-05 | -0.001586 | -0.001586 | -3.624
2019-03-28 15:34 | 1553801635.2763 | -266 | 0.20936 | 0.000454 | 0 | 100.3116 | 100.215 | 126.3758 | 21.2612 | 0.001848 | 0.001793 | -151.9258 | 0.23345 | 8.4015 | 0.002647 | 0.0039 | 0.00376 | 5.4e-05 | 0.001266 | 0.32101 | 1.3383 | 2.5474 | 100.63 | 21.16 | 25.886 | 66.3048 | 85.5069 | 64.9724 | 0.006468 | 1.0441 | 0.025249 | 24.1818 | 0.008613 | 0.13888 | -1.1e-05 | -0.000822 | -0.000822 | -6.9774
2019-03-28 15:34 | 1553801636.2761 | -265 | 0.20957 | 0.000454 | 0 | 100.3286 | 100.2561 | 126.2878 | 21.2612 | 0.001877 | 0.001802 | -151.8972 | 0.23356 | 8.4037 | 0.002695 | 0.003764 | 0.003696 | 5.5e-05 | 0.001263 | 0.32278 | 1.3343 | 2.5474 | 100.63 | 21.16 | 25.891 | 66.9966 | 86.6403 | 65.198 | 0.00647 | 1.0434 | 0.025232 | 24.1814 | 0.006459 | 0.13894 | -1.1e-05 | -0.000821 | -0.000821 | -6.9898
2019-03-28 15:34 | 1553801637.2772 | -264 | 0.20967 | 0.000457 | 0 | 100.3697 | 100.4841 | 126.0725 | 21.2612 | 0.001928 | 0.001753 | -151.9178 | 0.23253 | 8.402 | 0.002755 | 0.004499 | 0.004247 | 4.9e-05 | 0.00126 | 0.33859 | 1.3353 | 2.5475 | 100.63 | 21.16 | 25.901 | 68.219 | 88.0973 | 64.0075 | 0.006472 | 1.0471 | 0.025255 | 24.1183 | -0.010192 | 0.13833 | -1.5e-05 | -0.001155 | -0.001155 | -4.9486
2019-03-28 15:34 | 1553801638.2751 | -263 | 0.20968 | 0.000456 | 0 | 100.4394 | 100.5062 | 126.1042 | 21.2612 | 0.001909 | 0.001722 | -151.9022 | 0.2334 | 8.4048 | 0.002702 | 0.003969 | 0.003795 | 5.4e-05 | 0.001263 | 0.31996 | 1.2897 | 2.5475 | 100.63 | 21.16 | 25.907 | 67.7526 | 86.8191 | 63.2712 | 0.006474 | 1.0494 | 0.025286 | 24.095 | -0.005948 | 0.13885 | -1.3e-05 | -0.00104 | -0.00104 | -5.5148
2019-03-28 15:34 | 1553801639.2761 | -262 | 0.20963 | 0.000458 | 0 | 100.4855 | 100.5476 | 126.2129 | 21.2612 | 0.001899 | 0.001792 | -151.8999 | 0.23284 | 8.4026 | 0.002665 | 0.004274 | 0.004105 | 5e-05 | 0.001256 | 0.30934 | 1.3516 | 2.5475 | 100.63 | 21.16 | 25.917 | 67.5249 | 85.9252 | 64.9466 | 0.006476 | 1.0442 | 0.025234 | 24.1653 | -0.005526 | 0.13851 | -1.6e-05 | -0.001216 | -0.001216 | -4.7045
2019-03-28 15:34 | 1553801640.2755 | -261 | 0.20966 | 0.000452 | 0 | 100.5671 | 100.6332 | 126.4303 | 21.2612 | 0.001856 | 0.001791 | -151.9075 | 0.23251 | 8.4034 | 0.002726 | 0.003698 | 0.003576 | 5.2e-05 | 0.001255 | 0.34468 | 1.3186 | 2.5474 | 100.63 | 21.16 | 25.922 | 66.4772 | 87.4053 | 64.9233 | 0.006477 | 1.0443 | 0.025256 | 24.1852 | -0.005877 | 0.13832 | -9e-06 | -0.000665 | -0.000665 | -8.5909
2019-03-28 15:34 | 1553801641.2762 | -260 | 0.20934 | 0.000456 | 0 | 100.5871 | 100.6264 | 126.4939 | 21.2612 | 0.001849 | 0.001758 | -151.8995 | 0.23303 | 8.4017 | 0.002668 | 0.00424 | 0.004078 | 5.2e-05 | 0.001261 | 0.31996 | 1.3164 | 2.5474 | 100.63 | 21.16 | 25.927 | 66.31 | 86.0001 | 64.1261 | 0.006479 | 1.0468 | 0.025293 | 24.1628 | -0.003498 | 0.13863 | -1.3e-05 | -0.001012 | -0.001012 | -5.6586
2019-03-28 15:34 | 1553801642.2753 | -259 | 0.20941 | 0.000457 | 0 | 100.6353 | 100.6397 | 126.9175 | 21.2612 | 0.001886 | 0.001695 | -151.9021 | 0.23405 | 8.4066 | 0.002736 | 0.004036 | 0.003893 | 5.4e-05 | 0.001263 | 0.34683 | 1.3192 | 2.5475 | 100.63 | 21.16 | 25.934 | 67.2034 | 87.6474 | 62.6034 | 0.00648 | 1.0515 | 0.025392 | 24.1485 | -0.000386 | 0.13924 | -1.5e-05 | -0.001143 | -0.001143 | -5.0309
2019-03-28 15:34 | 1553801643.2752 | -258 | 0.20949 | 0.000456 | 0 | 100.6562 | 100.5368 | 127.1882 | 21.2612 | 0.001864 | 0.00178 | -151.9049 | 0.23301 | 8.4032 | 0.002752 | 0.003901 | 0.003724 | 4.6e-05 | 0.001265 | 0.30326 | 1.3317 | 2.5475 | 100.63 | 21.16 | 25.939 | 66.6698 | 88.0284 | 64.6579 | 0.006482 | 1.0451 | 0.025342 | 24.2481 | 0.01062 | 0.13861 | -1.3e-05 | -0.000989 | -0.000989 | -5.7903
2019-03-28 15:34 | 1553801644.2763 | -257 | 0.20934 | 0.000456 | 0 | 100.7378 | 100.9047 | 127.2994 | 21.2612 | 0.001845 | 0.001734 | -151.903 | 0.23371 | 8.4003 | 0.00274 | 0.004031 | 0.003866 | 5e-05 | 0.001265 | 0.31116 | 1.3636 | 2.5474 | 100.63 | 21.16 | 25.947 | 66.2179 | 87.7302 | 63.5389 | 0.006484 | 1.0486 | 0.025395 | 24.2185 | -0.014804 | 0.13903 | -1.3e-05 | -0.00103 | -0.00103 | -5.5754
2019-03-28 15:34 | 1553801645.2757 | -256 | 0.2095 | 0.000456 | 0 | 100.7765 | 100.814 | 127.214 | 21.2612 | 0.001896 | 0.00176 | -151.9145 | 0.23365 | 8.4003 | 0.00267 | 0.003741 | 0.00345 | 4.9e-05 | 0.001261 | 0.30165 | 1.3117 | 2.5473 | 100.63 | 21.16 | 25.949 | 67.4523 | 86.0484 | 64.1848 | 0.006484 | 1.0466 | 0.025362 | 24.2336 | -0.003322 | 0.139 | -1.3e-05 | -0.00101 | -0.00101 | -5.6836
2019-03-28 15:34 | 1553801646.2754 | -255 | 0.20938 | 0.000452 | 0 | 100.733 | 100.7002 | 127.1779 | 21.2612 | 0.001817 | 0.001698 | -151.9219 | 0.23332 | 8.4021 | 0.002626 | 0.004041 | 0.003823 | 5e-05 | 0.001262 | 0.33682 | 1.2797 | 2.5475 | 100.63 | 21.16 | 25.952 | 65.5569 | 84.9934 | 62.6737 | 0.006485 | 1.0513 | 0.025416 | 24.1758 | 0.00292 | 0.1388 | -8e-06 | -0.000593 | -0.000593 | -9.673
2019-03-28 15:34 | 1553801647.2766 | -254 | 0.20949 | 0.000451 | 0 | 100.7461 | 101.0215 | 126.992 | 21.2612 | 0.001883 | 0.00172 | -151.9149 | 0.23379 | 8.405 | 0.002682 | 0.003786 | 0.003635 | 5.2e-05 | 0.001262 | 0.31454 | 1.2881 | 2.5474 | 100.63 | 21.16 | 25.956 | 67.1349 | 86.3436 | 63.2031 | 0.006486 | 1.0496 | 0.025377 | 24.1772 | -0.024417 | 0.13908 | -7e-06 | -0.000506 | -0.000506 | -11.3501
2019-03-28 15:34 | 1553801648.2764 | -253 | 0.20957 | 0.000448 | 0 | 100.8176 | 100.8816 | 126.8265 | 21.2612 | 0.001859 | 0.001753 | -151.945 | 0.23329 | 8.4035 | 0.002639 | 0.004207 | 0.003972 | 5.3e-05 | 0.001264 | 0.31205 | 1.3452 | 2.5473 | 100.63 | 21.16 | 25.958 | 66.5597 | 85.3101 | 64.0099 | 0.006486 | 1.0471 | 0.02533 | 24.1904 | -0.005671 | 0.13878 | -3e-06 | -0.000226 | -0.000226 | -25.3696
2019-03-28 15:34 | 1553801649.2755 | -252 | 0.20946 | 0.000452 | 0 | 100.8111 | 100.7188 | 127.0253 | 21.2612 | 0.001904 | 0.001774 | -151.9287 | 0.2337 | 8.404 | 0.00266 | 0.00442 | 0.004238 | 5.2e-05 | 0.001269 | 0.32516 | 1.3292 | 2.5473 | 100.63 | 21.16 | 25.962 | 67.6434 | 85.8167 | 64.5118 | 0.006487 | 1.0456 | 0.025331 | 24.2273 | 0.008188 | 0.13903 | -8e-06 | -0.000615 | -0.000615 | -9.3314
2019-03-28 15:34 | 1553801650.2759 | -251 | 0.20937 | 0.000452 | 0 | 100.8233 | 100.8214 | 126.9004 | 21.2612 | 0.00191 | 0.00172 | -151.9217 | 0.23347 | 8.4029 | 0.002622 | 0.00415 | 0.004008 | 5.1e-05 | 0.001265 | 0.3246 | 1.3591 | 2.5475 | 100.63 | 21.16 | 25.965 | 67.787 | 84.9012 | 63.2149 | 0.006488 | 1.0496 | 0.025367 | 24.1689 | 0.000164 | 0.13889 | -8e-06 | -0.000581 | -0.000581 | -9.881
2019-03-28 15:34 | 1553801651.2755 | -250 | 0.20956 | 0.000451 | 1e-06 | 100.8769 | 101.0497 | 126.7265 | 21.2612 | 0.001875 | 0.001697 | -151.9373 | 0.23424 | 8.404 | 0.002634 | 0.00387 | 0.003721 | 5.5e-05 | 0.001261 | 0.31924 | 1.3087 | 2.5475 | 100.63 | 21.16 | 25.966 | 66.9538 | 85.1815 | 62.6633 | 0.006488 | 1.0513 | 0.025371 | 24.1325 | -0.015312 | 0.13935 | -7e-06 | -0.000605 | -0.000605 | -9.506
2019-03-28 15:34 | 1553801652.2765 | -249 | 0.20971 | 0.00045 | 0 | 100.8954 | 101.0321 | 126.2946 | 21.2612 | 0.001906 | 0.001748 | -151.9368 | 0.23409 | 8.4035 | 0.002615 | 0.003328 | 0.003121 | 5.5e-05 | 0.001257 | 0.34108 | 1.3442 | 2.5474 | 100.63 | 21.16 | 25.968 | 67.6848 | 84.7382 | 63.8864 | 0.006489 | 1.0475 | 0.025282 | 24.1352 | -0.01211 | 0.13926 | -6e-06 | -0.000454 | -0.000454 | -12.6663
2019-03-28 15:34 | 1553801653.2759 | -248 | 0.20963 | 0.000453 | 0 | 100.98 | 101.044 | 126.2082 | 21.2612 | 0.001864 | 0.001794 | -151.9463 | 0.23341 | 8.4025 | 0.002557 | 0.004027 | 0.003825 | 5.1e-05 | 0.001249 | 0.29446 | 1.2893 | 2.5474 | 100.63 | 21.16 | 25.969 | 66.6785 | 83.3353 | 64.9834 | 0.006489 | 1.0441 | 0.025232 | 24.1661 | -0.005667 | 0.13886 | -8e-06 | -0.000609 | -0.000609 | -9.4117
2019-03-28 15:34 | 1553801654.2746 | -247 | 0.20936 | 0.000453 | 0 | 100.991 | 101.1201 | 126.1568 | 21.2612 | 0.001852 | 0.001752 | -151.947 | 0.23391 | 8.4035 | 0.002573 | 0.004201 | 0.004021 | 5.1e-05 | 0.001247 | 0.30845 | 1.2961 | 2.5475 | 100.63 | 21.16 | 25.971 | 66.3889 | 83.7089 | 63.9935 | 0.00649 | 1.0472 | 0.025264 | 24.1258 | -0.011425 | 0.13915 | -9e-06 | -0.000664 | -0.000664 | -8.6531
2019-03-28 15:34 | 1553801655.2756 | -246 | 0.20936 | 0.000448 | 0 | 101.0148 | 101.1591 | 126.5188 | 21.2612 | 0.001857 | 0.001742 | -151.9615 | 0.23405 | 8.4038 | 0.00262 | 0.004288 | 0.004182 | 5.4e-05 | 0.001247 | 0.29535 | 1.2976 | 2.5474 | 100.63 | 21.16 | 25.971 | 66.5075 | 84.8582 | 63.7459 | 0.00649 | 1.0479 | 0.025309 | 24.1515 | -0.01277 | 0.13924 | -3e-06 | -0.00028 | -0.00028 | -20.5297
2019-03-28 15:34 | 1553801656.2747 | -245 | 0.20947 | 0.000452 | 1e-06 | 100.9966 | 101.0603 | 126.6251 | 21.2612 | 0.001891 | 0.001799 | -151.9546 | 0.23387 | 8.4024 | 0.002512 | 0.003419 | 0.003256 | 5.4e-05 | 0.00125 | 0.3147 | 1.3424 | 2.5474 | 100.63 | 21.16 | 25.972 | 67.3213 | 82.251 | 65.1222 | 0.00649 | 1.0437 | 0.025268 | 24.211 | -0.005633 | 0.13913 | -9e-06 | -0.000736 | -0.000736 | -7.8104
2019-03-28 15:34 | 1553801657.2747 | -244 | 0.20933 | 0.000455 | 1e-06 | 101.0047 | 100.8918 | 126.7018 | 21.2612 | 0.001861 | 0.00177 | -151.9643 | 0.2346 | 8.3998 | 0.00253 | 0.004131 | 0.003979 | 5.2e-05 | 0.001243 | 0.31094 | 1.3139 | 2.5474 | 100.63 | 21.17 | 25.975 | 66.6073 | 82.6902 | 64.4259 | 0.006495 | 1.0458 | 0.025302 | 24.1934 | 0.010005 | 0.13956 | -1.3e-05 | -0.000987 | -0.000987 | -5.8379
2019-03-28 15:34 | 1553801658.2759 | -243 | 0.2096 | 0.000453 | 0 | 101.022 | 101.1469 | 126.8542 | 21.2612 | 0.001896 | 0.001752 | -151.9452 | 0.23401 | 8.401 | 0.002524 | 0.003889 | 0.003639 | 5.2e-05 | 0.001242 | 0.3178 | 1.3315 | 2.5475 | 100.63 | 21.17 | 25.974 | 67.4547 | 82.5467 | 63.9771 | 0.006494 | 1.0472 | 0.025334 | 24.1918 | -0.011055 | 0.13922 | -9e-06 | -0.000711 | -0.000711 | -8.089
2019-03-28 15:34 | 1553801659.2748 | -242 | 0.20948 | 0.000458 | 0 | 101.0809 | 101.0902 | 126.7699 | 21.2612 | 0.001945 | 0.001782 | -151.9489 | 0.23348 | 8.4019 | 0.002566 | 0.003866 | 0.003745 | 5.1e-05 | 0.00124 | 0.32388 | 1.3196 | 2.5474 | 100.63 | 21.17 | 25.975 | 68.6286 | 83.5396 | 64.7094 | 0.006495 | 1.0449 | 0.025298 | 24.21 | -0.000817 | 0.1389 | -1.5e-05 | -0.001165 | -0.001165 | -4.9239
2019-03-28 15:34 | 1553801660.2747 | -241 | 0.20933 | 0.000446 | 0 | 101.0169 | 101.1275 | 127.0213 | 21.2612 | 0.001852 | 0.001824 | -151.9688 | 0.23365 | 8.4036 | 0.002678 | 0.004121 | 0.003919 | 5e-05 | 0.001241 | 0.31974 | 1.3534 | 2.5475 | 100.63 | 21.17 | 25.976 | 66.3952 | 86.2522 | 65.7131 | 0.006495 | 1.0419 | 0.025286 | 24.27 | -0.009791 | 0.139 | 0 | -4.2e-05 | -4.2e-05 | -137.8327
2019-03-28 15:34 | 1553801661.2748 | -240 | 0.20945 | 0.000456 | 0 | 101.083 | 101.2322 | 127.1645 | 21.2612 | 0.001918 | 0.001747 | -151.9623 | 0.2344 | 8.402 | 0.002549 | 0.004169 | 0.003988 | 5.2e-05 | 0.00124 | 0.28816 | 1.329 | 2.5474 | 100.63 | 21.17 | 25.977 | 67.9764 | 83.1513 | 63.8537 | 0.006495 | 1.0476 | 0.02537 | 24.217 | -0.013188 | 0.13945 | -1.3e-05 | -0.001054 | -0.001054 | -5.4648
2019-03-28 15:34 | 1553801662.2758 | -239 | 0.20945 | 0.000452 | 0 | 101.0363 | 101.1941 | 127.1483 | 21.2612 | 0.001912 | 0.001752 | -151.9635 | 0.23399 | 8.4014 | 0.002659 | 0.00386 | 0.003733 | 5.1e-05 | 0.001243 | 0.31437 | 1.3516 | 2.5475 | 100.63 | 21.17 | 25.978 | 67.8252 | 85.7754 | 63.9817 | 0.006495 | 1.0472 | 0.025363 | 24.22 | -0.013964 | 0.1392 | -8e-06 | -0.000584 | -0.000584 | -9.8464
2019-03-28 15:34 | 1553801663.2764 | -238 | 0.20952 | 0.000453 | 0 | 101.052 | 101.0861 | 127.2046 | 21.2612 | 0.001836 | 0.001738 | -151.961 | 0.23361 | 8.4025 | 0.002632 | 0.003927 | 0.003709 | 5.2e-05 | 0.001249 | 0.29463 | 1.3602 | 2.5474 | 100.63 | 21.17 | 25.979 | 65.9995 | 85.1346 | 63.6467 | 0.006496 | 1.0482 | 0.025382 | 24.2133 | -0.003013 | 0.13897 | -1e-05 | -0.000772 | -0.000772 | -7.432
2019-03-28 15:34 | 1553801664.275 | -237 | 0.20949 | 0.000449 | 0 | 101.0366 | 101.2509 | 127.1187 | 21.2612 | 0.001895 | 0.001747 | -151.9406 | 0.23461 | 8.4015 | 0.002629 | 0.004384 | 0.004201 | 5e-05 | 0.001253 | 0.32587 | 1.3467 | 2.5474 | 100.63 | 21.17 | 25.981 | 67.4283 | 85.0542 | 63.8669 | 0.006496 | 1.0476 | 0.025365 | 24.2131 | -0.018951 | 0.13957 | -4e-06 | -0.000333 | -0.000333 | -17.2899
2019-03-28 15:34 | 1553801665.2761 | -236 | 0.2097 | 0.000453 | 0 | 101.0154 | 100.9767 | 126.823 | 21.2612 | 0.001867 | 0.001787 | -151.9402 | 0.23409 | 8.4045 | 0.002606 | 0.00409 | 0.003945 | 4.9e-05 | 0.001251 | 0.335 | 1.3905 | 2.5474 | 100.63 | 21.16 | 25.983 | 66.7595 | 84.5056 | 64.8155 | 0.006493 | 1.0446 | 0.0253 | 24.2189 | 0.003428 | 0.13926 | -1e-05 | -0.0008 | -0.0008 | -7.1864
2019-03-28 15:34 | 1553801666.275 | -235 | 0.20939 | 0.000451 | 0 | 101.1179 | 101.2203 | 126.7312 | 21.2612 | 0.001874 | 0.001774 | -151.95 | 0.23359 | 8.4022 | 0.002596 | 0.004152 | 0.003971 | 4.9e-05 | 0.001254 | 0.32278 | 1.3052 | 2.5475 | 100.63 | 21.16 | 25.982 | 66.9117 | 84.28 | 64.5143 | 0.006492 | 1.0456 | 0.025302 | 24.1994 | -0.009054 | 0.13896 | -7e-06 | -0.000517 | -0.000517 | -11.1049
2019-03-28 15:34 | 1553801667.276 | -234 | 0.20947 | 0.000453 | 0 | 101.114 | 101.3008 | 126.9633 | 21.2612 | 0.001931 | 0.001762 | -151.931 | 0.23464 | 8.4026 | 0.002688 | 0.003834 | 0.003682 | 4.9e-05 | 0.00125 | 0.29607 | 1.2872 | 2.5474 | 100.63 | 21.16 | 25.983 | 68.2777 | 86.481 | 64.2127 | 0.006493 | 1.0465 | 0.025336 | 24.2107 | -0.016509 | 0.13959 | -1e-05 | -0.000755 | -0.000755 | -7.6319
2019-03-28 15:34 | 1553801668.276 | -233 | 0.2096 | 0.000454 | 0 | 101.0622 | 101.1058 | 127.0417 | 21.2612 | 0.001878 | 0.001762 | -151.9422 | 0.23377 | 8.4029 | 0.00264 | 0.003752 | 0.003571 | 5e-05 | 0.001254 | 0.30541 | 1.2859 | 2.5473 | 100.63 | 21.16 | 25.981 | 67.0264 | 85.3383 | 64.2291 | 0.006492 | 1.0464 | 0.025343 | 24.2188 | -0.00386 | 0.13907 | -1e-05 | -0.000764 | -0.000764 | -7.521
2019-03-28 15:34 | 1553801669.2741 | -232 | 0.20951 | 0.000449 | 0 | 101.0866 | 101.0813 | 127.2274 | 21.2612 | 0.001837 | 0.001807 | -151.9605 | 0.23442 | 8.4047 | 0.002779 | 0.004309 | 0.004145 | 5.1e-05 | 0.001253 | 0.30829 | 1.332 | 2.5474 | 100.63 | 21.16 | 25.982 | 66.0183 | 88.6753 | 65.2995 | 0.006492 | 1.0431 | 0.025322 | 24.2749 | 0.000466 | 0.13946 | -5e-06 | -0.00036 | -0.00036 | -15.98
2019-03-28 15:34 | 1553801670.2751 | -231 | 0.20946 | 0.000452 | 0 | 101.1122 | 101.2621 | 127.2325 | 21.2612 | 0.001861 | 0.001744 | -151.9391 | 0.2334 | 8.4025 | 0.002664 | 0.004326 | 0.004145 | 5.1e-05 | 0.001252 | 0.31708 | 1.3118 | 2.5475 | 100.63 | 21.16 | 25.981 | 66.6066 | 85.9072 | 63.7873 | 0.006492 | 1.0478 | 0.025379 | 24.2211 | -0.013247 | 0.13885 | -8e-06 | -0.000603 | -0.000603 | -9.503
2019-03-28 15:34 | 1553801671.2741 | -230 | 0.20949 | 0.000451 | 0 | 101.0744 | 101.1734 | 127.0861 | 21.2612 | 0.001909 | 0.001724 | -151.9298 | 0.23352 | 8.4029 | 0.002569 | 0.004279 | 0.004097 | 4.8e-05 | 0.001254 | 0.34811 | 1.3304 | 2.5475 | 100.63 | 21.16 | 25.979 | 67.7488 | 83.6099 | 63.3218 | 0.006492 | 1.0493 | 0.025382 | 24.1904 | -0.008758 | 0.13892 | -6e-06 | -0.000436 | -0.000436 | -13.1587
2019-03-28 15:35 | 1553801672.2752 | -229 | 0.20946 | 0.00045 | 0 | 101.117 | 101.1819 | 127.2645 | 21.2612 | 0.001829 | 0.001769 | -151.9317 | 0.23402 | 8.4052 | 0.002657 | 0.004094 | 0.003875 | 4.7e-05 | 0.001249 | 0.30574 | 1.3136 | 2.5474 | 100.63 | 21.16 | 25.976 | 65.8341 | 85.7479 | 64.3908 | 0.006491 | 1.0459 | 0.025359 | 24.2458 | -0.005741 | 0.13922 | -6e-06 | -0.000457 | -0.000457 | -12.5901
2019-03-28 15:35 | 1553801673.2752 | -228 | 0.20945 | 0.000448 | 0 | 101.1077 | 101.366 | 127.211 | 21.2612 | 0.001858 | 0.001715 | -151.9252 | 0.23406 | 8.4018 | 0.002665 | 0.003794 | 0.003543 | 4.5e-05 | 0.001247 | 0.31188 | 1.3467 | 2.5474 | 100.63 | 21.17 | 25.973 | 66.5364 | 85.9236 | 63.1039 | 0.006494 | 1.0499 | 0.025403 | 24.1944 | -0.02282 | 0.13924 | -3e-06 | -0.000279 | -0.000279 | -20.5872
2019-03-28 15:35 | 1553801674.2746 | -227 | 0.2095 | 0.000449 | 0 | 101.1063 | 101.2811 | 127.3159 | 21.2612 | 0.001909 | 0.001757 | -151.9663 | 0.23416 | 8.403 | 0.002655 | 0.004206 | 0.004039 | 4.8e-05 | 0.001246 | 0.31708 | 1.3314 | 2.5473 | 100.63 | 21.17 | 25.971 | 67.7588 | 85.6941 | 64.0989 | 0.006494 | 1.0468 | 0.025376 | 24.2402 | -0.015454 | 0.1393 | -4e-06 | -0.00034 | -0.00034 | -16.9216
2019-03-28 15:35 | 1553801675.2743 | -226 | 0.20947 | 0.000447 | 0 | 101.131 | 101.1741 | 127.3039 | 21.2612 | 0.001894 | 0.001717 | -151.9507 | 0.23317 | 8.403 | 0.002659 | 0.004177 | 0.003928 | 4.9e-05 | 0.001242 | 0.29551 | 1.3682 | 2.5475 | 100.63 | 21.17 | 25.967 | 67.4095 | 85.79 | 63.1509 | 0.006493 | 1.0498 | 0.02541 | 24.205 | -0.003812 | 0.13871 | -2e-06 | -0.000143 | -0.000143 | -40.1358
2019-03-28 15:35 | 1553801676.2757 | -225 | 0.20949 | 0.000457 | 0 | 101.1164 | 101.3674 | 126.9854 | 21.2612 | 0.001857 | 0.001733 | -151.952 | 0.23325 | 8.4051 | 0.002713 | 0.00349 | 0.003334 | 4.7e-05 | 0.001238 | 0.30359 | 1.3663 | 2.5475 | 100.63 | 21.17 | 25.966 | 66.5193 | 87.0821 | 63.5179 | 0.006492 | 1.0486 | 0.025364 | 24.1878 | -0.022176 | 0.13876 | -1.4e-05 | -0.001093 | -0.001093 | -5.2448
2019-03-28 15:35 | 1553801677.2753 | -224 | 0.20948 | 0.000451 | 0 | 101.1295 | 101.4292 | 126.884 | 21.2612 | 0.001823 | 0.001735 | -151.9753 | 0.23381 | 8.3998 | 0.00266 | 0.00396 | 0.003749 | 4.7e-05 | 0.001238 | 0.33373 | 1.3256 | 2.5474 | 100.63 | 21.17 | 25.965 | 65.7037 | 85.8088 | 63.5834 | 0.006492 | 1.0484 | 0.025352 | 24.1805 | -0.02647 | 0.13909 | -7e-06 | -0.00052 | -0.00052 | -11.0531
2019-03-28 15:35 | 1553801678.2743 | -223 | 0.20946 | 0.000453 | 0 | 101.1569 | 101.2295 | 127.0011 | 21.2612 | 0.001871 | 0.001708 | -151.9849 | 0.23422 | 8.4039 | 0.0027 | 0.004176 | 0.004071 | 4.6e-05 | 0.001245 | 0.29607 | 1.3583 | 2.5474 | 100.63 | 21.16 | 25.963 | 66.8406 | 86.7705 | 62.9166 | 0.006488 | 1.0505 | 0.025389 | 24.1677 | -0.006414 | 0.13934 | -9e-06 | -0.00065 | -0.00065 | -8.8515
2019-03-28 15:35 | 1553801679.2758 | -222 | 0.20948 | 0.000445 | 0 | 101.069 | 101.1085 | 127.1396 | 21.2612 | 0.001872 | 0.001741 | -151.9768 | 0.23341 | 8.4048 | 0.002691 | 0.003526 | 0.003373 | 4.9e-05 | 0.001246 | 0.29767 | 1.3285 | 2.5473 | 100.63 | 21.17 | 25.963 | 66.8718 | 86.5457 | 63.7264 | 0.006492 | 1.048 | 0.025372 | 24.21 | -0.003494 | 0.13886 | 1e-06 | 6.1e-05 | 6.1e-05 | 94.5367
2019-03-28 15:35 | 1553801680.2745 | -221 | 0.20943 | 0.000452 | 0 | 101.1313 | 101.0929 | 126.6955 | 21.2612 | 0.001836 | 0.001718 | -151.9925 | 0.23442 | 8.4043 | 0.002703 | 0.003831 | 0.003703 | 4.5e-05 | 0.001243 | 0.31078 | 1.304 | 2.5474 | 100.63 | 21.17 | 25.961 | 66.0128 | 86.8362 | 63.1766 | 0.006491 | 1.0497 | 0.025348 | 24.148 | 0.003396 | 0.13945 | -7e-06 | -0.000553 | -0.000553 | -10.415
2019-03-28 15:35 | 1553801681.2749 | -220 | 0.20949 | 0.000451 | 0 | 101.1441 | 101.0915 | 126.6568 | 21.2612 | 0.001864 | 0.00175 | -151.9909 | 0.23388 | 8.4019 | 0.002643 | 0.004185 | 0.003972 | 4.8e-05 | 0.001242 | 0.31758 | 1.2918 | 2.5475 | 100.63 | 21.17 | 25.961 | 66.6785 | 85.4116 | 63.9278 | 0.006491 | 1.0474 | 0.025316 | 24.1712 | 0.004649 | 0.13914 | -6e-06 | -0.000468 | -0.000468 | -12.2677
2019-03-28 15:35 | 1553801682.2756 | -219 | 0.20947 | 0.000452 | 0 | 101.0982 | 101.2349 | 126.7645 | 21.2612 | 0.001832 | 0.001703 | -152.0168 | 0.23456 | 8.4042 | 0.002639 | 0.003518 | 0.003321 | 4.8e-05 | 0.001245 | 0.28407 | 1.3339 | 2.5475 | 100.63 | 21.17 | 25.961 | 65.9019 | 85.3087 | 62.8119 | 0.006491 | 1.0508 | 0.025369 | 24.1414 | -0.012084 | 0.13954 | -8e-06 | -0.000569 | -0.000569 | -10.1342
2019-03-28 15:35 | 1553801683.2746 | -218 | 0.20946 | 0.000456 | 0 | 101.1864 | 101.2461 | 126.7725 | 21.2612 | 0.001894 | 0.001834 | -152.0248 | 0.23369 | 8.4029 | 0.002674 | 0.004018 | 0.003838 | 4.5e-05 | 0.001248 | 0.26344 | 1.335 | 2.5474 | 100.63 | 21.17 | 25.96 | 67.4064 | 86.1413 | 65.9675 | 0.006491 | 1.0411 | 0.025252 | 24.2553 | -0.005278 | 0.13902 | -1.3e-05 | -0.00097 | -0.00097 | -5.9174
2019-03-28 15:35 | 1553801684.2736 | -217 | 0.20948 | 0.000452 | 0 | 101.1655 | 101.2879 | 126.553 | 21.2612 | 0.001817 | 0.00172 | -152.0145 | 0.23447 | 8.402 | 0.002655 | 0.004078 | 0.003892 | 4.9e-05 | 0.001249 | 0.31293 | 1.3307 | 2.5473 | 100.63 | 21.16 | 25.959 | 65.5367 | 85.6919 | 63.211 | 0.006487 | 1.0496 | 0.025333 | 24.1356 | -0.010812 | 0.13949 | -8e-06 | -0.000624 | -0.000624 | -9.2292
2019-03-28 15:35 | 1553801685.2735 | -216 | 0.20948 | 0.000449 | 0 | 101.0714 | 101.3042 | 126.3918 | 21.2612 | 0.001873 | 0.001756 | -152.0266 | 0.23517 | 8.4039 | 0.002624 | 0.004411 | 0.004235 | 4.7e-05 | 0.00125 | 0.323 | 1.3168 | 2.5474 | 100.63 | 21.17 | 25.956 | 66.9038 | 84.945 | 64.0878 | 0.00649 | 1.0469 | 0.025284 | 24.1517 | -0.020578 | 0.1399 | -4e-06 | -0.000301 | -0.000301 | -19.2031
2019-03-28 15:35 | 1553801686.2739 | -215 | 0.20947 | 0.000453 | 1e-06 | 101.2176 | 101.4859 | 126.7192 | 21.2612 | 0.001912 | 0.001782 | -152.0402 | 0.23422 | 8.4055 | 0.002664 | 0.004429 | 0.004296 | 4.6e-05 | 0.00125 | 0.31045 | 1.3236 | 2.5473 | 100.63 | 21.16 | 25.955 | 67.8283 | 85.894 | 64.6937 | 0.006486 | 1.045 | 0.025294 | 24.2046 | -0.023678 | 0.13934 | -9e-06 | -0.000737 | -0.000737 | -7.805
2019-03-28 15:35 | 1553801687.2751 | -214 | 0.20948 | 0.000453 | 0 | 101.1634 | 101.1866 | 126.8518 | 21.2612 | 0.00189 | 0.001707 | -152.0303 | 0.23283 | 8.4031 | 0.002651 | 0.00362 | 0.003455 | 4.6e-05 | 0.001247 | 0.31492 | 1.3158 | 2.5475 | 100.63 | 21.16 | 25.953 | 67.3032 | 85.5848 | 62.9055 | 0.006485 | 1.0506 | 0.025374 | 24.1531 | -0.002052 | 0.13851 | -1e-05 | -0.000759 | -0.000759 | -7.5361
2019-03-28 15:35 | 1553801688.2747 | -213 | 0.20949 | 0.000455 | 0 | 101.1652 | 101.2682 | 126.5929 | 21.2612 | 0.001869 | 0.001734 | -152.0241 | 0.23294 | 8.4025 | 0.002722 | 0.004263 | 0.004034 | 4.7e-05 | 0.001247 | 0.30574 | 1.3365 | 2.5474 | 100.63 | 21.16 | 25.948 | 66.7891 | 87.2912 | 63.5461 | 0.006484 | 1.0486 | 0.025324 | 24.1515 | -0.009098 | 0.13858 | -1.2e-05 | -0.000916 | -0.000916 | -6.2511
2019-03-28 15:35 | 1553801689.2738 | -212 | 0.2095 | 0.000452 | 0 | 101.201 | 101.1663 | 126.1631 | 21.2612 | 0.001869 | 0.001717 | -152.0121 | 0.23304 | 8.4026 | 0.002634 | 0.00395 | 0.003896 | 4.6e-05 | 0.001246 | 0.30021 | 1.3433 | 2.5474 | 100.63 | 21.16 | 25.948 | 66.7869 | 85.1798 | 63.147 | 0.006484 | 1.0498 | 0.025296 | 24.0961 | 0.003068 | 0.13863 | -7e-06 | -0.000582 | -0.000582 | -9.8313
2019-03-28 15:35 | 1553801690.2748 | -211 | 0.2095 | 0.00045 | 0 | 101.2194 | 101.0657 | 126.2817 | 21.2612 | 0.00194 | 0.001786 | -152.0019 | 0.2331 | 8.4049 | 0.00266 | 0.00391 | 0.003661 | 4.6e-05 | 0.00125 | 0.30397 | 1.3313 | 2.5474 | 100.63 | 21.16 | 25.946 | 68.5008 | 85.8035 | 64.8054 | 0.006483 | 1.0447 | 0.025246 | 24.1668 | 0.013595 | 0.13867 | -5e-06 | -0.000375 | -0.000375 | -15.2532
2019-03-28 15:35 | 1553801691.2738 | -210 | 0.20948 | 0.000455 | 0 | 101.15 | 101.2488 | 126.7396 | 21.2612 | 0.001835 | 0.001772 | -151.9944 | 0.23319 | 8.4038 | 0.002637 | 0.004336 | 0.004155 | 4.7e-05 | 0.001253 | 0.33069 | 1.3131 | 2.5474 | 100.63 | 21.17 | 25.942 | 65.9784 | 85.2642 | 64.4604 | 0.006486 | 1.0457 | 0.025305 | 24.1982 | -0.008731 | 0.13873 | -1.1e-05 | -0.000879 | -0.000879 | -6.5215
2019-03-28 15:35 | 1553801692.2749 | -209 | 0.20949 | 0.00045 | 0 | 101.1888 | 101.2349 | 126.6399 | 21.2612 | 0.001912 | 0.001744 | -151.981 | 0.23315 | 8.3995 | 0.002588 | 0.004109 | 0.003967 | 4.7e-05 | 0.001259 | 0.2995 | 1.3542 | 2.5475 | 100.63 | 21.17 | 25.939 | 67.8384 | 84.0735 | 63.7873 | 0.006486 | 1.0478 | 0.02532 | 24.1646 | -0.004077 | 0.1387 | -5e-06 | -0.000423 | -0.000423 | -13.5414
2019-03-28 15:35 | 1553801693.2749 | -208 | 0.20949 | 0.00045 | 0 | 101.1795 | 101.224 | 126.6582 | 21.2612 | 0.001855 | 0.001729 | -151.9885 | 0.23317 | 8.4026 | 0.002583 | 0.004279 | 0.004058 | 4.7e-05 | 0.00126 | 0.30198 | 1.3643 | 2.5475 | 100.63 | 21.16 | 25.935 | 66.47 | 83.9657 | 63.4282 | 0.006481 | 1.0489 | 0.025335 | 24.1534 | -0.003932 | 0.13871 | -6e-06 | -0.000466 | -0.000466 | -12.3007
2019-03-28 15:35 | 1553801694.2752 | -207 | 0.20949 | 0.000451 | 0 | 101.2132 | 101.2447 | 126.7044 | 21.2612 | 0.001867 | 0.001758 | -151.98 | 0.23367 | 8.4025 | 0.00257 | 0.004216 | 0.004047 | 4.8e-05 | 0.001259 | 0.29645 | 1.3347 | 2.5474 | 100.63 | 21.16 | 25.934 | 66.7509 | 83.6522 | 64.1285 | 0.00648 | 1.0467 | 0.025313 | 24.183 | -0.002789 | 0.13901 | -7e-06 | -0.000536 | -0.000536 | -10.7079
2019-03-28 15:35 | 1553801695.2739 | -206 | 0.20946 | 0.000452 | 0 | 101.1649 | 101.208 | 126.6394 | 21.2612 | 0.001864 | 0.001742 | -151.9595 | 0.2334 | 8.4055 | 0.002584 | 0.003816 | 0.003693 | 4.9e-05 | 0.001255 | 0.30989 | 1.3439 | 2.5473 | 100.63 | 21.16 | 25.931 | 66.6806 | 83.9891 | 63.7466 | 0.00648 | 1.0479 | 0.025321 | 24.1631 | -0.003812 | 0.13885 | -8e-06 | -0.000612 | -0.000612 | -9.3648
2019-03-28 15:35 | 1553801696.2745 | -205 | 0.20949 | 0.000453 | 0 | 101.1754 | 101.3531 | 126.4962 | 21.2612 | 0.001896 | 0.001708 | -151.9456 | 0.23366 | 8.4037 | 0.002589 | 0.003548 | 0.00342 | 4.9e-05 | 0.001253 | 0.31222 | 1.2825 | 2.5473 | 100.63 | 21.16 | 25.928 | 67.4359 | 84.11 | 62.9352 | 0.006479 | 1.0505 | 0.025338 | 24.1203 | -0.015701 | 0.13901 | -9e-06 | -0.000709 | -0.000709 | -8.0969
2019-03-28 15:35 | 1553801697.275 | -204 | 0.20949 | 0.000451 | 0 | 101.2022 | 101.3796 | 126.4601 | 21.2612 | 0.00184 | 0.001719 | -151.9615 | 0.2328 | 8.4037 | 0.002641 | 0.004029 | 0.003856 | 4.9e-05 | 0.001249 | 0.28998 | 1.2684 | 2.5472 | 100.63 | 21.16 | 25.927 | 66.1001 | 85.3491 | 63.1954 | 0.006479 | 1.0497 | 0.025324 | 24.1262 | -0.01567 | 0.13849 | -6e-06 | -0.000425 | -0.000425 | -13.4732
2019-03-28 15:35 | 1553801698.2753 | -203 | 0.20951 | 0.00045 | 0 | 101.1533 | 101.1493 | 126.4786 | 21.2612 | 0.001841 | 0.00175 | -151.9629 | 0.23325 | 8.4013 | 0.002549 | 0.003357 | 0.003135 | 4.9e-05 | 0.001256 | 0.3194 | 1.2998 | 2.5473 | 100.63 | 21.16 | 25.925 | 66.1268 | 83.1326 | 63.9451 | 0.006478 | 1.0473 | 0.025298 | 24.1548 | 0.000357 | 0.13876 | -4e-06 | -0.000322 | -0.000322 | -17.7949
2019-03-28 15:35 | 1553801699.2754 | -202 | 0.20952 | 0.000452 | 0 | 101.184 | 101.2685 | 126.4899 | 21.2612 | 0.001876 | 0.001801 | -151.9776 | 0.23307 | 8.4025 | 0.002595 | 0.003789 | 0.003567 | 4.6e-05 | 0.001255 | 0.30326 | 1.3136 | 2.5474 | 100.63 | 21.16 | 25.923 | 66.9726 | 84.238 | 65.1723 | 0.006478 | 1.0435 | 0.025253 | 24.1999 | -0.007469 | 0.13865 | -7e-06 | -0.000554 | -0.000554 | -10.3422
2019-03-28 15:35 | 1553801700.2741 | -201 | 0.2095 | 0.00045 | 0 | 101.1557 | 101.3256 | 126.3948 | 21.2612 | 0.00187 | 0.001766 | -151.9841 | 0.23388 | 8.4016 | 0.002605 | 0.003698 | 0.00351 | 4.9e-05 | 0.001256 | 0.31078 | 1.2826 | 2.5475 | 100.63 | 21.16 | 25.92 | 66.8148 | 84.4782 | 64.3323 | 0.006477 | 1.0461 | 0.025275 | 24.1607 | -0.01501 | 0.13914 | -5e-06 | -0.000418 | -0.000418 | -13.7638
2019-03-28 15:35 | 1553801701.2741 | -200 | 0.20951 | 0.000451 | 0 | 101.1548 | 101.4258 | 126.0791 | 21.2612 | 0.001868 | 0.001697 | -151.9911 | 0.23364 | 8.4022 | 0.002674 | 0.003399 | 0.003272 | 4.7e-05 | 0.001259 | 0.32211 | 1.3185 | 2.5474 | 100.63 | 21.16 | 25.916 | 66.7665 | 86.143 | 62.6696 | 0.006476 | 1.0513 | 0.025306 | 24.071 | -0.023932 | 0.13899 | -6e-06 | -0.000474 | -0.000474 | -12.1132
2019-03-28 15:35 | 1553801702.2731 | -199 | 0.20949 | 0.000451 | 0 | 101.229 | 101.2692 | 125.9512 | 21.2612 | 0.001879 | 0.001734 | -151.9975 | 0.23382 | 8.4014 | 0.002544 | 0.004119 | 0.003905 | 4.6e-05 | 0.001262 | 0.3183 | 1.3504 | 2.5475 | 100.63 | 21.17 | 25.909 | 67.0286 | 83.0186 | 63.5514 | 0.006478 | 1.0485 | 0.02526 | 24.0904 | -0.003554 | 0.1391 | -6e-06 | -0.00046 | -0.00046 | -12.485
2019-03-28 15:35 | 1553801703.2736 | -198 | 0.20949 | 0.000453 | 0 | 101.1 | 101.3694 | 126.019 | 21.2612 | 0.001832 | 0.001729 | -151.9854 | 0.23308 | 8.4021 | 0.002572 | 0.00392 | 0.003754 | 4.7e-05 | 0.001264 | 0.32045 | 1.3574 | 2.5476 | 100.63 | 21.16 | 25.906 | 65.9144 | 83.7044 | 63.4383 | 0.006473 | 1.0489 | 0.025271 | 24.0928 | -0.023805 | 0.13866 | -9e-06 | -0.000702 | -0.000702 | -8.1554
2019-03-28 15:35 | 1553801704.2732 | -197 | 0.20952 | 0.000453 | 0 | 101.1474 | 101.2461 | 125.9765 | 21.2612 | 0.001828 | 0.001754 | -152.0035 | 0.23282 | 8.4026 | 0.002591 | 0.00405 | 0.003794 | 4.7e-05 | 0.001266 | 0.29336 | 1.3561 | 2.5475 | 100.63 | 21.16 | 25.903 | 65.8069 | 84.1497 | 64.0412 | 0.006473 | 1.047 | 0.025244 | 24.1103 | -0.008728 | 0.13851 | -9e-06 | -0.000733 | -0.000733 | -7.8075
2019-03-28 15:35 | 1553801705.2742 | -196 | 0.20951 | 0.000453 | 0 | 101.1313 | 101.2328 | 125.992 | 21.2612 | 0.001851 | 0.001763 | -152.0187 | 0.23314 | 8.4047 | 0.00269 | 0.003881 | 0.003715 | 5.1e-05 | 0.00127 | 0.30342 | 1.316 | 2.5475 | 100.63 | 21.17 | 25.894 | 66.3719 | 86.5308 | 64.2527 | 0.006474 | 1.0464 | 0.025238 | 24.1193 | -0.00898 | 0.1387 | -9e-06 | -0.000674 | -0.000674 | -8.501
2019-03-28 15:35 | 1553801706.2747 | -195 | 0.20948 | 0.000452 | 0 | 101.1572 | 101.1799 | 125.9038 | 21.2612 | 0.001846 | 0.001725 | -152.0119 | 0.23378 | 8.4044 | 0.002689 | 0.004256 | 0.004121 | 4.7e-05 | 0.001269 | 0.27367 | 1.3271 | 2.5475 | 100.63 | 21.16 | 25.891 | 66.2554 | 86.5043 | 63.3367 | 0.00647 | 1.0492 | 0.025263 | 24.0781 | -0.002004 | 0.13908 | -7e-06 | -0.00057 | -0.00057 | -10.0839
2019-03-28 15:35 | 1553801707.2743 | -194 | 0.20949 | 0.000449 | 1e-06 | 101.2045 | 101.4938 | 125.8082 | 21.2612 | 0.001842 | 0.00172 | -151.9964 | 0.23365 | 8.4008 | 0.002591 | 0.003197 | 0.002955 | 4.6e-05 | 0.001267 | 0.29607 | 1.3289 | 2.5474 | 100.63 | 21.17 | 25.887 | 66.158 | 84.1569 | 63.2163 | 0.006473 | 1.0496 | 0.025258 | 24.0647 | -0.025525 | 0.139 | -5e-06 | -0.000399 | -0.000399 | -14.3905
2019-03-28 15:35 | 1553801708.2737 | -193 | 0.20949 | 0.000453 | 0 | 101.1557 | 101.1996 | 126.1277 | 21.2612 | 0.001811 | 0.001695 | -151.9869 | 0.23448 | 8.4029 | 0.002632 | 0.004183 | 0.004021 | 4.8e-05 | 0.001273 | 0.3011 | 1.3382 | 2.5475 | 100.63 | 21.17 | 25.881 | 65.4133 | 85.1338 | 62.6219 | 0.006471 | 1.0514 | 0.025312 | 24.0739 | -0.003878 | 0.13949 | -9e-06 | -0.000712 | -0.000712 | -8.0969
2019-03-28 15:35 | 1553801709.2734 | -192 | 0.20949 | 0.000453 | 0 | 101.1128 | 101.1469 | 126.3218 | 21.2612 | 0.001845 | 0.001716 | -151.9904 | 0.23372 | 8.4048 | 0.002628 | 0.003691 | 0.003509 | 4.8e-05 | 0.001274 | 0.30845 | 1.32 | 2.5474 | 100.63 | 21.17 | 25.879 | 66.2158 | 85.0504 | 63.1251 | 0.006471 | 1.0499 | 0.025313 | 24.1105 | -0.003016 | 0.13904 | -9e-06 | -0.000677 | -0.000677 | -8.477
2019-03-28 15:35 | 1553801710.2744 | -191 | 0.2095 | 0.000452 | 1e-06 | 101.201 | 101.3001 | 126.2671 | 21.2612 | 0.001805 | 0.001705 | -152.0064 | 0.23418 | 8.4048 | 0.002649 | 0.003794 | 0.003733 | 4.8e-05 | 0.001269 | 0.27959 | 1.3406 | 2.5474 | 100.63 | 21.17 | 25.876 | 65.2644 | 85.5396 | 62.8603 | 0.00647 | 1.0507 | 0.025317 | 24.0958 | -0.008759 | 0.13931 | -8e-06 | -0.000628 | -0.000628 | -9.1584
2019-03-28 15:35 | 1553801711.2733 | -190 | 0.20947 | 0.000448 | 0 | 101.1596 | 101.1965 | 126.3627 | 21.2612 | 0.001857 | 0.001771 | -151.9994 | 0.23381 | 8.4029 | 0.00262 | 0.004291 | 0.004083 | 4.5e-05 | 0.001272 | 0.31061 | 1.3719 | 2.5474 | 100.63 | 21.17 | 25.875 | 66.5099 | 84.8428 | 64.4385 | 0.00647 | 1.0458 | 0.025268 | 24.1614 | -0.003265 | 0.13909 | -3e-06 | -0.00025 | -0.00025 | -23.0036
2019-03-28 15:35 | 1553801712.2744 | -189 | 0.20947 | 0.000453 | 0 | 101.1551 | 101.1846 | 126.7206 | 21.2612 | 0.001867 | 0.001722 | -151.9853 | 0.23368 | 8.4046 | 0.002661 | 0.003388 | 0.003284 | 4.7e-05 | 0.001271 | 0.29391 | 1.3858 | 2.5475 | 100.63 | 21.17 | 25.876 | 66.7501 | 85.8314 | 63.2727 | 0.00647 | 1.0494 | 0.025347 | 24.1538 | -0.002609 | 0.13902 | -9e-06 | -0.000685 | -0.000685 | -8.383
2019-03-28 15:35 | 1553801713.2739 | -188 | 0.20946 | 0.000449 | 0 | 101.1161 | 101.2777 | 126.7711 | 21.2612 | 0.001856 | 0.001714 | -151.9745 | 0.23429 | 8.4013 | 0.002686 | 0.003783 | 0.003568 | 4.4e-05 | 0.001271 | 0.29031 | 1.3003 | 2.5473 | 100.63 | 21.17 | 25.876 | 66.4895 | 86.4449 | 63.0611 | 0.00647 | 1.0501 | 0.02536 | 24.151 | -0.014284 | 0.13938 | -4e-06 | -0.000335 | -0.000335 | -17.1659
2019-03-28 15:35 | 1553801714.2745 | -187 | 0.2095 | 0.00045 | 0 | 101.1634 | 101.275 | 126.9044 | 21.2612 | 0.001866 | 0.001733 | -151.9677 | 0.23423 | 8.4043 | 0.00263 | 0.004669 | 0.004383 | 4.8e-05 | 0.00127 | 0.28694 | 1.3114 | 2.5474 | 100.63 | 21.17 | 25.877 | 66.7174 | 85.0994 | 63.5295 | 0.00647 | 1.0486 | 0.025356 | 24.1806 | -0.009856 | 0.13935 | -6e-06 | -0.000448 | -0.000448 | -12.8476
2019-03-28 15:35 | 1553801715.2739 | -186 | 0.20951 | 0.000453 | 0 | 101.1792 | 101.3888 | 126.8028 | 21.2612 | 0.001827 | 0.001775 | -151.9659 | 0.23411 | 8.4025 | 0.002636 | 0.00433 | 0.004163 | 4.6e-05 | 0.001269 | 0.30093 | 1.3087 | 2.5474 | 100.63 | 21.17 | 25.882 | 65.7881 | 85.2344 | 64.5352 | 0.006471 | 1.0455 | 0.025308 | 24.2069 | -0.018506 | 0.13927 | -9e-06 | -0.000655 | -0.000655 | -8.7842
2019-03-28 15:35 | 1553801716.2739 | -185 | 0.20949 | 0.000453 | 0 | 101.1438 | 101.1534 | 126.9262 | 21.2612 | 0.001885 | 0.001822 | -151.9652 | 0.23394 | 8.4027 | 0.002699 | 0.00443 | 0.004281 | 4.7e-05 | 0.001264 | 0.30989 | 1.3364 | 2.5474 | 100.63 | 21.17 | 25.886 | 67.1746 | 86.7471 | 65.657 | 0.006472 | 1.042 | 0.025278 | 24.2589 | -0.000847 | 0.13917 | -8e-06 | -0.000648 | -0.000648 | -8.8687
2019-03-28 15:35 | 1553801717.2727 | -184 | 0.20948 | 0.000453 | 0 | 101.1569 | 101.2695 | 126.8884 | 21.2612 | 0.001869 | 0.001728 | -151.9565 | 0.23407 | 8.4026 | 0.00266 | 0.004515 | 0.004339 | 4.7e-05 | 0.001261 | 0.31885 | 1.3361 | 2.5475 | 100.63 | 21.17 | 25.888 | 66.8079 | 85.8189 | 63.4014 | 0.006473 | 1.049 | 0.025359 | 24.1744 | -0.009955 | 0.13925 | -8e-06 | -0.000634 | -0.000634 | -9.0653
2019-03-28 15:35 | 1553801718.2731 | -183 | 0.20949 | 0.000449 | 0 | 101.2036 | 101.3382 | 126.9398 | 21.2612 | 0.001917 | 0.001769 | -151.9564 | 0.23397 | 8.4048 | 0.002754 | 0.0042 | 0.003982 | 4.8e-05 | 0.001254 | 0.30071 | 1.3106 | 2.5475 | 100.63 | 21.17 | 25.894 | 67.9586 | 88.0684 | 64.387 | 0.006474 | 1.0459 | 0.025327 | 24.2147 | -0.011881 | 0.13919 | -3e-06 | -0.000216 | -0.000216 | -26.6691
2019-03-28 15:35 | 1553801719.2726 | -182 | 0.20948 | 0.000448 | 0 | 101.2051 | 101.313 | 127.2039 | 21.2612 | 0.001885 | 0.001759 | -151.9548 | 0.23357 | 8.402 | 0.002646 | 0.004349 | 0.004128 | 4.7e-05 | 0.001254 | 0.32621 | 1.3745 | 2.5474 | 100.63 | 21.17 | 25.896 | 67.1825 | 85.4812 | 64.1574 | 0.006475 | 1.0467 | 0.025362 | 24.2316 | -0.00953 | 0.13895 | -3e-06 | -0.000223 | -0.000223 | -25.7031
2019-03-28 15:35 | 1553801720.274 | -181 | 0.20949 | 0.000451 | 0 | 101.1197 | 101.1085 | 127.2347 | 21.2612 | 0.001883 | 0.001719 | -151.9398 | 0.23393 | 8.403 | 0.002667 | 0.004326 | 0.00413 | 4.6e-05 | 0.001256 | 0.32421 | 1.3127 | 2.5475 | 100.63 | 21.17 | 25.9 | 67.1308 | 85.9851 | 63.1884 | 0.006476 | 1.0497 | 0.025402 | 24.1997 | 0.000986 | 0.13917 | -6e-06 | -0.00046 | -0.00046 | -12.501
2019-03-28 15:35 | 1553801721.2737 | -180 | 0.20951 | 0.000447 | 0 | 101.1158 | 101.1041 | 127.0194 | 21.2612 | 0.001851 | 0.001749 | -151.958 | 0.2337 | 8.4039 | 0.002791 | 0.00417 | 0.003965 | 4.4e-05 | 0.001258 | 0.32244 | 1.2926 | 2.5474 | 100.63 | 21.17 | 25.9 | 66.3615 | 88.9728 | 63.9069 | 0.006476 | 1.0474 | 0.025353 | 24.2051 | 0.001034 | 0.13903 | -2e-06 | -0.000133 | -0.000133 | -43.1565
2019-03-28 15:35 | 1553801722.2738 | -179 | 0.20949 | 0.000448 | 0 | 101.1542 | 101.2624 | 126.4634 | 21.2612 | 0.00185 | 0.001757 | -151.9639 | 0.23338 | 8.402 | 0.002711 | 0.003555 | 0.003358 | 4.7e-05 | 0.001255 | 0.31045 | 1.2974 | 2.5474 | 100.63 | 21.17 | 25.899 | 66.35 | 87.0407 | 64.0934 | 0.006476 | 1.0469 | 0.025291 | 24.1587 | -0.009562 | 0.13884 | -2e-06 | -0.000172 | -0.000172 | -33.3555
2019-03-28 15:35 | 1553801723.2733 | -178 | 0.20947 | 0.000447 | 0 | 101.2251 | 101.2169 | 126.143 | 21.2612 | 0.001811 | 0.001807 | -151.9724 | 0.233 | 8.4033 | 0.002686 | 0.004074 | 0.003861 | 5e-05 | 0.001258 | 0.27992 | 1.3047 | 2.5473 | 100.63 | 21.17 | 25.898 | 65.4025 | 86.4295 | 65.2986 | 0.006475 | 1.0431 | 0.025214 | 24.1712 | 0.000726 | 0.13861 | -2e-06 | -0.000165 | -0.000165 | -34.6076
2019-03-28 15:35 | 1553801724.2729 | -177 | 0.2095 | 0.00045 | 0 | 101.2042 | 101.4367 | 126.3427 | 21.2612 | 0.001896 | 0.001727 | -151.9764 | 0.23389 | 8.4029 | 0.002673 | 0.003797 | 0.003591 | 4.8e-05 | 0.001262 | 0.31852 | 1.3355 | 2.5474 | 100.63 | 21.17 | 25.896 | 67.4509 | 86.1218 | 63.3858 | 0.006475 | 1.0491 | 0.025305 | 24.1218 | -0.02052 | 0.13914 | -6e-06 | -0.000497 | -0.000497 | -11.5636
2019-03-28 15:35 | 1553801725.273 | -176 | 0.20946 | 0.000451 | 0 | 101.1882 | 101.2386 | 126.5603 | 21.2612 | 0.001834 | 0.001781 | -151.9956 | 0.23332 | 8.4021 | 0.002729 | 0.003963 | 0.0039 | 4.8e-05 | 0.001261 | 0.30431 | 1.2866 | 2.5475 | 100.63 | 21.17 | 25.895 | 65.9668 | 87.4679 | 64.6829 | 0.006475 | 1.045 | 0.025278 | 24.1891 | -0.00446 | 0.1388 | -7e-06 | -0.000572 | -0.000572 | -10.0305
2019-03-28 15:35 | 1553801726.2744 | -175 | 0.20947 | 0.00045 | 0 | 101.1688 | 101.2726 | 126.8633 | 21.2612 | 0.001897 | 0.001785 | -152.0023 | 0.23402 | 8.4024 | 0.002775 | 0.003948 | 0.003676 | 5e-05 | 0.001259 | 0.30198 | 1.2949 | 2.5473 | 100.63 | 21.17 | 25.888 | 67.4648 | 88.5751 | 64.7717 | 0.006473 | 1.0448 | 0.025305 | 24.2212 | -0.009172 | 0.13922 | -5e-06 | -0.000401 | -0.000401 | -14.3408
2019-03-28 15:35 | 1553801727.274 | -174 | 0.20951 | 0.00045 | 0 | 101.1581 | 101.3392 | 126.8659 | 21.2612 | 0.001842 | 0.001749 | -152.0103 | 0.23356 | 8.4022 | 0.002694 | 0.003897 | 0.003764 | 4.9e-05 | 0.001259 | 0.32493 | 1.3327 | 2.5473 | 100.63 | 21.17 | 25.884 | 66.1487 | 86.6299 | 63.9191 | 0.006472 | 1.0474 | 0.025337 | 24.1909 | -0.015999 | 0.13894 | -5e-06 | -0.000371 | -0.000371 | -15.4588
2019-03-28 15:35 | 1553801728.2734 | -173 | 0.20948 | 0.000447 | 0 | 101.2084 | 101.3844 | 126.4413 | 21.2612 | 0.001861 | 0.001755 | -152.0155 | 0.23351 | 8.4028 | 0.002643 | 0.003818 | 0.003648 | 4.9e-05 | 0.001262 | 0.30829 | 1.3299 | 2.5473 | 100.63 | 21.17 | 25.873 | 66.6114 | 85.3929 | 64.045 | 0.006469 | 1.047 | 0.02529 | 24.1549 | -0.015537 | 0.13891 | -2e-06 | -0.000141 | -0.000141 | -40.7602
2019-03-28 15:35 | 1553801729.274 | -172 | 0.20947 | 0.000446 | 0 | 101.1902 | 101.172 | 126.0284 | 21.2612 | 0.001858 | 0.0017 | -152.0156 | 0.23334 | 8.4032 | 0.002656 | 0.003913 | 0.003744 | 4.7e-05 | 0.001255 | 0.28351 | 1.3212 | 2.5476 | 100.63 | 21.17 | 25.867 | 66.5457 | 85.7152 | 62.7399 | 0.006468 | 1.0511 | 0.025298 | 24.0687 | 0.00161 | 0.13882 | -1e-06 | -7.4e-05 | -7.4e-05 | -77.1066
2019-03-28 15:35 | 1553801730.2741 | -171 | 0.20948 | 0.000447 | 0 | 101.1575 | 101.4037 | 126.4211 | 21.2612 | 0.001863 | 0.00172 | -152.0214 | 0.23426 | 8.4017 | 0.00273 | 0.003536 | 0.003381 | 4.6e-05 | 0.001246 | 0.31238 | 1.328 | 2.5475 | 100.63 | 21.17 | 25.861 | 66.6511 | 87.4989 | 63.2093 | 0.006466 | 1.0496 | 0.02532 | 24.123 | -0.021747 | 0.13936 | -1e-06 | -7e-05 | -7e-05 | -82.154
2019-03-28 15:36 | 1553801731.2737 | -170 | 0.2095 | 0.000451 | 0 | 101.2284 | 101.2977 | 126.8065 | 21.2612 | 0.001809 | 0.001724 | -152.0199 | 0.23418 | 8.4035 | 0.002675 | 0.003566 | 0.003399 | 4.9e-05 | 0.001241 | 0.28368 | 1.3528 | 2.5474 | 100.63 | 21.17 | 25.85 | 65.3667 | 86.1632 | 63.3008 | 0.006463 | 1.0493 | 0.025355 | 24.163 | -0.006127 | 0.13932 | -6e-06 | -0.000494 | -0.000494 | -11.6385
2019-03-28 15:36 | 1553801732.2731 | -169 | 0.20943 | 0.000447 | 0 | 101.187 | 101.3698 | 127.3166 | 21.2612 | 0.001857 | 0.001744 | -152.0131 | 0.23436 | 8.4038 | 0.002661 | 0.004226 | 0.004019 | 4.7e-05 | 0.001237 | 0.29789 | 1.3338 | 2.5473 | 100.63 | 21.17 | 25.845 | 66.5005 | 85.83 | 63.7959 | 0.006462 | 1.0478 | 0.025387 | 24.2294 | -0.016143 | 0.13942 | -1e-06 | -6e-05 | -6e-05 | -95.9054
2019-03-28 15:36 | 1553801733.2721 | -168 | 0.2095 | 0.000444 | 0 | 101.1902 | 101.2257 | 127.5081 | 21.2612 | 0.001895 | 0.001721 | -152.0252 | 0.23388 | 8.4031 | 0.002664 | 0.003597 | 0.003401 | 4.7e-05 | 0.001238 | 0.30663 | 1.2857 | 2.5475 | 100.63 | 21.17 | 25.838 | 67.4234 | 85.9026 | 63.2298 | 0.00646 | 1.0495 | 0.025428 | 24.2272 | -0.003135 | 0.13914 | -5.6e-05 | -0.00425 | -0.00425 | -1.3522
2019-03-28 15:36 | 1553801734.2732 | -167 | 0.20947 | 0.000448 | 0 | 101.2108 | 101.3147 | 127.4084 | 21.2612 | 0.00184 | 0.001715 | -152.0085 | 0.23534 | 8.4032 | 0.002662 | 0.004012 | 0.003893 | 4.5e-05 | 0.001237 | 0.3235 | 1.3166 | 2.5474 | 100.63 | 21.17 | 25.835 | 66.0931 | 85.8596 | 63.0938 | 0.00646 | 1.05 | 0.025423 | 24.2128 | -0.00918 | 0.14 | 0.00012 | 0.009204 | 0.009204 | 0.62826
2019-03-28 15:36 | 1553801735.2733 | -166 | 0.20947 | 0.000444 | 0 | 101.1837 | 101.2498 | 127.4764 | 21.2612 | 0.001885 | 0.00176 | -151.9963 | 0.23308 | 8.402 | 0.002608 | 0.0041 | 0.003835 | 4.6e-05 | 0.001228 | 0.28136 | 1.3026 | 2.5474 | 100.63 | 21.17 | 25.834 | 67.1878 | 84.5602 | 64.1855 | 0.006459 | 1.0466 | 0.025388 | 24.2586 | -0.005845 | 0.13866 | 0.000131 | 0.010013 | 0.010013 | 0.57193
2019-03-28 15:36 | 1553801736.2732 | -165 | 0.20949 | 0.000444 | 0 | 101.2185 | 101.2838 | 127.0415 | 21.2612 | 0.00189 | 0.001658 | -151.9963 | 0.23355 | 8.4014 | 0.002687 | 0.004035 | 0.003932 | 4.6e-05 | 0.001227 | 0.26416 | 1.3235 | 2.5474 | 100.63 | 21.17 | 25.833 | 67.3112 | 86.4678 | 61.7242 | 0.006459 | 1.0543 | 0.025438 | 24.1287 | -0.005766 | 0.13894 | -2.1e-05 | -0.001602 | -0.001602 | -3.5824
2019-03-28 15:36 | 1553801737.2738 | -164 | 0.20948 | 0.000447 | 0 | 101.1926 | 101.2481 | 126.864 | 21.2612 | 0.001896 | 0.001718 | -152.0103 | 0.23356 | 8.4011 | 0.00262 | 0.003721 | 0.003496 | 4.7e-05 | 0.001232 | 0.33644 | 1.3277 | 2.5476 | 100.63 | 21.17 | 25.832 | 67.4571 | 84.8529 | 63.1665 | 0.006459 | 1.0497 | 0.025366 | 24.1637 | -0.004905 | 0.13895 | 5.2e-05 | 0.003998 | 0.003998 | 1.4352
2019-03-28 15:36 | 1553801738.2724 | -163 | 0.20948 | 0.000445 | 0 | 101.2433 | 101.3674 | 126.8253 | 21.2612 | 0.001841 | 0.001762 | -152.003 | 0.23353 | 8.403 | 0.002613 | 0.004095 | 0.003882 | 5e-05 | 0.001237 | 0.27671 | 1.3006 | 2.5474 | 100.63 | 21.17 | 25.833 | 66.1229 | 84.6735 | 64.2176 | 0.006459 | 1.0465 | 0.025322 | 24.1977 | -0.010959 | 0.13893 | 3.5e-05 | 0.002678 | 0.002678 | 2.1422
2019-03-28 15:36 | 1553801739.2724 | -162 | 0.20948 | 0.000449 | 0 | 101.195 | 101.3742 | 126.8544 | 21.2612 | 0.001905 | 0.001752 | -151.9916 | 0.23414 | 8.4055 | 0.002656 | 0.004235 | 0.004085 | 4.5e-05 | 0.001243 | 0.32117 | 1.3045 | 2.5474 | 100.63 | 21.16 | 25.841 | 67.6598 | 85.7145 | 63.981 | 0.006457 | 1.0472 | 0.025334 | 24.192 | -0.015822 | 0.13929 | 2.9e-05 | 0.002219 | 0.002219 | 2.592
2019-03-28 15:36 | 1553801740.2738 | -161 | 0.20948 | 0.000449 | 0 | 101.2391 | 101.081 | 126.5629 | 21.2612 | 0.001908 | 0.001759 | -151.9828 | 0.23375 | 8.4036 | 0.002652 | 0.003614 | 0.003382 | 5.3e-05 | 0.001249 | 0.30574 | 1.3029 | 2.5474 | 100.63 | 21.16 | 25.846 | 67.7456 | 85.6075 | 64.1528 | 0.006458 | 1.0467 | 0.025298 | 24.1703 | 0.013979 | 0.13906 | 3e-05 | 0.002258 | 0.002258 | 2.5433
2019-03-28 15:36 | 1553801741.2725 | -160 | 0.20949 | 0.000447 | 0 | 101.1429 | 101.2325 | 126.6962 | 21.2612 | 0.001881 | 0.001775 | -151.9776 | 0.23318 | 8.4 | 0.002657 | 0.00424 | 0.004075 | 5.2e-05 | 0.001252 | 0.28744 | 1.3225 | 2.5475 | 100.63 | 21.16 | 25.861 | 67.0943 | 85.7354 | 64.5275 | 0.006462 | 1.0455 | 0.025298 | 24.1965 | -0.007922 | 0.13872 | -2.6e-05 | -0.001952 | -0.001952 | -2.9346
2019-03-28 15:36 | 1553801742.2726 | -159 | 0.2095 | 0.000444 | 0 | 101.187 | 101.2349 | 126.8349 | 21.2612 | 0.001919 | 0.001784 | -151.958 | 0.23423 | 8.4032 | 0.002676 | 0.004198 | 0.004007 | 4.9e-05 | 0.001249 | 0.33682 | 1.3091 | 2.5472 | 100.63 | 21.16 | 25.867 | 67.9921 | 86.2062 | 64.7647 | 0.006464 | 1.0448 | 0.025303 | 24.2182 | -0.004235 | 0.13934 | -6.4e-05 | -0.004881 | -0.004881 | -1.1791
2019-03-28 15:36 | 1553801743.2735 | -158 | 0.20948 | 0.000443 | 0 | 101.1241 | 101.2468 | 126.8337 | 21.2612 | 0.001899 | 0.001734 | -151.9352 | 0.23347 | 8.4014 | 0.002647 | 0.00447 | 0.004232 | 5e-05 | 0.001246 | 0.33699 | 1.3368 | 2.5474 | 100.63 | 21.17 | 25.882 | 67.5312 | 85.4927 | 63.542 | 0.006471 | 1.0486 | 0.025348 | 24.1743 | -0.010843 | 0.13889 | 2.5e-05 | 0.001893 | 0.001893 | 3.0302
2019-03-28 15:36 | 1553801744.2726 | -157 | 0.20949 | 0.000445 | 0 | 101.1709 | 101.3137 | 126.6439 | 21.2612 | 0.00189 | 0.001757 | -151.9412 | 0.23438 | 8.4029 | 0.002684 | 0.0031 | 0.003026 | 5.1e-05 | 0.001247 | 0.32853 | 1.3393 | 2.5474 | 100.63 | 21.16 | 25.889 | 67.3112 | 86.3763 | 64.0965 | 0.006469 | 1.0468 | 0.025309 | 24.176 | -0.012618 | 0.13943 | -3.1e-05 | -0.002352 | -0.002352 | -2.4487
2019-03-28 15:36 | 1553801745.274 | -156 | 0.20953 | 0.000445 | 0 | 101.2716 | 101.3378 | 126.2979 | 21.2612 | 0.001857 | 0.001761 | -151.9122 | 0.23323 | 8.4029 | 0.002636 | 0.003953 | 0.003752 | 4.9e-05 | 0.001256 | 0.32621 | 1.3539 | 2.5474 | 100.63 | 21.16 | 25.896 | 66.502 | 85.2385 | 64.1887 | 0.006471 | 1.0466 | 0.025271 | 24.1463 | -0.005851 | 0.13875 | -0.00024 | -0.018255 | -0.018255 | -0.3139
2019-03-28 15:36 | 1553801746.2736 | -155 | 0.20949 | 0.000442 | 0 | 101.1807 | 101.2142 | 126.2084 | 21.2612 | 0.001801 | 0.001731 | -151.9197 | 0.23381 | 8.4053 | 0.002681 | 0.003929 | 0.003696 | 5.1e-05 | 0.001256 | 0.2912 | 1.382 | 2.5474 | 100.63 | 21.16 | 25.909 | 65.1598 | 86.3116 | 63.4734 | 0.006474 | 1.0488 | 0.025288 | 24.1121 | -0.002956 | 0.13909 | -6e-06 | -0.00042 | -0.00042 | -13.6834
2019-03-28 15:36 | 1553801747.2721 | -154 | 0.20948 | 0.000443 | 0 | 101.2153 | 101.314 | 126.3512 | 21.2612 | 0.001865 | 0.001748 | -151.9227 | 0.23359 | 8.4008 | 0.002638 | 0.004411 | 0.004281 | 5.3e-05 | 0.001261 | 0.28567 | 1.3085 | 2.5475 | 100.63 | 21.16 | 25.915 | 66.7112 | 85.2914 | 63.8777 | 0.006476 | 1.0475 | 0.025288 | 24.1403 | -0.008725 | 0.13896 | 5.4e-05 | 0.004112 | 0.004112 | 1.3956
2019-03-28 15:36 | 1553801748.2727 | -153 | 0.2095 | 0.000446 | 0 | 101.2138 | 101.3151 | 126.7453 | 21.2612 | 0.001857 | 0.001716 | -151.9272 | 0.23331 | 8.4017 | 0.00258 | 0.00381 | 0.003624 | 5.2e-05 | 0.001263 | 0.30326 | 1.3253 | 2.5473 | 100.63 | 21.16 | 25.922 | 66.5099 | 83.8869 | 63.1133 | 0.006477 | 1.0499 | 0.025356 | 24.1504 | -0.008947 | 0.1388 | -6.8e-05 | -0.005205 | -0.005205 | -1.1014
2019-03-28 15:36 | 1553801749.2726 | -152 | 0.20951 | 0.000444 | -1e-06 | 101.1956 | 101.0103 | 127.287 | 21.2612 | 0.001852 | 0.001772 | -151.9385 | 0.23435 | 8.4055 | 0.002673 | 0.004207 | 0.003968 | 5.3e-05 | 0.001263 | 0.28119 | 1.3012 | 2.5475 | 100.63 | 21.16 | 25.933 | 66.3889 | 86.1124 | 64.4659 | 0.00648 | 1.0457 | 0.025359 | 24.2506 | 0.016392 | 0.13941 | -0.000118 | -0.008937 | -0.008937 | -0.64424
2019-03-28 15:36 | 1553801750.2732 | -151 | 0.20948 | 0.000442 | 0 | 101.2081 | 101.4625 | 127.499 | 21.2612 | 0.00186 | 0.00172 | -151.9393 | 0.23383 | 8.4033 | 0.00268 | 0.004046 | 0.003901 | 5.1e-05 | 0.001265 | 0.31382 | 1.2966 | 2.5473 | 100.63 | 21.16 | 25.937 | 66.5763 | 86.2873 | 63.2024 | 0.006481 | 1.0496 | 0.025428 | 24.2254 | -0.022454 | 0.1391 | 7.1e-05 | 0.005477 | 0.005477 | 1.0489
2019-03-28 15:36 | 1553801751.2718 | -150 | 0.20946 | 0.000444 | 0 | 101.2483 | 101.3144 | 127.014 | 21.2612 | 0.001937 | 0.001825 | -151.9338 | 0.23362 | 8.4041 | 0.002678 | 0.003637 | 0.003442 | 5.3e-05 | 0.001265 | 0.32853 | 1.3159 | 2.5474 | 100.63 | 21.16 | 25.943 | 68.4344 | 86.2467 | 65.752 | 0.006483 | 1.0417 | 0.025284 | 24.2707 | -0.005834 | 0.13898 | 0.000143 | 0.010926 | 0.010926 | 0.52537
2019-03-28 15:36 | 1553801752.2729 | -149 | 0.20951 | 0.000446 | 0 | 101.1944 | 101.1853 | 126.8323 | 21.2612 | 0.001799 | 0.001772 | -151.953 | 0.23364 | 8.4041 | 0.002676 | 0.003375 | 0.003168 | 5.3e-05 | 0.001266 | 0.30381 | 1.3288 | 2.5473 | 100.63 | 21.16 | 25.945 | 65.1114 | 86.1969 | 64.4705 | 0.006483 | 1.0457 | 0.025313 | 24.2074 | 0.000807 | 0.13899 | -0.000151 | -0.01147 | -0.01147 | -0.50048
2019-03-28 15:36 | 1553801753.3343 | -148 | 0.20949 | 0.000443 | 0 | 101.1575 | 101.2335 | 126.9678 | 21.2612 | 0.001868 | 0.001772 | -151.9533 | 0.2339 | 8.4031 | 0.002645 | 0.003777 | 0.003613 | 5.1e-05 | 0.00126 | 0.3063 | 1.3446 | 2.5474 | 100.63 | 21.16 | 25.946 | 66.7845 | 85.4552 | 64.4642 | 0.006483 | 1.0457 | 0.025327 | 24.2201 | -0.006722 | 0.13915 | -2.2e-05 | -0.001712 | -0.001712 | -3.3559
2019-03-28 15:36 | 1553801754.2817 | -147 | 0.20953 | 0.000444 | 0 | 101.1929 | 101.1717 | 127.0403 | 21.2612 | 0.001944 | 0.00172 | -151.9559 | 0.23366 | 8.403 | 0.00261 | 0.003778 | 0.003601 | 5.2e-05 | 0.001257 | 0.31277 | 1.28 | 2.5474 | 100.63 | 21.15 | 25.948 | 68.595 | 84.6134 | 63.2118 | 0.00648 | 1.0496 | 0.025382 | 24.1821 | 0.001877 | 0.13901 | -0.000244 | -0.018671 | -0.018671 | -0.30748
2019-03-28 15:36 | 1553801755.2818 | -146 | 0.20948 | 0.000449 | 0 | 101.1962 | 101.2675 | 127.5574 | 21.2612 | 0.001825 | 0.001752 | -151.9803 | 0.23418 | 8.4043 | 0.002644 | 0.003898 | 0.003735 | 5.1e-05 | 0.001253 | 0.31022 | 1.3779 | 2.5474 | 100.63 | 21.15 | 25.949 | 65.7335 | 85.4218 | 63.988 | 0.00648 | 1.0472 | 0.025404 | 24.2592 | -0.0063 | 0.13932 | 3.8e-05 | 0.002938 | 0.002938 | 1.9587
2019-03-28 15:36 | 1553801756.2817 | -145 | 0.20949 | 0.000448 | -1e-06 | 101.2156 | 101.136 | 127.9649 | 21.2612 | 0.001889 | 0.001762 | -151.9926 | 0.23431 | 8.4009 | 0.002634 | 0.003649 | 0.003357 | 5.1e-05 | 0.001252 | 0.31974 | 1.3633 | 2.5473 | 100.63 | 21.15 | 25.95 | 67.2799 | 85.1899 | 64.2291 | 0.00648 | 1.0464 | 0.025435 | 24.3066 | 0.007031 | 0.13939 | -1.7e-05 | -0.001203 | -0.001203 | -4.7861
2019-03-28 15:36 | 1553801757.2832 | -144 | 0.20949 | 0.000443 | -1e-06 | 101.1503 | 101.2546 | 128.01 | 21.2612 | 0.001852 | 0.00174 | -151.9711 | 0.23349 | 8.4042 | 0.002638 | 0.004225 | 0.004085 | 4.7e-05 | 0.001256 | 0.30807 | 1.3383 | 2.5473 | 100.63 | 21.15 | 25.949 | 66.3952 | 85.2765 | 63.7037 | 0.00648 | 1.0481 | 0.02546 | 24.2919 | -0.009215 | 0.1389 | -4.6e-05 | -0.003466 | -0.003466 | -1.655
2019-03-28 15:36 | 1553801758.2819 | -143 | 0.20947 | 0.000447 | 0 | 101.1423 | 101.2702 | 127.6731 | 21.2612 | 0.001834 | 0.001756 | -151.9771 | 0.2348 | 8.403 | 0.002711 | 0.003743 | 0.003424 | 5.1e-05 | 0.001261 | 0.29767 | 1.3505 | 2.5474 | 100.63 | 21.15 | 25.948 | 65.9488 | 87.0258 | 64.0864 | 0.00648 | 1.0469 | 0.025412 | 24.2737 | -0.011306 | 0.13968 | 0.000119 | 0.009128 | 0.009128 | 0.63202
2019-03-28 15:36 | 1553801759.2832 | -142 | 0.2095 | 0.000444 | 0 | 101.2185 | 101.4279 | 127.1164 | 21.2612 | 0.001869 | 0.001764 | -151.9787 | 0.23398 | 8.4041 | 0.002679 | 0.003626 | 0.003482 | 5.1e-05 | 0.001261 | 0.28727 | 1.3788 | 2.5476 | 100.63 | 21.15 | 25.943 | 66.7869 | 86.2741 | 64.2722 | 0.006479 | 1.0463 | 0.025349 | 24.2274 | -0.018477 | 0.1392 | -0.000108 | -0.008239 | -0.008239 | -0.69778
2019-03-28 15:36 | 1553801760.2829 | -141 | 0.20949 | 0.000442 | 0 | 101.2495 | 101.4265 | 126.7784 | 21.2612 | 0.001819 | 0.001751 | -151.9886 | 0.23458 | 8.4007 | 0.002762 | 0.004073 | 0.003908 | 4.8e-05 | 0.001262 | 0.28313 | 1.3503 | 2.5474 | 100.63 | 21.15 | 25.942 | 65.6 | 88.2713 | 63.9615 | 0.006478 | 1.0473 | 0.025327 | 24.184 | -0.015621 | 0.13955 | 1.3e-05 | 0.000988 | 0.000988 | 5.8341
2019-03-28 15:36 | 1553801761.282 | -140 | 0.20947 | 0.000445 | -1e-06 | 101.2066 | 101.5851 | 127.1293 | 21.2612 | 0.001854 | 0.001749 | -151.9864 | 0.23355 | 8.403 | 0.002698 | 0.004047 | 0.003828 | 5e-05 | 0.001271 | 0.30198 | 1.3409 | 2.5474 | 100.63 | 21.15 | 25.941 | 66.4279 | 86.7308 | 63.9138 | 0.006478 | 1.0474 | 0.025364 | 24.2158 | -0.03339 | 0.13894 | 8.1e-05 | 0.006251 | 0.006251 | 0.91795
2019-03-28 15:36 | 1553801762.2814 | -139 | 0.2095 | 0.000442 | 0 | 101.1301 | 101.3076 | 126.9262 | 21.2612 | 0.001858 | 0.001737 | -151.9922 | 0.2336 | 8.4028 | 0.002667 | 0.00469 | 0.004534 | 5.1e-05 | 0.001266 | 0.3183 | 1.3122 | 2.5473 | 100.63 | 21.15 | 25.937 | 66.5412 | 85.9736 | 63.6171 | 0.006477 | 1.0483 | 0.025355 | 24.1858 | -0.015686 | 0.13897 | -9.2e-05 | -0.007012 | -0.007012 | -0.8185
2019-03-28 15:36 | 1553801763.282 | -138 | 0.20949 | 0.000447 | -1e-06 | 101.145 | 101.1642 | 126.669 | 21.2612 | 0.001845 | 0.001724 | -151.9976 | 0.23441 | 8.4046 | 0.002764 | 0.003442 | 0.003278 | 4.9e-05 | 0.001264 | 0.29407 | 1.2953 | 2.5474 | 100.63 | 21.15 | 25.933 | 66.2321 | 88.3089 | 63.3085 | 0.006476 | 1.0493 | 0.025341 | 24.1502 | -0.001702 | 0.13945 | -1.1e-05 | -0.000822 | -0.000822 | -7.0035
2019-03-28 15:36 | 1553801764.2811 | -137 | 0.20949 | 0.000447 | 0 | 101.22 | 101.2417 | 126.5479 | 21.2612 | 0.001877 | 0.001755 | -152.0079 | 0.23422 | 8.4039 | 0.002746 | 0.004427 | 0.004221 | 4.8e-05 | 0.001265 | 0.3011 | 1.297 | 2.5474 | 100.63 | 21.15 | 25.929 | 66.9921 | 87.8715 | 64.0505 | 0.006475 | 1.047 | 0.025301 | 24.1652 | -0.001913 | 0.13934 | -1e-06 | -2.8e-05 | -2.8e-05 | -202.6022
2019-03-28 15:36 | 1553801765.2811 | -136 | 0.2095 | 0.000446 | 0 | 101.2236 | 101.5111 | 126.1838 | 21.2612 | 0.001835 | 0.001773 | -152.0118 | 0.23402 | 8.3986 | 0.002615 | 0.00455 | 0.004338 | 5e-05 | 0.00127 | 0.34307 | 1.2973 | 2.5474 | 100.63 | 21.15 | 25.927 | 65.97 | 84.7389 | 64.483 | 0.006475 | 1.0456 | 0.025248 | 24.1459 | -0.025367 | 0.13922 | -0.000105 | -0.007984 | -0.007984 | -0.72012
2019-03-28 15:36 | 1553801766.2825 | -135 | 0.20949 | 0.000447 | 0 | 101.2376 | 101.3463 | 126.1139 | 21.2612 | 0.00186 | 0.001787 | -152.0066 | 0.23324 | 8.4024 | 0.002716 | 0.003527 | 0.003416 | 4.8e-05 | 0.001273 | 0.30453 | 1.2965 | 2.5473 | 100.63 | 21.15 | 25.924 | 66.5722 | 87.1569 | 64.8304 | 0.006474 | 1.0446 | 0.025228 | 24.1516 | -0.0096 | 0.13875 | -1e-05 | -0.000764 | -0.000764 | -7.5037
2019-03-28 15:36 | 1553801767.2824 | -134 | 0.20947 | 0.000445 | 0 | 101.2236 | 101.3039 | 126.2775 | 21.2612 | 0.001889 | 0.001741 | -152.011 | 0.23358 | 8.4023 | 0.002704 | 0.004248 | 0.00399 | 4.7e-05 | 0.00127 | 0.30397 | 1.2918 | 2.5473 | 100.63 | 21.15 | 25.919 | 67.2885 | 86.8706 | 63.7256 | 0.006473 | 1.048 | 0.025286 | 24.1278 | -0.007088 | 0.13896 | 0.000101 | 0.007721 | 0.007721 | 0.74327
2019-03-28 15:36 | 1553801768.2812 | -133 | 0.2095 | 0.000448 | 0 | 101.2519 | 101.2665 | 126.5443 | 21.2612 | 0.001836 | 0.001779 | -151.9978 | 0.23364 | 8.4023 | 0.002636 | 0.003827 | 0.003661 | 5e-05 | 0.00127 | 0.30287 | 1.3254 | 2.5474 | 100.63 | 21.15 | 25.917 | 66.0104 | 85.243 | 64.6322 | 0.006472 | 1.0452 | 0.025279 | 24.1857 | -0.001288 | 0.13899 | -6.6e-05 | -0.005013 | -0.005013 | -1.1451
2019-03-28 15:36 | 1553801769.2822 | -132 | 0.20949 | 0.000441 | 0 | 101.2835 | 101.6718 | 126.3575 | 21.2612 | 0.001864 | 0.00177 | -151.9918 | 0.23392 | 8.3998 | 0.00272 | 0.004671 | 0.004454 | 4.8e-05 | 0.00127 | 0.32372 | 1.3178 | 2.5474 | 100.63 | 21.15 | 25.917 | 66.6713 | 87.2476 | 64.4214 | 0.006472 | 1.0458 | 0.025268 | 24.1603 | -0.034226 | 0.13916 | -1.6e-05 | -0.001262 | -0.001262 | -4.5532
2019-03-28 15:36 | 1553801770.2812 | -131 | 0.2095 | 0.000446 | 0 | 101.2299 | 101.2247 | 126.257 | 21.2612 | 0.001866 | 0.001769 | -151.9691 | 0.23343 | 8.4035 | 0.002717 | 0.004194 | 0.004026 | 5.2e-05 | 0.001271 | 0.30934 | 1.3224 | 2.5474 | 100.63 | 21.15 | 25.918 | 66.7213 | 87.1711 | 64.3901 | 0.006472 | 1.0459 | 0.025259 | 24.1496 | 0.000456 | 0.13887 | -9.5e-05 | -0.007208 | -0.007208 | -0.79563
2019-03-28 15:36 | 1553801771.2817 | -130 | 0.20951 | 0.00045 | 0 | 101.2162 | 101.3888 | 125.8153 | 21.2612 | 0.00188 | 0.001767 | -151.9917 | 0.23316 | 8.4026 | 0.002722 | 0.003734 | 0.003571 | 4.8e-05 | 0.001264 | 0.32421 | 1.3247 | 2.5474 | 100.63 | 21.15 | 25.92 | 67.0512 | 87.3117 | 64.3362 | 0.006473 | 1.0461 | 0.025217 | 24.1054 | -0.015242 | 0.13871 | -0.000148 | -0.011238 | -0.011238 | -0.50974
2019-03-28 15:36 | 1553801772.2834 | -129 | 0.2095 | 0.000446 | 0 | 101.2099 | 101.186 | 125.9495 | 21.2612 | 0.001862 | 0.001754 | -151.9913 | 0.23268 | 8.4028 | 0.002695 | 0.004071 | 0.003895 | 4.9e-05 | 0.001259 | 0.29479 | 1.3696 | 2.5474 | 100.63 | 21.15 | 25.922 | 66.6323 | 86.6449 | 64.0308 | 0.006473 | 1.0471 | 0.025242 | 24.1073 | 0.002116 | 0.13842 | -9.5e-05 | -0.007197 | -0.007197 | -0.79437
2019-03-28 15:36 | 1553801773.2812 | -128 | 0.20951 | 0.000448 | 0 | 101.2653 | 101.3776 | 126.0871 | 21.2612 | 0.001878 | 0.001762 | -151.9883 | 0.23399 | 8.4011 | 0.002678 | 0.003792 | 0.003648 | 5e-05 | 0.001252 | 0.29264 | 1.3588 | 2.5473 | 100.63 | 21.15 | 25.924 | 67.0139 | 86.2483 | 64.2127 | 0.006474 | 1.0465 | 0.025249 | 24.127 | -0.009911 | 0.1392 | -0.000151 | -0.011442 | -0.011442 | -0.50247
2019-03-28 15:36 | 1553801774.2823 | -127 | 0.20949 | 0.00045 | 0 | 101.2084 | 101.2974 | 126.2422 | 21.2612 | 0.001874 | 0.001805 | -151.9815 | 0.23404 | 8.4031 | 0.002676 | 0.004068 | 0.00387 | 4.9e-05 | 0.001253 | 0.28871 | 1.3435 | 2.5474 | 100.63 | 21.15 | 25.926 | 66.9288 | 86.2046 | 65.2481 | 0.006474 | 1.0433 | 0.025225 | 24.1789 | -0.007861 | 0.13923 | -1.5e-05 | -0.001142 | -0.001142 | -5.0363
2019-03-28 15:36 | 1553801775.2823 | -126 | 0.20952 | 0.000448 | 0 | 101.2022 | 101.365 | 126.3735 | 21.2612 | 0.001883 | 0.001739 | -151.9728 | 0.23364 | 8.4014 | 0.002709 | 0.004294 | 0.00407 | 4.7e-05 | 0.001252 | 0.3178 | 1.3105 | 2.5473 | 100.63 | 21.15 | 25.933 | 67.1363 | 86.9969 | 63.6749 | 0.006476 | 1.0482 | 0.025297 | 24.1351 | -0.014381 | 0.13899 | -0.000221 | -0.016874 | -0.016874 | -0.34019
2019-03-28 15:36 | 1553801776.2823 | -125 | 0.20949 | 0.000442 | 0 | 101.1581 | 101.345 | 126.284 | 21.2612 | 0.001878 | 0.001754 | -151.9449 | 0.23313 | 8.4044 | 0.002748 | 0.003913 | 0.003834 | 4.5e-05 | 0.001253 | 0.32621 | 1.3458 | 2.5474 | 100.63 | 21.15 | 25.934 | 67.0161 | 87.9278 | 64.038 | 0.006476 | 1.047 | 0.025275 | 24.1396 | -0.016509 | 0.13869 | -2e-05 | -0.001553 | -0.001553 | -3.6891
2019-03-28 15:36 | 1553801777.2817 | -124 | 0.20953 | 0.000447 | 0 | 101.1274 | 101.2064 | 126.412 | 21.2612 | 0.00186 | 0.001765 | -151.9533 | 0.23306 | 8.4025 | 0.002675 | 0.004184 | 0.00385 | 4.7e-05 | 0.001248 | 0.3246 | 1.3281 | 2.5474 | 100.63 | 21.15 | 25.939 | 66.5857 | 86.1812 | 64.3042 | 0.006478 | 1.0462 | 0.025278 | 24.1613 | -0.00698 | 0.13865 | -0.000237 | -0.018018 | -0.018018 | -0.31779
2019-03-28 15:36 | 1553801778.2805 | -123 | 0.20952 | 0.000448 | 0 | 101.2245 | 101.2882 | 126.4763 | 21.2612 | 0.001835 | 0.001725 | -151.9684 | 0.23344 | 8.4015 | 0.002659 | 0.003774 | 0.003711 | 4.9e-05 | 0.001244 | 0.31813 | 1.3115 | 2.5475 | 100.63 | 21.15 | 25.942 | 65.9902 | 85.7807 | 63.3321 | 0.006478 | 1.0492 | 0.025321 | 24.1327 | -0.005629 | 0.13887 | -0.000179 | -0.013683 | -0.013683 | -0.41917
2019-03-28 15:36 | 1553801779.2804 | -122 | 0.20952 | 0.000444 | 0 | 101.2352 | 101.3259 | 126.8401 | 21.2612 | 0.001865 | 0.001766 | -151.9728 | 0.23425 | 8.4018 | 0.002677 | 0.004095 | 0.003886 | 4.8e-05 | 0.001243 | 0.29336 | 1.2886 | 2.5474 | 100.63 | 21.15 | 25.949 | 66.6917 | 86.2115 | 64.3268 | 0.00648 | 1.0461 | 0.02532 | 24.203 | -0.008011 | 0.13935 | -0.000179 | -0.013669 | -0.013669 | -0.42105
2019-03-28 15:36 | 1553801780.2815 | -121 | 0.20951 | 0.000446 | 0 | 101.173 | 101.2227 | 127.1178 | 21.2612 | 0.001824 | 0.001747 | -151.9923 | 0.23338 | 8.4024 | 0.002613 | 0.004273 | 0.004154 | 4.9e-05 | 0.001253 | 0.33035 | 1.3296 | 2.5474 | 100.63 | 21.15 | 25.953 | 65.7248 | 84.6764 | 63.8575 | 0.006481 | 1.0476 | 0.025365 | 24.2127 | -0.004392 | 0.13884 | -0.000113 | -0.008656 | -0.008656 | -0.66245
2019-03-28 15:36 | 1553801781.2815 | -120 | 0.20952 | 0.000444 | 0 | 101.187 | 101.3816 | 127.3562 | 21.2612 | 0.001839 | 0.001751 | -151.994 | 0.23375 | 8.4024 | 0.002601 | 0.003839 | 0.003605 | 4.7e-05 | 0.001256 | 0.3178 | 1.3484 | 2.5475 | 100.63 | 21.14 | 25.956 | 66.076 | 84.3926 | 63.9661 | 0.006478 | 1.0473 | 0.025385 | 24.2393 | -0.017193 | 0.13906 | -0.000202 | -0.01545 | -0.01545 | -0.37173
2019-03-28 15:36 | 1553801782.2818 | -119 | 0.20953 | 0.000446 | 0 | 101.1944 | 101.2559 | 127.3567 | 21.2612 | 0.001893 | 0.001757 | -151.9914 | 0.23418 | 8.4031 | 0.002602 | 0.003531 | 0.003409 | 4.9e-05 | 0.001262 | 0.30215 | 1.3104 | 2.5475 | 100.63 | 21.14 | 25.965 | 67.3674 | 84.4246 | 64.1121 | 0.00648 | 1.0468 | 0.025379 | 24.2446 | -0.005437 | 0.13931 | -0.000257 | -0.019635 | -0.019635 | -0.29302
2019-03-28 15:36 | 1553801783.2806 | -118 | 0.20952 | 0.000446 | 0 | 101.131 | 101.3388 | 127.1624 | 21.2612 | 0.001883 | 0.00174 | -151.9916 | 0.23436 | 8.3983 | 0.0026 | 0.004356 | 0.004104 | 4.9e-05 | 0.00126 | 0.3178 | 1.3229 | 2.5473 | 100.63 | 21.14 | 25.97 | 67.1339 | 84.362 | 63.7021 | 0.006481 | 1.0481 | 0.025375 | 24.2113 | -0.018366 | 0.13942 | -0.000218 | -0.016651 | -0.016651 | -0.34581
2019-03-28 15:36 | 1553801784.2806 | -117 | 0.20947 | 0.000446 | 0 | 101.1456 | 101.2386 | 127.2375 | 21.2612 | 0.001868 | 0.001752 | -151.9919 | 0.23375 | 8.4021 | 0.002654 | 0.003658 | 0.003493 | 5.1e-05 | 0.001263 | 0.29878 | 1.3195 | 2.5473 | 100.63 | 21.14 | 25.975 | 66.7783 | 85.67 | 63.9942 | 0.006483 | 1.0472 | 0.025372 | 24.229 | -0.008225 | 0.13906 | 7.2e-05 | 0.005516 | 0.005516 | 1.0411
2019-03-28 15:36 | 1553801785.2807 | -116 | 0.2095 | 0.000445 | 0 | 101.2465 | 101.133 | 127.3271 | 21.2612 | 0.001942 | 0.001719 | -151.989 | 0.23412 | 8.4026 | 0.002641 | 0.004115 | 0.003887 | 5e-05 | 0.001265 | 0.29535 | 1.3188 | 2.5473 | 100.63 | 21.15 | 25.986 | 68.5647 | 85.3491 | 63.1898 | 0.006489 | 1.0497 | 0.025411 | 24.2086 | 0.010039 | 0.13928 | -5.4e-05 | -0.004116 | -0.004116 | -1.3976
2019-03-28 15:36 | 1553801786.2817 | -115 | 0.20955 | 0.000444 | 0 | 101.1929 | 101.3093 | 127.2877 | 21.2612 | 0.001891 | 0.001755 | -151.9941 | 0.23487 | 8.4021 | 0.0026 | 0.004172 | 0.004052 | 4.9e-05 | 0.001268 | 0.3194 | 1.2908 | 2.5473 | 100.63 | 21.14 | 25.991 | 67.3306 | 84.3604 | 64.0474 | 0.006487 | 1.047 | 0.025375 | 24.2357 | -0.010279 | 0.13973 | -0.000364 | -0.027814 | -0.027814 | -0.20747
2019-03-28 15:36 | 1553801787.2807 | -114 | 0.20953 | 0.000444 | 0 | 101.2319 | 101.2648 | 127.1889 | 21.2612 | 0.001819 | 0.001772 | -151.9742 | 0.23331 | 8.4029 | 0.002689 | 0.004535 | 0.004326 | 5.2e-05 | 0.001271 | 0.31636 | 1.297 | 2.5474 | 100.63 | 21.14 | 25.991 | 65.5969 | 86.4988 | 64.4572 | 0.006487 | 1.0457 | 0.025349 | 24.241 | -0.002901 | 0.1388 | -0.000249 | -0.019058 | -0.019058 | -0.30078
2019-03-28 15:36 | 1553801788.2821 | -113 | 0.20953 | 0.00045 | 0 | 101.2302 | 101.3613 | 126.7964 | 21.2612 | 0.001886 | 0.001737 | -151.9588 | 0.23386 | 8.404 | 0.002714 | 0.004279 | 0.004111 | 4.6e-05 | 0.001273 | 0.29247 | 1.303 | 2.5474 | 100.63 | 21.14 | 26.007 | 67.1957 | 87.1023 | 63.6132 | 0.006491 | 1.0483 | 0.025342 | 24.1733 | -0.011577 | 0.13912 | -0.000247 | -0.018823 | -0.018823 | -0.30525
2019-03-28 15:36 | 1553801789.2808 | -112 | 0.20952 | 0.000448 | 0 | 101.1477 | 101.2862 | 126.3983 | 21.2612 | 0.001863 | 0.001768 | -151.961 | 0.23389 | 8.4022 | 0.002645 | 0.003682 | 0.003513 | 4.6e-05 | 0.001269 | 0.28423 | 1.3431 | 2.5474 | 100.63 | 21.14 | 26.01 | 66.6535 | 85.4576 | 64.3588 | 0.006491 | 1.046 | 0.025274 | 24.162 | -0.012242 | 0.13914 | -0.000202 | -0.015343 | -0.015343 | -0.37455
2019-03-28 15:36 | 1553801790.2821 | -111 | 0.20952 | 0.000444 | 0 | 101.1995 | 101.4031 | 126.4568 | 21.2612 | 0.001842 | 0.001786 | -151.9428 | 0.23411 | 8.4033 | 0.002646 | 0.00403 | 0.003767 | 4.6e-05 | 0.001272 | 0.30525 | 1.2849 | 2.5475 | 100.63 | 21.14 | 26.012 | 66.147 | 85.4614 | 64.8047 | 0.006492 | 1.0447 | 0.025263 | 24.1835 | -0.017975 | 0.13927 | -0.000172 | -0.013086 | -0.013086 | -0.43955
2019-03-28 15:36 | 1553801791.2819 | -110 | 0.20949 | 0.000448 | 1e-06 | 101.181 | 101.3123 | 126.6096 | 21.2612 | 0.001872 | 0.001713 | -151.9299 | 0.23472 | 8.402 | 0.002648 | 0.004196 | 0.003991 | 4.6e-05 | 0.00127 | 0.3011 | 1.3061 | 2.5474 | 100.63 | 21.14 | 26.013 | 66.8774 | 85.5317 | 63.0454 | 0.006492 | 1.0501 | 0.025345 | 24.1351 | -0.011602 | 0.13963 | 1e-05 | 0.000706 | 0.000706 | 8.1735
2019-03-28 15:37 | 1553801792.2809 | -109 | 0.2095 | 0.000448 | 0 | 101.1596 | 101.188 | 126.7201 | 21.2612 | 0.001891 | 0.00175 | -151.926 | 0.23421 | 8.4028 | 0.002696 | 0.00423 | 0.004046 | 4.8e-05 | 0.001272 | 0.29087 | 1.3399 | 2.5475 | 100.63 | 21.14 | 26.011 | 67.3174 | 86.687 | 63.9302 | 0.006492 | 1.0474 | 0.025322 | 24.1774 | -0.002514 | 0.13933 | -8.6e-05 | -0.006602 | -0.006602 | -0.87164
2019-03-28 15:37 | 1553801793.2813 | -108 | 0.20951 | 0.000448 | 0 | 101.1855 | 100.9869 | 127.0995 | 21.2612 | 0.001845 | 0.001755 | -151.9186 | 0.23334 | 8.404 | 0.002645 | 0.004082 | 0.00386 | 4.4e-05 | 0.001275 | 0.31116 | 1.3211 | 2.5473 | 100.63 | 21.14 | 26.009 | 66.2283 | 85.4468 | 64.0597 | 0.006491 | 1.047 | 0.025355 | 24.2182 | 0.017572 | 0.13882 | -0.000162 | -0.012405 | -0.012405 | -0.46217
2019-03-28 15:37 | 1553801794.2809 | -107 | 0.20948 | 0.000446 | 0 | 101.122 | 101.1839 | 127.0567 | 21.2612 | 0.001886 | 0.001749 | -151.9181 | 0.23449 | 8.4006 | 0.002688 | 0.004373 | 0.004208 | 4.4e-05 | 0.001277 | 0.31222 | 1.318 | 2.5474 | 100.63 | 21.14 | 25.999 | 67.2145 | 86.48 | 63.909 | 0.006489 | 1.0474 | 0.025357 | 24.2087 | -0.005473 | 0.1395 | 3.4e-05 | 0.002595 | 0.002595 | 2.2205
2019-03-28 15:37 | 1553801795.2799 | -106 | 0.20951 | 0.000446 | 0 | 101.17 | 101.4323 | 126.9861 | 21.2612 | 0.001832 | 0.001707 | -151.9313 | 0.23401 | 8.4029 | 0.002684 | 0.0043 | 0.004098 | 5.1e-05 | 0.001279 | 0.33826 | 1.3265 | 2.5474 | 100.63 | 21.14 | 25.994 | 65.9012 | 86.3819 | 62.9125 | 0.006487 | 1.0505 | 0.025387 | 24.1662 | -0.023158 | 0.13922 | -0.000156 | -0.011963 | -0.011963 | -0.48062
2019-03-28 15:37 | 1553801796.2813 | -105 | 0.20948 | 0.000449 | 0 | 101.2156 | 101.225 | 126.9736 | 21.2612 | 0.001869 | 0.001777 | -151.9444 | 0.23389 | 8.4014 | 0.002631 | 0.004108 | 0.003991 | 4.7e-05 | 0.001277 | 0.32173 | 1.3223 | 2.5474 | 100.63 | 21.14 | 25.989 | 66.8047 | 85.1206 | 64.5758 | 0.006486 | 1.0454 | 0.025324 | 24.2247 | -0.000837 | 0.13914 | 4.4e-05 | 0.003331 | 0.003331 | 1.725
2019-03-28 15:37 | 1553801797.2805 | -104 | 0.20948 | 0.000444 | 0 | 101.0824 | 101.2743 | 126.7577 | 21.2612 | 0.001879 | 0.001796 | -151.9645 | 0.23416 | 8.4033 | 0.002697 | 0.003771 | 0.003588 | 4.9e-05 | 0.001275 | 0.31205 | 1.3141 | 2.5475 | 100.63 | 21.14 | 25.975 | 67.0404 | 86.6901 | 65.0342 | 0.006483 | 1.0439 | 0.025285 | 24.2205 | -0.016962 | 0.1393 | 4.1e-05 | 0.003128 | 0.003128 | 1.8391
2019-03-28 15:37 | 1553801798.2814 | -103 | 0.20952 | 0.000448 | 0 | 101.1554 | 101.3878 | 126.9476 | 21.2612 | 0.00188 | 0.001779 | -151.9737 | 0.23466 | 8.4027 | 0.002646 | 0.004479 | 0.004295 | 5.2e-05 | 0.001275 | 0.31924 | 1.2897 | 2.5473 | 100.63 | 21.14 | 25.968 | 67.0623 | 85.4749 | 64.636 | 0.006481 | 1.0452 | 0.025319 | 24.2244 | -0.020523 | 0.1396 | -0.000184 | -0.014047 | -0.014047 | -0.41043
2019-03-28 15:37 | 1553801799.2811 | -102 | 0.2095 | 0.000449 | 0 | 101.1724 | 101.3147 | 127.4161 | 21.2612 | 0.001921 | 0.001759 | -151.9765 | 0.23411 | 8.4007 | 0.002619 | 0.004967 | 0.004817 | 5e-05 | 0.001267 | 0.30646 | 1.32 | 2.5472 | 100.63 | 21.14 | 25.956 | 68.0435 | 84.8303 | 64.1581 | 0.006478 | 1.0467 | 0.025383 | 24.2519 | -0.012576 | 0.13927 | -5.6e-05 | -0.004306 | -0.004306 | -1.3357
2019-03-28 15:37 | 1553801800.2815 | -101 | 0.20952 | 0.000446 | 0 | 101.2171 | 101.3694 | 127.7503 | 21.2612 | 0.001872 | 0.0018 | -152.0024 | 0.23379 | 8.4044 | 0.002641 | 0.004224 | 0.004174 | 4.7e-05 | 0.001271 | 0.32228 | 1.3399 | 2.5474 | 100.63 | 21.14 | 25.95 | 66.8634 | 85.3623 | 65.1396 | 0.006476 | 1.0436 | 0.02538 | 24.319 | -0.013454 | 0.13908 | -0.00017 | -0.013019 | -0.013019 | -0.4412
2019-03-28 15:37 | 1553801801.2811 | -100 | 0.20951 | 0.000449 | 0 | 101.1203 | 101.186 | 127.6274 | 21.2612 | 0.001852 | 0.001762 | -152.0035 | 0.23397 | 8.4049 | 0.002665 | 0.003611 | 0.003396 | 4.8e-05 | 0.001269 | 0.31205 | 1.3341 | 2.5474 | 100.63 | 21.14 | 25.946 | 66.382 | 85.9377 | 64.22 | 0.006475 | 1.0465 | 0.025402 | 24.2742 | -0.005811 | 0.13919 | -0.000115 | -0.00882 | -0.00882 | -0.65175
2019-03-28 15:37 | 1553801802.2803 | -99 | 0.20953 | 0.00045 | 0 | 101.2209 | 101.2203 | 127.6762 | 21.2612 | 0.001912 | 0.001738 | -152.0156 | 0.23498 | 8.4034 | 0.002695 | 0.003885 | 0.003733 | 5.2e-05 | 0.00127 | 0.29983 | 1.3257 | 2.5473 | 100.63 | 21.14 | 25.939 | 67.8269 | 86.6526 | 63.639 | 0.006474 | 1.0483 | 0.025429 | 24.2579 | 5.7e-05 | 0.13979 | -0.000247 | -0.0189 | -0.0189 | -0.30546
2019-03-28 15:37 | 1553801803.2806 | -98 | 0.20951 | 0.000445 | 0 | 101.1578 | 101.2121 | 127.3053 | 21.2612 | 0.001861 | 0.001753 | -152.025 | 0.23425 | 8.4044 | 0.002675 | 0.00373 | 0.003517 | 4.9e-05 | 0.00127 | 0.34562 | 1.3449 | 2.5473 | 100.63 | 21.14 | 25.937 | 66.6167 | 86.1702 | 63.9988 | 0.006473 | 1.0472 | 0.025378 | 24.2356 | -0.004804 | 0.13935 | -0.000108 | -0.008262 | -0.008262 | -0.69662
2019-03-28 15:37 | 1553801804.2813 | -97 | 0.2095 | 0.000443 | 0 | 101.159 | 101.1659 | 127.2523 | 21.2612 | 0.001832 | 0.001729 | -152.0273 | 0.23411 | 8.4036 | 0.00256 | 0.003851 | 0.003688 | 4.7e-05 | 0.001273 | 0.32117 | 1.3078 | 2.5473 | 100.63 | 21.14 | 25.934 | 65.8973 | 83.3978 | 63.4282 | 0.006472 | 1.0489 | 0.025395 | 24.21 | -0.000615 | 0.13927 | -9.1e-05 | -0.006956 | -0.006956 | -0.82693
2019-03-28 15:37 | 1553801805.2817 | -96 | 0.20949 | 0.000449 | 0 | 101.1277 | 101.3134 | 127.0722 | 21.2612 | 0.001855 | 0.001793 | -152.022 | 0.23423 | 8.404 | 0.002566 | 0.004116 | 0.004006 | 4.9e-05 | 0.001277 | 0.31797 | 1.333 | 2.5474 | 100.63 | 21.14 | 25.929 | 66.4623 | 83.5445 | 64.9647 | 0.006471 | 1.0442 | 0.025319 | 24.2481 | -0.016407 | 0.13935 | -4.8e-05 | -0.003664 | -0.003664 | -1.5707
2019-03-28 15:37 | 1553801806.2808 | -95 | 0.20948 | 0.000445 | 0 | 101.1232 | 101.364 | 126.9258 | 21.2612 | 0.001883 | 0.00172 | -152.0122 | 0.23393 | 8.4034 | 0.002623 | 0.003884 | 0.003676 | 5.1e-05 | 0.001277 | 0.33483 | 1.3197 | 2.5474 | 100.63 | 21.14 | 25.926 | 67.1286 | 84.9253 | 63.2031 | 0.00647 | 1.0496 | 0.02537 | 24.1709 | -0.02127 | 0.13917 | 6.7e-05 | 0.005156 | 0.005156 | 1.1148
2019-03-28 15:37 | 1553801807.2804 | -94 | 0.2095 | 0.000447 | 0 | 101.1882 | 101.383 | 126.6178 | 21.2612 | 0.001874 | 0.001739 | -152.0027 | 0.23444 | 8.403 | 0.002664 | 0.00338 | 0.003227 | 5e-05 | 0.001279 | 0.30143 | 1.3161 | 2.5475 | 100.63 | 21.13 | 25.922 | 66.9163 | 85.8954 | 63.6763 | 0.006465 | 1.0482 | 0.025322 | 24.1585 | -0.017207 | 0.13947 | -8.1e-05 | -0.006206 | -0.006206 | -0.92807
2019-03-28 15:37 | 1553801808.2807 | -93 | 0.20951 | 0.000446 | 0 | 101.1292 | 101.2264 | 126.5462 | 21.2612 | 0.001898 | 0.001765 | -151.9776 | 0.23476 | 8.4043 | 0.002636 | 0.003851 | 0.003644 | 5.1e-05 | 0.001278 | 0.30414 | 1.3254 | 2.5474 | 100.63 | 21.13 | 25.92 | 67.5062 | 85.2274 | 64.2948 | 0.006465 | 1.0462 | 0.025291 | 24.1738 | -0.008594 | 0.13966 | -0.000108 | -0.008271 | -0.008271 | -0.69738
2019-03-28 15:37 | 1553801809.2808 | -92 | 0.20949 | 0.000447 | 0 | 101.1813 | 101.3368 | 126.311 | 21.2612 | 0.001907 | 0.001759 | -151.9902 | 0.23349 | 8.4035 | 0.002609 | 0.004002 | 0.003812 | 5.2e-05 | 0.001265 | 0.32228 | 1.3079 | 2.5473 | 100.63 | 21.13 | 25.919 | 67.7067 | 84.5908 | 64.1504 | 0.006465 | 1.0467 | 0.025273 | 24.1462 | -0.013735 | 0.1389 | -1.3e-05 | -0.000973 | -0.000973 | -5.8982
2019-03-28 15:37 | 1553801810.2804 | -91 | 0.20949 | 0.000447 | 0 | 101.1369 | 101.2064 | 125.9585 | 21.2612 | 0.001953 | 0.001774 | -151.9891 | 0.23411 | 8.4021 | 0.002723 | 0.003717 | 0.00354 | 5.4e-05 | 0.001254 | 0.27903 | 1.3063 | 2.5474 | 100.63 | 21.13 | 25.911 | 68.8135 | 87.3155 | 64.5111 | 0.006463 | 1.0456 | 0.025225 | 24.1254 | -0.006137 | 0.13927 | -4.5e-05 | -0.003437 | -0.003437 | -1.6734
2019-03-28 15:37 | 1553801811.2808 | -90 | 0.20948 | 0.000446 | 0 | 101.1557 | 101.1102 | 125.8714 | 21.2612 | 0.001851 | 0.001765 | -151.9821 | 0.23442 | 8.4051 | 0.002703 | 0.004201 | 0.004021 | 5.3e-05 | 0.001245 | 0.33196 | 1.3277 | 2.5474 | 100.63 | 21.14 | 25.907 | 66.3593 | 86.847 | 64.2878 | 0.006466 | 1.0463 | 0.025224 | 24.109 | 0.004023 | 0.13945 | 5.3e-05 | 0.004051 | 0.004051 | 1.4218
2019-03-28 15:37 | 1553801812.2806 | -89 | 0.20949 | 0.000448 | -1e-06 | 101.1048 | 101.3932 | 125.4475 | 21.2612 | 0.001915 | 0.001737 | -151.9738 | 0.23403 | 8.4005 | 0.002619 | 0.004191 | 0.004053 | 5.1e-05 | 0.00124 | 0.29878 | 1.3516 | 2.5475 | 100.63 | 21.13 | 25.9 | 67.9086 | 84.824 | 63.6318 | 0.00646 | 1.0483 | 0.025206 | 24.045 | -0.025481 | 0.13923 | 1e-05 | 0.00082 | 0.00082 | 7.0135
2019-03-28 15:37 | 1553801813.2799 | -88 | 0.20949 | 0.000447 | 0 | 101.1253 | 101.1428 | 125.4857 | 21.2612 | 0.001864 | 0.001773 | -151.9741 | 0.23408 | 8.4041 | 0.00273 | 0.00412 | 0.003978 | 5.1e-05 | 0.001233 | 0.3178 | 1.3177 | 2.5474 | 100.63 | 21.14 | 25.896 | 66.673 | 87.4905 | 64.4868 | 0.006463 | 1.0456 | 0.025178 | 24.0792 | -0.001548 | 0.13925 | 3e-06 | 0.000263 | 0.000263 | 21.8857
2019-03-28 15:37 | 1553801814.2809 | -87 | 0.2095 | 0.000448 | 0 | 101.1229 | 101.1839 | 125.5907 | 21.2612 | 0.001843 | 0.001753 | -151.9546 | 0.23411 | 8.4005 | 0.002661 | 0.0039 | 0.003689 | 5.2e-05 | 0.001231 | 0.29463 | 1.3354 | 2.5475 | 100.63 | 21.14 | 25.891 | 66.1657 | 85.8415 | 64.0161 | 0.006462 | 1.0471 | 0.025206 | 24.0724 | -0.005394 | 0.13927 | -5.8e-05 | -0.004366 | -0.004366 | -1.3175
2019-03-28 15:37 | 1553801815.2806 | -86 | 0.2095 | 0.000445 | -1e-06 | 101.1658 | 101.1826 | 125.7869 | 21.2612 | 0.001901 | 0.001754 | -151.9534 | 0.2343 | 8.4045 | 0.002697 | 0.004297 | 0.004053 | 5.5e-05 | 0.001231 | 0.323 | 1.3578 | 2.5474 | 100.63 | 21.14 | 25.882 | 67.5624 | 86.7019 | 64.0238 | 0.006459 | 1.0471 | 0.025226 | 24.0915 | -0.001481 | 0.13939 | -5.4e-05 | -0.004043 | -0.004043 | -1.4239
2019-03-28 15:37 | 1553801816.281 | -85 | 0.20952 | 0.000443 | 0 | 101.1077 | 101.2858 | 126.0624 | 21.2612 | 0.001852 | 0.001739 | -151.9608 | 0.2339 | 8.4047 | 0.002736 | 0.004619 | 0.004483 | 4.9e-05 | 0.001239 | 0.31293 | 1.3369 | 2.5475 | 100.63 | 21.14 | 25.879 | 66.399 | 87.638 | 63.671 | 0.006459 | 1.0482 | 0.025266 | 24.1053 | -0.015742 | 0.13915 | -0.00018 | -0.013733 | -0.013733 | -0.41846
2019-03-28 15:37 | 1553801817.2807 | -84 | 0.2095 | 0.000447 | 0 | 101.1113 | 101.118 | 126.4587 | 21.2612 | 0.001878 | 0.001727 | -151.9683 | 0.2345 | 8.4057 | 0.002739 | 0.004135 | 0.003946 | 5.3e-05 | 0.00124 | 0.30807 | 1.3348 | 2.5474 | 100.63 | 21.14 | 25.872 | 67.0233 | 87.7037 | 63.392 | 0.006457 | 1.049 | 0.025317 | 24.1331 | -0.000593 | 0.1395 | -0.000107 | -0.008162 | -0.008162 | -0.70584
2019-03-28 15:37 | 1553801818.2798 | -83 | 0.20952 | 0.000451 | 0 | 101.1554 | 101.2043 | 126.592 | 21.2612 | 0.001842 | 0.001745 | -151.9851 | 0.23321 | 8.4055 | 0.002693 | 0.004569 | 0.004454 | 5.5e-05 | 0.001244 | 0.32587 | 1.3544 | 2.5474 | 100.63 | 21.14 | 25.868 | 66.1549 | 86.6083 | 63.8099 | 0.006456 | 1.0477 | 0.025314 | 24.1608 | -0.004324 | 0.13874 | -0.000192 | -0.014595 | -0.014595 | -0.39258
2019-03-28 15:37 | 1553801819.2811 | -82 | 0.20949 | 0.000446 | 0 | 101.136 | 101.3524 | 126.8312 | 21.2612 | 0.001864 | 0.001747 | -151.9553 | 0.23366 | 8.4036 | 0.002732 | 0.003873 | 0.003666 | 5.4e-05 | 0.001243 | 0.2943 | 1.3626 | 2.5473 | 100.63 | 21.14 | 25.864 | 66.6691 | 87.5333 | 63.8583 | 0.006455 | 1.0476 | 0.025336 | 24.1854 | -0.019118 | 0.139 | -3.3e-05 | -0.002503 | -0.002503 | -2.2933
2019-03-28 15:37 | 1553801820.2798 | -81 | 0.20951 | 0.000444 | -1e-06 | 101.1453 | 101.3059 | 127.3753 | 21.2612 | 0.001907 | 0.001784 | -151.9459 | 0.2333 | 8.402 | 0.002688 | 0.00413 | 0.003909 | 5.4e-05 | 0.001243 | 0.30021 | 1.3442 | 2.5474 | 100.63 | 21.13 | 25.858 | 67.7223 | 86.4894 | 64.7633 | 0.006449 | 1.0448 | 0.025357 | 24.2697 | -0.014193 | 0.13879 | -0.000126 | -0.009553 | -0.009553 | -0.60002
2019-03-28 15:37 | 1553801821.2802 | -80 | 0.20948 | 0.000444 | 0 | 101.1751 | 101.1771 | 127.5994 | 21.2612 | 0.001862 | 0.001747 | -151.9435 | 0.23433 | 8.4015 | 0.002688 | 0.004076 | 0.003913 | 5.4e-05 | 0.001241 | 0.30735 | 1.3509 | 2.5475 | 100.63 | 21.13 | 25.855 | 66.6347 | 86.4935 | 63.8614 | 0.006449 | 1.0476 | 0.025413 | 24.2586 | -0.000184 | 0.13941 | 3.8e-05 | 0.002949 | 0.002949 | 1.9524
2019-03-28 15:37 | 1553801822.2798 | -79 | 0.20952 | 0.000447 | 0 | 101.1867 | 101.3035 | 127.9351 | 21.2612 | 0.00184 | 0.001757 | -151.9362 | 0.234 | 8.4029 | 0.002673 | 0.00349 | 0.003338 | 5.7e-05 | 0.001243 | 0.30934 | 1.309 | 2.5474 | 100.63 | 21.13 | 25.855 | 66.1126 | 86.1329 | 64.0989 | 0.006449 | 1.0468 | 0.025437 | 24.2991 | -0.010322 | 0.13921 | -0.000191 | -0.01461 | -0.01461 | -0.39353
2019-03-28 15:37 | 1553801823.2801 | -78 | 0.20949 | 0.000447 | -1e-06 | 101.1905 | 101.2383 | 127.9351 | 21.2612 | 0.001865 | 0.001742 | -151.9479 | 0.23386 | 8.4029 | 0.00266 | 0.004338 | 0.004158 | 5.6e-05 | 0.001235 | 0.29856 | 1.3007 | 2.5474 | 100.63 | 21.13 | 25.853 | 66.7011 | 85.8074 | 63.7326 | 0.006448 | 1.048 | 0.025451 | 24.2859 | -0.004219 | 0.13913 | -2.1e-05 | -0.001579 | -0.001579 | -3.6391
2019-03-28 15:37 | 1553801824.2799 | -77 | 0.20949 | 0.000447 | 0 | 101.1572 | 101.0915 | 127.7466 | 21.2612 | 0.001807 | 0.001761 | -151.9433 | 0.23366 | 8.4052 | 0.002655 | 0.004593 | 0.004436 | 5.5e-05 | 0.001233 | 0.31365 | 1.3097 | 2.5474 | 100.63 | 21.13 | 25.853 | 65.3142 | 85.6864 | 64.2074 | 0.006448 | 1.0465 | 0.025414 | 24.2851 | 0.005808 | 0.139 | 0 | 6.4e-05 | 6.4e-05 | 89.7222
2019-03-28 15:37 | 1553801825.28 | -76 | 0.20949 | 0.000447 | 0 | 101.1581 | 101.2216 | 127.4694 | 21.2612 | 0.001935 | 0.00179 | -151.9212 | 0.2337 | 8.4033 | 0.002669 | 0.004385 | 0.004178 | 5.5e-05 | 0.001231 | 0.31868 | 1.341 | 2.5474 | 100.63 | 21.13 | 25.851 | 68.3743 | 86.0313 | 64.899 | 0.006448 | 1.0444 | 0.025361 | 24.2836 | -0.005618 | 0.13903 | -4.1e-05 | -0.00317 | -0.00317 | -1.8113
2019-03-28 15:37 | 1553801826.2794 | -75 | 0.20949 | 0.000446 | 0 | 101.1387 | 101.1904 | 127.2917 | 21.2612 | 0.001858 | 0.001754 | -151.9306 | 0.23399 | 8.4053 | 0.002693 | 0.004078 | 0.003951 | 5.8e-05 | 0.00123 | 0.2855 | 1.3421 | 2.5475 | 100.63 | 21.13 | 25.849 | 66.5433 | 86.6104 | 64.0356 | 0.006447 | 1.047 | 0.025376 | 24.2356 | -0.004568 | 0.1392 | 7e-06 | 0.00059 | 0.00059 | 9.7413
2019-03-28 15:37 | 1553801827.2799 | -74 | 0.20947 | 0.000449 | 0 | 101.139 | 101.3039 | 126.9969 | 21.2612 | 0.001907 | 0.001778 | -151.9348 | 0.23364 | 8.4046 | 0.002686 | 0.003849 | 0.003703 | 5.7e-05 | 0.001231 | 0.32565 | 1.2681 | 2.5474 | 100.63 | 21.13 | 25.847 | 67.7081 | 86.4418 | 64.6211 | 0.006447 | 1.0452 | 0.025324 | 24.2285 | -0.014566 | 0.13899 | 0.000129 | 0.00987 | 0.00987 | 0.58162
2019-03-28 15:37 | 1553801828.28 | -73 | 0.2095 | 0.000441 | 0 | 101.1792 | 101.3708 | 126.5758 | 21.2612 | 0.001884 | 0.001772 | -151.9362 | 0.23405 | 8.4012 | 0.002581 | 0.004501 | 0.004302 | 5.8e-05 | 0.001234 | 0.27887 | 1.3105 | 2.5474 | 100.63 | 21.13 | 25.846 | 67.1537 | 83.9102 | 64.4767 | 0.006446 | 1.0457 | 0.025288 | 24.1832 | -0.016918 | 0.13923 | -4.9e-05 | -0.003744 | -0.003744 | -1.5358
2019-03-28 15:37 | 1553801829.2803 | -72 | 0.20947 | 0.000447 | 0 | 101.1852 | 101.3005 | 126.5845 | 21.2612 | 0.001877 | 0.001759 | -151.9497 | 0.23389 | 8.4022 | 0.002651 | 0.004162 | 0.004049 | 5.7e-05 | 0.001241 | 0.30071 | 1.3047 | 2.5475 | 100.63 | 21.13 | 25.845 | 66.9841 | 85.6019 | 64.155 | 0.006446 | 1.0467 | 0.0253 | 24.1725 | -0.010184 | 0.13914 | 0.000134 | 0.010214 | 0.010214 | 0.56263
2019-03-28 15:37 | 1553801830.2801 | -71 | 0.20948 | 0.000447 | 0 | 101.1417 | 101.2006 | 125.8979 | 21.2612 | 0.001955 | 0.001783 | -151.9479 | 0.23432 | 8.4035 | 0.002628 | 0.004094 | 0.003902 | 5.7e-05 | 0.001242 | 0.29806 | 1.2928 | 2.5474 | 100.63 | 21.14 | 25.845 | 68.8758 | 85.0487 | 64.7344 | 0.00645 | 1.0449 | 0.02521 | 24.1275 | -0.005206 | 0.1394 | 6.1e-05 | 0.00465 | 0.00465 | 1.238
2019-03-28 15:37 | 1553801831.2791 | -70 | 0.20951 | 0.000444 | 0 | 101.1411 | 101.2339 | 125.7589 | 21.2612 | 0.001868 | 0.001764 | -151.9684 | 0.23424 | 8.4059 | 0.002653 | 0.004075 | 0.003851 | 5.7e-05 | 0.001245 | 0.29231 | 1.316 | 2.5475 | 100.63 | 21.14 | 25.843 | 66.7651 | 85.6293 | 64.2736 | 0.00645 | 1.0463 | 0.025213 | 24.0978 | -0.0082 | 0.13935 | -0.000138 | -0.010475 | -0.010475 | -0.54942
2019-03-28 15:37 | 1553801832.2792 | -69 | 0.20947 | 0.000451 | 0 | 101.1488 | 101.1374 | 125.7017 | 21.2612 | 0.001904 | 0.001772 | -151.9671 | 0.23312 | 8.4012 | 0.002655 | 0.003543 | 0.003385 | 5.7e-05 | 0.001243 | 0.3183 | 1.316 | 2.5475 | 100.63 | 21.14 | 25.842 | 67.6341 | 85.6777 | 64.4596 | 0.006449 | 1.0457 | 0.025201 | 24.0989 | 0.001013 | 0.13869 | 0.000115 | 0.00876 | 0.00876 | 0.65386
2019-03-28 15:37 | 1553801833.2801 | -68 | 0.20947 | 0.000454 | 0 | 101.2209 | 101.3531 | 125.8258 | 21.2612 | 0.00189 | 0.00174 | -151.9816 | 0.23422 | 8.4013 | 0.002638 | 0.003628 | 0.003398 | 5.3e-05 | 0.001245 | 0.31996 | 1.357 | 2.5474 | 100.63 | 21.14 | 25.839 | 67.308 | 85.2883 | 63.7013 | 0.006449 | 1.0481 | 0.025242 | 24.0837 | -0.011673 | 0.13933 | 7.4e-05 | 0.00564 | 0.00564 | 1.0203
2019-03-28 15:37 | 1553801834.2799 | -67 | 0.20945 | 0.000453 | 0 | 101.1357 | 101.2125 | 125.7923 | 21.2612 | 0.001894 | 0.001806 | -151.9794 | 0.23427 | 8.4022 | 0.002634 | 0.004029 | 0.003819 | 5.6e-05 | 0.001247 | 0.3235 | 1.3159 | 2.5475 | 100.63 | 21.14 | 25.837 | 67.3907 | 85.1728 | 65.2791 | 0.006448 | 1.0432 | 0.025179 | 24.1368 | -0.006783 | 0.13937 | 0.000204 | 0.015492 | 0.015492 | 0.37154
2019-03-28 15:37 | 1553801835.2802 | -66 | 0.20948 | 0.000455 | -1e-06 | 101.1372 | 101.3259 | 125.9202 | 21.2612 | 0.001824 | 0.00175 | -152.013 | 0.23362 | 8.4042 | 0.002685 | 0.004046 | 0.003832 | 5.7e-05 | 0.001245 | 0.32068 | 1.2937 | 2.5473 | 100.63 | 21.14 | 25.835 | 65.7061 | 86.4052 | 63.941 | 0.006448 | 1.0473 | 0.025242 | 24.1013 | -0.016673 | 0.13898 | 5.8e-05 | 0.004412 | 0.004412 | 1.3009
2019-03-28 15:37 | 1553801836.2793 | -65 | 0.20947 | 0.000458 | 0 | 101.1277 | 101.2774 | 125.9174 | 21.2612 | 0.001886 | 0.001781 | -152.0078 | 0.23404 | 8.4033 | 0.002717 | 0.004161 | 0.004061 | 5.4e-05 | 0.001253 | 0.35369 | 1.3135 | 2.5474 | 100.63 | 21.14 | 25.829 | 67.2027 | 87.1812 | 64.6719 | 0.006446 | 1.0451 | 0.025214 | 24.1272 | -0.013227 | 0.13923 | 0.000104 | 0.0079 | 0.0079 | 0.72784
2019-03-28 15:37 | 1553801837.2795 | -64 | 0.20949 | 0.00046 | 0 | 101.181 | 101.1856 | 126.2023 | 21.2612 | 0.001893 | 0.001803 | -151.9889 | 0.23409 | 8.4063 | 0.002697 | 0.003894 | 0.003692 | 5.4e-05 | 0.00125 | 0.30917 | 1.3176 | 2.5475 | 100.63 | 21.14 | 25.828 | 67.3643 | 86.6963 | 65.2019 | 0.006446 | 1.0434 | 0.025223 | 24.1734 | -0.000408 | 0.13926 | -6e-06 | -0.000514 | -0.000514 | -11.1931
2019-03-28 15:37 | 1553801838.2807 | -63 | 0.20949 | 0.000458 | 0 | 101.1712 | 101.2896 | 126.4817 | 21.2612 | 0.001896 | 0.001784 | -152.0106 | 0.23401 | 8.4037 | 0.002689 | 0.004031 | 0.003907 | 5.5e-05 | 0.001243 | 0.29302 | 1.3468 | 2.5475 | 100.63 | 21.14 | 25.825 | 67.4453 | 86.5106 | 64.7616 | 0.006445 | 1.0448 | 0.025268 | 24.1844 | -0.010462 | 0.13921 | -4.6e-05 | -0.003567 | -0.003567 | -1.6117
2019-03-28 15:37 | 1553801839.2795 | -62 | 0.20949 | 0.00046 | 1e-06 | 101.1622 | 101.0915 | 126.7936 | 21.2612 | 0.001873 | 0.001776 | -152.0321 | 0.23353 | 8.4033 | 0.002768 | 0.003586 | 0.003366 | 5.5e-05 | 0.001241 | 0.32604 | 1.3473 | 2.5475 | 100.63 | 21.14 | 25.823 | 66.8891 | 88.4088 | 64.565 | 0.006445 | 1.0454 | 0.025306 | 24.2071 | 0.006256 | 0.13892 | -3.4e-05 | -0.002631 | -0.002631 | -2.1808
2019-03-28 15:37 | 1553801840.2794 | -61 | 0.20952 | 0.000459 | 0 | 101.15 | 101.3453 | 126.9959 | 21.2612 | 0.001876 | 0.001752 | -152.0278 | 0.23407 | 8.4028 | 0.002761 | 0.00425 | 0.00407 | 5.6e-05 | 0.001244 | 0.32117 | 1.2959 | 2.5474 | 100.63 | 21.14 | 25.82 | 66.957 | 88.2417 | 63.9887 | 0.006444 | 1.0472 | 0.025348 | 24.2058 | -0.01725 | 0.13925 | -0.000236 | -0.018066 | -0.018066 | -0.31834
2019-03-28 15:37 | 1553801841.2794 | -60 | 0.20948 | 0.000461 | 0 | 101.1334 | 101.2536 | 127.1438 | 21.2612 | 0.001889 | 0.001795 | -152.0178 | 0.23383 | 8.3994 | 0.002748 | 0.004 | 0.003819 | 5.2e-05 | 0.001242 | 0.31022 | 1.3241 | 2.5475 | 100.63 | 21.14 | 25.812 | 67.2837 | 87.9184 | 65.0294 | 0.006442 | 1.044 | 0.025324 | 24.2572 | -0.010626 | 0.1391 | 3.6e-05 | 0.002752 | 0.002752 | 2.0876
2019-03-28 15:37 | 1553801842.2794 | -59 | 0.20951 | 0.000461 | 0 | 101.1426 | 101.3066 | 127.1943 | 21.2612 | 0.001863 | 0.001708 | -152.0145 | 0.23452 | 8.4045 | 0.002754 | 0.004034 | 0.003943 | 5.4e-05 | 0.001242 | 0.32997 | 1.2738 | 2.5474 | 100.63 | 21.15 | 25.807 | 66.659 | 88.0799 | 62.9219 | 0.006445 | 1.0505 | 0.025408 | 24.1863 | -0.01449 | 0.13952 | -0.000128 | -0.009745 | -0.009745 | -0.59129
2019-03-28 15:37 | 1553801843.2795 | -58 | 0.20949 | 0.000456 | 0 | 101.1173 | 101.2791 | 127.241 | 21.2612 | 0.001938 | 0.001769 | -152.0046 | 0.23381 | 8.4036 | 0.002667 | 0.003638 | 0.003435 | 5.5e-05 | 0.001244 | 0.29734 | 1.3035 | 2.5475 | 100.63 | 21.15 | 25.796 | 68.4469 | 85.982 | 64.404 | 0.006442 | 1.0459 | 0.025357 | 24.244 | -0.014299 | 0.13909 | -4.7e-05 | -0.003535 | -0.003535 | -1.6248
2019-03-28 15:37 | 1553801844.2785 | -57 | 0.20954 | 0.000455 | 0 | 101.114 | 101.2318 | 127.3236 | 21.2612 | 0.001832 | 0.001822 | -152.02 | 0.23406 | 8.405 | 0.002753 | 0.004293 | 0.004084 | 5.5e-05 | 0.001247 | 0.33052 | 1.3335 | 2.5473 | 100.63 | 21.15 | 25.789 | 65.898 | 88.0472 | 65.6734 | 0.00644 | 1.042 | 0.025317 | 24.2974 | -0.010417 | 0.13924 | -0.000324 | -0.024709 | -0.024709 | -0.23273
2019-03-28 15:37 | 1553801845.2789 | -56 | 0.20952 | 0.000455 | 0 | 101.2075 | 101.1802 | 127.1894 | 21.2612 | 0.001834 | 0.001773 | -151.9965 | 0.23429 | 8.4019 | 0.002666 | 0.004182 | 0.004014 | 5.5e-05 | 0.00125 | 0.31293 | 1.3485 | 2.5475 | 100.63 | 21.15 | 25.783 | 65.9449 | 85.9541 | 64.4924 | 0.006439 | 1.0456 | 0.025348 | 24.2423 | 0.002416 | 0.13938 | -0.000225 | -0.017124 | -0.017124 | -0.33616
2019-03-28 15:37 | 1553801846.2796 | -55 | 0.20952 | 0.000456 | 0 | 101.1661 | 101.2954 | 127.0178 | 21.2612 | 0.001906 | 0.00177 | -151.9849 | 0.23371 | 8.4015 | 0.002668 | 0.00384 | 0.003643 | 5.5e-05 | 0.001258 | 0.33229 | 1.29 | 2.5474 | 100.63 | 21.15 | 25.77 | 67.6802 | 85.9986 | 64.4245 | 0.006435 | 1.0458 | 0.025334 | 24.2235 | -0.011419 | 0.13903 | -0.000234 | -0.017851 | -0.017851 | -0.32166
2019-03-28 15:37 | 1553801847.2792 | -54 | 0.20956 | 0.000455 | 0 | 101.1661 | 101.2719 | 126.9297 | 21.2612 | 0.001913 | 0.001769 | -151.9911 | 0.23356 | 8.4001 | 0.002675 | 0.004017 | 0.003814 | 5.4e-05 | 0.001261 | 0.32781 | 1.3424 | 2.5475 | 100.63 | 21.15 | 25.765 | 67.8579 | 86.1649 | 64.3901 | 0.006434 | 1.0459 | 0.025326 | 24.2139 | -0.009349 | 0.13894 | -0.000443 | -0.033828 | -0.033828 | -0.16964
2019-03-28 15:37 | 1553801848.2787 | -53 | 0.20953 | 0.000453 | 0 | 101.1095 | 101.3504 | 126.8666 | 21.2612 | 0.001895 | 0.001743 | -151.9599 | 0.23455 | 8.4026 | 0.002661 | 0.003818 | 0.003589 | 5.9e-05 | 0.001268 | 0.32084 | 1.3301 | 2.5474 | 100.63 | 21.15 | 25.759 | 67.4283 | 85.823 | 63.7646 | 0.006433 | 1.0479 | 0.025343 | 24.1854 | -0.021283 | 0.13953 | -0.000239 | -0.018209 | -0.018209 | -0.31647
2019-03-28 15:37 | 1553801849.2797 | -52 | 0.20952 | 0.000456 | -1e-06 | 101.1042 | 101.047 | 126.2331 | 21.2612 | 0.001872 | 0.001773 | -151.9432 | 0.23444 | 8.4047 | 0.002651 | 0.004334 | 0.004153 | 5.7e-05 | 0.001274 | 0.32029 | 1.33 | 2.5475 | 100.63 | 21.15 | 25.747 | 66.875 | 85.6036 | 64.4924 | 0.00643 | 1.0456 | 0.025253 | 24.151 | 0.005057 | 0.13947 | -0.000186 | -0.014128 | -0.014128 | -0.40771
2019-03-28 15:37 | 1553801850.2791 | -51 | 0.20955 | 0.000453 | 0 | 101.1524 | 101.4408 | 126.3167 | 21.2612 | 0.001834 | 0.001791 | -151.9528 | 0.2344 | 8.404 | 0.00269 | 0.003568 | 0.003536 | 5.6e-05 | 0.001269 | 0.32925 | 1.3476 | 2.5474 | 100.63 | 21.15 | 25.742 | 65.955 | 86.5339 | 64.9279 | 0.006428 | 1.0443 | 0.025245 | 24.1745 | -0.025461 | 0.13945 | -0.000352 | -0.026734 | -0.026734 | -0.21542
2019-03-28 15:38 | 1553801851.2799 | -50 | 0.20949 | 0.000446 | 0 | 101.0887 | 101.31 | 126.1526 | 21.2612 | 0.001868 | 0.001754 | -151.9394 | 0.23448 | 8.4035 | 0.002705 | 0.004122 | 0.003897 | 5.5e-05 | 0.001271 | 0.29302 | 1.2874 | 2.5473 | 100.63 | 21.15 | 25.729 | 66.7797 | 86.8908 | 64.0419 | 0.006425 | 1.047 | 0.025262 | 24.1272 | -0.019559 | 0.13949 | 1.5e-05 | 0.001154 | 0.001154 | 4.9911
2019-03-28 15:38 | 1553801852.2797 | -49 | 0.20948 | 0.000451 | 0 | 101.1194 | 101.222 | 126.3798 | 21.2612 | 0.001902 | 0.001782 | -151.9527 | 0.23373 | 8.403 | 0.002646 | 0.003529 | 0.003349 | 5.8e-05 | 0.001267 | 0.31526 | 1.29 | 2.5474 | 100.63 | 21.15 | 25.723 | 67.5872 | 85.4739 | 64.7094 | 0.006424 | 1.0449 | 0.025259 | 24.1728 | -0.009073 | 0.13905 | 7e-05 | 0.005328 | 0.005328 | 1.0779
2019-03-28 15:38 | 1553801853.2788 | -48 | 0.20953 | 0.000451 | 0 | 101.1256 | 101.2926 | 126.4706 | 21.2612 | 0.001854 | 0.001774 | -151.9464 | 0.23373 | 8.4033 | 0.002644 | 0.004071 | 0.003956 | 5.3e-05 | 0.001268 | 0.32372 | 1.3081 | 2.5473 | 100.63 | 21.15 | 25.71 | 66.4459 | 85.4141 | 64.522 | 0.00642 | 1.0455 | 0.025275 | 24.1747 | -0.014762 | 0.13905 | -0.000244 | -0.018561 | -0.018561 | -0.30938
2019-03-28 15:38 | 1553801854.2793 | -47 | 0.20947 | 0.000451 | 0 | 101.1786 | 101.1812 | 126.3566 | 21.2612 | 0.001937 | 0.001774 | -151.954 | 0.23374 | 8.4029 | 0.002751 | 0.004317 | 0.004032 | 5.4e-05 | 0.00127 | 0.30757 | 1.2783 | 2.5474 | 100.63 | 21.15 | 25.704 | 68.4445 | 87.9964 | 64.5195 | 0.006419 | 1.0455 | 0.025264 | 24.1638 | -0.000228 | 0.13905 | 9.9e-05 | 0.007557 | 0.007557 | 0.75991
2019-03-28 15:38 | 1553801855.2802 | -46 | 0.20954 | 0.000449 | 0 | 101.1843 | 101.3487 | 126.4906 | 21.2612 | 0.001917 | 0.001792 | -151.9718 | 0.23383 | 8.402 | 0.002756 | 0.004424 | 0.004304 | 5.3e-05 | 0.001272 | 0.30613 | 1.2835 | 2.5473 | 100.63 | 21.15 | 25.699 | 67.9531 | 88.1168 | 64.9466 | 0.006418 | 1.0442 | 0.025261 | 24.1918 | -0.014522 | 0.13911 | -0.000301 | -0.022895 | -0.022895 | -0.25093
2019-03-28 15:38 | 1553801856.2792 | -45 | 0.20953 | 0.000449 | 0 | 101.1462 | 101.205 | 126.761 | 21.2612 | 0.001842 | 0.00179 | -151.9995 | 0.23441 | 8.404 | 0.002743 | 0.003674 | 0.003516 | 5.2e-05 | 0.001268 | 0.32261 | 1.319 | 2.5474 | 100.63 | 21.15 | 25.686 | 66.1564 | 87.8012 | 64.8927 | 0.006414 | 1.0444 | 0.02529 | 24.2158 | -0.005201 | 0.13945 | -0.000244 | -0.018572 | -0.018572 | -0.31011
2019-03-28 15:38 | 1553801857.2791 | -44 | 0.2095 | 0.000448 | 0 | 101.1643 | 101.1972 | 126.4023 | 21.2612 | 0.001898 | 0.00176 | -151.9906 | 0.23461 | 8.4032 | 0.002816 | 0.004037 | 0.003848 | 5.1e-05 | 0.001261 | 0.31813 | 1.273 | 2.5473 | 100.63 | 21.15 | 25.677 | 67.4999 | 89.5532 | 64.1682 | 0.006412 | 1.0466 | 0.025282 | 24.1556 | -0.002904 | 0.13957 | -7.8e-05 | -0.005895 | -0.005895 | -0.97784
2019-03-28 15:38 | 1553801858.2793 | -43 | 0.20952 | 0.000449 | 0 | 101.1223 | 101.2712 | 126.2296 | 21.2612 | 0.001899 | 0.001766 | -152.0088 | 0.23401 | 8.401 | 0.002777 | 0.004183 | 0.003999 | 5.3e-05 | 0.001256 | 0.31437 | 1.2967 | 2.5474 | 100.63 | 21.15 | 25.67 | 67.5138 | 88.637 | 64.3097 | 0.00641 | 1.0462 | 0.025259 | 24.1441 | -0.013161 | 0.13921 | -0.000217 | -0.016502 | -0.016502 | -0.34842
2019-03-28 15:38 | 1553801859.279 | -42 | 0.20953 | 0.000447 | -1e-06 | 101.1408 | 101.2926 | 126.3239 | 21.2612 | 0.001926 | 0.0018 | -152.0076 | 0.23457 | 8.4044 | 0.002779 | 0.00391 | 0.003792 | 5.2e-05 | 0.001259 | 0.31924 | 1.3001 | 2.5475 | 100.63 | 21.15 | 25.654 | 68.1724 | 88.6627 | 65.1504 | 0.006406 | 1.0436 | 0.025237 | 24.1832 | -0.013418 | 0.13954 | -0.000247 | -0.01875 | -0.01875 | -0.30737
2019-03-28 15:38 | 1553801860.2792 | -41 | 0.20952 | 0.000445 | 0 | 101.1911 | 101.2797 | 125.7744 | 21.2612 | 0.001863 | 0.001783 | -152.0056 | 0.23388 | 8.4026 | 0.002744 | 0.004024 | 0.003783 | 5.5e-05 | 0.001264 | 0.32908 | 1.2923 | 2.5475 | 100.63 | 21.15 | 25.645 | 66.659 | 87.8332 | 64.7327 | 0.006404 | 1.0449 | 0.025198 | 24.1156 | -0.007827 | 0.13913 | -0.000211 | -0.016021 | -0.016021 | -0.35867
2019-03-28 15:38 | 1553801861.278 | -40 | 0.20952 | 0.000444 | 0 | 101.1759 | 101.0769 | 125.6705 | 21.2612 | 0.001912 | 0.001797 | -152.0308 | 0.23394 | 8.4041 | 0.002781 | 0.003698 | 0.003595 | 5.4e-05 | 0.001261 | 0.30541 | 1.3065 | 2.5474 | 100.63 | 21.15 | 25.627 | 67.8291 | 88.7126 | 65.0568 | 0.0064 | 1.0439 | 0.025175 | 24.1172 | 0.008759 | 0.13917 | -0.000211 | -0.01598 | -0.01598 | -0.3597
2019-03-28 15:38 | 1553801862.2782 | -39 | 0.20949 | 0.000446 | 0 | 101.1566 | 101.3803 | 125.239 | 21.2612 | 0.001921 | 0.001781 | -152.0327 | 0.23395 | 8.4023 | 0.002808 | 0.00342 | 0.003146 | 5.4e-05 | 0.001257 | 0.30663 | 1.3184 | 2.5474 | 100.63 | 21.15 | 25.619 | 68.0529 | 89.3651 | 64.6719 | 0.006398 | 1.0451 | 0.025146 | 24.0621 | -0.019758 | 0.13918 | -3.5e-05 | -0.002663 | -0.002663 | -2.1587
2019-03-28 15:38 | 1553801863.278 | -38 | 0.20949 | 0.000449 | 0 | 101.1995 | 101.2658 | 124.9186 | 21.2612 | 0.001908 | 0.001753 | -152.0217 | 0.23385 | 8.4027 | 0.002802 | 0.004223 | 0.004093 | 5.7e-05 | 0.001255 | 0.33572 | 1.3504 | 2.5475 | 100.63 | 21.15 | 25.613 | 67.7425 | 89.2235 | 64.0051 | 0.006396 | 1.0471 | 0.025139 | 24.0076 | -0.00586 | 0.13912 | -2.9e-05 | -0.002236 | -0.002236 | -2.5696
2019-03-28 15:38 | 1553801864.2795 | -37 | 0.2095 | 0.000443 | 0 | 101.1715 | 101.2689 | 125.0299 | 21.2612 | 0.00188 | 0.001752 | -152.0112 | 0.2342 | 8.4017 | 0.002765 | 0.003993 | 0.003797 | 5.5e-05 | 0.001252 | 0.31509 | 1.3417 | 2.5474 | 100.63 | 21.16 | 25.604 | 67.0724 | 88.3385 | 63.974 | 0.006398 | 1.0472 | 0.025151 | 24.0171 | -0.008605 | 0.13933 | -3.7e-05 | -0.002831 | -0.002831 | -2.0328
2019-03-28 15:38 | 1553801865.2792 | -36 | 0.20952 | 0.000447 | 0 | 101.1348 | 101.208 | 125.2034 | 21.2612 | 0.001855 | 0.001741 | -152.0043 | 0.2334 | 8.4031 | 0.002763 | 0.004088 | 0.003897 | 5.4e-05 | 0.001248 | 0.31238 | 1.3464 | 2.5474 | 100.63 | 21.15 | 25.6 | 66.4529 | 88.2894 | 63.7139 | 0.006393 | 1.048 | 0.025179 | 24.0245 | -0.006472 | 0.13885 | -0.000196 | -0.014919 | -0.014919 | -0.38436
2019-03-28 15:38 | 1553801866.2792 | -35 | 0.20952 | 0.000449 | 1e-06 | 101.1783 | 101.1676 | 125.2606 | 21.2612 | 0.001834 | 0.001775 | -151.9869 | 0.23325 | 8.4023 | 0.002783 | 0.004128 | 0.00398 | 5.9e-05 | 0.00125 | 0.31581 | 1.3225 | 2.5473 | 100.63 | 21.15 | 25.593 | 65.9651 | 88.7595 | 64.5337 | 0.006391 | 1.0455 | 0.025154 | 24.0592 | 0.000947 | 0.13876 | -0.00019 | -0.014443 | -0.014443 | -0.39678
2019-03-28 15:38 | 1553801867.2782 | -34 | 0.20951 | 0.000449 | 0 | 101.1715 | 101.2179 | 125.5984 | 21.2612 | 0.001864 | 0.001787 | -151.9949 | 0.23315 | 8.4012 | 0.002768 | 0.003789 | 0.003578 | 5.7e-05 | 0.001255 | 0.31421 | 1.3294 | 2.5472 | 100.63 | 21.15 | 25.591 | 66.6838 | 88.4189 | 64.8162 | 0.006391 | 1.0446 | 0.025177 | 24.1017 | -0.004103 | 0.1387 | -0.000107 | -0.008123 | -0.008123 | -0.7052
2019-03-28 15:38 | 1553801868.2782 | -33 | 0.20952 | 0.000443 | 0 | 101.1757 | 101.2515 | 125.7587 | 21.2612 | 0.0019 | 0.00176 | -151.9927 | 0.23384 | 8.4013 | 0.002793 | 0.004353 | 0.004086 | 5.6e-05 | 0.001256 | 0.32045 | 1.3377 | 2.5473 | 100.63 | 21.15 | 25.59 | 67.5413 | 89.0174 | 64.1762 | 0.00639 | 1.0466 | 0.025217 | 24.0943 | -0.006706 | 0.13911 | -0.000157 | -0.011911 | -0.011911 | -0.48236
2019-03-28 15:38 | 1553801869.2794 | -32 | 0.20956 | 0.000446 | 0 | 101.0967 | 101.2661 | 125.8463 | 21.2612 | 0.001927 | 0.001796 | -151.9925 | 0.23308 | 8.4016 | 0.002777 | 0.004114 | 0.003981 | 5.2e-05 | 0.001256 | 0.30901 | 1.3391 | 2.5473 | 100.63 | 21.16 | 25.587 | 68.1942 | 88.6213 | 65.038 | 0.006394 | 1.0439 | 0.025194 | 24.1334 | -0.014977 | 0.13866 | -0.000439 | -0.0333 | -0.0333 | -0.17197
2019-03-28 15:38 | 1553801870.2794 | -31 | 0.20956 | 0.000445 | 0 | 101.1104 | 101.0423 | 126.4462 | 21.2612 | 0.001971 | 0.001815 | -152.0067 | 0.2359 | 8.4029 | 0.002819 | 0.004312 | 0.004082 | 5.4e-05 | 0.001253 | 0.29302 | 1.3402 | 2.5475 | 100.63 | 21.16 | 25.584 | 69.2447 | 89.629 | 65.5023 | 0.006393 | 1.0425 | 0.025236 | 24.2075 | 0.006031 | 0.14034 | -0.000414 | -0.031495 | -0.031495 | -0.18403
2019-03-28 15:38 | 1553801871.2787 | -30 | 0.20954 | 0.000444 | 0 | 101.1512 | 101.1741 | 126.8199 | 21.2612 | 0.001865 | 0.00181 | -152.0028 | 0.23321 | 8.4027 | 0.002736 | 0.003756 | 0.003602 | 5.3e-05 | 0.001251 | 0.31006 | 1.2885 | 2.5474 | 100.63 | 21.15 | 25.581 | 66.7025 | 87.6373 | 65.3698 | 0.006388 | 1.0429 | 0.025279 | 24.2385 | -0.00202 | 0.13873 | -0.000293 | -0.022325 | -0.022325 | -0.25666
2019-03-28 15:38 | 1553801872.2784 | -29 | 0.20954 | 0.000441 | 0 | 101.1515 | 101.2091 | 127.1262 | 21.2612 | 0.001846 | 0.001774 | -152.0025 | 0.23388 | 8.4043 | 0.002733 | 0.004052 | 0.003779 | 5.6e-05 | 0.001254 | 0.30596 | 1.3087 | 2.5475 | 100.63 | 21.15 | 25.577 | 66.2336 | 87.5771 | 64.5188 | 0.006387 | 1.0455 | 0.025341 | 24.2372 | -0.005087 | 0.13913 | -0.000287 | -0.021906 | -0.021906 | -0.26232
2019-03-28 15:38 | 1553801873.2788 | -28 | 0.20954 | 0.000447 | 0 | 101.1411 | 101.2227 | 127.1208 | 21.2612 | 0.001934 | 0.001792 | -152.0007 | 0.23411 | 8.4035 | 0.002812 | 0.003716 | 0.003484 | 5.4e-05 | 0.001251 | 0.27455 | 1.3418 | 2.5476 | 100.63 | 21.15 | 25.574 | 68.3705 | 89.4588 | 64.9522 | 0.006386 | 1.0442 | 0.025324 | 24.2522 | -0.00721 | 0.13927 | -0.000303 | -0.023123 | -0.023123 | -0.24876
2019-03-28 15:38 | 1553801874.2778 | -27 | 0.20954 | 0.000449 | 0 | 101.1351 | 101.33 | 127.2769 | 21.2612 | 0.001881 | 0.001782 | -152.0073 | 0.23374 | 8.4017 | 0.002746 | 0.004579 | 0.004376 | 5.2e-05 | 0.001253 | 0.29358 | 1.3037 | 2.5474 | 100.63 | 21.15 | 25.563 | 67.0894 | 87.8746 | 64.7007 | 0.006384 | 1.045 | 0.025349 | 24.2581 | -0.017218 | 0.13905 | -0.000298 | -0.02279 | -0.02279 | -0.25199
2019-03-28 15:38 | 1553801875.2785 | -26 | 0.20956 | 0.000448 | 0 | 101.1465 | 101.2345 | 127.0795 | 21.2612 | 0.001885 | 0.001736 | -152.0227 | 0.23362 | 8.404 | 0.002813 | 0.004667 | 0.00451 | 5.4e-05 | 0.001259 | 0.30342 | 1.3314 | 2.5473 | 100.63 | 21.15 | 25.558 | 67.1714 | 89.4884 | 63.6077 | 0.006382 | 1.0484 | 0.025371 | 24.2 | -0.007786 | 0.13898 | -0.000439 | -0.033536 | -0.033536 | -0.17116
2019-03-28 15:38 | 1553801876.2775 | -25 | 0.20952 | 0.000446 | 0 | 101.1765 | 101.3749 | 127.1189 | 21.2612 | 0.001855 | 0.0018 | -152.0081 | 0.23328 | 8.4022 | 0.002763 | 0.003737 | 0.003576 | 5.5e-05 | 0.001256 | 0.32532 | 1.3576 | 2.5475 | 100.63 | 21.15 | 25.552 | 66.4567 | 88.2925 | 65.1292 | 0.006381 | 1.0437 | 0.025317 | 24.2584 | -0.017514 | 0.13878 | -0.000191 | -0.014604 | -0.014604 | -0.39246
2019-03-28 15:38 | 1553801877.2776 | -24 | 0.20952 | 0.000447 | 0 | 101.097 | 101.1931 | 127.4074 | 21.2612 | 0.001919 | 0.001817 | -152.0127 | 0.23457 | 8.4036 | 0.002766 | 0.004364 | 0.004251 | 5.3e-05 | 0.001257 | 0.3126 | 1.3466 | 2.5475 | 100.63 | 21.16 | 25.54 | 67.9913 | 88.3495 | 65.5555 | 0.006382 | 1.0423 | 0.02533 | 24.3012 | -0.008497 | 0.13954 | -0.00017 | -0.012979 | -0.012979 | -0.44405
2019-03-28 15:38 | 1553801878.2779 | -23 | 0.20954 | 0.000445 | 0 | 101.167 | 101.1489 | 127.7377 | 21.2612 | 0.001955 | 0.001777 | -152.0038 | 0.23403 | 8.4022 | 0.002749 | 0.003961 | 0.003777 | 5.5e-05 | 0.001253 | 0.32892 | 1.3361 | 2.5475 | 100.63 | 21.16 | 25.534 | 68.8657 | 87.9473 | 64.5946 | 0.00638 | 1.0453 | 0.025399 | 24.2982 | 0.001598 | 0.13923 | -0.000292 | -0.022291 | -0.022291 | -0.25795
2019-03-28 15:38 | 1553801879.2776 | -22 | 0.20952 | 0.000444 | 0 | 101.1164 | 101.3412 | 127.7729 | 21.2612 | 0.001942 | 0.001776 | -152.0063 | 0.23446 | 8.4025 | 0.0027 | 0.004519 | 0.004323 | 5.6e-05 | 0.001243 | 0.31974 | 1.404 | 2.5475 | 100.63 | 21.16 | 25.529 | 68.5421 | 86.769 | 64.565 | 0.006379 | 1.0454 | 0.025404 | 24.3004 | -0.019867 | 0.13948 | -0.000161 | -0.012297 | -0.012297 | -0.46846
2019-03-28 15:38 | 1553801880.2787 | -21 | 0.20951 | 0.000447 | 0 | 101.1864 | 101.2227 | 127.8938 | 21.2612 | 0.001864 | 0.001789 | -151.9848 | 0.23427 | 8.4047 | 0.002698 | 0.003999 | 0.003849 | 5.1e-05 | 0.001243 | 0.32156 | 1.3671 | 2.5475 | 100.63 | 21.16 | 25.523 | 66.6768 | 86.7199 | 64.8858 | 0.006378 | 1.0444 | 0.025403 | 24.3235 | -0.003207 | 0.13937 | -0.000111 | -0.00847 | -0.00847 | -0.67959
2019-03-28 15:38 | 1553801881.2787 | -20 | 0.20952 | 0.000447 | 0 | 101.1268 | 100.9621 | 127.879 | 21.2612 | 0.001905 | 0.00179 | -151.9775 | 0.235 | 8.402 | 0.002734 | 0.003862 | 0.003734 | 5.5e-05 | 0.001241 | 0.32587 | 1.3297 | 2.5474 | 100.63 | 21.16 | 25.511 | 67.6559 | 87.5795 | 64.8927 | 0.006375 | 1.0444 | 0.025402 | 24.3223 | 0.014581 | 0.1398 | -0.000206 | -0.015782 | -0.015782 | -0.36584
2019-03-28 15:38 | 1553801882.2788 | -19 | 0.20945 | 0.000446 | -1e-06 | 101.0848 | 100.8102 | 128.0658 | 21.2612 | 0.001979 | 0.001772 | -151.9517 | 0.23361 | 8.4041 | 0.002709 | 0.003872 | 0.003623 | 5.4e-05 | 0.001238 | 0.31614 | 1.308 | 2.5474 | 100.63 | 21.16 | 25.506 | 69.4351 | 87 | 64.4596 | 0.006373 | 1.0457 | 0.025437 | 24.3245 | 0.024329 | 0.13897 | 0.000224 | 0.01717 | 0.01717 | 0.33429
2019-03-28 15:38 | 1553801883.2778 | -18 | 0.20951 | 0.000448 | 0 | 101.1238 | 101.1982 | 128.3698 | 21.2612 | 0.001896 | 0.001717 | -151.9562 | 0.23398 | 8.4019 | 0.00267 | 0.003718 | 0.003641 | 5e-05 | 0.00123 | 0.31996 | 1.298 | 2.5473 | 100.63 | 21.16 | 25.497 | 67.4367 | 86.0414 | 63.1415 | 0.006371 | 1.0498 | 0.025517 | 24.3057 | -0.006576 | 0.1392 | -0.000138 | -0.010582 | -0.010582 | -0.54328
2019-03-28 15:38 | 1553801884.2778 | -17 | 0.20956 | 0.000446 | 0 | 101.1089 | 101.2199 | 128.3534 | 21.2612 | 0.001898 | 0.001728 | -151.9637 | 0.23416 | 8.4041 | 0.002622 | 0.004227 | 0.003998 | 5.4e-05 | 0.001228 | 0.34003 | 1.3239 | 2.5474 | 100.63 | 21.16 | 25.492 | 67.4929 | 84.8933 | 63.4139 | 0.00637 | 1.049 | 0.025505 | 24.314 | -0.009815 | 0.1393 | -0.000412 | -0.031656 | -0.031656 | -0.18174
2019-03-28 15:38 | 1553801885.2785 | -16 | 0.20954 | 0.000444 | 0 | 101.0762 | 101.2852 | 128.6026 | 21.2612 | 0.001831 | 0.001714 | -151.9336 | 0.23412 | 8.4027 | 0.002685 | 0.004081 | 0.00388 | 5.3e-05 | 0.001224 | 0.33019 | 1.3389 | 2.5473 | 100.63 | 21.16 | 25.489 | 65.8834 | 86.4006 | 63.0705 | 0.006369 | 1.05 | 0.025542 | 24.3252 | -0.018476 | 0.13928 | -0.000318 | -0.024429 | -0.024429 | -0.23546
2019-03-28 15:38 | 1553801886.2788 | -15 | 0.2095 | 0.000445 | 0 | 101.061 | 101.1591 | 128.7636 | 21.2612 | 0.001856 | 0.001746 | -151.9507 | 0.23377 | 8.4028 | 0.002657 | 0.004292 | 0.004141 | 5.3e-05 | 0.001219 | 0.30735 | 1.2677 | 2.5475 | 100.63 | 21.16 | 25.485 | 66.4911 | 85.7441 | 63.8419 | 0.006368 | 1.0476 | 0.025529 | 24.3683 | -0.008683 | 0.13907 | -7.5e-05 | -0.005808 | -0.005808 | -0.98883
2019-03-28 15:38 | 1553801887.2768 | -14 | 0.20952 | 0.000442 | 0 | 101.0732 | 101.223 | 128.8423 | 21.2612 | 0.001829 | 0.001699 | -151.9586 | 0.23327 | 8.4016 | 0.002691 | 0.004113 | 0.003855 | 5.6e-05 | 0.001224 | 0.31669 | 1.2733 | 2.5473 | 100.63 | 21.16 | 25.485 | 65.8434 | 86.5628 | 62.7141 | 0.006368 | 1.0512 | 0.02558 | 24.335 | -0.013247 | 0.13877 | -0.000182 | -0.014026 | -0.014026 | -0.40863
2019-03-28 15:38 | 1553801888.2769 | -13 | 0.20951 | 0.000448 | 0 | 101.0943 | 101.1204 | 128.682 | 21.2612 | 0.001883 | 0.001711 | -146.235 | 0.23374 | 8.4037 | 0.002684 | 0.004314 | 0.004088 | 5.4e-05 | 0.001225 | 0.31205 | 1.3109 | 2.5475 | 100.63 | 21.16 | 25.486 | 67.1419 | 86.3982 | 62.9884 | 0.006368 | 1.0503 | 0.025553 | 24.3297 | -0.002306 | 0.13905 | -0.000138 | -0.010632 | -0.010632 | -0.54014
2019-03-28 15:38 | 1553801889.2781 | -12 | 0.20949 | 0.000446 | 0 | 101.1012 | 101.1958 | 128.792 | 21.2612 | 0.001867 | 0.001734 | -122.897 | 0.23432 | 8.402 | 0.002709 | 0.004167 | 0.003977 | 5.8e-05 | 0.00123 | 0.335 | 1.3643 | 2.5475 | 100.63 | 21.16 | 25.496 | 66.7384 | 86.9844 | 63.5406 | 0.006371 | 1.0486 | 0.025543 | 24.3601 | -0.008368 | 0.1394 | -2.6e-05 | -0.001971 | -0.001971 | -2.9203
2019-03-28 15:38 | 1553801890.2773 | -11 | 0.20953 | 0.000448 | 0 | 101.0824 | 101.0932 | 128.8625 | 21.2612 | 0.001825 | 0.001704 | -89.9563 | 0.23408 | 8.4027 | 0.002694 | 0.0042 | 0.004055 | 5.7e-05 | 0.001233 | 0.32173 | 1.3378 | 2.5473 | 100.63 | 21.16 | 25.504 | 65.7429 | 86.6292 | 62.8244 | 0.006373 | 1.0508 | 0.025578 | 24.3409 | -0.000956 | 0.13925 | -0.000242 | -0.01866 | -0.01866 | -0.30821
2019-03-28 15:38 | 1553801891.2769 | -10 | 0.2095 | 0.000445 | -1e-06 | 101.0646 | 101.0936 | 128.9528 | 21.2612 | 0.00186 | 0.001748 | -56.9077 | 0.23366 | 8.4026 | 0.002676 | 0.004255 | 0.004071 | 5.3e-05 | 0.001237 | 0.32117 | 1.331 | 2.5473 | 100.63 | 21.16 | 25.536 | 66.5707 | 86.1836 | 63.8903 | 0.006381 | 1.0475 | 0.025546 | 24.388 | -0.002567 | 0.13901 | -6.9e-05 | -0.005301 | -0.005301 | -1.0831
2019-03-28 15:38 | 1553801892.2781 | -9 | 0.2095 | 0.000452 | 0 | 101.0985 | 101.1377 | 128.9294 | 21.2612 | 0.001786 | 0.001719 | -23.862 | 0.23381 | 8.4043 | 0.002679 | 0.003818 | 0.003628 | 5.5e-05 | 0.001232 | 0.32587 | 1.3716 | 2.5474 | 100.63 | 21.16 | 25.56 | 64.7929 | 86.2702 | 63.2009 | 0.006387 | 1.0496 | 0.02557 | 24.3608 | -0.003469 | 0.13909 | -7.7e-05 | -0.005913 | -0.005913 | -0.97147
2019-03-28 15:38 | 1553801893.2784 | -8 | 0.20952 | 0.00044 | 0 | 101.0053 | 101.2468 | 129.0845 | 21.2612 | 0.001838 | 0.001685 | 9.2066 | 0.23818 | 8.4021 | 0.002786 | 0.004262 | 0.004079 | 5.5e-05 | 0.001227 | 0.31653 | 1.3496 | 2.5474 | 100.63 | 21.16 | 25.588 | 66.0597 | 88.8368 | 62.3783 | 0.006394 | 1.0522 | 0.025617 | 24.3456 | -0.02136 | 0.14169 | -0.000171 | -0.013249 | -0.013249 | -0.4417
2019-03-28 15:38 | 1553801894.2781 | -7 | 0.20952 | 0.00044 | 0 | 101.0625 | 101.1955 | 129.3617 | 21.2612 | 0.001902 | 0.001729 | 42.2818 | 0.25199 | 8.4016 | 0.00283 | 0.004607 | 0.004439 | 5.2e-05 | 0.001231 | 0.33915 | 1.3291 | 2.5473 | 100.63 | 21.16 | 25.656 | 67.599 | 89.9065 | 63.4414 | 0.006411 | 1.0489 | 0.025604 | 24.4103 | -0.011764 | 0.14991 | -0.000165 | -0.012718 | -0.012718 | -0.48683
2019-03-28 15:38 | 1553801895.2785 | -6 | 0.20949 | 0.000446 | 0 | 101.0366 | 101.0861 | 129.5739 | 21.2612 | 0.001785 | 0.001709 | 75.3135 | 0.25953 | 8.4042 | 0.002883 | 0.004373 | 0.004261 | 4.9e-05 | 0.001231 | 0.29999 | 1.3178 | 2.5473 | 100.63 | 21.16 | 25.693 | 64.7883 | 91.1741 | 62.9438 | 0.00642 | 1.0504 | 0.025644 | 24.4123 | -0.004384 | 0.1544 | 2.2e-05 | 0.001685 | 0.001685 | 3.7844
2019-03-28 15:38 | 1553801896.2782 | -5 | 0.20951 | 0.000444 | 0 | 101.0783 | 101.0304 | 129.2723 | 21.2612 | 0.001842 | 0.001728 | 178.472 | 0.25932 | 8.4062 | 0.002842 | 0.00395 | 0.003918 | 5.2e-05 | 0.001226 | 0.32477 | 1.2978 | 2.5474 | 100.63 | 21.16 | 25.771 | 66.1383 | 90.1784 | 63.4007 | 0.00644 | 1.049 | 0.025596 | 24.4004 | 0.004238 | 0.15427 | -0.000134 | -0.010348 | -0.010348 | -0.61568
2019-03-28 15:38 | 1553801897.2782 | -4 | 0.20951 | 0.000443 | 0 | 100.9835 | 100.939 | 129.1925 | 21.2612 | 0.001844 | 0.001706 | 178.6748 | 0.25923 | 8.4029 | 0.002903 | 0.004546 | 0.004565 | 5e-05 | 0.001219 | 0.29607 | 1.316 | 2.5473 | 100.63 | 21.15 | 25.809 | 66.1876 | 91.6683 | 62.8742 | 0.006445 | 1.0507 | 0.025608 | 24.3738 | 0.003947 | 0.15422 | -0.000143 | -0.010986 | -0.010986 | -0.57978
2019-03-28 15:38 | 1553801898.2786 | -3 | 0.20952 | 0.00045 | 0 | 100.9844 | 101.1472 | 129.3439 | 21.2612 | 0.001864 | 0.00166 | 178.7747 | 0.26051 | 8.4012 | 0.002909 | 0.004052 | 0.003955 | 5.4e-05 | 0.001214 | 0.31509 | 1.3068 | 2.5473 | 100.63 | 21.15 | 25.846 | 66.6869 | 91.802 | 61.7579 | 0.006454 | 1.0542 | 0.025666 | 24.3475 | -0.014409 | 0.15498 | -0.000173 | -0.013353 | -0.013353 | -0.47933
2019-03-28 15:38 | 1553801899.278 | -2 | 0.20951 | 0.000445 | 0 | 101.0771 | 101.1075 | 129.3387 | 21.2612 | 0.001898 | 0.001711 | 178.7152 | 0.25991 | 8.4022 | 0.00286 | 0.004773 | 0.004652 | 5e-05 | 0.001212 | 0.30525 | 1.2944 | 2.5473 | 100.63 | 21.15 | 25.91 | 67.4898 | 90.6161 | 62.9977 | 0.00647 | 1.0503 | 0.025618 | 24.3921 | -0.002692 | 0.15462 | -0.000129 | -0.009928 | -0.009928 | -0.64324
2019-03-28 15:38 | 1553801900.2775 | -1 | 0.2095 | 0.000444 | 0 | 100.9809 | 101.0113 | 129.0181 | 21.2612 | 0.001823 | 0.001674 | 178.4709 | 0.25977 | 8.4042 | 0.002942 | 0.004275 | 0.004141 | 5e-05 | 0.001213 | 0.2995 | 1.2811 | 2.5472 | 100.63 | 21.15 | 25.937 | 65.6967 | 92.595 | 62.1095 | 0.006477 | 1.0531 | 0.02562 | 24.3296 | -0.0027 | 0.15454 | -8.7e-05 | -0.006714 | -0.006714 | -0.95068
2019-03-28 15:38 | 1553801901.2784 | 0 | 0.20951 | 0.000445 | 0 | 100.8826 | 100.9502 | 128.4874 | 21.2612 | 0.001856 | 0.001726 | 178.642 | 0.25976 | 8.4075 | 0.002963 | 0.005095 | 0.004895 | 5.1e-05 | 0.001217 | 0.32189 | 1.2876 | 2.6962 | 100.63 | 21.15 | 25.98 | 66.4911 | 93.1112 | 63.368 | 0.006488 | 1.0491 | 0.02552 | 24.3251 | -0.005994 | 0.15453 | -0.00012 | -0.009256 | -0.009256 | -0.68955
2019-03-28 15:38 | 1553801902.2783 | 1 | 0.20951 | 0.000447 | 0 | 96.9646 | 100.2436 | 127.072 | 21.2612 | 0.001829 | 0.00169 | 178.6457 | 0.25949 | 8.4017 | 0.002926 | 0.004081 | 0.00331 | 5.1e-05 | 0.001213 | 0.335 | 1.3059 | 2.5473 | 100.63 | 21.15 | 25.995 | 65.8396 | 92.2079 | 62.5018 | 0.006492 | 1.0518 | 0.025412 | 24.1596 | -0.29747 | 0.15437 | -0.000124 | -0.009509 | -0.009509 | -0.67046
2019-03-28 15:38 | 1553801903.2763 | 2 | 0.2095 | 0.000442 | 0 | 93.5861 | 98.4665 | 125.6099 | 21.2612 | 0.001801 | 0.001788 | 178.669 | 0.26011 | 8.4027 | 0.002881 | 0.005107 | 0.005366 | 5.3e-05 | 0.001217 | 0.32444 | 1.3314 | 2.5474 | 100.63 | 21.15 | 26.003 | 65.1636 | 91.1336 | 64.853 | 0.006494 | 1.0445 | 0.025177 | 24.1041 | -0.45469 | 0.15474 | -6.6e-05 | -0.00499 | -0.00499 | -1.2807
2019-03-28 15:38 | 1553801904.2769 | 3 | 0.20949 | 0.000448 | 2e-06 | 93.4253 | 99.414 | 123.6087 | 21.2612 | 0.001781 | 0.001818 | 178.6765 | 0.2593 | 8.4028 | 0.002914 | 0.004575 | 0.004628 | 4.8e-05 | 0.001218 | 0.34307 | 1.3005 | 2.5477 | 100.63 | 21.15 | 26 | 64.675 | 91.9271 | 65.5764 | 0.006493 | 1.0423 | 0.024949 | 23.9369 | -0.55573 | 0.15426 | 1.2e-05 | 0.00079 | 0.00079 | 8.0675
2019-03-28 15:38 | 1553801905.2776 | 4 | 0.20944 | 0.000445 | 5e-06 | 96.4809 | 100.0516 | 121.7003 | 21.2612 | 0.001865 | 0.001796 | 178.6115 | 0.25962 | 8.4008 | 0.003006 | 0.005199 | 0.005265 | 5.2e-05 | 0.001214 | 0.32156 | 1.2987 | 2.5474 | 100.63 | 21.15 | 25.991 | 66.7112 | 94.1417 | 65.0412 | 0.006491 | 1.0439 | 0.024775 | 23.7327 | -0.32505 | 0.15445 | 0.000274 | 0.02008 | 0.02008 | 0.31767
2019-03-28 15:38 | 1553801906.2768 | 5 | 0.2091 | 0.000446 | 1e-05 | 98.4798 | 100.1963 | 119.881 | 21.2612 | 0.001916 | 0.001826 | 178.7175 | 0.26003 | 8.4052 | 0.002933 | 0.004361 | 0.004271 | 5.1e-05 | 0.001207 | 0.37664 | 1.2979 | 2.5474 | 100.63 | 21.15 | 25.979 | 67.9267 | 92.3923 | 65.7756 | 0.006488 | 1.0417 | 0.024563 | 23.5802 | -0.15457 | 0.15469 | 0.002324 | 0.17126 | 0.17126 | 0.037305
2019-03-28 15:38 | 1553801907.2765 | 6 | 0.20836 | 0.000457 | 1.8e-05 | 99.4937 | 99.0053 | 117.4054 | 21.2612 | 0.001975 | 0.001919 | 178.5175 | 0.26013 | 8.4043 | 0.002916 | 0.004891 | 0.004795 | 5.2e-05 | 0.001203 | 0.36857 | 1.329 | 2.5478 | 100.63 | 21.15 | 25.944 | 69.3351 | 91.97 | 68.0046 | 0.006479 | 1.0349 | 0.024228 | 23.412 | 0.044017 | 0.15475 | 0.006809 | 0.49541 | 0.49541 | 0.012901
2019-03-28 15:38 | 1553801908.2775 | 7 | 0.2073 | 0.000474 | 2.8e-05 | 100.3342 | 96.9995 | 114.9478 | 21.2612 | 0.001949 | 0.001999 | 178.4705 | 0.25968 | 8.3994 | 0.002992 | 0.004007 | 0.003851 | 4.9e-05 | 0.001205 | 0.30558 | 1.3573 | 1.7922 | 100.63 | 21.15 | 25.924 | 68.7222 | 93.8122 | 69.9139 | 0.006474 | 1.0291 | 0.023906 | 23.2304 | 0.30233 | 0.15449 | 0.01312 | 0.94164 | 0.94164 | 0.006776
2019-03-28 15:38 | 1553801909.2765 | 8 | 0.20643 | 0.00055 | 4e-05 | 100.5296 | 93.6312 | 112.7039 | 21.2612 | 0.002118 | 0.002076 | 178.3884 | 0.26004 | 8.4029 | 0.002951 | 0.004688 | 0.004611 | 4.2e-05 | 0.001207 | 0.29623 | 1.342 | 2.5474 | 100.63 | 21.15 | 25.883 | 72.7821 | 92.8125 | 71.7672 | 0.006464 | 1.0236 | 0.023608 | 23.0646 | 0.63586 | 0.1547 | 0.018246 | 1.2925 | 1.2925 | 0.004943
2019-03-28 15:38 | 1553801910.2775 | 9 | 0.20581 | 0.000812 | 5.4e-05 | 100.7018 | 88.2136 | 110.7869 | 21.2612 | 0.002074 | 0.002103 | 178.3097 | 0.25925 | 8.4026 | 0.003012 | 0.004473 | 0.004458 | 4.8e-05 | 0.001204 | 0.26808 | 1.3152 | 2.5475 | 100.63 | 21.15 | 25.863 | 71.7157 | 94.2927 | 72.4111 | 0.006459 | 1.0217 | 0.023385 | 22.889 | 1.1843 | 0.15423 | 0.021602 | 1.5148 | 1.5148 | 0.004205
2019-03-28 15:38 | 1553801911.2771 | 10 | 0.20464 | 0.001442 | 6.9e-05 | 100.7923 | 84.3417 | 108.8628 | 21.2612 | 0.002241 | 0.002181 | 178.1932 | 0.2597 | 8.4039 | 0.002966 | 0.00428 | 0.004157 | 4.4e-05 | 0.00121 | 0.30773 | 1.3167 | 2.5474 | 100.63 | 21.15 | 25.843 | 75.7379 | 93.1864 | 74.3015 | 0.006454 | 1.0161 | 0.023118 | 22.7513 | 1.5938 | 0.1545 | 0.027842 | 1.9288 | 1.9288 | 0.003308
2019-03-28 15:39 | 1553801912.2771 | 11 | 0.20333 | 0.002147 | 8.4e-05 | 100.806 | 82.1851 | 107.6852 | 21.2612 | 0.002252 | 0.002279 | 178.0954 | 0.25898 | 8.4023 | 0.002987 | 0.003846 | 0.003805 | 4.8e-05 | 0.001215 | 0.30143 | 1.2865 | 2.5474 | 100.63 | 21.15 | 25.802 | 76.0039 | 93.6832 | 76.6499 | 0.006443 | 1.0093 | 0.022915 | 22.7043 | 1.8267 | 0.15407 | 0.034748 | 2.3847 | 2.3847 | 0.002668
2019-03-28 15:39 | 1553801913.2766 | 12 | 0.20256 | 0.002926 | 9.7e-05 | 100.8334 | 77.6766 | 106.5611 | 21.2612 | 0.002304 | 0.002358 | 177.9695 | 0.25928 | 8.4005 | 0.002996 | 0.004068 | 0.003935 | 4.8e-05 | 0.001213 | 0.30071 | 1.3175 | 2.5474 | 100.63 | 21.15 | 25.781 | 77.2573 | 93.8913 | 78.5528 | 0.006438 | 1.0038 | 0.022733 | 22.6468 | 2.3338 | 0.15425 | 0.038359 | 2.6104 | 2.6104 | 0.00244
2019-03-28 15:39 | 1553801914.2777 | 13 | 0.2015 | 0.003752 | 0.000108 | 100.8489 | 74.5145 | 106.1861 | 21.2612 | 0.002421 | 0.002421 | 177.8821 | 0.25926 | 8.4028 | 0.003075 | 0.004343 | 0.004333 | 4.6e-05 | 0.001212 | 0.28622 | 1.3008 | 2.5474 | 100.64 | 21.15 | 25.739 | 80.066 | 95.8201 | 80.0739 | 0.006427 | 0.9995 | 0.022644 | 22.6557 | 2.7069 | 0.15423 | 0.043663 | 2.9582 | 2.9582 | 0.002153
2019-03-28 15:39 | 1553801915.2761 | 14 | 0.20019 | 0.004575 | 0.000118 | 100.7995 | 72.3182 | 105.8999 | 21.2612 | 0.002471 | 0.00246 | 177.7581 | 0.25998 | 8.4044 | 0.003156 | 0.005097 | 0.004968 | 4.9e-05 | 0.001209 | 0.28224 | 1.3202 | 2.5475 | 100.64 | 21.15 | 25.72 | 81.2638 | 97.7662 | 80.9893 | 0.006422 | 0.99691 | 0.022585 | 22.6545 | 2.9701 | 0.15466 | 0.050466 | 3.4081 | 3.4081 | 0.001874
2019-03-28 15:39 | 1553801916.2772 | 15 | 0.19973 | 0.005217 | 0.000126 | 100.8438 | 71.4315 | 106.0774 | 21.2612 | 0.002531 | 0.002506 | 177.6377 | 0.25939 | 8.4035 | 0.003141 | 0.00439 | 0.004337 | 5.1e-05 | 0.001213 | 0.31924 | 1.2807 | 2.5473 | 100.64 | 21.15 | 25.702 | 82.7191 | 97.3949 | 82.1012 | 0.006418 | 0.99379 | 0.022568 | 22.709 | 3.0844 | 0.15431 | 0.052406 | 3.5356 | 3.5356 | 0.001803
2019-03-28 15:39 | 1553801917.2778 | 16 | 0.19962 | 0.005722 | 0.000131 | 100.6622 | 71.3092 | 106.7954 | 21.2612 | 0.002614 | 0.002508 | 177.4985 | 0.25944 | 8.4052 | 0.003196 | 0.00456 | 0.004454 | 5e-05 | 0.001211 | 0.29463 | 1.274 | 2.5473 | 100.64 | 21.15 | 25.671 | 84.71 | 98.7274 | 82.1592 | 0.00641 | 0.99363 | 0.022642 | 22.7876 | 3.0836 | 0.15434 | 0.052476 | 3.5517 | 3.5517 | 0.001795
2019-03-28 15:39 | 1553801918.2768 | 17 | 0.1996 | 0.005996 | 0.000135 | 100.7646 | 70.0084 | 107.4069 | 21.2612 | 0.002619 | 0.002543 | 177.3419 | 0.25965 | 8.4016 | 0.003151 | 0.004599 | 0.004416 | 5.3e-05 | 0.001209 | 0.30071 | 1.2945 | 2.5475 | 100.64 | 21.15 | 25.66 | 84.8334 | 97.6376 | 83.0061 | 0.006407 | 0.99127 | 0.02268 | 22.88 | 3.2574 | 0.15447 | 0.052265 | 3.5432 | 3.5432 | 0.001801
2019-03-28 15:39 | 1553801919.2761 | 18 | 0.19962 | 0.006225 | 0.000137 | 100.6651 | 68.8567 | 107.9279 | 21.2612 | 0.002659 | 0.00255 | 177.2204 | 0.25924 | 8.4046 | 0.003168 | 0.004561 | 0.004558 | 5.1e-05 | 0.001206 | 0.30558 | 1.2706 | 2.5473 | 100.64 | 21.14 | 25.65 | 85.7886 | 98.0602 | 83.1552 | 0.006401 | 0.99085 | 0.02273 | 22.9402 | 3.3969 | 0.15422 | 0.051849 | 3.5227 | 3.5227 | 0.001808
2019-03-28 15:39 | 1553801920.2769 | 19 | 0.19984 | 0.006432 | 0.000138 | 100.6931 | 69.1152 | 108.3857 | 21.2612 | 0.002697 | 0.00264 | 177.1082 | 0.25939 | 8.4043 | 0.003206 | 0.004617 | 0.00453 | 5.2e-05 | 0.001203 | 0.27113 | 1.3092 | 2.5475 | 100.64 | 21.15 | 25.635 | 86.6949 | 98.9799 | 85.3366 | 0.006401 | 0.98482 | 0.022709 | 23.0591 | 3.3659 | 0.15431 | 0.050323 | 3.4161 | 3.4161 | 0.001866
2019-03-28 15:39 | 1553801921.2769 | 20 | 0.20004 | 0.0066 | 0.000137 | 100.7265 | 70.0077 | 108.9971 | 21.2612 | 0.002747 | 0.002658 | 177.0122 | 0.26008 | 8.4025 | 0.003194 | 0.005075 | 0.004993 | 5.2e-05 | 0.001208 | 0.30807 | 1.3047 | 2.5474 | 100.64 | 21.15 | 25.63 | 87.9028 | 98.6819 | 85.7636 | 0.0064 | 0.98365 | 0.022759 | 23.1378 | 3.2541 | 0.15472 | 0.048917 | 3.3283 | 3.3283 | 0.00192
2019-03-28 15:39 | 1553801922.2764 | 21 | 0.20017 | 0.006728 | 0.000136 | 100.6717 | 70.3892 | 109.6806 | 21.2612 | 0.002777 | 0.002649 | 176.9178 | 0.25931 | 8.4056 | 0.003157 | 0.004641 | 0.004469 | 5.2e-05 | 0.001209 | 0.32211 | 1.2743 | 2.5474 | 100.64 | 21.14 | 25.623 | 88.6175 | 97.7992 | 85.5459 | 0.006394 | 0.98425 | 0.022838 | 23.2032 | 3.2006 | 0.15427 | 0.048005 | 3.2778 | 3.2778 | 0.001944
2019-03-28 15:39 | 1553801923.2764 | 22 | 0.20048 | 0.006818 | 0.000135 | 100.5978 | 71.2596 | 110.3836 | 21.2612 | 0.002813 | 0.002679 | 176.8319 | 0.25992 | 8.4018 | 0.00319 | 0.004544 | 0.004407 | 5.1e-05 | 0.001207 | 0.28512 | 1.3441 | 2.5474 | 100.64 | 21.14 | 25.622 | 89.5002 | 98.5792 | 86.2702 | 0.006394 | 0.98226 | 0.022888 | 23.3009 | 3.0841 | 0.15463 | 0.046011 | 3.1489 | 3.1489 | 0.002028
2019-03-28 15:39 | 1553801924.2761 | 23 | 0.20073 | 0.006862 | 0.000132 | 100.6398 | 71.5969 | 110.6871 | 21.2612 | 0.002845 | 0.002757 | 176.6899 | 0.25965 | 8.4042 | 0.003205 | 0.004713 | 0.004608 | 5.4e-05 | 0.001218 | 0.31852 | 1.3508 | 2.5475 | 100.64 | 21.14 | 25.62 | 90.2644 | 98.9422 | 88.1387 | 0.006393 | 0.97718 | 0.02286 | 23.3935 | 3.0456 | 0.15447 | 0.044468 | 3.04 | 3.04 | 0.002099
2019-03-28 15:39 | 1553801925.2771 | 24 | 0.20099 | 0.006876 | 0.000128 | 100.5832 | 72.045 | 110.9366 | 21.2612 | 0.002879 | 0.002753 | 176.5747 | 0.25906 | 8.3987 | 0.00323 | 0.005023 | 0.004909 | 5.4e-05 | 0.001216 | 0.30198 | 1.3042 | 2.5473 | 100.64 | 21.14 | 25.617 | 91.0803 | 99.5635 | 88.0395 | 0.006393 | 0.97745 | 0.022889 | 23.4166 | 2.9847 | 0.15411 | 0.042937 | 2.9395 | 2.9395 | 0.002165
2019-03-28 15:39 | 1553801926.2761 | 25 | 0.20136 | 0.006851 | 0.000125 | 100.5442 | 72.2988 | 111.2066 | 21.2612 | 0.002839 | 0.00277 | 176.4848 | 0.2594 | 8.403 | 0.003252 | 0.005133 | 0.005023 | 5.3e-05 | 0.001222 | 0.3194 | 1.3308 | 2.5473 | 100.64 | 21.14 | 25.615 | 90.1136 | 100.0749 | 88.4526 | 0.006392 | 0.97633 | 0.022903 | 23.4585 | 2.9498 | 0.15432 | 0.040737 | 2.7911 | 2.7911 | 0.002283
2019-03-28 15:39 | 1553801927.2766 | 26 | 0.20172 | 0.006803 | 0.000121 | 100.5177 | 73.1736 | 111.2958 | 21.2612 | 0.002922 | 0.002816 | 176.3909 | 0.25948 | 8.4034 | 0.003278 | 0.004849 | 0.00472 | 5e-05 | 0.001219 | 0.30381 | 1.3316 | 2.5474 | 100.64 | 21.14 | 25.611 | 92.1243 | 100.7192 | 89.5657 | 0.006391 | 0.97334 | 0.022877 | 23.504 | 2.8399 | 0.15436 | 0.038697 | 2.6487 | 2.6487 | 0.002407
2019-03-28 15:39 | 1553801928.2766 | 27 | 0.20192 | 0.006728 | 0.000118 | 100.5493 | 74.7197 | 111.3753 | 21.2612 | 0.002884 | 0.00278 | 176.3284 | 0.25983 | 8.4041 | 0.003277 | 0.003967 | 0.003937 | 5.2e-05 | 0.001215 | 0.29048 | 1.301 | 2.5474 | 100.64 | 21.14 | 25.608 | 91.2085 | 100.6863 | 88.708 | 0.00639 | 0.97565 | 0.022913 | 23.4846 | 2.6557 | 0.15457 | 0.037571 | 2.5759 | 2.5759 | 0.002478
2019-03-28 15:39 | 1553801929.2763 | 28 | 0.20207 | 0.006644 | 0.000115 | 100.4617 | 76.4417 | 111.6615 | 21.2612 | 0.002951 | 0.002792 | 176.2377 | 0.26007 | 8.4016 | 0.00325 | 0.0043 | 0.004107 | 5.2e-05 | 0.001206 | 0.28926 | 1.3345 | 2.5474 | 100.64 | 21.14 | 25.602 | 92.8202 | 100.0403 | 88.9948 | 0.006389 | 0.97487 | 0.022933 | 23.5241 | 2.4441 | 0.15472 | 0.036791 | 2.5249 | 2.5249 | 0.002531
2019-03-28 15:39 | 1553801930.2773 | 29 | 0.20217 | 0.006552 | 0.000112 | 100.448 | 77.324 | 111.7722 | 21.2612 | 0.002936 | 0.00279 | 176.1422 | 0.2602 | 8.4045 | 0.003234 | 0.004489 | 0.004369 | 5.3e-05 | 0.001199 | 0.30879 | 1.3203 | 2.5474 | 100.64 | 21.14 | 25.588 | 92.4511 | 99.6607 | 88.9415 | 0.006385 | 0.97502 | 0.022946 | 23.534 | 2.3402 | 0.15479 | 0.036304 | 2.4931 | 2.4931 | 0.002564
2019-03-28 15:39 | 1553801931.2773 | 30 | 0.20226 | 0.006462 | 0.000108 | 100.4048 | 78.297 | 112.0638 | 21.2612 | 0.002908 | 0.002876 | 176.0528 | 0.25962 | 8.4046 | 0.003288 | 0.004397 | 0.004373 | 4.6e-05 | 0.001193 | 0.31509 | 1.3062 | 2.5473 | 100.64 | 21.14 | 25.58 | 91.7753 | 100.9437 | 91.0044 | 0.006383 | 0.96949 | 0.022911 | 23.6317 | 2.2245 | 0.15445 | 0.035854 | 2.4587 | 2.4587 | 0.002594
2019-03-28 15:39 | 1553801932.2777 | 31 | 0.20235 | 0.00636 | 0.000105 | 100.3923 | 78.7719 | 112.2656 | 21.2612 | 0.002981 | 0.002834 | 175.9639 | 0.25975 | 8.403 | 0.00328 | 0.004326 | 0.004242 | 5.2e-05 | 0.001193 | 0.31277 | 1.2985 | 2.5472 | 100.63 | 21.14 | 25.57 | 93.5407 | 100.7608 | 89.9908 | 0.006381 | 0.9722 | 0.022963 | 23.62 | 2.1693 | 0.15452 | 0.035464 | 2.4378 | 2.4378 | 0.002618
2019-03-28 15:39 | 1553801933.2763 | 32 | 0.20251 | 0.006272 | 0.000102 | 100.3342 | 79.0342 | 112.7055 | 21.2612 | 0.002956 | 0.002818 | 175.8569 | 0.25987 | 8.4049 | 0.003314 | 0.004971 | 0.004745 | 5.3e-05 | 0.001191 | 0.32925 | 1.2948 | 2.5474 | 100.63 | 21.15 | 25.549 | 92.9415 | 101.5929 | 89.6064 | 0.00638 | 0.97323 | 0.023021 | 23.6537 | 2.1344 | 0.15459 | 0.034605 | 2.3849 | 2.3849 | 0.002677
2019-03-28 15:39 | 1553801934.2753 | 33 | 0.20266 | 0.006186 | 0.0001 | 100.3375 | 79.8111 | 113.2961 | 21.2612 | 0.002985 | 0.002816 | 175.7826 | 0.25929 | 8.4057 | 0.003293 | 0.0046 | 0.004578 | 5e-05 | 0.001188 | 0.30807 | 1.3006 | 2.5474 | 100.63 | 21.15 | 25.538 | 93.6463 | 101.0662 | 89.5712 | 0.006377 | 0.97332 | 0.023082 | 23.7145 | 2.0472 | 0.15425 | 0.03383 | 2.3379 | 2.3379 | 0.002725
2019-03-28 15:39 | 1553801935.2767 | 34 | 0.20293 | 0.0061 | 9.8e-05 | 100.2648 | 80.6112 | 113.7392 | 21.2612 | 0.003002 | 0.00282 | 175.7115 | 0.25968 | 8.4031 | 0.003225 | 0.004812 | 0.004703 | 4.7e-05 | 0.001185 | 0.28943 | 1.2951 | 2.5472 | 100.63 | 21.15 | 25.517 | 94.0375 | 99.423 | 89.6533 | 0.006372 | 0.9731 | 0.023124 | 23.7635 | 1.9515 | 0.15449 | 0.03233 | 2.2386 | 2.2386 | 0.00285
2019-03-28 15:39 | 1553801936.2765 | 35 | 0.2031 | 0.006015 | 9.4e-05 | 100.2249 | 80.0004 | 114.2169 | 21.2612 | 0.003055 | 0.002884 | 175.6087 | 0.25932 | 8.4014 | 0.003243 | 0.004756 | 0.004612 | 5.4e-05 | 0.001192 | 0.31902 | 1.304 | 2.5474 | 100.63 | 21.15 | 25.508 | 95.3237 | 99.8575 | 91.2015 | 0.00637 | 0.96897 | 0.023124 | 23.8641 | 2.016 | 0.15427 | 0.031382 | 2.1731 | 2.1731 | 0.002932
2019-03-28 15:39 | 1553801937.2764 | 36 | 0.20327 | 0.005921 | 9.1e-05 | 100.2475 | 79.2509 | 114.5671 | 21.2612 | 0.00298 | 0.002861 | 175.5489 | 0.25999 | 8.4013 | 0.003241 | 0.004585 | 0.004498 | 5e-05 | 0.001186 | 0.30055 | 1.319 | 2.5474 | 100.63 | 21.15 | 25.493 | 93.5281 | 99.8176 | 90.6363 | 0.006366 | 0.97047 | 0.023177 | 23.8821 | 2.1022 | 0.15467 | 0.030483 | 2.1159 | 2.1159 | 0.003019
2019-03-28 15:39 | 1553801938.2755 | 37 | 0.20346 | 0.005815 | 8.9e-05 | 100.0939 | 78.9578 | 115.1931 | 21.2612 | 0.003026 | 0.002849 | 175.4414 | 0.25907 | 8.405 | 0.003285 | 0.004355 | 0.004217 | 4.8e-05 | 0.001185 | 0.28584 | 1.3101 | 2.5474 | 100.63 | 21.15 | 25.485 | 94.6341 | 100.8871 | 90.3549 | 0.006364 | 0.97123 | 0.023249 | 23.938 | 2.1216 | 0.15412 | 0.029471 | 2.0522 | 2.0522 | 0.003102
2019-03-28 15:39 | 1553801939.2765 | 38 | 0.20371 | 0.005712 | 8.5e-05 | 100.0754 | 79.873 | 115.5797 | 21.2612 | 0.003115 | 0.00281 | 175.3745 | 0.25997 | 8.403 | 0.0032 | 0.004265 | 0.00421 | 4.7e-05 | 0.00119 | 0.29607 | 1.3329 | 2.5475 | 100.63 | 21.15 | 25.477 | 96.7648 | 98.8356 | 89.4299 | 0.006362 | 0.9737 | 0.023318 | 23.9476 | 2.0169 | 0.15466 | 0.02815 | 1.9663 | 1.9663 | 0.003248
2019-03-28 15:39 | 1553801940.2766 | 39 | 0.20399 | 0.005601 | 8.2e-05 | 100.0709 | 81.3239 | 116.0825 | 21.2612 | 0.003061 | 0.002831 | 175.3032 | 0.25927 | 8.4026 | 0.003205 | 0.005214 | 0.00504 | 4.7e-05 | 0.001196 | 0.30735 | 1.3264 | 2.5474 | 100.63 | 21.15 | 25.462 | 95.4757 | 98.9492 | 89.9361 | 0.006358 | 0.97235 | 0.023352 | 24.0164 | 1.8554 | 0.15424 | 0.026591 | 1.8603 | 1.8603 | 0.003424
2019-03-28 15:39 | 1553801941.2759 | 40 | 0.20425 | 0.005485 | 7.9e-05 | 100.0492 | 81.1531 | 116.4904 | 21.2612 | 0.00309 | 0.002852 | 175.225 | 0.26001 | 8.4009 | 0.00315 | 0.004939 | 0.00485 | 4.7e-05 | 0.001203 | 0.31293 | 1.3151 | 2.5475 | 100.63 | 21.15 | 25.453 | 96.1648 | 97.6299 | 90.4263 | 0.006356 | 0.97103 | 0.023377 | 24.0748 | 1.8723 | 0.15468 | 0.025162 | 1.7625 | 1.7625 | 0.003625
2019-03-28 15:39 | 1553801942.2767 | 41 | 0.2045 | 0.005367 | 7.5e-05 | 100.0322 | 81.1398 | 116.7838 | 21.2612 | 0.003087 | 0.002864 | 175.1464 | 0.25857 | 8.4014 | 0.003186 | 0.004628 | 0.004487 | 4.6e-05 | 0.001204 | 0.30901 | 1.31 | 2.5474 | 100.63 | 21.15 | 25.446 | 96.1006 | 98.4985 | 90.7136 | 0.006354 | 0.97027 | 0.023398 | 24.1146 | 1.8723 | 0.15382 | 0.023821 | 1.6702 | 1.6702 | 0.003804
2019-03-28 15:39 | 1553801943.2767 | 42 | 0.20466 | 0.00524 | 7.2e-05 | 99.9279 | 82.0418 | 117.1671 | 21.2612 | 0.003036 | 0.002833 | 175.0696 | 0.25957 | 8.4017 | 0.003115 | 0.004777 | 0.004677 | 4.7e-05 | 0.001207 | 0.29894 | 1.2989 | 2.5474 | 100.63 | 21.15 | 25.436 | 94.8579 | 96.7757 | 89.965 | 0.006352 | 0.97227 | 0.02346 | 24.1293 | 1.7641 | 0.15442 | 0.023028 | 1.6191 | 1.6191 | 0.003939
2019-03-28 15:39 | 1553801944.2767 | 43 | 0.20482 | 0.005112 | 6.9e-05 | 99.8877 | 83.2831 | 117.5584 | 21.2612 | 0.003107 | 0.002779 | 175.0129 | 0.25964 | 8.4052 | 0.003116 | 0.004399 | 0.004329 | 5e-05 | 0.001205 | 0.3115 | 1.3153 | 2.5474 | 100.63 | 21.15 | 25.432 | 96.5704 | 96.7938 | 88.6712 | 0.006351 | 0.97575 | 0.023541 | 24.1264 | 1.6261 | 0.15446 | 0.022202 | 1.5666 | 1.5666 | 0.004072
2019-03-28 15:39 | 1553801945.2767 | 44 | 0.20494 | 0.004978 | 6.6e-05 | 99.8979 | 84.1671 | 117.8546 | 21.2612 | 0.003055 | 0.002848 | 174.9188 | 0.25954 | 8.4 | 0.003095 | 0.004548 | 0.004418 | 5e-05 | 0.001203 | 0.30414 | 1.2668 | 2.7557 | 100.63 | 21.15 | 25.425 | 95.3142 | 96.2877 | 90.3378 | 0.006349 | 0.97127 | 0.023517 | 24.2124 | 1.5326 | 0.1544 | 0.021659 | 1.5268 | 1.5268 | 0.004176
2019-03-28 15:39 | 1553801946.2757 | 45 | 0.20505 | 0.004847 | 6.4e-05 | 99.7209 | 83.7873 | 117.9124 | 21.2612 | 0.003066 | 0.002787 | 174.8571 | 0.25997 | 8.4023 | 0.003072 | 0.004519 | 0.004396 | 4.8e-05 | 0.001203 | 0.3011 | 1.3127 | 2.5474 | 100.63 | 21.15 | 25.424 | 95.5899 | 95.7395 | 88.8667 | 0.006349 | 0.97522 | 0.02357 | 24.1693 | 1.5572 | 0.15466 | 0.021178 | 1.4964 | 1.4964 | 0.004268
2019-03-28 15:39 | 1553801947.2757 | 46 | 0.20524 | 0.004715 | 6.2e-05 | 99.7168 | 84.3648 | 117.9884 | 21.2612 | 0.00305 | 0.002813 | 174.7736 | 0.25961 | 8.4056 | 0.003005 | 0.005129 | 0.005018 | 5.1e-05 | 0.001201 | 0.31094 | 1.2911 | 2.5475 | 100.63 | 21.15 | 25.423 | 95.197 | 94.1111 | 89.4828 | 0.006349 | 0.97356 | 0.023558 | 24.1976 | 1.4954 | 0.15444 | 0.020189 | 1.4259 | 1.4259 | 0.004473
2019-03-28 15:39 | 1553801948.2753 | 47 | 0.20537 | 0.004593 | 6e-05 | 99.7212 | 85.1791 | 117.9689 | 21.2612 | 0.003026 | 0.002815 | 174.7342 | 0.25913 | 8.4036 | 0.002952 | 0.004469 | 0.004348 | 4.8e-05 | 0.001202 | 0.33019 | 1.3395 | 2.5474 | 100.63 | 21.15 | 25.425 | 94.6167 | 92.8436 | 89.5337 | 0.006349 | 0.97342 | 0.023554 | 24.1973 | 1.4099 | 0.15416 | 0.019585 | 1.3831 | 1.3831 | 0.004603
2019-03-28 15:39 | 1553801949.2765 | 48 | 0.20557 | 0.004474 | 5.8e-05 | 99.6739 | 85.6904 | 118.0847 | 21.2612 | 0.003058 | 0.002796 | 174.6666 | 0.25897 | 8.4044 | 0.002991 | 0.003639 | 0.003664 | 5.2e-05 | 0.001202 | 0.29662 | 1.3018 | 2.5474 | 100.63 | 21.15 | 25.426 | 95.4098 | 93.7801 | 89.0735 | 0.006349 | 0.97466 | 0.023581 | 24.1938 | 1.3521 | 0.15406 | 0.018524 | 1.3097 | 1.3097 | 0.004858
2019-03-28 15:39 | 1553801950.2759 | 49 | 0.20564 | 0.004355 | 5.6e-05 | 99.6241 | 86.1361 | 117.8272 | 21.2612 | 0.003032 | 0.002765 | 174.581 | 0.25965 | 8.4035 | 0.00298 | 0.00434 | 0.004161 | 5e-05 | 0.001199 | 0.29336 | 1.3504 | 2.5474 | 100.63 | 21.15 | 25.426 | 94.7812 | 93.5071 | 88.3416 | 0.006349 | 0.97664 | 0.023579 | 24.143 | 1.3012 | 0.15447 | 0.018244 | 1.2899 | 1.2899 | 0.004946
2019-03-28 15:39 | 1553801951.275 | 50 | 0.20569 | 0.004248 | 5.5e-05 | 99.6745 | 86.9559 | 117.6295 | 21.2612 | 0.002984 | 0.002818 | 174.5237 | 0.25923 | 8.4031 | 0.003002 | 0.004425 | 0.004282 | 4.9e-05 | 0.001201 | 0.29927 | 1.2992 | 2.5474 | 100.63 | 21.15 | 25.428 | 93.6034 | 94.0493 | 89.6102 | 0.00635 | 0.97322 | 0.023518 | 24.165 | 1.221 | 0.15422 | 0.018105 | 1.2768 | 1.2768 | 0.004988
2019-03-28 15:39 | 1553801952.2751 | 51 | 0.20574 | 0.004138 | 5.3e-05 | 99.5604 | 87.3884 | 117.4516 | 21.2612 | 0.003008 | 0.002767 | 174.4622 | 0.25987 | 8.4028 | 0.002954 | 0.004635 | 0.004493 | 4.8e-05 | 0.001206 | 0.30502 | 1.283 | 2.5473 | 100.63 | 21.15 | 25.431 | 94.1878 | 92.8993 | 88.3816 | 0.006351 | 0.97653 | 0.02354 | 24.1058 | 1.1664 | 0.1546 | 0.0179 | 1.2637 | 1.2637 | 0.005053
2019-03-28 15:39 | 1553801953.275 | 52 | 0.2058 | 0.004045 | 5.2e-05 | 99.6036 | 87.438 | 117.5479 | 21.2612 | 0.003033 | 0.002801 | 174.4084 | 0.25983 | 8.4026 | 0.002925 | 0.004604 | 0.004563 | 5e-05 | 0.001209 | 0.33196 | 1.282 | 2.5474 | 100.63 | 21.15 | 25.437 | 94.7834 | 92.1811 | 89.2103 | 0.006352 | 0.97429 | 0.023523 | 24.1433 | 1.1652 | 0.15457 | 0.017648 | 1.245 | 1.245 | 0.005128
2019-03-28 15:39 | 1553801954.275 | 53 | 0.20589 | 0.003952 | 5.1e-05 | 99.4928 | 87.5983 | 117.5265 | 21.2612 | 0.003008 | 0.002738 | 174.3431 | 0.25959 | 8.4036 | 0.002935 | 0.004175 | 0.004096 | 5.1e-05 | 0.001211 | 0.32068 | 1.2829 | 2.5473 | 100.63 | 21.15 | 25.44 | 94.187 | 92.4369 | 87.6801 | 0.006353 | 0.97842 | 0.02357 | 24.0901 | 1.1388 | 0.15443 | 0.017248 | 1.2193 | 1.2193 | 0.005231
2019-03-28 15:39 | 1553801955.275 | 54 | 0.2059 | 0.003865 | 4.9e-05 | 99.5407 | 87.9333 | 117.0979 | 21.2612 | 0.003046 | 0.002773 | 174.3075 | 0.25887 | 8.4012 | 0.002888 | 0.004706 | 0.004573 | 5e-05 | 0.001214 | 0.31614 | 1.2739 | 2.5473 | 100.63 | 21.15 | 25.444 | 95.1077 | 91.3015 | 88.5392 | 0.006354 | 0.9761 | 0.023499 | 24.0748 | 1.109 | 0.154 | 0.017294 | 1.2189 | 1.2189 | 0.005218
2019-03-28 15:39 | 1553801956.2751 | 55 | 0.20599 | 0.003791 | 4.7e-05 | 99.441 | 88.2856 | 117.1728 | 21.2612 | 0.003004 | 0.002768 | 174.2564 | 0.26005 | 8.3992 | 0.002915 | 0.00437 | 0.00431 | 5e-05 | 0.001209 | 0.31509 | 1.2636 | 2.5474 | 100.63 | 21.15 | 25.455 | 94.0893 | 91.9536 | 88.4143 | 0.006357 | 0.97644 | 0.023511 | 24.0783 | 1.0643 | 0.1547 | 0.016841 | 1.1877 | 1.1877 | 0.005379
2019-03-28 15:39 | 1553801957.2762 | 56 | 0.20599 | 0.003721 | 4.6e-05 | 99.376 | 88.2197 | 117.4206 | 21.2612 | 0.003033 | 0.002755 | 174.1977 | 0.25897 | 8.4048 | 0.002899 | 0.004452 | 0.004292 | 5.2e-05 | 0.001216 | 0.29645 | 1.284 | 2.5474 | 100.63 | 21.15 | 25.46 | 94.7834 | 91.5541 | 88.1057 | 0.006358 | 0.97727 | 0.023546 | 24.0935 | 1.0651 | 0.15406 | 0.016949 | 1.1972 | 1.1972 | 0.005315
2019-03-28 15:39 | 1553801958.2756 | 57 | 0.20603 | 0.003656 | 4.4e-05 | 99.3516 | 88.2034 | 117.7492 | 21.2612 | 0.003032 | 0.002743 | 174.1325 | 0.25972 | 8.4016 | 0.0028 | 0.004338 | 0.004234 | 5.4e-05 | 0.001212 | 0.29767 | 1.2631 | 2.5473 | 100.63 | 21.15 | 25.467 | 94.7749 | 89.1869 | 87.8043 | 0.00636 | 0.97809 | 0.023589 | 24.1171 | 1.0646 | 0.15451 | 0.01681 | 1.1897 | 1.1897 | 0.005364
2019-03-28 15:39 | 1553801959.2761 | 58 | 0.20608 | 0.0036 | 4.3e-05 | 99.3382 | 88.8509 | 117.4849 | 21.2612 | 0.003047 | 0.002765 | 174.0779 | 0.25891 | 8.4031 | 0.002837 | 0.003729 | 0.003643 | 5.4e-05 | 0.001212 | 0.30165 | 1.3198 | 2.5473 | 100.63 | 21.15 | 25.472 | 95.1366 | 90.0784 | 88.3426 | 0.006361 | 0.97663 | 0.023545 | 24.108 | 0.99795 | 0.15402 | 0.016548 | 1.1689 | 1.1689 | 0.005442
2019-03-28 15:39 | 1553801960.2758 | 59 | 0.20607 | 0.003548 | 4.3e-05 | 99.3007 | 88.9637 | 117.6739 | 21.2612 | 0.002997 | 0.002768 | 174.0369 | 0.25954 | 8.4012 | 0.002845 | 0.004271 | 0.004136 | 5.1e-05 | 0.001213 | 0.31437 | 1.3369 | 2.5473 | 100.63 | 21.15 | 25.477 | 93.9335 | 90.2557 | 88.408 | 0.006362 | 0.97646 | 0.023561 | 24.1295 | 0.98322 | 0.1544 | 0.016655 | 1.1774 | 1.1774 | 0.005416
2019-03-28 15:39 | 1553801961.2762 | 60 | 0.20607 | 0.003499 | 4.2e-05 | 99.3084 | 88.6087 | 117.5488 | 21.2612 | 0.003006 | 0.002748 | 173.96 | 0.25936 | 8.4009 | 0.002831 | 0.005161 | 0.005098 | 5.2e-05 | 0.001208 | 0.31238 | 1.3197 | 2.5474 | 100.63 | 21.15 | 25.489 | 94.1378 | 89.9236 | 87.9292 | 0.006365 | 0.97775 | 0.023564 | 24.1007 | 1.0197 | 0.15429 | 0.016744 | 1.1839 | 1.1839 | 0.005382
2019-03-28 15:39 | 1553801962.2763 | 61 | 0.20613 | 0.003457 | 4.2e-05 | 99.2888 | 88.4422 | 117.7028 | 21.2612 | 0.003053 | 0.002747 | 173.9147 | 0.2596 | 8.3988 | 0.002764 | 0.004808 | 0.004654 | 5.4e-05 | 0.001212 | 0.33356 | 1.3288 | 2.5474 | 100.63 | 21.15 | 25.496 | 95.2776 | 88.3064 | 87.898 | 0.006367 | 0.97784 | 0.023581 | 24.1154 | 1.0347 | 0.15444 | 0.01645 | 1.1639 | 1.1639 | 0.00548
2019-03-28 15:39 | 1553801963.2767 | 62 | 0.20613 | 0.003417 | 4.3e-05 | 99.2116 | 88.7894 | 117.4493 | 21.2612 | 0.003025 | 0.002739 | 173.8523 | 0.2597 | 8.4033 | 0.002838 | 0.00421 | 0.004171 | 5.4e-05 | 0.00121 | 0.28495 | 1.3042 | 2.5474 | 100.63 | 21.15 | 25.511 | 94.6097 | 90.0979 | 87.7044 | 0.006371 | 0.97836 | 0.023562 | 24.083 | 0.99274 | 0.15449 | 0.01651 | 1.1671 | 1.1671 | 0.005467
2019-03-28 15:39 | 1553801964.2767 | 63 | 0.20616 | 0.003385 | 4.4e-05 | 99.1654 | 89.1179 | 117.3645 | 21.2612 | 0.003001 | 0.002695 | 173.8122 | 0.25878 | 8.4036 | 0.002777 | 0.004664 | 0.004469 | 5e-05 | 0.001201 | 0.32278 | 1.3227 | 2.5474 | 100.63 | 21.15 | 25.519 | 94.0329 | 88.6237 | 86.6441 | 0.006373 | 0.98124 | 0.023588 | 24.0389 | 0.95554 | 0.15395 | 0.016362 | 1.1579 | 1.1579 | 0.005491
2019-03-28 15:39 | 1553801965.2744 | 64 | 0.20614 | 0.003355 | 4.4e-05 | 99.1115 | 88.9018 | 117.8039 | 21.2612 | 0.002956 | 0.002692 | 173.7612 | 0.25978 | 8.4 | 0.002822 | 0.00473 | 0.004644 | 5e-05 | 0.001208 | 0.31421 | 1.2767 | 2.5474 | 100.63 | 21.15 | 25.526 | 92.9453 | 89.7136 | 86.5715 | 0.006374 | 0.98144 | 0.023635 | 24.0814 | 0.97239 | 0.15454 | 0.016502 | 1.1701 | 1.1701 | 0.005455
2019-03-28 15:39 | 1553801966.2757 | 65 | 0.20615 | 0.003334 | 4.5e-05 | 99.1199 | 88.9073 | 117.596 | 21.2612 | 0.002972 | 0.002702 | 173.6955 | 0.25871 | 8.4041 | 0.002804 | 0.004663 | 0.004536 | 5.2e-05 | 0.001205 | 0.31222 | 1.2935 | 2.5474 | 100.63 | 21.15 | 25.537 | 93.3224 | 89.2868 | 86.8159 | 0.006377 | 0.98077 | 0.023606 | 24.0684 | 0.97259 | 0.15391 | 0.016482 | 1.1672 | 1.1672 | 0.005446
2019-03-28 15:39 | 1553801967.2744 | 66 | 0.2061 | 0.003303 | 4.5e-05 | 99.0368 | 89.085 | 117.2061 | 21.2612 | 0.003029 | 0.002731 | 173.6594 | 0.25886 | 8.4028 | 0.002808 | 0.003798 | 0.003755 | 5.4e-05 | 0.001205 | 0.30453 | 1.2739 | 2.5473 | 100.63 | 21.15 | 25.542 | 94.6965 | 89.3711 | 87.5249 | 0.006378 | 0.97885 | 0.023543 | 24.0521 | 0.94723 | 0.15399 | 0.016793 | 1.1861 | 1.1861 | 0.005362
2019-03-28 15:39 | 1553801968.2755 | 67 | 0.2061 | 0.00329 | 4.5e-05 | 99.0701 | 89.032 | 116.7263 | 21.2612 | 0.002993 | 0.002741 | 173.5971 | 0.2594 | 8.4018 | 0.002795 | 0.004609 | 0.004521 | 5e-05 | 0.001204 | 0.32156 | 1.2822 | 2.5474 | 100.63 | 21.15 | 25.545 | 93.8272 | 89.0533 | 87.7574 | 0.006379 | 0.97822 | 0.023487 | 24.0105 | 0.95557 | 0.15432 | 0.016818 | 1.185 | 1.185 | 0.005378
2019-03-28 15:39 | 1553801969.2748 | 68 | 0.20609 | 0.003271 | 4.5e-05 | 99.0156 | 88.8444 | 116.8829 | 21.2612 | 0.002936 | 0.002719 | 173.5454 | 0.2591 | 8.4018 | 0.002721 | 0.004565 | 0.004442 | 4.7e-05 | 0.001199 | 0.29645 | 1.2634 | 2.5473 | 100.63 | 21.15 | 25.547 | 92.455 | 87.2859 | 87.2202 | 0.00638 | 0.97967 | 0.023521 | 24.0087 | 0.96951 | 0.15414 | 0.016891 | 1.1918 | 1.1918 | 0.005341
2019-03-28 15:39 | 1553801970.2755 | 69 | 0.20613 | 0.003258 | 4.6e-05 | 98.9987 | 88.8227 | 116.6977 | 21.2612 | 0.003001 | 0.002697 | 173.4604 | 0.26006 | 8.4019 | 0.002715 | 0.004101 | 0.004029 | 4.6e-05 | 0.001204 | 0.31045 | 1.2711 | 2.5473 | 100.63 | 21.15 | 25.548 | 94.0196 | 87.1211 | 86.6939 | 0.00638 | 0.98111 | 0.023519 | 23.9722 | 0.97016 | 0.15471 | 0.016696 | 1.178 | 1.178 | 0.005424
2019-03-28 15:40 | 1553801971.2749 | 70 | 0.2061 | 0.003252 | 4.6e-05 | 98.973 | 88.806 | 116.7702 | 21.2612 | 0.002978 | 0.002724 | 173.4385 | 0.2592 | 8.4012 | 0.002687 | 0.004022 | 0.003926 | 4.9e-05 | 0.001204 | 0.31758 | 1.2789 | 2.5475 | 100.63 | 21.15 | 25.546 | 93.4741 | 86.4709 | 87.3507 | 0.006379 | 0.97932 | 0.023505 | 24.0015 | 0.96952 | 0.15419 | 0.016843 | 1.1876 | 1.1876 | 0.005362
2019-03-28 15:40 | 1553801972.2756 | 71 | 0.20608 | 0.003237 | 4.5e-05 | 98.9486 | 88.4541 | 116.8683 | 21.2612 | 0.003037 | 0.002746 | 173.3446 | 0.25968 | 8.4034 | 0.002667 | 0.004236 | 0.004075 | 5.1e-05 | 0.001205 | 0.31045 | 1.3323 | 2.5474 | 100.63 | 21.15 | 25.544 | 94.8922 | 85.9697 | 87.8871 | 0.006379 | 0.97786 | 0.023498 | 24.0295 | 1.0028 | 0.15448 | 0.016994 | 1.198 | 1.198 | 0.005326
2019-03-28 15:40 | 1553801973.2757 | 72 | 0.20606 | 0.003226 | 4.6e-05 | 98.8971 | 88.6379 | 116.9772 | 21.2612 | 0.003001 | 0.00275 | 173.2968 | 0.25884 | 8.4041 | 0.002743 | 0.004298 | 0.004246 | 4.6e-05 | 0.001201 | 0.30237 | 1.3113 | 2.5474 | 100.63 | 21.16 | 25.539 | 94.0158 | 87.8036 | 87.9856 | 0.006382 | 0.9776 | 0.023505 | 24.0439 | 0.97961 | 0.15399 | 0.017116 | 1.2069 | 1.2069 | 0.005269
2019-03-28 15:40 | 1553801974.2747 | 73 | 0.20604 | 0.003225 | 4.6e-05 | 98.8956 | 88.6138 | 117.0174 | 21.2612 | 0.003025 | 0.00271 | 173.2241 | 0.25894 | 8.4008 | 0.002646 | 0.00443 | 0.004246 | 5.1e-05 | 0.001194 | 0.29463 | 1.2955 | 2.5474 | 100.63 | 21.15 | 25.535 | 94.6097 | 85.4677 | 87.001 | 0.006377 | 0.98027 | 0.023541 | 24.0152 | 0.98191 | 0.15404 | 0.017241 | 1.2176 | 1.2176 | 0.005225
2019-03-28 15:40 | 1553801975.275 | 74 | 0.20604 | 0.003221 | 4.7e-05 | 98.8161 | 87.8874 | 116.7817 | 21.2612 | 0.003027 | 0.002721 | 173.1664 | 0.25842 | 8.4021 | 0.002651 | 0.004822 | 0.004778 | 4.6e-05 | 0.001191 | 0.31078 | 1.2894 | 2.5472 | 100.63 | 21.16 | 25.528 | 94.6488 | 85.5856 | 87.2874 | 0.006379 | 0.97949 | 0.023508 | 24.0006 | 1.0483 | 0.15374 | 0.017276 | 1.2183 | 1.2183 | 0.005211
2019-03-28 15:40 | 1553801976.2748 | 75 | 0.20605 | 0.003221 | 4.7e-05 | 98.8298 | 88.1106 | 116.7061 | 21.2612 | 0.003009 | 0.002718 | 173.1216 | 0.25975 | 8.4021 | 0.002722 | 0.003722 | 0.003606 | 5e-05 | 0.001188 | 0.31238 | 1.3297 | 2.5473 | 100.63 | 21.16 | 25.523 | 94.2261 | 87.3085 | 87.1993 | 0.006378 | 0.97973 | 0.023504 | 23.9899 | 1.0269 | 0.15452 | 0.017207 | 1.2132 | 1.2132 | 0.00526
2019-03-28 15:40 | 1553801977.2757 | 76 | 0.20604 | 0.003218 | 4.7e-05 | 98.8226 | 87.8062 | 116.8923 | 21.2612 | 0.002999 | 0.002698 | 173.0162 | 0.2596 | 8.4019 | 0.002715 | 0.003588 | 0.003491 | 4.8e-05 | 0.001187 | 0.32549 | 1.3081 | 2.5473 | 100.63 | 21.16 | 25.52 | 93.9625 | 87.1266 | 86.7207 | 0.006377 | 0.98103 | 0.023538 | 23.9931 | 1.0572 | 0.15443 | 0.01725 | 1.218 | 1.218 | 0.005236
2019-03-28 15:40 | 1553801978.2762 | 77 | 0.20599 | 0.003222 | 4.6e-05 | 98.8584 | 87.7515 | 116.8878 | 21.2612 | 0.002998 | 0.002754 | 172.9726 | 0.259 | 8.4036 | 0.002748 | 0.004184 | 0.004067 | 5.3e-05 | 0.001181 | 0.32211 | 1.312 | 2.5475 | 100.63 | 21.16 | 25.512 | 93.9429 | 87.9331 | 88.0799 | 0.006375 | 0.97734 | 0.023493 | 24.0379 | 1.066 | 0.15408 | 0.017592 | 1.2398 | 1.2398 | 0.005132
2019-03-28 15:40 | 1553801979.275 | 78 | 0.20598 | 0.003225 | 4.6e-05 | 98.7973 | 87.884 | 117.1181 | 21.2612 | 0.003048 | 0.002736 | 172.9 | 0.25808 | 8.4035 | 0.002679 | 0.004359 | 0.004244 | 5.2e-05 | 0.001179 | 0.31669 | 1.2974 | 2.5475 | 100.63 | 21.16 | 25.507 | 95.1678 | 86.2724 | 87.6286 | 0.006374 | 0.97857 | 0.023531 | 24.0465 | 1.047 | 0.15353 | 0.017636 | 1.2449 | 1.2449 | 0.005093
2019-03-28 15:40 | 1553801980.2738 | 79 | 0.20595 | 0.003219 | 4.6e-05 | 98.7776 | 87.8487 | 117.3749 | 21.2612 | 0.002992 | 0.002771 | 172.8856 | 0.26056 | 8.402 | 0.0027 | 0.004119 | 0.004061 | 4.4e-05 | 0.001178 | 0.31653 | 1.3188 | 2.5474 | 100.63 | 21.16 | 25.5 | 93.8084 | 86.7673 | 88.4697 | 0.006372 | 0.97629 | 0.023529 | 24.1009 | 1.0488 | 0.15501 | 0.017786 | 1.2554 | 1.2554 | 0.005099
2019-03-28 15:40 | 1553801981.2752 | 80 | 0.20594 | 0.003223 | 4.5e-05 | 98.761 | 87.776 | 117.5401 | 21.2612 | 0.002998 | 0.002723 | 172.8027 | 0.25918 | 8.401 | 0.002698 | 0.004765 | 0.004713 | 4.6e-05 | 0.001182 | 0.30021 | 1.3019 | 2.5475 | 100.63 | 21.16 | 25.497 | 93.9444 | 86.7284 | 87.3141 | 0.006371 | 0.97942 | 0.023584 | 24.0793 | 1.0547 | 0.15419 | 0.017878 | 1.2649 | 1.2649 | 0.005034
2019-03-28 15:40 | 1553801982.2739 | 81 | 0.20595 | 0.003228 | 4.6e-05 | 98.7038 | 87.384 | 117.5525 | 21.2612 | 0.003025 | 0.002763 | 172.7649 | 0.25959 | 8.4056 | 0.002654 | 0.004371 | 0.004239 | 5.2e-05 | 0.001185 | 0.28047 | 1.3299 | 2.5474 | 100.63 | 21.16 | 25.495 | 94.6027 | 85.6669 | 88.2949 | 0.006371 | 0.97676 | 0.023553 | 24.1133 | 1.0895 | 0.15443 | 0.017814 | 1.2587 | 1.2587 | 0.005067
2019-03-28 15:40 | 1553801983.2753 | 82 | 0.20592 | 0.003231 | 4.5e-05 | 98.7547 | 87.8123 | 117.2129 | 21.2612 | 0.002993 | 0.002726 | 172.7185 | 0.25917 | 8.4034 | 0.002663 | 0.004434 | 0.004291 | 4.8e-05 | 0.001186 | 0.30702 | 1.297 | 2.5474 | 100.63 | 21.16 | 25.492 | 93.8318 | 85.8745 | 87.4046 | 0.00637 | 0.97917 | 0.023548 | 24.0488 | 1.0504 | 0.15418 | 0.017959 | 1.2688 | 1.2688 | 0.005019
2019-03-28 15:40 | 1553801984.275 | 83 | 0.20591 | 0.003239 | 4.6e-05 | 98.6937 | 87.8701 | 116.8707 | 21.2612 | 0.003061 | 0.002736 | 172.6494 | 0.25994 | 8.401 | 0.002671 | 0.004307 | 0.004257 | 4.8e-05 | 0.001188 | 0.30055 | 1.2754 | 2.5474 | 100.63 | 21.15 | 25.49 | 95.481 | 86.0797 | 87.6474 | 0.006365 | 0.97851 | 0.023506 | 24.0217 | 1.039 | 0.15464 | 0.018036 | 1.2718 | 1.2718 | 0.005022
2019-03-28 15:40 | 1553801985.2744 | 84 | 0.20586 | 0.003247 | 4.6e-05 | 98.7565 | 87.7444 | 116.9416 | 21.2612 | 0.002976 | 0.002748 | 172.5891 | 0.25901 | 8.4034 | 0.002633 | 0.003521 | 0.003502 | 5.2e-05 | 0.001193 | 0.31277 | 1.314 | 2.5473 | 100.63 | 21.16 | 25.49 | 93.4256 | 85.1697 | 87.9324 | 0.006369 | 0.97774 | 0.023503 | 24.0385 | 1.0575 | 0.15409 | 0.018331 | 1.2925 | 1.2925 | 0.004924
2019-03-28 15:40 | 1553801986.2741 | 85 | 0.20584 | 0.003255 | 4.7e-05 | 98.716 | 87.1251 | 117.1106 | 21.2612 | 0.002989 | 0.002704 | 172.5168 | 0.25856 | 8.4019 | 0.002669 | 0.004113 | 0.004009 | 4.9e-05 | 0.001194 | 0.31006 | 1.3296 | 2.5474 | 100.63 | 21.16 | 25.487 | 93.7449 | 86.0234 | 86.8713 | 0.006369 | 0.98062 | 0.023555 | 24.0205 | 1.1172 | 0.15382 | 0.018424 | 1.3018 | 1.3018 | 0.00488
2019-03-28 15:40 | 1553801987.2751 | 86 | 0.20581 | 0.00327 | 4.8e-05 | 98.7056 | 87.1512 | 117.0235 | 21.2612 | 0.002984 | 0.002776 | 172.4831 | 0.25966 | 8.4028 | 0.002658 | 0.004355 | 0.004279 | 4.7e-05 | 0.001199 | 0.30431 | 1.3217 | 2.5474 | 100.63 | 21.16 | 25.486 | 93.6096 | 85.773 | 88.6095 | 0.006368 | 0.97591 | 0.02349 | 24.0694 | 1.1136 | 0.15447 | 0.018577 | 1.3089 | 1.3089 | 0.004874
2019-03-28 15:40 | 1553801988.2751 | 87 | 0.2058 | 0.003283 | 4.9e-05 | 98.6168 | 87.8426 | 116.8789 | 21.2612 | 0.003019 | 0.002684 | 172.4331 | 0.25892 | 8.4046 | 0.002669 | 0.00453 | 0.004444 | 5.2e-05 | 0.001201 | 0.30342 | 1.3668 | 2.5476 | 100.63 | 21.16 | 25.483 | 94.4625 | 86.015 | 86.3944 | 0.006368 | 0.98192 | 0.023547 | 23.9808 | 1.0348 | 0.15403 | 0.018649 | 1.3172 | 1.3172 | 0.00483
2019-03-28 15:40 | 1553801989.274 | 88 | 0.2058 | 0.003297 | 4.9e-05 | 98.6969 | 87.4563 | 116.6216 | 21.2612 | 0.003065 | 0.002749 | 172.3946 | 0.25968 | 8.4011 | 0.002647 | 0.004326 | 0.004244 | 4.9e-05 | 0.001204 | 0.30038 | 1.3311 | 2.5475 | 100.63 | 21.16 | 25.481 | 95.5657 | 85.4997 | 87.962 | 0.006367 | 0.97766 | 0.02347 | 24.0066 | 1.0815 | 0.15448 | 0.01861 | 1.3101 | 1.3101 | 0.00487
2019-03-28 15:40 | 1553801990.2745 | 89 | 0.2058 | 0.003306 | 4.9e-05 | 98.7071 | 87.5477 | 116.3472 | 21.2612 | 0.003 | 0.002761 | 172.2942 | 0.25918 | 8.4055 | 0.002707 | 0.003662 | 0.003561 | 4.6e-05 | 0.0012 | 0.31222 | 1.3206 | 2.5474 | 100.63 | 21.16 | 25.479 | 93.9953 | 86.9346 | 88.241 | 0.006367 | 0.97691 | 0.023434 | 23.9876 | 1.0731 | 0.15418 | 0.018591 | 1.3067 | 1.3067 | 0.004873
2019-03-28 15:40 | 1553801991.2744 | 90 | 0.20575 | 0.003316 | 4.9e-05 | 98.7321 | 87.4305 | 116.3306 | 21.2612 | 0.00298 | 0.002709 | 172.2286 | 0.2591 | 8.4014 | 0.002646 | 0.004822 | 0.004694 | 4.6e-05 | 0.001192 | 0.32875 | 1.2961 | 2.5472 | 100.63 | 21.15 | 25.474 | 93.5047 | 85.4739 | 86.9875 | 0.006361 | 0.98031 | 0.023473 | 23.9442 | 1.0874 | 0.15414 | 0.018859 | 1.3279 | 1.3279 | 0.004794
2019-03-28 15:40 | 1553801992.2745 | 91 | 0.20577 | 0.003325 | 4.9e-05 | 98.6779 | 87.7879 | 116.1531 | 21.2612 | 0.003038 | 0.002724 | 172.1891 | 0.2596 | 8.4012 | 0.002646 | 0.00384 | 0.003696 | 4.7e-05 | 0.001191 | 0.34058 | 1.2997 | 2.5474 | 100.63 | 21.15 | 25.471 | 94.9127 | 85.467 | 87.3538 | 0.006361 | 0.97931 | 0.023443 | 23.9381 | 1.0459 | 0.15444 | 0.018774 | 1.3202 | 1.3202 | 0.004831
2019-03-28 15:40 | 1553801993.2756 | 92 | 0.20575 | 0.003333 | 4.9e-05 | 98.7723 | 87.8236 | 116.2813 | 21.2612 | 0.002995 | 0.002788 | 172.1079 | 0.2591 | 8.4039 | 0.002591 | 0.004328 | 0.004244 | 4.5e-05 | 0.001181 | 0.31078 | 1.2674 | 2.5472 | 100.63 | 21.15 | 25.466 | 93.8795 | 84.1507 | 88.8838 | 0.006359 | 0.97517 | 0.023406 | 24.0021 | 1.0509 | 0.15414 | 0.018843 | 1.323 | 1.323 | 0.004812
2019-03-28 15:40 | 1553801994.2743 | 93 | 0.20576 | 0.003343 | 4.8e-05 | 98.7881 | 87.7621 | 116.2505 | 21.2612 | 0.003001 | 0.002749 | 172.0656 | 0.25918 | 8.403 | 0.002641 | 0.004246 | 0.004173 | 4.8e-05 | 0.001176 | 0.33069 | 1.2745 | 2.5472 | 100.63 | 21.15 | 25.464 | 94.0141 | 85.3508 | 87.9487 | 0.006359 | 0.9777 | 0.023433 | 23.9679 | 1.0586 | 0.15419 | 0.01878 | 1.3201 | 1.3201 | 0.004824
2019-03-28 15:40 | 1553801995.2743 | 94 | 0.20575 | 0.003352 | 4.9e-05 | 98.775 | 87.7081 | 116.4202 | 21.2612 | 0.003069 | 0.002768 | 171.9833 | 0.25949 | 8.4015 | 0.002659 | 0.003918 | 0.003769 | 4.7e-05 | 0.001167 | 0.32764 | 1.2595 | 2.5473 | 100.63 | 21.15 | 25.462 | 95.662 | 85.7862 | 88.4112 | 0.006358 | 0.97645 | 0.023435 | 24.0007 | 1.0629 | 0.15437 | 0.018812 | 1.3225 | 1.3225 | 0.004821
2019-03-28 15:40 | 1553801996.2743 | 95 | 0.2057 | 0.003366 | 4.9e-05 | 98.7726 | 87.636 | 116.2933 | 21.2612 | 0.003032 | 0.002799 | 171.9246 | 0.25957 | 8.4008 | 0.002682 | 0.004298 | 0.004232 | 4.5e-05 | 0.001165 | 0.32405 | 1.2736 | 2.5473 | 100.63 | 21.15 | 25.458 | 94.7819 | 86.335 | 89.1462 | 0.006357 | 0.97447 | 0.023399 | 24.012 | 1.07 | 0.15442 | 0.019091 | 1.34 | 1.34 | 0.004759
2019-03-28 15:40 | 1553801997.2738 | 96 | 0.20568 | 0.003375 | 5e-05 | 98.7511 | 87.4638 | 116.2979 | 21.2612 | 0.003025 | 0.002756 | 171.8651 | 0.25965 | 8.4061 | 0.002638 | 0.003927 | 0.003841 | 4.7e-05 | 0.001169 | 0.31045 | 1.2754 | 2.5474 | 100.63 | 21.15 | 25.455 | 94.6114 | 85.2875 | 88.1088 | 0.006357 | 0.97726 | 0.023433 | 23.9781 | 1.0857 | 0.15447 | 0.019232 | 1.3518 | 1.3518 | 0.004719
2019-03-28 15:40 | 1553801998.2744 | 97 | 0.20561 | 0.003385 | 5e-05 | 98.7616 | 87.5888 | 116.6101 | 21.2612 | 0.003082 | 0.002792 | 171.7978 | 0.25951 | 8.4029 | 0.002578 | 0.004285 | 0.004173 | 4.6e-05 | 0.001172 | 0.31526 | 1.2837 | 2.5472 | 100.63 | 21.15 | 25.453 | 95.9768 | 83.8277 | 88.9837 | 0.006356 | 0.9749 | 0.023436 | 24.0393 | 1.0738 | 0.15438 | 0.019672 | 1.3829 | 1.3829 | 0.004611
2019-03-28 15:40 | 1553801999.2735 | 98 | 0.20561 | 0.003399 | 5e-05 | 98.8736 | 87.3208 | 116.486 | 21.2612 | 0.003071 | 0.002822 | 171.7214 | 0.25845 | 8.4044 | 0.002588 | 0.004019 | 0.003975 | 4.4e-05 | 0.001167 | 0.31996 | 1.2932 | 2.5474 | 100.63 | 21.15 | 25.454 | 95.7067 | 84.0749 | 89.7189 | 0.006356 | 0.97293 | 0.0234 | 24.0509 | 1.1114 | 0.15375 | 0.01964 | 1.3785 | 1.3785 | 0.004606
2019-03-28 15:40 | 1553802000.2735 | 99 | 0.20558 | 0.003413 | 5e-05 | 98.8244 | 87.5922 | 116.5991 | 21.2612 | 0.00303 | 0.00276 | 171.687 | 0.25873 | 8.4038 | 0.002617 | 0.003991 | 0.003921 | 4.9e-05 | 0.001168 | 0.32836 | 1.2763 | 2.5475 | 100.63 | 21.15 | 25.454 | 94.7265 | 84.7709 | 88.2049 | 0.006356 | 0.97701 | 0.02346 | 24.0123 | 1.0792 | 0.15392 | 0.019797 | 1.3931 | 1.3931 | 0.004563
2019-03-28 15:40 | 1553802001.2745 | 100 | 0.20554 | 0.003425 | 5.1e-05 | 98.8986 | 87.654 | 116.4493 | 21.2612 | 0.003101 | 0.002756 | 171.6428 | 0.25906 | 8.4053 | 0.002648 | 0.003586 | 0.003463 | 4.6e-05 | 0.001168 | 0.32925 | 1.335 | 2.5474 | 100.63 | 21.16 | 25.459 | 96.4326 | 85.5216 | 88.1286 | 0.006362 | 0.97721 | 0.023448 | 23.9944 | 1.0796 | 0.15411 | 0.020014 | 1.4076 | 1.4076 | 0.004522
2019-03-28 15:40 | 1553802002.2749 | 101 | 0.2055 | 0.003443 | 5.1e-05 | 98.861 | 87.8412 | 116.4036 | 21.2612 | 0.003096 | 0.002743 | 171.5562 | 0.25842 | 8.4011 | 0.002681 | 0.003622 | 0.003536 | 4.9e-05 | 0.001163 | 0.29518 | 1.2887 | 2.5473 | 100.63 | 21.16 | 25.461 | 96.3184 | 86.321 | 87.7949 | 0.006362 | 0.97811 | 0.023454 | 23.9786 | 1.0571 | 0.15373 | 0.020205 | 1.4213 | 1.4213 | 0.004467
2019-03-28 15:40 | 1553802003.2736 | 102 | 0.20546 | 0.003461 | 5.2e-05 | 98.9242 | 87.7984 | 116.4219 | 21.2612 | 0.00307 | 0.002767 | 171.5071 | 0.25901 | 8.4045 | 0.002687 | 0.004137 | 0.00398 | 4.6e-05 | 0.001164 | 0.32388 | 1.2967 | 2.5475 | 100.63 | 21.15 | 25.464 | 95.6965 | 86.4678 | 88.3736 | 0.006359 | 0.97655 | 0.023437 | 23.9997 | 1.0672 | 0.15409 | 0.020421 | 1.4355 | 1.4355 | 0.004433
2019-03-28 15:40 | 1553802004.2736 | 103 | 0.20546 | 0.003481 | 5.3e-05 | 98.8718 | 87.4781 | 116.5601 | 21.2612 | 0.003024 | 0.002781 | 171.4295 | 0.25922 | 8.4007 | 0.00265 | 0.004324 | 0.004179 | 4.7e-05 | 0.00116 | 0.30397 | 1.3259 | 2.5475 | 100.63 | 21.16 | 25.475 | 94.5752 | 85.5584 | 88.7198 | 0.006366 | 0.97561 | 0.02344 | 24.0254 | 1.0951 | 0.15421 | 0.020438 | 1.4367 | 1.4367 | 0.004433
2019-03-28 15:40 | 1553802005.274 | 104 | 0.20543 | 0.0035 | 5.3e-05 | 98.8247 | 87.2742 | 116.0611 | 21.2612 | 0.003099 | 0.002801 | 171.3751 | 0.25923 | 8.4046 | 0.002717 | 0.003942 | 0.00392 | 4.5e-05 | 0.001168 | 0.32947 | 1.3081 | 2.5473 | 100.63 | 21.16 | 25.479 | 96.3973 | 87.1844 | 89.2134 | 0.006367 | 0.97428 | 0.023373 | 23.9903 | 1.1117 | 0.15421 | 0.020562 | 1.4414 | 1.4414 | 0.004419
2019-03-28 15:40 | 1553802006.2746 | 105 | 0.20545 | 0.003519 | 5.4e-05 | 98.8688 | 87.0374 | 115.9043 | 21.2612 | 0.003079 | 0.002772 | 171.3199 | 0.25955 | 8.3998 | 0.002769 | 0.003685 | 0.003657 | 4.6e-05 | 0.001168 | 0.3235 | 1.2926 | 2.5473 | 100.63 | 21.16 | 25.489 | 95.9101 | 88.4369 | 88.5128 | 0.006369 | 0.97617 | 0.02338 | 23.9509 | 1.14 | 0.1544 | 0.020436 | 1.4329 | 1.4329 | 0.00445
2019-03-28 15:40 | 1553802007.2748 | 106 | 0.20542 | 0.00354 | 5.4e-05 | 98.8602 | 87.3724 | 116.0222 | 21.2612 | 0.00308 | 0.002747 | 171.2554 | 0.259 | 8.4026 | 0.00269 | 0.004385 | 0.004216 | 4.3e-05 | 0.001169 | 0.2975 | 1.295 | 2.5474 | 100.63 | 21.16 | 25.493 | 95.9307 | 86.5339 | 87.9021 | 0.00637 | 0.97782 | 0.023412 | 23.9428 | 1.1049 | 0.15408 | 0.020588 | 1.4455 | 1.4455 | 0.004402
2019-03-28 15:40 | 1553802008.2747 | 107 | 0.20535 | 0.003559 | 5.5e-05 | 98.9346 | 87.9302 | 115.8684 | 21.2612 | 0.003047 | 0.002832 | 171.1818 | 0.25926 | 8.4012 | 0.002699 | 0.004673 | 0.004546 | 4.6e-05 | 0.001174 | 0.32405 | 1.3572 | 2.5475 | 100.63 | 21.16 | 25.497 | 95.1405 | 86.7502 | 89.9573 | 0.006371 | 0.97229 | 0.02333 | 23.9949 | 1.0547 | 0.15423 | 0.020979 | 1.4678 | 1.4678 | 0.00434
2019-03-28 15:40 | 1553802009.2748 | 108 | 0.2053 | 0.003575 | 5.6e-05 | 98.7738 | 87.9363 | 115.713 | 21.2612 | 0.003035 | 0.00276 | 171.1237 | 0.25958 | 8.4023 | 0.002684 | 0.004566 | 0.004473 | 4.5e-05 | 0.001171 | 0.31222 | 1.3235 | 2.5475 | 100.63 | 21.16 | 25.506 | 94.8493 | 86.3903 | 88.226 | 0.006373 | 0.97695 | 0.02337 | 23.9216 | 1.0395 | 0.15442 | 0.021258 | 1.4898 | 1.4898 | 0.004281
2019-03-28 15:40 | 1553802010.2738 | 109 | 0.20533 | 0.003593 | 5.6e-05 | 98.7976 | 88.042 | 115.4973 | 21.2612 | 0.003067 | 0.002798 | 171.0624 | 0.25906 | 8.4031 | 0.002675 | 0.004242 | 0.004152 | 4.5e-05 | 0.001169 | 0.32493 | 1.2966 | 2.5475 | 100.63 | 21.16 | 25.509 | 95.6125 | 86.1718 | 89.1361 | 0.006374 | 0.97449 | 0.023319 | 23.9294 | 1.0309 | 0.15412 | 0.021054 | 1.4722 | 1.4722 | 0.004323
2019-03-28 15:40 | 1553802011.2732 | 110 | 0.20533 | 0.003617 | 5.6e-05 | 98.7773 | 87.8694 | 115.2804 | 21.2612 | 0.003054 | 0.002776 | 170.9709 | 0.25849 | 8.4026 | 0.002778 | 0.004204 | 0.004067 | 4.2e-05 | 0.001168 | 0.34418 | 1.3462 | 2.5475 | 100.63 | 21.16 | 25.512 | 95.3128 | 88.6391 | 88.5987 | 0.006375 | 0.97594 | 0.023314 | 23.8892 | 1.0467 | 0.15378 | 0.021012 | 1.469 | 1.469 | 0.004323
2019-03-28 15:40 | 1553802012.2729 | 111 | 0.20535 | 0.003627 | 5.6e-05 | 98.8068 | 87.7108 | 115.517 | 21.2612 | 0.003109 | 0.002826 | 170.9002 | 0.25921 | 8.4025 | 0.002742 | 0.003677 | 0.00359 | 4.6e-05 | 0.001171 | 0.32493 | 1.3087 | 2.5473 | 100.63 | 21.16 | 25.513 | 96.6252 | 87.7894 | 89.8135 | 0.006375 | 0.97267 | 0.023299 | 23.9538 | 1.0655 | 0.1542 | 0.02089 | 1.4596 | 1.4596 | 0.004363
2019-03-28 15:40 | 1553802013.2729 | 112 | 0.20526 | 0.003647 | 5.6e-05 | 98.7425 | 87.7033 | 115.6121 | 21.2612 | 0.003069 | 0.00284 | 170.8271 | 0.26021 | 8.4022 | 0.002714 | 0.004767 | 0.004671 | 4.4e-05 | 0.001174 | 0.31205 | 1.2818 | 2.5474 | 100.63 | 21.16 | 25.512 | 95.6673 | 87.1172 | 90.1386 | 0.006375 | 0.9718 | 0.023298 | 23.9744 | 1.0604 | 0.1548 | 0.0214 | 1.4951 | 1.4951 | 0.004276
2019-03-28 15:40 | 1553802014.2729 | 113 | 0.20525 | 0.003661 | 5.6e-05 | 98.797 | 87.5096 | 115.2508 | 21.2612 | 0.003114 | 0.002795 | 170.7723 | 0.25933 | 8.4014 | 0.002718 | 0.004131 | 0.004039 | 4.9e-05 | 0.001177 | 0.34562 | 1.2808 | 2.5474 | 100.63 | 21.16 | 25.512 | 96.7414 | 87.2046 | 89.061 | 0.006375 | 0.97469 | 0.023297 | 23.9013 | 1.0851 | 0.15427 | 0.021428 | 1.4969 | 1.4969 | 0.004256
2019-03-28 15:40 | 1553802015.273 | 114 | 0.20531 | 0.003678 | 5.5e-05 | 98.6299 | 87.0914 | 115.3473 | 21.2612 | 0.003113 | 0.002865 | 170.7167 | 0.25957 | 8.4012 | 0.002722 | 0.004286 | 0.004248 | 4.6e-05 | 0.001176 | 0.32261 | 1.3095 | 2.5474 | 100.63 | 21.16 | 25.51 | 96.732 | 87.3117 | 90.7551 | 0.006374 | 0.97016 | 0.023252 | 23.9672 | 1.1128 | 0.15442 | 0.021056 | 1.4682 | 1.4682 | 0.004344
2019-03-28 15:40 | 1553802016.274 | 115 | 0.20527 | 0.003689 | 5.4e-05 | 98.5963 | 87.4159 | 115.3461 | 21.2612 | 0.003116 | 0.002772 | 170.6358 | 0.25874 | 8.4036 | 0.002721 | 0.004651 | 0.004591 | 4.5e-05 | 0.001177 | 0.32781 | 1.3628 | 2.5474 | 100.63 | 21.16 | 25.506 | 96.7984 | 87.2703 | 88.5034 | 0.006373 | 0.9762 | 0.023324 | 23.8928 | 1.0765 | 0.15392 | 0.021291 | 1.4893 | 1.4893 | 0.004269
2019-03-28 15:40 | 1553802017.2744 | 116 | 0.20514 | 0.003702 | 5.3e-05 | 98.582 | 86.8037 | 115.1809 | 21.2612 | 0.003093 | 0.002852 | 170.5575 | 0.25878 | 8.4017 | 0.002718 | 0.004144 | 0.004066 | 4.6e-05 | 0.001177 | 0.29336 | 1.3442 | 2.5475 | 100.63 | 21.16 | 25.503 | 96.2532 | 87.1969 | 90.4347 | 0.006373 | 0.97101 | 0.023245 | 23.9394 | 1.1381 | 0.15395 | 0.022087 | 1.5398 | 1.5398 | 0.004129
2019-03-28 15:40 | 1553802018.2741 | 117 | 0.20515 | 0.003713 | 5.3e-05 | 98.5569 | 86.9817 | 114.8882 | 21.2612 | 0.003064 | 0.002834 | 170.4708 | 0.25788 | 8.402 | 0.002792 | 0.003853 | 0.003789 | 5e-05 | 0.001177 | 0.31852 | 1.3034 | 2.5474 | 100.63 | 21.16 | 25.5 | 95.5515 | 88.9899 | 89.9908 | 0.006372 | 0.9722 | 0.02323 | 23.8943 | 1.1175 | 0.15341 | 0.021974 | 1.5309 | 1.5309 | 0.004139
2019-03-28 15:40 | 1553802019.2731 | 118 | 0.20511 | 0.003727 | 5.4e-05 | 98.4789 | 87.0796 | 114.5927 | 21.2612 | 0.003146 | 0.00289 | 170.4116 | 0.25827 | 8.4055 | 0.00276 | 0.004439 | 0.004341 | 5.2e-05 | 0.001171 | 0.36049 | 1.3055 | 2.5475 | 100.63 | 21.16 | 25.494 | 97.521 | 88.2198 | 91.3586 | 0.00637 | 0.96855 | 0.023157 | 23.9085 | 1.1004 | 0.15365 | 0.022226 | 1.5435 | 1.5435 | 0.004111
2019-03-28 15:40 | 1553802020.2735 | 119 | 0.20502 | 0.003737 | 5.5e-05 | 98.4735 | 87.2518 | 114.4678 | 21.2612 | 0.003134 | 0.002869 | 170.3382 | 0.25803 | 8.404 | 0.002733 | 0.003967 | 0.003815 | 4.9e-05 | 0.001163 | 0.32477 | 1.3077 | 2.5474 | 100.63 | 21.16 | 25.493 | 97.2294 | 87.5733 | 90.8522 | 0.00637 | 0.9699 | 0.02316 | 23.8788 | 1.0822 | 0.1535 | 0.022746 | 1.5798 | 1.5798 | 0.004013
2019-03-28 15:40 | 1553802021.2741 | 120 | 0.20497 | 0.003757 | 5.6e-05 | 98.3618 | 86.6631 | 114.6636 | 21.2612 | 0.003137 | 0.002911 | 170.2743 | 0.25857 | 8.4047 | 0.002719 | 0.004297 | 0.004282 | 5.4e-05 | 0.001156 | 0.28943 | 1.3157 | 2.5475 | 100.63 | 21.16 | 25.489 | 97.3164 | 87.2265 | 91.8575 | 0.006369 | 0.96723 | 0.023148 | 23.9322 | 1.1326 | 0.15382 | 0.022989 | 1.5957 | 1.5957 | 0.003981
2019-03-28 15:40 | 1553802022.2731 | 121 | 0.20495 | 0.003777 | 5.7e-05 | 98.3493 | 87.0364 | 114.4044 | 21.2612 | 0.003162 | 0.002929 | 170.1791 | 0.25873 | 8.4006 | 0.002734 | 0.004817 | 0.004673 | 5e-05 | 0.00115 | 0.31758 | 1.3264 | 2.5474 | 100.63 | 21.16 | 25.485 | 97.8978 | 87.6014 | 92.2814 | 0.006368 | 0.96611 | 0.023108 | 23.919 | 1.093 | 0.15392 | 0.023128 | 1.6026 | 1.6026 | 0.003967
2019-03-28 15:40 | 1553802023.2739 | 122 | 0.20499 | 0.003806 | 5.8e-05 | 98.3318 | 87.2212 | 114.3458 | 21.2612 | 0.003159 | 0.0029 | 170.1011 | 0.25829 | 8.3994 | 0.002767 | 0.004281 | 0.004181 | 4.9e-05 | 0.001153 | 0.36193 | 1.3019 | 2.5473 | 100.63 | 21.16 | 25.484 | 97.8485 | 88.383 | 91.5767 | 0.006368 | 0.96797 | 0.023125 | 23.8898 | 1.0724 | 0.15366 | 0.022867 | 1.5856 | 1.5856 | 0.004002
2019-03-28 15:40 | 1553802024.2736 | 123 | 0.20498 | 0.003831 | 5.8e-05 | 98.3595 | 87.2076 | 114.6166 | 21.2612 | 0.00318 | 0.002888 | 170.0173 | 0.25862 | 8.4026 | 0.00279 | 0.00426 | 0.00408 | 5e-05 | 0.001144 | 0.29767 | 1.3383 | 2.5474 | 100.63 | 21.16 | 25.484 | 98.3511 | 88.9329 | 91.2984 | 0.006368 | 0.96871 | 0.023161 | 23.909 | 1.0764 | 0.15385 | 0.02284 | 1.5862 | 1.5862 | 0.004006
2019-03-28 15:40 | 1553802025.2743 | 124 | 0.20493 | 0.003858 | 5.8e-05 | 98.3103 | 87.051 | 114.3953 | 21.2612 | 0.003171 | 0.002857 | 169.9108 | 0.25845 | 8.4013 | 0.002755 | 0.004676 | 0.004632 | 5.2e-05 | 0.001143 | 0.30414 | 1.2966 | 2.5474 | 100.63 | 21.16 | 25.485 | 98.1346 | 88.1057 | 90.5607 | 0.006368 | 0.97068 | 0.023162 | 23.8617 | 1.088 | 0.15375 | 0.02313 | 1.6064 | 1.6064 | 0.003953
2019-03-28 15:40 | 1553802026.2735 | 125 | 0.20496 | 0.003887 | 5.8e-05 | 98.2236 | 87.0252 | 114.1307 | 21.2612 | 0.00318 | 0.002941 | 169.8416 | 0.25911 | 8.4034 | 0.002796 | 0.004389 | 0.004272 | 5.1e-05 | 0.001141 | 0.31078 | 1.2852 | 2.5474 | 100.63 | 21.16 | 25.487 | 98.3354 | 89.0713 | 92.5856 | 0.006369 | 0.9653 | 0.023071 | 23.9004 | 1.0827 | 0.15415 | 0.022908 | 1.5847 | 1.5847 | 0.004017
2019-03-28 15:40 | 1553802027.2736 | 126 | 0.20494 | 0.003907 | 5.9e-05 | 98.3115 | 87.1427 | 113.9439 | 21.2612 | 0.003182 | 0.002911 | 169.7603 | 0.25817 | 8.4023 | 0.002744 | 0.004039 | 0.003948 | 5e-05 | 0.001135 | 0.31404 | 1.2873 | 2.5473 | 100.63 | 21.16 | 25.491 | 98.4035 | 87.8332 | 91.8481 | 0.00637 | 0.96725 | 0.023075 | 23.8567 | 1.0787 | 0.15359 | 0.022999 | 1.5913 | 1.5913 | 0.003986
2019-03-28 15:40 | 1553802028.2737 | 127 | 0.20497 | 0.003923 | 5.8e-05 | 98.2761 | 87.2005 | 114.0049 | 21.2612 | 0.003224 | 0.002886 | 169.7039 | 0.25776 | 8.4053 | 0.002708 | 0.004177 | 0.004066 | 5.3e-05 | 0.001134 | 0.3178 | 1.282 | 2.5473 | 100.63 | 21.16 | 25.495 | 99.4175 | 86.9659 | 91.2524 | 0.006371 | 0.96883 | 0.0231 | 23.8436 | 1.0695 | 0.15334 | 0.022809 | 1.5799 | 1.5799 | 0.004009
2019-03-28 15:40 | 1553802029.2734 | 128 | 0.20494 | 0.003936 | 5.8e-05 | 98.2606 | 86.9009 | 114.4838 | 21.2612 | 0.003248 | 0.002887 | 169.6185 | 0.25822 | 8.4029 | 0.002764 | 0.004439 | 0.004282 | 5.1e-05 | 0.001137 | 0.32101 | 1.2523 | 2.5474 | 100.63 | 21.16 | 25.504 | 99.9823 | 88.3064 | 91.2806 | 0.006373 | 0.96876 | 0.023148 | 23.8946 | 1.0989 | 0.15362 | 0.022973 | 1.5945 | 1.5945 | 0.003979
2019-03-28 15:40 | 1553802030.2724 | 129 | 0.20492 | 0.003949 | 5.8e-05 | 98.3645 | 87.4196 | 114.8288 | 21.2612 | 0.003197 | 0.002948 | 169.5246 | 0.25859 | 8.4047 | 0.002765 | 0.00399 | 0.00388 | 5.2e-05 | 0.001142 | 0.32244 | 1.3175 | 2.5474 | 100.63 | 21.16 | 25.509 | 98.7539 | 88.3322 | 92.7467 | 0.006374 | 0.96488 | 0.023136 | 23.9786 | 1.0551 | 0.15383 | 0.023074 | 1.6008 | 1.6008 | 0.003969
2019-03-28 15:40 | 1553802031.2738 | 130 | 0.2049 | 0.003955 | 5.8e-05 | 98.3353 | 87.1828 | 114.4338 | 21.2612 | 0.003294 | 0.002916 | 169.4705 | 0.25866 | 8.4038 | 0.002817 | 0.004354 | 0.004283 | 5.4e-05 | 0.001142 | 0.30973 | 1.3017 | 2.5475 | 100.63 | 21.16 | 25.513 | 101.0945 | 89.5862 | 91.9787 | 0.006375 | 0.96691 | 0.023121 | 23.9122 | 1.0767 | 0.15387 | 0.023182 | 1.6071 | 1.6071 | 0.003954
2019-03-28 15:41 | 1553802032.2725 | 131 | 0.2049 | 0.003961 | 5.8e-05 | 98.3005 | 86.9838 | 114.3326 | 21.2612 | 0.003242 | 0.002941 | 169.4167 | 0.25929 | 8.4038 | 0.002851 | 0.003608 | 0.003513 | 5.3e-05 | 0.001143 | 0.31957 | 1.2985 | 2.5474 | 100.63 | 21.16 | 25.519 | 99.8466 | 90.4043 | 92.5692 | 0.006377 | 0.96535 | 0.023092 | 23.921 | 1.094 | 0.15425 | 0.02321 | 1.6071 | 1.6071 | 0.003964
2019-03-28 15:41 | 1553802033.2735 | 132 | 0.20484 | 0.003979 | 5.8e-05 | 98.3541 | 87.5256 | 113.8331 | 21.2612 | 0.003197 | 0.002978 | 169.3643 | 0.25911 | 8.4027 | 0.002834 | 0.004393 | 0.004276 | 5.2e-05 | 0.001144 | 0.32659 | 1.2935 | 2.5473 | 100.63 | 21.17 | 25.521 | 98.7643 | 89.9908 | 93.4577 | 0.006381 | 0.96301 | 0.023014 | 23.8976 | 1.0433 | 0.15415 | 0.023519 | 1.6229 | 1.6229 | 0.003923
2019-03-28 15:41 | 1553802034.2736 | 133 | 0.20483 | 0.003982 | 5.8e-05 | 98.3604 | 87.1682 | 113.597 | 21.2612 | 0.00327 | 0.002936 | 169.2735 | 0.25912 | 8.4015 | 0.002822 | 0.004387 | 0.004281 | 5.3e-05 | 0.001146 | 0.34036 | 1.2855 | 2.5474 | 100.63 | 21.17 | 25.52 | 100.5199 | 89.6993 | 92.4543 | 0.006381 | 0.96565 | 0.023021 | 23.8401 | 1.0805 | 0.15415 | 0.023548 | 1.6254 | 1.6254 | 0.003917
2019-03-28 15:41 | 1553802035.2736 | 134 | 0.20484 | 0.003998 | 5.9e-05 | 98.3237 | 87.3941 | 113.9749 | 21.2612 | 0.00327 | 0.002951 | 169.2027 | 0.25907 | 8.4027 | 0.002862 | 0.004512 | 0.004406 | 5.4e-05 | 0.00115 | 0.33268 | 1.2552 | 2.5476 | 100.63 | 21.17 | 25.514 | 100.5252 | 90.6621 | 92.8147 | 0.006379 | 0.9647 | 0.023048 | 23.8915 | 1.054 | 0.15412 | 0.023473 | 1.6221 | 1.6221 | 0.003924
2019-03-28 15:41 | 1553802036.2727 | 135 | 0.20483 | 0.004005 | 5.8e-05 | 98.3925 | 87.5477 | 114.4439 | 21.2612 | 0.003198 | 0.002988 | 169.1152 | 0.25914 | 8.4002 | 0.002841 | 0.004367 | 0.004202 | 5.3e-05 | 0.001154 | 0.31653 | 1.2775 | 2.5474 | 100.63 | 21.17 | 25.508 | 98.787 | 90.1666 | 93.7145 | 0.006378 | 0.96233 | 0.023067 | 23.97 | 1.0446 | 0.15416 | 0.023548 | 1.6287 | 1.6287 | 0.003909
2019-03-28 15:41 | 1553802037.2727 | 136 | 0.20485 | 0.004021 | 5.8e-05 | 98.3639 | 87.8484 | 114.3941 | 21.2612 | 0.003291 | 0.002999 | 169.0419 | 0.25869 | 8.401 | 0.002836 | 0.004455 | 0.004435 | 5.3e-05 | 0.001155 | 0.32421 | 1.2937 | 2.5473 | 100.63 | 21.17 | 25.494 | 101.0221 | 90.037 | 93.968 | 0.006374 | 0.96167 | 0.023054 | 23.9731 | 1.0113 | 0.1539 | 0.023433 | 1.6199 | 1.6199 | 0.003924
2019-03-28 15:41 | 1553802038.2728 | 137 | 0.20484 | 0.004029 | 5.8e-05 | 98.4044 | 87.8256 | 114.1519 | 21.2612 | 0.003247 | 0.002971 | 168.9525 | 0.25949 | 8.4035 | 0.002825 | 0.004291 | 0.004267 | 5e-05 | 0.001157 | 0.32892 | 1.3395 | 2.5474 | 100.63 | 21.17 | 25.486 | 99.9595 | 89.7885 | 93.2966 | 0.006372 | 0.96343 | 0.023051 | 23.9258 | 1.0173 | 0.15437 | 0.023443 | 1.6203 | 1.6203 | 0.003935
2019-03-28 15:41 | 1553802039.2727 | 138 | 0.20484 | 0.004038 | 5.8e-05 | 98.2981 | 87.3055 | 114.208 | 21.2612 | 0.003283 | 0.003007 | 168.8674 | 0.25837 | 8.4053 | 0.00286 | 0.004498 | 0.004373 | 5.2e-05 | 0.001164 | 0.33588 | 1.3597 | 2.5474 | 100.63 | 21.17 | 25.476 | 100.8385 | 90.6255 | 94.1668 | 0.00637 | 0.96115 | 0.023029 | 23.9601 | 1.0607 | 0.1537 | 0.023471 | 1.6207 | 1.6207 | 0.003917
2019-03-28 15:41 | 1553802040.2732 | 139 | 0.20487 | 0.004049 | 5.8e-05 | 98.3475 | 87.8609 | 114.1477 | 21.2612 | 0.003255 | 0.002971 | 168.7932 | 0.2602 | 8.4008 | 0.002809 | 0.004285 | 0.004189 | 5e-05 | 0.001162 | 0.31581 | 1.3435 | 2.5473 | 100.63 | 21.17 | 25.455 | 100.162 | 89.3962 | 93.2959 | 0.006364 | 0.96343 | 0.02305 | 23.9253 | 1.0085 | 0.15479 | 0.023254 | 1.6073 | 1.6073 | 0.003978
2019-03-28 15:41 | 1553802041.2728 | 140 | 0.20491 | 0.00405 | 5.7e-05 | 98.3353 | 88.1725 | 114.3467 | 21.2612 | 0.003288 | 0.002976 | 168.6865 | 0.25884 | 8.4031 | 0.002852 | 0.003936 | 0.003799 | 5e-05 | 0.001165 | 0.3361 | 1.3208 | 2.5474 | 100.63 | 21.17 | 25.444 | 100.9452 | 90.4371 | 93.41 | 0.006362 | 0.96313 | 0.023067 | 23.9499 | 0.97575 | 0.15398 | 0.02304 | 1.5936 | 1.5936 | 0.003991
2019-03-28 15:41 | 1553802042.2738 | 141 | 0.20489 | 0.004056 | 5.7e-05 | 98.3803 | 87.7543 | 114.3155 | 21.2612 | 0.003305 | 0.003018 | 168.6275 | 0.25793 | 8.4059 | 0.002916 | 0.004685 | 0.004616 | 5e-05 | 0.001166 | 0.28313 | 1.3205 | 2.5473 | 100.63 | 21.17 | 25.422 | 101.3708 | 91.9717 | 94.4304 | 0.006356 | 0.96046 | 0.023032 | 23.9799 | 1.0224 | 0.15344 | 0.023102 | 1.5955 | 1.5955 | 0.003972
2019-03-28 15:41 | 1553802043.2722 | 142 | 0.20493 | 0.004059 | 5.7e-05 | 98.4438 | 88.2057 | 114.6192 | 21.2612 | 0.003279 | 0.003002 | 168.5458 | 0.2583 | 8.4037 | 0.002825 | 0.004639 | 0.00455 | 5e-05 | 0.001167 | 0.30901 | 1.3595 | 2.5474 | 100.63 | 21.17 | 25.414 | 100.7364 | 89.7877 | 94.0517 | 0.006354 | 0.96145 | 0.023074 | 23.9994 | 0.98223 | 0.15366 | 0.022913 | 1.5854 | 1.5854 | 0.004003
2019-03-28 15:41 | 1553802044.2721 | 143 | 0.20492 | 0.004054 | 5.6e-05 | 98.3586 | 87.7342 | 114.7417 | 21.2612 | 0.003279 | 0.002992 | 168.4753 | 0.25806 | 8.4036 | 0.002918 | 0.003957 | 0.00395 | 4.7e-05 | 0.00117 | 0.31996 | 1.3165 | 2.5474 | 100.63 | 21.17 | 25.408 | 100.7364 | 92.0163 | 93.817 | 0.006353 | 0.96206 | 0.023094 | 24.0046 | 1.0224 | 0.15352 | 0.022964 | 1.5903 | 1.5903 | 0.003987
2019-03-28 15:41 | 1553802045.2729 | 144 | 0.2049 | 0.004054 | 5.6e-05 | 98.4664 | 87.7057 | 114.435 | 21.2612 | 0.003273 | 0.003026 | 168.4039 | 0.25925 | 8.4045 | 0.002806 | 0.004028 | 0.00391 | 5e-05 | 0.00117 | 0.3063 | 1.342 | 2.5475 | 100.63 | 21.17 | 25.401 | 100.5867 | 89.3268 | 94.6341 | 0.006351 | 0.95993 | 0.023037 | 23.9991 | 1.0351 | 0.15423 | 0.023074 | 1.594 | 1.594 | 0.003996
2019-03-28 15:41 | 1553802046.2729 | 145 | 0.20492 | 0.00405 | 5.7e-05 | 98.3312 | 88.0063 | 114.2828 | 21.2612 | 0.003326 | 0.003025 | 168.3542 | 0.25937 | 8.4032 | 0.002761 | 0.004838 | 0.004743 | 4.9e-05 | 0.001169 | 0.32139 | 1.3274 | 2.7261 | 100.63 | 21.17 | 25.402 | 101.8638 | 88.2285 | 94.602 | 0.006351 | 0.96001 | 0.023023 | 23.9821 | 0.99224 | 0.1543 | 0.02294 | 1.5838 | 1.5838 | 0.004024
2019-03-28 15:41 | 1553802047.2723 | 146 | 0.20489 | 0.004047 | 5.7e-05 | 98.4667 | 88.2265 | 114.4896 | 21.2612 | 0.003275 | 0.003041 | 168.2904 | 0.25855 | 8.4031 | 0.00287 | 0.00465 | 0.004565 | 5e-05 | 0.001165 | 0.32659 | 1.326 | 2.5943 | 100.63 | 21.16 | 25.406 | 100.6336 | 90.8616 | 94.9839 | 0.006348 | 0.95901 | 0.023032 | 24.0163 | 0.98221 | 0.15381 | 0.023119 | 1.5967 | 1.5967 | 0.003978
2019-03-28 15:41 | 1553802048.2731 | 147 | 0.20491 | 0.004042 | 5.7e-05 | 98.3764 | 88.5489 | 114.8274 | 21.2612 | 0.003373 | 0.002983 | 168.2267 | 0.2591 | 8.4049 | 0.002923 | 0.004459 | 0.004394 | 4.7e-05 | 0.001165 | 0.35778 | 1.3227 | 2.5475 | 100.63 | 21.16 | 25.411 | 103.0068 | 92.1397 | 93.6002 | 0.00635 | 0.96263 | 0.023109 | 24.0064 | 0.94138 | 0.15414 | 0.023008 | 1.5944 | 1.5944 | 0.003993
2019-03-28 15:41 | 1553802049.2731 | 148 | 0.20487 | 0.004043 | 5.7e-05 | 98.4566 | 88.7802 | 114.7467 | 21.2612 | 0.003347 | 0.003005 | 168.1428 | 0.25814 | 8.404 | 0.002858 | 0.004572 | 0.004519 | 4.9e-05 | 0.001163 | 0.33019 | 1.2856 | 2.5475 | 100.63 | 21.16 | 25.416 | 102.3862 | 90.5754 | 94.1135 | 0.006351 | 0.96129 | 0.023085 | 24.0148 | 0.92533 | 0.15356 | 0.023249 | 1.6094 | 1.6094 | 0.003941
2019-03-28 15:41 | 1553802050.2735 | 149 | 0.20489 | 0.004044 | 5.7e-05 | 98.4503 | 88.7113 | 114.3918 | 21.2612 | 0.003276 | 0.003002 | 168.0689 | 0.25882 | 8.4006 | 0.002806 | 0.004732 | 0.004655 | 4.8e-05 | 0.001162 | 0.31293 | 1.3067 | 2.5474 | 100.63 | 21.16 | 25.427 | 100.6602 | 89.3331 | 94.0493 | 0.006354 | 0.96145 | 0.023051 | 23.9755 | 0.93171 | 0.15397 | 0.023137 | 1.5993 | 1.5993 | 0.003976
2019-03-28 15:41 | 1553802051.2731 | 150 | 0.2049 | 0.004048 | 5.7e-05 | 98.4104 | 88.6912 | 114.4612 | 21.2612 | 0.003321 | 0.003027 | 168.0049 | 0.25921 | 8.4063 | 0.002889 | 0.004849 | 0.004759 | 4.9e-05 | 0.001159 | 0.30287 | 1.2689 | 2.5474 | 100.63 | 21.16 | 25.434 | 101.7609 | 91.3242 | 94.6543 | 0.006355 | 0.95987 | 0.023039 | 24.0025 | 0.9301 | 0.15421 | 0.023062 | 1.5933 | 1.5933 | 0.003997
2019-03-28 15:41 | 1553802052.2732 | 151 | 0.20492 | 0.004055 | 5.7e-05 | 98.4607 | 88.7364 | 114.6727 | 21.2612 | 0.003343 | 0.003016 | 167.9182 | 0.25887 | 8.4042 | 0.002885 | 0.004001 | 0.003926 | 4.7e-05 | 0.001156 | 0.30596 | 1.2846 | 2.5472 | 100.63 | 21.16 | 25.44 | 102.292 | 91.2155 | 94.3952 | 0.006357 | 0.96055 | 0.023069 | 24.0162 | 0.93012 | 0.154 | 0.022922 | 1.5856 | 1.5856 | 0.004011
2019-03-28 15:41 | 1553802053.2722 | 152 | 0.20496 | 0.004052 | 5.7e-05 | 98.4247 | 88.8832 | 114.6941 | 21.2612 | 0.003336 | 0.003003 | 167.8271 | 0.25909 | 8.4007 | 0.002847 | 0.004009 | 0.003873 | 4.7e-05 | 0.001153 | 0.31797 | 1.2816 | 2.5472 | 100.63 | 21.16 | 25.453 | 102.1191 | 90.3166 | 94.0744 | 0.00636 | 0.96139 | 0.023081 | 24.008 | 0.91207 | 0.15413 | 0.022741 | 1.5739 | 1.5739 | 0.004044
2019-03-28 15:41 | 1553802054.2896 | 153 | 0.20496 | 0.004056 | 5.6e-05 | 98.5093 | 88.8553 | 114.7894 | 21.2612 | 0.003292 | 0.002977 | 167.7604 | 0.25929 | 8.4009 | 0.002801 | 0.004436 | 0.004295 | 5e-05 | 0.001157 | 0.29734 | 1.3046 | 2.5474 | 100.63 | 21.16 | 25.46 | 101.0536 | 89.2033 | 93.4493 | 0.006362 | 0.96303 | 0.02311 | 23.9975 | 0.92255 | 0.15425 | 0.02272 | 1.5745 | 1.5745 | 0.004046
2019-03-28 15:41 | 1553802055.2866 | 154 | 0.20493 | 0.004051 | 5.6e-05 | 98.5185 | 89.1067 | 114.9692 | 21.2612 | 0.003311 | 0.003028 | 167.6769 | 0.25853 | 8.4038 | 0.002876 | 0.004144 | 0.003993 | 4.5e-05 | 0.00116 | 0.31974 | 1.2885 | 2.5474 | 100.63 | 21.16 | 25.471 | 101.5136 | 91.0061 | 94.6763 | 0.006365 | 0.95982 | 0.02309 | 24.0565 | 0.89812 | 0.1538 | 0.022874 | 1.5838 | 1.5838 | 0.004011
2019-03-28 15:41 | 1553802056.2863 | 155 | 0.20493 | 0.004044 | 5.5e-05 | 98.5117 | 89.0788 | 114.7302 | 21.2612 | 0.003301 | 0.003022 | 167.5613 | 0.25851 | 8.4023 | 0.002791 | 0.004149 | 0.004105 | 4.9e-05 | 0.001154 | 0.32908 | 1.3189 | 2.5474 | 100.63 | 21.16 | 25.475 | 101.2774 | 88.9712 | 94.5245 | 0.006366 | 0.96021 | 0.023071 | 24.0265 | 0.9003 | 0.15379 | 0.022915 | 1.5854 | 1.5854 | 0.004006
2019-03-28 15:41 | 1553802057.2863 | 156 | 0.20499 | 0.004038 | 5.5e-05 | 98.5301 | 88.7996 | 114.4371 | 21.2612 | 0.003359 | 0.003039 | 167.4763 | 0.25949 | 8.4032 | 0.002809 | 0.004355 | 0.004242 | 5.2e-05 | 0.001154 | 0.32117 | 1.2989 | 2.5473 | 100.63 | 21.16 | 25.48 | 102.6854 | 89.4034 | 94.9393 | 0.006367 | 0.95913 | 0.023028 | 24.0093 | 0.93006 | 0.15437 | 0.022569 | 1.5586 | 1.5586 | 0.004091
2019-03-28 15:41 | 1553802058.2853 | 157 | 0.20497 | 0.004034 | 5.5e-05 | 98.6171 | 89.1665 | 114.5108 | 21.2612 | 0.003325 | 0.003033 | 167.3977 | 0.25855 | 8.4036 | 0.00291 | 0.004347 | 0.00424 | 4.9e-05 | 0.001157 | 0.30702 | 1.2866 | 2.7047 | 100.63 | 21.16 | 25.492 | 101.8459 | 91.8348 | 94.7906 | 0.00637 | 0.95952 | 0.02304 | 24.0122 | 0.90107 | 0.15381 | 0.022679 | 1.567 | 1.567 | 0.004054
2019-03-28 15:41 | 1553802059.2843 | 158 | 0.205 | 0.004031 | 5.4e-05 | 98.554 | 89.0333 | 114.1749 | 21.2612 | 0.003301 | 0.003025 | 167.3263 | 0.25873 | 8.4051 | 0.002783 | 0.004311 | 0.004189 | 5.1e-05 | 0.001158 | 0.31686 | 1.2895 | 2.5474 | 100.63 | 21.16 | 25.498 | 101.2805 | 88.7735 | 94.602 | 0.006371 | 0.96001 | 0.023012 | 23.9708 | 0.90871 | 0.15392 | 0.022486 | 1.5518 | 1.5518 | 0.004096
2019-03-28 15:41 | 1553802060.2847 | 159 | 0.205 | 0.004023 | 5.4e-05 | 98.601 | 89.3289 | 114.3084 | 21.2612 | 0.003338 | 0.003089 | 167.2722 | 0.25912 | 8.4038 | 0.002834 | 0.004144 | 0.00407 | 5.3e-05 | 0.001161 | 0.32189 | 1.3256 | 2.5474 | 100.63 | 21.16 | 25.505 | 102.1608 | 89.9949 | 96.1506 | 0.006373 | 0.95598 | 0.022977 | 24.0353 | 0.88333 | 0.15415 | 0.02254 | 1.5532 | 1.5532 | 0.004099
2019-03-28 15:41 | 1553802061.2859 | 160 | 0.20501 | 0.004016 | 5.4e-05 | 98.5948 | 89.4743 | 114.1791 | 21.2612 | 0.003367 | 0.003027 | 167.1837 | 0.2592 | 8.4057 | 0.00285 | 0.004088 | 0.003949 | 5.2e-05 | 0.00116 | 0.31996 | 1.3114 | 2.5474 | 100.63 | 21.16 | 25.512 | 102.8629 | 90.3925 | 94.6442 | 0.006375 | 0.9599 | 0.023011 | 23.9726 | 0.86822 | 0.15419 | 0.022487 | 1.5518 | 1.5518 | 0.004104
2019-03-28 15:41 | 1553802062.2844 | 161 | 0.20501 | 0.004008 | 5.5e-05 | 98.5897 | 89.763 | 113.9899 | 21.2612 | 0.003413 | 0.002989 | 167.1088 | 0.25873 | 8.4021 | 0.002885 | 0.004643 | 0.00458 | 5.2e-05 | 0.001159 | 0.32565 | 1.2832 | 2.5475 | 100.63 | 21.16 | 25.527 | 103.9932 | 91.2155 | 93.7268 | 0.006379 | 0.9623 | 0.023021 | 23.9228 | 0.83894 | 0.15392 | 0.022458 | 1.5504 | 1.5504 | 0.0041
2019-03-28 15:41 | 1553802063.2845 | 162 | 0.20501 | 0.004 | 5.6e-05 | 98.596 | 89.5606 | 113.7944 | 21.2612 | 0.003411 | 0.003033 | 167.02 | 0.25896 | 8.4007 | 0.002844 | 0.00436 | 0.00425 | 4.9e-05 | 0.001154 | 0.3147 | 1.2858 | 2.5473 | 100.63 | 21.16 | 25.535 | 103.9435 | 90.2463 | 94.7858 | 0.006381 | 0.95953 | 0.022968 | 23.9368 | 0.85971 | 0.15406 | 0.022464 | 1.5472 | 1.5472 | 0.004112
2019-03-28 15:41 | 1553802064.2848 | 163 | 0.205 | 0.003997 | 5.7e-05 | 98.624 | 89.5426 | 113.7453 | 21.2612 | 0.003363 | 0.003045 | 166.9467 | 0.2589 | 8.4036 | 0.002858 | 0.00368 | 0.003642 | 4.8e-05 | 0.001157 | 0.31708 | 1.3415 | 2.5475 | 100.63 | 21.16 | 25.553 | 102.775 | 90.5677 | 95.0794 | 0.006385 | 0.95876 | 0.022954 | 23.9412 | 0.86404 | 0.15402 | 0.022528 | 1.5506 | 1.5506 | 0.004102
2019-03-28 15:41 | 1553802065.2855 | 164 | 0.20501 | 0.003991 | 5.7e-05 | 98.6445 | 89.5256 | 113.9021 | 21.2612 | 0.003394 | 0.003031 | 166.8889 | 0.25895 | 8.4041 | 0.002893 | 0.004058 | 0.003961 | 5e-05 | 0.001157 | 0.30646 | 1.3096 | 2.5475 | 100.63 | 21.16 | 25.562 | 103.519 | 91.4189 | 94.7397 | 0.006387 | 0.95965 | 0.02298 | 23.9466 | 0.8676 | 0.15405 | 0.022476 | 1.5488 | 1.5488 | 0.004108
2019-03-28 15:41 | 1553802066.2845 | 165 | 0.20502 | 0.003983 | 5.8e-05 | 98.6076 | 89.2191 | 113.8995 | 21.2612 | 0.00337 | 0.003042 | 166.8303 | 0.25843 | 8.4046 | 0.002798 | 0.003948 | 0.003937 | 5.1e-05 | 0.001158 | 0.30397 | 1.3393 | 2.5476 | 100.63 | 21.16 | 25.571 | 102.9322 | 89.125 | 95.0215 | 0.00639 | 0.95891 | 0.022971 | 23.9555 | 0.89492 | 0.15374 | 0.022417 | 1.544 | 1.544 | 0.004112
2019-03-28 15:41 | 1553802067.2856 | 166 | 0.20502 | 0.003984 | 5.9e-05 | 98.6466 | 89.157 | 114.0434 | 21.2612 | 0.003325 | 0.00304 | 166.7542 | 0.25919 | 8.4021 | 0.002789 | 0.00439 | 0.00424 | 5.2e-05 | 0.001155 | 0.3183 | 1.3433 | 2.5474 | 100.63 | 21.16 | 25.587 | 101.856 | 88.9081 | 94.9644 | 0.006394 | 0.95906 | 0.022988 | 23.9688 | 0.9047 | 0.15419 | 0.022418 | 1.5451 | 1.5451 | 0.004121
2019-03-28 15:41 | 1553802068.2846 | 167 | 0.20505 | 0.003982 | 5.9e-05 | 98.674 | 89.3547 | 114.0369 | 21.2612 | 0.003353 | 0.003023 | 166.6534 | 0.25858 | 8.4003 | 0.002863 | 0.004023 | 0.004072 | 5e-05 | 0.00115 | 0.34075 | 1.3175 | 2.5473 | 100.63 | 21.16 | 25.595 | 102.5314 | 90.7028 | 94.5651 | 0.006396 | 0.96011 | 0.022999 | 23.9551 | 0.88737 | 0.15383 | 0.022274 | 1.536 | 1.536 | 0.004136
2019-03-28 15:41 | 1553802069.2846 | 168 | 0.20509 | 0.003978 | 5.9e-05 | 98.7127 | 89.1325 | 114.1906 | 21.2612 | 0.00339 | 0.003075 | 166.567 | 0.25889 | 8.4044 | 0.002914 | 0.004226 | 0.004063 | 4.9e-05 | 0.001148 | 0.30807 | 1.3335 | 2.5473 | 100.63 | 21.15 | 25.606 | 103.435 | 91.9286 | 95.8138 | 0.006394 | 0.95686 | 0.022976 | 24.0119 | 0.91314 | 0.15401 | 0.022031 | 1.5177 | 1.5177 | 0.004191
2019-03-28 15:41 | 1553802070.2852 | 169 | 0.20509 | 0.003975 | 6e-05 | 98.6999 | 89.1746 | 114.0277 | 21.2612 | 0.003299 | 0.00304 | 166.4616 | 0.25948 | 8.403 | 0.002849 | 0.004589 | 0.004481 | 4.8e-05 | 0.001147 | 0.33019 | 1.299 | 2.5474 | 100.63 | 21.15 | 25.61 | 101.2183 | 90.366 | 94.9581 | 0.006395 | 0.95908 | 0.022986 | 23.9669 | 0.90776 | 0.15437 | 0.022039 | 1.5189 | 1.5189 | 0.004197
2019-03-28 15:41 | 1553802071.2847 | 170 | 0.2051 | 0.003964 | 5.9e-05 | 98.6814 | 89.3156 | 114.2383 | 21.2612 | 0.003336 | 0.003085 | 166.4021 | 0.25917 | 8.4045 | 0.002882 | 0.00441 | 0.004389 | 4.8e-05 | 0.001145 | 0.30901 | 1.2942 | 2.5474 | 100.63 | 21.15 | 25.614 | 102.1199 | 91.1594 | 96.0567 | 0.006396 | 0.95623 | 0.022973 | 24.0248 | 0.89195 | 0.15418 | 0.021979 | 1.5139 | 1.5139 | 0.004206
2019-03-28 15:41 | 1553802072.2857 | 171 | 0.20513 | 0.003965 | 6e-05 | 98.696 | 89.4634 | 113.9979 | 21.2612 | 0.003338 | 0.003118 | 166.3126 | 0.25864 | 8.4006 | 0.002865 | 0.00488 | 0.004697 | 4.6e-05 | 0.001145 | 0.33826 | 1.2965 | 2.5474 | 100.63 | 21.15 | 25.619 | 102.1567 | 90.738 | 96.8346 | 0.006398 | 0.95422 | 0.022925 | 24.0248 | 0.87849 | 0.15387 | 0.021822 | 1.4999 | 1.4999 | 0.004237
2019-03-28 15:41 | 1553802073.2847 | 172 | 0.20515 | 0.003957 | 5.9e-05 | 98.719 | 89.1492 | 114.2542 | 21.2612 | 0.003343 | 0.003071 | 166.2358 | 0.25815 | 8.4035 | 0.002902 | 0.004299 | 0.004211 | 5.1e-05 | 0.001145 | 0.32549 | 1.319 | 2.5474 | 100.63 | 21.15 | 25.622 | 102.2942 | 91.6401 | 95.712 | 0.006398 | 0.95712 | 0.022986 | 24.0153 | 0.91204 | 0.15357 | 0.021712 | 1.4963 | 1.4963 | 0.004239
2019-03-28 15:41 | 1553802074.2838 | 173 | 0.20515 | 0.003947 | 5.9e-05 | 98.691 | 89.0157 | 114.3263 | 21.2612 | 0.003361 | 0.003049 | 166.1687 | 0.25869 | 8.4018 | 0.002867 | 0.004382 | 0.004211 | 4.8e-05 | 0.001144 | 0.31293 | 1.2935 | 2.5473 | 100.63 | 21.15 | 25.624 | 102.7113 | 90.7857 | 95.1772 | 0.006399 | 0.95851 | 0.023009 | 24.0054 | 0.92291 | 0.15389 | 0.021722 | 1.4986 | 1.4986 | 0.004241
2019-03-28 15:41 | 1553802075.2838 | 174 | 0.20519 | 0.003943 | 5.9e-05 | 98.7315 | 89.3951 | 114.13 | 21.2612 | 0.003379 | 0.003068 | 166.0828 | 0.25825 | 8.4049 | 0.002845 | 0.004432 | 0.004375 | 4.7e-05 | 0.001146 | 0.32068 | 1.2975 | 2.5474 | 100.63 | 21.15 | 25.624 | 103.1584 | 90.2706 | 95.6432 | 0.006399 | 0.9573 | 0.022975 | 24 | 0.88853 | 0.15363 | 0.021452 | 1.4778 | 1.4778 | 0.004294
2019-03-28 15:41 | 1553802076.2838 | 175 | 0.20522 | 0.003925 | 5.9e-05 | 98.7205 | 89.2949 | 114.4723 | 21.2612 | 0.003357 | 0.003044 | 166.0164 | 0.25898 | 8.4024 | 0.002817 | 0.003902 | 0.003894 | 4.9e-05 | 0.001146 | 0.31116 | 1.2692 | 2.5475 | 100.63 | 21.15 | 25.623 | 102.6249 | 89.5987 | 95.0575 | 0.006399 | 0.95882 | 0.023028 | 24.0168 | 0.89756 | 0.15407 | 0.021282 | 1.4694 | 1.4694 | 0.00433
2019-03-28 15:41 | 1553802077.2849 | 176 | 0.20522 | 0.003912 | 5.9e-05 | 98.7508 | 89.2042 | 114.6223 | 21.2612 | 0.00338 | 0.003065 | 165.9372 | 0.25874 | 8.4041 | 0.002789 | 0.004734 | 0.00458 | 5.2e-05 | 0.001149 | 0.30879 | 1.3104 | 2.5475 | 100.63 | 21.15 | 25.618 | 103.1758 | 88.9025 | 95.5679 | 0.006397 | 0.95749 | 0.023027 | 24.0492 | 0.90941 | 0.15392 | 0.02136 | 1.4747 | 1.4747 | 0.004311
2019-03-28 15:41 | 1553802078.2849 | 177 | 0.20524 | 0.0039 | 5.8e-05 | 98.7294 | 89.1665 | 114.5063 | 21.2612 | 0.003358 | 0.003053 | 165.8721 | 0.25849 | 8.404 | 0.002842 | 0.004205 | 0.00409 | 4.8e-05 | 0.00115 | 0.32173 | 1.2832 | 2.5473 | 100.63 | 21.15 | 25.613 | 102.6476 | 90.194 | 95.2884 | 0.006396 | 0.95822 | 0.023024 | 24.0279 | 0.91125 | 0.15378 | 0.021256 | 1.4675 | 1.4675 | 0.004328
2019-03-28 15:41 | 1553802079.2839 | 178 | 0.2053 | 0.003887 | 5.8e-05 | 98.6457 | 88.8206 | 114.6622 | 21.2612 | 0.003368 | 0.003068 | 165.7884 | 0.25847 | 8.4026 | 0.002797 | 0.004469 | 0.004363 | 4.7e-05 | 0.001147 | 0.32748 | 1.3111 | 2.5475 | 100.63 | 21.15 | 25.607 | 102.8881 | 89.1149 | 95.6275 | 0.006395 | 0.95734 | 0.023029 | 24.0554 | 0.93842 | 0.15376 | 0.020909 | 1.4438 | 1.4438 | 0.004398
2019-03-28 15:41 | 1553802080.2839 | 179 | 0.20534 | 0.003872 | 5.7e-05 | 98.6603 | 88.8658 | 114.8053 | 21.2612 | 0.003351 | 0.003059 | 165.73 | 0.25921 | 8.4057 | 0.00281 | 0.004197 | 0.004104 | 5.1e-05 | 0.001144 | 0.32764 | 1.3201 | 2.5475 | 100.63 | 21.15 | 25.592 | 102.4796 | 89.4205 | 95.4224 | 0.006391 | 0.95787 | 0.02305 | 24.0637 | 0.93519 | 0.15421 | 0.020652 | 1.4273 | 1.4273 | 0.004462
2019-03-28 15:41 | 1553802081.284 | 180 | 0.20541 | 0.003858 | 5.6e-05 | 98.7273 | 88.9878 | 114.7957 | 21.2612 | 0.003397 | 0.003032 | 165.6388 | 0.25842 | 8.4045 | 0.002746 | 0.004281 | 0.004163 | 4.9e-05 | 0.001146 | 0.33107 | 1.3104 | 2.5473 | 100.63 | 21.15 | 25.583 | 103.5978 | 87.8676 | 94.7631 | 0.006389 | 0.95959 | 0.02307 | 24.0411 | 0.929 | 0.15373 | 0.020289 | 1.4036 | 1.4036 | 0.004523
2019-03-28 15:41 | 1553802082.2844 | 181 | 0.20545 | 0.003842 | 5.4e-05 | 98.7219 | 89.0357 | 115.0307 | 21.2612 | 0.003373 | 0.003063 | 165.5657 | 0.25916 | 8.4038 | 0.002778 | 0.004375 | 0.004295 | 5e-05 | 0.001143 | 0.32764 | 1.3322 | 2.5474 | 100.63 | 21.15 | 25.564 | 103.0092 | 88.6509 | 95.5155 | 0.006384 | 0.95763 | 0.02307 | 24.0903 | 0.9237 | 0.15418 | 0.020056 | 1.3875 | 1.3875 | 0.004589
2019-03-28 15:41 | 1553802083.2841 | 182 | 0.20548 | 0.003822 | 5.3e-05 | 98.671 | 89.2599 | 114.826 | 21.2612 | 0.003375 | 0.00307 | 165.4877 | 0.25918 | 8.4016 | 0.002753 | 0.004144 | 0.004048 | 4.9e-05 | 0.001143 | 0.33284 | 1.283 | 2.5474 | 100.63 | 21.15 | 25.554 | 103.0603 | 88.0417 | 95.6799 | 0.006381 | 0.9572 | 0.023044 | 24.0742 | 0.89659 | 0.15419 | 0.019909 | 1.3759 | 1.3759 | 0.004628
2019-03-28 15:41 | 1553802084.2851 | 183 | 0.20552 | 0.003799 | 5.2e-05 | 98.6651 | 89.3163 | 114.7748 | 21.2612 | 0.003363 | 0.00299 | 165.3956 | 0.25915 | 8.4031 | 0.002795 | 0.004546 | 0.004473 | 5e-05 | 0.001145 | 0.32964 | 1.2898 | 2.5473 | 100.63 | 21.15 | 25.544 | 102.7789 | 89.0487 | 93.7589 | 0.006379 | 0.96222 | 0.023099 | 24.0061 | 0.8904 | 0.15417 | 0.019685 | 1.3638 | 1.3638 | 0.004669
2019-03-28 15:41 | 1553802085.2851 | 184 | 0.20553 | 0.003781 | 5.1e-05 | 98.6457 | 89.4451 | 114.7645 | 21.2612 | 0.003363 | 0.003039 | 165.3336 | 0.25882 | 8.4012 | 0.002711 | 0.004329 | 0.004294 | 4.7e-05 | 0.001145 | 0.30862 | 1.3033 | 2.5475 | 100.63 | 21.15 | 25.524 | 102.7774 | 87.0306 | 94.9315 | 0.006374 | 0.95915 | 0.023061 | 24.0434 | 0.87576 | 0.15397 | 0.019612 | 1.3565 | 1.3565 | 0.004688
2019-03-28 15:41 | 1553802086.2852 | 185 | 0.20554 | 0.003754 | 5.1e-05 | 98.6299 | 89.6119 | 114.6532 | 21.2612 | 0.003294 | 0.003028 | 165.2496 | 0.25823 | 8.4058 | 0.00277 | 0.004501 | 0.004323 | 5.3e-05 | 0.001146 | 0.32029 | 1.2896 | 2.5474 | 100.63 | 21.15 | 25.515 | 101.0945 | 88.4588 | 94.6731 | 0.006372 | 0.95982 | 0.023058 | 24.0233 | 0.85766 | 0.15362 | 0.01962 | 1.3569 | 1.3569 | 0.004676
2019-03-28 15:41 | 1553802087.2852 | 186 | 0.20553 | 0.003732 | 5.2e-05 | 98.6111 | 89.5881 | 114.5713 | 21.2612 | 0.003366 | 0.003074 | 165.1827 | 0.25878 | 8.4006 | 0.002725 | 0.00453 | 0.004465 | 5e-05 | 0.001146 | 0.3147 | 1.3012 | 2.5473 | 100.63 | 21.15 | 25.5 | 102.8496 | 87.3764 | 95.774 | 0.006368 | 0.95696 | 0.023015 | 24.0506 | 0.85834 | 0.15395 | 0.01967 | 1.3578 | 1.3578 | 0.004683
2019-03-28 15:41 | 1553802088.2843 | 187 | 0.20553 | 0.003709 | 5.3e-05 | 98.5635 | 89.4654 | 114.5077 | 21.2612 | 0.003366 | 0.00301 | 165.1016 | 0.25831 | 8.4035 | 0.002762 | 0.004389 | 0.004265 | 4.8e-05 | 0.001147 | 0.35258 | 1.315 | 2.5474 | 100.63 | 21.15 | 25.495 | 102.8433 | 88.2706 | 94.2505 | 0.006367 | 0.96093 | 0.023057 | 23.9942 | 0.86626 | 0.15367 | 0.019709 | 1.3628 | 1.3628 | 0.004657
2019-03-28 15:41 | 1553802089.2847 | 188 | 0.20555 | 0.00369 | 5.4e-05 | 98.5602 | 89.0574 | 114.4105 | 21.2612 | 0.003365 | 0.003063 | 165.0478 | 0.25903 | 8.4018 | 0.002757 | 0.004507 | 0.004423 | 4.9e-05 | 0.001145 | 0.30596 | 1.3304 | 2.5474 | 100.63 | 21.15 | 25.491 | 102.8191 | 88.1408 | 95.5131 | 0.006366 | 0.95764 | 0.023007 | 24.0252 | 0.90685 | 0.1541 | 0.019653 | 1.356 | 1.356 | 0.004693
2019-03-28 15:41 | 1553802090.2844 | 189 | 0.20549 | 0.003676 | 5.7e-05 | 98.5236 | 88.9678 | 114.614 | 21.2612 | 0.003332 | 0.003017 | 164.9551 | 0.25897 | 8.4043 | 0.00275 | 0.004315 | 0.004266 | 4.9e-05 | 0.001144 | 0.32405 | 1.3105 | 2.5475 | 100.63 | 21.15 | 25.491 | 102.0257 | 87.9745 | 94.4078 | 0.006366 | 0.96052 | 0.023062 | 24.0105 | 0.91254 | 0.15406 | 0.020034 | 1.3855 | 1.3855 | 0.004592
2019-03-28 15:42 | 1553802091.2833 | 190 | 0.20553 | 0.003657 | 5.7e-05 | 98.4688 | 88.3165 | 114.1568 | 21.2612 | 0.003344 | 0.003019 | 164.9024 | 0.2589 | 8.4021 | 0.002782 | 0.00439 | 0.00427 | 4.7e-05 | 0.001141 | 0.31006 | 1.2851 | 2.5474 | 100.63 | 21.15 | 25.493 | 102.3029 | 88.7487 | 94.4587 | 0.006366 | 0.96038 | 0.023015 | 23.9642 | 0.97328 | 0.15402 | 0.019775 | 1.3647 | 1.3647 | 0.004661
2019-03-28 15:42 | 1553802092.2843 | 191 | 0.20559 | 0.00365 | 5.8e-05 | 98.4238 | 88.1072 | 114.088 | 21.2612 | 0.003354 | 0.003041 | 164.8349 | 0.25865 | 8.4037 | 0.002759 | 0.004101 | 0.003996 | 5.1e-05 | 0.00114 | 0.3131 | 1.3008 | 2.5474 | 100.63 | 21.15 | 25.496 | 102.5479 | 88.1902 | 94.9933 | 0.006367 | 0.95899 | 0.022991 | 23.9744 | 0.99041 | 0.15387 | 0.019448 | 1.3408 | 1.3408 | 0.00474
2019-03-28 15:42 | 1553802093.2834 | 192 | 0.20551 | 0.003639 | 5.9e-05 | 98.456 | 88.699 | 114.1634 | 21.2612 | 0.003355 | 0.003073 | 164.7831 | 0.25905 | 8.4023 | 0.002744 | 0.004871 | 0.004811 | 5.1e-05 | 0.001136 | 0.30414 | 1.301 | 2.5474 | 100.63 | 21.15 | 25.506 | 102.5818 | 87.8238 | 95.7482 | 0.006369 | 0.95703 | 0.022975 | 24.0069 | 0.93346 | 0.15411 | 0.01991 | 1.3716 | 1.3716 | 0.00464
2019-03-28 15:42 | 1553802094.2834 | 193 | 0.20552 | 0.003631 | 5.9e-05 | 98.4679 | 88.5281 | 113.8941 | 21.2612 | 0.003395 | 0.003075 | 164.7471 | 0.258 | 8.4031 | 0.002778 | 0.004122 | 0.004003 | 5.1e-05 | 0.001141 | 0.31509 | 1.2951 | 2.5473 | 100.63 | 21.15 | 25.512 | 103.5411 | 88.6464 | 95.8114 | 0.006371 | 0.95686 | 0.022946 | 23.9806 | 0.95179 | 0.15349 | 0.019891 | 1.3685 | 1.3685 | 0.004632
2019-03-28 15:42 | 1553802095.2845 | 194 | 0.20552 | 0.003623 | 5.9e-05 | 98.4083 | 88.681 | 114.1728 | 21.2612 | 0.00334 | 0.003056 | 164.6515 | 0.25855 | 8.404 | 0.002721 | 0.004654 | 0.004568 | 4.8e-05 | 0.001143 | 0.33516 | 1.3137 | 2.5474 | 100.63 | 21.15 | 25.524 | 102.215 | 87.2671 | 95.3488 | 0.006374 | 0.95806 | 0.022989 | 23.9949 | 0.93094 | 0.15381 | 0.019918 | 1.3729 | 1.3729 | 0.004627
2019-03-28 15:42 | 1553802096.2834 | 195 | 0.20555 | 0.003625 | 6e-05 | 98.3982 | 89.1536 | 114.2786 | 21.2612 | 0.003386 | 0.003077 | 164.5728 | 0.25881 | 8.4008 | 0.002705 | 0.004565 | 0.004486 | 5e-05 | 0.001145 | 0.31293 | 1.3244 | 2.5474 | 100.63 | 21.15 | 25.531 | 103.3359 | 86.8994 | 95.864 | 0.006376 | 0.95673 | 0.022983 | 24.0228 | 0.88248 | 0.15396 | 0.019699 | 1.3575 | 1.3575 | 0.004684
2019-03-28 15:42 | 1553802097.2845 | 196 | 0.20555 | 0.003613 | 6e-05 | 98.4032 | 89.3109 | 114.4666 | 21.2612 | 0.003369 | 0.003051 | 164.5027 | 0.25772 | 8.4018 | 0.002698 | 0.00407 | 0.00392 | 5e-05 | 0.001143 | 0.32709 | 1.2836 | 2.5474 | 100.63 | 21.15 | 25.537 | 102.9179 | 86.7151 | 95.2197 | 0.006377 | 0.9584 | 0.023022 | 24.0215 | 0.86718 | 0.15332 | 0.019719 | 1.3611 | 1.3611 | 0.004652
2019-03-28 15:42 | 1553802098.2835 | 197 | 0.20557 | 0.003608 | 6e-05 | 98.319 | 88.6939 | 114.626 | 21.2612 | 0.003413 | 0.003042 | 164.4581 | 0.25898 | 8.4032 | 0.002671 | 0.005509 | 0.005431 | 5.3e-05 | 0.001143 | 0.29717 | 1.2705 | 2.5473 | 100.63 | 21.15 | 25.548 | 103.9837 | 86.0657 | 95.0167 | 0.00638 | 0.95893 | 0.023045 | 24.0316 | 0.92151 | 0.15407 | 0.019584 | 1.353 | 1.353 | 0.004703
2019-03-28 15:42 | 1553802099.2836 | 198 | 0.20558 | 0.003602 | 6e-05 | 98.2546 | 88.4459 | 114.6854 | 21.2612 | 0.003365 | 0.003089 | 164.3737 | 0.25836 | 8.4037 | 0.002677 | 0.003899 | 0.00377 | 4.9e-05 | 0.001137 | 0.32947 | 1.2627 | 2.5474 | 100.63 | 21.15 | 25.552 | 102.8276 | 86.2303 | 96.1467 | 0.006381 | 0.95599 | 0.023015 | 24.0747 | 0.94071 | 0.1537 | 0.019548 | 1.3489 | 1.3489 | 0.004706
2019-03-28 15:42 | 1553802100.2846 | 199 | 0.2056 | 0.003591 | 6.1e-05 | 98.3002 | 88.3681 | 114.5279 | 21.2612 | 0.003352 | 0.003042 | 164.3133 | 0.25921 | 8.4041 | 0.002664 | 0.004324 | 0.00428 | 5.2e-05 | 0.001138 | 0.30558 | 1.2642 | 2.5474 | 100.63 | 21.15 | 25.558 | 102.4922 | 85.9033 | 95.0191 | 0.006382 | 0.95892 | 0.023035 | 24.0214 | 0.95272 | 0.1542 | 0.019435 | 1.3422 | 1.3422 | 0.004745
2019-03-28 15:42 | 1553802101.2841 | 200 | 0.20561 | 0.003582 | 6e-05 | 98.1748 | 88.6759 | 114.9957 | 21.2612 | 0.003384 | 0.003078 | 164.2303 | 0.25808 | 8.4012 | 0.002693 | 0.004328 | 0.004171 | 4.9e-05 | 0.001143 | 0.31277 | 1.2647 | 2.5474 | 100.63 | 21.15 | 25.561 | 103.2841 | 86.5917 | 95.8759 | 0.006383 | 0.95669 | 0.023055 | 24.0984 | 0.91021 | 0.15353 | 0.019381 | 1.3397 | 1.3397 | 0.004733
2019-03-28 15:42 | 1553802102.2847 | 201 | 0.20558 | 0.003578 | 6.2e-05 | 98.1432 | 88.9548 | 115.2661 | 21.2612 | 0.00335 | 0.003069 | 164.1637 | 0.25718 | 8.4022 | 0.002627 | 0.003255 | 0.003208 | 5.1e-05 | 0.001141 | 0.32587 | 1.3228 | 2.5474 | 100.63 | 21.15 | 25.563 | 102.467 | 85.0268 | 95.6729 | 0.006384 | 0.95722 | 0.023088 | 24.1201 | 0.87924 | 0.153 | 0.019566 | 1.3543 | 1.3543 | 0.004666
2019-03-28 15:42 | 1553802103.2838 | 202 | 0.20559 | 0.003565 | 6.4e-05 | 98.0902 | 88.4578 | 115.1436 | 21.2612 | 0.003389 | 0.002983 | 164.0896 | 0.25783 | 8.4037 | 0.002714 | 0.003588 | 0.003505 | 5.1e-05 | 0.001142 | 0.32925 | 1.3232 | 2.5473 | 100.63 | 21.15 | 25.566 | 103.4028 | 87.1134 | 93.5807 | 0.006384 | 0.96268 | 0.023142 | 24.0388 | 0.92452 | 0.15338 | 0.01953 | 1.3548 | 1.3548 | 0.004676
2019-03-28 15:42 | 1553802104.2847 | 203 | 0.2056 | 0.003561 | 6.6e-05 | 98.0563 | 88.965 | 115.2816 | 21.2612 | 0.003357 | 0.003032 | 163.9762 | 0.25929 | 8.4021 | 0.002633 | 0.003448 | 0.003207 | 5.2e-05 | 0.001141 | 0.31116 | 1.2985 | 2.5473 | 100.63 | 21.15 | 25.566 | 102.6343 | 85.1603 | 94.7774 | 0.006384 | 0.95955 | 0.023118 | 24.0924 | 0.87028 | 0.15425 | 0.019499 | 1.3512 | 1.3512 | 0.004715
2019-03-28 15:42 | 1553802105.2832 | 204 | 0.20555 | 0.003552 | 6.6e-05 | 98.0613 | 88.6583 | 114.98 | 21.2612 | 0.003376 | 0.003055 | 163.9384 | 0.25848 | 8.4024 | 0.002668 | 0.00461 | 0.004492 | 5.1e-05 | 0.001136 | 0.31526 | 1.3136 | 2.5473 | 100.63 | 21.15 | 25.565 | 103.0877 | 86.0039 | 95.3316 | 0.006384 | 0.95811 | 0.02307 | 24.079 | 0.90164 | 0.15377 | 0.019805 | 1.3695 | 1.3695 | 0.004637
2019-03-28 15:42 | 1553802106.2828 | 205 | 0.20553 | 0.003555 | 6.7e-05 | 98.1048 | 88.6807 | 114.6183 | 21.2612 | 0.003335 | 0.003057 | 163.8636 | 0.25902 | 8.4029 | 0.002686 | 0.003355 | 0.003302 | 5e-05 | 0.001139 | 0.3183 | 1.3558 | 2.5474 | 100.63 | 21.15 | 25.564 | 102.0838 | 86.4449 | 95.3794 | 0.006384 | 0.95798 | 0.023032 | 24.0427 | 0.90335 | 0.15409 | 0.019922 | 1.3754 | 1.3754 | 0.004627
2019-03-28 15:42 | 1553802107.2838 | 206 | 0.20551 | 0.00355 | 6.7e-05 | 98.0554 | 88.8016 | 114.396 | 21.2612 | 0.003382 | 0.003011 | 163.8106 | 0.25825 | 8.4028 | 0.002659 | 0.003974 | 0.003873 | 4.8e-05 | 0.001136 | 0.31868 | 1.3733 | 2.5476 | 100.63 | 21.15 | 25.563 | 103.2424 | 85.7768 | 94.2637 | 0.006384 | 0.96089 | 0.023045 | 23.9829 | 0.88665 | 0.15363 | 0.020015 | 1.3825 | 1.3825 | 0.00459
2019-03-28 15:42 | 1553802108.2839 | 207 | 0.20551 | 0.003557 | 6.7e-05 | 98.1212 | 89.3136 | 114.3145 | 21.2612 | 0.003373 | 0.002986 | 163.7264 | 0.25881 | 8.4064 | 0.002714 | 0.0042 | 0.004077 | 4.9e-05 | 0.001132 | 0.30646 | 1.3633 | 2.5474 | 100.63 | 21.15 | 25.561 | 103.0043 | 87.1095 | 93.6637 | 0.006383 | 0.96247 | 0.023056 | 23.9548 | 0.84123 | 0.15397 | 0.020037 | 1.3847 | 1.3847 | 0.004592
2019-03-28 15:42 | 1553802109.2843 | 208 | 0.20553 | 0.003559 | 6.9e-05 | 98.0875 | 89.5497 | 114.2298 | 21.2612 | 0.003336 | 0.003071 | 163.6771 | 0.25871 | 8.4037 | 0.002618 | 0.004801 | 0.004708 | 4.8e-05 | 0.001135 | 0.31686 | 1.3363 | 2.5475 | 100.63 | 21.15 | 25.56 | 102.1049 | 84.7872 | 95.7081 | 0.006383 | 0.95713 | 0.022983 | 24.0126 | 0.81455 | 0.15391 | 0.019926 | 1.3725 | 1.3725 | 0.004631
2019-03-28 15:42 | 1553802110.2833 | 209 | 0.20552 | 0.003557 | 7e-05 | 98.086 | 89.4573 | 114.3796 | 21.2612 | 0.003379 | 0.003046 | 163.6201 | 0.25813 | 8.4049 | 0.002736 | 0.004173 | 0.004052 | 5e-05 | 0.001136 | 0.33892 | 1.3445 | 2.5476 | 100.63 | 21.15 | 25.556 | 103.1489 | 87.6334 | 95.1185 | 0.006382 | 0.95866 | 0.023017 | 24.0091 | 0.82365 | 0.15356 | 0.019958 | 1.3767 | 1.3767 | 0.004607
2019-03-28 15:42 | 1553802111.283 | 210 | 0.20551 | 0.003557 | 7.3e-05 | 98.1569 | 89.9883 | 114.3946 | 21.2612 | 0.003316 | 0.003077 | 163.5616 | 0.25848 | 8.3991 | 0.002719 | 0.00387 | 0.003766 | 4.9e-05 | 0.001136 | 0.3178 | 1.3045 | 2.5474 | 100.63 | 21.15 | 25.556 | 101.6433 | 87.2335 | 95.8539 | 0.006382 | 0.95675 | 0.022995 | 24.0346 | 0.77717 | 0.15377 | 0.020015 | 1.3791 | 1.3791 | 0.004605
2019-03-28 15:42 | 1553802112.283 | 211 | 0.20549 | 0.003555 | 7.6e-05 | 98.1644 | 89.5962 | 114.2986 | 21.2612 | 0.003409 | 0.003114 | 163.4937 | 0.25904 | 8.4044 | 0.002695 | 0.004512 | 0.004401 | 4.9e-05 | 0.001134 | 0.31045 | 1.2717 | 2.5474 | 100.63 | 21.15 | 25.554 | 103.8832 | 86.6557 | 96.7515 | 0.006381 | 0.95443 | 0.022958 | 24.0538 | 0.8169 | 0.1541 | 0.020138 | 1.3853 | 1.3853 | 0.004594
2019-03-28 15:42 | 1553802113.2831 | 212 | 0.20546 | 0.00356 | 7.9e-05 | 98.1986 | 89.7586 | 114.2333 | 21.2612 | 0.003394 | 0.003075 | 163.3901 | 0.25859 | 8.4016 | 0.00278 | 0.00365 | 0.003602 | 4.8e-05 | 0.001126 | 0.32532 | 1.2715 | 2.5474 | 100.63 | 21.15 | 25.551 | 103.5285 | 88.7049 | 95.8146 | 0.006381 | 0.95685 | 0.02298 | 24.0164 | 0.80383 | 0.15383 | 0.020283 | 1.3964 | 1.3964 | 0.00455
2019-03-28 15:42 | 1553802114.2831 | 213 | 0.20542 | 0.003563 | 8.1e-05 | 98.1718 | 89.4906 | 114.1786 | 21.2612 | 0.003384 | 0.003066 | 163.3166 | 0.25866 | 8.4041 | 0.00273 | 0.00451 | 0.00443 | 4.9e-05 | 0.001129 | 0.33826 | 1.255 | 2.5473 | 100.63 | 21.15 | 25.551 | 103.2676 | 87.4929 | 95.5978 | 0.006381 | 0.95742 | 0.022981 | 24.0036 | 0.82814 | 0.15387 | 0.020529 | 1.4133 | 1.4133 | 0.004496
2019-03-28 15:42 | 1553802115.2831 | 214 | 0.20543 | 0.00357 | 8.5e-05 | 98.2382 | 89.5439 | 114.0423 | 21.2612 | 0.003371 | 0.00305 | 163.2294 | 0.25847 | 8.4055 | 0.002793 | 0.004294 | 0.004218 | 4.8e-05 | 0.001121 | 0.3235 | 1.2871 | 2.5473 | 100.63 | 21.15 | 25.551 | 102.9729 | 89.0025 | 95.2071 | 0.006381 | 0.95843 | 0.02298 | 23.9766 | 0.82886 | 0.15376 | 0.020474 | 1.4092 | 1.4092 | 0.004506
2019-03-28 15:42 | 1553802116.2831 | 215 | 0.20541 | 0.003586 | 8.8e-05 | 98.1808 | 89.796 | 114.2622 | 21.2612 | 0.003394 | 0.00306 | 163.1727 | 0.25804 | 8.402 | 0.002754 | 0.004182 | 0.004103 | 4.4e-05 | 0.001123 | 0.32477 | 1.2977 | 2.5473 | 100.63 | 21.15 | 25.551 | 103.5106 | 88.0737 | 95.4364 | 0.006381 | 0.95784 | 0.022995 | 24.0071 | 0.79848 | 0.15351 | 0.020548 | 1.4151 | 1.4151 | 0.00448
2019-03-28 15:42 | 1553802117.2842 | 216 | 0.20538 | 0.003594 | 9.1e-05 | 98.0956 | 90.0739 | 114.4347 | 21.2612 | 0.003352 | 0.003084 | 163.0731 | 0.2591 | 8.3997 | 0.002744 | 0.003741 | 0.003591 | 4.6e-05 | 0.001127 | 0.33035 | 1.3211 | 2.5474 | 100.63 | 21.15 | 25.551 | 102.5 | 87.8262 | 96.0277 | 0.006381 | 0.9563 | 0.022994 | 24.0445 | 0.76307 | 0.15414 | 0.020739 | 1.428 | 1.428 | 0.004458
2019-03-28 15:42 | 1553802118.2832 | 217 | 0.20539 | 0.003596 | 9.5e-05 | 97.9475 | 89.759 | 114.476 | 21.2612 | 0.003359 | 0.003044 | 163.0072 | 0.25897 | 8.4044 | 0.002802 | 0.004095 | 0.003974 | 4.5e-05 | 0.001126 | 0.31116 | 1.3248 | 2.5474 | 100.63 | 21.15 | 25.554 | 102.6753 | 89.2252 | 95.0505 | 0.006381 | 0.95884 | 0.023028 | 24.017 | 0.78089 | 0.15406 | 0.02069 | 1.4265 | 1.4265 | 0.00446
2019-03-28 15:42 | 1553802119.2832 | 218 | 0.20538 | 0.003607 | 9.7e-05 | 98.0437 | 89.9479 | 114.5544 | 21.2612 | 0.003454 | 0.003111 | 162.9173 | 0.2586 | 8.4019 | 0.002784 | 0.004295 | 0.004176 | 4.5e-05 | 0.001126 | 0.32693 | 1.3095 | 2.5475 | 100.63 | 21.15 | 25.554 | 104.9759 | 88.8002 | 96.6803 | 0.006381 | 0.95461 | 0.022986 | 24.0783 | 0.77087 | 0.15384 | 0.020732 | 1.4266 | 1.4266 | 0.004454
2019-03-28 15:42 | 1553802120.2826 | 219 | 0.20538 | 0.003616 | 0.0001 | 98.0157 | 89.8476 | 114.3744 | 21.2612 | 0.003339 | 0.003091 | 162.8362 | 0.25763 | 8.4029 | 0.002854 | 0.004965 | 0.004907 | 4.4e-05 | 0.001124 | 0.32604 | 1.3195 | 2.5474 | 100.63 | 21.15 | 25.555 | 102.1867 | 90.4862 | 96.1931 | 0.006382 | 0.95587 | 0.022983 | 24.0436 | 0.77829 | 0.15326 | 0.020708 | 1.4247 | 1.4247 | 0.004443
2019-03-28 15:42 | 1553802121.2833 | 220 | 0.2054 | 0.003621 | 0.000101 | 98.0104 | 89.8837 | 114.1939 | 21.2612 | 0.003419 | 0.003096 | 162.7497 | 0.25897 | 8.4011 | 0.002764 | 0.004591 | 0.004466 | 4.9e-05 | 0.001133 | 0.32029 | 1.2774 | 2.5473 | 100.63 | 21.15 | 25.555 | 104.1291 | 88.3166 | 96.3135 | 0.006382 | 0.95556 | 0.022961 | 24.0285 | 0.77422 | 0.15406 | 0.02055 | 1.4123 | 1.4123 | 0.004505
2019-03-28 15:42 | 1553802122.2823 | 221 | 0.20539 | 0.003622 | 0.000103 | 98.0032 | 89.7865 | 114.3535 | 21.2612 | 0.003392 | 0.003081 | 162.6789 | 0.25899 | 8.4041 | 0.002732 | 0.004406 | 0.004284 | 4.8e-05 | 0.001128 | 0.30574 | 1.2814 | 2.5474 | 100.63 | 21.15 | 25.555 | 103.4689 | 87.5333 | 95.9493 | 0.006382 | 0.95651 | 0.022988 | 24.0334 | 0.78324 | 0.15407 | 0.020644 | 1.4204 | 1.4204 | 0.00448
2019-03-28 15:42 | 1553802123.2833 | 222 | 0.2055 | 0.003626 | 0.000103 | 97.9919 | 89.9886 | 114.7211 | 21.2612 | 0.003389 | 0.00307 | 162.6421 | 0.25858 | 8.4009 | 0.002736 | 0.004439 | 0.004302 | 4.8e-05 | 0.001128 | 0.31813 | 1.3058 | 2.5473 | 100.62 | 21.15 | 25.554 | 103.3911 | 87.6318 | 95.6886 | 0.006382 | 0.95718 | 0.023033 | 24.0635 | 0.76209 | 0.15383 | 0.019959 | 1.3758 | 1.3758 | 0.004618
2019-03-28 15:42 | 1553802124.2824 | 223 | 0.2054 | 0.003619 | 0.000104 | 97.977 | 90.5247 | 114.7969 | 21.2612 | 0.00342 | 0.003117 | 162.5749 | 0.25763 | 8.4043 | 0.002711 | 0.004074 | 0.003997 | 4.3e-05 | 0.001122 | 0.31349 | 1.2935 | 2.5474 | 100.62 | 21.15 | 25.552 | 104.163 | 87.0462 | 96.8314 | 0.006382 | 0.95422 | 0.023005 | 24.1087 | 0.7076 | 0.15327 | 0.020541 | 1.4143 | 1.4143 | 0.004476
2019-03-28 15:42 | 1553802125.2824 | 224 | 0.2055 | 0.003617 | 0.000104 | 97.9014 | 90.4374 | 114.9941 | 21.2612 | 0.003377 | 0.003064 | 162.522 | 0.25792 | 8.4064 | 0.002818 | 0.004701 | 0.004599 | 4.5e-05 | 0.001118 | 0.33339 | 1.3262 | 2.5474 | 100.62 | 21.15 | 25.549 | 103.109 | 89.6088 | 95.5351 | 0.006381 | 0.95758 | 0.023065 | 24.0871 | 0.70932 | 0.15344 | 0.019961 | 1.3779 | 1.3779 | 0.004599
2019-03-28 15:42 | 1553802126.2834 | 225 | 0.20556 | 0.003608 | 0.000104 | 97.8793 | 90.7958 | 115.1523 | 21.2612 | 0.003395 | 0.003081 | 162.4427 | 0.25756 | 8.4016 | 0.002759 | 0.003979 | 0.003915 | 4.6e-05 | 0.001113 | 0.31509 | 1.2918 | 2.5474 | 100.62 | 21.15 | 25.54 | 103.5435 | 88.1964 | 95.9485 | 0.006379 | 0.95651 | 0.023068 | 24.1172 | 0.67193 | 0.15322 | 0.019613 | 1.354 | 1.354 | 0.004674
2019-03-28 15:42 | 1553802127.2825 | 226 | 0.20545 | 0.003599 | 0.000105 | 97.8954 | 90.9864 | 114.9661 | 21.2612 | 0.00339 | 0.003066 | 162.3861 | 0.25734 | 8.4055 | 0.002743 | 0.004056 | 0.004041 | 4.7e-05 | 0.001116 | 0.34092 | 1.2836 | 2.5474 | 100.62 | 21.15 | 25.535 | 103.4248 | 87.805 | 95.5853 | 0.006377 | 0.95745 | 0.023061 | 24.0858 | 0.65465 | 0.15309 | 0.020275 | 1.3993 | 1.3993 | 0.004519
2019-03-28 15:42 | 1553802128.2825 | 227 | 0.20553 | 0.003595 | 0.000106 | 97.8763 | 90.5641 | 114.6718 | 21.2612 | 0.003439 | 0.003061 | 162.3051 | 0.25836 | 8.3998 | 0.002786 | 0.004453 | 0.004351 | 4.3e-05 | 0.001114 | 0.31564 | 1.2908 | 2.5473 | 100.62 | 21.15 | 25.525 | 104.6199 | 88.8416 | 95.474 | 0.006375 | 0.95774 | 0.023035 | 24.0514 | 0.69451 | 0.1537 | 0.019839 | 1.3675 | 1.3675 | 0.004642
2019-03-28 15:42 | 1553802129.2835 | 228 | 0.20555 | 0.003589 | 0.00011 | 97.8299 | 90.451 | 114.5033 | 21.2612 | 0.003382 | 0.003068 | 162.252 | 0.25831 | 8.4039 | 0.002732 | 0.004989 | 0.004815 | 4.5e-05 | 0.00112 | 0.31116 | 1.325 | 2.5473 | 100.62 | 21.15 | 25.52 | 103.2424 | 87.5357 | 95.6362 | 0.006374 | 0.95732 | 0.023013 | 24.039 | 0.70145 | 0.15367 | 0.019716 | 1.3575 | 1.3575 | 0.004675
2019-03-28 15:42 | 1553802130.2826 | 229 | 0.20547 | 0.003579 | 0.000113 | 97.7748 | 90.3137 | 114.5213 | 21.2612 | 0.003372 | 0.00307 | 162.1827 | 0.25851 | 8.4039 | 0.002746 | 0.004054 | 0.004079 | 4.7e-05 | 0.00112 | 0.32516 | 1.3132 | 2.5474 | 100.62 | 21.15 | 25.516 | 102.9816 | 87.8684 | 95.6973 | 0.006373 | 0.95716 | 0.023013 | 24.0429 | 0.70999 | 0.15379 | 0.020184 | 1.3896 | 1.3896 | 0.004571
2019-03-28 15:42 | 1553802131.283 | 230 | 0.20548 | 0.003566 | 0.000115 | 97.7986 | 89.7515 | 114.7131 | 21.2612 | 0.003356 | 0.003075 | 162.1065 | 0.25839 | 8.3984 | 0.002797 | 0.004315 | 0.004154 | 4.9e-05 | 0.001126 | 0.31924 | 1.3366 | 2.5474 | 100.62 | 21.15 | 25.507 | 102.613 | 89.1033 | 95.8013 | 0.00637 | 0.95689 | 0.023029 | 24.0664 | 0.76803 | 0.15371 | 0.020136 | 1.3872 | 1.3872 | 0.004577
2019-03-28 15:42 | 1553802132.2827 | 231 | 0.20547 | 0.003559 | 0.000116 | 97.6991 | 89.617 | 114.377 | 21.2612 | 0.003396 | 0.003128 | 162.0306 | 0.25872 | 8.4028 | 0.002709 | 0.004102 | 0.004055 | 5.1e-05 | 0.001131 | 0.30718 | 1.3291 | 2.5474 | 100.62 | 21.15 | 25.502 | 103.5772 | 86.9767 | 97.0797 | 0.006369 | 0.95359 | 0.022955 | 24.0727 | 0.77234 | 0.15391 | 0.020199 | 1.387 | 1.387 | 0.004583
2019-03-28 15:42 | 1553802133.2836 | 232 | 0.20542 | 0.003553 | 0.000117 | 97.681 | 90.0589 | 114.188 | 21.2612 | 0.003407 | 0.003133 | 161.9537 | 0.2578 | 8.4027 | 0.00273 | 0.004308 | 0.004245 | 4.7e-05 | 0.001141 | 0.31166 | 1.3076 | 2.5475 | 100.62 | 21.15 | 25.499 | 103.8359 | 87.4835 | 97.1988 | 0.006368 | 0.95328 | 0.022933 | 24.0567 | 0.72667 | 0.15337 | 0.020499 | 1.4062 | 1.4062 | 0.004504
2019-03-28 15:42 | 1553802134.2826 | 233 | 0.20548 | 0.003544 | 0.000119 | 97.5502 | 90.0498 | 114.3627 | 21.2612 | 0.003419 | 0.003058 | 161.9175 | 0.25769 | 8.4043 | 0.002706 | 0.003829 | 0.003691 | 5e-05 | 0.001142 | 0.32565 | 1.3262 | 2.5474 | 100.62 | 21.15 | 25.495 | 104.1267 | 86.9213 | 95.4043 | 0.006367 | 0.95792 | 0.023006 | 24.0166 | 0.7156 | 0.1533 | 0.020167 | 1.3877 | 1.3877 | 0.004562
2019-03-28 15:42 | 1553802135.2827 | 234 | 0.20543 | 0.003542 | 0.000121 | 97.597 | 90.4347 | 114.3112 | 21.2612 | 0.00345 | 0.003066 | 161.8056 | 0.25754 | 8.4024 | 0.002694 | 0.004047 | 0.003929 | 4.7e-05 | 0.001141 | 0.30718 | 1.3307 | 2.5475 | 100.62 | 21.15 | 25.492 | 104.8809 | 86.6191 | 95.5915 | 0.006367 | 0.95743 | 0.022995 | 24.0173 | 0.68174 | 0.15321 | 0.020447 | 1.4062 | 1.4062 | 0.0045
2019-03-28 15:42 | 1553802136.2832 | 235 | 0.20543 | 0.003542 | 0.000123 | 97.602 | 90.5227 | 113.9338 | 21.2612 | 0.003331 | 0.003112 | 161.7346 | 0.25834 | 8.4042 | 0.002732 | 0.004091 | 0.004044 | 4.7e-05 | 0.001138 | 0.31526 | 1.3078 | 2.5475 | 100.62 | 21.15 | 25.489 | 101.9988 | 87.549 | 96.6943 | 0.006366 | 0.95458 | 0.022923 | 24.0135 | 0.67351 | 0.15369 | 0.020466 | 1.403 | 1.403 | 0.004524
2019-03-28 15:42 | 1553802137.2828 | 236 | 0.20539 | 0.003542 | 0.000126 | 97.571 | 90.3977 | 113.9322 | 21.2612 | 0.003452 | 0.003164 | 161.6622 | 0.25871 | 8.4014 | 0.002694 | 0.004025 | 0.003907 | 5.1e-05 | 0.001143 | 0.33483 | 1.302 | 2.5474 | 100.62 | 21.15 | 25.489 | 104.9311 | 86.6379 | 97.9519 | 0.006366 | 0.95134 | 0.022884 | 24.0541 | 0.68303 | 0.1539 | 0.020715 | 1.4176 | 1.4176 | 0.004484
2019-03-28 15:42 | 1553802138.2819 | 237 | 0.2054 | 0.003535 | 0.000128 | 97.4436 | 90.2217 | 114.3368 | 21.2612 | 0.00341 | 0.003085 | 161.5968 | 0.25837 | 8.4024 | 0.002719 | 0.004584 | 0.004483 | 4.5e-05 | 0.001144 | 0.32477 | 1.321 | 2.5475 | 100.62 | 21.15 | 25.49 | 103.905 | 87.2397 | 96.0403 | 0.006366 | 0.95627 | 0.022984 | 24.0346 | 0.68876 | 0.1537 | 0.020642 | 1.4186 | 1.4186 | 0.004475
2019-03-28 15:42 | 1553802139.2828 | 238 | 0.20542 | 0.003546 | 0.00013 | 97.4096 | 90.5104 | 114.3927 | 21.2612 | 0.003426 | 0.003072 | 161.5128 | 0.25873 | 8.4031 | 0.002719 | 0.004059 | 0.004027 | 5e-05 | 0.001143 | 0.33395 | 1.3389 | 2.5475 | 100.62 | 21.15 | 25.494 | 104.2871 | 87.2233 | 95.7371 | 0.006367 | 0.95705 | 0.022999 | 24.0306 | 0.65706 | 0.15392 | 0.020473 | 1.4077 | 1.4077 | 0.004516
2019-03-28 15:42 | 1553802140.2823 | 239 | 0.20539 | 0.003547 | 0.000132 | 97.5002 | 90.6558 | 114.2666 | 21.2612 | 0.003414 | 0.003117 | 161.4268 | 0.25862 | 8.403 | 0.002714 | 0.004013 | 0.003803 | 4.9e-05 | 0.001141 | 0.33892 | 1.2768 | 2.5474 | 100.62 | 21.15 | 25.496 | 104.0026 | 87.1155 | 96.8203 | 0.006368 | 0.95425 | 0.022952 | 24.0526 | 0.65102 | 0.15385 | 0.020676 | 1.4187 | 1.4187 | 0.004479
2019-03-28 15:42 | 1553802141.2819 | 240 | 0.20543 | 0.003553 | 0.000133 | 97.3551 | 90.5519 | 114.0833 | 21.2612 | 0.003386 | 0.003104 | 161.3475 | 0.25851 | 8.404 | 0.002693 | 0.003618 | 0.00355 | 4.8e-05 | 0.00114 | 0.33411 | 1.293 | 2.5474 | 100.62 | 21.16 | 25.499 | 103.3369 | 86.6114 | 96.4968 | 0.006372 | 0.95509 | 0.022944 | 24.0228 | 0.64796 | 0.15379 | 0.020434 | 1.4015 | 1.4015 | 0.004532
2019-03-28 15:42 | 1553802142.2824 | 241 | 0.20545 | 0.003555 | 0.000134 | 97.362 | 90.4625 | 114.1533 | 21.2612 | 0.003452 | 0.003095 | 161.2427 | 0.25853 | 8.4045 | 0.002692 | 0.004467 | 0.004385 | 4.9e-05 | 0.001142 | 0.2995 | 1.3061 | 2.5473 | 100.62 | 21.16 | 25.501 | 104.9156 | 86.5871 | 96.3003 | 0.006373 | 0.9556 | 0.022957 | 24.0238 | 0.65742 | 0.1538 | 0.020294 | 1.3926 | 1.3926 | 0.004561
2019-03-28 15:42 | 1553802143.2831 | 242 | 0.20542 | 0.003556 | 0.000137 | 97.3396 | 90.5906 | 114.1197 | 21.2612 | 0.003403 | 0.003053 | 161.1627 | 0.25818 | 8.4022 | 0.002744 | 0.003705 | 0.003597 | 4.8e-05 | 0.001139 | 0.3402 | 1.2735 | 2.5475 | 100.62 | 21.16 | 25.503 | 103.7494 | 87.8214 | 95.2822 | 0.006373 | 0.95824 | 0.022985 | 23.9871 | 0.64271 | 0.15359 | 0.020505 | 1.4087 | 1.4087 | 0.004503
2019-03-28 15:42 | 1553802144.2831 | 243 | 0.20539 | 0.00356 | 0.00014 | 97.2976 | 90.8726 | 114.0275 | 21.2612 | 0.003381 | 0.003095 | 161.0753 | 0.25877 | 8.4021 | 0.002729 | 0.004269 | 0.0042 | 4.8e-05 | 0.001139 | 0.33157 | 1.2954 | 2.5474 | 100.62 | 21.16 | 25.509 | 103.1962 | 87.4638 | 96.2846 | 0.006375 | 0.95564 | 0.022945 | 24.01 | 0.61106 | 0.15394 | 0.020641 | 1.4153 | 1.4153 | 0.004492
2019-03-28 15:42 | 1553802145.2834 | 244 | 0.20543 | 0.003561 | 0.000145 | 97.3608 | 91.0747 | 114.0667 | 21.2612 | 0.003437 | 0.003105 | 161.0088 | 0.25881 | 8.4029 | 0.002725 | 0.004635 | 0.004563 | 5.1e-05 | 0.00114 | 0.31902 | 1.3042 | 2.5474 | 100.62 | 21.16 | 25.512 | 104.5537 | 87.375 | 96.5313 | 0.006375 | 0.955 | 0.022941 | 24.0222 | 0.59698 | 0.15397 | 0.020395 | 1.3979 | 1.3979 | 0.004549
2019-03-28 15:42 | 1553802146.282 | 245 | 0.20539 | 0.003559 | 0.00015 | 97.2553 | 91.2381 | 114.2826 | 21.2612 | 0.003411 | 0.003131 | 160.9312 | 0.2584 | 8.4005 | 0.002743 | 0.004348 | 0.004266 | 4.9e-05 | 0.001142 | 0.30685 | 1.2992 | 2.5475 | 100.62 | 21.16 | 25.515 | 103.9287 | 87.8082 | 97.1526 | 0.006376 | 0.9534 | 0.022944 | 24.0651 | 0.57126 | 0.15372 | 0.02061 | 1.4126 | 1.4126 | 0.004494
2019-03-28 15:42 | 1553802147.2832 | 246 | 0.20543 | 0.003558 | 0.000154 | 97.2726 | 90.948 | 114.3765 | 21.2612 | 0.003388 | 0.003116 | 160.8418 | 0.2581 | 8.4034 | 0.00272 | 0.003809 | 0.003704 | 4.8e-05 | 0.001144 | 0.32853 | 1.2946 | 2.5474 | 100.62 | 21.15 | 25.521 | 103.3691 | 87.2445 | 96.793 | 0.006374 | 0.95432 | 0.022964 | 24.0633 | 0.60134 | 0.15354 | 0.020383 | 1.398 | 1.398 | 0.004536
2019-03-28 15:42 | 1553802148.2826 | 247 | 0.20534 | 0.003559 | 0.000157 | 97.2646 | 90.7221 | 114.227 | 21.2612 | 0.003455 | 0.003075 | 160.7764 | 0.25869 | 8.4005 | 0.00269 | 0.004552 | 0.004421 | 5e-05 | 0.001143 | 0.33627 | 1.3004 | 2.5474 | 100.62 | 21.15 | 25.526 | 104.991 | 86.5308 | 95.8006 | 0.006375 | 0.95689 | 0.02298 | 24.0153 | 0.62284 | 0.15389 | 0.020919 | 1.4357 | 1.4357 | 0.004427
2019-03-28 15:42 | 1553802149.2832 | 248 | 0.20527 | 0.003559 | 0.000159 | 97.3116 | 90.4367 | 114.3389 | 21.2612 | 0.00341 | 0.00311 | 160.6973 | 0.2584 | 8.4029 | 0.002671 | 0.003572 | 0.003436 | 4.8e-05 | 0.001142 | 0.31868 | 1.3045 | 2.5476 | 100.62 | 21.15 | 25.53 | 103.905 | 86.0773 | 96.6482 | 0.006376 | 0.9547 | 0.022965 | 24.0546 | 0.65535 | 0.15372 | 0.021336 | 1.4635 | 1.4635 | 0.004338
2019-03-28 15:42 | 1553802150.2822 | 249 | 0.20522 | 0.003568 | 0.000159 | 97.2556 | 90.4445 | 114.5263 | 21.2612 | 0.003425 | 0.003127 | 160.64 | 0.25815 | 8.4029 | 0.002705 | 0.003792 | 0.00364 | 4.4e-05 | 0.001142 | 0.30558 | 1.2917 | 2.5475 | 100.62 | 21.15 | 25.533 | 104.2776 | 86.9002 | 97.0703 | 0.006377 | 0.95361 | 0.022971 | 24.0881 | 0.64943 | 0.15357 | 0.02164 | 1.4847 | 1.4847 | 0.004272
2019-03-28 15:42 | 1553802151.2816 | 250 | 0.20527 | 0.003581 | 0.000159 | 97.2526 | 90.3368 | 114.4631 | 21.2612 | 0.003394 | 0.003124 | 160.5742 | 0.25876 | 8.4085 | 0.002727 | 0.003772 | 0.003755 | 4.4e-05 | 0.001144 | 0.33035 | 1.3393 | 2.5475 | 100.62 | 21.15 | 25.542 | 103.5295 | 87.4202 | 96.9998 | 0.006379 | 0.95379 | 0.022966 | 24.0791 | 0.65981 | 0.15394 | 0.0213 | 1.461 | 1.461 | 0.004352
2019-03-28 15:43 | 1553802152.2823 | 251 | 0.2053 | 0.003588 | 0.000161 | 97.3125 | 90.8767 | 114.4328 | 21.2612 | 0.003453 | 0.003145 | 160.5194 | 0.25832 | 8.4025 | 0.002714 | 0.004243 | 0.004068 | 4.7e-05 | 0.001143 | 0.30845 | 1.3165 | 2.5474 | 100.62 | 21.15 | 25.545 | 104.9391 | 87.1117 | 97.4975 | 0.00638 | 0.95251 | 0.022948 | 24.0921 | 0.61202 | 0.15368 | 0.021106 | 1.4464 | 1.4464 | 0.004388
2019-03-28 15:43 | 1553802153.2827 | 252 | 0.2053 | 0.003598 | 0.000162 | 97.3039 | 90.769 | 114.6246 | 21.2612 | 0.003432 | 0.003133 | 160.4446 | 0.25825 | 8.4047 | 0.002738 | 0.003932 | 0.003857 | 4.3e-05 | 0.001147 | 0.29679 | 1.2624 | 2.5474 | 100.62 | 21.15 | 25.549 | 104.4524 | 87.6895 | 97.2176 | 0.006381 | 0.95323 | 0.022976 | 24.1032 | 0.62184 | 0.15363 | 0.021149 | 1.451 | 1.451 | 0.004373
2019-03-28 15:43 | 1553802154.2827 | 253 | 0.20523 | 0.0036 | 0.000163 | 97.3164 | 90.9144 | 114.5058 | 21.2612 | 0.003468 | 0.003166 | 160.3861 | 0.25786 | 8.4029 | 0.002706 | 0.004461 | 0.004344 | 4.6e-05 | 0.00115 | 0.3194 | 1.3133 | 2.5472 | 100.62 | 21.15 | 25.549 | 105.3032 | 86.9276 | 97.9952 | 0.006381 | 0.95123 | 0.02294 | 24.116 | 0.60867 | 0.1534 | 0.021538 | 1.4754 | 1.4754 | 0.004294
2019-03-28 15:43 | 1553802155.2824 | 254 | 0.20526 | 0.003604 | 0.000163 | 97.2809 | 90.7265 | 114.223 | 21.2612 | 0.003434 | 0.003164 | 160.3317 | 0.25803 | 8.4023 | 0.002714 | 0.003985 | 0.003886 | 4.1e-05 | 0.001152 | 0.32084 | 1.3545 | 2.5475 | 100.62 | 21.15 | 25.551 | 104.4868 | 87.1033 | 97.9505 | 0.006381 | 0.95135 | 0.022913 | 24.0847 | 0.62391 | 0.1535 | 0.021334 | 1.4597 | 1.4597 | 0.004343
2019-03-28 15:43 | 1553802156.2814 | 255 | 0.20528 | 0.003605 | 0.000162 | 97.3533 | 90.3725 | 114.1758 | 21.2612 | 0.003468 | 0.00311 | 160.2626 | 0.25769 | 8.4051 | 0.002838 | 0.003911 | 0.003837 | 4.4e-05 | 0.00115 | 0.29695 | 1.3217 | 2.5474 | 100.62 | 21.15 | 25.551 | 105.3183 | 90.0837 | 96.6544 | 0.006381 | 0.95468 | 0.022948 | 24.0377 | 0.66553 | 0.1533 | 0.02122 | 1.4542 | 1.4542 | 0.004354
2019-03-28 15:43 | 1553802157.2815 | 256 | 0.20518 | 0.003612 | 0.000162 | 97.3346 | 90.699 | 114.0667 | 21.2612 | 0.003469 | 0.003137 | 160.2005 | 0.25849 | 8.4035 | 0.002768 | 0.004086 | 0.004031 | 4.5e-05 | 0.001149 | 0.3178 | 1.3681 | 2.5476 | 100.62 | 21.16 | 25.552 | 105.3403 | 88.4136 | 97.2992 | 0.006385 | 0.95302 | 0.022917 | 24.0471 | 0.63155 | 0.15378 | 0.021847 | 1.4953 | 1.4953 | 0.004247
2019-03-28 15:43 | 1553802158.2815 | 257 | 0.20521 | 0.003613 | 0.000162 | 97.3518 | 91.0061 | 113.9026 | 21.2612 | 0.00343 | 0.003082 | 160.1053 | 0.25858 | 8.4041 | 0.002796 | 0.004357 | 0.004194 | 4.6e-05 | 0.001147 | 0.31293 | 1.3217 | 2.5474 | 100.62 | 21.15 | 25.554 | 104.3862 | 89.0783 | 95.9799 | 0.006382 | 0.95643 | 0.022942 | 23.987 | 0.60291 | 0.15383 | 0.021622 | 1.4814 | 1.4814 | 0.004289
2019-03-28 15:43 | 1553802159.2825 | 258 | 0.20529 | 0.003612 | 0.000165 | 97.3858 | 91.0971 | 113.7312 | 21.2612 | 0.003455 | 0.003135 | 160.0504 | 0.25806 | 8.4054 | 0.002839 | 0.00432 | 0.004269 | 4.5e-05 | 0.00115 | 0.30127 | 1.2944 | 2.5474 | 100.62 | 21.15 | 25.558 | 104.9885 | 90.1293 | 97.2451 | 0.006383 | 0.95316 | 0.022885 | 24.01 | 0.59708 | 0.15352 | 0.021147 | 1.445 | 1.445 | 0.004388
2019-03-28 15:43 | 1553802160.2829 | 259 | 0.20514 | 0.003616 | 0.000167 | 97.3092 | 90.9596 | 113.6636 | 21.2612 | 0.00346 | 0.003152 | 159.9721 | 0.25806 | 8.4024 | 0.002736 | 0.004163 | 0.00409 | 4.5e-05 | 0.001145 | 0.30973 | 1.308 | 2.5475 | 100.62 | 21.15 | 25.56 | 105.1325 | 87.6474 | 97.6715 | 0.006384 | 0.95206 | 0.022865 | 24.0167 | 0.60357 | 0.15352 | 0.022078 | 1.5075 | 1.5075 | 0.004206
2019-03-28 15:43 | 1553802161.2826 | 260 | 0.20524 | 0.00362 | 0.000168 | 97.3417 | 90.7139 | 113.9376 | 21.2612 | 0.003445 | 0.003189 | 159.9252 | 0.25728 | 8.4046 | 0.002764 | 0.004159 | 0.003945 | 4.6e-05 | 0.001148 | 0.32211 | 1.3303 | 2.5474 | 100.62 | 21.15 | 25.566 | 104.7503 | 88.3144 | 98.5674 | 0.006385 | 0.94977 | 0.022865 | 24.0746 | 0.63074 | 0.15305 | 0.021457 | 1.4648 | 1.4648 | 0.004315
2019-03-28 15:43 | 1553802162.282 | 261 | 0.20515 | 0.003625 | 0.000168 | 97.3938 | 90.8689 | 114.0704 | 21.2612 | 0.003433 | 0.003148 | 159.8309 | 0.25751 | 8.4004 | 0.002719 | 0.004476 | 0.004473 | 4.6e-05 | 0.001145 | 0.31238 | 1.2917 | 2.5473 | 100.62 | 21.15 | 25.569 | 104.4626 | 87.2219 | 97.5702 | 0.006386 | 0.95232 | 0.022909 | 24.0563 | 0.62026 | 0.15319 | 0.021954 | 1.5018 | 1.5018 | 0.004213
2019-03-28 15:43 | 1553802163.2831 | 262 | 0.20509 | 0.003631 | 0.00017 | 97.4311 | 90.9633 | 113.869 | 21.2612 | 0.003404 | 0.003142 | 159.7623 | 0.25802 | 8.4064 | 0.002726 | 0.004368 | 0.004211 | 4.4e-05 | 0.001141 | 0.3131 | 1.2975 | 2.5475 | 100.62 | 21.15 | 25.573 | 103.7739 | 87.3906 | 97.4198 | 0.006387 | 0.95271 | 0.022894 | 24.0302 | 0.61439 | 0.15349 | 0.022329 | 1.5264 | 1.5264 | 0.004153
2019-03-28 15:43 | 1553802164.2817 | 263 | 0.20516 | 0.003639 | 0.000169 | 97.3837 | 91.1158 | 113.9925 | 21.2612 | 0.003426 | 0.003122 | 159.6764 | 0.25799 | 8.4023 | 0.002681 | 0.003571 | 0.003605 | 4.8e-05 | 0.001143 | 0.30807 | 1.2952 | 2.5475 | 100.62 | 21.15 | 25.581 | 104.2871 | 86.3155 | 96.9334 | 0.006389 | 0.95396 | 0.022921 | 24.0274 | 0.59506 | 0.15348 | 0.021917 | 1.5 | 1.5 | 0.004226
2019-03-28 15:43 | 1553802165.2828 | 264 | 0.20511 | 0.003652 | 0.000168 | 97.4355 | 91.3849 | 114.2897 | 21.2612 | 0.003461 | 0.003209 | 159.5942 | 0.25783 | 8.4016 | 0.002729 | 0.003668 | 0.003513 | 4.7e-05 | 0.001143 | 0.32388 | 1.3042 | 2.5474 | 100.62 | 21.15 | 25.586 | 105.1444 | 87.4693 | 99.0348 | 0.00639 | 0.94858 | 0.022886 | 24.1269 | 0.57344 | 0.15338 | 0.022156 | 1.5141 | 1.5141 | 0.004184
2019-03-28 15:43 | 1553802166.2817 | 265 | 0.20514 | 0.003666 | 0.000166 | 97.4728 | 91.2385 | 114.2807 | 21.2612 | 0.003404 | 0.003173 | 159.5361 | 0.25793 | 8.4018 | 0.002745 | 0.003955 | 0.003888 | 4.7e-05 | 0.001148 | 0.32029 | 1.3069 | 2.5475 | 100.62 | 21.15 | 25.599 | 103.7518 | 87.8628 | 98.181 | 0.006393 | 0.95076 | 0.022912 | 24.0983 | 0.5912 | 0.15344 | 0.021988 | 1.5043 | 1.5043 | 0.004213
2019-03-28 15:43 | 1553802167.281 | 266 | 0.20511 | 0.003671 | 0.000163 | 97.543 | 91.6723 | 114.2256 | 21.2612 | 0.003468 | 0.003192 | 159.4592 | 0.25812 | 8.4051 | 0.002674 | 0.003943 | 0.003826 | 4.4e-05 | 0.001146 | 0.32189 | 1.3056 | 2.5474 | 100.62 | 21.15 | 25.605 | 105.319 | 86.1468 | 98.6326 | 0.006395 | 0.9496 | 0.022892 | 24.1071 | 0.55522 | 0.15356 | 0.022143 | 1.5139 | 1.5139 | 0.004189
2019-03-28 15:43 | 1553802168.2818 | 267 | 0.20511 | 0.003679 | 0.00016 | 97.5401 | 92.0059 | 114.1636 | 21.2612 | 0.003425 | 0.003199 | 159.4176 | 0.2583 | 8.4003 | 0.002669 | 0.003557 | 0.00342 | 4.4e-05 | 0.001145 | 0.32549 | 1.2741 | 2.5473 | 100.62 | 21.15 | 25.61 | 104.2612 | 86.0337 | 98.7955 | 0.006396 | 0.94919 | 0.022881 | 24.1059 | 0.52246 | 0.15366 | 0.022131 | 1.5125 | 1.5125 | 0.004196
2019-03-28 15:43 | 1553802169.2819 | 268 | 0.20515 | 0.003688 | 0.000157 | 97.5299 | 92.0548 | 114.1941 | 21.2612 | 0.003472 | 0.00317 | 159.3261 | 0.25858 | 8.4016 | 0.00266 | 0.004647 | 0.004551 | 4.1e-05 | 0.001145 | 0.32836 | 1.2922 | 2.5475 | 100.62 | 21.15 | 25.621 | 105.4102 | 85.8105 | 98.1119 | 0.006399 | 0.95093 | 0.022905 | 24.0869 | 0.51677 | 0.15383 | 0.021913 | 1.4993 | 1.4993 | 0.004237
2019-03-28 15:43 | 1553802170.2819 | 269 | 0.20518 | 0.003689 | 0.000156 | 97.6199 | 92.011 | 114.3514 | 21.2612 | 0.003496 | 0.003122 | 159.2523 | 0.25815 | 8.4013 | 0.002657 | 0.003845 | 0.003744 | 4.7e-05 | 0.001146 | 0.31708 | 1.293 | 2.5475 | 100.62 | 21.15 | 25.625 | 105.9963 | 85.7253 | 96.9341 | 0.0064 | 0.95396 | 0.022957 | 24.0653 | 0.52928 | 0.15357 | 0.021731 | 1.4903 | 1.4903 | 0.004256
2019-03-28 15:43 | 1553802171.2813 | 270 | 0.20517 | 0.003696 | 0.000154 | 97.6107 | 92.0796 | 114.3988 | 21.2612 | 0.003507 | 0.003152 | 159.169 | 0.25814 | 8.4034 | 0.002684 | 0.005139 | 0.005041 | 4.9e-05 | 0.001145 | 0.3126 | 1.3437 | 2.5474 | 100.62 | 21.15 | 25.632 | 106.2695 | 86.392 | 97.6676 | 0.006402 | 0.95207 | 0.022939 | 24.0941 | 0.52176 | 0.15356 | 0.021761 | 1.4913 | 1.4913 | 0.004253
2019-03-28 15:43 | 1553802172.2819 | 271 | 0.20516 | 0.003695 | 0.000151 | 97.6845 | 92.1663 | 114.6147 | 21.2612 | 0.003493 | 0.003205 | 159.1335 | 0.25792 | 8.4054 | 0.002688 | 0.004358 | 0.004248 | 4.6e-05 | 0.001145 | 0.31205 | 1.3443 | 2.5476 | 100.62 | 21.15 | 25.633 | 105.9209 | 86.4849 | 98.9461 | 0.006402 | 0.9488 | 0.022921 | 24.1584 | 0.52012 | 0.15344 | 0.021872 | 1.4979 | 1.4979 | 0.004231
2019-03-28 15:43 | 1553802173.281 | 272 | 0.20515 | 0.003696 | 0.000148 | 97.6291 | 92.1282 | 114.6227 | 21.2612 | 0.003485 | 0.003162 | 159.0764 | 0.2583 | 8.4048 | 0.002697 | 0.004022 | 0.003933 | 4.8e-05 | 0.001143 | 0.33069 | 1.3017 | 2.5473 | 100.62 | 21.15 | 25.635 | 105.7242 | 86.7043 | 97.9058 | 0.006402 | 0.95146 | 0.022954 | 24.1254 | 0.51874 | 0.15366 | 0.021936 | 1.5046 | 1.5046 | 0.004218
2019-03-28 15:43 | 1553802174.282 | 273 | 0.20516 | 0.003702 | 0.000146 | 97.7402 | 92.2369 | 114.1472 | 21.2612 | 0.003555 | 0.00314 | 159.009 | 0.25764 | 8.4024 | 0.002654 | 0.003881 | 0.003803 | 4.7e-05 | 0.001139 | 0.29999 | 1.3192 | 2.5474 | 100.62 | 21.15 | 25.635 | 107.4175 | 85.6536 | 97.3838 | 0.006402 | 0.9528 | 0.022923 | 24.0584 | 0.51836 | 0.15327 | 0.021884 | 1.4991 | 1.4991 | 0.004223
2019-03-28 15:43 | 1553802175.2822 | 274 | 0.20528 | 0.003707 | 0.000145 | 97.8162 | 93.1749 | 114.019 | 21.2612 | 0.003497 | 0.003178 | 158.9579 | 0.25808 | 8.4008 | 0.002698 | 0.004016 | 0.003855 | 4.5e-05 | 0.001138 | 0.2975 | 1.2866 | 2.5472 | 100.62 | 21.15 | 25.634 | 106.0225 | 86.7276 | 98.29 | 0.006402 | 0.95048 | 0.022882 | 24.0742 | 0.43481 | 0.15353 | 0.021128 | 1.4446 | 1.4446 | 0.004389
2019-03-28 15:43 | 1553802176.2811 | 275 | 0.20527 | 0.003703 | 0.000144 | 97.8341 | 93.3614 | 113.8453 | 21.2612 | 0.003434 | 0.00319 | 158.8749 | 0.25914 | 8.4024 | 0.002722 | 0.004437 | 0.00437 | 4.6e-05 | 0.001141 | 0.34667 | 1.3126 | 2.5474 | 100.62 | 21.15 | 25.633 | 104.4899 | 87.2943 | 98.5932 | 0.006402 | 0.9497 | 0.022855 | 24.0657 | 0.41856 | 0.15416 | 0.021223 | 1.4494 | 1.4494 | 0.004393
2019-03-28 15:43 | 1553802177.2811 | 276 | 0.20532 | 0.003698 | 0.000143 | 97.9326 | 93.2507 | 113.9444 | 21.2612 | 0.003443 | 0.003106 | 158.8181 | 0.25882 | 8.4024 | 0.002692 | 0.00323 | 0.003116 | 4.7e-05 | 0.001138 | 0.32228 | 1.2928 | 2.5474 | 100.62 | 21.15 | 25.628 | 104.7008 | 86.5847 | 96.5627 | 0.006401 | 0.95492 | 0.022928 | 24.0103 | 0.43818 | 0.15398 | 0.020876 | 1.4303 | 1.4303 | 0.004446
2019-03-28 15:43 | 1553802178.2811 | 277 | 0.20531 | 0.003687 | 0.000139 | 97.9496 | 93.1019 | 114.0972 | 21.2612 | 0.003487 | 0.003131 | 158.723 | 0.25797 | 8.4034 | 0.002686 | 0.004111 | 0.004027 | 4.5e-05 | 0.001137 | 0.29927 | 1.3141 | 2.5472 | 100.62 | 21.15 | 25.625 | 105.7785 | 86.4389 | 97.1502 | 0.0064 | 0.9534 | 0.022925 | 24.0455 | 0.45402 | 0.15347 | 0.021002 | 1.439 | 1.439 | 0.004405
2019-03-28 15:43 | 1553802179.2815 | 278 | 0.2054 | 0.003675 | 0.000135 | 98.0509 | 93.3329 | 114.3279 | 21.2612 | 0.003494 | 0.003135 | 158.6647 | 0.25838 | 8.4036 | 0.002645 | 0.004183 | 0.004097 | 4.4e-05 | 0.001139 | 0.32532 | 1.3022 | 2.5473 | 100.62 | 21.15 | 25.619 | 105.9422 | 85.46 | 97.257 | 0.006398 | 0.95313 | 0.022945 | 24.0733 | 0.44109 | 0.15371 | 0.020443 | 1.402 | 1.402 | 0.004528
2019-03-28 15:43 | 1553802180.2822 | 279 | 0.20541 | 0.003656 | 0.000131 | 98.1632 | 93.4348 | 113.9749 | 21.2612 | 0.003462 | 0.003122 | 158.6086 | 0.25759 | 8.402 | 0.002698 | 0.004643 | 0.004596 | 4.5e-05 | 0.001133 | 0.31205 | 1.2874 | 2.5472 | 100.62 | 21.15 | 25.615 | 105.1679 | 86.7353 | 96.9341 | 0.006397 | 0.95396 | 0.022919 | 24.0256 | 0.44157 | 0.15324 | 0.02045 | 1.4012 | 1.4012 | 0.004517
2019-03-28 15:43 | 1553802181.2823 | 280 | 0.20545 | 0.003643 | 0.000129 | 98.1483 | 94.1027 | 113.6697 | 21.2612 | 0.003517 | 0.003145 | 158.5415 | 0.25788 | 8.4043 | 0.002742 | 0.004244 | 0.004081 | 4.5e-05 | 0.001126 | 0.32908 | 1.2993 | 2.5472 | 100.62 | 21.16 | 25.61 | 106.5041 | 87.7872 | 97.5091 | 0.0064 | 0.95248 | 0.022871 | 24.0121 | 0.3765 | 0.15341 | 0.020174 | 1.3794 | 1.3794 | 0.004593
2019-03-28 15:43 | 1553802182.2817 | 281 | 0.20546 | 0.003624 | 0.000126 | 98.1995 | 94.6157 | 113.3817 | 21.2612 | 0.003441 | 0.003143 | 158.4816 | 0.2581 | 8.4013 | 0.002666 | 0.004534 | 0.004513 | 4.1e-05 | 0.001118 | 0.34772 | 1.3013 | 2.5474 | 100.62 | 21.15 | 25.595 | 104.6504 | 85.9556 | 97.4599 | 0.006392 | 0.95261 | 0.022844 | 23.98 | 0.33254 | 0.15354 | 0.020152 | 1.3764 | 1.3764 | 0.004607
2019-03-28 15:43 | 1553802183.2813 | 282 | 0.20549 | 0.00361 | 0.000125 | 98.2504 | 94.9208 | 113.4968 | 21.2612 | 0.003457 | 0.003166 | 158.415 | 0.25797 | 8.4027 | 0.002651 | 0.004224 | 0.004127 | 4.5e-05 | 0.001112 | 0.34202 | 1.3243 | 2.5474 | 100.62 | 21.16 | 25.587 | 105.0572 | 85.6029 | 98.0022 | 0.006394 | 0.95122 | 0.022838 | 24.0097 | 0.30839 | 0.15346 | 0.020039 | 1.3684 | 1.3684 | 0.004632
2019-03-28 15:43 | 1553802184.2803 | 283 | 0.20556 | 0.003595 | 0.000123 | 98.4119 | 94.9095 | 113.8629 | 21.2612 | 0.003478 | 0.003156 | 158.3518 | 0.25783 | 8.3996 | 0.002732 | 0.003993 | 0.003912 | 4.5e-05 | 0.001105 | 0.30973 | 1.3238 | 2.5473 | 100.62 | 21.16 | 25.568 | 105.5486 | 87.5381 | 97.7693 | 0.006389 | 0.95181 | 0.022882 | 24.0409 | 0.32413 | 0.15338 | 0.019617 | 1.3423 | 1.3423 | 0.004719
2019-03-28 15:43 | 1553802185.2817 | 284 | 0.20561 | 0.003576 | 0.000121 | 98.3648 | 94.8097 | 113.7507 | 21.2612 | 0.003474 | 0.003189 | 158.2977 | 0.25875 | 8.4036 | 0.002674 | 0.003986 | 0.003903 | 4.2e-05 | 0.001104 | 0.32029 | 1.3572 | 2.5472 | 100.62 | 21.16 | 25.557 | 105.4701 | 86.1562 | 98.5666 | 0.006387 | 0.94977 | 0.022847 | 24.0548 | 0.32927 | 0.15393 | 0.019326 | 1.3203 | 1.3203 | 0.004815
2019-03-28 15:43 | 1553802186.2814 | 285 | 0.20569 | 0.003565 | 0.000121 | 98.3925 | 95.0682 | 114.0639 | 21.2612 | 0.003431 | 0.003127 | 158.2473 | 0.25834 | 8.4026 | 0.002678 | 0.004686 | 0.004607 | 4.4e-05 | 0.001102 | 0.32693 | 1.3355 | 2.5475 | 100.62 | 21.16 | 25.547 | 104.4184 | 86.2351 | 97.0657 | 0.006384 | 0.95362 | 0.022924 | 24.0393 | 0.30743 | 0.15368 | 0.01887 | 1.2934 | 1.2934 | 0.004907
2019-03-28 15:43 | 1553802187.2805 | 286 | 0.20578 | 0.003546 | 0.00012 | 98.4741 | 95.3491 | 114.0221 | 21.2612 | 0.003466 | 0.00312 | 158.2059 | 0.25846 | 8.4052 | 0.002754 | 0.004176 | 0.004133 | 4.3e-05 | 0.001099 | 0.32388 | 1.373 | 2.5474 | 100.62 | 21.16 | 25.528 | 105.2695 | 88.066 | 96.9034 | 0.006379 | 0.95404 | 0.022925 | 24.0296 | 0.28845 | 0.15376 | 0.018347 | 1.2576 | 1.2576 | 0.00505
2019-03-28 15:43 | 1553802188.2824 | 287 | 0.20579 | 0.003521 | 0.00012 | 98.5575 | 95.341 | 114.3434 | 21.2612 | 0.00347 | 0.003094 | 158.1671 | 0.25782 | 8.4005 | 0.002706 | 0.004407 | 0.004324 | 4.6e-05 | 0.001097 | 0.32493 | 1.3204 | 2.5473 | 100.62 | 21.16 | 25.518 | 105.3607 | 86.9189 | 96.2674 | 0.006377 | 0.95568 | 0.022977 | 24.0427 | 0.29679 | 0.15338 | 0.018303 | 1.2574 | 1.2574 | 0.005038
2019-03-28 15:43 | 1553802189.2814 | 288 | 0.20581 | 0.003492 | 0.000119 | 98.6103 | 95.711 | 114.3732 | 21.2612 | 0.003439 | 0.003149 | 158.1188 | 0.259 | 8.4031 | 0.002694 | 0.004751 | 0.004659 | 4.5e-05 | 0.001099 | 0.33339 | 1.3267 | 2.5475 | 100.62 | 21.16 | 25.501 | 104.6143 | 86.6309 | 97.6033 | 0.006373 | 0.95224 | 0.022939 | 24.0893 | 0.26693 | 0.15408 | 0.018259 | 1.2523 | 1.2523 | 0.005081
2019-03-28 15:43 | 1553802190.2807 | 289 | 0.20584 | 0.003465 | 0.000118 | 98.733 | 95.9953 | 114.5479 | 21.2612 | 0.003453 | 0.003163 | 158.0569 | 0.25801 | 8.4036 | 0.002748 | 0.003799 | 0.00371 | 4.6e-05 | 0.001101 | 0.34164 | 1.3114 | 2.5476 | 100.62 | 21.16 | 25.493 | 104.943 | 87.9237 | 97.9246 | 0.006371 | 0.95141 | 0.022946 | 24.1181 | 0.25152 | 0.15349 | 0.018113 | 1.2428 | 1.2428 | 0.005101
2019-03-28 15:43 | 1553802191.2806 | 290 | 0.20591 | 0.003427 | 0.000117 | 98.6704 | 96.1974 | 115.0234 | 21.2612 | 0.003425 | 0.003163 | 158.0323 | 0.25839 | 8.4033 | 0.002704 | 0.004029 | 0.003919 | 4.7e-05 | 0.0011 | 0.30254 | 1.3304 | 2.5474 | 100.62 | 21.16 | 25.484 | 104.2629 | 86.8573 | 97.9372 | 0.006368 | 0.95138 | 0.022994 | 24.1685 | 0.22703 | 0.15372 | 0.017756 | 1.2208 | 1.2208 | 0.005201
2019-03-28 15:43 | 1553802192.2807 | 291 | 0.20596 | 0.0034 | 0.000116 | 98.7866 | 96.2542 | 115.2356 | 21.2612 | 0.003432 | 0.003104 | 157.9653 | 0.2579 | 8.4032 | 0.002694 | 0.002986 | 0.002948 | 4.9e-05 | 0.001099 | 0.32516 | 1.3536 | 2.5476 | 100.62 | 21.17 | 25.469 | 104.4461 | 86.6191 | 96.5103 | 0.006369 | 0.95505 | 0.023059 | 24.1443 | 0.23228 | 0.15343 | 0.017476 | 1.2049 | 1.2049 | 0.005259
2019-03-28 15:43 | 1553802193.2816 | 292 | 0.20599 | 0.003367 | 0.000115 | 98.8134 | 96.4804 | 115.1164 | 21.2612 | 0.003417 | 0.003106 | 157.9302 | 0.25897 | 8.4018 | 0.002678 | 0.004412 | 0.004381 | 4.9e-05 | 0.001099 | 0.32029 | 1.3197 | 2.5474 | 100.62 | 21.17 | 25.465 | 104.0835 | 86.2435 | 96.5682 | 0.006368 | 0.9549 | 0.023045 | 24.1337 | 0.21371 | 0.15406 | 0.017302 | 1.1922 | 1.1922 | 0.005337
2019-03-28 15:43 | 1553802194.282 | 293 | 0.20606 | 0.003338 | 0.000114 | 98.8557 | 96.8528 | 115.1762 | 21.2612 | 0.00336 | 0.003058 | 157.8553 | 0.25871 | 8.4038 | 0.002693 | 0.00464 | 0.004499 | 4.7e-05 | 0.001098 | 0.33019 | 1.2908 | 2.5473 | 100.62 | 21.16 | 25.462 | 102.6869 | 86.5965 | 95.4028 | 0.006363 | 0.95792 | 0.023088 | 24.1019 | 0.18309 | 0.1539 | 0.016912 | 1.1674 | 1.1674 | 0.005445
2019-03-28 15:43 | 1553802195.2821 | 294 | 0.20615 | 0.003302 | 0.000112 | 98.9185 | 97.0552 | 115.833 | 21.2612 | 0.003453 | 0.00314 | 157.8093 | 0.25875 | 8.3989 | 0.002653 | 0.004014 | 0.003926 | 4.7e-05 | 0.001099 | 0.31636 | 1.296 | 2.5473 | 100.62 | 21.17 | 25.459 | 104.9548 | 85.6363 | 97.3782 | 0.006366 | 0.95282 | 0.023092 | 24.2352 | 0.17009 | 0.15393 | 0.016462 | 1.1366 | 1.1366 | 0.005593
2019-03-28 15:43 | 1553802196.2807 | 295 | 0.20616 | 0.003256 | 0.00011 | 98.993 | 97.2734 | 115.8541 | 21.2612 | 0.003399 | 0.00308 | 157.7721 | 0.25728 | 8.4026 | 0.002592 | 0.004273 | 0.00409 | 4.7e-05 | 0.001097 | 0.32228 | 1.2911 | 2.5474 | 100.62 | 21.17 | 25.46 | 103.6377 | 84.1653 | 95.9297 | 0.006366 | 0.95656 | 0.023139 | 24.19 | 0.15674 | 0.15305 | 0.01643 | 1.1368 | 1.1368 | 0.00556
2019-03-28 15:43 | 1553802197.2811 | 296 | 0.20622 | 0.003223 | 0.000108 | 98.9578 | 97.575 | 116.013 | 21.2612 | 0.003417 | 0.003166 | 157.7078 | 0.25788 | 8.4041 | 0.002607 | 0.003715 | 0.003696 | 4.6e-05 | 0.001098 | 0.32477 | 1.2824 | 2.5473 | 100.62 | 21.17 | 25.462 | 104.0692 | 84.5237 | 98.0022 | 0.006367 | 0.95122 | 0.02309 | 24.2744 | 0.12587 | 0.15341 | 0.016145 | 1.1147 | 1.1147 | 0.005684
2019-03-28 15:43 | 1553802198.2798 | 297 | 0.20622 | 0.003183 | 0.000106 | 98.9671 | 97.5088 | 115.997 | 21.2612 | 0.003412 | 0.003094 | 157.6608 | 0.25874 | 8.4043 | 0.002648 | 0.004102 | 0.003928 | 4.9e-05 | 0.001098 | 0.32493 | 1.3489 | 2.5476 | 100.62 | 21.16 | 25.462 | 103.9476 | 85.5153 | 96.2783 | 0.006363 | 0.95565 | 0.023142 | 24.2163 | 0.13278 | 0.15392 | 0.016197 | 1.1209 | 1.1209 | 0.005671
2019-03-28 15:43 | 1553802199.2798 | 298 | 0.20627 | 0.003151 | 0.000104 | 99.0347 | 97.6022 | 116.2357 | 21.2612 | 0.003341 | 0.003109 | 157.6086 | 0.25874 | 8.4008 | 0.002605 | 0.003757 | 0.003647 | 4.8e-05 | 0.001095 | 0.31509 | 1.357 | 2.5475 | 100.62 | 21.17 | 25.463 | 102.2399 | 84.4938 | 96.6175 | 0.006367 | 0.95478 | 0.023156 | 24.2523 | 0.13032 | 0.15392 | 0.015942 | 1.104 | 1.104 | 0.005758
2019-03-28 15:43 | 1553802200.2802 | 299 | 0.20635 | 0.00312 | 0.000102 | 99.0597 | 97.8723 | 116.1515 | 21.2612 | 0.003425 | 0.003068 | 157.5805 | 0.25849 | 8.401 | 0.002572 | 0.004046 | 0.003911 | 4.8e-05 | 0.001094 | 0.32637 | 1.2884 | 2.5474 | 100.62 | 21.16 | 25.462 | 104.2706 | 83.6926 | 95.6282 | 0.006363 | 0.95734 | 0.023178 | 24.2111 | 0.10787 | 0.15378 | 0.0155 | 1.0744 | 1.0744 | 0.005911
2019-03-28 15:43 | 1553802201.2799 | 300 | 0.20645 | 0.003081 | 0.000101 | 99.1372 | 97.9331 | 115.8902 | 21.2612 | 0.003426 | 0.003129 | 157.5099 | 0.25818 | 8.4021 | 0.002614 | 0.004678 | 0.004529 | 4.8e-05 | 0.001096 | 0.33428 | 1.3126 | 2.5475 | 100.62 | 21.16 | 25.461 | 104.2864 | 84.7055 | 97.1009 | 0.006363 | 0.95353 | 0.023106 | 24.2321 | 0.1093 | 0.15359 | 0.01493 | 1.0316 | 1.0316 | 0.006149
2019-03-28 15:43 | 1553802202.2799 | 301 | 0.20646 | 0.00305 | 9.9e-05 | 99.1517 | 97.9664 | 115.9013 | 21.2612 | 0.003365 | 0.003049 | 157.4846 | 0.25851 | 8.4019 | 0.002606 | 0.004384 | 0.00434 | 5e-05 | 0.0011 | 0.32012 | 1.3326 | 2.5476 | 100.62 | 21.16 | 25.455 | 102.8222 | 84.5174 | 95.1835 | 0.006361 | 0.95849 | 0.023167 | 24.1704 | 0.10758 | 0.15379 | 0.014892 | 1.0318 | 1.0318 | 0.006156
2019-03-28 15:43 | 1553802203.2813 | 302 | 0.20647 | 0.00302 | 9.8e-05 | 99.1541 | 98.1801 | 116.0109 | 21.2612 | 0.003357 | 0.003092 | 157.4253 | 0.25889 | 8.4057 | 0.002662 | 0.004608 | 0.004496 | 5.1e-05 | 0.001101 | 0.32604 | 1.3337 | 2.5476 | 100.62 | 21.16 | 25.452 | 102.6273 | 85.8682 | 96.2235 | 0.00636 | 0.95579 | 0.023145 | 24.216 | 0.088302 | 0.15401 | 0.014896 | 1.0312 | 1.0312 | 0.006168
2019-03-28 15:43 | 1553802204.281 | 303 | 0.20656 | 0.002983 | 9.6e-05 | 99.2107 | 98.0955 | 115.9961 | 21.2612 | 0.003391 | 0.003023 | 157.4102 | 0.25799 | 8.4018 | 0.002712 | 0.004025 | 0.004026 | 5e-05 | 0.001098 | 0.31238 | 1.3052 | 2.5476 | 100.62 | 21.16 | 25.448 | 103.4556 | 87.0578 | 94.5511 | 0.006359 | 0.96014 | 0.023197 | 24.1595 | 0.10112 | 0.15348 | 0.014409 | 0.99962 | 0.99962 | 0.006341
2019-03-28 15:43 | 1553802205.28 | 304 | 0.20656 | 0.002951 | 9.5e-05 | 99.2203 | 98.157 | 115.7179 | 21.2612 | 0.003364 | 0.003041 | 157.3623 | 0.25816 | 8.403 | 0.002691 | 0.004022 | 0.003891 | 5e-05 | 0.001102 | 0.33069 | 1.3165 | 2.5472 | 100.62 | 21.16 | 25.438 | 102.7852 | 86.5628 | 94.9902 | 0.006357 | 0.959 | 0.023155 | 24.145 | 0.09637 | 0.15358 | 0.01446 | 1.0015 | 1.0015 | 0.006333
2019-03-28 15:43 | 1553802206.281 | 305 | 0.2067 | 0.002915 | 9.4e-05 | 99.2557 | 98.2178 | 116.1778 | 21.2612 | 0.003358 | 0.003021 | 157.305 | 0.2581 | 8.4046 | 0.002661 | 0.00445 | 0.004274 | 5e-05 | 0.0011 | 0.32516 | 1.3448 | 2.5472 | 100.62 | 21.16 | 25.433 | 102.6438 | 85.8314 | 94.504 | 0.006356 | 0.96027 | 0.023216 | 24.1769 | 0.094025 | 0.15354 | 0.013642 | 0.94716 | 0.94716 | 0.006695
2019-03-28 15:43 | 1553802207.2811 | 306 | 0.20649 | 0.002876 | 9.4e-05 | 99.2334 | 98.3843 | 116.247 | 21.2612 | 0.003344 | 0.003013 | 157.2659 | 0.25827 | 8.4005 | 0.002696 | 0.004148 | 0.004111 | 5e-05 | 0.001098 | 0.34811 | 1.3244 | 2.5473 | 100.62 | 21.16 | 25.425 | 102.3162 | 86.6668 | 94.3154 | 0.006354 | 0.96076 | 0.023229 | 24.1779 | 0.076865 | 0.15364 | 0.014927 | 1.0374 | 1.0374 | 0.006117
2019-03-28 15:43 | 1553802208.2812 | 307 | 0.20677 | 0.002839 | 9.2e-05 | 99.2137 | 98.1288 | 116.1113 | 21.2612 | 0.00329 | 0.003046 | 157.2429 | 0.25797 | 8.4026 | 0.002763 | 0.003656 | 0.00365 | 5.1e-05 | 0.001096 | 0.33428 | 1.29 | 2.5474 | 100.62 | 21.16 | 25.42 | 100.9916 | 88.2869 | 95.1014 | 0.006352 | 0.95871 | 0.023191 | 24.1896 | 0.098349 | 0.15346 | 0.013328 | 0.92436 | 0.92436 | 0.006857
2019-03-28 15:43 | 1553802209.2801 | 308 | 0.20674 | 0.0028 | 9.2e-05 | 99.262 | 98.3068 | 116.067 | 21.2612 | 0.003327 | 0.003021 | 157.2027 | 0.25839 | 8.4011 | 0.00267 | 0.004805 | 0.004624 | 5.2e-05 | 0.001098 | 0.31636 | 1.2991 | 2.5474 | 100.62 | 21.16 | 25.417 | 101.8952 | 86.0616 | 94.5088 | 0.006352 | 0.96025 | 0.023205 | 24.1655 | 0.086486 | 0.15372 | 0.013536 | 0.93944 | 0.93944 | 0.006758
2019-03-28 15:43 | 1553802210.2812 | 309 | 0.20702 | 0.002761 | 9e-05 | 99.2849 | 98.3938 | 116.6965 | 21.2612 | 0.003388 | 0.00307 | 157.1476 | 0.25807 | 8.403 | 0.002681 | 0.005038 | 0.004971 | 5.2e-05 | 0.001097 | 0.3115 | 1.2926 | 2.5474 | 100.62 | 21.16 | 25.411 | 103.3785 | 86.3232 | 95.6808 | 0.00635 | 0.9572 | 0.023231 | 24.2696 | 0.080643 | 0.15353 | 0.011932 | 0.82862 | 0.82862 | 0.007652
2019-03-28 15:43 | 1553802211.2812 | 310 | 0.20702 | 0.002726 | 8.9e-05 | 99.2152 | 98.3927 | 116.8594 | 21.2612 | 0.003335 | 0.00302 | 157.1272 | 0.25855 | 8.4015 | 0.002687 | 0.004106 | 0.004031 | 5.2e-05 | 0.0011 | 0.30973 | 1.3017 | 2.5475 | 100.62 | 21.16 | 25.41 | 102.1027 | 86.4605 | 94.4705 | 0.00635 | 0.96035 | 0.023285 | 24.2466 | 0.074455 | 0.15381 | 0.011959 | 0.83253 | 0.83253 | 0.00763
2019-03-28 15:44 | 1553802212.2817 | 311 | 0.20684 | 0.002695 | 8.8e-05 | 99.284 | 98.656 | 116.9237 | 21.2612 | 0.003339 | 0.003029 | 157.093 | 0.25791 | 8.4049 | 0.002662 | 0.003734 | 0.003642 | 5.3e-05 | 0.001101 | 0.32012 | 1.3271 | 2.5475 | 100.62 | 21.16 | 25.41 | 102.185 | 85.8509 | 94.6881 | 0.00635 | 0.95978 | 0.023285 | 24.2605 | 0.056753 | 0.15343 | 0.013074 | 0.91054 | 0.91054 | 0.006959
2019-03-28 15:44 | 1553802213.2802 | 312 | 0.20695 | 0.002661 | 8.7e-05 | 99.3477 | 98.7518 | 117.1829 | 21.2612 | 0.003311 | 0.003041 | 157.0613 | 0.25852 | 8.4039 | 0.002637 | 0.004275 | 0.00412 | 5.2e-05 | 0.001097 | 0.32405 | 1.2989 | 2.5475 | 100.62 | 21.16 | 25.417 | 101.5082 | 85.2673 | 94.994 | 0.006352 | 0.95899 | 0.023301 | 24.2974 | 0.053812 | 0.15379 | 0.012444 | 0.86719 | 0.86719 | 0.007324
2019-03-28 15:44 | 1553802214.2807 | 313 | 0.20696 | 0.002624 | 8.7e-05 | 99.2694 | 98.528 | 117.5279 | 21.2612 | 0.003333 | 0.003042 | 157.0172 | 0.25757 | 8.3986 | 0.002606 | 0.004284 | 0.004141 | 4.9e-05 | 0.001092 | 0.31653 | 1.2912 | 2.5474 | 100.62 | 21.16 | 25.423 | 102.0429 | 84.5196 | 95.0223 | 0.006353 | 0.95891 | 0.023334 | 24.3341 | 0.067058 | 0.15323 | 0.012455 | 0.86918 | 0.86918 | 0.007281
2019-03-28 15:44 | 1553802215.2793 | 314 | 0.20702 | 0.002592 | 8.6e-05 | 99.3745 | 98.5952 | 117.4044 | 21.2612 | 0.003353 | 0.002966 | 156.9782 | 0.25827 | 8.4027 | 0.00262 | 0.003907 | 0.003839 | 4.9e-05 | 0.001095 | 0.33931 | 1.294 | 2.5476 | 100.62 | 21.16 | 25.437 | 102.5377 | 84.8519 | 93.1739 | 0.006357 | 0.96375 | 0.023381 | 24.2602 | 0.070422 | 0.15364 | 0.012157 | 0.85004 | 0.85004 | 0.007465
2019-03-28 15:44 | 1553802216.2803 | 315 | 0.20696 | 0.002559 | 8.5e-05 | 99.4255 | 98.3126 | 117.1742 | 21.2612 | 0.003338 | 0.002982 | 156.9605 | 0.25751 | 8.4026 | 0.002617 | 0.003903 | 0.00388 | 5.2e-05 | 0.001086 | 0.32892 | 1.3116 | 2.5478 | 100.62 | 21.16 | 25.446 | 102.1755 | 84.7771 | 93.5532 | 0.006359 | 0.96276 | 0.023346 | 24.2489 | 0.10068 | 0.15319 | 0.012537 | 0.87544 | 0.87544 | 0.007227
2019-03-28 15:44 | 1553802217.2794 | 316 | 0.20708 | 0.002524 | 8.4e-05 | 99.3319 | 98.5575 | 117.0566 | 21.2612 | 0.003281 | 0.00297 | 156.9232 | 0.25844 | 8.405 | 0.002634 | 0.004328 | 0.004182 | 5.2e-05 | 0.001076 | 0.31222 | 1.3085 | 2.5472 | 100.62 | 21.16 | 25.455 | 100.7945 | 85.194 | 93.2723 | 0.006361 | 0.96349 | 0.023343 | 24.2275 | 0.070009 | 0.15375 | 0.011874 | 0.82891 | 0.82891 | 0.00766
2019-03-28 15:44 | 1553802218.2794 | 317 | 0.2071 | 0.002493 | 8.4e-05 | 99.4436 | 98.8266 | 117.1821 | 21.2612 | 0.00334 | 0.002934 | 156.8689 | 0.25836 | 8.4008 | 0.002631 | 0.004459 | 0.004348 | 5.1e-05 | 0.001072 | 0.29142 | 1.2707 | 2.5472 | 100.62 | 21.16 | 25.467 | 102.2203 | 85.1143 | 92.4174 | 0.006364 | 0.96575 | 0.023383 | 24.2122 | 0.055676 | 0.1537 | 0.011802 | 0.82531 | 0.82531 | 0.007691
2019-03-28 15:44 | 1553802219.2798 | 318 | 0.20714 | 0.002468 | 8.3e-05 | 99.3909 | 98.8551 | 116.9216 | 21.2612 | 0.003262 | 0.002943 | 156.8354 | 0.25793 | 8.4031 | 0.002587 | 0.003899 | 0.00385 | 5.2e-05 | 0.001078 | 0.31725 | 1.318 | 2.5475 | 100.62 | 21.16 | 25.474 | 100.3322 | 84.0578 | 92.6293 | 0.006366 | 0.96519 | 0.02335 | 24.1923 | 0.04835 | 0.15344 | 0.011587 | 0.8091 | 0.8091 | 0.007833
2019-03-28 15:44 | 1553802220.2798 | 319 | 0.2071 | 0.002439 | 8.3e-05 | 99.4836 | 98.7834 | 116.6404 | 21.2612 | 0.003322 | 0.002963 | 156.8327 | 0.25794 | 8.4047 | 0.002647 | 0.003617 | 0.003539 | 5.3e-05 | 0.00108 | 0.3413 | 1.3298 | 2.5473 | 100.62 | 21.15 | 25.486 | 101.7899 | 85.4896 | 93.116 | 0.006365 | 0.9639 | 0.023306 | 24.1792 | 0.063171 | 0.15345 | 0.011825 | 0.82425 | 0.82425 | 0.007689
2019-03-28 15:44 | 1553802221.28 | 320 | 0.20716 | 0.002411 | 8.3e-05 | 99.5092 | 98.8952 | 116.6977 | 21.2612 | 0.003239 | 0.002981 | 156.7974 | 0.25775 | 8.4023 | 0.002563 | 0.004396 | 0.00426 | 5.1e-05 | 0.001078 | 0.31509 | 1.3515 | 2.5475 | 100.62 | 21.15 | 25.491 | 99.7736 | 83.4875 | 93.5431 | 0.006366 | 0.96278 | 0.023299 | 24.1993 | 0.055359 | 0.15333 | 0.011486 | 0.80025 | 0.80025 | 0.007913
2019-03-28 15:44 | 1553802222.2805 | 321 | 0.20715 | 0.002383 | 8.2e-05 | 99.4136 | 98.8334 | 116.7435 | 21.2612 | 0.003231 | 0.002947 | 156.7463 | 0.25815 | 8.4044 | 0.002648 | 0.003909 | 0.003852 | 5.2e-05 | 0.001078 | 0.32316 | 1.2996 | 2.5475 | 100.62 | 21.15 | 25.496 | 99.5869 | 85.5209 | 92.71 | 0.006368 | 0.96497 | 0.02333 | 24.1765 | 0.052354 | 0.15357 | 0.011608 | 0.80995 | 0.80995 | 0.007831
2019-03-28 15:44 | 1553802223.2809 | 322 | 0.20725 | 0.002363 | 8e-05 | 99.4401 | 99.1259 | 116.7798 | 21.2612 | 0.003239 | 0.002997 | 156.7185 | 0.25791 | 8.4008 | 0.002604 | 0.004266 | 0.004114 | 5.3e-05 | 0.001077 | 0.31725 | 1.3175 | 2.5473 | 100.62 | 21.15 | 25.504 | 99.769 | 84.4628 | 93.9318 | 0.00637 | 0.96176 | 0.023294 | 24.2206 | 0.028306 | 0.15343 | 0.011066 | 0.77083 | 0.77083 | 0.008221
2019-03-28 15:44 | 1553802224.2807 | 323 | 0.20727 | 0.002342 | 8e-05 | 99.4002 | 98.8561 | 116.7369 | 21.2612 | 0.003241 | 0.002951 | 156.6884 | 0.2578 | 8.4039 | 0.002634 | 0.003882 | 0.003729 | 5e-05 | 0.001077 | 0.31885 | 1.3026 | 2.5474 | 100.62 | 21.15 | 25.508 | 99.8302 | 85.1735 | 92.8156 | 0.006371 | 0.9647 | 0.023326 | 24.1793 | 0.049089 | 0.15336 | 0.010972 | 0.76533 | 0.76533 | 0.008276
2019-03-28 15:44 | 1553802225.281 | 324 | 0.20728 | 0.002322 | 7.9e-05 | 99.3722 | 98.7861 | 117.0481 | 21.2612 | 0.003261 | 0.002941 | 156.6515 | 0.2581 | 8.4016 | 0.002601 | 0.004168 | 0.004069 | 5e-05 | 0.001072 | 0.3115 | 1.3342 | 2.5477 | 100.62 | 21.15 | 25.509 | 100.3049 | 84.3871 | 92.5849 | 0.006371 | 0.9653 | 0.023364 | 24.2039 | 0.052903 | 0.15354 | 0.010928 | 0.76351 | 0.76351 | 0.008305
2019-03-28 15:44 | 1553802226.2807 | 325 | 0.20734 | 0.0023 | 7.9e-05 | 99.3418 | 99.0583 | 117.1643 | 21.2612 | 0.003233 | 0.002901 | 156.5861 | 0.25802 | 8.4025 | 0.002632 | 0.003795 | 0.003722 | 5e-05 | 0.001076 | 0.33373 | 1.3003 | 2.5472 | 100.62 | 21.14 | 25.511 | 99.6215 | 85.1269 | 91.6129 | 0.006367 | 0.96788 | 0.023407 | 24.1837 | 0.025563 | 0.15349 | 0.010551 | 0.73843 | 0.73843 | 0.008585
2019-03-28 15:44 | 1553802227.2797 | 326 | 0.20736 | 0.002273 | 7.8e-05 | 99.3719 | 98.9013 | 117.3073 | 21.2612 | 0.00324 | 0.002949 | 156.5771 | 0.25746 | 8.4045 | 0.002559 | 0.004357 | 0.004265 | 4.8e-05 | 0.001075 | 0.3282 | 1.3027 | 2.5473 | 100.62 | 21.14 | 25.511 | 99.7886 | 83.3798 | 92.7662 | 0.006367 | 0.96483 | 0.023384 | 24.2367 | 0.042455 | 0.15317 | 0.010518 | 0.73543 | 0.73543 | 0.008601
2019-03-28 15:44 | 1553802228.2797 | 327 | 0.2074 | 0.002255 | 7.7e-05 | 99.2357 | 98.8324 | 117.5143 | 21.2612 | 0.003206 | 0.002913 | 156.5589 | 0.25803 | 8.4022 | 0.002544 | 0.004068 | 0.004022 | 4.9e-05 | 0.001075 | 0.32012 | 1.3642 | 2.5475 | 100.62 | 21.14 | 25.508 | 98.9611 | 83.0201 | 91.8951 | 0.006367 | 0.96713 | 0.023433 | 24.2291 | 0.036434 | 0.1535 | 0.010293 | 0.72115 | 0.72115 | 0.008791
2019-03-28 15:44 | 1553802229.2797 | 328 | 0.20741 | 0.002232 | 7.8e-05 | 99.278 | 98.7121 | 117.7408 | 21.2612 | 0.003155 | 0.002907 | 156.5435 | 0.25775 | 8.4043 | 0.002606 | 0.004117 | 0.004041 | 4.9e-05 | 0.001074 | 0.30663 | 1.344 | 2.5474 | 100.62 | 21.14 | 25.507 | 97.7427 | 84.5071 | 91.7449 | 0.006366 | 0.96753 | 0.02346 | 24.2475 | 0.051136 | 0.15333 | 0.01025 | 0.71892 | 0.71892 | 0.008809
2019-03-28 15:44 | 1553802230.2798 | 329 | 0.20743 | 0.002206 | 7.7e-05 | 99.2643 | 98.6615 | 118.0159 | 21.2612 | 0.003147 | 0.002916 | 156.4744 | 0.25773 | 8.4008 | 0.002531 | 0.00387 | 0.003757 | 4.9e-05 | 0.001075 | 0.3115 | 1.3332 | 2.5475 | 100.62 | 21.14 | 25.507 | 97.5524 | 82.7097 | 91.9724 | 0.006366 | 0.96692 | 0.02348 | 24.2834 | 0.054489 | 0.15332 | 0.010158 | 0.71311 | 0.71311 | 0.00888
2019-03-28 15:44 | 1553802231.2788 | 330 | 0.20747 | 0.002185 | 7.8e-05 | 99.1392 | 98.5086 | 118.029 | 21.2612 | 0.003168 | 0.002888 | 156.4351 | 0.25835 | 8.406 | 0.00254 | 0.003847 | 0.003769 | 4.8e-05 | 0.001083 | 0.32156 | 1.3161 | 2.5474 | 100.62 | 21.14 | 25.503 | 98.0602 | 82.9297 | 91.2984 | 0.006365 | 0.96871 | 0.023503 | 24.2623 | 0.05708 | 0.15369 | 0.009931 | 0.69769 | 0.69769 | 0.009098
2019-03-28 15:44 | 1553802232.2802 | 331 | 0.20744 | 0.002162 | 7.8e-05 | 99.1931 | 98.5419 | 118.2128 | 21.2612 | 0.003156 | 0.002888 | 156.4201 | 0.25827 | 8.4047 | 0.002547 | 0.004562 | 0.004496 | 4.5e-05 | 0.001083 | 0.32565 | 1.3 | 2.5474 | 100.62 | 21.13 | 25.503 | 97.7591 | 83.0903 | 91.2907 | 0.006361 | 0.96873 | 0.023522 | 24.2809 | 0.05892 | 0.15365 | 0.010181 | 0.71596 | 0.71596 | 0.008863
2019-03-28 15:44 | 1553802233.2791 | 332 | 0.20748 | 0.002143 | 7.8e-05 | 99.1446 | 98.494 | 118.1295 | 21.2612 | 0.003196 | 0.002882 | 156.3838 | 0.25756 | 8.4049 | 0.002542 | 0.004421 | 0.004332 | 4.4e-05 | 0.001086 | 0.32532 | 1.3034 | 2.5474 | 100.62 | 21.13 | 25.505 | 98.7228 | 82.9734 | 91.1594 | 0.006362 | 0.96908 | 0.023518 | 24.268 | 0.05889 | 0.15322 | 0.009946 | 0.69925 | 0.69925 | 0.00905
2019-03-28 15:44 | 1553802234.2793 | 333 | 0.20747 | 0.002123 | 7.7e-05 | 99.155 | 98.6795 | 118.2746 | 21.2612 | 0.003191 | 0.002878 | 156.3591 | 0.25763 | 8.404 | 0.002507 | 0.004031 | 0.00391 | 4.8e-05 | 0.001092 | 0.33179 | 1.3502 | 2.5475 | 100.62 | 21.13 | 25.506 | 98.6113 | 82.1294 | 91.0671 | 0.006362 | 0.96933 | 0.023535 | 24.2798 | 0.043001 | 0.15326 | 0.010006 | 0.704 | 0.704 | 0.008991
2019-03-28 15:44 | 1553802235.279 | 334 | 0.20751 | 0.002107 | 7.7e-05 | 99.1157 | 98.6632 | 118.2915 | 21.2612 | 0.003216 | 0.002877 | 156.325 | 0.25769 | 8.4035 | 0.002493 | 0.00405 | 0.003895 | 4.8e-05 | 0.001095 | 0.33196 | 1.3598 | 2.5476 | 100.62 | 21.13 | 25.507 | 99.219 | 81.7933 | 91.0427 | 0.006362 | 0.96939 | 0.023538 | 24.2807 | 0.040932 | 0.1533 | 0.009812 | 0.69042 | 0.69042 | 0.00917
2019-03-28 15:44 | 1553802236.2789 | 335 | 0.20757 | 0.002092 | 7.7e-05 | 99.1622 | 98.8014 | 118.3161 | 21.2612 | 0.003148 | 0.002826 | 156.3415 | 0.2576 | 8.4029 | 0.002506 | 0.004366 | 0.004309 | 5e-05 | 0.00109 | 0.31924 | 1.3731 | 2.5474 | 100.62 | 21.13 | 25.51 | 97.5828 | 82.1092 | 89.8056 | 0.006363 | 0.9727 | 0.02358 | 24.242 | 0.032598 | 0.15324 | 0.009463 | 0.66692 | 0.66692 | 0.00949
2019-03-28 15:44 | 1553802237.2794 | 336 | 0.2076 | 0.002074 | 7.6e-05 | 99.0535 | 98.5847 | 118.2933 | 21.2612 | 0.003165 | 0.002825 | 156.2832 | 0.25791 | 8.4046 | 0.002517 | 0.003914 | 0.003828 | 4.7e-05 | 0.001091 | 0.33268 | 1.3654 | 2.5476 | 100.62 | 21.12 | 25.512 | 97.9788 | 82.3594 | 89.7752 | 0.00636 | 0.97278 | 0.023579 | 24.2386 | 0.04243 | 0.15343 | 0.009297 | 0.65513 | 0.65513 | 0.009673
2019-03-28 15:44 | 1553802238.2801 | 337 | 0.20759 | 0.002054 | 7.6e-05 | 99.0627 | 98.6251 | 118.4788 | 21.2612 | 0.00311 | 0.002822 | 156.2454 | 0.25797 | 8.4023 | 0.002612 | 0.003468 | 0.003423 | 4.3e-05 | 0.001094 | 0.33373 | 1.31 | 2.5473 | 100.62 | 21.12 | 25.515 | 96.652 | 84.6663 | 89.7073 | 0.006361 | 0.97296 | 0.023599 | 24.2554 | 0.039597 | 0.15347 | 0.009368 | 0.66074 | 0.66074 | 0.009593
2019-03-28 15:44 | 1553802239.2791 | 338 | 0.20762 | 0.002031 | 7.6e-05 | 98.9013 | 98.6166 | 118.7163 | 21.2612 | 0.003154 | 0.002859 | 156.1949 | 0.25802 | 8.4037 | 0.002577 | 0.003921 | 0.003805 | 4.8e-05 | 0.001092 | 0.28119 | 1.2794 | 2.5475 | 100.62 | 21.12 | 25.519 | 97.7277 | 83.8121 | 90.5973 | 0.006362 | 0.97058 | 0.023594 | 24.3094 | 0.025779 | 0.15349 | 0.009219 | 0.65003 | 0.65003 | 0.009752
2019-03-28 15:44 | 1553802240.2792 | 339 | 0.20758 | 0.002021 | 7.6e-05 | 98.9153 | 98.3418 | 118.6762 | 21.2612 | 0.003082 | 0.002846 | 156.1745 | 0.25842 | 8.4049 | 0.002483 | 0.004285 | 0.004263 | 4.9e-05 | 0.001094 | 0.30165 | 1.2918 | 2.5473 | 100.62 | 21.12 | 25.53 | 95.9855 | 81.5633 | 90.2894 | 0.006364 | 0.9714 | 0.0236 | 24.295 | 0.052007 | 0.15374 | 0.009515 | 0.67123 | 0.67123 | 0.009459
2019-03-28 15:44 | 1553802241.2791 | 340 | 0.2076 | 0.002001 | 7.5e-05 | 98.7961 | 98.1267 | 118.7827 | 21.2612 | 0.003138 | 0.002828 | 156.1461 | 0.25804 | 8.4036 | 0.002508 | 0.003984 | 0.00389 | 4.9e-05 | 0.001096 | 0.33229 | 1.3214 | 2.5474 | 100.62 | 21.12 | 25.536 | 97.3181 | 82.1426 | 89.865 | 0.006366 | 0.97254 | 0.023625 | 24.2917 | 0.060808 | 0.15351 | 0.009406 | 0.66418 | 0.66418 | 0.009545
2019-03-28 15:44 | 1553802242.2805 | 341 | 0.20767 | 0.001985 | 7.5e-05 | 98.786 | 98.2952 | 118.7489 | 21.2612 | 0.003128 | 0.002831 | 156.1005 | 0.25808 | 8.4008 | 0.002505 | 0.004197 | 0.004067 | 4.8e-05 | 0.001098 | 0.30718 | 1.2933 | 2.5475 | 100.62 | 21.12 | 25.55 | 97.0993 | 82.0818 | 89.9228 | 0.006369 | 0.97238 | 0.023619 | 24.2902 | 0.044545 | 0.15353 | 0.009005 | 0.63562 | 0.63562 | 0.009976
2019-03-28 15:44 | 1553802243.2792 | 342 | 0.20768 | 0.001969 | 7.5e-05 | 98.7148 | 98.2868 | 119.0588 | 21.2612 | 0.0031 | 0.002803 | 156.0776 | 0.25737 | 8.4029 | 0.002518 | 0.00458 | 0.004576 | 4.6e-05 | 0.001096 | 0.32117 | 1.3 | 2.5474 | 100.62 | 21.12 | 25.56 | 96.4193 | 82.3969 | 89.2454 | 0.006372 | 0.9742 | 0.023672 | 24.2992 | 0.03887 | 0.15311 | 0.008962 | 0.63398 | 0.63398 | 0.009974
2019-03-28 15:44 | 1553802244.2806 | 343 | 0.20769 | 0.001955 | 7.5e-05 | 98.7285 | 98.1665 | 119.0771 | 21.2612 | 0.003072 | 0.00281 | 156.0366 | 0.25789 | 8.401 | 0.002528 | 0.003639 | 0.003383 | 4.5e-05 | 0.001097 | 0.31332 | 1.2943 | 2.5473 | 100.62 | 21.12 | 25.569 | 95.7255 | 82.6433 | 89.4306 | 0.006374 | 0.9737 | 0.023668 | 24.3073 | 0.051062 | 0.15342 | 0.008901 | 0.62952 | 0.62952 | 0.010065
2019-03-28 15:44 | 1553802245.2783 | 344 | 0.20773 | 0.001935 | 7.5e-05 | 98.6874 | 98.1352 | 119.0283 | 21.2612 | 0.003082 | 0.002785 | 156.0344 | 0.25778 | 8.402 | 0.002538 | 0.004028 | 0.004022 | 4.8e-05 | 0.0011 | 0.32405 | 1.342 | 2.5475 | 100.62 | 21.11 | 25.586 | 95.9744 | 82.8664 | 88.8079 | 0.006374 | 0.97538 | 0.023684 | 24.2814 | 0.050186 | 0.15336 | 0.008698 | 0.61549 | 0.61549 | 0.01029
2019-03-28 15:44 | 1553802246.2794 | 345 | 0.20772 | 0.001925 | 7.4e-05 | 98.7225 | 98.0751 | 119.005 | 21.2612 | 0.003096 | 0.002789 | 155.9713 | 0.25776 | 8.4038 | 0.002544 | 0.004228 | 0.004105 | 5.2e-05 | 0.001107 | 0.33323 | 1.3497 | 2.5475 | 100.62 | 21.11 | 25.594 | 96.3041 | 83.0186 | 88.9057 | 0.006376 | 0.97511 | 0.023678 | 24.2823 | 0.058853 | 0.15334 | 0.008753 | 0.61929 | 0.61929 | 0.010226
2019-03-28 15:44 | 1553802247.2787 | 346 | 0.20772 | 0.001914 | 7.4e-05 | 98.6254 | 97.9875 | 119.4496 | 21.2612 | 0.003061 | 0.002813 | 155.9625 | 0.2569 | 8.4052 | 0.002464 | 0.004486 | 0.004382 | 5e-05 | 0.001106 | 0.30182 | 1.3242 | 2.5474 | 100.62 | 21.11 | 25.603 | 95.4757 | 81.1064 | 89.4814 | 0.006379 | 0.97357 | 0.023703 | 24.3469 | 0.058048 | 0.15283 | 0.008765 | 0.62082 | 0.62082 | 0.010167
2019-03-28 15:44 | 1553802248.2784 | 347 | 0.20773 | 0.001895 | 7.4e-05 | 98.6415 | 98.0873 | 119.1724 | 21.2612 | 0.00301 | 0.002794 | 155.9374 | 0.25776 | 8.4003 | 0.002451 | 0.00407 | 0.003954 | 5.1e-05 | 0.001108 | 0.29463 | 1.3286 | 2.5474 | 100.62 | 21.11 | 25.619 | 94.2428 | 80.7959 | 89.0407 | 0.006383 | 0.97475 | 0.02369 | 24.3039 | 0.050395 | 0.15334 | 0.008769 | 0.62076 | 0.62076 | 0.010202
2019-03-28 15:44 | 1553802249.2784 | 348 | 0.20775 | 0.001885 | 7.4e-05 | 98.5989 | 97.9667 | 118.9215 | 21.2612 | 0.003054 | 0.002809 | 155.9281 | 0.25769 | 8.4022 | 0.002451 | 0.003803 | 0.003648 | 4.4e-05 | 0.001106 | 0.335 | 1.3627 | 2.5475 | 100.62 | 21.11 | 25.628 | 95.2986 | 80.7874 | 89.4002 | 0.006385 | 0.97378 | 0.023654 | 24.2903 | 0.057536 | 0.1533 | 0.008605 | 0.60813 | 0.60813 | 0.010411
2019-03-28 15:44 | 1553802250.2785 | 349 | 0.2078 | 0.001874 | 7.3e-05 | 98.5572 | 98.1723 | 119.0118 | 21.2612 | 0.003065 | 0.002756 | 155.916 | 0.25698 | 8.4025 | 0.002436 | 0.004018 | 0.003984 | 4.9e-05 | 0.001105 | 0.31725 | 1.3604 | 2.5476 | 100.62 | 21.11 | 25.642 | 95.5742 | 80.4255 | 88.1144 | 0.006388 | 0.97725 | 0.023705 | 24.2564 | 0.035006 | 0.15288 | 0.008363 | 0.5922 | 0.5922 | 0.010661
2019-03-28 15:44 | 1553802251.2796 | 350 | 0.20778 | 0.001865 | 7.3e-05 | 98.461 | 97.8998 | 118.9994 | 21.2612 | 0.003053 | 0.002778 | 155.8784 | 0.25726 | 8.4019 | 0.002468 | 0.003569 | 0.003356 | 4.5e-05 | 0.001092 | 0.33771 | 1.2951 | 2.5474 | 100.62 | 21.11 | 25.649 | 95.2776 | 81.1999 | 88.6485 | 0.00639 | 0.97581 | 0.023686 | 24.2731 | 0.051129 | 0.15305 | 0.008481 | 0.60017 | 0.60017 | 0.010532
2019-03-28 15:44 | 1553802252.2795 | 351 | 0.20781 | 0.001852 | 7.3e-05 | 98.4709 | 97.9324 | 118.8351 | 21.2612 | 0.003028 | 0.002734 | 155.8887 | 0.25728 | 8.4023 | 0.00246 | 0.00416 | 0.00411 | 4.7e-05 | 0.001084 | 0.3246 | 1.2988 | 2.5474 | 100.62 | 21.11 | 25.664 | 94.6746 | 80.9932 | 87.5959 | 0.006394 | 0.97865 | 0.023704 | 24.221 | 0.049043 | 0.15305 | 0.008338 | 0.59048 | 0.59048 | 0.010705
2019-03-28 15:44 | 1553802253.2795 | 352 | 0.20785 | 0.001836 | 7.3e-05 | 98.4503 | 97.7245 | 118.7844 | 21.2612 | 0.003011 | 0.002706 | 155.84 | 0.2588 | 8.4025 | 0.002475 | 0.004207 | 0.004131 | 4.8e-05 | 0.001078 | 0.34108 | 1.3078 | 2.5475 | 100.62 | 21.11 | 25.67 | 94.2599 | 81.3628 | 86.9165 | 0.006395 | 0.9805 | 0.023721 | 24.193 | 0.066186 | 0.15396 | 0.008122 | 0.57544 | 0.57544 | 0.01105
2019-03-28 15:44 | 1553802254.2799 | 353 | 0.20784 | 0.001826 | 7.3e-05 | 98.4187 | 97.8003 | 118.8607 | 21.2612 | 0.002984 | 0.002707 | 155.8119 | 0.25757 | 8.4024 | 0.002329 | 0.003742 | 0.003629 | 4.8e-05 | 0.001076 | 0.31598 | 1.2963 | 2.5474 | 100.62 | 21.11 | 25.678 | 93.6243 | 77.8512 | 86.9367 | 0.006397 | 0.98045 | 0.023728 | 24.2015 | 0.056385 | 0.15323 | 0.00819 | 0.5805 | 0.5805 | 0.010902
2019-03-28 15:44 | 1553802255.2785 | 354 | 0.20782 | 0.001806 | 7.4e-05 | 98.3758 | 97.824 | 118.8909 | 21.2612 | 0.002989 | 0.002741 | 155.7534 | 0.25775 | 8.405 | 0.002401 | 0.003763 | 0.003729 | 4.9e-05 | 0.001076 | 0.33323 | 1.3222 | 2.5473 | 100.62 | 21.11 | 25.693 | 93.7449 | 79.5709 | 87.7535 | 0.006401 | 0.97823 | 0.023704 | 24.232 | 0.050312 | 0.15334 | 0.008309 | 0.5883 | 0.5883 | 0.010765
2019-03-28 15:44 | 1553802256.2796 | 355 | 0.20779 | 0.001798 | 7.4e-05 | 98.3898 | 97.803 | 118.8421 | 21.2612 | 0.003014 | 0.002706 | 155.7297 | 0.25852 | 8.4025 | 0.002409 | 0.003809 | 0.003637 | 4.8e-05 | 0.001078 | 0.32421 | 1.357 | 2.5473 | 100.62 | 21.11 | 25.701 | 94.3335 | 79.7637 | 86.9071 | 0.006403 | 0.98053 | 0.023727 | 24.1986 | 0.053511 | 0.15379 | 0.00852 | 0.60391 | 0.60391 | 0.010517
2019-03-28 15:44 | 1553802257.2797 | 356 | 0.20781 | 0.00179 | 7.3e-05 | 98.335 | 97.819 | 118.582 | 21.2612 | 0.002978 | 0.002754 | 155.6979 | 0.25795 | 8.4035 | 0.002399 | 0.004084 | 0.003978 | 4.9e-05 | 0.001082 | 0.32372 | 1.3221 | 2.5475 | 100.62 | 21.11 | 25.709 | 93.4594 | 79.5383 | 88.0684 | 0.006405 | 0.97737 | 0.023663 | 24.2111 | 0.047067 | 0.15345 | 0.008364 | 0.59125 | 0.59125 | 0.010719
2019-03-28 15:44 | 1553802258.279 | 357 | 0.20783 | 0.001779 | 7.4e-05 | 98.3288 | 97.7306 | 118.8776 | 21.2612 | 0.003002 | 0.002675 | 155.6605 | 0.25786 | 8.4026 | 0.002357 | 0.00418 | 0.004139 | 5e-05 | 0.001089 | 0.30862 | 1.3407 | 2.5473 | 100.62 | 21.11 | 25.721 | 94.0392 | 78.5201 | 86.175 | 0.006408 | 0.98252 | 0.023755 | 24.1776 | 0.054578 | 0.1534 | 0.008269 | 0.5867 | 0.5867 | 0.010798
2019-03-28 15:44 | 1553802259.2801 | 358 | 0.20787 | 0.001769 | 7.3e-05 | 98.2951 | 97.734 | 118.8717 | 21.2612 | 0.003009 | 0.002709 | 155.6211 | 0.25836 | 8.4024 | 0.002441 | 0.00394 | 0.003803 | 4.8e-05 | 0.001084 | 0.33019 | 1.3438 | 2.5475 | 100.62 | 21.11 | 25.725 | 94.2145 | 80.5324 | 86.9813 | 0.006409 | 0.98032 | 0.023728 | 24.2041 | 0.051205 | 0.1537 | 0.008076 | 0.57234 | 0.57234 | 0.011091
2019-03-28 15:44 | 1553802260.2798 | 359 | 0.20789 | 0.00176 | 7.4e-05 | 98.1861 | 97.662 | 118.8118 | 21.2612 | 0.002903 | 0.002684 | 155.5905 | 0.25762 | 8.4008 | 0.00242 | 0.004094 | 0.004021 | 4.9e-05 | 0.001083 | 0.33339 | 1.3177 | 2.5475 | 100.62 | 21.11 | 25.733 | 91.648 | 80.034 | 86.3903 | 0.006411 | 0.98194 | 0.023741 | 24.1782 | 0.047874 | 0.15326 | 0.007959 | 0.56428 | 0.56428 | 0.011217
2019-03-28 15:44 | 1553802261.2778 | 360 | 0.20788 | 0.001745 | 7.4e-05 | 98.1084 | 97.6246 | 119.1064 | 21.2612 | 0.002987 | 0.002729 | 155.5804 | 0.25854 | 8.4045 | 0.002447 | 0.004093 | 0.004022 | 5.1e-05 | 0.001083 | 0.30973 | 1.3308 | 2.5475 | 100.62 | 21.11 | 25.735 | 93.6832 | 80.6766 | 87.4717 | 0.006411 | 0.97899 | 0.023735 | 24.2445 | 0.044212 | 0.1538 | 0.00806 | 0.57131 | 0.57131 | 0.011119
2019-03-28 15:44 | 1553802262.2787 | 361 | 0.20796 | 0.00173 | 7.3e-05 | 98.036 | 97.3974 | 119.305 | 21.2612 | 0.00292 | 0.002705 | 155.5358 | 0.25868 | 8.4037 | 0.002457 | 0.003564 | 0.003421 | 4.8e-05 | 0.001083 | 0.32261 | 1.3323 | 2.5474 | 100.62 | 21.11 | 25.737 | 92.078 | 80.934 | 86.8845 | 0.006412 | 0.98059 | 0.023774 | 24.2449 | 0.058459 | 0.15389 | 0.007585 | 0.53838 | 0.53838 | 0.011805
2019-03-28 15:44 | 1553802263.2792 | 362 | 0.20795 | 0.001724 | 7.4e-05 | 98.0169 | 97.4415 | 119.2672 | 21.2612 | 0.002927 | 0.002716 | 155.5271 | 0.25789 | 8.4027 | 0.002493 | 0.003757 | 0.003665 | 5.2e-05 | 0.001089 | 0.32587 | 1.2958 | 2.5474 | 100.62 | 21.11 | 25.739 | 92.2306 | 81.8034 | 87.1555 | 0.006412 | 0.97985 | 0.023762 | 24.2502 | 0.052664 | 0.15342 | 0.007661 | 0.54344 | 0.54344 | 0.011659
2019-03-28 15:44 | 1553802264.2779 | 363 | 0.20794 | 0.001708 | 7.5e-05 | 97.8769 | 97.1755 | 119.2496 | 21.2612 | 0.002973 | 0.002745 | 155.4971 | 0.25781 | 8.402 | 0.002463 | 0.003565 | 0.003459 | 4.7e-05 | 0.001084 | 0.30558 | 1.3188 | 2.5475 | 100.62 | 21.11 | 25.739 | 93.3444 | 81.0626 | 87.8457 | 0.006412 | 0.97798 | 0.023737 | 24.2716 | 0.064331 | 0.15337 | 0.007743 | 0.54865 | 0.54865 | 0.011545
2019-03-28 15:44 | 1553802265.2779 | 364 | 0.20794 | 0.001698 | 7.6e-05 | 97.7885 | 97.1181 | 119.4376 | 21.2612 | 0.002993 | 0.002717 | 155.4765 | 0.25862 | 8.4023 | 0.002501 | 0.003784 | 0.003791 | 5.2e-05 | 0.00109 | 0.32211 | 1.3003 | 2.5474 | 100.62 | 21.11 | 25.735 | 93.824 | 81.9859 | 87.1921 | 0.006411 | 0.97975 | 0.023777 | 24.2688 | 0.06153 | 0.15385 | 0.007741 | 0.54939 | 0.54939 | 0.011566
2019-03-28 15:44 | 1553802266.2784 | 365 | 0.20798 | 0.001688 | 7.7e-05 | 97.7227 | 97.0199 | 119.7589 | 21.2612 | 0.002973 | 0.002673 | 155.4464 | 0.25787 | 8.4027 | 0.002467 | 0.004022 | 0.003842 | 5.2e-05 | 0.001088 | 0.34147 | 1.3006 | 2.5474 | 100.62 | 21.11 | 25.732 | 93.3506 | 81.1756 | 86.1155 | 0.006411 | 0.98269 | 0.023845 | 24.2651 | 0.064554 | 0.15341 | 0.007513 | 0.53456 | 0.53456 | 0.011852
2019-03-28 15:44 | 1553802267.279 | 366 | 0.20786 | 0.001676 | 7.7e-05 | 97.5547 | 96.8691 | 119.8514 | 21.2612 | 0.002934 | 0.002741 | 155.4117 | 0.25727 | 8.4034 | 0.00244 | 0.004183 | 0.004078 | 5.1e-05 | 0.001084 | 0.31061 | 1.3021 | 2.5475 | 100.62 | 21.1 | 25.728 | 92.4111 | 80.5106 | 87.7685 | 0.006406 | 0.97819 | 0.023799 | 24.3302 | 0.063082 | 0.15305 | 0.008223 | 0.58433 | 0.58433 | 0.010817
2019-03-28 15:44 | 1553802268.2783 | 367 | 0.20797 | 0.001666 | 7.6e-05 | 97.5305 | 96.6945 | 119.7864 | 21.2612 | 0.002911 | 0.002691 | 155.3938 | 0.25793 | 8.4014 | 0.002533 | 0.003733 | 0.003607 | 5.2e-05 | 0.001079 | 0.34075 | 1.3073 | 2.5474 | 100.62 | 21.1 | 25.724 | 91.8481 | 82.7667 | 86.5652 | 0.006405 | 0.98146 | 0.023833 | 24.283 | 0.077008 | 0.15344 | 0.007559 | 0.53764 | 0.53764 | 0.011787
2019-03-28 15:44 | 1553802269.279 | 368 | 0.20794 | 0.001661 | 7.6e-05 | 97.2649 | 96.4108 | 120.2167 | 21.2612 | 0.002971 | 0.002663 | 155.3763 | 0.25728 | 8.4037 | 0.002538 | 0.003139 | 0.003042 | 4.8e-05 | 0.001076 | 0.3366 | 1.3129 | 2.5474 | 100.62 | 21.1 | 25.723 | 93.2879 | 82.8838 | 85.8916 | 0.006404 | 0.9833 | 0.023898 | 24.3038 | 0.078888 | 0.15306 | 0.007752 | 0.553 | 0.553 | 0.011431
2019-03-28 15:44 | 1553802270.279 | 369 | 0.208 | 0.001651 | 7.6e-05 | 97.174 | 96.4125 | 120.0812 | 21.2612 | 0.002937 | 0.002622 | 155.3483 | 0.25711 | 8.4061 | 0.002499 | 0.003599 | 0.003494 | 5.2e-05 | 0.001073 | 0.32444 | 1.2886 | 2.5473 | 100.62 | 21.1 | 25.724 | 92.4839 | 81.9289 | 84.888 | 0.006405 | 0.98606 | 0.023918 | 24.2561 | 0.070372 | 0.15296 | 0.007409 | 0.52878 | 0.52878 | 0.011947
2019-03-28 15:44 | 1553802271.2785 | 370 | 0.20798 | 0.001639 | 7.7e-05 | 97.1311 | 96.4757 | 120.0636 | 21.2612 | 0.002922 | 0.0027 | 155.3151 | 0.25726 | 8.4026 | 0.002463 | 0.004087 | 0.004014 | 5.5e-05 | 0.001067 | 0.32278 | 1.3476 | 2.5473 | 100.62 | 21.1 | 25.73 | 92.1156 | 81.0751 | 86.7729 | 0.006406 | 0.98089 | 0.023853 | 24.3181 | 0.060563 | 0.15305 | 0.007543 | 0.53693 | 0.53693 | 0.011772
2019-03-28 15:45 | 1553802272.2782 | 371 | 0.208 | 0.001629 | 7.8e-05 | 96.9932 | 96.1251 | 120.3347 | 21.2612 | 0.002909 | 0.002711 | 155.3056 | 0.25739 | 8.404 | 0.002497 | 0.003883 | 0.003804 | 5.3e-05 | 0.00106 | 0.31222 | 1.3534 | 2.5475 | 100.62 | 21.1 | 25.734 | 91.8122 | 81.8791 | 87.0267 | 0.006407 | 0.9802 | 0.023872 | 24.3541 | 0.080418 | 0.15312 | 0.007445 | 0.53026 | 0.53026 | 0.011926
2019-03-28 15:45 | 1553802273.2795 | 372 | 0.2081 | 0.001617 | 7.8e-05 | 96.7773 | 96.1115 | 120.2352 | 21.2612 | 0.002961 | 0.002665 | 155.2737 | 0.25681 | 8.4018 | 0.002532 | 0.003764 | 0.00364 | 5.1e-05 | 0.001059 | 0.32388 | 1.3575 | 2.5475 | 100.62 | 21.1 | 25.747 | 93.0548 | 82.7402 | 85.9346 | 0.00641 | 0.98318 | 0.023898 | 24.3071 | 0.061747 | 0.15277 | 0.006846 | 0.48773 | 0.48773 | 0.012937
2019-03-28 15:45 | 1553802274.2786 | 373 | 0.20805 | 0.001607 | 7.8e-05 | 96.7257 | 95.8781 | 120.4514 | 21.2612 | 0.002933 | 0.002609 | 155.2751 | 0.25718 | 8.4011 | 0.002482 | 0.004195 | 0.004071 | 5.3e-05 | 0.001056 | 0.35059 | 1.2945 | 2.5475 | 100.62 | 21.1 | 25.754 | 92.3713 | 81.5369 | 84.5845 | 0.006412 | 0.98689 | 0.023965 | 24.2832 | 0.07873 | 0.153 | 0.007169 | 0.51234 | 0.51234 | 0.012333
2019-03-28 15:45 | 1553802275.2792 | 374 | 0.208 | 0.001601 | 7.8e-05 | 96.5423 | 95.7133 | 120.6446 | 21.2612 | 0.002945 | 0.002606 | 155.2778 | 0.25711 | 8.4044 | 0.002491 | 0.004365 | 0.004332 | 5.5e-05 | 0.001054 | 0.31852 | 1.3326 | 2.5475 | 100.62 | 21.1 | 25.76 | 92.6787 | 81.7371 | 84.5001 | 0.006414 | 0.98713 | 0.023987 | 24.2998 | 0.077132 | 0.15295 | 0.007472 | 0.53471 | 0.53471 | 0.011814
2019-03-28 15:45 | 1553802276.2783 | 375 | 0.20814 | 0.001589 | 7.8e-05 | 96.4199 | 95.449 | 120.2235 | 21.2612 | 0.002879 | 0.002673 | 155.2728 | 0.25802 | 8.4 | 0.002498 | 0.003419 | 0.003329 | 4.9e-05 | 0.001052 | 0.30574 | 1.316 | 2.5475 | 100.62 | 21.1 | 25.775 | 91.0897 | 81.9118 | 86.1218 | 0.006417 | 0.98267 | 0.023891 | 24.3123 | 0.090517 | 0.1535 | 0.006663 | 0.47444 | 0.47444 | 0.013362
2019-03-28 15:45 | 1553802277.2773 | 376 | 0.20816 | 0.001579 | 7.9e-05 | 96.228 | 95.3689 | 120.0627 | 21.2612 | 0.00287 | 0.00262 | 155.2545 | 0.25693 | 8.4037 | 0.002464 | 0.003689 | 0.003612 | 5.3e-05 | 0.001055 | 0.31974 | 1.3327 | 2.5475 | 100.62 | 21.1 | 25.783 | 90.8529 | 81.1023 | 84.856 | 0.006419 | 0.98614 | 0.023917 | 24.2532 | 0.080221 | 0.15285 | 0.006573 | 0.46849 | 0.46849 | 0.013475
2019-03-28 15:45 | 1553802278.2783 | 377 | 0.20838 | 0.001574 | 7.9e-05 | 96.1881 | 95.2418 | 120.2713 | 21.2612 | 0.002904 | 0.002615 | 155.2239 | 0.25601 | 8.4027 | 0.002506 | 0.004648 | 0.004582 | 5.3e-05 | 0.00105 | 0.33157 | 1.3134 | 2.5475 | 100.62 | 21.1 | 25.795 | 91.6721 | 82.0998 | 84.7257 | 0.006422 | 0.9865 | 0.023942 | 24.2698 | 0.088435 | 0.1523 | 0.005236 | 0.37266 | 0.37266 | 0.016878
2019-03-28 15:45 | 1553802279.2787 | 378 | 0.20823 | 0.001563 | 7.9e-05 | 95.9969 | 94.8926 | 119.8498 | 21.2612 | 0.002858 | 0.002626 | 155.2312 | 0.25707 | 8.4025 | 0.00247 | 0.003898 | 0.003841 | 5.5e-05 | 0.001056 | 0.33212 | 1.3032 | 2.5475 | 100.62 | 21.1 | 25.801 | 90.5816 | 81.2318 | 84.9854 | 0.006424 | 0.98579 | 0.023892 | 24.236 | 0.10349 | 0.15293 | 0.00617 | 0.43898 | 0.43898 | 0.014388
2019-03-28 15:45 | 1553802280.2774 | 379 | 0.20815 | 0.001556 | 8e-05 | 95.8974 | 94.9547 | 120.1031 | 21.2612 | 0.00288 | 0.002641 | 155.2094 | 0.2582 | 8.4029 | 0.002549 | 0.003531 | 0.003393 | 5.6e-05 | 0.001059 | 0.31349 | 1.296 | 2.5474 | 100.62 | 21.1 | 25.804 | 91.1148 | 83.1357 | 85.3546 | 0.006425 | 0.98477 | 0.023905 | 24.2741 | 0.088362 | 0.1536 | 0.006629 | 0.47214 | 0.47214 | 0.013436
2019-03-28 15:45 | 1553802281.2783 | 380 | 0.20818 | 0.001544 | 8e-05 | 95.7961 | 94.6578 | 119.9155 | 21.2612 | 0.002838 | 0.00264 | 155.1894 | 0.25914 | 8.4024 | 0.002526 | 0.004255 | 0.004162 | 4.9e-05 | 0.001061 | 0.31222 | 1.331 | 2.5474 | 100.62 | 21.1 | 25.809 | 90.1025 | 82.5873 | 85.3352 | 0.006426 | 0.98483 | 0.023886 | 24.2545 | 0.10692 | 0.15416 | 0.006469 | 0.46026 | 0.46026 | 0.013834
2019-03-28 15:45 | 1553802282.2784 | 381 | 0.20808 | 0.001537 | 8.1e-05 | 95.6314 | 94.6857 | 119.3672 | 21.2612 | 0.00283 | 0.002647 | 155.1325 | 0.25817 | 8.4031 | 0.002486 | 0.003629 | 0.003568 | 5.3e-05 | 0.001067 | 0.3194 | 1.2863 | 2.5474 | 100.62 | 21.09 | 25.809 | 89.8956 | 81.6373 | 85.4975 | 0.006422 | 0.98438 | 0.023826 | 24.2045 | 0.0889 | 0.15359 | 0.007098 | 0.50417 | 0.50417 | 0.012581
2019-03-28 15:45 | 1553802283.2786 | 382 | 0.20812 | 0.001524 | 8.2e-05 | 95.4968 | 94.6412 | 119.5322 | 21.2612 | 0.002891 | 0.002658 | 155.1313 | 0.25728 | 8.4012 | 0.002436 | 0.00384 | 0.003764 | 5.3e-05 | 0.001075 | 0.32068 | 1.3251 | 2.5474 | 100.62 | 21.09 | 25.808 | 91.3634 | 80.4262 | 85.7573 | 0.006422 | 0.98367 | 0.023834 | 24.23 | 0.080504 | 0.15306 | 0.00687 | 0.4879 | 0.4879 | 0.012956
2019-03-28 15:45 | 1553802284.2785 | 383 | 0.20811 | 0.001513 | 8.2e-05 | 95.3553 | 94.3374 | 119.4529 | 21.2612 | 0.0029 | 0.002649 | 155.1255 | 0.25744 | 8.4012 | 0.002508 | 0.003981 | 0.003859 | 5.2e-05 | 0.001077 | 0.32731 | 1.3375 | 2.5475 | 100.62 | 21.09 | 25.804 | 91.5885 | 82.145 | 85.5545 | 0.006421 | 0.98422 | 0.023833 | 24.2151 | 0.095993 | 0.15315 | 0.006898 | 0.48992 | 0.48992 | 0.012911
2019-03-28 15:45 | 1553802285.2786 | 384 | 0.20816 | 0.001506 | 8.3e-05 | 95.2532 | 94.4326 | 119.0839 | 21.2612 | 0.002883 | 0.002582 | 155.0721 | 0.25814 | 8.3993 | 0.002451 | 0.004713 | 0.004506 | 5.2e-05 | 0.001081 | 0.30055 | 1.2933 | 2.5474 | 100.62 | 21.09 | 25.802 | 91.1733 | 80.7906 | 83.9282 | 0.00642 | 0.98871 | 0.02385 | 24.1228 | 0.07739 | 0.15357 | 0.006644 | 0.472 | 0.472 | 0.013437
2019-03-28 15:45 | 1553802286.278 | 385 | 0.20812 | 0.001502 | 8.3e-05 | 95.0557 | 93.9563 | 119.3369 | 21.2612 | 0.002845 | 0.00261 | 155.0568 | 0.25733 | 8.4015 | 0.002483 | 0.003903 | 0.003892 | 5.6e-05 | 0.001088 | 0.30973 | 1.3085 | 2.5474 | 100.62 | 21.09 | 25.801 | 90.269 | 81.5491 | 84.6093 | 0.00642 | 0.98682 | 0.023853 | 24.1714 | 0.10406 | 0.15308 | 0.006889 | 0.48963 | 0.48963 | 0.012912
2019-03-28 15:45 | 1553802287.2777 | 386 | 0.20813 | 0.001496 | 8.3e-05 | 94.8844 | 93.7552 | 119.6418 | 21.2612 | 0.002821 | 0.002608 | 155.0348 | 0.25695 | 8.402 | 0.00242 | 0.003797 | 0.003615 | 5.1e-05 | 0.001092 | 0.31669 | 1.291 | 2.5474 | 100.62 | 21.09 | 25.801 | 89.6861 | 80.0294 | 84.5508 | 0.00642 | 0.98699 | 0.023885 | 24.2003 | 0.10709 | 0.15286 | 0.006854 | 0.48772 | 0.48772 | 0.012944
2019-03-28 15:45 | 1553802288.2777 | 387 | 0.20814 | 0.001489 | 8.4e-05 | 94.6917 | 93.5615 | 119.2031 | 21.2612 | 0.002831 | 0.002625 | 155.0195 | 0.25733 | 8.4024 | 0.002499 | 0.003806 | 0.003713 | 4.7e-05 | 0.001098 | 0.30254 | 1.3132 | 2.5475 | 100.62 | 21.09 | 25.801 | 89.9368 | 81.9414 | 84.9566 | 0.00642 | 0.98587 | 0.023828 | 24.1696 | 0.1074 | 0.15309 | 0.006753 | 0.47929 | 0.47929 | 0.013191
2019-03-28 15:45 | 1553802289.2791 | 388 | 0.20818 | 0.001481 | 8.4e-05 | 94.5237 | 93.3393 | 119.5003 | 21.2612 | 0.002822 | 0.002588 | 154.9691 | 0.25826 | 8.4033 | 0.002493 | 0.004411 | 0.004432 | 5.4e-05 | 0.001101 | 0.33859 | 1.3146 | 2.5474 | 100.62 | 21.08 | 25.803 | 89.7143 | 81.7948 | 84.0764 | 0.006417 | 0.9883 | 0.023887 | 24.1699 | 0.11278 | 0.15364 | 0.00652 | 0.46374 | 0.46374 | 0.013683
2019-03-28 15:45 | 1553802290.278 | 389 | 0.20817 | 0.001475 | 8.5e-05 | 94.2834 | 93.0716 | 119.5315 | 21.2612 | 0.002835 | 0.002599 | 154.9351 | 0.25754 | 8.4031 | 0.00249 | 0.003462 | 0.003303 | 5e-05 | 0.001106 | 0.32908 | 1.3216 | 2.5474 | 100.62 | 21.08 | 25.804 | 90.0259 | 81.7231 | 84.3387 | 0.006417 | 0.98757 | 0.023881 | 24.182 | 0.1157 | 0.15321 | 0.006634 | 0.4717 | 0.4717 | 0.013414
2019-03-28 15:45 | 1553802291.2781 | 390 | 0.20821 | 0.001469 | 8.7e-05 | 94.1651 | 93.0057 | 119.5519 | 21.2612 | 0.002832 | 0.002571 | 154.9197 | 0.25753 | 8.4025 | 0.002527 | 0.00409 | 0.004082 | 4.9e-05 | 0.001107 | 0.34866 | 1.3011 | 2.5474 | 100.62 | 21.08 | 25.809 | 89.9399 | 82.6044 | 83.6582 | 0.006418 | 0.98945 | 0.023906 | 24.161 | 0.11081 | 0.15321 | 0.006351 | 0.45176 | 0.45176 | 0.014006
2019-03-28 15:45 | 1553802292.2778 | 391 | 0.2082 | 0.001459 | 8.8e-05 | 93.9659 | 92.9568 | 119.7744 | 21.2612 | 0.002834 | 0.002624 | 154.9094 | 0.25927 | 8.403 | 0.002418 | 0.004599 | 0.004467 | 5e-05 | 0.001104 | 0.32925 | 1.3235 | 2.5475 | 100.62 | 21.08 | 25.813 | 89.9949 | 79.9881 | 84.9426 | 0.006419 | 0.98591 | 0.023886 | 24.227 | 0.09657 | 0.15424 | 0.006426 | 0.45672 | 0.45672 | 0.013947
2019-03-28 15:45 | 1553802293.2771 | 392 | 0.20818 | 0.001452 | 9e-05 | 93.6424 | 92.4007 | 119.888 | 21.2612 | 0.002804 | 0.002599 | 154.8731 | 0.2581 | 8.4065 | 0.002521 | 0.003393 | 0.003288 | 5.4e-05 | 0.001104 | 0.33323 | 1.3544 | 2.5474 | 100.62 | 21.08 | 25.816 | 89.2844 | 82.4688 | 84.3503 | 0.00642 | 0.98754 | 0.023917 | 24.2184 | 0.1194 | 0.15354 | 0.006564 | 0.46709 | 0.46709 | 0.013576
2019-03-28 15:45 | 1553802294.2769 | 393 | 0.20822 | 0.001444 | 9.1e-05 | 93.4634 | 92.0385 | 119.7336 | 21.2612 | 0.002847 | 0.002561 | 154.8224 | 0.25826 | 8.4041 | 0.002482 | 0.004316 | 0.00426 | 5.1e-05 | 0.001094 | 0.34971 | 1.3112 | 2.5474 | 100.62 | 21.08 | 25.826 | 90.3166 | 81.5234 | 83.4187 | 0.006422 | 0.99012 | 0.023932 | 24.1713 | 0.13741 | 0.15364 | 0.006344 | 0.45152 | 0.45152 | 0.014053
2019-03-28 15:45 | 1553802295.2778 | 394 | 0.20819 | 0.001437 | 9.3e-05 | 93.1587 | 91.7613 | 119.6228 | 21.2612 | 0.002803 | 0.002635 | 154.8331 | 0.25762 | 8.4025 | 0.002509 | 0.004438 | 0.004251 | 4.9e-05 | 0.00109 | 0.31365 | 1.3149 | 2.5474 | 100.62 | 21.08 | 25.831 | 89.2541 | 82.1755 | 85.2048 | 0.006423 | 0.98518 | 0.023862 | 24.2205 | 0.13518 | 0.15326 | 0.006508 | 0.46183 | 0.46183 | 0.013705
2019-03-28 15:45 | 1553802296.2769 | 395 | 0.20821 | 0.001424 | 9.5e-05 | 92.9263 | 91.5524 | 119.8169 | 21.2612 | 0.002811 | 0.002572 | 154.8052 | 0.25801 | 8.402 | 0.002443 | 0.00372 | 0.003703 | 4.7e-05 | 0.001083 | 0.33085 | 1.3217 | 2.5475 | 100.62 | 21.08 | 25.843 | 89.4501 | 80.5932 | 83.6964 | 0.006426 | 0.98935 | 0.023931 | 24.1891 | 0.13324 | 0.15349 | 0.006429 | 0.45738 | 0.45738 | 0.01386
2019-03-28 15:45 | 1553802297.2773 | 396 | 0.20822 | 0.001418 | 9.6e-05 | 92.6604 | 91.2487 | 119.7775 | 21.2612 | 0.002825 | 0.0026 | 154.7806 | 0.25782 | 8.4019 | 0.002485 | 0.003876 | 0.003725 | 5e-05 | 0.001081 | 0.31542 | 1.3289 | 2.5474 | 100.62 | 21.08 | 25.85 | 89.7909 | 81.6085 | 84.362 | 0.006428 | 0.98751 | 0.023905 | 24.2076 | 0.13732 | 0.15338 | 0.006373 | 0.45274 | 0.45274 | 0.013991
2019-03-28 15:45 | 1553802298.277 | 397 | 0.20821 | 0.001412 | 9.7e-05 | 92.4373 | 90.7931 | 119.9561 | 21.2612 | 0.002754 | 0.002561 | 154.7518 | 0.2584 | 8.4065 | 0.002495 | 0.003763 | 0.003738 | 5e-05 | 0.001083 | 0.3178 | 1.3021 | 2.5474 | 100.62 | 21.07 | 25.859 | 88.0674 | 81.8527 | 83.4266 | 0.006426 | 0.9901 | 0.023954 | 24.194 | 0.16053 | 0.15372 | 0.006426 | 0.45747 | 0.45747 | 0.013878
2019-03-28 15:45 | 1553802299.278 | 398 | 0.20822 | 0.001403 | 9.8e-05 | 92.191 | 90.8104 | 120.1054 | 21.2612 | 0.002787 | 0.002587 | 154.7372 | 0.25835 | 8.3991 | 0.002482 | 0.004426 | 0.004287 | 5.2e-05 | 0.001086 | 0.33859 | 1.31 | 2.5474 | 100.62 | 21.07 | 25.88 | 88.8751 | 81.5369 | 84.0492 | 0.006432 | 0.98837 | 0.023948 | 24.2301 | 0.13495 | 0.15369 | 0.006396 | 0.4552 | 0.4552 | 0.013945
2019-03-28 15:45 | 1553802300.2774 | 399 | 0.20825 | 0.001404 | 9.9e-05 | 91.7954 | 90.3263 | 120.185 | 21.2612 | 0.002737 | 0.002559 | 154.7269 | 0.25838 | 8.4057 | 0.00262 | 0.004104 | 0.004043 | 4.9e-05 | 0.001073 | 0.323 | 1.341 | 2.5474 | 100.62 | 21.07 | 25.891 | 87.652 | 84.8568 | 83.3798 | 0.006434 | 0.99023 | 0.023979 | 24.2154 | 0.14431 | 0.15371 | 0.00623 | 0.44368 | 0.44368 | 0.014308
2019-03-28 15:45 | 1553802301.2781 | 400 | 0.20828 | 0.001397 | 9.9e-05 | 91.555 | 90.0087 | 120.1918 | 21.2612 | 0.00281 | 0.002564 | 154.6867 | 0.25785 | 8.4042 | 0.002582 | 0.004019 | 0.003945 | 5.1e-05 | 0.001069 | 0.3235 | 1.3919 | 2.5474 | 100.62 | 21.07 | 25.901 | 89.4104 | 83.9439 | 83.5045 | 0.006437 | 0.98988 | 0.023975 | 24.2204 | 0.15236 | 0.15339 | 0.006016 | 0.42815 | 0.42815 | 0.014797
2019-03-28 15:45 | 1553802302.2771 | 401 | 0.20829 | 0.001395 | 0.000101 | 91.4037 | 89.8673 | 120.3481 | 21.2612 | 0.00276 | 0.002554 | 154.6816 | 0.25824 | 8.4038 | 0.002557 | 0.003885 | 0.003741 | 5.2e-05 | 0.001071 | 0.3246 | 1.3687 | 2.5474 | 100.62 | 21.07 | 25.922 | 88.2174 | 83.3331 | 83.2605 | 0.006442 | 0.99056 | 0.023999 | 24.2278 | 0.15163 | 0.15363 | 0.005996 | 0.42706 | 0.42706 | 0.014857
2019-03-28 15:45 | 1553802303.2785 | 402 | 0.2083 | 0.001386 | 0.000101 | 91.1476 | 89.5911 | 120.2709 | 21.2612 | 0.00277 | 0.002547 | 154.6634 | 0.25783 | 8.4002 | 0.002545 | 0.003479 | 0.003395 | 5.2e-05 | 0.001071 | 0.30646 | 1.3722 | 2.5475 | 100.62 | 21.07 | 25.93 | 88.4564 | 83.0318 | 83.0857 | 0.006444 | 0.99104 | 0.023997 | 24.2141 | 0.15406 | 0.15338 | 0.005942 | 0.42314 | 0.42314 | 0.014971
2019-03-28 15:45 | 1553802304.2771 | 403 | 0.20827 | 0.001384 | 0.000102 | 90.8694 | 89.3914 | 120.3486 | 21.2612 | 0.002732 | 0.002522 | 154.637 | 0.25778 | 8.405 | 0.002608 | 0.004441 | 0.004292 | 5.1e-05 | 0.00107 | 0.33787 | 1.3913 | 2.5476 | 100.62 | 21.07 | 25.948 | 87.5521 | 84.5633 | 82.4921 | 0.006449 | 0.9927 | 0.024025 | 24.2017 | 0.14668 | 0.15335 | 0.006101 | 0.43506 | 0.43506 | 0.014557
2019-03-28 15:45 | 1553802305.2778 | 404 | 0.20827 | 0.001377 | 0.000103 | 90.6094 | 89.0907 | 120.8962 | 21.2612 | 0.002785 | 0.002562 | 154.6012 | 0.25771 | 8.4023 | 0.002579 | 0.004263 | 0.004169 | 5.1e-05 | 0.001074 | 0.34075 | 1.3501 | 2.5475 | 100.62 | 21.07 | 25.956 | 88.8166 | 83.8657 | 83.4485 | 0.006451 | 0.99004 | 0.024047 | 24.2893 | 0.15118 | 0.15331 | 0.006145 | 0.43855 | 0.43855 | 0.014438
2019-03-28 15:45 | 1553802306.2772 | 405 | 0.2083 | 0.001369 | 0.000103 | 90.4104 | 88.77 | 120.7345 | 21.2612 | 0.00279 | 0.002555 | 154.6015 | 0.25825 | 8.4031 | 0.002597 | 0.004662 | 0.00462 | 5e-05 | 0.001072 | 0.32156 | 1.3034 | 2.5473 | 100.62 | 21.07 | 25.965 | 88.944 | 84.2988 | 83.2939 | 0.006453 | 0.99047 | 0.024036 | 24.2678 | 0.16378 | 0.15363 | 0.005926 | 0.42255 | 0.42255 | 0.015016
2019-03-28 15:45 | 1553802307.2763 | 406 | 0.2083 | 0.001362 | 0.000103 | 90.1408 | 88.3389 | 120.7237 | 21.2612 | 0.002749 | 0.002534 | 154.5855 | 0.25886 | 8.4032 | 0.002562 | 0.004114 | 0.004011 | 4.9e-05 | 0.001075 | 0.32156 | 1.2845 | 2.5474 | 100.62 | 21.07 | 25.979 | 87.9512 | 83.4569 | 82.7885 | 0.006456 | 0.99187 | 0.024052 | 24.2495 | 0.18061 | 0.154 | 0.005921 | 0.4224 | 0.4224 | 0.015057
2019-03-28 15:45 | 1553802308.2777 | 407 | 0.20833 | 0.001358 | 0.000105 | 89.8966 | 88.3158 | 120.9457 | 21.2612 | 0.00273 | 0.002553 | 154.5721 | 0.25911 | 8.4014 | 0.00252 | 0.004039 | 0.003906 | 4.8e-05 | 0.001076 | 0.30502 | 1.2731 | 2.5473 | 100.62 | 21.07 | 25.984 | 87.4999 | 82.4532 | 83.2425 | 0.006458 | 0.99061 | 0.024059 | 24.2873 | 0.15868 | 0.15415 | 0.005774 | 0.4118 | 0.4118 | 0.015459
2019-03-28 15:45 | 1553802309.2764 | 408 | 0.20832 | 0.001349 | 0.000105 | 89.7926 | 88.0719 | 120.9042 | 21.2612 | 0.002726 | 0.002512 | 154.5513 | 0.25896 | 8.4058 | 0.002558 | 0.003681 | 0.003553 | 5e-05 | 0.001076 | 0.3449 | 1.3279 | 2.5474 | 100.62 | 21.07 | 25.988 | 87.3976 | 83.354 | 82.2613 | 0.006459 | 0.99334 | 0.024088 | 24.2497 | 0.17307 | 0.15406 | 0.005856 | 0.41819 | 0.41819 | 0.015214
2019-03-28 15:45 | 1553802310.2773 | 409 | 0.20833 | 0.001344 | 0.000107 | 89.7644 | 88.2948 | 120.6157 | 21.2612 | 0.002768 | 0.002525 | 154.5175 | 0.25772 | 8.4032 | 0.002555 | 0.004126 | 0.003976 | 4.9e-05 | 0.001075 | 0.32068 | 1.3212 | 2.5474 | 100.62 | 21.07 | 25.986 | 88.4143 | 83.2737 | 82.5537 | 0.006458 | 0.99253 | 0.02405 | 24.2307 | 0.14765 | 0.15332 | 0.00577 | 0.41122 | 0.41122 | 0.015398
2019-03-28 15:45 | 1553802311.2774 | 410 | 0.20833 | 0.001337 | 0.000109 | 89.9148 | 88.1422 | 120.5626 | 21.2612 | 0.002765 | 0.002536 | 154.5039 | 0.25808 | 8.4029 | 0.002556 | 0.0042 | 0.004116 | 5e-05 | 0.001075 | 0.34468 | 1.3162 | 2.5473 | 100.62 | 21.07 | 25.983 | 88.3377 | 83.2987 | 82.8386 | 0.006457 | 0.99173 | 0.024035 | 24.235 | 0.17809 | 0.15353 | 0.005823 | 0.41468 | 0.41468 | 0.015291
2019-03-28 15:45 | 1553802312.2768 | 411 | 0.20833 | 0.001333 | 0.00011 | 90.1533 | 88.3403 | 120.7936 | 21.2612 | 0.002762 | 0.002533 | 154.4832 | 0.25776 | 8.4018 | 0.002521 | 0.004029 | 0.003941 | 5e-05 | 0.001071 | 0.31797 | 1.3355 | 2.5474 | 100.62 | 21.07 | 25.977 | 88.2557 | 82.4553 | 82.7628 | 0.006456 | 0.99194 | 0.02406 | 24.2557 | 0.18172 | 0.15334 | 0.005769 | 0.41114 | 0.41114 | 0.015404
2019-03-28 15:45 | 1553802313.2775 | 412 | 0.20826 | 0.001326 | 0.000111 | 90.3577 | 88.5499 | 120.8255 | 21.2612 | 0.002778 | 0.002483 | 154.4575 | 0.25832 | 8.4006 | 0.002551 | 0.003917 | 0.003779 | 4.8e-05 | 0.001074 | 0.31492 | 1.3789 | 2.5474 | 100.62 | 21.07 | 25.973 | 88.6572 | 83.1864 | 81.5602 | 0.006455 | 0.99531 | 0.024104 | 24.2179 | 0.18077 | 0.15367 | 0.006202 | 0.44327 | 0.44327 | 0.014318
2019-03-28 15:45 | 1553802314.2774 | 413 | 0.2083 | 0.001319 | 0.000112 | 90.5638 | 88.9718 | 120.8964 | 21.2612 | 0.002758 | 0.002566 | 154.427 | 0.25837 | 8.4013 | 0.002553 | 0.003641 | 0.003567 | 4.9e-05 | 0.001075 | 0.33516 | 1.3706 | 2.5476 | 100.62 | 21.07 | 25.968 | 88.1558 | 83.2425 | 83.5375 | 0.006454 | 0.98979 | 0.024044 | 24.2924 | 0.15863 | 0.1537 | 0.006013 | 0.42839 | 0.42839 | 0.014818
2019-03-28 15:45 | 1553802315.2779 | 414 | 0.2083 | 0.00131 | 0.000113 | 90.8637 | 89.1988 | 120.7722 | 21.2612 | 0.002694 | 0.002531 | 154.4063 | 0.25897 | 8.4044 | 0.002519 | 0.004175 | 0.004011 | 4.9e-05 | 0.001063 | 0.32117 | 1.2987 | 2.5474 | 100.62 | 21.07 | 25.959 | 86.6331 | 82.4188 | 82.7034 | 0.006451 | 0.99211 | 0.02406 | 24.2515 | 0.16542 | 0.15406 | 0.006004 | 0.428 | 0.428 | 0.014866
2019-03-28 15:45 | 1553802316.2766 | 415 | 0.20831 | 0.00131 | 0.000113 | 91.217 | 89.7043 | 120.7366 | 21.2612 | 0.002699 | 0.00252 | 154.3599 | 0.25797 | 8.4006 | 0.002514 | 0.004267 | 0.004189 | 5.2e-05 | 0.001064 | 0.33644 | 1.3806 | 2.5475 | 100.62 | 21.06 | 25.955 | 86.7579 | 82.2986 | 82.4376 | 0.006446 | 0.99285 | 0.024066 | 24.2389 | 0.14958 | 0.15347 | 0.005951 | 0.42421 | 0.42421 | 0.014941
2019-03-28 15:45 | 1553802317.2769 | 416 | 0.20832 | 0.001307 | 0.000114 | 91.5008 | 89.9231 | 120.8572 | 21.2612 | 0.002712 | 0.002589 | 154.3294 | 0.25728 | 8.4052 | 0.002558 | 0.003901 | 0.003834 | 4.9e-05 | 0.001057 | 0.30397 | 1.3266 | 2.5474 | 100.62 | 21.06 | 25.948 | 87.0688 | 83.3526 | 84.0944 | 0.006445 | 0.98825 | 0.024022 | 24.3074 | 0.15558 | 0.15306 | 0.005884 | 0.4186 | 0.4186 | 0.015101
2019-03-28 15:45 | 1553802318.2776 | 417 | 0.20838 | 0.001302 | 0.000114 | 91.7326 | 90.2081 | 120.9093 | 21.2612 | 0.002731 | 0.002518 | 154.3197 | 0.25819 | 8.4024 | 0.0025 | 0.003484 | 0.003302 | 4.9e-05 | 0.001059 | 0.33931 | 1.2897 | 2.5473 | 100.62 | 21.06 | 25.945 | 87.5201 | 81.9609 | 82.3844 | 0.006444 | 0.993 | 0.024085 | 24.2544 | 0.1499 | 0.1536 | 0.005543 | 0.39492 | 0.39492 | 0.016063
2019-03-28 15:45 | 1553802319.2776 | 418 | 0.20839 | 0.001296 | 0.000115 | 91.9971 | 90.7296 | 121.025 | 21.2612 | 0.002744 | 0.002541 | 154.2869 | 0.25825 | 8.4015 | 0.002517 | 0.003789 | 0.003718 | 5.3e-05 | 0.001061 | 0.3126 | 1.2917 | 2.5474 | 100.62 | 21.06 | 25.941 | 87.8332 | 82.3813 | 82.9499 | 0.006443 | 0.99142 | 0.024077 | 24.2853 | 0.12409 | 0.15363 | 0.005501 | 0.39171 | 0.39171 | 0.016198
2019-03-28 15:45 | 1553802320.2781 | 419 | 0.20838 | 0.001285 | 0.000115 | 92.2118 | 90.8807 | 121.4443 | 21.2612 | 0.002717 | 0.00248 | 154.3046 | 0.25832 | 8.4013 | 0.002466 | 0.004476 | 0.00426 | 5.1e-05 | 0.001059 | 0.33483 | 1.2869 | 2.5474 | 100.62 | 21.06 | 25.94 | 87.1788 | 81.1328 | 81.4775 | 0.006443 | 0.99554 | 0.024169 | 24.277 | 0.13005 | 0.15368 | 0.005517 | 0.39438 | 0.39438 | 0.016093
2019-03-28 15:45 | 1553802321.2777 | 420 | 0.20838 | 0.001283 | 0.000115 | 92.4563 | 91.1308 | 121.5971 | 21.2612 | 0.002713 | 0.002489 | 154.2812 | 0.25635 | 8.4011 | 0.002516 | 0.003789 | 0.003794 | 5.4e-05 | 0.001061 | 0.34003 | 1.286 | 2.5473 | 100.62 | 21.07 | 25.939 | 87.0821 | 82.351 | 81.6967 | 0.006446 | 0.99492 | 0.024176 | 24.2997 | 0.12917 | 0.1525 | 0.00556 | 0.39761 | 0.39761 | 0.01584
2019-03-28 15:45 | 1553802322.2778 | 421 | 0.20838 | 0.001277 | 0.000116 | 92.7393 | 91.4848 | 121.4546 | 21.2612 | 0.002709 | 0.002544 | 154.2503 | 0.25834 | 8.4008 | 0.002507 | 0.00337 | 0.003255 | 5.4e-05 | 0.00106 | 0.31902 | 1.3017 | 2.5474 | 100.62 | 21.07 | 25.94 | 86.9885 | 82.1318 | 83.0148 | 0.006447 | 0.99124 | 0.024117 | 24.3306 | 0.12182 | 0.15368 | 0.005539 | 0.39507 | 0.39507 | 0.016066
2019-03-28 15:45 | 1553802323.2768 | 422 | 0.20838 | 0.001274 | 0.000117 | 92.9463 | 91.764 | 121.4788 | 21.2612 | 0.002709 | 0.002536 | 154.2175 | 0.25776 | 8.4027 | 0.002554 | 0.003532 | 0.003419 | 5.3e-05 | 0.001051 | 0.31924 | 1.3597 | 2.5475 | 100.62 | 21.07 | 25.939 | 86.9962 | 83.2518 | 82.8376 | 0.006446 | 0.99174 | 0.024126 | 24.3269 | 0.1145 | 0.15334 | 0.005553 | 0.39614 | 0.39614 | 0.015987
2019-03-28 15:45 | 1553802324.2771 | 423 | 0.20841 | 0.001264 | 0.000118 | 93.2176 | 92.0504 | 121.4703 | 21.2612 | 0.002741 | 0.002512 | 154.2097 | 0.25783 | 8.4015 | 0.002507 | 0.003153 | 0.003116 | 5.1e-05 | 0.001057 | 0.32709 | 1.3336 | 2.5476 | 100.62 | 21.07 | 25.939 | 87.7605 | 82.1318 | 82.2464 | 0.006446 | 0.99338 | 0.024145 | 24.3059 | 0.11271 | 0.15338 | 0.005383 | 0.38407 | 0.38407 | 0.016493
2019-03-28 15:45 | 1553802325.2759 | 424 | 0.2084 | 0.001257 | 0.00012 | 93.4032 | 91.9482 | 121.4579 | 21.2612 | 0.002694 | 0.002538 | 154.1779 | 0.25763 | 8.406 | 0.002543 | 0.004122 | 0.004031 | 5.2e-05 | 0.001057 | 0.31614 | 1.3013 | 2.5475 | 100.62 | 21.07 | 25.94 | 86.6222 | 83.0023 | 82.8744 | 0.006447 | 0.99163 | 0.024123 | 24.3261 | 0.14044 | 0.15326 | 0.005458 | 0.389 | 0.389 | 0.016272
2019-03-28 15:45 | 1553802326.2771 | 425 | 0.2084 | 0.001257 | 0.000122 | 93.643 | 92.3202 | 121.383 | 21.2612 | 0.002733 | 0.002485 | 154.1592 | 0.25843 | 8.4014 | 0.002434 | 0.004526 | 0.004407 | 5.3e-05 | 0.001058 | 0.31116 | 1.2979 | 2.5475 | 100.62 | 21.07 | 25.94 | 87.5663 | 80.3647 | 81.6061 | 0.006447 | 0.99518 | 0.024158 | 24.2752 | 0.12725 | 0.15374 | 0.005441 | 0.38829 | 0.38829 | 0.016353
2019-03-28 15:45 | 1553802327.2763 | 426 | 0.20841 | 0.001249 | 0.000123 | 93.9042 | 92.7941 | 121.4107 | 21.2612 | 0.002682 | 0.002496 | 154.1629 | 0.25792 | 8.4008 | 0.002432 | 0.004263 | 0.004196 | 5.3e-05 | 0.001061 | 0.35834 | 1.3323 | 2.5473 | 100.62 | 21.07 | 25.941 | 86.3475 | 80.3195 | 81.869 | 0.006447 | 0.99444 | 0.024152 | 24.287 | 0.10637 | 0.15344 | 0.005423 | 0.38674 | 0.38674 | 0.016386
2019-03-28 15:45 | 1553802328.2759 | 427 | 0.20841 | 0.001239 | 0.000125 | 94.1958 | 93.0546 | 121.3534 | 21.2612 | 0.002689 | 0.002488 | 154.132 | 0.25797 | 8.4003 | 0.002476 | 0.003801 | 0.003727 | 5.3e-05 | 0.001061 | 0.33787 | 1.3357 | 2.5475 | 100.62 | 21.07 | 25.941 | 86.5169 | 81.3847 | 81.6717 | 0.006447 | 0.99499 | 0.024153 | 24.2745 | 0.10902 | 0.15347 | 0.005426 | 0.3869 | 0.3869 | 0.016382
2019-03-28 15:45 | 1553802329.2764 | 428 | 0.20841 | 0.00124 | 0.000126 | 94.3685 | 93.2683 | 121.2006 | 21.2612 | 0.002722 | 0.002545 | 154.1194 | 0.25834 | 8.4039 | 0.002518 | 0.003867 | 0.003729 | 5e-05 | 0.001059 | 0.32388 | 1.3276 | 2.5475 | 100.62 | 21.07 | 25.939 | 87.3054 | 82.3861 | 83.0544 | 0.006446 | 0.99113 | 0.024091 | 24.3065 | 0.10489 | 0.15369 | 0.005407 | 0.38445 | 0.38445 | 0.01651
2019-03-28 15:45 | 1553802330.2774 | 429 | 0.20843 | 0.001234 | 0.000127 | 94.6348 | 93.515 | 121.0302 | 21.2612 | 0.002727 | 0.002491 | 154.0863 | 0.25754 | 8.4042 | 0.002486 | 0.00433 | 0.004241 | 5.6e-05 | 0.001057 | 0.32947 | 1.3501 | 2.5475 | 100.62 | 21.07 | 25.937 | 87.4178 | 81.621 | 81.7503 | 0.006446 | 0.99477 | 0.024118 | 24.2449 | 0.10648 | 0.15321 | 0.005298 | 0.37696 | 0.37696 | 0.016786
2019-03-28 15:46 | 1553802331.276 | 430 | 0.20844 | 0.001226 | 0.000127 | 94.8648 | 93.9862 | 120.9882 | 21.2612 | 0.002775 | 0.002514 | 154.0728 | 0.25847 | 8.4004 | 0.002457 | 0.004267 | 0.004121 | 5.3e-05 | 0.001061 | 0.33268 | 1.3192 | 2.5473 | 100.62 | 21.06 | 25.932 | 88.5729 | 80.9184 | 82.2871 | 0.006441 | 0.99327 | 0.024096 | 24.259 | 0.083226 | 0.15377 | 0.005266 | 0.37424 | 0.37424 | 0.016969
2019-03-28 15:46 | 1553802332.2764 | 431 | 0.20844 | 0.001224 | 0.000128 | 95.154 | 94.1581 | 120.8366 | 21.2612 | 0.002679 | 0.002463 | 154.0622 | 0.25808 | 8.4047 | 0.002553 | 0.003814 | 0.003752 | 5.7e-05 | 0.001059 | 0.33843 | 1.3053 | 2.5473 | 100.62 | 21.06 | 25.929 | 86.2702 | 83.2449 | 81.0672 | 0.00644 | 0.99669 | 0.024122 | 24.2021 | 0.09411 | 0.15353 | 0.005249 | 0.37341 | 0.37341 | 0.016981
2019-03-28 15:46 | 1553802333.2762 | 432 | 0.20848 | 0.001222 | 0.000129 | 95.2993 | 94.471 | 120.9612 | 21.2612 | 0.002716 | 0.002579 | 154.0468 | 0.25737 | 8.4054 | 0.002502 | 0.004364 | 0.004257 | 5.4e-05 | 0.001059 | 0.33251 | 1.3188 | 2.5473 | 100.62 | 21.06 | 25.927 | 87.1548 | 82.014 | 83.8688 | 0.006439 | 0.98887 | 0.02404 | 24.3102 | 0.078089 | 0.15311 | 0.005002 | 0.35419 | 0.35419 | 0.017853
2019-03-28 15:46 | 1553802334.2761 | 433 | 0.20847 | 0.001211 | 0.00013 | 95.4658 | 94.4971 | 120.8772 | 21.2612 | 0.002676 | 0.002431 | 154.0422 | 0.25825 | 8.4024 | 0.002471 | 0.004019 | 0.003953 | 5.2e-05 | 0.001061 | 0.3361 | 1.3418 | 2.5473 | 100.62 | 21.06 | 25.918 | 86.1858 | 81.2614 | 80.2931 | 0.006437 | 0.99887 | 0.024153 | 24.1797 | 0.091226 | 0.15363 | 0.005064 | 0.36033 | 0.36033 | 0.017609
2019-03-28 15:46 | 1553802335.2775 | 434 | 0.20848 | 0.001209 | 0.00013 | 95.6985 | 94.6972 | 121.29 | 21.2612 | 0.002708 | 0.002492 | 154.0343 | 0.25845 | 8.4009 | 0.00243 | 0.00439 | 0.004199 | 5.7e-05 | 0.001061 | 0.33826 | 1.3341 | 2.5473 | 100.62 | 21.06 | 25.915 | 86.9711 | 80.2734 | 81.7597 | 0.006436 | 0.99475 | 0.024144 | 24.2712 | 0.094075 | 0.15375 | 0.005034 | 0.35794 | 0.35794 | 0.01774
2019-03-28 15:46 | 1553802336.2776 | 435 | 0.20846 | 0.001199 | 0.000131 | 95.9338 | 95.0505 | 121.5973 | 21.2612 | 0.002625 | 0.002467 | 154.0339 | 0.25858 | 8.4037 | 0.002471 | 0.004083 | 0.004047 | 5.1e-05 | 0.001063 | 0.31653 | 1.3142 | 2.5475 | 100.62 | 21.06 | 25.912 | 84.9729 | 81.2693 | 81.1648 | 0.006436 | 0.99642 | 0.024195 | 24.2815 | 0.082731 | 0.15383 | 0.005152 | 0.36728 | 0.36728 | 0.017298
2019-03-28 15:46 | 1553802337.2772 | 436 | 0.20845 | 0.001191 | 0.000131 | 96.0571 | 95.0628 | 121.5393 | 21.2612 | 0.002661 | 0.002475 | 154.0052 | 0.25738 | 8.4052 | 0.002486 | 0.003962 | 0.00382 | 5.4e-05 | 0.001071 | 0.34396 | 1.3533 | 2.5475 | 100.62 | 21.06 | 25.909 | 85.8314 | 81.6217 | 81.3698 | 0.006435 | 0.99584 | 0.024182 | 24.2828 | 0.093069 | 0.15312 | 0.005202 | 0.37074 | 0.37074 | 0.017057
2019-03-28 15:46 | 1553802338.2753 | 437 | 0.20851 | 0.001181 | 0.000131 | 96.265 | 95.326 | 121.3868 | 21.2612 | 0.002656 | 0.002529 | 153.9968 | 0.25864 | 8.4061 | 0.002502 | 0.004292 | 0.004174 | 5.5e-05 | 0.001071 | 0.31957 | 1.3079 | 2.5474 | 100.62 | 21.06 | 25.906 | 85.7128 | 82.0178 | 82.6707 | 0.006434 | 0.9922 | 0.024122 | 24.312 | 0.087669 | 0.15387 | 0.004884 | 0.34672 | 0.34672 | 0.018328
2019-03-28 15:46 | 1553802339.2756 | 438 | 0.20847 | 0.001184 | 0.00013 | 96.4767 | 95.5414 | 121.4867 | 21.2612 | 0.002644 | 0.002485 | 154.0043 | 0.25744 | 8.4057 | 0.002538 | 0.004049 | 0.004077 | 5.3e-05 | 0.001069 | 0.34235 | 1.3369 | 2.5474 | 100.62 | 21.05 | 25.904 | 85.4225 | 82.8737 | 81.5921 | 0.00643 | 0.99522 | 0.024169 | 24.2851 | 0.087137 | 0.15315 | 0.005087 | 0.3622 | 0.3622 | 0.017463
2019-03-28 15:46 | 1553802340.2763 | 439 | 0.2085 | 0.001176 | 0.00013 | 96.6787 | 95.9977 | 121.7534 | 21.2612 | 0.002695 | 0.002503 | 153.9664 | 0.25839 | 8.4021 | 0.002554 | 0.003969 | 0.003754 | 5.3e-05 | 0.001068 | 0.32516 | 1.3368 | 2.5475 | 100.62 | 21.05 | 25.904 | 86.6629 | 83.2574 | 82.0226 | 0.00643 | 0.99401 | 0.024181 | 24.3265 | 0.063228 | 0.15371 | 0.00498 | 0.35457 | 0.35457 | 0.017904
2019-03-28 15:46 | 1553802341.2763 | 440 | 0.20851 | 0.001168 | 0.00013 | 96.7451 | 95.8903 | 122.0581 | 21.2612 | 0.00271 | 0.002485 | 153.9708 | 0.25863 | 8.4021 | 0.002479 | 0.004633 | 0.00452 | 5.6e-05 | 0.001069 | 0.31725 | 1.3428 | 2.5473 | 100.62 | 21.05 | 25.905 | 87.0048 | 81.4486 | 81.6109 | 0.00643 | 0.99516 | 0.024225 | 24.3428 | 0.079379 | 0.15386 | 0.004907 | 0.34991 | 0.34991 | 0.01816
2019-03-28 15:46 | 1553802342.2764 | 441 | 0.20854 | 0.001163 | 0.00013 | 97.0549 | 96.3282 | 122.2041 | 21.2612 | 0.002645 | 0.002467 | 153.9705 | 0.25868 | 8.4036 | 0.002612 | 0.002984 | 0.002969 | 5.3e-05 | 0.001069 | 0.3501 | 1.3395 | 2.5475 | 100.62 | 21.05 | 25.907 | 85.4545 | 84.6617 | 81.171 | 0.00643 | 0.9964 | 0.024255 | 24.3423 | 0.067219 | 0.15389 | 0.00473 | 0.33743 | 0.33743 | 0.018835
2019-03-28 15:46 | 1553802343.2754 | 442 | 0.2085 | 0.00116 | 0.00013 | 97.1758 | 96.3537 | 122.0135 | 21.2612 | 0.002646 | 0.00246 | 153.9723 | 0.25681 | 8.402 | 0.002547 | 0.003825 | 0.003807 | 4.9e-05 | 0.00107 | 0.335 | 1.3488 | 2.5476 | 100.62 | 21.05 | 25.909 | 85.4701 | 83.0811 | 81.0011 | 0.006431 | 0.99688 | 0.024242 | 24.3174 | 0.075991 | 0.15278 | 0.00497 | 0.35482 | 0.35482 | 0.017783
2019-03-28 15:46 | 1553802344.2764 | 443 | 0.20856 | 0.001158 | 0.000131 | 97.3646 | 96.4512 | 122.2092 | 21.2612 | 0.002675 | 0.002459 | 153.9405 | 0.25732 | 8.4017 | 0.002533 | 0.004629 | 0.00445 | 5.2e-05 | 0.001061 | 0.32493 | 1.3611 | 2.5476 | 100.62 | 21.05 | 25.913 | 86.1774 | 82.7667 | 80.9792 | 0.006432 | 0.99694 | 0.024262 | 24.3362 | 0.084309 | 0.15308 | 0.004606 | 0.32842 | 0.32842 | 0.01925
2019-03-28 15:46 | 1553802345.2769 | 444 | 0.20855 | 0.001145 | 0.000131 | 97.4748 | 96.7845 | 121.9698 | 21.2612 | 0.002644 | 0.002421 | 153.9124 | 0.2579 | 8.403 | 0.002515 | 0.003975 | 0.00384 | 5.2e-05 | 0.001064 | 0.29302 | 1.3792 | 2.5475 | 100.62 | 21.05 | 25.915 | 85.435 | 82.3121 | 80.0614 | 0.006432 | 0.99953 | 0.024269 | 24.2808 | 0.063574 | 0.15343 | 0.00466 | 0.33246 | 0.33246 | 0.019059
2019-03-28 15:46 | 1553802346.2765 | 445 | 0.20855 | 0.001141 | 0.000131 | 97.7158 | 96.9605 | 121.6217 | 21.2612 | 0.002611 | 0.002452 | 153.8789 | 0.25756 | 8.4054 | 0.002594 | 0.003988 | 0.003922 | 5.2e-05 | 0.001066 | 0.31078 | 1.3281 | 2.5475 | 100.62 | 21.05 | 25.917 | 84.6367 | 84.2286 | 80.7983 | 0.006433 | 0.99745 | 0.024209 | 24.2714 | 0.06941 | 0.15322 | 0.004712 | 0.33541 | 0.33541 | 0.018867
2019-03-28 15:46 | 1553802347.2755 | 446 | 0.20856 | 0.001131 | 0.000131 | 97.8793 | 96.8354 | 121.5576 | 21.2612 | 0.002635 | 0.002427 | 153.8621 | 0.25698 | 8.4032 | 0.00257 | 0.003774 | 0.003775 | 5.3e-05 | 0.001067 | 0.32803 | 1.304 | 2.5475 | 100.62 | 21.05 | 25.92 | 85.2187 | 83.6567 | 80.2167 | 0.006434 | 0.99909 | 0.024223 | 24.2451 | 0.095906 | 0.15288 | 0.004634 | 0.32995 | 0.32995 | 0.019136
2019-03-28 15:46 | 1553802348.2765 | 447 | 0.20857 | 0.001128 | 0.000132 | 98.0354 | 97.1239 | 121.7733 | 21.2612 | 0.002605 | 0.002474 | 153.8456 | 0.25729 | 8.4053 | 0.00249 | 0.004458 | 0.004366 | 5.4e-05 | 0.001063 | 0.33069 | 1.3038 | 2.5474 | 100.62 | 21.05 | 25.921 | 84.4869 | 81.72 | 81.3364 | 0.006434 | 0.99594 | 0.024206 | 24.305 | 0.083555 | 0.15306 | 0.004595 | 0.32678 | 0.32678 | 0.019344
2019-03-28 15:46 | 1553802349.2756 | 448 | 0.20857 | 0.001124 | 0.000131 | 98.109 | 97.3817 | 121.8773 | 21.2612 | 0.002613 | 0.00241 | 153.8041 | 0.2581 | 8.4052 | 0.002478 | 0.004039 | 0.003945 | 5.6e-05 | 0.001062 | 0.323 | 1.3063 | 2.5473 | 100.62 | 21.05 | 25.922 | 84.6882 | 81.4222 | 79.8088 | 0.006434 | 1.0002 | 0.024269 | 24.2629 | 0.066549 | 0.15354 | 0.004611 | 0.32889 | 0.32889 | 0.019281
2019-03-28 15:46 | 1553802350.277 | 449 | 0.20861 | 0.001124 | 0.000131 | 98.2102 | 97.6049 | 121.4133 | 21.2612 | 0.002632 | 0.002412 | 153.7963 | 0.2576 | 8.4057 | 0.002537 | 0.004379 | 0.004329 | 5e-05 | 0.001064 | 0.33196 | 1.2598 | 2.5473 | 100.62 | 21.05 | 25.925 | 85.1252 | 82.8626 | 79.8579 | 0.006435 | 1.0001 | 0.024221 | 24.2184 | 0.0553 | 0.15325 | 0.004319 | 0.30699 | 0.30699 | 0.020617
2019-03-28 15:46 | 1553802351.2757 | 450 | 0.2086 | 0.001116 | 0.00013 | 98.2817 | 97.6939 | 121.2232 | 21.2612 | 0.002558 | 0.002403 | 153.753 | 0.25686 | 8.4038 | 0.002519 | 0.003491 | 0.003355 | 5.1e-05 | 0.001059 | 0.31902 | 1.2785 | 2.5475 | 100.62 | 21.05 | 25.926 | 83.3572 | 82.4094 | 79.6202 | 0.006435 | 1.0008 | 0.02421 | 24.1912 | 0.053655 | 0.15281 | 0.004418 | 0.3141 | 0.3141 | 0.020092
2019-03-28 15:46 | 1553802352.2757 | 451 | 0.20861 | 0.001115 | 0.000129 | 98.4369 | 97.7697 | 121.3692 | 21.2612 | 0.002603 | 0.002432 | 153.7065 | 0.25811 | 8.4022 | 0.002454 | 0.004469 | 0.004452 | 5.4e-05 | 0.001063 | 0.3454 | 1.3003 | 2.5474 | 100.62 | 21.05 | 25.931 | 84.4455 | 80.867 | 80.321 | 0.006436 | 0.9988 | 0.024201 | 24.2298 | 0.060833 | 0.15355 | 0.004385 | 0.31162 | 0.31162 | 0.02035
2019-03-28 15:46 | 1553802353.2767 | 452 | 0.2086 | 0.001113 | 0.000127 | 98.6043 | 98.0085 | 121.4173 | 21.2612 | 0.00262 | 0.002423 | 153.6879 | 0.25649 | 8.4022 | 0.002459 | 0.003875 | 0.00369 | 5.1e-05 | 0.001064 | 0.31852 | 1.346 | 2.5474 | 100.62 | 21.06 | 25.933 | 84.838 | 80.9768 | 80.1066 | 0.006441 | 0.9994 | 0.024213 | 24.2273 | 0.054208 | 0.15258 | 0.004431 | 0.31527 | 0.31527 | 0.019988
2019-03-28 15:46 | 1553802354.2747 | 453 | 0.20863 | 0.001109 | 0.000125 | 98.6594 | 98.1943 | 121.3511 | 21.2612 | 0.002627 | 0.002383 | 153.6828 | 0.25729 | 8.4038 | 0.002544 | 0.003517 | 0.003365 | 5.2e-05 | 0.001068 | 0.32211 | 1.3301 | 2.5474 | 100.62 | 21.06 | 25.934 | 85.0198 | 83.0232 | 79.1384 | 0.006441 | 1.0021 | 0.024239 | 24.1875 | 0.042262 | 0.15306 | 0.004282 | 0.30481 | 0.30481 | 0.020739
2019-03-28 15:46 | 1553802355.3413 | 454 | 0.20857 | 0.001101 | 0.000124 | 98.8208 | 98.1699 | 121.2957 | 21.2612 | 0.002629 | 0.002449 | 153.6898 | 0.25835 | 8.4019 | 0.002454 | 0.003881 | 0.003852 | 5e-05 | 0.001069 | 0.31238 | 1.323 | 2.5476 | 100.62 | 21.05 | 25.935 | 85.0549 | 80.8615 | 80.7437 | 0.006437 | 0.9976 | 0.024179 | 24.237 | 0.059114 | 0.15369 | 0.004608 | 0.3279 | 0.3279 | 0.019358
2019-03-28 15:46 | 1553802356.2893 | 455 | 0.20858 | 0.001096 | 0.000123 | 98.8134 | 98.087 | 121.7827 | 21.2612 | 0.002555 | 0.002414 | 153.6582 | 0.25643 | 8.4036 | 0.002453 | 0.003985 | 0.00381 | 5.2e-05 | 0.001066 | 0.3183 | 1.3703 | 2.5474 | 100.62 | 21.06 | 25.936 | 83.2949 | 80.8357 | 79.9047 | 0.006442 | 0.99997 | 0.024256 | 24.2568 | 0.065996 | 0.15255 | 0.004565 | 0.32587 | 0.32587 | 0.019334
2019-03-28 15:46 | 1553802357.2887 | 456 | 0.20858 | 0.001094 | 0.000121 | 98.9364 | 98.3642 | 121.6705 | 21.2612 | 0.002641 | 0.002465 | 153.6518 | 0.25793 | 8.4015 | 0.002442 | 0.003667 | 0.003608 | 5.2e-05 | 0.001062 | 0.33107 | 1.3734 | 2.5475 | 100.62 | 21.05 | 25.938 | 85.3484 | 80.5683 | 81.1234 | 0.006438 | 0.99653 | 0.024203 | 24.2874 | 0.05188 | 0.15344 | 0.004576 | 0.32606 | 0.32606 | 0.019436
2019-03-28 15:46 | 1553802358.2883 | 457 | 0.20859 | 0.001095 | 0.000121 | 99.0305 | 98.2524 | 121.4839 | 21.2612 | 0.002656 | 0.00245 | 153.6265 | 0.25756 | 8.4057 | 0.002451 | 0.004429 | 0.004333 | 5.2e-05 | 0.001065 | 0.33445 | 1.3165 | 2.5474 | 100.62 | 21.05 | 25.94 | 85.702 | 80.7889 | 80.7608 | 0.006439 | 0.99755 | 0.024197 | 24.2564 | 0.070555 | 0.15322 | 0.004525 | 0.32228 | 0.32228 | 0.019636
2019-03-28 15:46 | 1553802359.2884 | 458 | 0.20862 | 0.00109 | 0.000121 | 99.1446 | 98.4981 | 121.5166 | 21.2612 | 0.002651 | 0.002448 | 153.6209 | 0.25686 | 8.4028 | 0.00244 | 0.003131 | 0.003058 | 5.1e-05 | 0.001066 | 0.32316 | 1.3375 | 2.5476 | 100.62 | 21.05 | 25.941 | 85.6012 | 80.5099 | 80.7132 | 0.006439 | 0.99769 | 0.024202 | 24.258 | 0.05852 | 0.15281 | 0.004372 | 0.31118 | 0.31118 | 0.020281
2019-03-28 15:46 | 1553802360.2884 | 459 | 0.20854 | 0.001089 | 0.00012 | 99.2337 | 98.6092 | 121.6271 | 21.2612 | 0.002602 | 0.002405 | 153.5957 | 0.25748 | 8.4012 | 0.002446 | 0.00365 | 0.003514 | 4.9e-05 | 0.001068 | 0.31166 | 1.2988 | 2.5474 | 100.62 | 21.05 | 25.943 | 84.4097 | 80.6656 | 79.6863 | 0.006439 | 1.0006 | 0.024248 | 24.2338 | 0.056469 | 0.15318 | 0.004855 | 0.34704 | 0.34704 | 0.018229
2019-03-28 15:46 | 1553802361.2895 | 460 | 0.20858 | 0.001092 | 0.000119 | 99.2858 | 98.6051 | 121.74 | 21.2612 | 0.002666 | 0.002433 | 153.5994 | 0.25739 | 8.4039 | 0.00252 | 0.00394 | 0.003914 | 5.3e-05 | 0.001073 | 0.34346 | 1.305 | 2.5474 | 100.62 | 21.06 | 25.944 | 85.9486 | 82.4484 | 80.3421 | 0.006444 | 0.99874 | 0.024237 | 24.2676 | 0.061535 | 0.15312 | 0.004591 | 0.32772 | 0.32772 | 0.019297
2019-03-28 15:46 | 1553802362.2894 | 461 | 0.20864 | 0.00109 | 0.000118 | 99.3516 | 98.6051 | 121.84 | 21.2612 | 0.002538 | 0.002419 | 153.5957 | 0.25802 | 8.4045 | 0.002453 | 0.004346 | 0.004205 | 5.3e-05 | 0.001071 | 0.29966 | 1.2795 | 2.5474 | 100.62 | 21.06 | 25.945 | 82.8657 | 80.8225 | 80.0045 | 0.006444 | 0.99969 | 0.024258 | 24.2659 | 0.067463 | 0.1535 | 0.004242 | 0.30256 | 0.30256 | 0.020953
2019-03-28 15:46 | 1553802363.2889 | 462 | 0.20863 | 0.001086 | 0.000119 | 99.4287 | 98.724 | 121.7381 | 21.2612 | 0.002636 | 0.002381 | 153.5977 | 0.25718 | 8.4024 | 0.002482 | 0.004609 | 0.004447 | 5.2e-05 | 0.001064 | 0.32421 | 1.2992 | 2.5474 | 100.62 | 21.05 | 25.942 | 85.2406 | 81.5296 | 79.1117 | 0.006439 | 1.0022 | 0.024279 | 24.2251 | 0.063626 | 0.153 | 0.004308 | 0.30766 | 0.30766 | 0.020538
2019-03-28 15:46 | 1553802364.2885 | 463 | 0.20862 | 0.001087 | 0.000119 | 99.5133 | 98.9944 | 121.7625 | 21.2612 | 0.002598 | 0.002466 | 153.5834 | 0.25771 | 8.4013 | 0.002473 | 0.003967 | 0.003903 | 5.4e-05 | 0.001065 | 0.30326 | 1.2526 | 2.5473 | 100.62 | 21.06 | 25.94 | 84.313 | 81.3191 | 81.1429 | 0.006443 | 0.99648 | 0.024212 | 24.2973 | 0.046766 | 0.15331 | 0.004343 | 0.30933 | 0.30933 | 0.02047
2019-03-28 15:46 | 1553802365.2885 | 464 | 0.20858 | 0.001087 | 0.000119 | 99.5351 | 98.9349 | 121.91 | 21.2612 | 0.00252 | 0.002424 | 153.5566 | 0.25738 | 8.4016 | 0.002494 | 0.003187 | 0.003086 | 5.1e-05 | 0.001063 | 0.31708 | 1.3017 | 2.5473 | 100.62 | 21.05 | 25.936 | 82.4407 | 81.8114 | 80.1448 | 0.006438 | 0.99929 | 0.024261 | 24.2777 | 0.054093 | 0.15312 | 0.004563 | 0.32597 | 0.32597 | 0.0194
2019-03-28 15:46 | 1553802366.2886 | 465 | 0.20862 | 0.00108 | 0.00012 | 99.5655 | 98.9312 | 121.8398 | 21.2612 | 0.00262 | 0.002424 | 153.5637 | 0.2573 | 8.4009 | 0.002509 | 0.004477 | 0.004379 | 5.2e-05 | 0.001067 | 0.33035 | 1.2943 | 2.5476 | 100.62 | 21.05 | 25.933 | 84.8481 | 82.1669 | 80.1299 | 0.006437 | 0.99934 | 0.024254 | 24.2702 | 0.057161 | 0.15307 | 0.004341 | 0.30963 | 0.30963 | 0.020417
2019-03-28 15:46 | 1553802367.2888 | 466 | 0.20862 | 0.001078 | 0.000121 | 99.5729 | 98.9346 | 121.96 | 21.2612 | 0.002557 | 0.002408 | 153.5308 | 0.25845 | 8.4024 | 0.002559 | 0.003838 | 0.003727 | 5e-05 | 0.001061 | 0.34916 | 1.2867 | 2.5474 | 100.62 | 21.05 | 25.931 | 83.3206 | 83.3798 | 79.7589 | 0.006436 | 1.0004 | 0.024279 | 24.2694 | 0.057523 | 0.15375 | 0.004344 | 0.31012 | 0.31012 | 0.020476
2019-03-28 15:46 | 1553802368.2897 | 467 | 0.20864 | 0.001075 | 0.000121 | 99.6801 | 98.991 | 121.9546 | 21.2612 | 0.002575 | 0.002381 | 153.5378 | 0.25675 | 8.4045 | 0.0026 | 0.004004 | 0.003978 | 5.1e-05 | 0.001058 | 0.32964 | 1.3167 | 2.5475 | 100.62 | 21.05 | 25.923 | 83.7645 | 84.3582 | 79.1134 | 0.006434 | 1.0022 | 0.024301 | 24.2467 | 0.062052 | 0.15274 | 0.004267 | 0.3048 | 0.3048 | 0.020696
2019-03-28 15:46 | 1553802369.2886 | 468 | 0.20861 | 0.001075 | 0.00012 | 99.6995 | 99.0233 | 121.7651 | 21.2612 | 0.002596 | 0.002406 | 153.512 | 0.25821 | 8.4036 | 0.002537 | 0.003635 | 0.003447 | 5.2e-05 | 0.001059 | 0.31542 | 1.3307 | 2.5474 | 100.62 | 21.06 | 25.92 | 84.2716 | 82.8446 | 79.6981 | 0.006438 | 1.0006 | 0.024262 | 24.2479 | 0.060874 | 0.15361 | 0.004408 | 0.31465 | 0.31465 | 0.020162
2019-03-28 15:46 | 1553802370.2877 | 469 | 0.20865 | 0.001068 | 0.00012 | 99.6938 | 99.007 | 121.813 | 21.2612 | 0.002555 | 0.002443 | 153.5161 | 0.25783 | 8.4031 | 0.002559 | 0.0038 | 0.003707 | 5.5e-05 | 0.001052 | 0.30414 | 1.301 | 2.5473 | 100.62 | 21.06 | 25.912 | 83.2862 | 83.3798 | 80.6026 | 0.006436 | 0.998 | 0.024235 | 24.2838 | 0.061839 | 0.15338 | 0.004196 | 0.29885 | 0.29885 | 0.021197
2019-03-28 15:46 | 1553802371.2887 | 470 | 0.20864 | 0.001068 | 0.000119 | 99.7671 | 99.1877 | 121.6644 | 21.2612 | 0.002578 | 0.002376 | 153.4948 | 0.25691 | 8.4027 | 0.002527 | 0.004256 | 0.004093 | 5.1e-05 | 0.001053 | 0.32659 | 1.3119 | 2.5474 | 100.62 | 21.06 | 25.909 | 83.8486 | 82.603 | 78.9801 | 0.006435 | 1.0026 | 0.024276 | 24.2132 | 0.052098 | 0.15284 | 0.004235 | 0.30221 | 0.30221 | 0.020887
2019-03-28 15:46 | 1553802372.2881 | 471 | 0.20866 | 0.001065 | 0.000121 | 99.8088 | 99.2489 | 121.6503 | 21.2612 | 0.002626 | 0.002456 | 153.4558 | 0.25768 | 8.4027 | 0.002551 | 0.004212 | 0.004164 | 5.5e-05 | 0.001054 | 0.33052 | 1.3169 | 2.5475 | 100.62 | 21.06 | 25.907 | 84.9833 | 83.1949 | 80.9011 | 0.006434 | 0.99716 | 0.024209 | 24.2778 | 0.050323 | 0.15329 | 0.004147 | 0.29491 | 0.29491 | 0.021468
2019-03-28 15:46 | 1553802373.2887 | 472 | 0.20867 | 0.00106 | 0.000123 | 99.9303 | 99.131 | 121.7637 | 21.2612 | 0.002558 | 0.002358 | 153.4176 | 0.25763 | 8.4055 | 0.002539 | 0.004118 | 0.004026 | 5.2e-05 | 0.001057 | 0.3361 | 1.3419 | 2.5475 | 100.62 | 21.06 | 25.898 | 83.3665 | 82.9102 | 78.5427 | 0.006432 | 1.0038 | 0.024301 | 24.2081 | 0.071837 | 0.15327 | 0.004079 | 0.29096 | 0.29096 | 0.021755
2019-03-28 15:46 | 1553802374.2878 | 473 | 0.2087 | 0.001058 | 0.000124 | 99.9348 | 99.3729 | 121.694 | 21.2612 | 0.00262 | 0.00241 | 153.4038 | 0.25752 | 8.4035 | 0.002437 | 0.003951 | 0.003808 | 5.2e-05 | 0.001058 | 0.31852 | 1.2968 | 2.5473 | 100.62 | 21.06 | 25.895 | 84.8418 | 80.4481 | 79.8019 | 0.006431 | 1.0003 | 0.024251 | 24.2444 | 0.050438 | 0.1532 | 0.00391 | 0.27791 | 0.27791 | 0.022766
2019-03-28 15:46 | 1553802375.2879 | 474 | 0.20871 | 0.001054 | 0.000124 | 99.9125 | 99.3501 | 121.5335 | 21.2612 | 0.002497 | 0.002355 | 153.401 | 0.25759 | 8.4037 | 0.002519 | 0.003885 | 0.003831 | 5.4e-05 | 0.001057 | 0.3366 | 1.2862 | 2.5473 | 100.62 | 21.06 | 25.893 | 81.8969 | 82.4195 | 78.4819 | 0.006431 | 1.004 | 0.02428 | 24.1831 | 0.050488 | 0.15324 | 0.00385 | 0.27385 | 0.27385 | 0.023111
2019-03-28 15:46 | 1553802376.2889 | 475 | 0.20864 | 0.001049 | 0.000124 | 99.9509 | 99.3525 | 121.3868 | 21.2612 | 0.002543 | 0.002414 | 153.3693 | 0.25803 | 8.4005 | 0.002494 | 0.004057 | 0.003965 | 5.5e-05 | 0.00106 | 0.32781 | 1.283 | 2.5473 | 100.62 | 21.06 | 25.891 | 82.996 | 81.8198 | 79.8874 | 0.00643 | 1 | 0.024217 | 24.2167 | 0.053713 | 0.1535 | 0.004244 | 0.30191 | 0.30191 | 0.020999
2019-03-28 15:46 | 1553802377.288 | 476 | 0.20869 | 0.001042 | 0.000124 | 99.9988 | 99.2638 | 121.6717 | 21.2612 | 0.002548 | 0.002379 | 153.3671 | 0.25725 | 8.4015 | 0.002539 | 0.004245 | 0.004117 | 5.4e-05 | 0.001057 | 0.31022 | 1.2767 | 2.5474 | 100.62 | 21.06 | 25.888 | 83.1177 | 82.9088 | 79.0541 | 0.00643 | 1.0024 | 0.024274 | 24.2165 | 0.065989 | 0.15304 | 0.003967 | 0.28234 | 0.28234 | 0.022386
2019-03-28 15:46 | 1553802378.2889 | 477 | 0.20878 | 0.001038 | 0.000123 | 99.9982 | 99.3216 | 121.9043 | 21.2612 | 0.002547 | 0.002374 | 153.3333 | 0.25704 | 8.4039 | 0.002482 | 0.003493 | 0.00348 | 5.5e-05 | 0.00106 | 0.33467 | 1.2902 | 2.5476 | 100.62 | 21.06 | 25.883 | 83.0888 | 81.5344 | 78.9457 | 0.006428 | 1.0027 | 0.024301 | 24.2359 | 0.060733 | 0.15291 | 0.003422 | 0.24285 | 0.24285 | 0.026005
2019-03-28 15:46 | 1553802379.2891 | 478 | 0.20867 | 0.001037 | 0.000121 | 100.1111 | 99.7129 | 121.6388 | 21.2612 | 0.00254 | 0.002377 | 153.3266 | 0.25858 | 8.4024 | 0.002509 | 0.004236 | 0.004065 | 5.5e-05 | 0.001056 | 0.31188 | 1.3266 | 2.5473 | 100.62 | 21.07 | 25.882 | 82.9304 | 82.1693 | 79.0034 | 0.006432 | 1.0025 | 0.024273 | 24.2115 | 0.03565 | 0.15383 | 0.004132 | 0.29457 | 0.29457 | 0.021568
2019-03-28 15:46 | 1553802380.288 | 479 | 0.20869 | 0.00103 | 0.000119 | 100.0933 | 99.4636 | 121.6222 | 21.2612 | 0.002522 | 0.002389 | 153.3137 | 0.25752 | 8.4048 | 0.002482 | 0.004573 | 0.004504 | 5.5e-05 | 0.001059 | 0.3095 | 1.3326 | 2.5473 | 100.62 | 21.07 | 25.882 | 82.4952 | 81.5181 | 79.2927 | 0.006432 | 1.0017 | 0.024261 | 24.2198 | 0.056449 | 0.1532 | 0.003978 | 0.28333 | 0.28333 | 0.022332
2019-03-28 15:46 | 1553802381.288 | 480 | 0.20872 | 0.001026 | 0.000117 | 100.1713 | 99.5274 | 121.9773 | 21.2612 | 0.002572 | 0.002374 | 153.2849 | 0.25763 | 8.4047 | 0.002487 | 0.004031 | 0.0039 | 5.5e-05 | 0.001055 | 0.30143 | 1.3508 | 2.5473 | 100.62 | 21.07 | 25.883 | 83.6911 | 81.6529 | 78.9418 | 0.006432 | 1.0027 | 0.024309 | 24.243 | 0.057678 | 0.15327 | 0.003816 | 0.27216 | 0.27216 | 0.023258
2019-03-28 15:46 | 1553802382.2887 | 481 | 0.2087 | 0.001024 | 0.000115 | 100.0656 | 99.5889 | 121.7306 | 21.2612 | 0.002554 | 0.00241 | 153.2909 | 0.25789 | 8.4042 | 0.002529 | 0.004042 | 0.003996 | 5.3e-05 | 0.001058 | 0.34667 | 1.314 | 2.5473 | 100.62 | 21.07 | 25.884 | 83.2526 | 82.6659 | 79.805 | 0.006433 | 1.0003 | 0.024254 | 24.2482 | 0.042707 | 0.15342 | 0.003966 | 0.28261 | 0.28261 | 0.02242
2019-03-28 15:46 | 1553802383.2891 | 482 | 0.20871 | 0.001023 | 0.000113 | 100.2029 | 99.7302 | 121.8745 | 21.2612 | 0.002596 | 0.002365 | 153.2608 | 0.25754 | 8.4017 | 0.002425 | 0.003795 | 0.003683 | 5.4e-05 | 0.00106 | 0.31045 | 1.3143 | 2.5475 | 100.62 | 21.07 | 25.883 | 84.2598 | 80.1535 | 78.7282 | 0.006432 | 1.0033 | 0.024306 | 24.2255 | 0.042288 | 0.15321 | 0.003923 | 0.28012 | 0.28012 | 0.022589
2019-03-28 15:46 | 1553802384.2882 | 483 | 0.20871 | 0.001025 | 0.000112 | 100.272 | 99.6861 | 122.1125 | 21.2612 | 0.002541 | 0.002345 | 153.2519 | 0.2572 | 8.4021 | 0.002486 | 0.003541 | 0.003439 | 5.3e-05 | 0.001061 | 0.31277 | 1.3224 | 2.5475 | 100.62 | 21.07 | 25.884 | 82.9585 | 81.6154 | 78.2441 | 0.006433 | 1.0047 | 0.024346 | 24.2324 | 0.052416 | 0.15301 | 0.003893 | 0.27844 | 0.27844 | 0.022695
2019-03-28 15:46 | 1553802385.2876 | 484 | 0.20873 | 0.001024 | 0.000112 | 100.3506 | 99.6674 | 121.9994 | 21.2612 | 0.002554 | 0.002365 | 153.2196 | 0.25764 | 8.4013 | 0.002447 | 0.00423 | 0.004104 | 5.3e-05 | 0.001059 | 0.3095 | 1.2887 | 2.5474 | 100.62 | 21.07 | 25.883 | 83.2494 | 80.6976 | 78.7181 | 0.006432 | 1.0033 | 0.024319 | 24.2375 | 0.061105 | 0.15327 | 0.003798 | 0.27123 | 0.27123 | 0.023338
2019-03-28 15:46 | 1553802386.2882 | 485 | 0.20871 | 0.001022 | 0.000111 | 100.2651 | 99.7452 | 122.066 | 21.2612 | 0.002478 | 0.002371 | 153.2263 | 0.25795 | 8.4053 | 0.002441 | 0.004456 | 0.004336 | 5.4e-05 | 0.001065 | 0.31061 | 1.3338 | 2.5475 | 100.62 | 21.07 | 25.881 | 81.4361 | 80.5377 | 78.8553 | 0.006432 | 1.003 | 0.024321 | 24.2489 | 0.046503 | 0.15345 | 0.003899 | 0.2787 | 0.2787 | 0.022739
2019-03-28 15:46 | 1553802387.2886 | 486 | 0.20871 | 0.001023 | 0.000112 | 100.2973 | 99.7523 | 122.2395 | 21.2612 | 0.002537 | 0.002357 | 153.2085 | 0.25766 | 8.4029 | 0.002472 | 0.004175 | 0.004056 | 5.5e-05 | 0.00107 | 0.32621 | 1.3098 | 2.5475 | 100.62 | 21.07 | 25.882 | 82.8626 | 81.2833 | 78.5216 | 0.006432 | 1.0039 | 0.024349 | 24.2546 | 0.048732 | 0.15328 | 0.003888 | 0.27819 | 0.27819 | 0.022756
2019-03-28 15:46 | 1553802388.2887 | 487 | 0.20872 | 0.001022 | 0.000112 | 100.4027 | 99.8281 | 122.3024 | 21.2612 | 0.002537 | 0.002397 | 153.1945 | 0.25883 | 8.4013 | 0.002462 | 0.005011 | 0.004944 | 5.1e-05 | 0.001073 | 0.30702 | 1.3172 | 2.5474 | 100.62 | 21.07 | 25.88 | 82.8619 | 81.0602 | 79.4782 | 0.006432 | 1.0012 | 0.024323 | 24.2938 | 0.05134 | 0.15398 | 0.00383 | 0.27359 | 0.27359 | 0.023244
2019-03-28 15:46 | 1553802389.2873 | 488 | 0.20873 | 0.001015 | 0.000112 | 100.4274 | 99.7815 | 122.6817 | 21.2612 | 0.002523 | 0.002361 | 153.2067 | 0.25789 | 8.402 | 0.002488 | 0.003791 | 0.003771 | 5.3e-05 | 0.001073 | 0.31365 | 1.3356 | 2.5475 | 100.62 | 21.07 | 25.879 | 82.521 | 81.6832 | 78.6285 | 0.006431 | 1.0036 | 0.02439 | 24.3021 | 0.057713 | 0.15342 | 0.003788 | 0.27126 | 0.27126 | 0.023358
2019-03-28 15:46 | 1553802390.2887 | 489 | 0.20871 | 0.001019 | 0.000112 | 100.4206 | 99.8722 | 122.7514 | 21.2612 | 0.002592 | 0.00238 | 153.1798 | 0.25815 | 8.4022 | 0.002506 | 0.004151 | 0.003974 | 5e-05 | 0.001077 | 0.32139 | 1.3163 | 2.5475 | 100.62 | 21.07 | 25.879 | 84.1841 | 82.1029 | 79.0791 | 0.006431 | 1.0023 | 0.024381 | 24.3246 | 0.048976 | 0.15357 | 0.003901 | 0.27949 | 0.27949 | 0.022693
2019-03-28 15:47 | 1553802391.2874 | 490 | 0.20872 | 0.001012 | 0.000112 | 100.4179 | 99.9171 | 122.6104 | 21.2612 | 0.002556 | 0.002357 | 153.1854 | 0.25825 | 8.4022 | 0.002461 | 0.004971 | 0.004823 | 5.3e-05 | 0.001069 | 0.34003 | 1.2868 | 2.5474 | 100.62 | 21.07 | 25.875 | 83.304 | 81.0345 | 78.531 | 0.00643 | 1.0039 | 0.024386 | 24.2917 | 0.044721 | 0.15363 | 0.003869 | 0.2772 | 0.2772 | 0.022889
2019-03-28 15:47 | 1553802392.2874 | 491 | 0.20872 | 0.001013 | 0.000112 | 100.4745 | 99.9704 | 122.3719 | 21.2612 | 0.002515 | 0.002391 | 153.172 | 0.2575 | 8.4025 | 0.002472 | 0.004024 | 0.004016 | 5.1e-05 | 0.001068 | 0.29894 | 1.2995 | 2.5473 | 100.62 | 21.07 | 25.873 | 82.3267 | 81.284 | 79.3364 | 0.00643 | 1.0016 | 0.024334 | 24.2958 | 0.044988 | 0.15319 | 0.003866 | 0.27637 | 0.27637 | 0.022892
2019-03-28 15:47 | 1553802393.2885 | 492 | 0.20872 | 0.001011 | 0.000113 | 100.4965 | 100.0173 | 122.3928 | 21.2612 | 0.002491 | 0.002369 | 153.1667 | 0.25834 | 8.4018 | 0.002524 | 0.003943 | 0.003812 | 5.3e-05 | 0.001066 | 0.33085 | 1.3341 | 2.5474 | 100.62 | 21.07 | 25.87 | 81.7558 | 82.5311 | 78.8227 | 0.006429 | 1.003 | 0.024354 | 24.2802 | 0.042756 | 0.15369 | 0.003825 | 0.27353 | 0.27353 | 0.023206
2019-03-28 15:47 | 1553802394.2878 | 493 | 0.20869 | 0.001007 | 0.000115 | 100.5338 | 99.8237 | 122.1419 | 21.2612 | 0.002536 | 0.002349 | 153.1313 | 0.25799 | 8.4031 | 0.002495 | 0.004287 | 0.004181 | 5.4e-05 | 0.001069 | 0.32748 | 1.3194 | 2.5474 | 100.62 | 21.07 | 25.869 | 82.8323 | 81.8323 | 78.3423 | 0.006429 | 1.0044 | 0.024346 | 24.2387 | 0.063403 | 0.15348 | 0.004033 | 0.28859 | 0.28859 | 0.021964
2019-03-28 15:47 | 1553802395.2884 | 494 | 0.20873 | 0.001002 | 0.000114 | 100.5391 | 99.7656 | 122.1792 | 21.2612 | 0.002553 | 0.002387 | 153.1215 | 0.25846 | 8.4029 | 0.00254 | 0.003797 | 0.003656 | 5.1e-05 | 0.001063 | 0.30613 | 1.3048 | 2.5473 | 100.61 | 21.07 | 25.869 | 83.2247 | 82.9335 | 79.2506 | 0.00643 | 1.0018 | 0.024318 | 24.2737 | 0.069087 | 0.15376 | 0.003795 | 0.27085 | 0.27085 | 0.023445
2019-03-28 15:47 | 1553802396.2875 | 495 | 0.20869 | 0.001001 | 0.000113 | 100.5517 | 99.9687 | 122.1428 | 21.2612 | 0.002507 | 0.002399 | 153.0795 | 0.25709 | 8.4034 | 0.002512 | 0.00397 | 0.003901 | 4.9e-05 | 0.001062 | 0.32477 | 1.3359 | 2.5474 | 100.61 | 21.07 | 25.87 | 82.1311 | 82.2558 | 79.5359 | 0.00643 | 1.001 | 0.024305 | 24.2799 | 0.052004 | 0.15294 | 0.004032 | 0.2881 | 0.2881 | 0.021925
2019-03-28 15:47 | 1553802397.289 | 496 | 0.20873 | 0.001002 | 0.000112 | 100.6175 | 99.9436 | 121.6905 | 21.2612 | 0.002553 | 0.002318 | 153.0591 | 0.25819 | 8.4039 | 0.00256 | 0.003779 | 0.003701 | 4.9e-05 | 0.001068 | 0.33539 | 1.3257 | 2.5473 | 100.61 | 21.07 | 25.87 | 83.2456 | 83.4117 | 77.5855 | 0.00643 | 1.0066 | 0.024327 | 24.1678 | 0.060108 | 0.1536 | 0.003836 | 0.27416 | 0.27416 | 0.023138
2019-03-28 15:47 | 1553802398.2877 | 497 | 0.20873 | 0.001003 | 0.00011 | 100.6428 | 99.9436 | 121.7069 | 21.2612 | 0.002513 | 0.002377 | 153.0354 | 0.25818 | 8.4026 | 0.002544 | 0.00432 | 0.004264 | 4.9e-05 | 0.001064 | 0.28888 | 1.2997 | 2.5474 | 100.61 | 21.08 | 25.871 | 82.2729 | 83.0117 | 79.0159 | 0.006434 | 1.0025 | 0.024279 | 24.2187 | 0.062358 | 0.15359 | 0.003829 | 0.27316 | 0.27316 | 0.023222
2019-03-28 15:47 | 1553802399.2877 | 498 | 0.20873 | 0.000999 | 0.000109 | 100.6365 | 99.984 | 121.2865 | 21.2612 | 0.002477 | 0.002347 | 152.9892 | 0.25738 | 8.4027 | 0.002517 | 0.003554 | 0.00347 | 4.9e-05 | 0.001067 | 0.3361 | 1.3392 | 2.5474 | 100.61 | 21.08 | 25.873 | 81.4025 | 82.3587 | 78.2792 | 0.006435 | 1.0046 | 0.024263 | 24.1515 | 0.058185 | 0.15312 | 0.003826 | 0.27287 | 0.27287 | 0.023174
2019-03-28 15:47 | 1553802400.2871 | 499 | 0.20874 | 0.000994 | 0.000108 | 100.6583 | 99.9912 | 121.2922 | 21.2612 | 0.00253 | 0.002361 | 152.9751 | 0.25815 | 8.4007 | 0.002501 | 0.003883 | 0.00376 | 4.5e-05 | 0.001066 | 0.30829 | 1.344 | 2.5475 | 100.61 | 21.08 | 25.875 | 82.6926 | 81.9907 | 78.6268 | 0.006435 | 1.0036 | 0.024251 | 24.164 | 0.059479 | 0.15357 | 0.003759 | 0.26786 | 0.26786 | 0.023678
2019-03-28 15:47 | 1553802401.2877 | 500 | 0.20871 | 0.000995 | 0.000107 | 100.6589 | 100.0516 | 121.2685 | 21.2612 | 0.002464 | 0.002354 | 152.9576 | 0.25821 | 8.4026 | 0.002497 | 0.004135 | 0.004036 | 4.7e-05 | 0.001061 | 0.31293 | 1.3077 | 2.5475 | 100.61 | 21.08 | 25.88 | 81.1023 | 81.89 | 78.449 | 0.006436 | 1.0041 | 0.024255 | 24.1556 | 0.054124 | 0.15361 | 0.003934 | 0.2808 | 0.2808 | 0.022593
2019-03-28 15:47 | 1553802402.2881 | 501 | 0.20873 | 0.000996 | 0.000106 | 100.7042 | 100.0571 | 121.5593 | 21.2612 | 0.002508 | 0.002369 | 152.9463 | 0.25821 | 8.4054 | 0.002497 | 0.003475 | 0.003348 | 4.6e-05 | 0.001057 | 0.35402 | 1.3195 | 2.5475 | 100.61 | 21.08 | 25.883 | 82.1443 | 81.8868 | 78.8039 | 0.006437 | 1.0031 | 0.024272 | 24.1967 | 0.05766 | 0.15361 | 0.003819 | 0.27263 | 0.27263 | 0.02327
2019-03-28 15:47 | 1553802403.2872 | 502 | 0.20872 | 0.000996 | 0.000105 | 100.7012 | 100.0717 | 121.4267 | 21.2612 | 0.002542 | 0.002384 | 152.9339 | 0.25784 | 8.4027 | 0.002477 | 0.00433 | 0.004212 | 4.8e-05 | 0.001057 | 0.30901 | 1.3155 | 2.5475 | 100.61 | 21.08 | 25.885 | 82.971 | 81.4198 | 79.1641 | 0.006438 | 1.0021 | 0.024246 | 24.1959 | 0.05609 | 0.15339 | 0.003894 | 0.27788 | 0.27788 | 0.022798
2019-03-28 15:47 | 1553802404.2868 | 503 | 0.20871 | 0.000997 | 0.000104 | 100.7301 | 100.0537 | 121.6464 | 21.2612 | 0.002504 | 0.002357 | 152.9381 | 0.25801 | 8.402 | 0.002489 | 0.004782 | 0.004704 | 4.6e-05 | 0.001055 | 0.31924 | 1.2674 | 2.5474 | 100.61 | 21.08 | 25.889 | 82.0498 | 81.7037 | 78.5317 | 0.006439 | 1.0039 | 0.02429 | 24.196 | 0.060265 | 0.15349 | 0.00398 | 0.28472 | 0.28472 | 0.022265
2019-03-28 15:47 | 1553802405.2868 | 504 | 0.20875 | 0.000997 | 0.000103 | 100.7146 | 100.1967 | 121.6844 | 21.2612 | 0.002543 | 0.002363 | 152.9176 | 0.25785 | 8.4039 | 0.002524 | 0.004478 | 0.004349 | 4.7e-05 | 0.001051 | 0.32892 | 1.3083 | 2.5474 | 100.61 | 21.08 | 25.89 | 82.9881 | 82.5484 | 78.6783 | 0.006439 | 1.0035 | 0.024289 | 24.2048 | 0.046113 | 0.15339 | 0.003719 | 0.26565 | 0.26565 | 0.023847
2019-03-28 15:47 | 1553802406.2868 | 505 | 0.20867 | 0.000997 | 0.000102 | 100.7107 | 100.0567 | 121.694 | 21.2612 | 0.002538 | 0.002301 | 152.8975 | 0.25785 | 8.4026 | 0.002585 | 0.003741 | 0.003719 | 4.6e-05 | 0.00105 | 0.31542 | 1.2849 | 2.5474 | 100.61 | 21.08 | 25.893 | 82.8782 | 84.0165 | 77.1823 | 0.00644 | 1.0077 | 0.024341 | 24.1543 | 0.058272 | 0.15339 | 0.004192 | 0.30096 | 0.30096 | 0.02105
2019-03-28 15:47 | 1553802407.2883 | 506 | 0.20872 | 0.000995 | 0.000102 | 100.7652 | 100.1719 | 122.0456 | 21.2612 | 0.002542 | 0.002341 | 152.8929 | 0.25751 | 8.4012 | 0.002531 | 0.003757 | 0.003733 | 4.6e-05 | 0.00105 | 0.3194 | 1.2915 | 2.5473 | 100.61 | 21.08 | 25.894 | 82.9734 | 82.7183 | 78.1357 | 0.00644 | 1.005 | 0.024343 | 24.222 | 0.052823 | 0.15319 | 0.003914 | 0.28062 | 0.28062 | 0.022546
2019-03-28 15:47 | 1553802408.2869 | 507 | 0.20871 | 0.000996 | 0.0001 | 100.7962 | 100.0673 | 121.8377 | 21.2612 | 0.00246 | 0.002322 | 152.8762 | 0.25767 | 8.4028 | 0.002473 | 0.003844 | 0.003728 | 4.9e-05 | 0.00104 | 0.34418 | 1.3418 | 2.5473 | 100.61 | 21.08 | 25.894 | 81.0049 | 81.316 | 77.6751 | 0.00644 | 1.0063 | 0.024339 | 24.1855 | 0.06492 | 0.15329 | 0.003968 | 0.28465 | 0.28465 | 0.02224
2019-03-28 15:47 | 1553802409.288 | 508 | 0.20875 | 0.000999 | 9.9e-05 | 100.7768 | 100.0526 | 122.132 | 21.2612 | 0.002472 | 0.002318 | 152.8626 | 0.25783 | 8.4048 | 0.002571 | 0.003872 | 0.003817 | 4.7e-05 | 0.001045 | 0.323 | 1.3092 | 2.5475 | 100.61 | 21.08 | 25.897 | 81.2826 | 83.6801 | 77.5915 | 0.006441 | 1.0066 | 0.024371 | 24.2118 | 0.064508 | 0.15338 | 0.003735 | 0.26801 | 0.26801 | 0.023636
2019-03-28 15:47 | 1553802410.288 | 509 | 0.20868 | 0.000994 | 9.7e-05 | 100.7167 | 100.001 | 122.0602 | 21.2612 | 0.00254 | 0.002335 | 152.8524 | 0.25806 | 8.4025 | 0.002481 | 0.004011 | 0.003958 | 4.7e-05 | 0.001044 | 0.32997 | 1.332 | 2.5474 | 100.61 | 21.08 | 25.897 | 82.9335 | 81.4955 | 77.9947 | 0.006441 | 1.0054 | 0.02435 | 24.2186 | 0.063784 | 0.15352 | 0.004144 | 0.29783 | 0.29783 | 0.021289
2019-03-28 15:47 | 1553802411.288 | 510 | 0.20871 | 0.000998 | 9.7e-05 | 100.7664 | 99.9344 | 121.8341 | 21.2612 | 0.002517 | 0.002365 | 152.8329 | 0.25811 | 8.4012 | 0.002529 | 0.004451 | 0.004284 | 5e-05 | 0.00104 | 0.30879 | 1.2925 | 2.5473 | 100.61 | 21.08 | 25.894 | 82.3635 | 82.6599 | 78.7244 | 0.00644 | 1.0033 | 0.024302 | 24.2213 | 0.074158 | 0.15355 | 0.003953 | 0.28332 | 0.28332 | 0.022384
2019-03-28 15:47 | 1553802412.288 | 511 | 0.20875 | 0.00099 | 9.7e-05 | 100.7286 | 100.0693 | 121.8191 | 21.2612 | 0.002533 | 0.002283 | 152.8086 | 0.25893 | 8.3991 | 0.002485 | 0.003977 | 0.003893 | 5e-05 | 0.001041 | 0.32372 | 1.2889 | 2.5473 | 100.61 | 21.08 | 25.892 | 82.7628 | 81.6099 | 76.7357 | 0.006439 | 1.009 | 0.024369 | 24.1513 | 0.058736 | 0.15404 | 0.003754 | 0.26946 | 0.26946 | 0.023609
2019-03-28 15:47 | 1553802413.2881 | 512 | 0.20871 | 0.000994 | 9.7e-05 | 100.7012 | 99.9901 | 121.829 | 21.2612 | 0.002537 | 0.002414 | 152.803 | 0.25857 | 8.4025 | 0.002479 | 0.004381 | 0.004269 | 5.1e-05 | 0.001044 | 0.33771 | 1.3304 | 2.5474 | 100.61 | 21.08 | 25.889 | 82.8501 | 81.4674 | 79.8905 | 0.006439 | 1 | 0.024261 | 24.2609 | 0.06338 | 0.15382 | 0.00398 | 0.28478 | 0.28478 | 0.022308
2019-03-28 15:47 | 1553802414.2871 | 513 | 0.20872 | 0.000991 | 9.7e-05 | 100.7152 | 100.1349 | 121.8557 | 21.2612 | 0.002482 | 0.002366 | 152.7723 | 0.2578 | 8.4021 | 0.002427 | 0.003798 | 0.003716 | 5.1e-05 | 0.001046 | 0.32516 | 1.2909 | 2.5474 | 100.61 | 21.08 | 25.879 | 81.5397 | 80.2049 | 78.7328 | 0.006436 | 1.0033 | 0.024304 | 24.2238 | 0.051687 | 0.15336 | 0.003919 | 0.28082 | 0.28082 | 0.022555
2019-03-28 15:47 | 1553802415.2875 | 514 | 0.20871 | 0.000992 | 9.6e-05 | 100.7625 | 100.034 | 121.7996 | 21.2612 | 0.002484 | 0.002351 | 152.7603 | 0.25785 | 8.4036 | 0.00246 | 0.004208 | 0.004126 | 5.2e-05 | 0.001049 | 0.33085 | 1.3328 | 2.5474 | 100.61 | 21.09 | 25.872 | 81.5695 | 80.9963 | 78.3844 | 0.006438 | 1.0043 | 0.02431 | 24.2062 | 0.064909 | 0.15339 | 0.003946 | 0.28294 | 0.28294 | 0.022391
2019-03-28 15:47 | 1553802416.2862 | 515 | 0.20871 | 0.00099 | 9.4e-05 | 100.7691 | 100.055 | 121.9496 | 21.2612 | 0.002532 | 0.002324 | 152.7836 | 0.25752 | 8.3989 | 0.002459 | 0.003846 | 0.003747 | 5.2e-05 | 0.00105 | 0.31526 | 1.3735 | 2.5475 | 100.61 | 21.09 | 25.857 | 82.7301 | 80.9824 | 77.7304 | 0.006435 | 1.0062 | 0.024348 | 24.1985 | 0.063608 | 0.1532 | 0.003978 | 0.28582 | 0.28582 | 0.022137
2019-03-28 15:47 | 1553802417.2862 | 516 | 0.20871 | 0.000991 | 9.3e-05 | 100.7319 | 99.8169 | 121.7515 | 21.2612 | 0.002489 | 0.002359 | 152.7515 | 0.25713 | 8.4056 | 0.002525 | 0.003647 | 0.003593 | 4.9e-05 | 0.00105 | 0.31598 | 1.3598 | 2.5475 | 100.61 | 21.09 | 25.846 | 81.7092 | 82.5686 | 78.5653 | 0.006432 | 1.0038 | 0.024299 | 24.2076 | 0.081617 | 0.15297 | 0.003964 | 0.28431 | 0.28431 | 0.02222
2019-03-28 15:47 | 1553802418.2877 | 517 | 0.20872 | 0.000996 | 9.3e-05 | 100.7506 | 99.9185 | 121.4842 | 21.2612 | 0.002469 | 0.002308 | 152.7741 | 0.25787 | 8.4016 | 0.002529 | 0.003669 | 0.003509 | 4.8e-05 | 0.001055 | 0.33284 | 1.3096 | 2.5473 | 100.61 | 21.09 | 25.837 | 81.2193 | 82.6599 | 77.3407 | 0.00643 | 1.0073 | 0.024315 | 24.1389 | 0.074186 | 0.15341 | 0.003913 | 0.28076 | 0.28076 | 0.022566
2019-03-28 15:47 | 1553802419.2873 | 518 | 0.20872 | 0.000994 | 9.3e-05 | 100.666 | 100.0934 | 121.6229 | 21.2612 | 0.002573 | 0.002357 | 152.7472 | 0.25728 | 8.4033 | 0.002591 | 0.00381 | 0.003625 | 5e-05 | 0.001051 | 0.31404 | 1.2937 | 2.5474 | 100.61 | 21.09 | 25.828 | 83.7058 | 84.1528 | 78.524 | 0.006427 | 1.0039 | 0.024288 | 24.1934 | 0.051024 | 0.15305 | 0.003917 | 0.28074 | 0.28074 | 0.022516
2019-03-28 15:47 | 1553802420.2873 | 519 | 0.20873 | 0.000989 | 9.3e-05 | 100.7226 | 99.8519 | 121.618 | 21.2612 | 0.002458 | 0.002365 | 152.7337 | 0.25723 | 8.4038 | 0.002558 | 0.004241 | 0.004227 | 5.1e-05 | 0.00105 | 0.32947 | 1.3335 | 2.5475 | 100.61 | 21.09 | 25.819 | 80.9417 | 83.3682 | 78.708 | 0.006425 | 1.0034 | 0.024281 | 24.1993 | 0.077662 | 0.15303 | 0.003837 | 0.27478 | 0.27478 | 0.023
2019-03-28 15:47 | 1553802421.2867 | 520 | 0.20874 | 0.00099 | 9.4e-05 | 100.6231 | 99.9212 | 121.5759 | 21.2612 | 0.002493 | 0.002377 | 152.7032 | 0.25899 | 8.4031 | 0.002512 | 0.003418 | 0.00331 | 4.7e-05 | 0.001048 | 0.31686 | 1.3263 | 2.5475 | 100.61 | 21.09 | 25.802 | 81.8051 | 82.252 | 79.0065 | 0.006421 | 1.0025 | 0.024266 | 24.2053 | 0.062618 | 0.15407 | 0.003764 | 0.26922 | 0.26922 | 0.023636
2019-03-28 15:47 | 1553802422.2874 | 521 | 0.20874 | 0.000987 | 9.4e-05 | 100.7119 | 99.9647 | 121.7698 | 21.2612 | 0.002516 | 0.002373 | 152.6861 | 0.25825 | 8.4023 | 0.002573 | 0.004122 | 0.004022 | 5.1e-05 | 0.00105 | 0.30685 | 1.3249 | 2.5474 | 100.61 | 21.09 | 25.794 | 82.3541 | 83.7262 | 78.9113 | 0.006419 | 1.0028 | 0.024289 | 24.2214 | 0.066612 | 0.15363 | 0.003791 | 0.27154 | 0.27154 | 0.023367
2019-03-28 15:47 | 1553802423.2864 | 522 | 0.20872 | 0.000987 | 9.4e-05 | 100.6824 | 100.0078 | 121.8623 | 21.2612 | 0.002521 | 0.002334 | 152.671 | 0.25831 | 8.4027 | 0.002477 | 0.003947 | 0.00383 | 5e-05 | 0.001049 | 0.30525 | 1.3183 | 2.5474 | 100.61 | 21.09 | 25.78 | 82.4734 | 81.3986 | 77.961 | 0.006415 | 1.0055 | 0.024331 | 24.1978 | 0.060133 | 0.15367 | 0.003897 | 0.27972 | 0.27972 | 0.022689
2019-03-28 15:47 | 1553802424.2865 | 523 | 0.20873 | 0.000991 | 9.4e-05 | 100.756 | 100.0696 | 121.648 | 21.2612 | 0.002491 | 0.002352 | 152.6622 | 0.25773 | 8.4044 | 0.002552 | 0.003824 | 0.003843 | 5.3e-05 | 0.001055 | 0.32388 | 1.2885 | 2.5475 | 100.61 | 21.1 | 25.774 | 81.7349 | 83.2215 | 78.407 | 0.006418 | 1.0042 | 0.024294 | 24.1919 | 0.061138 | 0.15332 | 0.00384 | 0.27513 | 0.27513 | 0.023016
2019-03-28 15:47 | 1553802425.2879 | 524 | 0.20871 | 0.000991 | 9.3e-05 | 100.6267 | 99.8519 | 122.0689 | 21.2612 | 0.002472 | 0.002345 | 152.6588 | 0.25807 | 8.4012 | 0.002507 | 0.00445 | 0.004281 | 5.2e-05 | 0.001056 | 0.32709 | 1.2741 | 2.5472 | 100.61 | 21.1 | 25.769 | 81.2902 | 82.1387 | 78.2441 | 0.006417 | 1.0047 | 0.024342 | 24.2281 | 0.069141 | 0.15353 | 0.003952 | 0.28389 | 0.28389 | 0.022335
2019-03-28 15:47 | 1553802426.2875 | 525 | 0.20871 | 0.000985 | 9.3e-05 | 100.669 | 100.0278 | 121.8217 | 21.2612 | 0.002472 | 0.002386 | 152.6244 | 0.25721 | 8.4027 | 0.002543 | 0.003889 | 0.003813 | 5.2e-05 | 0.001056 | 0.29927 | 1.2714 | 2.5475 | 100.61 | 21.1 | 25.763 | 81.3003 | 82.9943 | 79.2271 | 0.006415 | 1.0019 | 0.024283 | 24.2374 | 0.05715 | 0.15301 | 0.003975 | 0.28493 | 0.28493 | 0.022179
2019-03-28 15:47 | 1553802427.2875 | 526 | 0.20872 | 0.000984 | 9.3e-05 | 100.6422 | 99.8678 | 121.8609 | 21.2612 | 0.002439 | 0.002314 | 152.6419 | 0.25699 | 8.4035 | 0.002484 | 0.003565 | 0.003474 | 5.2e-05 | 0.001051 | 0.30558 | 1.2769 | 2.5474 | 100.61 | 21.1 | 25.762 | 80.5058 | 81.5772 | 77.4894 | 0.006415 | 1.0069 | 0.024347 | 24.1814 | 0.069087 | 0.15288 | 0.003933 | 0.2826 | 0.2826 | 0.022343
2019-03-28 15:47 | 1553802428.2865 | 527 | 0.20873 | 0.000985 | 9.2e-05 | 100.6118 | 99.9018 | 122.2585 | 21.2612 | 0.002531 | 0.002337 | 152.6015 | 0.25828 | 8.4052 | 0.002525 | 0.00355 | 0.003472 | 5.1e-05 | 0.001049 | 0.35673 | 1.2594 | 2.5473 | 100.61 | 21.1 | 25.762 | 82.7027 | 82.5554 | 78.0547 | 0.006415 | 1.0052 | 0.024367 | 24.2404 | 0.063345 | 0.15365 | 0.003848 | 0.27662 | 0.27662 | 0.022941
2019-03-28 15:47 | 1553802429.2876 | 528 | 0.20871 | 0.000981 | 9.3e-05 | 100.6404 | 99.895 | 122.0743 | 21.2612 | 0.00253 | 0.00235 | 152.6179 | 0.25788 | 8.4024 | 0.002426 | 0.003592 | 0.003464 | 5.3e-05 | 0.001045 | 0.32068 | 1.3266 | 2.5474 | 100.61 | 21.1 | 25.762 | 82.677 | 80.1823 | 78.3594 | 0.006415 | 1.0044 | 0.024338 | 24.2326 | 0.066495 | 0.15341 | 0.003974 | 0.28556 | 0.28556 | 0.022187
2019-03-28 15:47 | 1553802430.287 | 529 | 0.20872 | 0.000982 | 9.2e-05 | 100.6288 | 99.9986 | 122.0895 | 21.2612 | 0.002451 | 0.00232 | 152.611 | 0.25794 | 8.4036 | 0.002396 | 0.004427 | 0.004337 | 5.2e-05 | 0.001041 | 0.32068 | 1.3221 | 2.5473 | 100.61 | 21.1 | 25.762 | 80.7788 | 79.4549 | 77.6391 | 0.006415 | 1.0064 | 0.024365 | 24.2093 | 0.05619 | 0.15345 | 0.003955 | 0.28449 | 0.28449 | 0.022276
2019-03-28 15:47 | 1553802431.2866 | 530 | 0.20874 | 0.000979 | 9.2e-05 | 100.5817 | 99.9592 | 122.3423 | 21.2612 | 0.002531 | 0.002392 | 152.606 | 0.25804 | 8.4031 | 0.002453 | 0.004516 | 0.004422 | 4.8e-05 | 0.001041 | 0.32803 | 1.3142 | 2.5474 | 100.61 | 21.1 | 25.762 | 82.7121 | 80.8264 | 79.3621 | 0.006415 | 1.0015 | 0.024331 | 24.2938 | 0.055531 | 0.15351 | 0.003832 | 0.27511 | 0.27511 | 0.023045
2019-03-28 15:47 | 1553802432.2856 | 531 | 0.20871 | 0.000977 | 9.2e-05 | 100.5052 | 99.8441 | 122.1623 | 21.2612 | 0.002469 | 0.002352 | 152.5864 | 0.25784 | 8.4034 | 0.002494 | 0.003437 | 0.003349 | 5.4e-05 | 0.001048 | 0.32621 | 1.3056 | 2.5473 | 100.61 | 21.1 | 25.763 | 81.2287 | 81.8229 | 78.3986 | 0.006415 | 1.0043 | 0.024346 | 24.2427 | 0.059033 | 0.15339 | 0.003969 | 0.28536 | 0.28536 | 0.0222
2019-03-28 15:47 | 1553802433.2861 | 532 | 0.20872 | 0.000976 | 9.1e-05 | 100.5278 | 99.8689 | 122.086 | 21.2612 | 0.00252 | 0.002336 | 152.5599 | 0.25754 | 8.4012 | 0.002491 | 0.004315 | 0.004173 | 5.1e-05 | 0.001047 | 0.31797 | 1.3308 | 2.5473 | 100.61 | 21.1 | 25.762 | 82.439 | 81.7388 | 78.0149 | 0.006415 | 1.0054 | 0.024352 | 24.2219 | 0.058826 | 0.15321 | 0.003916 | 0.28152 | 0.28152 | 0.022477
2019-03-28 15:47 | 1553802434.2858 | 533 | 0.20875 | 0.000977 | 9.1e-05 | 100.4599 | 99.839 | 122.309 | 21.2612 | 0.002502 | 0.002335 | 152.5584 | 0.25855 | 8.4036 | 0.002529 | 0.003717 | 0.003675 | 5.4e-05 | 0.001046 | 0.32228 | 1.3332 | 2.5474 | 100.61 | 21.1 | 25.762 | 82.0109 | 82.6659 | 77.9939 | 0.006415 | 1.0054 | 0.024375 | 24.2433 | 0.055459 | 0.15381 | 0.003722 | 0.26752 | 0.26752 | 0.023745
2019-03-28 15:47 | 1553802435.2868 | 534 | 0.20877 | 0.000975 | 9.1e-05 | 100.4554 | 100.0102 | 121.8616 | 21.2612 | 0.002464 | 0.002362 | 152.5087 | 0.2574 | 8.4002 | 0.002525 | 0.003953 | 0.003858 | 5e-05 | 0.001049 | 0.33124 | 1.3116 | 2.5473 | 100.61 | 21.1 | 25.76 | 81.0915 | 82.5546 | 78.6386 | 0.006414 | 1.0036 | 0.024308 | 24.2211 | 0.039735 | 0.15313 | 0.003616 | 0.25903 | 0.25903 | 0.024415
2019-03-28 15:47 | 1553802436.2862 | 535 | 0.20877 | 0.000974 | 9.1e-05 | 100.498 | 99.6834 | 121.9308 | 21.2612 | 0.002473 | 0.002294 | 152.5134 | 0.25693 | 8.4005 | 0.002524 | 0.003827 | 0.003708 | 5.3e-05 | 0.001049 | 0.32637 | 1.3117 | 2.5473 | 100.61 | 21.1 | 25.757 | 81.3152 | 82.539 | 77.0069 | 0.006414 | 1.0082 | 0.024371 | 24.1717 | 0.072804 | 0.15285 | 0.003609 | 0.25927 | 0.25927 | 0.024348
2019-03-28 15:47 | 1553802437.287 | 536 | 0.20876 | 0.000976 | 9.1e-05 | 100.51 | 99.5566 | 121.9116 | 21.2612 | 0.002534 | 0.00232 | 152.4737 | 0.25758 | 8.4014 | 0.002439 | 0.0043 | 0.00418 | 5.5e-05 | 0.00105 | 0.31797 | 1.3292 | 2.5474 | 100.61 | 21.1 | 25.755 | 82.7768 | 80.4901 | 77.6297 | 0.006413 | 1.0065 | 0.024348 | 24.1913 | 0.085241 | 0.15324 | 0.003687 | 0.26468 | 0.26468 | 0.023911
2019-03-28 15:47 | 1553802438.287 | 537 | 0.20877 | 0.000971 | 9.1e-05 | 100.4132 | 99.6698 | 121.821 | 21.2612 | 0.002484 | 0.002353 | 152.4541 | 0.25752 | 8.4009 | 0.002445 | 0.003703 | 0.003665 | 5.4e-05 | 0.00105 | 0.30735 | 1.3957 | 2.5476 | 100.61 | 21.1 | 25.753 | 81.5828 | 80.6346 | 78.4212 | 0.006413 | 1.0042 | 0.024311 | 24.2096 | 0.066464 | 0.1532 | 0.003626 | 0.25983 | 0.25983 | 0.02435
2019-03-28 15:47 | 1553802439.287 | 538 | 0.20876 | 0.000971 | 9e-05 | 100.4254 | 99.6511 | 121.9764 | 21.2612 | 0.002444 | 0.002327 | 152.4082 | 0.2575 | 8.4021 | 0.002468 | 0.003842 | 0.00383 | 5.3e-05 | 0.001046 | 0.31868 | 1.2944 | 2.5475 | 100.61 | 21.1 | 25.749 | 80.6134 | 81.185 | 77.8082 | 0.006412 | 1.0059 | 0.024348 | 24.2039 | 0.069229 | 0.15319 | 0.003705 | 0.26608 | 0.26608 | 0.023778
2019-03-28 15:47 | 1553802440.2863 | 539 | 0.20875 | 0.000973 | 9e-05 | 100.353 | 99.7116 | 122.051 | 21.2612 | 0.002461 | 0.002336 | 152.3653 | 0.2585 | 8.4036 | 0.002515 | 0.004267 | 0.004145 | 5.3e-05 | 0.00105 | 0.32493 | 1.2925 | 2.5475 | 100.61 | 21.1 | 25.748 | 81.0126 | 82.3121 | 78.011 | 0.006411 | 1.0054 | 0.024348 | 24.2183 | 0.057354 | 0.15378 | 0.00376 | 0.27018 | 0.27018 | 0.023507
2019-03-28 15:47 | 1553802441.287 | 540 | 0.20877 | 0.000969 | 9e-05 | 100.3047 | 99.8111 | 121.6935 | 21.2612 | 0.002454 | 0.002323 | 152.3667 | 0.2572 | 8.4043 | 0.00254 | 0.004168 | 0.004059 | 5.4e-05 | 0.001051 | 0.32316 | 1.2776 | 2.5474 | 100.61 | 21.1 | 25.747 | 80.8615 | 82.9172 | 77.7071 | 0.006411 | 1.0062 | 0.024323 | 24.1723 | 0.044128 | 0.15301 | 0.003631 | 0.26039 | 0.26039 | 0.024268
2019-03-28 15:47 | 1553802442.2871 | 541 | 0.20877 | 0.000971 | 8.9e-05 | 100.3554 | 99.6708 | 121.6741 | 21.2612 | 0.002507 | 0.002313 | 152.3509 | 0.25745 | 8.4046 | 0.002525 | 0.004039 | 0.003998 | 5.5e-05 | 0.001047 | 0.30541 | 1.2968 | 2.5474 | 100.61 | 21.1 | 25.747 | 82.1404 | 82.5522 | 77.4615 | 0.006411 | 1.0069 | 0.02433 | 24.1619 | 0.061224 | 0.15316 | 0.003643 | 0.26142 | 0.26142 | 0.024196
2019-03-28 15:47 | 1553802443.2871 | 542 | 0.20878 | 0.000969 | 8.9e-05 | 100.2219 | 99.5978 | 121.475 | 21.2612 | 0.002478 | 0.002352 | 152.345 | 0.25775 | 8.4054 | 0.002512 | 0.003692 | 0.003596 | 5.4e-05 | 0.001051 | 0.32764 | 1.3157 | 2.5473 | 100.61 | 21.1 | 25.747 | 81.4275 | 82.2597 | 78.4125 | 0.006411 | 1.0042 | 0.024277 | 24.1749 | 0.05588 | 0.15334 | 0.003582 | 0.25639 | 0.25639 | 0.0247
2019-03-28 15:47 | 1553802444.2871 | 543 | 0.20877 | 0.000965 | 8.8e-05 | 100.2195 | 99.4551 | 121.3635 | 21.2612 | 0.002473 | 0.002344 | 152.318 | 0.258 | 8.4033 | 0.002518 | 0.003691 | 0.003585 | 5.2e-05 | 0.001049 | 0.32211 | 1.3825 | 2.5475 | 100.61 | 21.11 | 25.742 | 81.3239 | 82.3868 | 78.2044 | 0.006414 | 1.0048 | 0.024273 | 24.1566 | 0.068491 | 0.15349 | 0.003627 | 0.25973 | 0.25973 | 0.024406
2019-03-28 15:47 | 1553802445.2872 | 544 | 0.20877 | 0.000968 | 8.7e-05 | 100.202 | 99.363 | 121.3131 | 21.2612 | 0.002493 | 0.002322 | 152.3226 | 0.25735 | 8.4028 | 0.002487 | 0.00345 | 0.003393 | 5.3e-05 | 0.001054 | 0.3246 | 1.3314 | 2.5474 | 100.61 | 21.11 | 25.739 | 81.8027 | 81.645 | 77.6922 | 0.006413 | 1.0063 | 0.024286 | 24.134 | 0.075206 | 0.1531 | 0.003627 | 0.25988 | 0.25988 | 0.024331
2019-03-28 15:47 | 1553802446.2875 | 545 | 0.20876 | 0.000966 | 8.7e-05 | 100.1713 | 99.238 | 121.468 | 21.2612 | 0.002506 | 0.002299 | 152.2953 | 0.2569 | 8.4037 | 0.002518 | 0.003662 | 0.003504 | 5.3e-05 | 0.001056 | 0.33987 | 1.3308 | 2.5475 | 100.61 | 21.11 | 25.736 | 82.1075 | 82.3986 | 77.1201 | 0.006412 | 1.0079 | 0.024321 | 24.1297 | 0.083728 | 0.15283 | 0.003694 | 0.26517 | 0.26517 | 0.023804
2019-03-28 15:47 | 1553802447.2862 | 546 | 0.20876 | 0.000965 | 8.7e-05 | 100.1472 | 99.3667 | 121.6701 | 21.2612 | 0.002498 | 0.002327 | 152.2718 | 0.25829 | 8.4016 | 0.002546 | 0.004133 | 0.004029 | 5.1e-05 | 0.001057 | 0.32211 | 1.3133 | 2.5474 | 100.61 | 21.11 | 25.732 | 81.9188 | 83.0638 | 77.8099 | 0.006411 | 1.0059 | 0.024317 | 24.1735 | 0.069976 | 0.15366 | 0.003695 | 0.26524 | 0.26524 | 0.023926
2019-03-28 15:47 | 1553802448.2852 | 547 | 0.20877 | 0.000961 | 8.7e-05 | 100.0968 | 99.311 | 121.779 | 21.2612 | 0.002504 | 0.002335 | 152.2718 | 0.25804 | 8.403 | 0.00252 | 0.004201 | 0.004076 | 5.2e-05 | 0.001054 | 0.32244 | 1.308 | 2.5474 | 100.61 | 21.11 | 25.729 | 82.0623 | 82.4525 | 78.0033 | 0.006411 | 1.0054 | 0.024321 | 24.191 | 0.070496 | 0.15351 | 0.003662 | 0.26285 | 0.26285 | 0.024119
2019-03-28 15:47 | 1553802449.2856 | 548 | 0.20876 | 0.000964 | 8.6e-05 | 100.0745 | 99.4469 | 121.6353 | 21.2612 | 0.002517 | 0.002295 | 152.2447 | 0.25757 | 8.404 | 0.00246 | 0.004548 | 0.004429 | 5.1e-05 | 0.001048 | 0.30757 | 1.3066 | 2.5473 | 100.61 | 21.11 | 25.723 | 82.3799 | 81.0049 | 77.0351 | 0.006409 | 1.0082 | 0.024341 | 24.1434 | 0.056269 | 0.15323 | 0.003705 | 0.26626 | 0.26626 | 0.023768
2019-03-28 15:47 | 1553802450.2856 | 549 | 0.20877 | 0.000966 | 8.6e-05 | 100.0581 | 99.486 | 121.6234 | 21.2612 | 0.002457 | 0.00229 | 152.2383 | 0.25816 | 8.4019 | 0.00258 | 0.004043 | 0.004013 | 5.2e-05 | 0.001047 | 0.32244 | 1.2998 | 2.5472 | 100.61 | 21.11 | 25.721 | 80.9285 | 83.8837 | 76.9173 | 0.006409 | 1.0085 | 0.024343 | 24.1381 | 0.051291 | 0.15358 | 0.003662 | 0.26317 | 0.26317 | 0.024101
2019-03-28 15:48 | 1553802451.2853 | 550 | 0.20877 | 0.000963 | 8.4e-05 | 99.978 | 99.2081 | 121.6208 | 21.2612 | 0.002436 | 0.002371 | 152.2029 | 0.25822 | 8.404 | 0.002573 | 0.004218 | 0.004101 | 5e-05 | 0.00105 | 0.31758 | 1.317 | 2.5474 | 100.61 | 21.11 | 25.716 | 80.4293 | 83.7089 | 78.8592 | 0.006407 | 1.0029 | 0.024276 | 24.2047 | 0.069146 | 0.15362 | 0.003643 | 0.26111 | 0.26111 | 0.024297
2019-03-28 15:48 | 1553802452.2863 | 551 | 0.20878 | 0.000967 | 8.4e-05 | 99.9393 | 99.1609 | 121.5168 | 21.2612 | 0.002475 | 0.002297 | 152.1892 | 0.25718 | 8.4014 | 0.002541 | 0.003903 | 0.00379 | 4.7e-05 | 0.001046 | 0.31542 | 1.3116 | 2.5474 | 100.61 | 21.11 | 25.713 | 81.3729 | 82.946 | 77.0896 | 0.006407 | 1.008 | 0.024327 | 24.1335 | 0.06994 | 0.153 | 0.00357 | 0.25639 | 0.25639 | 0.024646
2019-03-28 15:48 | 1553802453.2867 | 552 | 0.20876 | 0.000961 | 8.3e-05 | 99.8416 | 99.1028 | 121.7952 | 21.2612 | 0.002498 | 0.002338 | 152.1701 | 0.25739 | 8.4037 | 0.00254 | 0.004225 | 0.004139 | 4.7e-05 | 0.001046 | 0.30397 | 1.2956 | 2.5473 | 100.61 | 21.11 | 25.711 | 81.9118 | 82.9164 | 78.0773 | 0.006406 | 1.0052 | 0.02432 | 24.1952 | 0.066433 | 0.15312 | 0.003723 | 0.26749 | 0.26749 | 0.023642
2019-03-28 15:48 | 1553802454.2864 | 553 | 0.20875 | 0.00096 | 8.3e-05 | 99.8258 | 99.3086 | 121.7095 | 21.2612 | 0.002513 | 0.002322 | 152.1461 | 0.25829 | 8.4009 | 0.002563 | 0.004531 | 0.004427 | 4.9e-05 | 0.001053 | 0.30502 | 1.3223 | 2.5475 | 100.61 | 21.11 | 25.703 | 82.2635 | 83.4735 | 77.6915 | 0.006404 | 1.0063 | 0.024325 | 24.1734 | 0.046457 | 0.15366 | 0.003782 | 0.27193 | 0.27193 | 0.023337
2019-03-28 15:48 | 1553802455.2864 | 554 | 0.20878 | 0.000954 | 8.3e-05 | 99.8011 | 98.9316 | 122.0858 | 21.2612 | 0.002487 | 0.002346 | 152.1413 | 0.25821 | 8.4027 | 0.002577 | 0.004384 | 0.004317 | 4.8e-05 | 0.001054 | 0.3126 | 1.3226 | 2.5475 | 100.61 | 21.11 | 25.701 | 81.6481 | 83.8229 | 78.252 | 0.006404 | 1.0047 | 0.024343 | 24.23 | 0.078271 | 0.15361 | 0.003617 | 0.25997 | 0.25997 | 0.024403
2019-03-28 15:48 | 1553802456.2858 | 555 | 0.20879 | 0.000951 | 8.4e-05 | 99.712 | 98.7542 | 121.7809 | 21.2612 | 0.002518 | 0.002311 | 152.1225 | 0.25781 | 8.4041 | 0.002574 | 0.004122 | 0.004051 | 5e-05 | 0.001051 | 0.31996 | 1.3144 | 2.5475 | 100.61 | 21.11 | 25.699 | 82.3938 | 83.744 | 77.4216 | 0.006403 | 1.0071 | 0.024342 | 24.1712 | 0.086333 | 0.15337 | 0.003562 | 0.25595 | 0.25595 | 0.024747
2019-03-28 15:48 | 1553802457.2865 | 556 | 0.20878 | 0.000954 | 8.5e-05 | 99.7084 | 98.8099 | 121.6841 | 21.2612 | 0.002468 | 0.002371 | 152.1399 | 0.25757 | 8.4018 | 0.002546 | 0.004629 | 0.004433 | 4.6e-05 | 0.001049 | 0.32068 | 1.3756 | 2.5474 | 100.61 | 21.11 | 25.695 | 81.1867 | 83.0732 | 78.8724 | 0.006402 | 1.0029 | 0.024282 | 24.2115 | 0.080967 | 0.15323 | 0.003592 | 0.25742 | 0.25742 | 0.024584
2019-03-28 15:48 | 1553802458.2865 | 557 | 0.20878 | 0.00095 | 8.6e-05 | 99.6706 | 98.7546 | 121.7621 | 21.2612 | 0.002478 | 0.002325 | 152.1127 | 0.25731 | 8.4019 | 0.002559 | 0.003506 | 0.003424 | 4.8e-05 | 0.001052 | 0.30021 | 1.3302 | 2.5473 | 100.61 | 21.11 | 25.692 | 81.4438 | 83.3721 | 77.7631 | 0.006401 | 1.0061 | 0.024328 | 24.1811 | 0.082588 | 0.15307 | 0.003634 | 0.26088 | 0.26088 | 0.024233
2019-03-28 15:48 | 1553802459.2866 | 558 | 0.20877 | 0.000951 | 8.7e-05 | 99.7058 | 98.7427 | 121.8839 | 21.2612 | 0.002534 | 0.002373 | 152.1112 | 0.25663 | 8.3981 | 0.002555 | 0.004306 | 0.004164 | 4.9e-05 | 0.001051 | 0.29878 | 1.3315 | 2.5475 | 100.61 | 21.11 | 25.684 | 82.7784 | 83.2761 | 78.9068 | 0.006399 | 1.0028 | 0.024301 | 24.2325 | 0.086818 | 0.15267 | 0.003643 | 0.26119 | 0.26119 | 0.02414
2019-03-28 15:48 | 1553802460.2856 | 559 | 0.20875 | 0.00095 | 8.7e-05 | 99.6203 | 98.796 | 122.1926 | 21.2612 | 0.002529 | 0.002291 | 152.0835 | 0.25662 | 8.4055 | 0.002455 | 0.004212 | 0.004138 | 5.1e-05 | 0.001053 | 0.31365 | 1.3064 | 2.5473 | 100.61 | 21.11 | 25.68 | 82.6645 | 80.874 | 76.9438 | 0.006398 | 1.0084 | 0.024399 | 24.1955 | 0.074316 | 0.15266 | 0.003767 | 0.27138 | 0.27138 | 0.023233
2019-03-28 15:48 | 1553802461.2866 | 560 | 0.20877 | 0.000949 | 8.8e-05 | 99.6405 | 98.5548 | 122.2895 | 21.2612 | 0.002461 | 0.002315 | 152.0635 | 0.25793 | 8.4039 | 0.002603 | 0.003966 | 0.003848 | 5.1e-05 | 0.001057 | 0.32781 | 1.285 | 2.5474 | 100.61 | 21.11 | 25.675 | 81.0321 | 84.4284 | 77.5245 | 0.006397 | 1.0068 | 0.024389 | 24.2251 | 0.097999 | 0.15344 | 0.003654 | 0.26296 | 0.26296 | 0.024099
2019-03-28 15:48 | 1553802462.2866 | 561 | 0.20876 | 0.000949 | 8.8e-05 | 99.5905 | 98.7413 | 122.2095 | 21.2612 | 0.002449 | 0.002347 | 152.0613 | 0.2585 | 8.4024 | 0.002433 | 0.004449 | 0.004311 | 4.8e-05 | 0.001055 | 0.32693 | 1.3194 | 2.5473 | 100.61 | 21.11 | 25.663 | 80.7374 | 80.3608 | 78.2962 | 0.006394 | 1.0046 | 0.024354 | 24.2438 | 0.076595 | 0.15378 | 0.003715 | 0.26702 | 0.26702 | 0.023785
2019-03-28 15:48 | 1553802463.2859 | 562 | 0.20876 | 0.000951 | 8.8e-05 | 99.4827 | 98.4566 | 122.1482 | 21.2612 | 0.002484 | 0.002352 | 152.049 | 0.25709 | 8.4043 | 0.002533 | 0.00421 | 0.004097 | 4.8e-05 | 0.001057 | 0.30469 | 1.2947 | 2.5474 | 100.61 | 21.11 | 25.656 | 81.5866 | 82.7433 | 78.4007 | 0.006392 | 1.0043 | 0.024344 | 24.2414 | 0.092732 | 0.15294 | 0.003723 | 0.2675 | 0.2675 | 0.023613
2019-03-28 15:48 | 1553802464.2851 | 563 | 0.20877 | 0.000957 | 8.7e-05 | 99.455 | 98.6683 | 122.1869 | 21.2612 | 0.00247 | 0.002322 | 152.0344 | 0.25725 | 8.4033 | 0.002487 | 0.003836 | 0.003778 | 5e-05 | 0.001055 | 0.30038 | 1.2755 | 2.5473 | 100.61 | 21.11 | 25.64 | 81.2333 | 81.6481 | 77.6859 | 0.006388 | 1.0063 | 0.024373 | 24.2205 | 0.071034 | 0.15304 | 0.003695 | 0.26583 | 0.26583 | 0.023776
2019-03-28 15:48 | 1553802465.2848 | 564 | 0.20878 | 0.000951 | 8.6e-05 | 99.4457 | 98.7233 | 122.1907 | 21.2612 | 0.00245 | 0.002313 | 152.0089 | 0.25742 | 8.4029 | 0.00242 | 0.00365 | 0.003558 | 4.8e-05 | 0.001056 | 0.31022 | 1.3008 | 2.5474 | 100.61 | 21.12 | 25.631 | 80.7593 | 80.0496 | 77.4598 | 0.00639 | 1.0069 | 0.024381 | 24.2131 | 0.065215 | 0.15314 | 0.003587 | 0.25803 | 0.25803 | 0.024511
2019-03-28 15:48 | 1553802466.2862 | 565 | 0.20881 | 0.000956 | 8.6e-05 | 99.4702 | 98.6122 | 122.2888 | 21.2612 | 0.002475 | 0.002343 | 151.9981 | 0.25813 | 8.401 | 0.002491 | 0.003924 | 0.0038 | 5.1e-05 | 0.001055 | 0.31188 | 1.3166 | 2.5474 | 100.61 | 21.12 | 25.621 | 81.3551 | 81.7489 | 78.1833 | 0.006388 | 1.0049 | 0.024366 | 24.2478 | 0.077483 | 0.15356 | 0.003441 | 0.24723 | 0.24723 | 0.025652
2019-03-28 15:48 | 1553802467.2852 | 566 | 0.20882 | 0.000954 | 8.5e-05 | 99.3427 | 98.7063 | 122.1543 | 21.2612 | 0.002509 | 0.002308 | 151.9554 | 0.25747 | 8.4056 | 0.002534 | 0.004594 | 0.00455 | 4.9e-05 | 0.00106 | 0.32925 | 1.2976 | 2.5473 | 100.61 | 21.12 | 25.601 | 82.1777 | 82.7909 | 77.3546 | 0.006383 | 1.0072 | 0.024381 | 24.2059 | 0.05748 | 0.15317 | 0.003362 | 0.24158 | 0.24158 | 0.026186
2019-03-28 15:48 | 1553802468.2848 | 567 | 0.20882 | 0.000954 | 8.5e-05 | 99.3451 | 98.6037 | 121.9738 | 21.2612 | 0.002539 | 0.002311 | 151.9347 | 0.25726 | 8.4035 | 0.002546 | 0.004116 | 0.003955 | 4.8e-05 | 0.001052 | 0.30917 | 1.3285 | 2.5474 | 100.61 | 21.12 | 25.592 | 82.9102 | 83.0583 | 77.4147 | 0.00638 | 1.0071 | 0.024361 | 24.1901 | 0.066997 | 0.15304 | 0.003397 | 0.24396 | 0.24396 | 0.025908
2019-03-28 15:48 | 1553802469.2859 | 568 | 0.20879 | 0.000946 | 8.5e-05 | 99.3364 | 98.3578 | 121.7494 | 21.2612 | 0.002487 | 0.002323 | 151.9351 | 0.25867 | 8.4034 | 0.002521 | 0.003822 | 0.003836 | 4.8e-05 | 0.001049 | 0.30829 | 1.3106 | 2.5474 | 100.61 | 21.12 | 25.584 | 81.6467 | 82.4609 | 77.7008 | 0.006378 | 1.0063 | 0.024329 | 24.1777 | 0.088559 | 0.15388 | 0.00357 | 0.2563 | 0.2563 | 0.024796
2019-03-28 15:48 | 1553802470.286 | 569 | 0.2088 | 0.000947 | 8.5e-05 | 99.3045 | 98.3645 | 121.4579 | 21.2612 | 0.002504 | 0.002283 | 151.964 | 0.25779 | 8.4004 | 0.002561 | 0.003901 | 0.003764 | 5.2e-05 | 0.001052 | 0.32764 | 1.298 | 2.5474 | 100.61 | 21.12 | 25.572 | 82.0561 | 83.4298 | 76.7364 | 0.006375 | 1.009 | 0.024333 | 24.1155 | 0.085071 | 0.15336 | 0.00349 | 0.25051 | 0.25051 | 0.025283
2019-03-28 15:48 | 1553802471.2849 | 570 | 0.20883 | 0.000943 | 8.5e-05 | 99.2992 | 98.2293 | 121.709 | 21.2612 | 0.002458 | 0.002281 | 151.9788 | 0.25794 | 8.4028 | 0.002565 | 0.004704 | 0.004699 | 4.7e-05 | 0.001047 | 0.33212 | 1.3666 | 2.5473 | 100.61 | 21.12 | 25.566 | 80.9557 | 83.532 | 76.7093 | 0.006374 | 1.0091 | 0.024359 | 24.1395 | 0.096891 | 0.15345 | 0.003318 | 0.23816 | 0.23816 | 0.026609
2019-03-28 15:48 | 1553802472.2854 | 571 | 0.2088 | 0.000944 | 8.5e-05 | 99.326 | 98.2256 | 121.8196 | 21.2612 | 0.00247 | 0.00234 | 151.9535 | 0.25783 | 8.4021 | 0.002566 | 0.004099 | 0.003978 | 5.1e-05 | 0.001042 | 0.33445 | 1.3281 | 2.5474 | 100.61 | 21.12 | 25.559 | 81.2311 | 83.5591 | 78.1086 | 0.006372 | 1.0051 | 0.024322 | 24.1987 | 0.099646 | 0.15338 | 0.003484 | 0.24996 | 0.24996 | 0.025343
2019-03-28 15:48 | 1553802473.285 | 572 | 0.2088 | 0.000944 | 8.4e-05 | 99.2548 | 98.3921 | 122.0541 | 21.2612 | 0.002446 | 0.002334 | 151.9489 | 0.25707 | 8.4018 | 0.002546 | 0.003969 | 0.003862 | 4.8e-05 | 0.001046 | 0.31581 | 1.3345 | 2.5473 | 100.61 | 21.12 | 25.556 | 80.6541 | 83.0686 | 77.9815 | 0.006371 | 1.0055 | 0.02435 | 24.2176 | 0.078087 | 0.15293 | 0.003495 | 0.25108 | 0.25108 | 0.025156
2019-03-28 15:48 | 1553802474.286 | 573 | 0.20882 | 0.000944 | 8.4e-05 | 99.2253 | 98.6486 | 122.2757 | 21.2612 | 0.002467 | 0.002297 | 151.9481 | 0.25796 | 8.402 | 0.002556 | 0.003681 | 0.003575 | 4.7e-05 | 0.001048 | 0.30685 | 1.3251 | 2.5473 | 100.61 | 21.12 | 25.555 | 81.1617 | 83.304 | 77.0913 | 0.006371 | 1.008 | 0.024403 | 24.2088 | 0.052143 | 0.15346 | 0.003405 | 0.24506 | 0.24506 | 0.025863
2019-03-28 15:48 | 1553802475.2864 | 574 | 0.20883 | 0.000941 | 8.3e-05 | 99.2283 | 98.4243 | 121.9682 | 21.2612 | 0.002479 | 0.002269 | 151.9729 | 0.258 | 8.4047 | 0.002608 | 0.004487 | 0.004401 | 5.2e-05 | 0.001045 | 0.30502 | 1.3063 | 2.5474 | 100.61 | 21.12 | 25.554 | 81.447 | 84.5564 | 76.4044 | 0.006371 | 1.01 | 0.024396 | 24.1546 | 0.072764 | 0.15348 | 0.003353 | 0.24118 | 0.24118 | 0.026282
2019-03-28 15:48 | 1553802476.2854 | 575 | 0.20883 | 0.00094 | 8.4e-05 | 99.2161 | 98.5976 | 121.9276 | 21.2612 | 0.002412 | 0.002275 | 151.9634 | 0.25748 | 8.4004 | 0.002614 | 0.004062 | 0.003892 | 4.8e-05 | 0.001047 | 0.33555 | 1.2948 | 2.5473 | 100.61 | 21.12 | 25.555 | 79.8415 | 84.7031 | 76.5603 | 0.006371 | 1.0095 | 0.024386 | 24.156 | 0.055932 | 0.15318 | 0.003305 | 0.23753 | 0.23753 | 0.026633
2019-03-28 15:48 | 1553802477.2851 | 576 | 0.20882 | 0.000941 | 8.3e-05 | 99.2042 | 98.6346 | 121.9841 | 21.2612 | 0.002455 | 0.002337 | 151.9566 | 0.25789 | 8.4002 | 0.002543 | 0.00472 | 0.004592 | 4.8e-05 | 0.001044 | 0.33411 | 1.273 | 2.5472 | 100.61 | 21.12 | 25.563 | 80.8716 | 82.9929 | 78.0406 | 0.006373 | 1.0053 | 0.024341 | 24.2127 | 0.051499 | 0.15342 | 0.003414 | 0.24509 | 0.24509 | 0.025852
2019-03-28 15:48 | 1553802478.2845 | 577 | 0.20885 | 0.000938 | 8.3e-05 | 99.1509 | 98.4298 | 121.9325 | 21.2612 | 0.002439 | 0.002286 | 151.9722 | 0.25737 | 8.4028 | 0.0026 | 0.003827 | 0.003737 | 5.1e-05 | 0.001045 | 0.31708 | 1.2714 | 2.5474 | 100.61 | 21.12 | 25.567 | 80.4901 | 84.3714 | 76.8292 | 0.006374 | 1.0088 | 0.024377 | 24.1658 | 0.065287 | 0.15311 | 0.003218 | 0.23115 | 0.23115 | 0.027356
2019-03-28 15:48 | 1553802479.2846 | 578 | 0.20885 | 0.000934 | 8.2e-05 | 99.1806 | 98.302 | 122.0381 | 21.2612 | 0.002464 | 0.002291 | 151.9308 | 0.25822 | 8.4024 | 0.002596 | 0.003846 | 0.003733 | 5.1e-05 | 0.001048 | 0.33572 | 1.2815 | 2.5474 | 100.61 | 21.12 | 25.572 | 81.104 | 84.28 | 76.9493 | 0.006375 | 1.0084 | 0.024384 | 24.1804 | 0.079589 | 0.15362 | 0.003243 | 0.23306 | 0.23306 | 0.027222
2019-03-28 15:48 | 1553802480.2843 | 579 | 0.20881 | 0.000932 | 8.1e-05 | 99.1705 | 98.5083 | 122.101 | 21.2612 | 0.002432 | 0.002351 | 151.896 | 0.25873 | 8.4006 | 0.002572 | 0.003864 | 0.003748 | 5e-05 | 0.001054 | 0.34252 | 1.3092 | 2.5474 | 100.61 | 21.11 | 25.585 | 80.3383 | 83.6933 | 78.3743 | 0.006375 | 1.0043 | 0.024341 | 24.2358 | 0.059932 | 0.15392 | 0.003453 | 0.24813 | 0.24813 | 0.025619
2019-03-28 15:48 | 1553802481.2843 | 580 | 0.20883 | 0.000928 | 8.1e-05 | 99.169 | 98.4322 | 121.8128 | 21.2612 | 0.002559 | 0.002301 | 151.8581 | 0.25764 | 8.402 | 0.002534 | 0.004254 | 0.004159 | 5e-05 | 0.001054 | 0.30414 | 1.2942 | 2.5474 | 100.61 | 21.11 | 25.591 | 83.3721 | 82.7869 | 77.1903 | 0.006376 | 1.0077 | 0.024353 | 24.1663 | 0.06671 | 0.15327 | 0.003337 | 0.23976 | 0.23976 | 0.026402
2019-03-28 15:48 | 1553802482.2847 | 581 | 0.20883 | 0.000928 | 8.1e-05 | 99.1982 | 98.2073 | 121.9499 | 21.2612 | 0.002508 | 0.002321 | 151.8542 | 0.25797 | 8.4032 | 0.002591 | 0.004062 | 0.004051 | 4.9e-05 | 0.001053 | 0.32316 | 1.3027 | 2.5474 | 100.61 | 21.11 | 25.598 | 82.1592 | 84.1521 | 77.6617 | 0.006378 | 1.0064 | 0.02435 | 24.1962 | 0.089802 | 0.15347 | 0.003371 | 0.24218 | 0.24218 | 0.026172
2019-03-28 15:48 | 1553802483.2843 | 582 | 0.2088 | 0.000926 | 8.1e-05 | 99.1601 | 98.1971 | 121.686 | 21.2612 | 0.002412 | 0.002346 | 151.8273 | 0.25755 | 8.4047 | 0.00248 | 0.003957 | 0.003867 | 5e-05 | 0.001049 | 0.32101 | 1.3038 | 2.5474 | 100.61 | 21.12 | 25.611 | 79.8446 | 81.4845 | 78.2534 | 0.006385 | 1.0047 | 0.024303 | 24.1904 | 0.087292 | 0.15322 | 0.003547 | 0.25458 | 0.25458 | 0.024856
2019-03-28 15:48 | 1553802484.2844 | 583 | 0.20883 | 0.000927 | 8.1e-05 | 99.1866 | 98.245 | 121.3469 | 21.2612 | 0.002475 | 0.002283 | 151.8022 | 0.25818 | 8.4026 | 0.002567 | 0.003785 | 0.00371 | 4.8e-05 | 0.001045 | 0.32278 | 1.3199 | 2.5473 | 100.61 | 21.12 | 25.618 | 81.3628 | 83.5639 | 76.759 | 0.006387 | 1.009 | 0.024321 | 24.1052 | 0.085321 | 0.15359 | 0.003371 | 0.24193 | 0.24193 | 0.02622
2019-03-28 15:48 | 1553802485.2858 | 584 | 0.20883 | 0.000924 | 8.1e-05 | 99.1095 | 97.9963 | 121.3241 | 21.2612 | 0.002484 | 0.002291 | 151.7697 | 0.25795 | 8.4027 | 0.002516 | 0.003741 | 0.003665 | 4.9e-05 | 0.001044 | 0.32444 | 1.3539 | 2.5475 | 100.61 | 21.11 | 25.631 | 81.5741 | 82.344 | 76.9305 | 0.006386 | 1.0085 | 0.024313 | 24.1089 | 0.10103 | 0.15345 | 0.003362 | 0.24116 | 0.24116 | 0.02628
2019-03-28 15:48 | 1553802486.2854 | 585 | 0.20884 | 0.000923 | 8.1e-05 | 99.1416 | 97.9966 | 121.3487 | 21.2612 | 0.002448 | 0.002304 | 151.7613 | 0.25772 | 8.4028 | 0.002577 | 0.003649 | 0.003488 | 4.9e-05 | 0.001038 | 0.31581 | 1.3058 | 2.5475 | 100.61 | 21.11 | 25.637 | 80.7149 | 83.8166 | 77.2635 | 0.006388 | 1.0075 | 0.024304 | 24.1228 | 0.1039 | 0.15332 | 0.003321 | 0.23811 | 0.23811 | 0.026593
2019-03-28 15:48 | 1553802487.2855 | 586 | 0.20881 | 0.000926 | 8e-05 | 99.0987 | 98.1889 | 121.5194 | 21.2612 | 0.002459 | 0.002305 | 151.767 | 0.2583 | 8.4047 | 0.002527 | 0.004049 | 0.003911 | 4.9e-05 | 0.001038 | 0.32659 | 1.3539 | 2.5475 | 100.61 | 21.12 | 25.649 | 80.9699 | 82.6114 | 77.2727 | 0.006395 | 1.0075 | 0.024321 | 24.1401 | 0.082498 | 0.15366 | 0.003455 | 0.24811 | 0.24811 | 0.025579
2019-03-28 15:48 | 1553802488.2845 | 587 | 0.20883 | 0.000926 | 8e-05 | 99.1392 | 98.3435 | 121.6619 | 21.2612 | 0.002487 | 0.002338 | 151.7316 | 0.25873 | 8.4016 | 0.002505 | 0.004211 | 0.004125 | 5.1e-05 | 0.001037 | 0.31382 | 1.3485 | 2.5473 | 100.61 | 21.12 | 25.654 | 81.6513 | 82.0928 | 78.0593 | 0.006396 | 1.0052 | 0.024308 | 24.1813 | 0.072084 | 0.15392 | 0.003336 | 0.23929 | 0.23929 | 0.026565
2019-03-28 15:48 | 1553802489.2855 | 588 | 0.20886 | 0.000921 | 8e-05 | 99.2203 | 98.1206 | 121.8191 | 21.2612 | 0.002429 | 0.002319 | 151.7603 | 0.25776 | 8.4043 | 0.00255 | 0.003673 | 0.003516 | 4.7e-05 | 0.001038 | 0.32316 | 1.3582 | 2.5473 | 100.61 | 21.11 | 25.659 | 80.2664 | 83.1646 | 77.6119 | 0.006393 | 1.0065 | 0.024339 | 24.1815 | 0.099683 | 0.15334 | 0.003185 | 0.22852 | 0.22852 | 0.027714
2019-03-28 15:48 | 1553802490.2845 | 589 | 0.20883 | 0.000919 | 8.1e-05 | 99.1744 | 98.4053 | 121.91 | 21.2612 | 0.002481 | 0.002319 | 151.7507 | 0.25785 | 8.4022 | 0.002564 | 0.004522 | 0.004429 | 5.1e-05 | 0.00104 | 0.32516 | 1.3263 | 2.5472 | 100.61 | 21.11 | 25.67 | 81.5171 | 83.5118 | 77.6158 | 0.006396 | 1.0065 | 0.024348 | 24.1907 | 0.069633 | 0.15339 | 0.003347 | 0.24043 | 0.24043 | 0.02635
2019-03-28 15:48 | 1553802491.2845 | 590 | 0.20884 | 0.000923 | 8.1e-05 | 99.1926 | 98.3598 | 122.2074 | 21.2612 | 0.002443 | 0.002274 | 151.7529 | 0.25791 | 8.4036 | 0.002563 | 0.003804 | 0.003657 | 4.9e-05 | 0.001033 | 0.33356 | 1.3212 | 2.5474 | 100.61 | 21.11 | 25.675 | 80.586 | 83.4673 | 76.5401 | 0.006397 | 1.0096 | 0.024415 | 24.183 | 0.07541 | 0.15343 | 0.003302 | 0.23775 | 0.23775 | 0.026653
2019-03-28 15:48 | 1553802492.2859 | 591 | 0.20885 | 0.000922 | 8.2e-05 | 99.1961 | 98.4152 | 122.1792 | 21.2612 | 0.002468 | 0.002316 | 151.7564 | 0.25791 | 8.4045 | 0.002526 | 0.004001 | 0.00399 | 5.3e-05 | 0.001028 | 0.30973 | 1.3395 | 2.5475 | 100.61 | 21.11 | 25.681 | 81.1874 | 82.5811 | 77.5338 | 0.006399 | 1.0067 | 0.024378 | 24.2145 | 0.070698 | 0.15343 | 0.003224 | 0.23164 | 0.23164 | 0.027355
2019-03-28 15:48 | 1553802493.2847 | 592 | 0.20887 | 0.000918 | 8.2e-05 | 99.1506 | 98.3146 | 122.1956 | 21.2612 | 0.002456 | 0.002258 | 151.7536 | 0.2562 | 8.4042 | 0.002554 | 0.003984 | 0.003891 | 4.9e-05 | 0.001023 | 0.32244 | 1.303 | 2.5474 | 100.61 | 21.11 | 25.691 | 80.9081 | 83.2651 | 76.1512 | 0.006401 | 1.0107 | 0.024427 | 24.1684 | 0.075732 | 0.15241 | 0.003104 | 0.22324 | 0.22324 | 0.028197
2019-03-28 15:48 | 1553802494.2837 | 593 | 0.20882 | 0.000918 | 8.3e-05 | 99.1461 | 98.3883 | 122.4261 | 21.2612 | 0.002445 | 0.002326 | 151.7535 | 0.25661 | 8.4021 | 0.002508 | 0.004325 | 0.004179 | 5e-05 | 0.001021 | 0.29911 | 1.3694 | 2.5475 | 100.61 | 21.11 | 25.695 | 80.6291 | 82.1561 | 77.7835 | 0.006402 | 1.006 | 0.024394 | 24.2476 | 0.068624 | 0.15266 | 0.003407 | 0.24515 | 0.24515 | 0.025719
2019-03-28 15:48 | 1553802495.2847 | 594 | 0.20889 | 0.00092 | 8.3e-05 | 99.1923 | 98.5392 | 122.2435 | 21.2612 | 0.002454 | 0.002292 | 151.7243 | 0.25756 | 8.4017 | 0.002497 | 0.003853 | 0.003791 | 5.1e-05 | 0.001016 | 0.31564 | 1.3148 | 2.5474 | 100.61 | 21.11 | 25.703 | 80.8591 | 81.8799 | 76.957 | 0.006404 | 1.0084 | 0.024404 | 24.201 | 0.059086 | 0.15322 | 0.002979 | 0.21382 | 0.21382 | 0.029595
2019-03-28 15:48 | 1553802496.2837 | 595 | 0.2088 | 0.000919 | 8.4e-05 | 99.1541 | 98.3741 | 122.143 | 21.2612 | 0.002434 | 0.002333 | 151.7187 | 0.25644 | 8.4049 | 0.002446 | 0.003762 | 0.003691 | 5e-05 | 0.001015 | 0.29927 | 1.3372 | 2.5474 | 100.61 | 21.11 | 25.707 | 80.3755 | 80.661 | 77.9589 | 0.006405 | 1.0055 | 0.024359 | 24.2256 | 0.070647 | 0.15255 | 0.003557 | 0.25575 | 0.25575 | 0.024635
2019-03-28 15:48 | 1553802497.2841 | 596 | 0.20888 | 0.000918 | 8.4e-05 | 99.0758 | 98.2935 | 121.8867 | 21.2612 | 0.002459 | 0.002359 | 151.7071 | 0.25815 | 8.405 | 0.002535 | 0.003672 | 0.003599 | 5e-05 | 0.001021 | 0.3178 | 1.3553 | 2.5476 | 100.61 | 21.11 | 25.71 | 80.9876 | 82.8035 | 78.5653 | 0.006406 | 1.0038 | 0.024313 | 24.2211 | 0.070902 | 0.15357 | 0.003086 | 0.22081 | 0.22081 | 0.028724
2019-03-28 15:48 | 1553802498.2839 | 597 | 0.20885 | 0.000919 | 8.4e-05 | 99.2027 | 98.4318 | 121.7426 | 21.2612 | 0.002471 | 0.002315 | 151.705 | 0.25696 | 8.4036 | 0.002459 | 0.004431 | 0.004329 | 5e-05 | 0.00102 | 0.27616 | 1.3573 | 2.5474 | 100.61 | 21.11 | 25.717 | 81.2544 | 80.9713 | 77.5175 | 0.006408 | 1.0068 | 0.024335 | 24.1707 | 0.069776 | 0.15286 | 0.003235 | 0.23186 | 0.23186 | 0.027229
2019-03-28 15:48 | 1553802499.2848 | 598 | 0.20884 | 0.00092 | 8.4e-05 | 99.2015 | 98.3938 | 121.6663 | 21.2612 | 0.002493 | 0.002326 | 151.6702 | 0.25761 | 8.4035 | 0.002508 | 0.003673 | 0.003578 | 4.6e-05 | 0.001019 | 0.31636 | 1.3447 | 2.5474 | 100.61 | 21.11 | 25.722 | 81.7989 | 82.145 | 77.7772 | 0.006409 | 1.006 | 0.024318 | 24.172 | 0.073127 | 0.15325 | 0.003312 | 0.23738 | 0.23738 | 0.026663
2019-03-28 15:48 | 1553802500.2841 | 599 | 0.20885 | 0.000923 | 8.3e-05 | 99.2768 | 98.3771 | 121.3973 | 21.2612 | 0.002468 | 0.002297 | 151.6635 | 0.25746 | 8.4017 | 0.002523 | 0.003859 | 0.003742 | 4.8e-05 | 0.001023 | 0.30989 | 1.3449 | 2.5474 | 100.61 | 21.11 | 25.731 | 81.2023 | 82.5092 | 77.0944 | 0.006411 | 1.008 | 0.024315 | 24.1218 | 0.081432 | 0.15316 | 0.003264 | 0.23386 | 0.23386 | 0.027049
2019-03-28 15:48 | 1553802501.285 | 600 | 0.20884 | 0.000917 | 8.3e-05 | 99.2125 | 98.4189 | 121.5126 | 21.2612 | 0.002427 | 0.0023 | 151.6457 | 0.25761 | 8.4031 | 0.002511 | 0.004135 | 0.004103 | 4.9e-05 | 0.001028 | 0.30596 | 1.3484 | 2.5475 | 100.61 | 21.11 | 25.736 | 80.201 | 82.2363 | 77.1489 | 0.006412 | 1.0078 | 0.024324 | 24.1351 | 0.071835 | 0.15326 | 0.003324 | 0.23838 | 0.23838 | 0.026551
2019-03-28 15:48 | 1553802502.2849 | 601 | 0.20885 | 0.000917 | 8.1e-05 | 99.2304 | 98.3676 | 121.2717 | 21.2612 | 0.002516 | 0.002325 | 151.588 | 0.25668 | 8.4035 | 0.002474 | 0.003195 | 0.003059 | 4.8e-05 | 0.001032 | 0.30757 | 1.3389 | 2.5474 | 100.61 | 21.11 | 25.74 | 82.3517 | 81.3292 | 77.7474 | 0.006413 | 1.0061 | 0.02428 | 24.1318 | 0.07811 | 0.1527 | 0.003262 | 0.23349 | 0.23349 | 0.02701
2019-03-28 15:48 | 1553802503.2843 | 602 | 0.20884 | 0.000916 | 8e-05 | 99.2101 | 98.37 | 121.4149 | 21.2612 | 0.002456 | 0.002351 | 151.5838 | 0.25802 | 8.4059 | 0.002512 | 0.004147 | 0.003997 | 4.9e-05 | 0.001038 | 0.28495 | 1.2907 | 2.5475 | 100.61 | 21.11 | 25.746 | 80.9112 | 82.2387 | 78.393 | 0.006415 | 1.0043 | 0.024272 | 24.1682 | 0.076068 | 0.1535 | 0.003327 | 0.2383 | 0.2383 | 0.026603
2019-03-28 15:48 | 1553802504.284 | 603 | 0.20879 | 0.000915 | 7.8e-05 | 99.2551 | 98.3489 | 121.1938 | 21.2612 | 0.002457 | 0.002318 | 151.5708 | 0.25789 | 8.402 | 0.002472 | 0.004059 | 0.003888 | 4.5e-05 | 0.001045 | 0.31996 | 1.2562 | 2.5473 | 100.61 | 21.11 | 25.748 | 80.9355 | 81.2787 | 77.5869 | 0.006415 | 1.0066 | 0.024277 | 24.1185 | 0.082037 | 0.15342 | 0.003601 | 0.25847 | 0.25847 | 0.024514
2019-03-28 15:48 | 1553802505.285 | 604 | 0.20881 | 0.000918 | 7.7e-05 | 99.2962 | 98.5721 | 121.5656 | 21.2612 | 0.00251 | 0.002281 | 151.5901 | 0.25733 | 8.4002 | 0.002528 | 0.003968 | 0.003836 | 4.8e-05 | 0.001049 | 0.29391 | 1.2816 | 2.5473 | 100.61 | 21.11 | 25.75 | 82.2128 | 82.6465 | 76.703 | 0.006416 | 1.0091 | 0.024345 | 24.125 | 0.065464 | 0.15309 | 0.003484 | 0.25071 | 0.25071 | 0.025218
2019-03-28 15:48 | 1553802506.2845 | 605 | 0.2088 | 0.000912 | 7.5e-05 | 99.2554 | 98.5059 | 121.5893 | 21.2612 | 0.002462 | 0.002304 | 151.5466 | 0.2578 | 8.4017 | 0.002456 | 0.003932 | 0.003827 | 4.7e-05 | 0.001049 | 0.32532 | 1.322 | 2.5474 | 100.61 | 21.11 | 25.755 | 81.0525 | 80.909 | 77.2501 | 0.006417 | 1.0075 | 0.024329 | 24.1462 | 0.067801 | 0.15336 | 0.003567 | 0.25664 | 0.25664 | 0.02468
2019-03-28 15:48 | 1553802507.2851 | 606 | 0.20882 | 0.000915 | 7.5e-05 | 99.2941 | 98.5327 | 121.6543 | 21.2612 | 0.002457 | 0.002298 | 151.5223 | 0.25724 | 8.4018 | 0.002516 | 0.00412 | 0.003987 | 4.7e-05 | 0.001051 | 0.32764 | 1.3305 | 2.5473 | 100.61 | 21.11 | 25.757 | 80.934 | 82.351 | 77.1052 | 0.006418 | 1.008 | 0.02434 | 24.1477 | 0.068853 | 0.15303 | 0.003418 | 0.24591 | 0.24591 | 0.025701
2019-03-28 15:48 | 1553802508.2851 | 607 | 0.20882 | 0.000913 | 7.5e-05 | 99.2798 | 98.5317 | 121.8569 | 21.2612 | 0.00243 | 0.002332 | 151.4967 | 0.25742 | 8.4025 | 0.002476 | 0.00454 | 0.004491 | 4.6e-05 | 0.001051 | 0.32444 | 1.3091 | 2.5473 | 100.61 | 21.11 | 25.76 | 80.282 | 81.3962 | 77.9159 | 0.006418 | 1.0056 | 0.024332 | 24.1957 | 0.067657 | 0.15314 | 0.003438 | 0.24729 | 0.24729 | 0.025575
2019-03-28 15:48 | 1553802509.2842 | 608 | 0.2088 | 0.000922 | 7.5e-05 | 99.3754 | 98.5602 | 121.9029 | 21.2612 | 0.002449 | 0.002346 | 151.4959 | 0.2575 | 8.402 | 0.002492 | 0.00402 | 0.003872 | 4.8e-05 | 0.00105 | 0.29966 | 1.3289 | 2.5475 | 100.61 | 21.11 | 25.762 | 80.7326 | 81.7808 | 78.2496 | 0.006419 | 1.0047 | 0.024325 | 24.2118 | 0.073677 | 0.15318 | 0.003545 | 0.25501 | 0.25501 | 0.024809
2019-03-28 15:48 | 1553802510.2846 | 609 | 0.2088 | 0.000919 | 7.6e-05 | 99.3069 | 98.7148 | 121.9712 | 21.2612 | 0.002447 | 0.002288 | 151.5016 | 0.25831 | 8.4011 | 0.002494 | 0.003931 | 0.003845 | 4.6e-05 | 0.001046 | 0.32045 | 1.3531 | 2.5475 | 100.61 | 21.11 | 25.769 | 80.6822 | 81.8159 | 76.8628 | 0.006421 | 1.0087 | 0.02438 | 24.1708 | 0.053492 | 0.15367 | 0.003566 | 0.25712 | 0.25712 | 0.024684
2019-03-28 15:48 | 1553802511.2832 | 610 | 0.20884 | 0.000918 | 7.6e-05 | 99.3093 | 98.6129 | 121.8187 | 21.2612 | 0.002484 | 0.002298 | 151.4659 | 0.25793 | 8.4015 | 0.00253 | 0.004068 | 0.004039 | 4.8e-05 | 0.001046 | 0.33467 | 1.3092 | 2.5474 | 100.61 | 21.11 | 25.772 | 81.582 | 82.6801 | 77.117 | 0.006421 | 1.0079 | 0.024356 | 24.1644 | 0.062946 | 0.15344 | 0.003301 | 0.23747 | 0.23747 | 0.026685
2019-03-28 15:49 | 1553802512.2833 | 611 | 0.20883 | 0.00092 | 7.6e-05 | 99.3537 | 98.8232 | 121.9487 | 21.2612 | 0.002477 | 0.002308 | 151.4909 | 0.25761 | 8.4035 | 0.002516 | 0.003699 | 0.0036 | 4.6e-05 | 0.001044 | 0.31996 | 1.3058 | 2.5474 | 100.61 | 21.11 | 25.774 | 81.4143 | 82.3424 | 77.3553 | 0.006422 | 1.0072 | 0.024361 | 24.1855 | 0.047889 | 0.15325 | 0.003361 | 0.24192 | 0.24192 | 0.026163
2019-03-28 15:49 | 1553802513.2847 | 612 | 0.20882 | 0.000919 | 7.5e-05 | 99.315 | 98.6435 | 121.84 | 21.2612 | 0.002467 | 0.002329 | 151.4627 | 0.2578 | 8.4016 | 0.002555 | 0.003317 | 0.003247 | 4.8e-05 | 0.001041 | 0.32211 | 1.3128 | 2.5473 | 100.61 | 21.11 | 25.778 | 81.178 | 83.2877 | 77.8519 | 0.006423 | 1.0058 | 0.024333 | 24.1919 | 0.060683 | 0.15336 | 0.003439 | 0.2474 | 0.2474 | 0.025602
2019-03-28 15:49 | 1553802514.2837 | 613 | 0.20883 | 0.000918 | 7.5e-05 | 99.3579 | 98.5283 | 121.8954 | 21.2612 | 0.002429 | 0.002348 | 151.4372 | 0.25767 | 8.4015 | 0.00257 | 0.003994 | 0.003848 | 4.8e-05 | 0.001038 | 0.3282 | 1.3451 | 2.5474 | 100.61 | 21.11 | 25.778 | 80.257 | 83.6397 | 78.3003 | 0.006423 | 1.0045 | 0.024323 | 24.2128 | 0.074994 | 0.15329 | 0.003405 | 0.24478 | 0.24478 | 0.025863
2019-03-28 15:49 | 1553802515.2833 | 614 | 0.20883 | 0.00091 | 7.4e-05 | 99.3275 | 98.7107 | 121.7853 | 21.2612 | 0.002448 | 0.002275 | 151.4112 | 0.25696 | 8.4042 | 0.00257 | 0.004337 | 0.004197 | 4.6e-05 | 0.001043 | 0.32621 | 1.328 | 2.5476 | 100.61 | 21.11 | 25.778 | 80.718 | 83.6498 | 76.5642 | 0.006423 | 1.0095 | 0.024372 | 24.142 | 0.055712 | 0.15287 | 0.003417 | 0.24621 | 0.24621 | 0.025642
2019-03-28 15:49 | 1553802516.2847 | 615 | 0.20883 | 0.000917 | 7.5e-05 | 99.3629 | 98.6791 | 122.0219 | 21.2612 | 0.002488 | 0.002319 | 151.4291 | 0.25751 | 8.4049 | 0.002587 | 0.003903 | 0.003836 | 4.9e-05 | 0.001041 | 0.31492 | 1.3154 | 2.5474 | 100.61 | 21.11 | 25.779 | 81.6787 | 84.0429 | 77.6189 | 0.006423 | 1.0065 | 0.024359 | 24.2019 | 0.061767 | 0.15319 | 0.00338 | 0.24332 | 0.24332 | 0.026002
2019-03-28 15:49 | 1553802517.2835 | 616 | 0.20884 | 0.000913 | 7.4e-05 | 99.3316 | 98.5565 | 122.0036 | 21.2612 | 0.002443 | 0.0023 | 151.4227 | 0.25852 | 8.4033 | 0.002646 | 0.00381 | 0.003732 | 4.8e-05 | 0.001044 | 0.31293 | 1.3517 | 2.5475 | 100.61 | 21.11 | 25.778 | 80.5947 | 85.4638 | 77.1458 | 0.006423 | 1.0078 | 0.024374 | 24.1837 | 0.070074 | 0.15379 | 0.003333 | 0.24006 | 0.24006 | 0.026459
2019-03-28 15:49 | 1553802518.2835 | 617 | 0.20881 | 0.000907 | 7.3e-05 | 99.4508 | 98.6098 | 121.8799 | 21.2612 | 0.002453 | 0.002296 | 151.4089 | 0.2577 | 8.4033 | 0.002558 | 0.004198 | 0.004092 | 4.9e-05 | 0.001046 | 0.31006 | 1.273 | 2.5474 | 100.61 | 21.11 | 25.776 | 80.8216 | 83.3485 | 77.0584 | 0.006422 | 1.0081 | 0.024364 | 24.1685 | 0.075957 | 0.1533 | 0.003513 | 0.25321 | 0.25321 | 0.025005
2019-03-28 15:49 | 1553802519.2845 | 618 | 0.20884 | 0.000907 | 7.3e-05 | 99.4153 | 98.7726 | 122.0123 | 21.2612 | 0.002453 | 0.002348 | 151.4075 | 0.25758 | 8.4055 | 0.002618 | 0.003108 | 0.002923 | 4.5e-05 | 0.001047 | 0.32173 | 1.2961 | 2.5473 | 100.61 | 21.11 | 25.773 | 80.8372 | 84.8014 | 78.3159 | 0.006422 | 1.0045 | 0.024334 | 24.225 | 0.05802 | 0.15324 | 0.003336 | 0.23997 | 0.23997 | 0.026373
2019-03-28 15:49 | 1553802520.2845 | 619 | 0.20882 | 0.000907 | 7.3e-05 | 99.399 | 98.6523 | 122.1726 | 21.2612 | 0.002472 | 0.00232 | 151.3609 | 0.25732 | 8.4021 | 0.00254 | 0.004431 | 0.004352 | 4.9e-05 | 0.001049 | 0.33627 | 1.2759 | 2.5473 | 100.61 | 21.11 | 25.767 | 81.2787 | 82.9126 | 77.6314 | 0.00642 | 1.0065 | 0.024374 | 24.2172 | 0.067444 | 0.15308 | 0.003439 | 0.24791 | 0.24791 | 0.025501
2019-03-28 15:49 | 1553802521.2835 | 620 | 0.20885 | 0.00091 | 7.2e-05 | 99.3489 | 98.7488 | 122.347 | 21.2612 | 0.002465 | 0.002349 | 151.3987 | 0.25735 | 8.4031 | 0.002551 | 0.003308 | 0.003218 | 4.7e-05 | 0.001052 | 0.30845 | 1.3259 | 2.5473 | 100.61 | 21.11 | 25.763 | 81.1102 | 83.1771 | 78.3438 | 0.006419 | 1.0044 | 0.024366 | 24.2591 | 0.054196 | 0.1531 | 0.003289 | 0.23687 | 0.23687 | 0.026693
2019-03-28 15:49 | 1553802522.2835 | 621 | 0.20887 | 0.000909 | 7.2e-05 | 99.3871 | 98.8031 | 122.4336 | 21.2612 | 0.002459 | 0.002316 | 151.3734 | 0.25631 | 8.4037 | 0.002578 | 0.004654 | 0.004532 | 4.9e-05 | 0.001054 | 0.29645 | 1.3077 | 2.5474 | 100.61 | 21.11 | 25.758 | 80.9675 | 83.8337 | 77.548 | 0.006418 | 1.0067 | 0.024402 | 24.2402 | 0.052706 | 0.15248 | 0.003163 | 0.22796 | 0.22796 | 0.027624
2019-03-28 15:49 | 1553802523.2851 | 622 | 0.20887 | 0.000903 | 7.2e-05 | 99.4201 | 98.6958 | 122.413 | 21.2612 | 0.002466 | 0.002294 | 151.3774 | 0.25679 | 8.401 | 0.002521 | 0.003942 | 0.003882 | 4.8e-05 | 0.001059 | 0.31133 | 1.3072 | 2.5474 | 100.61 | 21.11 | 25.748 | 81.1561 | 82.4553 | 77.0233 | 0.006415 | 1.0082 | 0.024419 | 24.22 | 0.065405 | 0.15277 | 0.003173 | 0.22893 | 0.22893 | 0.02756
2019-03-28 15:49 | 1553802524.284 | 623 | 0.20887 | 0.000902 | 7.2e-05 | 99.421 | 98.9465 | 122.3423 | 21.2612 | 0.002453 | 0.002294 | 151.3954 | 0.25811 | 8.4042 | 0.002559 | 0.003788 | 0.00362 | 4.6e-05 | 0.001057 | 0.3418 | 1.3141 | 2.5475 | 100.61 | 21.11 | 25.742 | 80.8208 | 83.3908 | 77.017 | 0.006414 | 1.0082 | 0.024412 | 24.2128 | 0.042792 | 0.15355 | 0.003182 | 0.22953 | 0.22953 | 0.027629
2019-03-28 15:49 | 1553802525.284 | 624 | 0.20886 | 0.000899 | 7.1e-05 | 99.3531 | 98.5762 | 122.2942 | 21.2612 | 0.002417 | 0.002308 | 151.392 | 0.25718 | 8.4033 | 0.002473 | 0.003999 | 0.003989 | 4.8e-05 | 0.001057 | 0.35153 | 1.3205 | 2.5474 | 100.61 | 21.1 | 25.73 | 79.9701 | 81.3066 | 77.3469 | 0.006407 | 1.0073 | 0.024396 | 24.2195 | 0.070217 | 0.153 | 0.003214 | 0.23171 | 0.23171 | 0.02727
2019-03-28 15:49 | 1553802526.2841 | 625 | 0.20886 | 0.000904 | 7.1e-05 | 99.4687 | 98.6594 | 121.982 | 21.2612 | 0.002405 | 0.002305 | 151.3652 | 0.25757 | 8.4012 | 0.002491 | 0.003721 | 0.003611 | 4.8e-05 | 0.001052 | 0.3011 | 1.3044 | 2.5473 | 100.61 | 21.1 | 25.726 | 79.6887 | 81.7371 | 77.2698 | 0.006406 | 1.0075 | 0.024367 | 24.1859 | 0.073066 | 0.15323 | 0.003235 | 0.23303 | 0.23303 | 0.027157
2019-03-28 15:49 | 1553802527.2828 | 626 | 0.20888 | 0.000899 | 7.1e-05 | 99.4523 | 98.6927 | 121.955 | 21.2612 | 0.002449 | 0.002274 | 151.3393 | 0.25756 | 8.4032 | 0.002522 | 0.003891 | 0.003741 | 4.9e-05 | 0.001053 | 0.32228 | 1.3115 | 2.5474 | 100.61 | 21.1 | 25.722 | 80.7281 | 82.4928 | 76.5329 | 0.006405 | 1.0096 | 0.02439 | 24.1578 | 0.068575 | 0.15322 | 0.003131 | 0.22559 | 0.22559 | 0.028051
2019-03-28 15:49 | 1553802528.2841 | 627 | 0.20886 | 0.000899 | 7.1e-05 | 99.3123 | 98.6771 | 121.7839 | 21.2612 | 0.002461 | 0.002328 | 151.3368 | 0.25717 | 8.4023 | 0.002506 | 0.003566 | 0.003536 | 4.5e-05 | 0.001052 | 0.31653 | 1.2983 | 2.5474 | 100.61 | 21.1 | 25.717 | 81.0182 | 82.1005 | 77.8217 | 0.006404 | 1.0059 | 0.024328 | 24.1852 | 0.057392 | 0.15299 | 0.003208 | 0.23067 | 0.23067 | 0.027391
2019-03-28 15:49 | 1553802529.2828 | 628 | 0.20884 | 0.0009 | 7.1e-05 | 99.4326 | 98.7342 | 121.5722 | 21.2612 | 0.002463 | 0.002369 | 151.319 | 0.25739 | 8.4031 | 0.002531 | 0.004065 | 0.003952 | 5e-05 | 0.00105 | 0.29231 | 1.3226 | 2.5473 | 100.61 | 21.1 | 25.714 | 81.072 | 82.709 | 78.8123 | 0.006403 | 1.0031 | 0.024273 | 24.1983 | 0.063052 | 0.15312 | 0.003325 | 0.23865 | 0.23865 | 0.026498
2019-03-28 15:49 | 1553802530.2838 | 629 | 0.20884 | 0.000898 | 7.1e-05 | 99.4806 | 98.6193 | 121.2058 | 21.2612 | 0.002491 | 0.002307 | 151.3129 | 0.25705 | 8.4023 | 0.002471 | 0.003752 | 0.00367 | 4.7e-05 | 0.001054 | 0.33826 | 1.3333 | 2.5474 | 100.61 | 21.1 | 25.714 | 81.734 | 81.2645 | 77.3337 | 0.006403 | 1.0073 | 0.024287 | 24.111 | 0.077773 | 0.15292 | 0.003333 | 0.23944 | 0.23944 | 0.026377
2019-03-28 15:49 | 1553802531.2842 | 630 | 0.20886 | 0.000896 | 7.1e-05 | 99.4055 | 98.9693 | 121.1419 | 21.2612 | 0.002476 | 0.002357 | 151.2935 | 0.25784 | 8.4014 | 0.002511 | 0.004034 | 0.004009 | 4.8e-05 | 0.001049 | 0.33301 | 1.331 | 2.5475 | 100.61 | 21.1 | 25.714 | 81.3792 | 82.2152 | 78.5192 | 0.006403 | 1.0039 | 0.02424 | 24.1454 | 0.039341 | 0.15339 | 0.003234 | 0.23173 | 0.23173 | 0.027338
2019-03-28 15:49 | 1553802532.2833 | 631 | 0.20887 | 0.000897 | 7.1e-05 | 99.4725 | 98.7804 | 121.1408 | 21.2612 | 0.002479 | 0.002264 | 151.2894 | 0.25794 | 8.4044 | 0.002507 | 0.004018 | 0.003912 | 4.6e-05 | 0.00105 | 0.34202 | 1.3215 | 2.5475 | 100.61 | 21.1 | 25.713 | 81.4462 | 82.1412 | 76.2999 | 0.006403 | 1.0103 | 0.024317 | 24.0689 | 0.062457 | 0.15345 | 0.003176 | 0.22816 | 0.22816 | 0.027776
2019-03-28 15:49 | 1553802533.2829 | 632 | 0.20888 | 0.000899 | 7.1e-05 | 99.4815 | 98.7627 | 121.1771 | 21.2612 | 0.002498 | 0.002319 | 151.2974 | 0.25834 | 8.4017 | 0.002478 | 0.004187 | 0.004099 | 4.5e-05 | 0.001047 | 0.31813 | 1.3282 | 2.5474 | 100.61 | 21.1 | 25.711 | 81.908 | 81.4417 | 77.615 | 0.006402 | 1.0065 | 0.024275 | 24.1178 | 0.06486 | 0.15369 | 0.00311 | 0.22302 | 0.22302 | 0.028461
2019-03-28 15:49 | 1553802534.283 | 633 | 0.20888 | 0.000901 | 7e-05 | 99.4463 | 98.8324 | 121.2424 | 21.2612 | 0.00239 | 0.002321 | 151.2517 | 0.2581 | 8.4018 | 0.00246 | 0.003933 | 0.003886 | 4.8e-05 | 0.001045 | 0.30342 | 1.2866 | 2.5472 | 100.61 | 21.1 | 25.71 | 79.3246 | 80.9939 | 77.6516 | 0.006402 | 1.0064 | 0.02428 | 24.1256 | 0.055394 | 0.15354 | 0.003105 | 0.22274 | 0.22274 | 0.028469
2019-03-28 15:49 | 1553802535.2841 | 634 | 0.20888 | 0.000898 | 7.1e-05 | 99.4448 | 98.921 | 121.264 | 21.2612 | 0.002488 | 0.002276 | 151.2314 | 0.25828 | 8.4026 | 0.002513 | 0.003744 | 0.003678 | 4.8e-05 | 0.001045 | 0.31022 | 1.2701 | 2.5473 | 100.61 | 21.1 | 25.71 | 81.6832 | 82.2791 | 76.5891 | 0.006402 | 1.0095 | 0.024319 | 24.0912 | 0.047239 | 0.15365 | 0.003087 | 0.22175 | 0.22175 | 0.028616
2019-03-28 15:49 | 1553802536.2831 | 635 | 0.20891 | 0.000897 | 7e-05 | 99.3725 | 98.8918 | 121.4196 | 21.2612 | 0.002494 | 0.002306 | 151.2171 | 0.25882 | 8.4035 | 0.002501 | 0.003976 | 0.00398 | 5e-05 | 0.001047 | 0.29662 | 1.3073 | 2.5473 | 100.61 | 21.1 | 25.709 | 81.8065 | 81.9789 | 77.2969 | 0.006402 | 1.0074 | 0.02431 | 24.131 | 0.043369 | 0.15397 | 0.002935 | 0.21057 | 0.21057 | 0.030199
2019-03-28 15:49 | 1553802537.2844 | 636 | 0.20887 | 0.000899 | 7e-05 | 99.4442 | 98.9064 | 121.5893 | 21.2612 | 0.002425 | 0.002334 | 151.2319 | 0.2577 | 8.4029 | 0.002489 | 0.003521 | 0.00328 | 5.1e-05 | 0.001045 | 0.2959 | 1.3371 | 2.5474 | 100.61 | 21.1 | 25.706 | 80.1479 | 81.7051 | 77.9805 | 0.006401 | 1.0055 | 0.024303 | 24.1714 | 0.048506 | 0.1533 | 0.003163 | 0.22723 | 0.22723 | 0.027864
2019-03-28 15:49 | 1553802538.2841 | 637 | 0.20889 | 0.000902 | 6.9e-05 | 99.4847 | 98.9652 | 121.5267 | 21.2612 | 0.002427 | 0.002312 | 151.2247 | 0.25817 | 8.4006 | 0.00253 | 0.003874 | 0.003765 | 4.6e-05 | 0.001043 | 0.30596 | 1.2994 | 2.5475 | 100.61 | 21.11 | 25.702 | 80.2049 | 82.677 | 77.4505 | 0.006404 | 1.007 | 0.024315 | 24.1469 | 0.046836 | 0.15358 | 0.00306 | 0.21982 | 0.21982 | 0.028855
2019-03-28 15:49 | 1553802539.2841 | 638 | 0.20887 | 0.000898 | 6.9e-05 | 99.4523 | 99.0026 | 121.4574 | 21.2612 | 0.002476 | 0.002293 | 151.2132 | 0.25783 | 8.4022 | 0.002468 | 0.003434 | 0.003349 | 4.7e-05 | 0.001042 | 0.30845 | 1.3243 | 2.5473 | 100.61 | 21.11 | 25.699 | 81.384 | 81.1943 | 76.9961 | 0.006403 | 1.0083 | 0.024324 | 24.1244 | 0.040539 | 0.15338 | 0.00315 | 0.22648 | 0.22648 | 0.02797
2019-03-28 15:49 | 1553802540.2831 | 639 | 0.20887 | 0.000898 | 6.9e-05 | 99.4767 | 98.8711 | 121.7325 | 21.2612 | 0.002445 | 0.002308 | 151.2078 | 0.25742 | 8.4028 | 0.002464 | 0.004473 | 0.004317 | 4.6e-05 | 0.001046 | 0.31813 | 1.3 | 2.5475 | 100.61 | 21.11 | 25.696 | 80.6346 | 81.0915 | 77.3407 | 0.006402 | 1.0073 | 0.02434 | 24.1636 | 0.054622 | 0.15314 | 0.003156 | 0.22706 | 0.22706 | 0.027854
2019-03-28 15:49 | 1553802541.2832 | 640 | 0.20889 | 0.000894 | 6.9e-05 | 99.5407 | 99.0039 | 121.5799 | 21.2612 | 0.002476 | 0.002338 | 151.1929 | 0.25826 | 8.4013 | 0.00246 | 0.004104 | 0.004028 | 5e-05 | 0.001046 | 0.31852 | 1.3421 | 2.5475 | 100.61 | 21.11 | 25.692 | 81.3753 | 80.9949 | 78.0648 | 0.006401 | 1.0052 | 0.024299 | 24.1734 | 0.048369 | 0.15364 | 0.00305 | 0.21898 | 0.21898 | 0.028977
2019-03-28 15:49 | 1553802542.2827 | 641 | 0.20883 | 0.000894 | 6.9e-05 | 99.648 | 99.0946 | 121.5313 | 21.2612 | 0.002433 | 0.002305 | 151.1909 | 0.25867 | 8.4051 | 0.002524 | 0.003137 | 0.003024 | 4.9e-05 | 0.001047 | 0.2975 | 1.3512 | 2.5475 | 100.61 | 21.11 | 25.69 | 80.3452 | 82.5404 | 77.2868 | 0.006401 | 1.0074 | 0.024321 | 24.1417 | 0.049808 | 0.15389 | 0.003398 | 0.24465 | 0.24465 | 0.025978
2019-03-28 15:49 | 1553802543.2822 | 642 | 0.20888 | 0.000896 | 6.9e-05 | 99.5348 | 98.9329 | 121.5933 | 21.2612 | 0.002432 | 0.002336 | 151.1517 | 0.2575 | 8.4048 | 0.002554 | 0.003845 | 0.003766 | 4.5e-05 | 0.001048 | 0.32045 | 1.291 | 2.5474 | 100.61 | 21.11 | 25.689 | 80.3234 | 83.2629 | 78.0103 | 0.006401 | 1.0054 | 0.024303 | 24.1728 | 0.054251 | 0.15319 | 0.003135 | 0.2252 | 0.2252 | 0.028093
2019-03-28 15:49 | 1553802544.2838 | 643 | 0.20884 | 0.000891 | 6.9e-05 | 99.5235 | 99.097 | 121.6975 | 21.2612 | 0.002462 | 0.00233 | 151.1538 | 0.25827 | 8.4011 | 0.00254 | 0.004057 | 0.003942 | 4.6e-05 | 0.001046 | 0.31974 | 1.3133 | 2.5473 | 100.61 | 21.11 | 25.689 | 81.0564 | 82.9297 | 77.8839 | 0.006401 | 1.0057 | 0.024317 | 24.1788 | 0.038411 | 0.15365 | 0.00335 | 0.24107 | 0.24107 | 0.026322
2019-03-28 15:49 | 1553802545.2823 | 644 | 0.20892 | 0.000894 | 6.9e-05 | 99.5741 | 99.0559 | 121.9196 | 21.2612 | 0.002469 | 0.002286 | 151.1264 | 0.25834 | 8.403 | 0.002535 | 0.003712 | 0.003683 | 5e-05 | 0.001045 | 0.31045 | 1.3156 | 2.5475 | 100.61 | 21.11 | 25.69 | 81.2256 | 82.7931 | 76.8308 | 0.006401 | 1.0088 | 0.024376 | 24.1645 | 0.046672 | 0.15368 | 0.002898 | 0.20849 | 0.20849 | 0.030444
2019-03-28 15:49 | 1553802546.2834 | 645 | 0.20887 | 0.000892 | 6.9e-05 | 99.6512 | 99.0569 | 122.0245 | 21.2612 | 0.002495 | 0.002328 | 151.1216 | 0.25813 | 8.4015 | 0.002476 | 0.004013 | 0.00398 | 4.8e-05 | 0.001041 | 0.31437 | 1.2891 | 2.5474 | 100.61 | 21.11 | 25.691 | 81.8416 | 81.3736 | 77.8176 | 0.006401 | 1.0059 | 0.024352 | 24.209 | 0.053506 | 0.15356 | 0.003173 | 0.22847 | 0.22847 | 0.027759
2019-03-28 15:49 | 1553802547.2837 | 646 | 0.20889 | 0.000893 | 6.8e-05 | 99.6652 | 99.1252 | 121.8527 | 21.2612 | 0.002526 | 0.002292 | 151.1202 | 0.25855 | 8.4025 | 0.00247 | 0.004511 | 0.004356 | 4.8e-05 | 0.001039 | 0.31277 | 1.3184 | 2.5474 | 100.61 | 21.11 | 25.692 | 82.5888 | 81.2333 | 76.9594 | 0.006401 | 1.0084 | 0.024365 | 24.1623 | 0.048599 | 0.15381 | 0.003047 | 0.21934 | 0.21934 | 0.028961
2019-03-28 15:49 | 1553802548.2824 | 647 | 0.20887 | 0.000898 | 6.8e-05 | 99.642 | 99.1011 | 122.0557 | 21.2612 | 0.002507 | 0.002286 | 151.0929 | 0.25903 | 8.4007 | 0.002448 | 0.004633 | 0.00453 | 4.4e-05 | 0.001046 | 0.31094 | 1.3145 | 2.5473 | 100.61 | 21.11 | 25.693 | 82.121 | 80.7038 | 76.8191 | 0.006402 | 1.0088 | 0.02439 | 24.1776 | 0.048691 | 0.1541 | 0.00319 | 0.23011 | 0.23011 | 0.027657
2019-03-28 15:49 | 1553802549.2828 | 648 | 0.20885 | 0.000895 | 6.8e-05 | 99.6179 | 98.9985 | 122.335 | 21.2612 | 0.002458 | 0.002341 | 151.0638 | 0.25863 | 8.405 | 0.002499 | 0.004112 | 0.004052 | 4.5e-05 | 0.001041 | 0.29717 | 1.3247 | 2.5474 | 100.61 | 21.11 | 25.697 | 80.9636 | 81.9383 | 78.1427 | 0.006403 | 1.005 | 0.024372 | 24.251 | 0.05579 | 0.15386 | 0.003328 | 0.24004 | 0.24004 | 0.026473
2019-03-28 15:49 | 1553802550.2829 | 649 | 0.20884 | 0.000894 | 6.7e-05 | 99.6277 | 99.1857 | 121.994 | 21.2612 | 0.002425 | 0.00233 | 151.0474 | 0.25818 | 8.4044 | 0.00241 | 0.003715 | 0.003588 | 4.8e-05 | 0.00104 | 0.33732 | 1.3059 | 2.5474 | 100.61 | 21.11 | 25.699 | 80.1479 | 79.7963 | 77.8721 | 0.006403 | 1.0058 | 0.024347 | 24.2078 | 0.039775 | 0.15359 | 0.003382 | 0.24385 | 0.24385 | 0.026013
2019-03-28 15:49 | 1553802551.2835 | 650 | 0.20894 | 0.000896 | 6.6e-05 | 99.6259 | 99.1299 | 122.1735 | 21.2612 | 0.002527 | 0.002307 | 151.0123 | 0.25864 | 8.4036 | 0.002461 | 0.003912 | 0.003815 | 4.9e-05 | 0.001036 | 0.27489 | 1.3545 | 2.5474 | 100.61 | 21.11 | 25.704 | 82.613 | 81.0174 | 77.3243 | 0.006404 | 1.0073 | 0.024384 | 24.2067 | 0.044641 | 0.15386 | 0.00277 | 0.19935 | 0.19935 | 0.031876
2019-03-28 15:49 | 1553802552.2829 | 651 | 0.20885 | 0.000891 | 6.5e-05 | 99.6974 | 99.0787 | 122.4153 | 21.2612 | 0.002466 | 0.002331 | 151.0141 | 0.25815 | 8.4011 | 0.00246 | 0.003246 | 0.003137 | 4.6e-05 | 0.001037 | 0.29767 | 1.3343 | 2.5473 | 100.61 | 21.11 | 25.706 | 81.1343 | 81.0025 | 77.8918 | 0.006405 | 1.0057 | 0.024389 | 24.2503 | 0.055686 | 0.15358 | 0.003301 | 0.23841 | 0.23841 | 0.026604
2019-03-28 15:49 | 1553802553.284 | 652 | 0.20888 | 0.000893 | 6.5e-05 | 99.6641 | 99.0936 | 122.2409 | 21.2612 | 0.002453 | 0.002342 | 151.0212 | 0.2578 | 8.4026 | 0.00244 | 0.003981 | 0.003955 | 4.6e-05 | 0.001038 | 0.31902 | 1.3584 | 2.5474 | 100.61 | 21.11 | 25.708 | 80.8225 | 80.5207 | 78.1537 | 0.006405 | 1.005 | 0.024362 | 24.242 | 0.051344 | 0.15336 | 0.003101 | 0.22348 | 0.22348 | 0.028343
2019-03-28 15:49 | 1553802554.2836 | 653 | 0.20887 | 0.000885 | 6.5e-05 | 99.7019 | 99.2961 | 122.1513 | 21.2612 | 0.002468 | 0.002277 | 151.025 | 0.25842 | 8.3975 | 0.002446 | 0.003292 | 0.003154 | 4.5e-05 | 0.001045 | 0.31437 | 1.2998 | 2.5473 | 100.61 | 21.1 | 25.712 | 81.2044 | 80.6579 | 76.5985 | 0.006402 | 1.0094 | 0.024407 | 24.1794 | 0.03648 | 0.15373 | 0.003192 | 0.23063 | 0.23063 | 0.02753
2019-03-28 15:49 | 1553802555.2829 | 654 | 0.20884 | 0.000887 | 6.4e-05 | 99.7257 | 99.2815 | 122.0304 | 21.2612 | 0.002481 | 0.002296 | 151.0201 | 0.25808 | 8.4031 | 0.002433 | 0.003964 | 0.003911 | 4.9e-05 | 0.001049 | 0.31797 | 1.3375 | 2.5475 | 100.61 | 21.1 | 25.713 | 81.4994 | 80.3476 | 77.0663 | 0.006403 | 1.0081 | 0.024379 | 24.1836 | 0.039934 | 0.15353 | 0.003355 | 0.24231 | 0.24231 | 0.026168
2019-03-28 15:49 | 1553802556.2826 | 655 | 0.2089 | 0.000885 | 6.4e-05 | 99.7859 | 99.3589 | 122.048 | 21.2612 | 0.002419 | 0.002336 | 151.0249 | 0.25845 | 8.403 | 0.002443 | 0.003896 | 0.003851 | 4.7e-05 | 0.001055 | 0.32892 | 1.3363 | 2.5475 | 100.61 | 21.1 | 25.715 | 80.0201 | 80.5954 | 78.0228 | 0.006403 | 1.0053 | 0.024348 | 24.2184 | 0.038353 | 0.15375 | 0.003025 | 0.21781 | 0.21781 | 0.029153
2019-03-28 15:49 | 1553802557.2827 | 656 | 0.20889 | 0.000886 | 6.3e-05 | 99.7892 | 99.2781 | 121.9006 | 21.2612 | 0.002456 | 0.002367 | 150.9973 | 0.25799 | 8.4017 | 0.002463 | 0.004627 | 0.004444 | 4.4e-05 | 0.001056 | 0.29446 | 1.3252 | 2.5475 | 100.61 | 21.1 | 25.717 | 80.9153 | 81.0804 | 78.7773 | 0.006404 | 1.0032 | 0.024307 | 24.2297 | 0.045928 | 0.15348 | 0.003095 | 0.22266 | 0.22266 | 0.028467
2019-03-28 15:49 | 1553802558.2827 | 657 | 0.20889 | 0.000888 | 6.3e-05 | 99.7183 | 99.3144 | 122.0653 | 21.2612 | 0.002496 | 0.002291 | 151.0186 | 0.25883 | 8.4012 | 0.002385 | 0.003652 | 0.003464 | 4.6e-05 | 0.001058 | 0.29286 | 1.3162 | 2.5475 | 100.61 | 21.1 | 25.72 | 81.8698 | 79.1999 | 76.9298 | 0.006404 | 1.0085 | 0.024387 | 24.1824 | 0.036297 | 0.15398 | 0.003065 | 0.22127 | 0.22127 | 0.02874
2019-03-28 15:49 | 1553802559.2827 | 658 | 0.20892 | 0.00089 | 6.2e-05 | 99.6751 | 99.0861 | 121.6935 | 21.2612 | 0.002459 | 0.002323 | 151.0304 | 0.25754 | 8.4078 | 0.002459 | 0.003784 | 0.003727 | 4.8e-05 | 0.001059 | 0.33428 | 1.3006 | 2.5473 | 100.61 | 21.1 | 25.726 | 80.9814 | 80.965 | 77.7015 | 0.006406 | 1.0063 | 0.024323 | 24.1721 | 0.053007 | 0.15321 | 0.002873 | 0.20666 | 0.20666 | 0.030618
2019-03-28 15:49 | 1553802560.2818 | 659 | 0.20883 | 0.000887 | 6.2e-05 | 99.7743 | 99.4116 | 122.1083 | 21.2612 | 0.002407 | 0.002323 | 151.0233 | 0.25847 | 8.4003 | 0.002467 | 0.003322 | 0.003288 | 4.5e-05 | 0.001053 | 0.33931 | 1.2879 | 2.5473 | 100.61 | 21.1 | 25.729 | 79.7223 | 81.1742 | 77.7023 | 0.006407 | 1.0063 | 0.024365 | 24.2133 | 0.032572 | 0.15376 | 0.003414 | 0.24667 | 0.24667 | 0.025745
2019-03-28 15:49 | 1553802561.2819 | 660 | 0.20886 | 0.00089 | 6.1e-05 | 99.8398 | 99.3779 | 121.6872 | 21.2612 | 0.002454 | 0.002315 | 151.047 | 0.25805 | 8.4005 | 0.002438 | 0.004016 | 0.003806 | 5e-05 | 0.001046 | 0.35601 | 1.3033 | 2.5473 | 100.61 | 21.1 | 25.739 | 80.8482 | 80.4755 | 77.506 | 0.006409 | 1.0068 | 0.024329 | 24.1648 | 0.041471 | 0.15351 | 0.003258 | 0.23495 | 0.23495 | 0.026985
2019-03-28 15:49 | 1553802562.283 | 661 | 0.20886 | 0.00089 | 6e-05 | 99.7531 | 99.2849 | 121.5255 | 21.2612 | 0.002484 | 0.00229 | 151.023 | 0.25792 | 8.4033 | 0.002365 | 0.003566 | 0.003521 | 4.6e-05 | 0.001041 | 0.3361 | 1.3146 | 2.5474 | 100.61 | 21.1 | 25.743 | 81.5828 | 78.7244 | 76.9135 | 0.00641 | 1.0085 | 0.024334 | 24.1283 | 0.042085 | 0.15344 | 0.003285 | 0.23705 | 0.23705 | 0.026733
2019-03-28 15:49 | 1553802563.282 | 662 | 0.20889 | 0.000889 | 6e-05 | 99.7659 | 99.3593 | 121.6619 | 21.2612 | 0.002513 | 0.002286 | 151.0189 | 0.25828 | 8.4043 | 0.002475 | 0.004103 | 0.004036 | 4.6e-05 | 0.001041 | 0.31205 | 1.2981 | 2.5474 | 100.61 | 21.1 | 25.749 | 82.2628 | 81.3698 | 76.8159 | 0.006412 | 1.0088 | 0.024351 | 24.1385 | 0.036533 | 0.15365 | 0.003064 | 0.22098 | 0.22098 | 0.028716
2019-03-28 15:49 | 1553802564.2819 | 663 | 0.20888 | 0.000884 | 5.9e-05 | 99.8008 | 99.258 | 121.6445 | 21.2612 | 0.002445 | 0.002323 | 150.9956 | 0.25778 | 8.4023 | 0.002401 | 0.003433 | 0.003221 | 4.6e-05 | 0.00104 | 0.34324 | 1.2942 | 2.5475 | 100.61 | 21.1 | 25.759 | 80.6509 | 79.5851 | 77.7133 | 0.006414 | 1.0062 | 0.024318 | 24.1677 | 0.048775 | 0.15335 | 0.003119 | 0.22472 | 0.22472 | 0.028183
2019-03-28 15:49 | 1553802565.2823 | 664 | 0.20889 | 0.000888 | 5.8e-05 | 99.8148 | 99.4228 | 121.537 | 21.2612 | 0.002464 | 0.00232 | 151.0005 | 0.25761 | 8.4019 | 0.002393 | 0.003938 | 0.003892 | 4.6e-05 | 0.001035 | 0.31238 | 1.278 | 2.5475 | 100.61 | 21.1 | 25.764 | 81.0939 | 79.3924 | 77.629 | 0.006415 | 1.0065 | 0.02431 | 24.1541 | 0.035194 | 0.15326 | 0.003067 | 0.22093 | 0.22093 | 0.02865
2019-03-28 15:49 | 1553802566.283 | 665 | 0.20888 | 0.000886 | 5.8e-05 | 99.8183 | 99.4017 | 121.4912 | 21.2612 | 0.002482 | 0.002274 | 150.973 | 0.25741 | 8.403 | 0.0025 | 0.003637 | 0.003564 | 4.8e-05 | 0.001037 | 0.31614 | 1.2711 | 2.5472 | 100.61 | 21.11 | 25.773 | 81.5275 | 81.9681 | 76.53 | 0.006422 | 1.0096 | 0.024344 | 24.1117 | 0.03741 | 0.15313 | 0.003119 | 0.22501 | 0.22501 | 0.028107
2019-03-28 15:49 | 1553802567.282 | 666 | 0.2089 | 0.000889 | 5.8e-05 | 99.8037 | 99.4829 | 121.509 | 21.2612 | 0.002475 | 0.002276 | 150.9661 | 0.25792 | 8.4006 | 0.002345 | 0.004384 | 0.004182 | 4.5e-05 | 0.001037 | 0.31332 | 1.3096 | 2.5472 | 100.61 | 21.11 | 25.779 | 81.3643 | 78.2417 | 76.5829 | 0.006423 | 1.0095 | 0.024344 | 24.1153 | 0.028799 | 0.15344 | 0.003037 | 0.21909 | 0.21909 | 0.028924
2019-03-28 15:49 | 1553802568.2823 | 667 | 0.20888 | 0.000888 | 5.8e-05 | 99.8204 | 99.5702 | 121.5818 | 21.2612 | 0.002434 | 0.00228 | 150.9304 | 0.25811 | 8.4026 | 0.002417 | 0.003999 | 0.00391 | 4.5e-05 | 0.001037 | 0.33356 | 1.3019 | 2.5473 | 100.61 | 21.11 | 25.785 | 80.3741 | 79.9631 | 76.6749 | 0.006425 | 1.0092 | 0.024348 | 24.1257 | 0.022447 | 0.15355 | 0.003127 | 0.22571 | 0.22571 | 0.028096
2019-03-28 15:49 | 1553802569.2821 | 668 | 0.20887 | 0.000881 | 5.8e-05 | 99.8934 | 99.451 | 121.8738 | 21.2612 | 0.002382 | 0.002256 | 150.9226 | 0.25721 | 8.4009 | 0.002388 | 0.004738 | 0.004682 | 4.5e-05 | 0.001039 | 0.31725 | 1.298 | 2.5473 | 100.61 | 21.11 | 25.794 | 79.1242 | 79.267 | 76.0935 | 0.006427 | 1.0109 | 0.024397 | 24.1345 | 0.039702 | 0.15301 | 0.003232 | 0.23386 | 0.23386 | 0.027022
2019-03-28 15:49 | 1553802570.2827 | 669 | 0.20889 | 0.000884 | 5.7e-05 | 99.9086 | 99.4591 | 121.8118 | 21.2612 | 0.002448 | 0.002259 | 150.9225 | 0.25825 | 8.4031 | 0.002402 | 0.004524 | 0.004451 | 4.7e-05 | 0.001034 | 0.31686 | 1.3072 | 2.5474 | 100.61 | 21.11 | 25.798 | 80.7187 | 79.6154 | 76.173 | 0.006428 | 1.0107 | 0.024388 | 24.1311 | 0.040329 | 0.15363 | 0.003082 | 0.22279 | 0.22279 | 0.028481
2019-03-28 15:50 | 1553802571.2825 | 670 | 0.20887 | 0.000885 | 5.7e-05 | 99.8895 | 99.627 | 121.9196 | 21.2612 | 0.002397 | 0.00228 | 150.8864 | 0.25811 | 8.4015 | 0.0024 | 0.003878 | 0.003717 | 4.6e-05 | 0.001032 | 0.30934 | 1.2884 | 2.5474 | 100.61 | 21.11 | 25.803 | 79.4962 | 79.5601 | 76.6804 | 0.006429 | 1.0092 | 0.024381 | 24.1593 | 0.023542 | 0.15355 | 0.0032 | 0.23143 | 0.23143 | 0.027402
2019-03-28 15:50 | 1553802572.2821 | 671 | 0.20889 | 0.000878 | 5.7e-05 | 99.8574 | 99.5923 | 122.0074 | 21.2612 | 0.002523 | 0.002354 | 150.8772 | 0.25807 | 8.405 | 0.002443 | 0.003432 | 0.003359 | 4.7e-05 | 0.001034 | 0.33052 | 1.3169 | 2.5474 | 100.61 | 21.11 | 25.811 | 82.5039 | 80.6026 | 78.4646 | 0.006431 | 1.0041 | 0.024328 | 24.2296 | 0.023772 | 0.15353 | 0.003103 | 0.22382 | 0.22382 | 0.02833
2019-03-28 15:50 | 1553802573.2826 | 672 | 0.20889 | 0.000882 | 5.6e-05 | 99.9476 | 99.4044 | 122.0297 | 21.2612 | 0.002428 | 0.002329 | 150.8652 | 0.25803 | 8.4044 | 0.002443 | 0.004021 | 0.003917 | 4.5e-05 | 0.001029 | 0.31382 | 1.3336 | 2.5474 | 100.61 | 21.11 | 25.813 | 80.2392 | 80.5901 | 77.8637 | 0.006432 | 1.0058 | 0.024351 | 24.2111 | 0.048741 | 0.1535 | 0.003066 | 0.22136 | 0.22136 | 0.028639
2019-03-28 15:50 | 1553802574.2822 | 673 | 0.20888 | 0.000876 | 5.7e-05 | 99.9661 | 99.4554 | 121.9252 | 21.2612 | 0.002498 | 0.002321 | 150.8791 | 0.25784 | 8.4029 | 0.002431 | 0.004572 | 0.00446 | 4.8e-05 | 0.001035 | 0.30038 | 1.3291 | 2.5475 | 100.61 | 21.11 | 25.817 | 81.9087 | 80.3085 | 77.6532 | 0.006432 | 1.0064 | 0.024348 | 24.1935 | 0.04581 | 0.15339 | 0.003133 | 0.22621 | 0.22621 | 0.028005
2019-03-28 15:50 | 1553802575.2813 | 674 | 0.20888 | 0.000878 | 5.6e-05 | 99.9309 | 99.4014 | 121.9304 | 21.2612 | 0.002492 | 0.002291 | 150.8624 | 0.25806 | 8.4043 | 0.002388 | 0.004667 | 0.004583 | 5e-05 | 0.001034 | 0.30917 | 1.2998 | 2.5472 | 100.61 | 21.11 | 25.819 | 81.7794 | 79.2732 | 76.9329 | 0.006433 | 1.0085 | 0.024374 | 24.1691 | 0.047524 | 0.15352 | 0.003143 | 0.22723 | 0.22723 | 0.027902
2019-03-28 15:50 | 1553802576.2824 | 675 | 0.20888 | 0.000881 | 5.6e-05 | 99.9911 | 99.5437 | 121.7919 | 21.2612 | 0.002473 | 0.002303 | 150.868 | 0.25735 | 8.4018 | 0.002425 | 0.003428 | 0.003356 | 4.7e-05 | 0.001034 | 0.31421 | 1.3031 | 2.5474 | 100.61 | 21.11 | 25.822 | 81.3083 | 80.1712 | 77.2191 | 0.006434 | 1.0076 | 0.02435 | 24.1653 | 0.040108 | 0.1531 | 0.003181 | 0.22973 | 0.22973 | 0.027523
2019-03-28 15:50 | 1553802577.2823 | 676 | 0.2089 | 0.000878 | 5.5e-05 | 99.9092 | 99.6766 | 121.7353 | 21.2612 | 0.002468 | 0.002216 | 150.8542 | 0.25854 | 8.4022 | 0.002373 | 0.004152 | 0.004023 | 5.2e-05 | 0.00104 | 0.31957 | 1.3314 | 2.5474 | 100.61 | 21.11 | 25.828 | 81.2006 | 78.9036 | 75.1455 | 0.006435 | 1.0136 | 0.024417 | 24.0881 | 0.02085 | 0.1538 | 0.003021 | 0.2187 | 0.2187 | 0.029045
2019-03-28 15:50 | 1553802578.2824 | 677 | 0.20887 | 0.000881 | 5.5e-05 | 100.031 | 99.7037 | 121.4821 | 21.2612 | 0.002426 | 0.002335 | 150.8583 | 0.25821 | 8.4014 | 0.002376 | 0.00357 | 0.003448 | 5e-05 | 0.001045 | 0.34379 | 1.3167 | 2.5474 | 100.61 | 21.11 | 25.833 | 80.1931 | 78.9752 | 77.9923 | 0.006436 | 1.0054 | 0.024292 | 24.1611 | 0.029311 | 0.15361 | 0.003216 | 0.23183 | 0.23183 | 0.027365
2019-03-28 15:50 | 1553802579.2813 | 678 | 0.20889 | 0.000879 | 5.5e-05 | 100.0271 | 99.5495 | 121.1034 | 21.2612 | 0.00242 | 0.002276 | 150.8549 | 0.25769 | 8.4016 | 0.002363 | 0.002973 | 0.002916 | 4.8e-05 | 0.001044 | 0.30071 | 1.3316 | 2.5474 | 100.61 | 21.11 | 25.837 | 80.0434 | 78.6783 | 76.5884 | 0.006437 | 1.0095 | 0.024303 | 24.0752 | 0.042812 | 0.1533 | 0.003094 | 0.22304 | 0.22304 | 0.028387
2019-03-28 15:50 | 1553802580.2824 | 679 | 0.20888 | 0.000881 | 5.4e-05 | 100.0358 | 99.7075 | 121.3623 | 21.2612 | 0.00249 | 0.002312 | 150.8395 | 0.25786 | 8.4014 | 0.002366 | 0.003945 | 0.003798 | 5e-05 | 0.001048 | 0.31725 | 1.3351 | 2.5474 | 100.61 | 21.11 | 25.847 | 81.7207 | 78.7297 | 77.4529 | 0.00644 | 1.007 | 0.024299 | 24.1307 | 0.029402 | 0.1534 | 0.003138 | 0.22626 | 0.22626 | 0.028
2019-03-28 15:50 | 1553802581.2825 | 680 | 0.2089 | 0.000879 | 5.5e-05 | 100.0102 | 99.7605 | 121.3584 | 21.2612 | 0.00241 | 0.002278 | 150.8593 | 0.25823 | 8.4004 | 0.002347 | 0.003864 | 0.003773 | 4.7e-05 | 0.001047 | 0.31238 | 1.3248 | 2.5474 | 100.61 | 21.11 | 25.852 | 79.8096 | 78.2892 | 76.6374 | 0.006441 | 1.0093 | 0.024327 | 24.1022 | 0.022358 | 0.15362 | 0.003018 | 0.21773 | 0.21773 | 0.029139
2019-03-28 15:50 | 1553802582.2815 | 681 | 0.20888 | 0.000876 | 5.5e-05 | 100.0504 | 99.7292 | 121.5412 | 21.2612 | 0.002501 | 0.002328 | 150.8467 | 0.25737 | 8.4015 | 0.002364 | 0.00361 | 0.003504 | 4.8e-05 | 0.001043 | 0.31741 | 1.3144 | 2.5473 | 100.61 | 21.11 | 25.863 | 81.996 | 78.6932 | 77.8169 | 0.006444 | 1.0059 | 0.024304 | 24.161 | 0.028756 | 0.15311 | 0.003145 | 0.2268 | 0.2268 | 0.027881
2019-03-28 15:50 | 1553802583.2815 | 682 | 0.20892 | 0.000878 | 5.4e-05 | 100.0614 | 99.8009 | 121.6921 | 21.2612 | 0.00242 | 0.002242 | 150.8367 | 0.25767 | 8.4022 | 0.002384 | 0.003974 | 0.003876 | 4.7e-05 | 0.001041 | 0.30829 | 1.3222 | 2.5473 | 100.61 | 21.11 | 25.867 | 80.0434 | 79.1663 | 75.7629 | 0.006445 | 1.0118 | 0.024391 | 24.1051 | 0.023315 | 0.15329 | 0.002916 | 0.21086 | 0.21086 | 0.030025
2019-03-28 15:50 | 1553802584.2825 | 683 | 0.20895 | 0.000875 | 5.4e-05 | 100.0224 | 99.7635 | 121.7834 | 21.2612 | 0.002458 | 0.002295 | 150.7964 | 0.25734 | 8.4026 | 0.002351 | 0.004546 | 0.004477 | 4.9e-05 | 0.001042 | 0.30093 | 1.3052 | 2.5473 | 100.61 | 21.11 | 25.875 | 80.9643 | 78.3914 | 77.0327 | 0.006447 | 1.0082 | 0.024355 | 24.158 | 0.023176 | 0.15309 | 0.002744 | 0.19788 | 0.19788 | 0.031953
2019-03-28 15:50 | 1553802585.2815 | 684 | 0.20894 | 0.000872 | 5.4e-05 | 100.0483 | 99.7197 | 121.6675 | 21.2612 | 0.002479 | 0.002283 | 150.7735 | 0.25802 | 8.4056 | 0.002443 | 0.003943 | 0.0039 | 4.9e-05 | 0.001045 | 0.30685 | 1.2955 | 2.5475 | 100.61 | 21.11 | 25.879 | 81.4611 | 80.5877 | 76.7395 | 0.006448 | 1.009 | 0.024354 | 24.1364 | 0.029423 | 0.15349 | 0.002813 | 0.20301 | 0.20301 | 0.031227
2019-03-28 15:50 | 1553802586.2816 | 685 | 0.20894 | 0.000878 | 5.4e-05 | 100.0939 | 99.666 | 121.829 | 21.2612 | 0.002445 | 0.002319 | 150.7684 | 0.25797 | 8.4038 | 0.002438 | 0.003732 | 0.003625 | 5e-05 | 0.001045 | 0.31669 | 1.3663 | 2.5474 | 100.61 | 21.11 | 25.881 | 80.6516 | 80.4762 | 77.6206 | 0.006448 | 1.0065 | 0.02434 | 24.1828 | 0.038312 | 0.15347 | 0.002765 | 0.19929 | 0.19929 | 0.031804
2019-03-28 15:50 | 1553802587.2817 | 686 | 0.20894 | 0.00087 | 5.4e-05 | 100.1665 | 100.0506 | 122.0376 | 21.2612 | 0.002421 | 0.002295 | 150.7445 | 0.25741 | 8.4012 | 0.002425 | 0.003474 | 0.003349 | 4.9e-05 | 0.001043 | 0.32421 | 1.3202 | 2.5474 | 100.61 | 21.11 | 25.884 | 80.0653 | 80.1573 | 77.0295 | 0.006449 | 1.0082 | 0.024381 | 24.1831 | 0.010358 | 0.15313 | 0.002804 | 0.20252 | 0.20252 | 0.03123
2019-03-28 15:50 | 1553802588.2817 | 687 | 0.20893 | 0.000872 | 5.4e-05 | 100.1585 | 99.9888 | 122.1339 | 21.2612 | 0.002501 | 0.002286 | 150.7407 | 0.25777 | 8.4043 | 0.002355 | 0.00323 | 0.003233 | 5e-05 | 0.001036 | 0.30829 | 1.3003 | 2.5474 | 100.61 | 21.11 | 25.885 | 81.9938 | 78.4733 | 76.8207 | 0.006449 | 1.0088 | 0.024398 | 24.1854 | 0.015169 | 0.15335 | 0.002867 | 0.20729 | 0.20729 | 0.030553
2019-03-28 15:50 | 1553802589.2817 | 688 | 0.20896 | 0.00087 | 5.4e-05 | 100.1698 | 99.9748 | 122.5848 | 21.2612 | 0.002419 | 0.002293 | 150.7047 | 0.25851 | 8.4033 | 0.002377 | 0.004128 | 0.003866 | 5.1e-05 | 0.001037 | 0.30773 | 1.3157 | 2.5473 | 100.61 | 21.11 | 25.884 | 80.0037 | 79.0065 | 76.9921 | 0.006449 | 1.0083 | 0.024437 | 24.2359 | 0.017426 | 0.15379 | 0.002703 | 0.19557 | 0.19557 | 0.032477
2019-03-28 15:50 | 1553802590.2817 | 689 | 0.20892 | 0.00087 | 5.4e-05 | 100.132 | 99.9028 | 122.2308 | 21.2612 | 0.002474 | 0.002293 | 150.6952 | 0.25785 | 8.4031 | 0.002362 | 0.004206 | 0.004143 | 5.2e-05 | 0.001035 | 0.31454 | 1.3476 | 2.5473 | 100.61 | 21.11 | 25.88 | 81.3261 | 78.6379 | 76.9952 | 0.006448 | 1.0083 | 0.024401 | 24.201 | 0.020493 | 0.15339 | 0.002926 | 0.21161 | 0.21161 | 0.029938
2019-03-28 15:50 | 1553802591.2822 | 690 | 0.20892 | 0.000864 | 5.4e-05 | 100.1966 | 99.9337 | 122.5045 | 21.2612 | 0.002425 | 0.002286 | 150.6812 | 0.25849 | 8.4032 | 0.002373 | 0.003616 | 0.003468 | 5.1e-05 | 0.001042 | 0.3178 | 1.3297 | 2.5474 | 100.61 | 21.11 | 25.878 | 80.1503 | 78.906 | 76.8308 | 0.006448 | 1.0088 | 0.024435 | 24.2224 | 0.023497 | 0.15378 | 0.002905 | 0.21041 | 0.21041 | 0.030184
2019-03-28 15:50 | 1553802592.2818 | 691 | 0.20892 | 0.000865 | 5.4e-05 | 100.2365 | 99.8407 | 121.9935 | 21.2612 | 0.002453 | 0.002331 | 150.684 | 0.25788 | 8.4038 | 0.002325 | 0.004072 | 0.004026 | 5.2e-05 | 0.001041 | 0.31116 | 1.318 | 2.5475 | 100.61 | 21.11 | 25.872 | 80.8365 | 77.746 | 77.8894 | 0.006446 | 1.0057 | 0.024347 | 24.2084 | 0.035395 | 0.15341 | 0.002909 | 0.20992 | 0.20992 | 0.030182
2019-03-28 15:50 | 1553802593.2808 | 692 | 0.20892 | 0.000865 | 5.4e-05 | 100.1883 | 100.1797 | 121.786 | 21.2612 | 0.002503 | 0.002336 | 150.6767 | 0.25837 | 8.4014 | 0.002357 | 0.004226 | 0.004032 | 4.9e-05 | 0.001043 | 0.3194 | 1.3023 | 2.5473 | 100.61 | 21.11 | 25.869 | 82.0241 | 78.5192 | 78.0173 | 0.006445 | 1.0053 | 0.024322 | 24.1922 | 0.000765 | 0.1537 | 0.002903 | 0.20925 | 0.20925 | 0.030336
2019-03-28 15:50 | 1553802594.2813 | 693 | 0.20892 | 0.000866 | 5.5e-05 | 100.2815 | 99.984 | 121.9191 | 21.2612 | 0.002474 | 0.002314 | 150.6913 | 0.25804 | 8.4021 | 0.002333 | 0.003919 | 0.0039 | 5.1e-05 | 0.001044 | 0.31166 | 1.2769 | 2.5473 | 100.61 | 21.11 | 25.868 | 81.3434 | 77.9502 | 77.4872 | 0.006445 | 1.0069 | 0.024353 | 24.1871 | 0.026573 | 0.15351 | 0.002924 | 0.21104 | 0.21104 | 0.030041
2019-03-28 15:50 | 1553802595.282 | 694 | 0.2089 | 0.000865 | 5.5e-05 | 100.3244 | 99.8542 | 122.0869 | 21.2612 | 0.002478 | 0.002239 | 150.7145 | 0.25815 | 8.4025 | 0.002395 | 0.003701 | 0.003577 | 5e-05 | 0.001041 | 0.29734 | 1.3667 | 2.5474 | 100.61 | 21.11 | 25.866 | 81.4354 | 79.4369 | 75.682 | 0.006445 | 1.0121 | 0.024433 | 24.1414 | 0.042015 | 0.15358 | 0.00308 | 0.22326 | 0.22326 | 0.028409
2019-03-28 15:50 | 1553802596.2819 | 695 | 0.20891 | 0.000867 | 5.5e-05 | 100.3032 | 100.019 | 122.1243 | 21.2612 | 0.002466 | 0.002259 | 150.7005 | 0.25806 | 8.4025 | 0.002399 | 0.003989 | 0.003967 | 4.8e-05 | 0.00104 | 0.30182 | 1.341 | 2.5475 | 100.61 | 21.11 | 25.867 | 81.1398 | 79.5383 | 76.1745 | 0.006445 | 1.0107 | 0.024419 | 24.1621 | 0.025382 | 0.15352 | 0.002976 | 0.21543 | 0.21543 | 0.02943
2019-03-28 15:50 | 1553802597.2819 | 696 | 0.2089 | 0.000867 | 5.5e-05 | 100.2303 | 99.9962 | 122.1674 | 21.2612 | 0.002456 | 0.002266 | 150.7193 | 0.25779 | 8.4042 | 0.002347 | 0.003596 | 0.003398 | 5.1e-05 | 0.001042 | 0.33019 | 1.3114 | 2.5475 | 100.61 | 21.12 | 25.869 | 80.9066 | 78.2861 | 76.3412 | 0.006449 | 1.0102 | 0.024418 | 24.1722 | 0.020909 | 0.15336 | 0.003034 | 0.21972 | 0.21972 | 0.028826
2019-03-28 15:50 | 1553802598.281 | 697 | 0.20891 | 0.000869 | 5.4e-05 | 100.2368 | 100.0751 | 122.0022 | 21.2612 | 0.002433 | 0.0023 | 150.7029 | 0.25739 | 8.4038 | 0.002318 | 0.003209 | 0.003093 | 4.5e-05 | 0.001036 | 0.33284 | 1.325 | 2.5472 | 100.61 | 21.11 | 25.871 | 80.3445 | 77.5956 | 77.1542 | 0.006446 | 1.0078 | 0.024373 | 24.1839 | 0.014446 | 0.15312 | 0.002973 | 0.21487 | 0.21487 | 0.029431
2019-03-28 15:50 | 1553802599.2821 | 698 | 0.20888 | 0.000867 | 5.5e-05 | 100.364 | 100.0594 | 121.7266 | 21.2612 | 0.002471 | 0.0023 | 150.709 | 0.25748 | 8.4041 | 0.002371 | 0.00343 | 0.003366 | 4.7e-05 | 0.001035 | 0.3115 | 1.369 | 2.5475 | 100.61 | 21.11 | 25.873 | 81.2645 | 78.8522 | 77.1598 | 0.006446 | 1.0078 | 0.024345 | 24.1567 | 0.027185 | 0.15318 | 0.003189 | 0.23038 | 0.23038 | 0.02746
2019-03-28 15:50 | 1553802600.2831 | 699 | 0.20891 | 0.000867 | 5.5e-05 | 100.3452 | 100.2065 | 121.7593 | 21.2612 | 0.002449 | 0.002264 | 150.6907 | 0.25755 | 8.4028 | 0.002323 | 0.004231 | 0.004014 | 4.6e-05 | 0.00103 | 0.31332 | 1.3688 | 2.5475 | 100.61 | 21.11 | 25.877 | 80.7367 | 77.7102 | 76.3008 | 0.006447 | 1.0103 | 0.024379 | 24.1303 | 0.012372 | 0.15322 | 0.002989 | 0.21604 | 0.21604 | 0.02929
2019-03-28 15:50 | 1553802601.2812 | 700 | 0.2089 | 0.000868 | 5.5e-05 | 100.3649 | 100.1386 | 121.5513 | 21.2612 | 0.002452 | 0.002278 | 150.6687 | 0.25837 | 8.4006 | 0.002405 | 0.004344 | 0.004264 | 4.9e-05 | 0.001036 | 0.30237 | 1.3202 | 2.5474 | 100.61 | 21.11 | 25.88 | 80.8194 | 79.6748 | 76.636 | 0.006448 | 1.0093 | 0.024346 | 24.1213 | 0.020191 | 0.1537 | 0.003068 | 0.22158 | 0.22158 | 0.028649
2019-03-28 15:50 | 1553802602.2811 | 701 | 0.2089 | 0.000871 | 5.5e-05 | 100.4578 | 100.0676 | 121.9555 | 21.2612 | 0.00245 | 0.002305 | 150.654 | 0.25709 | 8.4029 | 0.002444 | 0.004 | 0.003904 | 4.6e-05 | 0.001039 | 0.30879 | 1.3329 | 2.5474 | 100.61 | 21.11 | 25.884 | 80.7569 | 80.6142 | 77.2736 | 0.006449 | 1.0075 | 0.024364 | 24.1834 | 0.034814 | 0.15294 | 0.003055 | 0.22075 | 0.22075 | 0.028613
2019-03-28 15:50 | 1553802603.2812 | 702 | 0.20892 | 0.000863 | 5.4e-05 | 100.44 | 100.3193 | 121.7776 | 21.2612 | 0.002447 | 0.002238 | 150.6173 | 0.25827 | 8.4028 | 0.00232 | 0.00414 | 0.003997 | 4.5e-05 | 0.00104 | 0.31205 | 1.3342 | 2.5474 | 100.61 | 21.11 | 25.886 | 80.679 | 77.6376 | 75.657 | 0.00645 | 1.0122 | 0.024403 | 24.1099 | 0.010749 | 0.15364 | 0.002955 | 0.21381 | 0.21381 | 0.029678
2019-03-28 15:50 | 1553802604.2813 | 703 | 0.20893 | 0.000867 | 5.5e-05 | 100.4662 | 100.3662 | 121.9811 | 21.2612 | 0.002429 | 0.002315 | 150.614 | 0.25825 | 8.404 | 0.00241 | 0.00343 | 0.003372 | 4.8e-05 | 0.001036 | 0.29662 | 1.3137 | 2.5474 | 100.61 | 21.11 | 25.891 | 80.2626 | 79.7894 | 77.5223 | 0.006451 | 1.0068 | 0.024358 | 24.1945 | 0.008904 | 0.15363 | 0.002861 | 0.20649 | 0.20649 | 0.030728
2019-03-28 15:50 | 1553802605.2802 | 704 | 0.20894 | 0.000864 | 5.5e-05 | 100.5332 | 100.2358 | 121.9815 | 21.2612 | 0.002441 | 0.00229 | 150.5931 | 0.25757 | 8.4038 | 0.002437 | 0.003242 | 0.003104 | 4.6e-05 | 0.001037 | 0.31238 | 1.276 | 2.5473 | 100.61 | 21.11 | 25.893 | 80.5324 | 80.4366 | 76.926 | 0.006451 | 1.0085 | 0.024379 | 24.174 | 0.026502 | 0.15323 | 0.002789 | 0.20136 | 0.20136 | 0.031429
2019-03-28 15:50 | 1553802606.2812 | 705 | 0.20894 | 0.000862 | 5.5e-05 | 100.4828 | 100.2735 | 121.8431 | 21.2612 | 0.002472 | 0.002299 | 150.5732 | 0.25743 | 8.4024 | 0.002405 | 0.004554 | 0.004414 | 4.6e-05 | 0.001032 | 0.30845 | 1.2762 | 2.5473 | 100.61 | 21.11 | 25.895 | 81.3003 | 79.6841 | 77.141 | 0.006452 | 1.0079 | 0.024358 | 24.1676 | 0.018658 | 0.15315 | 0.002834 | 0.20447 | 0.20447 | 0.030934
2019-03-28 15:50 | 1553802607.2812 | 706 | 0.20893 | 0.000862 | 5.5e-05 | 100.5332 | 100.3496 | 121.7776 | 21.2612 | 0.002456 | 0.002302 | 150.5546 | 0.25882 | 8.4051 | 0.002431 | 0.0041 | 0.003978 | 4.7e-05 | 0.00103 | 0.32139 | 1.3227 | 2.5475 | 100.61 | 21.11 | 25.896 | 80.8949 | 80.2962 | 77.2136 | 0.006452 | 1.0077 | 0.024349 | 24.1637 | 0.016352 | 0.15397 | 0.002878 | 0.20764 | 0.20764 | 0.030624
2019-03-28 15:50 | 1553802608.2818 | 707 | 0.20893 | 0.000857 | 5.5e-05 | 100.5934 | 100.2775 | 121.7534 | 21.2612 | 0.002428 | 0.002328 | 150.5462 | 0.25782 | 8.4035 | 0.002497 | 0.0038 | 0.00377 | 4.5e-05 | 0.001029 | 0.33212 | 1.3019 | 2.5475 | 100.61 | 21.11 | 25.895 | 80.2229 | 81.8868 | 77.8262 | 0.006452 | 1.0059 | 0.024325 | 24.1824 | 0.028125 | 0.15338 | 0.002869 | 0.20679 | 0.20679 | 0.030632
2019-03-28 15:50 | 1553802609.2814 | 708 | 0.20896 | 0.000861 | 5.5e-05 | 100.6276 | 100.4549 | 121.9689 | 21.2612 | 0.002422 | 0.002277 | 150.5222 | 0.25853 | 8.4006 | 0.00248 | 0.004191 | 0.004095 | 4.2e-05 | 0.001029 | 0.30309 | 1.308 | 2.5473 | 100.61 | 21.11 | 25.893 | 80.0965 | 81.4789 | 76.5939 | 0.006451 | 1.0094 | 0.024389 | 24.1612 | 0.015366 | 0.1538 | 0.002684 | 0.1937 | 0.1937 | 0.032792
2019-03-28 15:50 | 1553802610.2803 | 709 | 0.20893 | 0.000853 | 5.5e-05 | 100.5945 | 100.336 | 122.1078 | 21.2612 | 0.00248 | 0.002291 | 150.5267 | 0.25743 | 8.4059 | 0.002504 | 0.004086 | 0.004047 | 4.6e-05 | 0.001026 | 0.31708 | 1.2887 | 2.5473 | 100.61 | 21.11 | 25.892 | 81.4876 | 82.064 | 76.9438 | 0.006451 | 1.0084 | 0.024391 | 24.1871 | 0.023021 | 0.15314 | 0.002865 | 0.20706 | 0.20706 | 0.030546
2019-03-28 15:50 | 1553802611.2814 | 710 | 0.20895 | 0.000859 | 5.5e-05 | 100.6601 | 100.1896 | 122.1097 | 21.2612 | 0.002434 | 0.002298 | 150.5356 | 0.25805 | 8.4027 | 0.002514 | 0.004144 | 0.004044 | 4.4e-05 | 0.00103 | 0.30613 | 1.2923 | 2.5473 | 100.61 | 21.11 | 25.891 | 80.3858 | 82.3041 | 77.1191 | 0.006451 | 1.0079 | 0.024385 | 24.1933 | 0.041908 | 0.15351 | 0.002777 | 0.20054 | 0.20054 | 0.031615
2019-03-28 15:50 | 1553802612.2819 | 711 | 0.20894 | 0.000856 | 5.5e-05 | 100.759 | 100.5201 | 122.0717 | 21.2612 | 0.002504 | 0.002313 | 150.5172 | 0.25772 | 8.4016 | 0.002436 | 0.003641 | 0.00354 | 4.4e-05 | 0.001027 | 0.30453 | 1.3082 | 2.5474 | 100.61 | 21.11 | 25.885 | 82.0609 | 80.4178 | 77.4731 | 0.006449 | 1.0069 | 0.024369 | 24.2018 | 0.021228 | 0.15332 | 0.002798 | 0.20194 | 0.20194 | 0.031356
2019-03-28 15:50 | 1553802613.2805 | 712 | 0.20895 | 0.000851 | 5.5e-05 | 100.7917 | 100.5048 | 122.4282 | 21.2612 | 0.002457 | 0.002333 | 150.5217 | 0.25786 | 8.4019 | 0.002405 | 0.003922 | 0.003752 | 4.4e-05 | 0.00103 | 0.30127 | 1.2962 | 2.5474 | 100.61 | 21.11 | 25.884 | 80.9215 | 79.6786 | 77.9408 | 0.006449 | 1.0056 | 0.024388 | 24.2533 | 0.025497 | 0.1534 | 0.002783 | 0.201 | 0.201 | 0.031519
2019-03-28 15:50 | 1553802614.2819 | 713 | 0.20894 | 0.000846 | 5.5e-05 | 100.7581 | 100.6281 | 122.4118 | 21.2612 | 0.002504 | 0.002315 | 150.5114 | 0.25782 | 8.4026 | 0.002341 | 0.003893 | 0.00382 | 4.6e-05 | 0.001029 | 0.30198 | 1.3215 | 2.5473 | 100.61 | 21.11 | 25.88 | 82.0515 | 78.1412 | 77.506 | 0.006448 | 1.0068 | 0.024402 | 24.2366 | 0.011541 | 0.15338 | 0.002866 | 0.20717 | 0.20717 | 0.030576
2019-03-28 15:50 | 1553802615.2819 | 714 | 0.20896 | 0.000847 | 5.5e-05 | 100.7801 | 100.6332 | 122.2947 | 21.2612 | 0.002479 | 0.002298 | 150.5104 | 0.25859 | 8.4013 | 0.00245 | 0.004357 | 0.004281 | 4.5e-05 | 0.001028 | 0.32549 | 1.3083 | 2.5475 | 100.61 | 21.11 | 25.878 | 81.4486 | 80.7562 | 77.1098 | 0.006448 | 1.008 | 0.024404 | 24.2113 | 0.013045 | 0.15383 | 0.002738 | 0.19779 | 0.19779 | 0.032122
2019-03-28 15:50 | 1553802616.2806 | 715 | 0.20894 | 0.000846 | 5.5e-05 | 100.8033 | 100.4878 | 122.197 | 21.2612 | 0.002419 | 0.002315 | 150.5001 | 0.25789 | 8.4038 | 0.002425 | 0.004185 | 0.004062 | 4.2e-05 | 0.001031 | 0.31238 | 1.3382 | 2.5475 | 100.61 | 21.11 | 25.876 | 80.0107 | 80.1698 | 77.5175 | 0.006447 | 1.0068 | 0.02438 | 24.2157 | 0.028039 | 0.15342 | 0.002858 | 0.20644 | 0.20644 | 0.030692
2019-03-28 15:50 | 1553802617.282 | 716 | 0.20893 | 0.00084 | 5.6e-05 | 100.8021 | 100.7012 | 122.1778 | 21.2612 | 0.002472 | 0.0023 | 150.5009 | 0.25759 | 8.4033 | 0.002394 | 0.003519 | 0.003454 | 4.5e-05 | 0.001035 | 0.33467 | 1.362 | 2.5474 | 100.61 | 21.11 | 25.872 | 81.295 | 79.4066 | 77.1535 | 0.006446 | 1.0078 | 0.024391 | 24.2013 | 0.008964 | 0.15324 | 0.002892 | 0.20898 | 0.20898 | 0.030284
2019-03-28 15:50 | 1553802618.2797 | 717 | 0.20895 | 0.000841 | 5.6e-05 | 100.831 | 100.6794 | 122.2069 | 21.2612 | 0.00246 | 0.002293 | 150.5133 | 0.258 | 8.402 | 0.002364 | 0.004375 | 0.004218 | 4.3e-05 | 0.001037 | 0.31293 | 1.3182 | 2.5475 | 100.61 | 21.11 | 25.87 | 81.0011 | 78.7008 | 76.9781 | 0.006446 | 1.0083 | 0.0244 | 24.1981 | 0.013458 | 0.15348 | 0.002807 | 0.20284 | 0.20284 | 0.031251
2019-03-28 15:50 | 1553802619.2807 | 718 | 0.20895 | 0.00084 | 5.6e-05 | 100.9091 | 100.4654 | 121.7226 | 21.2612 | 0.002462 | 0.002325 | 150.5023 | 0.25743 | 8.4031 | 0.002444 | 0.003809 | 0.003716 | 4.3e-05 | 0.001035 | 0.32836 | 1.2953 | 2.5474 | 100.61 | 21.11 | 25.867 | 81.0439 | 80.612 | 77.7655 | 0.006445 | 1.0061 | 0.024324 | 24.1772 | 0.039413 | 0.15315 | 0.002811 | 0.20246 | 0.20246 | 0.031241
2019-03-28 15:50 | 1553802620.2801 | 719 | 0.20896 | 0.000835 | 5.6e-05 | 100.8706 | 100.6638 | 121.6034 | 21.2612 | 0.002496 | 0.002331 | 150.524 | 0.25743 | 8.402 | 0.00242 | 0.002809 | 0.002786 | 4.9e-05 | 0.001039 | 0.32549 | 1.2833 | 2.5473 | 100.61 | 21.11 | 25.866 | 81.8722 | 80.0458 | 77.8902 | 0.006445 | 1.0057 | 0.024308 | 24.1697 | 0.01836 | 0.15314 | 0.002762 | 0.19869 | 0.19869 | 0.031833
2019-03-28 15:50 | 1553802621.2798 | 720 | 0.20896 | 0.000835 | 5.7e-05 | 100.896 | 100.6767 | 121.6809 | 21.2612 | 0.002479 | 0.002322 | 150.5026 | 0.25791 | 8.4028 | 0.002476 | 0.003732 | 0.003609 | 4.5e-05 | 0.001035 | 0.34219 | 1.2693 | 2.5474 | 100.61 | 21.11 | 25.867 | 81.4635 | 81.3746 | 77.6734 | 0.006445 | 1.0063 | 0.024323 | 24.1699 | 0.019458 | 0.15343 | 0.002709 | 0.19495 | 0.19495 | 0.032504
2019-03-28 15:50 | 1553802622.2798 | 721 | 0.20896 | 0.00084 | 5.7e-05 | 100.9207 | 100.728 | 121.5041 | 21.2612 | 0.00243 | 0.00229 | 150.5029 | 0.25795 | 8.4037 | 0.002492 | 0.003493 | 0.003422 | 4.4e-05 | 0.001033 | 0.3079 | 1.285 | 2.5475 | 100.61 | 21.11 | 25.869 | 80.2806 | 81.7739 | 76.9041 | 0.006445 | 1.0085 | 0.024332 | 24.1258 | 0.017093 | 0.15345 | 0.0027 | 0.19437 | 0.19437 | 0.032606
2019-03-28 15:50 | 1553802623.2811 | 722 | 0.20895 | 0.000835 | 5.7e-05 | 100.9913 | 100.7199 | 121.3628 | 21.2612 | 0.002449 | 0.002291 | 150.5026 | 0.25767 | 8.4024 | 0.002492 | 0.003863 | 0.00372 | 4.5e-05 | 0.001028 | 0.32012 | 1.2903 | 2.5474 | 100.61 | 21.11 | 25.872 | 80.7257 | 81.7739 | 76.9416 | 0.006446 | 1.0084 | 0.024317 | 24.1131 | 0.024071 | 0.15329 | 0.00281 | 0.20225 | 0.20225 | 0.031302
2019-03-28 15:50 | 1553802624.2802 | 723 | 0.20892 | 0.000831 | 5.7e-05 | 100.9213 | 100.8551 | 120.9758 | 21.2612 | 0.002459 | 0.002281 | 150.4863 | 0.25668 | 8.4017 | 0.002438 | 0.004279 | 0.004135 | 4.8e-05 | 0.001022 | 0.3235 | 1.3289 | 2.5475 | 100.61 | 21.11 | 25.875 | 80.973 | 80.4738 | 76.7076 | 0.006447 | 1.0091 | 0.024286 | 24.0666 | 0.00587 | 0.1527 | 0.002979 | 0.21438 | 0.21438 | 0.029418
2019-03-28 15:50 | 1553802625.2802 | 724 | 0.20896 | 0.000834 | 5.7e-05 | 101.0094 | 100.8707 | 121.2166 | 21.2612 | 0.002484 | 0.002251 | 150.4605 | 0.25767 | 8.4005 | 0.00243 | 0.003601 | 0.003521 | 4.6e-05 | 0.001019 | 0.31996 | 1.3069 | 2.5474 | 100.61 | 21.11 | 25.882 | 81.5703 | 80.2796 | 75.9695 | 0.006449 | 1.0112 | 0.024336 | 24.0651 | 0.012295 | 0.15329 | 0.00274 | 0.19732 | 0.19732 | 0.032084
2019-03-28 15:50 | 1553802626.2803 | 725 | 0.20891 | 0.000839 | 5.7e-05 | 101.0455 | 100.8571 | 121.4682 | 21.2612 | 0.002461 | 0.002297 | 150.4573 | 0.2578 | 8.4003 | 0.002481 | 0.003643 | 0.003514 | 4.7e-05 | 0.001019 | 0.33483 | 1.3167 | 2.5474 | 100.61 | 21.11 | 25.886 | 81.0314 | 81.4994 | 77.085 | 0.00645 | 1.008 | 0.024322 | 24.1285 | 0.016691 | 0.15337 | 0.003062 | 0.22075 | 0.22075 | 0.028693
2019-03-28 15:50 | 1553802627.281 | 726 | 0.20897 | 0.000834 | 5.7e-05 | 101.072 | 100.8425 | 121.6076 | 21.2612 | 0.002469 | 0.002278 | 150.4644 | 0.25718 | 8.4037 | 0.002456 | 0.004206 | 0.004133 | 4.4e-05 | 0.001024 | 0.31437 | 1.2779 | 2.5473 | 100.61 | 21.11 | 25.895 | 81.21 | 80.8966 | 76.6197 | 0.006452 | 1.0094 | 0.024352 | 24.1263 | 0.020333 | 0.153 | 0.002651 | 0.19086 | 0.19086 | 0.033107
2019-03-28 15:50 | 1553802628.2813 | 727 | 0.209 | 0.000828 | 5.6e-05 | 101.0669 | 100.8785 | 121.5403 | 21.2612 | 0.00246 | 0.00231 | 150.4353 | 0.25851 | 8.405 | 0.002465 | 0.003971 | 0.003884 | 4.7e-05 | 0.00103 | 0.33284 | 1.3067 | 2.5473 | 100.61 | 21.11 | 25.899 | 81.0081 | 81.1095 | 77.407 | 0.006453 | 1.0071 | 0.024318 | 24.1468 | 0.016691 | 0.15379 | 0.002496 | 0.17935 | 0.17935 | 0.035413
2019-03-28 15:50 | 1553802629.281 | 728 | 0.20898 | 0.000831 | 5.6e-05 | 101.0684 | 100.8687 | 121.8365 | 21.2612 | 0.002444 | 0.002331 | 150.4443 | 0.25689 | 8.4035 | 0.002382 | 0.003853 | 0.003752 | 4.7e-05 | 0.001031 | 0.30431 | 1.3185 | 2.5475 | 100.61 | 21.11 | 25.904 | 80.6142 | 79.1211 | 77.8909 | 0.006454 | 1.0057 | 0.024331 | 24.1928 | 0.017696 | 0.15282 | 0.002626 | 0.18896 | 0.18896 | 0.033401
2019-03-28 15:50 | 1553802630.2804 | 729 | 0.20899 | 0.000827 | 5.6e-05 | 101.0523 | 100.8479 | 121.8177 | 21.2612 | 0.002478 | 0.002338 | 150.4061 | 0.25809 | 8.4022 | 0.002411 | 0.004143 | 0.004007 | 5e-05 | 0.001031 | 0.3063 | 1.3135 | 2.5474 | 100.61 | 21.11 | 25.908 | 81.4361 | 79.8158 | 78.0742 | 0.006455 | 1.0052 | 0.024323 | 24.1973 | 0.018111 | 0.15354 | 0.002553 | 0.18356 | 0.18356 | 0.034544
2019-03-28 15:50 | 1553802631.2801 | 730 | 0.20899 | 0.000831 | 5.6e-05 | 101.0693 | 100.8911 | 121.7027 | 21.2612 | 0.00249 | 0.002303 | 150.3693 | 0.25802 | 8.4041 | 0.002449 | 0.00384 | 0.003787 | 4.8e-05 | 0.00103 | 0.32925 | 1.2981 | 2.5472 | 100.61 | 21.11 | 25.909 | 81.7183 | 80.7264 | 77.2198 | 0.006455 | 1.0076 | 0.024341 | 24.1564 | 0.015787 | 0.1535 | 0.002565 | 0.18454 | 0.18454 | 0.034352
2019-03-28 15:51 | 1553802632.281 | 731 | 0.20898 | 0.00083 | 5.6e-05 | 101.0791 | 100.8466 | 121.887 | 21.2612 | 0.002466 | 0.002351 | 150.3662 | 0.25647 | 8.4013 | 0.002444 | 0.003272 | 0.003148 | 4.8e-05 | 0.001027 | 0.32388 | 1.3443 | 2.5474 | 100.61 | 21.11 | 25.906 | 81.1523 | 80.612 | 78.3798 | 0.006455 | 1.0043 | 0.024319 | 24.2147 | 0.020604 | 0.15257 | 0.002632 | 0.18932 | 0.18932 | 0.033285
2019-03-28 15:51 | 1553802633.2805 | 732 | 0.20897 | 0.000827 | 5.6e-05 | 101.0911 | 100.922 | 121.9679 | 21.2612 | 0.002474 | 0.002323 | 150.3566 | 0.2578 | 8.4047 | 0.002504 | 0.004908 | 0.004706 | 4.9e-05 | 0.001026 | 0.29247 | 1.2833 | 2.5473 | 100.61 | 21.11 | 25.904 | 81.3378 | 82.0546 | 77.7155 | 0.006454 | 1.0062 | 0.02435 | 24.1998 | 0.014971 | 0.15336 | 0.002696 | 0.19422 | 0.19422 | 0.032611
2019-03-28 15:51 | 1553802634.2813 | 733 | 0.20895 | 0.000824 | 5.7e-05 | 101.1012 | 100.7926 | 122.1735 | 21.2612 | 0.002456 | 0.002349 | 150.3529 | 0.25701 | 8.4006 | 0.002463 | 0.00388 | 0.003852 | 4.7e-05 | 0.001031 | 0.34418 | 1.3145 | 2.5474 | 100.61 | 21.11 | 25.9 | 80.905 | 81.0727 | 78.3361 | 0.006453 | 1.0044 | 0.024349 | 24.2416 | 0.027346 | 0.1529 | 0.002819 | 0.20316 | 0.20316 | 0.031081
2019-03-28 15:51 | 1553802635.2791 | 734 | 0.20895 | 0.00082 | 5.7e-05 | 101.1244 | 101.0457 | 122.0259 | 21.2612 | 0.002516 | 0.002331 | 150.3458 | 0.25739 | 8.4017 | 0.002433 | 0.004345 | 0.00425 | 4.8e-05 | 0.001034 | 0.33284 | 1.3177 | 2.5473 | 100.61 | 21.11 | 25.89 | 82.3472 | 80.3608 | 77.9089 | 0.006451 | 1.0057 | 0.024349 | 24.2123 | 0.006969 | 0.15312 | 0.002796 | 0.20149 | 0.20149 | 0.031385
2019-03-28 15:51 | 1553802636.2802 | 735 | 0.20897 | 0.000822 | 5.7e-05 | 101.1092 | 101.0694 | 122.0311 | 21.2612 | 0.002453 | 0.002296 | 150.3456 | 0.25753 | 8.4037 | 0.002465 | 0.004621 | 0.00451 | 4.8e-05 | 0.001031 | 0.31454 | 1.3068 | 2.5475 | 100.61 | 21.11 | 25.884 | 80.8271 | 81.1148 | 77.0552 | 0.006449 | 1.0081 | 0.024379 | 24.1833 | 0.003521 | 0.15321 | 0.002719 | 0.19612 | 0.19612 | 0.032262
2019-03-28 15:51 | 1553802637.2807 | 736 | 0.20897 | 0.000819 | 5.6e-05 | 101.1226 | 100.9142 | 121.6487 | 21.2612 | 0.002466 | 0.00231 | 150.3088 | 0.25678 | 8.4036 | 0.002484 | 0.004035 | 0.003904 | 4.9e-05 | 0.00103 | 0.31022 | 1.3284 | 2.5474 | 100.61 | 21.11 | 25.879 | 81.1359 | 81.5866 | 77.3952 | 0.006448 | 1.0071 | 0.024329 | 24.1571 | 0.018457 | 0.15276 | 0.002662 | 0.19156 | 0.19156 | 0.032935
2019-03-28 15:51 | 1553802638.2803 | 737 | 0.20897 | 0.000824 | 5.6e-05 | 101.1286 | 100.8377 | 121.9508 | 21.2612 | 0.002475 | 0.002274 | 150.3125 | 0.25755 | 8.4019 | 0.002442 | 0.003515 | 0.003443 | 4.9e-05 | 0.001027 | 0.32156 | 1.3182 | 2.5474 | 100.61 | 21.11 | 25.868 | 81.3551 | 80.5728 | 76.5308 | 0.006445 | 1.0096 | 0.02439 | 24.1573 | 0.025761 | 0.15322 | 0.002688 | 0.19393 | 0.19393 | 0.03263
2019-03-28 15:51 | 1553802639.2804 | 738 | 0.20898 | 0.000823 | 5.6e-05 | 101.1235 | 101.098 | 121.978 | 21.2612 | 0.002446 | 0.002323 | 150.2963 | 0.25792 | 8.403 | 0.002486 | 0.003674 | 0.0036 | 4.5e-05 | 0.001027 | 0.3314 | 1.303 | 2.5474 | 100.61 | 21.11 | 25.863 | 80.6649 | 81.6349 | 77.7155 | 0.006444 | 1.0062 | 0.024351 | 24.2008 | 0.00226 | 0.15344 | 0.00264 | 0.19018 | 0.19018 | 0.033321
2019-03-28 15:51 | 1553802640.2797 | 739 | 0.20896 | 0.000824 | 5.6e-05 | 101.1125 | 100.9767 | 121.8879 | 21.2612 | 0.002436 | 0.002274 | 150.3034 | 0.25905 | 8.402 | 0.002494 | 0.004056 | 0.003834 | 4.4e-05 | 0.001028 | 0.31277 | 1.3404 | 2.5474 | 100.61 | 21.11 | 25.855 | 80.4192 | 81.8253 | 76.5214 | 0.006442 | 1.0096 | 0.024384 | 24.1507 | 0.012022 | 0.15411 | 0.002776 | 0.20039 | 0.20039 | 0.031761
2019-03-28 15:51 | 1553802641.2794 | 740 | 0.20895 | 0.000825 | 5.6e-05 | 101.1593 | 100.9546 | 121.5773 | 21.2612 | 0.002466 | 0.002288 | 150.3509 | 0.25815 | 8.4042 | 0.002523 | 0.003569 | 0.00355 | 4.6e-05 | 0.001029 | 0.30685 | 1.335 | 2.5474 | 100.61 | 21.11 | 25.852 | 81.1477 | 82.5178 | 76.8705 | 0.006441 | 1.0086 | 0.024341 | 24.132 | 0.018114 | 0.15358 | 0.002788 | 0.20087 | 0.20087 | 0.031576
2019-03-28 15:51 | 1553802642.2794 | 741 | 0.20899 | 0.000822 | 5.6e-05 | 101.1334 | 101.0874 | 121.5184 | 21.2612 | 0.002478 | 0.002292 | 150.3207 | 0.25777 | 8.404 | 0.002468 | 0.003979 | 0.003848 | 4.8e-05 | 0.001029 | 0.30342 | 1.3529 | 2.5475 | 100.61 | 21.1 | 25.852 | 81.4354 | 81.1835 | 76.968 | 0.006437 | 1.0084 | 0.024331 | 24.1295 | 0.004061 | 0.15335 | 0.002571 | 0.18491 | 0.18491 | 0.034251
2019-03-28 15:51 | 1553802643.2805 | 742 | 0.20895 | 0.000822 | 5.6e-05 | 101.142 | 100.9172 | 121.4572 | 21.2612 | 0.002427 | 0.002263 | 150.3261 | 0.25783 | 8.4038 | 0.002507 | 0.004069 | 0.004017 | 4.7e-05 | 0.001031 | 0.3246 | 1.3487 | 2.5475 | 100.61 | 21.1 | 25.854 | 80.2042 | 82.1186 | 76.2554 | 0.006438 | 1.0104 | 0.02435 | 24.0988 | 0.019898 | 0.15338 | 0.002828 | 0.20388 | 0.20388 | 0.031071
2019-03-28 15:51 | 1553802644.2805 | 743 | 0.20898 | 0.000821 | 5.7e-05 | 101.17 | 101.101 | 121.1687 | 21.2612 | 0.002414 | 0.002258 | 150.3321 | 0.25813 | 8.4043 | 0.002484 | 0.003554 | 0.00344 | 5e-05 | 0.001039 | 0.31758 | 1.3196 | 2.5473 | 100.61 | 21.1 | 25.855 | 79.8946 | 81.5828 | 76.1581 | 0.006438 | 1.0107 | 0.024324 | 24.0668 | 0.006099 | 0.15356 | 0.002626 | 0.18889 | 0.18889 | 0.033576
2019-03-28 15:51 | 1553802645.2799 | 744 | 0.20898 | 0.000821 | 5.7e-05 | 101.1462 | 100.9142 | 121.3431 | 21.2612 | 0.002392 | 0.002257 | 150.3282 | 0.25821 | 8.4019 | 0.002487 | 0.003266 | 0.003125 | 4.7e-05 | 0.001037 | 0.29214 | 1.364 | 2.5474 | 100.61 | 21.1 | 25.863 | 79.3737 | 81.6412 | 76.1339 | 0.00644 | 1.0108 | 0.024343 | 24.0833 | 0.020538 | 0.15361 | 0.002621 | 0.18865 | 0.18865 | 0.033629
2019-03-28 15:51 | 1553802646.2796 | 745 | 0.20896 | 0.000825 | 5.7e-05 | 101.1569 | 101.1483 | 121.3997 | 21.2612 | 0.002452 | 0.002303 | 150.3261 | 0.25768 | 8.4024 | 0.002447 | 0.003955 | 0.003868 | 4.9e-05 | 0.001038 | 0.32892 | 1.3101 | 2.5475 | 100.61 | 21.1 | 25.867 | 80.8091 | 80.6992 | 77.2205 | 0.006441 | 1.0076 | 0.024311 | 24.1264 | 0.000763 | 0.15329 | 0.002761 | 0.19864 | 0.19864 | 0.031872
2019-03-28 15:51 | 1553802647.2806 | 746 | 0.20894 | 0.00082 | 5.7e-05 | 101.1617 | 100.8432 | 121.6928 | 21.2612 | 0.002392 | 0.002278 | 150.3231 | 0.25808 | 8.4039 | 0.002461 | 0.003657 | 0.003573 | 4.8e-05 | 0.001035 | 0.32045 | 1.3184 | 2.5474 | 100.61 | 21.1 | 25.873 | 79.3559 | 81.0189 | 76.6297 | 0.006442 | 1.0093 | 0.02436 | 24.1351 | 0.028203 | 0.15353 | 0.002885 | 0.20816 | 0.20816 | 0.030462
2019-03-28 15:51 | 1553802648.2799 | 747 | 0.20896 | 0.00082 | 5.7e-05 | 101.1554 | 100.9916 | 121.395 | 21.2612 | 0.002444 | 0.002297 | 150.3025 | 0.25779 | 8.4019 | 0.002435 | 0.004114 | 0.004005 | 4.8e-05 | 0.001038 | 0.32388 | 1.3147 | 2.5474 | 100.61 | 21.1 | 25.883 | 80.6057 | 80.389 | 77.0944 | 0.006445 | 1.008 | 0.024314 | 24.1216 | 0.014491 | 0.15336 | 0.00278 | 0.20007 | 0.20007 | 0.031658
2019-03-28 15:51 | 1553802649.2806 | 748 | 0.20898 | 0.000825 | 5.6e-05 | 101.1542 | 101.1105 | 121.5396 | 21.2612 | 0.002453 | 0.002273 | 150.3158 | 0.2571 | 8.4025 | 0.002472 | 0.003172 | 0.003057 | 4.7e-05 | 0.001036 | 0.32029 | 1.3703 | 2.5474 | 100.61 | 21.1 | 25.889 | 80.8434 | 81.291 | 76.4988 | 0.006446 | 1.0097 | 0.02435 | 24.1154 | 0.003861 | 0.15295 | 0.002648 | 0.19073 | 0.19073 | 0.033118
2019-03-28 15:51 | 1553802650.2786 | 749 | 0.20894 | 0.000821 | 5.6e-05 | 101.1185 | 101.0497 | 121.7625 | 21.2612 | 0.002448 | 0.002264 | 150.308 | 0.25777 | 8.4051 | 0.002459 | 0.003746 | 0.003691 | 4.9e-05 | 0.001032 | 0.30829 | 1.3292 | 2.5475 | 100.61 | 21.1 | 25.898 | 80.71 | 80.9744 | 76.2922 | 0.006449 | 1.0103 | 0.024379 | 24.1304 | 0.006081 | 0.15335 | 0.002852 | 0.20593 | 0.20593 | 0.030755
2019-03-28 15:51 | 1553802651.2789 | 750 | 0.20898 | 0.000819 | 5.6e-05 | 101.2203 | 101.153 | 121.9212 | 21.2612 | 0.00246 | 0.002267 | 150.2739 | 0.25629 | 8.405 | 0.002427 | 0.004396 | 0.004251 | 5e-05 | 0.001028 | 0.31581 | 1.3304 | 2.5475 | 100.61 | 21.1 | 25.902 | 80.9939 | 80.2135 | 76.3693 | 0.00645 | 1.0101 | 0.024392 | 24.1488 | 0.00595 | 0.15247 | 0.002654 | 0.19147 | 0.19147 | 0.032887
2019-03-28 15:51 | 1553802652.2797 | 751 | 0.20899 | 0.000822 | 5.5e-05 | 101.1986 | 101.1924 | 121.4623 | 21.2612 | 0.002435 | 0.00226 | 150.2983 | 0.25791 | 8.4003 | 0.002306 | 0.003773 | 0.003661 | 4.6e-05 | 0.001029 | 0.30934 | 1.3112 | 2.5475 | 100.61 | 21.1 | 25.907 | 80.388 | 77.3063 | 76.187 | 0.006451 | 1.0106 | 0.024353 | 24.097 | 0.000545 | 0.15343 | 0.002604 | 0.18754 | 0.18754 | 0.033788
2019-03-28 15:51 | 1553802653.2801 | 752 | 0.20896 | 0.000821 | 5.5e-05 | 101.1065 | 101.2532 | 121.2508 | 21.2612 | 0.00242 | 0.002272 | 150.3008 | 0.25707 | 8.4018 | 0.002374 | 0.003232 | 0.003157 | 4.6e-05 | 0.001029 | 0.33251 | 1.3655 | 2.5475 | 100.61 | 21.1 | 25.914 | 80.0326 | 78.9349 | 76.4815 | 0.006453 | 1.0098 | 0.024321 | 24.0861 | -0.012967 | 0.15293 | 0.002743 | 0.19756 | 0.19756 | 0.031969
2019-03-28 15:51 | 1553802654.2792 | 753 | 0.20898 | 0.000816 | 5.6e-05 | 101.1599 | 101.0046 | 121.6468 | 21.2612 | 0.002424 | 0.002284 | 150.2739 | 0.25757 | 8.4027 | 0.002434 | 0.003602 | 0.003491 | 4.7e-05 | 0.001027 | 0.32676 | 1.3178 | 2.5474 | 100.61 | 21.1 | 25.916 | 80.1417 | 80.3664 | 76.7662 | 0.006453 | 1.0089 | 0.024351 | 24.1353 | 0.013743 | 0.15323 | 0.002649 | 0.19083 | 0.19083 | 0.033162
2019-03-28 15:51 | 1553802655.2788 | 754 | 0.20898 | 0.000814 | 5.6e-05 | 101.1533 | 101.0406 | 121.5907 | 21.2612 | 0.002454 | 0.0023 | 150.2767 | 0.25748 | 8.4019 | 0.002386 | 0.00423 | 0.004149 | 5.1e-05 | 0.001024 | 0.31116 | 1.3234 | 2.5475 | 100.61 | 21.1 | 25.921 | 80.8497 | 79.2249 | 77.1441 | 0.006454 | 1.0079 | 0.024332 | 24.1427 | 0.009975 | 0.15318 | 0.002657 | 0.19126 | 0.19126 | 0.033077
2019-03-28 15:51 | 1553802656.3465 | 755 | 0.20899 | 0.000812 | 5.7e-05 | 101.142 | 101.1139 | 121.3903 | 21.2612 | 0.002466 | 0.002277 | 150.2472 | 0.25864 | 8.4029 | 0.002425 | 0.004005 | 0.003931 | 5.1e-05 | 0.001024 | 0.29695 | 1.2806 | 2.5473 | 100.61 | 21.1 | 25.921 | 81.1374 | 80.1667 | 76.6086 | 0.006454 | 1.0094 | 0.024331 | 24.1044 | 0.002481 | 0.15386 | 0.002606 | 0.18745 | 0.18745 | 0.0339
2019-03-28 15:51 | 1553802657.2913 | 756 | 0.20897 | 0.000816 | 5.8e-05 | 101.122 | 101.1143 | 121.3203 | 21.2612 | 0.002428 | 0.002318 | 150.254 | 0.25785 | 8.4025 | 0.00248 | 0.002982 | 0.002886 | 5.2e-05 | 0.001026 | 0.3178 | 1.2898 | 2.5472 | 100.61 | 21.1 | 25.922 | 80.2243 | 81.472 | 77.5956 | 0.006455 | 1.0066 | 0.02429 | 24.1314 | 0.000686 | 0.15339 | 0.002679 | 0.19241 | 0.19241 | 0.032926
2019-03-28 15:51 | 1553802658.2913 | 757 | 0.20897 | 0.000814 | 5.7e-05 | 101.1345 | 100.9081 | 121.5252 | 21.2612 | 0.002483 | 0.00232 | 150.2605 | 0.25779 | 8.4025 | 0.002422 | 0.00414 | 0.004 | 4.8e-05 | 0.001026 | 0.29856 | 1.3048 | 2.5474 | 100.61 | 21.1 | 25.923 | 81.5585 | 80.0934 | 77.6259 | 0.006455 | 1.0065 | 0.024309 | 24.1528 | 0.020053 | 0.15336 | 0.002703 | 0.19434 | 0.19434 | 0.032591
2019-03-28 15:51 | 1553802659.2927 | 758 | 0.20898 | 0.000812 | 5.8e-05 | 101.1977 | 101.1 | 121.2837 | 21.2612 | 0.002447 | 0.002309 | 150.2423 | 0.25797 | 8.4016 | 0.002505 | 0.004276 | 0.004182 | 4.9e-05 | 0.001023 | 0.31332 | 1.2765 | 2.5474 | 100.61 | 21.1 | 25.923 | 80.6899 | 82.0724 | 77.381 | 0.006455 | 1.0072 | 0.024293 | 24.1204 | 0.008637 | 0.15347 | 0.002637 | 0.18933 | 0.18933 | 0.033476
2019-03-28 15:51 | 1553802660.2913 | 759 | 0.20896 | 0.00081 | 5.8e-05 | 101.1539 | 100.7929 | 121.0006 | 21.2612 | 0.002443 | 0.002289 | 150.2194 | 0.25764 | 8.4 | 0.002427 | 0.003924 | 0.003799 | 5e-05 | 0.00102 | 0.31133 | 1.3053 | 2.5473 | 100.61 | 21.1 | 25.924 | 80.5932 | 80.2049 | 76.8931 | 0.006455 | 1.0086 | 0.024282 | 24.0754 | 0.031979 | 0.15327 | 0.002747 | 0.19728 | 0.19728 | 0.032086
2019-03-28 15:51 | 1553802661.2919 | 760 | 0.20896 | 0.000808 | 5.9e-05 | 101.1655 | 101.1269 | 121.2835 | 21.2612 | 0.00243 | 0.002235 | 150.2128 | 0.25788 | 8.4058 | 0.002428 | 0.004104 | 0.003943 | 5.5e-05 | 0.001023 | 0.33915 | 1.3236 | 2.5474 | 100.61 | 21.1 | 25.928 | 80.2851 | 80.2376 | 75.6032 | 0.006456 | 1.0123 | 0.024355 | 24.0591 | 0.00342 | 0.15341 | 0.002761 | 0.19883 | 0.19883 | 0.031865
2019-03-28 15:51 | 1553802662.2917 | 761 | 0.20894 | 0.000815 | 5.9e-05 | 101.1733 | 101.1656 | 121.8682 | 21.2612 | 0.002429 | 0.002271 | 150.2096 | 0.25753 | 8.4002 | 0.002475 | 0.003567 | 0.003533 | 5.2e-05 | 0.001026 | 0.32045 | 1.3302 | 2.5473 | 100.61 | 21.1 | 25.93 | 80.2587 | 81.3604 | 76.4628 | 0.006457 | 1.0098 | 0.024384 | 24.1467 | 0.000679 | 0.15321 | 0.002878 | 0.20768 | 0.20768 | 0.030467
2019-03-28 15:51 | 1553802663.2924 | 762 | 0.20896 | 0.000812 | 6e-05 | 101.12 | 100.9376 | 121.924 | 21.2612 | 0.002427 | 0.002247 | 150.1902 | 0.25778 | 8.4045 | 0.00246 | 0.003736 | 0.003629 | 5.3e-05 | 0.001027 | 0.30502 | 1.3562 | 2.5474 | 100.61 | 21.1 | 25.936 | 80.2135 | 80.997 | 75.883 | 0.006458 | 1.0115 | 0.02441 | 24.1322 | 0.016142 | 0.15335 | 0.002792 | 0.20154 | 0.20154 | 0.031425
2019-03-28 15:51 | 1553802664.2918 | 763 | 0.20895 | 0.000814 | 6.1e-05 | 101.1905 | 100.9597 | 121.7442 | 21.2612 | 0.002417 | 0.002299 | 150.2109 | 0.25693 | 8.4021 | 0.002447 | 0.003584 | 0.003441 | 5.4e-05 | 0.001032 | 0.29192 | 1.3335 | 2.5475 | 100.61 | 21.09 | 25.94 | 79.9756 | 80.6812 | 77.1302 | 0.006455 | 1.0079 | 0.024348 | 24.1575 | 0.020427 | 0.15285 | 0.002807 | 0.20205 | 0.20205 | 0.031244
2019-03-28 15:51 | 1553802665.2905 | 764 | 0.20894 | 0.00082 | 6.1e-05 | 101.1253 | 100.8853 | 121.5893 | 21.2612 | 0.002464 | 0.002289 | 150.1849 | 0.25779 | 8.4013 | 0.0025 | 0.003454 | 0.003339 | 5.4e-05 | 0.00103 | 0.30397 | 1.3232 | 2.5474 | 100.61 | 21.09 | 25.945 | 81.0922 | 81.9671 | 76.8947 | 0.006456 | 1.0086 | 0.024341 | 24.134 | 0.021254 | 0.15336 | 0.002893 | 0.20827 | 0.20827 | 0.030411
2019-03-28 15:51 | 1553802666.2908 | 765 | 0.20895 | 0.000825 | 6.1e-05 | 101.1703 | 101.2869 | 121.6403 | 21.2612 | 0.002474 | 0.002311 | 150.2032 | 0.25821 | 8.4038 | 0.002477 | 0.003543 | 0.003443 | 5.3e-05 | 0.001026 | 0.31492 | 1.3661 | 2.5474 | 100.61 | 21.1 | 25.957 | 81.3489 | 81.4198 | 77.4317 | 0.006463 | 1.007 | 0.024327 | 24.1575 | -0.010301 | 0.15361 | 0.002811 | 0.20217 | 0.20217 | 0.031381
2019-03-28 15:51 | 1553802667.2919 | 766 | 0.20894 | 0.000825 | 6.2e-05 | 101.1667 | 101.1805 | 121.6928 | 21.2612 | 0.002421 | 0.00227 | 150.1838 | 0.25855 | 8.4022 | 0.00245 | 0.003156 | 0.003023 | 5.4e-05 | 0.001022 | 0.31996 | 1.3275 | 2.5474 | 100.61 | 21.1 | 25.964 | 80.059 | 80.7646 | 76.4426 | 0.006465 | 1.0099 | 0.024367 | 24.1287 | -0.001221 | 0.15381 | 0.002834 | 0.20412 | 0.20412 | 0.031121
2019-03-28 15:51 | 1553802668.2915 | 767 | 0.20894 | 0.000824 | 6.2e-05 | 101.1491 | 101.0756 | 121.8557 | 21.2612 | 0.002406 | 0.002265 | 150.1802 | 0.25859 | 8.401 | 0.002462 | 0.004218 | 0.004123 | 5.6e-05 | 0.001018 | 0.32068 | 1.3265 | 2.5473 | 100.61 | 21.1 | 25.981 | 79.7144 | 81.0564 | 76.3186 | 0.006469 | 1.0102 | 0.024388 | 24.1405 | 0.006509 | 0.15383 | 0.002875 | 0.20725 | 0.20725 | 0.030655
2019-03-28 15:51 | 1553802669.2919 | 768 | 0.20895 | 0.000828 | 6.3e-05 | 101.1798 | 101.0936 | 122.082 | 21.2612 | 0.002413 | 0.002291 | 150.1816 | 0.25674 | 8.4005 | 0.002423 | 0.00401 | 0.00386 | 5.7e-05 | 0.001013 | 0.30845 | 1.2984 | 2.5473 | 100.61 | 21.1 | 25.991 | 79.8634 | 80.1097 | 76.9361 | 0.006472 | 1.0085 | 0.024389 | 24.1843 | 0.007629 | 0.15273 | 0.002774 | 0.19979 | 0.19979 | 0.031572
2019-03-28 15:51 | 1553802670.2906 | 769 | 0.20893 | 0.00083 | 6.4e-05 | 101.1313 | 101.0976 | 122.2179 | 21.2612 | 0.002463 | 0.00227 | 150.1861 | 0.25855 | 8.4036 | 0.002527 | 0.003601 | 0.003538 | 5.3e-05 | 0.00101 | 0.31277 | 1.2783 | 2.5473 | 100.61 | 21.1 | 26 | 81.0696 | 82.6155 | 76.4263 | 0.006474 | 1.0099 | 0.02442 | 24.1801 | 0.002975 | 0.15381 | 0.00293 | 0.21153 | 0.21153 | 0.030031
2019-03-28 15:51 | 1553802671.2916 | 770 | 0.20896 | 0.000832 | 6.4e-05 | 101.164 | 101.0249 | 122.1968 | 21.2612 | 0.00248 | 0.002272 | 150.1838 | 0.25728 | 8.4027 | 0.00249 | 0.003839 | 0.003812 | 5.3e-05 | 0.001015 | 0.31437 | 1.2979 | 2.5475 | 100.61 | 21.1 | 26.015 | 81.4806 | 81.72 | 76.4762 | 0.006478 | 1.0098 | 0.024416 | 24.1797 | 0.012307 | 0.15305 | 0.002725 | 0.19644 | 0.19644 | 0.032179
2019-03-28 15:51 | 1553802672.292 | 771 | 0.20897 | 0.000827 | 6.5e-05 | 101.2078 | 101.0599 | 122.2238 | 21.2612 | 0.002447 | 0.002243 | 150.1529 | 0.25761 | 8.4029 | 0.002483 | 0.003442 | 0.003295 | 5.7e-05 | 0.001017 | 0.33019 | 1.3001 | 2.5474 | 100.61 | 21.1 | 26.023 | 80.7 | 81.5469 | 75.7903 | 0.00648 | 1.0118 | 0.024443 | 24.1587 | 0.01308 | 0.15325 | 0.00271 | 0.19545 | 0.19545 | 0.032382
2019-03-28 15:51 | 1553802673.2917 | 772 | 0.20895 | 0.000823 | 6.5e-05 | 101.1298 | 101.1275 | 122.2618 | 21.2612 | 0.002514 | 0.002257 | 150.1333 | 0.2582 | 8.4021 | 0.002436 | 0.004128 | 0.003967 | 5.6e-05 | 0.001017 | 0.31061 | 1.3906 | 2.5476 | 100.61 | 21.1 | 26.031 | 82.2916 | 80.4154 | 76.1293 | 0.006482 | 1.0108 | 0.024435 | 24.1742 | 0.000199 | 0.1536 | 0.002803 | 0.20222 | 0.20222 | 0.031371
2019-03-28 15:51 | 1553802674.2908 | 773 | 0.20901 | 0.000825 | 6.5e-05 | 101.1911 | 101.0236 | 122.1574 | 21.2612 | 0.002424 | 0.002284 | 150.1101 | 0.25775 | 8.4014 | 0.002499 | 0.004167 | 0.004067 | 5.4e-05 | 0.001023 | 0.32045 | 1.3399 | 2.5474 | 100.61 | 21.1 | 26.043 | 80.1441 | 81.9383 | 76.7684 | 0.006485 | 1.0089 | 0.024402 | 24.1859 | 0.014823 | 0.15334 | 0.002435 | 0.17491 | 0.17491 | 0.036205
2019-03-28 15:51 | 1553802675.2907 | 774 | 0.209 | 0.000821 | 6.5e-05 | 101.1405 | 101.1877 | 122.2348 | 21.2612 | 0.002435 | 0.002253 | 150.1198 | 0.25792 | 8.4024 | 0.002466 | 0.004515 | 0.004439 | 5.7e-05 | 0.001024 | 0.29966 | 1.3198 | 2.5474 | 100.61 | 21.1 | 26.048 | 80.388 | 81.1499 | 76.0154 | 0.006486 | 1.0111 | 0.024436 | 24.1675 | -0.00417 | 0.15344 | 0.00252 | 0.18144 | 0.18144 | 0.034926
2019-03-28 15:51 | 1553802676.2912 | 775 | 0.20899 | 0.000816 | 6.5e-05 | 101.1491 | 101.1826 | 122.4343 | 21.2612 | 0.002379 | 0.002241 | 150.1099 | 0.25807 | 8.4019 | 0.002497 | 0.003794 | 0.003756 | 5.8e-05 | 0.001022 | 0.30862 | 1.3474 | 2.5473 | 100.61 | 21.1 | 26.054 | 79.0658 | 81.8823 | 75.7403 | 0.006488 | 1.0119 | 0.024466 | 24.1777 | -0.002955 | 0.15353 | 0.002553 | 0.18404 | 0.18404 | 0.034452
2019-03-28 15:51 | 1553802677.2908 | 776 | 0.20903 | 0.000813 | 6.5e-05 | 101.1905 | 101.0025 | 122.1165 | 21.2612 | 0.002425 | 0.002343 | 150.1271 | 0.25734 | 8.4047 | 0.002504 | 0.004211 | 0.0041 | 5.6e-05 | 0.001021 | 0.31797 | 1.3542 | 2.5473 | 100.61 | 21.1 | 26.056 | 80.159 | 82.0633 | 78.1802 | 0.006488 | 1.0049 | 0.024349 | 24.2306 | 0.016636 | 0.15309 | 0.002325 | 0.1665 | 0.1665 | 0.037976
2019-03-28 15:51 | 1553802678.2918 | 777 | 0.20903 | 0.000815 | 6.5e-05 | 101.1777 | 101.1574 | 121.8935 | 21.2612 | 0.002446 | 0.002292 | 150.1255 | 0.2578 | 8.4016 | 0.00249 | 0.00386 | 0.003748 | 5.5e-05 | 0.00102 | 0.33843 | 1.343 | 2.5475 | 100.61 | 21.1 | 26.059 | 80.6603 | 81.7308 | 76.9517 | 0.006489 | 1.0084 | 0.024369 | 24.1661 | 0.001795 | 0.15336 | 0.002306 | 0.16517 | 0.16517 | 0.038347
2019-03-28 15:51 | 1553802679.2919 | 778 | 0.20906 | 0.000812 | 6.6e-05 | 101.1926 | 100.8877 | 122.0853 | 21.2612 | 0.002511 | 0.002285 | 150.1082 | 0.25778 | 8.4029 | 0.002565 | 0.003631 | 0.003549 | 5.7e-05 | 0.001026 | 0.31653 | 1.3445 | 2.5475 | 100.61 | 21.1 | 26.06 | 82.2222 | 83.5146 | 76.8003 | 0.006489 | 1.0088 | 0.024394 | 24.1799 | 0.026995 | 0.15336 | 0.002166 | 0.155 | 0.155 | 0.040862
2019-03-28 15:51 | 1553802680.2913 | 779 | 0.20905 | 0.000802 | 6.8e-05 | 101.1468 | 101.1102 | 122.1078 | 21.2612 | 0.002428 | 0.002289 | 150.1066 | 0.25795 | 8.401 | 0.002483 | 0.004137 | 0.0041 | 5.9e-05 | 0.001027 | 0.30845 | 1.3531 | 2.5474 | 100.61 | 21.1 | 26.062 | 80.2236 | 81.5609 | 76.8799 | 0.00649 | 1.0086 | 0.024393 | 24.1849 | 0.003233 | 0.15345 | 0.002229 | 0.15955 | 0.15955 | 0.039721
2019-03-28 15:51 | 1553802681.2912 | 780 | 0.20903 | 0.000804 | 6.9e-05 | 101.0863 | 101.0698 | 122.048 | 21.2612 | 0.002383 | 0.002323 | 150.1049 | 0.2584 | 8.4071 | 0.002626 | 0.003882 | 0.003737 | 5.6e-05 | 0.001032 | 0.31188 | 1.3107 | 2.5475 | 100.61 | 21.09 | 26.066 | 79.16 | 84.991 | 77.7102 | 0.006487 | 1.0062 | 0.024358 | 24.2076 | 0.001461 | 0.15372 | 0.002358 | 0.16872 | 0.16872 | 0.03763
2019-03-28 15:51 | 1553802682.29 | 781 | 0.209 | 0.000798 | 7e-05 | 101.156 | 100.9621 | 121.8635 | 21.2612 | 0.002494 | 0.002323 | 150.1068 | 0.2586 | 8.4032 | 0.00265 | 0.004449 | 0.004335 | 5.4e-05 | 0.001032 | 0.31006 | 1.3022 | 2.5474 | 100.61 | 21.1 | 26.071 | 81.8159 | 85.574 | 77.6999 | 0.006492 | 1.0063 | 0.02434 | 24.1889 | 0.017162 | 0.15384 | 0.002512 | 0.17984 | 0.17984 | 0.035329
2019-03-28 15:51 | 1553802683.291 | 782 | 0.20904 | 0.000797 | 7e-05 | 101.1054 | 100.9594 | 121.7201 | 21.2612 | 0.002442 | 0.002308 | 150.1069 | 0.25765 | 8.4029 | 0.002601 | 0.00354 | 0.003513 | 5.6e-05 | 0.001039 | 0.3011 | 1.3358 | 2.5474 | 100.61 | 21.1 | 26.075 | 80.5721 | 84.3996 | 77.3438 | 0.006493 | 1.0073 | 0.024338 | 24.1624 | 0.012924 | 0.15328 | 0.002304 | 0.16451 | 0.16451 | 0.03848
2019-03-28 15:51 | 1553802684.2904 | 783 | 0.20901 | 0.000795 | 7.1e-05 | 101.2245 | 101.3001 | 121.7494 | 21.2612 | 0.002423 | 0.002285 | 150.1028 | 0.25789 | 8.4023 | 0.002579 | 0.004124 | 0.003935 | 5.3e-05 | 0.001041 | 0.3449 | 1.3115 | 2.5474 | 100.61 | 21.1 | 26.08 | 80.1035 | 83.8628 | 76.7934 | 0.006494 | 1.0089 | 0.02436 | 24.1464 | -0.006679 | 0.15342 | 0.002485 | 0.17792 | 0.17792 | 0.035613
2019-03-28 15:51 | 1553802685.291 | 784 | 0.20901 | 0.00079 | 7.2e-05 | 101.2179 | 101.2794 | 121.6766 | 21.2612 | 0.002419 | 0.002289 | 150.1012 | 0.25797 | 8.4017 | 0.002527 | 0.003919 | 0.003819 | 5.5e-05 | 0.001044 | 0.33157 | 1.3264 | 2.5474 | 100.61 | 21.1 | 26.081 | 80.0239 | 82.6005 | 76.8885 | 0.006494 | 1.0086 | 0.02435 | 24.1424 | -0.005428 | 0.15347 | 0.002469 | 0.17662 | 0.17662 | 0.035887
2019-03-28 15:51 | 1553802686.2901 | 785 | 0.20902 | 0.000787 | 7.2e-05 | 101.1911 | 101.3123 | 121.7015 | 21.2612 | 0.002441 | 0.002323 | 150.0782 | 0.25793 | 8.4028 | 0.002533 | 0.004308 | 0.00427 | 5.6e-05 | 0.001045 | 0.3366 | 1.3286 | 2.5473 | 100.61 | 21.1 | 26.082 | 80.5464 | 82.7472 | 77.7047 | 0.006495 | 1.0062 | 0.024324 | 24.173 | -0.010707 | 0.15344 | 0.002411 | 0.17213 | 0.17213 | 0.036816
2019-03-28 15:51 | 1553802687.2905 | 786 | 0.209 | 0.000786 | 7.1e-05 | 101.2054 | 101.2743 | 121.9564 | 21.2612 | 0.002453 | 0.002278 | 150.0512 | 0.25761 | 8.4058 | 0.002504 | 0.004296 | 0.004146 | 5.7e-05 | 0.001049 | 0.33395 | 1.3237 | 2.5474 | 100.61 | 21.1 | 26.084 | 80.841 | 82.0529 | 76.6297 | 0.006495 | 1.0093 | 0.024387 | 24.1612 | -0.006084 | 0.15325 | 0.002527 | 0.18123 | 0.18123 | 0.034924
2019-03-28 15:51 | 1553802688.2902 | 787 | 0.20902 | 0.000784 | 6.9e-05 | 101.1545 | 101.117 | 122.1789 | 21.2612 | 0.002462 | 0.002291 | 150.0325 | 0.25735 | 8.4046 | 0.002553 | 0.004405 | 0.004309 | 5.3e-05 | 0.001051 | 0.32947 | 1.3444 | 2.5475 | 100.61 | 21.1 | 26.086 | 81.0578 | 83.2473 | 76.9416 | 0.006496 | 1.0084 | 0.024398 | 24.194 | 0.003317 | 0.1531 | 0.002413 | 0.17303 | 0.17303 | 0.036543
2019-03-28 15:51 | 1553802689.2912 | 788 | 0.20903 | 0.000783 | 6.9e-05 | 101.156 | 100.9903 | 121.878 | 21.2612 | 0.002389 | 0.002304 | 150.0303 | 0.25753 | 8.4029 | 0.002561 | 0.003608 | 0.003499 | 5.4e-05 | 0.001048 | 0.30359 | 1.3195 | 2.5473 | 100.61 | 21.1 | 26.089 | 79.284 | 83.4197 | 77.2611 | 0.006496 | 1.0075 | 0.024357 | 24.1753 | 0.014664 | 0.15321 | 0.002347 | 0.16791 | 0.16791 | 0.037682
2019-03-28 15:51 | 1553802690.2902 | 789 | 0.20902 | 0.000784 | 6.8e-05 | 101.1387 | 100.886 | 121.9409 | 21.2612 | 0.002481 | 0.002264 | 150.0118 | 0.25789 | 8.4046 | 0.002536 | 0.004041 | 0.00393 | 5.3e-05 | 0.001049 | 0.32676 | 1.3034 | 2.5475 | 100.61 | 21.1 | 26.092 | 81.5133 | 82.822 | 76.2898 | 0.006497 | 1.0103 | 0.024397 | 24.148 | 0.022379 | 0.15342 | 0.002425 | 0.17399 | 0.17399 | 0.036418
2019-03-28 15:52 | 1553802691.2903 | 790 | 0.20902 | 0.00078 | 6.8e-05 | 101.1634 | 101.1153 | 121.9909 | 21.2612 | 0.002395 | 0.002269 | 150.0086 | 0.25769 | 8.4038 | 0.002558 | 0.004494 | 0.004433 | 5.4e-05 | 0.001051 | 0.31382 | 1.3304 | 2.5474 | 100.61 | 21.1 | 26.093 | 79.4292 | 83.3533 | 76.4145 | 0.006497 | 1.01 | 0.024398 | 24.1572 | 0.004257 | 0.1533 | 0.00244 | 0.17503 | 0.17503 | 0.036173
2019-03-28 15:52 | 1553802692.2912 | 791 | 0.20899 | 0.00078 | 6.9e-05 | 101.1536 | 100.8367 | 121.9848 | 21.2612 | 0.002435 | 0.002297 | 150.014 | 0.25778 | 8.4037 | 0.002537 | 0.004096 | 0.00404 | 5.2e-05 | 0.001055 | 0.31797 | 1.3429 | 2.5473 | 100.61 | 21.1 | 26.092 | 80.4092 | 82.8407 | 77.0858 | 0.006497 | 1.008 | 0.024374 | 24.1798 | 0.028064 | 0.15335 | 0.002585 | 0.18544 | 0.18544 | 0.034154
2019-03-28 15:52 | 1553802693.2917 | 792 | 0.20899 | 0.000783 | 6.9e-05 | 101.1795 | 101.1962 | 122.1461 | 21.2612 | 0.002427 | 0.002283 | 149.9989 | 0.25754 | 8.4018 | 0.002483 | 0.004243 | 0.004124 | 5.3e-05 | 0.001055 | 0.33395 | 1.3269 | 2.5474 | 100.61 | 21.1 | 26.085 | 80.201 | 81.5602 | 76.7427 | 0.006495 | 1.009 | 0.024402 | 24.1839 | -0.001471 | 0.15321 | 0.002629 | 0.18889 | 0.18889 | 0.033499
2019-03-28 15:52 | 1553802694.2914 | 793 | 0.20897 | 0.000784 | 6.8e-05 | 101.1411 | 100.9855 | 122.3771 | 21.2612 | 0.002437 | 0.002272 | 149.9966 | 0.25788 | 8.404 | 0.002506 | 0.003939 | 0.003879 | 4.9e-05 | 0.001053 | 0.31564 | 1.3166 | 2.5474 | 100.61 | 21.1 | 26.079 | 80.438 | 82.0991 | 76.4714 | 0.006494 | 1.0098 | 0.024434 | 24.1974 | 0.013769 | 0.15341 | 0.002709 | 0.19511 | 0.19511 | 0.032473
2019-03-28 15:52 | 1553802695.2904 | 794 | 0.20898 | 0.000781 | 6.8e-05 | 101.2147 | 101.027 | 122.2857 | 21.2612 | 0.002477 | 0.002269 | 149.9737 | 0.25737 | 8.4046 | 0.002481 | 0.003357 | 0.003256 | 4.9e-05 | 0.001052 | 0.34827 | 1.3574 | 2.5475 | 100.61 | 21.1 | 26.062 | 81.3994 | 81.4955 | 76.4217 | 0.00649 | 1.0099 | 0.024427 | 24.1866 | 0.016602 | 0.15311 | 0.002662 | 0.1916 | 0.1916 | 0.033004
2019-03-28 15:52 | 1553802696.2895 | 795 | 0.20898 | 0.00078 | 6.7e-05 | 101.1798 | 101.1989 | 122.0879 | 21.2612 | 0.002391 | 0.00226 | 149.9878 | 0.25816 | 8.4023 | 0.002571 | 0.004024 | 0.003928 | 5e-05 | 0.001049 | 0.31006 | 1.3696 | 2.5474 | 100.61 | 21.1 | 26.051 | 79.3527 | 83.6668 | 76.1971 | 0.006487 | 1.0106 | 0.024415 | 24.1593 | -0.001684 | 0.15358 | 0.002656 | 0.19116 | 0.19116 | 0.033182
2019-03-28 15:52 | 1553802697.2895 | 796 | 0.20898 | 0.000783 | 6.8e-05 | 101.1989 | 101.2182 | 122.1273 | 21.2612 | 0.002426 | 0.002269 | 149.991 | 0.25823 | 8.4052 | 0.002548 | 0.004187 | 0.004122 | 5.2e-05 | 0.001044 | 0.33732 | 1.357 | 2.5474 | 100.61 | 21.1 | 26.041 | 80.1722 | 83.1184 | 76.4068 | 0.006484 | 1.01 | 0.024412 | 24.1705 | -0.001711 | 0.15362 | 0.002675 | 0.19245 | 0.19245 | 0.032968
2019-03-28 15:52 | 1553802698.2905 | 797 | 0.20898 | 0.00078 | 6.8e-05 | 101.0428 | 101.1785 | 121.7731 | 21.2612 | 0.002428 | 0.002274 | 149.9987 | 0.25763 | 8.401 | 0.002553 | 0.003759 | 0.003589 | 5e-05 | 0.00104 | 0.32211 | 1.345 | 2.5474 | 100.61 | 21.1 | 26.017 | 80.2236 | 83.2386 | 76.5284 | 0.006478 | 1.0096 | 0.024372 | 24.1396 | -0.012003 | 0.15327 | 0.002672 | 0.19195 | 0.19195 | 0.032977
2019-03-28 15:52 | 1553802699.2895 | 798 | 0.209 | 0.000773 | 6.9e-05 | 101.1754 | 101.2077 | 121.8585 | 21.2612 | 0.002437 | 0.002289 | 149.9887 | 0.25778 | 8.4031 | 0.002552 | 0.00384 | 0.003746 | 4.8e-05 | 0.00104 | 0.32604 | 1.3331 | 2.5475 | 100.61 | 21.1 | 26.002 | 80.445 | 83.2018 | 76.8962 | 0.006475 | 1.0086 | 0.024368 | 24.1607 | -0.00286 | 0.15336 | 0.002538 | 0.18199 | 0.18199 | 0.034801
2019-03-28 15:52 | 1553802700.2896 | 799 | 0.20899 | 0.000777 | 6.9e-05 | 101.2054 | 101.1836 | 121.9865 | 21.2612 | 0.00243 | 0.002303 | 149.9888 | 0.25806 | 8.4052 | 0.002573 | 0.003604 | 0.003497 | 5.2e-05 | 0.00104 | 0.30901 | 1.2787 | 2.5474 | 100.61 | 21.1 | 25.986 | 80.2806 | 83.7075 | 77.2275 | 0.006471 | 1.0076 | 0.024369 | 24.1849 | 0.001931 | 0.15352 | 0.002639 | 0.1894 | 0.1894 | 0.033476
2019-03-28 15:52 | 1553802701.2896 | 800 | 0.20902 | 0.000778 | 7.1e-05 | 101.1319 | 101.1319 | 121.9625 | 21.2612 | 0.002434 | 0.002286 | 149.9626 | 0.25801 | 8.4005 | 0.002551 | 0.003853 | 0.003745 | 4.8e-05 | 0.001041 | 0.31868 | 1.2961 | 2.5473 | 100.61 | 21.1 | 25.95 | 80.3866 | 83.1816 | 76.8176 | 0.006462 | 1.0088 | 0.024381 | 24.1683 | -7e-06 | 0.15349 | 0.002445 | 0.17517 | 0.17517 | 0.036188
2019-03-28 15:52 | 1553802702.2907 | 801 | 0.20903 | 0.00077 | 7.2e-05 | 101.153 | 101.1245 | 121.9499 | 21.2612 | 0.00249 | 0.002229 | 149.962 | 0.25724 | 8.4015 | 0.002507 | 0.003437 | 0.003316 | 4.8e-05 | 0.001039 | 0.30182 | 1.281 | 2.5474 | 100.61 | 21.1 | 25.931 | 81.7301 | 82.1412 | 75.4511 | 0.006457 | 1.0127 | 0.024427 | 24.1198 | 0.002524 | 0.15303 | 0.002352 | 0.16858 | 0.16858 | 0.037492
2019-03-28 15:52 | 1553802703.2896 | 802 | 0.20905 | 0.000772 | 7.3e-05 | 101.1724 | 101.1092 | 121.8968 | 21.2612 | 0.002455 | 0.002298 | 149.9846 | 0.25683 | 8.4045 | 0.002479 | 0.003656 | 0.003648 | 4.7e-05 | 0.001035 | 0.3314 | 1.3174 | 2.5474 | 100.61 | 21.1 | 25.895 | 80.873 | 81.4674 | 77.0958 | 0.006448 | 1.008 | 0.024365 | 24.1714 | 0.005588 | 0.15279 | 0.002244 | 0.16015 | 0.16015 | 0.039402
2019-03-28 15:52 | 1553802704.2907 | 803 | 0.20903 | 0.000769 | 7.4e-05 | 101.1899 | 101.0083 | 121.9184 | 21.2612 | 0.002396 | 0.002312 | 149.9721 | 0.25728 | 8.4037 | 0.002516 | 0.003603 | 0.003594 | 5.1e-05 | 0.00103 | 0.30381 | 1.3345 | 2.5474 | 100.61 | 21.1 | 25.878 | 79.4594 | 82.34 | 77.4404 | 0.006444 | 1.007 | 0.024355 | 24.1854 | 0.016072 | 0.15306 | 0.002355 | 0.16819 | 0.16819 | 0.037585
2019-03-28 15:52 | 1553802705.2901 | 804 | 0.20901 | 0.000766 | 7.3e-05 | 101.1462 | 101.2125 | 122.0973 | 21.2612 | 0.002472 | 0.002278 | 149.9672 | 0.25814 | 8.4007 | 0.0025 | 0.003956 | 0.003733 | 5e-05 | 0.00103 | 0.32803 | 1.3208 | 2.5474 | 100.61 | 21.1 | 25.863 | 81.2958 | 81.9549 | 76.6235 | 0.00644 | 1.0094 | 0.024401 | 24.175 | -0.005861 | 0.15356 | 0.002513 | 0.18013 | 0.18013 | 0.03521
2019-03-28 15:52 | 1553802706.2898 | 805 | 0.20902 | 0.000768 | 7.2e-05 | 101.164 | 101.1211 | 121.8489 | 21.2612 | 0.002415 | 0.002295 | 149.9512 | 0.25734 | 8.4044 | 0.002538 | 0.003623 | 0.003542 | 5e-05 | 0.001029 | 0.33179 | 1.3397 | 2.5474 | 100.61 | 21.1 | 25.84 | 79.9256 | 82.8797 | 77.0295 | 0.006434 | 1.0082 | 0.024362 | 24.1644 | 0.003799 | 0.15309 | 0.002435 | 0.17422 | 0.17422 | 0.036291
2019-03-28 15:52 | 1553802707.2898 | 806 | 0.20901 | 0.000758 | 7e-05 | 101.1605 | 101.0871 | 121.5872 | 21.2612 | 0.002422 | 0.002276 | 149.9383 | 0.25747 | 8.402 | 0.002485 | 0.004413 | 0.004292 | 4.9e-05 | 0.001034 | 0.33588 | 1.357 | 2.5474 | 100.61 | 21.1 | 25.831 | 80.0965 | 81.5921 | 76.5774 | 0.006432 | 1.0095 | 0.024352 | 24.1228 | 0.006488 | 0.15317 | 0.00254 | 0.18191 | 0.18191 | 0.034775
2019-03-28 15:52 | 1553802708.291 | 807 | 0.20903 | 0.000762 | 6.9e-05 | 101.1992 | 101.0902 | 121.4161 | 21.2612 | 0.00241 | 0.00227 | 149.9219 | 0.25723 | 8.4036 | 0.002548 | 0.003814 | 0.003809 | 5.1e-05 | 0.001037 | 0.32621 | 1.3044 | 2.5474 | 100.61 | 21.1 | 25.822 | 79.8002 | 83.1138 | 76.4238 | 0.00643 | 1.0099 | 0.02434 | 24.1006 | 0.009641 | 0.15303 | 0.002422 | 0.17325 | 0.17325 | 0.036479
2019-03-28 15:52 | 1553802709.2899 | 808 | 0.20902 | 0.000759 | 6.9e-05 | 101.0905 | 101.0348 | 121.537 | 21.2612 | 0.002482 | 0.002259 | 149.9065 | 0.25721 | 8.4025 | 0.002436 | 0.004429 | 0.004235 | 5e-05 | 0.001034 | 0.31924 | 1.3021 | 2.5474 | 100.61 | 21.1 | 25.821 | 81.5234 | 80.4161 | 76.1689 | 0.00643 | 1.0107 | 0.024361 | 24.1038 | 0.004928 | 0.15302 | 0.002479 | 0.17761 | 0.17761 | 0.035581
2019-03-28 15:52 | 1553802710.2899 | 809 | 0.20903 | 0.000758 | 6.9e-05 | 101.1125 | 101.0776 | 121.3924 | 21.2612 | 0.00242 | 0.00231 | 149.9161 | 0.25704 | 8.4044 | 0.002455 | 0.004354 | 0.004296 | 4.9e-05 | 0.001034 | 0.31094 | 1.2664 | 2.5473 | 100.61 | 21.1 | 25.823 | 80.0319 | 80.8855 | 77.389 | 0.00643 | 1.0071 | 0.024304 | 24.1314 | 0.003089 | 0.15291 | 0.002424 | 0.17322 | 0.17322 | 0.036459
2019-03-28 15:52 | 1553802711.2913 | 810 | 0.20903 | 0.000764 | 7e-05 | 101.1104 | 101.0939 | 121.3473 | 21.2612 | 0.00243 | 0.002244 | 149.9276 | 0.25781 | 8.4016 | 0.002521 | 0.003493 | 0.003469 | 5e-05 | 0.001038 | 0.32173 | 1.2452 | 2.5473 | 100.61 | 21.1 | 25.835 | 80.2758 | 82.4741 | 75.8057 | 0.006433 | 1.0117 | 0.024355 | 24.0724 | 0.001462 | 0.15337 | 0.002362 | 0.16892 | 0.16892 | 0.037499
2019-03-28 15:52 | 1553802712.289 | 811 | 0.20904 | 0.000762 | 7.1e-05 | 101.0699 | 101.1065 | 121.1631 | 21.2612 | 0.00241 | 0.002297 | 149.9342 | 0.25754 | 8.4051 | 0.002555 | 0.003995 | 0.003768 | 4.9e-05 | 0.001039 | 0.31116 | 1.2704 | 2.5473 | 100.61 | 21.1 | 25.844 | 79.7987 | 83.28 | 77.0802 | 0.006435 | 1.008 | 0.024292 | 24.098 | -0.003235 | 0.15321 | 0.002355 | 0.16795 | 0.16795 | 0.037675
2019-03-28 15:52 | 1553802713.2896 | 812 | 0.20902 | 0.000766 | 7.1e-05 | 101.1143 | 101.0474 | 121.126 | 21.2612 | 0.002418 | 0.002259 | 149.9058 | 0.25864 | 8.4043 | 0.002609 | 0.003419 | 0.003375 | 5.1e-05 | 0.001038 | 0.3298 | 1.2984 | 2.5473 | 100.61 | 21.1 | 25.855 | 79.9843 | 84.5915 | 76.1644 | 0.006438 | 1.0107 | 0.02432 | 24.0628 | 0.005923 | 0.15386 | 0.002473 | 0.17671 | 0.17671 | 0.035961
2019-03-28 15:52 | 1553802714.2904 | 813 | 0.20904 | 0.000761 | 7e-05 | 101.1402 | 101.133 | 120.965 | 21.2612 | 0.002437 | 0.002245 | 149.8903 | 0.25692 | 8.404 | 0.002557 | 0.003085 | 0.003042 | 4.7e-05 | 0.001036 | 0.32173 | 1.3131 | 2.5475 | 100.61 | 21.1 | 25.88 | 80.4457 | 83.3276 | 75.8331 | 0.006444 | 1.0116 | 0.024315 | 24.0354 | 0.00064 | 0.15284 | 0.002354 | 0.16804 | 0.16804 | 0.037563
2019-03-28 15:52 | 1553802715.29 | 814 | 0.20906 | 0.000765 | 6.9e-05 | 101.0985 | 101.0222 | 120.8762 | 21.2612 | 0.002431 | 0.002287 | 149.8908 | 0.258 | 8.4027 | 0.002513 | 0.003627 | 0.003422 | 4.9e-05 | 0.001037 | 0.30453 | 1.3358 | 2.5474 | 100.61 | 21.1 | 25.892 | 80.3094 | 82.2645 | 76.8417 | 0.006447 | 1.0087 | 0.024271 | 24.0613 | 0.006752 | 0.15349 | 0.002234 | 0.15903 | 0.15903 | 0.039861
2019-03-28 15:52 | 1553802716.289 | 815 | 0.20905 | 0.000765 | 6.8e-05 | 101.092 | 100.9247 | 120.9159 | 21.2612 | 0.002429 | 0.002254 | 149.8869 | 0.25748 | 8.4019 | 0.002471 | 0.003418 | 0.003313 | 5.2e-05 | 0.001041 | 0.32709 | 1.3046 | 2.5475 | 100.61 | 21.1 | 25.916 | 80.2493 | 81.2676 | 76.0522 | 0.006453 | 1.011 | 0.024303 | 24.0381 | 0.014809 | 0.15318 | 0.002291 | 0.16353 | 0.16353 | 0.038686
2019-03-28 15:52 | 1553802717.2891 | 816 | 0.20906 | 0.000763 | 6.8e-05 | 101.027 | 100.8704 | 121.0457 | 21.2612 | 0.002466 | 0.002268 | 149.8913 | 0.2583 | 8.4041 | 0.002567 | 0.004246 | 0.004163 | 5.2e-05 | 0.001039 | 0.3246 | 1.3143 | 2.5474 | 100.61 | 21.09 | 25.926 | 81.1343 | 83.5647 | 76.3849 | 0.006452 | 1.01 | 0.024304 | 24.0624 | 0.013881 | 0.15366 | 0.002245 | 0.16014 | 0.16014 | 0.039629
2019-03-28 15:52 | 1553802718.2897 | 817 | 0.20908 | 0.000767 | 6.8e-05 | 101.0237 | 100.9587 | 121.3652 | 21.2612 | 0.002445 | 0.0023 | 149.8783 | 0.258 | 8.4042 | 0.002534 | 0.003281 | 0.003201 | 5.3e-05 | 0.001041 | 0.3235 | 1.3175 | 2.5473 | 100.61 | 21.09 | 25.936 | 80.6399 | 82.7698 | 77.1583 | 0.006454 | 1.0078 | 0.024309 | 24.1208 | 0.005762 | 0.15349 | 0.002116 | 0.15078 | 0.15078 | 0.042041
2019-03-28 15:52 | 1553802719.2901 | 818 | 0.20909 | 0.000761 | 7e-05 | 101.0663 | 100.9458 | 121.1112 | 21.2612 | 0.002426 | 0.002287 | 149.8966 | 0.25784 | 8.4038 | 0.00252 | 0.003804 | 0.00375 | 5e-05 | 0.001042 | 0.30613 | 1.3327 | 2.5475 | 100.61 | 21.1 | 25.952 | 80.19 | 82.4508 | 76.851 | 0.006462 | 1.0087 | 0.024294 | 24.085 | 0.010676 | 0.15339 | 0.00206 | 0.14642 | 0.14642 | 0.043266
2019-03-28 15:52 | 1553802720.2902 | 819 | 0.20909 | 0.000761 | 7.2e-05 | 101.0791 | 101.1078 | 120.8173 | 21.2612 | 0.002422 | 0.002274 | 149.887 | 0.25719 | 8.4032 | 0.002527 | 0.00444 | 0.004326 | 4.9e-05 | 0.001038 | 0.31094 | 1.3316 | 2.5476 | 100.61 | 21.1 | 25.958 | 80.0857 | 82.6207 | 76.5197 | 0.006464 | 1.0097 | 0.024276 | 24.0444 | -0.002538 | 0.153 | 0.002048 | 0.14538 | 0.14538 | 0.043466
2019-03-28 15:52 | 1553802721.2901 | 820 | 0.20908 | 0.000759 | 7.3e-05 | 101.0604 | 100.8952 | 120.4406 | 21.2612 | 0.002489 | 0.002298 | 149.8705 | 0.25685 | 8.4012 | 0.002514 | 0.004111 | 0.004 | 5.2e-05 | 0.00104 | 0.31686 | 1.3016 | 2.5474 | 100.61 | 21.1 | 25.966 | 81.7037 | 82.3027 | 77.1014 | 0.006466 | 1.008 | 0.024218 | 24.0268 | 0.014635 | 0.1528 | 0.002084 | 0.1475 | 0.1475 | 0.042785
2019-03-28 15:52 | 1553802722.2893 | 821 | 0.20905 | 0.000755 | 7.5e-05 | 101.0443 | 101.0175 | 120.65 | 21.2612 | 0.002452 | 0.002326 | 149.8475 | 0.25819 | 8.4041 | 0.00253 | 0.004075 | 0.003965 | 4.9e-05 | 0.001043 | 0.32421 | 1.2854 | 2.5474 | 100.61 | 21.1 | 25.968 | 80.8014 | 82.6878 | 77.7717 | 0.006466 | 1.0061 | 0.024216 | 24.0707 | 0.002376 | 0.1536 | 0.00225 | 0.15949 | 0.15949 | 0.039772
2019-03-28 15:52 | 1553802723.2903 | 822 | 0.2091 | 0.000756 | 7.5e-05 | 101.0285 | 100.8568 | 120.7162 | 21.2612 | 0.002505 | 0.002258 | 149.8277 | 0.25794 | 8.4036 | 0.002478 | 0.003677 | 0.003623 | 5.2e-05 | 0.001042 | 0.32565 | 1.2789 | 2.5472 | 100.61 | 21.1 | 25.968 | 82.0811 | 81.4393 | 76.148 | 0.006466 | 1.0107 | 0.024279 | 24.0215 | 0.015218 | 0.15345 | 0.001982 | 0.14029 | 0.14029 | 0.045173
2019-03-28 15:52 | 1553802724.2893 | 823 | 0.20908 | 0.000758 | 7.4e-05 | 101.0794 | 100.8975 | 120.5856 | 21.2612 | 0.002482 | 0.002323 | 149.8349 | 0.25713 | 8.4012 | 0.002483 | 0.003635 | 0.003577 | 5.1e-05 | 0.001047 | 0.33179 | 1.3114 | 2.5474 | 100.61 | 21.1 | 25.966 | 81.5258 | 81.5626 | 77.6999 | 0.006466 | 1.0063 | 0.024212 | 24.0618 | 0.016111 | 0.15296 | 0.002126 | 0.15052 | 0.15052 | 0.04197
2019-03-28 15:52 | 1553802725.2904 | 824 | 0.20908 | 0.000754 | 7.2e-05 | 101.0654 | 100.8439 | 120.6291 | 21.2612 | 0.002459 | 0.002289 | 149.8337 | 0.25788 | 8.4062 | 0.00251 | 0.003966 | 0.003885 | 5.1e-05 | 0.001051 | 0.34075 | 1.3208 | 2.5474 | 100.61 | 21.1 | 25.964 | 80.9862 | 82.2037 | 76.9017 | 0.006465 | 1.0086 | 0.024244 | 24.0387 | 0.019632 | 0.15341 | 0.002108 | 0.14953 | 0.14953 | 0.042371
2019-03-28 15:52 | 1553802726.2903 | 825 | 0.20904 | 0.000749 | 7e-05 | 101.0172 | 100.8636 | 120.7781 | 21.2612 | 0.002452 | 0.002283 | 149.8428 | 0.25758 | 8.4005 | 0.002503 | 0.004165 | 0.004024 | 5.2e-05 | 0.001048 | 0.29407 | 1.2609 | 2.5475 | 100.61 | 21.1 | 25.96 | 80.8132 | 82.0311 | 76.759 | 0.006464 | 1.009 | 0.024264 | 24.0487 | 0.013613 | 0.15324 | 0.002342 | 0.16679 | 0.16679 | 0.037943
2019-03-28 15:52 | 1553802727.2908 | 826 | 0.2091 | 0.000748 | 7e-05 | 100.9731 | 100.9118 | 120.8389 | 21.2612 | 0.002464 | 0.00232 | 149.8158 | 0.25777 | 8.4009 | 0.002486 | 0.003642 | 0.003544 | 4.8e-05 | 0.001047 | 0.32637 | 1.3299 | 2.5476 | 100.61 | 21.1 | 25.958 | 81.1085 | 81.6162 | 77.6477 | 0.006464 | 1.0064 | 0.02424 | 24.0853 | 0.005432 | 0.15335 | 0.002011 | 0.14252 | 0.14252 | 0.044439
2019-03-28 15:52 | 1553802728.2895 | 827 | 0.2091 | 0.000743 | 7.1e-05 | 100.9943 | 101.0508 | 120.7551 | 21.2612 | 0.002501 | 0.002256 | 149.8051 | 0.25783 | 8.4025 | 0.002464 | 0.003555 | 0.003469 | 4.9e-05 | 0.001045 | 0.30431 | 1.369 | 2.5475 | 100.61 | 21.1 | 25.955 | 81.9883 | 81.0922 | 76.0949 | 0.006463 | 1.0109 | 0.024285 | 24.0236 | -0.005002 | 0.15338 | 0.001986 | 0.14096 | 0.14096 | 0.044941
2019-03-28 15:52 | 1553802729.2901 | 828 | 0.20909 | 0.000743 | 7.2e-05 | 100.9975 | 100.8894 | 120.8694 | 21.2612 | 0.002446 | 0.002362 | 149.8074 | 0.2579 | 8.4022 | 0.002471 | 0.004092 | 0.003963 | 4.8e-05 | 0.001043 | 0.31974 | 1.3432 | 2.5475 | 100.61 | 21.1 | 25.951 | 80.6759 | 81.2739 | 78.6348 | 0.006462 | 1.0036 | 0.024209 | 24.1222 | 0.009583 | 0.15343 | 0.002085 | 0.1476 | 0.1476 | 0.04293
2019-03-28 15:52 | 1553802730.2899 | 829 | 0.20906 | 0.000748 | 7.4e-05 | 101.0035 | 101.0079 | 121.2408 | 21.2612 | 0.002438 | 0.002326 | 149.7863 | 0.25801 | 8.4023 | 0.00242 | 0.003872 | 0.003755 | 4.6e-05 | 0.001043 | 0.33212 | 1.3248 | 2.5475 | 100.61 | 21.09 | 25.949 | 80.4808 | 80.0294 | 77.7818 | 0.006457 | 1.006 | 0.024275 | 24.1299 | -0.000395 | 0.15349 | 0.002227 | 0.15825 | 0.15825 | 0.040058
2019-03-28 15:52 | 1553802731.2886 | 830 | 0.20905 | 0.000742 | 7.6e-05 | 100.991 | 101.0463 | 121.5147 | 21.2612 | 0.002503 | 0.002294 | 149.7782 | 0.2577 | 8.4026 | 0.002388 | 0.004551 | 0.004428 | 5e-05 | 0.001033 | 0.30829 | 1.3424 | 2.5476 | 100.61 | 21.09 | 25.948 | 82.0241 | 79.26 | 77.0086 | 0.006457 | 1.0082 | 0.024329 | 24.1305 | -0.004901 | 0.1533 | 0.002282 | 0.16254 | 0.16254 | 0.038954
2019-03-28 15:52 | 1553802732.2896 | 831 | 0.20908 | 0.000747 | 7.8e-05 | 100.9648 | 100.7399 | 121.3013 | 21.2612 | 0.00244 | 0.002273 | 149.7627 | 0.25744 | 8.4042 | 0.002478 | 0.003631 | 0.003511 | 5e-05 | 0.001025 | 0.34899 | 1.311 | 2.5474 | 100.61 | 21.09 | 25.945 | 80.5137 | 81.4369 | 76.4957 | 0.006456 | 1.0097 | 0.024326 | 24.0916 | 0.019943 | 0.15315 | 0.002096 | 0.14875 | 0.14875 | 0.042523
2019-03-28 15:52 | 1553802733.2896 | 832 | 0.20907 | 0.000741 | 8e-05 | 100.9755 | 100.8894 | 121.5011 | 21.2612 | 0.002446 | 0.002268 | 149.7893 | 0.25839 | 8.4041 | 0.002466 | 0.00445 | 0.004431 | 5e-05 | 0.001021 | 0.31382 | 1.2836 | 2.5474 | 100.61 | 21.09 | 25.943 | 80.6617 | 81.1436 | 76.3772 | 0.006456 | 1.0101 | 0.02435 | 24.1074 | 0.007631 | 0.15372 | 0.002186 | 0.15538 | 0.15538 | 0.040859
2019-03-28 15:52 | 1553802734.2887 | 833 | 0.20909 | 0.000741 | 8.2e-05 | 100.9353 | 100.833 | 121.8088 | 21.2612 | 0.002428 | 0.002289 | 149.7776 | 0.25821 | 8.4022 | 0.002454 | 0.003654 | 0.003512 | 4.7e-05 | 0.001016 | 0.30254 | 1.3301 | 2.5475 | 100.61 | 21.09 | 25.939 | 80.2219 | 80.8451 | 76.9034 | 0.006455 | 1.0085 | 0.024363 | 24.1561 | 0.009069 | 0.15361 | 0.002055 | 0.14568 | 0.14568 | 0.043547
2019-03-28 15:52 | 1553802735.2886 | 834 | 0.20908 | 0.00074 | 8.3e-05 | 100.9201 | 100.9193 | 121.756 | 21.2612 | 0.00245 | 0.002299 | 149.766 | 0.2572 | 8.4021 | 0.002465 | 0.003746 | 0.003579 | 4.7e-05 | 0.001013 | 0.32693 | 1.3155 | 2.5475 | 100.61 | 21.09 | 25.937 | 80.767 | 81.1117 | 77.1247 | 0.006454 | 1.0079 | 0.02435 | 24.1584 | 7.2e-05 | 0.15301 | 0.002114 | 0.14992 | 0.14992 | 0.042152
2019-03-28 15:52 | 1553802736.2897 | 835 | 0.20907 | 0.000736 | 8.5e-05 | 100.9225 | 100.9488 | 122.0306 | 21.2612 | 0.002445 | 0.002329 | 149.7473 | 0.25759 | 8.403 | 0.002518 | 0.004209 | 0.004125 | 4.9e-05 | 0.001013 | 0.32709 | 1.3001 | 2.5472 | 100.61 | 21.09 | 25.935 | 80.6408 | 82.4001 | 77.8589 | 0.006454 | 1.0058 | 0.024351 | 24.211 | -0.002336 | 0.15324 | 0.002168 | 0.15372 | 0.15372 | 0.041171
2019-03-28 15:52 | 1553802737.2894 | 836 | 0.2091 | 0.000732 | 8.6e-05 | 100.8897 | 100.7793 | 122.3475 | 21.2612 | 0.00245 | 0.002343 | 149.7361 | 0.25819 | 8.4032 | 0.002504 | 0.004218 | 0.004154 | 5e-05 | 0.001016 | 0.32278 | 1.2733 | 2.5474 | 100.61 | 21.09 | 25.93 | 80.7514 | 82.0686 | 78.1972 | 0.006453 | 1.0048 | 0.024371 | 24.2541 | 0.009793 | 0.1536 | 0.001992 | 0.14094 | 0.14094 | 0.045009
2019-03-28 15:52 | 1553802738.2887 | 837 | 0.20912 | 0.000732 | 8.6e-05 | 100.9058 | 100.7477 | 122.4111 | 21.2612 | 0.002419 | 0.002262 | 149.7495 | 0.2581 | 8.4036 | 0.002522 | 0.003724 | 0.00366 | 4.9e-05 | 0.001013 | 0.33428 | 1.3196 | 2.5473 | 100.61 | 21.09 | 25.927 | 80.0186 | 82.4883 | 76.2463 | 0.006452 | 1.0104 | 0.024446 | 24.193 | 0.014023 | 0.15354 | 0.0019 | 0.13453 | 0.13453 | 0.047136
2019-03-28 15:52 | 1553802739.2902 | 838 | 0.2091 | 0.000738 | 8.7e-05 | 100.8763 | 100.8371 | 122.393 | 21.2612 | 0.002443 | 0.002304 | 149.6973 | 0.25761 | 8.4013 | 0.002503 | 0.004114 | 0.003939 | 4.6e-05 | 0.001015 | 0.32925 | 1.2879 | 2.5474 | 100.61 | 21.09 | 25.926 | 80.5985 | 82.0342 | 77.2518 | 0.006452 | 1.0075 | 0.024409 | 24.226 | 0.00348 | 0.15326 | 0.001966 | 0.13917 | 0.13917 | 0.04548
2019-03-28 15:52 | 1553802740.2898 | 839 | 0.20915 | 0.000729 | 8.8e-05 | 100.8742 | 100.9037 | 122.3742 | 21.2612 | 0.002463 | 0.002263 | 149.7102 | 0.25818 | 8.4042 | 0.002534 | 0.004322 | 0.004186 | 4.7e-05 | 0.001014 | 0.32372 | 1.27 | 2.5474 | 100.61 | 21.09 | 25.925 | 81.079 | 82.7854 | 76.2554 | 0.006451 | 1.0104 | 0.024442 | 24.1896 | -0.002609 | 0.15359 | 0.001715 | 0.12081 | 0.12081 | 0.052506
2019-03-28 15:52 | 1553802741.2888 | 840 | 0.20912 | 0.000729 | 8.8e-05 | 100.8212 | 100.7348 | 122.4336 | 21.2612 | 0.002461 | 0.002291 | 149.7112 | 0.25825 | 8.4018 | 0.002559 | 0.004691 | 0.004566 | 4.5e-05 | 0.001016 | 0.32117 | 1.2761 | 2.5473 | 100.61 | 21.09 | 25.923 | 81.0227 | 83.3853 | 76.9462 | 0.006451 | 1.0084 | 0.024423 | 24.2194 | 0.007668 | 0.15363 | 0.001889 | 0.1335 | 0.1335 | 0.047529
2019-03-28 15:52 | 1553802742.2903 | 841 | 0.20912 | 0.000733 | 8.9e-05 | 100.8147 | 100.8755 | 122.6275 | 21.2612 | 0.002449 | 0.002323 | 149.7009 | 0.25771 | 8.4036 | 0.002543 | 0.003926 | 0.003879 | 5.1e-05 | 0.001019 | 0.31509 | 1.3169 | 2.5474 | 100.61 | 21.09 | 25.918 | 80.7257 | 82.9881 | 77.7023 | 0.00645 | 1.0063 | 0.024416 | 24.2647 | -0.005393 | 0.15331 | 0.001874 | 0.13233 | 0.13233 | 0.047849
2019-03-28 15:52 | 1553802743.2889 | 842 | 0.20912 | 0.000723 | 8.8e-05 | 100.8715 | 100.7318 | 122.4684 | 21.2612 | 0.002409 | 0.002283 | 149.6603 | 0.25778 | 8.403 | 0.002461 | 0.004292 | 0.004219 | 5e-05 | 0.001024 | 0.31222 | 1.3032 | 2.5474 | 100.61 | 21.09 | 25.914 | 79.7699 | 81.0307 | 76.7465 | 0.006449 | 1.009 | 0.024434 | 24.2159 | 0.012404 | 0.15335 | 0.001901 | 0.13442 | 0.13442 | 0.047117
2019-03-28 15:52 | 1553802744.2879 | 843 | 0.20911 | 0.000725 | 8.9e-05 | 100.8293 | 100.7498 | 122.3113 | 21.2612 | 0.002435 | 0.002344 | 149.6632 | 0.25802 | 8.4023 | 0.002551 | 0.003545 | 0.003411 | 4.6e-05 | 0.001025 | 0.33395 | 1.2839 | 2.5472 | 100.61 | 21.09 | 25.91 | 80.3904 | 83.198 | 78.2128 | 0.006448 | 1.0048 | 0.024367 | 24.2511 | 0.007054 | 0.15349 | 0.00197 | 0.1391 | 0.1391 | 0.045573
2019-03-28 15:52 | 1553802745.2884 | 844 | 0.20913 | 0.000725 | 8.8e-05 | 100.8352 | 100.763 | 122.1921 | 21.2612 | 0.002417 | 0.002296 | 149.6715 | 0.25717 | 8.4018 | 0.002511 | 0.004126 | 0.004089 | 5e-05 | 0.001028 | 0.34235 | 1.3297 | 2.5476 | 100.61 | 21.09 | 25.903 | 79.967 | 82.2378 | 77.0521 | 0.006446 | 1.0081 | 0.024396 | 24.1992 | 0.006407 | 0.15299 | 0.001849 | 0.13037 | 0.13037 | 0.048463
2019-03-28 15:52 | 1553802746.2891 | 845 | 0.20911 | 0.000719 | 8.8e-05 | 100.8295 | 100.7735 | 121.7691 | 21.2612 | 0.002434 | 0.002271 | 149.6656 | 0.25796 | 8.4013 | 0.002554 | 0.00384 | 0.003761 | 4.7e-05 | 0.001028 | 0.32803 | 1.3089 | 2.5474 | 100.61 | 21.09 | 25.9 | 80.3664 | 83.2518 | 76.4599 | 0.006445 | 1.0098 | 0.024374 | 24.1368 | 0.00497 | 0.15346 | 0.001941 | 0.13709 | 0.13709 | 0.046231
2019-03-28 15:52 | 1553802747.2894 | 846 | 0.20914 | 0.000721 | 8.9e-05 | 100.829 | 100.4674 | 121.7327 | 21.2612 | 0.002494 | 0.002329 | 149.6756 | 0.25738 | 8.4046 | 0.002543 | 0.003948 | 0.003818 | 4.9e-05 | 0.001025 | 0.31078 | 1.3087 | 2.5475 | 100.61 | 21.09 | 25.891 | 81.8205 | 82.9929 | 77.8637 | 0.006443 | 1.0058 | 0.024322 | 24.1816 | 0.032127 | 0.15312 | 0.001789 | 0.12558 | 0.12558 | 0.050357
2019-03-28 15:52 | 1553802748.2894 | 847 | 0.20911 | 0.000719 | 8.9e-05 | 100.7777 | 100.7175 | 121.5874 | 21.2612 | 0.00242 | 0.002295 | 149.6576 | 0.25768 | 8.4032 | 0.002502 | 0.004151 | 0.004087 | 4.8e-05 | 0.001026 | 0.34164 | 1.2982 | 2.5476 | 100.61 | 21.09 | 25.887 | 80.0451 | 82.0053 | 77.0365 | 0.006442 | 1.0082 | 0.024336 | 24.1387 | 0.005348 | 0.15329 | 0.001976 | 0.13931 | 0.13931 | 0.045447
2019-03-28 15:52 | 1553802749.2881 | 848 | 0.20913 | 0.000719 | 9e-05 | 100.6744 | 100.6726 | 121.4161 | 21.2612 | 0.002437 | 0.002316 | 149.6945 | 0.25718 | 8.4082 | 0.002538 | 0.003794 | 0.003681 | 4.9e-05 | 0.001024 | 0.31564 | 1.3236 | 2.5475 | 100.61 | 21.09 | 25.874 | 80.4575 | 82.8775 | 77.541 | 0.006439 | 1.0067 | 0.024301 | 24.139 | 0.000153 | 0.153 | 0.001844 | 0.1295 | 0.1295 | 0.048792
2019-03-28 15:52 | 1553802750.2881 | 849 | 0.20908 | 0.000717 | 9.1e-05 | 100.7173 | 100.6213 | 121.2443 | 21.2612 | 0.002445 | 0.002357 | 149.6754 | 0.25759 | 8.4026 | 0.002549 | 0.003616 | 0.00352 | 4.9e-05 | 0.00102 | 0.28799 | 1.2697 | 2.5473 | 100.61 | 21.09 | 25.866 | 80.6283 | 83.1357 | 78.5192 | 0.006437 | 1.0039 | 0.02425 | 24.1556 | 0.008522 | 0.15324 | 0.002097 | 0.14764 | 0.14764 | 0.042865
2019-03-28 15:52 | 1553802751.2883 | 850 | 0.20911 | 0.000711 | 9e-05 | 100.6967 | 100.4477 | 121.2663 | 21.2612 | 0.002445 | 0.002259 | 149.6831 | 0.25754 | 8.4044 | 0.00249 | 0.003891 | 0.003838 | 4.9e-05 | 0.00102 | 0.30901 | 1.2577 | 2.5474 | 100.61 | 21.09 | 25.857 | 80.6447 | 81.7207 | 76.1675 | 0.006435 | 1.0107 | 0.024334 | 24.0769 | 0.022142 | 0.15321 | 0.001945 | 0.13705 | 0.13705 | 0.046169
2019-03-28 15:53 | 1553802752.2882 | 851 | 0.20911 | 0.000715 | 9e-05 | 100.6249 | 100.6601 | 121.1833 | 21.2612 | 0.002463 | 0.002254 | 149.691 | 0.25786 | 8.4021 | 0.002601 | 0.003408 | 0.003275 | 4.6e-05 | 0.001021 | 0.32709 | 1.2933 | 2.5475 | 100.61 | 21.09 | 25.84 | 81.0835 | 84.3916 | 76.0512 | 0.00643 | 1.011 | 0.02433 | 24.0646 | -0.003124 | 0.1534 | 0.001924 | 0.13542 | 0.13542 | 0.046784
2019-03-28 15:53 | 1553802753.2886 | 852 | 0.20911 | 0.000709 | 8.9e-05 | 100.6255 | 100.6349 | 121.1879 | 21.2612 | 0.002392 | 0.002284 | 149.7087 | 0.25698 | 8.4032 | 0.00259 | 0.004339 | 0.004295 | 4.8e-05 | 0.001022 | 0.3246 | 1.2973 | 2.5475 | 100.61 | 21.09 | 25.831 | 79.3559 | 84.1208 | 76.7684 | 0.006428 | 1.0089 | 0.024305 | 24.0898 | -0.000837 | 0.15288 | 0.001988 | 0.14001 | 0.14001 | 0.045096
2019-03-28 15:53 | 1553802754.2893 | 853 | 0.20912 | 0.000717 | 9e-05 | 100.6062 | 100.4549 | 121.4224 | 21.2612 | 0.002409 | 0.002279 | 149.6998 | 0.25733 | 8.4007 | 0.002531 | 0.003823 | 0.0037 | 4.6e-05 | 0.001023 | 0.32565 | 1.2682 | 2.5472 | 100.61 | 21.09 | 25.821 | 79.7637 | 82.7159 | 76.6593 | 0.006426 | 1.0093 | 0.024332 | 24.1093 | 0.01346 | 0.15308 | 0.001906 | 0.13421 | 0.13421 | 0.047109
2019-03-28 15:53 | 1553802755.2883 | 854 | 0.20911 | 0.000709 | 9e-05 | 100.6458 | 100.6451 | 121.4783 | 21.2612 | 0.002422 | 0.00226 | 149.6867 | 0.25844 | 8.4026 | 0.002571 | 0.004028 | 0.003919 | 4.6e-05 | 0.001027 | 0.29999 | 1.3072 | 2.5474 | 100.61 | 21.09 | 25.807 | 80.0785 | 83.6582 | 76.204 | 0.006422 | 1.0106 | 0.024354 | 24.0992 | 5.8e-05 | 0.15375 | 0.001952 | 0.13769 | 0.13769 | 0.046117
2019-03-28 15:53 | 1553802756.2887 | 855 | 0.20911 | 0.000711 | 9.1e-05 | 100.5916 | 100.5446 | 121.8489 | 21.2612 | 0.002416 | 0.002273 | 149.6754 | 0.25846 | 8.405 | 0.002559 | 0.003912 | 0.003899 | 4.7e-05 | 0.001028 | 0.33964 | 1.3131 | 2.5475 | 100.61 | 21.09 | 25.801 | 79.9485 | 83.3759 | 76.5135 | 0.006421 | 1.0097 | 0.02438 | 24.1466 | 0.00418 | 0.15376 | 0.001936 | 0.1366 | 0.1366 | 0.046488
2019-03-28 15:53 | 1553802757.2884 | 856 | 0.20914 | 0.000706 | 9.1e-05 | 100.6076 | 100.3947 | 121.6506 | 21.2612 | 0.002413 | 0.002283 | 149.6477 | 0.25751 | 8.4039 | 0.002566 | 0.004034 | 0.003911 | 4.8e-05 | 0.001027 | 0.33859 | 1.2892 | 2.5474 | 100.61 | 21.09 | 25.786 | 79.8742 | 83.5584 | 76.7506 | 0.006417 | 1.009 | 0.024352 | 24.1351 | 0.018948 | 0.15319 | 0.001807 | 0.12694 | 0.12694 | 0.049843
2019-03-28 15:53 | 1553802758.2884 | 857 | 0.20909 | 0.000709 | 9.2e-05 | 100.5686 | 100.407 | 121.833 | 21.2612 | 0.002465 | 0.002333 | 149.6486 | 0.25772 | 8.403 | 0.002487 | 0.003638 | 0.003539 | 4.9e-05 | 0.001025 | 0.30934 | 1.3008 | 2.5474 | 100.61 | 21.09 | 25.779 | 81.1186 | 81.6522 | 77.9517 | 0.006415 | 1.0055 | 0.024329 | 24.1946 | 0.014389 | 0.15332 | 0.0021 | 0.14828 | 0.14828 | 0.042705
2019-03-28 15:53 | 1553802759.2875 | 858 | 0.20913 | 0.00071 | 9.2e-05 | 100.5469 | 100.5653 | 121.913 | 21.2612 | 0.002461 | 0.002296 | 149.6287 | 0.2575 | 8.4031 | 0.002538 | 0.004173 | 0.004083 | 5.1e-05 | 0.001024 | 0.31238 | 1.29 | 2.5474 | 100.61 | 21.09 | 25.772 | 81.022 | 82.8768 | 77.0576 | 0.006413 | 1.0081 | 0.024368 | 24.1717 | -0.001637 | 0.15318 | 0.001825 | 0.12831 | 0.12831 | 0.049305
2019-03-28 15:53 | 1553802760.2878 | 859 | 0.20913 | 0.000706 | 9.1e-05 | 100.5537 | 100.3125 | 122.0116 | 21.2612 | 0.00245 | 0.002297 | 149.6312 | 0.25782 | 8.4043 | 0.002596 | 0.004055 | 0.003914 | 4.6e-05 | 0.001022 | 0.32372 | 1.2973 | 2.5473 | 100.61 | 21.09 | 25.762 | 80.7593 | 84.263 | 77.0778 | 0.006411 | 1.008 | 0.024377 | 24.1822 | 0.021481 | 0.15338 | 0.001848 | 0.13006 | 0.13006 | 0.048705
2019-03-28 15:53 | 1553802761.2874 | 860 | 0.20917 | 0.000712 | 9.1e-05 | 100.5052 | 100.2269 | 122.4085 | 21.2612 | 0.002423 | 0.002292 | 149.6217 | 0.25784 | 8.4024 | 0.002607 | 0.003727 | 0.00368 | 4.5e-05 | 0.001024 | 0.30879 | 1.3464 | 2.5474 | 100.61 | 21.1 | 25.757 | 80.1035 | 84.5251 | 76.9611 | 0.006414 | 1.0084 | 0.02442 | 24.2174 | 0.024799 | 0.15339 | 0.001601 | 0.11217 | 0.11217 | 0.056476
2019-03-28 15:53 | 1553802762.2886 | 861 | 0.20913 | 0.000702 | 9.1e-05 | 100.4513 | 100.3288 | 122.5885 | 21.2612 | 0.002483 | 0.002297 | 149.6023 | 0.25821 | 8.4041 | 0.00261 | 0.003808 | 0.003636 | 4.6e-05 | 0.001022 | 0.31974 | 1.3285 | 2.5475 | 100.61 | 21.09 | 25.751 | 81.5515 | 84.5953 | 77.0826 | 0.006408 | 1.008 | 0.024434 | 24.2395 | 0.010909 | 0.15361 | 0.001842 | 0.1299 | 0.1299 | 0.04884
2019-03-28 15:53 | 1553802763.2887 | 862 | 0.20916 | 0.000707 | 9.1e-05 | 100.4864 | 100.2867 | 122.6592 | 21.2612 | 0.002477 | 0.002291 | 149.604 | 0.25825 | 8.4017 | 0.002589 | 0.004397 | 0.004317 | 4.9e-05 | 0.001021 | 0.32997 | 1.2948 | 2.5475 | 100.61 | 21.09 | 25.749 | 81.401 | 84.0929 | 76.9486 | 0.006408 | 1.0084 | 0.024446 | 24.2418 | 0.017794 | 0.15363 | 0.001695 | 0.11921 | 0.11921 | 0.053224
2019-03-28 15:53 | 1553802764.2885 | 863 | 0.2091 | 0.000706 | 9.1e-05 | 100.4301 | 100.2894 | 122.3794 | 21.2612 | 0.002441 | 0.002257 | 149.5975 | 0.25697 | 8.4024 | 0.002559 | 0.003988 | 0.003865 | 5.1e-05 | 0.001028 | 0.32532 | 1.3781 | 2.5473 | 100.61 | 21.1 | 25.747 | 80.5332 | 83.3752 | 76.1214 | 0.006411 | 1.0108 | 0.024447 | 24.1855 | 0.012539 | 0.15287 | 0.002042 | 0.14471 | 0.14471 | 0.043628
2019-03-28 15:53 | 1553802765.2876 | 864 | 0.20909 | 0.000702 | 9e-05 | 100.3071 | 100.1437 | 122.5294 | 21.2612 | 0.002452 | 0.002309 | 149.6171 | 0.25743 | 8.4016 | 0.002667 | 0.00378 | 0.00374 | 5e-05 | 0.001033 | 0.3194 | 1.3112 | 2.5474 | 100.61 | 21.09 | 25.745 | 80.8146 | 85.9851 | 77.382 | 0.006407 | 1.0072 | 0.024418 | 24.244 | 0.014585 | 0.15315 | 0.00208 | 0.14742 | 0.14742 | 0.042904
2019-03-28 15:53 | 1553802766.2886 | 865 | 0.2091 | 0.000705 | 8.9e-05 | 100.3819 | 100.0044 | 122.4163 | 21.2612 | 0.002419 | 0.002291 | 149.6143 | 0.25713 | 8.4067 | 0.002696 | 0.004545 | 0.004492 | 4.6e-05 | 0.001032 | 0.3079 | 1.326 | 2.5474 | 100.61 | 21.09 | 25.745 | 80.0256 | 86.6761 | 76.9385 | 0.006407 | 1.0084 | 0.024422 | 24.2174 | 0.033698 | 0.15296 | 0.002049 | 0.1452 | 0.1452 | 0.04351
2019-03-28 15:53 | 1553802767.2887 | 866 | 0.20911 | 0.000705 | 8.9e-05 | 100.3342 | 100.1046 | 122.2245 | 21.2612 | 0.002449 | 0.002312 | 149.6229 | 0.25839 | 8.4049 | 0.002602 | 0.003835 | 0.003761 | 5e-05 | 0.001029 | 0.31902 | 1.2962 | 2.5474 | 100.61 | 21.09 | 25.745 | 80.7389 | 84.4065 | 77.4521 | 0.006407 | 1.007 | 0.024385 | 24.2162 | 0.020491 | 0.15371 | 0.00195 | 0.1377 | 0.1377 | 0.046104
2019-03-28 15:53 | 1553802768.2877 | 867 | 0.20918 | 0.000697 | 8.9e-05 | 100.3211 | 100.2993 | 122.0912 | 21.2612 | 0.002455 | 0.002286 | 149.6217 | 0.25757 | 8.4032 | 0.002577 | 0.004674 | 0.0045 | 4.8e-05 | 0.00103 | 0.35242 | 1.2907 | 2.5475 | 100.61 | 21.09 | 25.747 | 80.8723 | 83.8111 | 76.8114 | 0.006407 | 1.0088 | 0.024394 | 24.1809 | 0.001947 | 0.15323 | 0.001538 | 0.1075 | 0.1075 | 0.058869
2019-03-28 15:53 | 1553802769.2877 | 868 | 0.20925 | 0.000705 | 8.9e-05 | 100.3679 | 100.2456 | 122.1412 | 21.2612 | 0.002494 | 0.002302 | 149.634 | 0.25779 | 8.4025 | 0.002585 | 0.004773 | 0.00476 | 5.1e-05 | 0.00103 | 0.35203 | 1.2898 | 2.5474 | 100.61 | 21.09 | 25.749 | 81.8198 | 83.9984 | 77.1965 | 0.006408 | 1.0077 | 0.024386 | 24.1991 | 0.010903 | 0.15336 | 0.001092 | 0.074764 | 0.074764 | 0.084716
2019-03-28 15:53 | 1553802770.2878 | 869 | 0.20904 | 0.000699 | 9e-05 | 100.2803 | 100.109 | 122.0461 | 21.2612 | 0.002494 | 0.002327 | 149.6318 | 0.25812 | 8.4018 | 0.00259 | 0.004005 | 0.003923 | 5.3e-05 | 0.001032 | 0.30486 | 1.3022 | 2.5475 | 100.61 | 21.09 | 25.753 | 81.8058 | 84.1319 | 77.795 | 0.006409 | 1.006 | 0.024355 | 24.2103 | 0.01529 | 0.15356 | 0.002401 | 0.17059 | 0.17059 | 0.037176
2019-03-28 15:53 | 1553802771.2888 | 870 | 0.20902 | 0.000695 | 9.1e-05 | 100.2809 | 100.2996 | 121.6079 | 21.2612 | 0.002412 | 0.002274 | 149.6089 | 0.25728 | 8.4024 | 0.00258 | 0.004293 | 0.00417 | 5e-05 | 0.001033 | 0.32781 | 1.2939 | 2.5473 | 100.61 | 21.09 | 25.755 | 79.8446 | 83.8893 | 76.536 | 0.006409 | 1.0096 | 0.024355 | 24.1235 | -0.001669 | 0.15305 | 0.002542 | 0.18084 | 0.18084 | 0.034955
2019-03-28 15:53 | 1553802772.2878 | 871 | 0.20905 | 0.000697 | 9.1e-05 | 100.2785 | 100.0768 | 121.3696 | 21.2612 | 0.002395 | 0.002245 | 149.6244 | 0.2582 | 8.4016 | 0.002563 | 0.004205 | 0.004088 | 4.9e-05 | 0.001038 | 0.31636 | 1.283 | 2.5472 | 100.61 | 21.09 | 25.756 | 79.4407 | 83.4663 | 75.8439 | 0.006409 | 1.0116 | 0.024355 | 24.0759 | 0.018014 | 0.1536 | 0.002372 | 0.1684 | 0.1684 | 0.03767
2019-03-28 15:53 | 1553802773.2888 | 872 | 0.20909 | 0.000701 | 9.1e-05 | 100.2335 | 100.0856 | 121.3767 | 21.2612 | 0.00243 | 0.002296 | 149.6085 | 0.25779 | 8.4003 | 0.002552 | 0.003713 | 0.003658 | 5e-05 | 0.001034 | 0.31526 | 1.3589 | 2.5474 | 100.61 | 21.09 | 25.758 | 80.2727 | 83.2184 | 77.0545 | 0.00641 | 1.0081 | 0.024314 | 24.1184 | 0.013212 | 0.15336 | 0.002119 | 0.14962 | 0.14962 | 0.042332
2019-03-28 15:53 | 1553802774.2879 | 873 | 0.2091 | 0.0007 | 9.1e-05 | 100.1394 | 100.1899 | 121.3565 | 21.2612 | 0.002469 | 0.002289 | 149.5951 | 0.25718 | 8.4023 | 0.002549 | 0.003893 | 0.003729 | 5.1e-05 | 0.00103 | 0.31813 | 1.3707 | 2.5474 | 100.61 | 21.09 | 25.758 | 81.2232 | 83.1427 | 76.8808 | 0.00641 | 1.0086 | 0.024318 | 24.1104 | -0.004507 | 0.153 | 0.002038 | 0.14369 | 0.14369 | 0.043975
2019-03-28 15:53 | 1553802775.287 | 874 | 0.20906 | 0.000696 | 9.1e-05 | 100.2121 | 99.9799 | 121.103 | 21.2612 | 0.00241 | 0.002269 | 149.5825 | 0.25775 | 8.403 | 0.002502 | 0.004088 | 0.003911 | 4.9e-05 | 0.001032 | 0.32764 | 1.3243 | 2.5474 | 100.61 | 21.09 | 25.757 | 79.7971 | 82.0178 | 76.4162 | 0.00641 | 1.01 | 0.024309 | 24.0692 | 0.020746 | 0.15334 | 0.00226 | 0.15991 | 0.15991 | 0.039602
2019-03-28 15:53 | 1553802776.2871 | 875 | 0.20913 | 0.000697 | 9.1e-05 | 100.2666 | 100.1702 | 120.8255 | 21.2612 | 0.002488 | 0.00231 | 149.5709 | 0.25761 | 8.4015 | 0.002483 | 0.004041 | 0.003959 | 4.8e-05 | 0.00103 | 0.34578 | 1.3121 | 2.5473 | 100.61 | 21.09 | 25.756 | 81.6794 | 81.5484 | 77.3959 | 0.006409 | 1.0071 | 0.024247 | 24.0753 | 0.008605 | 0.15325 | 0.001883 | 0.13196 | 0.13196 | 0.047962
2019-03-28 15:53 | 1553802777.288 | 876 | 0.20911 | 0.000696 | 9e-05 | 100.1811 | 99.914 | 120.8981 | 21.2612 | 0.002446 | 0.002265 | 149.5794 | 0.25665 | 8.4029 | 0.002617 | 0.003471 | 0.003391 | 5.1e-05 | 0.001019 | 0.33339 | 1.3271 | 2.5473 | 100.61 | 21.09 | 25.756 | 80.6665 | 84.7834 | 76.3092 | 0.006409 | 1.0103 | 0.024292 | 24.0451 | 0.023878 | 0.15268 | 0.002007 | 0.14127 | 0.14127 | 0.044637
2019-03-28 15:53 | 1553802778.2884 | 877 | 0.20912 | 0.000696 | 9e-05 | 100.2145 | 100.0703 | 120.957 | 21.2612 | 0.002491 | 0.002278 | 149.5742 | 0.25748 | 8.4029 | 0.002539 | 0.004034 | 0.003814 | 5.1e-05 | 0.001015 | 0.33588 | 1.3323 | 2.5475 | 100.61 | 21.09 | 25.755 | 81.7513 | 82.889 | 76.629 | 0.006409 | 1.0093 | 0.024287 | 24.062 | 0.012877 | 0.15317 | 0.001937 | 0.13615 | 0.13615 | 0.046465
2019-03-28 15:53 | 1553802779.2881 | 878 | 0.20913 | 0.000701 | 9e-05 | 100.2335 | 100.0183 | 121.107 | 21.2612 | 0.00251 | 0.002288 | 149.5732 | 0.25656 | 8.4035 | 0.002571 | 0.004452 | 0.004464 | 4.9e-05 | 0.001009 | 0.3314 | 1.311 | 2.5475 | 100.61 | 21.09 | 25.753 | 82.2058 | 83.6592 | 76.8621 | 0.006409 | 1.0087 | 0.024294 | 24.0849 | 0.019226 | 0.15263 | 0.001845 | 0.12951 | 0.12951 | 0.048673
2019-03-28 15:53 | 1553802780.2881 | 879 | 0.20913 | 0.000701 | 8.9e-05 | 100.0819 | 99.9932 | 121.1187 | 21.2612 | 0.002455 | 0.002235 | 149.5662 | 0.25746 | 8.4013 | 0.002536 | 0.00372 | 0.003611 | 4.8e-05 | 0.001007 | 0.30342 | 1.3739 | 2.5475 | 100.61 | 21.09 | 25.751 | 80.8917 | 82.8261 | 75.589 | 0.006408 | 1.0123 | 0.024339 | 24.0423 | 0.007935 | 0.15317 | 0.001856 | 0.13056 | 0.13056 | 0.048452
2019-03-28 15:53 | 1553802781.2871 | 880 | 0.20912 | 0.000702 | 8.9e-05 | 100.1519 | 100.0336 | 121.0035 | 21.2612 | 0.002436 | 0.002282 | 149.5817 | 0.25782 | 8.4022 | 0.002553 | 0.004233 | 0.004116 | 4.8e-05 | 0.001008 | 0.32101 | 1.3308 | 2.5474 | 100.61 | 21.09 | 25.749 | 80.4303 | 83.2278 | 76.7239 | 0.006408 | 1.0091 | 0.024288 | 24.0699 | 0.010573 | 0.15338 | 0.001894 | 0.13312 | 0.13312 | 0.047585
2019-03-28 15:53 | 1553802782.2882 | 881 | 0.20912 | 0.000695 | 8.8e-05 | 100.0447 | 99.9144 | 121.0649 | 21.2612 | 0.002455 | 0.002267 | 149.5506 | 0.2578 | 8.4015 | 0.002548 | 0.003839 | 0.003694 | 5e-05 | 0.001009 | 0.34683 | 1.3118 | 2.5474 | 100.61 | 21.09 | 25.748 | 80.8684 | 83.1076 | 76.3717 | 0.006407 | 1.0101 | 0.024306 | 24.0639 | 0.01166 | 0.15336 | 0.001931 | 0.13599 | 0.13599 | 0.046575
2019-03-28 15:53 | 1553802783.2882 | 882 | 0.20911 | 0.000695 | 8.8e-05 | 100.1108 | 99.9327 | 121.0044 | 21.2612 | 0.002449 | 0.00226 | 149.5561 | 0.25717 | 8.4053 | 0.002533 | 0.004144 | 0.004084 | 5e-05 | 0.001009 | 0.3246 | 1.2944 | 2.5475 | 100.61 | 21.09 | 25.747 | 80.7336 | 82.7573 | 76.2026 | 0.006407 | 1.0106 | 0.024306 | 24.052 | 0.015928 | 0.15299 | 0.001989 | 0.14022 | 0.14022 | 0.045062
2019-03-28 15:53 | 1553802784.2886 | 883 | 0.20911 | 0.00069 | 8.9e-05 | 100.0271 | 99.769 | 121.16 | 21.2612 | 0.002453 | 0.002307 | 149.559 | 0.258 | 8.4041 | 0.002485 | 0.003832 | 0.003669 | 5.1e-05 | 0.001006 | 0.32388 | 1.3141 | 2.5474 | 100.61 | 21.09 | 25.746 | 80.824 | 81.5936 | 77.325 | 0.006407 | 1.0073 | 0.024283 | 24.1061 | 0.023115 | 0.15349 | 0.001959 | 0.13781 | 0.13781 | 0.045997
2019-03-28 15:53 | 1553802785.2883 | 884 | 0.20912 | 0.000692 | 8.8e-05 | 100.0906 | 100.003 | 121.1262 | 21.2612 | 0.002455 | 0.002312 | 149.5495 | 0.25775 | 8.4039 | 0.002474 | 0.003936 | 0.003841 | 5.1e-05 | 0.001009 | 0.34739 | 1.304 | 2.5472 | 100.61 | 21.09 | 25.744 | 80.8684 | 81.3458 | 77.4452 | 0.006406 | 1.007 | 0.024275 | 24.1069 | 0.007826 | 0.15333 | 0.001918 | 0.13481 | 0.13481 | 0.046973
2019-03-28 15:53 | 1553802786.2882 | 885 | 0.20913 | 0.000696 | 8.8e-05 | 100.0435 | 99.9032 | 121.2703 | 21.2612 | 0.002472 | 0.002315 | 149.5718 | 0.258 | 8.4019 | 0.002541 | 0.004076 | 0.004048 | 5.1e-05 | 0.001008 | 0.32421 | 1.305 | 2.5474 | 100.61 | 21.09 | 25.743 | 81.2864 | 82.9453 | 77.5192 | 0.006406 | 1.0068 | 0.024287 | 24.1238 | 0.012557 | 0.15348 | 0.001858 | 0.13052 | 0.13052 | 0.048567
2019-03-28 15:53 | 1553802787.2883 | 886 | 0.20911 | 0.00069 | 8.9e-05 | 100.0492 | 100.0346 | 121.3546 | 21.2612 | 0.002507 | 0.002285 | 149.5473 | 0.25789 | 8.3999 | 0.002517 | 0.003941 | 0.003849 | 4.7e-05 | 0.001005 | 0.30917 | 1.3368 | 2.5473 | 100.61 | 21.09 | 25.739 | 82.1186 | 82.3657 | 76.802 | 0.006405 | 1.0088 | 0.024321 | 24.1075 | 0.0013 | 0.15342 | 0.002013 | 0.142 | 0.142 | 0.04462
2019-03-28 15:53 | 1553802788.2873 | 887 | 0.20914 | 0.000692 | 8.8e-05 | 100.003 | 99.8913 | 121.2422 | 21.2612 | 0.002455 | 0.002311 | 149.5516 | 0.25819 | 8.4037 | 0.002497 | 0.00377 | 0.003746 | 4.9e-05 | 0.001003 | 0.32659 | 1.3376 | 2.5474 | 100.61 | 21.09 | 25.736 | 80.8872 | 81.8823 | 77.4147 | 0.006404 | 1.0071 | 0.024288 | 24.1174 | 0.009999 | 0.1536 | 0.001828 | 0.12834 | 0.12834 | 0.049429
2019-03-28 15:53 | 1553802789.2874 | 888 | 0.20915 | 0.00069 | 8.9e-05 | 99.9679 | 99.6783 | 121.2706 | 21.2612 | 0.002499 | 0.002295 | 149.535 | 0.25807 | 8.4046 | 0.002542 | 0.004019 | 0.003953 | 4.8e-05 | 0.001008 | 0.34346 | 1.3236 | 2.5474 | 100.61 | 21.09 | 25.732 | 81.9351 | 82.9821 | 77.0444 | 0.006403 | 1.0081 | 0.024304 | 24.1075 | 0.025949 | 0.15353 | 0.001727 | 0.12094 | 0.12094 | 0.052427
2019-03-28 15:53 | 1553802790.2874 | 889 | 0.20914 | 0.000692 | 8.9e-05 | 99.9354 | 99.8593 | 121.5945 | 21.2612 | 0.002443 | 0.002289 | 149.5367 | 0.2586 | 8.4013 | 0.002503 | 0.003953 | 0.003836 | 5.1e-05 | 0.001008 | 0.34219 | 1.3145 | 2.5473 | 100.61 | 21.09 | 25.721 | 80.6033 | 82.0296 | 76.9017 | 0.006401 | 1.0086 | 0.024341 | 24.1347 | 0.00681 | 0.15384 | 0.001789 | 0.12573 | 0.12573 | 0.050537
2019-03-28 15:53 | 1553802791.2864 | 890 | 0.20914 | 0.000694 | 8.8e-05 | 99.8782 | 99.6494 | 121.7644 | 21.2612 | 0.002464 | 0.002254 | 149.5268 | 0.25802 | 8.4008 | 0.002458 | 0.004565 | 0.004482 | 5e-05 | 0.001009 | 0.3282 | 1.3096 | 2.5476 | 100.61 | 21.09 | 25.717 | 81.1023 | 80.9494 | 76.0575 | 0.0064 | 1.011 | 0.024388 | 24.1225 | 0.020515 | 0.1535 | 0.00178 | 0.12529 | 0.12529 | 0.050598
2019-03-28 15:53 | 1553802792.2875 | 891 | 0.20911 | 0.000691 | 8.8e-05 | 99.9372 | 99.911 | 121.7212 | 21.2612 | 0.002409 | 0.002271 | 149.5459 | 0.25798 | 8.4034 | 0.002525 | 0.003213 | 0.003122 | 4.9e-05 | 0.001013 | 0.31404 | 1.3272 | 2.5475 | 100.61 | 21.09 | 25.709 | 79.7846 | 82.5599 | 76.4536 | 0.006398 | 1.0098 | 0.024369 | 24.1319 | 0.002346 | 0.15347 | 0.002018 | 0.14274 | 0.14274 | 0.044404
2019-03-28 15:53 | 1553802793.2865 | 892 | 0.20916 | 0.000692 | 8.8e-05 | 99.8702 | 99.7897 | 121.709 | 21.2612 | 0.002451 | 0.002295 | 149.5259 | 0.25758 | 8.4031 | 0.002508 | 0.004147 | 0.004035 | 4.8e-05 | 0.001018 | 0.33179 | 1.345 | 2.5475 | 100.61 | 21.1 | 25.706 | 80.7927 | 82.1575 | 77.0427 | 0.006401 | 1.0081 | 0.024348 | 24.151 | 0.00721 | 0.15324 | 0.001688 | 0.11836 | 0.11836 | 0.053467
2019-03-28 15:53 | 1553802794.2875 | 893 | 0.20913 | 0.000687 | 8.8e-05 | 99.8711 | 99.7088 | 121.7217 | 21.2612 | 0.002396 | 0.002227 | 149.5212 | 0.25754 | 8.4011 | 0.002489 | 0.003986 | 0.003858 | 4.6e-05 | 0.00102 | 0.33627 | 1.308 | 2.5474 | 100.61 | 21.09 | 25.704 | 79.4695 | 81.6967 | 75.4013 | 0.006396 | 1.0129 | 0.024406 | 24.0956 | 0.01454 | 0.15321 | 0.001844 | 0.13012 | 0.13012 | 0.048627
2019-03-28 15:53 | 1553802795.2865 | 894 | 0.20911 | 0.000694 | 8.7e-05 | 99.8112 | 99.6372 | 121.8806 | 21.2612 | 0.002461 | 0.002305 | 149.4935 | 0.25806 | 8.4032 | 0.002529 | 0.003618 | 0.00352 | 4.7e-05 | 0.001012 | 0.31133 | 1.3257 | 2.5474 | 100.61 | 21.09 | 25.702 | 81.0321 | 82.6575 | 77.2861 | 0.006396 | 1.0074 | 0.024356 | 24.1764 | 0.01561 | 0.15352 | 0.001974 | 0.1394 | 0.1394 | 0.045484
2019-03-28 15:53 | 1553802796.2876 | 895 | 0.20915 | 0.000689 | 8.8e-05 | 99.8758 | 99.7516 | 121.944 | 21.2612 | 0.002421 | 0.002253 | 149.4645 | 0.25717 | 8.4019 | 0.002467 | 0.004627 | 0.004501 | 4.5e-05 | 0.001004 | 0.33987 | 1.3091 | 2.5474 | 100.61 | 21.1 | 25.704 | 80.0614 | 81.1593 | 76.0342 | 0.0064 | 1.0111 | 0.024406 | 24.1394 | 0.011128 | 0.15299 | 0.001737 | 0.12228 | 0.12228 | 0.051675
2019-03-28 15:53 | 1553802797.2866 | 896 | 0.20912 | 0.000686 | 8.8e-05 | 99.8091 | 99.4422 | 121.7529 | 21.2612 | 0.00245 | 0.002327 | 149.4986 | 0.25776 | 8.402 | 0.002485 | 0.003351 | 0.003294 | 4.9e-05 | 0.001003 | 0.31222 | 1.2985 | 2.5475 | 100.61 | 21.1 | 25.708 | 80.7562 | 81.5945 | 77.7943 | 0.006401 | 1.006 | 0.024326 | 24.1812 | 0.032945 | 0.15334 | 0.001965 | 0.13859 | 0.13859 | 0.045695
2019-03-28 15:53 | 1553802798.2877 | 897 | 0.20914 | 0.000691 | 8.8e-05 | 99.8606 | 99.5835 | 121.5217 | 21.2612 | 0.002417 | 0.002264 | 149.4841 | 0.25838 | 8.4034 | 0.002476 | 0.004411 | 0.004277 | 4.9e-05 | 0.000999 | 0.33035 | 1.3284 | 2.5473 | 100.61 | 21.1 | 25.719 | 79.9701 | 81.3861 | 76.3008 | 0.006404 | 1.0103 | 0.024355 | 24.1068 | 0.024859 | 0.15371 | 0.001831 | 0.12893 | 0.12893 | 0.049236
2019-03-28 15:53 | 1553802799.2867 | 898 | 0.20914 | 0.000687 | 8.7e-05 | 99.7844 | 99.7099 | 121.2915 | 21.2612 | 0.002529 | 0.002329 | 149.4951 | 0.25708 | 8.4022 | 0.002506 | 0.003946 | 0.00388 | 4.9e-05 | 0.000999 | 0.32211 | 1.3091 | 2.5475 | 100.61 | 21.1 | 25.727 | 82.6659 | 82.1147 | 77.8512 | 0.006406 | 1.0058 | 0.024278 | 24.1373 | 0.006683 | 0.15294 | 0.001822 | 0.12787 | 0.12787 | 0.049396
2019-03-28 15:53 | 1553802800.2866 | 899 | 0.20916 | 0.000693 | 8.7e-05 | 99.7772 | 99.9483 | 121.084 | 21.2612 | 0.002436 | 0.00228 | 149.4823 | 0.25751 | 8.403 | 0.002471 | 0.003666 | 0.003529 | 5e-05 | 0.001003 | 0.32764 | 1.3268 | 2.5473 | 100.61 | 21.09 | 25.75 | 80.4349 | 81.2763 | 76.6703 | 0.006408 | 1.0092 | 0.024298 | 24.0761 | -0.015326 | 0.15319 | 0.001695 | 0.11867 | 0.11867 | 0.053314
2019-03-28 15:53 | 1553802801.2877 | 900 | 0.20915 | 0.000694 | 8.7e-05 | 99.785 | 99.6056 | 121.0157 | 21.2612 | 0.002511 | 0.002308 | 149.4881 | 0.25784 | 8.4024 | 0.002542 | 0.003719 | 0.003643 | 4.7e-05 | 0.001004 | 0.31022 | 1.3103 | 2.5475 | 100.61 | 21.09 | 25.764 | 82.2371 | 82.9648 | 77.3414 | 0.006411 | 1.0073 | 0.024268 | 24.0923 | 0.016097 | 0.15339 | 0.001759 | 0.12327 | 0.12327 | 0.05139
2019-03-28 15:53 | 1553802802.2878 | 901 | 0.20916 | 0.000696 | 8.7e-05 | 99.7909 | 99.5451 | 121.0084 | 21.2612 | 0.002429 | 0.002303 | 149.4679 | 0.25743 | 8.4008 | 0.002511 | 0.003463 | 0.003362 | 4.8e-05 | 0.001011 | 0.31708 | 1.3087 | 2.5474 | 100.61 | 21.09 | 25.777 | 80.258 | 82.2222 | 77.2174 | 0.006415 | 1.0076 | 0.024271 | 24.0874 | 0.022063 | 0.15314 | 0.001709 | 0.11959 | 0.11959 | 0.052887
2019-03-28 15:53 | 1553802803.2869 | 902 | 0.20913 | 0.000697 | 8.6e-05 | 99.8338 | 99.6735 | 121.0241 | 21.2612 | 0.002424 | 0.002272 | 149.4713 | 0.25866 | 8.4048 | 0.002495 | 0.004318 | 0.004182 | 4.7e-05 | 0.001013 | 0.33196 | 1.3161 | 2.5473 | 100.61 | 21.09 | 25.804 | 80.1253 | 81.8315 | 76.4762 | 0.006421 | 1.0098 | 0.024299 | 24.0634 | 0.014375 | 0.15388 | 0.001842 | 0.12947 | 0.12947 | 0.049084
2019-03-28 15:53 | 1553802804.2871 | 903 | 0.20915 | 0.000693 | 8.7e-05 | 99.8487 | 99.7214 | 120.9227 | 21.2612 | 0.002507 | 0.002337 | 149.4656 | 0.25689 | 8.4045 | 0.002499 | 0.003582 | 0.003507 | 4.8e-05 | 0.001011 | 0.34252 | 1.3306 | 2.5475 | 100.61 | 21.09 | 25.818 | 82.1241 | 81.94 | 78.0478 | 0.006425 | 1.0053 | 0.024234 | 24.1073 | 0.011412 | 0.15282 | 0.00173 | 0.12095 | 0.12095 | 0.052181
2019-03-28 15:53 | 1553802805.2878 | 904 | 0.20913 | 0.000696 | 8.7e-05 | 99.7519 | 99.6545 | 120.8779 | 21.2612 | 0.002449 | 0.002283 | 149.4828 | 0.25724 | 8.4021 | 0.002456 | 0.003908 | 0.003804 | 5.1e-05 | 0.001013 | 0.31332 | 1.3674 | 2.5475 | 100.61 | 21.09 | 25.842 | 80.7257 | 80.8941 | 76.7443 | 0.006431 | 1.009 | 0.024275 | 24.0581 | 0.008741 | 0.15303 | 0.001867 | 0.13117 | 0.13117 | 0.048183
2019-03-28 15:53 | 1553802806.2869 | 905 | 0.20913 | 0.000697 | 8.7e-05 | 99.8317 | 99.5417 | 121.0347 | 21.2612 | 0.002495 | 0.002296 | 149.5083 | 0.25751 | 8.4055 | 0.002503 | 0.003844 | 0.00377 | 5e-05 | 0.001017 | 0.33915 | 1.3417 | 2.5474 | 100.61 | 21.09 | 25.854 | 81.8323 | 82.0383 | 77.05 | 0.006434 | 1.0081 | 0.02428 | 24.0842 | 0.026025 | 0.15319 | 0.001883 | 0.13242 | 0.13242 | 0.04778
2019-03-28 15:53 | 1553802807.2869 | 906 | 0.20914 | 0.000695 | 8.6e-05 | 99.8231 | 99.5913 | 121.1196 | 21.2612 | 0.002448 | 0.002278 | 149.482 | 0.25732 | 8.4031 | 0.002572 | 0.004119 | 0.003991 | 5e-05 | 0.001015 | 0.31277 | 1.3217 | 2.5474 | 100.61 | 21.09 | 25.866 | 80.7048 | 83.6832 | 76.6211 | 0.006437 | 1.0094 | 0.024303 | 24.0779 | 0.020795 | 0.15308 | 0.001805 | 0.12683 | 0.12683 | 0.049848
2019-03-28 15:53 | 1553802808.286 | 907 | 0.20914 | 0.000693 | 8.6e-05 | 99.7802 | 99.6079 | 121.3656 | 21.2612 | 0.002468 | 0.002314 | 149.5023 | 0.2573 | 8.4029 | 0.002461 | 0.00397 | 0.003845 | 5e-05 | 0.001016 | 0.31924 | 1.3227 | 2.5474 | 100.61 | 21.09 | 25.885 | 81.2013 | 81.0338 | 77.5028 | 0.006442 | 1.0068 | 0.024297 | 24.1327 | 0.015456 | 0.15307 | 0.001812 | 0.12728 | 0.12728 | 0.049667
2019-03-28 15:53 | 1553802809.2861 | 908 | 0.20915 | 0.000692 | 8.6e-05 | 99.7972 | 99.8009 | 121.33 | 21.2612 | 0.002466 | 0.002297 | 149.4731 | 0.25829 | 8.4006 | 0.002439 | 0.004081 | 0.004016 | 4.6e-05 | 0.001017 | 0.32045 | 1.3253 | 2.5475 | 100.61 | 21.09 | 25.895 | 81.1499 | 80.4974 | 77.0788 | 0.006444 | 1.008 | 0.024308 | 24.1146 | -0.000333 | 0.15366 | 0.00178 | 0.12502 | 0.12502 | 0.050761
2019-03-28 15:53 | 1553802810.287 | 909 | 0.20915 | 0.000696 | 8.6e-05 | 99.8017 | 99.6164 | 121.3046 | 21.2612 | 0.002423 | 0.002294 | 149.4697 | 0.25666 | 8.402 | 0.002478 | 0.004037 | 0.003971 | 4.8e-05 | 0.00102 | 0.33339 | 1.3239 | 2.5475 | 100.61 | 21.09 | 25.904 | 80.1059 | 81.4275 | 77.0014 | 0.006446 | 1.0083 | 0.024309 | 24.1094 | 0.016616 | 0.15269 | 0.001751 | 0.12291 | 0.12291 | 0.051305
2019-03-28 15:54 | 1553802811.2861 | 910 | 0.20916 | 0.000692 | 8.6e-05 | 99.883 | 99.5227 | 121.3811 | 21.2612 | 0.002406 | 0.002274 | 149.4595 | 0.25722 | 8.4023 | 0.002468 | 0.003921 | 0.003751 | 4.9e-05 | 0.001022 | 0.32012 | 1.3066 | 2.5474 | 100.61 | 21.09 | 25.918 | 79.6966 | 81.2013 | 76.526 | 0.00645 | 1.0096 | 0.024333 | 24.1006 | 0.032323 | 0.15302 | 0.001663 | 0.11657 | 0.11657 | 0.054215
2019-03-28 15:54 | 1553802812.2871 | 911 | 0.20917 | 0.000697 | 8.6e-05 | 99.8639 | 99.7255 | 121.54 | 21.2612 | 0.002459 | 0.002284 | 149.4588 | 0.25711 | 8.4053 | 0.002507 | 0.004319 | 0.004214 | 5e-05 | 0.00102 | 0.32731 | 1.3461 | 2.5474 | 100.61 | 21.09 | 25.923 | 80.9792 | 82.1241 | 76.7677 | 0.006451 | 1.0089 | 0.02434 | 24.1247 | 0.012407 | 0.15296 | 0.00164 | 0.11495 | 0.11495 | 0.054957
2019-03-28 15:54 | 1553802813.2871 | 912 | 0.20914 | 0.000692 | 8.6e-05 | 99.9157 | 99.6555 | 121.7933 | 21.2612 | 0.002484 | 0.002302 | 149.4711 | 0.25694 | 8.4032 | 0.002501 | 0.003976 | 0.003916 | 4.7e-05 | 0.001022 | 0.32997 | 1.354 | 2.5474 | 100.61 | 21.09 | 25.932 | 81.5741 | 81.9827 | 77.1955 | 0.006453 | 1.0077 | 0.024351 | 24.1646 | 0.023327 | 0.15286 | 0.001807 | 0.12718 | 0.12718 | 0.049636
2019-03-28 15:54 | 1553802814.2861 | 913 | 0.20915 | 0.000686 | 8.6e-05 | 99.8785 | 99.6219 | 122.0621 | 21.2612 | 0.002468 | 0.0023 | 149.4703 | 0.25708 | 8.4005 | 0.002419 | 0.004451 | 0.004337 | 4.5e-05 | 0.001021 | 0.31686 | 1.3255 | 2.5474 | 100.61 | 21.09 | 25.935 | 81.1898 | 80.027 | 77.1629 | 0.006454 | 1.0078 | 0.024379 | 24.1901 | 0.023012 | 0.15294 | 0.001745 | 0.12281 | 0.12281 | 0.051431
2019-03-28 15:54 | 1553802815.2862 | 914 | 0.20916 | 0.000691 | 8.6e-05 | 99.8457 | 99.825 | 122.224 | 21.2612 | 0.002438 | 0.002326 | 149.4581 | 0.25666 | 8.4013 | 0.002465 | 0.003926 | 0.003823 | 4.8e-05 | 0.001023 | 0.35651 | 1.3283 | 2.5475 | 100.61 | 21.1 | 25.937 | 80.4637 | 81.1249 | 77.7888 | 0.006458 | 1.006 | 0.024373 | 24.2278 | 0.001856 | 0.15269 | 0.001698 | 0.11931 | 0.11931 | 0.052853
2019-03-28 15:54 | 1553802816.2866 | 915 | 0.20916 | 0.000693 | 8.6e-05 | 99.8642 | 99.7218 | 121.8719 | 21.2612 | 0.002409 | 0.002338 | 149.4445 | 0.25645 | 8.4007 | 0.00245 | 0.003712 | 0.003607 | 5e-05 | 0.001022 | 0.30541 | 1.3017 | 2.5473 | 100.61 | 21.1 | 25.939 | 79.7863 | 80.7523 | 78.0773 | 0.006459 | 1.0052 | 0.024328 | 24.2028 | 0.012769 | 0.15256 | 0.001724 | 0.12105 | 0.12105 | 0.052052
2019-03-28 15:54 | 1553802817.2867 | 916 | 0.20914 | 0.000687 | 8.6e-05 | 99.8928 | 99.6997 | 121.7241 | 21.2612 | 0.002484 | 0.002303 | 149.4749 | 0.25806 | 8.4054 | 0.002437 | 0.003886 | 0.003794 | 4.8e-05 | 0.001021 | 0.3115 | 1.3146 | 2.5474 | 100.61 | 21.1 | 25.94 | 81.5695 | 80.4536 | 77.2275 | 0.006459 | 1.0076 | 0.024343 | 24.1588 | 0.01731 | 0.15352 | 0.001801 | 0.12676 | 0.12676 | 0.050017
2019-03-28 15:54 | 1553802818.2869 | 917 | 0.20916 | 0.000691 | 8.5e-05 | 99.88 | 99.7054 | 121.7072 | 21.2612 | 0.002469 | 0.002333 | 149.4579 | 0.25851 | 8.405 | 0.002456 | 0.003978 | 0.003869 | 4.7e-05 | 0.001022 | 0.31277 | 1.3474 | 2.5474 | 100.61 | 21.1 | 25.938 | 81.2225 | 80.9153 | 77.9408 | 0.006459 | 1.0056 | 0.024316 | 24.1817 | 0.015645 | 0.15379 | 0.001681 | 0.11785 | 0.11785 | 0.053896
2019-03-28 15:54 | 1553802819.2863 | 918 | 0.20913 | 0.000681 | 8.5e-05 | 99.9271 | 99.7516 | 121.5818 | 21.2612 | 0.002486 | 0.00229 | 149.4604 | 0.25726 | 8.4009 | 0.002362 | 0.004655 | 0.004511 | 4.7e-05 | 0.001025 | 0.30917 | 1.316 | 2.5474 | 100.61 | 21.1 | 25.937 | 81.6373 | 78.6511 | 76.9056 | 0.006458 | 1.0085 | 0.02434 | 24.1336 | 0.015715 | 0.15305 | 0.001872 | 0.132 | 0.132 | 0.047884
2019-03-28 15:54 | 1553802820.2878 | 919 | 0.20913 | 0.000685 | 8.6e-05 | 99.9455 | 99.8298 | 121.3898 | 21.2612 | 0.002488 | 0.002278 | 149.4702 | 0.25783 | 8.4047 | 0.002473 | 0.004458 | 0.004384 | 4.8e-05 | 0.001027 | 0.32659 | 1.304 | 2.5476 | 100.61 | 21.1 | 25.936 | 81.6748 | 81.3035 | 76.6165 | 0.006458 | 1.0094 | 0.024331 | 24.1046 | 0.010364 | 0.15338 | 0.001859 | 0.13095 | 0.13095 | 0.048375
2019-03-28 15:54 | 1553802821.2864 | 920 | 0.20913 | 0.000687 | 8.5e-05 | 99.9348 | 99.7014 | 121.4203 | 21.2612 | 0.002402 | 0.002283 | 149.4639 | 0.2573 | 8.4019 | 0.002448 | 0.00345 | 0.003327 | 4.7e-05 | 0.001029 | 0.30071 | 1.2991 | 2.5474 | 100.61 | 21.1 | 25.932 | 79.6053 | 80.7125 | 76.7513 | 0.006457 | 1.009 | 0.024329 | 24.1123 | 0.020917 | 0.15307 | 0.00188 | 0.13257 | 0.13257 | 0.047686
2019-03-28 15:54 | 1553802822.2864 | 921 | 0.20913 | 0.000681 | 8.5e-05 | 99.9116 | 99.7503 | 121.1725 | 21.2612 | 0.002498 | 0.002276 | 149.4628 | 0.25777 | 8.4021 | 0.002425 | 0.00388 | 0.003853 | 4.8e-05 | 0.001031 | 0.32516 | 1.3097 | 2.5473 | 100.61 | 21.1 | 25.929 | 81.9174 | 80.1621 | 76.5728 | 0.006456 | 1.0095 | 0.02431 | 24.0815 | 0.01445 | 0.15335 | 0.001898 | 0.13375 | 0.13375 | 0.047353
2019-03-28 15:54 | 1553802823.2858 | 922 | 0.2092 | 0.000683 | 8.4e-05 | 99.9339 | 99.8777 | 120.7253 | 21.2612 | 0.002515 | 0.002288 | 149.4386 | 0.25826 | 8.403 | 0.002415 | 0.004049 | 0.003867 | 4.7e-05 | 0.001032 | 0.33069 | 1.3483 | 2.5473 | 100.61 | 21.1 | 25.918 | 82.3128 | 79.9148 | 76.8777 | 0.006454 | 1.0086 | 0.024255 | 24.0475 | 0.005034 | 0.15364 | 0.00149 | 0.10369 | 0.10369 | 0.061193
2019-03-28 15:54 | 1553802824.2855 | 923 | 0.20918 | 0.00068 | 8.4e-05 | 99.9312 | 99.737 | 120.3434 | 21.2612 | 0.002377 | 0.002343 | 149.4258 | 0.25782 | 8.4026 | 0.002466 | 0.003855 | 0.003792 | 4.8e-05 | 0.001032 | 0.32676 | 1.3593 | 2.5475 | 100.61 | 21.1 | 25.91 | 79.005 | 81.1383 | 78.1833 | 0.006452 | 1.0049 | 0.024171 | 24.0542 | 0.017398 | 0.15338 | 0.001588 | 0.11047 | 0.11047 | 0.057341
2019-03-28 15:54 | 1553802825.2855 | 924 | 0.2092 | 0.000684 | 8.4e-05 | 99.9122 | 99.6582 | 120.0728 | 21.2612 | 0.002458 | 0.002263 | 149.4236 | 0.25817 | 8.3998 | 0.002351 | 0.003528 | 0.003353 | 4.6e-05 | 0.001038 | 0.33987 | 1.3479 | 2.5475 | 100.61 | 21.1 | 25.901 | 80.9425 | 78.3774 | 76.2679 | 0.006449 | 1.0104 | 0.02421 | 23.9615 | 0.022763 | 0.15359 | 0.001443 | 0.10013 | 0.10013 | 0.063353
2019-03-28 15:54 | 1553802826.2855 | 925 | 0.20916 | 0.00068 | 8.4e-05 | 99.8487 | 99.6878 | 120.3202 | 21.2612 | 0.002461 | 0.002277 | 149.4065 | 0.25792 | 8.4019 | 0.00241 | 0.004076 | 0.003985 | 4.9e-05 | 0.001035 | 0.31885 | 1.3759 | 2.5474 | 100.61 | 21.1 | 25.881 | 81.029 | 79.8033 | 76.6103 | 0.006444 | 1.0094 | 0.024223 | 23.9979 | 0.014429 | 0.15344 | 0.001713 | 0.11984 | 0.11984 | 0.052879
2019-03-28 15:54 | 1553802827.2856 | 926 | 0.20918 | 0.000677 | 8.4e-05 | 99.7886 | 99.7078 | 120.7809 | 21.2612 | 0.002432 | 0.002284 | 149.3705 | 0.2582 | 8.4041 | 0.002421 | 0.004752 | 0.004619 | 4.5e-05 | 0.001034 | 0.31564 | 1.3405 | 2.5474 | 100.61 | 21.1 | 25.87 | 80.3265 | 80.0576 | 76.7825 | 0.006442 | 1.0089 | 0.024264 | 24.0498 | 0.007239 | 0.1536 | 0.001578 | 0.11017 | 0.11017 | 0.057583
2019-03-28 15:54 | 1553802828.2866 | 927 | 0.20917 | 0.000682 | 8.4e-05 | 99.8472 | 99.8563 | 121.1837 | 21.2612 | 0.002466 | 0.002274 | 149.3412 | 0.25813 | 8.4011 | 0.002398 | 0.00411 | 0.004023 | 4.8e-05 | 0.001033 | 0.32549 | 1.3503 | 2.5474 | 100.61 | 21.1 | 25.859 | 81.1523 | 79.5008 | 76.5392 | 0.006439 | 1.0096 | 0.024313 | 24.0815 | -0.000811 | 0.15356 | 0.001637 | 0.11467 | 0.11467 | 0.055306
2019-03-28 15:54 | 1553802829.2857 | 928 | 0.20919 | 0.000679 | 8.4e-05 | 99.874 | 99.733 | 121.3363 | 21.2612 | 0.002448 | 0.002259 | 149.3628 | 0.25755 | 8.4022 | 0.002415 | 0.003731 | 0.003691 | 4.6e-05 | 0.001029 | 0.30646 | 1.3413 | 2.5474 | 100.61 | 21.1 | 25.832 | 80.7024 | 79.9211 | 76.1699 | 0.006432 | 1.0107 | 0.024341 | 24.0839 | 0.012643 | 0.15322 | 0.001528 | 0.10684 | 0.10684 | 0.059225
2019-03-28 15:54 | 1553802830.2871 | 929 | 0.20918 | 0.000677 | 8.5e-05 | 99.8145 | 99.6161 | 121.5473 | 21.2612 | 0.002435 | 0.00231 | 149.3341 | 0.25769 | 8.4048 | 0.002448 | 0.003602 | 0.003427 | 4.7e-05 | 0.001029 | 0.30558 | 1.3282 | 2.5473 | 100.61 | 21.1 | 25.819 | 80.3967 | 80.7117 | 77.4022 | 0.006429 | 1.0071 | 0.024319 | 24.1473 | 0.017794 | 0.1533 | 0.001574 | 0.11011 | 0.11011 | 0.057499
2019-03-28 15:54 | 1553802831.2871 | 930 | 0.20916 | 0.000677 | 8.4e-05 | 99.8067 | 99.6446 | 121.2135 | 21.2612 | 0.002496 | 0.002303 | 149.3258 | 0.25836 | 8.4017 | 0.00248 | 0.004528 | 0.00439 | 4.5e-05 | 0.001027 | 0.30215 | 1.3232 | 2.5474 | 100.61 | 21.1 | 25.79 | 81.876 | 81.4893 | 77.2306 | 0.006422 | 1.0076 | 0.024292 | 24.1082 | 0.014538 | 0.1537 | 0.001712 | 0.12006 | 0.12006 | 0.05287
2019-03-28 15:54 | 1553802832.2858 | 931 | 0.20916 | 0.000678 | 8.4e-05 | 99.6882 | 99.4201 | 121.2295 | 21.2612 | 0.002457 | 0.002302 | 149.3274 | 0.25808 | 8.4049 | 0.002511 | 0.003938 | 0.003899 | 4.7e-05 | 0.001027 | 0.34252 | 1.3266 | 2.5473 | 100.61 | 21.1 | 25.774 | 80.9254 | 82.2222 | 77.198 | 0.006418 | 1.0077 | 0.024294 | 24.1087 | 0.024088 | 0.15353 | 0.001747 | 0.1227 | 0.1227 | 0.051675
2019-03-28 15:54 | 1553802833.2872 | 932 | 0.20916 | 0.000682 | 8.4e-05 | 99.7355 | 99.4775 | 121.5407 | 21.2612 | 0.002519 | 0.002313 | 149.3215 | 0.25917 | 8.4017 | 0.002425 | 0.003668 | 0.003563 | 4.8e-05 | 0.001026 | 0.31349 | 1.3308 | 2.5473 | 100.61 | 21.1 | 25.746 | 82.407 | 80.1628 | 77.4685 | 0.006411 | 1.0069 | 0.024316 | 24.1489 | 0.023172 | 0.15418 | 0.001735 | 0.12197 | 0.12197 | 0.052207
2019-03-28 15:54 | 1553802834.2868 | 933 | 0.20914 | 0.000679 | 8.3e-05 | 99.6334 | 99.433 | 121.7353 | 21.2612 | 0.002504 | 0.002218 | 149.3233 | 0.25862 | 8.4035 | 0.002478 | 0.004274 | 0.004172 | 4.9e-05 | 0.001032 | 0.31581 | 1.3283 | 2.5474 | 100.61 | 21.1 | 25.734 | 82.0609 | 81.433 | 75.1854 | 0.006408 | 1.0135 | 0.024415 | 24.0894 | 0.018008 | 0.15385 | 0.001806 | 0.12763 | 0.12763 | 0.049785
2019-03-28 15:54 | 1553802835.2868 | 934 | 0.20917 | 0.000678 | 8.4e-05 | 99.6084 | 99.5753 | 121.97 | 21.2612 | 0.002485 | 0.002245 | 149.3076 | 0.25767 | 8.4028 | 0.002349 | 0.003559 | 0.003498 | 4.8e-05 | 0.00103 | 0.29574 | 1.3113 | 2.5472 | 100.61 | 21.1 | 25.724 | 81.6085 | 78.3431 | 75.8417 | 0.006405 | 1.0116 | 0.024416 | 24.1354 | 0.002966 | 0.15329 | 0.001676 | 0.11812 | 0.11812 | 0.053598
2019-03-28 15:54 | 1553802836.2859 | 935 | 0.20918 | 0.000677 | 8.3e-05 | 99.6051 | 99.5152 | 121.9468 | 21.2612 | 0.002419 | 0.002307 | 149.3195 | 0.25855 | 8.4036 | 0.002404 | 0.004327 | 0.004201 | 4.6e-05 | 0.001028 | 0.34827 | 1.3184 | 2.5474 | 100.61 | 21.1 | 25.709 | 80.0194 | 79.6459 | 77.3306 | 0.006402 | 1.0073 | 0.024361 | 24.1845 | 0.008075 | 0.15381 | 0.001625 | 0.11415 | 0.11415 | 0.055649
2019-03-28 15:54 | 1553802837.2849 | 936 | 0.20916 | 0.000678 | 8.3e-05 | 99.6343 | 99.469 | 121.6466 | 21.2612 | 0.002441 | 0.002309 | 149.3057 | 0.25757 | 8.4017 | 0.00236 | 0.003537 | 0.003463 | 4.9e-05 | 0.001026 | 0.32587 | 1.3075 | 2.5474 | 100.61 | 21.1 | 25.704 | 80.5502 | 78.5879 | 77.381 | 0.0064 | 1.0072 | 0.02433 | 24.1564 | 0.014849 | 0.15323 | 0.001711 | 0.12029 | 0.12029 | 0.052611
2019-03-28 15:54 | 1553802838.2849 | 937 | 0.20915 | 0.000678 | 8.2e-05 | 99.6682 | 99.2747 | 121.3919 | 21.2612 | 0.002429 | 0.002304 | 149.323 | 0.25791 | 8.4013 | 0.002401 | 0.004223 | 0.004082 | 4.9e-05 | 0.001024 | 0.34252 | 1.2947 | 2.5474 | 100.61 | 21.1 | 25.699 | 80.2438 | 79.5928 | 77.2604 | 0.006399 | 1.0075 | 0.024308 | 24.127 | 0.035388 | 0.15343 | 0.0018 | 0.12671 | 0.12671 | 0.050009
2019-03-28 15:54 | 1553802839.2849 | 938 | 0.20915 | 0.000674 | 8.2e-05 | 99.569 | 99.3739 | 121.6041 | 21.2612 | 0.00246 | 0.002257 | 149.3041 | 0.25755 | 8.4037 | 0.002373 | 0.00417 | 0.004138 | 4.8e-05 | 0.001025 | 0.29247 | 1.3148 | 2.5474 | 100.61 | 21.11 | 25.695 | 81.0112 | 78.8988 | 76.1324 | 0.006402 | 1.0108 | 0.024369 | 24.1092 | 0.017549 | 0.15322 | 0.001805 | 0.12744 | 0.12744 | 0.049654
2019-03-28 15:54 | 1553802840.285 | 939 | 0.20917 | 0.000676 | 8.2e-05 | 99.4028 | 99.2978 | 121.6374 | 21.2612 | 0.002441 | 0.00228 | 149.3123 | 0.25832 | 8.4038 | 0.002342 | 0.003714 | 0.003585 | 5e-05 | 0.00103 | 0.32444 | 1.2909 | 2.5474 | 100.61 | 21.11 | 25.694 | 80.5346 | 78.1653 | 76.6764 | 0.006402 | 1.0092 | 0.024353 | 24.1312 | 0.009459 | 0.15368 | 0.001655 | 0.11634 | 0.11634 | 0.054553
2019-03-28 15:54 | 1553802841.285 | 940 | 0.20916 | 0.000678 | 8.1e-05 | 99.5306 | 99.3871 | 121.9919 | 21.2612 | 0.002444 | 0.002372 | 149.3079 | 0.25838 | 8.4004 | 0.00237 | 0.003721 | 0.003646 | 4.8e-05 | 0.001026 | 0.33229 | 1.29 | 2.5473 | 100.61 | 21.11 | 25.695 | 80.6274 | 78.8342 | 78.8943 | 0.006402 | 1.0028 | 0.024312 | 24.2428 | 0.012905 | 0.15371 | 0.001715 | 0.12058 | 0.12058 | 0.052644
2019-03-28 15:54 | 1553802842.286 | 941 | 0.20919 | 0.000675 | 8.1e-05 | 99.4142 | 99.2451 | 122.0855 | 21.2612 | 0.002432 | 0.002336 | 149.3185 | 0.25892 | 8.4008 | 0.002328 | 0.003466 | 0.003371 | 4.8e-05 | 0.001029 | 0.29999 | 1.2801 | 2.5474 | 100.61 | 21.11 | 25.696 | 80.3296 | 77.8356 | 78.0266 | 0.006402 | 1.0053 | 0.024351 | 24.2222 | 0.015222 | 0.15403 | 0.00156 | 0.1094 | 0.1094 | 0.05815
2019-03-28 15:54 | 1553802843.2851 | 942 | 0.2092 | 0.000672 | 8.2e-05 | 99.497 | 99.1948 | 121.9999 | 21.2612 | 0.002498 | 0.002285 | 149.3374 | 0.25761 | 8.4018 | 0.002337 | 0.004143 | 0.003981 | 5e-05 | 0.001032 | 0.32549 | 1.3285 | 2.5475 | 100.61 | 21.11 | 25.696 | 81.925 | 78.0384 | 76.7965 | 0.006402 | 1.0089 | 0.024385 | 24.1713 | 0.027201 | 0.15325 | 0.001481 | 0.10377 | 0.10377 | 0.060991
2019-03-28 15:54 | 1553802844.2853 | 943 | 0.20926 | 0.000669 | 8.1e-05 | 99.4034 | 99.1782 | 122.0116 | 21.2612 | 0.002418 | 0.00226 | 149.3363 | 0.25733 | 8.4048 | 0.002388 | 0.003925 | 0.0038 | 5e-05 | 0.001033 | 0.32228 | 1.3019 | 2.5475 | 100.61 | 21.11 | 25.697 | 79.9864 | 79.2795 | 76.1846 | 0.006403 | 1.0106 | 0.024408 | 24.1513 | 0.020291 | 0.15309 | 0.001139 | 0.078731 | 0.078731 | 0.080306
2019-03-28 15:54 | 1553802845.2854 | 944 | 0.20916 | 0.000672 | 8.1e-05 | 99.4225 | 99.0012 | 122.0632 | 21.2612 | 0.002492 | 0.002249 | 149.3537 | 0.25792 | 8.4025 | 0.002417 | 0.003786 | 0.003674 | 4.8e-05 | 0.00103 | 0.33035 | 1.2797 | 2.5474 | 100.61 | 21.11 | 25.696 | 81.7714 | 79.9773 | 75.9407 | 0.006402 | 1.0113 | 0.024422 | 24.148 | 0.037983 | 0.15344 | 0.001723 | 0.12173 | 0.12173 | 0.052055
2019-03-28 15:54 | 1553802846.2875 | 945 | 0.20912 | 0.000669 | 8e-05 | 99.2661 | 99.0837 | 121.8104 | 21.2612 | 0.002443 | 0.002275 | 149.3439 | 0.25823 | 8.4044 | 0.002354 | 0.003921 | 0.003814 | 5e-05 | 0.001031 | 0.30055 | 1.3163 | 2.5474 | 100.61 | 21.11 | 25.696 | 80.6017 | 78.4514 | 76.5666 | 0.006402 | 1.0095 | 0.024374 | 24.1446 | 0.016448 | 0.15362 | 0.001954 | 0.1385 | 0.1385 | 0.045808
2019-03-28 15:54 | 1553802847.2853 | 946 | 0.20917 | 0.000666 | 8e-05 | 99.248 | 99.1184 | 122.1583 | 21.2612 | 0.002473 | 0.002271 | 149.3457 | 0.25844 | 8.4039 | 0.002382 | 0.003878 | 0.003793 | 5.1e-05 | 0.001029 | 0.33251 | 1.3138 | 2.5473 | 100.61 | 21.11 | 25.696 | 81.3246 | 79.1197 | 76.4505 | 0.006402 | 1.0099 | 0.024413 | 24.175 | 0.011683 | 0.15375 | 0.001657 | 0.11693 | 0.11693 | 0.054302
2019-03-28 15:54 | 1553802848.2853 | 947 | 0.20921 | 0.000671 | 8e-05 | 99.1988 | 99.0199 | 122.225 | 21.2612 | 0.002464 | 0.002308 | 149.3338 | 0.25734 | 8.4039 | 0.002323 | 0.003947 | 0.003861 | 5.2e-05 | 0.001025 | 0.33069 | 1.308 | 2.5475 | 100.61 | 21.11 | 25.695 | 81.0876 | 77.7039 | 77.3476 | 0.006402 | 1.0073 | 0.024389 | 24.2126 | 0.016148 | 0.15309 | 0.001409 | 0.098568 | 0.098568 | 0.064146
2019-03-28 15:54 | 1553802849.2853 | 948 | 0.20916 | 0.000666 | 7.9e-05 | 99.0949 | 98.83 | 122.3454 | 21.2612 | 0.002461 | 0.002278 | 149.3161 | 0.25815 | 8.4031 | 0.002379 | 0.00375 | 0.00363 | 5.2e-05 | 0.001023 | 0.33196 | 1.2893 | 2.5473 | 100.61 | 21.11 | 25.695 | 81.0251 | 79.0627 | 76.6367 | 0.006402 | 1.0093 | 0.024425 | 24.2 | 0.023941 | 0.15357 | 0.001718 | 0.12148 | 0.12148 | 0.052209
2019-03-28 15:54 | 1553802850.2867 | 949 | 0.20916 | 0.00067 | 7.9e-05 | 99.0412 | 98.901 | 122.339 | 21.2612 | 0.002483 | 0.002287 | 149.3085 | 0.2581 | 8.4014 | 0.002478 | 0.002796 | 0.002724 | 5e-05 | 0.001021 | 0.3183 | 1.3009 | 2.5474 | 100.61 | 21.11 | 25.694 | 81.5546 | 81.4292 | 76.8457 | 0.006402 | 1.0087 | 0.024417 | 24.2066 | 0.012677 | 0.15354 | 0.001707 | 0.12064 | 0.12064 | 0.052563
2019-03-28 15:54 | 1553802851.2853 | 950 | 0.20918 | 0.000672 | 7.9e-05 | 98.9534 | 98.602 | 122.4461 | 21.2612 | 0.002456 | 0.002289 | 149.2996 | 0.25823 | 8.4047 | 0.002478 | 0.003629 | 0.003436 | 5e-05 | 0.001021 | 0.33052 | 1.3058 | 2.5474 | 100.61 | 21.11 | 25.693 | 80.905 | 81.4345 | 76.8962 | 0.006402 | 1.0086 | 0.024426 | 24.2189 | 0.031817 | 0.15362 | 0.001585 | 0.11169 | 0.11169 | 0.056805
2019-03-28 15:54 | 1553802852.286 | 951 | 0.20916 | 0.00067 | 8e-05 | 98.9677 | 98.7005 | 122.5158 | 21.2612 | 0.002469 | 0.002256 | 149.2893 | 0.25738 | 8.4044 | 0.002402 | 0.0041 | 0.00402 | 5.2e-05 | 0.001022 | 0.3361 | 1.3278 | 2.5474 | 100.61 | 21.11 | 25.691 | 81.2054 | 79.6185 | 76.0887 | 0.006401 | 1.0109 | 0.024462 | 24.1978 | 0.024176 | 0.15312 | 0.001713 | 0.12128 | 0.12128 | 0.052139
2019-03-28 15:54 | 1553802853.2844 | 952 | 0.20917 | 0.00067 | 8e-05 | 98.7866 | 98.4637 | 122.5444 | 21.2612 | 0.00249 | 0.002281 | 149.2504 | 0.25759 | 8.4035 | 0.002332 | 0.003691 | 0.003595 | 5.3e-05 | 0.001021 | 0.29983 | 1.2982 | 2.5474 | 100.61 | 21.11 | 25.691 | 81.7325 | 77.9276 | 76.7006 | 0.006401 | 1.0091 | 0.024443 | 24.2219 | 0.029278 | 0.15324 | 0.001662 | 0.11742 | 0.11742 | 0.053898
2019-03-28 15:54 | 1553802854.2855 | 953 | 0.20917 | 0.000671 | 7.9e-05 | 98.791 | 98.5742 | 122.4778 | 21.2612 | 0.002441 | 0.002296 | 149.2566 | 0.25777 | 8.4035 | 0.002451 | 0.003703 | 0.003618 | 5.3e-05 | 0.001013 | 0.30182 | 1.3566 | 2.5475 | 100.61 | 21.11 | 25.692 | 80.5519 | 80.7843 | 77.0632 | 0.006401 | 1.0081 | 0.024424 | 24.2278 | 0.019659 | 0.15335 | 0.001643 | 0.11598 | 0.11598 | 0.054605
2019-03-28 15:54 | 1553802855.2857 | 954 | 0.20919 | 0.00067 | 7.9e-05 | 98.6934 | 98.5813 | 122.3724 | 21.2612 | 0.002449 | 0.002245 | 149.2407 | 0.25754 | 8.4046 | 0.002469 | 0.003884 | 0.003838 | 5.1e-05 | 0.00101 | 0.33157 | 1.3407 | 2.5474 | 100.61 | 21.11 | 25.692 | 80.7288 | 81.2107 | 75.8463 | 0.006401 | 1.0116 | 0.024456 | 24.1753 | 0.010162 | 0.15321 | 0.001536 | 0.10829 | 0.10829 | 0.058433
2019-03-28 15:54 | 1553802856.2845 | 955 | 0.20922 | 0.000667 | 7.8e-05 | 98.6385 | 98.2052 | 122.7897 | 21.2612 | 0.00252 | 0.002269 | 149.235 | 0.25847 | 8.4005 | 0.002436 | 0.003417 | 0.003354 | 5.2e-05 | 0.001012 | 0.3413 | 1.2922 | 2.5473 | 100.61 | 21.11 | 25.692 | 82.4532 | 80.4154 | 76.4106 | 0.006401 | 1.01 | 0.024478 | 24.236 | 0.039381 | 0.15377 | 0.001359 | 0.095386 | 0.095386 | 0.066577
2019-03-28 15:54 | 1553802857.2849 | 956 | 0.2092 | 0.000667 | 7.7e-05 | 98.5206 | 98.2008 | 122.7892 | 21.2612 | 0.002471 | 0.002288 | 149.2324 | 0.25791 | 8.4034 | 0.002408 | 0.003434 | 0.003353 | 5.3e-05 | 0.001009 | 0.32316 | 1.3216 | 2.5474 | 100.61 | 21.11 | 25.694 | 81.2693 | 79.755 | 76.8635 | 0.006402 | 1.0087 | 0.024462 | 24.2517 | 0.029081 | 0.15343 | 0.001476 | 0.10393 | 0.10393 | 0.06097
2019-03-28 15:54 | 1553802858.2849 | 957 | 0.20918 | 0.000666 | 7.7e-05 | 98.5694 | 98.3245 | 122.2761 | 21.2612 | 0.002435 | 0.00228 | 149.2439 | 0.25743 | 8.4031 | 0.002339 | 0.003823 | 0.003702 | 5.3e-05 | 0.001011 | 0.32925 | 1.2862 | 2.5475 | 100.61 | 21.11 | 25.693 | 80.3974 | 78.1045 | 76.6742 | 0.006402 | 1.0092 | 0.024417 | 24.1944 | 0.022259 | 0.15314 | 0.001632 | 0.11528 | 0.11528 | 0.054865
2019-03-28 15:54 | 1553802859.2846 | 958 | 0.20918 | 0.000662 | 7.7e-05 | 98.4038 | 97.7911 | 121.924 | 21.2612 | 0.002443 | 0.002286 | 149.261 | 0.25622 | 8.4025 | 0.002326 | 0.003513 | 0.003382 | 4.9e-05 | 0.001013 | 0.3115 | 1.2939 | 2.5475 | 100.61 | 21.11 | 25.693 | 80.5971 | 77.7873 | 76.8301 | 0.006402 | 1.0088 | 0.024377 | 24.1649 | 0.055871 | 0.15243 | 0.001621 | 0.11429 | 0.11429 | 0.055081
2019-03-28 15:54 | 1553802860.285 | 959 | 0.2092 | 0.000662 | 7.7e-05 | 98.3803 | 97.838 | 121.8992 | 21.2612 | 0.00242 | 0.002267 | 149.2695 | 0.25698 | 8.4006 | 0.00241 | 0.004446 | 0.004309 | 4.7e-05 | 0.001008 | 0.32516 | 1.2447 | 2.5473 | 100.61 | 21.11 | 25.692 | 80.0364 | 79.7925 | 76.3592 | 0.006401 | 1.0101 | 0.024391 | 24.1462 | 0.049445 | 0.15288 | 0.001499 | 0.10541 | 0.10541 | 0.059896
2019-03-28 15:54 | 1553802861.285 | 960 | 0.2092 | 0.000663 | 7.7e-05 | 98.3529 | 98.0197 | 122.0111 | 21.2612 | 0.00247 | 0.002278 | 149.2661 | 0.25704 | 8.4018 | 0.002372 | 0.00323 | 0.003097 | 5.1e-05 | 0.001012 | 0.34418 | 1.2906 | 2.5473 | 100.61 | 21.11 | 25.691 | 81.2412 | 78.8811 | 76.6336 | 0.006401 | 1.0093 | 0.024392 | 24.1668 | 0.030352 | 0.15291 | 0.001507 | 0.106 | 0.106 | 0.059578
2019-03-28 15:54 | 1553802862.2848 | 961 | 0.20918 | 0.00066 | 7.7e-05 | 98.2469 | 97.7639 | 121.9414 | 21.2612 | 0.002425 | 0.002317 | 149.2404 | 0.25783 | 8.4029 | 0.002366 | 0.003748 | 0.003595 | 5.3e-05 | 0.001015 | 0.33754 | 1.3212 | 2.5474 | 100.61 | 21.11 | 25.69 | 80.1674 | 78.7525 | 77.5581 | 0.006401 | 1.0067 | 0.024353 | 24.1918 | 0.044078 | 0.15338 | 0.001628 | 0.1147 | 0.1147 | 0.055229
2019-03-28 15:54 | 1553802863.2848 | 962 | 0.20925 | 0.000658 | 7.7e-05 | 98.1748 | 97.8652 | 121.8855 | 21.2612 | 0.002428 | 0.002262 | 149.2632 | 0.25683 | 8.4029 | 0.002444 | 0.004439 | 0.004373 | 5.2e-05 | 0.001013 | 0.33682 | 1.2955 | 2.5472 | 100.61 | 21.11 | 25.689 | 80.2414 | 80.6057 | 76.2494 | 0.006401 | 1.0104 | 0.024393 | 24.1411 | 0.028255 | 0.15279 | 0.00119 | 0.082692 | 0.082692 | 0.076309
2019-03-28 15:54 | 1553802864.2851 | 963 | 0.20921 | 0.000659 | 7.7e-05 | 98.128 | 97.625 | 121.7787 | 21.2612 | 0.0025 | 0.002249 | 149.25 | 0.25774 | 8.4026 | 0.002441 | 0.003898 | 0.003809 | 5e-05 | 0.001016 | 0.32045 | 1.3044 | 2.5474 | 100.61 | 21.11 | 25.688 | 81.9681 | 80.5534 | 75.9335 | 0.0064 | 1.0113 | 0.024393 | 24.1196 | 0.045973 | 0.15333 | 0.001479 | 0.10393 | 0.10393 | 0.06093
2019-03-28 15:54 | 1553802865.2848 | 964 | 0.20919 | 0.000659 | 7.7e-05 | 98.1289 | 97.6226 | 121.9905 | 21.2612 | 0.002437 | 0.002317 | 149.2675 | 0.25676 | 8.4007 | 0.002393 | 0.003879 | 0.00379 | 4.7e-05 | 0.001018 | 0.35402 | 1.3273 | 2.5474 | 100.61 | 21.11 | 25.685 | 80.4498 | 79.3823 | 77.5752 | 0.0064 | 1.0066 | 0.024357 | 24.1972 | 0.046273 | 0.15275 | 0.001552 | 0.10915 | 0.10915 | 0.057794
2019-03-28 15:54 | 1553802866.2849 | 965 | 0.20921 | 0.000653 | 7.6e-05 | 98.0622 | 97.6165 | 121.9583 | 21.2612 | 0.002483 | 0.002314 | 149.2507 | 0.25861 | 8.4033 | 0.002441 | 0.003882 | 0.003768 | 5.6e-05 | 0.001017 | 0.28479 | 1.3284 | 2.5474 | 100.61 | 21.11 | 25.685 | 81.5626 | 80.5377 | 77.4966 | 0.0064 | 1.0068 | 0.024357 | 24.1913 | 0.040749 | 0.15385 | 0.001484 | 0.10418 | 0.10418 | 0.06099
2019-03-28 15:54 | 1553802867.2849 | 966 | 0.20915 | 0.000653 | 7.6e-05 | 98.0795 | 97.6678 | 121.8607 | 21.2612 | 0.002415 | 0.00234 | 149.2525 | 0.25735 | 8.4048 | 0.00236 | 0.004294 | 0.004161 | 5.5e-05 | 0.001025 | 0.30486 | 1.3043 | 2.5474 | 100.61 | 21.11 | 25.683 | 79.9103 | 78.6042 | 78.1256 | 0.006399 | 1.005 | 0.024325 | 24.2033 | 0.037625 | 0.1531 | 0.001819 | 0.12858 | 0.12858 | 0.049176
2019-03-28 15:54 | 1553802868.2849 | 967 | 0.20918 | 0.000652 | 7.5e-05 | 98.0283 | 97.7829 | 122.4841 | 21.2612 | 0.002514 | 0.002286 | 149.2437 | 0.2575 | 8.4057 | 0.002416 | 0.0041 | 0.004118 | 5.2e-05 | 0.001023 | 0.30414 | 1.3396 | 2.5475 | 100.61 | 21.11 | 25.682 | 82.2964 | 79.9413 | 76.827 | 0.006399 | 1.0088 | 0.024433 | 24.2203 | 0.022411 | 0.15318 | 0.00164 | 0.11607 | 0.11607 | 0.054505
2019-03-28 15:54 | 1553802869.2843 | 968 | 0.20922 | 0.000655 | 7.5e-05 | 97.9571 | 97.4517 | 122.669 | 21.2612 | 0.002457 | 0.002271 | 149.2086 | 0.25691 | 8.4012 | 0.002403 | 0.003721 | 0.00363 | 5.1e-05 | 0.00102 | 0.31492 | 1.3215 | 2.5475 | 100.61 | 21.11 | 25.682 | 80.9223 | 79.6397 | 76.4637 | 0.006399 | 1.0098 | 0.024464 | 24.226 | 0.046264 | 0.15284 | 0.001419 | 0.099917 | 0.099917 | 0.063174
2019-03-28 15:54 | 1553802870.2851 | 969 | 0.20918 | 0.000654 | 7.6e-05 | 97.9133 | 97.3974 | 122.9777 | 21.2612 | 0.002392 | 0.002283 | 149.2198 | 0.25724 | 8.4037 | 0.002386 | 0.003708 | 0.003602 | 5.6e-05 | 0.001024 | 0.33212 | 1.2854 | 2.5475 | 100.61 | 21.11 | 25.681 | 79.3737 | 79.2249 | 76.7537 | 0.006399 | 1.009 | 0.024484 | 24.2665 | 0.047256 | 0.15303 | 0.001637 | 0.11602 | 0.11602 | 0.054476
2019-03-28 15:54 | 1553802871.284 | 970 | 0.2092 | 0.000648 | 7.6e-05 | 97.8698 | 97.4588 | 122.6282 | 21.2612 | 0.002471 | 0.002308 | 149.2266 | 0.25788 | 8.4034 | 0.002412 | 0.003893 | 0.003769 | 5.5e-05 | 0.001023 | 0.323 | 1.3074 | 2.5474 | 100.61 | 21.11 | 25.679 | 81.2715 | 79.8415 | 77.339 | 0.006398 | 1.0073 | 0.024429 | 24.2522 | 0.037637 | 0.15341 | 0.00151 | 0.10641 | 0.10641 | 0.059542
2019-03-28 15:55 | 1553802872.2844 | 971 | 0.20919 | 0.000648 | 7.5e-05 | 97.8236 | 97.2693 | 122.8303 | 21.2612 | 0.002448 | 0.002221 | 149.2575 | 0.25765 | 8.4048 | 0.002465 | 0.003511 | 0.003524 | 5.6e-05 | 0.001029 | 0.31868 | 1.3369 | 2.5473 | 100.61 | 21.11 | 25.678 | 80.7069 | 81.1234 | 75.2562 | 0.006398 | 1.0133 | 0.024522 | 24.2 | 0.050832 | 0.15327 | 0.001587 | 0.11254 | 0.11254 | 0.056246
2019-03-28 15:55 | 1553802873.284 | 972 | 0.20921 | 0.000654 | 7.5e-05 | 97.7766 | 97.2353 | 122.8474 | 21.2612 | 0.002482 | 0.002271 | 149.2277 | 0.25696 | 8.4056 | 0.002382 | 0.002736 | 0.002667 | 5.6e-05 | 0.00103 | 0.33196 | 1.3467 | 2.5475 | 100.61 | 21.11 | 25.677 | 81.5181 | 79.1281 | 76.4481 | 0.006398 | 1.0099 | 0.024482 | 24.243 | 0.049652 | 0.15286 | 0.001447 | 0.1021 | 0.1021 | 0.061836
2019-03-28 15:55 | 1553802874.2841 | 973 | 0.2092 | 0.000652 | 7.4e-05 | 97.7325 | 97.2543 | 122.9265 | 21.2612 | 0.002448 | 0.002304 | 149.2238 | 0.25718 | 8.4044 | 0.002367 | 0.003245 | 0.003096 | 5.5e-05 | 0.001028 | 0.2959 | 1.3374 | 2.5475 | 100.61 | 21.11 | 25.674 | 80.7079 | 78.7657 | 77.2563 | 0.006397 | 1.0075 | 0.024462 | 24.2789 | 0.043869 | 0.153 | 0.001538 | 0.10869 | 0.10869 | 0.058135
2019-03-28 15:55 | 1553802875.2855 | 974 | 0.20922 | 0.00065 | 7.4e-05 | 97.5949 | 97.2683 | 122.9545 | 21.2612 | 0.002465 | 0.00231 | 149.2141 | 0.25744 | 8.4046 | 0.002415 | 0.003582 | 0.003437 | 5.3e-05 | 0.001029 | 0.33699 | 1.272 | 2.5474 | 100.61 | 21.11 | 25.674 | 81.1155 | 79.9242 | 77.407 | 0.006397 | 1.0071 | 0.024459 | 24.2868 | 0.029985 | 0.15315 | 0.001408 | 0.099206 | 0.099206 | 0.063757
2019-03-28 15:55 | 1553802876.2853 | 975 | 0.2092 | 0.000651 | 7.4e-05 | 97.5582 | 97.0998 | 123.1085 | 21.2612 | 0.002421 | 0.002279 | 149.2155 | 0.25697 | 8.4034 | 0.002455 | 0.003879 | 0.00379 | 5.6e-05 | 0.001032 | 0.31669 | 1.2866 | 2.5474 | 100.61 | 21.11 | 25.673 | 80.0638 | 80.8747 | 76.6468 | 0.006397 | 1.0093 | 0.024501 | 24.2757 | 0.042135 | 0.15287 | 0.001538 | 0.10892 | 0.10892 | 0.057967
2019-03-28 15:55 | 1553802877.2856 | 976 | 0.20921 | 0.000647 | 7.4e-05 | 97.5621 | 97.0352 | 122.7815 | 21.2612 | 0.002413 | 0.002302 | 149.2279 | 0.25671 | 8.4031 | 0.002396 | 0.003879 | 0.003776 | 5.7e-05 | 0.001032 | 0.31205 | 1.2441 | 2.5473 | 100.61 | 21.1 | 25.672 | 79.8821 | 79.4573 | 77.2073 | 0.006392 | 1.0077 | 0.024449 | 24.2628 | 0.048438 | 0.15272 | 0.001484 | 0.10474 | 0.10474 | 0.060217
2019-03-28 15:55 | 1553802878.2843 | 977 | 0.20921 | 0.00064 | 7.3e-05 | 97.4683 | 97.0495 | 122.5289 | 21.2612 | 0.002477 | 0.002231 | 149.2219 | 0.25746 | 8.4024 | 0.002445 | 0.00359 | 0.003496 | 5.6e-05 | 0.001035 | 0.323 | 1.3023 | 2.5474 | 100.61 | 21.1 | 25.671 | 81.4066 | 80.6367 | 75.4893 | 0.006392 | 1.0126 | 0.024484 | 24.1784 | 0.038517 | 0.15316 | 0.00147 | 0.10392 | 0.10392 | 0.06087
2019-03-28 15:55 | 1553802879.2846 | 978 | 0.20921 | 0.000645 | 7.3e-05 | 97.3441 | 96.796 | 122.4409 | 21.2612 | 0.002411 | 0.002242 | 127.8548 | 0.25668 | 8.3997 | 0.002383 | 0.003819 | 0.003726 | 5.4e-05 | 0.001034 | 0.30309 | 1.3283 | 2.5474 | 100.61 | 21.1 | 25.67 | 79.8245 | 79.1507 | 75.7598 | 0.006392 | 1.0119 | 0.024466 | 24.1791 | 0.0505 | 0.1527 | 0.001484 | 0.10486 | 0.10486 | 0.060141
2019-03-28 15:55 | 1553802880.2843 | 979 | 0.20919 | 0.000648 | 7.2e-05 | 97.3322 | 96.8511 | 122.6489 | 21.2612 | 0.002453 | 0.002229 | 97.7086 | 0.25787 | 8.4039 | 0.002463 | 0.004041 | 0.003891 | 5.9e-05 | 0.001034 | 0.33085 | 1.3434 | 2.5474 | 100.61 | 21.1 | 25.669 | 80.8216 | 81.0696 | 75.458 | 0.006392 | 1.0127 | 0.024497 | 24.1891 | 0.044321 | 0.15341 | 0.001571 | 0.11148 | 0.11148 | 0.05683
2019-03-28 15:55 | 1553802881.2844 | 980 | 0.20916 | 0.000645 | 7.1e-05 | 97.311 | 96.7573 | 123.1925 | 21.2612 | 0.002441 | 0.00227 | 67.5786 | 0.25784 | 8.4058 | 0.002422 | 0.004135 | 0.004075 | 5.9e-05 | 0.001033 | 0.33467 | 1.3251 | 2.5474 | 100.61 | 21.1 | 25.669 | 80.544 | 80.084 | 76.4411 | 0.006392 | 1.0099 | 0.024517 | 24.2768 | 0.051041 | 0.15339 | 0.001779 | 0.12694 | 0.12694 | 0.049904
2019-03-28 15:55 | 1553802882.2857 | 981 | 0.20923 | 0.000641 | 7e-05 | 97.2967 | 96.6663 | 123.158 | 21.2612 | 0.002447 | 0.002251 | 37.4039 | 0.25772 | 8.405 | 0.002504 | 0.003931 | 0.003897 | 5.7e-05 | 0.001031 | 0.31332 | 1.33 | 2.5475 | 100.61 | 21.1 | 25.669 | 80.6829 | 82.0647 | 75.9827 | 0.006392 | 1.0112 | 0.024529 | 24.2575 | 0.058148 | 0.15332 | 0.001352 | 0.095597 | 0.095597 | 0.066238
2019-03-28 15:55 | 1553802883.2848 | 982 | 0.20923 | 0.00064 | 6.7e-05 | 97.2586 | 96.7373 | 123.0967 | 21.2612 | 0.002437 | 0.00224 | 7.2797 | 0.2576 | 8.4009 | 0.002505 | 0.003909 | 0.003721 | 5e-05 | 0.00102 | 0.31061 | 1.3255 | 2.5475 | 100.61 | 21.1 | 25.669 | 80.4411 | 82.0765 | 75.706 | 0.006392 | 1.012 | 0.024533 | 24.2419 | 0.048074 | 0.15325 | 0.001363 | 0.096559 | 0.096559 | 0.065546
2019-03-28 15:55 | 1553802884.2844 | 983 | 0.20925 | 0.000641 | 6.3e-05 | 97.2684 | 96.7689 | 123.7159 | 21.2612 | 0.002404 | 0.002234 | -22.8587 | 0.24043 | 8.3992 | 0.002394 | 0.003955 | 0.003815 | 5.3e-05 | 0.001021 | 0.34468 | 1.3231 | 2.5474 | 100.61 | 21.1 | 25.67 | 79.6553 | 79.4087 | 75.5587 | 0.006392 | 1.0124 | 0.0246 | 24.2976 | 0.046058 | 0.14303 | 0.001257 | 0.089276 | 0.089276 | 0.066168
2019-03-28 15:55 | 1553802885.2834 | 984 | 0.20923 | 0.000644 | 5.8e-05 | 97.2768 | 96.7128 | 124.0997 | 21.2612 | 0.002385 | 0.002213 | -52.9953 | 0.23481 | 8.4004 | 0.002428 | 0.003803 | 0.003639 | 5.4e-05 | 0.001021 | 0.31404 | 1.2917 | 2.5474 | 100.61 | 21.1 | 25.671 | 79.2038 | 80.2323 | 75.0566 | 0.006392 | 1.0139 | 0.024656 | 24.3177 | 0.052007 | 0.13969 | 0.001371 | 0.098191 | 0.098191 | 0.058755
2019-03-28 15:55 | 1553802886.2839 | 985 | 0.20923 | 0.000636 | 5.4e-05 | 97.2059 | 96.635 | 124.0842 | 21.2612 | 0.002435 | 0.002176 | -83.1294 | 0.23418 | 8.4003 | 0.002486 | 0.00337 | 0.00331 | 5.4e-05 | 0.001021 | 0.32516 | 1.2998 | 2.5474 | 100.61 | 21.09 | 25.67 | 80.3904 | 81.621 | 74.1665 | 0.006388 | 1.0165 | 0.024686 | 24.2851 | 0.052685 | 0.13931 | 0.00137 | 0.09847 | 0.09847 | 0.05843
2019-03-28 15:55 | 1553802887.2846 | 986 | 0.20932 | 0.000632 | 5e-05 | 97.1535 | 96.756 | 124.683 | 21.2612 | 0.00238 | 0.002202 | -113.2865 | 0.23468 | 8.3997 | 0.002425 | 0.003613 | 0.003361 | 5.3e-05 | 0.001022 | 0.3361 | 1.3242 | 2.5474 | 100.61 | 21.09 | 25.669 | 79.0735 | 80.1712 | 74.8103 | 0.006388 | 1.0146 | 0.024722 | 24.3662 | 0.036672 | 0.13961 | 0.000857 | 0.060691 | 0.060691 | 0.095006
2019-03-28 15:55 | 1553802888.2839 | 987 | 0.20935 | 0.000622 | 4.7e-05 | 97.1079 | 96.5277 | 125.1275 | 21.2612 | 0.002397 | 0.002221 | -143.4317 | 0.23383 | 8.4023 | 0.002489 | 0.00373 | 0.003659 | 5.6e-05 | 0.001015 | 0.34219 | 1.2997 | 2.5473 | 100.61 | 21.09 | 25.668 | 79.4931 | 81.7005 | 75.257 | 0.006388 | 1.0133 | 0.02475 | 24.4253 | 0.053605 | 0.1391 | 0.000713 | 0.050236 | 0.050236 | 0.11436
2019-03-28 15:55 | 1553802889.2846 | 988 | 0.20943 | 0.000621 | 4.2e-05 | 97.026 | 96.6391 | 125.3555 | 21.2612 | 0.002293 | 0.002183 | -152.183 | 0.23321 | 8.4022 | 0.002485 | 0.003759 | 0.003551 | 5.4e-05 | 0.001015 | 0.30525 | 1.3018 | 2.5474 | 100.61 | 21.09 | 25.667 | 76.9829 | 81.6085 | 74.3428 | 0.006387 | 1.016 | 0.024806 | 24.4154 | 0.035738 | 0.13873 | 0.000223 | 0.014059 | 0.014059 | 0.40756
2019-03-28 15:55 | 1553802890.285 | 989 | 0.20939 | 0.000612 | 3.8e-05 | 96.9765 | 96.4386 | 125.2196 | 21.2612 | 0.002359 | 0.002164 | -152.2104 | 0.2344 | 8.4028 | 0.002556 | 0.004249 | 0.004042 | 5.8e-05 | 0.001013 | 0.3063 | 1.3077 | 2.5474 | 100.61 | 21.09 | 25.667 | 78.5615 | 83.3144 | 73.8914 | 0.006387 | 1.0173 | 0.024808 | 24.3863 | 0.04975 | 0.13945 | 0.000474 | 0.033037 | 0.033037 | 0.17433
2019-03-28 15:55 | 1553802891.2851 | 990 | 0.20938 | 0.000599 | 3.3e-05 | 96.9286 | 96.4957 | 125.4167 | 21.2612 | 0.002419 | 0.002144 | -152.2109 | 0.2336 | 8.4031 | 0.002533 | 0.003855 | 0.003641 | 5.7e-05 | 0.001011 | 0.34092 | 1.3395 | 2.5474 | 100.61 | 21.09 | 25.666 | 80.0107 | 82.7635 | 73.4027 | 0.006387 | 1.0187 | 0.024845 | 24.3883 | 0.040034 | 0.13897 | 0.000537 | 0.038108 | 0.038108 | 0.15061
2019-03-28 15:55 | 1553802892.2851 | 991 | 0.2094 | 0.000594 | 2.9e-05 | 96.9476 | 96.369 | 125.5714 | 21.2612 | 0.002307 | 0.002177 | -152.204 | 0.23335 | 8.4048 | 0.002522 | 0.003907 | 0.00368 | 5.5e-05 | 0.001011 | 0.33268 | 1.2951 | 2.5475 | 100.61 | 21.09 | 25.663 | 77.3219 | 82.4803 | 74.1953 | 0.006386 | 1.0164 | 0.024832 | 24.4312 | 0.053546 | 0.13882 | 0.000437 | 0.030897 | 0.030897 | 0.18556
2019-03-28 15:55 | 1553802893.2837 | 992 | 0.20947 | 0.00058 | 2.5e-05 | 96.9578 | 96.4794 | 125.8322 | 21.2612 | 0.002362 | 0.002141 | -152.1985 | 0.23413 | 8.4009 | 0.002524 | 0.004098 | 0.003929 | 5.8e-05 | 0.001005 | 0.31669 | 1.3156 | 2.5475 | 100.61 | 21.09 | 25.661 | 78.648 | 82.5421 | 73.3433 | 0.006386 | 1.0189 | 0.024888 | 24.4266 | 0.044239 | 0.13928 | 6.9e-05 | 0.003604 | 0.003604 | 1.5962
2019-03-28 15:55 | 1553802894.2837 | 993 | 0.2095 | 0.00057 | 2.1e-05 | 96.8589 | 96.1978 | 125.5294 | 21.2612 | 0.002333 | 0.00216 | -152.1922 | 0.2339 | 8.4048 | 0.002616 | 0.003733 | 0.003458 | 5.5e-05 | 0.001006 | 0.30414 | 1.3293 | 2.5473 | 100.61 | 21.09 | 25.658 | 77.9394 | 84.7444 | 73.8042 | 0.006385 | 1.0176 | 0.024842 | 24.4134 | 0.06126 | 0.13914 | -0.000119 | -0.010186 | -0.010186 | -0.56418
2019-03-28 15:55 | 1553802895.2838 | 994 | 0.20947 | 0.000559 | 1.8e-05 | 96.822 | 96.2888 | 125.1987 | 21.2612 | 0.002355 | 0.002106 | -152.1934 | 0.23391 | 8.4032 | 0.002525 | 0.003966 | 0.003771 | 5.6e-05 | 0.000999 | 0.33229 | 1.288 | 2.5473 | 100.61 | 21.09 | 25.655 | 78.4803 | 82.57 | 72.4959 | 0.006384 | 1.0214 | 0.024856 | 24.3352 | 0.049387 | 0.13915 | 8.8e-05 | 0.005433 | 0.005433 | 1.0578
2019-03-28 15:55 | 1553802896.2838 | 995 | 0.20949 | 0.000546 | 1.5e-05 | 96.7707 | 96.3045 | 125.528 | 21.2612 | 0.00224 | 0.002116 | -152.1821 | 0.2341 | 8.4016 | 0.002531 | 0.004048 | 0.003849 | 5.5e-05 | 0.000996 | 0.33931 | 1.3196 | 2.5473 | 100.61 | 21.09 | 25.652 | 75.724 | 82.7104 | 72.7261 | 0.006384 | 1.0207 | 0.024881 | 24.3753 | 0.043201 | 0.13926 | 7e-06 | -0.00046 | -0.00046 | -12.5144
2019-03-28 15:55 | 1553802897.2839 | 996 | 0.20946 | 0.000536 | 1.3e-05 | 96.7287 | 96.122 | 125.5792 | 21.2612 | 0.002284 | 0.002095 | -152.1553 | 0.23412 | 8.4021 | 0.002547 | 0.004033 | 0.003851 | 5.6e-05 | 0.00099 | 0.32659 | 1.3212 | 2.5474 | 100.61 | 21.09 | 25.648 | 76.7756 | 83.1013 | 72.241 | 0.006383 | 1.0222 | 0.024903 | 24.3632 | 0.056279 | 0.13928 | 0.000176 | 0.012363 | 0.012363 | 0.46528
2019-03-28 15:55 | 1553802898.2848 | 997 | 0.2095 | 0.000526 | 1.1e-05 | 96.6757 | 96.0187 | 125.4707 | 21.2612 | 0.002321 | 0.002153 | -152.1497 | 0.23371 | 8.406 | 0.002555 | 0.003668 | 0.003557 | 5.9e-05 | 0.00099 | 0.3126 | 1.2975 | 2.5473 | 100.61 | 21.09 | 25.646 | 77.6672 | 83.2932 | 73.6177 | 0.006382 | 1.0181 | 0.024843 | 24.4011 | 0.06099 | 0.13903 | -7e-05 | -0.005936 | -0.005936 | -0.96732
2019-03-28 15:55 | 1553802899.2839 | 998 | 0.20948 | 0.000517 | 9e-06 | 96.7555 | 96.0935 | 125.3667 | 21.2612 | 0.002273 | 0.002101 | -152.1462 | 0.23385 | 8.4033 | 0.002596 | 0.003701 | 0.003505 | 5.8e-05 | 0.000991 | 0.32477 | 1.3112 | 2.5474 | 100.61 | 21.09 | 25.643 | 76.4995 | 84.2622 | 72.3736 | 0.006381 | 1.0218 | 0.024877 | 24.3472 | 0.061413 | 0.13912 | 9.6e-05 | 0.006619 | 0.006619 | 0.86808
2019-03-28 15:55 | 1553802900.2829 | 999 | 0.20949 | 0.000515 | 8e-06 | 96.7162 | 96.2222 | 125.2961 | 21.2612 | 0.002235 | 0.002082 | -152.133 | 0.23353 | 8.4046 | 0.002507 | 0.004323 | 0.004095 | 5.6e-05 | 0.000989 | 0.32333 | 1.3094 | 2.5475 | 100.61 | 21.09 | 25.643 | 75.5844 | 82.1272 | 71.9269 | 0.006381 | 1.0231 | 0.024886 | 24.3246 | 0.045801 | 0.13892 | 3.9e-05 | 0.00243 | 0.00243 | 2.3612
2019-03-28 15:55 | 1553802901.2833 | 1000 | 0.20948 | 0.000504 | 6e-06 | 96.7198 | 96.2977 | 125.1285 | 21.2612 | 0.002251 | 0.002098 | -152.1381 | 0.23379 | 8.403 | 0.002509 | 0.004248 | 0.004106 | 5.7e-05 | 0.000989 | 0.32173 | 1.3267 | 2.5474 | 100.61 | 21.09 | 25.641 | 75.9727 | 82.1717 | 72.2948 | 0.006381 | 1.022 | 0.024856 | 24.3213 | 0.039123 | 0.13908 | 9e-05 | 0.006365 | 0.006365 | 0.90238
2019-03-28 15:55 | 1553802902.284 | 1001 | 0.20951 | 0.000498 | 5e-06 | 96.7463 | 96.2114 | 125.5108 | 21.2612 | 0.002235 | 0.00209 | -152.1251 | 0.23392 | 8.4021 | 0.002488 | 0.00393 | 0.003784 | 5.6e-05 | 0.000987 | 0.29789 | 1.3135 | 2.5474 | 100.61 | 21.09 | 25.636 | 75.5844 | 81.6662 | 72.1069 | 0.00638 | 1.0226 | 0.024901 | 24.3518 | 0.049593 | 0.13916 | -8.3e-05 | -0.006532 | -0.006532 | -0.87978
2019-03-28 15:55 | 1553802903.284 | 1002 | 0.20952 | 0.000495 | 5e-06 | 96.7609 | 96.048 | 125.7977 | 21.2612 | 0.00226 | 0.002116 | -152.1319 | 0.23476 | 8.4022 | 0.002512 | 0.004187 | 0.003946 | 5.4e-05 | 0.000991 | 0.32676 | 1.2916 | 2.5474 | 100.61 | 21.09 | 25.633 | 76.1932 | 82.2479 | 72.7307 | 0.006379 | 1.0207 | 0.024907 | 24.4016 | 0.066147 | 0.13966 | -0.00014 | -0.010823 | -0.010823 | -0.53293
2019-03-28 15:55 | 1553802904.2841 | 1003 | 0.20952 | 0.000483 | 3e-06 | 96.7993 | 96.1533 | 125.7247 | 21.2612 | 0.002308 | 0.002065 | -152.1492 | 0.23397 | 8.4031 | 0.002542 | 0.003454 | 0.003244 | 5.6e-05 | 0.000994 | 0.31614 | 1.2796 | 2.5474 | 100.61 | 21.09 | 25.631 | 77.3476 | 82.9609 | 71.5067 | 0.006378 | 1.0243 | 0.024944 | 24.3514 | 0.059896 | 0.13919 | -9.3e-05 | -0.007221 | -0.007221 | -0.79608
2019-03-28 15:55 | 1553802905.2833 | 1004 | 0.20953 | 0.000478 | 4e-06 | 96.7764 | 96.3085 | 125.6789 | 21.2612 | 0.002242 | 0.002058 | -152.1455 | 0.23496 | 8.4031 | 0.002554 | 0.00352 | 0.003296 | 5.5e-05 | 0.000997 | 0.32173 | 1.3254 | 2.5473 | 100.61 | 21.09 | 25.626 | 75.7612 | 83.2658 | 71.343 | 0.006377 | 1.0248 | 0.024945 | 24.3411 | 0.043346 | 0.13978 | -0.000147 | -0.011269 | -0.011269 | -0.51226
2019-03-28 15:55 | 1553802906.2832 | 1005 | 0.20955 | 0.000476 | 3e-06 | 96.8357 | 96.3204 | 125.6451 | 21.2612 | 0.002195 | 0.00206 | -152.1644 | 0.23478 | 8.4015 | 0.002472 | 0.004418 | 0.004232 | 5.5e-05 | 0.000993 | 0.3178 | 1.34 | 2.5475 | 100.61 | 21.09 | 25.622 | 74.6287 | 81.2826 | 71.3945 | 0.006376 | 1.0247 | 0.02494 | 24.3397 | 0.047718 | 0.13967 | -0.000238 | -0.018108 | -0.018108 | -0.31854
2019-03-28 15:55 | 1553802907.2846 | 1006 | 0.20953 | 0.00048 | 3e-06 | 96.8231 | 96.3248 | 125.6702 | 21.2612 | 0.002224 | 0.002087 | -152.1855 | 0.23455 | 8.3999 | 0.002505 | 0.003564 | 0.003328 | 5.3e-05 | 0.000998 | 0.32532 | 1.308 | 2.5473 | 100.61 | 21.09 | 25.614 | 75.3403 | 82.0765 | 72.0461 | 0.006374 | 1.0227 | 0.024919 | 24.3651 | 0.046152 | 0.13954 | -0.000148 | -0.011273 | -0.011273 | -0.51121
2019-03-28 15:55 | 1553802908.2833 | 1007 | 0.20951 | 0.000473 | 2e-06 | 96.8634 | 96.1254 | 125.9174 | 21.2612 | 0.002183 | 0.002043 | -152.1791 | 0.23377 | 8.4024 | 0.002427 | 0.003805 | 0.003672 | 5.3e-05 | 0.001002 | 0.34235 | 1.3394 | 2.5474 | 100.61 | 21.09 | 25.608 | 74.3334 | 80.2063 | 70.9915 | 0.006373 | 1.0259 | 0.024982 | 24.3518 | 0.068403 | 0.13907 | -8e-06 | -0.000738 | -0.000738 | -7.7862
2019-03-28 15:55 | 1553802909.2842 | 1008 | 0.20952 | 0.000472 | 2e-06 | 96.9244 | 96.3422 | 125.9993 | 21.2612 | 0.002258 | 0.002038 | -152.1907 | 0.23476 | 8.4025 | 0.00251 | 0.004058 | 0.003875 | 5.6e-05 | 0.001002 | 0.32803 | 1.2878 | 2.5474 | 100.61 | 21.09 | 25.602 | 76.1394 | 82.2137 | 70.8581 | 0.006371 | 1.0263 | 0.024995 | 24.355 | 0.053894 | 0.13966 | -7.3e-05 | -0.005625 | -0.005625 | -1.0254
2019-03-28 15:55 | 1553802910.2846 | 1009 | 0.2095 | 0.000464 | 1e-06 | 96.8589 | 96.4726 | 125.9721 | 21.2612 | 0.002213 | 0.002024 | -152.1925 | 0.23472 | 8.4066 | 0.002446 | 0.003548 | 0.003356 | 5.5e-05 | 0.001004 | 0.32333 | 1.328 | 2.5472 | 100.61 | 21.09 | 25.588 | 75.0583 | 80.6735 | 70.5267 | 0.006368 | 1.0273 | 0.025004 | 24.3406 | 0.035741 | 0.13963 | 4.4e-05 | 0.003201 | 0.003201 | 1.8014
2019-03-28 15:55 | 1553802911.2843 | 1010 | 0.20952 | 0.000472 | 1e-06 | 96.8333 | 96.3772 | 125.7989 | 21.2612 | 0.00218 | 0.002001 | -152.2107 | 0.23489 | 8.4017 | 0.002509 | 0.003264 | 0.003134 | 5.5e-05 | 0.001007 | 0.30702 | 1.3291 | 2.5474 | 100.61 | 21.09 | 25.58 | 74.2609 | 82.1731 | 69.9723 | 0.006366 | 1.0289 | 0.025007 | 24.3043 | 0.042231 | 0.13973 | -8.8e-05 | -0.006686 | -0.006686 | -0.86312
2019-03-28 15:55 | 1553802912.2833 | 1011 | 0.20951 | 0.000465 | 1e-06 | 96.9366 | 96.4155 | 125.6975 | 21.2612 | 0.002149 | 0.00208 | -152.2454 | 0.23389 | 8.4044 | 0.002501 | 0.003976 | 0.003868 | 5.1e-05 | 0.001014 | 0.31188 | 1.3123 | 2.5474 | 100.61 | 21.1 | 25.562 | 73.5312 | 81.975 | 71.877 | 0.006365 | 1.0232 | 0.024928 | 24.3618 | 0.048211 | 0.13914 | -2.8e-05 | -0.002122 | -0.002122 | -2.7082
2019-03-28 15:55 | 1553802913.2833 | 1012 | 0.20952 | 0.000467 | 1e-06 | 96.7984 | 96.3245 | 125.7888 | 21.2612 | 0.002196 | 0.002018 | -152.2487 | 0.2346 | 8.4045 | 0.002551 | 0.004086 | 0.003794 | 5.2e-05 | 0.001021 | 0.32964 | 1.3179 | 2.5473 | 100.61 | 21.1 | 25.553 | 74.653 | 83.1872 | 70.3731 | 0.006363 | 1.0277 | 0.024992 | 24.3175 | 0.0439 | 0.13956 | -8.9e-05 | -0.006728 | -0.006728 | -0.8567
2019-03-28 15:55 | 1553802914.2843 | 1013 | 0.20953 | 0.00046 | 1e-06 | 96.7779 | 96.1862 | 125.8587 | 21.2612 | 0.002195 | 0.002027 | -152.2308 | 0.23429 | 8.4053 | 0.002523 | 0.003655 | 0.00354 | 5.5e-05 | 0.001025 | 0.31669 | 1.353 | 2.5475 | 100.61 | 21.1 | 25.543 | 74.6297 | 82.5203 | 70.5993 | 0.00636 | 1.027 | 0.02499 | 24.3322 | 0.054849 | 0.13938 | -0.000125 | -0.009453 | -0.009453 | -0.60898
2019-03-28 15:55 | 1553802915.2834 | 1014 | 0.20954 | 0.000456 | 1e-06 | 96.813 | 96.2233 | 126.127 | 21.2612 | 0.002216 | 0.002004 | -152.2435 | 0.23463 | 8.404 | 0.002536 | 0.004318 | 0.004129 | 5.5e-05 | 0.001028 | 0.31454 | 1.3434 | 2.5477 | 100.61 | 21.1 | 25.524 | 75.1299 | 82.8292 | 70.04 | 0.006356 | 1.0287 | 0.025037 | 24.3383 | 0.054654 | 0.13958 | -0.000157 | -0.011902 | -0.011902 | -0.48435
2019-03-28 15:55 | 1553802916.2825 | 1015 | 0.20954 | 0.00046 | 1e-06 | 96.8711 | 96.4519 | 126.0127 | 21.2612 | 0.002191 | 0.001963 | -152.2335 | 0.23376 | 8.4046 | 0.002581 | 0.003209 | 0.003155 | 5.3e-05 | 0.001025 | 0.3194 | 1.3066 | 2.5472 | 100.61 | 21.09 | 25.513 | 74.5439 | 83.9112 | 69.0599 | 0.006349 | 1.0317 | 0.025062 | 24.2925 | 0.038791 | 0.13906 | -0.00017 | -0.012857 | -0.012857 | -0.4467
2019-03-28 15:55 | 1553802917.2829 | 1016 | 0.20953 | 0.000457 | 1e-06 | 96.9161 | 96.5283 | 126.0362 | 21.2612 | 0.002211 | 0.002021 | -152.2284 | 0.23416 | 8.4032 | 0.002544 | 0.003717 | 0.003426 | 5.5e-05 | 0.00103 | 0.31741 | 1.3085 | 2.5472 | 100.61 | 21.09 | 25.496 | 75.0246 | 83.0242 | 70.4527 | 0.006345 | 1.0275 | 0.025013 | 24.3442 | 0.035857 | 0.1393 | -0.000111 | -0.008429 | -0.008429 | -0.68258
2019-03-28 15:55 | 1553802918.2835 | 1017 | 0.20953 | 0.000461 | 1e-06 | 96.9313 | 96.4261 | 126.0117 | 21.2612 | 0.002133 | 0.001979 | -152.2227 | 0.23314 | 8.4013 | 0.002548 | 0.003456 | 0.003297 | 5.4e-05 | 0.001031 | 0.31238 | 1.3382 | 2.5474 | 100.61 | 21.09 | 25.489 | 73.1555 | 83.1045 | 69.4358 | 0.006343 | 1.0305 | 0.025048 | 24.3058 | 0.046739 | 0.13869 | -0.000113 | -0.008565 | -0.008565 | -0.66873
2019-03-28 15:55 | 1553802919.2826 | 1018 | 0.20952 | 0.000457 | 1e-06 | 96.9506 | 96.3235 | 126.0479 | 21.2612 | 0.002156 | 0.002003 | -152.2197 | 0.23364 | 8.4031 | 0.002464 | 0.004171 | 0.004012 | 5.4e-05 | 0.001032 | 0.33644 | 1.3053 | 2.5472 | 100.61 | 21.09 | 25.482 | 73.7052 | 81.0929 | 70.016 | 0.006341 | 1.0288 | 0.02503 | 24.3299 | 0.058047 | 0.13899 | -6.9e-05 | -0.005234 | -0.005234 | -1.0967
2019-03-28 15:55 | 1553802920.2829 | 1019 | 0.20951 | 0.000457 | 0 | 97.0144 | 96.5535 | 126.1784 | 21.2612 | 0.002134 | 0.001993 | -152.2105 | 0.23392 | 8.4048 | 0.002518 | 0.003997 | 0.003809 | 5.3e-05 | 0.001034 | 0.29446 | 1.3068 | 2.5474 | 100.61 | 21.09 | 25.474 | 73.1742 | 82.3924 | 69.7774 | 0.006339 | 1.0295 | 0.025052 | 24.334 | 0.042595 | 0.13916 | -2.6e-05 | -0.001958 | -0.001958 | -2.9358
2019-03-28 15:55 | 1553802921.2836 | 1020 | 0.20951 | 0.000459 | 1e-06 | 97.0302 | 96.4679 | 126.4425 | 21.2612 | 0.002123 | 0.002016 | -152.2075 | 0.2337 | 8.4028 | 0.002562 | 0.004089 | 0.003917 | 5.6e-05 | 0.001037 | 0.33516 | 1.2891 | 2.5473 | 100.61 | 21.09 | 25.471 | 72.9092 | 83.4507 | 70.3412 | 0.006338 | 1.0278 | 0.025058 | 24.3794 | 0.051984 | 0.13903 | 2.4e-05 | 0.001803 | 0.001803 | 3.1842
2019-03-28 15:55 | 1553802922.2837 | 1021 | 0.20953 | 0.000455 | 1e-06 | 97.1374 | 96.597 | 126.2397 | 21.2612 | 0.002107 | 0.001989 | -152.188 | 0.2344 | 8.4016 | 0.002519 | 0.004208 | 0.00396 | 5.6e-05 | 0.001044 | 0.33771 | 1.2837 | 2.5475 | 100.61 | 21.09 | 25.471 | 72.5226 | 82.4306 | 69.6745 | 0.006338 | 1.0298 | 0.025062 | 24.3362 | 0.049901 | 0.13944 | -0.000124 | -0.009432 | -0.009432 | -0.61059
2019-03-28 15:55 | 1553802923.2837 | 1022 | 0.20952 | 0.000455 | 0 | 97.1329 | 96.8106 | 126.0693 | 21.2612 | 0.00213 | 0.002016 | -152.1791 | 0.23388 | 8.4039 | 0.002504 | 0.004319 | 0.004178 | 5.4e-05 | 0.001041 | 0.32189 | 1.2982 | 2.5473 | 100.61 | 21.09 | 25.468 | 73.0791 | 82.0577 | 70.3381 | 0.006338 | 1.0278 | 0.025021 | 24.3433 | 0.029725 | 0.13913 | -8.4e-05 | -0.006379 | -0.006379 | -0.9008
2019-03-28 15:55 | 1553802924.2839 | 1023 | 0.20952 | 0.000453 | 0 | 97.2282 | 96.7879 | 126.0791 | 21.2612 | 0.002113 | 0.001987 | -152.1951 | 0.23407 | 8.4032 | 0.002462 | 0.003838 | 0.003631 | 5.3e-05 | 0.00104 | 0.35706 | 1.3068 | 2.5474 | 100.61 | 21.09 | 25.466 | 72.6559 | 81.0477 | 69.6307 | 0.006337 | 1.0299 | 0.025047 | 24.3192 | 0.040601 | 0.13925 | -7.4e-05 | -0.005608 | -0.005608 | -1.0255
2019-03-28 15:55 | 1553802925.2828 | 1024 | 0.20954 | 0.000452 | 0 | 97.3593 | 96.6177 | 126.0002 | 21.2612 | 0.002149 | 0.001997 | -152.1803 | 0.23411 | 8.4036 | 0.002424 | 0.003481 | 0.003286 | 5.7e-05 | 0.001034 | 0.32084 | 1.3131 | 2.5473 | 100.61 | 21.09 | 25.468 | 73.5375 | 80.1239 | 69.8819 | 0.006338 | 1.0292 | 0.02503 | 24.3205 | 0.068392 | 0.13927 | -0.000167 | -0.012623 | -0.012623 | -0.45568
2019-03-28 15:55 | 1553802926.2827 | 1025 | 0.20954 | 0.000453 | 0 | 97.32 | 96.9961 | 125.751 | 21.2612 | 0.002138 | 0.001932 | -152.1762 | 0.23427 | 8.4016 | 0.002403 | 0.00341 | 0.003235 | 5.6e-05 | 0.001031 | 0.34755 | 1.3102 | 2.5472 | 100.61 | 21.09 | 25.47 | 73.2725 | 79.631 | 68.3183 | 0.006338 | 1.0339 | 0.025063 | 24.241 | 0.029812 | 0.13937 | -0.000149 | -0.011289 | -0.011289 | -0.50986
2019-03-28 15:55 | 1553802927.2828 | 1026 | 0.20955 | 0.000453 | 0 | 97.3867 | 96.9462 | 125.8643 | 21.2612 | 0.002122 | 0.001941 | -152.184 | 0.23364 | 8.4033 | 0.002497 | 0.003756 | 0.003547 | 5.7e-05 | 0.001032 | 0.30469 | 1.3082 | 2.5474 | 100.61 | 21.09 | 25.473 | 72.8703 | 81.8861 | 68.5241 | 0.006339 | 1.0333 | 0.025067 | 24.2592 | 0.040548 | 0.13899 | -0.000238 | -0.018002 | -0.018002 | -0.31888
2019-03-28 15:55 | 1553802928.2838 | 1027 | 0.20954 | 0.000453 | 1e-06 | 97.4677 | 97.1395 | 125.9805 | 21.2612 | 0.002115 | 0.001958 | -152.1902 | 0.23305 | 8.4042 | 0.002482 | 0.003673 | 0.003486 | 5.6e-05 | 0.001035 | 0.31437 | 1.282 | 2.5474 | 100.61 | 21.1 | 25.48 | 72.7064 | 81.5244 | 68.9366 | 0.006345 | 1.032 | 0.025063 | 24.2851 | 0.030169 | 0.13864 | -0.000167 | -0.01269 | -0.01269 | -0.45123
2019-03-28 15:55 | 1553802929.2833 | 1028 | 0.20954 | 0.000453 | 0 | 97.5496 | 97.0189 | 126.0329 | 21.2612 | 0.002128 | 0.001958 | -152.1791 | 0.2334 | 8.4042 | 0.002484 | 0.00371 | 0.003596 | 5.6e-05 | 0.001032 | 0.30862 | 1.2936 | 2.5474 | 100.61 | 21.1 | 25.484 | 73.0221 | 81.5765 | 68.9282 | 0.006346 | 1.0321 | 0.025069 | 24.2898 | 0.048794 | 0.13885 | -0.000147 | -0.011164 | -0.011164 | -0.51368
2019-03-28 15:55 | 1553802930.2829 | 1029 | 0.20952 | 0.000451 | 1e-06 | 97.4939 | 97.1371 | 126.0164 | 21.2612 | 0.002103 | 0.001973 | -152.1887 | 0.23354 | 8.4042 | 0.002444 | 0.004229 | 0.003977 | 6.1e-05 | 0.001027 | 0.32604 | 1.3264 | 2.5474 | 100.61 | 21.1 | 25.492 | 72.4125 | 80.6182 | 69.2885 | 0.006348 | 1.031 | 0.025054 | 24.301 | 0.032793 | 0.13893 | -3.7e-05 | -0.002868 | -0.002868 | -2.0008
2019-03-28 15:56 | 1553802931.2819 | 1030 | 0.20949 | 0.00045 | 1e-06 | 97.5463 | 97.2669 | 125.9974 | 21.2612 | 0.002061 | 0.001935 | -152.1905 | 0.23388 | 8.4035 | 0.002517 | 0.003187 | 0.00308 | 5.8e-05 | 0.001023 | 0.32731 | 1.3084 | 2.5474 | 100.61 | 21.1 | 25.494 | 71.4187 | 82.3705 | 68.3854 | 0.006348 | 1.0337 | 0.025085 | 24.2671 | 0.025659 | 0.13914 | 0.000115 | 0.008678 | 0.008678 | 0.66221
2019-03-28 15:56 | 1553802932.282 | 1031 | 0.2095 | 0.000449 | 1e-06 | 97.6422 | 97.3168 | 125.9923 | 21.2612 | 0.002126 | 0.00195 | -152.1852 | 0.23368 | 8.4034 | 0.002502 | 0.003939 | 0.003737 | 5.6e-05 | 0.001019 | 0.31382 | 1.296 | 2.5473 | 100.61 | 21.1 | 25.496 | 72.9793 | 82.0046 | 68.7517 | 0.006349 | 1.0326 | 0.025071 | 24.2796 | 0.029858 | 0.13902 | 6.6e-05 | 0.004928 | 0.004928 | 1.165
2019-03-28 15:56 | 1553802933.2824 | 1032 | 0.20949 | 0.000446 | 1e-06 | 97.7703 | 97.3759 | 126.2052 | 21.2612 | 0.002063 | 0.001919 | -152.181 | 0.23388 | 8.4019 | 0.002485 | 0.003648 | 0.003461 | 6.1e-05 | 0.001021 | 0.33539 | 1.336 | 2.5474 | 100.61 | 21.1 | 25.498 | 71.467 | 81.6078 | 68.0038 | 0.006349 | 1.0349 | 0.02512 | 24.2735 | 0.036151 | 0.13913 | 0.000126 | 0.009455 | 0.009455 | 0.60773
2019-03-28 15:56 | 1553802934.2819 | 1033 | 0.20949 | 0.000449 | 0 | 97.8731 | 97.4072 | 126.3406 | 21.2612 | 0.002125 | 0.001914 | -152.1587 | 0.23289 | 8.401 | 0.002471 | 0.0032 | 0.003018 | 6.1e-05 | 0.001021 | 0.31708 | 1.3257 | 2.5474 | 100.61 | 21.1 | 25.499 | 72.9551 | 81.259 | 67.8735 | 0.006349 | 1.0353 | 0.025138 | 24.2819 | 0.042677 | 0.13855 | 0.000141 | 0.010656 | 0.010656 | 0.537
2019-03-28 15:56 | 1553802935.283 | 1034 | 0.20949 | 0.000448 | 0 | 97.8147 | 97.3053 | 126.2228 | 21.2612 | 0.002093 | 0.001957 | -152.1846 | 0.23293 | 8.4024 | 0.00246 | 0.003573 | 0.003388 | 6e-05 | 0.001025 | 0.31277 | 1.3673 | 2.5476 | 100.61 | 21.1 | 25.499 | 72.1874 | 81.0102 | 68.9032 | 0.006349 | 1.0321 | 0.025088 | 24.3072 | 0.046704 | 0.13857 | 0.000125 | 0.009398 | 0.009398 | 0.60893
2019-03-28 15:56 | 1553802936.2824 | 1035 | 0.2095 | 0.00045 | 0 | 97.8519 | 97.4836 | 126.4554 | 21.2612 | 0.002097 | 0.001936 | -152.2044 | 0.2334 | 8.4029 | 0.002485 | 0.003867 | 0.003691 | 6e-05 | 0.001027 | 0.34523 | 1.3476 | 2.5475 | 100.61 | 21.1 | 25.498 | 72.2878 | 81.6123 | 68.4101 | 0.006349 | 1.0336 | 0.02513 | 24.3121 | 0.033728 | 0.13885 | 6.9e-05 | 0.005179 | 0.005179 | 1.1072
2019-03-28 15:56 | 1553802937.2821 | 1036 | 0.20952 | 0.000448 | 1e-06 | 97.8734 | 97.6267 | 126.1584 | 21.2612 | 0.00209 | 0.001925 | -152.2082 | 0.23455 | 8.4017 | 0.002494 | 0.004576 | 0.004434 | 6.1e-05 | 0.001023 | 0.3095 | 1.3324 | 2.5475 | 100.61 | 21.1 | 25.497 | 72.1062 | 81.8097 | 68.1498 | 0.006349 | 1.0344 | 0.02511 | 24.2742 | 0.022574 | 0.13954 | -1.3e-05 | -0.001059 | -0.001059 | -5.4406
2019-03-28 15:56 | 1553802938.2822 | 1037 | 0.2095 | 0.000447 | 1e-06 | 97.9282 | 97.732 | 125.8012 | 21.2612 | 0.002068 | 0.001941 | -152.226 | 0.2339 | 8.4018 | 0.002489 | 0.003994 | 0.003792 | 5.9e-05 | 0.001023 | 0.32875 | 1.3537 | 2.5475 | 100.61 | 21.1 | 25.493 | 71.5761 | 81.6981 | 68.5404 | 0.006348 | 1.0332 | 0.02506 | 24.2537 | 0.017938 | 0.13914 | 7.6e-05 | 0.005682 | 0.005682 | 1.0113
2019-03-28 15:56 | 1553802939.2822 | 1038 | 0.20951 | 0.000452 | 1e-06 | 97.983 | 97.6311 | 125.9561 | 21.2612 | 0.002075 | 0.001949 | -152.2397 | 0.23377 | 8.4035 | 0.002506 | 0.004067 | 0.00388 | 5.6e-05 | 0.001022 | 0.31614 | 1.309 | 2.5472 | 100.61 | 21.1 | 25.49 | 71.7467 | 82.1169 | 68.7332 | 0.006347 | 1.0326 | 0.025068 | 24.2755 | 0.032181 | 0.13907 | -1.3e-05 | -0.001035 | -0.001035 | -5.5476
2019-03-28 15:56 | 1553802940.2832 | 1039 | 0.20952 | 0.000447 | 0 | 98.0327 | 97.6712 | 126.1458 | 21.2612 | 0.002087 | 0.001906 | -152.2296 | 0.23416 | 8.4035 | 0.002447 | 0.003665 | 0.003501 | 5.8e-05 | 0.001025 | 0.31868 | 1.3018 | 2.5474 | 100.61 | 21.1 | 25.488 | 72.0305 | 80.6906 | 67.698 | 0.006347 | 1.0358 | 0.025125 | 24.257 | 0.033049 | 0.1393 | -3.6e-05 | -0.002677 | -0.002677 | -2.1489
2019-03-28 15:56 | 1553802941.2824 | 1040 | 0.20951 | 0.000453 | 0 | 98.0622 | 97.8787 | 126.0153 | 21.2612 | 0.002113 | 0.001966 | -152.2056 | 0.23459 | 8.3999 | 0.002402 | 0.003795 | 0.003621 | 6.2e-05 | 0.001019 | 0.2975 | 1.277 | 2.5473 | 100.61 | 21.1 | 25.481 | 72.6535 | 79.5959 | 69.1207 | 0.006345 | 1.0315 | 0.02506 | 24.2949 | 0.01675 | 0.13956 | 3.4e-05 | 0.002579 | 0.002579 | 2.2345
2019-03-28 15:56 | 1553802942.2833 | 1041 | 0.20949 | 0.000448 | 0 | 98.137 | 97.8903 | 126.0502 | 21.2612 | 0.002075 | 0.001903 | -152.1983 | 0.23517 | 8.4003 | 0.002419 | 0.004123 | 0.003924 | 6e-05 | 0.001022 | 0.32444 | 1.301 | 2.5473 | 100.61 | 21.1 | 25.479 | 71.7592 | 80.0208 | 67.6069 | 0.006344 | 1.0361 | 0.025119 | 24.2445 | 0.022511 | 0.1399 | 0.00012 | 0.009036 | 0.009036 | 0.63943
2019-03-28 15:56 | 1553802943.2823 | 1042 | 0.20951 | 0.000445 | 0 | 98.1727 | 97.8692 | 125.6989 | 21.2612 | 0.002097 | 0.001921 | -152.1929 | 0.23385 | 8.405 | 0.002472 | 0.003823 | 0.003676 | 6.2e-05 | 0.00103 | 0.30718 | 1.3271 | 2.5474 | 100.61 | 21.1 | 25.475 | 72.2871 | 81.2965 | 68.0483 | 0.006343 | 1.0347 | 0.025068 | 24.2264 | 0.027693 | 0.13912 | -8e-06 | -0.00063 | -0.00063 | -9.1145
2019-03-28 15:56 | 1553802944.2834 | 1043 | 0.20951 | 0.000446 | 0 | 98.2085 | 97.8726 | 125.8296 | 21.2612 | 0.002064 | 0.001861 | -152.1987 | 0.23417 | 8.4012 | 0.002367 | 0.004199 | 0.003924 | 6e-05 | 0.001031 | 0.33229 | 1.3359 | 2.5475 | 100.61 | 21.1 | 25.475 | 71.4889 | 78.7741 | 66.6176 | 0.006343 | 1.0391 | 0.025133 | 24.1881 | 0.030638 | 0.13931 | 4e-05 | 0.003015 | 0.003015 | 1.9085
2019-03-28 15:56 | 1553802945.2824 | 1044 | 0.20951 | 0.000453 | 0 | 98.1942 | 98.0126 | 126.1054 | 21.2612 | 0.002059 | 0.001944 | -152.1755 | 0.23455 | 8.4024 | 0.002315 | 0.003818 | 0.003628 | 6e-05 | 0.00103 | 0.3282 | 1.3733 | 2.5475 | 100.61 | 21.1 | 25.476 | 71.3577 | 77.523 | 68.5998 | 0.006344 | 1.0331 | 0.025088 | 24.2851 | 0.016554 | 0.13954 | 2.2e-05 | 0.001645 | 0.001645 | 3.504
2019-03-28 15:56 | 1553802946.2814 | 1045 | 0.20952 | 0.000452 | 1e-06 | 98.2781 | 97.9229 | 125.8458 | 21.2612 | 0.002001 | 0.001856 | -152.1618 | 0.23422 | 8.4033 | 0.002451 | 0.003116 | 0.002941 | 6.1e-05 | 0.001029 | 0.31725 | 1.3379 | 2.5473 | 100.61 | 21.1 | 25.479 | 69.9754 | 80.7874 | 66.4856 | 0.006344 | 1.0395 | 0.02514 | 24.185 | 0.03239 | 0.13933 | -4.4e-05 | -0.003345 | -0.003345 | -1.7205
2019-03-28 15:56 | 1553802947.2827 | 1046 | 0.20951 | 0.000448 | 1e-06 | 98.3994 | 98.0337 | 126.012 | 21.2612 | 0.002072 | 0.001893 | -152.1701 | 0.23358 | 8.4035 | 0.002386 | 0.00364 | 0.003475 | 6.4e-05 | 0.001031 | 0.33019 | 1.3209 | 2.5472 | 100.61 | 21.1 | 25.481 | 71.6674 | 79.2225 | 67.3828 | 0.006345 | 1.0367 | 0.025123 | 24.2329 | 0.033306 | 0.13896 | 1.1e-05 | 0.00076 | 0.00076 | 7.5483
2019-03-28 15:56 | 1553802948.2815 | 1047 | 0.20952 | 0.000451 | 1e-06 | 98.436 | 98.1149 | 126.1807 | 21.2612 | 0.001988 | 0.00196 | -152.1538 | 0.23342 | 8.4022 | 0.00236 | 0.003216 | 0.002998 | 6.2e-05 | 0.001032 | 0.32477 | 1.3213 | 2.5476 | 100.61 | 21.1 | 25.482 | 69.6651 | 78.5879 | 68.9911 | 0.006345 | 1.0319 | 0.025081 | 24.3063 | 0.02923 | 0.13886 | -2.4e-05 | -0.001875 | -0.001875 | -3.0588
2019-03-28 15:56 | 1553802949.2825 | 1048 | 0.20949 | 0.000452 | 0 | 98.4971 | 98.1403 | 126.2591 | 21.2612 | 0.002018 | 0.001912 | -152.142 | 0.23318 | 8.401 | 0.002452 | 0.002504 | 0.002354 | 6.3e-05 | 0.001033 | 0.31852 | 1.3364 | 2.5475 | 100.61 | 21.1 | 25.482 | 70.3825 | 80.8115 | 67.8269 | 0.006345 | 1.0354 | 0.025132 | 24.2724 | 0.032454 | 0.13872 | 0.0001 | 0.007595 | 0.007595 | 0.75431
2019-03-28 15:56 | 1553802950.2818 | 1049 | 0.2095 | 0.000449 | 0 | 98.6308 | 98.334 | 126.4169 | 21.2612 | 0.002026 | 0.001957 | -152.155 | 0.23373 | 8.4024 | 0.00244 | 0.004165 | 0.003867 | 6.1e-05 | 0.001034 | 0.31957 | 1.3088 | 2.5475 | 100.61 | 21.1 | 25.482 | 70.5712 | 80.5269 | 68.9063 | 0.006345 | 1.0321 | 0.025108 | 24.326 | 0.026959 | 0.13905 | 9.9e-05 | 0.007514 | 0.007514 | 0.76428
2019-03-28 15:56 | 1553802951.2815 | 1050 | 0.20951 | 0.000448 | 0 | 98.7196 | 98.4974 | 126.5702 | 21.2612 | 0.002037 | 0.001926 | -152.1611 | 0.23353 | 8.4032 | 0.002378 | 0.00375 | 0.003535 | 5.8e-05 | 0.001036 | 0.32764 | 1.3282 | 2.5474 | 100.61 | 21.1 | 25.483 | 70.8355 | 79.0207 | 68.1613 | 0.006345 | 1.0344 | 0.02515 | 24.3142 | 0.020154 | 0.13893 | 4.2e-05 | 0.003144 | 0.003144 | 1.8248
2019-03-28 15:56 | 1553802952.2815 | 1051 | 0.20951 | 0.000453 | 1e-06 | 98.8077 | 98.5803 | 126.5335 | 21.2612 | 0.002055 | 0.001897 | -152.1593 | 0.23296 | 8.4023 | 0.002425 | 0.004397 | 0.004262 | 5.8e-05 | 0.00103 | 0.32421 | 1.3164 | 2.5474 | 100.61 | 21.1 | 25.482 | 71.269 | 80.158 | 67.4811 | 0.006345 | 1.0364 | 0.025172 | 24.2865 | 0.020614 | 0.13859 | 1e-06 | 2.4e-05 | 2.4e-05 | 242.3664
2019-03-28 15:56 | 1553802953.282 | 1052 | 0.2095 | 0.000449 | 1e-06 | 98.8634 | 98.4943 | 126.4988 | 21.2612 | 0.00207 | 0.001916 | -152.1799 | 0.23442 | 8.4037 | 0.002449 | 0.004693 | 0.004497 | 6.1e-05 | 0.001027 | 0.33085 | 1.2948 | 2.5475 | 100.61 | 21.1 | 25.478 | 71.6268 | 80.7399 | 67.9312 | 0.006344 | 1.0351 | 0.025152 | 24.2992 | 0.033457 | 0.13946 | 6.2e-05 | 0.004649 | 0.004649 | 1.2389
2019-03-28 15:56 | 1553802954.2818 | 1053 | 0.20952 | 0.000449 | 0 | 98.9501 | 98.6978 | 126.3648 | 21.2612 | 0.002071 | 0.001906 | -152.1871 | 0.23472 | 8.4052 | 0.002455 | 0.00419 | 0.004002 | 6.2e-05 | 0.001028 | 0.31636 | 1.2952 | 2.5474 | 100.61 | 21.1 | 25.476 | 71.665 | 80.891 | 67.6848 | 0.006344 | 1.0358 | 0.025147 | 24.2775 | 0.022834 | 0.13964 | -2.7e-05 | -0.002056 | -0.002056 | -2.8056
2019-03-28 15:56 | 1553802955.2817 | 1054 | 0.20952 | 0.000453 | 1e-06 | 99.0651 | 98.8317 | 126.2608 | 21.2612 | 0.002013 | 0.001904 | -152.1814 | 0.23433 | 8.4022 | 0.002453 | 0.003329 | 0.003209 | 6e-05 | 0.001028 | 0.34794 | 1.285 | 2.5474 | 100.61 | 21.1 | 25.474 | 70.2686 | 80.8232 | 67.6504 | 0.006343 | 1.0359 | 0.025138 | 24.2663 | 0.021099 | 0.13941 | -3.7e-05 | -0.002873 | -0.002873 | -2.0042
2019-03-28 15:56 | 1553802956.2818 | 1055 | 0.20951 | 0.00045 | 0 | 99.0823 | 98.7943 | 126.692 | 21.2612 | 0.002062 | 0.001866 | -152.1939 | 0.23447 | 8.405 | 0.002463 | 0.00342 | 0.003215 | 6.4e-05 | 0.00103 | 0.34451 | 1.2594 | 2.5473 | 100.61 | 21.1 | 25.465 | 71.4303 | 81.0828 | 66.7338 | 0.006341 | 1.0387 | 0.025215 | 24.275 | 0.026041 | 0.13949 | 1.4e-05 | 0.001027 | 0.001027 | 5.6099
2019-03-28 15:56 | 1553802957.346 | 1056 | 0.20953 | 0.000449 | 0 | 99.1532 | 98.8932 | 126.8553 | 21.2612 | 0.002086 | 0.001871 | -152.1735 | 0.23472 | 8.404 | 0.002509 | 0.003372 | 0.003253 | 6.5e-05 | 0.001034 | 0.3402 | 1.2523 | 2.5474 | 100.61 | 21.1 | 25.46 | 72.0079 | 82.17 | 66.8579 | 0.00634 | 1.0383 | 0.025227 | 24.2951 | 0.023492 | 0.13963 | -9.7e-05 | -0.007373 | -0.007373 | -0.78218
2019-03-28 15:56 | 1553802958.2942 | 1057 | 0.2095 | 0.000444 | 1e-06 | 99.1902 | 99.1731 | 126.7889 | 21.2612 | 0.001976 | 0.001898 | -152.182 | 0.23452 | 8.4008 | 0.002466 | 0.003921 | 0.003715 | 5.9e-05 | 0.001035 | 0.3282 | 1.2793 | 2.5472 | 100.61 | 21.1 | 25.453 | 69.3632 | 81.1484 | 67.5037 | 0.006338 | 1.0364 | 0.025196 | 24.3118 | 0.00154 | 0.13951 | 5.5e-05 | 0.004122 | 0.004122 | 1.3979
2019-03-28 15:56 | 1553802959.2947 | 1058 | 0.2095 | 0.000449 | 1e-06 | 99.2569 | 99.1531 | 126.7861 | 21.2612 | 0.001994 | 0.001895 | -152.1837 | 0.23383 | 8.4023 | 0.002412 | 0.003768 | 0.00356 | 6.2e-05 | 0.001039 | 0.34108 | 1.304 | 2.5474 | 100.61 | 21.1 | 25.447 | 69.7961 | 79.8509 | 67.4148 | 0.006336 | 1.0366 | 0.025199 | 24.3083 | 0.009362 | 0.1391 | 4.6e-05 | 0.003464 | 0.003464 | 1.6585
2019-03-28 15:56 | 1553802960.2945 | 1059 | 0.20952 | 0.00045 | 0 | 99.3873 | 99.2838 | 126.4577 | 21.2612 | 0.002009 | 0.001906 | -152.1908 | 0.23445 | 8.403 | 0.002463 | 0.004186 | 0.004018 | 6.2e-05 | 0.001042 | 0.30021 | 1.3092 | 2.5474 | 100.61 | 21.1 | 25.434 | 70.1564 | 81.0626 | 67.691 | 0.006333 | 1.0358 | 0.025156 | 24.2867 | 0.00932 | 0.13947 | -3.4e-05 | -0.00257 | -0.00257 | -2.2417
2019-03-28 15:56 | 1553802961.2942 | 1060 | 0.20952 | 0.00045 | 1e-06 | 99.3677 | 99.2176 | 126.4767 | 21.2612 | 0.002087 | 0.001889 | -152.1989 | 0.23491 | 8.4038 | 0.002451 | 0.003636 | 0.003541 | 6.2e-05 | 0.001043 | 0.32372 | 1.3159 | 2.5474 | 100.61 | 21.1 | 25.429 | 72.0408 | 80.7725 | 67.2847 | 0.006332 | 1.037 | 0.025173 | 24.274 | 0.013521 | 0.13975 | -4.2e-05 | -0.003196 | -0.003196 | -1.8061
2019-03-28 15:56 | 1553802962.2946 | 1061 | 0.2095 | 0.000447 | 0 | 99.4847 | 99.2648 | 126.6065 | 21.2612 | 0.002077 | 0.001904 | -152.2093 | 0.23472 | 8.4034 | 0.002435 | 0.004022 | 0.003789 | 6.1e-05 | 0.001044 | 0.32947 | 1.2789 | 2.5474 | 100.61 | 21.11 | 25.423 | 71.8037 | 80.3974 | 67.6302 | 0.006334 | 1.036 | 0.025173 | 24.2988 | 0.019795 | 0.13963 | 7.9e-05 | 0.005948 | 0.005948 | 0.96961
2019-03-28 15:56 | 1553802963.2943 | 1062 | 0.20951 | 0.000449 | 1e-06 | 99.5354 | 99.3355 | 126.7645 | 21.2612 | 0.002011 | 0.001888 | -152.2055 | 0.2347 | 8.4035 | 0.002439 | 0.003582 | 0.003517 | 6.4e-05 | 0.001046 | 0.31564 | 1.3152 | 2.5474 | 100.61 | 21.11 | 25.413 | 70.2008 | 80.487 | 67.2597 | 0.006332 | 1.0371 | 0.025203 | 24.3007 | 0.017982 | 0.13962 | 1.1e-05 | 0.000789 | 0.000789 | 7.3081
2019-03-28 15:56 | 1553802964.2943 | 1063 | 0.20952 | 0.000447 | 0 | 99.5479 | 99.5203 | 126.7101 | 21.2612 | 0.00199 | 0.001884 | -152.2288 | 0.2348 | 8.402 | 0.002364 | 0.004022 | 0.0038 | 5.9e-05 | 0.001044 | 0.31598 | 1.2912 | 2.5473 | 100.61 | 21.11 | 25.408 | 69.6992 | 78.6869 | 67.1645 | 0.006331 | 1.0374 | 0.025201 | 24.2921 | 0.00248 | 0.13968 | -5.1e-05 | -0.003935 | -0.003935 | -1.466
2019-03-28 15:56 | 1553802965.2937 | 1064 | 0.20947 | 0.000452 | 0 | 99.592 | 99.3086 | 126.4462 | 21.2612 | 0.002027 | 0.001911 | -152.2586 | 0.23325 | 8.4027 | 0.002471 | 0.003965 | 0.003792 | 6.4e-05 | 0.001042 | 0.29983 | 1.3087 | 2.5472 | 100.61 | 21.11 | 25.398 | 70.6039 | 81.2559 | 67.8166 | 0.006328 | 1.0354 | 0.025151 | 24.29 | 0.025482 | 0.13876 | 0.000283 | 0.021405 | 0.021405 | 0.26774
2019-03-28 15:56 | 1553802966.2944 | 1065 | 0.20949 | 0.00045 | 0 | 99.6417 | 99.4106 | 126.3113 | 21.2612 | 0.002007 | 0.001831 | -152.2485 | 0.23388 | 8.4017 | 0.002395 | 0.003346 | 0.003178 | 5.9e-05 | 0.001043 | 0.31885 | 1.2971 | 2.5475 | 100.61 | 21.11 | 25.391 | 70.1049 | 79.4455 | 65.8771 | 0.006326 | 1.0413 | 0.025209 | 24.2079 | 0.020774 | 0.13913 | 0.000161 | 0.012227 | 0.012227 | 0.46997
2019-03-28 15:56 | 1553802967.2944 | 1066 | 0.20948 | 0.00045 | 0 | 99.6501 | 99.523 | 126.2913 | 21.2612 | 0.002047 | 0.001876 | -152.2713 | 0.23505 | 8.4026 | 0.002483 | 0.003747 | 0.003571 | 6.5e-05 | 0.001039 | 0.32748 | 1.2953 | 2.5474 | 100.61 | 21.11 | 25.385 | 71.0787 | 81.5602 | 66.9594 | 0.006325 | 1.038 | 0.025167 | 24.2446 | 0.01141 | 0.13983 | 0.000173 | 0.013102 | 0.013102 | 0.44079
2019-03-28 15:56 | 1553802968.2934 | 1067 | 0.20951 | 0.000451 | 0 | 99.7275 | 99.4574 | 126.5192 | 21.2612 | 0.001918 | 0.001873 | -152.2797 | 0.23452 | 8.402 | 0.002454 | 0.002826 | 0.002606 | 6.4e-05 | 0.001034 | 0.3194 | 1.3016 | 2.5474 | 100.61 | 21.11 | 25.372 | 67.9726 | 80.8528 | 66.893 | 0.006322 | 1.0382 | 0.025192 | 24.2641 | 0.024254 | 0.13951 | 2.3e-05 | 0.001728 | 0.001728 | 3.3339
2019-03-28 15:56 | 1553802969.2939 | 1068 | 0.20951 | 0.000448 | 0 | 99.8073 | 99.8155 | 126.5737 | 21.2612 | 0.002008 | 0.001879 | -152.2919 | 0.23363 | 8.4004 | 0.002362 | 0.003445 | 0.003253 | 6.2e-05 | 0.001032 | 0.29518 | 1.3185 | 2.5474 | 100.61 | 21.11 | 25.367 | 70.1283 | 78.64 | 67.0411 | 0.00632 | 1.0378 | 0.025192 | 24.2746 | -0.000734 | 0.13899 | 2.7e-05 | 0.002018 | 0.002018 | 2.8447
2019-03-28 15:56 | 1553802970.2945 | 1069 | 0.20956 | 0.000451 | 1e-06 | 99.8014 | 99.8152 | 126.5943 | 21.2612 | 0.001991 | 0.001794 | -152.2867 | 0.23437 | 8.4049 | 0.002462 | 0.003948 | 0.003753 | 6.1e-05 | 0.001035 | 0.29966 | 1.3044 | 2.5474 | 100.61 | 21.1 | 25.363 | 69.7367 | 81.0501 | 64.9928 | 0.006315 | 1.0441 | 0.02527 | 24.2034 | -0.001238 | 0.13943 | -0.000295 | -0.022538 | -0.022538 | -0.2555
2019-03-28 15:56 | 1553802971.2941 | 1070 | 0.20951 | 0.000446 | 1e-06 | 99.8285 | 99.4897 | 126.2411 | 21.2612 | 0.002006 | 0.001828 | -152.2823 | 0.23392 | 8.4033 | 0.002509 | 0.004305 | 0.004142 | 6.5e-05 | 0.001034 | 0.34274 | 1.2509 | 2.5473 | 100.61 | 21.1 | 25.357 | 70.0939 | 82.1801 | 65.809 | 0.006314 | 1.0416 | 0.025204 | 24.1988 | 0.030403 | 0.13916 | 4e-05 | 0.003007 | 0.003007 | 1.9114
2019-03-28 15:56 | 1553802972.295 | 1071 | 0.20949 | 0.000444 | 1e-06 | 99.8761 | 99.771 | 126.4277 | 21.2612 | 0.001954 | 0.001913 | -152.2782 | 0.23442 | 8.4019 | 0.002459 | 0.004597 | 0.004469 | 6.1e-05 | 0.001035 | 0.31636 | 1.2508 | 2.5472 | 100.61 | 21.1 | 25.357 | 68.8323 | 80.9667 | 67.8509 | 0.006314 | 1.0353 | 0.025147 | 24.2895 | 0.009418 | 0.13946 | 0.000123 | 0.009284 | 0.009284 | 0.62038
2019-03-28 15:56 | 1553802973.2936 | 1072 | 0.20949 | 0.00045 | 0 | 99.894 | 99.7187 | 126.4911 | 21.2612 | 0.001958 | 0.001817 | -152.2744 | 0.23474 | 8.4053 | 0.002458 | 0.003617 | 0.003412 | 6.2e-05 | 0.001042 | 0.32139 | 1.3043 | 2.5473 | 100.61 | 21.1 | 25.356 | 68.939 | 80.9542 | 65.5406 | 0.006314 | 1.0424 | 0.025239 | 24.2131 | 0.01571 | 0.13965 | 0.000156 | 0.011873 | 0.011873 | 0.48578
2019-03-28 15:56 | 1553802974.294 | 1073 | 0.20958 | 0.000449 | 0 | 99.961 | 99.7724 | 126.3629 | 21.2612 | 0.002002 | 0.0019 | -152.2494 | 0.235 | 8.403 | 0.002471 | 0.003103 | 0.002893 | 6.2e-05 | 0.001045 | 0.33732 | 1.2804 | 2.5473 | 100.61 | 21.1 | 25.356 | 70.0042 | 81.2621 | 67.5468 | 0.006314 | 1.0362 | 0.025152 | 24.2724 | 0.016895 | 0.1398 | -0.000401 | -0.030349 | -0.030349 | -0.19025
2019-03-28 15:56 | 1553802975.2947 | 1074 | 0.20953 | 0.000448 | 0 | 99.9813 | 99.9477 | 126.3744 | 21.2612 | 0.001973 | 0.001844 | -152.2467 | 0.23441 | 8.4026 | 0.00239 | 0.003909 | 0.003741 | 6.2e-05 | 0.001048 | 0.31708 | 1.3105 | 2.5475 | 100.61 | 21.1 | 25.356 | 69.2916 | 79.323 | 66.197 | 0.006314 | 1.0404 | 0.025203 | 24.2254 | 0.003006 | 0.13945 | -9.5e-05 | -0.007218 | -0.007218 | -0.79789
2019-03-28 15:56 | 1553802976.2937 | 1075 | 0.20947 | 0.000448 | 0 | 100.0784 | 100.0673 | 126.266 | 21.2612 | 0.001992 | 0.001858 | -152.2427 | 0.23436 | 8.4024 | 0.002431 | 0.003936 | 0.003799 | 6.6e-05 | 0.001051 | 0.34219 | 1.339 | 2.5475 | 100.61 | 21.1 | 25.359 | 69.7478 | 80.2976 | 66.5238 | 0.006314 | 1.0394 | 0.02518 | 24.2267 | 0.000993 | 0.13942 | 0.000232 | 0.017617 | 0.017617 | 0.32685
2019-03-28 15:56 | 1553802977.2936 | 1076 | 0.20954 | 0.000452 | 0 | 100.2017 | 100.1729 | 126.1453 | 21.2612 | 0.001942 | 0.001843 | -152.24 | 0.23519 | 8.4029 | 0.002466 | 0.004104 | 0.003923 | 6.3e-05 | 0.001053 | 0.33212 | 1.3516 | 2.5474 | 100.61 | 21.1 | 25.361 | 68.556 | 81.1391 | 66.1681 | 0.006315 | 1.0405 | 0.025182 | 24.2024 | 0.002568 | 0.13991 | -0.000142 | -0.010764 | -0.010764 | -0.53683
2019-03-28 15:56 | 1553802978.2938 | 1077 | 0.20948 | 0.000453 | 0 | 100.1781 | 100.1141 | 126.1038 | 21.2612 | 0.001975 | 0.0019 | -152.2336 | 0.2342 | 8.4059 | 0.002519 | 0.004321 | 0.004096 | 6.3e-05 | 0.001057 | 0.32139 | 1.3311 | 2.5474 | 100.61 | 21.11 | 25.366 | 69.3483 | 82.4111 | 67.5429 | 0.00632 | 1.0363 | 0.025127 | 24.2474 | 0.005717 | 0.13932 | 0.00019 | 0.014364 | 0.014364 | 0.40058
2019-03-28 15:56 | 1553802979.2931 | 1078 | 0.20957 | 0.00045 | 1e-06 | 100.252 | 100.2011 | 125.9132 | 21.2612 | 0.001994 | 0.001871 | -152.2332 | 0.23475 | 8.4011 | 0.002418 | 0.003713 | 0.003572 | 6.6e-05 | 0.001058 | 0.30879 | 1.3279 | 2.5475 | 100.61 | 21.11 | 25.37 | 69.7985 | 79.9989 | 66.8447 | 0.006321 | 1.0384 | 0.025133 | 24.2042 | 0.004543 | 0.13965 | -0.000321 | -0.024395 | -0.024395 | -0.23643
2019-03-28 15:56 | 1553802980.2942 | 1079 | 0.20951 | 0.000457 | 0 | 100.2115 | 100.0143 | 125.9763 | 21.2612 | 0.001988 | 0.001903 | -152.2237 | 0.23349 | 8.4047 | 0.00245 | 0.003785 | 0.003566 | 6e-05 | 0.001061 | 0.33019 | 1.3534 | 2.5475 | 100.61 | 21.11 | 25.374 | 69.6478 | 80.7639 | 67.6153 | 0.006322 | 1.036 | 0.025111 | 24.2377 | 0.017623 | 0.1389 | -1.5e-05 | -0.001156 | -0.001156 | -4.9617
2019-03-28 15:56 | 1553802981.2928 | 1080 | 0.2095 | 0.000453 | 0 | 100.3527 | 100.2775 | 126.0704 | 21.2612 | 0.001957 | 0.001832 | -152.2394 | 0.23321 | 8.4031 | 0.002488 | 0.004437 | 0.004233 | 6.2e-05 | 0.001058 | 0.32836 | 1.3161 | 2.5475 | 100.61 | 21.11 | 25.378 | 68.9063 | 81.6748 | 65.9192 | 0.006323 | 1.0412 | 0.025183 | 24.1863 | 0.0067 | 0.13874 | 5.9e-05 | 0.004483 | 0.004483 | 1.2782
2019-03-28 15:56 | 1553802982.2928 | 1081 | 0.2095 | 0.000451 | 1e-06 | 100.3601 | 100.3081 | 125.8059 | 21.2612 | 0.001954 | 0.00182 | -152.2254 | 0.23416 | 8.4032 | 0.002513 | 0.00394 | 0.003772 | 6.2e-05 | 0.001063 | 0.32781 | 1.2901 | 2.5474 | 100.61 | 21.11 | 25.379 | 68.8438 | 82.2683 | 65.6219 | 0.006323 | 1.0421 | 0.025168 | 24.1504 | 0.004637 | 0.1393 | 6e-05 | 0.004524 | 0.004524 | 1.2716
2019-03-28 15:56 | 1553802983.2929 | 1082 | 0.20945 | 0.000455 | 1e-06 | 100.4227 | 100.5007 | 125.6205 | 21.2612 | 0.001946 | 0.001861 | -152.2301 | 0.2348 | 8.4016 | 0.002481 | 0.003363 | 0.003138 | 6.1e-05 | 0.001061 | 0.33539 | 1.2714 | 2.5474 | 100.61 | 21.11 | 25.381 | 68.6395 | 81.4955 | 66.602 | 0.006324 | 1.0391 | 0.025113 | 24.1674 | -0.00695 | 0.13968 | 0.000371 | 0.028007 | 0.028007 | 0.20598
2019-03-28 15:56 | 1553802984.2929 | 1083 | 0.20952 | 0.000453 | 1e-06 | 100.4033 | 100.4267 | 126.1746 | 21.2612 | 0.001949 | 0.001873 | -152.2368 | 0.2345 | 8.4079 | 0.00248 | 0.004017 | 0.003877 | 6e-05 | 0.001064 | 0.34634 | 1.2804 | 2.5474 | 100.61 | 21.11 | 25.382 | 68.7159 | 81.4845 | 66.8851 | 0.006324 | 1.0383 | 0.025158 | 24.2308 | -0.002081 | 0.1395 | -5.9e-05 | -0.004492 | -0.004492 | -1.2827
2019-03-28 15:56 | 1553802985.294 | 1084 | 0.20948 | 0.000454 | 0 | 100.4543 | 100.4973 | 126.4807 | 21.2612 | 0.001967 | 0.001875 | -152.247 | 0.23407 | 8.4041 | 0.002445 | 0.003925 | 0.003743 | 5.9e-05 | 0.001064 | 0.32211 | 1.3139 | 2.5474 | 100.61 | 21.11 | 25.384 | 69.1426 | 80.6478 | 66.9545 | 0.006325 | 1.038 | 0.025186 | 24.2626 | -0.003836 | 0.13925 | 0.000183 | 0.01387 | 0.01387 | 0.41462
2019-03-28 15:56 | 1553802986.2929 | 1085 | 0.20952 | 0.000453 | 0 | 100.4739 | 100.3937 | 126.3199 | 21.2612 | 0.002004 | 0.001863 | -152.221 | 0.2339 | 8.4034 | 0.002458 | 0.003468 | 0.003248 | 5.9e-05 | 0.00107 | 0.32781 | 1.3266 | 2.5474 | 100.61 | 21.11 | 25.388 | 70.041 | 80.9581 | 66.6566 | 0.006326 | 1.039 | 0.025181 | 24.2366 | 0.007141 | 0.13915 | -6.1e-05 | -0.004602 | -0.004602 | -1.2487
2019-03-28 15:56 | 1553802987.293 | 1086 | 0.20951 | 0.000447 | 1e-06 | 100.5001 | 100.4145 | 126.0315 | 21.2612 | 0.001995 | 0.001891 | -152.2099 | 0.23358 | 8.4003 | 0.002393 | 0.003523 | 0.003324 | 5.8e-05 | 0.001074 | 0.33085 | 1.3005 | 2.5474 | 100.61 | 21.11 | 25.39 | 69.8336 | 79.3893 | 67.3167 | 0.006326 | 1.0369 | 0.025128 | 24.2324 | 0.007628 | 0.13896 | 1.7e-05 | 0.001228 | 0.001228 | 4.6734
2019-03-28 15:56 | 1553802988.294 | 1087 | 0.20947 | 0.000448 | 0 | 100.5549 | 100.5578 | 126.1026 | 21.2612 | 0.001941 | 0.001796 | -152.218 | 0.23407 | 8.4021 | 0.002474 | 0.00366 | 0.003527 | 5.9e-05 | 0.001076 | 0.32748 | 1.2759 | 2.5472 | 100.61 | 21.11 | 25.394 | 68.5303 | 81.3386 | 65.0544 | 0.006327 | 1.0439 | 0.025219 | 24.1586 | -0.000257 | 0.13925 | 0.000278 | 0.021115 | 0.021115 | 0.27237
2019-03-28 15:56 | 1553802989.2931 | 1088 | 0.20947 | 0.000451 | 1e-06 | 100.5615 | 100.5238 | 126.2467 | 21.2612 | 0.001916 | 0.001879 | -152.2194 | 0.23416 | 8.4014 | 0.002481 | 0.002954 | 0.002779 | 5.8e-05 | 0.001076 | 0.33516 | 1.3049 | 2.5474 | 100.61 | 21.11 | 25.403 | 67.9329 | 81.4986 | 67.0317 | 0.006329 | 1.0378 | 0.02516 | 24.2429 | 0.003348 | 0.1393 | 0.000292 | 0.022042 | 0.022042 | 0.26102
2019-03-28 15:56 | 1553802990.2936 | 1089 | 0.2095 | 0.00045 | 0 | 100.5561 | 100.6115 | 126.1692 | 21.2612 | 0.00195 | 0.001885 | -152.2287 | 0.23448 | 8.4024 | 0.002511 | 0.004458 | 0.004282 | 5.8e-05 | 0.001081 | 0.33428 | 1.2913 | 2.5472 | 100.61 | 21.11 | 25.409 | 68.744 | 82.227 | 67.1864 | 0.006331 | 1.0373 | 0.025146 | 24.241 | -0.004924 | 0.13949 | 6.6e-05 | 0.004942 | 0.004942 | 1.1657
2019-03-28 15:56 | 1553802991.2931 | 1090 | 0.20951 | 0.000449 | 1e-06 | 100.5606 | 100.5578 | 126.1913 | 21.2612 | 0.00197 | 0.001851 | -152.22 | 0.2337 | 8.4024 | 0.002483 | 0.003804 | 0.003701 | 5.7e-05 | 0.001085 | 0.33483 | 1.309 | 2.5475 | 100.61 | 21.11 | 25.42 | 69.2253 | 81.5501 | 66.3646 | 0.006334 | 1.0399 | 0.025179 | 24.2138 | 0.000247 | 0.13903 | 1.8e-05 | 0.001281 | 0.001281 | 4.4839
2019-03-28 15:57 | 1553802992.2927 | 1091 | 0.20952 | 0.000451 | 1e-06 | 100.6362 | 100.6502 | 126.265 | 21.2612 | 0.001941 | 0.001822 | -152.2366 | 0.23442 | 8.402 | 0.002483 | 0.003771 | 0.003619 | 5.8e-05 | 0.001082 | 0.28639 | 1.3943 | 2.5475 | 100.61 | 21.11 | 25.426 | 68.538 | 81.5477 | 65.6609 | 0.006335 | 1.042 | 0.025212 | 24.1958 | -0.001242 | 0.13945 | -7.4e-05 | -0.005675 | -0.005675 | -1.015
2019-03-28 15:57 | 1553802993.2928 | 1092 | 0.20947 | 0.000452 | 1e-06 | 100.7381 | 100.8072 | 126.3488 | 21.2612 | 0.002002 | 0.001868 | -152.2292 | 0.2346 | 8.4021 | 0.002494 | 0.004016 | 0.003793 | 5.7e-05 | 0.001085 | 0.3298 | 1.3582 | 2.5474 | 100.61 | 21.11 | 25.433 | 69.998 | 81.8159 | 66.786 | 0.006337 | 1.0386 | 0.025179 | 24.244 | -0.00613 | 0.13956 | 0.000252 | 0.019066 | 0.019066 | 0.30231
2019-03-28 15:57 | 1553802994.2932 | 1093 | 0.20956 | 0.000448 | 1e-06 | 100.7286 | 100.8191 | 125.9772 | 21.2612 | 0.001966 | 0.001858 | -152.2229 | 0.23448 | 8.4048 | 0.002571 | 0.003656 | 0.003466 | 5.7e-05 | 0.001088 | 0.31686 | 1.3069 | 2.5473 | 100.61 | 21.11 | 25.45 | 69.1262 | 83.6592 | 66.5246 | 0.006341 | 1.0394 | 0.025152 | 24.199 | -0.008032 | 0.13949 | -0.000273 | -0.020762 | -0.020762 | -0.27747
2019-03-28 15:57 | 1553802995.2936 | 1094 | 0.20955 | 0.000448 | 0 | 100.7354 | 100.7647 | 126.0141 | 21.2612 | 0.002022 | 0.001864 | -152.2175 | 0.23453 | 8.4045 | 0.00247 | 0.00411 | 0.003911 | 5.7e-05 | 0.001092 | 0.31188 | 1.3077 | 2.5474 | 100.61 | 21.11 | 25.459 | 70.4861 | 81.2499 | 66.6792 | 0.006343 | 1.0389 | 0.025149 | 24.208 | -0.0026 | 0.13952 | -0.00025 | -0.01896 | -0.01896 | -0.30391
2019-03-28 15:57 | 1553802996.2924 | 1095 | 0.20957 | 0.000444 | 0 | 100.8001 | 100.7949 | 126.38 | 21.2612 | 0.001965 | 0.001825 | -152.2524 | 0.23407 | 8.4042 | 0.002481 | 0.003388 | 0.003183 | 6e-05 | 0.001094 | 0.31332 | 1.2957 | 2.5474 | 100.61 | 21.11 | 25.479 | 69.1152 | 81.5039 | 65.7499 | 0.006348 | 1.0417 | 0.02522 | 24.21 | 0.000454 | 0.13925 | -0.000323 | -0.024585 | -0.024585 | -0.23392
2019-03-28 15:57 | 1553802997.2923 | 1096 | 0.20952 | 0.000447 | 0 | 100.8072 | 100.7406 | 126.4253 | 21.2612 | 0.001885 | 0.001877 | -152.2568 | 0.23279 | 8.4031 | 0.002532 | 0.003887 | 0.003771 | 5.7e-05 | 0.001092 | 0.30558 | 1.3076 | 2.5474 | 100.61 | 21.11 | 25.488 | 67.1926 | 82.7371 | 66.9858 | 0.006351 | 1.038 | 0.025179 | 24.2584 | 0.005913 | 0.13849 | -2.8e-05 | -0.002122 | -0.002122 | -2.6955
2019-03-28 15:57 | 1553802998.2924 | 1097 | 0.20952 | 0.000445 | 0 | 100.8045 | 100.7888 | 126.4159 | 21.2612 | 0.001965 | 0.001874 | -152.2417 | 0.23476 | 8.402 | 0.002562 | 0.003536 | 0.003332 | 5.9e-05 | 0.001091 | 0.32029 | 1.311 | 2.5474 | 100.61 | 21.11 | 25.498 | 69.0974 | 83.4632 | 66.9194 | 0.006353 | 1.0382 | 0.025181 | 24.2552 | 0.001393 | 0.13966 | -2e-05 | -0.001539 | -0.001539 | -3.747
2019-03-28 15:57 | 1553802999.2924 | 1098 | 0.20948 | 0.000442 | 1e-06 | 100.854 | 100.9967 | 126.3958 | 21.2612 | 0.00196 | 0.001814 | -152.2518 | 0.23316 | 8.4015 | 0.002529 | 0.004057 | 0.003805 | 6e-05 | 0.001089 | 0.32068 | 1.3003 | 2.5474 | 100.61 | 21.11 | 25.515 | 68.9748 | 82.6707 | 65.4829 | 0.006357 | 1.0426 | 0.025232 | 24.2019 | -0.012653 | 0.13871 | 0.000192 | 0.014586 | 0.014586 | 0.39276
2019-03-28 15:57 | 1553803000.2924 | 1099 | 0.20955 | 0.000443 | 0 | 100.8349 | 101.0423 | 126.527 | 21.2612 | 0.001983 | 0.001863 | -152.2542 | 0.23389 | 8.4027 | 0.002564 | 0.004178 | 0.003994 | 5.7e-05 | 0.001091 | 0.32964 | 1.2862 | 2.5473 | 100.61 | 21.11 | 25.523 | 69.53 | 83.4961 | 66.6455 | 0.006359 | 1.039 | 0.025202 | 24.256 | -0.018374 | 0.13914 | -0.000188 | -0.014313 | -0.014313 | -0.40149
2019-03-28 15:57 | 1553803001.2934 | 1100 | 0.20951 | 0.000448 | 0 | 100.9165 | 100.7803 | 126.896 | 21.2612 | 0.001933 | 0.001876 | -152.2596 | 0.23415 | 8.4025 | 0.002535 | 0.003324 | 0.003146 | 5.7e-05 | 0.001097 | 0.32532 | 1.317 | 2.5474 | 100.61 | 21.11 | 25.538 | 68.3267 | 82.7962 | 66.9781 | 0.006363 | 1.038 | 0.025226 | 24.3033 | 0.012079 | 0.1393 | 3.2e-05 | 0.002425 | 0.002425 | 2.3719
2019-03-28 15:57 | 1553803002.2925 | 1101 | 0.20949 | 0.000444 | 1e-06 | 100.8483 | 100.8969 | 126.8922 | 21.2612 | 0.001947 | 0.001824 | -152.2451 | 0.23414 | 8.404 | 0.002495 | 0.004428 | 0.004227 | 5.6e-05 | 0.001101 | 0.37233 | 1.3097 | 2.5474 | 100.61 | 21.11 | 25.547 | 68.6825 | 81.8299 | 65.7186 | 0.006365 | 1.0418 | 0.025273 | 24.2579 | -0.004305 | 0.13929 | 0.000145 | 0.010976 | 0.010976 | 0.5241
2019-03-28 15:57 | 1553803003.2935 | 1102 | 0.20948 | 0.000443 | 1e-06 | 100.8468 | 100.8194 | 126.7551 | 21.2612 | 0.001954 | 0.001823 | -152.2631 | 0.23444 | 8.4044 | 0.002564 | 0.003877 | 0.003718 | 5.5e-05 | 0.001107 | 0.33229 | 1.2748 | 2.5473 | 100.61 | 21.11 | 25.556 | 68.8532 | 83.493 | 65.6818 | 0.006367 | 1.0419 | 0.02526 | 24.2434 | 0.002432 | 0.13947 | 0.000204 | 0.015473 | 0.015473 | 0.37227
2019-03-28 15:57 | 1553803004.2936 | 1103 | 0.20951 | 0.000445 | 1e-06 | 100.8415 | 100.8799 | 126.5662 | 21.2612 | 0.001975 | 0.001828 | -152.2625 | 0.23442 | 8.4023 | 0.002547 | 0.00392 | 0.003754 | 5.7e-05 | 0.001103 | 0.32892 | 1.3277 | 2.5474 | 100.61 | 21.11 | 25.575 | 69.3406 | 83.0903 | 65.8107 | 0.006372 | 1.0416 | 0.025237 | 24.23 | -0.003407 | 0.13945 | 3.5e-05 | 0.002629 | 0.002629 | 2.1904
2019-03-28 15:57 | 1553803005.2926 | 1104 | 0.20951 | 0.000447 | 0 | 100.8278 | 100.8823 | 126.6406 | 21.2612 | 0.001927 | 0.00181 | -152.2712 | 0.23453 | 8.401 | 0.002548 | 0.003164 | 0.002987 | 5.8e-05 | 0.001106 | 0.3314 | 1.3382 | 2.5474 | 100.61 | 21.11 | 25.584 | 68.2019 | 83.1177 | 65.3837 | 0.006374 | 1.0429 | 0.02526 | 24.2218 | -0.004833 | 0.13953 | 4.2e-05 | 0.003146 | 0.003146 | 1.8315
2019-03-28 15:57 | 1553803006.2935 | 1105 | 0.20952 | 0.000445 | 0 | 100.8319 | 100.831 | 126.5603 | 21.2612 | 0.001905 | 0.001833 | -152.272 | 0.23344 | 8.4026 | 0.002459 | 0.003578 | 0.003402 | 5.6e-05 | 0.001105 | 0.33284 | 1.3285 | 2.5473 | 100.61 | 21.11 | 25.593 | 67.6629 | 80.9667 | 65.9262 | 0.006377 | 1.0412 | 0.025232 | 24.2335 | 8.7e-05 | 0.13887 | -7e-06 | -0.000524 | -0.000524 | -10.9365
2019-03-28 15:57 | 1553803007.2936 | 1106 | 0.20952 | 0.00044 | 1e-06 | 100.9007 | 101.0334 | 127.0488 | 21.2612 | 0.00195 | 0.00186 | -152.2665 | 0.23507 | 8.4038 | 0.002464 | 0.004233 | 0.004063 | 6e-05 | 0.001101 | 0.33716 | 1.3023 | 2.5473 | 100.61 | 21.11 | 25.61 | 68.7565 | 81.0852 | 66.5825 | 0.006381 | 1.0392 | 0.025256 | 24.3038 | -0.011755 | 0.13985 | -6e-06 | -0.000522 | -0.000522 | -11.0575
2019-03-28 15:57 | 1553803008.2926 | 1107 | 0.2095 | 0.000444 | 1e-06 | 100.9654 | 100.8486 | 126.8244 | 21.2612 | 0.001959 | 0.001874 | -152.263 | 0.23425 | 8.4057 | 0.002522 | 0.003836 | 0.003694 | 5.8e-05 | 0.001106 | 0.32189 | 1.2971 | 2.5475 | 100.61 | 21.11 | 25.617 | 68.9741 | 82.489 | 66.918 | 0.006383 | 1.0382 | 0.025221 | 24.2943 | 0.010349 | 0.13935 | 7.1e-05 | 0.005309 | 0.005309 | 1.084
2019-03-28 15:57 | 1553803009.2927 | 1108 | 0.2095 | 0.000447 | 1e-06 | 100.8599 | 100.9417 | 126.6645 | 21.2612 | 0.001961 | 0.001848 | -152.2421 | 0.23511 | 8.403 | 0.002495 | 0.00375 | 0.003582 | 5.9e-05 | 0.001107 | 0.35153 | 1.3205 | 2.5474 | 100.61 | 21.11 | 25.629 | 69.0185 | 81.8479 | 66.2898 | 0.006386 | 1.0401 | 0.025229 | 24.2565 | -0.007249 | 0.13987 | 7.6e-05 | 0.00571 | 0.00571 | 1.0117
2019-03-28 15:57 | 1553803010.2927 | 1109 | 0.2095 | 0.000446 | 0 | 100.9246 | 101.0063 | 126.3819 | 21.2612 | 0.001899 | 0.001856 | -152.2489 | 0.23408 | 8.4001 | 0.002456 | 0.003838 | 0.003685 | 5.5e-05 | 0.001102 | 0.3095 | 1.3046 | 2.5474 | 100.61 | 21.11 | 25.634 | 67.5179 | 80.9074 | 66.4786 | 0.006387 | 1.0395 | 0.025194 | 24.2362 | -0.007237 | 0.13925 | 9.4e-05 | 0.007107 | 0.007107 | 0.80917
2019-03-28 15:57 | 1553803011.2928 | 1110 | 0.2095 | 0.000449 | 1e-06 | 100.9174 | 101.2199 | 125.9495 | 21.2612 | 0.001957 | 0.001841 | -152.2362 | 0.23505 | 8.4008 | 0.002532 | 0.004207 | 0.004064 | 5.4e-05 | 0.001106 | 0.3454 | 1.2928 | 2.5474 | 100.61 | 21.11 | 25.639 | 68.9181 | 82.7393 | 66.1299 | 0.006388 | 1.0406 | 0.025163 | 24.1823 | -0.026773 | 0.13983 | 0.000103 | 0.007741 | 0.007741 | 0.74605
2019-03-28 15:57 | 1553803012.2918 | 1111 | 0.2095 | 0.00044 | 1e-06 | 100.9231 | 100.9074 | 125.8451 | 21.2612 | 0.00195 | 0.00182 | -152.2605 | 0.23394 | 8.4022 | 0.002553 | 0.003699 | 0.003521 | 5.6e-05 | 0.001106 | 0.32587 | 1.2662 | 2.5474 | 100.61 | 21.11 | 25.647 | 68.7549 | 83.2456 | 65.6282 | 0.00639 | 1.0421 | 0.025172 | 24.1544 | 0.00139 | 0.13917 | 8e-05 | 0.00602 | 0.00602 | 0.95484
2019-03-28 15:57 | 1553803013.2918 | 1112 | 0.20952 | 0.000446 | 1e-06 | 100.901 | 100.8259 | 126.0693 | 21.2612 | 0.001953 | 0.001844 | -152.2571 | 0.23404 | 8.4053 | 0.002554 | 0.00383 | 0.003719 | 6.1e-05 | 0.001109 | 0.34451 | 1.3325 | 2.5475 | 100.61 | 21.11 | 25.649 | 68.8087 | 83.2566 | 66.2047 | 0.006391 | 1.0403 | 0.025173 | 24.1964 | 0.006666 | 0.13923 | -2.4e-05 | -0.001872 | -0.001872 | -3.0712
2019-03-28 15:57 | 1553803014.2933 | 1113 | 0.20951 | 0.000447 | 1e-06 | 100.9132 | 100.9295 | 126.408 | 21.2612 | 0.001965 | 0.001823 | -152.2494 | 0.23417 | 8.4039 | 0.002559 | 0.003613 | 0.003428 | 6.2e-05 | 0.001107 | 0.32892 | 1.368 | 2.5474 | 100.61 | 21.11 | 25.648 | 69.0996 | 83.3877 | 65.688 | 0.00639 | 1.0419 | 0.025226 | 24.2104 | -0.001439 | 0.13931 | 1.5e-05 | 0.001136 | 0.001136 | 5.0669
2019-03-28 15:57 | 1553803015.292 | 1114 | 0.20951 | 0.000443 | 0 | 100.8876 | 100.8802 | 126.4957 | 21.2612 | 0.00196 | 0.001828 | -152.2389 | 0.23457 | 8.402 | 0.002482 | 0.003956 | 0.003746 | 5.6e-05 | 0.001107 | 0.33987 | 1.3475 | 2.5473 | 100.61 | 21.11 | 25.647 | 68.9789 | 81.5282 | 65.8052 | 0.00639 | 1.0416 | 0.02523 | 24.223 | 0.000657 | 0.13955 | 2.5e-05 | 0.001888 | 0.001888 | 3.0527
2019-03-28 15:57 | 1553803016.2919 | 1115 | 0.20952 | 0.000445 | 1e-06 | 100.946 | 101.0518 | 126.4051 | 21.2612 | 0.001861 | 0.001817 | -152.2188 | 0.23432 | 8.4024 | 0.002531 | 0.003628 | 0.00344 | 5.8e-05 | 0.001107 | 0.33682 | 1.344 | 2.5475 | 100.61 | 21.11 | 25.645 | 66.6136 | 82.709 | 65.5562 | 0.00639 | 1.0423 | 0.02523 | 24.2055 | -0.009367 | 0.1394 | -4.8e-05 | -0.003671 | -0.003671 | -1.5684
2019-03-28 15:57 | 1553803017.2929 | 1116 | 0.20952 | 0.000448 | 1e-06 | 100.9695 | 100.9522 | 126.1343 | 21.2612 | 0.001922 | 0.001866 | -152.2335 | 0.23412 | 8.3997 | 0.002482 | 0.004301 | 0.004088 | 5.7e-05 | 0.00111 | 0.33323 | 1.3207 | 2.5474 | 100.61 | 21.11 | 25.639 | 68.0803 | 81.532 | 66.7369 | 0.006388 | 1.0387 | 0.025159 | 24.2216 | 0.001533 | 0.13928 | -6.7e-05 | -0.005171 | -0.005171 | -1.1125
2019-03-28 15:57 | 1553803018.292 | 1117 | 0.20952 | 0.00045 | 1e-06 | 100.8838 | 100.9832 | 125.8932 | 21.2612 | 0.00189 | 0.001865 | -152.2193 | 0.23436 | 8.4017 | 0.002496 | 0.003771 | 0.003557 | 5.8e-05 | 0.001109 | 0.31581 | 1.2921 | 2.5473 | 100.61 | 21.11 | 25.635 | 67.3063 | 81.8642 | 66.7119 | 0.006387 | 1.0388 | 0.025136 | 24.1976 | -0.008808 | 0.13942 | -6.2e-05 | -0.004726 | -0.004726 | -1.2185
2019-03-28 15:57 | 1553803019.293 | 1118 | 0.20953 | 0.000447 | 1e-06 | 100.9121 | 100.9584 | 125.7963 | 21.2612 | 0.001981 | 0.001847 | -152.2449 | 0.23464 | 8.4011 | 0.002443 | 0.003801 | 0.003626 | 5.8e-05 | 0.001111 | 0.32676 | 1.3261 | 2.5474 | 100.61 | 21.11 | 25.631 | 69.4897 | 80.5832 | 66.2742 | 0.006386 | 1.0401 | 0.025143 | 24.1727 | -0.004103 | 0.13959 | -0.000103 | -0.007842 | -0.007842 | -0.73517
2019-03-28 15:57 | 1553803020.292 | 1119 | 0.20953 | 0.000449 | 1e-06 | 100.9475 | 100.6995 | 125.7275 | 21.2612 | 0.002 | 0.001858 | -152.2259 | 0.23351 | 8.4061 | 0.002522 | 0.004027 | 0.003844 | 5.7e-05 | 0.001106 | 0.31885 | 1.2922 | 2.5472 | 100.61 | 21.11 | 25.622 | 69.9574 | 82.4904 | 66.5412 | 0.006384 | 1.0393 | 0.025126 | 24.1756 | 0.022003 | 0.13891 | -0.000122 | -0.009243 | -0.009243 | -0.62068
2019-03-28 15:57 | 1553803021.2936 | 1120 | 0.20953 | 0.000447 | 1e-06 | 100.9374 | 100.9172 | 125.6343 | 21.2612 | 0.001933 | 0.001848 | -152.2264 | 0.23395 | 8.4038 | 0.002467 | 0.003531 | 0.003344 | 5.4e-05 | 0.001106 | 0.34562 | 1.2909 | 2.5474 | 100.61 | 21.11 | 25.617 | 68.3291 | 81.1773 | 66.2891 | 0.006383 | 1.0401 | 0.025126 | 24.1576 | 0.001784 | 0.13918 | -8.2e-05 | -0.006241 | -0.006241 | -0.92102
2019-03-28 15:57 | 1553803022.2925 | 1121 | 0.20953 | 0.000446 | 1e-06 | 100.9055 | 100.8575 | 125.6531 | 21.2612 | 0.001929 | 0.001827 | -152.2205 | 0.2349 | 8.4002 | 0.002522 | 0.004062 | 0.003911 | 5.6e-05 | 0.001108 | 0.35579 | 1.314 | 2.5473 | 100.61 | 21.11 | 25.607 | 68.2356 | 82.4796 | 65.792 | 0.00638 | 1.0416 | 0.025146 | 24.1418 | 0.00426 | 0.13974 | -0.000104 | -0.007919 | -0.007919 | -0.72879
2019-03-28 15:57 | 1553803023.2922 | 1122 | 0.20953 | 0.00045 | 1e-06 | 100.8638 | 100.8996 | 125.5428 | 21.2612 | 0.001988 | 0.001847 | -152.2127 | 0.23531 | 8.4018 | 0.002569 | 0.004263 | 0.004036 | 5.4e-05 | 0.001109 | 0.35187 | 1.3667 | 2.5474 | 100.61 | 21.11 | 25.602 | 69.6509 | 83.6147 | 66.2812 | 0.006379 | 1.0401 | 0.025117 | 24.1486 | -0.003172 | 0.13999 | -8.3e-05 | -0.006318 | -0.006318 | -0.91503
2019-03-28 15:57 | 1553803024.2912 | 1123 | 0.20953 | 0.000444 | 1e-06 | 100.8689 | 100.8632 | 125.4212 | 21.2612 | 0.001964 | 0.00181 | -152.2202 | 0.23439 | 8.4004 | 0.002474 | 0.00415 | 0.003978 | 6e-05 | 0.001114 | 0.32228 | 1.3435 | 2.5475 | 100.61 | 21.11 | 25.592 | 69.0762 | 81.3292 | 65.3823 | 0.006376 | 1.0429 | 0.025138 | 24.1049 | 0.0005 | 0.13944 | -6.5e-05 | -0.004957 | -0.004957 | -1.1618
2019-03-28 15:57 | 1553803025.2912 | 1124 | 0.20953 | 0.00045 | 1e-06 | 100.943 | 100.9274 | 125.1998 | 21.2612 | 0.001946 | 0.001809 | -152.2233 | 0.23379 | 8.4031 | 0.002469 | 0.003624 | 0.003468 | 5.5e-05 | 0.001115 | 0.33035 | 1.2899 | 2.5473 | 100.61 | 21.11 | 25.588 | 68.645 | 81.2162 | 65.3667 | 0.006375 | 1.0429 | 0.025117 | 24.083 | 0.001382 | 0.13908 | -8.5e-05 | -0.006472 | -0.006472 | -0.88751
2019-03-28 15:57 | 1553803026.2913 | 1125 | 0.20954 | 0.000445 | 0 | 100.9275 | 100.9502 | 125.3874 | 21.2612 | 0.001965 | 0.00184 | -152.2384 | 0.23435 | 8.4008 | 0.002446 | 0.004197 | 0.004046 | 5.8e-05 | 0.001114 | 0.33859 | 1.3098 | 2.5474 | 100.61 | 21.12 | 25.585 | 69.1169 | 80.6541 | 66.0955 | 0.006379 | 1.0407 | 0.025108 | 24.127 | -0.002008 | 0.13941 | -0.000173 | -0.013104 | -0.013104 | -0.43938
2019-03-28 15:57 | 1553803027.2923 | 1126 | 0.20953 | 0.00045 | 1e-06 | 100.8465 | 100.9013 | 125.5019 | 21.2612 | 0.001929 | 0.001905 | -152.2238 | 0.23427 | 8.4019 | 0.002506 | 0.003456 | 0.003342 | 5.4e-05 | 0.001111 | 0.32587 | 1.3475 | 2.5475 | 100.61 | 21.12 | 25.577 | 68.2495 | 82.1137 | 67.6535 | 0.006377 | 1.0359 | 0.025062 | 24.1934 | -0.004855 | 0.13937 | -0.000103 | -0.007813 | -0.007813 | -0.73667
2019-03-28 15:57 | 1553803028.2912 | 1127 | 0.20954 | 0.000443 | 1e-06 | 100.8516 | 101.0446 | 125.7399 | 21.2612 | 0.002027 | 0.001801 | -152.2739 | 0.23419 | 8.4029 | 0.002443 | 0.004191 | 0.003896 | 5.8e-05 | 0.001109 | 0.34396 | 1.3278 | 2.5474 | 100.61 | 21.12 | 25.572 | 70.6046 | 80.5892 | 65.173 | 0.006375 | 1.0435 | 0.025178 | 24.128 | -0.017105 | 0.13932 | -0.000162 | -0.012286 | -0.012286 | -0.46832
2019-03-28 15:57 | 1553803029.2923 | 1128 | 0.20954 | 0.000446 | 0 | 100.8495 | 100.8768 | 125.9559 | 21.2612 | 0.002024 | 0.001836 | -152.2768 | 0.23509 | 8.4047 | 0.002566 | 0.004289 | 0.00412 | 5.6e-05 | 0.001101 | 0.33339 | 1.301 | 2.5473 | 100.61 | 21.12 | 25.567 | 70.5173 | 83.5459 | 66.012 | 0.006374 | 1.0409 | 0.025168 | 24.1787 | -0.002422 | 0.13985 | -0.000166 | -0.012564 | -0.012564 | -0.45972
2019-03-28 15:57 | 1553803030.2914 | 1129 | 0.20953 | 0.000443 | 0 | 100.8203 | 100.8952 | 126.0789 | 21.2612 | 0.001916 | 0.001855 | -152.3 | 0.23436 | 8.402 | 0.002435 | 0.003195 | 0.003015 | 6e-05 | 0.001097 | 0.32731 | 1.3105 | 2.5474 | 100.61 | 21.12 | 25.558 | 67.9243 | 80.4053 | 66.4529 | 0.006372 | 1.0396 | 0.025164 | 24.2062 | -0.006638 | 0.13942 | -0.000117 | -0.008894 | -0.008894 | -0.64745
2019-03-28 15:57 | 1553803031.2914 | 1130 | 0.20953 | 0.000448 | 1e-06 | 100.8555 | 100.7848 | 126.2136 | 21.2612 | 0.001949 | 0.001858 | -152.3152 | 0.23502 | 8.4016 | 0.00247 | 0.003775 | 0.003568 | 5.7e-05 | 0.001092 | 0.33516 | 1.3067 | 2.5474 | 100.61 | 21.12 | 25.552 | 68.7301 | 81.2506 | 66.527 | 0.00637 | 1.0394 | 0.025175 | 24.2218 | 0.006273 | 0.13981 | -9.3e-05 | -0.00712 | -0.00712 | -0.81096
2019-03-28 15:57 | 1553803032.2914 | 1131 | 0.20954 | 0.000446 | 0 | 100.7995 | 100.8065 | 126.3284 | 21.2612 | 0.00196 | 0.001842 | -152.3199 | 0.23388 | 8.4024 | 0.002495 | 0.00408 | 0.003884 | 6e-05 | 0.001095 | 0.33019 | 1.3178 | 2.5475 | 100.61 | 21.12 | 25.539 | 68.982 | 81.8402 | 66.1501 | 0.006367 | 1.0405 | 0.0252 | 24.2193 | -0.000624 | 0.13914 | -0.000145 | -0.01106 | -0.01106 | -0.51959
2019-03-28 15:57 | 1553803033.2925 | 1132 | 0.20953 | 0.00045 | 0 | 100.8483 | 100.9111 | 126.5089 | 21.2612 | 0.002006 | 0.001832 | -152.3175 | 0.23433 | 8.4034 | 0.002472 | 0.00404 | 0.003836 | 5.9e-05 | 0.001093 | 0.33588 | 1.3039 | 2.5475 | 100.61 | 21.12 | 25.533 | 70.0862 | 81.2826 | 65.9043 | 0.006366 | 1.0413 | 0.025228 | 24.2278 | -0.00557 | 0.13941 | -0.000118 | -0.008979 | -0.008979 | -0.64124
2019-03-28 15:57 | 1553803034.2929 | 1133 | 0.20953 | 0.00045 | 1e-06 | 100.868 | 100.9278 | 126.4138 | 21.2612 | 0.001959 | 0.001817 | -152.3204 | 0.23466 | 8.4001 | 0.002462 | 0.003565 | 0.003422 | 5.6e-05 | 0.001093 | 0.34683 | 1.327 | 2.5474 | 100.61 | 21.12 | 25.527 | 68.9647 | 81.0384 | 65.5586 | 0.006364 | 1.0423 | 0.025231 | 24.2064 | -0.005302 | 0.1396 | -0.00012 | -0.009145 | -0.009145 | -0.63046
2019-03-28 15:57 | 1553803035.2925 | 1134 | 0.20953 | 0.000446 | 0 | 100.8978 | 100.8503 | 126.2972 | 21.2612 | 0.001926 | 0.0018 | -152.3148 | 0.2341 | 8.4031 | 0.002442 | 0.0039 | 0.003733 | 5.9e-05 | 0.001095 | 0.34418 | 1.2969 | 2.5472 | 100.61 | 21.12 | 25.513 | 68.1707 | 80.5651 | 65.1348 | 0.006361 | 1.0436 | 0.025235 | 24.1801 | 0.004206 | 0.13926 | -9.4e-05 | -0.007145 | -0.007145 | -0.80503
2019-03-28 15:57 | 1553803036.2925 | 1135 | 0.20953 | 0.000454 | 0 | 100.8513 | 100.8449 | 125.865 | 21.2612 | 0.001941 | 0.00182 | -152.317 | 0.23422 | 8.3998 | 0.002416 | 0.00343 | 0.003261 | 5.6e-05 | 0.001094 | 0.31454 | 1.3168 | 2.5473 | 100.61 | 21.12 | 25.507 | 68.5349 | 79.9513 | 65.6241 | 0.006359 | 1.0421 | 0.025174 | 24.1561 | 0.000568 | 0.13934 | -8e-05 | -0.006079 | -0.006079 | -0.94671
2019-03-28 15:57 | 1553803037.2926 | 1136 | 0.20954 | 0.000447 | 0 | 100.8564 | 101.0253 | 125.7348 | 21.2612 | 0.001915 | 0.001836 | -152.3319 | 0.23372 | 8.4044 | 0.002473 | 0.004099 | 0.003804 | 6e-05 | 0.001095 | 0.34291 | 1.2804 | 2.5473 | 100.61 | 21.12 | 25.495 | 67.8993 | 81.3128 | 66.0159 | 0.006356 | 1.0409 | 0.025146 | 24.1576 | -0.014969 | 0.13904 | -0.000142 | -0.010797 | -0.010797 | -0.53184
2019-03-28 15:57 | 1553803038.2916 | 1137 | 0.20952 | 0.000447 | 0 | 100.9123 | 100.7066 | 125.7308 | 21.2612 | 0.001908 | 0.001821 | -152.3176 | 0.23383 | 8.402 | 0.002498 | 0.004267 | 0.004164 | 5.7e-05 | 0.001099 | 0.34468 | 1.3089 | 2.5475 | 100.61 | 21.12 | 25.489 | 67.7286 | 81.9166 | 65.6491 | 0.006355 | 1.042 | 0.025159 | 24.1441 | 0.018254 | 0.1391 | -3.9e-05 | -0.002941 | -0.002941 | -1.9534
2019-03-28 15:57 | 1553803039.2916 | 1138 | 0.20954 | 0.000447 | 0 | 100.8549 | 101.0484 | 125.7411 | 21.2612 | 0.001998 | 0.001802 | -152.2975 | 0.23388 | 8.4017 | 0.002467 | 0.004052 | 0.003868 | 5.8e-05 | 0.001102 | 0.33035 | 1.3491 | 2.5475 | 100.61 | 21.12 | 25.484 | 69.9083 | 81.1749 | 65.1793 | 0.006353 | 1.0435 | 0.025178 | 24.1284 | -0.017146 | 0.13914 | -0.000134 | -0.010147 | -0.010147 | -0.56634
2019-03-28 15:57 | 1553803040.2907 | 1139 | 0.20953 | 0.000449 | 0 | 100.9546 | 100.906 | 126.1239 | 21.2612 | 0.001875 | 0.001812 | -152.2906 | 0.23352 | 8.4008 | 0.002466 | 0.003236 | 0.003055 | 5.7e-05 | 0.001094 | 0.32659 | 1.3691 | 2.5475 | 100.61 | 21.12 | 25.474 | 66.9399 | 81.1383 | 65.4157 | 0.006351 | 1.0428 | 0.025207 | 24.1735 | 0.004308 | 0.13892 | -9.8e-05 | -0.007472 | -0.007472 | -0.7679
2019-03-28 15:57 | 1553803041.2917 | 1140 | 0.20954 | 0.000447 | 0 | 100.8441 | 100.8928 | 126.4375 | 21.2612 | 0.001872 | 0.001805 | -152.2756 | 0.2345 | 8.4015 | 0.002497 | 0.003447 | 0.0033 | 5.6e-05 | 0.001094 | 0.32278 | 1.347 | 2.5474 | 100.61 | 21.12 | 25.468 | 66.8649 | 81.8837 | 65.2495 | 0.006349 | 1.0433 | 0.025245 | 24.1976 | -0.004313 | 0.1395 | -0.000151 | -0.01149 | -0.01149 | -0.50142
2019-03-28 15:57 | 1553803042.2917 | 1141 | 0.20952 | 0.000447 | 0 | 100.9261 | 101.0341 | 126.4793 | 21.2612 | 0.001898 | 0.001868 | -152.2767 | 0.23394 | 8.3999 | 0.002518 | 0.003813 | 0.003599 | 5.6e-05 | 0.001096 | 0.34844 | 1.3052 | 2.5475 | 100.61 | 21.12 | 25.463 | 67.485 | 82.3986 | 66.7713 | 0.006348 | 1.0386 | 0.025192 | 24.256 | -0.009572 | 0.13917 | -1.7e-05 | -0.001345 | -0.001345 | -4.2726
2019-03-28 15:57 | 1553803043.2922 | 1142 | 0.20952 | 0.000445 | 0 | 100.8608 | 100.96 | 126.116 | 21.2612 | 0.001937 | 0.001802 | -152.2721 | 0.23432 | 8.4023 | 0.002487 | 0.003629 | 0.003433 | 5.9e-05 | 0.0011 | 0.32659 | 1.3439 | 2.5476 | 100.61 | 21.12 | 25.457 | 68.4344 | 81.6537 | 65.1848 | 0.006347 | 1.0435 | 0.025215 | 24.1645 | -0.008795 | 0.1394 | -3.8e-05 | -0.00286 | -0.00286 | -2.0128
2019-03-28 15:57 | 1553803044.2908 | 1143 | 0.20955 | 0.000447 | 1e-06 | 101.0041 | 101.0096 | 126.1228 | 21.2612 | 0.001911 | 0.001821 | -152.2579 | 0.23411 | 8.4055 | 0.00247 | 0.003814 | 0.003603 | 6.1e-05 | 0.001107 | 0.34219 | 1.3102 | 2.5474 | 100.61 | 21.12 | 25.454 | 67.8019 | 81.2395 | 65.6508 | 0.006346 | 1.042 | 0.025198 | 24.1818 | -0.000492 | 0.13927 | -0.000196 | -0.014907 | -0.014907 | -0.38585
2019-03-28 15:57 | 1553803045.2918 | 1144 | 0.20955 | 0.000447 | 1e-06 | 100.9085 | 100.9329 | 126.1857 | 21.2612 | 0.001916 | 0.001828 | -152.274 | 0.23421 | 8.4031 | 0.002558 | 0.003909 | 0.003787 | 5.7e-05 | 0.001109 | 0.34435 | 1.313 | 2.5473 | 100.61 | 21.12 | 25.448 | 67.9329 | 83.3651 | 65.8177 | 0.006344 | 1.0415 | 0.025199 | 24.1938 | -0.002162 | 0.13933 | -0.000209 | -0.015909 | -0.015909 | -0.3617
2019-03-28 15:57 | 1553803046.2908 | 1145 | 0.20956 | 0.000446 | 0 | 100.9633 | 101.1425 | 126.687 | 21.2612 | 0.001964 | 0.001827 | -152.2811 | 0.23401 | 8.4006 | 0.002456 | 0.004173 | 0.003903 | 5.9e-05 | 0.001105 | 0.35314 | 1.3211 | 2.5475 | 100.61 | 21.12 | 25.446 | 69.081 | 80.9028 | 65.7912 | 0.006344 | 1.0416 | 0.02525 | 24.2408 | -0.015862 | 0.13921 | -0.000253 | -0.019266 | -0.019266 | -0.29842
2019-03-28 15:57 | 1553803047.2909 | 1146 | 0.20955 | 0.000446 | 0 | 100.9612 | 100.9037 | 126.8443 | 21.2612 | 0.001921 | 0.001811 | -152.277 | 0.23406 | 8.4019 | 0.002466 | 0.003649 | 0.003527 | 5.7e-05 | 0.001104 | 0.34523 | 1.3218 | 2.5474 | 100.61 | 21.12 | 25.445 | 68.0483 | 81.1391 | 65.4025 | 0.006344 | 1.0428 | 0.02528 | 24.242 | 0.005099 | 0.13924 | -0.000178 | -0.013527 | -0.013527 | -0.42512
2019-03-28 15:57 | 1553803048.2919 | 1147 | 0.20954 | 0.000448 | 0 | 100.9362 | 100.8904 | 126.9042 | 21.2612 | 0.001931 | 0.001859 | -152.2826 | 0.23416 | 8.4022 | 0.002551 | 0.003791 | 0.00365 | 5.7e-05 | 0.001102 | 0.32925 | 1.3126 | 2.5474 | 100.61 | 21.12 | 25.445 | 68.2832 | 83.1872 | 66.5606 | 0.006344 | 1.0393 | 0.025243 | 24.2891 | 0.004057 | 0.1393 | -0.000172 | -0.013113 | -0.013113 | -0.43874
2019-03-28 15:57 | 1553803049.291 | 1148 | 0.20954 | 0.000446 | 0 | 100.9749 | 101.0803 | 126.6603 | 21.2612 | 0.001944 | 0.001836 | -152.287 | 0.23397 | 8.402 | 0.002538 | 0.00382 | 0.003599 | 5.6e-05 | 0.001096 | 0.32565 | 1.3163 | 2.5474 | 100.61 | 21.12 | 25.443 | 68.5991 | 82.8751 | 65.9957 | 0.006343 | 1.041 | 0.025239 | 24.2456 | -0.009333 | 0.13919 | -0.000161 | -0.012284 | -0.012284 | -0.46798
2019-03-28 15:57 | 1553803050.291 | 1149 | 0.20953 | 0.000443 | 0 | 100.9606 | 101.1374 | 126.4176 | 21.2612 | 0.002009 | 0.001799 | -152.2769 | 0.23417 | 8.4022 | 0.002514 | 0.003509 | 0.003329 | 5.6e-05 | 0.001085 | 0.32637 | 1.3049 | 2.5474 | 100.61 | 21.12 | 25.444 | 70.1578 | 82.2979 | 65.109 | 0.006343 | 1.0437 | 0.025248 | 24.1907 | -0.015648 | 0.13931 | -8e-05 | -0.006126 | -0.006126 | -0.93925
2019-03-28 15:58 | 1553803051.291 | 1150 | 0.20952 | 0.000447 | 0 | 100.9481 | 101.1472 | 126.5124 | 21.2612 | 0.001916 | 0.001798 | -152.2835 | 0.23451 | 8.4041 | 0.002505 | 0.004046 | 0.00386 | 5.9e-05 | 0.001083 | 0.3366 | 1.2875 | 2.5474 | 100.61 | 21.12 | 25.444 | 67.9267 | 82.0748 | 65.0794 | 0.006343 | 1.0438 | 0.025259 | 24.1987 | -0.017628 | 0.13951 | -1e-06 | -9e-05 | -9e-05 | -64.0311
2019-03-28 15:58 | 1553803052.2911 | 1151 | 0.20952 | 0.00045 | 0 | 100.9189 | 101.0416 | 126.3624 | 21.2612 | 0.001873 | 0.001804 | -152.268 | 0.23456 | 8.4032 | 0.002544 | 0.003585 | 0.003391 | 5.5e-05 | 0.001079 | 0.33107 | 1.3075 | 2.5474 | 100.61 | 21.12 | 25.445 | 66.8985 | 83.0141 | 65.2377 | 0.006344 | 1.0433 | 0.025238 | 24.19 | -0.010867 | 0.13954 | -1.1e-05 | -0.000815 | -0.000815 | -7.0725
2019-03-28 15:58 | 1553803053.2921 | 1152 | 0.20952 | 0.000446 | 0 | 100.9573 | 101.0786 | 126.4814 | 21.2612 | 0.00194 | 0.001765 | -152.2682 | 0.23446 | 8.4007 | 0.002531 | 0.003909 | 0.003679 | 5.5e-05 | 0.00108 | 0.33931 | 1.2772 | 2.5474 | 100.61 | 21.12 | 25.446 | 68.5092 | 82.7152 | 64.2979 | 0.006344 | 1.0462 | 0.025285 | 24.1678 | -0.010739 | 0.13948 | -4.3e-05 | -0.003275 | -0.003275 | -1.7591
2019-03-28 15:58 | 1553803054.2911 | 1153 | 0.20952 | 0.000443 | 1e-06 | 100.8829 | 101.046 | 126.5967 | 21.2612 | 0.00189 | 0.001791 | -152.3064 | 0.23448 | 8.4031 | 0.002486 | 0.003923 | 0.003709 | 5.7e-05 | 0.001083 | 0.33787 | 1.294 | 2.5473 | 100.61 | 21.12 | 25.447 | 67.3056 | 81.6217 | 64.9226 | 0.006344 | 1.0443 | 0.025273 | 24.2011 | -0.014452 | 0.13949 | -3.3e-05 | -0.002582 | -0.002582 | -2.2308
2019-03-28 15:58 | 1553803055.2902 | 1154 | 0.20949 | 0.000447 | 0 | 100.9415 | 100.9913 | 126.7929 | 21.2612 | 0.001916 | 0.001875 | -152.2945 | 0.23384 | 8.4043 | 0.002494 | 0.003573 | 0.003476 | 5.8e-05 | 0.001084 | 0.33284 | 1.3085 | 2.5473 | 100.61 | 21.12 | 25.45 | 67.9281 | 81.8229 | 66.9545 | 0.006345 | 1.038 | 0.025217 | 24.2926 | -0.004409 | 0.13911 | 0.000162 | 0.012279 | 0.012279 | 0.4679
2019-03-28 15:58 | 1553803056.2902 | 1155 | 0.20949 | 0.000451 | 0 | 101.0374 | 101.0868 | 126.4047 | 21.2612 | 0.001915 | 0.001838 | -152.2987 | 0.2345 | 8.4019 | 0.002517 | 0.003548 | 0.003414 | 5.6e-05 | 0.001085 | 0.31222 | 1.3005 | 2.5474 | 100.61 | 21.12 | 25.452 | 67.9094 | 82.3813 | 66.0486 | 0.006345 | 1.0408 | 0.025212 | 24.223 | -0.004365 | 0.1395 | 0.000163 | 0.012374 | 0.012374 | 0.4656
2019-03-28 15:58 | 1553803057.2916 | 1156 | 0.20952 | 0.000448 | 1e-06 | 100.9835 | 101.0762 | 126.0756 | 21.2612 | 0.001904 | 0.001773 | -152.286 | 0.23375 | 8.4027 | 0.00254 | 0.003626 | 0.003381 | 5.8e-05 | 0.001087 | 0.3183 | 1.3015 | 2.5474 | 100.61 | 21.12 | 25.456 | 67.6317 | 82.9265 | 64.483 | 0.006346 | 1.0456 | 0.025237 | 24.1356 | -0.008207 | 0.13906 | -3.5e-05 | -0.002717 | -0.002717 | -2.1139
2019-03-28 15:58 | 1553803058.2902 | 1157 | 0.20949 | 0.000443 | 1e-06 | 100.9463 | 101.0402 | 126.0833 | 21.2612 | 0.001896 | 0.001835 | -152.2697 | 0.2347 | 8.4017 | 0.00252 | 0.003947 | 0.003739 | 5.9e-05 | 0.001091 | 0.31078 | 1.293 | 2.5475 | 100.61 | 21.12 | 25.458 | 67.4376 | 82.4383 | 65.9902 | 0.006347 | 1.041 | 0.025182 | 24.1901 | -0.008318 | 0.13962 | 0.000127 | 0.0096 | 0.0096 | 0.60066
2019-03-28 15:58 | 1553803059.2907 | 1158 | 0.20952 | 0.000447 | 0 | 101.0011 | 101.2522 | 126.3425 | 21.2612 | 0.001913 | 0.00184 | -152.279 | 0.23488 | 8.4031 | 0.002581 | 0.004438 | 0.004309 | 5.5e-05 | 0.001092 | 0.33804 | 1.3026 | 2.5474 | 100.61 | 21.12 | 25.459 | 67.868 | 83.9018 | 66.1063 | 0.006347 | 1.0406 | 0.025203 | 24.2191 | -0.02221 | 0.13973 | -3.8e-05 | -0.002897 | -0.002897 | -1.9921
2019-03-28 15:58 | 1553803060.2903 | 1159 | 0.20951 | 0.000446 | 1e-06 | 101.0497 | 100.9716 | 126.0845 | 21.2612 | 0.00188 | 0.001823 | -152.2887 | 0.23379 | 8.3997 | 0.002471 | 0.003422 | 0.003244 | 5.7e-05 | 0.001096 | 0.3366 | 1.283 | 2.5473 | 100.61 | 21.12 | 25.46 | 67.0623 | 81.2756 | 65.6912 | 0.006347 | 1.0419 | 0.025193 | 24.1796 | 0.006912 | 0.13908 | 5.9e-05 | 0.004406 | 0.004406 | 1.3037
2019-03-28 15:58 | 1553803061.2903 | 1160 | 0.20952 | 0.000445 | 0 | 100.9871 | 100.9342 | 125.888 | 21.2612 | 0.00193 | 0.001766 | -152.2796 | 0.23407 | 8.4026 | 0.002573 | 0.003747 | 0.003592 | 5.7e-05 | 0.001097 | 0.33035 | 1.2901 | 2.5474 | 100.61 | 21.12 | 25.461 | 68.2558 | 83.7214 | 64.3112 | 0.006348 | 1.0462 | 0.025225 | 24.1115 | 0.004685 | 0.13925 | -3.9e-05 | -0.002969 | -0.002969 | -1.9372
2019-03-28 15:58 | 1553803062.2904 | 1161 | 0.20952 | 0.000449 | 1e-06 | 101.0226 | 100.9546 | 126.1108 | 21.2612 | 0.001933 | 0.001889 | -152.2949 | 0.23409 | 8.4026 | 0.002512 | 0.004054 | 0.003878 | 5.9e-05 | 0.001098 | 0.31349 | 1.2958 | 2.5474 | 100.61 | 21.12 | 25.46 | 68.3253 | 82.2503 | 67.2768 | 0.006347 | 1.0371 | 0.025137 | 24.2386 | 0.006017 | 0.13926 | -3.4e-05 | -0.002631 | -0.002631 | -2.186
2019-03-28 15:58 | 1553803063.2914 | 1162 | 0.20954 | 0.000447 | 0 | 101.0884 | 101.293 | 126.3033 | 21.2612 | 0.001866 | 0.001851 | -152.307 | 0.23351 | 8.4004 | 0.002533 | 0.003607 | 0.003446 | 5.9e-05 | 0.001103 | 0.33157 | 1.328 | 2.5474 | 100.61 | 21.13 | 25.458 | 66.7198 | 82.7534 | 66.3757 | 0.006351 | 1.0398 | 0.02519 | 24.225 | -0.018085 | 0.13891 | -0.000125 | -0.009495 | -0.009495 | -0.60421
2019-03-28 15:58 | 1553803064.2905 | 1163 | 0.20954 | 0.000448 | 1e-06 | 101.0681 | 101.0463 | 126.0643 | 21.2612 | 0.001925 | 0.001807 | -152.2864 | 0.23427 | 8.4025 | 0.002572 | 0.003838 | 0.003676 | 5.3e-05 | 0.001101 | 0.32244 | 1.2823 | 2.5474 | 100.61 | 21.13 | 25.457 | 68.1466 | 83.7005 | 65.2986 | 0.006351 | 1.0431 | 0.025206 | 24.1636 | 0.001928 | 0.13937 | -0.000146 | -0.011147 | -0.011147 | -0.51635
2019-03-28 15:58 | 1553803065.2908 | 1164 | 0.20953 | 0.000445 | 0 | 100.9877 | 100.9679 | 125.8744 | 21.2612 | 0.001971 | 0.001842 | -152.2711 | 0.23438 | 8.4042 | 0.002532 | 0.00421 | 0.003995 | 6e-05 | 0.001104 | 0.31885 | 1.2688 | 2.5473 | 100.61 | 21.13 | 25.452 | 69.2541 | 82.7246 | 66.1431 | 0.006349 | 1.0405 | 0.025155 | 24.1755 | 0.001758 | 0.13943 | -0.000115 | -0.008717 | -0.008717 | -0.6606
2019-03-28 15:58 | 1553803066.2915 | 1165 | 0.20949 | 0.000447 | 0 | 101.0303 | 101.0501 | 126.0838 | 21.2612 | 0.001879 | 0.001836 | -152.2738 | 0.23414 | 8.4029 | 0.002496 | 0.00421 | 0.004107 | 6e-05 | 0.001099 | 0.30453 | 1.287 | 2.5472 | 100.61 | 21.13 | 25.447 | 67.0279 | 81.8597 | 66.001 | 0.006348 | 1.041 | 0.025182 | 24.1905 | -0.001751 | 0.13929 | 0.000126 | 0.009547 | 0.009547 | 0.60253
2019-03-28 15:58 | 1553803067.2916 | 1166 | 0.20954 | 0.000445 | 1e-06 | 101.022 | 101.2288 | 126.2336 | 21.2612 | 0.001962 | 0.001807 | -152.2793 | 0.23409 | 8.401 | 0.002488 | 0.003575 | 0.003449 | 5.6e-05 | 0.0011 | 0.33964 | 1.3084 | 2.5474 | 100.61 | 21.13 | 25.437 | 69.0428 | 81.6678 | 65.3027 | 0.006346 | 1.0431 | 0.025223 | 24.18 | -0.018293 | 0.13926 | -0.000159 | -0.012091 | -0.012091 | -0.47566
2019-03-28 15:58 | 1553803068.2906 | 1167 | 0.20951 | 0.000443 | 0 | 101.0794 | 100.8469 | 126.2376 | 21.2612 | 0.00194 | 0.001838 | -152.2848 | 0.23377 | 8.4033 | 0.002527 | 0.0035 | 0.003289 | 5.4e-05 | 0.001092 | 0.30757 | 1.358 | 2.5474 | 100.61 | 21.13 | 25.432 | 68.4998 | 82.6037 | 66.0573 | 0.006344 | 1.0408 | 0.025195 | 24.2073 | 0.0206 | 0.13907 | 3.4e-05 | 0.002592 | 0.002592 | 2.2158
2019-03-28 15:58 | 1553803069.291 | 1168 | 0.20951 | 0.000446 | 0 | 101.0032 | 101.0266 | 126.1176 | 21.2612 | 0.001904 | 0.001812 | -152.2779 | 0.23457 | 8.404 | 0.002478 | 0.00347 | 0.003277 | 5.7e-05 | 0.001095 | 0.33948 | 1.3908 | 2.5475 | 100.61 | 21.13 | 25.425 | 67.6389 | 81.4438 | 65.4237 | 0.006343 | 1.0427 | 0.025206 | 24.1732 | -0.002076 | 0.13954 | 5.3e-05 | 0.003991 | 0.003991 | 1.4439
2019-03-28 15:58 | 1553803070.2897 | 1169 | 0.20949 | 0.000444 | 0 | 101.0774 | 101.1568 | 126.6279 | 21.2612 | 0.001942 | 0.00179 | -152.2885 | 0.23403 | 8.4015 | 0.002456 | 0.003417 | 0.003219 | 5.8e-05 | 0.001085 | 0.33572 | 1.338 | 2.5473 | 100.61 | 21.13 | 25.411 | 68.5592 | 80.9143 | 64.8975 | 0.006339 | 1.0444 | 0.025277 | 24.2032 | -0.007023 | 0.13923 | 0.000146 | 0.011092 | 0.011092 | 0.51841
2019-03-28 15:58 | 1553803071.2901 | 1170 | 0.2095 | 0.000441 | 1e-06 | 101.0297 | 101.0769 | 126.43 | 21.2612 | 0.001917 | 0.001774 | -152.3038 | 0.23417 | 8.4013 | 0.002511 | 0.003663 | 0.003436 | 5.2e-05 | 0.001083 | 0.33588 | 1.3188 | 2.5473 | 100.61 | 21.13 | 25.403 | 67.9507 | 82.2356 | 64.5104 | 0.006337 | 1.0456 | 0.025272 | 24.1704 | -0.004179 | 0.13931 | 0.000126 | 0.009578 | 0.009578 | 0.6007
2019-03-28 15:58 | 1553803072.2907 | 1171 | 0.20947 | 0.000447 | 1e-06 | 101.0419 | 101.2709 | 126.2333 | 21.2612 | 0.001932 | 0.001857 | -152.2933 | 0.2332 | 8.4031 | 0.002466 | 0.003585 | 0.00344 | 5.4e-05 | 0.001082 | 0.34202 | 1.3003 | 2.5473 | 100.61 | 21.13 | 25.395 | 68.3027 | 81.1429 | 66.5154 | 0.006335 | 1.0394 | 0.025177 | 24.2232 | -0.020247 | 0.13873 | 0.000274 | 0.020733 | 0.020733 | 0.27635
2019-03-28 15:58 | 1553803073.2897 | 1172 | 0.20947 | 0.000444 | 0 | 101.0464 | 100.9964 | 126.3075 | 21.2612 | 0.001875 | 0.001844 | -152.2877 | 0.2342 | 8.4036 | 0.002546 | 0.003833 | 0.003669 | 5.4e-05 | 0.001083 | 0.33157 | 1.2881 | 2.5473 | 100.61 | 21.13 | 25.378 | 66.9459 | 83.0749 | 66.1915 | 0.006331 | 1.0404 | 0.025197 | 24.2188 | 0.004425 | 0.13932 | 0.000279 | 0.021135 | 0.021135 | 0.27226
2019-03-28 15:58 | 1553803074.2898 | 1173 | 0.20949 | 0.00045 | 0 | 101.1453 | 101.0973 | 126.4742 | 21.2612 | 0.001915 | 0.001761 | -152.2929 | 0.23309 | 8.4026 | 0.00249 | 0.003448 | 0.003228 | 5.5e-05 | 0.001084 | 0.33356 | 1.2981 | 2.5474 | 100.61 | 21.13 | 25.37 | 67.9 | 81.7231 | 64.194 | 0.006329 | 1.0465 | 0.025288 | 24.1633 | 0.004243 | 0.13867 | 0.000173 | 0.013129 | 0.013129 | 0.43619
2019-03-28 15:58 | 1553803075.2908 | 1174 | 0.2095 | 0.000445 | 0 | 101.0517 | 101.2193 | 126.4436 | 21.2612 | 0.001933 | 0.001833 | -152.3174 | 0.23388 | 8.4043 | 0.002512 | 0.003679 | 0.003503 | 5.7e-05 | 0.001081 | 0.30934 | 1.3225 | 2.5474 | 100.61 | 21.14 | 25.353 | 68.3346 | 82.2582 | 65.9293 | 0.006329 | 1.0412 | 0.02522 | 24.2225 | -0.014815 | 0.13914 | 0.000112 | 0.008512 | 0.008512 | 0.67508
2019-03-28 15:58 | 1553803076.2898 | 1175 | 0.20953 | 0.000443 | 0 | 101.1149 | 101.2712 | 126.1415 | 21.2612 | 0.001925 | 0.001873 | -152.3012 | 0.23366 | 8.4026 | 0.002541 | 0.003753 | 0.003578 | 5.5e-05 | 0.001085 | 0.34252 | 1.3099 | 2.5474 | 100.61 | 21.13 | 25.346 | 68.1341 | 82.9499 | 66.8829 | 0.006323 | 1.0383 | 0.025155 | 24.2275 | -0.01382 | 0.13901 | -9.4e-05 | -0.00718 | -0.00718 | -0.79953
2019-03-28 15:58 | 1553803077.2909 | 1176 | 0.20955 | 0.000446 | 0 | 101.0908 | 101.1218 | 126.2312 | 21.2612 | 0.001912 | 0.001814 | -152.3003 | 0.2345 | 8.4028 | 0.002663 | 0.00411 | 0.003998 | 5.3e-05 | 0.001088 | 0.33644 | 1.2979 | 2.5474 | 100.61 | 21.13 | 25.34 | 67.8283 | 85.8901 | 65.4853 | 0.006321 | 1.0426 | 0.025215 | 24.1863 | -0.002742 | 0.1395 | -0.000212 | -0.016152 | -0.016152 | -0.3567
2019-03-28 15:58 | 1553803078.2909 | 1177 | 0.20955 | 0.000443 | 0 | 101.1086 | 101.2447 | 126.2629 | 21.2612 | 0.001954 | 0.001843 | -152.2912 | 0.23459 | 8.4031 | 0.00254 | 0.003892 | 0.00374 | 5.7e-05 | 0.001087 | 0.32781 | 1.332 | 2.5474 | 100.61 | 21.13 | 25.329 | 68.8448 | 82.9203 | 66.1727 | 0.006319 | 1.0404 | 0.025193 | 24.2139 | -0.012032 | 0.13956 | -0.000235 | -0.01783 | -0.01783 | -0.32327
2019-03-28 15:58 | 1553803079.2909 | 1178 | 0.20956 | 0.000447 | 0 | 101.0514 | 101.1398 | 126.212 | 21.2612 | 0.001909 | 0.001816 | -152.2855 | 0.2347 | 8.4045 | 0.002561 | 0.003843 | 0.003665 | 5.5e-05 | 0.00109 | 0.32101 | 1.3595 | 2.5475 | 100.61 | 21.13 | 25.325 | 67.7536 | 83.4266 | 65.5235 | 0.006318 | 1.0424 | 0.025212 | 24.1858 | -0.007814 | 0.13962 | -0.000258 | -0.019617 | -0.019617 | -0.29396
2019-03-28 15:58 | 1553803080.2909 | 1179 | 0.20948 | 0.000445 | 1e-06 | 101.0234 | 101.2576 | 125.9521 | 21.2612 | 0.001909 | 0.001835 | -152.2823 | 0.23522 | 8.4024 | 0.002508 | 0.004044 | 0.003825 | 6e-05 | 0.001095 | 0.33964 | 1.3099 | 2.5474 | 100.61 | 21.14 | 25.32 | 67.7689 | 82.1529 | 65.9738 | 0.00632 | 1.0411 | 0.025169 | 24.1769 | -0.020712 | 0.13993 | 0.000214 | 0.016136 | 0.016136 | 0.35814
2019-03-28 15:58 | 1553803081.2899 | 1180 | 0.20954 | 0.000443 | 1e-06 | 101.1086 | 101.1751 | 126.1092 | 21.2612 | 0.001931 | 0.001819 | -152.2758 | 0.23459 | 8.4019 | 0.002532 | 0.004376 | 0.004167 | 5.7e-05 | 0.001093 | 0.31614 | 1.3523 | 2.5473 | 100.61 | 21.14 | 25.318 | 68.2846 | 82.7308 | 65.5945 | 0.00632 | 1.0422 | 0.025199 | 24.1785 | -0.005878 | 0.13956 | -0.000149 | -0.011332 | -0.011332 | -0.50861
2019-03-28 15:58 | 1553803082.29 | 1181 | 0.20953 | 0.000446 | 0 | 101.1092 | 101.2291 | 126.0042 | 21.2612 | 0.00188 | 0.001828 | -152.267 | 0.23402 | 8.4035 | 0.002547 | 0.003818 | 0.003651 | 5.8e-05 | 0.001096 | 0.33627 | 1.3408 | 2.5474 | 100.61 | 21.14 | 25.318 | 67.0606 | 83.0857 | 65.8146 | 0.00632 | 1.0415 | 0.025181 | 24.1763 | -0.010599 | 0.13922 | -5.9e-05 | -0.004497 | -0.004497 | -1.2784
2019-03-28 15:58 | 1553803083.2911 | 1182 | 0.20954 | 0.000447 | 0 | 101.1063 | 101.1316 | 125.7357 | 21.2612 | 0.00188 | 0.001843 | -152.2667 | 0.23461 | 8.4006 | 0.00259 | 0.004056 | 0.003824 | 5.3e-05 | 0.001099 | 0.33284 | 1.3092 | 2.5475 | 100.61 | 21.14 | 25.32 | 67.07 | 84.117 | 66.1633 | 0.00632 | 1.0405 | 0.025141 | 24.1629 | -0.002243 | 0.13957 | -0.000147 | -0.011176 | -0.011176 | -0.51579
2019-03-28 15:58 | 1553803084.2911 | 1183 | 0.20953 | 0.000445 | 0 | 101.0089 | 101.3371 | 125.65 | 21.2612 | 0.001953 | 0.001842 | -152.2525 | 0.2342 | 8.4023 | 0.002578 | 0.004013 | 0.003825 | 5.7e-05 | 0.001101 | 0.31365 | 1.3315 | 2.5473 | 100.61 | 21.14 | 25.321 | 68.8135 | 83.8267 | 66.1438 | 0.006321 | 1.0405 | 0.025133 | 24.154 | -0.029024 | 0.13932 | -7.1e-05 | -0.005392 | -0.005392 | -1.0671
2019-03-28 15:58 | 1553803085.2901 | 1184 | 0.2095 | 0.000442 | 0 | 101.1164 | 101.0192 | 125.7604 | 21.2612 | 0.001903 | 0.001792 | -152.2358 | 0.23403 | 8.4026 | 0.002524 | 0.004229 | 0.004082 | 5.4e-05 | 0.001099 | 0.30486 | 1.3296 | 2.5473 | 100.61 | 21.13 | 25.327 | 67.624 | 82.539 | 64.9358 | 0.006318 | 1.0442 | 0.025189 | 24.1215 | 0.008604 | 0.13923 | 0.000125 | 0.009484 | 0.009484 | 0.6063
2019-03-28 15:58 | 1553803086.2891 | 1185 | 0.20948 | 0.000443 | 0 | 101.0759 | 101.0759 | 126.0014 | 21.2612 | 0.001918 | 0.001797 | -152.2567 | 0.23418 | 8.4036 | 0.002561 | 0.003472 | 0.003233 | 5.8e-05 | 0.001097 | 0.32731 | 1.3372 | 2.5474 | 100.61 | 21.13 | 25.331 | 67.9687 | 83.4173 | 65.0607 | 0.006319 | 1.0439 | 0.025208 | 24.1491 | -2e-06 | 0.13932 | 0.000245 | 0.018644 | 0.018644 | 0.30861
2019-03-28 15:58 | 1553803087.2902 | 1186 | 0.2095 | 0.000444 | 0 | 101.1611 | 101.2821 | 125.8571 | 21.2612 | 0.001869 | 0.001804 | -152.2532 | 0.23431 | 8.4028 | 0.002583 | 0.003787 | 0.003639 | 5.7e-05 | 0.001098 | 0.30342 | 1.3338 | 2.5475 | 100.61 | 21.13 | 25.335 | 66.8047 | 83.9689 | 65.2339 | 0.00632 | 1.0433 | 0.025187 | 24.1414 | -0.010697 | 0.13939 | 0.000109 | 0.008236 | 0.008236 | 0.699
2019-03-28 15:58 | 1553803088.2892 | 1187 | 0.20951 | 0.000441 | 0 | 101.0961 | 101.1306 | 126.0441 | 21.2612 | 0.001897 | 0.001838 | -152.2688 | 0.23423 | 8.4024 | 0.002554 | 0.004468 | 0.004231 | 5.8e-05 | 0.0011 | 0.32084 | 1.3297 | 2.5474 | 100.61 | 21.13 | 25.34 | 67.4624 | 83.2494 | 66.0416 | 0.006321 | 1.0408 | 0.025176 | 24.1882 | -0.003049 | 0.13935 | 6.4e-05 | 0.004852 | 0.004852 | 1.1862
2019-03-28 15:58 | 1553803089.2892 | 1188 | 0.20952 | 0.000447 | 1e-06 | 101.136 | 101.2797 | 125.9338 | 21.2612 | 0.001918 | 0.001823 | -152.2717 | 0.23388 | 8.4052 | 0.002531 | 0.00411 | 0.003925 | 5.9e-05 | 0.001103 | 0.33843 | 1.2787 | 2.5473 | 100.61 | 21.13 | 25.344 | 67.9663 | 82.702 | 65.6967 | 0.006322 | 1.0419 | 0.025178 | 24.1653 | -0.012699 | 0.13914 | -1.1e-05 | -0.000926 | -0.000926 | -6.2089
2019-03-28 15:58 | 1553803090.2903 | 1189 | 0.20954 | 0.000448 | 0 | 101.2034 | 101.1883 | 125.9756 | 21.2612 | 0.001919 | 0.001816 | -152.2656 | 0.23397 | 8.4033 | 0.002536 | 0.004294 | 0.004137 | 6.1e-05 | 0.001107 | 0.3298 | 1.3061 | 2.5475 | 100.61 | 21.13 | 25.352 | 67.9976 | 82.8212 | 65.5235 | 0.006324 | 1.0424 | 0.025189 | 24.1631 | 0.001326 | 0.13919 | -0.000167 | -0.012684 | -0.012684 | -0.4532
2019-03-28 15:58 | 1553803091.2893 | 1190 | 0.20952 | 0.000448 | 0 | 101.114 | 101.3015 | 126.066 | 21.2612 | 0.001912 | 0.001871 | -152.2558 | 0.23505 | 8.401 | 0.002519 | 0.003925 | 0.003736 | 5.7e-05 | 0.001107 | 0.35546 | 1.3147 | 2.5473 | 100.61 | 21.13 | 25.355 | 67.8221 | 82.4265 | 66.8415 | 0.006325 | 1.0384 | 0.025149 | 24.2188 | -0.016569 | 0.13983 | -2.8e-05 | -0.002117 | -0.002117 | -2.7276
2019-03-28 15:58 | 1553803092.2903 | 1191 | 0.20956 | 0.000445 | 0 | 101.0908 | 101.1554 | 126.2317 | 21.2612 | 0.001943 | 0.001837 | -152.275 | 0.23474 | 8.403 | 0.002488 | 0.003844 | 0.003645 | 5.8e-05 | 0.001111 | 0.33411 | 1.3563 | 2.5475 | 100.61 | 21.13 | 25.359 | 68.5803 | 81.6616 | 66.0253 | 0.006326 | 1.0409 | 0.025195 | 24.2056 | -0.005717 | 0.13965 | -0.000265 | -0.020079 | -0.020079 | -0.28724
2019-03-28 15:58 | 1553803093.2894 | 1192 | 0.20955 | 0.000442 | 1e-06 | 101.1313 | 101.2539 | 126.5108 | 21.2612 | 0.001953 | 0.001823 | -152.2683 | 0.23424 | 8.4014 | 0.002525 | 0.003515 | 0.003314 | 6.1e-05 | 0.001112 | 0.33395 | 1.3284 | 2.5474 | 100.61 | 21.13 | 25.367 | 68.8135 | 82.5669 | 65.696 | 0.006328 | 1.0419 | 0.025236 | 24.2206 | -0.01084 | 0.13935 | -0.000217 | -0.016509 | -0.016509 | -0.34861
2019-03-28 15:58 | 1553803094.2898 | 1193 | 0.20956 | 0.000445 | 1e-06 | 101.0991 | 101.2023 | 126.5845 | 21.2612 | 0.001915 | 0.001773 | -152.2831 | 0.23423 | 8.4037 | 0.00265 | 0.003962 | 0.003863 | 6e-05 | 0.001108 | 0.3413 | 1.3322 | 2.5474 | 100.61 | 21.13 | 25.372 | 67.894 | 85.5771 | 64.4924 | 0.006329 | 1.0456 | 0.025288 | 24.1846 | -0.009123 | 0.13934 | -0.000272 | -0.020793 | -0.020793 | -0.27676
2019-03-28 15:58 | 1553803095.2904 | 1194 | 0.20957 | 0.000444 | 0 | 101.1337 | 101.1479 | 126.3606 | 21.2612 | 0.001963 | 0.00182 | -152.2838 | 0.23444 | 8.4028 | 0.002596 | 0.003776 | 0.003531 | 5.7e-05 | 0.001106 | 0.32621 | 1.314 | 2.5473 | 100.61 | 21.13 | 25.375 | 69.0693 | 84.2654 | 65.6233 | 0.00633 | 1.0421 | 0.025223 | 24.2036 | -0.001261 | 0.13947 | -0.000291 | -0.022104 | -0.022104 | -0.2606
2019-03-28 15:58 | 1553803096.2895 | 1195 | 0.20951 | 0.000441 | 1e-06 | 101.1762 | 101.2278 | 126.3458 | 21.2612 | 0.001893 | 0.001791 | -152.2607 | 0.23406 | 8.403 | 0.002461 | 0.00398 | 0.003856 | 6.1e-05 | 0.0011 | 0.33157 | 1.2803 | 2.5474 | 100.61 | 21.13 | 25.381 | 67.3821 | 81.0338 | 64.9334 | 0.006332 | 1.0443 | 0.025248 | 24.1775 | -0.004552 | 0.13924 | 2.3e-05 | 0.001691 | 0.001691 | 3.4001
2019-03-28 15:58 | 1553803097.2895 | 1196 | 0.20955 | 0.000446 | 1e-06 | 101.1024 | 101.1968 | 126.465 | 21.2612 | 0.001977 | 0.001838 | -152.2692 | 0.23401 | 8.4025 | 0.002495 | 0.004547 | 0.004393 | 5.7e-05 | 0.001105 | 0.31725 | 1.3101 | 2.5473 | 100.61 | 21.13 | 25.385 | 69.3865 | 81.8424 | 66.044 | 0.006333 | 1.0408 | 0.025218 | 24.2286 | -0.008353 | 0.13921 | -0.000196 | -0.014965 | -0.014965 | -0.3842
2019-03-28 15:58 | 1553803098.2905 | 1197 | 0.20952 | 0.000445 | 0 | 101.1485 | 101.1058 | 126.4453 | 21.2612 | 0.00194 | 0.001827 | -152.2597 | 0.23444 | 8.4055 | 0.002503 | 0.004131 | 0.003873 | 5.9e-05 | 0.001105 | 0.32333 | 1.3457 | 2.5474 | 100.61 | 21.13 | 25.391 | 68.5147 | 82.0296 | 65.7826 | 0.006334 | 1.0416 | 0.025226 | 24.2174 | 0.003782 | 0.13947 | -5.2e-05 | -0.003959 | -0.003959 | -1.4547
2019-03-28 15:58 | 1553803099.2905 | 1198 | 0.20953 | 0.000443 | 0 | 101.0926 | 101.0358 | 126.546 | 21.2612 | 0.001915 | 0.001762 | -152.2537 | 0.23446 | 8.403 | 0.002509 | 0.00411 | 0.00402 | 5.7e-05 | 0.001107 | 0.33196 | 1.3042 | 2.5473 | 100.61 | 21.13 | 25.395 | 67.8971 | 82.1871 | 64.2231 | 0.006335 | 1.0465 | 0.025294 | 24.1712 | 0.005022 | 0.13948 | -7.2e-05 | -0.005541 | -0.005541 | -1.0397
2019-03-28 15:58 | 1553803100.2896 | 1199 | 0.20951 | 0.000446 | 0 | 101.125 | 101.0895 | 126.5619 | 21.2612 | 0.001919 | 0.001796 | -152.2491 | 0.23345 | 8.4011 | 0.002482 | 0.004564 | 0.004357 | 5.6e-05 | 0.001109 | 0.33019 | 1.2998 | 2.5474 | 100.61 | 21.13 | 25.399 | 68.0053 | 81.5244 | 65.0349 | 0.006336 | 1.0439 | 0.025265 | 24.2018 | 0.003143 | 0.13888 | 5e-05 | 0.003817 | 0.003817 | 1.5028
2019-03-28 15:58 | 1553803101.2906 | 1200 | 0.20952 | 0.000445 | 0 | 101.1191 | 101.1452 | 126.4354 | 21.2612 | 0.001932 | 0.001794 | -152.2786 | 0.23332 | 8.403 | 0.002488 | 0.003827 | 0.003617 | 5.7e-05 | 0.001107 | 0.31022 | 1.2674 | 2.5474 | 100.61 | 21.13 | 25.404 | 68.3034 | 81.6801 | 65.0012 | 0.006337 | 1.044 | 0.025254 | 24.1885 | -0.002312 | 0.13881 | 1.1e-05 | 0.000852 | 0.000852 | 6.7249
2019-03-28 15:58 | 1553803102.29 | 1201 | 0.20952 | 0.000444 | 1e-06 | 101.114 | 101.1809 | 126.3744 | 21.2612 | 0.00191 | 0.001848 | -152.2727 | 0.23406 | 8.4015 | 0.002559 | 0.003554 | 0.003383 | 5.7e-05 | 0.001101 | 0.32211 | 1.258 | 2.5472 | 100.61 | 21.13 | 25.409 | 67.7762 | 83.3899 | 66.285 | 0.006339 | 1.0401 | 0.0252 | 24.2285 | -0.005914 | 0.13924 | 8e-06 | 0.000553 | 0.000553 | 10.3903
2019-03-28 15:58 | 1553803103.2887 | 1202 | 0.20952 | 0.000443 | 0 | 101.1021 | 101.0953 | 126.1596 | 21.2612 | 0.001919 | 0.0018 | -152.2907 | 0.23406 | 8.4035 | 0.002549 | 0.003948 | 0.003708 | 5.7e-05 | 0.001102 | 0.29535 | 1.2784 | 2.5475 | 100.61 | 21.13 | 25.42 | 67.9976 | 83.1396 | 65.1292 | 0.006341 | 1.0437 | 0.025222 | 24.1667 | 0.000603 | 0.13924 | -5e-06 | -0.000394 | -0.000394 | -14.5963
2019-03-28 15:58 | 1553803104.2898 | 1203 | 0.20956 | 0.000442 | 1e-06 | 101.1679 | 101.2203 | 125.9662 | 21.2612 | 0.001921 | 0.001841 | -152.3011 | 0.23384 | 8.4031 | 0.002557 | 0.003552 | 0.00344 | 5.8e-05 | 0.001106 | 0.33483 | 1.2971 | 2.5475 | 100.61 | 21.13 | 25.425 | 68.049 | 83.3213 | 66.1362 | 0.006343 | 1.0406 | 0.025165 | 24.1841 | -0.004629 | 0.13911 | -0.00026 | -0.01977 | -0.01977 | -0.2906
2019-03-28 15:58 | 1553803105.2888 | 1204 | 0.20953 | 0.000446 | 1e-06 | 101.1265 | 101.1782 | 126.0824 | 21.2612 | 0.001958 | 0.001819 | -152.2978 | 0.23378 | 8.4021 | 0.002607 | 0.003628 | 0.003424 | 5.5e-05 | 0.001101 | 0.33627 | 1.2847 | 2.5474 | 100.61 | 21.13 | 25.43 | 68.9272 | 84.5369 | 65.5844 | 0.006344 | 1.0422 | 0.025197 | 24.1756 | -0.004567 | 0.13908 | -9.2e-05 | -0.007005 | -0.007005 | -0.82
2019-03-28 15:58 | 1553803106.2888 | 1205 | 0.20949 | 0.000446 | 0 | 101.0598 | 101.3514 | 126.2516 | 21.2612 | 0.001899 | 0.001802 | -152.2975 | 0.23407 | 8.4 | 0.002572 | 0.003648 | 0.003435 | 5.5e-05 | 0.001085 | 0.29999 | 1.2972 | 2.5474 | 100.61 | 21.13 | 25.443 | 67.5232 | 83.688 | 65.1963 | 0.006347 | 1.0434 | 0.025228 | 24.1779 | -0.025774 | 0.13925 | 0.00014 | 0.010615 | 0.010615 | 0.54178
2019-03-28 15:58 | 1553803107.2888 | 1206 | 0.20948 | 0.000444 | 0 | 101.0705 | 101.1139 | 126.1711 | 21.2612 | 0.001885 | 0.001767 | -152.3045 | 0.23431 | 8.4031 | 0.002615 | 0.004104 | 0.004005 | 5.8e-05 | 0.001077 | 0.31188 | 1.3557 | 2.5473 | 100.61 | 21.13 | 25.45 | 67.1878 | 84.7187 | 64.3501 | 0.006349 | 1.0461 | 0.025252 | 24.14 | -0.003843 | 0.13939 | 0.000211 | 0.016022 | 0.016022 | 0.35931
2019-03-28 15:58 | 1553803108.2889 | 1207 |  | 0.000443 |  | 101.1414 | 101.1517 | 125.9404 | 21.2612 | 0.001902 | 0.001758 | -152.3261 | 0.23501 | 8.4009 | 0.002516 | 0.00433 | 0.004044 | 5.6e-05 | 0.001072 | 0.33196 | 1.3606 | 2.5475 | 100.61 | 21.13 | 25.457 | 67.5896 | 82.3368 | 64.1348 | 0.006351 | 1.0467 | 0.025237 | 24.1102 | -0.000907 | 0.13981
2019-03-28 15:58 | 1553803109.2889 | 1208 |  | 0.000445 |  | 101.1265 | 101.2427 | 126.1223 | 21.2612 | 0.001902 | 0.001816 | -152.3456 | 0.23472 | 8.4028 | 0.002637 | 0.003255 | 0.003049 | 5.6e-05 | 0.001072 | 0.33157 | 1.3506 | 2.5476 | 100.61 | 21.13 | 25.472 | 67.5951 | 85.2673 | 65.5242 | 0.006354 | 1.0424 | 0.025203 | 24.1772 | -0.010271 | 0.13963
2019-03-28 15:58 | 1553803110.2889 | 1209 |  |  |  | 101.1307 | 101.3161 | 126.0094 | 21.2612 | 0.001864 | 0.001849 | -152.3529 | 0.23446 | 8.3983 | 0.002614 | 0.004331 | 0.004295 | 5.8e-05 | 0.001076 | 0.30663 | 1.339 | 2.5474 | 100.61 | 21.13 | 25.478 | 66.683 | 84.692 | 66.3242 | 0.006356 | 1.04 | 0.025162 | 24.1949 | -0.016383 | 0.13948
2019-03-28 15:58 | 1553803111.2889 | 1210 |  |  |  | 101.1054 | 101.0304 | 126.0108 | 21.2612 | 0.001924 | 0.00184 | -152.3424 | 0.23427 | 8.4028 | 0.002576 | 0.00414 | 0.003886 | 5.8e-05 | 0.001076 | 0.33069 | 1.3411 | 2.5475 | 100.61 | 21.13 | 25.491 | 68.1137 | 83.7864 | 66.0917 | 0.006359 | 1.0407 | 0.025171 | 24.1868 | 0.006636 | 0.13937
2019-03-28 15:59 | 1553803112.2889 | 1211 |  |  |  | 101.1208 | 101.4333 | 125.9728 | 21.2612 | 0.001928 | 0.001826 | -152.3423 | 0.23473 | 8.401 | 0.002587 | 0.004224 | 0.004073 | 6.1e-05 | 0.001077 | 0.33212 | 1.3294 | 2.5474 | 100.61 | 21.13 | 25.498 | 68.2099 | 84.0554 | 65.7614 | 0.006361 | 1.0417 | 0.025179 | 24.1714 | -0.027594 | 0.13964
2019-03-28 15:59 | 1553803113.2889 | 1212 |  |  |  | 101.142 | 101.1038 |  | 21.2612 | 0.001946 | 0.001796 | -152.354 | 0.2344 | 8.4028 | 0.002563 | 0.004105 | 0.003895 | 6.1e-05 | 0.001079 | 0.32803 | 1.3176 | 2.5474 | 100.61 | 21.13 | 25.505 | 68.6558 | 83.4812 | 65.0506 | 0.006363 | 1.0439 |  |  | 0.003382 | 0.13944
2019-03-28 15:59 | 1553803114.2889 | 1213 |  |  |  | 101.0291 | 101.117 |  | 21.2612 | 0.001924 |  | -152.3295 | 0.23412 | 8.4028 | 0.002524 | 0.003693 | 0.00353 | 6e-05 | 0.001081 | 0.32875 | 1.2864 | 2.5473 | 100.61 | 21.13 | 25.516 | 68.131 | 82.539 |  | 0.006365 |  |  |  | -0.007778 | 0.13928
2019-03-28 15:59 | 1553803115.2889 | 1214 |  |  |  | 101.1182 | 101.203 |  | 21.2612 | 0.00193 |  |  | 0.23366 | 8.4027 | 0.00254 | 0.004022 | 0.003758 | 6.2e-05 | 0.001082 | 0.33212 | 1.3181 | 2.5473 | 100.61 | 21.13 | 25.52 | 68.2659 | 82.921 |  | 0.006366 |  |  |  | -0.007497 | 0.13901
2019-03-28 15:59 | 1553803116.2889 | 1215 |  |  |  | 101.0273 | 101.223 |  | 21.2612 |  |  |  | 0.23388 | 8.4008 | 0.002476 | 0.004215 | 0.00399 | 5.9e-05 | 0.001079 | 0.3063 | 1.2855 | 2.5472 | 100.61 | 21.13 | 25.53 |  | 81.3792 |  | 0.006369 |  |  |  | -0.017308 | 0.13913
2019-03-28 15:59 | 1553803117.2889 | 1216 |  |  |  | 101.1 | 101.2193 |  | 21.2612 |  |  |  | 0.23433 | 8.4042 | 0.002541 | 0.004081 | 0.003937 | 6.2e-05 | 0.00108 | 0.33483 | 1.2844 | 2.5474 | 100.61 | 21.13 | 25.534 |  | 82.9499 |  | 0.00637 |  |  |  | -0.010545 | 0.13941
2019-03-28 15:59 | 1553803118.2889 | 1217 |  |  |  | 101.1194 | 101.1853 |  | 21.2612 |  |  |  | 0.23415 | 8.4026 | 0.002531 | 0.003227 | 0.003063 | 6.1e-05 | 0.001081 | 0.32853 | 1.288 | 2.5473 | 100.61 | 21.13 | 25.538 |  | 82.7191 |  | 0.006371 |  |  |  | -0.00583 | 0.1393
2019-03-28 15:59 | 1553803119.2889 | 1218 |  |  |  | 101.1533 | 101.1228 |  | 21.2612 |  |  |  | 0.23366 | 8.4029 | 0.00247 | 0.003679 | 0.003541 | 6.3e-05 | 0.001086 | 0.32444 | 1.2769 | 2.5474 | 100.61 | 21.13 | 25.543 |  | 81.2482 |  | 0.006372 |  |  |  | 0.0027 | 0.139
2019-03-28 15:59 | 1553803120.2889

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-28 15:38 | 0 | 0 | 0 | 0 | 0 | 45.9 | -0.009256
2019-03-28 15:38 | 1 | 0 | 0 | 0 | 0 | 45.9036 | -0.009509
2019-03-28 15:38 | 2 | 0 | 0 | 0 | 0 | 45.927 | -0.00499
2019-03-28 15:38 | 3 | 0 | 0 | 0 | 0 | 45.9345 | 0.00079
2019-03-28 15:38 | 4 | 0 | 0 | 0 | 0 | 45.8695 | 0.02008
2019-03-28 15:38 | 5 | 0 | 0 | 0 | 0 | 45.9755 | 0.17126
2019-03-28 15:38 | 6 | 0 | 0 | 0 | 0 | 45.7755 | 0.49541
2019-03-28 15:38 | 7 | 0 | 0 | 0 | 0 | 45.7285 | 0.94164
2019-03-28 15:38 | 8 | 0 | 0 | 0 | 0 | 45.6463 | 1.2925
2019-03-28 15:38 | 9 | 0 | 0 | 0 | 0 | 45.5677 | 1.5148
2019-03-28 15:38 | 10 | 0 | 0 | 0 | 0 | 45.4512 | 1.9288
2019-03-28 15:38 | 11 | 0 | 0 | 0 | 0 | 45.3534 | 2.3847
2019-03-28 15:39 | 12 | 0 | 0 | 0 | 0 | 45.2275 | 2.6104
2019-03-28 15:39 | 13 | 0 | 0 | 0 | 0 | 45.14 | 2.9582
2019-03-28 15:39 | 14 | 0 | 0 | 0 | 0 | 45.0161 | 3.4081
2019-03-28 15:39 | 15 | 0.12807 | 0 | 0 | 0 | 44.8957 | 3.5356
2019-03-28 15:39 | 16 | 0.15239 | 0 | 0 | 0 | 44.7565 | 3.5517
2019-03-28 15:39 | 17 | 0.14128 | 0 | 0 | 0 | 44.5999 | 3.5432
2019-03-28 15:39 | 18 | 0.1153 | 0 | 0 | 0 | 44.4784 | 3.5227
2019-03-28 15:39 | 19 | 0.10347 | 0 | 0 | 0 | 44.3662 | 3.4161
2019-03-28 15:39 | 20 | 0.094576 | 0 | 0 | 0 | 44.2701 | 3.3283
2019-03-28 15:39 | 21 | 0.085318 | 0 | 0 | 0 | 44.1757 | 3.2778
2019-03-28 15:39 | 22 | 0.11548 | 0 | 0 | 0 | 44.0899 | 3.1489
2019-03-28 15:39 | 23 | 0.13535 | 0 | 0 | 0 | 43.9478 | 3.04
2019-03-28 15:39 | 24 | 0.099937 | 0 | 0 | 0 | 43.8327 | 2.9395
2019-03-28 15:39 | 25 | 0.092448 | 0 | 0 | 0 | 43.7428 | 2.7911
2019-03-28 15:39 | 26 | 0.076186 | 0 | 0 | 0 | 43.6489 | 2.6487
2019-03-28 15:39 | 27 | 0.073543 | 0 | 0 | 0 | 43.5864 | 2.5759
2019-03-28 15:39 | 28 | 0.09599 | 0.026555 | 2.4108 | 26.3039 | 43.4957 | 2.5249
2019-03-28 15:39 | 29 | 0.092919 | 0.026732 | 2.4574 | 26.8313 | 43.4002 | 2.4931
2019-03-28 15:39 | 30 | 0.087115 | 0.027453 | 2.5812 | 28.2236 | 43.3108 | 2.4587
2019-03-28 15:39 | 31 | 0.10061 | 0.023162 | 2.2046 | 24.2292 | 43.2219 | 2.4378
2019-03-28 15:39 | 32 | 0.09244 | 0.024551 | 2.3723 | 25.7991 | 43.1149 | 2.3849
2019-03-28 15:39 | 33 | 0.067349 | 0.033125 | 3.22 | 34.713 | 43.0406 | 2.3379
2019-03-28 15:39 | 34 | 0.090237 | 0.024272 | 2.3742 | 24.8074 | 42.9695 | 2.2386
2019-03-28 15:39 | 35 | 0.080007 | 0.026259 | 2.6404 | 27.1618 | 42.8667 | 2.1731
2019-03-28 15:39 | 36 | 0.083432 | 0.024433 | 2.4982 | 25.3609 | 42.8068 | 2.1159
2019-03-28 15:39 | 37 | 0.090773 | 0.022032 | 2.2621 | 22.6086 | 42.6994 | 2.0522
2019-03-28 15:39 | 38 | 0.065158 | 0.029401 | 3.1047 | 30.178 | 42.6325 | 1.9663
2019-03-28 15:39 | 39 | 0.075072 | 0.024653 | 2.6462 | 24.781 | 42.5612 | 1.8603
2019-03-28 15:39 | 40 | 0.07913 | 0.022557 | 2.4611 | 22.2739 | 42.483 | 1.7625
2019-03-28 15:39 | 41 | 0.079412 | 0.021358 | 2.4017 | 21.0315 | 42.4044 | 1.6702
2019-03-28 15:39 | 42 | 0.063484 | 0.025716 | 2.941 | 25.5038 | 42.3276 | 1.6191
2019-03-28 15:39 | 43 | 0.076414 | 0.020545 | 2.3919 | 20.501 | 42.2709 | 1.5666
2019-03-28 15:39 | 44 | 0.079229 | 0.018934 | 2.2442 | 19.2713 | 42.1768 | 1.5268
2019-03-28 15:39 | 45 | 0.073603 | 0.019809 | 2.3574 | 20.3309 | 42.1151 | 1.4964
2019-03-28 15:39 | 46 | 0.06091 | 0.023177 | 2.7697 | 23.4094 | 42.0316 | 1.4259
2019-03-28 15:39 | 47 | 0.048345 | 0.028254 | 3.3987 | 28.6096 | 41.9922 | 1.3831
2019-03-28 15:39 | 48 | 0.081313 | 0.016257 | 1.9706 | 16.1067 | 41.9245 | 1.3097
2019-03-28 15:39 | 49 | 0.072554 | 0.01759 | 2.1496 | 17.779 | 41.839 | 1.2899
2019-03-28 15:39 | 50 | 0.057703 | 0.021666 | 2.6296 | 22.1277 | 41.7817 | 1.2768
2019-03-28 15:39 | 51 | 0.057028 | 0.021145 | 2.5942 | 22.1584 | 41.7201 | 1.2637
2019-03-28 15:39 | 52 | 0.061321 | 0.01928 | 2.3567 | 20.3037 | 41.6664 | 1.245
2019-03-28 15:39 | 53 | 0.050153 | 0.023166 | 2.8209 | 24.3126 | 41.6011 | 1.2193
2019-03-28 15:39 | 54 | 0.040297 | 0.027617 | 3.4231 | 30.2488 | 41.5655 | 1.2189
2019-03-28 15:39 | 55 | 0.055636 | 0.019197 | 2.4332 | 21.3485 | 41.5143 | 1.1877
2019-03-28 15:39 | 56 | 0.063465 | 0.016495 | 2.0967 | 18.8639 | 41.4557 | 1.1972
2019-03-28 15:39 | 57 | 0.061571 | 0.016293 | 2.1274 | 19.3215 | 41.3904 | 1.1897
2019-03-28 15:39 | 58 | 0.043926 | 0.022277 | 2.9308 | 26.611 | 41.3359 | 1.1689
2019-03-28 15:39 | 59 | 0.060479 | 0.016191 | 2.0993 | 19.4672 | 41.2948 | 1.1774
2019-03-28 15:39 | 60 | 0.062617 | 0.015276 | 1.9999 | 18.9069 | 41.218 | 1.1839
2019-03-28 15:39 | 61 | 0.053091 | 0.01803 | 2.3321 | 21.9224 | 41.1727 | 1.1639
2019-03-28 15:39 | 62 | 0.05179 | 0.018908 | 2.3611 | 22.5357 | 41.1102 | 1.1671
2019-03-28 15:39 | 63 | 0.042404 | 0.023656 | 2.8599 | 27.3059 | 41.0702 | 1.1579
2019-03-28 15:39 | 64 | 0.06171 | 0.016288 | 1.9516 | 18.9618 | 41.0192 | 1.1701
2019-03-28 15:39 | 65 | 0.049969 | 0.020547 | 2.3922 | 23.3586 | 40.9535 | 1.1672
2019-03-28 15:39 | 66 | 0.047603 | 0.02151 | 2.4811 | 24.9161 | 40.9174 | 1.1861
2019-03-28 15:39 | 67 | 0.056414 | 0.018108 | 2.0804 | 21.0061 | 40.8551 | 1.185
2019-03-28 15:39 | 68 | 0.072725 | 0.014067 | 1.6068 | 16.3882 | 40.8034 | 1.1918
2019-03-28 15:39 | 69 | 0.050216 | 0.020823 | 2.3176 | 23.4592 | 40.7184 | 1.178
2019-03-28 15:39 | 70 | 0.056525 | 0.018488 | 2.0539 | 21.0107 | 40.6965 | 1.1876
2019-03-28 15:40 | 71 | 0.074789 | 0.013665 | 1.5447 | 16.0178 | 40.6026 | 1.198
2019-03-28 15:40 | 72 | 0.057625 | 0.018135 | 1.9986 | 20.9442 | 40.5548 | 1.2069
2019-03-28 15:40 | 73 | 0.068321 | 0.015319 | 1.6877 | 17.8216 | 40.4821 | 1.2176
2019-03-28 15:40 | 74 | 0.044982 | 0.02374 | 2.5567 | 27.0847 | 40.4244 | 1.2183
2019-03-28 15:40 | 75 | 0.079181 | 0.013484 | 1.4522 | 15.3218 | 40.3796 | 1.2132
2019-03-28 15:40 | 76 | 0.077126 | 0.013864 | 1.4916 | 15.7928 | 40.2742 | 1.218
2019-03-28 15:40 | 77 | 0.057785 | 0.018076 | 1.9896 | 21.4563 | 40.2306 | 1.2398
2019-03-28 15:40 | 78 | 0.04019 | 0.026031 | 2.8679 | 30.9761 | 40.158 | 1.2449
2019-03-28 15:40 | 79 | 0.047621 | 0.021967 | 2.4156 | 26.3635 | 40.1436 | 1.2554
2019-03-28 15:40 | 80 | 0.065324 | 0.015703 | 1.7673 | 19.3639 | 40.0606 | 1.2649
2019-03-28 15:40 | 81 | 0.036415 | 0.028757 | 3.1711 | 34.5662 | 40.0229 | 1.2587
2019-03-28 15:40 | 82 | 0.059201 | 0.0173 | 1.9519 | 21.4313 | 39.9765 | 1.2688
2019-03-28 15:40 | 83 | 0.065627 | 0.015924 | 1.762 | 19.3793 | 39.9074 | 1.2718
2019-03-28 15:40 | 84 | 0.068791 | 0.01519 | 1.6849 | 18.7889 | 39.847 | 1.2925
2019-03-28 15:40 | 85 | 0.052635 | 0.020329 | 2.2124 | 24.733 | 39.7748 | 1.3018
2019-03-28 15:40 | 86 | 0.039609 | 0.027513 | 2.9454 | 33.0468 | 39.741 | 1.3089
2019-03-28 15:40 | 87 | 0.040431 | 0.027582 | 2.904 | 32.5794 | 39.6911 | 1.3172
2019-03-28 15:40 | 88 | 0.071398 | 0.015568 | 1.6461 | 18.3499 | 39.6526 | 1.3101
2019-03-28 15:40 | 89 | 0.090329 | 0.012286 | 1.3027 | 14.4665 | 39.5522 | 1.3067
2019-03-28 15:40 | 90 | 0.04615 | 0.024088 | 2.5616 | 28.7733 | 39.4866 | 1.3279
2019-03-28 15:40 | 91 | 0.061387 | 0.018086 | 1.9285 | 21.5057 | 39.4471 | 1.3202
2019-03-28 15:40 | 92 | 0.061929 | 0.017899 | 1.9132 | 21.3629 | 39.3659 | 1.323
2019-03-28 15:40 | 93 | 0.061027 | 0.017814 | 1.9496 | 21.6321 | 39.3236 | 1.3201
2019-03-28 15:40 | 94 | 0.073748 | 0.015049 | 1.6178 | 17.9322 | 39.2413 | 1.3225
2019-03-28 15:40 | 95 | 0.056486 | 0.019618 | 2.1178 | 23.7221 | 39.1826 | 1.34
2019-03-28 15:40 | 96 | 0.062728 | 0.018053 | 1.9149 | 21.5499 | 39.1231 | 1.3518
2019-03-28 15:40 | 97 | 0.075998 | 0.014903 | 1.5854 | 18.1967 | 39.0558 | 1.3829
2019-03-28 15:40 | 98 | 0.055317 | 0.020443 | 2.1838 | 24.9201 | 38.9794 | 1.3785
2019-03-28 15:40 | 99 | 0.032294 | 0.035107 | 3.7658 | 43.1378 | 38.945 | 1.3931
2019-03-28 15:40 | 100 | 0.069375 | 0.01666 | 1.7582 | 20.2892 | 38.9008 | 1.4076
2019-03-28 15:40 | 101 | 0.068996 | 0.016756 | 1.7776 | 20.6001 | 38.8141 | 1.4213
2019-03-28 15:40 | 102 | 0.062145 | 0.018954 | 1.9825 | 23.0988 | 38.7651 | 1.4355
2019-03-28 15:40 | 103 | 0.068299 | 0.01758 | 1.8145 | 21.036 | 38.6875 | 1.4367
2019-03-28 15:40 | 104 | 0.052103 | 0.022979 | 2.3847 | 27.6642 | 38.6331 | 1.4414
2019-03-28 15:40 | 105 | 0.059171 | 0.020623 | 2.1118 | 24.2166 | 38.5779 | 1.4329
2019-03-28 15:40 | 106 | 0.0711 | 0.017186 | 1.7704 | 20.3307 | 38.5134 | 1.4455
2019-03-28 15:40 | 107 | 0.066355 | 0.01869 | 1.9005 | 22.1196 | 38.4398 | 1.4678
2019-03-28 15:40 | 108 | 0.055886 | 0.022633 | 2.2706 | 26.6576 | 38.3817 | 1.4898
2019-03-28 15:40 | 109 | 0.078392 | 0.0161 | 1.6233 | 18.7807 | 38.3204 | 1.4722
2019-03-28 15:40 | 110 | 0.083387 | 0.015133 | 1.5359 | 17.6168 | 38.2289 | 1.469
2019-03-28 15:40 | 111 | 0.071703 | 0.017587 | 1.79 | 20.3557 | 38.1582 | 1.4596
2019-03-28 15:40 | 112 | 0.06408 | 0.019679 | 2.0139 | 23.3323 | 38.0851 | 1.4951
2019-03-28 15:40 | 113 | 0.05159 | 0.024442 | 2.511 | 29.0162 | 38.0303 | 1.4969
2019-03-28 15:40 | 114 | 0.068569 | 0.018026 | 1.8943 | 21.4121 | 37.9746 | 1.4682
2019-03-28 15:40 | 115 | 0.080945 | 0.015039 | 1.6145 | 18.3988 | 37.8938 | 1.4893
2019-03-28 15:40 | 116 | 0.084556 | 0.014082 | 1.5457 | 18.2101 | 37.8155 | 1.5398
2019-03-28 15:40 | 117 | 0.072528 | 0.016407 | 1.8062 | 21.1076 | 37.7288 | 1.5309
2019-03-28 15:40 | 118 | 0.06483 | 0.018643 | 2.022 | 23.8083 | 37.6695 | 1.5435
2019-03-28 15:40 | 119 | 0.067216 | 0.018316 | 1.9556 | 23.5032 | 37.5961 | 1.5798
2019-03-28 15:40 | 120 | 0.080197 | 0.015623 | 1.647 | 19.8975 | 37.5322 | 1.5957
2019-03-28 15:40 | 121 | 0.088728 | 0.014348 | 1.494 | 18.0618 | 37.437 | 1.6026
2019-03-28 15:40 | 122 | 0.077517 | 0.016723 | 1.7245 | 20.4545 | 37.359 | 1.5856
2019-03-28 15:40 | 123 | 0.098749 | 0.013148 | 1.3647 | 16.0631 | 37.2753 | 1.5862
2019-03-28 15:40 | 124 | 0.088792 | 0.014623 | 1.5285 | 18.0919 | 37.1687 | 1.6064
2019-03-28 15:40 | 125 | 0.074179 | 0.017435 | 1.8361 | 21.3636 | 37.0995 | 1.5847
2019-03-28 15:40 | 126 | 0.06739 | 0.019526 | 2.0319 | 23.6138 | 37.0183 | 1.5913
2019-03-28 15:40 | 127 | 0.068141 | 0.019004 | 2.0199 | 23.1852 | 36.9619 | 1.5799
2019-03-28 15:40 | 128 | 0.095442 | 0.013596 | 1.4499 | 16.707 | 36.8765 | 1.5945
2019-03-28 15:40 | 129 | 0.074699 | 0.017362 | 1.8577 | 21.4298 | 36.7825 | 1.6008
2019-03-28 15:40 | 130 | 0.05071 | 0.02556 | 2.7388 | 31.6928 | 36.7285 | 1.6071
2019-03-28 15:40 | 131 | 0.049882 | 0.025951 | 2.785 | 32.2171 | 36.6747 | 1.6071
2019-03-28 15:41 | 132 | 0.073158 | 0.017635 | 1.9011 | 22.184 | 36.6223 | 1.6229
2019-03-28 15:41 | 133 | 0.082635 | 0.015617 | 1.6849 | 19.6702 | 36.5315 | 1.6254
2019-03-28 15:41 | 134 | 0.078643 | 0.016712 | 1.7796 | 20.6266 | 36.4607 | 1.6221
2019-03-28 15:41 | 135 | 0.080422 | 0.016079 | 1.7447 | 20.2517 | 36.3732 | 1.6287
2019-03-28 15:41 | 136 | 0.080537 | 0.016047 | 1.7482 | 20.1135 | 36.2999 | 1.6199
2019-03-28 15:41 | 137 | 0.08948 | 0.014441 | 1.5764 | 18.1082 | 36.2105 | 1.6203
2019-03-28 15:41 | 138 | 0.076618 | 0.016849 | 1.8434 | 21.1531 | 36.1254 | 1.6207
2019-03-28 15:41 | 139 | 0.093552 | 0.013812 | 1.5152 | 17.1804 | 36.0512 | 1.6073
2019-03-28 15:41 | 140 | 0.083652 | 0.015191 | 1.6962 | 19.0505 | 35.9445 | 1.5936
2019-03-28 15:41 | 141 | 0.067293 | 0.018856 | 2.1085 | 23.7101 | 35.8855 | 1.5955
2019-03-28 15:41 | 142 | 0.077881 | 0.016322 | 1.8265 | 20.3572 | 35.8038 | 1.5854
2019-03-28 15:41 | 143 | 0.071859 | 0.017395 | 1.9788 | 22.1304 | 35.7333 | 1.5903
2019-03-28 15:41 | 144 | 0.059489 | 0.02096 | 2.3844 | 26.7957 | 35.6619 | 1.594
2019-03-28 15:41 | 145 | 0.0549 | 0.023103 | 2.5796 | 28.8481 | 35.6122 | 1.5838
2019-03-28 15:41 | 146 | 0.063247 | 0.020062 | 2.2384 | 25.2462 | 35.5484 | 1.5967
2019-03-28 15:41 | 147 | 0.07463 | 0.017059 | 1.9009 | 21.3643 | 35.4847 | 1.5944
2019-03-28 15:41 | 148 | 0.081395 | 0.015625 | 1.7416 | 19.7728 | 35.4008 | 1.6094
2019-03-28 15:41 | 149 | 0.066245 | 0.01917 | 2.1372 | 24.1421 | 35.3269 | 1.5993
2019-03-28 15:41 | 150 | 0.074151 | 0.017117 | 1.9103 | 21.4871 | 35.2629 | 1.5933
2019-03-28 15:41 | 151 | 0.092825 | 0.013691 | 1.5306 | 17.0818 | 35.1762 | 1.5856
2019-03-28 15:41 | 152 | 0.077841 | 0.016335 | 1.8248 | 20.22 | 35.0851 | 1.5739
2019-03-28 15:41 | 153 | 0.070412 | 0.017764 | 2.0219 | 22.3614 | 35.0184 | 1.5745
2019-03-28 15:41 | 154 | 0.10353 | 0.012071 | 1.3722 | 15.2978 | 34.9349 | 1.5838
2019-03-28 15:41 | 155 | 0.1035 | 0.01185 | 1.3692 | 15.3182 | 34.8193 | 1.5854
2019-03-28 15:41 | 156 | 0.079843 | 0.015332 | 1.7689 | 19.5208 | 34.7343 | 1.5586
2019-03-28 15:41 | 157 | 0.075935 | 0.016129 | 1.859 | 20.6358 | 34.6557 | 1.567
2019-03-28 15:41 | 158 | 0.059291 | 0.020257 | 2.3762 | 26.1728 | 34.5842 | 1.5518
2019-03-28 15:41 | 159 | 0.070939 | 0.016905 | 1.9791 | 21.8945 | 34.5302 | 1.5532
2019-03-28 15:41 | 160 | 0.083441 | 0.014393 | 1.6821 | 18.598 | 34.4417 | 1.5518
2019-03-28 15:41 | 161 | 0.082004 | 0.014923 | 1.7089 | 18.9065 | 34.3667 | 1.5504
2019-03-28 15:41 | 162 | 0.083446 | 0.014898 | 1.6722 | 18.5418 | 34.278 | 1.5472
2019-03-28 15:41 | 163 | 0.064212 | 0.019694 | 2.1701 | 24.1482 | 34.2047 | 1.5506
2019-03-28 15:41 | 164 | 0.055461 | 0.022827 | 2.5116 | 27.926 | 34.1469 | 1.5488
2019-03-28 15:41 | 165 | 0.065374 | 0.019697 | 2.1256 | 23.6183 | 34.0883 | 1.544
2019-03-28 15:41 | 166 | 0.0911 | 0.014389 | 1.5269 | 16.9608 | 34.0122 | 1.5451
2019-03-28 15:41 | 167 | 0.09409 | 0.013939 | 1.4783 | 16.3251 | 33.9114 | 1.536
2019-03-28 15:41 | 168 | 0.098542 | 0.013296 | 1.4087 | 15.402 | 33.8249 | 1.5177
2019-03-28 15:41 | 169 | 0.081512 | 0.016353 | 1.7025 | 18.634 | 33.7196 | 1.5189
2019-03-28 15:41 | 170 | 0.071763 | 0.018255 | 1.9273 | 21.0958 | 33.6601 | 1.5139
2019-03-28 15:41 | 171 | 0.086471 | 0.015374 | 1.5966 | 17.346 | 33.5705 | 1.4999
2019-03-28 15:41 | 172 | 0.069267 | 0.018923 | 1.9944 | 21.6025 | 33.4938 | 1.4963
2019-03-28 15:41 | 173 | 0.077304 | 0.016973 | 1.7843 | 19.3859 | 33.4267 | 1.4986
2019-03-28 15:41 | 174 | 0.076666 | 0.017089 | 1.7947 | 19.2757 | 33.3408 | 1.4778
2019-03-28 15:41 | 175 | 0.072358 | 0.018148 | 1.8972 | 20.3078 | 33.2744 | 1.4694
2019-03-28 15:41 | 176 | 0.071661 | 0.018324 | 1.9092 | 20.5794 | 33.1952 | 1.4747
2019-03-28 15:41 | 177 | 0.075313 | 0.017138 | 1.8108 | 19.4849 | 33.1301 | 1.4675
2019-03-28 15:41 | 178 | 0.069907 | 0.018467 | 1.9448 | 20.6529 | 33.0464 | 1.4438
2019-03-28 15:41 | 179 | 0.074222 | 0.017109 | 1.8263 | 19.2309 | 32.988 | 1.4273
2019-03-28 15:41 | 180 | 0.084464 | 0.014783 | 1.6005 | 16.6177 | 32.8968 | 1.4036
2019-03-28 15:41 | 181 | 0.07284 | 0.01653 | 1.8482 | 19.0493 | 32.8237 | 1.3875
2019-03-28 15:41 | 182 | 0.08796 | 0.01342 | 1.5208 | 15.6426 | 32.7457 | 1.3759
2019-03-28 15:41 | 183 | 0.076418 | 0.015192 | 1.7441 | 17.8461 | 32.6536 | 1.3638
2019-03-28 15:41 | 184 | 0.071942 | 0.015801 | 1.8408 | 18.8557 | 32.5916 | 1.3565
2019-03-28 15:41 | 185 | 0.076049 | 0.014945 | 1.7287 | 17.843 | 32.5075 | 1.3569
2019-03-28 15:41 | 186 | 0.074843 | 0.015455 | 1.743 | 18.1419 | 32.4407 | 1.3578
2019-03-28 15:41 | 187 | 0.065448 | 0.018047 | 1.9846 | 20.8234 | 32.3596 | 1.3628
2019-03-28 15:41 | 188 | 0.074266 | 0.016169 | 1.7362 | 18.2592 | 32.3058 | 1.356
2019-03-28 15:41 | 189 | 0.074663 | 0.017017 | 1.7245 | 18.5569 | 32.2131 | 1.3855
2019-03-28 15:41 | 190 | 0.05809 | 0.021827 | 2.2006 | 23.4926 | 32.1604 | 1.3647
2019-03-28 15:42 | 191 | 0.062223 | 0.020713 | 2.0483 | 21.5476 | 32.0929 | 1.3408
2019-03-28 15:42 | 192 | 0.037612 | 0.034833 | 3.3761 | 36.4666 | 32.0411 | 1.3716
2019-03-28 15:42 | 193 | 0.065895 | 0.019857 | 1.9204 | 20.7683 | 32.0051 | 1.3685
2019-03-28 15:42 | 194 | 0.092827 | 0.014122 | 1.3628 | 14.7902 | 31.9095 | 1.3729
2019-03-28 15:42 | 195 | 0.075099 | 0.017747 | 1.6849 | 18.0757 | 31.8308 | 1.3575
2019-03-28 15:42 | 196 | 0.05334 | 0.025029 | 2.3684 | 25.517 | 31.7607 | 1.3611
2019-03-28 15:42 | 197 | 0.064384 | 0.020757 | 1.9614 | 21.0152 | 31.7161 | 1.353
2019-03-28 15:42 | 198 | 0.07381 | 0.018082 | 1.7059 | 18.2749 | 31.6317 | 1.3489
2019-03-28 15:42 | 199 | 0.071085 | 0.019105 | 1.7674 | 18.8815 | 31.5713 | 1.3422
2019-03-28 15:42 | 200 | 0.076079 | 0.017574 | 1.6487 | 17.6088 | 31.4883 | 1.3397
2019-03-28 15:42 | 201 | 0.065721 | 0.021051 | 1.9091 | 20.6067 | 31.4217 | 1.3543
2019-03-28 15:42 | 202 | 0.10069 | 0.014218 | 1.2445 | 13.4559 | 31.3476 | 1.3548
2019-03-28 15:42 | 203 | 0.075758 | 0.019466 | 1.6504 | 17.8358 | 31.2342 | 1.3512
2019-03-28 15:42 | 204 | 0.051828 | 0.028395 | 2.4014 | 26.4243 | 31.1964 | 1.3695
2019-03-28 15:42 | 205 | 0.064365 | 0.023172 | 1.9321 | 21.3679 | 31.1216 | 1.3754
2019-03-28 15:42 | 206 | 0.069663 | 0.021422 | 1.7836 | 19.8459 | 31.0686 | 1.3825
2019-03-28 15:42 | 207 | 0.068748 | 0.021717 | 1.8118 | 20.1413 | 30.9844 | 1.3847
2019-03-28 15:42 | 208 | 0.050129 | 0.030576 | 2.4783 | 27.3803 | 30.9351 | 1.3725
2019-03-28 15:42 | 209 | 0.057583 | 0.027043 | 2.1594 | 23.9074 | 30.8781 | 1.3767
2019-03-28 15:42 | 210 | 0.060346 | 0.026885 | 2.0586 | 22.854 | 30.8196 | 1.3791
2019-03-28 15:42 | 211 | 0.089031 | 0.018942 | 1.3923 | 15.5594 | 30.7517 | 1.3853
2019-03-28 15:42 | 212 | 0.090412 | 0.019407 | 1.3743 | 15.4451 | 30.6481 | 1.3964
2019-03-28 15:42 | 213 | 0.080392 | 0.022379 | 1.5469 | 17.5806 | 30.5746 | 1.4133
2019-03-28 15:42 | 214 | 0.06951 | 0.02716 | 1.7926 | 20.2735 | 30.4874 | 1.4092
2019-03-28 15:42 | 215 | 0.078402 | 0.024946 | 1.5974 | 18.0493 | 30.4307 | 1.4151
2019-03-28 15:42 | 216 | 0.084294 | 0.023992 | 1.489 | 16.9407 | 30.3311 | 1.428
2019-03-28 15:42 | 217 | 0.075837 | 0.027881 | 1.6584 | 18.8106 | 30.2652 | 1.4265
2019-03-28 15:42 | 218 | 0.08704 | 0.024759 | 1.4468 | 16.3908 | 30.1753 | 1.4266
2019-03-28 15:42 | 219 | 0.084398 | 0.02632 | 1.4956 | 16.8804 | 30.0942 | 1.4247
2019-03-28 15:42 | 220 | 0.081948 | 0.027352 | 1.5409 | 17.2343 | 30.0077 | 1.4123
2019-03-28 15:42 | 221 | 0.049947 | 0.045818 | 2.5319 | 28.4385 | 29.9369 | 1.4204
2019-03-28 15:42 | 222 | 0.050355 | 0.045536 | 2.5191 | 27.3228 | 29.9001 | 1.3758
2019-03-28 15:42 | 223 | 0.06042 | 0.038272 | 2.0928 | 23.408 | 29.8329 | 1.4143
2019-03-28 15:42 | 224 | 0.066823 | 0.034695 | 1.8962 | 20.6197 | 29.7799 | 1.3779
2019-03-28 15:42 | 225 | 0.068077 | 0.034061 | 1.8569 | 19.8886 | 29.7007 | 1.354
2019-03-28 15:42 | 226 | 0.069199 | 0.03382 | 1.8217 | 20.2209 | 29.6441 | 1.3993
2019-03-28 15:42 | 227 | 0.067752 | 0.034832 | 1.8564 | 20.1838 | 29.5631 | 1.3675
2019-03-28 15:42 | 228 | 0.058354 | 0.041928 | 2.1497 | 23.2631 | 29.51 | 1.3575
2019-03-28 15:42 | 229 | 0.074081 | 0.033928 | 1.6886 | 18.7579 | 29.4406 | 1.3896
2019-03-28 15:42 | 230 | 0.076529 | 0.033447 | 1.6298 | 18.1263 | 29.3645 | 1.3872
2019-03-28 15:42 | 231 | 0.079772 | 0.032262 | 1.5555 | 17.3869 | 29.2886 | 1.387
2019-03-28 15:42 | 232 | 0.050301 | 0.051556 | 2.4603 | 27.956 | 29.2117 | 1.4062
2019-03-28 15:42 | 233 | 0.074088 | 0.035715 | 1.6715 | 18.731 | 29.1755 | 1.3877
2019-03-28 15:42 | 234 | 0.097618 | 0.027548 | 1.2672 | 14.4054 | 29.0636 | 1.4062
2019-03-28 15:42 | 235 | 0.068906 | 0.039548 | 1.7896 | 20.3614 | 28.9926 | 1.403
2019-03-28 15:42 | 236 | 0.067486 | 0.041295 | 1.8242 | 21.0052 | 28.9202 | 1.4176
2019-03-28 15:42 | 237 | 0.073925 | 0.038464 | 1.6693 | 19.1893 | 28.8548 | 1.4186
2019-03-28 15:42 | 238 | 0.087129 | 0.033166 | 1.4216 | 16.1566 | 28.7708 | 1.4077
2019-03-28 15:42 | 239 | 0.080718 | 0.036277 | 1.5319 | 17.5763 | 28.6848 | 1.4187
2019-03-28 15:42 | 240 | 0.09352 | 0.031537 | 1.3239 | 14.9859 | 28.6055 | 1.4015
2019-03-28 15:42 | 241 | 0.093893 | 0.031666 | 1.3202 | 14.8314 | 28.5007 | 1.3926
2019-03-28 15:42 | 242 | 0.083382 | 0.036501 | 1.4888 | 16.8945 | 28.4207 | 1.4087
2019-03-28 15:42 | 243 | 0.076632 | 0.040515 | 1.619 | 18.4693 | 28.3333 | 1.4153
2019-03-28 15:42 | 244 | 0.069362 | 0.046352 | 1.7888 | 20.1543 | 28.2668 | 1.3979
2019-03-28 15:42 | 245 | 0.086441 | 0.038481 | 1.4348 | 16.3419 | 28.1892 | 1.4126
2019-03-28 15:42 | 246 | 0.0772 | 0.044275 | 1.6075 | 18.1089 | 28.0998 | 1.398
2019-03-28 15:42 | 247 | 0.072046 | 0.0484 | 1.7241 | 19.9282 | 28.0344 | 1.4357
2019-03-28 15:42 | 248 | 0.068646 | 0.051411 | 1.8084 | 21.319 | 27.9553 | 1.4635
2019-03-28 15:42 | 249 | 0.060639 | 0.058215 | 2.0529 | 24.484 | 27.898 | 1.4847
2019-03-28 15:42 | 250 | 0.059381 | 0.059435 | 2.1035 | 24.6033 | 27.8322 | 1.461
2019-03-28 15:42 | 251 | 0.065229 | 0.054744 | 1.9172 | 22.1736 | 27.7774 | 1.4464
2019-03-28 15:43 | 252 | 0.06866 | 0.052395 | 1.8287 | 21.1331 | 27.7026 | 1.451
2019-03-28 15:43 | 253 | 0.053901 | 0.067049 | 2.327 | 27.3732 | 27.644 | 1.4754
2019-03-28 15:43 | 254 | 0.061983 | 0.058238 | 2.0235 | 23.5505 | 27.5897 | 1.4597
2019-03-28 15:43 | 255 | 0.064067 | 0.056084 | 1.9612 | 22.6979 | 27.5206 | 1.4542
2019-03-28 15:43 | 256 | 0.081429 | 0.044066 | 1.544 | 18.3633 | 27.4585 | 1.4953
2019-03-28 15:43 | 257 | 0.075864 | 0.04735 | 1.6595 | 19.5267 | 27.3633 | 1.4814
2019-03-28 15:43 | 258 | 0.065807 | 0.055459 | 1.9078 | 21.9582 | 27.3084 | 1.445
2019-03-28 15:43 | 259 | 0.060565 | 0.060936 | 2.0734 | 24.8904 | 27.2301 | 1.5075
2019-03-28 15:43 | 260 | 0.070154 | 0.052922 | 1.792 | 20.8801 | 27.1832 | 1.4648
2019-03-28 15:43 | 261 | 0.083992 | 0.044288 | 1.5017 | 17.8802 | 27.0889 | 1.5018
2019-03-28 15:43 | 262 | 0.075431 | 0.049869 | 1.6738 | 20.2354 | 27.0203 | 1.5264
2019-03-28 15:43 | 263 | 0.087481 | 0.042797 | 1.4481 | 17.1463 | 26.9344 | 1.5
2019-03-28 15:43 | 264 | 0.068293 | 0.054414 | 1.8588 | 22.1715 | 26.8522 | 1.5141
2019-03-28 15:43 | 265 | 0.06843 | 0.053719 | 1.8643 | 21.9836 | 26.794 | 1.5043
2019-03-28 15:43 | 266 | 0.056637 | 0.063677 | 2.2536 | 26.7298 | 26.7172 | 1.5139
2019-03-28 15:43 | 267 | 0.065109 | 0.054345 | 1.9637 | 23.2302 | 26.6756 | 1.5125
2019-03-28 15:43 | 268 | 0.086011 | 0.04041 | 1.4917 | 17.4314 | 26.5841 | 1.4993
2019-03-28 15:43 | 269 | 0.081069 | 0.042697 | 1.5866 | 18.3828 | 26.5103 | 1.4903
2019-03-28 15:43 | 270 | 0.05838 | 0.058484 | 2.2057 | 25.5439 | 26.427 | 1.4913
2019-03-28 15:43 | 271 | 0.041429 | 0.080745 | 3.1049 | 36.1561 | 26.3915 | 1.4979
2019-03-28 15:43 | 272 | 0.065443 | 0.050173 | 1.9689 | 22.9915 | 26.3344 | 1.5046
2019-03-28 15:43 | 273 | 0.057456 | 0.056298 | 2.2432 | 26.0908 | 26.2669 | 1.4991
2019-03-28 15:43 | 274 | 0.067844 | 0.047267 | 1.8989 | 21.2928 | 26.2159 | 1.4446
2019-03-28 15:43 | 275 | 0.069378 | 0.045849 | 1.8528 | 20.8921 | 26.1329 | 1.4494
2019-03-28 15:43 | 276 | 0.076808 | 0.041258 | 1.6766 | 18.622 | 26.0761 | 1.4303
2019-03-28 15:43 | 277 | 0.080068 | 0.038466 | 1.6034 | 17.9722 | 25.981 | 1.439
2019-03-28 15:43 | 278 | 0.053273 | 0.056198 | 2.404 | 26.3175 | 25.9227 | 1.402
2019-03-28 15:43 | 279 | 0.06202 | 0.046789 | 2.052 | 22.5931 | 25.8665 | 1.4012
2019-03-28 15:43 | 280 | 0.063812 | 0.044687 | 1.9831 | 21.6166 | 25.7995 | 1.3794
2019-03-28 15:43 | 281 | 0.062899 | 0.044229 | 1.999 | 21.8832 | 25.7396 | 1.3764
2019-03-28 15:43 | 282 | 0.06626 | 0.041642 | 1.8898 | 20.6526 | 25.673 | 1.3684
2019-03-28 15:43 | 283 | 0.058672 | 0.046364 | 2.1294 | 22.8774 | 25.6098 | 1.3423
2019-03-28 15:43 | 284 | 0.052237 | 0.05115 | 2.3755 | 25.2748 | 25.5557 | 1.3203
2019-03-28 15:43 | 285 | 0.045842 | 0.058482 | 2.7076 | 28.215 | 25.5053 | 1.2934
2019-03-28 15:43 | 286 | 0.03855 | 0.068973 | 3.2028 | 32.6226 | 25.4639 | 1.2576
2019-03-28 15:43 | 287 | 0.042209 | 0.063137 | 2.9111 | 29.7902 | 25.4251 | 1.2574
2019-03-28 15:43 | 288 | 0.059009 | 0.044711 | 2.0617 | 21.2225 | 25.3768 | 1.2523
2019-03-28 15:43 | 289 | 0.04083 | 0.064095 | 2.9576 | 30.4378 | 25.3149 | 1.2428
2019-03-28 15:43 | 290 | 0.045316 | 0.057379 | 2.6411 | 26.9387 | 25.2903 | 1.2208
2019-03-28 15:43 | 291 | 0.051262 | 0.050432 | 2.3229 | 23.5044 | 25.2233 | 1.2049
2019-03-28 15:43 | 292 | 0.054735 | 0.046797 | 2.1531 | 21.7821 | 25.1882 | 1.1922
2019-03-28 15:43 | 293 | 0.064526 | 0.039424 | 1.814 | 18.0924 | 25.1133 | 1.1674
2019-03-28 15:43 | 294 | 0.03697 | 0.067614 | 3.1325 | 30.7438 | 25.0673 | 1.1366
2019-03-28 15:43 | 295 | 0.05146 | 0.047805 | 2.2236 | 22.0908 | 25.0301 | 1.1368
2019-03-28 15:43 | 296 | 0.057468 | 0.04194 | 1.9668 | 19.3969 | 24.9658 | 1.1147
2019-03-28 15:43 | 297 | 0.050135 | 0.047291 | 2.2315 | 22.3586 | 24.9188 | 1.1209
2019-03-28 15:43 | 298 | 0.037051 | 0.062821 | 2.991 | 29.797 | 24.8666 | 1.104
2019-03-28 15:43 | 299 | 0.051127 | 0.044693 | 2.1483 | 21.0147 | 24.8385 | 1.0744
2019-03-28 15:43 | 300 | 0.04863 | 0.046382 | 2.2234 | 21.2132 | 24.7679 | 1.0316
2019-03-28 15:43 | 301 | 0.042186 | 0.052547 | 2.544 | 24.4583 | 24.7426 | 1.0318
2019-03-28 15:43 | 302 | 0.03733 | 0.058726 | 2.8439 | 27.6246 | 24.6833 | 1.0312
2019-03-28 15:43 | 303 | 0.027056 | 0.079553 | 3.8845 | 36.9468 | 24.6682 | 0.99962
2019-03-28 15:43 | 304 | 0.056853 | 0.037396 | 1.8254 | 17.6153 | 24.6203 | 1.0015
2019-03-28 15:43 | 305 | 0.05032 | 0.041916 | 2.0426 | 18.8227 | 24.563 | 0.94716
2019-03-28 15:43 | 306 | 0.028088 | 0.075135 | 3.6124 | 36.9322 | 24.5239 | 1.0374
2019-03-28 15:43 | 307 | 0.02903 | 0.071035 | 3.4446 | 31.8415 | 24.5009 | 0.92436
2019-03-28 15:43 | 308 | 0.051951 | 0.039718 | 1.8995 | 18.0832 | 24.4607 | 0.93944
2019-03-28 15:43 | 309 | 0.037821 | 0.053431 | 2.5758 | 21.9093 | 24.4056 | 0.82862
2019-03-28 15:43 | 310 | 0.024606 | 0.081403 | 3.918 | 33.8349 | 24.3852 | 0.83253
2019-03-28 15:43 | 311 | 0.033111 | 0.059813 | 2.8785 | 27.4996 | 24.351 | 0.91054
2019-03-28 15:44 | 312 | 0.038156 | 0.05135 | 2.4681 | 22.7277 | 24.3193 | 0.86719
2019-03-28 15:44 | 313 | 0.044339 | 0.044252 | 2.0974 | 19.6032 | 24.2752 | 0.86918
2019-03-28 15:44 | 314 | 0.026251 | 0.074033 | 3.5064 | 32.3812 | 24.2362 | 0.85004
2019-03-28 15:44 | 315 | 0.024326 | 0.078843 | 3.73 | 35.9872 | 24.2185 | 0.87544
2019-03-28 15:44 | 316 | 0.049218 | 0.038506 | 1.8181 | 16.8416 | 24.1812 | 0.82891
2019-03-28 15:44 | 317 | 0.047854 | 0.039671 | 1.8502 | 17.2465 | 24.1269 | 0.82531
2019-03-28 15:44 | 318 | 0.013655 | 0.13718 | 6.4097 | 59.2522 | 24.0934 | 0.8091
2019-03-28 15:44 | 319 | 0.015141 | 0.12348 | 5.7022 | 54.4396 | 24.0906 | 0.82425
2019-03-28 15:44 | 320 | 0.047811 | 0.039093 | 1.7845 | 16.7379 | 24.0554 | 0.80025
2019-03-28 15:44 | 321 | 0.040556 | 0.045591 | 2.082 | 19.9712 | 24.0043 | 0.80995
2019-03-28 15:44 | 322 | 0.026499 | 0.06797 | 3.1549 | 29.0894 | 23.9765 | 0.77083
2019-03-28 15:44 | 323 | 0.031308 | 0.057609 | 2.6502 | 24.4453 | 23.9463 | 0.76533
2019-03-28 15:44 | 324 | 0.056372 | 0.031646 | 1.4617 | 13.5442 | 23.9095 | 0.76351
2019-03-28 15:44 | 325 | 0.038839 | 0.046017 | 2.1053 | 19.0126 | 23.8441 | 0.73843
2019-03-28 15:44 | 326 | 0.0091717 | 0.19221 | 8.8019 | 80.185 | 23.8351 | 0.73543
2019-03-28 15:44 | 327 | 0.013056 | 0.13357 | 6.1469 | 55.2334 | 23.8168 | 0.72115
2019-03-28 15:44 | 328 | 0.044447 | 0.039792 | 1.7893 | 16.1748 | 23.8015 | 0.71892
2019-03-28 15:44 | 329 | 0.060697 | 0 | 0 | 0 | 23.7324 | 0.71311
2019-03-28 15:44 | 330 | 0.022918 | 0 | 0 | 0 | 23.6931 | 0.69769
2019-03-28 15:44 | 331 | 0.024598 | 0 | 0 | 0 | 23.6781 | 0.71596
2019-03-28 15:44 | 332 | 0.031466 | 0 | 0 | 0 | 23.6418 | 0.69925
2019-03-28 15:44 | 333 | 0.032691 | 0 | 0 | 0 | 23.6171 | 0.704
2019-03-28 15:44 | 334 | 0.003372 | 0 | 0 | 0 | 23.5829 | 0.69042
2019-03-28 15:44 | 335 | 0.018349 | 0 | 0 | 0 | 23.5995 | 0.66692
2019-03-28 15:44 | 336 | 0.053256 | 0 | 0 | 0 | 23.5412 | 0.65513
2019-03-28 15:44 | 337 | 0.044968 | 0 | 0 | 0 | 23.5033 | 0.66074
2019-03-28 15:44 | 338 | 0.035788 | 0 | 0 | 0 | 23.4529 | 0.65003
2019-03-28 15:44 | 339 | 0.020457 | 0 | 0 | 0 | 23.4325 | 0.67123
2019-03-28 15:44 | 340 | 0.039615 | 0 | 0 | 0 | 23.4041 | 0.66418
2019-03-28 15:44 | 341 | 0.034148 | 0 | 0 | 0 | 23.3584 | 0.63562
2019-03-28 15:44 | 342 | 0.033271 | 0 | 0 | 0 | 23.3356 | 0.63398
2019-03-28 15:44 | 343 | 0.018038 | 0 | 0 | 0 | 23.2946 | 0.62952
2019-03-28 15:44 | 344 | 0.033945 | 0 | 0 | 0 | 23.2924 | 0.61549
2019-03-28 15:44 | 345 | 0.039651 | 0 | 0 | 0 | 23.2293 | 0.61929
2019-03-28 15:44 | 346 | 0.013743 | 0 | 0 | 0 | 23.2205 | 0.62082
2019-03-28 15:44 | 347 | 0.018367 | 0 | 0 | 0 | 23.1954 | 0.62076
2019-03-28 15:44 | 348 | 0.0072512 | 0 | 0 | 0 | 23.1861 | 0.60813
2019-03-28 15:44 | 349 | 0.029037 | 0 | 0 | 0 | 23.174 | 0.5922
2019-03-28 15:44 | 350 | 0.010864 | 0 | 0 | 0 | 23.1364 | 0.60017
2019-03-28 15:44 | 351 | 0.016987 | 0 | 0 | 0 | 23.1467 | 0.59048
2019-03-28 15:44 | 352 | 0.040742 | 0 | 0 | 0 | 23.0979 | 0.57544
2019-03-28 15:44 | 353 | 0.044469 | 0 | 0 | 0 | 23.0699 | 0.5805
2019-03-28 15:44 | 354 | 0.042993 | 0 | 0 | 0 | 23.0114 | 0.5883
2019-03-28 15:44 | 355 | 0.024361 | 0 | 0 | 0 | 22.9877 | 0.60391
2019-03-28 15:44 | 356 | 0.035096 | 0 | 0 | 0 | 22.9559 | 0.59125
2019-03-28 15:44 | 357 | 0.039587 | 0 | 0 | 0 | 22.9185 | 0.5867
2019-03-28 15:44 | 358 | 0.036913 | 0 | 0 | 0 | 22.8791 | 0.57234
2019-03-28 15:44 | 359 | 0.016744 | 0 | 0 | 0 | 22.8484 | 0.56428
2019-03-28 15:44 | 360 | 0.028631 | 0 | 0 | 0 | 22.8384 | 0.57131
2019-03-28 15:44 | 361 | 0.027783 | 0 | 0 | 0 | 22.7937 | 0.53838
2019-03-28 15:44 | 362 | 0.017099 | 0 | 0 | 0 | 22.7851 | 0.54344
2019-03-28 15:44 | 363 | 0.026288 | 0 | 0 | 0 | 22.7551 | 0.54865
2019-03-28 15:44 | 364 | 0.024198 | 0 | 0 | 0 | 22.7345 | 0.54939
2019-03-28 15:44 | 365 | 0.034555 | 0 | 0 | 0 | 22.7044 | 0.53456
2019-03-28 15:44 | 366 | 0.026719 | 0 | 0 | 0 | 22.6697 | 0.58433
2019-03-28 15:44 | 367 | 0.01541 | 0 | 0 | 0 | 22.6518 | 0.53764
2019-03-28 15:44 | 368 | 0.022309 | 0 | 0 | 0 | 22.6343 | 0.553
2019-03-28 15:44 | 369 | 0.033477 | 0 | 0 | 0 | 22.6062 | 0.52878
2019-03-28 15:44 | 370 | 0.01988 | 0 | 0 | 0 | 22.5731 | 0.53693
2019-03-28 15:44 | 371 | 0.021488 | 0 | 0 | 0 | 22.5636 | 0.53026
2019-03-28 15:45 | 372 | 0.017257 | 0 | 0 | 0 | 22.5317 | 0.48773
2019-03-28 15:45 | 373 | -0.0054273 | 0 | 0 | 0 | 22.533 | 0.51234
2019-03-28 15:45 | 374 | -0.00012167 | 0 | 0 | 0 | 22.5357 | 0.53471
2019-03-28 15:45 | 375 | 0.011264 | 0 | 0 | 0 | 22.5308 | 0.47444
2019-03-28 15:45 | 376 | 0.028726 | 0 | 0 | 0 | 22.5125 | 0.46849
2019-03-28 15:45 | 377 | 0.010248 | 0 | 0 | 0 | 22.4819 | 0.37266
2019-03-28 15:45 | 378 | 0.0042854 | 0 | 0 | 0 | 22.4892 | 0.43898
2019-03-28 15:45 | 379 | 0.020249 | 0 | 0 | 0 | 22.4673 | 0.47214
2019-03-28 15:45 | 380 | 0.042933 | 0 | 0 | 0 | 22.4473 | 0.46026
2019-03-28 15:45 | 381 | 0.031702 | 0 | 0 | 0 | 22.3905 | 0.50417
2019-03-28 15:45 | 382 | -0.0051065 | 0 | 0 | 0 | 22.3893 | 0.4879
2019-03-28 15:45 | 383 | 0.033169 | 0 | 0 | 0 | 22.3835 | 0.48992
2019-03-28 15:45 | 384 | 0.037713 | 0 | 0 | 0 | 22.3301 | 0.472
2019-03-28 15:45 | 385 | 0.016057 | 0 | 0 | 0 | 22.3148 | 0.48963
2019-03-28 15:45 | 386 | 0.016335 | 0 | 0 | 0 | 22.2928 | 0.48772
2019-03-28 15:45 | 387 | 0.033623 | 0 | 0 | 0 | 22.2775 | 0.47929
2019-03-28 15:45 | 388 | 0.04663 | 0 | 0 | 0 | 22.2271 | 0.46374
2019-03-28 15:45 | 389 | 0.023807 | 0 | 0 | 0 | 22.1931 | 0.4717
2019-03-28 15:45 | 390 | 0.0091558 | 0 | 0 | 0 | 22.1776 | 0.45176
2019-03-28 15:45 | 391 | 0.021666 | 0 | 0 | 0 | 22.1674 | 0.45672
2019-03-28 15:45 | 392 | 0.050815 | 0 | 0 | 0 | 22.1311 | 0.46709
2019-03-28 15:45 | 393 | 0.017932 | 0 | 0 | 0 | 22.0803 | 0.45152
2019-03-28 15:45 | 394 | 0.0037426 | 0 | 0 | 0 | 22.0911 | 0.46183
2019-03-28 15:45 | 395 | 0.029159 | 0 | 0 | 0 | 22.0632 | 0.45738
2019-03-28 15:45 | 396 | 0.027583 | 0 | 0 | 0 | 22.0386 | 0.45274
2019-03-28 15:45 | 397 | 0.022365 | 0 | 0 | 0 | 22.0098 | 0.45747
2019-03-28 15:45 | 398 | 0.0088171 | 0 | 0 | 0 | 21.9952 | 0.4552
2019-03-28 15:45 | 399 | 0.027821 | 0 | 0 | 0 | 21.9848 | 0.44368
2019-03-28 15:45 | 400 | 0.024006 | 0 | 0 | 0 | 21.9447 | 0.42815
2019-03-28 15:45 | 401 | 0.0080516 | 0 | 0 | 0 | 21.9396 | 0.42706
2019-03-28 15:45 | 402 | 0.022604 | 0 | 0 | 0 | 21.9214 | 0.42314
2019-03-28 15:45 | 403 | 0.034827 | 0 | 0 | 0 | 21.895 | 0.43506
2019-03-28 15:45 | 404 | 0.017181 | 0 | 0 | 0 | 21.8592 | 0.43855
2019-03-28 15:45 | 405 | 0.0050554 | 0 | 0 | 0 | 21.8595 | 0.42255
2019-03-28 15:45 | 406 | 0.015441 | 0 | 0 | 0 | 21.8435 | 0.4224
2019-03-28 15:45 | 407 | 0.015824 | 0 | 0 | 0 | 21.8301 | 0.4118
2019-03-28 15:45 | 408 | 0.029581 | 0 | 0 | 0 | 21.8092 | 0.41819
2019-03-28 15:45 | 409 | 0.024157 | 0 | 0 | 0 | 21.7755 | 0.41122
2019-03-28 15:45 | 410 | 0.015049 | 0 | 0 | 0 | 21.7619 | 0.41468
2019-03-28 15:45 | 411 | 0.023374 | 0 | 0 | 0 | 21.7412 | 0.41114
2019-03-28 15:45 | 412 | 0.029398 | 0 | 0 | 0 | 21.7155 | 0.44327
2019-03-28 15:45 | 413 | 0.023851 | 0 | 0 | 0 | 21.6849 | 0.42839
2019-03-28 15:45 | 414 | 0.034031 | 0 | 0 | 0 | 21.6643 | 0.428
2019-03-28 15:45 | 415 | 0.042367 | 0 | 0 | 0 | 21.6179 | 0.42421
2019-03-28 15:45 | 416 | 0.016828 | 0 | 0 | 0 | 21.5874 | 0.4186
2019-03-28 15:45 | 417 | 0.023726 | 0 | 0 | 0 | 21.5777 | 0.39492
2019-03-28 15:45 | 418 | 0.0060317 | 0 | 0 | 0 | 21.5449 | 0.39171
2019-03-28 15:45 | 419 | -0.002016 | 0 | 0 | 0 | 21.5626 | 0.39438
2019-03-28 15:45 | 420 | 0.03046 | 0 | 0 | 0 | 21.5392 | 0.39761
2019-03-28 15:45 | 421 | 0.034549 | 0 | 0 | 0 | 21.5083 | 0.39507
2019-03-28 15:45 | 422 | 0.018453 | 0 | 0 | 0 | 21.4755 | 0.39614
2019-03-28 15:45 | 423 | 0.018835 | 0 | 0 | 0 | 21.4677 | 0.38407
2019-03-28 15:45 | 424 | 0.029102 | 0 | 0 | 0 | 21.4359 | 0.389
2019-03-28 15:45 | 425 | 0.0035475 | 0 | 0 | 0 | 21.4172 | 0.38829
2019-03-28 15:45 | 426 | 0.013229 | 0 | 0 | 0 | 21.4209 | 0.38674
2019-03-28 15:45 | 427 | 0.022929 | 0 | 0 | 0 | 21.39 | 0.3869
2019-03-28 15:45 | 428 | 0.023007 | 0 | 0 | 0 | 21.3773 | 0.38445
2019-03-28 15:45 | 429 | 0.02523 | 0 | 0 | 0 | 21.3443 | 0.37696
2019-03-28 15:45 | 430 | 0.010004 | 0 | 0 | 0 | 21.3308 | 0.37424

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062482 | 0.062688 | 0 | 0.2095 | 2019-03-28 11:01 | Group A | 0.1 | 0.22 | Do Nothing | 0.20949 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025283 | 0.02541 | 0 | 0.0261 | 2019-03-28 11:01 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044597 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002743 | 0.002756 | 0 | 0.0011 | 2019-03-28 11:01 | Group A | -0.01 | 0.1 | Do Nothing | -1.2523e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.6589 | 183.6167 | 0 | 100 | 2019-03-28 11:01 | None | 0 | 1000 | Do Nothing | 101.202 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.412 | 209.4304 | 0 | 100 | 2019-03-28 11:01 | None | 0 | 1000 | Do Nothing | 101.2955 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -16.6878 | 1447 | 0 |  | 2019-03-28 10:45 | None | -10 | 1000 | Do Nothing | 126.7778 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2612 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001866 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017502 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.9737 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23377 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4027 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0026686 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0039749 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0037907 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 4.9382e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0012555 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.30797 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3247 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5474 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.63 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 21.1552 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 25.9138 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 66.7258 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 86.0172 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 63.9413 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0064733 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0473 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025319 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.1743 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN