## Cone Summary
Date: | 2019-04-09
Time: | 15:15
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | C0S-Test3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\04_Apr\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0403 |  | Start Test | 0 | 2019-04-09 15:26
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 5 | 2019-04-09 15:26
Sample Area | 1 | m^2 | Wetting | 92 | 2019-04-09 15:27
Cone Heat Flux | 75 | kW/m^2 | Dripping | 93 | 2019-04-09 15:28
Cone Orientation |  |  | Flame Out | 270 | 2019-04-09 15:30
Sample Grid
Sample Thickness | 106 | mm
Sample Initial Mass | 43.6 | g
Cone Separation Distance | 25 | mm
Exhaust Flow Rate | 23.1937 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 12.5942 | kW/m2 | Average Heat Release Rate | 5.9423 | kW/m2
Total Heat Release | 1.0972 | MJ/m2 | Average Mass Loss Rate | 0.23805 | g/m2s
Peak Mass Loss Rate | 0.49934 | g/m2s | Effective Heat of Combustion | 25.0514 | MJ/kg
Total Mass Lost | 43.8825 | g | Average Specific Extinction Area | 147.8211 | m2/kg
Time to Ignition | 5 | s | Peak CO Yield | 0.036929 | kg/kg
Time to Peak Heat Release | 122 | s | Peak CO2 Yield | 3.4462 | kg/kg
Time to Flame Out | 270 | s | Average CO Yield | 0.01793 | kg/kg
Average Heat Release Rate | 4.1249 | kW/m2 | Average CO2 Yield | 2.144 | kg/kg
Average Mass Loss Rate | 0.16722 | g/m2s
Effective Heat of Combustion | 25.0039 | MJ/kg
Dripping | 93 | s | Time to dripping
HRRD | 2.6585 | kW | HRR at dripping time
THRD | 0.23821 | MJ | Heat released from Ignition to Dripping
AHRRD | 2.7069 | kW | AHRR from Ignition to Dripping
Wetting | 92 | s | Time to wetting
HRRW | 2.6932 | kW | HRR at wetting time
THRW | 0.23555 | MJ | Heat released from Ignition to Wetting
AHRRW | 2.6805 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-04-09 15:15 | 1554837277.7048 | -681 | 0.20951 | 0.00045 | 3e-06 | 95.3546 | 95.2731 | 122.4538 | 20.6926 | 0.001912 | 0.001823 | -154.6834 | 0.23222 | 8.4036 | 0.002669 | 0.003734 | 0.003478 | 3.5e-05 | 0.001305 | 0.28335 | 1.314 | 2.5473 | 98.49 | 20.07 | 47.736 | 67.2865 | 85.4705 | 65.1493 | 0.011393 | 1.0436 | 0.024246 | 23.2333 | 0.007656 | 0.13815 | -3e-06 | -0.000359 | -0.000359 | -15.5046
2019-04-09 15:15 | 1554837278.7049 | -680 | 0.20951 | 0.00045 | 3e-06 | 95.3546 | 95.2731 | 122.4538 | 20.6926 | 0.001912 | 0.001823 | -154.6834 | 0.23222 | 8.4036 | 0.002669 | 0.003734 | 0.003478 | 3.5e-05 | 0.001305 | 0.28335 | 1.314 | 2.5473 | 98.49 | 20.07 | 47.739 | 67.2865 | 85.4705 | 65.1493 | 0.011394 | 1.0436 | 0.024246 | 23.2333 | 0.007656 | 0.13815 | -3e-06 | -0.000359 | -0.000359 | -15.5038
2019-04-09 15:15 | 1554837279.7038 | -679 | 0.20949 | 0.000443 | 2e-06 | 95.2737 | 95.1736 | 122.4538 | 20.6934 | 0.001934 | 0.001809 | -154.6834 | 0.23271 | 8.4024 | 0.00261 | 0.004309 | 0.004178 | 3.2e-05 | 0.001308 | 0.30469 | 1.3219 | 2.5475 | 98.49 | 20.07 | 47.742 | 67.8177 | 84.0653 | 64.8104 | 0.011395 | 1.0446 | 0.024258 | 23.2217 | 0.009397 | 0.13844 | 6e-06 | 0.000281 | 0.000281 | 19.8293
2019-04-09 15:15 | 1554837280.7053 | -678 | 0.20952 | 0.000446 | 3e-06 | 95.2605 | 95.0807 | 122.4538 | 20.6926 | 0.001905 | 0.001756 | -154.6834 | 0.23229 | 8.4019 | 0.002538 | 0.004314 | 0.004142 | 3.3e-05 | 0.00131 | 0.28766 | 1.2873 | 2.5475 | 98.49 | 20.07 | 47.745 | 67.1016 | 82.3239 | 63.5159 | 0.011395 | 1.0487 | 0.024305 | 23.1772 | 0.016904 | 0.13819 | 1e-06 | -0.000106 | -0.000106 | -52.3805
2019-04-09 15:15 | 1554837281.7053 | -677 | 0.20948 | 0.000445 | 2e-06 | 95.1861 | 95.2552 | 122.4538 | 20.6928 | 0.001869 | 0.001768 | -154.6834 | 0.23223 | 8.4028 | 0.00262 | 0.003701 | 0.003474 | 3.4e-05 | 0.001307 | 0.2907 | 1.293 | 2.5475 | 98.49 | 20.07 | 47.748 | 66.2358 | 84.2948 | 63.8191 | 0.011396 | 1.0477 | 0.024294 | 23.1876 | -0.006488 | 0.13816 | 3e-06 | 9.3e-05 | 9.3e-05 | 59.7626
2019-04-09 15:15 | 1554837282.7049 | -676 | 0.20953 | 0.000443 | 2e-06 | 95.156 | 94.9444 | 122.4538 | 20.6934 | 0.001912 | 0.001785 | -154.6834 | 0.23353 | 8.4026 | 0.002651 | 0.00435 | 0.004151 | 3.7e-05 | 0.001309 | 0.28672 | 1.3363 | 2.5475 | 98.49 | 20.07 | 47.75 | 67.2921 | 85.0523 | 64.2171 | 0.011397 | 1.0465 | 0.02428 | 23.2013 | 0.019915 | 0.13893 | 5e-06 | 0.000239 | 0.000239 | 23.4308
2019-04-09 15:15 | 1554837283.7054 | -675 | 0.20951 | 0.000443 | 2e-06 | 95.0997 | 94.9571 | 122.4538 | 20.6928 | 0.001818 | 0.001772 | -154.6834 | 0.23234 | 8.3997 | 0.002602 | 0.004399 | 0.004187 | 3.4e-05 | 0.001305 | 0.27815 | 1.343 | 2.5474 | 98.49 | 20.07 | 47.754 | 65.0299 | 83.868 | 63.92 | 0.011397 | 1.0474 | 0.02429 | 23.1911 | 0.013423 | 0.13822 | 6e-06 | 0.000344 | 0.000344 | 16.2094
2019-04-09 15:15 | 1554837284.7045 | -674 | 0.20954 | 0.000443 | 2e-06 | 95.0667 | 94.7496 | 122.4538 | 20.6934 | 0.001919 | 0.001761 | -154.6834 | 0.23317 | 8.4012 | 0.002647 | 0.003547 | 0.003438 | 3.6e-05 | 0.001302 | 0.28982 | 1.3643 | 2.5475 | 98.49 | 20.07 | 47.756 | 67.4613 | 84.9524 | 63.6572 | 0.011398 | 1.0482 | 0.0243 | 23.182 | 0.029883 | 0.13871 | 6e-06 | 0.000307 | 0.000307 | 18.2382
2019-04-09 15:15 | 1554837285.7041 | -673 | 0.20953 | 0.000446 | 2e-06 | 95.0586 | 94.9117 | 122.4538 | 20.6932 | 0.001934 | 0.001831 | -154.6834 | 0.23256 | 8.4006 | 0.002664 | 0.00444 | 0.00414 | 3.3e-05 | 0.001298 | 0.30735 | 1.3576 | 2.5476 | 98.49 | 20.07 | 47.76 | 67.8129 | 85.3657 | 65.3339 | 0.011399 | 1.043 | 0.024239 | 23.2397 | 0.013837 | 0.13835 | 3e-06 | 9.7e-05 | 9.7e-05 | 57.5487
2019-04-09 15:15 | 1554837286.7041 | -672 | 0.20954 | 0.000441 | 2e-06 | 94.9881 | 94.9193 | 122.4538 | 20.6928 | 0.001926 | 0.001758 | -154.6834 | 0.23268 | 8.4021 | 0.002666 | 0.004482 | 0.004286 | 3.1e-05 | 0.001301 | 0.31094 | 1.3657 | 2.5476 | 98.49 | 20.07 | 47.762 | 67.6183 | 85.3935 | 63.5655 | 0.011399 | 1.0485 | 0.024303 | 23.1789 | 0.00648 | 0.13842 | 9e-06 | 0.000529 | 0.000529 | 10.5416
2019-04-09 15:15 | 1554837287.7041 | -671 | 0.20955 | 0.000446 | 2e-06 | 94.8934 | 94.6811 | 122.4538 | 20.6939 | 0.00192 | 0.001765 | -154.6834 | 0.23312 | 8.401 | 0.002681 | 0.003581 | 0.003315 | 3.8e-05 | 0.001303 | 0.28368 | 1.3689 | 2.5476 | 98.49 | 20.07 | 47.763 | 67.465 | 85.7596 | 63.7568 | 0.0114 | 1.0479 | 0.024296 | 23.1855 | 0.020029 | 0.13868 | 2e-06 | 1.6e-05 | 1.6e-05 | 342.7678
2019-04-09 15:15 | 1554837288.7046 | -670 | 0.20953 | 0.000442 | 2e-06 | 94.8422 | 94.6433 | 122.7493 | 20.6929 | 0.001874 | 0.001726 | -154.6869 | 0.23248 | 8.4014 | 0.002762 | 0.00445 | 0.004315 | 3.8e-05 | 0.001296 | 0.28008 | 1.3295 | 2.5475 | 98.49 | 20.07 | 47.766 | 66.3701 | 87.7044 | 62.8149 | 0.0114 | 1.0508 | 0.024359 | 23.1809 | 0.018777 | 0.1383 | 7e-06 | 0.000349 | 0.000349 | 15.9838
2019-04-09 15:15 | 1554837289.7048 | -669 | 0.20954 | 0.000444 | 2e-06 | 94.8049 | 94.4581 | 122.7805 | 20.6922 | 0.001854 | 0.001778 | -154.6934 | 0.23314 | 8.4015 | 0.002679 | 0.004177 | 0.004004 | 3.1e-05 | 0.00129 | 0.29767 | 1.3142 | 2.5475 | 98.49 | 20.07 | 47.766 | 65.8761 | 85.7089 | 64.0575 | 0.0114 | 1.047 | 0.024318 | 23.2267 | 0.032781 | 0.1387 | 5e-06 | 0.000223 | 0.000223 | 25.0618
2019-04-09 15:15 | 1554837290.7046 | -668 | 0.20954 | 0.000449 | 2e-06 | 94.6367 | 94.6247 | 122.6489 | 20.693 | 0.001869 | 0.001784 | -154.6738 | 0.23252 | 8.4003 | 0.002661 | 0.003817 | 0.00355 | 3.6e-05 | 0.001296 | 0.31166 | 1.3529 | 2.5476 | 98.49 | 20.07 | 47.767 | 66.2485 | 85.2727 | 64.2051 | 0.011401 | 1.0465 | 0.024299 | 23.2193 | 0.001134 | 0.13833 | -1e-06 | -0.000209 | -0.000209 | -26.7008
2019-04-09 15:15 | 1554837291.7042 | -667 | 0.20953 | 0.000443 | 2e-06 | 94.6161 | 94.4113 | 122.4205 | 20.6932 | 0.001954 | 0.001794 | -154.6664 | 0.23263 | 8.4036 | 0.002648 | 0.004328 | 0.004155 | 3.5e-05 | 0.0013 | 0.27328 | 1.369 | 2.5476 | 98.49 | 20.07 | 47.767 | 68.2918 | 84.9794 | 64.4519 | 0.011401 | 1.0457 | 0.024268 | 23.2062 | 0.019384 | 0.13839 | 6e-06 | 0.000316 | 0.000316 | 17.6427
2019-04-09 15:15 | 1554837292.7046 | -666 | 0.20954 | 0.000443 | 2e-06 | 94.531 | 94.222 | 122.6482 | 20.6924 | 0.001891 | 0.00174 | -154.6467 | 0.23203 | 8.4008 | 0.002616 | 0.003734 | 0.003503 | 3.4e-05 | 0.001293 | 0.28871 | 1.3326 | 2.5475 | 98.49 | 20.07 | 47.767 | 66.7683 | 84.1999 | 63.1315 | 0.011401 | 1.0499 | 0.024338 | 23.1823 | 0.02928 | 0.13803 | 6e-06 | 0.000318 | 0.000318 | 17.4976
2019-04-09 15:15 | 1554837293.7043 | -665 | 0.20951 | 0.000445 | 2e-06 | 94.4333 | 94.074 | 122.6766 | 20.6939 | 0.001925 | 0.00181 | -154.6465 | 0.23157 | 8.4009 | 0.002659 | 0.004081 | 0.003983 | 3.6e-05 | 0.001291 | 0.29391 | 1.3121 | 2.5475 | 98.49 | 20.07 | 47.766 | 67.5999 | 85.2228 | 64.8366 | 0.0114 | 1.0446 | 0.024279 | 23.2437 | 0.034092 | 0.13776 | 3e-06 | 5.2e-05 | 5.2e-05 | 107.4749
2019-04-09 15:15 | 1554837294.7047 | -664 | 0.20956 | 0.000445 | 3e-06 | 94.391 | 94.2774 | 122.6573 | 20.6938 | 0.00193 | 0.001755 | -154.6471 | 0.23225 | 8.4018 | 0.002624 | 0.00432 | 0.004068 | 3.2e-05 | 0.001293 | 0.2896 | 1.328 | 2.5475 | 98.49 | 20.07 | 47.765 | 67.7034 | 84.3932 | 63.493 | 0.0114 | 1.0487 | 0.024326 | 23.1956 | 0.010768 | 0.13816 | 2e-06 | 9e-06 | 9e-06 | 648.9683
2019-04-09 15:15 | 1554837295.7034 | -663 | 0.20957 | 0.000443 | 2e-06 | 94.4029 | 94.2933 | 122.7367 | 20.6929 | 0.001929 | 0.001725 | -154.6648 | 0.23256 | 8.4025 | 0.002731 | 0.004088 | 0.003952 | 3.2e-05 | 0.00129 | 0.27721 | 1.3319 | 2.5476 | 98.49 | 20.07 | 47.763 | 67.7004 | 86.9697 | 62.7758 | 0.0114 | 1.051 | 0.02436 | 23.1784 | 0.010395 | 0.13835 | 6e-06 | 0.000321 | 0.000321 | 17.3877
2019-04-09 15:15 | 1554837296.7044 | -662 | 0.20958 | 0.000445 | 2e-06 | 94.3609 | 94.1828 | 123.0096 | 20.6922 | 0.001889 | 0.001724 | -154.681 | 0.23226 | 8.4034 | 0.002657 | 0.004105 | 0.003899 | 3e-05 | 0.001297 | 0.29214 | 1.2992 | 2.5476 | 98.49 | 20.07 | 47.759 | 66.7188 | 85.1867 | 62.7485 | 0.011399 | 1.051 | 0.024388 | 23.2032 | 0.016901 | 0.13817 | 3e-06 | 0.000115 | 0.000115 | 48.3341
2019-04-09 15:15 | 1554837297.7044 | -661 | 0.20955 | 0.000441 | 2e-06 | 94.2814 | 94.1594 | 122.9087 | 20.693 | 0.001908 | 0.001784 | -154.6773 | 0.23279 | 8.4025 | 0.002648 | 0.003882 | 0.003713 | 3.4e-05 | 0.001294 | 0.26576 | 1.3485 | 2.5474 | 98.49 | 20.07 | 47.756 | 67.1903 | 84.9667 | 64.1957 | 0.011398 | 1.0465 | 0.024325 | 23.2436 | 0.011584 | 0.13849 | 9e-06 | 0.000502 | 0.000502 | 11.1157
2019-04-09 15:15 | 1554837298.7046 | -660 | 0.20955 | 0.000443 | 2e-06 | 94.2522 | 94.0376 | 122.6353 | 20.6937 | 0.001919 | 0.001763 | -154.6721 | 0.232 | 8.4051 | 0.002677 | 0.004018 | 0.003855 | 3.1e-05 | 0.001299 | 0.27433 | 1.3555 | 2.5475 | 98.49 | 20.07 | 47.748 | 67.4429 | 85.6714 | 63.6965 | 0.011396 | 1.0481 | 0.024316 | 23.2005 | 0.020396 | 0.13802 | 6e-06 | 0.000331 | 0.000331 | 16.7873
2019-04-09 15:15 | 1554837299.7035 | -659 | 0.20955 | 0.000446 | 2e-06 | 94.1688 | 93.8827 | 122.708 | 20.693 | 0.001928 | 0.001805 | -154.653 | 0.23237 | 8.4013 | 0.002685 | 0.004139 | 0.003947 | 3.2e-05 | 0.001296 | 0.29407 | 1.3645 | 2.5475 | 98.49 | 20.07 | 47.744 | 67.6559 | 85.8579 | 64.714 | 0.011395 | 1.0449 | 0.024287 | 23.2425 | 0.027222 | 0.13824 | 3e-06 | 0.000107 | 0.000107 | 52.1359
2019-04-09 15:15 | 1554837300.7039 | -658 | 0.20956 | 0.000445 | 2e-06 | 94.0756 | 93.6968 | 122.9049 | 20.6933 | 0.001884 | 0.001757 | -154.6524 | 0.23273 | 8.4007 | 0.002692 | 0.003926 | 0.003759 | 3.2e-05 | 0.001295 | 0.28224 | 1.3792 | 2.5475 | 98.49 | 20.07 | 47.739 | 66.6038 | 86.0247 | 63.5479 | 0.011394 | 1.0486 | 0.024348 | 23.2209 | 0.036088 | 0.13845 | 3e-06 | 0.000107 | 0.000107 | 51.935
2019-04-09 15:15 | 1554837301.7046 | -657 | 0.20956 | 0.000446 | 2e-06 | 94.0557 | 93.8025 | 122.9575 | 20.6933 | 0.001903 | 0.001815 | -154.6409 | 0.23238 | 8.4002 | 0.00269 | 0.004289 | 0.004112 | 3.2e-05 | 0.001291 | 0.28943 | 1.4251 | 2.5476 | 98.49 | 20.07 | 47.734 | 67.072 | 85.9684 | 64.9368 | 0.011393 | 1.0442 | 0.024303 | 23.2737 | 0.02411 | 0.13824 | 2e-06 | 3.1e-05 | 3.1e-05 | 177.2525
2019-04-09 15:15 | 1554837302.7036 | -656 | 0.20957 | 0.000448 | 3e-06 | 94.0661 | 93.9529 | 122.666 | 20.6931 | 0.001906 | 0.001757 | -154.6436 | 0.23261 | 8.4019 | 0.002692 | 0.00395 | 0.003735 | 3.4e-05 | 0.001287 | 0.28838 | 1.3806 | 2.5475 | 98.49 | 20.07 | 47.731 | 67.1458 | 86.0165 | 63.5633 | 0.011392 | 1.0485 | 0.024324 | 23.1989 | 0.010772 | 0.13838 | 0 | -0.000184 | -0.000184 | -30.2711
2019-04-09 15:15 | 1554837303.7046 | -655 | 0.20955 | 0.000446 | 3e-06 | 94.0191 | 93.7856 | 122.6651 | 20.6936 | 0.001843 | 0.001759 | -154.6327 | 0.2319 | 8.4035 | 0.00272 | 0.00491 | 0.004746 | 3.1e-05 | 0.001289 | 0.28296 | 1.3804 | 2.5475 | 98.49 | 20.07 | 47.729 | 65.6202 | 86.7001 | 63.5983 | 0.011392 | 1.0484 | 0.024323 | 23.2 | 0.022234 | 0.13796 | 2e-06 | -3.2e-05 | -3.2e-05 | -175.8294
2019-04-09 15:15 | 1554837304.7026 | -654 | 0.20956 | 0.000442 | 2e-06 | 93.9565 | 93.6435 | 122.6057 | 20.6934 | 0.001957 | 0.001671 | -154.6317 | 0.23198 | 8.4022 | 0.002725 | 0.004727 | 0.004563 | 3.3e-05 | 0.001291 | 0.2855 | 1.4149 | 2.5476 | 98.49 | 20.07 | 47.724 | 68.3536 | 86.8161 | 61.4826 | 0.01139 | 1.055 | 0.024394 | 23.1214 | 0.029849 | 0.13801 | 7e-06 | 0.000354 | 0.000354 | 15.7044
2019-04-09 15:15 | 1554837305.7027 | -653 | 0.20954 | 0.000443 | 2e-06 | 93.9708 | 93.68 | 122.4475 | 20.6942 | 0.001863 | 0.001765 | -154.6229 | 0.23251 | 8.4003 | 0.002659 | 0.004222 | 0.004014 | 3.3e-05 | 0.001294 | 0.27616 | 1.3598 | 2.5476 | 98.49 | 20.07 | 47.723 | 66.0967 | 85.2363 | 63.7541 | 0.01139 | 1.0479 | 0.024296 | 23.1848 | 0.027726 | 0.13832 | 6e-06 | 0.000307 | 0.000307 | 18.1558
2019-04-09 15:15 | 1554837306.7027 | -652 | 0.20956 | 0.000442 | 2e-06 | 93.9473 | 93.7058 | 122.1085 | 20.6932 | 0.001887 | 0.001755 | -154.5993 | 0.23206 | 8.4012 | 0.002681 | 0.003906 | 0.003779 | 2.9e-05 | 0.001296 | 0.28639 | 1.3648 | 2.5476 | 98.49 | 20.07 | 47.722 | 66.6715 | 85.7527 | 63.4924 | 0.01139 | 1.0487 | 0.024271 | 23.1437 | 0.023022 | 0.13805 | 7e-06 | 0.000378 | 0.000378 | 14.7293
2019-04-09 15:15 | 1554837307.7027 | -651 | 0.20956 | 0.000448 | 2e-06 | 93.8794 | 93.5833 | 122.4428 | 20.6934 | 0.001888 | 0.001779 | -154.6088 | 0.23247 | 8.4021 | 0.002663 | 0.004246 | 0.004008 | 3.5e-05 | 0.001291 | 0.27527 | 1.3222 | 2.5476 | 98.49 | 20.07 | 47.723 | 66.6952 | 85.3269 | 64.0921 | 0.01139 | 1.0469 | 0.024283 | 23.1959 | 0.028258 | 0.13829 | 0 | -0.000179 | -0.000179 | -31.2103
2019-04-09 15:15 | 1554837308.7037 | -650 | 0.20956 | 0.000441 | 3e-06 | 93.9789 | 93.7609 | 122.543 | 20.6938 | 0.00184 | 0.001798 | -154.6298 | 0.23179 | 8.4018 | 0.002704 | 0.003188 | 0.003069 | 3.3e-05 | 0.001276 | 0.30397 | 1.2966 | 2.5474 | 98.49 | 20.07 | 47.725 | 65.5407 | 86.3257 | 64.5391 | 0.011391 | 1.0455 | 0.024277 | 23.2208 | 0.020772 | 0.13789 | 8e-06 | 0.000432 | 0.000432 | 12.872
2019-04-09 15:15 | 1554837309.7028 | -649 | 0.20953 | 0.000444 | 3e-06 | 93.8952 | 93.7798 | 122.4939 | 20.694 | 0.001871 | 0.001782 | -154.6277 | 0.23245 | 8.4026 | 0.002689 | 0.004263 | 0.004027 | 3.2e-05 | 0.001267 | 0.28727 | 1.3288 | 2.5475 | 98.49 | 20.07 | 47.727 | 66.3048 | 85.9642 | 64.1473 | 0.011391 | 1.0467 | 0.024286 | 23.2027 | 0.010999 | 0.13829 | 4e-06 | 0.000152 | 0.000152 | 36.5505
2019-04-09 15:15 | 1554837310.7028 | -648 | 0.20953 | 0.000445 | 3e-06 | 93.7707 | 93.5289 | 122.5303 | 20.6933 | 0.001787 | 0.001744 | -154.6324 | 0.23331 | 8.4008 | 0.002596 | 0.003945 | 0.003743 | 3.4e-05 | 0.001257 | 0.2724 | 1.3774 | 2.5476 | 98.49 | 20.07 | 47.729 | 64.267 | 83.7131 | 63.2309 | 0.011392 | 1.0495 | 0.024323 | 23.1746 | 0.023095 | 0.1388 | 3e-06 | 4.4e-05 | 4.4e-05 | 127.3952
2019-04-09 15:15 | 1554837311.7039 | -647 | 0.20958 | 0.000444 | 2e-06 | 93.7168 | 93.4036 | 122.9488 | 20.6946 | 0.001883 | 0.001747 | -154.6398 | 0.23249 | 8.4026 | 0.002743 | 0.004078 | 0.003909 | 3.6e-05 | 0.00126 | 0.29878 | 1.3697 | 2.5476 | 98.49 | 20.07 | 47.732 | 66.5849 | 87.2664 | 63.3157 | 0.011392 | 1.0493 | 0.024361 | 23.2171 | 0.02994 | 0.13831 | 5e-06 | 0.000275 | 0.000275 | 20.2572
2019-04-09 15:15 | 1554837312.7029 | -646 | 0.20958 | 0.000446 | 2e-06 | 93.69 | 93.4053 | 123.2204 | 20.6941 | 0.001921 | 0.001743 | -154.6446 | 0.23298 | 8.4007 | 0.002725 | 0.00434 | 0.004134 | 3.5e-05 | 0.001267 | 0.28766 | 1.3474 | 2.5475 | 98.49 | 20.07 | 47.733 | 67.4917 | 86.8303 | 63.2042 | 0.011392 | 1.0496 | 0.024392 | 23.2388 | 0.027217 | 0.1386 | 3e-06 | 5.2e-05 | 5.2e-05 | 106.8131
2019-04-09 15:15 | 1554837313.7029 | -645 | 0.20956 | 0.000444 | 2e-06 | 93.6534 | 93.0743 | 123.2573 | 20.6936 | 0.001923 | 0.001707 | -154.6535 | 0.2331 | 8.3998 | 0.002681 | 0.00448 | 0.004298 | 3.9e-05 | 0.001266 | 0.27671 | 1.3395 | 2.5475 | 98.49 | 20.07 | 47.732 | 67.5426 | 85.7751 | 62.3476 | 0.011392 | 1.0523 | 0.024427 | 23.2127 | 0.05548 | 0.13867 | 5e-06 | 0.000205 | 0.000205 | 27.2854
2019-04-09 15:15 | 1554837314.7029 | -644 | 0.20956 | 0.000446 | 1e-06 | 93.6322 | 93.4535 | 123.5291 | 20.6951 | 0.001929 | 0.001778 | -154.6357 | 0.2318 | 8.4029 | 0.002753 | 0.004443 | 0.004283 | 3.6e-05 | 0.001274 | 0.27345 | 1.3736 | 2.5474 | 98.49 | 20.07 | 47.732 | 67.682 | 87.4956 | 64.0586 | 0.011392 | 1.047 | 0.024392 | 23.2975 | 0.017087 | 0.13789 | 3e-06 | 0.000159 | 0.000159 | 34.9406
2019-04-09 15:15 | 1554837315.7033 | -643 | 0.2096 | 0.000444 | 2e-06 | 93.567 | 93.3544 | 123.5441 | 20.6936 | 0.001863 | 0.001796 | -154.6394 | 0.23247 | 8.4041 | 0.00272 | 0.003606 | 0.003392 | 3.4e-05 | 0.001276 | 0.29806 | 1.3471 | 2.5475 | 98.49 | 20.07 | 47.733 | 66.0961 | 86.707 | 64.4961 | 0.011392 | 1.0456 | 0.024377 | 23.314 | 0.020347 | 0.1383 | 4e-06 | 0.00019 | 0.00019 | 29.3481
2019-04-09 15:15 | 1554837316.703 | -642 | 0.20956 | 0.000446 | 2e-06 | 93.5873 | 93.3245 | 123.4145 | 20.6939 | 0.001855 | 0.001787 | -154.6488 | 0.23185 | 8.4011 | 0.002707 | 0.00356 | 0.003429 | 3.3e-05 | 0.001284 | 0.28694 | 1.3232 | 2.5475 | 98.49 | 20.07 | 47.732 | 65.9153 | 86.3919 | 64.2793 | 0.011392 | 1.0463 | 0.024372 | 23.2943 | 0.025151 | 0.13793 | 2e-06 | 2.6e-05 | 2.6e-05 | 216.1814
2019-04-09 15:15 | 1554837317.703 | -641 | 0.20957 | 0.000438 | 2e-06 | 93.559 | 93.1341 | 123.3002 | 20.6945 | 0.001906 | 0.001759 | -154.6498 | 0.23201 | 8.4038 | 0.002668 | 0.00398 | 0.003792 | 3.5e-05 | 0.001283 | 0.2677 | 1.3522 | 2.5475 | 98.49 | 20.07 | 47.73 | 67.1302 | 85.4536 | 63.596 | 0.011392 | 1.0484 | 0.024386 | 23.2599 | 0.040707 | 0.13802 | 1.2e-05 | 0.000771 | 0.000771 | 7.21
2019-04-09 15:15 | 1554837318.7027 | -640 | 0.20959 | 0.000437 | 2e-06 | 93.4952 | 93.2185 | 123.2174 | 20.6931 | 0.001875 | 0.001798 | -154.627 | 0.23261 | 8.4003 | 0.002752 | 0.004076 | 0.003831 | 3.5e-05 | 0.001289 | 0.28102 | 1.3627 | 2.5475 | 98.49 | 20.07 | 47.73 | 66.3867 | 87.4608 | 64.5464 | 0.011392 | 1.0455 | 0.024343 | 23.2849 | 0.026518 | 0.13838 | 1.4e-05 | 0.000915 | 0.000915 | 6.0947
2019-04-09 15:15 | 1554837319.7031 | -639 | 0.20958 | 0.000443 | 2e-06 | 93.5304 | 93.3565 | 123.0225 | 20.6944 | 0.001896 | 0.001741 | -154.6238 | 0.23198 | 8.4009 | 0.002737 | 0.004351 | 0.004211 | 3.1e-05 | 0.001291 | 0.28694 | 1.3488 | 2.5476 | 98.49 | 20.07 | 47.726 | 66.8912 | 87.1123 | 63.1599 | 0.011391 | 1.0498 | 0.024374 | 23.2186 | 0.016647 | 0.138 | 6e-06 | 0.000316 | 0.000316 | 17.5744
2019-04-09 15:15 | 1554837320.7036 | -638 | 0.20953 | 0.000442 | 2e-06 | 93.4818 | 93.2821 | 123.0002 | 20.6942 | 0.001897 | 0.001758 | -154.6244 | 0.23252 | 8.4004 | 0.002689 | 0.003837 | 0.003677 | 3.3e-05 | 0.001292 | 0.28816 | 1.3736 | 2.5477 | 98.49 | 20.07 | 47.724 | 66.9314 | 85.9599 | 63.5783 | 0.01139 | 1.0485 | 0.024357 | 23.231 | 0.019128 | 0.13833 | 7e-06 | 0.000394 | 0.000394 | 14.1562
2019-04-09 15:15 | 1554837321.7022 | -637 | 0.20953 | 0.000438 | 2e-06 | 93.4964 | 93.2092 | 122.5681 | 20.6935 | 0.001914 | 0.001673 | -154.6234 | 0.23271 | 8.4015 | 0.002676 | 0.004212 | 0.004031 | 3.2e-05 | 0.001297 | 0.29015 | 1.3462 | 2.5476 | 98.49 | 20.07 | 47.723 | 67.3225 | 85.6438 | 61.5263 | 0.01139 | 1.0549 | 0.024388 | 23.1194 | 0.027523 | 0.13844 | 1.2e-05 | 0.000771 | 0.000771 | 7.2318
2019-04-09 15:15 | 1554837322.7032 | -636 | 0.20956 | 0.000443 | 3e-06 | 93.4262 | 93.13 | 122.4709 | 20.6931 | 0.001857 | 0.00177 | -154.6375 | 0.23199 | 8.402 | 0.002727 | 0.003608 | 0.003426 | 3e-05 | 0.001297 | 0.29103 | 1.3405 | 2.5475 | 98.49 | 20.07 | 47.722 | 65.966 | 86.8809 | 63.8539 | 0.01139 | 1.0476 | 0.024294 | 23.1904 | 0.028396 | 0.13801 | 5e-06 | 0.000159 | 0.000159 | 35.0294
2019-04-09 15:15 | 1554837323.7022 | -635 | 0.20956 | 0.000439 | 2e-06 | 93.4765 | 93.0478 | 122.4921 | 20.6942 | 0.00189 | 0.001797 | -154.6315 | 0.23248 | 8.4018 | 0.00268 | 0.004351 | 0.004205 | 3.1e-05 | 0.001298 | 0.27599 | 1.3739 | 2.5476 | 98.49 | 20.07 | 47.721 | 66.7468 | 85.7484 | 64.5154 | 0.01139 | 1.0455 | 0.024273 | 23.2152 | 0.041117 | 0.13831 | 1.1e-05 | 0.000645 | 0.000645 | 8.6477
2019-04-09 15:15 | 1554837324.7033 | -634 | 0.20955 | 0.00044 | 2e-06 | 93.4247 | 93.0987 | 122.2491 | 20.6937 | 0.001949 | 0.001831 | -154.6355 | 0.23241 | 8.4008 | 0.002692 | 0.004129 | 0.003886 | 3.5e-05 | 0.001303 | 0.30973 | 1.3527 | 2.5476 | 98.49 | 20.07 | 47.718 | 68.1815 | 86.0327 | 65.3337 | 0.011389 | 1.043 | 0.024219 | 23.2202 | 0.031262 | 0.13826 | 1e-05 | 0.000564 | 0.000564 | 9.8876
2019-04-09 15:15 | 1554837325.7032 | -633 | 0.20951 | 0.000442 | 2e-06 | 93.3949 | 92.9769 | 122.3754 | 20.6939 | 0.001904 | 0.001795 | -154.6234 | 0.23189 | 8.4001 | 0.002663 | 0.004151 | 0.003942 | 3.3e-05 | 0.001305 | 0.28495 | 1.3946 | 2.5476 | 98.49 | 20.07 | 47.717 | 67.0803 | 85.3288 | 64.4596 | 0.011389 | 1.0457 | 0.024263 | 23.2022 | 0.040124 | 0.13795 | 8e-06 | 0.000418 | 0.000418 | 13.2936
2019-04-09 15:15 | 1554837326.7025 | -632 | 0.20957 | 0.000441 | 2e-06 | 93.2874 | 92.9924 | 122.8327 | 20.6946 | 0.001891 | 0.001785 | -154.6311 | 0.23215 | 8.4034 | 0.002763 | 0.004417 | 0.004232 | 3.3e-05 | 0.0013 | 0.28296 | 1.3855 | 2.5477 | 98.49 | 20.07 | 47.716 | 66.7657 | 87.7402 | 64.223 | 0.011388 | 1.0465 | 0.024317 | 23.2373 | 0.028332 | 0.13811 | 9e-06 | 0.00048 | 0.00048 | 11.6046
2019-04-09 15:15 | 1554837327.7038 | -631 | 0.20951 | 0.000435 | 2e-06 | 93.2799 | 92.7934 | 122.5871 | 20.6943 | 0.001832 | 0.001764 | -154.6331 | 0.23265 | 8.3993 | 0.002675 | 0.004049 | 0.003918 | 3.2e-05 | 0.001301 | 0.29192 | 1.3875 | 2.5477 | 98.49 | 20.07 | 47.714 | 65.3687 | 85.6306 | 63.7183 | 0.011388 | 1.048 | 0.024311 | 23.1967 | 0.046772 | 0.1384 | 1.5e-05 | 0.000986 | 0.000986 | 5.655
2019-04-09 15:15 | 1554837328.7024 | -630 | 0.20957 | 0.000436 | 2e-06 | 93.2022 | 92.7955 | 122.7852 | 20.6939 | 0.00186 | 0.001753 | -154.6179 | 0.23207 | 8.4012 | 0.002767 | 0.003631 | 0.003496 | 3e-05 | 0.001303 | 0.28838 | 1.3655 | 2.5476 | 98.49 | 20.07 | 47.714 | 66.0231 | 87.8419 | 63.4549 | 0.011388 | 1.0488 | 0.02434 | 23.2064 | 0.039117 | 0.13806 | 1.5e-05 | 0.000945 | 0.000945 | 5.8903
2019-04-09 15:15 | 1554837329.7034 | -629 | 0.20953 | 0.000442 | 2e-06 | 93.1712 | 92.7346 | 123.1268 | 20.6937 | 0.001903 | 0.00173 | -154.6281 | 0.23215 | 8.4044 | 0.002743 | 0.004071 | 0.003856 | 3.3e-05 | 0.001309 | 0.27129 | 1.3931 | 2.5475 | 98.49 | 20.07 | 47.713 | 67.0567 | 87.2491 | 62.9087 | 0.011388 | 1.0505 | 0.024393 | 23.2198 | 0.042018 | 0.13811 | 8e-06 | 0.000449 | 0.000449 | 12.3967
2019-04-09 15:16 | 1554837330.7025 | -628 | 0.20953 | 0.000444 | 2e-06 | 93.1251 | 92.7318 | 123.0786 | 20.694 | 0.001887 | 0.001782 | -154.6383 | 0.23234 | 8.3973 | 0.002767 | 0.003871 | 0.003696 | 3.5e-05 | 0.001311 | 0.28263 | 1.3362 | 2.5475 | 98.49 | 20.07 | 47.713 | 66.6701 | 87.8272 | 64.1536 | 0.011388 | 1.0467 | 0.024344 | 23.2582 | 0.037851 | 0.13822 | 4e-06 | 0.000152 | 0.000152 | 36.6793
2019-04-09 15:16 | 1554837331.7036 | -627 | 0.20955 | 0.000442 | 3e-06 | 93.1164 | 92.728 | 123.4333 | 20.6946 | 0.001939 | 0.001762 | -154.6355 | 0.23285 | 8.4029 | 0.002743 | 0.004429 | 0.00422 | 3.2e-05 | 0.001303 | 0.28871 | 1.3488 | 2.5477 | 98.49 | 20.07 | 47.712 | 67.9359 | 87.2538 | 63.6772 | 0.011387 | 1.0481 | 0.024396 | 23.2752 | 0.037387 | 0.13852 | 6e-06 | 0.000316 | 0.000316 | 17.6387
2019-04-09 15:16 | 1554837332.7035 | -626 | 0.20956 | 0.000443 | 3e-06 | 93.0703 | 92.598 | 123.5129 | 20.6941 | 0.001873 | 0.00175 | -154.6247 | 0.23254 | 8.4022 | 0.002718 | 0.004428 | 0.00422 | 3.2e-05 | 0.001294 | 0.28495 | 1.3207 | 2.5475 | 98.49 | 20.06 | 47.712 | 66.3361 | 86.663 | 63.3799 | 0.01138 | 1.0491 | 0.024415 | 23.2725 | 0.045509 | 0.13834 | 5e-06 | 0.000225 | 0.000225 | 24.8006
2019-04-09 15:16 | 1554837333.7036 | -625 | 0.20953 | 0.000442 | 2e-06 | 93.0727 | 92.6709 | 123.3885 | 20.6948 | 0.001942 | 0.001849 | -154.6163 | 0.23306 | 8.404 | 0.002676 | 0.003206 | 0.003051 | 2.9e-05 | 0.001283 | 0.27743 | 1.3319 | 2.5476 | 98.49 | 20.06 | 47.711 | 68.0143 | 85.6458 | 65.7593 | 0.01138 | 1.0417 | 0.024316 | 23.3428 | 0.038693 | 0.13865 | 7e-06 | 0.000378 | 0.000378 | 14.7714
2019-04-09 15:16 | 1554837334.7016 | -624 | 0.20961 | 0.000445 | 3e-06 | 92.9976 | 92.619 | 123.4864 | 20.6943 | 0.001826 | 0.001733 | -154.6303 | 0.23204 | 8.4003 | 0.002718 | 0.004369 | 0.00414 | 3.3e-05 | 0.00127 | 0.28335 | 1.3402 | 2.5476 | 98.49 | 20.07 | 47.709 | 65.2157 | 86.6611 | 62.9653 | 0.011387 | 1.0504 | 0.024427 | 23.2556 | 0.036496 | 0.13804 | 2e-06 | -6e-06 | -6e-06 | -993.4373
2019-04-09 15:16 | 1554837335.702 | -623 | 0.20957 | 0.000448 | 2e-06 | 92.9416 | 92.5505 | 123.4288 | 20.6951 | 0.001943 | 0.001824 | -154.6406 | 0.23237 | 8.3995 | 0.002761 | 0.004297 | 0.00411 | 3.5e-05 | 0.001265 | 0.2959 | 1.3451 | 2.5475 | 98.49 | 20.06 | 47.709 | 68.0199 | 87.6955 | 65.1775 | 0.01138 | 1.0435 | 0.024341 | 23.3266 | 0.037726 | 0.13824 | 0 | -0.000166 | -0.000166 | -33.5821
2019-04-09 15:16 | 1554837336.702 | -622 | 0.20948 | 0.000445 | 2e-06 | 92.9467 | 92.525 | 122.8822 | 20.6944 | 0.001929 | 0.001733 | -154.6194 | 0.23265 | 8.4025 | 0.002686 | 0.004159 | 0.003921 | 3.5e-05 | 0.001259 | 0.29087 | 1.3667 | 2.5477 | 98.49 | 20.06 | 47.708 | 67.6994 | 85.8781 | 62.9734 | 0.011379 | 1.0503 | 0.024367 | 23.199 | 0.040675 | 0.13841 | 4e-06 | 0.000131 | 0.000131 | 42.6147
2019-04-09 15:16 | 1554837337.7021 | -621 | 0.2095 | 0.000444 | 2e-06 | 92.9723 | 92.4379 | 122.8719 | 20.6949 | 0.001894 | 0.001789 | -154.6144 | 0.23283 | 8.4027 | 0.002674 | 0.004049 | 0.003917 | 3.2e-05 | 0.001259 | 0.30237 | 1.3645 | 2.5477 | 98.49 | 20.06 | 47.704 | 66.8386 | 85.593 | 64.3225 | 0.011378 | 1.0461 | 0.024317 | 23.2445 | 0.05156 | 0.13851 | 5e-06 | 0.000201 | 0.000201 | 27.8062
2019-04-09 15:16 | 1554837338.7027 | -620 | 0.20952 | 0.000443 | 2e-06 | 92.8773 | 92.4575 | 122.6167 | 20.6949 | 0.0019 | 0.001741 | -154.6516 | 0.23276 | 8.4 | 0.002722 | 0.004004 | 0.003809 | 3.1e-05 | 0.001251 | 0.27295 | 1.3425 | 2.5475 | 98.49 | 20.07 | 47.704 | 66.9993 | 86.7428 | 63.1636 | 0.011386 | 1.0498 | 0.024334 | 23.1804 | 0.040518 | 0.13847 | 6e-06 | 0.000259 | 0.000259 | 21.5528
2019-04-09 15:16 | 1554837339.7022 | -619 | 0.20953 | 0.000445 | 2e-06 | 92.8267 | 92.3529 | 122.5712 | 20.6944 | 0.001876 | 0.001762 | -154.6557 | 0.23204 | 8.4014 | 0.002742 | 0.003752 | 0.003592 | 3.2e-05 | 0.001252 | 0.30574 | 1.3551 | 2.5476 | 98.49 | 20.07 | 47.703 | 66.4074 | 87.2248 | 63.6693 | 0.011385 | 1.0482 | 0.024311 | 23.1935 | 0.045769 | 0.13804 | 4e-06 | 0.000144 | 0.000144 | 38.5732
2019-04-09 15:16 | 1554837340.7018 | -618 | 0.20954 | 0.000447 | 2e-06 | 92.8115 | 92.2992 | 122.455 | 20.6935 | 0.001915 | 0.001791 | -154.6423 | 0.23155 | 8.4009 | 0.002693 | 0.003724 | 0.003616 | 3.4e-05 | 0.001253 | 0.30215 | 1.3604 | 2.5476 | 98.49 | 20.07 | 47.702 | 67.3444 | 86.0482 | 64.3711 | 0.011385 | 1.046 | 0.024274 | 23.2067 | 0.049507 | 0.13775 | 0 | -0.000138 | -0.000138 | -40.2076
2019-04-09 15:16 | 1554837341.7018 | -617 | 0.20957 | 0.000441 | 2e-06 | 92.7654 | 92.3113 | 122.3949 | 20.6945 | 0.001873 | 0.00177 | -154.6506 | 0.23255 | 8.4032 | 0.002737 | 0.003994 | 0.003757 | 3.3e-05 | 0.001254 | 0.28567 | 1.3561 | 2.5475 | 98.49 | 20.07 | 47.702 | 66.3404 | 87.118 | 63.8636 | 0.011385 | 1.0476 | 0.024286 | 23.1835 | 0.04389 | 0.13834 | 8e-06 | 0.000469 | 0.000469 | 11.8796
2019-04-09 15:16 | 1554837342.7028 | -616 | 0.20954 | 0.000447 | 2e-06 | 92.8204 | 92.2641 | 122.1355 | 20.695 | 0.00191 | 0.001791 | -154.6551 | 0.23292 | 8.4006 | 0.002779 | 0.004406 | 0.004286 | 3.3e-05 | 0.001252 | 0.29999 | 1.4068 | 2.5476 | 98.49 | 20.07 | 47.702 | 67.228 | 88.1334 | 64.3685 | 0.011385 | 1.046 | 0.024242 | 23.1763 | 0.05377 | 0.13856 | 1e-06 | -8.1e-05 | -8.1e-05 | -69.3255
2019-04-09 15:16 | 1554837343.7028 | -615 | 0.20953 | 0.000444 | 2e-06 | 92.7397 | 92.302 | 122.1205 | 20.6944 | 0.001928 | 0.001839 | -154.6501 | 0.2326 | 8.4036 | 0.002762 | 0.004568 | 0.004357 | 3.5e-05 | 0.001245 | 0.28838 | 1.3551 | 2.5475 | 98.49 | 20.07 | 47.701 | 67.6549 | 87.7058 | 65.5365 | 0.011385 | 1.0424 | 0.024199 | 23.2149 | 0.042321 | 0.13837 | 5e-06 | 0.000238 | 0.000238 | 23.3851
2019-04-09 15:16 | 1554837344.7019 | -614 | 0.20954 | 0.000445 | 2e-06 | 92.6891 | 92.0686 | 122.1569 | 20.6942 | 0.001879 | 0.001771 | -154.6388 | 0.23163 | 8.4035 | 0.002759 | 0.003905 | 0.003781 | 3.7e-05 | 0.001253 | 0.30021 | 1.3524 | 2.5474 | 98.49 | 20.07 | 47.702 | 66.4854 | 87.6392 | 63.8893 | 0.011385 | 1.0475 | 0.024262 | 23.1619 | 0.060077 | 0.1378 | 4e-06 | 0.000161 | 0.000161 | 34.5969
2019-04-09 15:16 | 1554837345.7019 | -613 | 0.20959 | 0.00044 | 2e-06 | 92.6644 | 92.2228 | 122.3177 | 20.6951 | 0.001883 | 0.001757 | -154.6475 | 0.23226 | 8.4027 | 0.002771 | 0.004137 | 0.003874 | 3.5e-05 | 0.001252 | 0.31293 | 1.368 | 2.5475 | 98.49 | 20.07 | 47.704 | 66.5954 | 87.9414 | 63.544 | 0.011386 | 1.0486 | 0.02429 | 23.1652 | 0.042725 | 0.13817 | 9e-06 | 0.00052 | 0.00052 | 10.7088
2019-04-09 15:16 | 1554837346.7034 | -612 | 0.2096 | 0.000445 | 2e-06 | 92.6299 | 92.0339 | 122.4799 | 20.6947 | 0.001868 | 0.001764 | -154.6584 | 0.2329 | 8.4018 | 0.00278 | 0.003909 | 0.003767 | 3.1e-05 | 0.001256 | 0.31708 | 1.335 | 2.5474 | 98.49 | 20.07 | 47.705 | 66.2314 | 88.1543 | 63.7336 | 0.011386 | 1.048 | 0.0243 | 23.1871 | 0.057734 | 0.13855 | 3e-06 | 9e-05 | 9e-05 | 61.9838
2019-04-09 15:16 | 1554837347.702 | -611 | 0.20958 | 0.000445 | 2e-06 | 92.5566 | 92.1554 | 122.4897 | 20.6952 | 0.001898 | 0.001807 | -154.6624 | 0.23248 | 8.4022 | 0.00276 | 0.004276 | 0.004113 | 3.4e-05 | 0.001266 | 0.26792 | 1.3977 | 2.5475 | 98.49 | 20.07 | 47.705 | 66.9428 | 87.676 | 64.7599 | 0.011386 | 1.0448 | 0.024264 | 23.2233 | 0.038858 | 0.1383 | 4e-06 | 0.000174 | 0.000174 | 32.0139
2019-04-09 15:16 | 1554837348.7034 | -610 | 0.20957 | 0.000446 | 2e-06 | 92.5295 | 92.1836 | 122.6514 | 20.6947 | 0.001895 | 0.001766 | -154.6415 | 0.23273 | 8.4036 | 0.002824 | 0.004253 | 0.004049 | 3.2e-05 | 0.001266 | 0.28672 | 1.3876 | 2.5476 | 98.49 | 20.06 | 47.708 | 66.8805 | 89.2165 | 63.7781 | 0.011379 | 1.0478 | 0.024315 | 23.2049 | 0.0335 | 0.13845 | 2e-06 | 2.8e-05 | 2.8e-05 | 195.9383
2019-04-09 15:16 | 1554837349.7021 | -609 | 0.20957 | 0.000445 | 2e-06 | 92.5024 | 92.2077 | 122.6186 | 20.6943 | 0.001841 | 0.001779 | -154.6451 | 0.23265 | 8.403 | 0.002773 | 0.004083 | 0.003931 | 3.3e-05 | 0.001268 | 0.30414 | 1.394 | 2.5476 | 98.49 | 20.06 | 47.71 | 65.5676 | 87.9688 | 64.0735 | 0.01138 | 1.0469 | 0.024301 | 23.2119 | 0.028543 | 0.13841 | 3e-06 | 9.7e-05 | 9.7e-05 | 57.3452
2019-04-09 15:16 | 1554837350.7011 | -608 | 0.20958 | 0.000447 | 2e-06 | 92.33 | 91.7235 | 122.8054 | 20.6947 | 0.001861 | 0.001743 | -154.6525 | 0.23277 | 8.3994 | 0.002782 | 0.003749 | 0.003605 | 3.3e-05 | 0.001266 | 0.28854 | 1.3772 | 2.5477 | 98.49 | 20.06 | 47.709 | 66.0677 | 88.2013 | 63.2236 | 0.01138 | 1.0496 | 0.02435 | 23.2003 | 0.058951 | 0.13848 | 1e-06 | -5.2e-05 | -5.2e-05 | -106.4498
2019-04-09 15:16 | 1554837351.7021 | -607 | 0.2096 | 0.000442 | 2e-06 | 92.2752 | 91.8267 | 122.7045 | 20.6952 | 0.001864 | 0.001771 | -154.6487 | 0.23241 | 8.4009 | 0.002735 | 0.003986 | 0.003799 | 3.4e-05 | 0.001268 | 0.30093 | 1.3555 | 2.5476 | 98.49 | 20.06 | 47.708 | 66.1373 | 87.0529 | 63.9025 | 0.011379 | 1.0474 | 0.024316 | 23.2142 | 0.043578 | 0.13826 | 7e-06 | 0.000347 | 0.000347 | 16.054
2019-04-09 15:16 | 1554837352.7025 | -606 | 0.20958 | 0.000443 | 3e-06 | 92.2451 | 91.7772 | 122.7083 | 20.6949 | 0.00192 | 0.001814 | -154.6463 | 0.23252 | 8.4008 | 0.002725 | 0.004196 | 0.003986 | 3.1e-05 | 0.001275 | 0.28926 | 1.3663 | 2.5476 | 98.49 | 20.06 | 47.705 | 67.4768 | 86.8238 | 64.936 | 0.011379 | 1.0442 | 0.024279 | 23.2501 | 0.045491 | 0.13833 | 5e-06 | 0.000233 | 0.000233 | 23.9208
2019-04-09 15:16 | 1554837353.7025 | -605 | 0.20954 | 0.000444 | 2e-06 | 92.2636 | 91.8866 | 122.8099 | 20.6943 | 0.001924 | 0.001795 | -154.6349 | 0.2319 | 8.4024 | 0.002817 | 0.003924 | 0.003803 | 3.5e-05 | 0.001279 | 0.29679 | 1.3869 | 2.5476 | 98.49 | 20.06 | 47.701 | 67.5635 | 89.0459 | 64.4648 | 0.011378 | 1.0457 | 0.024306 | 23.2435 | 0.036621 | 0.13796 | 4e-06 | 0.000178 | 0.000178 | 31.212
2019-04-09 15:16 | 1554837354.7016 | -604 | 0.20959 | 0.000446 | 2e-06 | 92.2406 | 91.8277 | 122.5742 | 20.6948 | 0.001875 | 0.001778 | -154.651 | 0.23253 | 8.4034 | 0.002753 | 0.004577 | 0.004361 | 3.4e-05 | 0.001286 | 0.29175 | 1.3873 | 2.5476 | 98.49 | 20.06 | 47.693 | 66.3859 | 87.486 | 64.0562 | 0.011376 | 1.047 | 0.024297 | 23.2072 | 0.040129 | 0.13833 | 2e-06 | 4e-05 | 4e-05 | 140.6948
2019-04-09 15:16 | 1554837355.7012 | -603 | 0.20957 | 0.000444 | 2e-06 | 92.2141 | 91.6794 | 122.3153 | 20.6952 | 0.00186 | 0.001726 | -154.6438 | 0.23223 | 8.4017 | 0.002695 | 0.004334 | 0.004094 | 3.4e-05 | 0.00129 | 0.28783 | 1.3891 | 2.5477 | 98.49 | 20.06 | 47.688 | 66.0352 | 86.092 | 62.8155 | 0.011375 | 1.0508 | 0.024316 | 23.1399 | 0.052018 | 0.13816 | 4e-06 | 0.000131 | 0.000131 | 42.4605
2019-04-09 15:16 | 1554837356.7016 | -602 | 0.20957 | 0.000445 | 2e-06 | 92.1397 | 91.5362 | 122.5371 | 20.6956 | 0.001889 | 0.001711 | -154.6481 | 0.23225 | 8.4023 | 0.002658 | 0.004354 | 0.004169 | 3.3e-05 | 0.001295 | 0.29518 | 1.367 | 2.5476 | 98.49 | 20.06 | 47.683 | 66.7378 | 85.216 | 62.4504 | 0.011373 | 1.052 | 0.024352 | 23.1483 | 0.058773 | 0.13816 | 4e-06 | 0.000115 | 0.000115 | 48.3168
2019-04-09 15:16 | 1554837357.7017 | -601 | 0.20958 | 0.000443 | 2e-06 | 92.0149 | 91.4994 | 122.7378 | 20.6952 | 0.001878 | 0.001773 | -154.6514 | 0.23275 | 8.4023 | 0.00269 | 0.004051 | 0.003946 | 3.4e-05 | 0.001298 | 0.28926 | 1.3627 | 2.5476 | 98.49 | 20.06 | 47.673 | 66.4729 | 85.9835 | 63.9314 | 0.011371 | 1.0474 | 0.024318 | 23.2183 | 0.050251 | 0.13846 | 5e-06 | 0.000225 | 0.000225 | 24.7741
2019-04-09 15:16 | 1554837358.7023 | -600 | 0.20956 | 0.000448 | 3e-06 | 92.0227 | 91.2554 | 122.8251 | 20.6961 | 0.001882 | 0.001771 | -154.6513 | 0.23242 | 8.3999 | 0.00271 | 0.00424 | 0.004022 | 3.4e-05 | 0.001299 | 0.28744 | 1.3589 | 2.5476 | 98.49 | 20.06 | 47.67 | 66.5644 | 86.466 | 63.894 | 0.01137 | 1.0475 | 0.024328 | 23.2253 | 0.074888 | 0.13827 | -1e-06 | -0.000211 | -0.000211 | -26.4221
2019-04-09 15:16 | 1554837359.7013 | -599 | 0.20959 | 0.000444 | 2e-06 | 91.9006 | 91.4647 | 122.7827 | 20.6944 | 0.001917 | 0.001772 | -154.6489 | 0.23289 | 8.4016 | 0.002729 | 0.004309 | 0.004091 | 3.2e-05 | 0.001305 | 0.28727 | 1.3192 | 2.5475 | 98.49 | 20.07 | 47.665 | 67.3967 | 86.9131 | 63.9198 | 0.011376 | 1.0474 | 0.024323 | 23.2222 | 0.042527 | 0.13855 | 5e-06 | 0.000205 | 0.000205 | 27.2548
2019-04-09 15:16 | 1554837360.7024 | -598 | 0.20956 | 0.000446 | 3e-06 | 91.8252 | 91.2021 | 122.7397 | 20.6959 | 0.001856 | 0.001768 | -154.6275 | 0.23264 | 8.4021 | 0.002685 | 0.004213 | 0.00408 | 3.4e-05 | 0.001311 | 0.29645 | 1.3076 | 2.5474 | 98.49 | 20.07 | 47.662 | 65.9252 | 85.8648 | 63.8183 | 0.011376 | 1.0477 | 0.024322 | 23.2146 | 0.060907 | 0.13839 | 1e-06 | -7.8e-05 | -7.8e-05 | -71.3669
2019-04-09 15:16 | 1554837361.7018 | -597 | 0.2096 | 0.000441 | 3e-06 | 91.7189 | 91.0885 | 122.3508 | 20.6944 | 0.001916 | 0.001808 | -154.6249 | 0.23257 | 8.4015 | 0.002764 | 0.003915 | 0.003733 | 3.3e-05 | 0.001314 | 0.26272 | 1.3119 | 2.5474 | 98.49 | 20.07 | 47.66 | 67.3765 | 87.7597 | 64.7762 | 0.011375 | 1.0447 | 0.024249 | 23.2107 | 0.061691 | 0.13836 | 7e-06 | 0.000386 | 0.000386 | 14.4416
2019-04-09 15:16 | 1554837362.7015 | -596 | 0.20956 | 0.00044 | 3e-06 | 91.5423 | 91.0386 | 122.5463 | 20.6956 | 0.001886 | 0.001815 | -154.6403 | 0.23262 | 8.4018 | 0.002704 | 0.003652 | 0.00347 | 3.5e-05 | 0.001314 | 0.30702 | 1.362 | 2.5476 | 98.49 | 20.07 | 47.655 | 66.6686 | 86.3234 | 64.9461 | 0.011374 | 1.0442 | 0.024262 | 23.2351 | 0.049354 | 0.13839 | 9e-06 | 0.000479 | 0.000479 | 11.6393
2019-04-09 15:16 | 1554837363.7015 | -595 | 0.20959 | 0.000438 | 3e-06 | 91.4723 | 90.8882 | 122.3782 | 20.6956 | 0.001876 | 0.001779 | -154.6471 | 0.23183 | 8.4024 | 0.002703 | 0.004111 | 0.003915 | 3.2e-05 | 0.001313 | 0.28639 | 1.3769 | 2.5476 | 98.49 | 20.07 | 47.654 | 66.4266 | 86.3039 | 64.0796 | 0.011374 | 1.0469 | 0.024277 | 23.1894 | 0.057303 | 0.13792 | 1.1e-05 | 0.000672 | 0.000672 | 8.2673
2019-04-09 15:16 | 1554837364.7016 | -594 | 0.20953 | 0.000443 | 3e-06 | 91.3723 | 90.6139 | 122.4712 | 20.6963 | 0.001887 | 0.001749 | -154.6297 | 0.23197 | 8.4018 | 0.002674 | 0.004407 | 0.004239 | 3.3e-05 | 0.001313 | 0.29878 | 1.3873 | 2.5477 | 98.49 | 20.07 | 47.656 | 66.6762 | 85.592 | 63.3633 | 0.011374 | 1.0491 | 0.024312 | 23.1736 | 0.074548 | 0.138 | 5e-06 | 0.000184 | 0.000184 | 30.2063
2019-04-09 15:16 | 1554837365.7006 | -593 | 0.20955 | 0.000443 | 3e-06 | 91.3547 | 90.5991 | 122.8578 | 20.6955 | 0.001936 | 0.001755 | -154.6254 | 0.23167 | 8.4007 | 0.002648 | 0.004002 | 0.003797 | 3.1e-05 | 0.001316 | 0.28727 | 1.3408 | 2.5476 | 98.49 | 20.07 | 47.657 | 67.8534 | 84.9817 | 63.5055 | 0.011374 | 1.0487 | 0.024345 | 23.215 | 0.074289 | 0.13782 | 5e-06 | 0.000166 | 0.000166 | 33.4942
2019-04-09 15:16 | 1554837366.7008 | -592 | 0.20955 | 0.000441 | 2e-06 | 91.2928 | 90.6838 | 122.6181 | 20.695 | 0.001835 | 0.001796 | -154.6195 | 0.23161 | 8.4012 | 0.00268 | 0.0042 | 0.003983 | 3.5e-05 | 0.001323 | 0.29142 | 1.345 | 2.5477 | 98.49 | 20.07 | 47.659 | 65.4264 | 85.735 | 64.4903 | 0.011375 | 1.0456 | 0.024286 | 23.2262 | 0.059867 | 0.13778 | 8e-06 | 0.000451 | 0.000451 | 12.3172
2019-04-09 15:16 | 1554837367.7017 | -591 | 0.20951 | 0.000441 | 3e-06 | 91.174 | 90.4639 | 122.339 | 20.6959 | 0.001927 | 0.001768 | -154.6065 | 0.23151 | 8.4029 | 0.00268 | 0.004001 | 0.003906 | 3.2e-05 | 0.001321 | 0.28727 | 1.333 | 2.5477 | 98.49 | 20.07 | 47.662 | 67.6525 | 85.7476 | 63.8259 | 0.011376 | 1.0477 | 0.024282 | 23.177 | 0.069936 | 0.13772 | 8e-06 | 0.000423 | 0.000423 | 13.1137
2019-04-09 15:16 | 1554837368.7006 | -590 | 0.20954 | 0.000441 | 2e-06 | 91.1719 | 90.5544 | 122.2465 | 20.6959 | 0.001906 | 0.001759 | -154.6019 | 0.23144 | 8.4045 | 0.002727 | 0.004215 | 0.003934 | 3.3e-05 | 0.001321 | 0.28136 | 1.3602 | 2.5477 | 98.49 | 20.07 | 47.663 | 67.1322 | 86.8623 | 63.6089 | 0.011376 | 1.0484 | 0.024281 | 23.1607 | 0.060786 | 0.13769 | 8e-06 | 0.000469 | 0.000469 | 11.8194
2019-04-09 15:16 | 1554837369.7017 | -589 | 0.20955 | 0.000442 | 2e-06 | 91.1516 | 90.5926 | 122.052 | 20.6956 | 0.001871 | 0.001757 | -154.6098 | 0.23174 | 8.4045 | 0.002644 | 0.004044 | 0.003876 | 3.3e-05 | 0.001322 | 0.30381 | 1.3829 | 2.5476 | 98.49 | 20.07 | 47.667 | 66.3064 | 84.8726 | 63.5578 | 0.011377 | 1.0485 | 0.024263 | 23.1405 | 0.055027 | 0.13786 | 8e-06 | 0.000411 | 0.000411 | 13.534
2019-04-09 15:16 | 1554837370.7008 | -588 | 0.2095 | 0.000444 | 2e-06 | 91.1361 | 90.5874 | 121.9095 | 20.6956 | 0.001857 | 0.001782 | -154.6414 | 0.23229 | 8.4008 | 0.002685 | 0.004626 | 0.004465 | 3.2e-05 | 0.001317 | 0.28512 | 1.3373 | 2.5475 | 98.49 | 20.07 | 47.668 | 65.9615 | 85.8583 | 64.1544 | 0.011377 | 1.0467 | 0.024228 | 23.1475 | 0.054017 | 0.13819 | 6e-06 | 0.000293 | 0.000293 | 19.016
2019-04-09 15:16 | 1554837371.7008 | -587 | 0.20955 | 0.000443 | 3e-06 | 91.1186 | 90.6562 | 121.9428 | 20.6952 | 0.001931 | 0.00179 | -154.6157 | 0.23183 | 8.4005 | 0.002666 | 0.003791 | 0.003524 | 3.3e-05 | 0.001322 | 0.29927 | 1.3485 | 2.5475 | 98.49 | 20.07 | 47.668 | 67.7446 | 85.3972 | 64.3446 | 0.011377 | 1.0461 | 0.024224 | 23.1572 | 0.045499 | 0.13792 | 6e-06 | 0.000233 | 0.000233 | 23.904
2019-04-09 15:16 | 1554837372.7018 | -586 | 0.20953 | 0.000443 | 3e-06 | 91.1296 | 90.3685 | 121.97 | 20.6954 | 0.001905 | 0.001724 | -154.6036 | 0.23176 | 8.4019 | 0.002676 | 0.004297 | 0.004174 | 3.3e-05 | 0.001323 | 0.28495 | 1.3635 | 2.5476 | 98.49 | 20.07 | 47.669 | 67.1161 | 85.655 | 62.7548 | 0.011377 | 1.051 | 0.024284 | 23.1052 | 0.075012 | 0.13787 | 5e-06 | 0.000196 | 0.000196 | 28.3936
2019-04-09 15:16 | 1554837373.7021 | -585 | 0.20952 | 0.000441 | 2e-06 | 91.1448 | 90.4528 | 121.897 | 20.6957 | 0.0019 | 0.001774 | -154.6165 | 0.23164 | 8.3996 | 0.002681 | 0.004002 | 0.003811 | 3.1e-05 | 0.001331 | 0.27383 | 1.3464 | 2.5476 | 98.49 | 20.07 | 47.67 | 66.9846 | 85.7537 | 63.9538 | 0.011377 | 1.0473 | 0.024234 | 23.1394 | 0.068161 | 0.1378 | 7e-06 | 0.000389 | 0.000389 | 14.2882
2019-04-09 15:16 | 1554837374.7009 | -584 | 0.20953 | 0.000441 | 3e-06 | 91.0813 | 90.4494 | 121.9271 | 20.6962 | 0.001927 | 0.001786 | -154.6155 | 0.23307 | 8.402 | 0.002696 | 0.004026 | 0.003827 | 3.3e-05 | 0.001333 | 0.28512 | 1.3097 | 2.5475 | 98.49 | 20.07 | 47.671 | 67.648 | 86.1225 | 64.2542 | 0.011378 | 1.0464 | 0.024226 | 23.1526 | 0.062275 | 0.13865 | 8e-06 | 0.000434 | 0.000434 | 12.8686
2019-04-09 15:16 | 1554837375.7023 | -583 | 0.20953 | 0.000438 | 2e-06 | 91.0274 | 90.2591 | 121.6773 | 20.6957 | 0.001878 | 0.001802 | -154.611 | 0.23267 | 8.4002 | 0.002645 | 0.00445 | 0.004321 | 3.4e-05 | 0.001333 | 0.28263 | 1.3768 | 2.5474 | 98.49 | 20.06 | 47.669 | 66.4633 | 84.9085 | 64.6471 | 0.01137 | 1.0451 | 0.024187 | 23.1423 | 0.07582 | 0.13841 | 1.1e-05 | 0.00069 | 0.00069 | 8.0807
2019-04-09 15:16 | 1554837376.7013 | -582 | 0.20959 | 0.000441 | 2e-06 | 90.9286 | 90.3623 | 121.6457 | 20.6966 | 0.001886 | 0.001804 | -154.6123 | 0.23165 | 8.4026 | 0.002679 | 0.004318 | 0.004102 | 3.5e-05 | 0.001338 | 0.31564 | 1.3415 | 2.5475 | 98.49 | 20.06 | 47.668 | 66.6592 | 85.7289 | 64.6817 | 0.01137 | 1.045 | 0.024183 | 23.1405 | 0.055874 | 0.13781 | 8e-06 | 0.000461 | 0.000461 | 12.0589
2019-04-09 15:16 | 1554837377.7023 | -581 | 0.20954 | 0.000439 | 3e-06 | 90.8654 | 90.1145 | 121.6931 | 20.696 | 0.001875 | 0.001792 | -154.6183 | 0.23225 | 8.4021 | 0.002663 | 0.004666 | 0.004452 | 3.3e-05 | 0.001339 | 0.27632 | 1.3628 | 2.5475 | 98.49 | 20.06 | 47.666 | 66.3887 | 85.3388 | 64.3969 | 0.011369 | 1.0459 | 0.024198 | 23.1353 | 0.074223 | 0.13817 | 1.1e-05 | 0.000629 | 0.000629 | 8.8568
2019-04-09 15:16 | 1554837378.7024 | -580 | 0.20957 | 0.000445 | 2e-06 | 90.8175 | 90.0997 | 121.5992 | 20.6962 | 0.001911 | 0.001792 | -154.6271 | 0.2328 | 8.4039 | 0.002712 | 0.004436 | 0.004244 | 3.4e-05 | 0.00134 | 0.29302 | 1.3428 | 2.5475 | 98.49 | 20.07 | 47.666 | 67.2681 | 86.5176 | 64.4033 | 0.011376 | 1.0459 | 0.024188 | 23.1266 | 0.070972 | 0.1385 | 3e-06 | 4.8e-05 | 4.8e-05 | 115.1071
2019-04-09 15:16 | 1554837379.702 | -579 | 0.20953 | 0.00044 | 3e-06 | 90.7097 | 90.0027 | 121.6797 | 20.6948 | 0.001855 | 0.001738 | -154.6239 | 0.23204 | 8.4004 | 0.002654 | 0.004174 | 0.00404 | 3.1e-05 | 0.001343 | 0.27793 | 1.3644 | 2.5476 | 98.49 | 20.06 | 47.664 | 65.9068 | 85.1176 | 63.1043 | 0.011369 | 1.0499 | 0.024243 | 23.0897 | 0.069988 | 0.13804 | 9e-06 | 0.00051 | 0.00051 | 10.9132
2019-04-09 15:16 | 1554837380.7011 | -578 | 0.20958 | 0.000444 | 2e-06 | 90.6674 | 89.9738 | 121.9325 | 20.6958 | 0.001879 | 0.001802 | -154.5969 | 0.23246 | 8.4034 | 0.002705 | 0.003851 | 0.00374 | 3.1e-05 | 0.001344 | 0.28174 | 1.3353 | 2.5474 | 98.49 | 20.07 | 47.661 | 66.4822 | 86.3416 | 64.6287 | 0.011375 | 1.0452 | 0.024213 | 23.166 | 0.068691 | 0.13829 | 5e-06 | 0.000181 | 0.000181 | 30.8657
2019-04-09 15:16 | 1554837381.7011 | -577 | 0.20957 | 0.000444 | 3e-06 | 90.5697 | 89.7508 | 122.3292 | 20.6947 | 0.001891 | 0.001793 | -154.637 | 0.23229 | 8.403 | 0.002636 | 0.004681 | 0.004451 | 3.2e-05 | 0.001342 | 0.28191 | 1.3966 | 2.5476 | 98.49 | 20.06 | 47.659 | 66.7769 | 84.6843 | 64.4315 | 0.011368 | 1.0458 | 0.024259 | 23.1968 | 0.081247 | 0.13819 | 4e-06 | 0.000152 | 0.000152 | 36.6148
2019-04-09 15:16 | 1554837382.7003 | -576 | 0.20956 | 0.000442 | 3e-06 | 90.5632 | 89.6716 | 122.4369 | 20.6952 | 0.001904 | 0.001791 | -154.6254 | 0.23282 | 8.4022 | 0.002646 | 0.003917 | 0.003808 | 3.3e-05 | 0.001338 | 0.3011 | 1.357 | 2.5477 | 98.49 | 20.06 | 47.655 | 67.0916 | 84.9205 | 64.3624 | 0.011367 | 1.046 | 0.024273 | 23.2047 | 0.088492 | 0.13851 | 7e-06 | 0.000373 | 0.000373 | 14.9588
2019-04-09 15:16 | 1554837383.7012 | -575 | 0.20958 | 0.000444 | 2e-06 | 90.5647 | 89.791 | 122.5968 | 20.6961 | 0.001925 | 0.00178 | -154.6105 | 0.23194 | 8.3987 | 0.002697 | 0.003952 | 0.003679 | 3e-05 | 0.00133 | 0.29789 | 1.3318 | 2.5476 | 98.49 | 20.06 | 47.653 | 67.5989 | 86.1474 | 64.1166 | 0.011366 | 1.0468 | 0.024297 | 23.2114 | 0.076735 | 0.13798 | 5e-06 | 0.000217 | 0.000217 | 25.6111
2019-04-09 15:16 | 1554837384.7012 | -574 | 0.20955 | 0.000446 | 3e-06 | 90.5614 | 89.8788 | 122.5937 | 20.6952 | 0.001886 | 0.001763 | -154.6427 | 0.23316 | 8.4023 | 0.002678 | 0.004311 | 0.004092 | 2.9e-05 | 0.001334 | 0.27184 | 1.2991 | 2.5475 | 98.49 | 20.06 | 47.651 | 66.6619 | 85.6993 | 63.6902 | 0.011366 | 1.0481 | 0.024312 | 23.1964 | 0.067674 | 0.1387 | 2e-06 | -3.8e-05 | -3.8e-05 | -146.8992
2019-04-09 15:16 | 1554837385.7016 | -573 | 0.20952 | 0.000444 | 2e-06 | 90.503 | 89.8702 | 122.4611 | 20.6964 | 0.001882 | 0.001776 | -154.648 | 0.23223 | 8.4029 | 0.002708 | 0.004413 | 0.004244 | 2.9e-05 | 0.00134 | 0.28766 | 1.2932 | 2.5475 | 98.49 | 20.06 | 47.647 | 66.5671 | 86.4047 | 64.0007 | 0.011365 | 1.0471 | 0.024288 | 23.1945 | 0.062764 | 0.13816 | 4e-06 | 0.000183 | 0.000183 | 30.3885
2019-04-09 15:16 | 1554837386.7012 | -572 | 0.20954 | 0.00044 | 3e-06 | 90.5358 | 89.7914 | 122.5874 | 20.6952 | 0.00191 | 0.00183 | -154.6607 | 0.23201 | 8.4021 | 0.002779 | 0.004217 | 0.003954 | 2.9e-05 | 0.001333 | 0.29142 | 1.3313 | 2.5475 | 98.49 | 20.06 | 47.644 | 67.2446 | 88.1336 | 65.3032 | 0.011364 | 1.0431 | 0.024254 | 23.2513 | 0.073848 | 0.13802 | 9e-06 | 0.000484 | 0.000484 | 11.4988
2019-04-09 15:16 | 1554837387.7007 | -571 | 0.20954 | 0.000441 | 2e-06 | 90.4313 | 89.6733 | 122.8172 | 20.6951 | 0.001925 | 0.001755 | -154.6706 | 0.23217 | 8.4012 | 0.002746 | 0.003737 | 0.003619 | 3.1e-05 | 0.001328 | 0.28622 | 1.3586 | 2.5476 | 98.49 | 20.06 | 47.641 | 67.5948 | 87.3294 | 63.5159 | 0.011363 | 1.0487 | 0.024341 | 23.2115 | 0.075282 | 0.13812 | 8e-06 | 0.000485 | 0.000485 | 11.4762
2019-04-09 15:16 | 1554837388.7013 | -570 | 0.20954 | 0.000442 | 2e-06 | 90.4581 | 89.5412 | 122.8345 | 20.6967 | 0.001872 | 0.001728 | -154.6632 | 0.2327 | 8.4009 | 0.002774 | 0.003946 | 0.003828 | 3.3e-05 | 0.001327 | 0.28982 | 1.3777 | 2.5475 | 98.49 | 20.06 | 47.638 | 66.3286 | 88.0101 | 62.8458 | 0.011363 | 1.0507 | 0.024367 | 23.19 | 0.091124 | 0.13843 | 7e-06 | 0.000373 | 0.000373 | 14.9536
2019-04-09 15:16 | 1554837389.7003 | -569 | 0.20957 | 0.000441 | 2e-06 | 90.3803 | 89.5694 | 122.9608 | 20.6952 | 0.001919 | 0.001779 | -154.6755 | 0.23264 | 8.4009 | 0.002864 | 0.003487 | 0.003289 | 3.1e-05 | 0.001325 | 0.27256 | 1.3691 | 2.5476 | 98.49 | 20.06 | 47.638 | 67.4451 | 90.1789 | 64.0939 | 0.011363 | 1.0469 | 0.024334 | 23.245 | 0.080617 | 0.1384 | 8e-06 | 0.000418 | 0.000418 | 13.3356
2019-04-09 15:16 | 1554837390.7004 | -568 | 0.20955 | 0.000439 | 2e-06 | 90.375 | 89.6338 | 122.8294 | 20.6946 | 0.001889 | 0.001821 | -154.6859 | 0.23238 | 8.4022 | 0.002772 | 0.004249 | 0.00405 | 3.3e-05 | 0.001328 | 0.29286 | 1.3498 | 2.5475 | 98.49 | 20.06 | 47.636 | 66.7378 | 87.9472 | 65.095 | 0.011362 | 1.0438 | 0.024285 | 23.267 | 0.073662 | 0.13824 | 1.1e-05 | 0.000642 | 0.000642 | 8.6828
2019-04-09 15:17 | 1554837391.7004 | -567 | 0.20958 | 0.00044 | 2e-06 | 90.3848 | 89.5756 | 122.8355 | 20.6955 | 0.001912 | 0.001771 | -154.6583 | 0.23156 | 8.4048 | 0.002737 | 0.004047 | 0.003912 | 2.9e-05 | 0.001329 | 0.28495 | 1.3697 | 2.5476 | 98.49 | 20.06 | 47.634 | 67.2817 | 87.1103 | 63.8952 | 0.011362 | 1.0475 | 0.024329 | 23.2263 | 0.08044 | 0.13775 | 1e-05 | 0.000588 | 0.000588 | 9.4443
2019-04-09 15:17 | 1554837392.7014 | -566 | 0.20957 | 0.000443 | 2e-06 | 90.3625 | 89.7143 | 122.6214 | 20.6955 | 0.001838 | 0.001783 | -154.6753 | 0.23307 | 8.4018 | 0.002788 | 0.004174 | 0.003966 | 3e-05 | 0.001336 | 0.29695 | 1.3653 | 2.5475 | 98.49 | 20.06 | 47.634 | 65.5127 | 88.3503 | 64.1769 | 0.011362 | 1.0466 | 0.024298 | 23.2158 | 0.064391 | 0.13865 | 6e-06 | 0.000322 | 0.000322 | 17.3725
2019-04-09 15:17 | 1554837393.7008 | -565 | 0.20954 | 0.000447 | 2e-06 | 90.361 | 89.4693 | 122.5381 | 20.6953 | 0.001881 | 0.001839 | -154.6716 | 0.23311 | 8.3986 | 0.002761 | 0.004028 | 0.003902 | 3.2e-05 | 0.001337 | 0.30829 | 1.3395 | 2.5476 | 98.49 | 20.06 | 47.636 | 66.534 | 87.6809 | 65.5366 | 0.011362 | 1.0424 | 0.02424 | 23.2546 | 0.088708 | 0.13868 | 1e-06 | -3.3e-05 | -3.3e-05 | -171.213
2019-04-09 15:17 | 1554837394.7005 | -564 | 0.20954 | 0.000441 | 2e-06 | 90.2076 | 89.3378 | 122.7531 | 20.6948 | 0.001916 | 0.001771 | -154.6536 | 0.2326 | 8.3991 | 0.002781 | 0.004004 | 0.003711 | 3.1e-05 | 0.001336 | 0.29695 | 1.3673 | 2.5476 | 98.49 | 20.06 | 47.637 | 67.3769 | 88.1787 | 63.885 | 0.011362 | 1.0475 | 0.024321 | 23.2182 | 0.086667 | 0.13837 | 8e-06 | 0.000421 | 0.000421 | 13.2541
2019-04-09 15:17 | 1554837395.7019 | -563 | 0.20955 | 0.000442 | 2e-06 | 90.2076 | 89.3268 | 122.7498 | 20.6951 | 0.001874 | 0.001737 | -154.6419 | 0.23172 | 8.4008 | 0.002778 | 0.003759 | 0.003639 | 3e-05 | 0.001336 | 0.27793 | 1.3332 | 2.5476 | 98.49 | 20.06 | 47.636 | 66.3574 | 88.0914 | 63.067 | 0.011362 | 1.0501 | 0.02435 | 23.1897 | 0.087769 | 0.13785 | 8e-06 | 0.000428 | 0.000428 | 12.9807
2019-04-09 15:17 | 1554837396.701 | -562 | 0.20955 | 0.000439 | 3e-06 | 90.1219 | 89.2621 | 122.5576 | 20.6949 | 0.001938 | 0.001852 | -154.6213 | 0.23284 | 8.4014 | 0.002744 | 0.004271 | 0.00402 | 3.2e-05 | 0.001332 | 0.29031 | 1.3466 | 2.5476 | 98.49 | 20.06 | 47.637 | 67.8966 | 87.2761 | 65.8343 | 0.011362 | 1.0415 | 0.024232 | 23.2667 | 0.085742 | 0.13852 | 9e-06 | 0.000503 | 0.000503 | 11.1077
2019-04-09 15:17 | 1554837397.6996 | -561 | 0.20957 | 0.000443 | 2e-06 | 90.0379 | 89.2889 | 122.6038 | 20.6968 | 0.001924 | 0.001778 | -154.6118 | 0.23308 | 8.4008 | 0.002854 | 0.00393 | 0.003856 | 2.9e-05 | 0.001329 | 0.27975 | 1.347 | 2.5477 | 98.49 | 20.06 | 47.638 | 67.5613 | 89.9227 | 64.0682 | 0.011363 | 1.0469 | 0.0243 | 23.2104 | 0.074714 | 0.13866 | 6e-06 | 0.000288 | 0.000288 | 19.4018
2019-04-09 15:17 | 1554837398.701 | -560 | 0.20957 | 0.000446 | 2e-06 | 90.1555 | 89.4073 | 122.5407 | 20.6945 | 0.001889 | 0.001813 | -154.6229 | 0.23249 | 8.4018 | 0.002847 | 0.003834 | 0.003599 | 3.4e-05 | 0.001323 | 0.28208 | 1.352 | 2.5477 | 98.49 | 20.06 | 47.643 | 66.736 | 89.7593 | 64.9012 | 0.011364 | 1.0444 | 0.024264 | 23.233 | 0.07454 | 0.13831 | 2e-06 | 1.8e-05 | 1.8e-05 | 305.2276
2019-04-09 15:17 | 1554837399.7006 | -559 | 0.20955 | 0.000439 | 2e-06 | 90.1079 | 89.3048 | 122.6902 | 20.6959 | 0.001905 | 0.001756 | -154.606 | 0.23247 | 8.4018 | 0.002812 | 0.004253 | 0.004067 | 3.3e-05 | 0.00132 | 0.30021 | 1.3219 | 2.5476 | 98.49 | 20.06 | 47.645 | 67.1103 | 88.9217 | 63.5309 | 0.011364 | 1.0486 | 0.024328 | 23.2 | 0.080078 | 0.13829 | 1.1e-05 | 0.000705 | 0.000705 | 7.9071
2019-04-09 15:17 | 1554837400.6997 | -558 | 0.20957 | 0.000443 | 2e-06 | 90.1332 | 89.6293 | 122.6489 | 20.6952 | 0.001965 | 0.001769 | -154.6145 | 0.23182 | 8.4028 | 0.002769 | 0.004251 | 0.004024 | 3e-05 | 0.001323 | 0.27793 | 1.3124 | 2.5475 | 98.49 | 20.06 | 47.648 | 68.5527 | 87.883 | 63.8402 | 0.011365 | 1.0476 | 0.024312 | 23.2068 | 0.050144 | 0.13791 | 6e-06 | 0.000275 | 0.000275 | 20.1948
2019-04-09 15:17 | 1554837401.7008 | -557 | 0.20958 | 0.000444 | 2e-06 | 90.1898 | 89.4847 | 122.5951 | 20.6952 | 0.001917 | 0.001822 | -154.6156 | 0.23292 | 8.4024 | 0.002805 | 0.004101 | 0.003947 | 3.2e-05 | 0.001322 | 0.29214 | 1.3158 | 2.5476 | 98.49 | 20.06 | 47.65 | 67.3991 | 88.7523 | 65.1198 | 0.011366 | 1.0437 | 0.024261 | 23.2457 | 0.070194 | 0.13856 | 4e-06 | 0.000157 | 0.000157 | 35.5009
2019-04-09 15:17 | 1554837402.7008 | -556 | 0.20958 | 0.000442 | 2e-06 | 90.1466 | 89.5085 | 122.9336 | 20.6939 | 0.001948 | 0.001766 | -154.5925 | 0.23241 | 8.4035 | 0.002754 | 0.00392 | 0.003695 | 3.1e-05 | 0.001326 | 0.28224 | 1.2999 | 2.5476 | 98.49 | 20.06 | 47.652 | 68.1451 | 87.5303 | 63.7624 | 0.011366 | 1.0479 | 0.024343 | 23.231 | 0.063537 | 0.13826 | 7e-06 | 0.000367 | 0.000367 | 15.1737
2019-04-09 15:17 | 1554837403.7012 | -555 | 0.20954 | 0.000443 | 3e-06 | 90.1478 | 89.186 | 123.0599 | 20.6942 | 0.0019 | 0.001768 | -154.6125 | 0.23243 | 8.4024 | 0.002869 | 0.003877 | 0.003692 | 3.2e-05 | 0.001324 | 0.27649 | 1.3232 | 2.5476 | 98.49 | 20.06 | 47.658 | 67.0027 | 90.2991 | 63.8268 | 0.011367 | 1.0477 | 0.024354 | 23.2452 | 0.095938 | 0.13827 | 5e-06 | 0.000214 | 0.000214 | 26.0843
2019-04-09 15:17 | 1554837404.6998 | -554 | 0.20953 | 0.000441 | 2e-06 | 90.0757 | 89.3794 | 123 | 20.6942 | 0.001909 | 0.001764 | -154.6193 | 0.23165 | 8.4013 | 0.002784 | 0.004443 | 0.004185 | 3.6e-05 | 0.001322 | 0.2844 | 1.3296 | 2.5475 | 98.49 | 20.06 | 47.66 | 67.2032 | 88.235 | 63.7307 | 0.011368 | 1.048 | 0.024351 | 23.2362 | 0.069408 | 0.13781 | 9e-06 | 0.000481 | 0.000481 | 11.5391
2019-04-09 15:17 | 1554837405.7 | -553 | 0.20961 | 0.00044 | 3e-06 | 90.0757 | 89.2848 | 123.0573 | 20.695 | 0.001888 | 0.001803 | -154.6325 | 0.23317 | 8.4013 | 0.002756 | 0.003942 | 0.003841 | 2.9e-05 | 0.001323 | 0.28335 | 1.3441 | 2.5475 | 98.49 | 20.05 | 47.663 | 66.7014 | 87.5666 | 64.6722 | 0.011362 | 1.0451 | 0.024323 | 23.2741 | 0.078884 | 0.13871 | 9e-06 | 0.000457 | 0.000457 | 12.2301
2019-04-09 15:17 | 1554837406.7013 | -552 | 0.20958 | 0.000443 | 2e-06 | 90.1564 | 89.3639 | 123.2324 | 20.6945 | 0.0019 | 0.001771 | -154.6524 | 0.23212 | 8.4022 | 0.002738 | 0.003891 | 0.003659 | 3.2e-05 | 0.001325 | 0.29734 | 1.3409 | 2.5472 | 98.49 | 20.05 | 47.671 | 66.995 | 87.1281 | 63.9012 | 0.011364 | 1.0475 | 0.024368 | 23.264 | 0.078968 | 0.13809 | 7e-06 | 0.000365 | 0.000365 | 15.2424
2019-04-09 15:17 | 1554837407.6999 | -551 | 0.20956 | 0.000444 | 3e-06 | 90.0751 | 89.1285 | 122.9869 | 20.6958 | 0.001873 | 0.001767 | -154.6654 | 0.23206 | 8.402 | 0.002737 | 0.003959 | 0.003811 | 3.6e-05 | 0.001323 | 0.27184 | 1.3431 | 2.5474 | 98.49 | 20.05 | 47.676 | 66.3559 | 87.1231 | 63.7987 | 0.011365 | 1.0478 | 0.024347 | 23.2373 | 0.094492 | 0.13805 | 4e-06 | 0.000123 | 0.000123 | 45.1287
2019-04-09 15:17 | 1554837408.7 | -550 | 0.20961 | 0.000442 | 2e-06 | 90.0266 | 89.2545 | 123.3098 | 20.6951 | 0.001898 | 0.001759 | -154.6581 | 0.23201 | 8.4012 | 0.002845 | 0.003889 | 0.003599 | 3.3e-05 | 0.001318 | 0.26936 | 1.331 | 2.5474 | 98.49 | 20.05 | 47.683 | 66.9559 | 89.7163 | 63.5972 | 0.011366 | 1.0484 | 0.024387 | 23.2608 | 0.077038 | 0.13802 | 7e-06 | 0.000353 | 0.000353 | 15.764
2019-04-09 15:17 | 1554837409.7009 | -549 | 0.20953 | 0.00044 | 3e-06 | 90.1034 | 89.4655 | 123.276 | 20.6954 | 0.001888 | 0.001772 | -154.6577 | 0.23299 | 8.4013 | 0.002736 | 0.004186 | 0.004035 | 3.3e-05 | 0.001318 | 0.28943 | 1.3589 | 2.5476 | 98.49 | 20.05 | 47.686 | 66.7112 | 87.089 | 63.9059 | 0.011367 | 1.0474 | 0.024372 | 23.2683 | 0.063552 | 0.13861 | 9e-06 | 0.000446 | 0.000446 | 12.5106
2019-04-09 15:17 | 1554837410.7 | -548 | 0.20959 | 0.000443 | 3e-06 | 90.0918 | 89.2542 | 123.2812 | 20.695 | 0.001888 | 0.001705 | -154.6373 | 0.23232 | 8.4033 | 0.002721 | 0.00449 | 0.004242 | 3.3e-05 | 0.001319 | 0.30287 | 1.3635 | 2.5476 | 98.49 | 20.05 | 47.69 | 66.7038 | 86.7296 | 62.3044 | 0.011368 | 1.0524 | 0.024431 | 23.2135 | 0.08355 | 0.1382 | 5e-06 | 0.000213 | 0.000213 | 26.183
2019-04-09 15:17 | 1554837411.7001 | -547 | 0.20955 | 0.000438 | 2e-06 | 90.1945 | 89.5157 | 123.3976 | 20.6949 | 0.001896 | 0.001814 | -154.6514 | 0.23265 | 8.4014 | 0.002816 | 0.004364 | 0.004198 | 2.9e-05 | 0.001322 | 0.28368 | 1.3689 | 2.5476 | 98.49 | 20.05 | 47.698 | 66.8893 | 89.0207 | 64.936 | 0.01137 | 1.0442 | 0.024347 | 23.3153 | 0.067571 | 0.13841 | 1.2e-05 | 0.00074 | 0.00074 | 7.5352
2019-04-09 15:17 | 1554837412.6998 | -546 | 0.20955 | 0.000444 | 2e-06 | 90.2052 | 89.3674 | 123.0026 | 20.6948 | 0.001915 | 0.001758 | -154.6683 | 0.23312 | 8.3994 | 0.00281 | 0.003752 | 0.003579 | 3.1e-05 | 0.001323 | 0.29391 | 1.3524 | 2.5477 | 98.49 | 20.05 | 47.702 | 67.3557 | 88.8605 | 63.575 | 0.011371 | 1.0485 | 0.024357 | 23.2311 | 0.083467 | 0.13869 | 4e-06 | 0.000151 | 0.000151 | 36.9312
2019-04-09 15:17 | 1554837413.6995 | -545 | 0.20953 | 0.000442 | 3e-06 | 90.3389 | 89.5825 | 123.1587 | 20.6947 | 0.001851 | 0.001734 | -154.6708 | 0.2317 | 8.4 | 0.002777 | 0.004673 | 0.004487 | 3.2e-05 | 0.001322 | 0.28816 | 1.3096 | 2.5476 | 98.49 | 20.05 | 47.709 | 65.8093 | 88.0684 | 63.0026 | 0.011373 | 1.0503 | 0.024393 | 23.226 | 0.075213 | 0.13784 | 6e-06 | 0.00031 | 0.00031 | 17.9154
2019-04-09 15:17 | 1554837414.6991 | -544 | 0.20947 | 0.000445 | 2e-06 | 90.3488 | 89.5477 | 123.0244 | 20.6952 | 0.001881 | 0.001806 | -154.6659 | 0.23161 | 8.4022 | 0.002793 | 0.003853 | 0.003691 | 2.8e-05 | 0.001323 | 0.28528 | 1.3084 | 2.5475 | 98.49 | 20.04 | 47.712 | 66.5308 | 88.4579 | 64.7419 | 0.011366 | 1.0448 | 0.024317 | 23.2734 | 0.079658 | 0.13778 | 3e-06 | 0.000116 | 0.000116 | 48.0072
2019-04-09 15:17 | 1554837415.7003 | -543 | 0.2095 | 0.000441 | 2e-06 | 90.3708 | 89.6744 | 123.1507 | 20.6946 | 0.001885 | 0.001765 | -154.6295 | 0.23177 | 8.4044 | 0.002751 | 0.004241 | 0.004113 | 3.5e-05 | 0.001328 | 0.29856 | 1.3241 | 2.5475 | 98.49 | 20.04 | 47.716 | 66.6332 | 87.4577 | 63.7538 | 0.011367 | 1.0479 | 0.024365 | 23.2512 | 0.069198 | 0.13788 | 8e-06 | 0.000457 | 0.000457 | 12.1507
2019-04-09 15:17 | 1554837416.7002 | -542 | 0.20954 | 0.000438 | 3e-06 | 90.3669 | 89.5216 | 123.2284 | 20.6946 | 0.001933 | 0.001814 | -154.6247 | 0.23209 | 8.4029 | 0.002783 | 0.003672 | 0.003373 | 3.1e-05 | 0.001329 | 0.28838 | 1.3358 | 2.5476 | 98.49 | 20.04 | 47.719 | 67.7909 | 88.2158 | 64.935 | 0.011368 | 1.0443 | 0.02433 | 23.2993 | 0.084069 | 0.13807 | 1.1e-05 | 0.00065 | 0.00065 | 8.5605
2019-04-09 15:17 | 1554837417.7002 | -541 | 0.20949 | 0.00044 | 2e-06 | 90.2666 | 89.4207 | 123.1885 | 20.6954 | 0.001878 | 0.001762 | -154.6122 | 0.23187 | 8.403 | 0.002765 | 0.004654 | 0.004447 | 3.3e-05 | 0.001332 | 0.2855 | 1.3248 | 2.5476 | 98.49 | 20.04 | 47.72 | 66.4576 | 87.7872 | 63.6803 | 0.011368 | 1.0481 | 0.024372 | 23.2523 | 0.084211 | 0.13794 | 1e-05 | 0.000583 | 0.000583 | 9.5377
2019-04-09 15:17 | 1554837418.7002 | -540 | 0.20953 | 0.000438 | 3e-06 | 90.2166 | 89.3461 | 122.9951 | 20.6948 | 0.001837 | 0.001741 | -154.6189 | 0.23236 | 8.4009 | 0.002693 | 0.004195 | 0.004033 | 3.7e-05 | 0.001329 | 0.30486 | 1.3562 | 2.5476 | 98.49 | 20.04 | 47.724 | 65.469 | 86.0464 | 63.1722 | 0.011369 | 1.0497 | 0.024371 | 23.2165 | 0.086725 | 0.13823 | 1.1e-05 | 0.00064 | 0.00064 | 8.706
2019-04-09 15:17 | 1554837419.6992 | -539 | 0.20951 | 0.000439 | 2e-06 | 90.1603 | 89.2631 | 122.8106 | 20.6941 | 0.001897 | 0.001754 | -154.6357 | 0.2314 | 8.4005 | 0.002756 | 0.004152 | 0.003959 | 3.4e-05 | 0.001332 | 0.29142 | 1.3378 | 2.5477 | 98.49 | 20.04 | 47.725 | 66.9182 | 87.5672 | 63.4767 | 0.011369 | 1.0488 | 0.024342 | 23.2095 | 0.089451 | 0.13766 | 1.1e-05 | 0.000668 | 0.000668 | 8.3089
2019-04-09 15:17 | 1554837420.7007 | -538 | 0.20952 | 0.000439 | 2e-06 | 90.1317 | 89.451 | 122.6254 | 20.6955 | 0.001849 | 0.001763 | -154.6433 | 0.23241 | 8.4019 | 0.002708 | 0.004369 | 0.004174 | 3.3e-05 | 0.001334 | 0.30215 | 1.3492 | 2.5477 | 98.49 | 20.04 | 47.727 | 65.7694 | 86.4028 | 63.7062 | 0.01137 | 1.0481 | 0.024315 | 23.1999 | 0.067805 | 0.13826 | 1.2e-05 | 0.000711 | 0.000711 | 7.8403
2019-04-09 15:17 | 1554837421.7004 | -537 | 0.20957 | 0.000441 | 2e-06 | 90.0522 | 89.155 | 122.4958 | 20.6957 | 0.001942 | 0.00177 | -154.6412 | 0.23377 | 8.4005 | 0.002731 | 0.003799 | 0.003638 | 3.2e-05 | 0.001337 | 0.29247 | 1.3311 | 2.5476 | 98.49 | 20.04 | 47.73 | 67.9955 | 86.9568 | 63.8602 | 0.011371 | 1.0476 | 0.024297 | 23.193 | 0.089555 | 0.13907 | 8e-06 | 0.000437 | 0.000437 | 12.8159
2019-04-09 15:17 | 1554837422.6998 | -536 | 0.20958 | 0.000437 | 2e-06 | 90.0385 | 89.144 | 122.4022 | 20.6963 | 0.001903 | 0.001775 | -154.6347 | 0.2325 | 8.3999 | 0.002703 | 0.004619 | 0.004424 | 3.3e-05 | 0.001339 | 0.29175 | 1.35 | 2.5476 | 98.49 | 20.04 | 47.732 | 67.0732 | 86.2857 | 63.9801 | 0.011371 | 1.0472 | 0.024283 | 23.1882 | 0.089299 | 0.13832 | 1.3e-05 | 0.000772 | 0.000772 | 7.2169
2019-04-09 15:17 | 1554837423.7008 | -535 | 0.20951 | 0.000444 | 3e-06 | 89.9071 | 89.1929 | 122.2808 | 20.6953 | 0.001889 | 0.001793 | -154.6381 | 0.23232 | 8.3997 | 0.002775 | 0.004143 | 0.003925 | 3.3e-05 | 0.001338 | 0.28102 | 1.3796 | 2.5475 | 98.49 | 20.04 | 47.733 | 66.7399 | 88.0269 | 64.4119 | 0.011371 | 1.0459 | 0.024255 | 23.1916 | 0.071342 | 0.1382 | 4e-06 | 0.00012 | 0.00012 | 46.2867
2019-04-09 15:17 | 1554837424.6995 | -534 | 0.20955 | 0.000434 | 2e-06 | 89.9018 | 88.8247 | 122.1355 | 20.6956 | 0.001918 | 0.001803 | -154.6388 | 0.23281 | 8.4028 | 0.002749 | 0.004176 | 0.004046 | 3.4e-05 | 0.001346 | 0.28943 | 1.3685 | 2.5476 | 98.49 | 20.04 | 47.738 | 67.4212 | 87.4001 | 64.6532 | 0.011372 | 1.0451 | 0.024232 | 23.1861 | 0.10781 | 0.1385 | 1.7e-05 | 0.00106 | 0.00106 | 5.2635
2019-04-09 15:17 | 1554837425.7009 | -533 | 0.20949 | 0.000444 | 2e-06 | 89.9581 | 89.1038 | 122.3719 | 20.6964 | 0.001892 | 0.001819 | -154.6649 | 0.23251 | 8.4011 | 0.002717 | 0.004131 | 0.003955 | 3.3e-05 | 0.001338 | 0.29623 | 1.3435 | 2.5477 | 98.49 | 20.04 | 47.74 | 66.8112 | 86.6364 | 65.056 | 0.011373 | 1.0439 | 0.024241 | 23.2223 | 0.08535 | 0.13832 | 5e-06 | 0.000279 | 0.000279 | 19.9968
2019-04-09 15:17 | 1554837426.6995 | -532 | 0.20949 | 0.000437 | 2e-06 | 89.8994 | 89.0404 | 122.3879 | 20.6951 | 0.001855 | 0.001749 | -154.6645 | 0.23276 | 8.4019 | 0.002724 | 0.004217 | 0.004024 | 2.9e-05 | 0.001345 | 0.27865 | 1.3578 | 2.5477 | 98.49 | 20.04 | 47.742 | 65.9196 | 86.8031 | 63.3568 | 0.011373 | 1.0491 | 0.024304 | 23.1654 | 0.085874 | 0.13847 | 1.3e-05 | 0.000786 | 0.000786 | 7.1
2019-04-09 15:17 | 1554837427.7001 | -531 | 0.20955 | 0.000437 | 2e-06 | 89.9015 | 89.0731 | 122.4167 | 20.6956 | 0.001898 | 0.001794 | -154.662 | 0.23269 | 8.4007 | 0.002699 | 0.003955 | 0.003713 | 3.1e-05 | 0.001348 | 0.29391 | 1.3537 | 2.5477 | 98.49 | 20.04 | 47.744 | 66.9453 | 86.1922 | 64.4403 | 0.011374 | 1.0458 | 0.024268 | 23.2054 | 0.082797 | 0.13843 | 1.4e-05 | 0.000841 | 0.000841 | 6.6304
2019-04-09 15:17 | 1554837428.6996 | -530 | 0.20951 | 0.000442 | 2e-06 | 89.897 | 89.2352 | 122.5944 | 20.6955 | 0.001913 | 0.00175 | -154.6592 | 0.23274 | 8.4036 | 0.002806 | 0.003689 | 0.003487 | 2.8e-05 | 0.001349 | 0.28102 | 1.3598 | 2.5475 | 98.49 | 20.04 | 47.745 | 67.3112 | 88.7634 | 63.3913 | 0.011374 | 1.049 | 0.024323 | 23.1862 | 0.06609 | 0.13846 | 8e-06 | 0.000446 | 0.000446 | 12.5117
2019-04-09 15:17 | 1554837429.6996 | -529 | 0.20955 | 0.00044 | 2e-06 | 89.8711 | 89.1475 | 122.5195 | 20.6953 | 0.001894 | 0.001799 | -154.6399 | 0.23277 | 8.4011 | 0.002705 | 0.003585 | 0.00341 | 3e-05 | 0.00134 | 0.25227 | 1.3709 | 2.5475 | 98.49 | 20.04 | 47.747 | 66.8414 | 86.3505 | 64.5702 | 0.011375 | 1.0454 | 0.024273 | 23.2197 | 0.072313 | 0.13848 | 1e-05 | 0.000592 | 0.000592 | 9.4255
2019-04-09 15:17 | 1554837430.6993 | -528 | 0.20952 | 0.000442 | 2e-06 | 89.8235 | 88.9534 | 122.6712 | 20.6963 | 0.001908 | 0.001799 | -154.6417 | 0.23276 | 8.4009 | 0.002727 | 0.004453 | 0.00427 | 3.2e-05 | 0.001335 | 0.28655 | 1.3139 | 2.5475 | 98.49 | 20.04 | 47.749 | 67.1856 | 86.8746 | 64.5604 | 0.011375 | 1.0454 | 0.024289 | 23.2337 | 0.087066 | 0.13847 | 7e-06 | 0.000352 | 0.000352 | 15.8669
2019-04-09 15:17 | 1554837431.6986 | -527 | 0.20952 | 0.000446 | 2e-06 | 89.9131 | 89.1361 | 122.8242 | 20.6954 | 0.001902 | 0.00179 | -154.6497 | 0.23226 | 8.4014 | 0.002728 | 0.004127 | 0.003924 | 3.3e-05 | 0.001333 | 0.28351 | 1.3725 | 2.5475 | 98.49 | 20.03 | 47.752 | 67.0358 | 86.9009 | 64.3378 | 0.011369 | 1.0461 | 0.024312 | 23.2405 | 0.077631 | 0.13817 | 2e-06 | -1.5e-05 | -1.5e-05 | -375.0889
2019-04-09 15:17 | 1554837432.6991 | -526 | 0.2095 | 0.000441 | 2e-06 | 89.8512 | 89.0838 | 122.7608 | 20.6948 | 0.001908 | 0.001783 | -154.6421 | 0.23332 | 8.4006 | 0.00266 | 0.003938 | 0.003809 | 2.9e-05 | 0.001332 | 0.30182 | 1.4285 | 2.5477 | 98.49 | 20.03 | 47.753 | 67.1982 | 85.2535 | 64.17 | 0.011369 | 1.0466 | 0.024312 | 23.2287 | 0.076718 | 0.13881 | 8e-06 | 0.000445 | 0.000445 | 12.5819
2019-04-09 15:17 | 1554837433.6991 | -525 | 0.20958 | 0.000442 | 2e-06 | 89.8515 | 89.0425 | 122.7538 | 20.6957 | 0.001969 | 0.001782 | -154.6346 | 0.23252 | 8.4015 | 0.00272 | 0.004795 | 0.004609 | 3.1e-05 | 0.001324 | 0.28943 | 1.4242 | 2.5477 | 98.49 | 20.03 | 47.755 | 68.647 | 86.7029 | 64.1452 | 0.011369 | 1.0467 | 0.024312 | 23.2272 | 0.080895 | 0.13833 | 6e-06 | 0.000331 | 0.000331 | 16.8338
2019-04-09 15:17 | 1554837434.6998 | -524 | 0.20955 | 0.000443 | 3e-06 | 89.8509 | 89.0735 | 122.723 | 20.6955 | 0.001893 | 0.001775 | -154.6267 | 0.23261 | 8.4006 | 0.002687 | 0.004101 | 0.00394 | 3.2e-05 | 0.00133 | 0.28816 | 1.3828 | 2.5476 | 98.49 | 20.04 | 47.757 | 66.8221 | 85.9198 | 63.9981 | 0.011377 | 1.0472 | 0.024314 | 23.2192 | 0.077725 | 0.13838 | 5e-06 | 0.000211 | 0.000211 | 26.4058
2019-04-09 15:17 | 1554837435.7002 | -523 | 0.20955 | 0.000441 | 2e-06 | 89.8705 | 89.049 | 122.3426 | 20.6948 | 0.00189 | 0.001818 | -154.6183 | 0.23251 | 8.4023 | 0.002699 | 0.003586 | 0.003404 | 3e-05 | 0.001328 | 0.28296 | 1.394 | 2.5478 | 98.49 | 20.04 | 47.758 | 66.7558 | 86.1938 | 65.0211 | 0.011377 | 1.044 | 0.02424 | 23.2184 | 0.082135 | 0.13832 | 8e-06 | 0.00048 | 0.00048 | 11.6179
2019-04-09 15:17 | 1554837436.6999 | -522 | 0.2095 | 0.000441 | 2e-06 | 89.9122 | 88.9971 | 122.54 | 20.6948 | 0.001861 | 0.001797 | -154.6119 | 0.23278 | 8.4024 | 0.002682 | 0.004157 | 0.003925 | 3.6e-05 | 0.001331 | 0.32211 | 1.3548 | 2.5477 | 98.49 | 20.04 | 47.763 | 66.0583 | 85.7887 | 64.5081 | 0.011378 | 1.0456 | 0.024278 | 23.2195 | 0.091505 | 0.13848 | 8e-06 | 0.000489 | 0.000489 | 11.4018
2019-04-09 15:17 | 1554837437.6999 | -521 | 0.2095 | 0.000438 | 2e-06 | 89.9226 | 89.1137 | 122.5238 | 20.696 | 0.001868 | 0.001773 | -154.5905 | 0.23226 | 8.4028 | 0.002665 | 0.004303 | 0.004078 | 3.5e-05 | 0.001337 | 0.28174 | 1.3647 | 2.5475 | 98.49 | 20.04 | 47.766 | 66.2163 | 85.3809 | 63.9315 | 0.011379 | 1.0474 | 0.024297 | 23.1981 | 0.080824 | 0.13817 | 1.2e-05 | 0.000734 | 0.000734 | 7.5853
2019-04-09 15:17 | 1554837438.6999 | -520 | 0.20951 | 0.000439 | 3e-06 | 89.83 | 89.0621 | 122.4855 | 20.6948 | 0.001838 | 0.001768 | -154.5971 | 0.23254 | 8.4012 | 0.002697 | 0.003867 | 0.003729 | 3.3e-05 | 0.001337 | 0.28584 | 1.3457 | 2.5475 | 98.49 | 20.04 | 47.77 | 65.4955 | 86.1407 | 63.8242 | 0.01138 | 1.0477 | 0.024297 | 23.1908 | 0.07679 | 0.13834 | 1e-05 | 0.000578 | 0.000578 | 9.651
2019-04-09 15:17 | 1554837439.699 | -519 | 0.20952 | 0.000442 | 2e-06 | 89.8935 | 88.9644 | 122.3646 | 20.6952 | 0.001881 | 0.001823 | -154.5902 | 0.23332 | 8.4018 | 0.002623 | 0.004543 | 0.004378 | 3.3e-05 | 0.001337 | 0.31078 | 1.3658 | 2.5475 | 98.49 | 20.04 | 47.777 | 66.5479 | 84.3712 | 65.1447 | 0.011382 | 1.0436 | 0.024237 | 23.2247 | 0.092925 | 0.1388 | 7e-06 | 0.000417 | 0.000417 | 13.4006
2019-04-09 15:17 | 1554837440.6993 | -518 | 0.20953 | 0.000446 | 2e-06 | 89.9039 | 89.1571 | 122.5348 | 20.6956 | 0.001863 | 0.001817 | -154.5929 | 0.23282 | 8.4021 | 0.002702 | 0.004056 | 0.003969 | 3.6e-05 | 0.00134 | 0.28313 | 1.3241 | 2.5476 | 98.49 | 20.03 | 47.78 | 66.0981 | 86.261 | 64.9889 | 0.011375 | 1.0441 | 0.02426 | 23.2355 | 0.074606 | 0.13851 | 2e-06 | 2.7e-05 | 2.7e-05 | 206.2266
2019-04-09 15:17 | 1554837441.6991 | -517 | 0.20952 | 0.000442 | 2e-06 | 89.8726 | 88.9451 | 122.6158 | 20.6961 | 0.001886 | 0.001761 | -154.5772 | 0.23319 | 8.4035 | 0.002694 | 0.004356 | 0.004036 | 3.5e-05 | 0.001339 | 0.27599 | 1.3116 | 2.5475 | 98.49 | 20.03 | 47.785 | 66.6524 | 86.0726 | 63.6574 | 0.011377 | 1.0482 | 0.024316 | 23.1974 | 0.092788 | 0.13872 | 7e-06 | 0.000417 | 0.000417 | 13.4094
2019-04-09 15:17 | 1554837442.6994 | -516 | 0.20959 | 0.000438 | 2e-06 | 89.8116 | 89.0597 | 122.8988 | 20.6954 | 0.001895 | 0.001775 | -154.5724 | 0.23299 | 8.4012 | 0.002653 | 0.004185 | 0.004077 | 3.5e-05 | 0.001338 | 0.2907 | 1.3155 | 2.5476 | 98.49 | 20.04 | 47.788 | 66.8672 | 85.0932 | 63.9834 | 0.011384 | 1.0472 | 0.024332 | 23.2354 | 0.075193 | 0.13861 | 1.2e-05 | 0.00071 | 0.00071 | 7.8689
2019-04-09 15:17 | 1554837443.6995 | -515 | 0.20955 | 0.000441 | 2e-06 | 89.7812 | 89.0219 | 122.8254 | 20.6948 | 0.001935 | 0.001809 | -154.5725 | 0.23314 | 8.4021 | 0.002692 | 0.004195 | 0.004076 | 3.3e-05 | 0.001344 | 0.29679 | 1.3164 | 2.5476 | 98.49 | 20.03 | 47.794 | 67.8239 | 86.022 | 64.7954 | 0.011379 | 1.0447 | 0.024296 | 23.2564 | 0.07597 | 0.13869 | 1e-05 | 0.000588 | 0.000588 | 9.5124
2019-04-09 15:17 | 1554837444.6981 | -514 | 0.20953 | 0.000443 | 2e-06 | 89.7922 | 89.0301 | 123.0085 | 20.6951 | 0.001917 | 0.001719 | -154.5807 | 0.23293 | 8.4021 | 0.00267 | 0.003684 | 0.00346 | 3.5e-05 | 0.001334 | 0.25857 | 1.3608 | 2.5475 | 98.49 | 20.04 | 47.798 | 67.4037 | 85.5025 | 62.6445 | 0.011387 | 1.0514 | 0.024391 | 23.1995 | 0.076238 | 0.13857 | 6e-06 | 0.000325 | 0.000325 | 17.1842
2019-04-09 15:17 | 1554837445.6982 | -513 | 0.20955 | 0.000437 | 2e-06 | 89.8607 | 88.984 | 122.9824 | 20.695 | 0.001819 | 0.00179 | -154.5887 | 0.23305 | 8.4021 | 0.002626 | 0.003887 | 0.003673 | 3.3e-05 | 0.00133 | 0.30071 | 1.3252 | 2.5476 | 98.49 | 20.04 | 47.801 | 65.0502 | 84.4451 | 64.3406 | 0.011387 | 1.0461 | 0.024327 | 23.2556 | 0.087693 | 0.13864 | 1.4e-05 | 0.000885 | 0.000885 | 6.3108
2019-04-09 15:17 | 1554837446.6984 | -512 | 0.20958 | 0.00044 | 2e-06 | 89.8803 | 88.9967 | 122.4547 | 20.6944 | 0.001931 | 0.001811 | -154.6009 | 0.23256 | 8.4023 | 0.002615 | 0.003277 | 0.003129 | 3.2e-05 | 0.001326 | 0.27328 | 1.3436 | 2.5475 | 98.49 | 20.03 | 47.807 | 67.729 | 84.1706 | 64.8527 | 0.011382 | 1.0445 | 0.024257 | 23.2232 | 0.088369 | 0.13835 | 1.1e-05 | 0.000658 | 0.000658 | 8.4779
2019-04-09 15:17 | 1554837447.6992 | -511 | 0.20957 | 0.000442 | 2e-06 | 89.8836 | 88.9964 | 122.504 | 20.6945 | 0.001873 | 0.001709 | -154.6201 | 0.23289 | 8.401 | 0.002615 | 0.003899 | 0.003681 | 3.1e-05 | 0.001322 | 0.26001 | 1.3555 | 2.5478 | 98.49 | 20.03 | 47.81 | 66.3474 | 84.17 | 62.3937 | 0.011383 | 1.0522 | 0.02435 | 23.1432 | 0.088729 | 0.13855 | 8e-06 | 0.000427 | 0.000427 | 13.0643
2019-04-09 15:17 | 1554837448.6984 | -510 | 0.20951 | 0.00044 | 2e-06 | 89.8139 | 89.0546 | 122.5388 | 20.6956 | 0.001851 | 0.001789 | -154.6056 | 0.23313 | 8.4013 | 0.002595 | 0.004001 | 0.003718 | 3.2e-05 | 0.001317 | 0.29607 | 1.3768 | 2.5476 | 98.49 | 20.03 | 47.815 | 65.8195 | 83.685 | 64.32 | 0.011384 | 1.0462 | 0.024284 | 23.2129 | 0.075949 | 0.13869 | 9e-06 | 0.000502 | 0.000502 | 11.1334
2019-04-09 15:17 | 1554837449.6993 | -509 | 0.20957 | 0.000444 | 2e-06 | 89.8976 | 89.0332 | 122.8014 | 20.6945 | 0.001942 | 0.001776 | -154.6102 | 0.23299 | 8.3986 | 0.002686 | 0.003781 | 0.003703 | 3.5e-05 | 0.001324 | 0.29391 | 1.3879 | 2.5476 | 98.49 | 20.03 | 47.818 | 68.0147 | 85.8812 | 64.0121 | 0.011384 | 1.0471 | 0.024321 | 23.2271 | 0.086422 | 0.13861 | 4e-06 | 0.000161 | 0.000161 | 34.6177
2019-04-09 15:18 | 1554837450.6997 | -508 | 0.20956 | 0.000441 | 2e-06 | 89.8655 | 89.0043 | 122.7892 | 20.6948 | 0.00188 | 0.001749 | -154.6162 | 0.23283 | 8.3979 | 0.002657 | 0.003444 | 0.003214 | 3.3e-05 | 0.001315 | 0.27312 | 1.3475 | 2.5476 | 98.49 | 20.03 | 47.821 | 66.5193 | 85.1854 | 63.3705 | 0.011385 | 1.0491 | 0.024343 | 23.2039 | 0.086126 | 0.13851 | 8e-06 | 0.000488 | 0.000488 | 11.4444
2019-04-09 15:18 | 1554837451.6997 | -507 | 0.20956 | 0.000437 | 2e-06 | 89.878 | 89.0487 | 122.6022 | 20.6952 | 0.001885 | 0.001769 | -154.6239 | 0.23234 | 8.4022 | 0.002667 | 0.003821 | 0.003685 | 3.3e-05 | 0.001316 | 0.28727 | 1.3805 | 2.5477 | 98.49 | 20.03 | 47.826 | 66.6401 | 85.4341 | 63.8403 | 0.011386 | 1.0476 | 0.024308 | 23.2024 | 0.082911 | 0.13822 | 1.3e-05 | 0.000827 | 0.000827 | 6.7385
2019-04-09 15:18 | 1554837452.6988 | -506 | 0.20955 | 0.000439 | 2e-06 | 89.8777 | 89.0243 | 122.5909 | 20.6952 | 0.001942 | 0.001744 | -154.634 | 0.23273 | 8.4033 | 0.002643 | 0.004064 | 0.003899 | 3.3e-05 | 0.00131 | 0.28584 | 1.369 | 2.5475 | 98.49 | 20.03 | 47.828 | 68.0116 | 84.8621 | 63.239 | 0.011387 | 1.0495 | 0.024328 | 23.1806 | 0.085336 | 0.13845 | 1.2e-05 | 0.000746 | 0.000746 | 7.4811
2019-04-09 15:18 | 1554837453.6994 | -505 | 0.20954 | 0.000438 | 1e-06 | 89.858 | 89.2311 | 122.6597 | 20.6953 | 0.001889 | 0.001771 | -154.6516 | 0.23271 | 8.4018 | 0.002642 | 0.003935 | 0.003743 | 3e-05 | 0.00132 | 0.27223 | 1.4037 | 2.5477 | 98.49 | 20.03 | 47.828 | 66.7298 | 84.8246 | 63.8988 | 0.011387 | 1.0475 | 0.024311 | 23.2098 | 0.062622 | 0.13844 | 1.3e-05 | 0.000835 | 0.000835 | 6.6803
2019-04-09 15:18 | 1554837454.6994 | -504 | 0.20953 | 0.00044 | 2e-06 | 89.9143 | 88.9258 | 122.5148 | 20.6953 | 0.001884 | 0.00183 | -154.6508 | 0.23225 | 8.4005 | 0.002661 | 0.003979 | 0.00377 | 3.1e-05 | 0.001323 | 0.27151 | 1.3576 | 2.5474 | 98.49 | 20.03 | 47.828 | 66.6207 | 85.2944 | 65.3088 | 0.011387 | 1.0431 | 0.024246 | 23.2446 | 0.098875 | 0.13817 | 1.1e-05 | 0.000657 | 0.000657 | 8.4693
2019-04-09 15:18 | 1554837455.6999 | -503 | 0.20955 | 0.000434 | 1e-06 | 89.8658 | 89.1168 | 122.4188 | 20.694 | 0.001932 | 0.00173 | -154.6624 | 0.23245 | 8.4018 | 0.002651 | 0.004587 | 0.004427 | 3.7e-05 | 0.001319 | 0.26576 | 1.3362 | 2.5476 | 98.49 | 20.03 | 47.827 | 67.7706 | 85.042 | 62.9048 | 0.011387 | 1.0506 | 0.024323 | 23.1528 | 0.074853 | 0.13829 | 1.9e-05 | 0.001284 | 0.001284 | 4.3393
2019-04-09 15:18 | 1554837456.6995 | -502 | 0.20954 | 0.00044 | 1e-06 | 89.8568 | 89.078 | 122.8219 | 20.6949 | 0.00193 | 0.00182 | -154.6498 | 0.23211 | 8.4006 | 0.002588 | 0.004442 | 0.004265 | 3.4e-05 | 0.001311 | 0.28407 | 1.3855 | 2.5476 | 98.49 | 20.03 | 47.822 | 67.7257 | 83.5283 | 65.0758 | 0.011385 | 1.0438 | 0.024285 | 23.2657 | 0.077868 | 0.13808 | 1.1e-05 | 0.000678 | 0.000678 | 8.209
2019-04-09 15:18 | 1554837457.6985 | -501 | 0.20953 | 0.000439 | 2e-06 | 89.8768 | 88.9468 | 123.108 | 20.6946 | 0.00193 | 0.001755 | -154.6346 | 0.23248 | 8.4043 | 0.002637 | 0.004045 | 0.003881 | 3.4e-05 | 0.001305 | 0.29214 | 1.3466 | 2.5475 | 98.49 | 20.04 | 47.821 | 67.7073 | 84.7061 | 63.514 | 0.011392 | 1.0487 | 0.02437 | 23.2389 | 0.09303 | 0.1383 | 1.2e-05 | 0.000776 | 0.000776 | 7.1801
2019-04-09 15:18 | 1554837458.6988 | -500 | 0.20953 | 0.000443 | 2e-06 | 89.9203 | 89.0676 | 122.9087 | 20.6951 | 0.001932 | 0.001751 | -154.6338 | 0.23241 | 8.4005 | 0.002686 | 0.003876 | 0.003627 | 3.5e-05 | 0.001305 | 0.27865 | 1.3559 | 2.5476 | 98.49 | 20.04 | 47.82 | 67.7533 | 85.8929 | 63.4129 | 0.011392 | 1.049 | 0.024354 | 23.2166 | 0.085217 | 0.13826 | 6e-06 | 0.000328 | 0.000328 | 17.0128
2019-04-09 15:18 | 1554837459.699 | -499 | 0.20954 | 0.000439 | 2e-06 | 89.8866 | 89.0917 | 122.8108 | 20.6942 | 0.001953 | 0.001802 | -154.6329 | 0.23227 | 8.4034 | 0.002661 | 0.004182 | 0.004018 | 3.2e-05 | 0.001303 | 0.28423 | 1.3615 | 2.5475 | 98.49 | 20.04 | 47.814 | 68.2563 | 85.2919 | 64.6456 | 0.011391 | 1.0451 | 0.024299 | 23.2498 | 0.07945 | 0.13818 | 1.1e-05 | 0.000691 | 0.000691 | 8.0632
2019-04-09 15:18 | 1554837460.6985 | -498 | 0.20958 | 0.000442 | 2e-06 | 89.9447 | 89.0659 | 122.7528 | 20.6948 | 0.001922 | 0.001812 | -154.6243 | 0.23239 | 8.4008 | 0.002646 | 0.004091 | 0.003891 | 3.3e-05 | 0.001303 | 0.28224 | 1.373 | 2.5477 | 98.49 | 20.04 | 47.81 | 67.5142 | 84.9295 | 64.8868 | 0.01139 | 1.0444 | 0.024285 | 23.2526 | 0.087818 | 0.13825 | 8e-06 | 0.000465 | 0.000465 | 11.9692
2019-04-09 15:18 | 1554837461.699 | -497 | 0.20953 | 0.000438 | 2e-06 | 89.833 | 88.9214 | 122.6766 | 20.6955 | 0.001928 | 0.001768 | -154.6114 | 0.23286 | 8.4033 | 0.00269 | 0.003878 | 0.003707 | 3.2e-05 | 0.001303 | 0.28119 | 1.4003 | 2.5477 | 98.49 | 20.04 | 47.802 | 67.664 | 85.9831 | 63.8232 | 0.011388 | 1.0477 | 0.024316 | 23.2088 | 0.091234 | 0.13853 | 1.3e-05 | 0.000863 | 0.000863 | 6.4666
2019-04-09 15:18 | 1554837462.6977 | -496 | 0.20955 | 0.000441 | 2e-06 | 89.9578 | 89.1382 | 122.3876 | 20.6947 | 0.001894 | 0.001835 | -154.6151 | 0.23313 | 8.4037 | 0.002651 | 0.004284 | 0.004104 | 3.6e-05 | 0.001302 | 0.28672 | 1.37 | 2.5477 | 98.49 | 20.04 | 47.798 | 66.8509 | 85.0496 | 65.4306 | 0.011387 | 1.0427 | 0.024229 | 23.2367 | 0.081866 | 0.13869 | 9e-06 | 0.000576 | 0.000576 | 9.698
2019-04-09 15:18 | 1554837463.6987 | -495 | 0.20955 | 0.000442 | 2e-06 | 89.8914 | 89.1843 | 122.2212 | 20.6947 | 0.001896 | 0.001806 | -154.6012 | 0.23271 | 8.402 | 0.002657 | 0.004048 | 0.003857 | 3.5e-05 | 0.001297 | 0.29535 | 1.3518 | 2.5476 | 98.49 | 20.04 | 47.793 | 66.9055 | 85.1947 | 64.7344 | 0.011386 | 1.0449 | 0.024238 | 23.197 | 0.070634 | 0.13844 | 7e-06 | 0.000431 | 0.000431 | 12.9492
2019-04-09 15:18 | 1554837464.6987 | -494 | 0.20954 | 0.000438 | 2e-06 | 89.8378 | 89.0883 | 122.3592 | 20.6946 | 0.001902 | 0.001777 | -154.5816 | 0.23278 | 8.401 | 0.002633 | 0.003842 | 0.003622 | 3.4e-05 | 0.001297 | 0.28567 | 1.3623 | 2.5476 | 98.49 | 20.04 | 47.788 | 67.0482 | 84.6171 | 64.0372 | 0.011384 | 1.047 | 0.024277 | 23.1861 | 0.074934 | 0.13848 | 1.2e-05 | 0.000793 | 0.000793 | 7.0341
2019-04-09 15:18 | 1554837465.6978 | -493 | 0.20953 | 0.000439 | 2e-06 | 89.9086 | 89.2442 | 122.647 | 20.6949 | 0.0019 | 0.001765 | -154.5773 | 0.23271 | 8.4017 | 0.002631 | 0.004194 | 0.003993 | 3.3e-05 | 0.001302 | 0.29214 | 1.3492 | 2.5477 | 98.49 | 20.04 | 47.787 | 66.9884 | 84.5645 | 63.7484 | 0.011384 | 1.0479 | 0.024316 | 23.2034 | 0.066349 | 0.13844 | 1.1e-05 | 0.00067 | 0.00067 | 8.3317
2019-04-09 15:18 | 1554837466.6988 | -492 | 0.20956 | 0.000441 | 2e-06 | 89.9063 | 89.0621 | 122.359 | 20.695 | 0.001861 | 0.001785 | -154.5853 | 0.23187 | 8.402 | 0.002712 | 0.004255 | 0.004037 | 3.5e-05 | 0.00131 | 0.27959 | 1.3775 | 2.5477 | 98.49 | 20.04 | 47.788 | 66.0585 | 86.5109 | 64.2227 | 0.011384 | 1.0465 | 0.02427 | 23.1925 | 0.084377 | 0.13794 | 9e-06 | 0.000526 | 0.000526 | 10.5606
2019-04-09 15:18 | 1554837467.6979 | -491 | 0.20952 | 0.000441 | 2e-06 | 89.897 | 88.9393 | 122.0552 | 20.694 | 0.001826 | 0.001766 | -154.5939 | 0.23253 | 8.4016 | 0.002637 | 0.003774 | 0.003653 | 3.3e-05 | 0.001309 | 0.28871 | 1.3567 | 2.5477 | 98.49 | 20.04 | 47.791 | 65.2154 | 84.7093 | 63.7733 | 0.011385 | 1.0479 | 0.024256 | 23.1483 | 0.095807 | 0.13833 | 9e-06 | 0.000564 | 0.000564 | 9.8929
2019-04-09 15:18 | 1554837468.6992 | -490 | 0.20955 | 0.000437 | 1e-06 | 89.8997 | 88.9916 | 122.347 | 20.6947 | 0.001888 | 0.001727 | -154.6 | 0.23247 | 8.4033 | 0.002619 | 0.004169 | 0.003917 | 3.6e-05 | 0.001308 | 0.28407 | 1.3314 | 2.5477 | 98.49 | 20.04 | 47.795 | 66.6955 | 84.281 | 62.8321 | 0.011386 | 1.0508 | 0.024319 | 23.1435 | 0.090814 | 0.13829 | 1.5e-05 | 0.000999 | 0.000999 | 5.5796
2019-04-09 15:18 | 1554837469.6991 | -489 | 0.2095 | 0.000441 | 2e-06 | 89.9444 | 89.1034 | 122.436 | 20.6951 | 0.001919 | 0.00176 | -154.5908 | 0.23222 | 8.4027 | 0.002631 | 0.003729 | 0.003547 | 3.6e-05 | 0.001306 | 0.27721 | 1.4069 | 2.5476 | 98.49 | 20.03 | 47.801 | 67.4412 | 84.5505 | 63.6175 | 0.01138 | 1.0483 | 0.024299 | 23.179 | 0.084022 | 0.13815 | 8e-06 | 0.000503 | 0.000503 | 11.0604
2019-04-09 15:18 | 1554837470.6993 | -488 | 0.20954 | 0.000443 | 2e-06 | 90.0147 | 89.1027 | 122.3071 | 20.6939 | 0.001879 | 0.001765 | -154.5935 | 0.23217 | 8.4037 | 0.002607 | 0.003464 | 0.003313 | 3.1e-05 | 0.001312 | 0.2844 | 1.388 | 2.5475 | 98.49 | 20.03 | 47.804 | 66.489 | 83.972 | 63.7437 | 0.011381 | 1.0479 | 0.024282 | 23.1711 | 0.091077 | 0.13812 | 7e-06 | 0.000437 | 0.000437 | 12.7467
2019-04-09 15:18 | 1554837471.6979 | -487 | 0.20953 | 0.000437 | 2e-06 | 90.0581 | 89.0972 | 122.2423 | 20.6949 | 0.001891 | 0.001825 | -154.5967 | 0.23271 | 8.4024 | 0.002563 | 0.003587 | 0.003386 | 3.2e-05 | 0.001315 | 0.26825 | 1.3749 | 2.5475 | 98.49 | 20.03 | 47.806 | 66.7895 | 82.9291 | 65.196 | 0.011382 | 1.0434 | 0.024223 | 23.2148 | 0.095949 | 0.13844 | 1.4e-05 | 0.000885 | 0.000885 | 6.3035
2019-04-09 15:18 | 1554837472.699 | -486 | 0.2095 | 0.00044 | 2e-06 | 90.0837 | 89.2703 | 122.8315 | 20.6942 | 0.001968 | 0.001805 | -154.5967 | 0.23255 | 8.4027 | 0.00259 | 0.004629 | 0.004413 | 3.3e-05 | 0.001322 | 0.28567 | 1.3723 | 2.5478 | 98.49 | 20.03 | 47.812 | 68.6267 | 83.5783 | 64.7012 | 0.011383 | 1.045 | 0.024299 | 23.2537 | 0.081131 | 0.13834 | 1e-05 | 0.000625 | 0.000625 | 8.9237
2019-04-09 15:18 | 1554837473.6994 | -485 | 0.20953 | 0.000438 | 2e-06 | 90.1186 | 89.4792 | 123.2965 | 20.6941 | 0.001876 | 0.001772 | -154.5879 | 0.23328 | 8.4035 | 0.002597 | 0.004057 | 0.003796 | 3.4e-05 | 0.001331 | 0.27848 | 1.3435 | 2.5475 | 98.49 | 20.03 | 47.815 | 66.4172 | 83.7406 | 63.9181 | 0.011384 | 1.0474 | 0.024374 | 23.2707 | 0.063683 | 0.13878 | 1.2e-05 | 0.000754 | 0.000754 | 7.4153
2019-04-09 15:18 | 1554837474.6981 | -484 | 0.20951 | 0.000441 | 2e-06 | 90.1457 | 89.2445 | 123.0735 | 20.6942 | 0.001865 | 0.001777 | -154.5973 | 0.23265 | 8.4032 | 0.002595 | 0.004421 | 0.004249 | 3.4e-05 | 0.001337 | 0.27057 | 1.3876 | 2.5476 | 98.49 | 20.03 | 47.821 | 66.1559 | 83.6868 | 64.0258 | 0.011385 | 1.0471 | 0.024348 | 23.2533 | 0.089865 | 0.13841 | 8e-06 | 0.000487 | 0.000487 | 11.4458
2019-04-09 15:18 | 1554837475.6974 | -483 | 0.20952 | 0.000443 | 2e-06 | 90.3345 | 89.4486 | 123.0265 | 20.6954 | 0.001944 | 0.001718 | -154.5972 | 0.23198 | 8.4045 | 0.002672 | 0.004518 | 0.004247 | 3.3e-05 | 0.001336 | 0.28711 | 1.365 | 2.5475 | 98.49 | 20.03 | 47.823 | 68.0444 | 85.5427 | 62.6142 | 0.011386 | 1.0515 | 0.024394 | 23.2002 | 0.088148 | 0.13801 | 7e-06 | 0.000362 | 0.000362 | 15.3621
2019-04-09 15:18 | 1554837476.6971 | -482 | 0.20952 | 0.000439 | 2e-06 | 90.3854 | 89.6541 | 123.3216 | 20.6948 | 0.001852 | 0.00179 | -154.6049 | 0.23246 | 8.3999 | 0.002644 | 0.004094 | 0.003925 | 3.4e-05 | 0.001336 | 0.28047 | 1.3524 | 2.5477 | 98.49 | 20.03 | 47.825 | 65.8303 | 84.8764 | 64.3467 | 0.011386 | 1.0461 | 0.024361 | 23.2878 | 0.072667 | 0.13829 | 1.2e-05 | 0.000752 | 0.000752 | 7.4069
2019-04-09 15:18 | 1554837477.6985 | -481 | 0.20954 | 0.000443 | 2e-06 | 90.5129 | 89.7291 | 123.434 | 20.6951 | 0.001877 | 0.001756 | -154.6146 | 0.23202 | 8.4014 | 0.002651 | 0.003779 | 0.003599 | 3.4e-05 | 0.001337 | 0.28407 | 1.3638 | 2.5476 | 98.49 | 20.03 | 47.826 | 66.4314 | 85.0351 | 63.5284 | 0.011386 | 1.0486 | 0.024401 | 23.2702 | 0.077789 | 0.13803 | 6e-06 | 0.000343 | 0.000343 | 16.2235
2019-04-09 15:18 | 1554837478.6972 | -480 | 0.20949 | 0.000438 | 2e-06 | 90.6585 | 89.9518 | 123.0423 | 20.6949 | 0.001818 | 0.001787 | -154.6133 | 0.23301 | 8.4055 | 0.002672 | 0.003647 | 0.003497 | 3.2e-05 | 0.001342 | 0.27168 | 1.3397 | 2.5477 | 98.49 | 20.03 | 47.826 | 65.0306 | 85.5398 | 64.2734 | 0.011386 | 1.0463 | 0.024336 | 23.2589 | 0.069999 | 0.13862 | 1.2e-05 | 0.000746 | 0.000746 | 7.4848
2019-04-09 15:18 | 1554837479.6986 | -479 | 0.2095 | 0.000441 | 2e-06 | 90.8011 | 90.234 | 122.8012 | 20.694 | 0.001846 | 0.001732 | -154.6282 | 0.23291 | 8.4032 | 0.002657 | 0.004503 | 0.004275 | 3.3e-05 | 0.001344 | 0.2943 | 1.3537 | 2.5476 | 98.49 | 20.03 | 47.826 | 65.6961 | 85.1839 | 62.9574 | 0.011386 | 1.0504 | 0.024359 | 23.1908 | 0.056044 | 0.13856 | 8e-06 | 0.000454 | 0.000454 | 12.2913
2019-04-09 15:18 | 1554837480.6987 | -478 | 0.20955 | 0.000443 | 2e-06 | 90.9375 | 90.1476 | 123.2359 | 20.6935 | 0.001837 | 0.001761 | -154.6292 | 0.23302 | 8.4016 | 0.002635 | 0.003511 | 0.003324 | 3.5e-05 | 0.001344 | 0.27704 | 1.3385 | 2.5476 | 98.49 | 20.03 | 47.826 | 65.4664 | 84.6459 | 63.6518 | 0.011386 | 1.0482 | 0.024377 | 23.2558 | 0.078036 | 0.13862 | 6e-06 | 0.000364 | 0.000364 | 15.3562
2019-04-09 15:18 | 1554837481.6972 | -477 | 0.20951 | 0.000438 | 2e-06 | 90.9232 | 90.2381 | 123.3343 | 20.6938 | 0.001872 | 0.001772 | -154.6278 | 0.2335 | 8.4019 | 0.002496 | 0.004038 | 0.003765 | 3.4e-05 | 0.001343 | 0.27887 | 1.3462 | 2.5474 | 98.49 | 20.03 | 47.826 | 66.3296 | 81.3192 | 63.9178 | 0.011386 | 1.0474 | 0.024377 | 23.2742 | 0.067653 | 0.13891 | 1.2e-05 | 0.000789 | 0.000789 | 7.0966
2019-04-09 15:18 | 1554837482.6973 | -476 | 0.20953 | 0.000444 | 2e-06 | 90.9521 | 90.2649 | 122.901 | 20.6947 | 0.001875 | 0.00178 | -154.6205 | 0.23254 | 8.4004 | 0.00259 | 0.004199 | 0.00402 | 3.6e-05 | 0.001344 | 0.30541 | 1.3302 | 2.5475 | 98.49 | 20.03 | 47.828 | 66.3976 | 83.5758 | 64.1076 | 0.011387 | 1.0468 | 0.024328 | 23.2398 | 0.067834 | 0.13834 | 5e-06 | 0.000256 | 0.000256 | 21.7595
2019-04-09 15:18 | 1554837483.6984 | -475 | 0.20954 | 0.00044 | 2e-06 | 91.0659 | 90.2718 | 122.5928 | 20.6946 | 0.001891 | 0.001754 | -154.6342 | 0.23297 | 8.3999 | 0.002574 | 0.004111 | 0.00395 | 3.6e-05 | 0.001337 | 0.27759 | 1.3291 | 2.5478 | 98.49 | 20.02 | 47.828 | 66.7688 | 83.1957 | 63.4813 | 0.01138 | 1.0488 | 0.02432 | 23.1891 | 0.078332 | 0.1386 | 1e-05 | 0.000643 | 0.000643 | 8.6818
2019-04-09 15:18 | 1554837484.6974 | -474 | 0.20951 | 0.000441 | 2e-06 | 91.0897 | 90.382 | 122.9498 | 20.6945 | 0.00191 | 0.001735 | -154.6338 | 0.23275 | 8.3993 | 0.002543 | 0.003929 | 0.003756 | 3.5e-05 | 0.00133 | 0.27759 | 1.3431 | 2.5475 | 98.49 | 20.02 | 47.828 | 67.226 | 82.4412 | 63.0211 | 0.01138 | 1.0502 | 0.024372 | 23.207 | 0.069766 | 0.13846 | 9e-06 | 0.000528 | 0.000528 | 10.5755
2019-04-09 15:18 | 1554837485.6974 | -473 | 0.20953 | 0.000441 | 2e-06 | 91.143 | 90.4081 | 122.8503 | 20.6953 | 0.001923 | 0.001814 | -154.6399 | 0.23294 | 8.4021 | 0.002594 | 0.003899 | 0.003724 | 3.6e-05 | 0.001329 | 0.30663 | 1.348 | 2.5475 | 98.49 | 20.02 | 47.829 | 67.5466 | 83.6621 | 64.9349 | 0.01138 | 1.0443 | 0.024293 | 23.2635 | 0.072411 | 0.13858 | 9e-06 | 0.000543 | 0.000543 | 10.2936
2019-04-09 15:18 | 1554837486.6984 | -472 | 0.20956 | 0.000441 | 2e-06 | 91.2043 | 90.6397 | 122.8162 | 20.6948 | 0.001891 | 0.001771 | -154.6112 | 0.23255 | 8.402 | 0.00261 | 0.003575 | 0.003484 | 3.3e-05 | 0.001326 | 0.2756 | 1.3304 | 2.5476 | 98.49 | 20.02 | 47.83 | 66.7722 | 84.0534 | 63.8899 | 0.01138 | 1.0475 | 0.024327 | 23.2243 | 0.055544 | 0.13834 | 8e-06 | 0.000442 | 0.000442 | 12.6187
2019-04-09 15:18 | 1554837487.6979 | -471 | 0.20953 | 0.000441 | 2e-06 | 91.2627 | 90.6015 | 122.9303 | 20.6937 | 0.001858 | 0.001795 | -154.6154 | 0.23283 | 8.4001 | 0.00258 | 0.003801 | 0.003509 | 3.4e-05 | 0.001331 | 0.286 | 1.3648 | 2.5476 | 98.49 | 20.02 | 47.832 | 65.9909 | 83.3408 | 64.4664 | 0.011381 | 1.0457 | 0.024318 | 23.255 | 0.065037 | 0.13851 | 9e-06 | 0.000561 | 0.000561 | 9.9579
2019-04-09 15:18 | 1554837488.6985 | -470 | 0.20955 | 0.000437 | 2e-06 | 91.2758 | 90.8128 | 123.0653 | 20.6935 | 0.001918 | 0.001769 | -154.6086 | 0.23258 | 8.4019 | 0.002531 | 0.003326 | 0.003216 | 3.1e-05 | 0.001334 | 0.25713 | 1.3584 | 2.5477 | 98.49 | 20.02 | 47.834 | 67.4208 | 82.1518 | 63.8306 | 0.011381 | 1.0477 | 0.024354 | 23.2458 | 0.045484 | 0.13836 | 1.3e-05 | 0.000846 | 0.000846 | 6.5923
2019-04-09 15:18 | 1554837489.6977 | -469 | 0.20955 | 0.000444 | 2e-06 | 91.3622 | 90.6938 | 123.4277 | 20.6949 | 0.001881 | 0.001747 | -154.5886 | 0.23248 | 8.4025 | 0.002647 | 0.004798 | 0.004613 | 3.1e-05 | 0.001331 | 0.28064 | 1.3413 | 2.5476 | 98.49 | 20.02 | 47.836 | 66.5476 | 84.9484 | 63.3143 | 0.011382 | 1.0493 | 0.024409 | 23.2622 | 0.065679 | 0.1383 | 5e-06 | 0.000265 | 0.000265 | 21.0405
2019-04-09 15:18 | 1554837490.6976 | -468 | 0.20954 | 0.000437 | 2e-06 | 91.4443 | 90.9643 | 122.9953 | 20.6949 | 0.001879 | 0.001821 | -154.5941 | 0.23275 | 8.4021 | 0.002613 | 0.004296 | 0.004125 | 3.5e-05 | 0.001326 | 0.27433 | 1.3528 | 2.5477 | 98.49 | 20.02 | 47.838 | 66.4844 | 84.1174 | 65.0883 | 0.011382 | 1.0438 | 0.024302 | 23.2825 | 0.047083 | 0.13846 | 1.4e-05 | 0.000949 | 0.000949 | 5.8776
2019-04-09 15:18 | 1554837491.6976 | -467 | 0.20954 | 0.000442 | 2e-06 | 91.4619 | 90.9147 | 122.9125 | 20.6947 | 0.001894 | 0.001763 | -154.6032 | 0.23253 | 8.4011 | 0.002631 | 0.003563 | 0.003329 | 3.1e-05 | 0.001321 | 0.29103 | 1.3784 | 2.5476 | 98.49 | 20.02 | 47.838 | 66.8391 | 84.5611 | 63.7016 | 0.011382 | 1.0481 | 0.024344 | 23.2269 | 0.053676 | 0.13833 | 7e-06 | 0.000389 | 0.000389 | 14.3392
2019-04-09 15:18 | 1554837492.6966 | -466 | 0.20951 | 0.000446 | 2e-06 | 91.6323 | 91.0444 | 123.1244 | 20.694 | 0.001825 | 0.001786 | -154.5979 | 0.23188 | 8.4034 | 0.002665 | 0.00419 | 0.003991 | 3.2e-05 | 0.001324 | 0.28584 | 1.3935 | 2.5476 | 98.49 | 20.02 | 47.838 | 65.1849 | 85.3908 | 64.2412 | 0.011382 | 1.0464 | 0.024345 | 23.2656 | 0.057563 | 0.13795 | 3e-06 | 8.2e-05 | 8.2e-05 | 67.7208
2019-04-09 15:18 | 1554837493.6976 | -465 | 0.20954 | 0.000437 | 2e-06 | 91.7389 | 91.0266 | 123.1998 | 20.6944 | 0.001928 | 0.001727 | -154.5876 | 0.23293 | 8.4015 | 0.00269 | 0.00357 | 0.00341 | 3.3e-05 | 0.001324 | 0.27671 | 1.3701 | 2.5477 | 98.49 | 20.02 | 47.837 | 67.672 | 85.975 | 62.8349 | 0.011382 | 1.0508 | 0.024403 | 23.2241 | 0.069722 | 0.13857 | 1.5e-05 | 0.000958 | 0.000958 | 5.8274
2019-04-09 15:18 | 1554837494.6967 | -464 | 0.20952 | 0.000439 | 2e-06 | 91.8297 | 91.3153 | 122.9237 | 20.6944 | 0.0019 | 0.001777 | -154.5774 | 0.23279 | 8.4009 | 0.002706 | 0.004071 | 0.00395 | 3.3e-05 | 0.001323 | 0.29214 | 1.3166 | 2.5476 | 98.49 | 20.02 | 47.836 | 66.9928 | 86.3559 | 64.0308 | 0.011382 | 1.0471 | 0.024333 | 23.2393 | 0.050245 | 0.13849 | 1.1e-05 | 0.000622 | 0.000622 | 8.973
2019-04-09 15:18 | 1554837495.719 | -463 | 0.20956 | 0.00044 | 2e-06 | 91.919 | 91.4096 | 122.947 | 20.695 | 0.00183 | 0.001774 | -154.5655 | 0.23295 | 8.4009 | 0.002701 | 0.004207 | 0.003943 | 3.1e-05 | 0.001316 | 0.27096 | 1.3612 | 2.5476 | 98.49 | 20.02 | 47.835 | 65.3201 | 86.2354 | 63.965 | 0.011381 | 1.0473 | 0.024337 | 23.2393 | 0.04971 | 0.13858 | 1e-05 | 0.000669 | 0.000669 | 8.3496
2019-04-09 15:18 | 1554837496.7091 | -462 | 0.20956 | 0.000439 | 2e-06 | 92.0131 | 91.5648 | 122.951 | 20.6932 | 0.001852 | 0.001779 | -154.581 | 0.23289 | 8.4039 | 0.002662 | 0.003954 | 0.003725 | 3.1e-05 | 0.001315 | 0.26178 | 1.3105 | 2.5475 | 98.49 | 20.02 | 47.833 | 65.85 | 85.2986 | 64.0763 | 0.011381 | 1.0469 | 0.024334 | 23.2435 | 0.043688 | 0.13855 | 1.2e-05 | 0.000767 | 0.000767 | 7.2775
2019-04-09 15:18 | 1554837497.7097 | -461 | 0.20954 | 0.000441 | 2e-06 | 92.2025 | 91.7348 | 122.8911 | 20.6947 | 0.001831 | 0.001735 | -154.5664 | 0.23291 | 8.4013 | 0.002649 | 0.003547 | 0.003382 | 3.2e-05 | 0.001321 | 0.27472 | 1.3574 | 2.5476 | 98.49 | 20.02 | 47.83 | 65.3246 | 84.9832 | 63.0353 | 0.01138 | 1.0502 | 0.024366 | 23.2019 | 0.045487 | 0.13856 | 9e-06 | 0.000571 | 0.000571 | 9.778
2019-04-09 15:18 | 1554837498.7101 | -460 | 0.20958 | 0.000443 | 2e-06 | 92.2317 | 91.8759 | 122.7139 | 20.6941 | 0.001858 | 0.001763 | -154.5674 | 0.23228 | 8.4032 | 0.002642 | 0.003657 | 0.003529 | 3.2e-05 | 0.001321 | 0.28512 | 1.3515 | 2.5476 | 98.49 | 20.02 | 47.829 | 65.9802 | 84.8305 | 63.6875 | 0.01138 | 1.0481 | 0.024324 | 23.2077 | 0.034571 | 0.13818 | 6e-06 | 0.000269 | 0.000269 | 20.7028
2019-04-09 15:18 | 1554837499.7102 | -459 | 0.20953 | 0.000437 | 2e-06 | 92.3449 | 91.9141 | 122.5829 | 20.6951 | 0.001913 | 0.00177 | -154.5655 | 0.23209 | 8.3991 | 0.002632 | 0.003788 | 0.003605 | 3.5e-05 | 0.001324 | 0.29717 | 1.3258 | 2.5475 | 98.49 | 20.02 | 47.829 | 67.2952 | 84.5762 | 63.8745 | 0.01138 | 1.0475 | 0.024305 | 23.2017 | 0.04182 | 0.13807 | 1.4e-05 | 0.000892 | 0.000892 | 6.2378
2019-04-09 15:18 | 1554837500.7102 | -458 | 0.20952 | 0.00044 | 1e-06 | 92.4512 | 91.9695 | 122.6815 | 20.6942 | 0.001901 | 0.001754 | -154.5695 | 0.23308 | 8.402 | 0.0026 | 0.004237 | 0.004017 | 3.4e-05 | 0.001322 | 0.27865 | 1.3719 | 2.5476 | 98.49 | 20.02 | 47.826 | 67.0267 | 83.8171 | 63.4854 | 0.011379 | 1.0487 | 0.024328 | 23.1976 | 0.046721 | 0.13866 | 1e-05 | 0.000671 | 0.000671 | 8.3321
2019-04-09 15:18 | 1554837501.7092 | -457 | 0.20953 | 0.000443 | 1e-06 | 92.5381 | 92.0759 | 122.7712 | 20.6941 | 0.001906 | 0.001732 | -154.561 | 0.23206 | 8.4023 | 0.002639 | 0.003965 | 0.003826 | 3.2e-05 | 0.001322 | 0.28711 | 1.3702 | 2.5476 | 98.49 | 20.02 | 47.826 | 67.1421 | 84.7627 | 62.9558 | 0.011379 | 1.0504 | 0.024357 | 23.1879 | 0.044793 | 0.13805 | 7e-06 | 0.000453 | 0.000453 | 12.286
2019-04-09 15:18 | 1554837502.7107 | -456 | 0.20953 | 0.000444 | 2e-06 | 92.6534 | 92.2438 | 122.3219 | 20.6936 | 0.00188 | 0.001799 | -154.5756 | 0.23221 | 8.4024 | 0.002691 | 0.00381 | 0.00361 | 3.5e-05 | 0.001311 | 0.28744 | 1.3495 | 2.5475 | 98.49 | 20.01 | 47.825 | 66.512 | 85.9976 | 64.5585 | 0.011372 | 1.0454 | 0.024254 | 23.2005 | 0.039625 | 0.13814 | 5e-06 | 0.000257 | 0.000257 | 21.6408
2019-04-09 15:18 | 1554837503.7105 | -455 | 0.20954 | 0.000443 | 1e-06 | 92.7225 | 92.5432 | 121.9968 | 20.6945 | 0.001917 | 0.001759 | -154.5647 | 0.23224 | 8.4031 | 0.00261 | 0.003567 | 0.003427 | 3e-05 | 0.001314 | 0.28279 | 1.3715 | 2.5477 | 98.49 | 20.02 | 47.825 | 67.4069 | 84.049 | 63.6061 | 0.011379 | 1.0484 | 0.024256 | 23.137 | 0.017307 | 0.13816 | 7e-06 | 0.000428 | 0.000428 | 13.0091
2019-04-09 15:18 | 1554837504.7097 | -454 | 0.20955 | 0.000439 | 2e-06 | 92.8157 | 92.2655 | 122.4536 | 20.6933 | 0.001907 | 0.001791 | -154.5814 | 0.23222 | 8.4015 | 0.002664 | 0.003764 | 0.003542 | 2.9e-05 | 0.001306 | 0.27489 | 1.3494 | 2.5476 | 98.49 | 20.02 | 47.825 | 67.1701 | 85.3549 | 64.3817 | 0.011379 | 1.046 | 0.024274 | 23.2069 | 0.053177 | 0.13815 | 1.2e-05 | 0.000786 | 0.000786 | 7.0792
2019-04-09 15:18 | 1554837505.7097 | -453 | 0.20954 | 0.000446 | 2e-06 | 92.9464 | 92.5615 | 122.5982 | 20.6944 | 0.001862 | 0.001804 | -154.587 | 0.23304 | 8.4007 | 0.002635 | 0.003895 | 0.003728 | 3.3e-05 | 0.001306 | 0.29159 | 1.3334 | 2.5476 | 98.49 | 20.01 | 47.825 | 66.0827 | 84.6693 | 64.674 | 0.011372 | 1.0451 | 0.024277 | 23.2307 | 0.03712 | 0.13863 | 2e-06 | 5.1e-05 | 5.1e-05 | 108.8635
2019-04-09 15:18 | 1554837506.7094 | -452 | 0.20954 | 0.000442 | 2e-06 | 93.0676 | 92.7628 | 122.4435 | 20.6938 | 0.001858 | 0.001769 | -154.5815 | 0.23361 | 8.4019 | 0.002575 | 0.003919 | 0.003751 | 3.3e-05 | 0.001301 | 0.26145 | 1.348 | 2.5475 | 98.49 | 20.01 | 47.827 | 65.9715 | 83.2159 | 63.8358 | 0.011372 | 1.0477 | 0.024292 | 23.1872 | 0.029342 | 0.13897 | 7e-06 | 0.000358 | 0.000358 | 15.6313
2019-04-09 15:18 | 1554837507.7094 | -451 | 0.20953 | 0.000442 | 2e-06 | 93.1519 | 92.6726 | 122.1383 | 20.6939 | 0.001834 | 0.001788 | -154.5999 | 0.23215 | 8.4037 | 0.002726 | 0.00409 | 0.003879 | 3.1e-05 | 0.0013 | 0.28888 | 1.3265 | 2.5475 | 98.49 | 20.01 | 47.828 | 65.4135 | 86.8486 | 64.2988 | 0.011373 | 1.0462 | 0.024245 | 23.1742 | 0.046137 | 0.1381 | 7e-06 | 0.000405 | 0.000405 | 13.7487
2019-04-09 15:18 | 1554837508.7094 | -450 | 0.20951 | 0.000441 | 2e-06 | 93.3228 | 92.9707 | 122.3303 | 20.694 | 0.001878 | 0.001746 | -154.6044 | 0.23162 | 8.4 | 0.002607 | 0.004322 | 0.004144 | 3.3e-05 | 0.001298 | 0.29336 | 1.3147 | 2.5475 | 98.49 | 20.01 | 47.83 | 66.4585 | 83.9877 | 63.2823 | 0.011373 | 1.0494 | 0.024301 | 23.1574 | 0.033815 | 0.13779 | 9e-06 | 0.000546 | 0.000546 | 10.1728
2019-04-09 15:18 | 1554837509.7098 | -449 | 0.20952 | 0.000441 | 2e-06 | 93.3675 | 92.9772 | 122.3238 | 20.6946 | 0.001878 | 0.001746 | -154.6338 | 0.23226 | 8.4038 | 0.002744 | 0.003905 | 0.003731 | 3.6e-05 | 0.001297 | 0.28208 | 1.3502 | 2.5476 | 98.49 | 20.01 | 47.832 | 66.466 | 87.275 | 63.2899 | 0.011374 | 1.0494 | 0.0243 | 23.1571 | 0.037466 | 0.13817 | 9e-06 | 0.000493 | 0.000493 | 11.2832
2019-04-09 15:18 | 1554837510.7085 | -448 | 0.20951 | 0.000438 | 2e-06 | 93.4479 | 93.2636 | 122.1365 | 20.6936 | 0.001898 | 0.001762 | -154.6293 | 0.23163 | 8.4031 | 0.002699 | 0.003783 | 0.003625 | 3.3e-05 | 0.001298 | 0.28998 | 1.3566 | 2.5476 | 98.49 | 20.01 | 47.835 | 66.9381 | 86.1857 | 63.6839 | 0.011374 | 1.0481 | 0.024267 | 23.1529 | 0.017662 | 0.1378 | 1.2e-05 | 0.000764 | 0.000764 | 7.2657
2019-04-09 15:19 | 1554837511.7096 | -447 | 0.20953 | 0.000444 | 2e-06 | 93.5307 | 93.2821 | 122.0684 | 20.6943 | 0.001896 | 0.001768 | -154.6254 | 0.23226 | 8.4018 | 0.002665 | 0.004206 | 0.003919 | 3.4e-05 | 0.001305 | 0.28854 | 1.3669 | 2.5477 | 98.49 | 20.01 | 47.835 | 66.9076 | 85.3675 | 63.8252 | 0.011374 | 1.0477 | 0.024255 | 23.1513 | 0.0238 | 0.13817 | 6e-06 | 0.000285 | 0.000285 | 19.5046
2019-04-09 15:19 | 1554837512.7099 | -446 | 0.2095 | 0.00044 | 1e-06 | 93.7516 | 93.596 | 122.4536 | 20.6945 | 0.001878 | 0.001794 | -154.6299 | 0.23247 | 8.3994 | 0.002743 | 0.004237 | 0.004099 | 3.4e-05 | 0.001311 | 0.28871 | 1.3495 | 2.5476 | 98.49 | 20.01 | 47.836 | 66.4659 | 87.2585 | 64.4399 | 0.011375 | 1.0458 | 0.024271 | 23.2089 | 0.014861 | 0.1383 | 1.1e-05 | 0.000691 | 0.000691 | 8.0707
2019-04-09 15:19 | 1554837513.7096 | -445 | 0.20954 | 0.00044 | 2e-06 | 93.7799 | 93.5513 | 122.7676 | 20.694 | 0.00182 | 0.001785 | -154.6356 | 0.23146 | 8.4028 | 0.002665 | 0.00388 | 0.003641 | 3.2e-05 | 0.001316 | 0.27975 | 1.307 | 2.5476 | 98.49 | 20.01 | 47.837 | 65.0655 | 85.3703 | 64.2311 | 0.011375 | 1.0464 | 0.02431 | 23.2315 | 0.021836 | 0.13769 | 1e-05 | 0.000608 | 0.000608 | 9.1236
2019-04-09 15:19 | 1554837514.7102 | -444 | 0.2095 | 0.000443 | 2e-06 | 93.8261 | 93.5578 | 122.4346 | 20.6941 | 0.001962 | 0.001734 | -154.6248 | 0.23187 | 8.4039 | 0.002719 | 0.004458 | 0.004291 | 3.4e-05 | 0.001315 | 0.28241 | 1.3846 | 2.5477 | 98.49 | 20.01 | 47.838 | 68.4738 | 86.6732 | 63.0052 | 0.011375 | 1.0502 | 0.024321 | 23.1578 | 0.025612 | 0.13794 | 7e-06 | 0.000384 | 0.000384 | 14.4581
2019-04-09 15:19 | 1554837515.7096 | -443 | 0.2095 | 0.000443 | 2e-06 | 93.9244 | 93.7849 | 122.2146 | 20.6936 | 0.001873 | 0.00175 | -154.6328 | 0.23174 | 8.4007 | 0.002704 | 0.003753 | 0.003537 | 3.3e-05 | 0.001319 | 0.29463 | 1.3455 | 2.5476 | 98.49 | 20.01 | 47.839 | 66.3371 | 86.3183 | 63.3857 | 0.011375 | 1.0491 | 0.024286 | 23.15 | 0.013286 | 0.13786 | 7e-06 | 0.000381 | 0.000381 | 14.5899
2019-04-09 15:19 | 1554837516.7091 | -442 | 0.20951 | 0.000444 | 2e-06 | 93.9631 | 93.6724 | 122.2421 | 20.6937 | 0.001887 | 0.001737 | -154.6339 | 0.23203 | 8.4011 | 0.002626 | 0.004024 | 0.003856 | 3.6e-05 | 0.001325 | 0.30342 | 1.3237 | 2.5475 | 98.49 | 20.01 | 47.84 | 66.683 | 84.4298 | 63.0788 | 0.011376 | 1.05 | 0.0243 | 23.1421 | 0.027712 | 0.13803 | 5e-06 | 0.000223 | 0.000223 | 24.9366
2019-04-09 15:19 | 1554837517.7098 | -441 | 0.20953 | 0.000442 | 2e-06 | 94.1334 | 93.9581 | 122.7974 | 20.694 | 0.001903 | 0.001745 | -154.6227 | 0.23344 | 8.4002 | 0.002662 | 0.004314 | 0.004175 | 3.5e-05 | 0.001322 | 0.29214 | 1.3381 | 2.5475 | 98.49 | 20.01 | 47.841 | 67.0594 | 85.301 | 63.2542 | 0.011376 | 1.0495 | 0.024348 | 23.2006 | 0.016678 | 0.13887 | 8e-06 | 0.000482 | 0.000482 | 11.6096
2019-04-09 15:19 | 1554837518.7097 | -440 | 0.20951 | 0.000438 | 1e-06 | 94.1316 | 93.8379 | 122.7594 | 20.6937 | 0.001914 | 0.001793 | -154.6291 | 0.23303 | 8.4009 | 0.002661 | 0.004068 | 0.003863 | 3.2e-05 | 0.00132 | 0.28047 | 1.3658 | 2.5476 | 98.49 | 20.01 | 47.841 | 67.3235 | 85.2813 | 64.4212 | 0.011376 | 1.0458 | 0.024302 | 23.2372 | 0.027949 | 0.13863 | 1.3e-05 | 0.000862 | 0.000862 | 6.4789
2019-04-09 15:19 | 1554837519.7099 | -439 | 0.20951 | 0.000444 | 2e-06 | 94.2376 | 93.9842 | 122.7787 | 20.6945 | 0.001913 | 0.001763 | -154.6257 | 0.23315 | 8.399 | 0.002623 | 0.003948 | 0.003802 | 3.8e-05 | 0.001322 | 0.30182 | 1.385 | 2.5477 | 98.49 | 20.01 | 47.843 | 67.2953 | 84.3736 | 63.6966 | 0.011376 | 1.0481 | 0.02433 | 23.2141 | 0.024086 | 0.1387 | 5e-06 | 0.000233 | 0.000233 | 24.035
2019-04-09 15:19 | 1554837520.7084 | -438 | 0.2095 | 0.000441 | 2e-06 | 94.1799 | 94.0272 | 122.8139 | 20.6938 | 0.001889 | 0.001777 | -154.6038 | 0.23274 | 8.4019 | 0.002627 | 0.003492 | 0.003291 | 3.4e-05 | 0.001323 | 0.28888 | 1.396 | 2.5477 | 98.49 | 20.01 | 47.843 | 66.7307 | 84.4727 | 64.0323 | 0.011376 | 1.047 | 0.024322 | 23.229 | 0.014508 | 0.13846 | 8e-06 | 0.000428 | 0.000428 | 13.0245
2019-04-09 15:19 | 1554837521.7089 | -437 | 0.20952 | 0.000448 | 2e-06 | 94.2213 | 94.0069 | 122.8399 | 20.694 | 0.001868 | 0.001731 | -154.6265 | 0.23236 | 8.4014 | 0.002671 | 0.003928 | 0.003682 | 3.6e-05 | 0.001324 | 0.30646 | 1.3016 | 2.5475 | 98.49 | 20.01 | 47.843 | 66.2227 | 85.5288 | 62.9368 | 0.011376 | 1.0505 | 0.024364 | 23.1937 | 0.02037 | 0.13823 | -1e-06 | -0.000238 | -0.000238 | -23.3691
2019-04-09 15:19 | 1554837522.7079 | -436 | 0.2095 | 0.000439 | 2e-06 | 94.3332 | 94.3659 | 122.7852 | 20.6937 | 0.001931 | 0.001762 | -154.623 | 0.23254 | 8.4021 | 0.002659 | 0.003706 | 0.003542 | 3.5e-05 | 0.001331 | 0.30071 | 1.3675 | 2.5477 | 98.49 | 20.01 | 47.843 | 67.7377 | 85.2375 | 63.67 | 0.011376 | 1.0482 | 0.024332 | 23.2138 | -0.003095 | 0.13834 | 1.1e-05 | 0.000693 | 0.000693 | 8.0498
2019-04-09 15:19 | 1554837523.7089 | -435 | 0.20952 | 0.000441 | 2e-06 | 94.4217 | 94.2964 | 122.7505 | 20.6944 | 0.00188 | 0.001788 | -154.613 | 0.23258 | 8.4031 | 0.002715 | 0.003913 | 0.003731 | 3.8e-05 | 0.001334 | 0.28031 | 1.3551 | 2.5477 | 98.49 | 20.01 | 47.846 | 66.5152 | 86.5931 | 64.3049 | 0.011377 | 1.0462 | 0.024306 | 23.2324 | 0.011879 | 0.13836 | 9e-06 | 0.000514 | 0.000514 | 10.8419
2019-04-09 15:19 | 1554837524.7093 | -434 | 0.20951 | 0.000442 | 2e-06 | 94.5217 | 94.4158 | 122.4848 | 20.6935 | 0.001895 | 0.001816 | -154.6375 | 0.23217 | 8.4028 | 0.002641 | 0.003794 | 0.003641 | 3.9e-05 | 0.001342 | 0.27776 | 1.3329 | 2.5476 | 98.49 | 20.01 | 47.849 | 66.8747 | 84.801 | 64.969 | 0.011378 | 1.0441 | 0.024256 | 23.2301 | 0.010031 | 0.13812 | 6e-06 | 0.000316 | 0.000316 | 17.614
2019-04-09 15:19 | 1554837525.709 | -433 | 0.20952 | 0.000439 | 2e-06 | 94.6385 | 94.4911 | 122.4064 | 20.693 | 0.001884 | 0.001784 | -154.6131 | 0.23333 | 8.4016 | 0.002656 | 0.004166 | 0.003952 | 3.6e-05 | 0.001344 | 0.29015 | 1.3536 | 2.5476 | 98.49 | 20.01 | 47.852 | 66.602 | 85.1524 | 64.2012 | 0.011378 | 1.0465 | 0.024275 | 23.1962 | 0.013934 | 0.13881 | 1.1e-05 | 0.000672 | 0.000672 | 8.3225
2019-04-09 15:19 | 1554837526.708 | -432 | 0.20953 | 0.00044 | 1e-06 | 94.7516 | 94.7173 | 122.5285 | 20.6938 | 0.001892 | 0.001795 | -154.604 | 0.23279 | 8.4024 | 0.002585 | 0.004338 | 0.004111 | 3.5e-05 | 0.001346 | 0.26985 | 1.3439 | 2.5475 | 98.49 | 20.01 | 47.855 | 66.7978 | 83.457 | 64.4602 | 0.011379 | 1.0457 | 0.024278 | 23.2167 | 0.003245 | 0.13849 | 1.1e-05 | 0.000699 | 0.000699 | 7.9861
2019-04-09 15:19 | 1554837527.7091 | -431 | 0.2095 | 0.000445 | 2e-06 | 94.8684 | 94.7038 | 122.0482 | 20.6935 | 0.001883 | 0.001739 | -154.6144 | 0.23203 | 8.4038 | 0.002659 | 0.004514 | 0.004284 | 3.5e-05 | 0.001346 | 0.25498 | 1.3916 | 2.5476 | 98.49 | 20.01 | 47.857 | 66.5914 | 85.2442 | 63.1277 | 0.01138 | 1.0499 | 0.024278 | 23.1254 | 0.015525 | 0.13803 | 3e-06 | 9.7e-05 | 9.7e-05 | 57.2454
2019-04-09 15:19 | 1554837528.7081 | -430 | 0.20952 | 0.000442 | 2e-06 | 94.8987 | 94.9623 | 121.9215 | 20.6931 | 0.001929 | 0.001747 | -154.6205 | 0.23219 | 8.4025 | 0.002592 | 0.004612 | 0.004422 | 3.5e-05 | 0.001343 | 0.2959 | 1.3895 | 2.5476 | 98.49 | 20.01 | 47.861 | 67.6888 | 83.6263 | 63.3166 | 0.01138 | 1.0493 | 0.024259 | 23.1199 | -0.005989 | 0.13813 | 7e-06 | 0.000411 | 0.000411 | 13.5592
2019-04-09 15:19 | 1554837529.7082 | -429 | 0.20953 | 0.00044 | 2e-06 | 94.9505 | 95.0084 | 121.9438 | 20.6922 | 0.001911 | 0.001769 | -154.6233 | 0.2317 | 8.4035 | 0.002628 | 0.003934 | 0.003736 | 3.5e-05 | 0.001336 | 0.28982 | 1.3929 | 2.5476 | 98.49 | 20.01 | 47.864 | 67.2612 | 84.4939 | 63.8491 | 0.011381 | 1.0476 | 0.024242 | 23.1403 | -0.005449 | 0.13784 | 1e-05 | 0.000615 | 0.000615 | 9.0259
2019-04-09 15:19 | 1554837530.7081 | -428 | 0.20951 | 0.000447 | 2e-06 | 95.1441 | 95.1106 | 122.2412 | 20.693 | 0.001901 | 0.001738 | -154.621 | 0.23219 | 8.4014 | 0.002578 | 0.003051 | 0.002944 | 3.5e-05 | 0.001335 | 0.29031 | 1.3438 | 2.5476 | 98.49 | 20.01 | 47.867 | 67.013 | 83.2853 | 63.0845 | 0.011382 | 1.05 | 0.024299 | 23.1422 | 0.003149 | 0.13813 | 1e-06 | -6.3e-05 | -6.3e-05 | -88.0904
2019-04-09 15:19 | 1554837531.7092 | -427 | 0.2095 | 0.000441 | 2e-06 | 95.1265 | 95.0187 | 122.2843 | 20.6928 | 0.001905 | 0.001735 | -154.5992 | 0.23206 | 8.4011 | 0.002727 | 0.003853 | 0.003681 | 4.1e-05 | 0.00133 | 0.29463 | 1.3497 | 2.5476 | 98.49 | 20.01 | 47.877 | 67.115 | 86.8687 | 63.0279 | 0.011384 | 1.0502 | 0.024306 | 23.1443 | 0.010143 | 0.13805 | 9e-06 | 0.000504 | 0.000504 | 11.0475
2019-04-09 15:19 | 1554837532.7092 | -426 | 0.20953 | 0.000445 | 2e-06 | 95.1953 | 95.2376 | 122.3043 | 20.6928 | 0.001889 | 0.00181 | -154.5929 | 0.2327 | 8.402 | 0.002766 | 0.003995 | 0.003788 | 3.6e-05 | 0.001323 | 0.2756 | 1.3452 | 2.5475 | 98.49 | 20.01 | 47.884 | 66.7288 | 87.8181 | 64.8192 | 0.011386 | 1.0446 | 0.024243 | 23.2078 | -0.003972 | 0.13843 | 4e-06 | 0.00016 | 0.00016 | 34.9762
2019-04-09 15:19 | 1554837533.7092 | -425 | 0.20952 | 0.000448 | 2e-06 | 95.1733 | 95.1843 | 122.0402 | 20.6934 | 0.001899 | 0.001798 | -154.585 | 0.23214 | 8.404 | 0.002686 | 0.004697 | 0.004467 | 3.7e-05 | 0.001322 | 0.28335 | 1.3645 | 2.5475 | 98.49 | 20.01 | 47.89 | 66.9605 | 85.8818 | 64.5488 | 0.011387 | 1.0454 | 0.024227 | 23.1735 | -0.001032 | 0.1381 | 0 | -9.4e-05 | -9.4e-05 | -59.0759
2019-04-09 15:19 | 1554837534.7086 | -424 | 0.20954 | 0.000447 | 2e-06 | 95.2123 | 95.1313 | 122.278 | 20.6925 | 0.001926 | 0.001775 | -154.5833 | 0.23258 | 8.4023 | 0.002687 | 0.004365 | 0.004208 | 3.6e-05 | 0.001327 | 0.28816 | 1.3396 | 2.5476 | 98.49 | 20 | 47.898 | 67.6086 | 85.9169 | 63.9765 | 0.011382 | 1.0472 | 0.024271 | 23.1764 | 0.007616 | 0.13836 | 1e-06 | -6.4e-05 | -6.4e-05 | -86.8345
2019-04-09 15:19 | 1554837535.7087 | -423 | 0.20954 | 0.000444 | 2e-06 | 95.2349 | 95.2125 | 122.366 | 20.6932 | 0.001895 | 0.001768 | -154.6052 | 0.2326 | 8.402 | 0.002678 | 0.003604 | 0.003428 | 3.6e-05 | 0.00133 | 0.28279 | 1.3232 | 2.5475 | 98.49 | 20 | 47.913 | 66.8745 | 85.688 | 63.8226 | 0.011386 | 1.0477 | 0.024285 | 23.1794 | 0.002108 | 0.13837 | 5e-06 | 0.000216 | 0.000216 | 25.8071
2019-04-09 15:19 | 1554837536.7074 | -422 | 0.20954 | 0.000448 | 2e-06 | 95.2769 | 95.1282 | 122.3179 | 20.6931 | 0.001829 | 0.001773 | -154.6241 | 0.23322 | 8.4011 | 0.002757 | 0.004406 | 0.004181 | 4e-05 | 0.001333 | 0.27721 | 1.3046 | 2.5475 | 98.49 | 20 | 47.92 | 65.2768 | 87.5936 | 63.9373 | 0.011387 | 1.0473 | 0.024276 | 23.1788 | 0.013976 | 0.13874 | 0 | -0.000105 | -0.000105 | -53.0516
2019-04-09 15:19 | 1554837537.7084 | -421 | 0.20952 | 0.00045 | 2e-06 | 95.3755 | 95.283 | 122.6388 | 20.6927 | 0.001851 | 0.001786 | -154.6024 | 0.23224 | 8.4032 | 0.00273 | 0.003764 | 0.003562 | 3.5e-05 | 0.001334 | 0.29015 | 1.3412 | 2.5477 | 98.49 | 20 | 47.935 | 65.822 | 86.9521 | 64.257 | 0.011391 | 1.0463 | 0.024296 | 23.2202 | 0.008674 | 0.13816 | -3e-06 | -0.000301 | -0.000301 | -18.4757
2019-04-09 15:19 | 1554837538.7074 | -420 | 0.20953 | 0.00045 | 2e-06 | 95.3713 | 95.3804 | 122.7092 | 20.6939 | 0.001954 | 0.001757 | -154.5868 | 0.2325 | 8.4014 | 0.002681 | 0.004318 | 0.004113 | 3.3e-05 | 0.001336 | 0.29446 | 1.3416 | 2.5476 | 98.49 | 20 | 47.942 | 68.2808 | 85.7583 | 63.5523 | 0.011393 | 1.0485 | 0.024329 | 23.2026 | -0.000855 | 0.13832 | -3e-06 | -0.000351 | -0.000351 | -15.8612
2019-04-09 15:19 | 1554837539.7074 | -419 | 0.20954 | 0.000451 | 2e-06 | 95.3907 | 95.2111 | 122.3548 | 20.6939 | 0.001877 | 0.001762 | -154.5746 | 0.23289 | 8.4023 | 0.002726 | 0.003617 | 0.003403 | 3.7e-05 | 0.001335 | 0.28296 | 1.4201 | 2.5476 | 98.49 | 20 | 47.95 | 66.4499 | 86.8495 | 63.6702 | 0.011395 | 1.0482 | 0.024289 | 23.1731 | 0.016853 | 0.13855 | -4e-06 | -0.000355 | -0.000355 | -15.7322
2019-04-09 15:19 | 1554837540.7085 | -418 | 0.20951 | 0.000453 | 2e-06 | 95.3916 | 95.3853 | 122.1365 | 20.6939 | 0.00193 | 0.001801 | -154.5857 | 0.23265 | 8.4024 | 0.002752 | 0.003606 | 0.003423 | 3.6e-05 | 0.001332 | 0.26073 | 1.3741 | 2.5475 | 98.49 | 19.99 | 47.968 | 67.7028 | 87.4656 | 64.6001 | 0.011392 | 1.0453 | 0.024234 | 23.1843 | 0.000592 | 0.13841 | -7e-06 | -0.000594 | -0.000594 | -9.3837
2019-04-09 15:19 | 1554837541.7087 | -417 | 0.20956 | 0.000452 | 1e-06 | 95.4854 | 95.5174 | 122.4449 | 20.6932 | 0.001937 | 0.001763 | -154.5932 | 0.23213 | 8.4009 | 0.002677 | 0.004298 | 0.004194 | 3.6e-05 | 0.001336 | 0.28512 | 1.407 | 2.5476 | 98.49 | 19.99 | 47.977 | 67.8901 | 85.6577 | 63.6915 | 0.011394 | 1.0481 | 0.024298 | 23.1824 | -0.003 | 0.13809 | -5e-06 | -0.000421 | -0.000421 | -13.2039
2019-04-09 15:19 | 1554837542.7075 | -416 | 0.20954 | 0.000454 | 2e-06 | 95.5819 | 95.705 | 122.2766 | 20.6928 | 0.001884 | 0.001805 | -154.5894 | 0.23296 | 8.4021 | 0.002678 | 0.004463 | 0.004238 | 3.5e-05 | 0.001339 | 0.29479 | 1.4007 | 2.5476 | 98.49 | 19.99 | 47.985 | 66.6103 | 85.6963 | 64.703 | 0.011396 | 1.045 | 0.024244 | 23.2012 | -0.011514 | 0.13859 | -8e-06 | -0.000689 | -0.000689 | -8.101
2019-04-09 15:19 | 1554837543.7085 | -415 | 0.20952 | 0.000443 | 2e-06 | 95.5979 | 95.7036 | 122.355 | 20.6927 | 0.001949 | 0.001772 | -154.6077 | 0.23236 | 8.4017 | 0.002725 | 0.004356 | 0.004221 | 3.4e-05 | 0.001337 | 0.28584 | 1.4021 | 2.5476 | 98.49 | 20 | 48.001 | 68.1743 | 86.8272 | 63.9049 | 0.011407 | 1.0474 | 0.024281 | 23.1812 | -0.00988 | 0.13823 | 6e-06 | 0.000283 | 0.000283 | 19.6794
2019-04-09 15:19 | 1554837544.708 | -414 | 0.20953 | 0.000445 | 2e-06 | 95.6888 | 95.5701 | 122.3984 | 20.6938 | 0.001878 | 0.001798 | -154.5901 | 0.23222 | 8.4023 | 0.002774 | 0.004254 | 0.004074 | 3.5e-05 | 0.001341 | 0.28528 | 1.3843 | 2.5477 | 98.49 | 19.99 | 48.009 | 66.452 | 88.0112 | 64.5408 | 0.011402 | 1.0455 | 0.024262 | 23.2072 | 0.011103 | 0.13815 | 4e-06 | 0.000184 | 0.000184 | 30.2727
2019-04-09 15:19 | 1554837545.7075 | -413 | 0.20958 | 0.000441 | 2e-06 | 95.729 | 95.6513 | 122.6813 | 20.6929 | 0.001893 | 0.001797 | -154.5742 | 0.23286 | 8.4013 | 0.002748 | 0.004147 | 0.003845 | 3.7e-05 | 0.001347 | 0.27743 | 1.3865 | 2.5476 | 98.49 | 20 | 48.024 | 66.8352 | 87.37 | 64.5038 | 0.011412 | 1.0456 | 0.024292 | 23.2327 | 0.007262 | 0.13853 | 9e-06 | 0.000552 | 0.000552 | 10.1188
2019-04-09 15:19 | 1554837546.708 | -412 | 0.20952 | 0.000448 | 2e-06 | 95.81 | 95.9342 | 122.4745 | 20.6931 | 0.001917 | 0.001797 | -154.5466 | 0.2322 | 8.401 | 0.002644 | 0.003835 | 0.003683 | 3.7e-05 | 0.001347 | 0.28047 | 1.3821 | 2.5475 | 98.49 | 20 | 48.032 | 67.4017 | 84.8694 | 64.5103 | 0.011414 | 1.0456 | 0.024271 | 23.2133 | -0.011588 | 0.13814 | 0 | -0.000126 | -0.000126 | -44.1568
2019-04-09 15:19 | 1554837547.7091 | -411 | 0.20956 | 0.000445 | 2e-06 | 95.9044 | 95.9493 | 122.274 | 20.6927 | 0.001874 | 0.001767 | -154.5588 | 0.23173 | 8.4014 | 0.002806 | 0.004386 | 0.004254 | 3.8e-05 | 0.001348 | 0.2688 | 1.4041 | 2.5476 | 98.49 | 20 | 48.049 | 66.355 | 88.7795 | 63.7808 | 0.011418 | 1.0478 | 0.024277 | 23.1692 | -0.004192 | 0.13786 | 3e-06 | 0.000107 | 0.000107 | 52.1148
2019-04-09 15:19 | 1554837548.7077 | -410 | 0.20952 | 0.000441 | 2e-06 | 95.9842 | 96.0567 | 122.3698 | 20.6928 | 0.00186 | 0.001792 | -154.5602 | 0.23204 | 8.4013 | 0.002643 | 0.004055 | 0.003762 | 3.6e-05 | 0.001354 | 0.29999 | 1.4067 | 2.5476 | 98.49 | 19.99 | 48.058 | 66.0328 | 84.8385 | 64.3983 | 0.011413 | 1.0459 | 0.024265 | 23.1996 | -0.006756 | 0.13804 | 8e-06 | 0.000455 | 0.000455 | 12.2285
2019-04-09 15:19 | 1554837549.7088 | -409 | 0.20952 | 0.000448 | 2e-06 | 96.0384 | 96.1248 | 122.4911 | 20.6927 | 0.001929 | 0.00178 | -154.5534 | 0.23271 | 8.3982 | 0.002609 | 0.00387 | 0.003732 | 3.9e-05 | 0.001351 | 0.28102 | 1.3914 | 2.5475 | 98.49 | 19.99 | 48.066 | 67.6822 | 84.0341 | 64.1057 | 0.011415 | 1.0468 | 0.024287 | 23.201 | -0.008049 | 0.13844 | 0 | -0.000124 | -0.000124 | -44.8164
2019-04-09 15:19 | 1554837550.7088 | -408 | 0.20957 | 0.000446 | 2e-06 | 96.1274 | 96.3138 | 122.4562 | 20.6933 | 0.001838 | 0.001776 | -154.5378 | 0.23258 | 8.4019 | 0.00271 | 0.004131 | 0.003937 | 3.7e-05 | 0.001355 | 0.27129 | 1.3847 | 2.5477 | 98.49 | 19.99 | 48.083 | 65.4894 | 86.4578 | 63.9984 | 0.011419 | 1.0472 | 0.024288 | 23.194 | -0.017324 | 0.13836 | 2e-06 | -4e-06 | -4e-06 | -1268.8014
2019-04-09 15:19 | 1554837551.7078 | -407 | 0.20948 | 0.000445 | 2e-06 | 96.3189 | 96.4883 | 122.2869 | 20.6927 | 0.001872 | 0.00178 | -154.526 | 0.23247 | 8.4033 | 0.00272 | 0.004227 | 0.004032 | 3.5e-05 | 0.001357 | 0.2636 | 1.3654 | 2.5476 | 98.49 | 19.99 | 48.089 | 66.3122 | 86.7101 | 64.0939 | 0.011421 | 1.0469 | 0.024267 | 23.1812 | -0.015715 | 0.1383 | 3e-06 | 0.000122 | 0.000122 | 45.6634
2019-04-09 15:19 | 1554837552.7069 | -406 | 0.2095 | 0.000449 | 2e-06 | 96.4475 | 96.5564 | 122.3421 | 20.692 | 0.00195 | 0.0018 | -154.5346 | 0.23264 | 8.4009 | 0.002638 | 0.004554 | 0.004374 | 3.6e-05 | 0.001353 | 0.29214 | 1.3808 | 2.5477 | 98.49 | 19.99 | 48.095 | 68.1847 | 84.7262 | 64.5857 | 0.011422 | 1.0453 | 0.024255 | 23.2034 | -0.010091 | 0.13839 | -1e-06 | -0.000208 | -0.000208 | -26.8216
2019-04-09 15:19 | 1554837553.7079 | -405 | 0.20954 | 0.000447 | 2e-06 | 96.5783 | 96.712 | 122.563 | 20.6918 | 0.001912 | 0.001792 | -154.5198 | 0.23227 | 8.3992 | 0.002648 | 0.003955 | 0.003784 | 3.8e-05 | 0.001354 | 0.29175 | 1.4219 | 2.5477 | 98.49 | 19.99 | 48.105 | 67.2679 | 84.9756 | 64.3865 | 0.011424 | 1.0459 | 0.024284 | 23.2174 | -0.012373 | 0.13818 | 1e-06 | -6.2e-05 | -6.2e-05 | -89.5831
2019-04-09 15:19 | 1554837554.707 | -404 | 0.20954 | 0.000446 | 2e-06 | 96.6873 | 96.8483 | 122.7531 | 20.6926 | 0.001922 | 0.001834 | -154.5181 | 0.23238 | 8.4014 | 0.002687 | 0.003916 | 0.003733 | 3.7e-05 | 0.001354 | 0.25155 | 1.3913 | 2.5476 | 98.49 | 19.99 | 48.109 | 67.5121 | 85.9131 | 65.4022 | 0.011425 | 1.0428 | 0.024267 | 23.2704 | -0.014881 | 0.13824 | 2e-06 | 4.6e-05 | 4.6e-05 | 121.382
2019-04-09 15:19 | 1554837555.7084 | -403 | 0.20953 | 0.000448 | 1e-06 | 96.8302 | 96.8796 | 122.7031 | 20.6913 | 0.001896 | 0.001844 | -154.5103 | 0.23206 | 8.4032 | 0.002733 | 0.003915 | 0.003799 | 3.6e-05 | 0.001354 | 0.27704 | 1.3557 | 2.5477 | 98.49 | 19.99 | 48.114 | 66.8834 | 87.0116 | 65.6405 | 0.011426 | 1.0421 | 0.024253 | 23.2738 | -0.004559 | 0.13805 | 0 | -7.2e-05 | -7.2e-05 | -77.6498
2019-04-09 15:19 | 1554837556.7073 | -402 | 0.20954 | 0.000447 | 2e-06 | 96.935 | 97.0379 | 122.8824 | 20.6926 | 0.00188 | 0.001742 | -154.5277 | 0.23249 | 8.4021 | 0.002673 | 0.003823 | 0.003631 | 3.4e-05 | 0.00136 | 0.27975 | 1.3632 | 2.5477 | 98.49 | 19.99 | 48.115 | 66.5159 | 85.5658 | 63.2004 | 0.011427 | 1.0496 | 0.024359 | 23.2068 | -0.009486 | 0.13831 | 1e-06 | 1e-06 | 1e-06 | 5687.2658
2019-04-09 15:19 | 1554837557.707 | -401 | 0.20953 | 0.000447 | 2e-06 | 96.9979 | 97.1824 | 123.0709 | 20.6916 | 0.001847 | 0.001775 | -154.5192 | 0.2331 | 8.4012 | 0.002703 | 0.004301 | 0.004027 | 3.6e-05 | 0.001356 | 0.27527 | 1.4069 | 2.5476 | 98.49 | 19.99 | 48.117 | 65.7211 | 86.286 | 63.9756 | 0.011427 | 1.0472 | 0.024349 | 23.2513 | -0.017003 | 0.13867 | 1e-06 | -2.6e-05 | -2.6e-05 | -214.5299
2019-04-09 15:19 | 1554837558.7071 | -400 | 0.20955 | 0.000443 | 2e-06 | 97.0217 | 97.1332 | 123.1932 | 20.6916 | 0.00189 | 0.001753 | -154.5109 | 0.23239 | 8.4017 | 0.002694 | 0.004269 | 0.004093 | 3.4e-05 | 0.001354 | 0.2636 | 1.3204 | 2.5475 | 98.49 | 19.99 | 48.116 | 66.7503 | 86.076 | 63.4627 | 0.011427 | 1.0488 | 0.02438 | 23.2452 | -0.010276 | 0.13825 | 6e-06 | 0.000294 | 0.000294 | 18.9594
2019-04-09 15:19 | 1554837559.7071 | -399 | 0.20952 | 0.000444 | 2e-06 | 97.1188 | 97.3913 | 123.3786 | 20.6927 | 0.001947 | 0.00177 | -154.5336 | 0.23242 | 8.4024 | 0.00266 | 0.003998 | 0.003769 | 3.7e-05 | 0.001349 | 0.28191 | 1.3433 | 2.5476 | 98.49 | 19.99 | 48.114 | 68.1151 | 85.2615 | 63.8743 | 0.011426 | 1.0475 | 0.024383 | 23.2769 | -0.025068 | 0.13827 | 5e-06 | 0.000209 | 0.000209 | 26.6997
2019-04-09 15:19 | 1554837560.7085 | -398 | 0.20952 | 0.000445 | 2e-06 | 97.2227 | 97.5479 | 123.2319 | 20.693 | 0.001901 | 0.00179 | -154.5438 | 0.23275 | 8.4029 | 0.002707 | 0.003767 | 0.003587 | 3.6e-05 | 0.00135 | 0.29839 | 1.3405 | 2.5476 | 98.49 | 19.99 | 48.113 | 67.0109 | 86.3872 | 64.3519 | 0.011426 | 1.0461 | 0.024352 | 23.2795 | -0.029871 | 0.13846 | 4e-06 | 0.000149 | 0.000149 | 37.3766
2019-04-09 15:19 | 1554837561.7072 | -397 | 0.20955 | 0.000445 | 2e-06 | 97.3219 | 97.7293 | 123.1397 | 20.6918 | 0.001864 | 0.001772 | -154.5684 | 0.23285 | 8.3998 | 0.002635 | 0.003807 | 0.003668 | 3.5e-05 | 0.001346 | 0.27433 | 1.3806 | 2.5477 | 98.49 | 19.99 | 48.111 | 66.1185 | 84.6652 | 63.9017 | 0.011426 | 1.0475 | 0.024359 | 23.2553 | -0.037367 | 0.13852 | 4e-06 | 0.000138 | 0.000138 | 40.343
2019-04-09 15:19 | 1554837562.7081 | -396 | 0.20955 | 0.000441 | 2e-06 | 97.4672 | 97.751 | 123.5854 | 20.6924 | 0.001871 | 0.00169 | -154.5625 | 0.23282 | 8.4013 | 0.002658 | 0.004404 | 0.004176 | 3.6e-05 | 0.00135 | 0.28982 | 1.339 | 2.5475 | 98.49 | 19.99 | 48.11 | 66.304 | 85.2043 | 61.9371 | 0.011425 | 1.0536 | 0.024474 | 23.2293 | -0.026005 | 0.13851 | 8e-06 | 0.000516 | 0.000516 | 10.8143
2019-04-09 15:19 | 1554837563.7076 | -395 | 0.20954 | 0.000444 | 2e-06 | 97.561 | 97.9864 | 123.7253 | 20.6927 | 0.001913 | 0.001797 | -154.5764 | 0.23211 | 8.4029 | 0.002668 | 0.003738 | 0.003617 | 3.5e-05 | 0.001354 | 0.26985 | 1.3355 | 2.5476 | 98.49 | 19.99 | 48.108 | 67.3054 | 85.448 | 64.5078 | 0.011425 | 1.0456 | 0.024395 | 23.3315 | -0.038915 | 0.13808 | 5e-06 | 0.000248 | 0.000248 | 22.4619
2019-04-09 15:19 | 1554837564.7073 | -394 | 0.20953 | 0.000445 | 1e-06 | 97.647 | 98.1227 | 123.2915 | 20.692 | 0.001957 | 0.001716 | -154.5688 | 0.23197 | 8.4008 | 0.00273 | 0.003947 | 0.003748 | 3.6e-05 | 0.001354 | 0.27776 | 1.3695 | 2.5476 | 98.49 | 19.99 | 48.108 | 68.3703 | 86.9421 | 62.5743 | 0.011425 | 1.0516 | 0.024422 | 23.2238 | -0.04346 | 0.138 | 4e-06 | 0.000224 | 0.000224 | 24.8775
2019-04-09 15:19 | 1554837565.7073 | -393 | 0.20956 | 0.000446 | 1e-06 | 97.7412 | 97.9898 | 123.0709 | 20.6919 | 0.001863 | 0.001788 | -154.5765 | 0.23198 | 8.4002 | 0.002611 | 0.003541 | 0.003512 | 4e-05 | 0.001354 | 0.27721 | 1.3666 | 2.5477 | 98.49 | 19.99 | 48.106 | 66.0959 | 84.0777 | 64.3087 | 0.011425 | 1.0462 | 0.024337 | 23.2628 | -0.022727 | 0.138 | 3e-06 | 0.000124 | 0.000124 | 44.8295
2019-04-09 15:19 | 1554837566.7087 | -392 | 0.20952 | 0.000441 | 2e-06 | 97.8227 | 98.2049 | 122.8488 | 20.6924 | 0.001914 | 0.001757 | -154.5898 | 0.23289 | 8.4025 | 0.002686 | 0.003605 | 0.003326 | 3.7e-05 | 0.001362 | 0.27815 | 1.3755 | 2.5475 | 98.49 | 19.99 | 48.103 | 67.3402 | 85.8761 | 63.5477 | 0.011424 | 1.0486 | 0.024343 | 23.2156 | -0.034876 | 0.13855 | 9e-06 | 0.000499 | 0.000499 | 11.1923
2019-04-09 15:19 | 1554837567.7083 | -391 | 0.20952 | 0.000444 | 1e-06 | 97.9761 | 98.346 | 122.7712 | 20.6924 | 0.00187 | 0.001744 | -154.5995 | 0.23245 | 8.4014 | 0.002639 | 0.004039 | 0.00393 | 3.4e-05 | 0.001362 | 0.27113 | 1.3495 | 2.5477 | 98.49 | 19.99 | 48.102 | 66.2626 | 84.7624 | 63.2478 | 0.011424 | 1.0495 | 0.024346 | 23.1979 | -0.033706 | 0.13829 | 5e-06 | 0.000288 | 0.000288 | 19.3231
2019-04-09 15:19 | 1554837568.7074 | -390 | 0.20955 | 0.000445 | 1e-06 | 98.0166 | 98.3481 | 122.5961 | 20.6928 | 0.001892 | 0.001731 | -154.5952 | 0.23233 | 8.4027 | 0.002639 | 0.003856 | 0.003657 | 3.6e-05 | 0.001359 | 0.28622 | 1.3532 | 2.5476 | 98.49 | 19.99 | 48.102 | 66.8062 | 84.7637 | 62.9335 | 0.011424 | 1.0505 | 0.02434 | 23.1706 | -0.030197 | 0.13821 | 5e-06 | 0.000331 | 0.000331 | 16.8415
2019-04-09 15:19 | 1554837569.7064 | -389 | 0.20951 | 0.000443 | 2e-06 | 98.1485 | 98.2872 | 122.3573 | 20.6927 | 0.001863 | 0.001742 | -154.5942 | 0.23185 | 8.4025 | 0.00259 | 0.003359 | 0.003196 | 3.5e-05 | 0.001359 | 0.28351 | 1.3548 | 2.5475 | 98.49 | 19.99 | 48.103 | 66.0991 | 83.5707 | 63.1848 | 0.011424 | 1.0497 | 0.024307 | 23.1566 | -0.012625 | 0.13793 | 7e-06 | 0.00039 | 0.00039 | 14.2532
2019-04-09 15:20 | 1554837570.7074 | -388 | 0.20956 | 0.000444 | 2e-06 | 98.3022 | 98.588 | 122.3505 | 20.6925 | 0.001911 | 0.001795 | -154.5908 | 0.23267 | 8.4021 | 0.002629 | 0.004239 | 0.004115 | 3.5e-05 | 0.001364 | 0.28766 | 1.4457 | 2.5477 | 98.49 | 19.99 | 48.103 | 67.253 | 84.5107 | 64.4669 | 0.011424 | 1.0457 | 0.02426 | 23.2001 | -0.025964 | 0.13841 | 5e-06 | 0.000256 | 0.000256 | 21.7497
2019-04-09 15:20 | 1554837571.7064 | -387 | 0.20952 | 0.000445 | 2e-06 | 98.36 | 98.6802 | 122.3559 | 20.6935 | 0.001918 | 0.001784 | -154.5719 | 0.23208 | 8.4022 | 0.002689 | 0.004703 | 0.004446 | 3.7e-05 | 0.001367 | 0.29446 | 1.4268 | 2.5476 | 98.49 | 19.99 | 48.102 | 67.4177 | 85.9616 | 64.2 | 0.011424 | 1.0465 | 0.02427 | 23.1914 | -0.029073 | 0.13806 | 4e-06 | 0.000228 | 0.000228 | 24.4082
2019-04-09 15:20 | 1554837572.7065 | -386 | 0.20952 | 0.000441 | 2e-06 | 98.4064 | 98.7886 | 121.9125 | 20.6919 | 0.001872 | 0.001752 | -154.5585 | 0.23215 | 8.4012 | 0.002671 | 0.003867 | 0.003738 | 3.3e-05 | 0.001371 | 0.27312 | 1.4 | 2.5477 | 98.49 | 19.99 | 48.1 | 66.3292 | 85.5136 | 63.4262 | 0.011423 | 1.0489 | 0.024254 | 23.1228 | -0.03467 | 0.1381 | 8e-06 | 0.000455 | 0.000455 | 12.2427
2019-04-09 15:20 | 1554837573.7079 | -385 | 0.20952 | 0.000445 | 2e-06 | 98.5925 | 98.9896 | 121.4809 | 20.6918 | 0.001835 | 0.001781 | -154.5571 | 0.23275 | 8.4035 | 0.00261 | 0.004029 | 0.003889 | 3.7e-05 | 0.001372 | 0.27975 | 1.3962 | 2.5477 | 98.48 | 19.99 | 48.099 | 65.4358 | 84.0574 | 64.1305 | 0.011424 | 1.0467 | 0.024186 | 23.106 | -0.035948 | 0.13846 | 3e-06 | 0.000134 | 0.000134 | 41.5363
2019-04-09 15:20 | 1554837574.7079 | -384 | 0.20952 | 0.00044 | 2e-06 | 98.6515 | 99.0932 | 121.6011 | 20.6917 | 0.001959 | 0.001741 | -154.5742 | 0.23312 | 8.4023 | 0.00264 | 0.003946 | 0.003749 | 3.6e-05 | 0.001368 | 0.30093 | 1.3738 | 2.5476 | 98.48 | 19.99 | 48.098 | 68.4065 | 84.7641 | 63.1542 | 0.011424 | 1.0498 | 0.024233 | 23.0839 | -0.039956 | 0.13869 | 1e-05 | 0.000604 | 0.000604 | 9.2556
2019-04-09 15:20 | 1554837575.7065 | -383 | 0.20952 | 0.000437 | 2e-06 | 98.8206 | 99.1266 | 121.7102 | 20.6918 | 0.001905 | 0.001775 | -154.58 | 0.23238 | 8.4014 | 0.002668 | 0.00429 | 0.004094 | 3.3e-05 | 0.001372 | 0.28119 | 1.3604 | 2.5477 | 98.48 | 19.99 | 48.097 | 67.1119 | 85.4488 | 63.982 | 0.011424 | 1.0472 | 0.024214 | 23.1227 | -0.027646 | 0.13824 | 1.4e-05 | 0.000893 | 0.000893 | 6.2394
2019-04-09 15:20 | 1554837576.7076 | -382 | 0.20955 | 0.000442 | 2e-06 | 98.9234 | 99.4852 | 121.8726 | 20.692 | 0.001905 | 0.001835 | -154.5771 | 0.23353 | 8.4005 | 0.002593 | 0.003787 | 0.003645 | 3.6e-05 | 0.001373 | 0.29319 | 1.3925 | 2.5476 | 98.48 | 19.99 | 48.094 | 67.108 | 83.6433 | 65.4359 | 0.011423 | 1.0427 | 0.024178 | 23.1879 | -0.050652 | 0.13892 | 7e-06 | 0.000337 | 0.000337 | 16.6337
2019-04-09 15:20 | 1554837577.708 | -381 | 0.20949 | 0.000438 | 2e-06 | 99.0449 | 99.6641 | 121.6867 | 20.6924 | 0.001915 | 0.001778 | -154.559 | 0.23211 | 8.4034 | 0.002629 | 0.004273 | 0.003981 | 3.5e-05 | 0.001375 | 0.27527 | 1.3668 | 2.5477 | 98.48 | 19.99 | 48.093 | 67.3457 | 84.5034 | 64.0632 | 0.011423 | 1.047 | 0.024209 | 23.1232 | -0.055748 | 0.13808 | 1.4e-05 | 0.000888 | 0.000888 | 6.2661
2019-04-09 15:20 | 1554837578.7066 | -380 | 0.20953 | 0.00044 | 2e-06 | 99.2018 | 99.8809 | 121.6597 | 20.6919 | 0.001888 | 0.00182 | -154.5807 | 0.23211 | 8.4029 | 0.002639 | 0.003052 | 0.002929 | 3.3e-05 | 0.001372 | 0.26648 | 1.3846 | 2.5476 | 98.48 | 19.99 | 48.092 | 66.7022 | 84.7472 | 65.0603 | 0.011422 | 1.0439 | 0.02417 | 23.1548 | -0.061024 | 0.13808 | 1.1e-05 | 0.0007 | 0.0007 | 7.9538
2019-04-09 15:20 | 1554837579.7078 | -379 | 0.20956 | 0.000437 | 2e-06 | 99.228 | 99.6913 | 122.0341 | 20.6923 | 0.001854 | 0.001832 | -154.5666 | 0.23236 | 8.4011 | 0.002694 | 0.004645 | 0.004441 | 3.5e-05 | 0.001372 | 0.28368 | 1.4265 | 2.5475 | 98.48 | 19.99 | 48.087 | 65.8918 | 86.0814 | 65.3674 | 0.011421 | 1.0429 | 0.024197 | 23.2009 | -0.041664 | 0.13823 | 1.4e-05 | 0.000947 | 0.000947 | 5.8812
2019-04-09 15:20 | 1554837580.7077 | -378 | 0.20956 | 0.00044 | 1e-06 | 99.2775 | 99.9542 | 122.4057 | 20.6926 | 0.001821 | 0.001724 | -154.5751 | 0.23292 | 8.401 | 0.002603 | 0.004261 | 0.004081 | 3.3e-05 | 0.001366 | 0.2896 | 1.3884 | 2.5477 | 98.48 | 19.99 | 48.084 | 65.086 | 83.8889 | 62.767 | 0.01142 | 1.051 | 0.024327 | 23.1468 | -0.060769 | 0.13856 | 1e-05 | 0.000655 | 0.000655 | 8.5306
2019-04-09 15:20 | 1554837581.7081 | -377 | 0.20954 | 0.000438 | 2e-06 | 99.4177 | 99.9715 | 122.574 | 20.6912 | 0.00186 | 0.00172 | -154.5918 | 0.23245 | 8.4006 | 0.00264 | 0.003295 | 0.003152 | 3.5e-05 | 0.001367 | 0.2912 | 1.4264 | 2.5477 | 98.48 | 19.99 | 48.077 | 66.0376 | 84.7769 | 62.6655 | 0.011419 | 1.0513 | 0.024347 | 23.1592 | -0.049681 | 0.13829 | 1.3e-05 | 0.000854 | 0.000854 | 6.5295
2019-04-09 15:20 | 1554837582.7064 | -376 | 0.2095 | 0.000444 | 2e-06 | 99.4424 | 99.9904 | 122.5569 | 20.6918 | 0.001875 | 0.001779 | -154.5863 | 0.2322 | 8.4023 | 0.002618 | 0.003975 | 0.003787 | 3.3e-05 | 0.001369 | 0.28998 | 1.3811 | 2.5477 | 98.48 | 19.99 | 48.074 | 66.3986 | 84.2455 | 64.0742 | 0.011418 | 1.0469 | 0.024295 | 23.2061 | -0.049151 | 0.13814 | 5e-06 | 0.000271 | 0.000271 | 20.5178
2019-04-09 15:20 | 1554837583.7073 | -375 | 0.20956 | 0.000439 | 2e-06 | 99.5327 | 100.0571 | 122.8805 | 20.6915 | 0.001897 | 0.001803 | -154.5929 | 0.23257 | 8.4021 | 0.002589 | 0.004391 | 0.004184 | 3.3e-05 | 0.001367 | 0.30525 | 1.3613 | 2.5476 | 98.48 | 19.99 | 48.071 | 66.9282 | 83.5493 | 64.6587 | 0.011417 | 1.0451 | 0.024306 | 23.2569 | -0.047009 | 0.13836 | 1.1e-05 | 0.000691 | 0.000691 | 8.0643
2019-04-09 15:20 | 1554837584.7073 | -374 | 0.20955 | 0.000436 | 2e-06 | 99.5872 | 100.2292 | 123.1836 | 20.6923 | 0.001909 | 0.001745 | -154.5886 | 0.23211 | 8.4034 | 0.002548 | 0.003647 | 0.00344 | 3.3e-05 | 0.001361 | 0.29247 | 1.3822 | 2.5476 | 98.48 | 19.99 | 48.064 | 67.2177 | 82.5742 | 63.2533 | 0.011416 | 1.0495 | 0.024387 | 23.237 | -0.057482 | 0.13809 | 1.5e-05 | 0.000948 | 0.000948 | 5.8694
2019-04-09 15:20 | 1554837585.7063 | -373 | 0.20951 | 0.000439 | 2e-06 | 99.7408 | 100.2144 | 123.1157 | 20.6919 | 0.001857 | 0.001785 | -154.5767 | 0.23197 | 8.4016 | 0.002568 | 0.004045 | 0.003889 | 3.4e-05 | 0.001355 | 0.28102 | 1.3588 | 2.5476 | 98.48 | 19.99 | 48.062 | 65.9665 | 83.0447 | 64.2203 | 0.011415 | 1.0465 | 0.024345 | 23.264 | -0.042371 | 0.138 | 1.1e-05 | 0.000707 | 0.000707 | 7.8614
2019-04-09 15:20 | 1554837586.7059 | -372 | 0.20952 | 0.000438 | 3e-06 | 99.8409 | 100.4626 | 122.923 | 20.6923 | 0.00184 | 0.001767 | -154.5679 | 0.23208 | 8.4001 | 0.002615 | 0.004137 | 0.00396 | 3.3e-05 | 0.001351 | 0.29214 | 1.3727 | 2.5475 | 98.48 | 19.99 | 48.06 | 65.557 | 84.1827 | 63.7921 | 0.011415 | 1.0478 | 0.024341 | 23.231 | -0.055522 | 0.13807 | 1.2e-05 | 0.000696 | 0.000696 | 7.9965
2019-04-09 15:20 | 1554837587.706 | -371 | 0.20955 | 0.00044 | 2e-06 | 99.9013 | 100.4577 | 122.8545 | 20.6922 | 0.001945 | 0.001762 | -154.5874 | 0.23161 | 8.3977 | 0.00259 | 0.003537 | 0.00336 | 3.2e-05 | 0.001347 | 0.29911 | 1.3381 | 2.5475 | 98.48 | 19.99 | 48.06 | 68.0616 | 83.5851 | 63.66 | 0.011415 | 1.0482 | 0.024339 | 23.22 | -0.049679 | 0.13778 | 1.1e-05 | 0.000691 | 0.000691 | 8.0353
2019-04-09 15:20 | 1554837588.7071 | -370 | 0.20952 | 0.00044 | 1e-06 | 99.9383 | 100.5999 | 122.9045 | 20.6914 | 0.00187 | 0.001789 | -154.5839 | 0.23218 | 8.4023 | 0.00261 | 0.003964 | 0.003707 | 3.3e-05 | 0.001351 | 0.28495 | 1.3366 | 2.5476 | 98.48 | 19.99 | 48.062 | 66.2758 | 84.0484 | 64.3315 | 0.011415 | 1.0461 | 0.02432 | 23.2479 | -0.05902 | 0.13812 | 1.1e-05 | 0.000693 | 0.000693 | 8.0271
2019-04-09 15:20 | 1554837589.7071 | -369 | 0.20955 | 0.00044 | 2e-06 | 100.1178 | 100.7011 | 122.7505 | 20.6912 | 0.001933 | 0.001709 | -154.5904 | 0.23192 | 8.4033 | 0.002669 | 0.003592 | 0.003493 | 3.3e-05 | 0.001353 | 0.27383 | 1.384 | 2.5477 | 98.48 | 19.99 | 48.066 | 67.7822 | 85.4615 | 62.3891 | 0.011416 | 1.0522 | 0.024375 | 23.1664 | -0.051955 | 0.13797 | 1.1e-05 | 0.000692 | 0.000692 | 8.04
2019-04-09 15:20 | 1554837590.7071 | -368 | 0.20955 | 0.00044 | 2e-06 | 100.1717 | 100.9306 | 122.4583 | 20.6918 | 0.001813 | 0.001806 | -154.5813 | 0.23185 | 8.4028 | 0.002638 | 0.004049 | 0.003788 | 3.1e-05 | 0.001349 | 0.30845 | 1.4105 | 2.5477 | 98.48 | 19.99 | 48.068 | 64.8916 | 84.737 | 64.7292 | 0.011417 | 1.0449 | 0.024262 | 23.2193 | -0.067507 | 0.13793 | 1e-05 | 0.000638 | 0.000638 | 8.7105
2019-04-09 15:20 | 1554837591.7062 | -367 | 0.20954 | 0.00044 | 2e-06 | 100.2134 | 100.9702 | 122.1963 | 20.6918 | 0.001917 | 0.00172 | -154.5752 | 0.23306 | 8.4002 | 0.002607 | 0.004077 | 0.003924 | 3.1e-05 | 0.001354 | 0.2891 | 1.3821 | 2.5478 | 98.48 | 19.99 | 48.074 | 67.4067 | 83.9887 | 62.6483 | 0.011418 | 1.0514 | 0.024311 | 23.1229 | -0.067292 | 0.13865 | 1e-05 | 0.000631 | 0.000631 | 8.8587
2019-04-09 15:20 | 1554837592.7071 | -366 | 0.20953 | 0.000436 | 2e-06 | 100.3111 | 100.9389 | 122.4714 | 20.6925 | 0.001865 | 0.001768 | -154.6122 | 0.23261 | 8.4023 | 0.002637 | 0.003944 | 0.00373 | 3.1e-05 | 0.00136 | 0.29175 | 1.4019 | 2.5477 | 98.48 | 19.99 | 48.078 | 66.1615 | 84.7089 | 63.8088 | 0.011419 | 1.0477 | 0.024296 | 23.1889 | -0.055803 | 0.13838 | 1.5e-05 | 0.001004 | 0.001004 | 5.5542
2019-04-09 15:20 | 1554837593.7067 | -365 | 0.20951 | 0.00044 | 1e-06 | 100.359 | 101.0404 | 122.2888 | 20.692 | 0.001858 | 0.001804 | -154.6094 | 0.23295 | 8.4011 | 0.002528 | 0.004317 | 0.004127 | 3.3e-05 | 0.001363 | 0.27992 | 1.3692 | 2.5476 | 98.48 | 19.99 | 48.081 | 65.975 | 82.0846 | 64.681 | 0.01142 | 1.045 | 0.024246 | 23.2016 | -0.060521 | 0.13858 | 1.1e-05 | 0.000711 | 0.000711 | 7.8521
2019-04-09 15:20 | 1554837594.7076 | -364 | 0.20953 | 0.000436 | 2e-06 | 100.4192 | 100.9822 | 122.1003 | 20.692 | 0.001925 | 0.001796 | -154.6006 | 0.2326 | 8.4013 | 0.00259 | 0.004362 | 0.004121 | 2.9e-05 | 0.001365 | 0.29822 | 1.3891 | 2.5474 | 98.48 | 19.99 | 48.088 | 67.6019 | 83.584 | 64.4804 | 0.011421 | 1.0457 | 0.024235 | 23.1768 | -0.050011 | 0.13837 | 1.6e-05 | 0.001063 | 0.001063 | 5.2436
2019-04-09 15:20 | 1554837595.7072 | -363 | 0.20954 | 0.00044 | 2e-06 | 100.527 | 101.1467 | 122.1341 | 20.6918 | 0.001883 | 0.00177 | -154.5926 | 0.23314 | 8.4015 | 0.002595 | 0.004024 | 0.003877 | 2.9e-05 | 0.001367 | 0.2943 | 1.4007 | 2.5475 | 98.48 | 19.99 | 48.092 | 66.5766 | 83.6962 | 63.8595 | 0.011422 | 1.0476 | 0.024261 | 23.1587 | -0.054973 | 0.13869 | 9e-06 | 0.00056 | 0.00056 | 9.9828
2019-04-09 15:20 | 1554837596.7076 | -362 | 0.20953 | 0.000442 | 1e-06 | 100.5416 | 101.2803 | 122.3132 | 20.6912 | 0.001855 | 0.00174 | -154.5977 | 0.23244 | 8.4011 | 0.002564 | 0.004142 | 0.003947 | 2.8e-05 | 0.001366 | 0.26305 | 1.4305 | 2.5477 | 98.48 | 19.99 | 48.099 | 65.9033 | 82.9536 | 63.1342 | 0.011424 | 1.0498 | 0.024305 | 23.1507 | -0.065476 | 0.13828 | 8e-06 | 0.000531 | 0.000531 | 10.4924
2019-04-09 15:20 | 1554837597.7063 | -361 | 0.20943 | 0.000442 | 1e-06 | 100.5523 | 101.2995 | 122.4731 | 20.6921 | 0.001856 | 0.001769 | -154.5682 | 0.23243 | 8.3997 | 0.002606 | 0.003853 | 0.003694 | 3.4e-05 | 0.001365 | 0.28766 | 1.4357 | 2.5477 | 98.48 | 19.99 | 48.1 | 65.9254 | 83.9524 | 63.831 | 0.011424 | 1.0477 | 0.024295 | 23.1898 | -0.066224 | 0.13828 | 8e-06 | 0.000531 | 0.000531 | 10.4961
2019-04-09 15:20 | 1554837598.7064 | -360 | 0.20684 | 0.000442 | 2e-06 | 100.65 | 101.2854 | 122.3273 | 20.6918 | 0.001873 | 0.001792 | -154.5668 | 0.23202 | 8.4017 | 0.002591 | 0.004805 | 0.004546 | 3e-05 | 0.001367 | 0.28368 | 1.4393 | 2.5477 | 98.48 | 19.99 | 48.102 | 66.3448 | 83.5893 | 64.3873 | 0.011425 | 1.0459 | 0.024261 | 23.1951 | -0.056293 | 0.13803 | 7e-06 | 0.000428 | 0.000428 | 13.0092
2019-04-09 15:20 | 1554837599.7064 | -359 | 0.20638 | 0.000441 | 1e-06 | 100.6419 | 101.315 | 122.4768 | 20.692 | 0.001857 | 0.001852 | -154.5573 | 0.23253 | 8.402 | 0.002588 | 0.004184 | 0.004054 | 2.7e-05 | 0.001359 | 0.26322 | 1.3656 | 2.5475 | 98.48 | 19.99 | 48.102 | 65.9691 | 83.527 | 65.8339 | 0.011425 | 1.0415 | 0.024224 | 23.259 | -0.059621 | 0.13833 | 1e-05 | 0.000624 | 0.000624 | 8.929
2019-04-09 15:20 | 1554837600.7064 | -358 | 0.20636 | 0.00044 | 2e-06 | 100.6756 | 101.4868 | 123.0017 | 20.6911 | 0.001795 | 0.001756 | -154.5842 | 0.23275 | 8.4029 | 0.002611 | 0.003734 | 0.003502 | 2.8e-05 | 0.001359 | 0.26432 | 1.4282 | 2.5477 | 98.48 | 19.99 | 48.101 | 64.4662 | 84.0745 | 63.5238 | 0.011425 | 1.0486 | 0.024359 | 23.2292 | -0.07178 | 0.13846 | 1e-05 | 0.000594 | 0.000594 | 9.4011
2019-04-09 15:20 | 1554837601.7068 | -357 | 0.20635 | 0.000442 | 2e-06 | 100.7956 | 101.3388 | 122.8667 | 20.6915 | 0.001833 | 0.001776 | -154.5995 | 0.23273 | 8.4033 | 0.002609 | 0.003999 | 0.003837 | 2.9e-05 | 0.001361 | 0.27201 | 1.4151 | 2.5478 | 98.48 | 19.99 | 48.1 | 65.3674 | 84.0251 | 64.0012 | 0.011424 | 1.0471 | 0.024328 | 23.2329 | -0.048072 | 0.13845 | 8e-06 | 0.000433 | 0.000433 | 12.8884
2019-04-09 15:20 | 1554837602.7067 | -356 | 0.20796 | 0.00044 | 2e-06 | 100.8245 | 101.5948 | 122.6343 | 20.6916 | 0.001942 | 0.001735 | -154.5804 | 0.23276 | 8.4018 | 0.002645 | 0.003037 | 0.002858 | 3.2e-05 | 0.001366 | 0.26825 | 1.4168 | 2.5476 | 98.48 | 19.99 | 48.098 | 68.0126 | 84.8874 | 63.0268 | 0.011424 | 1.0502 | 0.02434 | 23.1774 | -0.06808 | 0.13847 | 1e-05 | 0.000646 | 0.000646 | 8.6326
2019-04-09 15:20 | 1554837603.7069 | -355 | 0.2095 | 0.000438 | 1e-06 | 100.8179 | 101.5208 | 122.8155 | 20.6926 | 0.001848 | 0.001745 | -154.5887 | 0.23241 | 8.4043 | 0.002518 | 0.003616 | 0.0035 | 2.8e-05 | 0.001369 | 0.30143 | 1.4333 | 2.5478 | 98.48 | 19.99 | 48.096 | 65.7495 | 81.8427 | 63.2636 | 0.011423 | 1.0494 | 0.02435 | 23.2027 | -0.062145 | 0.13826 | 1.3e-05 | 0.000865 | 0.000865 | 6.4434
2019-04-09 15:20 | 1554837604.7055 | -354 | 0.2095 | 0.000443 | 2e-06 | 100.8102 | 101.398 | 122.8003 | 20.6924 | 0.001854 | 0.001782 | -154.5973 | 0.23283 | 8.4037 | 0.002493 | 0.003172 | 0.002972 | 3.1e-05 | 0.001361 | 0.27024 | 1.4035 | 2.5475 | 98.48 | 19.99 | 48.092 | 65.8943 | 81.2306 | 64.1638 | 0.011422 | 1.0466 | 0.024316 | 23.2322 | -0.052 | 0.13851 | 6e-06 | 0.00035 | 0.00035 | 15.9302
2019-04-09 15:20 | 1554837605.7065 | -353 | 0.20951 | 0.000441 | 1e-06 | 100.8382 | 101.6891 | 122.9136 | 20.6908 | 0.001888 | 0.001753 | -154.6078 | 0.23291 | 8.399 | 0.002522 | 0.004157 | 0.003989 | 3e-05 | 0.001359 | 0.26825 | 1.3807 | 2.5474 | 98.48 | 19.99 | 48.09 | 66.6966 | 81.9472 | 63.4477 | 0.011422 | 1.0489 | 0.024353 | 23.2183 | -0.075163 | 0.13856 | 1e-05 | 0.000617 | 0.000617 | 9.0553
2019-04-09 15:20 | 1554837606.7055 | -352 | 0.20948 | 0.000437 | 2e-06 | 100.9418 | 101.6031 | 122.9725 | 20.6911 | 0.001914 | 0.001751 | -154.6143 | 0.23212 | 8.4018 | 0.002601 | 0.004339 | 0.004165 | 3.3e-05 | 0.001353 | 0.27721 | 1.3648 | 2.5474 | 98.48 | 19.99 | 48.09 | 67.3264 | 83.8328 | 63.4013 | 0.011422 | 1.049 | 0.02436 | 23.2222 | -0.058405 | 0.13809 | 1.4e-05 | 0.000915 | 0.000915 | 6.0803
2019-04-09 15:20 | 1554837607.7066 | -351 | 0.2095 | 0.00044 | 2e-06 | 100.9341 | 101.6578 | 123.2446 | 20.6901 | 0.0019 | 0.00178 | -154.621 | 0.23288 | 8.4008 | 0.002599 | 0.003891 | 0.003793 | 3.1e-05 | 0.001358 | 0.27959 | 1.4043 | 2.5477 | 98.48 | 19.99 | 48.09 | 66.9854 | 83.7888 | 64.0961 | 0.011422 | 1.0468 | 0.024362 | 23.2719 | -0.063907 | 0.13854 | 1e-05 | 0.000575 | 0.000575 | 9.702
2019-04-09 15:20 | 1554837608.7057 | -350 | 0.20951 | 0.000437 | 2e-06 | 100.9603 | 101.8774 | 123.481 | 20.6915 | 0.001934 | 0.001753 | -154.6233 | 0.23255 | 8.4039 | 0.002633 | 0.003658 | 0.003449 | 3e-05 | 0.001363 | 0.29175 | 1.37 | 2.5476 | 98.48 | 19.99 | 48.09 | 67.8137 | 84.612 | 63.4595 | 0.011422 | 1.0488 | 0.024409 | 23.2722 | -0.080883 | 0.13834 | 1.4e-05 | 0.000917 | 0.000917 | 6.0782
2019-04-09 15:20 | 1554837609.7071 | -349 | 0.20953 | 0.000438 | 2e-06 | 100.9683 | 101.6747 | 123.396 | 20.6915 | 0.001848 | 0.001763 | -154.6283 | 0.23249 | 8.4017 | 0.002628 | 0.003794 | 0.003652 | 3.5e-05 | 0.001365 | 0.27865 | 1.3365 | 2.5476 | 98.48 | 19.99 | 48.089 | 65.7287 | 84.4855 | 63.6874 | 0.011422 | 1.0481 | 0.024392 | 23.2721 | -0.062356 | 0.13831 | 1.2e-05 | 0.000765 | 0.000765 | 7.2896
2019-04-09 15:20 | 1554837610.7067 | -348 | 0.20949 | 0.000436 | 2e-06 | 101.0014 | 101.8227 | 122.9366 | 20.6916 | 0.001923 | 0.00171 | -154.6193 | 0.23241 | 8.4053 | 0.002627 | 0.004904 | 0.004761 | 3e-05 | 0.001372 | 0.27599 | 1.3532 | 2.5474 | 98.48 | 19.99 | 48.09 | 67.5344 | 84.4653 | 62.4121 | 0.011422 | 1.0521 | 0.024393 | 23.1847 | -0.072437 | 0.13826 | 1.5e-05 | 0.001012 | 0.001012 | 5.5034
2019-04-09 15:20 | 1554837611.7071 | -347 | 0.20952 | 0.000438 | 2e-06 | 101.0702 | 101.8746 | 122.8824 | 20.6909 | 0.001866 | 0.001802 | -154.5966 | 0.23243 | 8.4014 | 0.002591 | 0.003427 | 0.003249 | 3.1e-05 | 0.001373 | 0.28224 | 1.3466 | 2.5476 | 98.48 | 19.99 | 48.091 | 66.1731 | 83.6033 | 64.6246 | 0.011422 | 1.0452 | 0.024307 | 23.2559 | -0.070911 | 0.13828 | 1.3e-05 | 0.000818 | 0.000818 | 6.8158
2019-04-09 15:20 | 1554837612.7057 | -346 | 0.20951 | 0.000437 | 2e-06 | 101.1294 | 102.0281 | 123.289 | 20.6908 | 0.001903 | 0.001729 | -154.5836 | 0.23323 | 8.4005 | 0.002568 | 0.003886 | 0.003808 | 3.2e-05 | 0.001371 | 0.29031 | 1.3426 | 2.5475 | 98.48 | 19.98 | 48.094 | 67.0679 | 83.0454 | 62.8753 | 0.011416 | 1.0507 | 0.024411 | 23.2339 | -0.079135 | 0.13875 | 1.4e-05 | 0.00087 | 0.00087 | 6.4247
2019-04-09 15:20 | 1554837613.7072 | -345 | 0.2095 | 0.000435 | 1e-06 | 101.1473 | 101.9483 | 123.292 | 20.6909 | 0.001835 | 0.001771 | -154.5918 | 0.23238 | 8.4009 | 0.002565 | 0.00384 | 0.00368 | 3.3e-05 | 0.001372 | 0.28136 | 1.377 | 2.5476 | 98.48 | 19.99 | 48.096 | 65.4332 | 82.9799 | 63.8876 | 0.011423 | 1.0475 | 0.024374 | 23.2692 | -0.070552 | 0.13824 | 1.7e-05 | 0.001149 | 0.001149 | 4.8475
2019-04-09 15:20 | 1554837614.7068 | -344 | 0.20953 | 0.000435 | 2e-06 | 101.1467 | 102.0536 | 123.1124 | 20.6903 | 0.00188 | 0.001789 | -154.5847 | 0.2336 | 8.4017 | 0.00258 | 0.003747 | 0.003565 | 3.4e-05 | 0.001367 | 0.29407 | 1.3762 | 2.5477 | 98.48 | 19.98 | 48.097 | 66.5136 | 83.325 | 64.3257 | 0.011416 | 1.0461 | 0.024341 | 23.2673 | -0.07984 | 0.13897 | 1.6e-05 | 0.001071 | 0.001071 | 5.2274
2019-04-09 15:20 | 1554837615.7058 | -343 | 0.2095 | 0.000436 | 1e-06 | 101.1586 | 101.8798 | 122.9951 | 20.6902 | 0.00186 | 0.001713 | -154.5878 | 0.23298 | 8.4012 | 0.002569 | 0.004123 | 0.003905 | 3.1e-05 | 0.001373 | 0.27831 | 1.415 | 2.5476 | 98.48 | 19.98 | 48.103 | 66.025 | 83.0782 | 62.4998 | 0.011418 | 1.0518 | 0.024395 | 23.1932 | -0.063541 | 0.1386 | 1.5e-05 | 0.00103 | 0.00103 | 5.4212
2019-04-09 15:20 | 1554837616.705 | -342 | 0.20953 | 0.000437 | 2e-06 | 101.2253 | 102.0608 | 123.0707 | 20.6907 | 0.001835 | 0.00175 | -154.5872 | 0.23269 | 8.3996 | 0.002575 | 0.003971 | 0.003868 | 3e-05 | 0.00137 | 0.27992 | 1.3771 | 2.5475 | 98.48 | 19.99 | 48.105 | 65.4222 | 83.2107 | 63.3909 | 0.011425 | 1.049 | 0.02437 | 23.2311 | -0.073524 | 0.13843 | 1.4e-05 | 0.000926 | 0.000926 | 6.025
2019-04-09 15:20 | 1554837617.7049 | -341 | 0.20955 | 0.000435 | 1e-06 | 101.2423 | 102.0532 | 122.7773 | 20.6907 | 0.001876 | 0.001766 | -154.6027 | 0.23276 | 8.4013 | 0.002671 | 0.003407 | 0.00325 | 3.6e-05 | 0.00136 | 0.28512 | 1.3763 | 2.5476 | 98.48 | 19.99 | 48.109 | 66.418 | 85.5242 | 63.7749 | 0.011426 | 1.0478 | 0.024327 | 23.2167 | -0.07136 | 0.13847 | 1.7e-05 | 0.001132 | 0.001132 | 4.9318
2019-04-09 15:20 | 1554837618.7064 | -340 | 0.20955 | 0.000436 | 1e-06 | 101.2259 | 102.1303 | 122.5972 | 20.6905 | 0.001918 | 0.001732 | -154.5962 | 0.2329 | 8.4012 | 0.002682 | 0.004062 | 0.003874 | 3.3e-05 | 0.001356 | 0.27041 | 1.3843 | 2.5476 | 98.48 | 19.98 | 48.111 | 67.4141 | 85.7956 | 62.9554 | 0.01142 | 1.0504 | 0.024339 | 23.1714 | -0.079561 | 0.13855 | 1.6e-05 | 0.001101 | 0.001101 | 5.0695
2019-04-09 15:20 | 1554837619.706 | -339 | 0.20953 | 0.000437 | 2e-06 | 101.2625 | 101.9675 | 122.7317 | 20.6906 | 0.001915 | 0.001733 | -154.5965 | 0.23317 | 8.4021 | 0.00265 | 0.004429 | 0.004202 | 3.4e-05 | 0.001354 | 0.28174 | 1.4011 | 2.5476 | 98.48 | 19.98 | 48.113 | 67.3401 | 85.0082 | 62.9712 | 0.01142 | 1.0504 | 0.024352 | 23.1847 | -0.062058 | 0.13872 | 1.4e-05 | 0.000926 | 0.000926 | 6.0357
2019-04-09 15:20 | 1554837620.7052 | -338 | 0.20952 | 0.000435 | 1e-06 | 101.2625 | 102.1661 | 122.9244 | 20.6912 | 0.001901 | 0.001763 | -154.5941 | 0.23359 | 8.4001 | 0.002638 | 0.003749 | 0.003546 | 3.4e-05 | 0.001354 | 0.30182 | 1.3879 | 2.5478 | 98.48 | 19.98 | 48.119 | 67.0262 | 84.7271 | 63.6941 | 0.011422 | 1.0481 | 0.024345 | 23.2278 | -0.07946 | 0.13896 | 1.7e-05 | 0.001133 | 0.001133 | 4.9449
2019-04-09 15:20 | 1554837621.7061 | -337 | 0.20953 | 0.000434 | 1e-06 | 101.2485 | 101.9744 | 123.0272 | 20.6909 | 0.001873 | 0.00175 | -154.6058 | 0.23292 | 8.3998 | 0.002681 | 0.004401 | 0.004198 | 3.2e-05 | 0.001351 | 0.26322 | 1.404 | 2.5477 | 98.48 | 19.98 | 48.122 | 66.3455 | 85.7505 | 63.3722 | 0.011422 | 1.0491 | 0.024367 | 23.2264 | -0.063898 | 0.13856 | 1.9e-05 | 0.001273 | 0.001273 | 4.3878
2019-04-09 15:20 | 1554837622.706 | -336 | 0.20951 | 0.000438 | 2e-06 | 101.2158 | 102.1524 | 122.6953 | 20.691 | 0.001875 | 0.001772 | -154.6113 | 0.23258 | 8.4007 | 0.002651 | 0.004143 | 0.003953 | 3.5e-05 | 0.001354 | 0.27328 | 1.3784 | 2.5477 | 98.48 | 19.98 | 48.125 | 66.3808 | 85.043 | 63.9033 | 0.011423 | 1.0474 | 0.024315 | 23.2134 | -0.082385 | 0.13836 | 1.2e-05 | 0.000749 | 0.000749 | 7.4435
2019-04-09 15:20 | 1554837623.7052 | -335 | 0.20952 | 0.000438 | 1e-06 | 101.2205 | 102.1627 | 122.7925 | 20.6907 | 0.001883 | 0.001793 | -154.5959 | 0.23149 | 8.4005 | 0.002606 | 0.003635 | 0.003585 | 3.3e-05 | 0.001354 | 0.29856 | 1.3573 | 2.5474 | 98.48 | 19.98 | 48.128 | 66.5926 | 83.9479 | 64.423 | 0.011424 | 1.0458 | 0.024306 | 23.2404 | -0.082868 | 0.13771 | 1.2e-05 | 0.000815 | 0.000815 | 6.8136
2019-04-09 15:20 | 1554837624.7065 | -334 | 0.20951 | 0.000436 | 2e-06 | 101.2661 | 102.2897 | 123.0423 | 20.6901 | 0.001881 | 0.001769 | -154.6078 | 0.23286 | 8.4022 | 0.002595 | 0.004139 | 0.003951 | 3.5e-05 | 0.001356 | 0.28816 | 1.398 | 2.5475 | 98.48 | 19.98 | 48.135 | 66.5437 | 83.7023 | 63.8367 | 0.011425 | 1.0477 | 0.024352 | 23.2438 | -0.089954 | 0.13853 | 1.5e-05 | 0.000966 | 0.000966 | 5.7785
2019-04-09 15:20 | 1554837625.7066 | -333 | 0.2095 | 0.000437 | 2e-06 | 101.2631 | 102.1269 | 123.0275 | 20.6904 | 0.001881 | 0.001741 | -154.61 | 0.23204 | 8.4017 | 0.002621 | 0.004279 | 0.004097 | 3.6e-05 | 0.001353 | 0.28567 | 1.3675 | 2.5475 | 98.48 | 19.98 | 48.139 | 66.5213 | 84.3196 | 63.1654 | 0.011426 | 1.0497 | 0.024374 | 23.2193 | -0.075971 | 0.13804 | 1.3e-05 | 0.000859 | 0.000859 | 6.4725
2019-04-09 15:20 | 1554837626.7056 | -332 | 0.20951 | 0.000439 | 2e-06 | 101.2744 | 102.1678 | 122.8625 | 20.6908 | 0.00186 | 0.001743 | -154.6152 | 0.23213 | 8.4027 | 0.002655 | 0.004566 | 0.004495 | 3.1e-05 | 0.001357 | 0.29446 | 1.3833 | 2.5476 | 98.48 | 19.98 | 48.146 | 66.0308 | 85.1286 | 63.2159 | 0.011428 | 1.0496 | 0.024356 | 23.2055 | -0.078558 | 0.1381 | 1.1e-05 | 0.000708 | 0.000708 | 7.8618
2019-04-09 15:20 | 1554837627.7062 | -331 | 0.20949 | 0.000436 | 2e-06 | 101.3158 | 102.088 | 123.1411 | 20.6915 | 0.001874 | 0.001783 | -154.6069 | 0.2326 | 8.4023 | 0.002658 | 0.004037 | 0.003936 | 3.3e-05 | 0.001357 | 0.28982 | 1.3679 | 2.5475 | 98.48 | 19.98 | 48.149 | 66.3563 | 85.2121 | 64.1832 | 0.011429 | 1.0466 | 0.024349 | 23.2651 | -0.06791 | 0.13837 | 1.6e-05 | 0.001017 | 0.001017 | 5.4858
2019-04-09 15:20 | 1554837628.7053 | -330 | 0.20951 | 0.000436 | 2e-06 | 101.2381 | 102.0694 | 123.2606 | 20.6904 | 0.001894 | 0.001817 | -154.612 | 0.23184 | 8.4011 | 0.002652 | 0.004015 | 0.003812 | 3.1e-05 | 0.001358 | 0.29839 | 1.39 | 2.5477 | 98.48 | 19.98 | 48.154 | 66.8554 | 85.0664 | 65.0018 | 0.01143 | 1.044 | 0.024331 | 23.3046 | -0.073146 | 0.13792 | 1.6e-05 | 0.00103 | 0.00103 | 5.3967
2019-04-09 15:20 | 1554837629.7057 | -329 | 0.20951 | 0.000433 | 1e-06 | 101.2881 | 102.0684 | 122.9272 | 20.6904 | 0.00187 | 0.001741 | -154.594 | 0.23191 | 8.4042 | 0.002659 | 0.004329 | 0.004146 | 3.2e-05 | 0.001354 | 0.28047 | 1.3792 | 2.5476 | 98.48 | 19.98 | 48.161 | 66.2631 | 85.2225 | 63.1554 | 0.011432 | 1.0498 | 0.024365 | 23.2095 | -0.068637 | 0.13796 | 2e-05 | 0.001362 | 0.001362 | 4.0821
2019-04-09 15:20 | 1554837630.7053 | -328 | 0.20953 | 0.000433 | 2e-06 | 101.2721 | 101.9975 | 122.993 | 20.6903 | 0.001873 | 0.001779 | -154.5812 | 0.23155 | 8.4045 | 0.002669 | 0.003939 | 0.003842 | 3.5e-05 | 0.001353 | 0.24491 | 1.417 | 2.5476 | 98.48 | 19.98 | 48.165 | 66.3411 | 85.462 | 64.0814 | 0.011433 | 1.0469 | 0.024338 | 23.2476 | -0.063842 | 0.13775 | 1.9e-05 | 0.001253 | 0.001253 | 4.4296
2019-04-09 15:21 | 1554837631.7053 | -327 | 0.20946 | 0.000438 | 2e-06 | 101.2765 | 102.0901 | 123.1467 | 20.6895 | 0.001887 | 0.001737 | -154.5767 | 0.23266 | 8.4006 | 0.00264 | 0.002963 | 0.002804 | 3.2e-05 | 0.001356 | 0.28495 | 1.4296 | 2.5478 | 98.48 | 19.98 | 48.17 | 66.6835 | 84.7745 | 63.0692 | 0.011434 | 1.05 | 0.02439 | 23.2272 | -0.071562 | 0.13841 | 1.2e-05 | 0.000784 | 0.000784 | 7.1141
2019-04-09 15:21 | 1554837632.7044 | -326 | 0.2095 | 0.000437 | 2e-06 | 101.2917 | 102.1128 | 123.0582 | 20.6909 | 0.001851 | 0.001792 | -154.571 | 0.23243 | 8.402 | 0.002611 | 0.004543 | 0.004316 | 3.3e-05 | 0.001357 | 0.29142 | 1.3909 | 2.5476 | 98.48 | 19.98 | 48.178 | 65.8171 | 84.0666 | 64.3926 | 0.011436 | 1.0459 | 0.024333 | 23.2645 | -0.072211 | 0.13827 | 1.4e-05 | 0.000955 | 0.000955 | 5.8325
2019-04-09 15:21 | 1554837633.7044 | -325 | 0.20951 | 0.000438 | 1e-06 | 101.2864 | 102.0082 | 122.8988 | 20.6905 | 0.001874 | 0.001785 | -154.5726 | 0.23283 | 8.4023 | 0.002703 | 0.003962 | 0.00384 | 3.5e-05 | 0.00136 | 0.29015 | 1.3688 | 2.5475 | 98.48 | 19.98 | 48.183 | 66.3584 | 86.2967 | 64.2347 | 0.011437 | 1.0464 | 0.024323 | 23.244 | -0.063515 | 0.13851 | 1.3e-05 | 0.000858 | 0.000858 | 6.5078
2019-04-09 15:21 | 1554837634.7047 | -324 | 0.20948 | 0.000436 | 1e-06 | 101.2718 | 102.1434 | 122.7364 | 20.6903 | 0.001866 | 0.001737 | -154.5529 | 0.23311 | 8.4027 | 0.002622 | 0.00477 | 0.004567 | 3.2e-05 | 0.001356 | 0.27776 | 1.3368 | 2.5475 | 98.48 | 19.98 | 48.193 | 66.1663 | 84.3508 | 63.063 | 0.011439 | 1.0501 | 0.024349 | 23.1883 | -0.076656 | 0.13868 | 1.6e-05 | 0.001055 | 0.001055 | 5.2973
2019-04-09 15:21 | 1554837635.7046 | -323 | 0.20949 | 0.000434 | 2e-06 | 101.3697 | 102.1971 | 122.8449 | 20.6908 | 0.001924 | 0.001773 | -154.5592 | 0.23217 | 8.4025 | 0.00272 | 0.004303 | 0.004244 | 3.7e-05 | 0.001344 | 0.27184 | 1.3638 | 2.5475 | 98.48 | 19.98 | 48.198 | 67.5704 | 86.6887 | 63.939 | 0.01144 | 1.0473 | 0.024328 | 23.2287 | -0.072707 | 0.13812 | 1.8e-05 | 0.001231 | 0.001231 | 4.5205
2019-04-09 15:21 | 1554837636.7059 | -322 | 0.2095 | 0.00043 | 2e-06 | 101.267 | 102.1857 | 123.1258 | 20.6913 | 0.001939 | 0.001773 | -154.5438 | 0.23281 | 8.4002 | 0.002726 | 0.004295 | 0.004149 | 3.7e-05 | 0.001346 | 0.27367 | 1.3457 | 2.5476 | 98.48 | 19.98 | 48.204 | 67.9282 | 86.8502 | 63.9349 | 0.011442 | 1.0473 | 0.024356 | 23.2551 | -0.080782 | 0.1385 | 2.3e-05 | 0.001551 | 0.001551 | 3.598
2019-04-09 15:21 | 1554837637.7049 | -321 | 0.20949 | 0.000435 | 2e-06 | 101.2873 | 102.0901 | 122.6953 | 20.691 | 0.001896 | 0.001772 | -154.5468 | 0.23275 | 8.4016 | 0.00272 | 0.004381 | 0.004133 | 3.4e-05 | 0.001344 | 0.29927 | 1.3637 | 2.5477 | 98.48 | 19.98 | 48.214 | 66.8894 | 86.6898 | 63.9001 | 0.011444 | 1.0475 | 0.024315 | 23.2132 | -0.070615 | 0.13846 | 1.7e-05 | 0.001104 | 0.001104 | 5.054
2019-04-09 15:21 | 1554837638.7046 | -320 | 0.20949 | 0.000435 | 2e-06 | 101.3146 | 102.1341 | 122.7963 | 20.69 | 0.001926 | 0.001771 | -154.5568 | 0.23259 | 8.4011 | 0.002669 | 0.004174 | 0.004099 | 3.7e-05 | 0.001339 | 0.27704 | 1.4078 | 2.5478 | 98.48 | 19.98 | 48.22 | 67.6256 | 85.4703 | 63.8773 | 0.011446 | 1.0475 | 0.024326 | 23.222 | -0.072055 | 0.13837 | 1.7e-05 | 0.001138 | 0.001138 | 4.9002
2019-04-09 15:21 | 1554837639.7056 | -319 | 0.20946 | 0.000435 | 2e-06 | 101.2995 | 102.0656 | 123.3258 | 20.6906 | 0.001888 | 0.001769 | -154.5777 | 0.23308 | 8.4013 | 0.002688 | 0.004055 | 0.00388 | 3.9e-05 | 0.001336 | 0.28838 | 1.4241 | 2.5477 | 98.48 | 19.98 | 48.231 | 66.7026 | 85.9198 | 63.8483 | 0.011448 | 1.0476 | 0.024379 | 23.271 | -0.067396 | 0.13866 | 1.6e-05 | 0.001092 | 0.001092 | 5.1187
2019-04-09 15:21 | 1554837640.7047 | -318 | 0.20951 | 0.00043 | 1e-06 | 101.2968 | 101.9355 | 123.0803 | 20.691 | 0.001897 | 0.001727 | -154.5893 | 0.23267 | 8.4026 | 0.002658 | 0.004121 | 0.003887 | 3.7e-05 | 0.001331 | 0.29358 | 1.3614 | 2.5475 | 98.48 | 19.98 | 48.235 | 66.92 | 85.2061 | 62.8233 | 0.011449 | 1.0508 | 0.024392 | 23.2125 | -0.056225 | 0.13842 | 2.3e-05 | 0.001611 | 0.001611 | 3.4627
2019-04-09 15:21 | 1554837641.7047 | -317 | 0.20945 | 0.000438 | 1e-06 | 101.2158 | 102.1 | 123.0345 | 20.6902 | 0.001866 | 0.001731 | -154.5971 | 0.23254 | 8.3992 | 0.002706 | 0.004176 | 0.004093 | 3.6e-05 | 0.001334 | 0.2891 | 1.4253 | 2.5476 | 98.48 | 19.98 | 48.239 | 66.1654 | 86.3643 | 62.9146 | 0.01145 | 1.0505 | 0.024384 | 23.2113 | -0.077803 | 0.13834 | 1.4e-05 | 0.000921 | 0.000921 | 6.0565
2019-04-09 15:21 | 1554837642.7057 | -316 | 0.20949 | 0.000436 | 1e-06 | 101.3045 | 101.9627 | 123.303 | 20.6903 | 0.001945 | 0.001779 | -154.6034 | 0.23242 | 8.4024 | 0.002616 | 0.004127 | 0.003865 | 3.5e-05 | 0.001333 | 0.28407 | 1.4514 | 2.5477 | 98.48 | 19.98 | 48.247 | 68.0823 | 84.194 | 64.0713 | 0.011452 | 1.0469 | 0.024369 | 23.2766 | -0.057927 | 0.13827 | 1.5e-05 | 0.001047 | 0.001047 | 5.3219
2019-04-09 15:21 | 1554837643.7051 | -315 | 0.20949 | 0.000437 | 2e-06 | 101.2533 | 102.1317 | 123.4619 | 20.6904 | 0.001881 | 0.001779 | -154.6009 | 0.23303 | 8.3994 | 0.002634 | 0.004148 | 0.004027 | 3.8e-05 | 0.001332 | 0.28174 | 1.3927 | 2.5476 | 98.48 | 19.98 | 48.251 | 66.5346 | 84.6318 | 64.0784 | 0.011453 | 1.0469 | 0.024384 | 23.2918 | -0.077261 | 0.13863 | 1.4e-05 | 0.000886 | 0.000886 | 6.3092
2019-04-09 15:21 | 1554837644.7047 | -314 | 0.20946 | 0.000432 | 1e-06 | 101.2694 | 102.0725 | 123.622 | 20.691 | 0.001864 | 0.001714 | -154.6174 | 0.23295 | 8.4005 | 0.002623 | 0.003655 | 0.003444 | 3.5e-05 | 0.001336 | 0.30702 | 1.3826 | 2.5476 | 98.48 | 19.98 | 48.256 | 66.1365 | 84.364 | 62.5045 | 0.011454 | 1.0518 | 0.024457 | 23.2524 | -0.070656 | 0.13858 | 2.1e-05 | 0.001487 | 0.001487 | 3.7566
2019-04-09 15:21 | 1554837645.7039 | -313 | 0.20951 | 0.000437 | 1e-06 | 101.2089 | 102.078 | 123.653 | 20.6914 | 0.001884 | 0.00174 | -154.6148 | 0.23273 | 8.4029 | 0.002645 | 0.005191 | 0.004966 | 3.7e-05 | 0.001333 | 0.27616 | 1.3746 | 2.5475 | 98.48 | 19.98 | 48.258 | 66.6137 | 84.8865 | 63.1445 | 0.011455 | 1.0498 | 0.024437 | 23.2775 | -0.076479 | 0.13845 | 1.4e-05 | 0.000928 | 0.000928 | 6.0106
2019-04-09 15:21 | 1554837646.7049 | -312 | 0.2095 | 0.000438 | 2e-06 | 101.231 | 102.0271 | 123.6511 | 20.6898 | 0.001925 | 0.001749 | -154.6274 | 0.23269 | 8.4001 | 0.002687 | 0.003896 | 0.003773 | 3.6e-05 | 0.001342 | 0.27256 | 1.3501 | 2.5475 | 98.48 | 19.98 | 48.258 | 67.6022 | 85.8955 | 63.3493 | 0.011455 | 1.0492 | 0.024429 | 23.2844 | -0.070067 | 0.13843 | 1.2e-05 | 0.000739 | 0.000739 | 7.5475
2019-04-09 15:21 | 1554837647.7039 | -311 | 0.20949 | 0.000442 | 2e-06 | 101.1568 | 101.9844 | 123.9225 | 20.6908 | 0.001879 | 0.001781 | -154.6331 | 0.23315 | 8.3988 | 0.002623 | 0.003822 | 0.00375 | 3.8e-05 | 0.001333 | 0.29999 | 1.3789 | 2.5476 | 98.48 | 19.98 | 48.257 | 66.4852 | 84.3686 | 64.1247 | 0.011454 | 1.0468 | 0.024428 | 23.3368 | -0.072879 | 0.1387 | 8e-06 | 0.00045 | 0.00045 | 12.411
2019-04-09 15:21 | 1554837648.7039 | -310 | 0.20947 | 0.000436 | 2e-06 | 101.2146 | 101.7056 | 123.9687 | 20.6901 | 0.001854 | 0.00172 | -154.6121 | 0.23267 | 8.4021 | 0.00269 | 0.003374 | 0.003201 | 3.3e-05 | 0.001332 | 0.27201 | 1.3954 | 2.5477 | 98.48 | 19.98 | 48.256 | 65.8866 | 85.9833 | 62.6637 | 0.011454 | 1.0513 | 0.024486 | 23.2906 | -0.04329 | 0.13842 | 1.5e-05 | 0.000964 | 0.000964 | 5.7892
2019-04-09 15:21 | 1554837649.7039 | -309 | 0.20945 | 0.000436 | 2e-06 | 101.1512 | 101.982 | 123.7152 | 20.6896 | 0.001873 | 0.001773 | -154.6197 | 0.23234 | 8.4029 | 0.002662 | 0.004505 | 0.004301 | 3.3e-05 | 0.001333 | 0.27616 | 1.3895 | 2.5476 | 98.48 | 19.98 | 48.255 | 66.3318 | 85.3006 | 63.9221 | 0.011454 | 1.0474 | 0.024415 | 23.3103 | -0.073168 | 0.13822 | 1.6e-05 | 0.001042 | 0.001042 | 5.3473
2019-04-09 15:21 | 1554837650.7044 | -308 | 0.20945 | 0.000438 | 2e-06 | 101.1485 | 101.9317 | 123.7312 | 20.6902 | 0.00186 | 0.001749 | -154.6161 | 0.23217 | 8.4023 | 0.002621 | 0.004077 | 0.003901 | 3.3e-05 | 0.001331 | 0.29264 | 1.3623 | 2.5476 | 98.48 | 19.98 | 48.25 | 66.0366 | 84.325 | 63.3629 | 0.011453 | 1.0491 | 0.024437 | 23.2924 | -0.068997 | 0.13812 | 1.2e-05 | 0.00079 | 0.00079 | 7.0441
2019-04-09 15:21 | 1554837651.7044 | -307 | 0.20951 | 0.000435 | 2e-06 | 101.1342 | 101.697 | 123.8873 | 20.6906 | 0.001879 | 0.001752 | -154.5924 | 0.23245 | 8.4007 | 0.002617 | 0.004029 | 0.003903 | 3.3e-05 | 0.001334 | 0.29087 | 1.3976 | 2.5477 | 98.48 | 19.98 | 48.248 | 66.4959 | 84.2263 | 63.4298 | 0.011452 | 1.0489 | 0.02445 | 23.3094 | -0.049642 | 0.13829 | 1.7e-05 | 0.00111 | 0.00111 | 5.0212
2019-04-09 15:21 | 1554837652.705 | -306 | 0.20949 | 0.000436 | 1e-06 | 101.0964 | 102.0161 | 123.7126 | 20.6909 | 0.001897 | 0.001759 | -154.6048 | 0.23254 | 8.4004 | 0.002523 | 0.004 | 0.003822 | 3.2e-05 | 0.001337 | 0.29623 | 1.3299 | 2.5476 | 98.48 | 19.98 | 48.247 | 66.9088 | 81.9668 | 63.6057 | 0.011452 | 1.0484 | 0.024426 | 23.2991 | -0.081003 | 0.13834 | 1.6e-05 | 0.001075 | 0.001075 | 5.1844
2019-04-09 15:21 | 1554837653.7051 | -305 | 0.20949 | 0.000433 | 2e-06 | 101.1154 | 101.8165 | 123.4157 | 20.6897 | 0.001893 | 0.001689 | -154.6082 | 0.23295 | 8.4033 | 0.002608 | 0.00394 | 0.003818 | 3.5e-05 | 0.00134 | 0.27936 | 1.3707 | 2.5478 | 98.48 | 19.97 | 48.247 | 66.821 | 84.0053 | 61.918 | 0.011445 | 1.0537 | 0.024458 | 23.2127 | -0.061798 | 0.13858 | 1.9e-05 | 0.00128 | 0.00128 | 4.3616
2019-04-09 15:21 | 1554837654.7051 | -304 | 0.20947 | 0.000436 | 2e-06 | 101.1163 | 102.0102 | 123.6659 | 20.6907 | 0.001854 | 0.001735 | -154.5983 | 0.23251 | 8.4012 | 0.002608 | 0.004198 | 0.004009 | 3.4e-05 | 0.001331 | 0.27632 | 1.4073 | 2.5476 | 98.48 | 19.97 | 48.246 | 65.8826 | 84.0158 | 63.0151 | 0.011445 | 1.0502 | 0.024443 | 23.2743 | -0.078725 | 0.13832 | 1.6e-05 | 0.001027 | 0.001027 | 5.4262
2019-04-09 15:21 | 1554837655.7041 | -303 | 0.20951 | 0.000437 | 2e-06 | 101.0574 | 101.7542 | 123.918 | 20.69 | 0.001938 | 0.001745 | -154.6016 | 0.23282 | 8.4005 | 0.002605 | 0.003355 | 0.00317 | 3.6e-05 | 0.00133 | 0.29319 | 1.3943 | 2.5478 | 98.48 | 19.97 | 48.248 | 67.8957 | 83.9417 | 63.2691 | 0.011445 | 1.0494 | 0.024459 | 23.3068 | -0.061463 | 0.13851 | 1.4e-05 | 0.000905 | 0.000905 | 6.171
2019-04-09 15:21 | 1554837656.7042 | -302 | 0.20953 | 0.000441 | 2e-06 | 101.041 | 101.9032 | 123.7037 | 20.6905 | 0.001876 | 0.001734 | -154.5855 | 0.23278 | 8.3998 | 0.002605 | 0.0039 | 0.003769 | 3.3e-05 | 0.001335 | 0.2625 | 1.4057 | 2.5476 | 98.48 | 19.97 | 48.249 | 66.413 | 83.9314 | 63.0024 | 0.011445 | 1.0503 | 0.024447 | 23.2774 | -0.076002 | 0.13848 | 9e-06 | 0.00056 | 0.00056 | 9.9638
2019-04-09 15:21 | 1554837657.7052 | -301 | 0.20949 | 0.000444 | 2e-06 | 101.0261 | 101.7531 | 123.5821 | 20.6901 | 0.001941 | 0.001709 | -154.5882 | 0.23261 | 8.4013 | 0.002655 | 0.004135 | 0.003975 | 3.2e-05 | 0.001339 | 0.28871 | 1.4007 | 2.5476 | 98.48 | 19.97 | 48.254 | 67.9848 | 85.1364 | 62.4028 | 0.011447 | 1.0521 | 0.024457 | 23.2452 | -0.06414 | 0.13838 | 5e-06 | 0.000258 | 0.000258 | 21.6152
2019-04-09 15:21 | 1554837658.7052 | -300 | 0.20948 | 0.00044 | 2e-06 | 101.0189 | 101.6788 | 123.6483 | 20.6894 | 0.00186 | 0.001793 | -154.5905 | 0.2326 | 8.4012 | 0.002635 | 0.004201 | 0.004025 | 3.5e-05 | 0.00134 | 0.27975 | 1.3413 | 2.5476 | 98.48 | 19.97 | 48.257 | 66.0264 | 84.6472 | 64.4264 | 0.011447 | 1.0458 | 0.02439 | 23.3214 | -0.058236 | 0.13837 | 1.1e-05 | 0.000669 | 0.000669 | 8.3343
2019-04-09 15:21 | 1554837659.7047 | -299 | 0.20947 | 0.000442 | 2e-06 | 101.0562 | 101.79 | 123.6237 | 20.6901 | 0.001875 | 0.001754 | -154.5766 | 0.23301 | 8.4005 | 0.002668 | 0.003498 | 0.003277 | 3.4e-05 | 0.001342 | 0.28639 | 1.3964 | 2.5475 | 98.48 | 19.97 | 48.261 | 66.3977 | 85.4392 | 63.4651 | 0.011448 | 1.0488 | 0.024422 | 23.2859 | -0.064714 | 0.13862 | 7e-06 | 0.000399 | 0.000399 | 14.0174
2019-04-09 15:21 | 1554837660.7037 | -298 | 0.2095 | 0.000441 | 2e-06 | 100.9138 | 101.7122 | 123.8161 | 20.6902 | 0.001949 | 0.00171 | -154.5734 | 0.23271 | 8.404 | 0.002668 | 0.003889 | 0.00381 | 3.3e-05 | 0.001335 | 0.28136 | 1.4143 | 2.5477 | 98.48 | 19.97 | 48.27 | 68.1704 | 85.4487 | 62.4209 | 0.01145 | 1.0521 | 0.024479 | 23.2678 | -0.070484 | 0.13844 | 8e-06 | 0.000467 | 0.000467 | 11.9496
2019-04-09 15:21 | 1554837661.7033 | -297 | 0.20945 | 0.000437 | 2e-06 | 100.9633 | 101.6065 | 123.7795 | 20.6893 | 0.001933 | 0.001738 | -154.583 | 0.23254 | 8.4042 | 0.002689 | 0.003971 | 0.00385 | 2.9e-05 | 0.001333 | 0.27776 | 1.3565 | 2.5474 | 98.48 | 19.97 | 48.273 | 67.7755 | 85.9576 | 63.0886 | 0.011451 | 1.05 | 0.024452 | 23.2875 | -0.056808 | 0.13834 | 1.4e-05 | 0.000916 | 0.000916 | 6.0872
2019-04-09 15:21 | 1554837662.7033 | -296 | 0.20949 | 0.000439 | 2e-06 | 100.9195 | 101.7449 | 123.6007 | 20.6913 | 0.001919 | 0.001774 | -154.5714 | 0.23353 | 8.402 | 0.002583 | 0.004124 | 0.00392 | 3.5e-05 | 0.001332 | 0.28385 | 1.3817 | 2.5475 | 98.48 | 19.97 | 48.277 | 67.4406 | 83.4126 | 63.9558 | 0.011452 | 1.0473 | 0.024402 | 23.3007 | -0.072858 | 0.13892 | 1.2e-05 | 0.000792 | 0.000792 | 7.0696
2019-04-09 15:21 | 1554837663.7044 | -295 | 0.20948 | 0.000439 | 2e-06 | 100.9588 | 101.8268 | 123.6547 | 20.6903 | 0.001891 | 0.001738 | -154.5833 | 0.23286 | 8.4012 | 0.002594 | 0.004537 | 0.004361 | 3.2e-05 | 0.001326 | 0.27848 | 1.4264 | 2.5477 | 98.48 | 19.97 | 48.284 | 66.764 | 83.6768 | 63.0905 | 0.011454 | 1.05 | 0.024439 | 23.2758 | -0.076572 | 0.13853 | 1.1e-05 | 0.000694 | 0.000694 | 8.0423
2019-04-09 15:21 | 1554837664.7044 | -294 | 0.20945 | 0.00044 | 1e-06 | 100.9525 | 101.6767 | 123.3706 | 20.6902 | 0.001849 | 0.001758 | -154.5722 | 0.23193 | 8.4032 | 0.002651 | 0.004059 | 0.003977 | 3.5e-05 | 0.001321 | 0.27273 | 1.3603 | 2.5476 | 98.48 | 19.97 | 48.287 | 65.7524 | 85.032 | 63.5713 | 0.011454 | 1.0485 | 0.024394 | 23.2657 | -0.063936 | 0.13798 | 1e-05 | 0.000666 | 0.000666 | 8.3478
2019-04-09 15:21 | 1554837665.7037 | -293 | 0.20948 | 0.000441 | 2e-06 | 100.8876 | 101.6378 | 123.1641 | 20.6905 | 0.001891 | 0.001777 | -154.5719 | 0.23219 | 8.4037 | 0.002635 | 0.004223 | 0.004016 | 3.6e-05 | 0.001322 | 0.28407 | 1.3638 | 2.5475 | 98.48 | 19.97 | 48.29 | 66.7797 | 84.6632 | 64.0277 | 0.011455 | 1.0471 | 0.024357 | 23.2619 | -0.066268 | 0.13813 | 9e-06 | 0.000529 | 0.000529 | 10.5152
2019-04-09 15:21 | 1554837666.7035 | -292 | 0.2095 | 0.000443 | 1e-06 | 100.9519 | 101.6148 | 123.3331 | 20.6898 | 0.001945 | 0.001735 | -154.5812 | 0.23226 | 8.4027 | 0.002646 | 0.004042 | 0.003887 | 3.3e-05 | 0.001322 | 0.28174 | 1.3236 | 2.5474 | 98.48 | 19.97 | 48.291 | 68.0647 | 84.9145 | 63.0298 | 0.011455 | 1.0502 | 0.02441 | 23.2434 | -0.058537 | 0.13817 | 8e-06 | 0.000483 | 0.000483 | 11.5282
2019-04-09 15:21 | 1554837667.7039 | -291 | 0.2095 | 0.000442 | 1e-06 | 100.9242 | 101.5711 | 123.2336 | 20.6907 | 0.001882 | 0.001721 | -154.572 | 0.23221 | 8.4057 | 0.002628 | 0.004392 | 0.004275 | 3.2e-05 | 0.001324 | 0.29192 | 1.345 | 2.5477 | 98.48 | 19.97 | 48.292 | 66.5631 | 84.484 | 62.6908 | 0.011456 | 1.0512 | 0.024412 | 23.2223 | -0.057143 | 0.13814 | 8e-06 | 0.000494 | 0.000494 | 11.279
2019-04-09 15:21 | 1554837668.7039 | -290 | 0.21001 | 0.00044 | 2e-06 | 100.9097 | 101.7555 | 122.9979 | 20.6903 | 0.001893 | 0.001764 | -154.5749 | 0.23235 | 8.4012 | 0.002635 | 0.00421 | 0.004014 | 3.4e-05 | 0.001323 | 0.28351 | 1.4404 | 2.5477 | 98.48 | 19.97 | 48.293 | 66.8132 | 84.6466 | 63.7067 | 0.011456 | 1.0481 | 0.024352 | 23.2352 | -0.074667 | 0.13822 | 1e-05 | 0.000645 | 0.000645 | 8.6297
2019-04-09 15:21 | 1554837669.705 | -289 | 0.21089 | 0.000441 | 2e-06 | 100.8558 | 101.5391 | 122.8878 | 20.6908 | 0.001958 | 0.001749 | -154.5768 | 0.23185 | 8.4002 | 0.00264 | 0.003382 | 0.003272 | 3.3e-05 | 0.001328 | 0.28241 | 1.419 | 2.5476 | 98.48 | 19.97 | 48.294 | 68.373 | 84.7758 | 63.3534 | 0.011456 | 1.0492 | 0.024354 | 23.2126 | -0.060397 | 0.13793 | 9e-06 | 0.00056 | 0.00056 | 9.9287
2019-04-09 15:21 | 1554837670.704 | -288 | 0.21239 | 0.000439 | 2e-06 | 100.8572 | 101.7666 | 122.7078 | 20.6903 | 0.001879 | 0.001782 | -154.5818 | 0.23163 | 8.4027 | 0.002655 | 0.003754 | 0.00361 | 3.2e-05 | 0.001325 | 0.28313 | 1.4462 | 2.5477 | 98.48 | 19.97 | 48.294 | 66.4806 | 85.1313 | 64.1447 | 0.011456 | 1.0467 | 0.024307 | 23.2228 | -0.080281 | 0.1378 | 1.1e-05 | 0.000655 | 0.000655 | 8.4833
2019-04-09 15:21 | 1554837671.7046 | -287 | 0.21242 | 0.00044 | 1e-06 | 100.8721 | 101.6898 | 122.7766 | 20.6906 | 0.001905 | 0.001798 | -154.5811 | 0.23271 | 8.4039 | 0.002653 | 0.004018 | 0.003778 | 3.4e-05 | 0.001327 | 0.27688 | 1.41 | 2.5477 | 98.48 | 19.97 | 48.294 | 67.1201 | 85.0909 | 64.5461 | 0.011456 | 1.0455 | 0.0243 | 23.2432 | -0.072213 | 0.13844 | 1.1e-05 | 0.00072 | 0.00072 | 7.7517
2019-04-09 15:21 | 1554837672.7036 | -286 | 0.21265 | 0.000448 | 2e-06 | 100.873 | 101.5983 | 123.1925 | 20.69 | 0.001917 | 0.001829 | -154.5869 | 0.23278 | 8.4016 | 0.002658 | 0.004472 | 0.004461 | 3.1e-05 | 0.00133 | 0.26233 | 1.3617 | 2.5475 | 98.48 | 19.97 | 48.294 | 67.4073 | 85.2166 | 65.2825 | 0.011456 | 1.0432 | 0.024314 | 23.3079 | -0.064078 | 0.13848 | 0 | -0.000139 | -0.000139 | -40.0771
2019-04-09 15:21 | 1554837673.7037 | -285 | 0.20981 | 0.000445 | 2e-06 | 100.8748 | 101.5928 | 123.3972 | 20.6906 | 0.00191 | 0.00174 | -154.5659 | 0.23144 | 8.4007 | 0.002633 | 0.004303 | 0.004111 | 2.9e-05 | 0.001334 | 0.29391 | 1.4008 | 2.5476 | 98.48 | 19.97 | 48.295 | 67.2316 | 84.6055 | 63.15 | 0.011456 | 1.0498 | 0.024411 | 23.2536 | -0.063435 | 0.13769 | 3e-06 | 6.8e-05 | 6.8e-05 | 81.335
2019-04-09 15:21 | 1554837674.7042 | -284 | 0.20945 | 0.000446 | 2e-06 | 100.8784 | 101.5466 | 123.3843 | 20.6903 | 0.001895 | 0.001719 | -154.5645 | 0.23276 | 8.4028 | 0.002692 | 0.004236 | 0.004056 | 3e-05 | 0.001338 | 0.28385 | 1.3756 | 2.5476 | 98.48 | 19.97 | 48.296 | 66.8731 | 86.0249 | 62.6367 | 0.011457 | 1.0514 | 0.024429 | 23.2346 | -0.059057 | 0.13847 | 2e-06 | 7.6e-05 | 7.6e-05 | 73.636
2019-04-09 15:21 | 1554837675.7041 | -283 | 0.20954 | 0.000459 | 2e-06 | 100.8346 | 101.4699 | 123.2028 | 20.6907 | 0.001914 | 0.001801 | -154.5605 | 0.23099 | 8.4039 | 0.00272 | 0.00432 | 0.004136 | 3.2e-05 | 0.001334 | 0.27649 | 1.3823 | 2.5476 | 98.48 | 19.97 | 48.297 | 67.3246 | 86.7004 | 64.6112 | 0.011457 | 1.0453 | 0.024339 | 23.2857 | -0.056176 | 0.13742 | -1.4e-05 | -0.001133 | -0.001133 | -4.8866
2019-04-09 15:21 | 1554837676.7048 | -282 | 0.2095 | 0.000458 | 2e-06 | 100.8492 | 101.4661 | 123.1012 | 20.6905 | 0.001895 | 0.001749 | -154.5567 | 0.23191 | 8.4018 | 0.002607 | 0.00419 | 0.004023 | 3e-05 | 0.001344 | 0.27959 | 1.3649 | 2.5476 | 98.48 | 19.97 | 48.298 | 66.8774 | 83.985 | 63.3468 | 0.011457 | 1.0492 | 0.024375 | 23.2325 | -0.054548 | 0.13796 | -1.3e-05 | -0.001063 | -0.001063 | -5.2297
2019-04-09 15:21 | 1554837677.7032 | -281 | 0.20948 | 0.000468 | 2e-06 | 100.806 | 101.3766 | 123.3779 | 20.6904 | 0.001868 | 0.001784 | -154.5684 | 0.23224 | 8.403 | 0.002685 | 0.004155 | 0.00396 | 3e-05 | 0.001353 | 0.29999 | 1.3654 | 2.5477 | 98.48 | 19.97 | 48.298 | 66.2162 | 85.8585 | 64.1876 | 0.011457 | 1.0466 | 0.024372 | 23.2876 | -0.050487 | 0.13816 | -2.6e-05 | -0.002011 | -0.002011 | -2.7684
2019-04-09 15:21 | 1554837678.7033 | -280 | 0.20949 | 0.000472 | 2e-06 | 100.7328 | 101.526 | 123.5917 | 20.6895 | 0.001925 | 0.001734 | -154.5677 | 0.23191 | 8.4041 | 0.002636 | 0.003799 | 0.003685 | 3e-05 | 0.001355 | 0.29407 | 1.3711 | 2.5477 | 98.48 | 19.97 | 48.3 | 67.5807 | 84.6778 | 63.0021 | 0.011458 | 1.0503 | 0.024436 | 23.2668 | -0.070159 | 0.13797 | -3e-05 | -0.002301 | -0.002301 | -2.4158
2019-04-09 15:21 | 1554837679.7038 | -279 | 0.20948 | 0.000474 | 2e-06 | 100.742 | 101.271 | 123.5129 | 20.6904 | 0.001892 | 0.001759 | -154.5547 | 0.23132 | 8.4025 | 0.002619 | 0.004288 | 0.004155 | 3e-05 | 0.001358 | 0.28838 | 1.399 | 2.5476 | 98.48 | 19.97 | 48.302 | 66.7946 | 84.2598 | 63.6014 | 0.011458 | 1.0484 | 0.024407 | 23.2801 | -0.046843 | 0.13761 | -3.3e-05 | -0.002556 | -0.002556 | -2.17
2019-04-09 15:21 | 1554837680.7029 | -278 | 0.2095 | 0.00047 | 2e-06 | 100.8153 | 101.5635 | 123.3561 | 20.6903 | 0.001894 | 0.001793 | -154.5583 | 0.232 | 8.4004 | 0.002636 | 0.003632 | 0.003459 | 3.4e-05 | 0.00135 | 0.27688 | 1.4236 | 2.5476 | 98.48 | 19.97 | 48.306 | 66.8553 | 84.6848 | 64.4053 | 0.011459 | 1.0459 | 0.024362 | 23.2931 | -0.066141 | 0.13801 | -2.8e-05 | -0.00215 | -0.00215 | -2.587
2019-04-09 15:21 | 1554837681.703 | -277 | 0.20947 | 0.000475 | 1e-06 | 100.7563 | 101.6103 | 123.0643 | 20.6901 | 0.001827 | 0.001756 | -154.5552 | 0.23169 | 8.4033 | 0.002681 | 0.003858 | 0.003667 | 3.2e-05 | 0.001347 | 0.2912 | 1.3926 | 2.5476 | 98.48 | 19.97 | 48.307 | 65.2374 | 85.7539 | 63.5361 | 0.011459 | 1.0486 | 0.024365 | 23.2356 | -0.075495 | 0.13783 | -3.3e-05 | -0.002527 | -0.002527 | -2.1978
2019-04-09 15:21 | 1554837682.704 | -276 | 0.20951 | 0.000476 | 2e-06 | 100.7762 | 101.4169 | 122.8106 | 20.6899 | 0.001862 | 0.00177 | -154.5544 | 0.23139 | 8.4003 | 0.002578 | 0.004598 | 0.004413 | 3.3e-05 | 0.00135 | 0.28152 | 1.3872 | 2.5477 | 98.48 | 19.97 | 48.309 | 66.0856 | 83.2715 | 63.8623 | 0.01146 | 1.0476 | 0.024328 | 23.2228 | -0.056682 | 0.13765 | -3.5e-05 | -0.00264 | -0.00264 | -2.1016
2019-04-09 15:21 | 1554837683.7034 | -275 | 0.20953 | 0.000476 | 2e-06 | 100.8611 | 101.5869 | 122.7003 | 20.6899 | 0.00184 | 0.001767 | -154.5533 | 0.23225 | 8.4024 | 0.002704 | 0.003622 | 0.003564 | 3.1e-05 | 0.001353 | 0.28528 | 1.3989 | 2.5477 | 98.48 | 19.97 | 48.311 | 65.5578 | 86.3187 | 63.785 | 0.01146 | 1.0478 | 0.02432 | 23.2097 | -0.064134 | 0.13817 | -3.6e-05 | -0.002711 | -0.002711 | -2.0541
2019-04-09 15:21 | 1554837684.703 | -274 | 0.20947 | 0.000472 | 1e-06 | 100.7896 | 101.4971 | 122.6815 | 20.6907 | 0.001865 | 0.001767 | -154.5435 | 0.23224 | 8.4024 | 0.002637 | 0.004331 | 0.00414 | 3e-05 | 0.001356 | 0.27096 | 1.4249 | 2.5477 | 98.48 | 19.97 | 48.313 | 66.1371 | 84.6977 | 63.7874 | 0.011461 | 1.0478 | 0.024318 | 23.2081 | -0.062562 | 0.13816 | -3e-05 | -0.002286 | -0.002286 | -2.4355
2019-04-09 15:21 | 1554837685.7034 | -273 | 0.2095 | 0.000467 | 2e-06 | 100.6938 | 101.5601 | 122.6388 | 20.6901 | 0.001893 | 0.001756 | -154.5426 | 0.23109 | 8.4019 | 0.002713 | 0.00414 | 0.004016 | 3.3e-05 | 0.001354 | 0.26449 | 1.4312 | 2.5477 | 98.47 | 19.98 | 48.314 | 66.812 | 86.5218 | 63.5277 | 0.011469 | 1.0486 | 0.024323 | 23.1951 | -0.076624 | 0.13747 | -2.3e-05 | -0.001813 | -0.001813 | -3.055
2019-04-09 15:21 | 1554837686.7036 | -272 | 0.20949 | 0.000467 | 1e-06 | 100.6783 | 101.495 | 122.5186 | 20.6898 | 0.001865 | 0.001758 | -154.5356 | 0.23195 | 8.4019 | 0.002758 | 0.004506 | 0.004363 | 3.4e-05 | 0.001364 | 0.27649 | 1.3767 | 2.5476 | 98.47 | 19.97 | 48.315 | 66.1509 | 87.6023 | 63.5741 | 0.011462 | 1.0485 | 0.024309 | 23.1853 | -0.07227 | 0.13799 | -2.4e-05 | -0.001837 | -0.001837 | -3.0266
2019-04-09 15:21 | 1554837687.7032 | -271 | 0.2095 | 0.00047 | 2e-06 | 100.745 | 101.3525 | 122.235 | 20.6915 | 0.00194 | 0.001784 | -154.5222 | 0.23204 | 8.4027 | 0.002627 | 0.003846 | 0.003755 | 3.5e-05 | 0.001363 | 0.27328 | 1.3717 | 2.5475 | 98.47 | 19.97 | 48.317 | 67.9423 | 84.4604 | 64.2089 | 0.011463 | 1.0465 | 0.024258 | 23.1803 | -0.05378 | 0.13804 | -2.7e-05 | -0.002085 | -0.002085 | -2.6679
2019-04-09 15:21 | 1554837688.7043 | -270 | 0.20949 | 0.000464 | 2e-06 | 100.7539 | 101.5129 | 122.0372 | 20.6908 | 0.001881 | 0.001713 | -154.5231 | 0.23321 | 8.4004 | 0.00259 | 0.004386 | 0.004192 | 3.2e-05 | 0.001362 | 0.28943 | 1.337 | 2.5474 | 98.47 | 19.98 | 48.321 | 66.5405 | 83.565 | 62.4935 | 0.011471 | 1.0518 | 0.0243 | 23.1025 | -0.067129 | 0.13874 | -2.1e-05 | -0.001651 | -0.001651 | -3.3874
2019-04-09 15:21 | 1554837689.7043 | -269 | 0.20947 | 0.000465 | 2e-06 | 100.7265 | 101.4813 | 122.0595 | 20.6897 | 0.001861 | 0.00172 | -154.5401 | 0.23221 | 8.4017 | 0.002686 | 0.003842 | 0.003629 | 3.2e-05 | 0.001355 | 0.27903 | 1.3755 | 2.5476 | 98.47 | 19.98 | 48.322 | 66.0524 | 85.8884 | 62.6462 | 0.011471 | 1.0514 | 0.024297 | 23.1099 | -0.066771 | 0.13814 | -2.1e-05 | -0.001625 | -0.001625 | -3.4269
2019-04-09 15:22 | 1554837690.7043 | -268 | 0.20949 | 0.000465 | 2e-06 | 100.7298 | 101.4826 | 122.1266 | 20.691 | 0.001945 | 0.001784 | -154.5378 | 0.23247 | 8.4022 | 0.002667 | 0.003745 | 0.003604 | 3.7e-05 | 0.00134 | 0.27544 | 1.3714 | 2.5477 | 98.47 | 19.98 | 48.323 | 68.062 | 85.4338 | 64.1906 | 0.011471 | 1.0466 | 0.024248 | 23.1694 | -0.066602 | 0.13829 | -2.2e-05 | -0.001676 | -0.001676 | -3.3245
2019-04-09 15:22 | 1554837691.7033 | -267 | 0.20947 | 0.000459 | 2e-06 | 100.7605 | 101.3821 | 122.3893 | 20.6901 | 0.001877 | 0.001741 | -154.5403 | 0.23254 | 8.4036 | 0.002584 | 0.004647 | 0.004462 | 3.3e-05 | 0.00134 | 0.26825 | 1.3772 | 2.5477 | 98.47 | 19.98 | 48.325 | 66.4344 | 83.419 | 63.1658 | 0.011472 | 1.0497 | 0.024311 | 23.159 | -0.055017 | 0.13834 | -1.4e-05 | -0.001133 | -0.001133 | -4.9225
2019-04-09 15:22 | 1554837692.7024 | -266 | 0.20949 | 0.000462 | 1e-06 | 100.795 | 101.4317 | 122.5566 | 20.6904 | 0.001889 | 0.001733 | -154.5551 | 0.23275 | 8.4016 | 0.0027 | 0.003603 | 0.003485 | 3.3e-05 | 0.001338 | 0.2808 | 1.395 | 2.5476 | 98.47 | 19.97 | 48.325 | 66.7133 | 86.2076 | 62.9608 | 0.011465 | 1.0504 | 0.024335 | 23.1678 | -0.056322 | 0.13846 | -1.6e-05 | -0.001262 | -0.001262 | -4.4225
2019-04-09 15:22 | 1554837693.7023 | -265 | 0.20946 | 0.000456 | 2e-06 | 100.7343 | 101.3636 | 122.4052 | 20.6898 | 0.001895 | 0.001807 | -154.5667 | 0.23238 | 8.4004 | 0.002686 | 0.003811 | 0.003588 | 3.8e-05 | 0.001333 | 0.28152 | 1.3814 | 2.5477 | 98.47 | 19.97 | 48.327 | 66.8579 | 85.869 | 64.7592 | 0.011465 | 1.0448 | 0.024255 | 23.2153 | -0.055704 | 0.13824 | -1.1e-05 | -0.000965 | -0.000965 | -5.7743
2019-04-09 15:22 | 1554837694.7034 | -264 | 0.20948 | 0.000457 | 2e-06 | 100.756 | 101.5343 | 122.5749 | 20.6905 | 0.001858 | 0.00178 | -154.5968 | 0.2317 | 8.4024 | 0.002711 | 0.003735 | 0.003567 | 3.4e-05 | 0.001334 | 0.27383 | 1.4075 | 2.5476 | 98.47 | 19.97 | 48.328 | 65.9759 | 86.4831 | 64.1119 | 0.011465 | 1.0468 | 0.024295 | 23.2091 | -0.068824 | 0.13784 | -1.1e-05 | -0.000896 | -0.000896 | -6.1963
2019-04-09 15:22 | 1554837695.7039 | -263 | 0.20948 | 0.000458 | 2e-06 | 100.6982 | 101.5308 | 122.9437 | 20.6906 | 0.001849 | 0.001801 | -154.6078 | 0.23181 | 8.3995 | 0.00272 | 0.004061 | 0.003896 | 3.9e-05 | 0.001323 | 0.28495 | 1.3899 | 2.5478 | 98.47 | 19.97 | 48.329 | 65.7677 | 86.6878 | 64.6173 | 0.011466 | 1.0452 | 0.024314 | 23.2614 | -0.073651 | 0.1379 | -1.3e-05 | -0.001033 | -0.001033 | -5.3782
2019-04-09 15:22 | 1554837696.7025 | -262 | 0.2095 | 0.000458 | 2e-06 | 100.7435 | 101.5566 | 123.2047 | 20.6903 | 0.001882 | 0.001698 | -154.6084 | 0.2323 | 8.4018 | 0.002708 | 0.004194 | 0.004021 | 3.4e-05 | 0.001322 | 0.29192 | 1.3667 | 2.5476 | 98.47 | 19.97 | 48.33 | 66.5508 | 86.4173 | 62.139 | 0.011466 | 1.053 | 0.024429 | 23.2005 | -0.071905 | 0.1382 | -1.3e-05 | -0.001089 | -0.001089 | -5.1133
2019-04-09 15:22 | 1554837697.7024 | -261 | 0.20952 | 0.000453 | 2e-06 | 100.7125 | 101.4868 | 123.4404 | 20.6909 | 0.001875 | 0.001754 | -154.6178 | 0.23227 | 8.4026 | 0.002725 | 0.00394 | 0.003658 | 3.7e-05 | 0.001316 | 0.27129 | 1.3579 | 2.5475 | 98.47 | 19.97 | 48.331 | 66.3821 | 86.8209 | 63.4892 | 0.011466 | 1.0487 | 0.024403 | 23.2694 | -0.0685 | 0.13818 | -7e-06 | -0.000589 | -0.000589 | -9.454
2019-04-09 15:22 | 1554837698.7036 | -260 | 0.20951 | 0.000453 | 2e-06 | 100.7235 | 101.4062 | 123.6331 | 20.6906 | 0.001894 | 0.001798 | -154.626 | 0.23211 | 8.4023 | 0.002706 | 0.003959 | 0.003854 | 3.8e-05 | 0.00131 | 0.2672 | 1.3586 | 2.5476 | 98.47 | 19.97 | 48.332 | 66.8438 | 86.3616 | 64.5385 | 0.011466 | 1.0455 | 0.024385 | 23.3238 | -0.060421 | 0.13809 | -6e-06 | -0.00054 | -0.00054 | -10.309
2019-04-09 15:22 | 1554837699.7035 | -259 | 0.20952 | 0.000455 | 1e-06 | 100.7116 | 101.4668 | 123.7403 | 20.6899 | 0.001889 | 0.001806 | -154.6115 | 0.2321 | 8.4009 | 0.002722 | 0.00487 | 0.004703 | 4e-05 | 0.001312 | 0.27113 | 1.332 | 2.5476 | 98.47 | 19.97 | 48.333 | 66.7354 | 86.7393 | 64.7367 | 0.011467 | 1.0449 | 0.024388 | 23.3408 | -0.06682 | 0.13808 | -9e-06 | -0.000724 | -0.000724 | -7.6903
2019-04-09 15:22 | 1554837700.7037 | -258 | 0.20954 | 0.000457 | 1e-06 | 100.7316 | 101.3801 | 123.8007 | 20.6898 | 0.001906 | 0.00178 | -154.5995 | 0.23204 | 8.4008 | 0.002722 | 0.003695 | 0.003469 | 4e-05 | 0.001323 | 0.28385 | 1.3599 | 2.5476 | 98.47 | 19.98 | 48.334 | 67.1239 | 86.7549 | 64.098 | 0.011474 | 1.0468 | 0.024417 | 23.3244 | -0.057399 | 0.13804 | -1e-05 | -0.000843 | -0.000843 | -6.5999
2019-04-09 15:22 | 1554837701.703 | -257 | 0.20952 | 0.000455 | 2e-06 | 100.801 | 101.5673 | 123.7809 | 20.6904 | 0.001918 | 0.001737 | -154.6057 | 0.23234 | 8.4027 | 0.002693 | 0.004261 | 0.004098 | 3.9e-05 | 0.001327 | 0.2943 | 1.3537 | 2.5475 | 98.47 | 19.98 | 48.335 | 67.4279 | 86.056 | 63.0725 | 0.011474 | 1.05 | 0.024452 | 23.2871 | -0.067743 | 0.13822 | -9e-06 | -0.000739 | -0.000739 | -7.5407
2019-04-09 15:22 | 1554837702.7041 | -256 | 0.2095 | 0.000455 | 1e-06 | 100.708 | 101.3119 | 123.7765 | 20.6898 | 0.001845 | 0.001785 | -154.5868 | 0.23199 | 8.4005 | 0.002653 | 0.004243 | 0.004094 | 3.6e-05 | 0.001331 | 0.28871 | 1.3171 | 2.5476 | 98.47 | 19.98 | 48.335 | 65.6554 | 85.0947 | 64.2208 | 0.011474 | 1.0465 | 0.02441 | 23.3264 | -0.053475 | 0.13801 | -8e-06 | -0.000636 | -0.000636 | -8.7394
2019-04-09 15:22 | 1554837703.7041 | -255 | 0.20952 | 0.000452 | 1e-06 | 100.7378 | 101.5817 | 123.776 | 20.6901 | 0.001894 | 0.001786 | -154.5844 | 0.23213 | 8.4007 | 0.002688 | 0.003681 | 0.003497 | 3.7e-05 | 0.001336 | 0.27649 | 1.3556 | 2.5474 | 98.47 | 19.98 | 48.335 | 66.8387 | 85.9333 | 64.253 | 0.011474 | 1.0464 | 0.024409 | 23.3274 | -0.07462 | 0.13809 | -5e-06 | -0.00043 | -0.00043 | -12.9437
2019-04-09 15:22 | 1554837704.7027 | -254 | 0.20952 | 0.000452 | 2e-06 | 100.8164 | 101.4933 | 123.4274 | 20.6907 | 0.00182 | 0.001782 | -154.5804 | 0.23241 | 8.4016 | 0.002696 | 0.003664 | 0.003499 | 3.9e-05 | 0.00133 | 0.28047 | 1.3372 | 2.5476 | 98.47 | 19.98 | 48.335 | 65.0623 | 86.1235 | 64.1591 | 0.011474 | 1.0467 | 0.024378 | 23.2913 | -0.059851 | 0.13826 | -5e-06 | -0.000477 | -0.000477 | -11.6719
2019-04-09 15:22 | 1554837705.7042 | -253 | 0.20951 | 0.000448 | 2e-06 | 100.7235 | 101.5446 | 123.3235 | 20.6907 | 0.001866 | 0.001753 | -154.5692 | 0.2319 | 8.3999 | 0.002604 | 0.004249 | 0.004064 | 3.8e-05 | 0.001325 | 0.28528 | 1.3454 | 2.5477 | 98.47 | 19.98 | 48.335 | 66.1611 | 83.8986 | 63.4486 | 0.011474 | 1.0489 | 0.024393 | 23.257 | -0.072615 | 0.13796 | 0 | -0.000125 | -0.000125 | -44.3611
2019-04-09 15:22 | 1554837706.7038 | -252 | 0.2095 | 0.000446 | 2e-06 | 100.7563 | 101.4609 | 123.2167 | 20.6909 | 0.001923 | 0.001786 | -154.5691 | 0.23232 | 8.4011 | 0.002698 | 0.00377 | 0.003654 | 3.8e-05 | 0.001327 | 0.24889 | 1.3283 | 2.5476 | 98.47 | 19.98 | 48.336 | 67.5541 | 86.1674 | 64.2374 | 0.011474 | 1.0464 | 0.024354 | 23.2741 | -0.062337 | 0.13821 | 3e-06 | 0.00015 | 0.00015 | 37.2022
2019-04-09 15:22 | 1554837707.7018 | -251 | 0.20951 | 0.000451 | 2e-06 | 100.6744 | 101.7204 | 123.2981 | 20.6895 | 0.00184 | 0.001754 | -154.5595 | 0.23241 | 8.403 | 0.00268 | 0.004234 | 0.004049 | 3.7e-05 | 0.001322 | 0.28031 | 1.315 | 2.5476 | 98.47 | 19.98 | 48.335 | 65.5442 | 85.7408 | 63.484 | 0.011474 | 1.0488 | 0.02439 | 23.2558 | -0.092457 | 0.13826 | -4e-06 | -0.000427 | -0.000427 | -13.0493
2019-04-09 15:22 | 1554837708.7018 | -250 | 0.2095 | 0.00045 | 1e-06 | 100.7232 | 101.5133 | 123.2143 | 20.6901 | 0.0019 | 0.001729 | -154.56 | 0.23152 | 8.4037 | 0.002651 | 0.004003 | 0.003905 | 3.7e-05 | 0.001326 | 0.27599 | 1.3278 | 2.5476 | 98.47 | 19.98 | 48.336 | 66.9831 | 85.0373 | 62.8638 | 0.011474 | 1.0507 | 0.024404 | 23.2265 | -0.069883 | 0.13773 | -2e-06 | -0.000251 | -0.000251 | -22.1292
2019-04-09 15:22 | 1554837709.7018 | -249 | 0.20953 | 0.00045 | 2e-06 | 100.7733 | 101.3907 | 122.8965 | 20.6907 | 0.001893 | 0.001784 | -154.5594 | 0.23309 | 8.3999 | 0.002608 | 0.003798 | 0.003674 | 3.5e-05 | 0.001328 | 0.28456 | 1.3887 | 2.5475 | 98.47 | 19.98 | 48.335 | 66.8331 | 83.9931 | 64.2067 | 0.011474 | 1.0465 | 0.024324 | 23.2428 | -0.054639 | 0.13867 | -2e-06 | -0.000271 | -0.000271 | -20.6264
2019-04-09 15:22 | 1554837710.7023 | -248 | 0.20952 | 0.00045 | 1e-06 | 100.7247 | 101.4647 | 122.8848 | 20.69 | 0.001907 | 0.001733 | -154.5668 | 0.23179 | 8.4013 | 0.002572 | 0.004079 | 0.003895 | 3.9e-05 | 0.001327 | 0.26825 | 1.3897 | 2.5477 | 98.47 | 19.98 | 48.336 | 67.1491 | 83.1357 | 62.9823 | 0.011474 | 1.0503 | 0.024367 | 23.1995 | -0.065474 | 0.13789 | -3e-06 | -0.000285 | -0.000285 | -19.5229
2019-04-09 15:22 | 1554837711.7033 | -247 | 0.20956 | 0.000452 | 1e-06 | 100.6863 | 101.7459 | 122.8167 | 20.6898 | 0.001919 | 0.001772 | -154.5601 | 0.2315 | 8.4053 | 0.002686 | 0.004198 | 0.004078 | 3.6e-05 | 0.001332 | 0.26698 | 1.3763 | 2.5477 | 98.47 | 19.98 | 48.337 | 67.443 | 85.8822 | 63.9161 | 0.011475 | 1.0474 | 0.024326 | 23.2253 | -0.093638 | 0.13772 | -4e-06 | -0.000377 | -0.000377 | -14.7322
2019-04-09 15:22 | 1554837712.703 | -246 | 0.20951 | 0.000448 | 2e-06 | 100.7411 | 101.5195 | 122.559 | 20.6897 | 0.001865 | 0.001765 | -154.5526 | 0.23305 | 8.402 | 0.002663 | 0.004 | 0.003841 | 3.7e-05 | 0.001345 | 0.27903 | 1.3593 | 2.5476 | 98.47 | 19.98 | 48.336 | 66.1453 | 85.3303 | 63.7504 | 0.011474 | 1.0479 | 0.024307 | 23.1952 | -0.068842 | 0.13864 | 0 | -0.000116 | -0.000116 | -48.0604
2019-04-09 15:22 | 1554837713.702 | -245 | 0.20949 | 0.000449 | 2e-06 | 100.7426 | 101.5546 | 122.8108 | 20.6897 | 0.001899 | 0.001821 | -154.5445 | 0.23279 | 8.4025 | 0.002676 | 0.004523 | 0.004305 | 3.3e-05 | 0.001349 | 0.29623 | 1.3631 | 2.5476 | 98.47 | 19.98 | 48.337 | 66.9631 | 85.6363 | 65.0895 | 0.011475 | 1.0438 | 0.024283 | 23.2651 | -0.071802 | 0.13849 | -1e-06 | -0.000205 | -0.000205 | -27.2679
2019-04-09 15:22 | 1554837714.703 | -244 | 0.20952 | 0.000449 | 2e-06 | 100.7774 | 101.589 | 122.9294 | 20.6901 | 0.001876 | 0.001803 | -154.5339 | 0.23282 | 8.4023 | 0.002632 | 0.00463 | 0.004485 | 3.4e-05 | 0.001351 | 0.28313 | 1.3554 | 2.5476 | 98.47 | 19.98 | 48.339 | 66.4205 | 84.5831 | 64.6565 | 0.011475 | 1.0451 | 0.024311 | 23.2614 | -0.07174 | 0.13851 | -1e-06 | -0.000243 | -0.000243 | -22.948
2019-04-09 15:22 | 1554837715.702 | -243 | 0.20952 | 0.000446 | 2e-06 | 100.7018 | 101.5002 | 122.6275 | 20.6902 | 0.001837 | 0.001796 | -154.5303 | 0.23327 | 8.401 | 0.002654 | 0.004335 | 0.004243 | 3.6e-05 | 0.00135 | 0.26665 | 1.3805 | 2.5476 | 98.47 | 19.98 | 48.339 | 65.4746 | 85.1054 | 64.4817 | 0.011475 | 1.0457 | 0.024287 | 23.2268 | -0.070635 | 0.13877 | 3e-06 | 7.2e-05 | 7.2e-05 | 77.4516
2019-04-09 15:22 | 1554837716.7023 | -242 | 0.20952 | 0.000451 | 2e-06 | 100.7551 | 101.591 | 122.493 | 20.6905 | 0.001849 | 0.001776 | -154.5386 | 0.232 | 8.4038 | 0.002681 | 0.004373 | 0.004204 | 3.4e-05 | 0.001358 | 0.28152 | 1.3406 | 2.5476 | 98.47 | 19.98 | 48.34 | 65.7669 | 85.7667 | 64.0183 | 0.011475 | 1.0471 | 0.024291 | 23.1982 | -0.073904 | 0.13801 | -5e-06 | -0.000449 | -0.000449 | -12.3853
2019-04-09 15:22 | 1554837717.7031 | -241 | 0.20953 | 0.000452 | 2e-06 | 100.7343 | 101.5353 | 122.194 | 20.6893 | 0.001857 | 0.001815 | -154.5398 | 0.23248 | 8.4028 | 0.002645 | 0.003869 | 0.003743 | 3.4e-05 | 0.001362 | 0.27759 | 1.3449 | 2.5475 | 98.47 | 19.98 | 48.34 | 65.9648 | 84.8883 | 64.9376 | 0.011475 | 1.0442 | 0.024228 | 23.2014 | -0.070845 | 0.1383 | -5e-06 | -0.000481 | -0.000481 | -11.5859
2019-04-09 15:22 | 1554837718.7025 | -240 | 0.2095 | 0.000457 | 2e-06 | 100.7524 | 101.4107 | 122.1125 | 20.6908 | 0.001942 | 0.001786 | -154.5302 | 0.23208 | 8.4034 | 0.002707 | 0.004116 | 0.003872 | 3.3e-05 | 0.001358 | 0.27201 | 1.3596 | 2.5474 | 98.47 | 19.98 | 48.342 | 67.9916 | 86.3781 | 64.2458 | 0.011476 | 1.0464 | 0.024245 | 23.1699 | -0.05825 | 0.13807 | -1.1e-05 | -0.00093 | -0.00093 | -5.985
2019-04-09 15:22 | 1554837719.7031 | -239 | 0.20953 | 0.000446 | 1e-06 | 100.7777 | 101.4616 | 122.0686 | 20.6904 | 0.001865 | 0.001797 | -154.5343 | 0.23322 | 8.4002 | 0.00261 | 0.004581 | 0.004501 | 3.3e-05 | 0.001363 | 0.2672 | 1.3622 | 2.5475 | 98.47 | 19.98 | 48.343 | 66.149 | 84.0481 | 64.5007 | 0.011476 | 1.0456 | 0.024231 | 23.1745 | -0.060495 | 0.13874 | 3e-06 | 0.000164 | 0.000164 | 34.0835
2019-04-09 15:22 | 1554837720.7038 | -238 | 0.20952 | 0.00045 | 2e-06 | 100.7039 | 101.4434 | 122.1041 | 20.6899 | 0.001844 | 0.001827 | -154.5354 | 0.23165 | 8.4028 | 0.002629 | 0.004011 | 0.003875 | 3.4e-05 | 0.00137 | 0.25984 | 1.3275 | 2.5476 | 98.47 | 19.98 | 48.344 | 65.6415 | 84.512 | 65.2333 | 0.011476 | 1.0433 | 0.024208 | 23.203 | -0.065444 | 0.13781 | -3e-06 | -0.000311 | -0.000311 | -17.8479
2019-04-09 15:22 | 1554837721.7026 | -237 | 0.20952 | 0.000451 | 2e-06 | 100.7313 | 101.5298 | 122.0719 | 20.6899 | 0.001868 | 0.001758 | -154.5404 | 0.23275 | 8.4023 | 0.002624 | 0.003698 | 0.003561 | 3.4e-05 | 0.001374 | 0.28136 | 1.3526 | 2.5477 | 98.47 | 19.98 | 48.347 | 66.2275 | 84.381 | 63.5735 | 0.011477 | 1.0485 | 0.024265 | 23.143 | -0.070624 | 0.13846 | -4e-06 | -0.000381 | -0.000381 | -14.6499
2019-04-09 15:22 | 1554837722.7032 | -236 | 0.20955 | 0.000445 | 2e-06 | 100.7539 | 101.6045 | 122.3367 | 20.6892 | 0.001824 | 0.001792 | -154.5534 | 0.23254 | 8.4005 | 0.002606 | 0.004052 | 0.003786 | 3.4e-05 | 0.001375 | 0.27129 | 1.349 | 2.5476 | 98.47 | 19.99 | 48.349 | 65.153 | 83.9472 | 64.3824 | 0.011485 | 1.046 | 0.024262 | 23.1959 | -0.075192 | 0.13834 | 3e-06 | 0.000112 | 0.000112 | 49.6153
2019-04-09 15:22 | 1554837723.7023 | -235 | 0.20954 | 0.000449 | 2e-06 | 100.8313 | 101.5112 | 121.9524 | 20.6905 | 0.001894 | 0.001768 | -154.5414 | 0.23263 | 8.4025 | 0.002667 | 0.003473 | 0.003348 | 3.3e-05 | 0.001376 | 0.28119 | 1.3609 | 2.5477 | 98.47 | 19.99 | 48.351 | 66.8538 | 85.4208 | 63.8051 | 0.011485 | 1.0478 | 0.024245 | 23.1396 | -0.060107 | 0.13839 | -2e-06 | -0.000214 | -0.000214 | -26.0559
2019-04-09 15:22 | 1554837724.7024 | -234 | 0.20951 | 0.000449 | 2e-06 | 100.7908 | 101.6113 | 121.9752 | 20.6897 | 0.001918 | 0.001741 | -154.54 | 0.23179 | 8.4004 | 0.002629 | 0.005048 | 0.004942 | 3.3e-05 | 0.001374 | 0.29717 | 1.3585 | 2.5475 | 98.47 | 19.98 | 48.352 | 67.4164 | 84.5094 | 63.153 | 0.011478 | 1.0498 | 0.02427 | 23.1194 | -0.07252 | 0.13789 | -1e-06 | -0.000157 | -0.000157 | -35.3075
2019-04-09 15:22 | 1554837725.7014 | -233 | 0.20955 | 0.00045 | 2e-06 | 100.7614 | 101.44 | 122.3611 | 20.6911 | 0.00181 | 0.001772 | -154.5485 | 0.23235 | 8.4009 | 0.002565 | 0.003803 | 0.00359 | 3.4e-05 | 0.001376 | 0.26089 | 1.4029 | 2.5478 | 98.47 | 19.99 | 48.352 | 64.8307 | 82.9793 | 63.9204 | 0.011485 | 1.0474 | 0.024281 | 23.1823 | -0.060037 | 0.13822 | -2e-06 | -0.000289 | -0.000289 | -19.2709
2019-04-09 15:22 | 1554837726.7014 | -232 | 0.20956 | 0.000451 | 1e-06 | 100.7226 | 101.4011 | 122.413 | 20.6902 | 0.001932 | 0.001794 | -154.5457 | 0.23308 | 8.3983 | 0.00261 | 0.003676 | 0.003472 | 3.7e-05 | 0.001374 | 0.28119 | 1.3627 | 2.5476 | 98.47 | 19.98 | 48.355 | 67.7663 | 84.059 | 64.4341 | 0.011479 | 1.0458 | 0.024268 | 23.2049 | -0.060045 | 0.13866 | -3e-06 | -0.000337 | -0.000337 | -16.5624
2019-04-09 15:22 | 1554837727.7015 | -231 | 0.20953 | 0.000449 | 2e-06 | 100.7435 | 101.4214 | 122.4301 | 20.6903 | 0.001894 | 0.001734 | -154.5426 | 0.23229 | 8.4004 | 0.002629 | 0.004358 | 0.004173 | 3.3e-05 | 0.001373 | 0.2912 | 1.3614 | 2.5476 | 98.47 | 19.98 | 48.356 | 66.8505 | 84.5014 | 63.0077 | 0.011479 | 1.0502 | 0.024321 | 23.1574 | -0.059985 | 0.13819 | -1e-06 | -0.000207 | -0.000207 | -26.928
2019-04-09 15:22 | 1554837728.7029 | -230 | 0.20953 | 0.000449 | 2e-06 | 100.7554 | 101.2896 | 122.3057 | 20.6908 | 0.00195 | 0.001798 | -154.5484 | 0.23189 | 8.403 | 0.002607 | 0.003975 | 0.003883 | 3.2e-05 | 0.001374 | 0.27704 | 1.3784 | 2.5475 | 98.47 | 19.99 | 48.361 | 68.1866 | 83.9838 | 64.5369 | 0.011487 | 1.0455 | 0.024253 | 23.1982 | -0.047295 | 0.13795 | -1e-06 | -0.000158 | -0.000158 | -35.2254
2019-04-09 15:22 | 1554837729.7025 | -229 | 0.20951 | 0.000452 | 2e-06 | 100.7813 | 101.6881 | 122.6737 | 20.69 | 0.001886 | 0.001739 | -154.5545 | 0.23183 | 8.4023 | 0.002635 | 0.004152 | 0.004014 | 3.3e-05 | 0.001368 | 0.27848 | 1.3581 | 2.5474 | 98.47 | 19.99 | 48.362 | 66.6458 | 84.647 | 63.1205 | 0.011488 | 1.0499 | 0.024341 | 23.1843 | -0.080119 | 0.13792 | -6e-06 | -0.00054 | -0.00054 | -10.2834
2019-04-09 15:22 | 1554837730.7027 | -228 | 0.20952 | 0.000452 | 2e-06 | 100.7414 | 101.3852 | 123.0535 | 20.6902 | 0.001884 | 0.001754 | -154.5546 | 0.23209 | 8.4038 | 0.002585 | 0.004191 | 0.003989 | 3.3e-05 | 0.001366 | 0.2844 | 1.3531 | 2.5476 | 98.47 | 19.99 | 48.362 | 66.614 | 83.4435 | 63.4691 | 0.011488 | 1.0488 | 0.024366 | 23.2322 | -0.056982 | 0.13807 | -5e-06 | -0.000513 | -0.000513 | -10.8513
2019-04-09 15:22 | 1554837731.7015 | -227 | 0.20951 | 0.00045 | 2e-06 | 100.7247 | 101.3432 | 122.9129 | 20.6902 | 0.00185 | 0.001746 | -154.5541 | 0.23267 | 8.4037 | 0.002604 | 0.003677 | 0.00357 | 3.5e-05 | 0.001367 | 0.29048 | 1.3439 | 2.5476 | 98.47 | 19.99 | 48.364 | 65.7831 | 83.9145 | 63.2825 | 0.011488 | 1.0494 | 0.024359 | 23.2125 | -0.054758 | 0.13841 | -2e-06 | -0.00029 | -0.00029 | -19.2134
2019-04-09 15:22 | 1554837732.702 | -226 | 0.20954 | 0.000446 | 1e-06 | 100.7715 | 101.686 | 122.6329 | 20.6903 | 0.001919 | 0.0017 | -154.5452 | 0.2328 | 8.402 | 0.002606 | 0.004315 | 0.00405 | 3.4e-05 | 0.00136 | 0.28816 | 1.3522 | 2.5476 | 98.47 | 19.99 | 48.364 | 67.4389 | 83.9458 | 62.1811 | 0.011488 | 1.0528 | 0.024371 | 23.1481 | -0.08081 | 0.1385 | 3e-06 | 0.000159 | 0.000159 | 35.0125
2019-04-09 15:22 | 1554837733.7027 | -225 | 0.20955 | 0.000452 | 2e-06 | 100.7664 | 101.5022 | 122.8728 | 20.6896 | 0.001866 | 0.001748 | -154.5632 | 0.23223 | 8.401 | 0.002639 | 0.004509 | 0.004383 | 3.7e-05 | 0.001365 | 0.28495 | 1.339 | 2.5474 | 98.47 | 19.99 | 48.364 | 66.168 | 84.7559 | 63.3255 | 0.011488 | 1.0492 | 0.024353 | 23.2102 | -0.065079 | 0.13816 | -4e-06 | -0.000428 | -0.000428 | -13.0093
2019-04-09 15:22 | 1554837734.7016 | -224 | 0.20957 | 0.00045 | 2e-06 | 100.7682 | 101.4665 | 122.9491 | 20.6905 | 0.001891 | 0.00174 | -154.5655 | 0.23299 | 8.4008 | 0.002723 | 0.003161 | 0.003099 | 3.5e-05 | 0.001366 | 0.27649 | 1.3843 | 2.5473 | 98.47 | 19.99 | 48.361 | 66.7836 | 86.7604 | 63.1304 | 0.011487 | 1.0499 | 0.024368 | 23.2107 | -0.061766 | 0.1386 | -2e-06 | -0.000283 | -0.000283 | -19.7463
2019-04-09 15:22 | 1554837735.7017 | -223 | 0.20953 | 0.000451 | 2e-06 | 100.7173 | 101.5009 | 122.5066 | 20.6902 | 0.001853 | 0.001795 | -154.5635 | 0.23238 | 8.4016 | 0.002646 | 0.004362 | 0.004169 | 3.5e-05 | 0.001367 | 0.28982 | 1.3522 | 2.5475 | 98.47 | 19.99 | 48.361 | 65.8526 | 84.9275 | 64.4577 | 0.011487 | 1.0457 | 0.024276 | 23.2146 | -0.06932 | 0.13824 | -4e-06 | -0.000394 | -0.000394 | -14.1318
2019-04-09 15:22 | 1554837736.7027 | -222 | 0.20955 | 0.000449 | 2e-06 | 100.7774 | 101.4355 | 122.2707 | 20.6896 | 0.001952 | 0.001764 | -154.5692 | 0.23273 | 8.4025 | 0.002602 | 0.004019 | 0.003866 | 3.4e-05 | 0.001368 | 0.28998 | 1.3699 | 2.5475 | 98.47 | 19.99 | 48.361 | 68.2408 | 83.861 | 63.706 | 0.011487 | 1.0481 | 0.02428 | 23.1664 | -0.058215 | 0.13845 | -1e-06 | -0.000183 | -0.000183 | -30.4441
2019-04-09 15:22 | 1554837737.7028 | -221 | 0.20953 | 0.00045 | 2e-06 | 100.7822 | 101.546 | 122.3517 | 20.6905 | 0.001876 | 0.00171 | -154.5677 | 0.23245 | 8.4035 | 0.002675 | 0.00392 | 0.0038 | 3.3e-05 | 0.001363 | 0.29374 | 1.4098 | 2.5477 | 98.47 | 19.99 | 48.36 | 66.42 | 85.6176 | 62.4125 | 0.011487 | 1.0521 | 0.024335 | 23.1295 | -0.067529 | 0.13829 | -3e-06 | -0.000321 | -0.000321 | -17.3828
2019-04-09 15:22 | 1554837738.7028 | -220 | 0.20953 | 0.00045 | 1e-06 | 100.7873 | 101.5621 | 122.4993 | 20.691 | 0.001882 | 0.001733 | -154.5828 | 0.23224 | 8.402 | 0.002673 | 0.00404 | 0.003812 | 3.4e-05 | 0.001369 | 0.3126 | 1.3774 | 2.5476 | 98.47 | 19.99 | 48.359 | 66.5664 | 85.5704 | 62.9638 | 0.011487 | 1.0504 | 0.024329 | 23.1625 | -0.068504 | 0.13816 | -2e-06 | -0.000186 | -0.000186 | -29.897
2019-04-09 15:22 | 1554837739.7018 | -219 | 0.20955 | 0.000447 | 1e-06 | 100.7899 | 101.4506 | 122.6104 | 20.6897 | 0.001924 | 0.001759 | -154.5688 | 0.23224 | 8.4018 | 0.002663 | 0.003998 | 0.003842 | 3.7e-05 | 0.001363 | 0.29662 | 1.3517 | 2.5477 | 98.47 | 19.99 | 48.357 | 67.5614 | 85.3351 | 63.606 | 0.011486 | 1.0484 | 0.024317 | 23.1951 | -0.058441 | 0.13816 | 1e-06 | 2.3e-05 | 2.3e-05 | 246.0861
2019-04-09 15:22 | 1554837741.4436 | -218 | 0.20952 | 0.000447 | 2e-06 | 100.7768 | 101.5026 | 122.212 | 20.6909 | 0.001943 | 0.001777 | -154.5807 | 0.23286 | 8.4034 | 0.002681 | 0.003966 | 0.003829 | 3.6e-05 | 0.001366 | 0.29407 | 1.3325 | 2.5476 | 98.47 | 19.99 | 48.356 | 68.0321 | 85.7522 | 64.0336 | 0.011486 | 1.047 | 0.024262 | 23.1721 | -0.064184 | 0.13853 | 2e-06 | 1.3e-05 | 1.3e-05 | 427.1832
2019-04-09 15:22 | 1554837741.7008 | -217 | 0.20957 | 0.000451 | 2e-06 | 100.7825 | 101.5924 | 122.1158 | 20.6903 | 0.001859 | 0.001833 | -154.5871 | 0.23228 | 8.402 | 0.002662 | 0.004014 | 0.003952 | 3.7e-05 | 0.001373 | 0.27367 | 1.3317 | 2.5477 | 98.47 | 19.99 | 48.356 | 66.0064 | 85.2927 | 65.3773 | 0.011486 | 1.0429 | 0.024204 | 23.209 | -0.071594 | 0.13818 | -4e-06 | -0.000384 | -0.000384 | -14.496
2019-04-09 15:22 | 1554837742.7013 | -216 | 0.20954 | 0.000452 | 2e-06 | 100.8379 | 101.3952 | 122.1212 | 20.6899 | 0.001858 | 0.00177 | -154.5918 | 0.23258 | 8.4022 | 0.002656 | 0.003653 | 0.003418 | 3.6e-05 | 0.001373 | 0.29551 | 1.3443 | 2.5476 | 98.47 | 19.99 | 48.355 | 65.9761 | 85.1698 | 63.8615 | 0.011486 | 1.0476 | 0.024259 | 23.1575 | -0.0493 | 0.13836 | -5e-06 | -0.00046 | -0.00046 | -12.1249
2019-04-09 15:22 | 1554837743.7009 | -215 | 0.20955 | 0.00045 | 2e-06 | 100.7795 | 101.4696 | 122.1123 | 20.6901 | 0.001843 | 0.001766 | -154.5909 | 0.23177 | 8.4001 | 0.002681 | 0.003838 | 0.00369 | 3.3e-05 | 0.001379 | 0.28479 | 1.3261 | 2.5476 | 98.47 | 19.99 | 48.354 | 65.6198 | 85.7483 | 63.7664 | 0.011486 | 1.0479 | 0.024262 | 23.1534 | -0.061034 | 0.13788 | -2e-06 | -0.000265 | -0.000265 | -20.9511
2019-04-09 15:22 | 1554837744.7013 | -214 | 0.20954 | 0.000447 | 2e-06 | 100.8233 | 101.559 | 122.0527 | 20.6896 | 0.001918 | 0.001773 | -154.6053 | 0.23272 | 8.401 | 0.002633 | 0.003694 | 0.003482 | 3.4e-05 | 0.001379 | 0.30397 | 1.4099 | 2.5477 | 98.47 | 19.99 | 48.354 | 67.4241 | 84.6101 | 63.9291 | 0.011486 | 1.0474 | 0.02425 | 23.1534 | -0.065034 | 0.13845 | 2e-06 | 3.7e-05 | 3.7e-05 | 148.8227
2019-04-09 15:22 | 1554837745.701 | -213 | 0.20957 | 0.000451 | 2e-06 | 100.7995 | 101.474 | 121.9822 | 20.6904 | 0.001955 | 0.001729 | -154.5933 | 0.23265 | 8.401 | 0.002673 | 0.004362 | 0.004269 | 3.5e-05 | 0.001384 | 0.30287 | 1.4101 | 2.5477 | 98.47 | 19.99 | 48.356 | 68.3187 | 85.5667 | 62.8874 | 0.011486 | 1.0506 | 0.024281 | 23.1109 | -0.059658 | 0.13841 | -3e-06 | -0.000361 | -0.000361 | -15.4311
2019-04-09 15:22 | 1554837746.702 | -212 | 0.20958 | 0.000448 | 2e-06 | 100.8158 | 101.5659 | 122.3254 | 20.6902 | 0.001968 | 0.001815 | -154.5887 | 0.2323 | 8.4036 | 0.002597 | 0.003584 | 0.003388 | 3.8e-05 | 0.001379 | 0.28584 | 1.3967 | 2.5478 | 98.47 | 19.99 | 48.356 | 68.6337 | 83.732 | 64.9415 | 0.011486 | 1.0442 | 0.024241 | 23.214 | -0.066301 | 0.1382 | 0 | -8.1e-05 | -8.1e-05 | -68.577
2019-04-09 15:22 | 1554837747.7011 | -211 | 0.20955 | 0.000444 | 2e-06 | 100.7685 | 101.4348 | 122.3935 | 20.6905 | 0.001922 | 0.001784 | -154.6032 | 0.2325 | 8.402 | 0.002683 | 0.004391 | 0.004266 | 3.6e-05 | 0.001364 | 0.29679 | 1.4149 | 2.5477 | 98.47 | 19.99 | 48.355 | 67.5163 | 85.8134 | 64.2072 | 0.011486 | 1.0465 | 0.024274 | 23.1952 | -0.058948 | 0.13832 | 5e-06 | 0.000264 | 0.000264 | 21.104
2019-04-09 15:22 | 1554837748.7021 | -210 | 0.2096 | 0.000445 | 2e-06 | 100.7649 | 101.4079 | 122.1686 | 20.6912 | 0.001875 | 0.001775 | -154.5851 | 0.23202 | 8.4012 | 0.002656 | 0.004513 | 0.004329 | 4e-05 | 0.001355 | 0.29502 | 1.4316 | 2.5477 | 98.47 | 19.99 | 48.355 | 66.3904 | 85.1656 | 63.9918 | 0.011486 | 1.0472 | 0.024259 | 23.1665 | -0.056897 | 0.13803 | 3e-06 | 9.2e-05 | 9.2e-05 | 60.4189
2019-04-09 15:22 | 1554837749.7025 | -209 | 0.20957 | 0.000442 | 2e-06 | 100.8114 | 101.7297 | 121.8266 | 20.6905 | 0.001874 | 0.001793 | -154.5934 | 0.23211 | 8.4017 | 0.002661 | 0.003603 | 0.003502 | 4.1e-05 | 0.001359 | 0.26698 | 1.3959 | 2.5476 | 98.47 | 19.99 | 48.354 | 66.3733 | 85.2678 | 64.4086 | 0.011486 | 1.0459 | 0.02421 | 23.1484 | -0.081112 | 0.13808 | 8e-06 | 0.000457 | 0.000457 | 12.1789
2019-04-09 15:22 | 1554837750.7015 | -208 | 0.20957 | 0.000446 | 2e-06 | 100.7435 | 101.5222 | 121.9853 | 20.6895 | 0.001944 | 0.001765 | -154.5754 | 0.23147 | 8.4044 | 0.002625 | 0.0037 | 0.003467 | 3.6e-05 | 0.00135 | 0.3027 | 1.3734 | 2.5475 | 98.47 | 19.99 | 48.353 | 68.0401 | 84.4241 | 63.7386 | 0.011486 | 1.048 | 0.02425 | 23.1404 | -0.068873 | 0.1377 | 3e-06 | 8.7e-05 | 8.7e-05 | 64.1442
2019-04-09 15:23 | 1554837751.7022 | -207 | 0.20955 | 0.000442 | 1e-06 | 100.7792 | 101.4152 | 122.3266 | 20.6906 | 0.001915 | 0.001811 | -154.568 | 0.2318 | 8.4033 | 0.00259 | 0.003936 | 0.003732 | 3.9e-05 | 0.001348 | 0.28335 | 1.3653 | 2.5476 | 98.47 | 19.99 | 48.351 | 67.3461 | 83.5638 | 64.8507 | 0.011485 | 1.0445 | 0.024244 | 23.211 | -0.056266 | 0.1379 | 8e-06 | 0.000513 | 0.000513 | 10.8355
2019-04-09 15:23 | 1554837752.7025 | -206 | 0.2096 | 0.000445 | 2e-06 | 100.8063 | 101.5329 | 122.4153 | 20.69 | 0.001797 | 0.001804 | -154.5431 | 0.2323 | 8.402 | 0.002653 | 0.004203 | 0.004144 | 4e-05 | 0.001352 | 0.27887 | 1.3405 | 2.5475 | 98.47 | 20 | 48.346 | 64.5114 | 85.0973 | 64.6902 | 0.011491 | 1.045 | 0.024259 | 23.2139 | -0.064236 | 0.1382 | 3e-06 | 0.000124 | 0.000124 | 44.9667
2019-04-09 15:23 | 1554837753.7013 | -205 | 0.20956 | 0.000445 | 2e-06 | 100.8066 | 101.4998 | 122.5698 | 20.6904 | 0.001847 | 0.001808 | -154.5448 | 0.23195 | 8.4001 | 0.002648 | 0.003656 | 0.003441 | 4.2e-05 | 0.001343 | 0.28313 | 1.3769 | 2.5475 | 98.47 | 20 | 48.346 | 65.7231 | 84.9596 | 64.7692 | 0.011491 | 1.0448 | 0.024271 | 23.2313 | -0.061299 | 0.13799 | 4e-06 | 0.000214 | 0.000214 | 25.9767
2019-04-09 15:23 | 1554837754.7006 | -204 | 0.20959 | 0.000445 | 2e-06 | 100.7867 | 101.7311 | 122.7622 | 20.6901 | 0.001901 | 0.001759 | -154.5378 | 0.23176 | 8.4015 | 0.002642 | 0.00439 | 0.00426 | 4.2e-05 | 0.00134 | 0.30127 | 1.3424 | 2.5474 | 98.47 | 20 | 48.342 | 67.0196 | 84.8102 | 63.5854 | 0.01149 | 1.0484 | 0.024333 | 23.2087 | -0.083426 | 0.13788 | 3e-06 | 0.000143 | 0.000143 | 38.8434
2019-04-09 15:23 | 1554837755.7019 | -203 | 0.20954 | 0.000448 | 2e-06 | 100.8641 | 101.5308 | 122.8634 | 20.6905 | 0.001883 | 0.001759 | -154.5462 | 0.23147 | 8.4025 | 0.002676 | 0.002966 | 0.002782 | 4.2e-05 | 0.001338 | 0.28241 | 1.3355 | 2.5476 | 98.47 | 20 | 48.338 | 66.584 | 85.6433 | 63.5921 | 0.011489 | 1.0484 | 0.024343 | 23.2185 | -0.05893 | 0.1377 | 0 | -9.4e-05 | -9.4e-05 | -58.8202
2019-04-09 15:23 | 1554837756.7017 | -202 | 0.20958 | 0.000451 | 2e-06 | 100.7429 | 101.3725 | 122.6179 | 20.6903 | 0.001862 | 0.001729 | -154.544 | 0.23228 | 8.4009 | 0.00263 | 0.004114 | 0.004018 | 3.9e-05 | 0.001339 | 0.31006 | 1.3146 | 2.5476 | 98.47 | 20 | 48.333 | 66.0672 | 84.5326 | 62.8647 | 0.011488 | 1.0507 | 0.024345 | 23.1702 | -0.055727 | 0.13818 | -3e-06 | -0.000305 | -0.000305 | -18.2706
2019-04-09 15:23 | 1554837757.7008 | -201 | 0.20955 | 0.000447 | 1e-06 | 100.8394 | 101.5659 | 122.705 | 20.6911 | 0.001865 | 0.001811 | -154.5431 | 0.23178 | 8.4034 | 0.002647 | 0.003823 | 0.003643 | 4.3e-05 | 0.001341 | 0.30143 | 1.325 | 2.5476 | 98.47 | 20 | 48.322 | 66.1389 | 84.9516 | 64.8449 | 0.011485 | 1.0445 | 0.024282 | 23.2467 | -0.064214 | 0.13788 | 1e-06 | 1.3e-05 | 1.3e-05 | 415.6354
2019-04-09 15:23 | 1554837758.7014 | -200 | 0.20956 | 0.000445 | 1e-06 | 100.8158 | 101.5797 | 122.935 | 20.6901 | 0.001849 | 0.001744 | -154.5359 | 0.23139 | 8.4028 | 0.002605 | 0.003687 | 0.003574 | 4e-05 | 0.00134 | 0.29192 | 1.3955 | 2.5475 | 98.47 | 20 | 48.315 | 65.7587 | 83.931 | 63.2473 | 0.011484 | 1.0495 | 0.024362 | 23.2134 | -0.067513 | 0.13765 | 5e-06 | 0.000243 | 0.000243 | 22.8468
2019-04-09 15:23 | 1554837759.7015 | -199 | 0.20956 | 0.000448 | 2e-06 | 100.8367 | 101.5666 | 123.2594 | 20.69 | 0.001903 | 0.001763 | -154.5392 | 0.2323 | 8.4004 | 0.002656 | 0.003904 | 0.00368 | 4.1e-05 | 0.00134 | 0.30198 | 1.3312 | 2.5475 | 98.47 | 20 | 48.3 | 67.0586 | 85.1619 | 63.7023 | 0.01148 | 1.0481 | 0.024378 | 23.2597 | -0.064512 | 0.13819 | 0 | -0.000111 | -0.000111 | -50.3413
2019-04-09 15:23 | 1554837760.7015 | -198 | 0.20956 | 0.000452 | 2e-06 | 100.798 | 101.5377 | 123.2277 | 20.6907 | 0.00191 | 0.001774 | -154.5387 | 0.23215 | 8.4018 | 0.002615 | 0.004706 | 0.00456 | 3.8e-05 | 0.00133 | 0.27743 | 1.3335 | 2.5476 | 98.47 | 20 | 48.291 | 67.2301 | 84.1764 | 63.9521 | 0.011478 | 1.0473 | 0.024366 | 23.2653 | -0.0654 | 0.1381 | -5e-06 | -0.000479 | -0.000479 | -11.6082
2019-04-09 15:23 | 1554837761.7005 | -197 | 0.20958 | 0.000453 | 2e-06 | 100.8069 | 101.5232 | 122.7275 | 20.6895 | 0.001857 | 0.001706 | -154.5394 | 0.23218 | 8.4014 | 0.002679 | 0.003719 | 0.003513 | 4e-05 | 0.001326 | 0.29031 | 1.3891 | 2.5477 | 98.47 | 20 | 48.283 | 65.9524 | 85.707 | 62.3092 | 0.011476 | 1.0524 | 0.024376 | 23.1614 | -0.063334 | 0.13812 | -6e-06 | -0.00055 | -0.00055 | -10.1205
2019-04-09 15:23 | 1554837762.7007 | -196 | 0.20954 | 0.00045 | 2e-06 | 100.7575 | 101.4665 | 122.6723 | 20.6903 | 0.001864 | 0.001763 | -154.562 | 0.23138 | 8.4014 | 0.002674 | 0.004309 | 0.004174 | 4e-05 | 0.001332 | 0.28456 | 1.3944 | 2.5476 | 98.47 | 20 | 48.266 | 66.1242 | 85.5837 | 63.6901 | 0.011472 | 1.0481 | 0.02432 | 23.2038 | -0.062717 | 0.13765 | -3e-06 | -0.000322 | -0.000322 | -17.2522
2019-04-09 15:23 | 1554837763.701 | -195 | 0.20956 | 0.000452 | 2e-06 | 100.8117 | 101.6492 | 122.4146 | 20.69 | 0.001952 | 0.001766 | -154.5712 | 0.23188 | 8.4026 | 0.002688 | 0.004356 | 0.004159 | 3.7e-05 | 0.001325 | 0.28136 | 1.3635 | 2.5476 | 98.47 | 20 | 48.257 | 68.2287 | 85.9277 | 63.7625 | 0.01147 | 1.0479 | 0.024292 | 23.1819 | -0.074002 | 0.13795 | -5e-06 | -0.000464 | -0.000464 | -11.9929
2019-04-09 15:23 | 1554837764.7016 | -194 | 0.20958 | 0.000448 | 2e-06 | 100.778 | 101.7139 | 122.3832 | 20.6906 | 0.001949 | 0.001797 | -154.5622 | 0.23189 | 8.4001 | 0.002664 | 0.003987 | 0.003818 | 4.1e-05 | 0.001328 | 0.29302 | 1.34 | 2.5476 | 98.47 | 20.01 | 48.248 | 68.1622 | 85.343 | 64.5235 | 0.011475 | 1.0455 | 0.024261 | 23.2051 | -0.08268 | 0.13795 | 0 | -0.000135 | -0.000135 | -41.043
2019-04-09 15:23 | 1554837765.702 | -193 | 0.20955 | 0.000452 | 2e-06 | 100.8013 | 101.4747 | 122.4857 | 20.6905 | 0.001978 | 0.001825 | -154.5731 | 0.2314 | 8.4033 | 0.002702 | 0.003671 | 0.00355 | 4e-05 | 0.001326 | 0.28136 | 1.3362 | 2.5476 | 98.47 | 20 | 48.23 | 68.8657 | 86.2756 | 65.18 | 0.011463 | 1.0435 | 0.024248 | 23.2374 | -0.05956 | 0.13766 | -5e-06 | -0.000509 | -0.000509 | -10.8915
2019-04-09 15:23 | 1554837766.7006 | -192 | 0.20956 | 0.000449 | 2e-06 | 100.7271 | 101.4097 | 122.57 | 20.6908 | 0.001855 | 0.001852 | -154.5667 | 0.23152 | 8.4024 | 0.002689 | 0.004197 | 0.003958 | 3.9e-05 | 0.001326 | 0.28567 | 1.3504 | 2.5475 | 98.47 | 20.01 | 48.221 | 65.906 | 85.9489 | 65.842 | 0.011468 | 1.0415 | 0.024233 | 23.2681 | -0.060407 | 0.13773 | -2e-06 | -0.000274 | -0.000274 | -20.2426
2019-04-09 15:23 | 1554837767.7021 | -191 | 0.20957 | 0.00045 | 2e-06 | 100.798 | 101.4854 | 122.6477 | 20.69 | 0.001812 | 0.001736 | -154.5619 | 0.23221 | 8.4026 | 0.002657 | 0.004245 | 0.004114 | 4e-05 | 0.001328 | 0.28766 | 1.3459 | 2.5476 | 98.47 | 20.01 | 48.208 | 64.8791 | 85.1878 | 63.0426 | 0.011465 | 1.0501 | 0.024341 | 23.1792 | -0.060791 | 0.13814 | -2e-06 | -0.000289 | -0.000289 | -19.2332
2019-04-09 15:23 | 1554837768.7008 | -190 | 0.20956 | 0.000447 | 2e-06 | 100.781 | 101.4169 | 122.3012 | 20.6905 | 0.001873 | 0.001759 | -154.5697 | 0.2323 | 8.4024 | 0.002649 | 0.003791 | 0.003627 | 4e-05 | 0.001329 | 0.29839 | 1.3522 | 2.5476 | 98.47 | 20.01 | 48.2 | 66.3334 | 85.0025 | 63.5928 | 0.011463 | 1.0484 | 0.024287 | 23.1654 | -0.056259 | 0.1382 | 2e-06 | 5e-06 | 5e-06 | 1145.5617
2019-04-09 15:23 | 1554837769.7015 | -189 | 0.20955 | 0.000449 | 2e-06 | 100.7771 | 101.5725 | 122.2402 | 20.6899 | 0.001932 | 0.001786 | -154.5706 | 0.23154 | 8.4021 | 0.002623 | 0.003799 | 0.003604 | 4e-05 | 0.001321 | 0.29015 | 1.4112 | 2.5475 | 98.47 | 20.01 | 48.195 | 67.7638 | 84.3653 | 64.2488 | 0.011462 | 1.0464 | 0.024257 | 23.1821 | -0.070312 | 0.13774 | -1e-06 | -0.000179 | -0.000179 | -31.0763
2019-04-09 15:23 | 1554837770.7009 | -188 | 0.20953 | 0.000451 | 2e-06 | 100.7155 | 101.5776 | 122.2559 | 20.6899 | 0.001921 | 0.001854 | -154.5839 | 0.23153 | 8.4025 | 0.002674 | 0.004012 | 0.003824 | 4e-05 | 0.001319 | 0.27256 | 1.3475 | 2.5476 | 98.47 | 20.01 | 48.188 | 67.4931 | 85.6014 | 65.877 | 0.011461 | 1.0413 | 0.0242 | 23.2395 | -0.076239 | 0.13774 | -4e-06 | -0.000361 | -0.000361 | -15.3824
2019-04-09 15:23 | 1554837771.7009 | -187 | 0.20949 | 0.000448 | 2e-06 | 100.7658 | 101.5831 | 122.5975 | 20.6908 | 0.001934 | 0.001808 | -154.5849 | 0.2322 | 8.4001 | 0.002603 | 0.003972 | 0.003775 | 4.1e-05 | 0.001319 | 0.31636 | 1.3713 | 2.5476 | 98.47 | 20.01 | 48.185 | 67.7957 | 83.8879 | 64.7845 | 0.01146 | 1.0447 | 0.024273 | 23.2344 | -0.072256 | 0.13814 | 0 | -8.9e-05 | -8.9e-05 | -62.7218
2019-04-09 15:23 | 1554837772.7002 | -186 | 0.2095 | 0.00045 | 2e-06 | 100.7759 | 101.4723 | 122.4228 | 20.6905 | 0.001926 | 0.001764 | -154.5691 | 0.23137 | 8.4005 | 0.002682 | 0.003806 | 0.003695 | 4.3e-05 | 0.001315 | 0.30973 | 1.3266 | 2.5476 | 98.47 | 20.01 | 48.184 | 67.606 | 85.787 | 63.7247 | 0.01146 | 1.048 | 0.024294 | 23.1814 | -0.061594 | 0.13764 | -2e-06 | -0.000232 | -0.000232 | -23.8681
2019-04-09 15:23 | 1554837773.7009 | -185 | 0.20948 | 0.000446 | 2e-06 | 100.7646 | 101.5263 | 122.4109 | 20.6896 | 0.001866 | 0.001771 | -154.569 | 0.23169 | 8.4004 | 0.002623 | 0.004065 | 0.003903 | 4.2e-05 | 0.001315 | 0.29374 | 1.3145 | 2.5476 | 98.47 | 20.01 | 48.183 | 66.1608 | 84.3626 | 63.8839 | 0.011459 | 1.0475 | 0.024287 | 23.1858 | -0.06736 | 0.13783 | 3e-06 | 0.000121 | 0.000121 | 46.0732
2019-04-09 15:23 | 1554837774.7011 | -184 | 0.2095 | 0.000446 | 2e-06 | 100.862 | 101.4513 | 122.2658 | 20.6906 | 0.001923 | 0.00178 | -154.5832 | 0.23161 | 8.4014 | 0.002669 | 0.004102 | 0.003887 | 4.2e-05 | 0.001316 | 0.29623 | 1.3113 | 2.5476 | 98.47 | 20.01 | 48.185 | 67.5383 | 85.4609 | 64.1089 | 0.01146 | 1.0468 | 0.024265 | 23.1797 | -0.052108 | 0.13778 | 3e-06 | 0.000126 | 0.000126 | 44.1427
2019-04-09 15:23 | 1554837775.7004 | -183 | 0.2095 | 0.000449 | 2e-06 | 100.7578 | 101.6365 | 122.0459 | 20.6905 | 0.001869 | 0.001756 | -154.5627 | 0.23202 | 8.4011 | 0.002652 | 0.00392 | 0.003757 | 4.2e-05 | 0.001322 | 0.30326 | 1.3347 | 2.5476 | 98.47 | 20.01 | 48.188 | 66.2398 | 85.072 | 63.5162 | 0.011461 | 1.0487 | 0.024264 | 23.1385 | -0.077665 | 0.13803 | -2e-06 | -0.000222 | -0.000222 | -25.0925
2019-04-09 15:23 | 1554837776.7001 | -182 | 0.20949 | 0.000452 | 1e-06 | 100.7116 | 101.622 | 122.2456 | 20.6895 | 0.001889 | 0.001791 | -154.567 | 0.23225 | 8.4006 | 0.002631 | 0.003685 | 0.003539 | 4.2e-05 | 0.001326 | 0.29103 | 1.3906 | 2.5476 | 98.47 | 20.01 | 48.189 | 66.7266 | 84.5646 | 64.3774 | 0.011461 | 1.046 | 0.024253 | 23.1871 | -0.080492 | 0.13817 | -5e-06 | -0.000462 | -0.000462 | -12.0417
2019-04-09 15:23 | 1554837777.7011 | -181 | 0.20952 | 0.000455 | 2e-06 | 100.8578 | 101.4885 | 122.1691 | 20.6901 | 0.001938 | 0.001775 | -154.5508 | 0.23256 | 8.4013 | 0.00266 | 0.003961 | 0.003745 | 4.3e-05 | 0.001331 | 0.29142 | 1.3731 | 2.5476 | 98.47 | 20.01 | 48.19 | 67.9066 | 85.2473 | 63.9766 | 0.011461 | 1.0472 | 0.02426 | 23.166 | -0.055754 | 0.13835 | -8e-06 | -0.000722 | -0.000722 | -7.7215
2019-04-09 15:23 | 1554837778.7001 | -180 | 0.20952 | 0.000453 | 1e-06 | 100.7974 | 101.6716 | 122.467 | 20.69 | 0.001946 | 0.001722 | -154.553 | 0.23131 | 8.4029 | 0.002719 | 0.003976 | 0.003767 | 4.2e-05 | 0.001328 | 0.30093 | 1.3936 | 2.5477 | 98.47 | 20.01 | 48.193 | 68.0868 | 86.6786 | 62.6973 | 0.011462 | 1.0512 | 0.024336 | 23.1502 | -0.077239 | 0.13761 | -6e-06 | -0.000497 | -0.000497 | -11.1646
2019-04-09 15:23 | 1554837779.7011 | -179 | 0.20955 | 0.000449 | 1e-06 | 100.8367 | 101.5377 | 122.109 | 20.6898 | 0.001896 | 0.001804 | -154.5417 | 0.23295 | 8.4021 | 0.002646 | 0.003984 | 0.00386 | 4.6e-05 | 0.001333 | 0.27865 | 1.3779 | 2.5476 | 98.47 | 20.01 | 48.194 | 66.9046 | 84.916 | 64.6914 | 0.011462 | 1.045 | 0.024228 | 23.1849 | -0.061966 | 0.13858 | -1e-06 | -0.000148 | -0.000148 | -37.7959
2019-04-09 15:23 | 1554837780.7012 | -178 | 0.20956 | 0.000451 | 1e-06 | 100.8662 | 101.5239 | 122.0991 | 20.6906 | 0.001987 | 0.001782 | -154.5408 | 0.23236 | 8.401 | 0.002672 | 0.003702 | 0.003555 | 4.4e-05 | 0.001332 | 0.29175 | 1.3813 | 2.5475 | 98.47 | 20.01 | 48.199 | 69.0732 | 85.5443 | 64.1512 | 0.011463 | 1.0467 | 0.024247 | 23.1654 | -0.058138 | 0.13823 | -3e-06 | -0.000285 | -0.000285 | -19.5792
2019-04-09 15:23 | 1554837781.7012 | -177 | 0.20955 | 0.000454 | 2e-06 | 100.856 | 101.5308 | 122.086 | 20.6896 | 0.001932 | 0.001853 | -154.5388 | 0.2332 | 8.4016 | 0.002657 | 0.003805 | 0.003618 | 4.2e-05 | 0.001332 | 0.30773 | 1.3625 | 2.5475 | 98.47 | 20.01 | 48.201 | 67.7501 | 85.1897 | 65.8503 | 0.011464 | 1.0414 | 0.024185 | 23.2224 | -0.059643 | 0.13873 | -8e-06 | -0.000704 | -0.000704 | -7.9374
2019-04-09 15:23 | 1554837782.7018 | -176 | 0.20951 | 0.000451 | 2e-06 | 100.8543 | 101.633 | 122.2667 | 20.6902 | 0.001925 | 0.001854 | -154.5549 | 0.23152 | 8.4023 | 0.002688 | 0.003794 | 0.003692 | 4.2e-05 | 0.001328 | 0.30165 | 1.3542 | 2.5476 | 98.47 | 20.01 | 48.203 | 67.5987 | 85.931 | 65.8797 | 0.011464 | 1.0413 | 0.024201 | 23.2406 | -0.068801 | 0.13773 | -3e-06 | -0.000346 | -0.000346 | -16.0205
2019-04-09 15:23 | 1554837783.7012 | -175 | 0.20951 | 0.000453 | 2e-06 | 100.8846 | 101.8096 | 122.3104 | 20.6897 | 0.001892 | 0.001768 | -154.5486 | 0.23245 | 8.4004 | 0.002614 | 0.004524 | 0.00438 | 4.4e-05 | 0.001315 | 0.29839 | 1.3512 | 2.5476 | 98.47 | 20.01 | 48.205 | 66.8086 | 84.1427 | 63.8084 | 0.011465 | 1.0477 | 0.02428 | 23.1736 | -0.081632 | 0.13829 | -7e-06 | -0.000588 | -0.000588 | -9.476
2019-04-09 15:23 | 1554837784.7013 | -174 | 0.2095 | 0.000456 | 1e-06 | 100.8614 | 101.5621 | 122.3679 | 20.6896 | 0.001896 | 0.001792 | -154.549 | 0.23196 | 8.4041 | 0.002719 | 0.003666 | 0.003458 | 4.4e-05 | 0.001317 | 0.29192 | 1.3845 | 2.5476 | 98.47 | 20.01 | 48.207 | 66.8967 | 86.6813 | 64.4015 | 0.011465 | 1.0459 | 0.024264 | 23.1995 | -0.061926 | 0.13799 | -9e-06 | -0.000709 | -0.000709 | -7.8398
2019-04-09 15:23 | 1554837785.7007 | -173 | 0.20953 | 0.000458 | 2e-06 | 100.9046 | 101.6915 | 122.4585 | 20.6897 | 0.00186 | 0.001762 | -154.5605 | 0.23182 | 8.4035 | 0.002651 | 0.004214 | 0.004031 | 4.5e-05 | 0.001318 | 0.30309 | 1.341 | 2.5475 | 98.47 | 20.01 | 48.207 | 66.0361 | 85.0494 | 63.6739 | 0.011465 | 1.0482 | 0.0243 | 23.183 | -0.069487 | 0.13791 | -1.3e-05 | -0.001048 | -0.001048 | -5.3039
2019-04-09 15:23 | 1554837786.7003 | -172 | 0.20952 | 0.000458 | 2e-06 | 100.856 | 101.7769 | 122.1325 | 20.6899 | 0.001918 | 0.001812 | -154.5654 | 0.23252 | 8.4011 | 0.002656 | 0.003559 | 0.003455 | 4.2e-05 | 0.001322 | 0.29142 | 1.3696 | 2.5476 | 98.47 | 20.01 | 48.207 | 67.4156 | 85.1628 | 64.8726 | 0.011465 | 1.0444 | 0.024224 | 23.1933 | -0.081294 | 0.13833 | -1.3e-05 | -0.001054 | -0.001054 | -5.2867
2019-04-09 15:23 | 1554837787.7007 | -171 | 0.20953 | 0.000462 | 1e-06 | 100.8855 | 101.5429 | 122.0158 | 20.6899 | 0.001907 | 0.001791 | -154.5813 | 0.23289 | 8.401 | 0.002659 | 0.003695 | 0.00348 | 4.3e-05 | 0.001316 | 0.29159 | 1.3492 | 2.5474 | 98.47 | 20.01 | 48.207 | 67.1535 | 85.2415 | 64.3635 | 0.011465 | 1.046 | 0.024231 | 23.1648 | -0.058089 | 0.13855 | -1.7e-05 | -0.001316 | -0.001316 | -4.2428
2019-04-09 15:23 | 1554837788.6994 | -170 | 0.2095 | 0.000462 | 2e-06 | 100.8811 | 101.5363 | 122.0405 | 20.6905 | 0.001905 | 0.001831 | -154.5776 | 0.23261 | 8.4028 | 0.002697 | 0.003997 | 0.003817 | 4.1e-05 | 0.001317 | 0.29966 | 1.3928 | 2.5477 | 98.47 | 20.01 | 48.207 | 67.1065 | 86.1334 | 65.3383 | 0.011465 | 1.043 | 0.024198 | 23.2005 | -0.057909 | 0.13838 | -1.8e-05 | -0.001411 | -0.001411 | -3.9534
2019-04-09 15:23 | 1554837789.7008 | -169 | 0.20948 | 0.000467 | 2e-06 | 100.8933 | 101.6557 | 122.1282 | 20.6896 | 0.001861 | 0.001808 | -154.5804 | 0.23193 | 8.4025 | 0.002653 | 0.003647 | 0.003638 | 3.9e-05 | 0.001314 | 0.28208 | 1.3266 | 2.5475 | 98.47 | 20.01 | 48.207 | 66.0398 | 85.0892 | 64.7747 | 0.011465 | 1.0447 | 0.024227 | 23.1896 | -0.067341 | 0.13798 | -2.3e-05 | -0.001794 | -0.001794 | -3.1003
2019-04-09 15:23 | 1554837790.7005 | -168 | 0.20953 | 0.000465 | 2e-06 | 100.8927 | 101.7105 | 122.0409 | 20.6895 | 0.001875 | 0.001837 | -154.585 | 0.23241 | 8.4024 | 0.002659 | 0.004036 | 0.00375 | 3.8e-05 | 0.001315 | 0.27616 | 1.329 | 2.5476 | 98.47 | 20.01 | 48.207 | 66.3925 | 85.2396 | 65.4739 | 0.011465 | 1.0426 | 0.024194 | 23.2052 | -0.072207 | 0.13826 | -2.2e-05 | -0.001686 | -0.001686 | -3.3051
2019-04-09 15:23 | 1554837791.7005 | -167 | 0.20947 | 0.000462 | 2e-06 | 101.0201 | 101.7373 | 122.2414 | 20.6902 | 0.001859 | 0.001779 | -154.5729 | 0.23209 | 8.3993 | 0.00266 | 0.004471 | 0.004326 | 4.5e-05 | 0.001321 | 0.28385 | 1.3507 | 2.5475 | 98.47 | 20.01 | 48.206 | 66.0046 | 85.2637 | 64.0907 | 0.011465 | 1.0469 | 0.024263 | 23.1768 | -0.063275 | 0.13807 | -1.7e-05 | -0.00132 | -0.00132 | -4.2159
2019-04-09 15:23 | 1554837792.6995 | -166 | 0.20954 | 0.000467 | 2e-06 | 100.9061 | 101.6423 | 122.4296 | 20.6903 | 0.001895 | 0.00179 | -154.5776 | 0.23286 | 8.4026 | 0.002637 | 0.003757 | 0.003604 | 4.2e-05 | 0.001315 | 0.27721 | 1.3327 | 2.5476 | 98.47 | 20.01 | 48.204 | 66.8616 | 84.7027 | 64.3351 | 0.011464 | 1.0461 | 0.024273 | 23.203 | -0.065025 | 0.13853 | -2.5e-05 | -0.00197 | -0.00197 | -2.8343
2019-04-09 15:23 | 1554837793.7006 | -165 | 0.20949 | 0.000465 | 2e-06 | 100.9335 | 101.6303 | 122.5728 | 20.6899 | 0.001925 | 0.001768 | -154.5677 | 0.23189 | 8.4034 | 0.002713 | 0.003549 | 0.003351 | 4.1e-05 | 0.001313 | 0.27721 | 1.3886 | 2.5477 | 98.47 | 20.01 | 48.2 | 67.5953 | 86.5404 | 63.8202 | 0.011463 | 1.0477 | 0.024306 | 23.1989 | -0.061537 | 0.13795 | -2.2e-05 | -0.001688 | -0.001688 | -3.2942
2019-04-09 15:23 | 1554837794.7006 | -164 | 0.20954 | 0.000467 | 1e-06 | 100.9671 | 101.8787 | 122.5864 | 20.6901 | 0.001951 | 0.001735 | -154.5496 | 0.23204 | 8.4009 | 0.002718 | 0.004371 | 0.004205 | 4.1e-05 | 0.001312 | 0.30165 | 1.3829 | 2.5477 | 98.47 | 20.01 | 48.198 | 68.2141 | 86.6445 | 63.0152 | 0.011463 | 1.0502 | 0.024336 | 23.1725 | -0.080397 | 0.13804 | -2.3e-05 | -0.001778 | -0.001778 | -3.1292
2019-04-09 15:23 | 1554837795.701 | -163 | 0.2095 | 0.000464 | 2e-06 | 100.9478 | 101.7056 | 122.7132 | 20.6903 | 0.001949 | 0.001749 | -154.5444 | 0.23261 | 8.4009 | 0.002641 | 0.004091 | 0.003923 | 4.3e-05 | 0.001315 | 0.30663 | 1.3836 | 2.5475 | 98.47 | 20.01 | 48.196 | 68.1674 | 84.7947 | 63.3536 | 0.011462 | 1.0492 | 0.024336 | 23.1961 | -0.066901 | 0.13838 | -2e-05 | -0.001544 | -0.001544 | -3.6107
2019-04-09 15:23 | 1554837796.7653 | -162 | 0.20953 | 0.000462 | 2e-06 | 100.9627 | 101.8867 | 122.7329 | 20.6898 | 0.001883 | 0.001811 | -154.5326 | 0.23327 | 8.4001 | 0.002664 | 0.004199 | 0.004091 | 4e-05 | 0.001309 | 0.2808 | 1.3867 | 2.5476 | 98.47 | 20.01 | 48.192 | 66.5708 | 85.3422 | 64.8545 | 0.011462 | 1.0445 | 0.024284 | 23.2496 | -0.081488 | 0.13877 | -1.7e-05 | -0.001346 | -0.001346 | -4.1563
2019-04-09 15:23 | 1554837797.7134 | -161 | 0.20952 | 0.000461 | 1e-06 | 100.9856 | 101.7142 | 122.7242 | 20.6899 | 0.001964 | 0.001728 | -154.53 | 0.23266 | 8.4019 | 0.00262 | 0.003799 | 0.003605 | 3.9e-05 | 0.001321 | 0.29214 | 1.3644 | 2.5475 | 98.47 | 20.01 | 48.19 | 68.5288 | 84.2951 | 62.8581 | 0.011461 | 1.0507 | 0.024355 | 23.1801 | -0.064307 | 0.13841 | -1.6e-05 | -0.001225 | -0.001225 | -4.5547
2019-04-09 15:23 | 1554837798.7145 | -160 | 0.20952 | 0.000467 | 1e-06 | 100.9451 | 101.7019 | 122.7024 | 20.6905 | 0.001901 | 0.00184 | -154.5217 | 0.23189 | 8.4002 | 0.002698 | 0.00437 | 0.004188 | 4.1e-05 | 0.001326 | 0.29142 | 1.3554 | 2.5477 | 98.47 | 20.01 | 48.186 | 67.024 | 86.1615 | 65.542 | 0.01146 | 1.0424 | 0.024257 | 23.2704 | -0.066805 | 0.13795 | -2.4e-05 | -0.001841 | -0.001841 | -3.019
2019-04-09 15:23 | 1554837799.7131 | -159 | 0.20951 | 0.000466 | 2e-06 | 100.9874 | 101.8736 | 122.8667 | 20.689 | 0.001963 | 0.001751 | -154.5208 | 0.23258 | 8.4016 | 0.002687 | 0.004389 | 0.004314 | 4.4e-05 | 0.00133 | 0.27671 | 1.3514 | 2.5475 | 98.47 | 20 | 48.185 | 68.5133 | 85.894 | 63.3983 | 0.011453 | 1.049 | 0.02435 | 23.2121 | -0.07815 | 0.13836 | -2.2e-05 | -0.001726 | -0.001726 | -3.2311
2019-04-09 15:23 | 1554837800.7122 | -158 | 0.20951 | 0.000454 | 1e-06 | 100.9743 | 101.6268 | 122.9268 | 20.6906 | 0.001935 | 0.001767 | -154.514 | 0.23226 | 8.4023 | 0.002718 | 0.004258 | 0.004126 | 4.2e-05 | 0.001333 | 0.262 | 1.343 | 2.5477 | 98.47 | 20 | 48.182 | 67.826 | 86.6544 | 63.7919 | 0.011452 | 1.0478 | 0.024342 | 23.2314 | -0.057619 | 0.13817 | -8e-06 | -0.000643 | -0.000643 | -8.655
2019-04-09 15:23 | 1554837801.7122 | -157 | 0.20952 | 0.000457 | 2e-06 | 101.0079 | 101.8003 | 122.5362 | 20.6899 | 0.001903 | 0.001786 | -154.518 | 0.2318 | 8.4008 | 0.002687 | 0.003639 | 0.003458 | 4.3e-05 | 0.001337 | 0.27367 | 1.3073 | 2.5474 | 98.47 | 20.01 | 48.178 | 67.0733 | 85.9057 | 64.2466 | 0.011458 | 1.0464 | 0.024287 | 23.2101 | -0.069892 | 0.13789 | -1.1e-05 | -0.000896 | -0.000896 | -6.1999
2019-04-09 15:23 | 1554837802.7132 | -156 | 0.20957 | 0.00046 | 1e-06 | 100.9936 | 101.7641 | 122.42 | 20.6907 | 0.0019 | 0.00174 | -154.5171 | 0.23168 | 8.4029 | 0.002703 | 0.004433 | 0.004279 | 4.5e-05 | 0.001332 | 0.30143 | 1.3349 | 2.5477 | 98.47 | 20 | 48.177 | 66.9983 | 86.2991 | 63.1323 | 0.011451 | 1.0498 | 0.024315 | 23.1608 | -0.067982 | 0.13783 | -1.5e-05 | -0.001193 | -0.001193 | -4.6549
2019-04-09 15:23 | 1554837803.7136 | -155 | 0.20952 | 0.000458 | 1e-06 | 100.9731 | 101.8512 | 122.4911 | 20.6906 | 0.001911 | 0.00178 | -154.5285 | 0.2318 | 8.4021 | 0.002779 | 0.004274 | 0.004102 | 4.3e-05 | 0.001331 | 0.27489 | 1.3236 | 2.5475 | 98.47 | 20.01 | 48.178 | 67.2463 | 88.1097 | 64.1137 | 0.011458 | 1.0468 | 0.024287 | 23.2013 | -0.077452 | 0.1379 | -1.3e-05 | -0.001008 | -0.001008 | -5.5141
2019-04-09 15:23 | 1554837804.7128 | -154 | 0.20953 | 0.000454 | 2e-06 | 101.0416 | 101.7927 | 122.5261 | 20.6902 | 0.001907 | 0.001834 | -154.5364 | 0.23197 | 8.4018 | 0.002691 | 0.003838 | 0.003705 | 4.4e-05 | 0.00133 | 0.29911 | 1.3393 | 2.5475 | 98.47 | 20.01 | 48.18 | 67.161 | 85.9957 | 65.4137 | 0.011459 | 1.0428 | 0.024244 | 23.2493 | -0.066247 | 0.138 | -8e-06 | -0.00069 | -0.00069 | -8.0623
2019-04-09 15:23 | 1554837805.7137 | -153 | 0.20955 | 0.000451 | 1e-06 | 101.0082 | 101.7483 | 122.317 | 20.6904 | 0.001905 | 0.001758 | -154.5403 | 0.23169 | 8.4023 | 0.002723 | 0.004349 | 0.004212 | 4.4e-05 | 0.001331 | 0.29175 | 1.353 | 2.5476 | 98.47 | 20.01 | 48.189 | 67.1182 | 86.7617 | 63.5833 | 0.011461 | 1.0484 | 0.024289 | 23.1665 | -0.065298 | 0.13783 | -3e-06 | -0.000296 | -0.000296 | -18.7952
2019-04-09 15:23 | 1554837806.7134 | -152 | 0.20949 | 0.000455 | 1e-06 | 101.0044 | 101.6981 | 122.3881 | 20.69 | 0.001926 | 0.001741 | -154.5544 | 0.23124 | 8.4009 | 0.002674 | 0.004653 | 0.004463 | 4.5e-05 | 0.001335 | 0.30021 | 1.3577 | 2.5475 | 98.47 | 20 | 48.196 | 67.6149 | 85.5813 | 63.1699 | 0.011455 | 1.0497 | 0.024311 | 23.159 | -0.061223 | 0.13756 | -8e-06 | -0.000661 | -0.000661 | -8.3901
2019-04-09 15:23 | 1554837807.7137 | -151 | 0.2095 | 0.000453 | 2e-06 | 100.9966 | 101.7727 | 122.482 | 20.6904 | 0.001926 | 0.001785 | -154.5537 | 0.23215 | 8.4005 | 0.002675 | 0.004379 | 0.004169 | 4.5e-05 | 0.00134 | 0.27848 | 1.3512 | 2.5476 | 98.47 | 20 | 48.218 | 67.6051 | 85.6042 | 64.2244 | 0.011461 | 1.0465 | 0.024282 | 23.2042 | -0.068474 | 0.13811 | -6e-06 | -0.000585 | -0.000585 | -9.5195
2019-04-09 15:23 | 1554837808.7134 | -150 | 0.20953 | 0.000449 | 2e-06 | 100.9394 | 101.6736 | 122.3468 | 20.6901 | 0.001959 | 0.00179 | -154.5672 | 0.23187 | 8.4001 | 0.002731 | 0.004033 | 0.00386 | 4.3e-05 | 0.001344 | 0.29336 | 1.3383 | 2.5476 | 98.47 | 20 | 48.231 | 68.4028 | 86.9606 | 64.3449 | 0.011464 | 1.0461 | 0.024264 | 23.1955 | -0.064825 | 0.13794 | -2e-06 | -0.000248 | -0.000248 | -22.4366
2019-04-09 15:23 | 1554837809.7124 | -149 | 0.2095 | 0.000454 | 2e-06 | 100.9531 | 101.7101 | 122.3472 | 20.6896 | 0.001955 | 0.001761 | -154.5704 | 0.2313 | 8.4022 | 0.002619 | 0.004201 | 0.004038 | 4.2e-05 | 0.001348 | 0.28456 | 1.3701 | 2.5477 | 98.47 | 20 | 48.245 | 68.3189 | 84.2722 | 63.6403 | 0.011467 | 1.0483 | 0.02429 | 23.1714 | -0.066819 | 0.1376 | -7e-06 | -0.000588 | -0.000588 | -9.4227
2019-04-09 15:24 | 1554837810.7134 | -148 | 0.20948 | 0.00045 | 1e-06 | 101.0347 | 101.7466 | 122.2242 | 20.6906 | 0.001904 | 0.001794 | -154.5728 | 0.23248 | 8.4022 | 0.002697 | 0.00426 | 0.004027 | 4.4e-05 | 0.001348 | 0.28512 | 1.3765 | 2.5477 | 98.47 | 20.01 | 48.277 | 67.0802 | 86.139 | 64.4354 | 0.011482 | 1.0458 | 0.024249 | 23.187 | -0.0628 | 0.1383 | -2e-06 | -0.000251 | -0.000251 | -22.1708
2019-04-09 15:24 | 1554837811.7125 | -147 | 0.20951 | 0.000458 | 2e-06 | 100.9365 | 101.7686 | 122.7428 | 20.6908 | 0.001924 | 0.001729 | -154.5744 | 0.23175 | 8.4029 | 0.002659 | 0.004451 | 0.004271 | 4.5e-05 | 0.001348 | 0.27079 | 1.358 | 2.5475 | 98.47 | 20.01 | 48.298 | 67.5555 | 85.226 | 62.8801 | 0.011487 | 1.0506 | 0.024356 | 23.1826 | -0.073441 | 0.13787 | -1.2e-05 | -0.000995 | -0.000995 | -5.5858
2019-04-09 15:24 | 1554837812.7137 | -146 | 0.20952 | 0.000452 | 1e-06 | 101.018 | 101.6647 | 122.7589 | 20.6916 | 0.001905 | 0.001821 | -154.5769 | 0.23098 | 8.3993 | 0.002664 | 0.004267 | 0.004095 | 4.7e-05 | 0.001351 | 0.28982 | 1.3727 | 2.5475 | 98.47 | 20.01 | 48.346 | 67.1085 | 85.3557 | 65.0827 | 0.011498 | 1.0438 | 0.024279 | 23.26 | -0.057073 | 0.13741 | -5e-06 | -0.00043 | -0.00043 | -12.8838
2019-04-09 15:24 | 1554837813.7128 | -145 | 0.2095 | 0.000453 | 2e-06 | 101.0136 | 101.7242 | 122.8345 | 20.6905 | 0.001949 | 0.001846 | -154.5805 | 0.23209 | 8.4014 | 0.002621 | 0.003466 | 0.003283 | 4.5e-05 | 0.001352 | 0.2855 | 1.3654 | 2.5475 | 98.47 | 20.01 | 48.372 | 68.1683 | 84.3065 | 65.702 | 0.011504 | 1.0419 | 0.024264 | 23.2884 | -0.062706 | 0.13807 | -5e-06 | -0.000483 | -0.000483 | -11.5199
2019-04-09 15:24 | 1554837814.7136 | -144 | 0.20951 | 0.000455 | 1e-06 | 101.0148 | 101.835 | 122.8939 | 20.6912 | 0.001962 | 0.001767 | -154.5758 | 0.23228 | 8.4017 | 0.002644 | 0.003785 | 0.003603 | 4.1e-05 | 0.00135 | 0.28208 | 1.3808 | 2.5476 | 98.47 | 20 | 48.398 | 68.4692 | 84.8815 | 63.7894 | 0.011503 | 1.0478 | 0.024339 | 23.2282 | -0.072339 | 0.13819 | -8e-06 | -0.000649 | -0.000649 | -8.575
2019-04-09 15:24 | 1554837815.7126 | -143 | 0.20953 | 0.000459 | 1e-06 | 100.9927 | 101.738 | 123.1789 | 20.6902 | 0.001843 | 0.001742 | -154.5689 | 0.23222 | 8.4025 | 0.002631 | 0.003832 | 0.0037 | 4.3e-05 | 0.001349 | 0.30558 | 1.3932 | 2.5476 | 98.47 | 20 | 48.449 | 65.6074 | 84.5496 | 63.1965 | 0.011515 | 1.0496 | 0.024388 | 23.2346 | -0.065762 | 0.13815 | -1.4e-05 | -0.001102 | -0.001102 | -5.054
2019-04-09 15:24 | 1554837816.712 | -142 | 0.20951 | 0.000458 | 1e-06 | 100.9665 | 101.8812 | 123.3718 | 20.6905 | 0.001956 | 0.001745 | -154.5632 | 0.23218 | 8.4013 | 0.002663 | 0.003988 | 0.003824 | 4.5e-05 | 0.001341 | 0.27759 | 1.3377 | 2.5475 | 98.47 | 20 | 48.471 | 68.3469 | 85.3335 | 63.2522 | 0.011521 | 1.0495 | 0.024405 | 23.2548 | -0.080661 | 0.13812 | -1.3e-05 | -0.000999 | -0.000999 | -5.5732
2019-04-09 15:24 | 1554837817.7127 | -141 | 0.20956 | 0.000456 | 2e-06 | 100.9939 | 101.7893 | 123.3638 | 20.6905 | 0.001967 | 0.001807 | -154.5635 | 0.23195 | 8.4004 | 0.002615 | 0.004198 | 0.004055 | 4.3e-05 | 0.001342 | 0.26753 | 1.3507 | 2.5476 | 98.47 | 20 | 48.49 | 68.6036 | 84.1817 | 64.7529 | 0.011525 | 1.0448 | 0.02435 | 23.3058 | -0.070164 | 0.13799 | -1e-05 | -0.000844 | -0.000844 | -6.591
2019-04-09 15:24 | 1554837818.7117 | -140 | 0.2095 | 0.000461 | 1e-06 | 101.018 | 101.7542 | 123.6488 | 20.6905 | 0.001935 | 0.001764 | -154.5499 | 0.23187 | 8.4004 | 0.002678 | 0.003931 | 0.003801 | 4.5e-05 | 0.00134 | 0.29463 | 1.3468 | 2.5478 | 98.47 | 20 | 48.52 | 67.8315 | 85.6785 | 63.7153 | 0.011532 | 1.048 | 0.024416 | 23.2969 | -0.064942 | 0.13794 | -1.6e-05 | -0.00121 | -0.00121 | -4.5937
2019-04-09 15:24 | 1554837819.7127 | -139 | 0.20951 | 0.000451 | 1e-06 | 101.0246 | 101.8089 | 123.7011 | 20.6915 | 0.001905 | 0.001811 | -154.5595 | 0.23195 | 8.4017 | 0.002658 | 0.003821 | 0.003588 | 4.2e-05 | 0.001336 | 0.30397 | 1.349 | 2.5475 | 98.47 | 20 | 48.532 | 67.1161 | 85.2127 | 64.8475 | 0.011535 | 1.0445 | 0.02438 | 23.3409 | -0.069171 | 0.13799 | -4e-06 | -0.000373 | -0.000373 | -14.9195
2019-04-09 15:24 | 1554837820.7117 | -138 | 0.20949 | 0.000456 | 2e-06 | 100.9907 | 101.7993 | 123.9995 | 20.6911 | 0.001955 | 0.001769 | -154.5602 | 0.23264 | 8.401 | 0.002691 | 0.003749 | 0.003606 | 4.8e-05 | 0.001331 | 0.28423 | 1.3644 | 2.5476 | 98.47 | 20 | 48.553 | 68.3054 | 85.992 | 63.8494 | 0.01154 | 1.0476 | 0.024446 | 23.3345 | -0.071331 | 0.1384 | -1.1e-05 | -0.000878 | -0.000878 | -6.3539
2019-04-09 15:24 | 1554837821.7134 | -137 | 0.20952 | 0.000454 | 2e-06 | 101.0175 | 101.7146 | 123.7535 | 20.6908 | 0.001902 | 0.001789 | -154.5615 | 0.23367 | 8.4031 | 0.002682 | 0.003848 | 0.003674 | 4.2e-05 | 0.001317 | 0.29015 | 1.3858 | 2.5476 | 98.47 | 20 | 48.563 | 67.0313 | 85.7887 | 64.3161 | 0.011543 | 1.0462 | 0.024404 | 23.3275 | -0.061516 | 0.13901 | -8e-06 | -0.000665 | -0.000665 | -8.4261
2019-04-09 15:24 | 1554837822.7128 | -136 | 0.20951 | 0.000456 | 2e-06 | 100.9862 | 101.8354 | 123.6746 | 20.6908 | 0.001951 | 0.001816 | -154.5522 | 0.23241 | 8.4001 | 0.00264 | 0.003892 | 0.003711 | 4.5e-05 | 0.001313 | 0.27544 | 1.3585 | 2.5476 | 98.46 | 20 | 48.573 | 68.2075 | 84.7626 | 64.9734 | 0.011546 | 1.0441 | 0.024373 | 23.3428 | -0.074901 | 0.13826 | -1e-05 | -0.000804 | -0.000804 | -6.9265
2019-04-09 15:24 | 1554837823.7123 | -135 | 0.20953 | 0.000455 | 1e-06 | 101.0571 | 101.738 | 123.4573 | 20.6905 | 0.001874 | 0.001765 | -154.5355 | 0.23236 | 8.4029 | 0.002695 | 0.004039 | 0.003903 | 4.1e-05 | 0.001304 | 0.28119 | 1.3852 | 2.5476 | 98.46 | 20 | 48.584 | 66.3538 | 86.1023 | 63.7528 | 0.011549 | 1.0479 | 0.024396 | 23.2801 | -0.060067 | 0.13823 | -8e-06 | -0.000683 | -0.000683 | -8.1573
2019-04-09 15:24 | 1554837824.7131 | -134 | 0.2095 | 0.00045 | 3e-06 | 101.0541 | 101.8667 | 123.4908 | 20.6912 | 0.001925 | 0.001766 | -154.5422 | 0.23215 | 8.4015 | 0.002606 | 0.002986 | 0.002856 | 4.5e-05 | 0.001312 | 0.26681 | 1.3722 | 2.5477 | 98.46 | 20 | 48.586 | 67.5896 | 83.953 | 63.7605 | 0.011549 | 1.0479 | 0.024399 | 23.2835 | -0.07164 | 0.13811 | -3e-06 | -0.000405 | -0.000405 | -13.7535
2019-04-09 15:24 | 1554837825.7129 | -133 | 0.20952 | 0.000457 | 2e-06 | 101.032 | 101.7284 | 123.426 | 20.6913 | 0.001948 | 0.001754 | -154.5465 | 0.23265 | 8.4008 | 0.002675 | 0.003577 | 0.003452 | 4.4e-05 | 0.00132 | 0.28208 | 1.4035 | 2.5476 | 98.46 | 20 | 48.586 | 68.145 | 85.6277 | 63.4764 | 0.011549 | 1.0488 | 0.024403 | 23.2676 | -0.061434 | 0.13841 | -1.3e-05 | -0.001086 | -0.001086 | -5.1352
2019-04-09 15:24 | 1554837826.7135 | -132 | 0.20954 | 0.000457 | 2e-06 | 101.1062 | 101.7318 | 123.1901 | 20.6909 | 0.001906 | 0.001765 | -154.536 | 0.23206 | 8.4015 | 0.002719 | 0.004043 | 0.003874 | 4.5e-05 | 0.001315 | 0.29407 | 1.3913 | 2.5476 | 98.46 | 20 | 48.585 | 67.1336 | 86.6794 | 63.7369 | 0.011549 | 1.048 | 0.02437 | 23.2544 | -0.055174 | 0.13805 | -1.2e-05 | -0.000961 | -0.000961 | -5.789
2019-04-09 15:24 | 1554837827.7123 | -131 | 0.20952 | 0.00045 | 2e-06 | 101.0344 | 101.928 | 123.0066 | 20.6905 | 0.001923 | 0.001782 | -154.531 | 0.23253 | 8.4009 | 0.002676 | 0.004172 | 0.00408 | 4.3e-05 | 0.001317 | 0.27903 | 1.3669 | 2.5476 | 98.46 | 20 | 48.584 | 67.542 | 85.6339 | 64.1605 | 0.011549 | 1.0466 | 0.024336 | 23.2516 | -0.078757 | 0.13833 | -3e-06 | -0.000367 | -0.000367 | -15.1977
2019-04-09 15:24 | 1554837828.7126 | -130 | 0.2095 | 0.000457 | 1e-06 | 101.0865 | 101.8419 | 122.6172 | 20.6915 | 0.001887 | 0.001784 | -154.5325 | 0.23271 | 8.4021 | 0.00261 | 0.004136 | 0.003879 | 4.3e-05 | 0.001313 | 0.30038 | 1.3471 | 2.5475 | 98.46 | 20 | 48.582 | 66.6786 | 84.0548 | 64.203 | 0.011548 | 1.0465 | 0.024296 | 23.2163 | -0.066591 | 0.13844 | -1.1e-05 | -0.000868 | -0.000868 | -6.4248
2019-04-09 15:24 | 1554837829.7135 | -129 | 0.20952 | 0.000453 | 2e-06 | 101.13 | 101.7287 | 122.4519 | 20.6907 | 0.001875 | 0.001832 | -154.513 | 0.23257 | 8.4016 | 0.002758 | 0.003403 | 0.003182 | 4.4e-05 | 0.001318 | 0.28926 | 1.361 | 2.5477 | 98.46 | 20 | 48.582 | 66.3961 | 87.6084 | 65.3551 | 0.011548 | 1.043 | 0.024238 | 23.2402 | -0.052795 | 0.13836 | -6e-06 | -0.000603 | -0.000603 | -9.2445
2019-04-09 15:24 | 1554837830.7131 | -128 | 0.2095 | 0.000459 | 2e-06 | 101.0883 | 101.6433 | 122.1667 | 20.6902 | 0.001957 | 0.001769 | -154.5239 | 0.23214 | 8.4028 | 0.002702 | 0.004303 | 0.004147 | 4.2e-05 | 0.001323 | 0.30879 | 1.3454 | 2.5475 | 98.46 | 20 | 48.579 | 68.3522 | 86.2651 | 63.8337 | 0.011548 | 1.0477 | 0.024265 | 23.1609 | -0.048975 | 0.1381 | -1.5e-05 | -0.001202 | -0.001202 | -4.6314
2019-04-09 15:24 | 1554837831.7125 | -127 | 0.2095 | 0.000456 | 2e-06 | 101.1276 | 101.9128 | 121.9686 | 20.6905 | 0.001876 | 0.001751 | -154.5308 | 0.23234 | 8.4025 | 0.002744 | 0.004562 | 0.004346 | 4e-05 | 0.001327 | 0.27599 | 1.2964 | 2.5476 | 98.46 | 20 | 48.578 | 66.4106 | 87.2764 | 63.4171 | 0.011547 | 1.049 | 0.02426 | 23.1278 | -0.06918 | 0.13822 | -1e-05 | -0.00085 | -0.00085 | -6.5505
2019-04-09 15:24 | 1554837832.7111 | -126 | 0.20953 | 0.00046 | 2e-06 | 101.0231 | 101.8264 | 121.7649 | 20.6916 | 0.001934 | 0.001814 | -154.5319 | 0.2318 | 8.4026 | 0.002753 | 0.003425 | 0.003353 | 4.4e-05 | 0.001334 | 0.27455 | 1.3376 | 2.5476 | 98.46 | 20 | 48.575 | 67.8201 | 87.4861 | 64.9117 | 0.011547 | 1.0443 | 0.024186 | 23.1597 | -0.070845 | 0.1379 | -1.5e-05 | -0.001231 | -0.001231 | -4.516
2019-04-09 15:24 | 1554837833.7118 | -125 | 0.20951 | 0.000461 | 1e-06 | 101.069 | 101.9847 | 121.8224 | 20.6908 | 0.001911 | 0.001828 | -154.5298 | 0.23245 | 8.3995 | 0.002722 | 0.004612 | 0.004416 | 4.3e-05 | 0.001339 | 0.28279 | 1.3432 | 2.5475 | 98.46 | 20 | 48.565 | 67.252 | 86.7426 | 65.2536 | 0.011544 | 1.0433 | 0.02418 | 23.1769 | -0.080681 | 0.13828 | -1.6e-05 | -0.001188 | -0.001188 | -4.6909
2019-04-09 15:24 | 1554837834.7122 | -124 | 0.20952 | 0.000459 | 1e-06 | 101.083 | 101.7993 | 121.6095 | 20.6911 | 0.001942 | 0.001771 | -154.532 | 0.23197 | 8.4007 | 0.002755 | 0.003705 | 0.003629 | 4.4e-05 | 0.001327 | 0.27959 | 1.3366 | 2.5476 | 98.46 | 20 | 48.559 | 68.0098 | 87.5332 | 63.8972 | 0.011543 | 1.0475 | 0.024207 | 23.1102 | -0.063158 | 0.138 | -1.3e-05 | -0.001031 | -0.001031 | -5.3937
2019-04-09 15:24 | 1554837835.7123 | -123 | 0.20951 | 0.00046 | 2e-06 | 101.1401 | 101.7466 | 121.847 | 20.6911 | 0.001832 | 0.001742 | -154.5253 | 0.23283 | 8.3991 | 0.002642 | 0.004557 | 0.00439 | 4.3e-05 | 0.001319 | 0.28279 | 1.3311 | 2.5476 | 98.46 | 20.01 | 48.545 | 65.3444 | 84.8244 | 63.1981 | 0.011547 | 1.0496 | 0.024256 | 23.1088 | -0.053473 | 0.13851 | -1.5e-05 | -0.001181 | -0.001181 | -4.7276
2019-04-09 15:24 | 1554837836.7114 | -122 | 0.20952 | 0.00046 | 2e-06 | 101.0907 | 101.8282 | 121.9644 | 20.6913 | 0.001887 | 0.001858 | -154.5338 | 0.2323 | 8.4007 | 0.002698 | 0.003867 | 0.003676 | 4.3e-05 | 0.001316 | 0.29302 | 1.3741 | 2.5476 | 98.46 | 20 | 48.539 | 66.6767 | 86.1686 | 65.9924 | 0.011538 | 1.041 | 0.024167 | 23.2157 | -0.065013 | 0.13819 | -1.5e-05 | -0.001236 | -0.001236 | -4.5056
2019-04-09 15:24 | 1554837837.7113 | -121 | 0.20953 | 0.000462 | 2e-06 | 101.1821 | 101.9813 | 122.3627 | 20.6909 | 0.001899 | 0.001732 | -154.5252 | 0.23222 | 8.4011 | 0.00269 | 0.003875 | 0.003743 | 4.1e-05 | 0.001318 | 0.29407 | 1.3608 | 2.5478 | 98.46 | 20.01 | 48.532 | 66.972 | 85.98 | 62.9379 | 0.011544 | 1.0505 | 0.024317 | 23.1486 | -0.07037 | 0.13815 | -1.8e-05 | -0.001414 | -0.001414 | -3.9365
2019-04-09 15:24 | 1554837838.7114 | -120 | 0.20954 | 0.000462 | 2e-06 | 101.1732 | 101.9211 | 122.8141 | 20.6915 | 0.001949 | 0.001795 | -154.5409 | 0.23234 | 8.4014 | 0.002672 | 0.004042 | 0.003889 | 4.1e-05 | 0.001313 | 0.2959 | 1.3982 | 2.5477 | 98.46 | 20.01 | 48.517 | 68.1678 | 85.5528 | 64.4652 | 0.01154 | 1.0457 | 0.024306 | 23.2439 | -0.065876 | 0.13822 | -1.8e-05 | -0.001423 | -0.001423 | -3.9142
2019-04-09 15:24 | 1554837839.7118 | -119 | 0.20952 | 0.000467 | 1e-06 | 101.1169 | 101.8027 | 123.0427 | 20.6903 | 0.001888 | 0.001749 | -154.5454 | 0.23241 | 8.4031 | 0.00272 | 0.003187 | 0.002978 | 4.1e-05 | 0.001315 | 0.29031 | 1.366 | 2.5475 | 98.46 | 20.01 | 48.512 | 66.6953 | 86.693 | 63.3459 | 0.011539 | 1.0492 | 0.024369 | 23.227 | -0.060457 | 0.13826 | -2.3e-05 | -0.001765 | -0.001765 | -3.1564
2019-04-09 15:24 | 1554837840.7114 | -118 | 0.20952 | 0.000465 | 2e-06 | 101.136 | 101.9617 | 123.1467 | 20.6911 | 0.001952 | 0.00176 | -154.5357 | 0.23227 | 8.4001 | 0.002629 | 0.004914 | 0.004794 | 4.2e-05 | 0.001312 | 0.28672 | 1.3969 | 2.5476 | 98.46 | 20.01 | 48.505 | 68.2415 | 84.5 | 63.6278 | 0.011537 | 1.0483 | 0.024369 | 23.2465 | -0.072731 | 0.13818 | -2.1e-05 | -0.001604 | -0.001604 | -3.4713
2019-04-09 15:24 | 1554837841.7128 | -117 | 0.20953 | 0.000462 | 1e-06 | 101.1494 | 101.8474 | 123.2045 | 20.6908 | 0.00198 | 0.0018 | -154.5365 | 0.23187 | 8.4017 | 0.002646 | 0.003996 | 0.003818 | 4.1e-05 | 0.001316 | 0.29734 | 1.374 | 2.5476 | 98.46 | 20.01 | 48.494 | 68.9237 | 84.9169 | 64.5783 | 0.011535 | 1.0454 | 0.024341 | 23.2847 | -0.061517 | 0.13794 | -1.7e-05 | -0.001305 | -0.001305 | -4.2603
2019-04-09 15:24 | 1554837842.7125 | -116 | 0.20951 | 0.000459 | 2e-06 | 101.0752 | 101.7438 | 123.5763 | 20.6915 | 0.001884 | 0.001739 | -154.5372 | 0.2325 | 8.4009 | 0.002674 | 0.004385 | 0.004232 | 4.2e-05 | 0.001321 | 0.2672 | 1.326 | 2.5476 | 98.46 | 20.01 | 48.488 | 66.6121 | 85.6005 | 63.1157 | 0.011533 | 1.0499 | 0.02443 | 23.2693 | -0.058974 | 0.13831 | -1.4e-05 | -0.001112 | -0.001112 | -5.0143
2019-04-09 15:24 | 1554837843.7124 | -115 | 0.20954 | 0.000459 | 1e-06 | 101.055 | 101.674 | 123.7551 | 20.6904 | 0.001921 | 0.001729 | -154.5417 | 0.23205 | 8.3995 | 0.002671 | 0.003703 | 0.00357 | 4e-05 | 0.001325 | 0.27577 | 1.3476 | 2.5475 | 98.46 | 20.01 | 48.476 | 67.4864 | 85.523 | 62.8662 | 0.01153 | 1.0507 | 0.024457 | 23.2775 | -0.054622 | 0.13805 | -1.3e-05 | -0.001046 | -0.001046 | -5.3178
2019-04-09 15:24 | 1554837844.7123 | -114 | 0.20951 | 0.000458 | 2e-06 | 101.1193 | 101.867 | 123.9112 | 20.6911 | 0.001941 | 0.001793 | -154.5532 | 0.23243 | 8.4012 | 0.002671 | 0.003935 | 0.003726 | 3.7e-05 | 0.001333 | 0.28456 | 1.3451 | 2.5476 | 98.46 | 20.01 | 48.471 | 67.967 | 85.5285 | 64.4235 | 0.011529 | 1.0458 | 0.024416 | 23.3461 | -0.065899 | 0.13827 | -1.3e-05 | -0.001044 | -0.001044 | -5.3393
2019-04-09 15:24 | 1554837845.7116 | -113 | 0.20954 | 0.000455 | 1e-06 | 101.0996 | 101.8237 | 123.9044 | 20.6918 | 0.001888 | 0.001849 | -154.5627 | 0.23257 | 8.4002 | 0.002626 | 0.003921 | 0.003796 | 3.9e-05 | 0.001341 | 0.29103 | 1.365 | 2.5476 | 98.46 | 20.01 | 48.466 | 66.7045 | 84.4405 | 65.7689 | 0.011528 | 1.0417 | 0.024367 | 23.3919 | -0.063829 | 0.13836 | -8e-06 | -0.000707 | -0.000707 | -7.887
2019-04-09 15:24 | 1554837846.7127 | -112 | 0.2095 | 0.00045 | 1e-06 | 101.077 | 101.738 | 124.0342 | 20.6915 | 0.001885 | 0.001783 | -154.5561 | 0.23318 | 8.4017 | 0.002597 | 0.003899 | 0.003618 | 3.9e-05 | 0.001348 | 0.29679 | 1.3808 | 2.5477 | 98.46 | 20.01 | 48.46 | 66.6364 | 83.7364 | 64.1723 | 0.011526 | 1.0466 | 0.024437 | 23.349 | -0.058301 | 0.13872 | -1e-06 | -0.000172 | -0.000172 | -32.5667
2019-04-09 15:24 | 1554837847.7117 | -111 | 0.20955 | 0.000454 | 1e-06 | 101.1241 | 101.9197 | 123.8866 | 20.6908 | 0.00187 | 0.001757 | -154.5592 | 0.23231 | 8.4014 | 0.002552 | 0.00414 | 0.004008 | 3.6e-05 | 0.001353 | 0.29679 | 1.3522 | 2.5476 | 98.46 | 20.01 | 48.458 | 66.2666 | 82.6694 | 63.5447 | 0.011526 | 1.0486 | 0.024446 | 23.3134 | -0.0701 | 0.1382 | -7e-06 | -0.00058 | -0.00058 | -9.5994
2019-04-09 15:24 | 1554837848.7117 | -110 | 0.20953 | 0.000456 | 2e-06 | 101.1005 | 101.9603 | 123.558 | 20.6913 | 0.001892 | 0.001767 | -154.5494 | 0.2324 | 8.4012 | 0.00259 | 0.004133 | 0.004005 | 4.2e-05 | 0.001357 | 0.29839 | 1.3372 | 2.5476 | 98.46 | 20.01 | 48.449 | 66.8048 | 83.5801 | 63.7871 | 0.011524 | 1.0478 | 0.024404 | 23.2908 | -0.075744 | 0.13825 | -1e-05 | -0.000806 | -0.000806 | -6.9128
2019-04-09 15:24 | 1554837849.7122 | -109 | 0.20955 | 0.000455 | 2e-06 | 101.0544 | 101.7638 | 123.6532 | 20.6915 | 0.001933 | 0.001786 | -154.5593 | 0.23226 | 8.4034 | 0.002557 | 0.00372 | 0.003584 | 3.4e-05 | 0.001363 | 0.28672 | 1.3712 | 2.5475 | 98.46 | 20.01 | 48.444 | 67.7754 | 82.7816 | 64.2394 | 0.011523 | 1.0464 | 0.024397 | 23.3154 | -0.062573 | 0.13817 | -9e-06 | -0.000775 | -0.000775 | -7.1856
2019-04-09 15:24 | 1554837850.7119 | -108 | 0.20953 | 0.000447 | 1e-06 | 101.0571 | 101.8099 | 123.4864 | 20.6908 | 0.001897 | 0.001773 | -154.5345 | 0.2318 | 8.4007 | 0.002641 | 0.003284 | 0.003089 | 3.9e-05 | 0.001369 | 0.27367 | 1.3556 | 2.5476 | 98.46 | 20.01 | 48.438 | 66.9135 | 84.8015 | 63.9288 | 0.011521 | 1.0474 | 0.024392 | 23.2889 | -0.066388 | 0.1379 | 2e-06 | 3.7e-05 | 3.7e-05 | 148.3025
2019-04-09 15:24 | 1554837851.7108 | -107 | 0.20954 | 0.000449 | 2e-06 | 101.0809 | 101.8856 | 123.538 | 20.6914 | 0.001864 | 0.001747 | -154.5314 | 0.23207 | 8.402 | 0.002625 | 0.003916 | 0.003776 | 3.6e-05 | 0.001378 | 0.30038 | 1.3733 | 2.5476 | 98.46 | 20.01 | 48.423 | 66.1229 | 84.4176 | 63.3078 | 0.011518 | 1.0493 | 0.02442 | 23.2723 | -0.070929 | 0.13806 | -1e-06 | -0.000155 | -0.000155 | -35.9052
2019-04-09 15:24 | 1554837852.7123 | -106 | 0.20955 | 0.000446 | 2e-06 | 101.0487 | 101.7108 | 123.4497 | 20.6913 | 0.001931 | 0.001796 | -154.5274 | 0.23187 | 8.3989 | 0.002597 | 0.00368 | 0.003574 | 3.7e-05 | 0.001383 | 0.27831 | 1.4084 | 2.5478 | 98.46 | 20.01 | 48.414 | 67.7253 | 83.7355 | 64.4939 | 0.011515 | 1.0456 | 0.024368 | 23.305 | -0.058414 | 0.13794 | 2e-06 | 2.9e-05 | 2.9e-05 | 191.0856
2019-04-09 15:24 | 1554837853.7119 | -105 | 0.20955 | 0.00045 | 2e-06 | 101.1014 | 101.9049 | 123.3875 | 20.6912 | 0.001888 | 0.001731 | -154.5318 | 0.23194 | 8.403 | 0.002684 | 0.004625 | 0.004482 | 3.4e-05 | 0.001383 | 0.2995 | 1.4042 | 2.5476 | 98.46 | 20.01 | 48.405 | 66.7001 | 85.8283 | 62.9171 | 0.011513 | 1.0505 | 0.024419 | 23.2447 | -0.070803 | 0.13798 | -3e-06 | -0.000312 | -0.000312 | -17.8128
2019-04-09 15:24 | 1554837854.7119 | -104 | 0.20954 | 0.000447 | 1e-06 | 100.9963 | 101.8708 | 123.2779 | 20.6909 | 0.00189 | 0.001747 | -154.5081 | 0.2326 | 8.3999 | 0.002601 | 0.003425 | 0.003224 | 3.8e-05 | 0.00139 | 0.26936 | 1.3782 | 2.5475 | 98.46 | 20.01 | 48.386 | 66.7558 | 83.8295 | 63.3159 | 0.011509 | 1.0493 | 0.024394 | 23.2481 | -0.077117 | 0.13837 | 1e-06 | 2e-05 | 2e-05 | 277.721
2019-04-09 15:24 | 1554837855.7109 | -103 | 0.20953 | 0.000449 | 2e-06 | 101.0484 | 101.7772 | 122.9782 | 20.6913 | 0.001911 | 0.001765 | -154.4963 | 0.23248 | 8.403 | 0.002614 | 0.00359 | 0.003416 | 3.3e-05 | 0.001392 | 0.27383 | 1.3506 | 2.5474 | 98.46 | 20.01 | 48.376 | 67.2634 | 84.1387 | 63.7325 | 0.011506 | 1.048 | 0.024349 | 23.2342 | -0.06428 | 0.13831 | -1.9e-05 | -0.001501 | -0.001501 | -3.7145
2019-04-09 15:24 | 1554837856.712 | -102 | 0.20954 | 0.000453 | 1e-06 | 100.9689 | 101.8405 | 122.7237 | 20.691 | 0.001904 | 0.00172 | -154.4958 | 0.2326 | 8.4047 | 0.002647 | 0.003658 | 0.003547 | 3.1e-05 | 0.001389 | 0.26808 | 1.3353 | 2.5476 | 98.46 | 20.01 | 48.356 | 67.0813 | 84.946 | 62.6653 | 0.011502 | 1.0513 | 0.024362 | 23.1734 | -0.076884 | 0.13837 | -8.4e-05 | -0.006218 | -0.006218 | -0.89688
2019-04-09 15:24 | 1554837857.7124 | -101 | 0.20953 | 0.000446 | 1e-06 | 100.9466 | 101.7463 | 122.6376 | 20.6911 | 0.001843 | 0.001787 | -154.5037 | 0.23225 | 8.4014 | 0.002577 | 0.004173 | 0.004012 | 3.5e-05 | 0.001381 | 0.28031 | 1.3464 | 2.5475 | 98.46 | 20.01 | 48.345 | 65.6215 | 83.2539 | 64.2624 | 0.011499 | 1.0463 | 0.024296 | 23.2202 | -0.070577 | 0.13816 | -3e-05 | -0.002252 | -0.002252 | -2.4723
2019-04-09 15:24 | 1554837858.7113 | -100 | 0.20956 | 0.000449 | 2e-06 | 100.9912 | 101.6609 | 123.0545 | 20.6922 | 0.001956 | 0.0018 | -154.4997 | 0.2323 | 8.4018 | 0.002624 | 0.003865 | 0.00373 | 3.5e-05 | 0.001367 | 0.2808 | 1.3598 | 2.5475 | 98.46 | 20.01 | 48.333 | 68.3347 | 84.3933 | 64.5853 | 0.011496 | 1.0453 | 0.024326 | 23.2708 | -0.059114 | 0.13819 | -0.00019 | -0.01396 | -0.01396 | -0.39894
2019-04-09 15:24 | 1554837859.7121 | -99 | 0.20954 | 0.000448 | 1e-06 | 101.0404 | 101.7132 | 122.6543 | 20.6907 | 0.001905 | 0.001767 | -154.506 | 0.2318 | 8.4035 | 0.002696 | 0.003968 | 0.003807 | 3.6e-05 | 0.001357 | 0.27328 | 1.3457 | 2.5474 | 98.46 | 20.01 | 48.312 | 67.1202 | 86.1181 | 63.8014 | 0.011491 | 1.0478 | 0.024314 | 23.206 | -0.059364 | 0.13789 | -4e-05 | -0.002986 | -0.002986 | -1.8613
2019-04-09 15:24 | 1554837860.7122 | -98 | 0.20956 | 0.000455 | 1e-06 | 101.0356 | 101.8946 | 122.6139 | 20.6906 | 0.001845 | 0.001777 | -154.549 | 0.23246 | 8.4012 | 0.002599 | 0.004263 | 0.004104 | 3.7e-05 | 0.001357 | 0.29015 | 1.3962 | 2.5477 | 98.46 | 20.01 | 48.304 | 65.667 | 83.7894 | 64.0341 | 0.011489 | 1.047 | 0.024302 | 23.2101 | -0.075721 | 0.13829 | -0.000167 | -0.012212 | -0.012212 | -0.45636
2019-04-09 15:24 | 1554837861.7111 | -97 | 0.20957 | 0.000447 | 1e-06 | 100.9576 | 101.7865 | 122.547 | 20.6915 | 0.001843 | 0.001764 | -154.5539 | 0.2325 | 8.4006 | 0.00264 | 0.003699 | 0.003524 | 3.7e-05 | 0.001357 | 0.27367 | 1.3429 | 2.5475 | 98.46 | 20.01 | 48.289 | 65.608 | 84.7709 | 63.7124 | 0.011486 | 1.048 | 0.024307 | 23.1927 | -0.07314 | 0.13832 | -0.000218 | -0.015933 | -0.015933 | -0.34985
2019-04-09 15:24 | 1554837862.7116 | -96 | 0.20957 | 0.000446 | 2e-06 | 101.0124 | 101.9238 | 122.5489 | 20.6918 | 0.001905 | 0.0018 | -154.5682 | 0.23208 | 8.4009 | 0.00267 | 0.003849 | 0.003746 | 3.6e-05 | 0.001355 | 0.28982 | 1.3992 | 2.5477 | 98.46 | 20.01 | 48.283 | 67.1001 | 85.4964 | 64.5754 | 0.011484 | 1.0454 | 0.024276 | 23.2226 | -0.080345 | 0.13807 | -0.000249 | -0.018235 | -0.018235 | -0.30514
2019-04-09 15:24 | 1554837863.7102 | -95 | 0.20956 | 0.000451 | 1e-06 | 101.0052 | 101.9056 | 122.2907 | 20.6915 | 0.00187 | 0.001813 | -154.5467 | 0.23269 | 8.4048 | 0.002658 | 0.003993 | 0.003744 | 3.6e-05 | 0.001353 | 0.28263 | 1.3852 | 2.5476 | 98.46 | 20.01 | 48.279 | 66.2766 | 85.2103 | 64.9077 | 0.011483 | 1.0443 | 0.024239 | 23.2095 | -0.079377 | 0.13843 | -0.000186 | -0.013567 | -0.013567 | -0.4112
2019-04-09 15:24 | 1554837864.7102 | -94 | 0.20953 | 0.00045 | 1e-06 | 101.0145 | 101.7872 | 121.8029 | 20.6909 | 0.00179 | 0.001763 | -154.5472 | 0.23224 | 8.4016 | 0.002569 | 0.003453 | 0.003361 | 3.9e-05 | 0.001353 | 0.29319 | 1.3715 | 2.5476 | 98.46 | 20.01 | 48.277 | 64.3357 | 83.0664 | 63.7042 | 0.011483 | 1.0481 | 0.024233 | 23.1219 | -0.068163 | 0.13816 | -1e-06 | -0.000104 | -0.000104 | -53.5856
2019-04-09 15:24 | 1554837865.7113 | -93 | 0.20954 | 0.000447 | 1e-06 | 100.9731 | 101.8058 | 121.5018 | 20.692 | 0.001909 | 0.001836 | -154.5572 | 0.23247 | 8.4011 | 0.002564 | 0.003714 | 0.003578 | 3.7e-05 | 0.001348 | 0.29214 | 1.3595 | 2.5475 | 98.46 | 20.01 | 48.277 | 67.2112 | 82.9381 | 65.4406 | 0.011483 | 1.0427 | 0.024141 | 23.1528 | -0.073462 | 0.1383 | -4.4e-05 | -0.003246 | -0.003246 | -1.7171
2019-04-09 15:24 | 1554837866.712 | -92 | 0.20955 | 0.00045 | 1e-06 | 100.9368 | 101.7614 | 121.7095 | 20.6916 | 0.001849 | 0.001773 | -154.5653 | 0.23222 | 8.4021 | 0.002592 | 0.004504 | 0.004262 | 3.9e-05 | 0.001339 | 0.27992 | 1.356 | 2.5474 | 98.46 | 20.01 | 48.277 | 65.7524 | 83.6304 | 63.9342 | 0.011483 | 1.0474 | 0.024216 | 23.121 | -0.072779 | 0.13815 | -0.000126 | -0.009195 | -0.009195 | -0.60551
2019-04-09 15:24 | 1554837867.7104 | -91 | 0.20955 | 0.000448 | 1e-06 | 100.9859 | 101.8099 | 121.4321 | 20.6917 | 0.001844 | 0.001794 | -154.5755 | 0.23268 | 8.4013 | 0.002662 | 0.003541 | 0.003464 | 3.8e-05 | 0.001332 | 0.27041 | 1.3325 | 2.5476 | 98.46 | 20.01 | 48.277 | 65.651 | 85.2978 | 64.4466 | 0.011483 | 1.0458 | 0.02417 | 23.1121 | -0.07269 | 0.13842 | -0.000106 | -0.007787 | -0.007787 | -0.71637
2019-04-09 15:24 | 1554837868.7108 | -90 | 0.20954 | 0.000447 | 1e-06 | 100.8995 | 101.7469 | 120.8089 | 20.6909 | 0.001887 | 0.001792 | -154.5804 | 0.23245 | 8.4029 | 0.002618 | 0.003967 | 0.003765 | 3.6e-05 | 0.001325 | 0.2896 | 1.3638 | 2.5476 | 98.46 | 20.01 | 48.279 | 66.6818 | 84.2415 | 64.3996 | 0.011483 | 1.0459 | 0.024109 | 23.0512 | -0.074808 | 0.13828 | -5e-05 | -0.003663 | -0.003663 | -1.5213
2019-04-09 15:24 | 1554837869.7109 | -89 | 0.20951 | 0.000451 | 1e-06 | 100.9793 | 101.6709 | 120.8746 | 20.6912 | 0.001915 | 0.001815 | -154.5822 | 0.23252 | 8.4032 | 0.002671 | 0.003693 | 0.003551 | 3.8e-05 | 0.001324 | 0.28351 | 1.3641 | 2.5477 | 98.46 | 20.01 | 48.283 | 67.346 | 85.5167 | 64.9543 | 0.011484 | 1.0442 | 0.024096 | 23.0764 | -0.061047 | 0.13833 | 9.4e-05 | 0.006742 | 0.006742 | 0.82678
2019-04-09 15:24 | 1554837870.7114 | -88 | 0.2095 | 0.00045 | 1e-06 | 100.8989 | 101.7222 | 121.3189 | 20.6918 | 0.001928 | 0.001765 | -154.5487 | 0.23142 | 8.4034 | 0.00273 | 0.004042 | 0.003915 | 3.7e-05 | 0.001315 | 0.27688 | 1.3525 | 2.5475 | 98.46 | 20.01 | 48.285 | 67.6663 | 86.945 | 63.7315 | 0.011485 | 1.048 | 0.024184 | 23.0769 | -0.072682 | 0.13767 | 0.000185 | 0.013349 | 0.013349 | 0.41561
2019-04-09 15:25 | 1554837871.7119 | -87 | 0.2095 | 0.000449 | 1e-06 | 100.9144 | 101.6705 | 121.4863 | 20.6911 | 0.001974 | 0.001753 | -154.5378 | 0.23241 | 8.4016 | 0.002664 | 0.004476 | 0.004305 | 4.1e-05 | 0.001309 | 0.29214 | 1.3877 | 2.5477 | 98.46 | 20.01 | 48.287 | 68.7663 | 85.3567 | 63.4598 | 0.011485 | 1.0488 | 0.024211 | 23.0835 | -0.066769 | 0.13826 | 0.000203 | 0.014633 | 0.014633 | 0.38078
2019-04-09 15:25 | 1554837872.7116 | -86 | 0.2095 | 0.000444 | 2e-06 | 100.9108 | 101.7208 | 121.1821 | 20.6912 | 0.001862 | 0.001781 | -154.5245 | 0.23289 | 8.4014 | 0.002728 | 0.003284 | 0.0032 | 4e-05 | 0.001307 | 0.28456 | 1.378 | 2.5477 | 98.46 | 20.01 | 48.291 | 66.0688 | 86.8884 | 64.1324 | 0.011486 | 1.0467 | 0.024156 | 23.0776 | -0.071505 | 0.13855 | 0.000191 | 0.013746 | 0.013746 | 0.40619
2019-04-09 15:25 | 1554837873.7105 | -85 | 0.20947 | 0.000448 | 2e-06 | 100.9105 | 101.7752 | 121.1363 | 20.6907 | 0.001956 | 0.001781 | -154.5387 | 0.23261 | 8.4019 | 0.002717 | 0.00439 | 0.004199 | 4.3e-05 | 0.001303 | 0.29214 | 1.3691 | 2.5476 | 98.46 | 20.01 | 48.292 | 68.33 | 86.627 | 64.1184 | 0.011486 | 1.0468 | 0.024152 | 23.0728 | -0.076311 | 0.13838 | 0.000365 | 0.026294 | 0.026294 | 0.2121
2019-04-09 15:25 | 1554837874.7105 | -84 | 0.20949 | 0.000445 | 1e-06 | 100.926 | 101.6754 | 121.6229 | 20.6911 | 0.001857 | 0.00178 | -154.5487 | 0.23226 | 8.4 | 0.002675 | 0.004223 | 0.004007 | 4.4e-05 | 0.0013 | 0.28456 | 1.3831 | 2.5476 | 98.46 | 20.01 | 48.295 | 65.9588 | 85.6165 | 64.094 | 0.011487 | 1.0469 | 0.024201 | 23.1182 | -0.066163 | 0.13817 | 0.000236 | 0.017054 | 0.017054 | 0.32651
2019-04-09 15:25 | 1554837875.7106 | -83 | 0.20949 | 0.000447 | 1e-06 | 100.9061 | 101.7469 | 122.0092 | 20.6916 | 0.0019 | 0.001841 | -154.5399 | 0.23185 | 8.4031 | 0.002639 | 0.004256 | 0.004181 | 4.2e-05 | 0.001296 | 0.29159 | 1.3956 | 2.5476 | 98.46 | 20.01 | 48.295 | 66.9891 | 84.7595 | 65.5619 | 0.011487 | 1.0423 | 0.024187 | 23.2052 | -0.074227 | 0.13793 | 0.00022 | 0.015912 | 0.015912 | 0.34932
2019-04-09 15:25 | 1554837876.712 | -82 | 0.2095 | 0.000447 | 2e-06 | 100.8462 | 101.6537 | 121.663 | 20.6919 | 0.001874 | 0.001797 | -154.542 | 0.23152 | 8.4026 | 0.002752 | 0.003858 | 0.003654 | 4.3e-05 | 0.001285 | 0.29734 | 1.404 | 2.5477 | 98.46 | 20.01 | 48.297 | 66.3723 | 87.4745 | 64.5154 | 0.011488 | 1.0455 | 0.02419 | 23.1365 | -0.071331 | 0.13773 | 0.000168 | 0.012043 | 0.012043 | 0.46089
2019-04-09 15:25 | 1554837877.7116 | -81 | 0.20952 | 0.000455 | 1e-06 | 100.9892 | 101.6998 | 121.5485 | 20.691 | 0.001898 | 0.001826 | -154.5453 | 0.23252 | 8.4001 | 0.002762 | 0.003921 | 0.003766 | 4.6e-05 | 0.001283 | 0.29463 | 1.3686 | 2.5476 | 98.46 | 20.01 | 48.302 | 66.9471 | 87.7126 | 65.2218 | 0.011489 | 1.0434 | 0.024154 | 23.1498 | -0.062719 | 0.13833 | 3.2e-05 | 0.0022 | 0.0022 | 2.5342
2019-04-09 15:25 | 1554837878.7111 | -80 | 0.2095 | 0.000451 | 2e-06 | 100.9186 | 101.8147 | 121.9292 | 20.6909 | 0.001946 | 0.00182 | -154.5413 | 0.23211 | 8.4034 | 0.002778 | 0.003841 | 0.003712 | 4.5e-05 | 0.001278 | 0.29015 | 1.3286 | 2.5476 | 98.46 | 20.01 | 48.305 | 68.0961 | 88.0922 | 65.0594 | 0.01149 | 1.0439 | 0.024197 | 23.1804 | -0.079076 | 0.13809 | 0.00015 | 0.010784 | 0.010784 | 0.51601
2019-04-09 15:25 | 1554837879.7112 | -79 | 0.2095 | 0.000451 | 1e-06 | 100.896 | 101.517 | 122.0839 | 20.6913 | 0.001889 | 0.001837 | -154.5323 | 0.23199 | 8.399 | 0.002695 | 0.003842 | 0.003659 | 4.3e-05 | 0.001273 | 0.30558 | 1.3003 | 2.5475 | 98.46 | 20.01 | 48.312 | 66.7166 | 86.0906 | 65.4758 | 0.011491 | 1.0426 | 0.024198 | 23.2094 | -0.054892 | 0.13801 | 0.000172 | 0.012415 | 0.012415 | 0.448
2019-04-09 15:25 | 1554837880.7097 | -78 | 0.20951 | 0.00045 | 2e-06 | 100.887 | 101.4441 | 121.6959 | 20.6924 | 0.001907 | 0.00177 | -154.5204 | 0.23197 | 8.4002 | 0.002678 | 0.003919 | 0.003795 | 4.3e-05 | 0.00127 | 0.29479 | 1.3217 | 2.5475 | 98.46 | 20.02 | 48.316 | 67.1483 | 85.6834 | 63.8639 | 0.011499 | 1.0476 | 0.024217 | 23.1173 | -0.049253 | 0.138 | 0.000105 | 0.007517 | 0.007517 | 0.73979
2019-04-09 15:25 | 1554837881.71 | -77 | 0.2095 | 0.000453 | 1e-06 | 100.89 | 101.4788 | 121.7027 | 20.6915 | 0.001922 | 0.001749 | -154.5329 | 0.23208 | 8.3995 | 0.002669 | 0.003932 | 0.003668 | 4.8e-05 | 0.001273 | 0.28512 | 1.3435 | 2.5474 | 98.46 | 20.02 | 48.321 | 67.5235 | 85.4773 | 63.3461 | 0.011501 | 1.0492 | 0.024236 | 23.1001 | -0.052053 | 0.13806 | 0.000148 | 0.010713 | 0.010713 | 0.51936
2019-04-09 15:25 | 1554837882.7108 | -76 | 0.20949 | 0.000453 | 1e-06 | 100.87 | 101.5487 | 122.0721 | 20.6911 | 0.001947 | 0.001806 | -154.5386 | 0.2321 | 8.4035 | 0.002742 | 0.003832 | 0.003685 | 4.6e-05 | 0.001278 | 0.28296 | 1.3537 | 2.5476 | 98.46 | 20.02 | 48.33 | 68.1292 | 87.2254 | 64.7295 | 0.011503 | 1.0449 | 0.024223 | 23.1827 | -0.059978 | 0.13808 | 0.000207 | 0.014966 | 0.014966 | 0.3718
2019-04-09 15:25 | 1554837883.7103 | -75 | 0.20952 | 0.000452 | 1e-06 | 100.84 | 101.6182 | 122.1538 | 20.6917 | 0.001918 | 0.001733 | -154.534 | 0.23226 | 8.3995 | 0.00274 | 0.004211 | 0.004053 | 4.8e-05 | 0.001287 | 0.29407 | 1.3443 | 2.5476 | 98.46 | 20.02 | 48.336 | 67.434 | 87.1855 | 62.9847 | 0.011504 | 1.0503 | 0.024294 | 23.1305 | -0.068767 | 0.13817 | 8.3e-05 | 0.006018 | 0.006018 | 0.9253
2019-04-09 15:25 | 1554837884.7109 | -74 | 0.2095 | 0.000453 | 1e-06 | 100.8864 | 101.5126 | 122.1625 | 20.6914 | 0.001904 | 0.001804 | -154.5352 | 0.2315 | 8.4014 | 0.002781 | 0.004249 | 0.004104 | 4.6e-05 | 0.001292 | 0.30757 | 1.3927 | 2.5476 | 98.46 | 20.02 | 48.345 | 67.0771 | 88.1619 | 64.693 | 0.011506 | 1.045 | 0.024234 | 23.19 | -0.055342 | 0.13772 | 0.000148 | 0.010664 | 0.010664 | 0.52046
2019-04-09 15:25 | 1554837885.7103 | -73 | 0.2095 | 0.000449 | 1e-06 | 100.8102 | 101.6093 | 121.9858 | 20.6908 | 0.001917 | 0.001758 | -154.5336 | 0.23228 | 8.4001 | 0.002725 | 0.003857 | 0.003695 | 4.5e-05 | 0.001297 | 0.26416 | 1.3744 | 2.5477 | 98.46 | 20.02 | 48.348 | 67.387 | 86.8262 | 63.5775 | 0.011507 | 1.0485 | 0.024256 | 23.1349 | -0.070621 | 0.13818 | 0.000164 | 0.011821 | 0.011821 | 0.47108
2019-04-09 15:25 | 1554837886.7109 | -72 | 0.20951 | 0.000451 | 2e-06 | 100.7932 | 101.7177 | 122.2771 | 20.6909 | 0.0019 | 0.001803 | -154.5227 | 0.23219 | 8.401 | 0.002776 | 0.004543 | 0.004389 | 5.2e-05 | 0.001303 | 0.29734 | 1.362 | 2.5475 | 98.46 | 20.02 | 48.35 | 66.9807 | 88.049 | 64.6535 | 0.011507 | 1.0451 | 0.024246 | 23.1995 | -0.081665 | 0.13813 | 0.0001 | 0.007151 | 0.007151 | 0.77841
2019-04-09 15:25 | 1554837887.71 | -71 | 0.20956 | 0.000451 | 1e-06 | 100.7596 | 101.5016 | 122.7273 | 20.6909 | 0.001948 | 0.001759 | -154.5167 | 0.23187 | 8.4007 | 0.002735 | 0.003901 | 0.003703 | 4.5e-05 | 0.001315 | 0.28152 | 1.3621 | 2.5476 | 98.46 | 20.02 | 48.351 | 68.153 | 87.0543 | 63.5987 | 0.011508 | 1.0484 | 0.024329 | 23.2059 | -0.065626 | 0.13794 | -0.000165 | -0.012067 | -0.012067 | -0.46068
2019-04-09 15:25 | 1554837888.711 | -70 | 0.20956 | 0.000452 | 1e-06 | 100.7825 | 101.5735 | 122.8646 | 20.6908 | 0.001823 | 0.001814 | -154.5145 | 0.2325 | 8.4029 | 0.002832 | 0.004585 | 0.004425 | 4.7e-05 | 0.001317 | 0.30397 | 1.3625 | 2.5476 | 98.46 | 20.02 | 48.349 | 65.1404 | 89.3996 | 64.9313 | 0.011507 | 1.0443 | 0.024294 | 23.2648 | -0.069927 | 0.13831 | -0.000199 | -0.01457 | -0.01457 | -0.38257
2019-04-09 15:25 | 1554837889.7111 | -69 | 0.20954 | 0.000457 | 1e-06 | 100.7899 | 101.3642 | 122.7399 | 20.6913 | 0.001933 | 0.001795 | -154.5221 | 0.23183 | 8.4018 | 0.002804 | 0.004434 | 0.004263 | 4.6e-05 | 0.001328 | 0.28208 | 1.3587 | 2.5476 | 98.46 | 20.02 | 48.348 | 67.7815 | 88.7314 | 64.4657 | 0.011507 | 1.0457 | 0.024299 | 23.2369 | -0.050821 | 0.13792 | -9.1e-05 | -0.006684 | -0.006684 | -0.83153
2019-04-09 15:25 | 1554837890.7115 | -68 | 0.20958 | 0.000449 | 1e-06 | 100.8492 | 101.3931 | 122.8578 | 20.6916 | 0.001944 | 0.001811 | -154.546 | 0.23206 | 8.4005 | 0.002725 | 0.003574 | 0.003447 | 4.6e-05 | 0.001326 | 0.29894 | 1.3254 | 2.5476 | 98.46 | 20.02 | 48.348 | 68.047 | 86.8216 | 64.8462 | 0.011507 | 1.0445 | 0.024297 | 23.2612 | -0.048114 | 0.13805 | -0.000307 | -0.022408 | -0.022408 | -0.24828
2019-04-09 15:25 | 1554837891.7101 | -67 | 0.20957 | 0.000448 | 1e-06 | 100.7679 | 101.2916 | 122.686 | 20.6925 | 0.00198 | 0.001802 | -154.5237 | 0.23315 | 8.4032 | 0.002738 | 0.004593 | 0.00442 | 4.5e-05 | 0.001323 | 0.29551 | 1.3319 | 2.5475 | 98.46 | 20.02 | 48.348 | 68.9048 | 87.1244 | 64.6261 | 0.011507 | 1.0452 | 0.024288 | 23.2373 | -0.046367 | 0.1387 | -0.000225 | -0.016467 | -0.016467 | -0.33944
2019-04-09 15:25 | 1554837892.7111 | -66 | 0.20954 | 0.000452 | 1e-06 | 100.7477 | 101.419 | 122.7068 | 20.6917 | 0.001937 | 0.00181 | -154.5312 | 0.23195 | 8.4031 | 0.002734 | 0.004718 | 0.004501 | 4.5e-05 | 0.001319 | 0.27312 | 1.3366 | 2.5477 | 98.46 | 20.02 | 48.349 | 67.8918 | 87.0402 | 64.8337 | 0.011507 | 1.0446 | 0.024282 | 23.2465 | -0.059402 | 0.13798 | -5.9e-05 | -0.004379 | -0.004379 | -1.2699
2019-04-09 15:25 | 1554837893.7101 | -65 | 0.20956 | 0.000444 | 1e-06 | 100.7974 | 101.4231 | 122.8984 | 20.6922 | 0.001939 | 0.001801 | -154.5246 | 0.23269 | 8.4017 | 0.002765 | 0.003745 | 0.003642 | 4.2e-05 | 0.00132 | 0.29767 | 1.3168 | 2.5475 | 98.46 | 20.02 | 48.349 | 67.9391 | 87.7855 | 64.6001 | 0.011507 | 1.0453 | 0.02431 | 23.2566 | -0.055352 | 0.13843 | -0.000195 | -0.014326 | -0.014326 | -0.38941
2019-04-09 15:25 | 1554837894.7099 | -64 | 0.20954 | 0.000448 | 1e-06 | 100.8081 | 101.5824 | 123.0277 | 20.6919 | 0.001871 | 0.001828 | -154.5337 | 0.23268 | 8.4009 | 0.00283 | 0.003908 | 0.00372 | 4.5e-05 | 0.001321 | 0.286 | 1.3351 | 2.5476 | 98.46 | 20.02 | 48.351 | 66.2996 | 89.3593 | 65.2499 | 0.011508 | 1.0433 | 0.024299 | 23.2912 | -0.068442 | 0.13842 | -4.2e-05 | -0.003129 | -0.003129 | -1.7829
2019-04-09 15:25 | 1554837895.7091 | -63 | 0.20954 | 0.000447 | 1e-06 | 100.7727 | 101.4991 | 123.0538 | 20.6915 | 0.001933 | 0.001821 | -154.5373 | 0.2319 | 8.4014 | 0.002864 | 0.004037 | 0.003794 | 4.2e-05 | 0.00132 | 0.28064 | 1.3698 | 2.5477 | 98.46 | 20.03 | 48.352 | 67.78 | 90.17 | 65.0864 | 0.011515 | 1.0438 | 0.024308 | 23.288 | -0.064251 | 0.13796 | -4.1e-05 | -0.003042 | -0.003042 | -1.8275
2019-04-09 15:25 | 1554837896.7092 | -62 | 0.20954 | 0.000447 | 2e-06 | 100.7992 | 101.5893 | 122.7773 | 20.6913 | 0.001855 | 0.001756 | -154.5487 | 0.23235 | 8.4009 | 0.002658 | 0.004197 | 0.004034 | 4.1e-05 | 0.001319 | 0.29463 | 1.3912 | 2.5477 | 98.46 | 20.03 | 48.354 | 65.9121 | 85.1993 | 63.5271 | 0.011516 | 1.0486 | 0.024336 | 23.2081 | -0.069841 | 0.13822 | -5.4e-05 | -0.004049 | -0.004049 | -1.3756
2019-04-09 15:25 | 1554837897.7102 | -61 | 0.20953 | 0.000444 | 1e-06 | 100.8489 | 101.4582 | 122.8437 | 20.6913 | 0.001913 | 0.001803 | -154.5425 | 0.23299 | 8.4008 | 0.002713 | 0.003977 | 0.003885 | 4e-05 | 0.001317 | 0.28799 | 1.3665 | 2.5475 | 98.46 | 20.03 | 48.355 | 67.3089 | 86.541 | 64.6679 | 0.011516 | 1.0451 | 0.024302 | 23.2537 | -0.053877 | 0.13861 | 2.2e-05 | 0.001489 | 0.001489 | 3.7511
2019-04-09 15:25 | 1554837898.7097 | -60 | 0.20957 | 0.000448 | 1e-06 | 100.7396 | 101.568 | 122.6667 | 20.6915 | 0.001957 | 0.001809 | -154.5218 | 0.23174 | 8.4032 | 0.0028 | 0.004417 | 0.004256 | 4.4e-05 | 0.001308 | 0.29518 | 1.3866 | 2.5476 | 98.46 | 20.03 | 48.356 | 68.3705 | 88.63 | 64.804 | 0.011516 | 1.0447 | 0.024279 | 23.2416 | -0.073249 | 0.13786 | -0.000244 | -0.017859 | -0.017859 | -0.31109
2019-04-09 15:25 | 1554837899.7107 | -59 | 0.20956 | 0.000444 | 2e-06 | 100.7354 | 101.5501 | 122.8578 | 20.6922 | 0.00194 | 0.001781 | -154.501 | 0.23192 | 8.4008 | 0.0028 | 0.003686 | 0.003538 | 4.6e-05 | 0.001305 | 0.29142 | 1.364 | 2.5476 | 98.46 | 20.03 | 48.358 | 67.9594 | 88.6213 | 64.1271 | 0.011516 | 1.0468 | 0.024323 | 23.2364 | -0.072043 | 0.13797 | -0.000182 | -0.013349 | -0.013349 | -0.41653
2019-04-09 15:25 | 1554837900.7103 | -58 | 0.20954 | 0.000448 | 1e-06 | 100.7971 | 101.5866 | 123.0401 | 20.6917 | 0.001928 | 0.00179 | -154.4977 | 0.23237 | 8.4015 | 0.002771 | 0.003578 | 0.003436 | 4.6e-05 | 0.001299 | 0.30525 | 1.3273 | 2.5475 | 98.46 | 20.03 | 48.358 | 67.6655 | 87.9367 | 64.3349 | 0.011516 | 1.0461 | 0.024333 | 23.2608 | -0.069784 | 0.13824 | -6.2e-05 | -0.004569 | -0.004569 | -1.2192
2019-04-09 15:25 | 1554837901.7103 | -57 | 0.20959 | 0.000451 | 1e-06 | 100.8126 | 101.314 | 123.2117 | 20.6905 | 0.001902 | 0.001842 | -154.5198 | 0.23209 | 8.4001 | 0.002677 | 0.004005 | 0.00376 | 4.7e-05 | 0.001304 | 0.29623 | 1.3612 | 2.5475 | 98.46 | 20.03 | 48.357 | 67.0435 | 85.6527 | 65.6046 | 0.011516 | 1.0422 | 0.024305 | 23.3208 | -0.044378 | 0.13807 | -0.000389 | -0.028418 | -0.028418 | -0.1958
2019-04-09 15:25 | 1554837902.7103 | -56 | 0.20954 | 0.000449 | 1e-06 | 100.764 | 101.2744 | 123.2185 | 20.6909 | 0.001935 | 0.001797 | -154.5 | 0.23219 | 8.4025 | 0.002798 | 0.004512 | 0.004395 | 4.6e-05 | 0.001307 | 0.28296 | 1.3567 | 2.5475 | 98.46 | 20.03 | 48.352 | 67.8287 | 88.5685 | 64.5223 | 0.011515 | 1.0455 | 0.024344 | 23.2841 | -0.045191 | 0.13813 | -3.6e-05 | -0.002725 | -0.002725 | -2.0431
2019-04-09 15:25 | 1554837903.7103 | -55 | 0.20954 | 0.000451 | 2e-06 | 100.7328 | 101.4296 | 123.157 | 20.6914 | 0.001893 | 0.001788 | -154.4998 | 0.23216 | 8.4008 | 0.002838 | 0.004027 | 0.00382 | 4.8e-05 | 0.001308 | 0.28998 | 1.3823 | 2.5476 | 98.46 | 20.03 | 48.346 | 66.822 | 89.5422 | 64.2916 | 0.011514 | 1.0462 | 0.024346 | 23.2704 | -0.061665 | 0.13811 | -3.7e-05 | -0.002773 | -0.002773 | -2.0071
2019-04-09 15:25 | 1554837904.7108 | -54 | 0.20955 | 0.000451 | 1e-06 | 100.6982 | 101.2961 | 123.0937 | 20.6909 | 0.001904 | 0.001805 | -154.4885 | 0.23277 | 8.4049 | 0.002741 | 0.004456 | 0.00424 | 4.7e-05 | 0.001315 | 0.2688 | 1.3506 | 2.5476 | 98.46 | 20.03 | 48.333 | 67.0905 | 87.2025 | 64.7011 | 0.011511 | 1.045 | 0.024325 | 23.2785 | -0.05295 | 0.13848 | -0.00014 | -0.010283 | -0.010283 | -0.5427
2019-04-09 15:25 | 1554837905.7099 | -53 | 0.20954 | 0.000449 | 1e-06 | 100.6768 | 101.3419 | 123.073 | 20.6917 | 0.001894 | 0.001825 | -154.497 | 0.23226 | 8.4012 | 0.002742 | 0.004056 | 0.003989 | 4.7e-05 | 0.001307 | 0.27759 | 1.4028 | 2.5474 | 98.46 | 20.03 | 48.326 | 66.8449 | 87.2277 | 65.1913 | 0.011509 | 1.0435 | 0.024306 | 23.2934 | -0.058895 | 0.13817 | -4.3e-05 | -0.00318 | -0.00318 | -1.7508
2019-04-09 15:25 | 1554837906.7095 | -52 | 0.20954 | 0.000446 | 1e-06 | 100.7295 | 101.4499 | 123.2108 | 20.6917 | 0.001949 | 0.001755 | -154.4782 | 0.23228 | 8.4021 | 0.002807 | 0.004762 | 0.004525 | 4.5e-05 | 0.001298 | 0.28136 | 1.4056 | 2.5476 | 98.46 | 20.03 | 48.317 | 68.1733 | 88.7855 | 63.5104 | 0.011507 | 1.0487 | 0.02438 | 23.2485 | -0.063746 | 0.13819 | -7e-05 | -0.005164 | -0.005164 | -1.0785
2019-04-09 15:25 | 1554837907.7109 | -51 | 0.20956 | 0.000448 | 1e-06 | 100.7212 | 101.3814 | 122.7636 | 20.6918 | 0.001903 | 0.001851 | -154.4729 | 0.23165 | 8.4003 | 0.002803 | 0.004084 | 0.003946 | 5e-05 | 0.001293 | 0.27649 | 1.3517 | 2.5472 | 98.46 | 20.03 | 48.294 | 67.0713 | 88.7069 | 65.8142 | 0.011501 | 1.0415 | 0.024253 | 23.2855 | -0.058446 | 0.13781 | -0.000156 | -0.011418 | -0.011418 | -0.48639
2019-04-09 15:25 | 1554837908.7106 | -50 | 0.20953 | 0.000449 | 1e-06 | 100.661 | 101.2524 | 122.6437 | 20.6922 | 0.001886 | 0.001772 | -154.5017 | 0.23299 | 8.4018 | 0.002821 | 0.003807 | 0.003614 | 4.6e-05 | 0.001299 | 0.27096 | 1.3567 | 2.5473 | 98.46 | 20.03 | 48.283 | 66.6441 | 89.1399 | 63.9028 | 0.011499 | 1.0474 | 0.02431 | 23.2084 | -0.052396 | 0.1386 | 1.6e-05 | 0.001113 | 0.001113 | 5.019
2019-04-09 15:25 | 1554837909.7096 | -49 | 0.20958 | 0.000445 | 1e-06 | 100.6815 | 101.4121 | 122.6456 | 20.6914 | 0.00192 | 0.001824 | -154.5182 | 0.2326 | 8.401 | 0.002817 | 0.004382 | 0.004273 | 4.1e-05 | 0.001309 | 0.28744 | 1.355 | 2.5476 | 98.46 | 20.03 | 48.263 | 67.4647 | 89.0368 | 65.1604 | 0.011494 | 1.0436 | 0.024265 | 23.2519 | -0.064666 | 0.13837 | -0.000289 | -0.021081 | -0.021081 | -0.26452
2019-04-09 15:25 | 1554837910.7087 | -48 | 0.20953 | 0.000449 | 2e-06 | 100.7435 | 101.4224 | 122.5162 | 20.6918 | 0.001954 | 0.001858 | -154.5074 | 0.23265 | 8.4017 | 0.002841 | 0.004798 | 0.004542 | 4.2e-05 | 0.001308 | 0.31509 | 1.3258 | 2.5477 | 98.46 | 20.03 | 48.255 | 68.2849 | 89.6108 | 65.9803 | 0.011492 | 1.041 | 0.024222 | 23.2678 | -0.060076 | 0.1384 | 5e-06 | 0.000253 | 0.000253 | 22.075
2019-04-09 15:25 | 1554837911.7093 | -47 | 0.20955 | 0.000454 | 1e-06 | 100.6589 | 101.4854 | 122.4209 | 20.6919 | 0.001923 | 0.001826 | -154.4787 | 0.23264 | 8.4012 | 0.002831 | 0.004036 | 0.00386 | 4.3e-05 | 0.001313 | 0.27184 | 1.3651 | 2.5475 | 98.46 | 20.03 | 48.248 | 67.552 | 89.3795 | 65.2148 | 0.01149 | 1.0434 | 0.02424 | 23.2324 | -0.07314 | 0.13839 | -0.000103 | -0.007564 | -0.007564 | -0.73732
2019-04-09 15:25 | 1554837912.7097 | -46 | 0.20953 | 0.00045 | 1e-06 | 100.6684 | 101.3787 | 122.0919 | 20.6918 | 0.001913 | 0.00185 | -154.4976 | 0.2324 | 8.404 | 0.002852 | 0.004495 | 0.004381 | 4.2e-05 | 0.001318 | 0.30685 | 1.3547 | 2.5475 | 98.46 | 20.03 | 48.234 | 67.3138 | 89.8786 | 65.7891 | 0.011487 | 1.0416 | 0.024187 | 23.2209 | -0.062885 | 0.13825 | 5e-06 | 0.000245 | 0.000245 | 22.7485
2019-04-09 15:25 | 1554837913.7087 | -45 | 0.20953 | 0.000446 | 2e-06 | 100.6074 | 101.5095 | 121.8935 | 20.6922 | 0.001951 | 0.001779 | -154.5159 | 0.23238 | 8.4057 | 0.002848 | 0.00376 | 0.003662 | 4.1e-05 | 0.001313 | 0.27577 | 1.3399 | 2.5476 | 98.46 | 20.03 | 48.229 | 68.2137 | 89.7735 | 64.0879 | 0.011486 | 1.0469 | 0.024229 | 23.1437 | -0.079842 | 0.13824 | -1.8e-05 | -0.001399 | -0.001399 | -3.9835
2019-04-09 15:25 | 1554837914.7088 | -44 | 0.20951 | 0.000451 | 2e-06 | 100.6187 | 101.3611 | 121.7203 | 20.692 | 0.001911 | 0.00172 | -154.5206 | 0.23062 | 8.4044 | 0.002809 | 0.003834 | 0.0037 | 4.2e-05 | 0.001306 | 0.29535 | 1.3649 | 2.5475 | 98.46 | 20.03 | 48.222 | 67.2469 | 88.8303 | 62.6648 | 0.011484 | 1.0513 | 0.024263 | 23.0784 | -0.065756 | 0.1372 | 9.3e-05 | 0.006645 | 0.006645 | 0.83202
2019-04-09 15:25 | 1554837915.7102 | -43 | 0.20954 | 0.00045 | 2e-06 | 100.6434 | 101.3811 | 121.8388 | 20.6924 | 0.001964 | 0.001822 | -154.5142 | 0.23231 | 8.4006 | 0.002866 | 0.004067 | 0.003902 | 4.2e-05 | 0.001294 | 0.27328 | 1.3422 | 2.5476 | 98.46 | 20.03 | 48.214 | 68.5329 | 90.2162 | 65.1232 | 0.011482 | 1.0437 | 0.024186 | 23.174 | -0.065321 | 0.1382 | -8.5e-05 | -0.006268 | -0.006268 | -0.8886
2019-04-09 15:25 | 1554837916.7092 | -42 | 0.20951 | 0.000451 | 1e-06 | 100.6738 | 101.4744 | 121.8304 | 20.6927 | 0.001897 | 0.001883 | -154.524 | 0.23219 | 8.4023 | 0.002892 | 0.004627 | 0.004471 | 3.9e-05 | 0.001281 | 0.29535 | 1.4012 | 2.5477 | 98.46 | 20.04 | 48.21 | 66.9262 | 90.8377 | 66.583 | 0.011488 | 1.0392 | 0.024133 | 23.2232 | -0.070847 | 0.13813 | 0.00011 | 0.007899 | 0.007899 | 0.70471
2019-04-09 15:25 | 1554837917.7089 | -41 | 0.20951 | 0.000451 | 2e-06 | 100.6425 | 101.5814 | 121.8956 | 20.692 | 0.001916 | 0.001808 | -154.529 | 0.23219 | 8.4009 | 0.002867 | 0.004482 | 0.004302 | 4.1e-05 | 0.001267 | 0.25841 | 1.3635 | 2.5477 | 98.46 | 20.03 | 48.204 | 67.3711 | 90.2509 | 64.7669 | 0.01148 | 1.0448 | 0.024204 | 23.1672 | -0.083054 | 0.13813 | 9.3e-05 | 0.006622 | 0.006622 | 0.84066
2019-04-09 15:25 | 1554837918.7103 | -40 | 0.20949 | 0.000451 | 2e-06 | 100.5946 | 101.3942 | 121.9247 | 20.6921 | 0.001902 | 0.00185 | -154.5098 | 0.23222 | 8.4024 | 0.002825 | 0.003713 | 0.003607 | 3.6e-05 | 0.001266 | 0.28479 | 1.4028 | 2.5477 | 98.45 | 20.04 | 48.204 | 67.0403 | 89.2289 | 65.7988 | 0.011488 | 1.0416 | 0.02417 | 23.2053 | -0.070816 | 0.13815 | 0.000217 | 0.015652 | 0.015652 | 0.35571
2019-04-09 15:25 | 1554837919.7093 | -39 | 0.20953 | 0.00045 | 1e-06 | 100.6705 | 101.4465 | 121.9456 | 20.6922 | 0.00196 | 0.001863 | -154.507 | 0.23204 | 8.3999 | 0.002778 | 0.00437 | 0.004192 | 4.2e-05 | 0.001261 | 0.27544 | 1.3479 | 2.5475 | 98.45 | 20.04 | 48.205 | 68.4236 | 88.1035 | 66.0931 | 0.011488 | 1.0407 | 0.024162 | 23.2174 | -0.06868 | 0.13804 | 6e-06 | 0.000342 | 0.000342 | 16.256
2019-04-09 15:25 | 1554837920.7105 | -38 | 0.20951 | 0.00045 | 1e-06 | 100.546 | 101.2138 | 122.0881 | 20.6923 | 0.001915 | 0.001855 | -154.5166 | 0.23157 | 8.4036 | 0.002826 | 0.004199 | 0.003999 | 3.8e-05 | 0.001264 | 0.2912 | 1.3446 | 2.5475 | 98.45 | 20.03 | 48.206 | 67.3572 | 89.2479 | 65.9145 | 0.011481 | 1.0412 | 0.024182 | 23.2248 | -0.05921 | 0.13776 | 0.0001 | 0.007187 | 0.007187 | 0.77247
2019-04-09 15:25 | 1554837921.71 | -37 | 0.20956 | 0.000451 | 1e-06 | 100.6765 | 101.4183 | 122.0344 | 20.6919 | 0.001908 | 0.00175 | -154.519 | 0.23226 | 8.4026 | 0.002792 | 0.005101 | 0.004922 | 4e-05 | 0.001265 | 0.29142 | 1.3606 | 2.5476 | 98.45 | 20.03 | 48.205 | 67.1938 | 88.4227 | 63.3865 | 0.011481 | 1.0491 | 0.024268 | 23.133 | -0.065662 | 0.13817 | -0.000205 | -0.014966 | -0.014966 | -0.37207
2019-04-09 15:25 | 1554837922.71 | -36 | 0.20954 | 0.00045 | 1e-06 | 100.6038 | 101.294 | 122.0675 | 20.6926 | 0.001978 | 0.001832 | -154.5054 | 0.23189 | 8.3999 | 0.002852 | 0.004352 | 0.004211 | 3.6e-05 | 0.001265 | 0.2855 | 1.3527 | 2.5475 | 98.45 | 20.03 | 48.2 | 68.8723 | 89.8685 | 65.3538 | 0.01148 | 1.043 | 0.0242 | 23.2036 | -0.061156 | 0.13795 | -4e-05 | -0.00296 | -0.00296 | -1.8782
2019-04-09 15:25 | 1554837923.71 | -35 | 0.20952 | 0.000449 | 1e-06 | 100.6014 | 101.3419 | 122.0726 | 20.6925 | 0.001985 | 0.001784 | -154.4944 | 0.2317 | 8.4 | 0.002755 | 0.004163 | 0.004021 | 3.9e-05 | 0.00127 | 0.30541 | 1.3302 | 2.5476 | 98.45 | 20.03 | 48.199 | 69.0267 | 87.536 | 64.2084 | 0.01148 | 1.0465 | 0.024242 | 23.1648 | -0.065591 | 0.13784 | 3.6e-05 | 0.00251 | 0.00251 | 2.2131
2019-04-09 15:25 | 1554837924.7091 | -34 | 0.20951 | 0.000452 | 1e-06 | 100.6175 | 101.4709 | 122.2393 | 20.6928 | 0.001875 | 0.001831 | -154.4902 | 0.232 | 8.4015 | 0.002819 | 0.004705 | 0.004527 | 3.7e-05 | 0.001286 | 0.28335 | 1.3744 | 2.5477 | 98.45 | 20.03 | 48.198 | 66.3919 | 89.0865 | 65.3352 | 0.01148 | 1.043 | 0.024218 | 23.2193 | -0.075545 | 0.13801 | 9.4e-05 | 0.006748 | 0.006748 | 0.82426
2019-04-09 15:25 | 1554837925.7085 | -33 | 0.20952 | 0.000449 | 1e-06 | 100.6541 | 101.3704 | 121.978 | 20.6927 | 0.001916 | 0.001823 | -154.5127 | 0.2318 | 8.3993 | 0.002782 | 0.00364 | 0.003495 | 4e-05 | 0.001304 | 0.27793 | 1.3983 | 2.5476 | 98.45 | 20.03 | 48.194 | 67.3856 | 88.1836 | 65.1376 | 0.011479 | 1.0436 | 0.024199 | 23.1877 | -0.063427 | 0.1379 | 5.2e-05 | 0.003677 | 0.003677 | 1.5112
2019-04-09 15:25 | 1554837926.7083 | -32 | 0.20954 | 0.00045 | 2e-06 | 100.6038 | 101.4045 | 121.7865 | 20.693 | 0.001908 | 0.001834 | -154.5109 | 0.23245 | 8.3996 | 0.002801 | 0.004116 | 0.003926 | 3.8e-05 | 0.001309 | 0.27936 | 1.3543 | 2.5476 | 98.45 | 20.03 | 48.191 | 67.1927 | 88.6619 | 65.4088 | 0.011478 | 1.0428 | 0.024171 | 23.1788 | -0.070906 | 0.13829 | -4.4e-05 | -0.00331 | -0.00331 | -1.6839
2019-04-09 15:25 | 1554837927.7091 | -31 | 0.20952 | 0.000451 | 2e-06 | 100.6479 | 101.4523 | 122.1097 | 20.6924 | 0.001949 | 0.001861 | -154.5207 | 0.23147 | 8.4029 | 0.002752 | 0.004067 | 0.003875 | 3.9e-05 | 0.001311 | 0.2855 | 1.3252 | 2.5475 | 98.45 | 20.03 | 48.189 | 68.1788 | 87.4632 | 66.055 | 0.011477 | 1.0408 | 0.02418 | 23.2317 | -0.071207 | 0.1377 | 3.5e-05 | 0.002431 | 0.002431 | 2.2825
2019-04-09 15:25 | 1554837928.7086 | -30 | 0.20953 | 0.000447 | 1e-06 | 100.5904 | 101.4878 | 122.2733 | 20.6934 | 0.001948 | 0.001834 | -154.5332 | 0.23205 | 8.4006 | 0.002825 | 0.004054 | 0.003835 | 3.9e-05 | 0.001307 | 0.30342 | 1.3701 | 2.5477 | 98.45 | 20.03 | 48.184 | 68.143 | 89.2388 | 65.395 | 0.011476 | 1.0428 | 0.024219 | 23.2246 | -0.079441 | 0.13805 | 3e-06 | 0.000104 | 0.000104 | 53.2581
2019-04-09 15:25 | 1554837929.7083 | -29 | 0.20956 | 0.00045 | 2e-06 | 100.5565 | 101.2655 | 121.9731 | 20.6928 | 0.001908 | 0.001746 | -154.5277 | 0.23137 | 8.402 | 0.002797 | 0.003994 | 0.003762 | 3.8e-05 | 0.001312 | 0.28102 | 1.3075 | 2.5476 | 98.45 | 20.03 | 48.182 | 67.1776 | 88.5572 | 63.2889 | 0.011476 | 1.0494 | 0.024265 | 23.1238 | -0.062844 | 0.13764 | -0.000191 | -0.013968 | -0.013968 | -0.39712
2019-04-09 15:26 | 1554837930.7083 | -28 | 0.20954 | 0.000451 | 2e-06 | 100.6664 | 101.3811 | 121.5072 | 20.6926 | 0.00192 | 0.001811 | -154.5342 | 0.23158 | 8.4018 | 0.00279 | 0.004105 | 0.004032 | 3.6e-05 | 0.001311 | 0.31852 | 1.3306 | 2.5475 | 98.45 | 20.03 | 48.177 | 67.4691 | 88.3803 | 64.8455 | 0.011475 | 1.0445 | 0.024163 | 23.1329 | -0.063283 | 0.13777 | -3.7e-05 | -0.002809 | -0.002809 | -1.9764
2019-04-09 15:26 | 1554837931.7094 | -27 | 0.20958 | 0.000451 | 2e-06 | 100.6628 | 101.3198 | 121.6043 | 20.6937 | 0.001952 | 0.001778 | -154.5514 | 0.23209 | 8.4013 | 0.002711 | 0.00446 | 0.004332 | 3.7e-05 | 0.001308 | 0.30502 | 1.3835 | 2.5476 | 98.45 | 20.03 | 48.175 | 68.251 | 86.4887 | 64.0683 | 0.011474 | 1.0469 | 0.0242 | 23.1156 | -0.058194 | 0.13807 | -0.000284 | -0.020676 | -0.020676 | -0.26912
2019-04-09 15:26 | 1554837932.7084 | -26 | 0.20955 | 0.000451 | 2e-06 | 100.6169 | 101.3102 | 121.8762 | 20.693 | 0.00193 | 0.001811 | -154.5636 | 0.2319 | 8.4011 | 0.002781 | 0.004669 | 0.004483 | 3.4e-05 | 0.001309 | 0.2844 | 1.3803 | 2.5476 | 98.45 | 20.03 | 48.173 | 67.7099 | 88.1729 | 64.8552 | 0.011474 | 1.0445 | 0.024199 | 23.1684 | -0.06142 | 0.13796 | -9.3e-05 | -0.006863 | -0.006863 | -0.81005
2019-04-09 15:26 | 1554837933.7083 | -25 | 0.20955 | 0.000451 | 1e-06 | 100.6154 | 101.3601 | 122.4029 | 20.6932 | 0.001963 | 0.001806 | -154.5616 | 0.23287 | 8.4041 | 0.0028 | 0.00387 | 0.003732 | 3.5e-05 | 0.001306 | 0.27831 | 1.3548 | 2.5476 | 98.45 | 20.03 | 48.171 | 68.5087 | 88.6247 | 64.7323 | 0.011473 | 1.0449 | 0.024256 | 23.2142 | -0.065957 | 0.13853 | -0.000139 | -0.010212 | -0.010212 | -0.54669
2019-04-09 15:26 | 1554837934.7098 | -24 | 0.20949 | 0.000451 | 1e-06 | 100.6705 | 101.3966 | 122.2616 | 20.6942 | 0.001939 | 0.00174 | -154.553 | 0.23158 | 8.4014 | 0.002784 | 0.004463 | 0.004263 | 3.9e-05 | 0.001311 | 0.30773 | 1.3576 | 2.5475 | 98.45 | 20.03 | 48.171 | 67.9379 | 88.253 | 63.1364 | 0.011473 | 1.0498 | 0.024299 | 23.1459 | -0.064279 | 0.13776 | 0.000231 | 0.016732 | 0.016732 | 0.33182
2019-04-09 15:26 | 1554837935.7095 | -23 | 0.20954 | 0.00045 | 1e-06 | 100.6732 | 101.4706 | 122.3864 | 20.6936 | 0.001923 | 0.001776 | -154.539 | 0.23151 | 8.4003 | 0.002718 | 0.00351 | 0.003406 | 3.6e-05 | 0.001304 | 0.28224 | 1.3682 | 2.5476 | 98.45 | 20.02 | 48.17 | 67.5457 | 86.6478 | 64.0018 | 0.011466 | 1.0471 | 0.024281 | 23.1875 | -0.070566 | 0.13772 | -7.3e-05 | -0.00535 | -0.00535 | -1.0375
2019-04-09 15:26 | 1554837936.7088 | -22 | 0.20954 | 0.00045 | 2e-06 | 100.658 | 101.4837 | 122.9967 | 20.6939 | 0.001927 | 0.001796 | -154.549 | 0.23246 | 8.4027 | 0.002758 | 0.004362 | 0.004184 | 3.5e-05 | 0.001295 | 0.27417 | 1.3855 | 2.5475 | 98.45 | 20.02 | 48.17 | 67.6343 | 87.6115 | 64.4839 | 0.011466 | 1.0456 | 0.024324 | 23.2619 | -0.073068 | 0.13829 | -7.6e-05 | -0.005674 | -0.005674 | -0.9822
2019-04-09 15:26 | 1554837937.7089 | -21 | 0.20955 | 0.000452 | 1e-06 | 100.6381 | 101.4331 | 123.0066 | 20.6938 | 0.001874 | 0.001788 | -154.539 | 0.23153 | 8.403 | 0.002699 | 0.004036 | 0.00384 | 3.5e-05 | 0.001296 | 0.28888 | 1.3607 | 2.5476 | 98.45 | 20.02 | 48.17 | 66.3592 | 86.2068 | 64.305 | 0.011466 | 1.0462 | 0.024331 | 23.2566 | -0.070381 | 0.13774 | -0.000121 | -0.008877 | -0.008877 | -0.62527
2019-04-09 15:26 | 1554837938.7085 | -20 | 0.20954 | 0.000452 | 2e-06 | 100.7048 | 101.3887 | 123.2068 | 20.6941 | 0.001921 | 0.001785 | -154.5399 | 0.23225 | 8.4022 | 0.002717 | 0.003073 | 0.002966 | 3.7e-05 | 0.001284 | 0.30326 | 1.3725 | 2.5477 | 98.45 | 20.02 | 48.171 | 67.4962 | 86.6233 | 64.2343 | 0.011466 | 1.0464 | 0.024353 | 23.2731 | -0.060539 | 0.13817 | -5.2e-05 | -0.003938 | -0.003938 | -1.4141
2019-04-09 15:26 | 1554837939.7086 | -19 | 0.20956 | 0.000449 | 1e-06 | 100.594 | 101.3298 | 123.4406 | 20.6939 | 0.001898 | 0.001764 | -154.5474 | 0.23204 | 8.3997 | 0.002764 | 0.003457 | 0.003331 | 3.9e-05 | 0.001272 | 0.29319 | 1.3438 | 2.5476 | 98.45 | 20.02 | 48.17 | 66.9384 | 87.7607 | 63.7193 | 0.011466 | 1.048 | 0.024395 | 23.2774 | -0.06519 | 0.13804 | -0.000196 | -0.014409 | -0.014409 | -0.3861
2019-04-09 15:26 | 1554837940.7096 | -18 | 0.20953 | 0.00045 | 2e-06 | 100.655 | 101.3336 | 123.692 | 20.6925 | 0.001941 | 0.001748 | -154.5334 | 0.23209 | 8.4017 | 0.002768 | 0.003714 | 0.003538 | 3.4e-05 | 0.001265 | 0.27544 | 1.343 | 2.5476 | 98.45 | 20.02 | 48.168 | 67.9791 | 87.8578 | 63.3262 | 0.011465 | 1.0492 | 0.024434 | 23.2875 | -0.060098 | 0.13807 | -1.2e-05 | -0.00099 | -0.00099 | -5.6183
2019-04-09 15:26 | 1554837941.7086 | -17 | 0.20955 | 0.00045 | 1e-06 | 100.608 | 101.4172 | 123.7845 | 20.6937 | 0.001898 | 0.00175 | -154.5138 | 0.23119 | 8.4014 | 0.002836 | 0.004667 | 0.004509 | 3.6e-05 | 0.001257 | 0.3115 | 1.33 | 2.5475 | 98.45 | 20.02 | 48.168 | 66.9539 | 89.4899 | 63.3938 | 0.011465 | 1.049 | 0.024441 | 23.2985 | -0.071658 | 0.13753 | -8.8e-05 | -0.006527 | -0.006527 | -0.84914
2019-04-09 15:26 | 1554837942.7091 | -16 | 0.20955 | 0.000452 | 1e-06 | 100.644 | 101.3226 | 123.7267 | 20.6934 | 0.001966 | 0.001761 | -154.509 | 0.23219 | 8.4015 | 0.002797 | 0.004595 | 0.00438 | 3.6e-05 | 0.001253 | 0.29983 | 1.3205 | 2.5477 | 98.45 | 20.01 | 48.17 | 68.5846 | 88.5567 | 63.6394 | 0.011459 | 1.0483 | 0.024426 | 23.3016 | -0.060105 | 0.13813 | -0.000115 | -0.008507 | -0.008507 | -0.65438
2019-04-09 15:26 | 1554837943.7091 | -15 | 0.20955 | 0.000454 | 2e-06 | 100.6208 | 101.2892 | 123.5934 | 20.6941 | 0.001898 | 0.001748 | -154.5241 | 0.23158 | 8.4017 | 0.002846 | 0.003651 | 0.003611 | 3.6e-05 | 0.001246 | 0.27815 | 1.355 | 2.5476 | 98.45 | 20.01 | 48.184 | 66.9431 | 89.7442 | 63.3339 | 0.011462 | 1.0492 | 0.024424 | 23.2785 | -0.059222 | 0.13777 | -0.000142 | -0.010524 | -0.010524 | -0.52757
2019-04-09 15:26 | 1554837944.7087 | -14 | 0.20953 | 0.000455 | 1e-06 | 100.5901 | 101.1808 | 123.7492 | 20.6935 | 0.001873 | 0.001791 | -154.5324 | 0.23147 | 8.4032 | 0.002922 | 0.004444 | 0.004243 | 3.7e-05 | 0.001249 | 0.26864 | 1.361 | 2.5477 | 98.45 | 20.01 | 48.196 | 66.3464 | 91.5584 | 64.3718 | 0.011465 | 1.046 | 0.024402 | 23.329 | -0.05237 | 0.1377 | -1.6e-05 | -0.001252 | -0.001252 | -4.4331
2019-04-09 15:26 | 1554837945.7082 | -13 | 0.20954 | 0.00045 | 2e-06 | 100.5833 | 101.2569 | 124.0894 | 20.6944 | 0.001889 | 0.001713 | -154.5473 | 0.23153 | 8.4006 | 0.00295 | 0.004694 | 0.004574 | 3.8e-05 | 0.001242 | 0.30397 | 1.3386 | 2.5475 | 98.45 | 20.01 | 48.207 | 66.732 | 92.2505 | 62.4991 | 0.011467 | 1.0518 | 0.024504 | 23.2962 | -0.059701 | 0.13774 | -3.8e-05 | -0.0029 | -0.0029 | -1.9141
2019-04-09 15:26 | 1554837946.7078 | -12 | 0.20955 | 0.000448 | 1e-06 | 100.6324 | 101.3973 | 124.0384 | 20.6932 | 0.001814 | 0.001755 | -153.0754 | 0.2313 | 8.4001 | 0.002929 | 0.00415 | 0.004061 | 3.6e-05 | 0.001241 | 0.29175 | 1.3491 | 2.5476 | 98.45 | 20.01 | 48.233 | 64.9329 | 91.7452 | 63.5025 | 0.011474 | 1.0487 | 0.024462 | 23.3262 | -0.067727 | 0.1376 | -0.000116 | -0.008605 | -0.008605 | -0.64442
2019-04-09 15:26 | 1554837947.7082 | -11 | 0.20953 | 0.00045 | 1e-06 | 100.6247 | 101.3326 | 124.1842 | 20.6938 | 0.00189 | 0.001736 | -121.2929 | 0.2322 | 8.3986 | 0.002917 | 0.004255 | 0.00405 | 3.7e-05 | 0.001244 | 0.29574 | 1.3713 | 2.5476 | 98.45 | 20.01 | 48.244 | 66.7478 | 91.4375 | 63.0421 | 0.011476 | 1.0501 | 0.024493 | 23.3239 | -0.062706 | 0.13814 | -9e-06 | -0.000733 | -0.000733 | -7.5947
2019-04-09 15:26 | 1554837948.7078 | -10 | 0.20957 | 0.000452 | 2e-06 | 100.6125 | 101.2286 | 124.2886 | 20.6936 | 0.001876 | 0.001784 | -88.4645 | 0.23197 | 8.4004 | 0.002882 | 0.004354 | 0.004159 | 3.6e-05 | 0.001242 | 0.29856 | 1.3618 | 2.5476 | 98.45 | 20.01 | 48.263 | 66.4059 | 90.6056 | 64.2119 | 0.011481 | 1.0465 | 0.024461 | 23.3743 | -0.054612 | 0.138 | -0.000231 | -0.01702 | -0.01702 | -0.32675
2019-04-09 15:26 | 1554837949.7079 | -9 | 0.20961 | 0.000454 | 2e-06 | 100.6116 | 101.3722 | 124.5011 | 20.6935 | 0.00183 | 0.001698 | -55.6285 | 0.23235 | 8.4026 | 0.002932 | 0.004467 | 0.004353 | 3.9e-05 | 0.001244 | 0.29142 | 1.3625 | 2.5477 | 98.45 | 20.01 | 48.272 | 65.3061 | 91.8103 | 62.139 | 0.011483 | 1.053 | 0.024557 | 23.3223 | -0.067364 | 0.13822 | -0.000477 | -0.035197 | -0.035197 | -0.15826
2019-04-09 15:26 | 1554837950.7078 | -8 | 0.2096 | 0.000445 | 2e-06 | 100.5717 | 101.2992 | 124.6595 | 20.6937 | 0.001858 | 0.001747 | -22.8016 | 0.23202 | 8.4018 | 0.002934 | 0.004637 | 0.004471 | 3.8e-05 | 0.001247 | 0.28512 | 1.3563 | 2.5476 | 98.45 | 20.01 | 48.279 | 65.9799 | 91.8474 | 63.3062 | 0.011485 | 1.0493 | 0.02453 | 23.3777 | -0.064472 | 0.13803 | -0.000432 | -0.031914 | -0.031914 | -0.17429
2019-04-09 15:26 | 1554837951.7089 | -7 | 0.20957 | 0.000446 | 2e-06 | 100.6211 | 101.2046 | 124.6515 | 20.6947 | 0.001873 | 0.001711 | 10.0383 | 0.23323 | 8.4015 | 0.002941 | 0.004379 | 0.004267 | 3.9e-05 | 0.001243 | 0.26897 | 1.3335 | 2.5475 | 98.45 | 20.01 | 48.292 | 66.3337 | 92.0289 | 62.4371 | 0.011488 | 1.052 | 0.024561 | 23.3467 | -0.051716 | 0.13875 | -0.000208 | -0.01541 | -0.01541 | -0.36285
2019-04-09 15:26 | 1554837952.7079 | -6 | 0.20961 | 0.000449 | 2e-06 | 100.6119 | 101.0755 | 124.6335 | 20.6933 | 0.001908 | 0.00174 | 42.8792 | 0.25663 | 8.4009 | 0.002918 | 0.004589 | 0.004451 | 3.8e-05 | 0.001238 | 0.286 | 1.3448 | 2.5476 | 98.45 | 20.01 | 48.299 | 67.1968 | 91.459 | 63.1349 | 0.011489 | 1.0498 | 0.024534 | 23.3693 | -0.041123 | 0.15267 | -0.000451 | -0.033237 | -0.033237 | -0.18511
2019-04-09 15:26 | 1554837953.7089 | -5 | 0.20959 | 0.00045 | 2e-06 | 100.6131 | 101.3636 | 124.3262 | 20.6934 | 0.001876 | 0.001749 | 173.7373 | 0.25725 | 8.4031 | 0.002999 | 0.004951 | 0.004853 | 3.8e-05 | 0.001243 | 0.28528 | 1.3698 | 2.5476 | 98.45 | 20.01 | 48.307 | 66.4103 | 93.4142 | 63.3528 | 0.011491 | 1.0492 | 0.024496 | 23.348 | -0.066472 | 0.15304 | -0.000355 | -0.026169 | -0.026169 | -0.23568
2019-04-09 15:26 | 1554837954.709 | -4 | 0.20959 | 0.000451 | 2e-06 | 100.6568 | 101.2613 | 124.3973 | 20.6941 | 0.001809 | 0.001672 | 173.7812 | 0.25864 | 8.4018 | 0.002993 | 0.005041 | 0.00494 | 4.2e-05 | 0.001247 | 0.28982 | 1.3295 | 2.5475 | 98.45 | 20.01 | 48.307 | 64.8017 | 93.2835 | 61.5112 | 0.011491 | 1.0549 | 0.02457 | 23.2907 | -0.053557 | 0.15387 | -0.000364 | -0.026926 | -0.026926 | -0.23029
2019-04-09 15:26 | 1554837955.709 | -3 | 0.20961 | 0.000453 | 1e-06 | 100.6497 | 101.3488 | 124.3213 | 20.693 | 0.001873 | 0.001701 | 173.7856 | 0.25779 | 8.4043 | 0.003023 | 0.004788 | 0.004749 | 3.9e-05 | 0.001242 | 0.27759 | 1.3343 | 2.5476 | 98.45 | 20.01 | 48.304 | 66.3476 | 93.9939 | 62.204 | 0.011491 | 1.0528 | 0.024537 | 23.3077 | -0.061911 | 0.15336 | -0.00045 | -0.033147 | -0.033147 | -0.18645
2019-04-09 15:26 | 1554837956.7084 | -2 | 0.20954 | 0.000449 | 1e-06 | 100.5809 | 101.2489 | 123.7535 | 20.6927 | 0.001866 | 0.001748 | 174.049 | 0.25705 | 8.4045 | 0.003055 | 0.004674 | 0.004651 | 3.4e-05 | 0.00123 | 0.30397 | 1.3223 | 2.5476 | 98.45 | 20.01 | 48.296 | 66.1833 | 94.7621 | 63.3357 | 0.011489 | 1.0492 | 0.02444 | 23.2936 | -0.059213 | 0.15292 | -8.2e-05 | -0.006092 | -0.006092 | -1.0117
2019-04-09 15:26 | 1554837957.7071 | -1 | 0.20954 | 0.000451 | 1e-06 | 100.6053 | 101.4444 | 122.3951 | 20.6938 | 0.001853 | 0.001702 | 173.8675 | 0.258 | 8.4031 | 0.003025 | 0.004301 | 0.004274 | 3.7e-05 | 0.001219 | 0.27256 | 1.3377 | 2.5476 | 98.45 | 20.01 | 48.291 | 65.8645 | 94.04 | 62.2337 | 0.011487 | 1.0527 | 0.024345 | 23.1274 | -0.074294 | 0.15348 | -6.3e-05 | -0.004706 | -0.004706 | -1.3144
2019-04-09 15:26 | 1554837958.7071 | 0 | 0.20954 | 0.000453 | 1e-06 | 100.6005 | 101.2366 | 120.718 | 20.6932 | 0.001873 | 0.001708 | 173.7285 | 0.25735 | 8.4022 | 0.003027 | 0.004127 | 0.003785 | 3.6e-05 | 0.00122 | 0.29789 | 1.3514 | 2.6964 | 98.45 | 20 | 48.283 | 66.3461 | 94.105 | 62.3629 | 0.011478 | 1.0523 | 0.024173 | 22.9729 | -0.056371 | 0.1531 | -5.7e-05 | -0.004237 | -0.004237 | -1.4563
2019-04-09 15:26 | 1554837959.7082 | 1 | 0.20949 | 0.00045 | 3e-06 | 96.2677 | 100.7819 | 119.0024 | 20.6786 | 0.001864 | 0.001821 | 174.2416 | 0.25789 | 8.4033 | 0.003069 | 0.005327 | 0.005224 | 3.8e-05 | 0.001216 | 0.2943 | 1.3934 | 2.5481 | 98.45 | 20 | 48.281 | 66.1041 | 95.0905 | 65.084 | 0.011478 | 1.0438 | 0.023904 | 22.9013 | -0.4099 | 0.15342 | 0.00025 | 0.017714 | 0.017714 | 0.34903
2019-04-09 15:26 | 1554837960.7081 | 2 | 0.20953 | 0.000448 | 7e-06 | 91.8249 | 98.1041 | 117.1164 | 20.6794 | 0.00189 | 0.00181 | 174.0737 | 0.25842 | 8.4008 | 0.003046 | 0.005773 | 0.007357 | 3.9e-05 | 0.001207 | 0.32676 | 1.3498 | 2.5482 | 98.45 | 20 | 48.281 | 66.737 | 94.5417 | 64.811 | 0.011478 | 1.0446 | 0.023724 | 22.71 | -0.59164 | 0.15374 | 9e-06 | 0.000274 | 0.000274 | 22.5871
2019-04-09 15:26 | 1554837961.7076 | 3 | 0.20936 | 0.000449 | 1.2e-05 | 88.115 | 97.529 | 115.0013 | 20.6805 | 0.001828 | 0.001883 | 173.7299 | 0.25829 | 8.4003 | 0.003013 | 0.005311 | 0.005334 | 4e-05 | 0.001213 | 0.33787 | 1.3539 | 2.5476 | 98.45 | 20 | 48.284 | 65.2552 | 93.7548 | 66.5743 | 0.011479 | 1.0392 | 0.023447 | 22.5626 | -0.90793 | 0.15366 | 0.000989 | 0.068778 | 0.068778 | 0.090034
2019-04-09 15:26 | 1554837962.7073 | 4 | 0.20859 | 0.000455 | 2e-05 | 92.076 | 99.1269 | 113.0745 | 20.6821 | 0.001889 | 0.001893 | 174.109 | 0.25744 | 8.403 | 0.002975 | 0.005079 | 0.005728 | 3.7e-05 | 0.001219 | 0.32604 | 1.3334 | 2.5476 | 98.45 | 20 | 48.288 | 66.7146 | 92.84 | 66.8113 | 0.01148 | 1.0385 | 0.023242 | 22.3806 | -0.65999 | 0.15315 | 0.005622 | 0.39025 | 0.39025 | 0.015816
2019-04-09 15:26 | 1554837963.7074 | 5 | 0.20726 | 0.000482 | 3e-05 | 96.2397 | 98.3546 | 111.1641 | 20.6918 | 0.002027 | 0.001973 | 173.9298 | 0.25822 | 8.4026 | 0.002982 | 0.004793 | 0.005041 | 3.8e-05 | 0.001222 | 0.27992 | 1.3172 | 1.8007 | 98.45 | 20 | 48.292 | 70.0491 | 93.0132 | 68.735 | 0.011481 | 1.0326 | 0.02298 | 22.2534 | -0.19443 | 0.15362 | 0.01361 | 0.9343 | 0.9343 | 0.006626
2019-04-09 15:26 | 1554837964.7088 | 6 | 0.20563 | 0.000566 | 4.2e-05 | 98.5413 | 93.7911 | 109.1018 | 20.6935 | 0.002042 | 0.002045 | 173.6914 | 0.25751 | 8.4009 | 0.003037 | 0.004529 | 0.004396 | 3.8e-05 | 0.001218 | 0.28313 | 1.3342 | 2.5474 | 98.45 | 20 | 48.299 | 70.411 | 94.3362 | 70.4703 | 0.011482 | 1.0274 | 0.022708 | 22.1019 | 0.44191 | 0.15319 | 0.023269 | 1.5774 | 1.5774 | 0.003914
2019-04-09 15:26 | 1554837965.7074 | 7 | 0.20483 | 0.001015 | 5.6e-05 | 99.432 | 88.9682 | 106.8107 | 20.6934 | 0.002092 | 0.002114 | 173.5798 | 0.25815 | 8.4025 | 0.003058 | 0.004639 | 0.004665 | 3.3e-05 | 0.001218 | 0.27417 | 1.3353 | 2.5475 | 98.45 | 20 | 48.303 | 71.5992 | 94.8512 | 72.1449 | 0.011483 | 1.0224 | 0.022414 | 21.9218 | 0.99459 | 0.15357 | 0.02748 | 1.8376 | 1.8376 | 0.003368
2019-04-09 15:26 | 1554837966.7084 | 8 | 0.20368 | 0.001736 | 7e-05 | 100.063 | 85.1271 | 105.1615 | 20.6936 | 0.002207 | 0.002178 | 173.504 | 0.25832 | 8.4025 | 0.00303 | 0.004333 | 0.004198 | 3.7e-05 | 0.001218 | 0.27793 | 1.3204 | 2.5476 | 98.45 | 20 | 48.307 | 74.3651 | 94.1807 | 73.6635 | 0.011484 | 1.018 | 0.022191 | 21.7997 | 1.4459 | 0.15367 | 0.033422 | 2.2113 | 2.2113 | 0.002801
2019-04-09 15:26 | 1554837967.7089 | 9 | 0.20281 | 0.002568 | 8.4e-05 | 100.3004 | 81.7034 | 104.5324 | 20.6944 | 0.002259 | 0.002219 | 173.3859 | 0.2575 | 8.3997 | 0.003114 | 0.004111 | 0.004065 | 3.7e-05 | 0.001222 | 0.25442 | 1.3549 | 2.5474 | 98.45 | 20 | 48.307 | 75.6108 | 96.2068 | 74.6707 | 0.011484 | 1.015 | 0.022093 | 21.766 | 1.8343 | 0.15318 | 0.037592 | 2.4747 | 2.4747 | 0.002495
2019-04-09 15:26 | 1554837968.7084 | 10 | 0.20199 | 0.003467 | 9.7e-05 | 100.3805 | 76.7053 | 103.9193 | 20.6949 | 0.002315 | 0.002321 | 173.2521 | 0.25772 | 8.4026 | 0.003139 | 0.004266 | 0.0042 | 4.1e-05 | 0.001216 | 0.28047 | 1.3297 | 2.5474 | 98.45 | 20 | 48.308 | 76.9691 | 96.812 | 77.1116 | 0.011484 | 1.0079 | 0.021951 | 21.778 | 2.4061 | 0.15332 | 0.041374 | 2.7049 | 2.7049 | 0.002284
2019-04-09 15:26 | 1554837969.7085 | 11 | 0.20142 | 0.004366 | 0.000108 | 100.4633 | 73.0965 | 103.4184 | 20.6939 | 0.002397 | 0.002361 | 173.1436 | 0.25852 | 8.4018 | 0.003188 | 0.004956 | 0.004899 | 4e-05 | 0.001211 | 0.26073 | 1.3787 | 2.5474 | 98.45 | 20 | 48.306 | 78.9459 | 97.9749 | 78.0742 | 0.011484 | 1.0052 | 0.021868 | 21.7553 | 2.8445 | 0.15379 | 0.043662 | 2.8427 | 2.8427 | 0.00218
2019-04-09 15:26 | 1554837970.7089 | 12 | 0.20079 | 0.005105 | 0.000117 | 100.5139 | 71.608 | 103.7334 | 20.6941 | 0.002365 | 0.002356 | 173.056 | 0.2575 | 8.4022 | 0.003173 | 0.004309 | 0.004221 | 4.4e-05 | 0.00121 | 0.2656 | 1.4053 | 2.5475 | 98.45 | 19.99 | 48.304 | 78.1781 | 97.6254 | 77.9537 | 0.011476 | 1.0055 | 0.021905 | 21.7847 | 3.033 | 0.15318 | 0.0465 | 3.0317 | 3.0317 | 0.002036
2019-04-09 15:26 | 1554837971.707 | 13 | 0.20038 | 0.00564 | 0.000124 | 100.4877 | 71.4655 | 103.8386 | 20.6943 | 0.002429 | 0.002374 | 172.8986 | 0.25835 | 8.404 | 0.003152 | 0.00503 | 0.004876 | 4.4e-05 | 0.001205 | 0.27079 | 1.4163 | 2.5475 | 98.45 | 19.99 | 48.298 | 79.7009 | 97.118 | 78.3781 | 0.011475 | 1.0043 | 0.021903 | 21.8089 | 3.0485 | 0.15369 | 0.048274 | 3.1463 | 3.1463 | 0.001969
2019-04-09 15:26 | 1554837972.7066 | 14 | 0.20029 | 0.006036 | 0.000129 | 100.513 | 70.5525 | 103.9534 | 20.6946 | 0.002526 | 0.002459 | 172.7669 | 0.25788 | 8.399 | 0.003142 | 0.004807 | 0.004702 | 4.1e-05 | 0.001209 | 0.27328 | 1.3322 | 2.5474 | 98.45 | 19.99 | 48.295 | 82.0389 | 96.8784 | 80.4326 | 0.011474 | 0.99848 | 0.021851 | 21.8846 | 3.1657 | 0.15341 | 0.04835 | 3.1435 | 3.1435 | 0.001967
2019-04-09 15:26 | 1554837973.7066 | 15 | 0.20028 | 0.006268 | 0.000134 | 100.4522 | 70.6761 | 104.2857 | 20.6939 | 0.002546 | 0.002531 | 172.6421 | 0.25717 | 8.401 | 0.003148 | 0.004392 | 0.004385 | 3.9e-05 | 0.001209 | 0.27743 | 1.3215 | 2.5473 | 98.45 | 19.99 | 48.292 | 82.5226 | 97.0195 | 82.149 | 0.011473 | 0.99366 | 0.021833 | 21.9727 | 3.1447 | 0.15299 | 0.048141 | 3.1271 | 3.1271 | 0.001972
2019-04-09 15:26 | 1554837974.7068 | 16 | 0.2002 | 0.006487 | 0.000138 | 100.521 | 71.9708 | 104.6005 | 20.6933 | 0.002562 | 0.002512 | 172.5116 | 0.2584 | 8.4002 | 0.00313 | 0.004561 | 0.004515 | 4.2e-05 | 0.001209 | 0.29175 | 1.3329 | 2.5475 | 98.45 | 19.99 | 48.289 | 82.9105 | 96.5841 | 81.7093 | 0.011473 | 0.99489 | 0.02188 | 21.9922 | 2.9884 | 0.15372 | 0.048326 | 3.1456 | 3.1456 | 0.001969
2019-04-09 15:26 | 1554837975.7067 | 17 | 0.20022 | 0.006683 | 0.000142 | 100.5079 | 72.2747 | 105.0859 | 20.6942 | 0.002654 | 0.002516 | 172.3868 | 0.25729 | 8.4017 | 0.003089 | 0.004628 | 0.004516 | 3.8e-05 | 0.001211 | 0.28528 | 1.3615 | 2.5474 | 98.45 | 19.98 | 48.29 | 85.1184 | 95.5824 | 81.789 | 0.011466 | 0.99467 | 0.021928 | 22.0457 | 2.9496 | 0.15306 | 0.04796 | 3.1285 | 3.1285 | 0.001972
2019-04-09 15:26 | 1554837976.7083 | 18 | 0.20023 | 0.006855 | 0.000145 | 100.4889 | 72.1456 | 105.3378 | 20.6946 | 0.002693 | 0.002573 | 172.2775 | 0.25736 | 8.4022 | 0.003185 | 0.004562 | 0.004511 | 4.2e-05 | 0.001198 | 0.27041 | 1.3292 | 2.5474 | 98.45 | 19.98 | 48.291 | 86.0517 | 97.9089 | 83.1707 | 0.011466 | 0.99081 | 0.021912 | 22.115 | 2.9639 | 0.1531 | 0.047731 | 3.1111 | 3.1111 | 0.001983
2019-04-09 15:26 | 1554837977.7072 | 19 | 0.20023 | 0.006997 | 0.000146 | 100.5032 | 72.7775 | 105.4263 | 20.694 | 0.002739 | 0.002609 | 172.1266 | 0.25666 | 8.4018 | 0.003227 | 0.004228 | 0.004212 | 4.2e-05 | 0.001197 | 0.26106 | 1.3229 | 2.5475 | 98.45 | 19.98 | 48.293 | 87.1642 | 98.9377 | 84.0283 | 0.011466 | 0.98843 | 0.021895 | 22.1509 | 2.8871 | 0.15269 | 0.047561 | 3.0976 | 3.0976 | 0.001986
2019-04-09 15:26 | 1554837978.7069 | 20 | 0.20025 | 0.007116 | 0.000147 | 100.4773 | 73.0562 | 105.4225 | 20.6944 | 0.002745 | 0.002677 | 171.9978 | 0.25753 | 8.4008 | 0.003233 | 0.004659 | 0.004523 | 4.4e-05 | 0.001196 | 0.2808 | 1.3076 | 2.5475 | 98.45 | 19.98 | 48.295 | 87.3152 | 99.0666 | 85.6798 | 0.011467 | 0.98388 | 0.021844 | 22.2017 | 2.8506 | 0.15321 | 0.047259 | 3.0708 | 3.0708 | 0.002011
2019-04-09 15:26 | 1554837979.7069 | 21 | 0.20033 | 0.007218 | 0.000147 | 100.4394 | 71.9763 | 105.8457 | 20.6943 | 0.002764 | 0.002689 | 171.8781 | 0.25787 | 8.4036 | 0.003205 | 0.005577 | 0.005448 | 4.2e-05 | 0.001195 | 0.28241 | 1.3276 | 2.5476 | 98.45 | 19.98 | 48.298 | 87.7596 | 98.3939 | 85.9499 | 0.011468 | 0.98314 | 0.021879 | 22.2545 | 2.9805 | 0.15341 | 0.046649 | 3.0361 | 3.0361 | 0.002036
2019-04-09 15:26 | 1554837980.7079 | 22 | 0.2003 | 0.007297 | 0.000147 | 100.4055 | 71.4473 | 106.0161 | 20.6941 | 0.002788 | 0.002673 | 171.7315 | 0.25763 | 8.4021 | 0.003289 | 0.00471 | 0.004742 | 4.2e-05 | 0.001197 | 0.26842 | 1.3083 | 2.5476 | 98.45 | 19.98 | 48.302 | 88.3506 | 100.4337 | 85.564 | 0.011469 | 0.9842 | 0.021909 | 22.2605 | 3.0434 | 0.15326 | 0.046771 | 3.0482 | 3.0482 | 0.002026
2019-04-09 15:26 | 1554837981.7079 | 23 | 0.20046 | 0.007365 | 0.000146 | 100.4648 | 71.0839 | 106.3342 | 20.6947 | 0.002847 | 0.002759 | 171.6003 | 0.25711 | 8.4016 | 0.003329 | 0.004983 | 0.004924 | 4.2e-05 | 0.001204 | 0.27577 | 1.3195 | 2.5475 | 98.45 | 19.98 | 48.305 | 89.7699 | 101.3935 | 87.6452 | 0.011469 | 0.97852 | 0.021878 | 22.3584 | 3.0943 | 0.15296 | 0.045727 | 2.9762 | 2.9762 | 0.002071
2019-04-09 15:26 | 1554837982.7083 | 24 | 0.20053 | 0.007411 | 0.000145 | 100.4091 | 72.7998 | 106.6344 | 20.6939 | 0.00293 | 0.002772 | 171.5212 | 0.25843 | 8.4039 | 0.003374 | 0.004672 | 0.004488 | 4.4e-05 | 0.001211 | 0.2656 | 1.3395 | 2.5476 | 98.45 | 19.98 | 48.31 | 91.7647 | 102.4782 | 87.9567 | 0.011471 | 0.97768 | 0.0219 | 22.3996 | 2.876 | 0.15374 | 0.045265 | 2.9491 | 2.9491 | 0.002101
2019-04-09 15:26 | 1554837983.7083 | 25 | 0.20064 | 0.007444 | 0.000145 | 100.4469 | 73.3856 | 106.8095 | 20.6953 | 0.002935 | 0.002771 | 171.3863 | 0.25711 | 8.4033 | 0.00338 | 0.005054 | 0.004995 | 4.6e-05 | 0.001217 | 0.30469 | 1.3317 | 2.5475 | 98.45 | 19.98 | 48.318 | 91.8848 | 102.6226 | 87.9424 | 0.011472 | 0.97772 | 0.021918 | 22.4176 | 2.8077 | 0.15295 | 0.044562 | 2.9058 | 2.9058 | 0.002121
2019-04-09 15:26 | 1554837984.7085 | 26 | 0.20072 | 0.007464 | 0.000144 | 100.4225 | 73.3095 | 106.6924 | 20.6952 | 0.002936 | 0.002799 | 171.2496 | 0.25756 | 8.3999 | 0.003337 | 0.004593 | 0.004531 | 4.2e-05 | 0.001215 | 0.28174 | 1.374 | 2.5476 | 98.45 | 19.98 | 48.322 | 91.9081 | 101.5809 | 88.6125 | 0.011473 | 0.9759 | 0.021886 | 22.4261 | 2.8148 | 0.15322 | 0.044029 | 2.867 | 2.867 | 0.002154
2019-04-09 15:26 | 1554837985.7075 | 27 | 0.20089 | 0.007471 | 0.000141 | 100.4505 | 74.225 | 106.7327 | 20.6949 | 0.002897 | 0.002748 | 171.15 | 0.25678 | 8.4019 | 0.00339 | 0.004942 | 0.004833 | 4.8e-05 | 0.001222 | 0.26842 | 1.3298 | 2.5476 | 98.45 | 19.97 | 48.325 | 90.9728 | 102.8798 | 87.3877 | 0.011467 | 0.97922 | 0.021927 | 22.3923 | 2.7063 | 0.15276 | 0.043061 | 2.8095 | 2.8095 | 0.002191
2019-04-09 15:26 | 1554837986.7073 | 28 | 0.20104 | 0.007467 | 0.00014 | 100.4135 | 74.3299 | 106.6564 | 20.6947 | 0.003014 | 0.002836 | 171.0039 | 0.25665 | 8.4017 | 0.003413 | 0.004237 | 0.004162 | 4.8e-05 | 0.001221 | 0.26737 | 1.3402 | 2.5475 | 98.45 | 19.97 | 48.33 | 93.7919 | 103.4216 | 89.497 | 0.011468 | 0.97352 | 0.021855 | 22.4497 | 2.6904 | 0.15268 | 0.042119 | 2.7392 | 2.7392 | 0.002246
2019-04-09 15:26 | 1554837987.7071 | 29 | 0.20116 | 0.007449 | 0.00014 | 100.3692 | 73.7339 | 106.5966 | 20.6942 | 0.003002 | 0.002862 | 170.8727 | 0.25756 | 8.3981 | 0.00336 | 0.004647 | 0.004571 | 4.4e-05 | 0.001226 | 0.29839 | 1.3566 | 2.5476 | 98.45 | 19.98 | 48.33 | 93.4833 | 102.1359 | 90.1156 | 0.011475 | 0.97187 | 0.021831 | 22.4625 | 2.7584 | 0.15322 | 0.041432 | 2.6915 | 2.6915 | 0.002294
2019-04-09 15:26 | 1554837988.7072 | 30 | 0.20117 | 0.007415 | 0.00014 | 100.4067 | 73.0996 | 107.0698 | 20.6943 | 0.003004 | 0.002902 | 170.7799 | 0.25761 | 8.4012 | 0.003333 | 0.005104 | 0.004991 | 4.4e-05 | 0.001228 | 0.28335 | 1.36 | 2.5475 | 98.45 | 19.97 | 48.332 | 93.5468 | 101.4859 | 91.0925 | 0.011469 | 0.96926 | 0.02185 | 22.5426 | 2.8391 | 0.15326 | 0.041443 | 2.6946 | 2.6946 | 0.002292
2019-04-09 15:26 | 1554837989.7062 | 31 | 0.2012 | 0.007372 | 0.000139 | 100.3444 | 74.0415 | 106.6334 | 20.6945 | 0.003084 | 0.002891 | 170.6587 | 0.2579 | 8.4013 | 0.003408 | 0.005035 | 0.004975 | 4.6e-05 | 0.001216 | 0.27489 | 1.3386 | 2.5475 | 98.45 | 19.98 | 48.332 | 95.4786 | 103.3122 | 90.8247 | 0.011476 | 0.96997 | 0.021813 | 22.4883 | 2.719 | 0.15343 | 0.041327 | 2.6826 | 2.6826 | 0.002305
2019-04-09 15:26 | 1554837990.7066 | 32 | 0.20113 | 0.007324 | 0.00014 | 100.4171 | 73.9176 | 106.3814 | 20.6936 | 0.003074 | 0.002943 | 170.5113 | 0.25805 | 8.4014 | 0.00337 | 0.004635 | 0.004585 | 4.5e-05 | 0.001206 | 0.29518 | 1.3644 | 2.5475 | 98.45 | 19.98 | 48.332 | 95.2287 | 102.3899 | 92.0598 | 0.011476 | 0.96669 | 0.02175 | 22.4998 | 2.7404 | 0.15351 | 0.04178 | 2.7042 | 2.7042 | 0.002288
2019-04-09 15:27 | 1554837991.7072 | 33 | 0.20108 | 0.007281 | 0.00014 | 100.3721 | 73.4479 | 106.4243 | 20.6946 | 0.003073 | 0.002919 | 170.4015 | 0.25687 | 8.4018 | 0.00341 | 0.005231 | 0.005156 | 4.4e-05 | 0.001199 | 0.27688 | 1.3821 | 2.5475 | 98.45 | 19.98 | 48.335 | 95.2039 | 103.3603 | 91.494 | 0.011476 | 0.96819 | 0.021772 | 22.4869 | 2.7935 | 0.15281 | 0.042161 | 2.7315 | 2.7315 | 0.002255
2019-04-09 15:27 | 1554837992.7067 | 34 | 0.20106 | 0.007246 | 0.00014 | 100.3915 | 73.6313 | 106.3307 | 20.6938 | 0.003081 | 0.002969 | 170.2551 | 0.25851 | 8.4051 | 0.003389 | 0.004932 | 0.004852 | 4.5e-05 | 0.001198 | 0.28694 | 1.3474 | 2.5475 | 98.45 | 19.97 | 48.335 | 95.3918 | 102.841 | 92.7029 | 0.011469 | 0.96499 | 0.021726 | 22.5142 | 2.7729 | 0.15379 | 0.042279 | 2.7333 | 2.7333 | 0.002267
2019-04-09 15:27 | 1554837993.7076 | 35 | 0.20106 | 0.007218 | 0.000141 | 100.3191 | 73.7968 | 106.0328 | 20.6944 | 0.00321 | 0.002974 | 170.1339 | 0.25782 | 8.4027 | 0.003375 | 0.004994 | 0.004935 | 4.5e-05 | 0.001202 | 0.28766 | 1.301 | 2.5475 | 98.45 | 19.97 | 48.336 | 98.5202 | 102.4968 | 92.8115 | 0.01147 | 0.96471 | 0.021692 | 22.486 | 2.7463 | 0.15338 | 0.042358 | 2.7341 | 2.7341 | 0.002261
2019-04-09 15:27 | 1554837994.7064 | 36 | 0.20105 | 0.007199 | 0.00014 | 100.3022 | 73.9607 | 106.1213 | 20.6945 | 0.003177 | 0.002944 | 170.0317 | 0.25719 | 8.4013 | 0.003381 | 0.004749 | 0.004648 | 4.4e-05 | 0.001203 | 0.26842 | 1.3268 | 2.5475 | 98.45 | 19.97 | 48.336 | 97.7293 | 102.6523 | 92.0966 | 0.01147 | 0.96659 | 0.021723 | 22.4734 | 2.725 | 0.153 | 0.042391 | 2.7402 | 2.7402 | 0.00225
2019-04-09 15:27 | 1554837995.7064 | 37 | 0.20115 | 0.007179 | 0.000139 | 100.323 | 74.056 | 106.1293 | 20.6943 | 0.003224 | 0.002959 | 169.9358 | 0.25782 | 8.4022 | 0.003439 | 0.005009 | 0.004949 | 4.5e-05 | 0.001197 | 0.28407 | 1.3409 | 2.5475 | 98.45 | 19.97 | 48.336 | 98.8509 | 104.0572 | 92.4625 | 0.01147 | 0.96563 | 0.021713 | 22.4855 | 2.7153 | 0.15338 | 0.04185 | 2.704 | 2.704 | 0.002286
2019-04-09 15:27 | 1554837996.7063 | 38 | 0.20125 | 0.007163 | 0.000137 | 100.3099 | 74.6534 | 106.5163 | 20.696 | 0.003232 | 0.002989 | 169.8093 | 0.25795 | 8.399 | 0.003388 | 0.005332 | 0.005227 | 4.6e-05 | 0.001197 | 0.28639 | 1.3236 | 2.5475 | 98.45 | 19.97 | 48.335 | 99.0588 | 102.8354 | 93.1947 | 0.011469 | 0.9637 | 0.02173 | 22.549 | 2.6423 | 0.15345 | 0.041275 | 2.6692 | 2.6692 | 0.002317
2019-04-09 15:27 | 1554837997.7064 | 39 | 0.20139 | 0.00714 | 0.000136 | 100.3126 | 74.5199 | 106.3616 | 20.695 | 0.003248 | 0.002986 | 169.6986 | 0.25749 | 8.4017 | 0.003376 | 0.005178 | 0.00516 | 4.2e-05 | 0.001196 | 0.28943 | 1.3995 | 2.5476 | 98.45 | 19.97 | 48.332 | 99.4454 | 102.5286 | 93.1084 | 0.011469 | 0.96392 | 0.021717 | 22.53 | 2.6585 | 0.15318 | 0.040475 | 2.616 | 2.616 | 0.00236
2019-04-09 15:27 | 1554837998.7065 | 40 | 0.20151 | 0.007107 | 0.000134 | 100.3075 | 74.7292 | 106.3621 | 20.6951 | 0.003275 | 0.003081 | 169.5863 | 0.25716 | 8.4007 | 0.003374 | 0.004748 | 0.004668 | 4.5e-05 | 0.001192 | 0.26521 | 1.364 | 2.5475 | 98.45 | 19.97 | 48.328 | 100.0785 | 102.4801 | 95.3938 | 0.011468 | 0.95795 | 0.02165 | 22.6002 | 2.633 | 0.15298 | 0.03976 | 2.562 | 2.562 | 0.002406
2019-04-09 15:27 | 1554837999.7078 | 41 | 0.20168 | 0.007064 | 0.000132 | 100.3203 | 76.1078 | 106.7797 | 20.6945 | 0.00327 | 0.003045 | 169.4689 | 0.25767 | 8.4013 | 0.003464 | 0.004997 | 0.004927 | 4.3e-05 | 0.001189 | 0.28208 | 1.3424 | 2.5475 | 98.45 | 19.97 | 48.318 | 99.9596 | 104.6732 | 94.5205 | 0.011465 | 0.96022 | 0.021718 | 22.6177 | 2.4706 | 0.15329 | 0.038805 | 2.5085 | 2.5085 | 0.002463
2019-04-09 15:27 | 1554838000.7069 | 42 | 0.20175 | 0.007011 | 0.00013 | 100.3718 | 76.8722 | 106.7459 | 20.6959 | 0.003269 | 0.003041 | 169.3721 | 0.25806 | 8.4008 | 0.003463 | 0.004986 | 0.004899 | 4.5e-05 | 0.001185 | 0.2912 | 1.3546 | 2.5475 | 98.45 | 19.97 | 48.312 | 99.9404 | 104.6518 | 94.434 | 0.011464 | 0.96045 | 0.021717 | 22.6114 | 2.3858 | 0.15352 | 0.038494 | 2.4885 | 2.4885 | 0.002486
2019-04-09 15:27 | 1554838001.7075 | 43 | 0.20191 | 0.006947 | 0.000126 | 100.3462 | 77.3454 | 106.7837 | 20.6947 | 0.003319 | 0.003042 | 169.2585 | 0.25716 | 8.4015 | 0.003519 | 0.004906 | 0.00487 | 4.6e-05 | 0.00118 | 0.26305 | 1.3358 | 2.5474 | 98.45 | 19.97 | 48.299 | 101.148 | 106.0044 | 94.4628 | 0.011461 | 0.96037 | 0.02172 | 22.6163 | 2.3287 | 0.15298 | 0.037611 | 2.432 | 2.432 | 0.002535
2019-04-09 15:27 | 1554838002.7068 | 44 | 0.20211 | 0.006885 | 0.000123 | 100.415 | 78.0182 | 106.9184 | 20.6947 | 0.003254 | 0.00306 | 169.1483 | 0.25737 | 8.405 | 0.00351 | 0.004773 | 0.004651 | 4.2e-05 | 0.001176 | 0.28672 | 1.3135 | 2.5476 | 98.45 | 19.97 | 48.292 | 99.58 | 105.7888 | 94.8964 | 0.011459 | 0.95924 | 0.021721 | 22.6439 | 2.2573 | 0.15311 | 0.036474 | 2.3588 | 2.3588 | 0.002616
2019-04-09 15:27 | 1554838003.7057 | 45 | 0.20217 | 0.006813 | 0.00012 | 100.3039 | 79.6055 | 106.7334 | 20.695 | 0.003291 | 0.003117 | 169.0213 | 0.25726 | 8.4012 | 0.003492 | 0.005106 | 0.00502 | 4.5e-05 | 0.001175 | 0.29518 | 1.3141 | 2.5475 | 98.45 | 19.97 | 48.284 | 100.4677 | 105.3336 | 96.2663 | 0.011457 | 0.95568 | 0.021662 | 22.6664 | 2.0673 | 0.15304 | 0.036214 | 2.3358 | 2.3358 | 0.00264
2019-04-09 15:27 | 1554838004.7056 | 46 | 0.2022 | 0.006736 | 0.000118 | 100.3998 | 79.9995 | 106.7179 | 20.6955 | 0.003316 | 0.003064 | 168.9248 | 0.25738 | 8.4017 | 0.003563 | 0.004957 | 0.004937 | 4.6e-05 | 0.001172 | 0.27328 | 1.3221 | 2.5476 | 98.45 | 19.98 | 48.267 | 101.0819 | 107.0678 | 94.9926 | 0.01146 | 0.95899 | 0.021698 | 22.6257 | 2.0317 | 0.15312 | 0.036106 | 2.3328 | 2.3328 | 0.002645
2019-04-09 15:27 | 1554838005.7061 | 47 | 0.20223 | 0.00666 | 0.000117 | 100.2888 | 81.415 | 106.8682 | 20.6956 | 0.003329 | 0.003081 | 168.7866 | 0.25687 | 8.401 | 0.003539 | 0.005186 | 0.005111 | 4.3e-05 | 0.001176 | 0.29264 | 1.315 | 2.5476 | 98.45 | 19.98 | 48.258 | 101.3913 | 106.4712 | 95.3928 | 0.011458 | 0.95795 | 0.021701 | 22.6539 | 1.8649 | 0.15281 | 0.036031 | 2.3284 | 2.3284 | 0.002645
2019-04-09 15:27 | 1554838006.7057 | 48 | 0.20227 | 0.006584 | 0.000115 | 100.3704 | 81.1387 | 106.2417 | 20.695 | 0.003374 | 0.003153 | 168.6671 | 0.2579 | 8.4028 | 0.00353 | 0.00525 | 0.005138 | 4.6e-05 | 0.001173 | 0.29319 | 1.3218 | 2.5475 | 98.45 | 19.98 | 48.238 | 102.4911 | 106.2644 | 97.1281 | 0.011453 | 0.95346 | 0.021587 | 22.6405 | 1.9026 | 0.15343 | 0.035921 | 2.3092 | 2.3092 | 0.002678
2019-04-09 15:27 | 1554838007.7069 | 49 | 0.20222 | 0.006519 | 0.000115 | 100.3129 | 80.8437 | 106.1112 | 20.6963 | 0.003373 | 0.003157 | 168.5616 | 0.25835 | 8.401 | 0.003541 | 0.005011 | 0.004902 | 4.2e-05 | 0.001177 | 0.2907 | 1.3118 | 2.5476 | 98.45 | 19.98 | 48.228 | 102.4545 | 106.5372 | 97.232 | 0.011451 | 0.95319 | 0.021571 | 22.6298 | 1.93 | 0.15369 | 0.036261 | 2.3293 | 2.3293 | 0.002659
2019-04-09 15:27 | 1554838008.7063 | 50 | 0.20224 | 0.006463 | 0.000115 | 100.3483 | 81.2378 | 105.8755 | 20.6949 | 0.0034 | 0.003141 | 168.443 | 0.25743 | 8.4003 | 0.003547 | 0.00537 | 0.005331 | 4.5e-05 | 0.00118 | 0.27632 | 1.3638 | 2.5476 | 98.45 | 19.98 | 48.216 | 103.1099 | 106.669 | 96.8505 | 0.011448 | 0.95417 | 0.021558 | 22.593 | 1.8897 | 0.15315 | 0.036215 | 2.325 | 2.325 | 0.002655
2019-04-09 15:27 | 1554838009.7072 | 51 | 0.20218 | 0.00642 | 0.000115 | 100.3337 | 81.5327 | 105.7408 | 20.6945 | 0.003396 | 0.003195 | 168.3132 | 0.25748 | 8.3999 | 0.003539 | 0.00496 | 0.00489 | 4.3e-05 | 0.001181 | 0.29894 | 1.3661 | 2.5477 | 98.45 | 19.98 | 48.201 | 103.0262 | 106.4828 | 98.1518 | 0.011445 | 0.95083 | 0.021506 | 22.6183 | 1.856 | 0.15317 | 0.036621 | 2.3453 | 2.3453 | 0.002632
2019-04-09 15:27 | 1554838010.7058 | 52 | 0.20214 | 0.00639 | 0.000115 | 100.2685 | 81.7929 | 105.5131 | 20.6949 | 0.003381 | 0.003188 | 168.1981 | 0.25628 | 8.4017 | 0.00353 | 0.00458 | 0.004455 | 4.4e-05 | 0.001179 | 0.28136 | 1.3592 | 2.5477 | 98.45 | 19.98 | 48.194 | 102.6502 | 106.2738 | 97.9867 | 0.011443 | 0.95125 | 0.021488 | 22.5889 | 1.8217 | 0.15246 | 0.036934 | 2.3634 | 2.3634 | 0.0026
2019-04-09 15:27 | 1554838011.7061 | 53 | 0.20209 | 0.006373 | 0.000114 | 100.306 | 81.4457 | 105.3622 | 20.6945 | 0.00347 | 0.003187 | 168.101 | 0.25786 | 8.4019 | 0.00353 | 0.004662 | 0.004538 | 4.3e-05 | 0.001185 | 0.30718 | 1.3439 | 2.5477 | 98.45 | 19.98 | 48.188 | 104.8085 | 106.2569 | 97.9684 | 0.011442 | 0.9513 | 0.021473 | 22.5722 | 1.8631 | 0.1534 | 0.037215 | 2.3797 | 2.3797 | 0.002598
2019-04-09 15:27 | 1554838012.7065 | 54 | 0.20183 | 0.006365 | 0.000114 | 100.3686 | 81.7905 | 105.039 | 20.6945 | 0.003435 | 0.003208 | 167.9832 | 0.25706 | 8.4005 | 0.003518 | 0.005243 | 0.005248 | 4.5e-05 | 0.001179 | 0.28351 | 1.3478 | 2.5475 | 98.45 | 19.98 | 48.179 | 103.9646 | 105.9806 | 98.4763 | 0.011439 | 0.95 | 0.021425 | 22.5529 | 1.8308 | 0.15293 | 0.0388 | 2.4754 | 2.4754 | 0.00249
2019-04-09 15:27 | 1554838013.706 | 55 | 0.20166 | 0.006375 | 0.000113 | 100.32 | 81.2192 | 105.2171 | 20.6952 | 0.003584 | 0.003262 | 167.8793 | 0.25696 | 8.4022 | 0.003565 | 0.005563 | 0.005482 | 4.7e-05 | 0.001186 | 0.28479 | 1.34 | 2.5476 | 98.45 | 19.98 | 48.175 | 107.5668 | 107.1204 | 99.7845 | 0.011438 | 0.94667 | 0.021406 | 22.6117 | 1.8892 | 0.15287 | 0.039815 | 2.5378 | 2.5378 | 0.002427
2019-04-09 15:27 | 1554838014.707 | 56 | 0.20156 | 0.006399 | 0.000112 | 100.3191 | 81.2722 | 104.9338 | 20.6951 | 0.003499 | 0.003198 | 167.7534 | 0.25681 | 8.4017 | 0.003551 | 0.004717 | 0.004685 | 4.5e-05 | 0.001186 | 0.27959 | 1.3332 | 2.5475 | 98.45 | 19.98 | 48.171 | 105.5185 | 106.7731 | 98.2292 | 0.011438 | 0.95063 | 0.021422 | 22.5341 | 1.8833 | 0.15278 | 0.040349 | 2.5737 | 2.5737 | 0.002392
2019-04-09 15:27 | 1554838015.707 | 57 | 0.20125 | 0.00644 | 0.000111 | 100.2965 | 82.1126 | 104.3705 | 20.6955 | 0.003548 | 0.003278 | 167.6144 | 0.25734 | 8.4045 | 0.00355 | 0.005038 | 0.0049 | 4.2e-05 | 0.001184 | 0.28119 | 1.3279 | 2.5475 | 98.45 | 19.98 | 48.171 | 106.7113 | 106.7436 | 100.1504 | 0.011438 | 0.94574 | 0.021309 | 22.5316 | 1.7893 | 0.15309 | 0.042193 | 2.6771 | 2.6771 | 0.002305
2019-04-09 15:27 | 1554838016.7061 | 58 | 0.20121 | 0.006501 | 0.00011 | 100.2876 | 82.2393 | 104.4766 | 20.6952 | 0.003559 | 0.003208 | 167.5007 | 0.25739 | 8.4017 | 0.003545 | 0.004836 | 0.004769 | 4.2e-05 | 0.001182 | 0.28744 | 1.3599 | 2.5477 | 98.45 | 19.98 | 48.171 | 106.9777 | 106.6164 | 98.4756 | 0.011438 | 0.95 | 0.021368 | 22.4924 | 1.7747 | 0.15312 | 0.042345 | 2.6941 | 2.6941 | 0.002291
2019-04-09 15:27 | 1554838017.7072 | 59 | 0.20117 | 0.006565 | 0.00011 | 100.4043 | 81.3245 | 104.597 | 20.695 | 0.003585 | 0.003307 | 167.401 | 0.25661 | 8.4043 | 0.003484 | 0.005523 | 0.005438 | 4.2e-05 | 0.00118 | 0.28351 | 1.3488 | 2.5476 | 98.45 | 19.98 | 48.173 | 107.5936 | 105.1401 | 100.8577 | 0.011438 | 0.94395 | 0.021312 | 22.5774 | 1.8851 | 0.15266 | 0.042501 | 2.697 | 2.697 | 0.002281
2019-04-09 15:27 | 1554838018.7061 | 60 | 0.20114 | 0.006644 | 0.00011 | 100.3358 | 81.127 | 104.447 | 20.6953 | 0.00359 | 0.003307 | 167.2772 | 0.25735 | 8.401 | 0.003447 | 0.004743 | 0.004656 | 4.2e-05 | 0.001178 | 0.27831 | 1.3345 | 2.5476 | 98.45 | 19.98 | 48.175 | 107.7237 | 104.2477 | 100.8674 | 0.011438 | 0.94393 | 0.021296 | 22.5615 | 1.9008 | 0.1531 | 0.042556 | 2.6985 | 2.6985 | 0.002286
2019-04-09 15:27 | 1554838019.7051 | 61 | 0.20107 | 0.006721 | 0.000109 | 100.3712 | 82.0187 | 104.1928 | 20.6943 | 0.003601 | 0.003344 | 167.1467 | 0.25753 | 8.4015 | 0.003452 | 0.005412 | 0.005244 | 4.2e-05 | 0.001178 | 0.31078 | 1.347 | 2.5475 | 98.45 | 19.98 | 48.179 | 107.9922 | 104.3827 | 101.7472 | 0.011439 | 0.94171 | 0.021246 | 22.5605 | 1.8062 | 0.15321 | 0.042899 | 2.7137 | 2.7137 | 0.002275
2019-04-09 15:27 | 1554838020.7054 | 62 | 0.20091 | 0.006792 | 0.000109 | 100.3123 | 81.3379 | 104.1789 | 20.6947 | 0.003603 | 0.003335 | 167.045 | 0.25655 | 8.4028 | 0.003502 | 0.00479 | 0.004763 | 4.5e-05 | 0.001172 | 0.28263 | 1.3653 | 2.5475 | 98.45 | 19.98 | 48.181 | 108.0334 | 105.5906 | 101.5498 | 0.01144 | 0.94221 | 0.02125 | 22.5531 | 1.8755 | 0.15262 | 0.043807 | 2.7716 | 2.7716 | 0.002219
2019-04-09 15:27 | 1554838021.7062 | 63 | 0.20077 | 0.006863 | 0.000109 | 100.3287 | 80.8699 | 103.9097 | 20.6948 | 0.003682 | 0.003346 | 166.9112 | 0.25732 | 8.4011 | 0.003457 | 0.004953 | 0.004944 | 4.5e-05 | 0.001171 | 0.28263 | 1.3395 | 2.5476 | 98.45 | 19.98 | 48.182 | 109.945 | 104.4847 | 101.805 | 0.01144 | 0.94157 | 0.021215 | 22.5316 | 1.9285 | 0.15308 | 0.044503 | 2.8109 | 2.8109 | 0.002195
2019-04-09 15:27 | 1554838022.7066 | 64 | 0.20069 | 0.006936 | 0.000108 | 100.3855 | 81.6542 | 103.4198 | 20.695 | 0.003679 | 0.003344 | 166.7698 | 0.25724 | 8.4024 | 0.003536 | 0.005076 | 0.004979 | 4.5e-05 | 0.001178 | 0.27367 | 1.3549 | 2.5476 | 98.45 | 19.98 | 48.189 | 109.8932 | 106.4201 | 101.7621 | 0.011442 | 0.94168 | 0.021166 | 22.4771 | 1.8473 | 0.15303 | 0.044924 | 2.8309 | 2.8309 | 0.002179
2019-04-09 15:27 | 1554838023.7066 | 65 | 0.2007 | 0.007011 | 0.000109 | 100.3802 | 82.0748 | 103.1431 | 20.6954 | 0.003715 | 0.003418 | 166.6644 | 0.25733 | 8.4037 | 0.003488 | 0.004561 | 0.00446 | 4.6e-05 | 0.001169 | 0.28655 | 1.3516 | 2.5477 | 98.45 | 19.98 | 48.194 | 110.7613 | 105.253 | 103.5396 | 0.011443 | 0.93723 | 0.021088 | 22.5002 | 1.8008 | 0.15309 | 0.044745 | 2.8092 | 2.8092 | 0.002196
2019-04-09 15:27 | 1554838024.7063 | 66 | 0.2007 | 0.007082 | 0.000108 | 100.3248 | 82.1618 | 103.0605 | 20.695 | 0.003757 | 0.003373 | 166.5396 | 0.257 | 8.4021 | 0.003469 | 0.004971 | 0.004927 | 4.7e-05 | 0.00116 | 0.27096 | 1.365 | 2.5478 | 98.45 | 19.98 | 48.203 | 111.7839 | 104.7933 | 102.4502 | 0.011445 | 0.93995 | 0.02111 | 22.4586 | 1.7864 | 0.15289 | 0.04471 | 2.81 | 2.81 | 0.002193
2019-04-09 15:27 | 1554838025.7057 | 67 | 0.20068 | 0.007148 | 0.000108 | 100.3617 | 82.3047 | 103.2673 | 20.6949 | 0.003698 | 0.003477 | 166.4204 | 0.25668 | 8.4003 | 0.00352 | 0.004951 | 0.004927 | 4.3e-05 | 0.001159 | 0.30431 | 1.3667 | 2.5476 | 98.45 | 19.98 | 48.208 | 110.3372 | 106.0195 | 104.9904 | 0.011446 | 0.93364 | 0.02106 | 22.5571 | 1.7742 | 0.1527 | 0.044744 | 2.8055 | 2.8055 | 0.002194
2019-04-09 15:27 | 1554838026.7054 | 68 | 0.20062 | 0.007207 | 0.000108 | 100.362 | 82.0607 | 103.03 | 20.6944 | 0.003769 | 0.003416 | 166.2981 | 0.25617 | 8.4015 | 0.003467 | 0.005101 | 0.004978 | 4.5e-05 | 0.001158 | 0.27256 | 1.3763 | 2.5477 | 98.45 | 19.98 | 48.212 | 112.0721 | 104.7366 | 103.5008 | 0.011447 | 0.93733 | 0.021077 | 22.4867 | 1.8008 | 0.1524 | 0.045024 | 2.8253 | 2.8253 | 0.002174
2019-04-09 15:27 | 1554838027.7053 | 69 | 0.20058 | 0.007255 | 0.000107 | 100.337 | 82.6192 | 102.9908 | 20.6952 | 0.003788 | 0.003458 | 166.1644 | 0.25715 | 8.4007 | 0.003449 | 0.005208 | 0.005182 | 4.2e-05 | 0.001153 | 0.28528 | 1.3387 | 2.5476 | 98.45 | 19.98 | 48.222 | 112.5261 | 104.2956 | 104.5142 | 0.01145 | 0.93481 | 0.021045 | 22.5126 | 1.7379 | 0.15298 | 0.045167 | 2.83 | 2.83 | 0.002178
2019-04-09 15:27 | 1554838028.7059 | 70 | 0.20059 | 0.007314 | 0.000107 | 100.3906 | 82.9486 | 103.2037 | 20.6952 | 0.003813 | 0.003502 | 166.0083 | 0.25657 | 8.4029 | 0.003446 | 0.004951 | 0.004835 | 4.4e-05 | 0.001151 | 0.28727 | 1.341 | 2.5475 | 98.45 | 19.98 | 48.228 | 113.1352 | 104.2382 | 105.5887 | 0.011451 | 0.93216 | 0.021037 | 22.5679 | 1.707 | 0.15264 | 0.045028 | 2.8202 | 2.8202 | 0.002181
2019-04-09 15:27 | 1554838029.7055 | 71 | 0.20056 | 0.007362 | 0.000107 | 100.284 | 83.556 | 103.1142 | 20.695 | 0.003775 | 0.003472 | 165.9087 | 0.25638 | 8.4015 | 0.003438 | 0.004878 | 0.004774 | 4.3e-05 | 0.001146 | 0.27688 | 1.3662 | 2.5475 | 98.45 | 19.98 | 48.234 | 112.2133 | 104.0264 | 104.8521 | 0.011452 | 0.93398 | 0.021048 | 22.5362 | 1.6323 | 0.15252 | 0.045176 | 2.831 | 2.831 | 0.002171
2019-04-09 15:27 | 1554838030.7064 | 72 | 0.20058 | 0.00741 | 0.000106 | 100.3948 | 83.8121 | 103.1269 | 20.695 | 0.003845 | 0.003454 | 165.7823 | 0.25706 | 8.4032 | 0.003526 | 0.004588 | 0.004567 | 4.2e-05 | 0.001142 | 0.30254 | 1.3324 | 2.5475 | 98.45 | 19.98 | 48.247 | 113.9283 | 106.1785 | 104.4267 | 0.011456 | 0.93503 | 0.021061 | 22.5249 | 1.6148 | 0.15293 | 0.045021 | 2.8231 | 2.8231 | 0.002183
2019-04-09 15:27 | 1554838031.7056 | 73 | 0.20061 | 0.007442 | 0.000104 | 100.3197 | 82.9668 | 103.1513 | 20.6944 | 0.003929 | 0.00351 | 165.6407 | 0.2575 | 8.4015 | 0.003526 | 0.004647 | 0.004569 | 4.4e-05 | 0.001139 | 0.2975 | 1.33 | 2.5476 | 98.45 | 19.98 | 48.254 | 115.97 | 106.1758 | 105.7681 | 0.011457 | 0.93172 | 0.021027 | 22.5676 | 1.6988 | 0.15318 | 0.04477 | 2.8028 | 2.8028 | 0.002203
2019-04-09 15:27 | 1554838032.7055 | 74 | 0.20063 | 0.007469 | 0.000103 | 100.2777 | 83.1988 | 103.5996 | 20.6946 | 0.003815 | 0.003482 | 165.5309 | 0.25718 | 8.3992 | 0.003475 | 0.004896 | 0.004807 | 4.6e-05 | 0.001133 | 0.28351 | 1.3333 | 2.5477 | 98.45 | 19.98 | 48.266 | 113.1829 | 104.9226 | 105.0949 | 0.01146 | 0.93338 | 0.021091 | 22.5964 | 1.67 | 0.153 | 0.044646 | 2.8037 | 2.8037 | 0.002199
2019-04-09 15:27 | 1554838033.7056 | 75 | 0.20066 | 0.007477 | 0.000103 | 100.25 | 82.5838 | 103.5954 | 20.6955 | 0.003827 | 0.003505 | 165.4171 | 0.25623 | 8.4033 | 0.003544 | 0.004912 | 0.004762 | 4.5e-05 | 0.001122 | 0.28495 | 1.3678 | 2.5477 | 98.45 | 19.98 | 48.272 | 113.4896 | 106.6143 | 105.651 | 0.011461 | 0.93201 | 0.021075 | 22.6126 | 1.7339 | 0.15243 | 0.044434 | 2.7883 | 2.7883 | 0.002203
2019-04-09 15:27 | 1554838034.706 | 76 | 0.20068 | 0.007493 | 0.000102 | 100.3394 | 82.7545 | 103.5351 | 20.695 | 0.003822 | 0.003512 | 165.2771 | 0.25675 | 8.4001 | 0.003546 | 0.004744 | 0.004754 | 4.4e-05 | 0.001118 | 0.30613 | 1.3535 | 2.5475 | 98.45 | 19.98 | 48.278 | 113.3674 | 106.6524 | 105.837 | 0.011463 | 0.93155 | 0.021064 | 22.6116 | 1.7234 | 0.15274 | 0.04431 | 2.7791 | 2.7791 | 0.002215
2019-04-09 15:27 | 1554838035.7056 | 77 | 0.20077 | 0.0075 | 0.000101 | 100.3558 | 83.6386 | 103.3269 | 20.6953 | 0.003936 | 0.003529 | 165.1607 | 0.25722 | 8.4025 | 0.003559 | 0.00473 | 0.004595 | 4.4e-05 | 0.001113 | 0.29103 | 1.3856 | 2.5476 | 98.45 | 19.98 | 48.285 | 116.1504 | 106.9591 | 106.2324 | 0.011465 | 0.93058 | 0.021032 | 22.6006 | 1.6298 | 0.15302 | 0.043795 | 2.7428 | 2.7428 | 0.002248
2019-04-09 15:27 | 1554838036.706 | 78 | 0.20092 | 0.007502 | 9.9e-05 | 100.2634 | 84.032 | 103.7365 | 20.6947 | 0.00384 | 0.003538 | 165.0715 | 0.25718 | 8.4047 | 0.003556 | 0.005026 | 0.004939 | 4.4e-05 | 0.00111 | 0.29479 | 1.3564 | 2.5478 | 98.45 | 19.98 | 48.289 | 113.8025 | 106.9041 | 106.4499 | 0.011466 | 0.93005 | 0.021067 | 22.6518 | 1.5796 | 0.153 | 0.042866 | 2.6894 | 2.6894 | 0.002293
2019-04-09 15:27 | 1554838037.7054 | 79 | 0.20093 | 0.007493 | 9.8e-05 | 100.2956 | 85.2221 | 103.755 | 20.6949 | 0.003889 | 0.003524 | 164.9499 | 0.25739 | 8.3999 | 0.003523 | 0.005473 | 0.005296 | 4.1e-05 | 0.001107 | 0.28385 | 1.371 | 2.5476 | 98.45 | 19.98 | 48.296 | 115.0006 | 106.0982 | 106.1188 | 0.011467 | 0.93086 | 0.021078 | 22.644 | 1.4567 | 0.15312 | 0.042822 | 2.688 | 2.688 | 0.002296
2019-04-09 15:27 | 1554838038.7059 | 80 | 0.20092 | 0.007477 | 9.6e-05 | 100.3004 | 85.3026 | 103.1926 | 20.6958 | 0.003901 | 0.003574 | 164.8505 | 0.25696 | 8.4051 | 0.003547 | 0.005358 | 0.005265 | 4.3e-05 | 0.001112 | 0.28296 | 1.3548 | 2.5477 | 98.45 | 19.98 | 48.299 | 115.2876 | 106.6801 | 107.3305 | 0.011468 | 0.92789 | 0.020988 | 22.6186 | 1.4487 | 0.15287 | 0.042875 | 2.6799 | 2.6799 | 0.002299
2019-04-09 15:27 | 1554838039.7058 | 81 | 0.20092 | 0.007441 | 9.5e-05 | 100.3209 | 85.5817 | 103.3006 | 20.6947 | 0.003907 | 0.003558 | 164.7387 | 0.25633 | 8.4032 | 0.003547 | 0.005213 | 0.005198 | 3.8e-05 | 0.001109 | 0.28152 | 1.3127 | 2.5476 | 98.45 | 19.98 | 48.302 | 115.4369 | 106.6797 | 106.9396 | 0.011469 | 0.92885 | 0.021009 | 22.6188 | 1.4213 | 0.15249 | 0.042961 | 2.6881 | 2.6881 | 0.002286
2019-04-09 15:27 | 1554838040.7047 | 82 | 0.20093 | 0.007415 | 9.4e-05 | 100.278 | 85.7906 | 102.8314 | 20.6945 | 0.003925 | 0.003584 | 164.6163 | 0.25551 | 8.4025 | 0.003543 | 0.005241 | 0.00513 | 4.3e-05 | 0.001116 | 0.30038 | 1.3263 | 2.5474 | 98.45 | 19.97 | 48.307 | 115.8718 | 106.5851 | 107.5773 | 0.011463 | 0.92729 | 0.020944 | 22.5863 | 1.3957 | 0.15201 | 0.042921 | 2.6773 | 2.6773 | 0.002288
2019-04-09 15:27 | 1554838041.7058 | 83 | 0.20092 | 0.007389 | 9.3e-05 | 100.3456 | 86.1134 | 103.0708 | 20.694 | 0.003929 | 0.003592 | 164.5149 | 0.25634 | 8.4043 | 0.00364 | 0.004868 | 0.004785 | 4.4e-05 | 0.001116 | 0.28888 | 1.3275 | 2.5474 | 98.45 | 19.97 | 48.311 | 115.9633 | 108.9273 | 107.7625 | 0.011464 | 0.92684 | 0.020963 | 22.618 | 1.3681 | 0.1525 | 0.043022 | 2.6861 | 2.6861 | 0.002288
2019-04-09 15:27 | 1554838042.7058 | 84 | 0.20091 | 0.007365 | 9.3e-05 | 100.2569 | 85.814 | 102.9809 | 20.6939 | 0.003939 | 0.003617 | 164.4129 | 0.25676 | 8.4003 | 0.00355 | 0.004871 | 0.004725 | 4.3e-05 | 0.001118 | 0.29319 | 1.3137 | 2.5476 | 98.45 | 19.97 | 48.316 | 116.2188 | 106.7483 | 108.3864 | 0.011465 | 0.92532 | 0.020937 | 22.6267 | 1.3914 | 0.15275 | 0.04311 | 2.6882 | 2.6882 | 0.00229
2019-04-09 15:27 | 1554838043.7048 | 85 | 0.20089 | 0.007355 | 9.4e-05 | 100.2363 | 86.3505 | 102.8058 | 20.6946 | 0.004004 | 0.003582 | 164.2775 | 0.25626 | 8.4028 | 0.003528 | 0.005248 | 0.0052 | 4.4e-05 | 0.001122 | 0.27936 | 1.3216 | 2.5476 | 98.45 | 19.97 | 48.32 | 117.811 | 106.2199 | 107.5167 | 0.011466 | 0.92744 | 0.020943 | 22.5817 | 1.3338 | 0.15245 | 0.04325 | 2.6976 | 2.6976 | 0.002277
2019-04-09 15:27 | 1554838044.7051 | 86 | 0.20089 | 0.007347 | 9.6e-05 | 100.2885 | 86.0962 | 103.2492 | 20.6941 | 0.003966 | 0.003663 | 164.1725 | 0.25714 | 8.4038 | 0.003511 | 0.004909 | 0.004918 | 4.2e-05 | 0.001119 | 0.27936 | 1.3311 | 2.5476 | 98.45 | 19.97 | 48.324 | 116.8705 | 105.7962 | 109.4855 | 0.011467 | 0.92267 | 0.020934 | 22.6887 | 1.3648 | 0.15297 | 0.043269 | 2.6976 | 2.6976 | 0.002285
2019-04-09 15:27 | 1554838045.7053 | 87 | 0.20076 | 0.007346 | 9.7e-05 | 100.2366 | 85.2358 | 103.5339 | 20.6941 | 0.004054 | 0.003588 | 164.0635 | 0.25694 | 8.4026 | 0.003515 | 0.004963 | 0.004766 | 4.4e-05 | 0.001124 | 0.28479 | 1.3204 | 2.5475 | 98.45 | 19.97 | 48.329 | 119.0079 | 105.8968 | 107.6784 | 0.011468 | 0.92705 | 0.021013 | 22.6663 | 1.45 | 0.15285 | 0.04403 | 2.7552 | 2.7552 | 0.002236
2019-04-09 15:27 | 1554838046.7059 | 88 | 0.20075 | 0.007345 | 9.7e-05 | 100.2882 | 85.1477 | 103.4574 | 20.6952 | 0.003984 | 0.003602 | 163.9355 | 0.25638 | 8.4019 | 0.003534 | 0.005782 | 0.005757 | 4.2e-05 | 0.001118 | 0.29645 | 1.3642 | 2.5476 | 98.45 | 19.97 | 48.333 | 117.3143 | 106.3495 | 108.0001 | 0.011469 | 0.92626 | 0.020996 | 22.6675 | 1.4639 | 0.15252 | 0.044062 | 2.755 | 2.755 | 0.002231
2019-04-09 15:27 | 1554838047.7062 | 89 | 0.20076 | 0.00736 | 9.7e-05 | 100.3138 | 83.9256 | 103.694 | 20.6937 | 0.00404 | 0.003638 | 163.8247 | 0.25647 | 8.4015 | 0.003541 | 0.004955 | 0.004909 | 4.2e-05 | 0.001116 | 0.30198 | 1.3538 | 2.5477 | 98.45 | 19.97 | 48.336 | 118.6762 | 106.5409 | 108.8915 | 0.01147 | 0.9241 | 0.020996 | 22.7199 | 1.5955 | 0.15257 | 0.043996 | 2.7508 | 2.7508 | 0.002235
2019-04-09 15:27 | 1554838048.706 | 90 | 0.20078 | 0.007377 | 9.7e-05 | 100.2626 | 83.9638 | 104.0405 | 20.6946 | 0.004043 | 0.00363 | 163.7084 | 0.25655 | 8.3998 | 0.003566 | 0.004648 | 0.004602 | 4.4e-05 | 0.001113 | 0.28512 | 1.3342 | 2.5477 | 98.45 | 19.97 | 48.344 | 118.7431 | 107.1466 | 108.6804 | 0.011471 | 0.92461 | 0.021036 | 22.7515 | 1.5868 | 0.15262 | 0.043856 | 2.7473 | 2.7473 | 0.002239
2019-04-09 15:27 | 1554838049.7051 | 91 | 0.20093 | 0.007394 | 9.7e-05 | 100.292 | 84.7289 | 104.0541 | 20.6939 | 0.004025 | 0.003676 | 163.5773 | 0.25694 | 8.4017 | 0.003384 | 0.005205 | 0.005185 | 4e-05 | 0.001108 | 0.27815 | 1.3921 | 2.5476 | 98.45 | 19.97 | 48.347 | 118.308 | 102.7193 | 109.8131 | 0.011472 | 0.92188 | 0.021007 | 22.7868 | 1.5083 | 0.15286 | 0.042938 | 2.6862 | 2.6862 | 0.002293
2019-04-09 15:27 | 1554838050.704 | 92 | 0.20092 | 0.007405 | 9.6e-05 | 100.2224 | 85.2513 | 104.1566 | 20.6946 | 0.004055 | 0.003633 | 163.4485 | 0.25678 | 8.4021 | 0.003567 | 0.004682 | 0.00458 | 4.2e-05 | 0.001102 | 0.29446 | 1.3861 | 2.5476 | 98.45 | 19.97 | 48.354 | 119.044 | 107.1701 | 108.7622 | 0.011474 | 0.92441 | 0.021046 | 22.7667 | 1.4471 | 0.15276 | 0.042969 | 2.6932 | 2.6932 | 0.002286
2019-04-09 15:28 | 1554838051.7041 | 93 | 0.20097 | 0.007403 | 9.7e-05 | 100.2858 | 84.8669 | 103.0351 | 20.6944 | 0.004078 | 0.003656 | 163.3098 | 0.25678 | 8.3988 | 0.003476 | 0.004937 | 0.004888 | 3.7e-05 | 0.001102 | 0.29231 | 1.351 | 2.7475 | 98.45 | 19.97 | 48.358 | 119.5991 | 104.9483 | 109.3274 | 0.011475 | 0.92305 | 0.020917 | 22.6605 | 1.4932 | 0.15276 | 0.042677 | 2.6585 | 2.6585 | 0.002316
2019-04-09 15:28 | 1554838052.7041 | 94 | 0.20092 | 0.007392 | 9.8e-05 | 100.3239 | 85.2203 | 100.8961 | 20.6947 | 0.004116 | 0.003687 | 163.189 | 0.25644 | 8.4033 | 0.003501 | 0.00528 | 0.005195 | 3.9e-05 | 0.001098 | 0.27793 | 1.3245 | 2.5476 | 98.45 | 19.97 | 48.362 | 120.5392 | 105.5552 | 110.0648 | 0.011476 | 0.92127 | 0.020679 | 22.4457 | 1.4594 | 0.15256 | 0.042998 | 2.6479 | 2.6479 | 0.002322
2019-04-09 15:28 | 1554838053.7041 | 95 | 0.20099 | 0.007369 | 0.000101 | 100.1821 | 85.4692 | 97.5318 | 20.695 | 0.00413 | 0.003766 | 163.0604 | 0.25637 | 8.4021 | 0.00348 | 0.004634 | 0.004534 | 4e-05 | 0.001095 | 0.27992 | 1.3503 | 2.5476 | 98.45 | 19.96 | 48.369 | 120.8722 | 105.0505 | 112.0041 | 0.01147 | 0.91663 | 0.02028 | 22.1241 | 1.4207 | 0.15252 | 0.042592 | 2.5722 | 2.5722 | 0.00239
2019-04-09 15:28 | 1554838054.7052 | 96 | 0.2009 | 0.007344 | 0.000109 | 100.2911 | 85.1962 | 94.82 | 20.695 | 0.004137 | 0.003898 | 162.8997 | 0.25633 | 8.4045 | 0.003478 | 0.004801 | 0.004752 | 4.2e-05 | 0.001094 | 0.2688 | 1.3588 | 2.5475 | 98.45 | 19.96 | 48.374 | 121.0529 | 104.9944 | 115.2198 | 0.011471 | 0.90904 | 0.019913 | 21.9052 | 1.459 | 0.15249 | 0.043184 | 2.5603 | 2.5603 | 0.0024
2019-04-09 15:28 | 1554838055.7052 | 97 | 0.20072 | 0.007336 | 0.000121 | 100.2444 | 85.1171 | 92.7589 | 20.6939 | 0.004257 | 0.004088 | 162.7176 | 0.25686 | 8.4057 | 0.003528 | 0.004993 | 0.004956 | 3.9e-05 | 0.001091 | 0.26178 | 1.3227 | 2.5476 | 98.45 | 19.96 | 48.383 | 123.9912 | 106.2223 | 119.8601 | 0.011474 | 0.89831 | 0.019579 | 21.7949 | 1.4632 | 0.15281 | 0.044253 | 2.5789 | 2.5789 | 0.002388
2019-04-09 15:28 | 1554838056.7042 | 98 | 0.19986 | 0.007345 | 0.000138 | 100.2456 | 84.8841 | 90.165 | 20.6945 | 0.004443 | 0.004214 | 162.5188 | 0.25678 | 8.4005 | 0.003475 | 0.004801 | 0.004705 | 4e-05 | 0.001091 | 0.28943 | 1.3448 | 2.5476 | 98.45 | 19.96 | 48.389 | 128.5607 | 104.9312 | 122.9291 | 0.011475 | 0.89135 | 0.019228 | 21.5717 | 1.4878 | 0.15276 | 0.049367 | 2.8237 | 2.8237 | 0.00218
2019-04-09 15:28 | 1554838057.7043 | 99 | 0.19503 | 0.007413 | 0.000158 | 100.2524 | 82.6918 | 87.5366 | 20.6949 | 0.004653 | 0.004331 | 162.2802 | 0.25713 | 8.4003 | 0.00352 | 0.004664 | 0.004508 | 4.1e-05 | 0.001101 | 0.29336 | 1.3476 | 2.5475 | 98.45 | 19.96 | 48.394 | 133.7118 | 106.033 | 125.8122 | 0.011476 | 0.88491 | 0.018877 | 21.3322 | 1.7225 | 0.15296 | 0.077804 | 4.359 | 4.359 | 0.001414
2019-04-09 15:28 | 1554838058.7043 | 100 | 0.18892 | 0.007888 | 0.000215 | 100.2748 | 77.6631 | 85.1684 | 20.6947 | 0.00478 | 0.004413 | 162.0354 | 0.25686 | 8.3983 | 0.003492 | 0.005047 | 0.005033 | 4.2e-05 | 0.001101 | 0.28854 | 1.3368 | 2.5475 | 98.45 | 19.96 | 48.407 | 136.8425 | 105.3427 | 127.8028 | 0.011479 | 0.88052 | 0.018574 | 21.0941 | 2.2856 | 0.15281 | 0.11293 | 6.2033 | 6.2033 | 0.000993
2019-04-09 15:28 | 1554838059.7044 | 101 | 0.18375 | 0.009388 | 0.000302 | 100.2441 | 75.1924 | 83.3698 | 20.6937 | 0.004891 | 0.004517 | 161.7885 | 0.25679 | 8.4055 | 0.003577 | 0.004983 | 0.004931 | 3.9e-05 | 0.001101 | 0.28694 | 1.3438 | 2.5475 | 98.45 | 19.96 | 48.413 | 139.5942 | 107.3993 | 130.38 | 0.011481 | 0.87489 | 0.018318 | 20.9371 | 2.5721 | 0.15277 | 0.14104 | 7.6169 | 7.6169 | 0.000808
2019-04-09 15:28 | 1554838060.7054 | 102 | 0.18203 | 0.011688 | 0.000399 | 100.2762 | 78.2464 | 81.8187 | 20.6945 | 0.005083 | 0.004625 | 161.4912 | 0.2562 | 8.4011 | 0.00366 | 0.005152 | 0.005034 | 3.9e-05 | 0.001106 | 0.27599 | 1.3349 | 2.5474 | 98.45 | 19.96 | 48.419 | 144.3415 | 109.4079 | 133.0304 | 0.011482 | 0.86918 | 0.018087 | 20.8094 | 2.2188 | 0.15241 | 0.14838 | 7.9026 | 7.9026 | 0.000777
2019-04-09 15:28 | 1554838061.7044 | 103 | 0.18068 | 0.014515 | 0.000482 | 100.2852 | 80.694 | 80.7319 | 20.6946 | 0.005166 | 0.004704 | 161.2154 | 0.2565 | 8.4031 | 0.003709 | 0.00501 | 0.004977 | 3.8e-05 | 0.00111 | 0.28385 | 1.3543 | 2.5474 | 98.45 | 19.96 | 48.43 | 146.3937 | 110.6168 | 134.9702 | 0.011485 | 0.86505 | 0.017924 | 20.7201 | 1.9441 | 0.15259 | 0.15304 | 8.0702 | 8.0702 | 0.000762
2019-04-09 15:28 | 1554838062.7045 | 104 | 0.17949 | 0.01738 | 0.000542 | 100.2533 | 80.6183 | 80.8847 | 20.6955 | 0.005188 | 0.004739 | 160.885 | 0.25625 | 8.4024 | 0.003738 | 0.005568 | 0.005561 | 3.8e-05 | 0.001116 | 0.29518 | 1.3334 | 2.5474 | 98.45 | 19.96 | 48.436 | 146.9521 | 111.3247 | 135.8498 | 0.011486 | 0.86319 | 0.017922 | 20.762 | 1.9497 | 0.15244 | 0.15679 | 8.2619 | 8.2619 | 0.000744
2019-04-09 15:28 | 1554838063.7044 | 105 | 0.17848 | 0.019967 | 0.000588 | 100.2355 | 79.8657 | 82.5348 | 20.6947 | 0.005252 | 0.004735 | 160.5278 | 0.25663 | 8.4031 | 0.003823 | 0.005682 | 0.00565 | 3.8e-05 | 0.001109 | 0.27632 | 1.3544 | 2.5477 | 98.45 | 19.96 | 48.447 | 148.5445 | 113.3797 | 135.7368 | 0.011489 | 0.86343 | 0.018106 | 20.9698 | 2.032 | 0.15267 | 0.15984 | 8.505 | 8.505 | 0.000723
2019-04-09 15:28 | 1554838064.7055 | 106 | 0.17762 | 0.02209 | 0.000639 | 100.3409 | 78.0712 | 82.9635 | 20.6952 | 0.005392 | 0.004858 | 160.1692 | 0.25594 | 8.4039 | 0.003881 | 0.005705 | 0.005557 | 3.9e-05 | 0.001108 | 0.28943 | 1.3906 | 2.5477 | 98.45 | 19.96 | 48.455 | 152.0198 | 114.7943 | 138.7852 | 0.011491 | 0.85704 | 0.018086 | 21.1024 | 2.2446 | 0.15226 | 0.16249 | 8.632 | 8.632 | 0.000711
2019-04-09 15:28 | 1554838065.7046 | 107 | 0.17738 | 0.023387 | 0.000691 | 100.2623 | 76.0022 | 82.5452 | 20.6946 | 0.005406 | 0.004873 | 159.8377 | 0.25644 | 8.4007 | 0.003893 | 0.005571 | 0.005465 | 3.9e-05 | 0.001106 | 0.26609 | 1.4085 | 2.5477 | 98.45 | 19.96 | 48.468 | 152.3685 | 115.0944 | 139.1501 | 0.011494 | 0.85628 | 0.018032 | 21.0585 | 2.4779 | 0.15255 | 0.16247 | 8.6029 | 8.6029 | 0.000715
2019-04-09 15:28 | 1554838066.7046 | 108 | 0.17726 | 0.023842 | 0.000743 | 100.2891 | 74.047 | 82.2544 | 20.6941 | 0.005496 | 0.005011 | 159.4591 | 0.2568 | 8.4015 | 0.003966 | 0.005471 | 0.005405 | 4e-05 | 0.001096 | 0.29479 | 1.3707 | 2.5475 | 98.45 | 19.96 | 48.476 | 154.5943 | 116.8698 | 142.567 | 0.011496 | 0.84924 | 0.017926 | 21.1083 | 2.7134 | 0.15277 | 0.16267 | 8.5603 | 8.5603 | 0.000719
2019-04-09 15:28 | 1554838067.7046 | 109 | 0.17695 | 0.024255 | 0.000791 | 100.362 | 72.0072 | 81.4129 | 20.6949 | 0.005597 | 0.005088 | 159.0833 | 0.25633 | 8.4024 | 0.004096 | 0.005343 | 0.005335 | 4.1e-05 | 0.001094 | 0.30469 | 1.3718 | 2.5476 | 98.45 | 19.96 | 48.482 | 157.1156 | 120.0396 | 144.4697 | 0.011497 | 0.84537 | 0.017793 | 21.048 | 2.9697 | 0.15249 | 0.16396 | 8.5618 | 8.5618 | 0.000718
2019-04-09 15:28 | 1554838068.7046 | 110 | 0.17644 | 0.024618 | 0.000821 | 100.2432 | 70.5143 | 80.6822 | 20.6943 | 0.005652 | 0.005191 | 158.7154 | 0.25578 | 8.4016 | 0.004108 | 0.005587 | 0.00541 | 4.2e-05 | 0.001099 | 0.2844 | 1.3419 | 2.5476 | 98.45 | 19.96 | 48.498 | 158.4846 | 120.3366 | 147.0306 | 0.011501 | 0.84022 | 0.017659 | 21.0175 | 3.1465 | 0.15216 | 0.16649 | 8.625 | 8.625 | 0.000711
2019-04-09 15:28 | 1554838069.7047 | 111 | 0.175 | 0.024932 | 0.000841 | 100.2864 | 69.38 | 79.2292 | 20.6944 | 0.005787 | 0.005332 | 158.3082 | 0.25657 | 8.3997 | 0.004193 | 0.005629 | 0.005509 | 4.1e-05 | 0.001099 | 0.27527 | 1.3579 | 2.5477 | 98.45 | 19.96 | 48.506 | 161.844 | 122.4238 | 150.5173 | 0.011503 | 0.83331 | 0.017428 | 20.9136 | 3.2955 | 0.15264 | 0.17443 | 8.9111 | 8.9111 | 0.00069
2019-04-09 15:28 | 1554838070.7037 | 112 | 0.17359 | 0.025234 | 0.000853 | 100.2989 | 68.9901 | 77.45 | 20.6952 | 0.005873 | 0.005498 | 157.8895 | 0.25646 | 8.3998 | 0.004236 | 0.005752 | 0.005726 | 4.1e-05 | 0.001103 | 0.26593 | 1.3807 | 2.5477 | 98.45 | 19.96 | 48.512 | 163.9961 | 123.4669 | 154.6469 | 0.011504 | 0.82526 | 0.017147 | 20.778 | 3.347 | 0.15257 | 0.18223 | 9.1537 | 9.1537 | 0.000672
2019-04-09 15:28 | 1554838071.7037 | 113 | 0.17176 | 0.025584 | 0.000856 | 100.3942 | 67.2782 | 76.2173 | 20.6941 | 0.006053 | 0.005595 | 157.4474 | 0.25544 | 8.4024 | 0.004389 | 0.00572 | 0.005551 | 4e-05 | 0.001117 | 0.29983 | 1.3703 | 2.5476 | 98.45 | 19.96 | 48.526 | 168.4784 | 127.2311 | 157.0597 | 0.011507 | 0.82064 | 0.016963 | 20.67 | 3.5802 | 0.15196 | 0.19233 | 9.549 | 9.549 | 0.000641
2019-04-09 15:28 | 1554838072.7037 | 114 | 0.16867 | 0.026338 | 0.000848 | 100.2953 | 67.3863 | 74.9291 | 20.6946 | 0.006179 | 0.005709 | 156.9931 | 0.25578 | 8.4022 | 0.004506 | 0.006105 | 0.005992 | 4.2e-05 | 0.001125 | 0.28567 | 1.3157 | 2.5476 | 98.45 | 19.96 | 48.535 | 171.622 | 130.0983 | 159.903 | 0.01151 | 0.81525 | 0.016763 | 20.5622 | 3.557 | 0.15216 | 0.20915 | 10.248 | 10.248 | 0.000598
2019-04-09 15:28 | 1554838073.7038 | 115 | 0.16674 | 0.027719 | 0.000828 | 100.3427 | 70.0655 | 72.9117 | 20.6946 | 0.006322 | 0.005837 | 156.5318 | 0.25599 | 8.4009 | 0.004519 | 0.005732 | 0.0057 | 3.9e-05 | 0.00113 | 0.24707 | 1.3246 | 2.5475 | 98.45 | 19.96 | 48.544 | 175.1975 | 130.4214 | 163.1042 | 0.011512 | 0.80927 | 0.016475 | 20.3584 | 3.2125 | 0.15229 | 0.21878 | 10.5273 | 10.5273 | 0.000583
2019-04-09 15:28 | 1554838074.7039 | 116 | 0.1637 | 0.02939 | 0.000801 | 100.3966 | 71.8706 | 71.2123 | 20.6933 | 0.006515 | 0.005934 | 156.0776 | 0.25577 | 8.403 | 0.004668 | 0.006074 | 0.005991 | 4e-05 | 0.001142 | 0.28711 | 1.3147 | 2.5476 | 98.45 | 19.96 | 48.552 | 180.0273 | 134.082 | 165.5116 | 0.011514 | 0.80482 | 0.016237 | 20.1751 | 2.9898 | 0.15216 | 0.23431 | 11.0977 | 11.0977 | 0.000553
2019-04-09 15:28 | 1554838075.7042 | 117 | 0.16267 | 0.031147 | 0.000772 | 100.4216 | 73.0246 | 69.8365 | 20.6944 | 0.006687 | 0.006101 | 155.5934 | 0.25499 | 8.4023 | 0.004729 | 0.006193 | 0.006067 | 4.1e-05 | 0.001139 | 0.28241 | 1.3437 | 2.5476 | 98.45 | 19.95 | 48.569 | 184.3391 | 135.5814 | 169.6953 | 0.011511 | 0.79722 | 0.016004 | 20.0744 | 2.8496 | 0.15169 | 0.23852 | 11.1312 | 11.1312 | 0.000549
2019-04-09 15:28 | 1554838076.7052 | 118 | 0.16112 | 0.032938 | 0.000745 | 100.387 | 74.9825 | 68.9108 | 20.6943 | 0.00691 | 0.006184 | 155.1101 | 0.25499 | 8.4023 | 0.004819 | 0.005726 | 0.005657 | 4.4e-05 | 0.00113 | 0.28385 | 1.3678 | 2.5476 | 98.45 | 19.95 | 48.576 | 189.9224 | 137.8168 | 171.7571 | 0.011512 | 0.79353 | 0.01586 | 19.9872 | 2.6098 | 0.15169 | 0.24555 | 11.3506 | 11.3506 | 0.000539
2019-04-09 15:28 | 1554838077.7039 | 119 | 0.15943 | 0.034686 | 0.000717 | 100.4284 | 76.6499 | 68.8317 | 20.6937 | 0.007044 | 0.006334 | 154.6389 | 0.25562 | 8.4017 | 0.004852 | 0.00644 | 0.006321 | 4e-05 | 0.001123 | 0.28672 | 1.3534 | 2.5476 | 98.45 | 19.95 | 48.591 | 193.2814 | 138.6312 | 175.5195 | 0.011516 | 0.78687 | 0.015785 | 20.06 | 2.4168 | 0.15207 | 0.25346 | 11.6525 | 11.6525 | 0.000526
2019-04-09 15:28 | 1554838078.7044 | 120 | 0.15794 | 0.036213 | 0.000682 | 100.4853 | 77.026 | 68.4233 | 20.6951 | 0.00714 | 0.006481 | 154.1466 | 0.25633 | 8.4031 | 0.004944 | 0.006008 | 0.005967 | 4.2e-05 | 0.001121 | 0.29407 | 1.3541 | 2.5477 | 98.45 | 19.95 | 48.598 | 195.6852 | 140.8934 | 179.1746 | 0.011517 | 0.78051 | 0.015674 | 20.0817 | 2.3781 | 0.15249 | 0.26041 | 11.8819 | 11.8819 | 0.000517
2019-04-09 15:28 | 1554838079.7043 | 121 | 0.15635 | 0.03768 | 0.000639 | 100.4639 | 77.2332 | 68.1299 | 20.6951 | 0.007319 | 0.00658 | 153.6489 | 0.25601 | 8.4036 | 0.005031 | 0.007008 | 0.006839 | 4.3e-05 | 0.001121 | 0.28313 | 1.3445 | 2.5477 | 98.45 | 19.95 | 48.605 | 200.1598 | 143.0515 | 181.6617 | 0.011519 | 0.77624 | 0.015598 | 20.0937 | 2.3521 | 0.1523 | 0.268 | 12.1613 | 12.1613 | 0.000505
2019-04-09 15:28 | 1554838080.704 | 122 | 0.15429 | 0.039058 | 0.000596 | 100.4522 | 78.902 | 68.3266 | 20.6947 | 0.007539 | 0.006698 | 153.1822 | 0.25517 | 8.402 | 0.005179 | 0.006806 | 0.006692 | 4.7e-05 | 0.001122 | 0.28799 | 1.367 | 2.5477 | 98.45 | 19.95 | 48.616 | 205.6588 | 146.7125 | 184.6197 | 0.011522 | 0.77123 | 0.01557 | 20.188 | 2.1599 | 0.1518 | 0.27827 | 12.5942 | 12.5942 | 0.000486
2019-04-09 15:28 | 1554838081.7034 | 123 | 0.1542 | 0.04013 | 0.00056 | 100.4251 | 81.1765 | 68.2416 | 20.6949 | 0.00765 | 0.006785 | 152.7174 | 0.25602 | 8.4031 | 0.005144 | 0.006713 | 0.006608 | 4.1e-05 | 0.001121 | 0.27776 | 1.3805 | 2.5477 | 98.45 | 19.95 | 48.622 | 208.4161 | 145.8439 | 186.8055 | 0.011523 | 0.76756 | 0.015523 | 20.2235 | 1.9033 | 0.15231 | 0.27789 | 12.5408 | 12.5408 | 0.000489
2019-04-09 15:28 | 1554838082.7032 | 124 | 0.15413 | 0.040797 | 0.000524 | 100.4698 | 82.7259 | 68.535 | 20.6938 | 0.007758 | 0.006813 | 152.2447 | 0.25604 | 8.4022 | 0.005233 | 0.006282 | 0.006182 | 4.4e-05 | 0.001124 | 0.29518 | 1.3298 | 2.5476 | 98.45 | 19.95 | 48.629 | 211.1137 | 148.0518 | 187.4968 | 0.011525 | 0.76641 | 0.015545 | 20.2822 | 1.7381 | 0.15232 | 0.27768 | 12.5504 | 12.5504 | 0.000489
2019-04-09 15:28 | 1554838083.7041 | 125 | 0.15459 | 0.04146 | 0.000491 | 100.5055 | 83.8754 | 69.5232 | 20.6947 | 0.007858 | 0.006879 | 151.7983 | 0.25525 | 8.4019 | 0.005246 | 0.007422 | 0.007297 | 4.6e-05 | 0.001133 | 0.26913 | 1.3481 | 2.5476 | 98.45 | 19.95 | 48.641 | 213.6183 | 148.3794 | 189.1546 | 0.011528 | 0.76366 | 0.015628 | 20.4646 | 1.6179 | 0.15185 | 0.27455 | 12.4807 | 12.4807 | 0.00049
2019-04-09 15:28 | 1554838084.7032 | 126 | 0.1556 | 0.042047 | 0.000453 | 100.5824 | 84.3379 | 70.3872 | 20.6947 | 0.007901 | 0.006855 | 151.3686 | 0.25573 | 8.4013 | 0.005325 | 0.006525 | 0.006429 | 4.3e-05 | 0.00114 | 0.28008 | 1.3432 | 2.5477 | 98.45 | 19.95 | 48.646 | 214.6702 | 150.3474 | 188.5336 | 0.011529 | 0.76469 | 0.015735 | 20.5775 | 1.5755 | 0.15214 | 0.26838 | 12.2927 | 12.2927 | 0.000499
2019-04-09 15:28 | 1554838085.7033 | 127 | 0.15654 | 0.042493 | 0.000427 | 100.5088 | 83.9759 | 71.4014 | 20.6944 | 0.007977 | 0.006983 | 150.9055 | 0.25551 | 8.4023 | 0.00527 | 0.007351 | 0.007305 | 4.2e-05 | 0.001133 | 0.27743 | 1.3736 | 2.5477 | 98.45 | 19.95 | 48.654 | 216.5594 | 148.9776 | 191.7513 | 0.011531 | 0.7594 | 0.015793 | 20.7974 | 1.6075 | 0.15201 | 0.26274 | 12.0862 | 12.0862 | 0.000507
2019-04-09 15:28 | 1554838086.7033 | 128 | 0.15872 | 0.04276 | 0.000401 | 100.5264 | 84.2275 | 72.1695 | 20.6944 | 0.008041 | 0.006985 | 150.484 | 0.25529 | 8.4018 | 0.005445 | 0.006863 | 0.006713 | 4.1e-05 | 0.001138 | 0.28064 | 1.3445 | 2.5476 | 98.45 | 19.95 | 48.658 | 218.1581 | 153.3218 | 191.8008 | 0.011532 | 0.75932 | 0.015877 | 20.91 | 1.5823 | 0.15187 | 0.25021 | 11.5857 | 11.5857 | 0.000528
2019-04-09 15:28 | 1554838087.7047 | 129 | 0.16126 | 0.042798 | 0.000378 | 100.5282 | 84.1259 | 72.6971 | 20.6938 | 0.008093 | 0.006984 | 150.071 | 0.25515 | 8.3986 | 0.005402 | 0.006696 | 0.006595 | 4.4e-05 | 0.001139 | 0.28152 | 1.3717 | 2.5476 | 98.45 | 19.95 | 48.666 | 219.4527 | 152.2606 | 191.7775 | 0.011534 | 0.75936 | 0.015936 | 20.9858 | 1.5932 | 0.15179 | 0.23574 | 10.9723 | 10.9723 | 0.000558
2019-04-09 15:28 | 1554838088.7037 | 130 | 0.16351 | 0.042598 | 0.000356 | 100.6437 | 84.2464 | 73.4271 | 20.6947 | 0.008122 | 0.007002 | 149.6546 | 0.25464 | 8.4029 | 0.005467 | 0.00727 | 0.007214 | 4.2e-05 | 0.001143 | 0.32565 | 1.3512 | 2.5477 | 98.45 | 19.95 | 48.669 | 220.1857 | 153.874 | 192.2192 | 0.011534 | 0.75863 | 0.016008 | 21.1009 | 1.5907 | 0.15148 | 0.22313 | 10.4458 | 10.4458 | 0.000584
2019-04-09 15:28 | 1554838089.7034 | 131 | 0.16453 | 0.042158 | 0.000338 | 100.6536 | 84.9873 | 73.9191 | 20.6947 | 0.008227 | 0.007059 | 149.2474 | 0.2549 | 8.4022 | 0.005509 | 0.006689 | 0.006573 | 4.1e-05 | 0.001141 | 0.28639 | 1.3381 | 2.5475 | 98.45 | 19.95 | 48.672 | 222.786 | 154.9144 | 193.6552 | 0.011535 | 0.7563 | 0.016037 | 21.2041 | 1.5133 | 0.15164 | 0.21771 | 10.2165 | 10.2165 | 0.000598
2019-04-09 15:28 | 1554838090.7047 | 132 | 0.16559 | 0.041289 | 0.000321 | 100.6857 | 84.9368 | 74.3977 | 20.695 | 0.008247 | 0.007088 | 148.872 | 0.25521 | 8.4032 | 0.005513 | 0.006779 | 0.006683 | 4.4e-05 | 0.001131 | 0.27433 | 1.3094 | 2.5476 | 98.45 | 19.95 | 48.679 | 223.2838 | 155.0108 | 194.3827 | 0.011537 | 0.75513 | 0.016076 | 21.2892 | 1.5214 | 0.15182 | 0.21249 | 10.0017 | 10.0017 | 0.000612
2019-04-09 15:28 | 1554838091.7035 | 133 | 0.1663 | 0.039739 | 0.000306 | 100.6967 | 84.9099 | 74.6298 | 20.6951 | 0.008228 | 0.007115 | 148.4709 | 0.25484 | 8.4011 | 0.005501 | 0.00654 | 0.006518 | 4.3e-05 | 0.001132 | 0.2808 | 1.3946 | 2.5475 | 98.45 | 19.95 | 48.682 | 222.8122 | 154.713 | 195.0553 | 0.011537 | 0.75404 | 0.01609 | 21.3378 | 1.5252 | 0.1516 | 0.20999 | 9.8955 | 9.8955 | 0.000617
2019-04-09 15:28 | 1554838092.7044 | 134 | 0.16681 | 0.038576 | 0.000291 | 100.6518 | 84.2484 | 74.7143 | 20.6948 | 0.008261 | 0.007118 | 148.0894 | 0.25573 | 8.4015 | 0.00556 | 0.007097 | 0.006963 | 4.1e-05 | 0.001137 | 0.25984 | 1.3835 | 2.5476 | 98.45 | 19.95 | 48.685 | 223.635 | 156.201 | 195.1161 | 0.011538 | 0.75394 | 0.016098 | 21.3512 | 1.5912 | 0.15213 | 0.20822 | 9.8188 | 9.8188 | 0.000624
2019-04-09 15:28 | 1554838093.7044 | 135 | 0.16783 | 0.037638 | 0.000277 | 100.7396 | 84.988 | 74.8228 | 20.6945 | 0.008314 | 0.007177 | 147.7275 | 0.25542 | 8.4008 | 0.005539 | 0.006405 | 0.006252 | 4.5e-05 | 0.001143 | 0.26936 | 1.3292 | 2.5477 | 98.45 | 19.95 | 48.689 | 224.9518 | 155.6782 | 196.6111 | 0.011539 | 0.75154 | 0.016084 | 21.4008 | 1.5208 | 0.15195 | 0.20328 | 9.5828 | 9.5828 | 0.000639
2019-04-09 15:28 | 1554838094.7048 | 136 | 0.16867 | 0.036894 | 0.000263 | 100.7596 | 85.2974 | 74.7777 | 20.6939 | 0.008404 | 0.00714 | 147.3409 | 0.25528 | 8.4009 | 0.005599 | 0.006889 | 0.00678 | 4.4e-05 | 0.001148 | 0.28639 | 1.3611 | 2.5476 | 98.45 | 19.95 | 48.69 | 227.1831 | 157.1701 | 195.6832 | 0.011539 | 0.75303 | 0.016095 | 21.3732 | 1.4901 | 0.15187 | 0.19922 | 9.4023 | 9.4023 | 0.000651
2019-04-09 15:28 | 1554838095.7039 | 137 | 0.16901 | 0.036295 | 0.000249 | 100.7977 | 85.3367 | 74.5923 | 20.6953 | 0.008454 | 0.007224 | 146.9734 | 0.25497 | 8.4017 | 0.005572 | 0.007491 | 0.007371 | 4.6e-05 | 0.001151 | 0.26609 | 1.4248 | 2.5476 | 98.45 | 19.94 | 48.695 | 228.4355 | 156.4962 | 197.7665 | 0.011533 | 0.7497 | 0.016039 | 21.3941 | 1.4894 | 0.15168 | 0.19787 | 9.3084 | 9.3084 | 0.000657
2019-04-09 15:28 | 1554838096.7036 | 138 | 0.16955 | 0.035702 | 0.000238 | 100.7941 | 85.3618 | 74.7606 | 20.694 | 0.008548 | 0.007235 | 146.6149 | 0.25482 | 8.402 | 0.00561 | 0.007389 | 0.007285 | 4.7e-05 | 0.001151 | 0.27024 | 1.4096 | 2.5477 | 98.45 | 19.94 | 48.696 | 230.7689 | 157.4264 | 198.0402 | 0.011533 | 0.74926 | 0.016053 | 21.4244 | 1.4864 | 0.15159 | 0.1954 | 9.2022 | 9.2022 | 0.000664
2019-04-09 15:28 | 1554838097.7752 | 139 | 0.1699 | 0.035138 | 0.000227 | 100.7736 | 86.4834 | 74.9138 | 20.6943 | 0.008514 | 0.007261 | 146.2583 | 0.25504 | 8.4021 | 0.005593 | 0.00727 | 0.007159 | 4.2e-05 | 0.001155 | 0.28766 | 1.3201 | 2.5477 | 98.45 | 19.94 | 48.699 | 229.9239 | 157.0066 | 198.7008 | 0.011534 | 0.74821 | 0.016058 | 21.4614 | 1.3678 | 0.15172 | 0.19396 | 9.1389 | 9.1389 | 0.000669
2019-04-09 15:28 | 1554838098.7178 | 140 | 0.17061 | 0.034598 | 0.000218 | 100.8424 | 87.3407 | 75.0354 | 20.6941 | 0.008531 | 0.007275 | 145.9024 | 0.25487 | 8.4003 | 0.005613 | 0.006855 | 0.006759 | 4.6e-05 | 0.001149 | 0.27455 | 1.3431 | 2.5476 | 98.45 | 19.94 | 48.704 | 230.3427 | 157.5065 | 199.0567 | 0.011535 | 0.74765 | 0.016065 | 21.4869 | 1.2857 | 0.15162 | 0.19042 | 8.9796 | 8.9796 | 0.00068
2019-04-09 15:28 | 1554838099.7178 | 141 | 0.17089 | 0.034088 | 0.00021 | 100.8787 | 88.5053 | 75.6129 | 20.694 | 0.008548 | 0.007281 | 145.5735 | 0.25559 | 8.4013 | 0.005611 | 0.006939 | 0.006951 | 4.7e-05 | 0.001149 | 0.28456 | 1.3271 | 2.5476 | 98.45 | 19.94 | 48.706 | 230.7503 | 157.4653 | 199.2123 | 0.011536 | 0.7474 | 0.016124 | 21.573 | 1.1705 | 0.15205 | 0.18932 | 8.9619 | 8.9619 | 0.000684
2019-04-09 15:28 | 1554838100.7173 | 142 | 0.17151 | 0.0336 | 0.000202 | 100.8927 | 89.0501 | 76.0445 | 20.6937 | 0.008561 | 0.007225 | 145.2335 | 0.25493 | 8.4025 | 0.005634 | 0.006793 | 0.006578 | 4.3e-05 | 0.001142 | 0.28047 | 1.3296 | 2.5476 | 98.45 | 19.94 | 48.709 | 231.0756 | 158.0216 | 197.7973 | 0.011537 | 0.74965 | 0.016194 | 21.602 | 1.1168 | 0.15166 | 0.18626 | 8.8584 | 8.8584 | 0.00069
2019-04-09 15:28 | 1554838101.7169 | 143 | 0.17212 | 0.033107 | 0.000196 | 100.9073 | 88.3071 | 76.9196 | 20.6934 | 0.008663 | 0.007293 | 144.8857 | 0.25557 | 8.401 | 0.00558 | 0.007756 | 0.007603 | 4.3e-05 | 0.001144 | 0.26936 | 1.3153 | 2.5475 | 98.45 | 19.94 | 48.71 | 233.5956 | 156.6843 | 199.5077 | 0.011537 | 0.74694 | 0.016257 | 21.7654 | 1.193 | 0.15204 | 0.18326 | 8.7527 | 8.7527 | 0.0007
2019-04-09 15:28 | 1554838102.718 | 144 | 0.17261 | 0.03264 | 0.00019 | 100.9463 | 88.4158 | 77.3716 | 20.6936 | 0.008628 | 0.007278 | 144.5622 | 0.25525 | 8.4021 | 0.005544 | 0.006939 | 0.00687 | 4.9e-05 | 0.001143 | 0.30735 | 1.327 | 2.5476 | 98.45 | 19.94 | 48.712 | 232.7508 | 155.7821 | 199.1269 | 0.011537 | 0.74754 | 0.016312 | 21.8205 | 1.1855 | 0.15185 | 0.18095 | 8.6732 | 8.6732 | 0.000706
2019-04-09 15:28 | 1554838103.7179 | 145 | 0.17292 | 0.03219 | 0.000184 | 101.0085 | 88.5242 | 78.0195 | 20.6936 | 0.008675 | 0.007297 | 144.2405 | 0.25561 | 8.4011 | 0.005631 | 0.006572 | 0.006492 | 4.3e-05 | 0.001143 | 0.26609 | 1.3374 | 2.5477 | 98.45 | 19.94 | 48.716 | 233.8986 | 157.9625 | 199.5931 | 0.011538 | 0.7468 | 0.016372 | 21.9224 | 1.18 | 0.15206 | 0.17958 | 8.6408 | 8.6408 | 0.000709
2019-04-09 15:28 | 1554838104.7174 | 146 | 0.1734 | 0.031759 | 0.000179 | 100.9942 | 88.0307 | 78.5565 | 20.6927 | 0.008586 | 0.007251 | 143.8999 | 0.25432 | 8.4053 | 0.005635 | 0.007426 | 0.007256 | 4.3e-05 | 0.001139 | 0.28224 | 1.3364 | 2.5477 | 98.45 | 19.93 | 48.718 | 231.6864 | 158.0642 | 198.4527 | 0.011531 | 0.74861 | 0.016448 | 21.9712 | 1.2288 | 0.15129 | 0.1773 | 8.5729 | 8.5729 | 0.000711
2019-04-09 15:28 | 1554838105.7182 | 147 | 0.17396 | 0.031343 | 0.000175 | 101.027 | 87.1569 | 78.8687 | 20.6927 | 0.008636 | 0.007254 | 143.583 | 0.25507 | 8.4008 | 0.005664 | 0.007182 | 0.007136 | 4.3e-05 | 0.001136 | 0.27721 | 1.369 | 2.5476 | 98.45 | 19.93 | 48.72 | 232.9346 | 158.7765 | 198.5149 | 0.011532 | 0.74851 | 0.016479 | 22.0163 | 1.3209 | 0.15174 | 0.1745 | 8.4562 | 8.4562 | 0.000723
2019-04-09 15:28 | 1554838106.717 | 148 | 0.17456 | 0.03093 | 0.000172 | 101.1127 | 88.063 | 79.3382 | 20.6931 | 0.008679 | 0.007321 | 143.279 | 0.25493 | 8.4016 | 0.005709 | 0.006737 | 0.006586 | 4.4e-05 | 0.001135 | 0.27793 | 1.3581 | 2.5476 | 98.45 | 19.93 | 48.724 | 234.008 | 159.9022 | 200.2 | 0.011533 | 0.74584 | 0.016499 | 22.1211 | 1.236 | 0.15166 | 0.17144 | 8.3206 | 8.3206 | 0.000735
2019-04-09 15:28 | 1554838107.7171 | 149 | 0.17527 | 0.030524 | 0.000169 | 101.0609 | 88.7145 | 79.6339 | 20.6927 | 0.00866 | 0.007274 | 142.9662 | 0.25424 | 8.4018 | 0.005689 | 0.006837 | 0.006758 | 4.5e-05 | 0.00113 | 0.25912 | 1.3505 | 2.5475 | 98.45 | 19.93 | 48.727 | 233.5363 | 159.3949 | 199.0277 | 0.011534 | 0.7477 | 0.01655 | 22.1348 | 1.1655 | 0.15125 | 0.16778 | 8.1709 | 8.1709 | 0.000746
2019-04-09 15:28 | 1554838108.7181 | 150 | 0.17599 | 0.030102 | 0.000166 | 101.1124 | 88.5359 | 79.9313 | 20.6928 | 0.008628 | 0.007249 | 142.6666 | 0.25491 | 8.4025 | 0.005619 | 0.007254 | 0.007115 | 4.3e-05 | 0.00113 | 0.29087 | 1.3753 | 2.5476 | 98.45 | 19.93 | 48.732 | 232.7442 | 157.6513 | 198.3885 | 0.011535 | 0.74871 | 0.016592 | 22.1611 | 1.1881 | 0.15165 | 0.16401 | 8.0111 | 8.0111 | 0.000763
2019-04-09 15:28 | 1554838109.7185 | 151 | 0.17652 | 0.029679 | 0.000164 | 101.0931 | 87.8548 | 80.3592 | 20.6931 | 0.008685 | 0.007239 | 142.3749 | 0.25422 | 8.4007 | 0.005699 | 0.007326 | 0.007166 | 4.4e-05 | 0.001129 | 0.28064 | 1.3632 | 2.5477 | 98.45 | 19.94 | 48.734 | 234.1575 | 159.6508 | 198.1406 | 0.011542 | 0.7491 | 0.016641 | 22.2145 | 1.2554 | 0.15124 | 0.1614 | 7.9087 | 7.9087 | 0.000771
2019-04-09 15:28 | 1554838110.7181 | 152 | 0.17697 | 0.02924 | 0.000163 | 101.1381 | 87.6026 | 80.8789 | 20.6925 | 0.008666 | 0.007203 | 142.0687 | 0.25429 | 8.401 | 0.005677 | 0.00617 | 0.006083 | 4.2e-05 | 0.00113 | 0.28423 | 1.4142 | 2.5477 | 98.45 | 19.94 | 48.736 | 233.6678 | 159.1124 | 197.2533 | 0.011543 | 0.75052 | 0.01671 | 22.2652 | 1.2851 | 0.15128 | 0.15923 | 7.8365 | 7.8365 | 0.000778
2019-04-09 15:29 | 1554838111.7181 | 153 | 0.17754 | 0.028804 | 0.000162 | 101.166 | 88.1143 | 81.341 | 20.6933 | 0.00862 | 0.00716 | 141.7862 | 0.25449 | 8.3999 | 0.005616 | 0.006967 | 0.006879 | 4.6e-05 | 0.001135 | 0.28943 | 1.3642 | 2.5476 | 98.45 | 19.94 | 48.742 | 232.5325 | 157.5914 | 196.1653 | 0.011544 | 0.75226 | 0.016778 | 22.3029 | 1.2355 | 0.15139 | 0.1564 | 7.7306 | 7.7306 | 0.000789
2019-04-09 15:29 | 1554838112.7162 | 154 | 0.17831 | 0.028362 | 0.000162 | 101.1666 | 89.9569 | 81.8039 | 20.6923 | 0.008586 | 0.007209 | 141.5163 | 0.25488 | 8.4008 | 0.005644 | 0.00672 | 0.006608 | 4.4e-05 | 0.001137 | 0.26936 | 1.3171 | 2.5477 | 98.45 | 19.94 | 48.743 | 231.6892 | 158.2748 | 197.3976 | 0.011545 | 0.75029 | 0.016803 | 22.3956 | 1.0504 | 0.15163 | 0.1524 | 7.5473 | 7.5473 | 0.00081
2019-04-09 15:29 | 1554838113.7173 | 155 | 0.17889 | 0.027918 | 0.000162 | 101.2247 | 90.3355 | 81.7748 | 20.6919 | 0.008594 | 0.007172 | 141.2294 | 0.25362 | 8.4015 | 0.005648 | 0.00705 | 0.006916 | 4.7e-05 | 0.001135 | 0.29358 | 1.3291 | 2.5475 | 98.45 | 19.94 | 48.748 | 231.8955 | 158.3763 | 196.4747 | 0.011546 | 0.75176 | 0.016817 | 22.3697 | 1.018 | 0.15088 | 0.14952 | 7.4123 | 7.4123 | 0.00082
2019-04-09 15:29 | 1554838114.7163 | 156 | 0.17945 | 0.02746 | 0.000161 | 101.1672 | 89.8148 | 81.8203 | 20.6931 | 0.00858 | 0.007101 | 140.9709 | 0.25391 | 8.4036 | 0.005673 | 0.006979 | 0.0069 | 4.4e-05 | 0.001137 | 0.26217 | 1.3293 | 2.5475 | 98.45 | 19.94 | 48.75 | 231.5461 | 159.0153 | 194.7102 | 0.011546 | 0.7546 | 0.016853 | 22.3338 | 1.0646 | 0.15105 | 0.14671 | 7.291 | 7.291 | 0.000835
2019-04-09 15:29 | 1554838115.7167 | 157 | 0.18005 | 0.02699 | 0.000162 | 101.2051 | 88.2441 | 82.3921 | 20.6924 | 0.008494 | 0.007144 | 140.7286 | 0.25381 | 8.4036 | 0.005644 | 0.006599 | 0.006506 | 4.7e-05 | 0.001143 | 0.26648 | 1.3622 | 2.5478 | 98.45 | 19.93 | 48.752 | 229.4272 | 158.2685 | 195.775 | 0.01154 | 0.75288 | 0.016893 | 22.4372 | 1.2258 | 0.15099 | 0.14369 | 7.16 | 7.16 | 0.00085
2019-04-09 15:29 | 1554838116.7164 | 158 | 0.18054 | 0.026528 | 0.000163 | 101.18 | 87.0898 | 82.8233 | 20.6924 | 0.008551 | 0.007103 | 140.4631 | 0.25399 | 8.401 | 0.00559 | 0.006569 | 0.00655 | 4.9e-05 | 0.001144 | 0.25912 | 1.3477 | 2.5476 | 98.45 | 19.93 | 48.757 | 230.834 | 156.9425 | 194.7485 | 0.011541 | 0.75453 | 0.016955 | 22.4712 | 1.3413 | 0.1511 | 0.14135 | 7.0707 | 7.0707 | 0.000861
2019-04-09 15:29 | 1554838117.7174 | 159 | 0.18094 | 0.026065 | 0.000165 | 101.1774 | 87.2808 | 83.0916 | 20.6928 | 0.008483 | 0.007088 | 140.1959 | 0.25388 | 8.4022 | 0.005647 | 0.007183 | 0.006941 | 4.4e-05 | 0.001147 | 0.27151 | 1.3278 | 2.5476 | 98.45 | 19.93 | 48.759 | 229.1347 | 158.3578 | 194.3747 | 0.011541 | 0.75514 | 0.01699 | 22.4986 | 1.3215 | 0.15103 | 0.13954 | 6.9954 | 6.9954 | 0.00087
2019-04-09 15:29 | 1554838118.7165 | 160 | 0.18108 | 0.025613 | 0.000168 | 101.2211 | 88.2991 | 83.5338 | 20.6929 | 0.008526 | 0.007078 | 139.9542 | 0.25496 | 8.4023 | 0.005645 | 0.007386 | 0.00722 | 4.8e-05 | 0.001145 | 0.28296 | 1.336 | 2.5477 | 98.45 | 19.93 | 48.762 | 230.2114 | 158.3021 | 194.1119 | 0.011542 | 0.75556 | 0.017039 | 22.552 | 1.2216 | 0.15168 | 0.1392 | 6.9993 | 6.9993 | 0.000873
2019-04-09 15:29 | 1554838119.7165 | 161 | 0.1815 | 0.025181 | 0.00017 | 101.1982 | 88.7662 | 83.6756 | 20.6931 | 0.008527 | 0.007039 | 139.7072 | 0.25411 | 8.3992 | 0.005551 | 0.007319 | 0.007271 | 5e-05 | 0.001151 | 0.25857 | 1.3339 | 2.5475 | 98.45 | 19.93 | 48.764 | 230.242 | 155.9545 | 193.1549 | 0.011542 | 0.75711 | 0.017071 | 22.548 | 1.1724 | 0.15117 | 0.13722 | 6.9138 | 6.9138 | 0.000881
2019-04-09 15:29 | 1554838120.7164 | 162 | 0.18228 | 0.02477 | 0.000172 | 101.1934 | 89.4737 | 83.7258 | 20.6931 | 0.008514 | 0.00707 | 139.4605 | 0.25449 | 8.401 | 0.005549 | 0.007586 | 0.007481 | 4.8e-05 | 0.001147 | 0.27616 | 1.3258 | 2.5476 | 98.45 | 19.93 | 48.766 | 229.9124 | 155.9061 | 193.9318 | 0.011543 | 0.75585 | 0.017062 | 22.5736 | 1.101 | 0.15139 | 0.13309 | 6.7049 | 6.7049 | 0.00091
2019-04-09 15:29 | 1554838121.7175 | 163 | 0.18302 | 0.024383 | 0.000174 | 101.1985 | 90.2901 | 83.3738 | 20.6922 | 0.008462 | 0.007107 | 139.2412 | 0.25456 | 8.3997 | 0.005564 | 0.006569 | 0.00644 | 4.9e-05 | 0.001149 | 0.274 | 1.3854 | 2.5479 | 98.45 | 19.93 | 48.77 | 228.6336 | 156.2992 | 194.8513 | 0.011544 | 0.75437 | 0.01701 | 22.5482 | 1.0202 | 0.15144 | 0.1292 | 6.4908 | 6.4908 | 0.00094
2019-04-09 15:29 | 1554838122.7175 | 164 | 0.18362 | 0.023978 | 0.000177 | 101.2235 | 90.7185 | 83.2611 | 20.6929 | 0.008434 | 0.007024 | 139.0065 | 0.255 | 8.3989 | 0.005503 | 0.007107 | 0.007025 | 5.4e-05 | 0.001147 | 0.26665 | 1.3906 | 2.5476 | 98.45 | 19.93 | 48.773 | 227.9293 | 154.7626 | 192.7783 | 0.011544 | 0.75772 | 0.017036 | 22.483 | 0.98005 | 0.1517 | 0.12613 | 6.3482 | 6.3482 | 0.000963
2019-04-09 15:29 | 1554838123.718 | 165 | 0.18393 | 0.023569 | 0.000178 | 101.1622 | 90.9794 | 83.7127 | 20.6932 | 0.008388 | 0.006957 | 138.7606 | 0.2545 | 8.4024 | 0.005546 | 0.007148 | 0.007086 | 5.2e-05 | 0.001149 | 0.2672 | 1.3142 | 2.5474 | 98.45 | 19.93 | 48.775 | 226.7935 | 155.8478 | 191.103 | 0.011545 | 0.76046 | 0.017113 | 22.5034 | 0.94894 | 0.1514 | 0.12474 | 6.3074 | 6.3074 | 0.000967
2019-04-09 15:29 | 1554838124.717 | 166 | 0.18435 | 0.023149 | 0.000179 | 101.1419 | 90.2698 | 84.1197 | 20.6933 | 0.008437 | 0.006956 | 138.5421 | 0.25468 | 8.4034 | 0.005473 | 0.007346 | 0.007239 | 5.4e-05 | 0.001152 | 0.27345 | 1.3683 | 2.5475 | 98.45 | 19.93 | 48.781 | 228.0103 | 154.0243 | 191.0717 | 0.011546 | 0.76051 | 0.017155 | 22.5572 | 1.0172 | 0.15151 | 0.12273 | 6.2223 | 6.2223 | 0.000981
2019-04-09 15:29 | 1554838125.7171 | 167 | 0.18478 | 0.022729 | 0.000181 | 101.091 | 91.0272 | 84.4356 | 20.6937 | 0.008482 | 0.006918 | 138.3318 | 0.25419 | 8.4002 | 0.00536 | 0.007017 | 0.006939 | 5.4e-05 | 0.001155 | 0.27328 | 1.3495 | 2.5475 | 98.45 | 19.93 | 48.784 | 229.1133 | 151.2154 | 190.1183 | 0.011547 | 0.76207 | 0.017205 | 22.5763 | 0.93795 | 0.15122 | 0.12071 | 6.1387 | 6.1387 | 0.000993
2019-04-09 15:29 | 1554838126.7176 | 168 | 0.18497 | 0.022307 | 0.000182 | 101.1205 | 91.0937 | 84.6751 | 20.6928 | 0.008358 | 0.006889 | 138.1052 | 0.25454 | 8.3995 | 0.00536 | 0.006759 | 0.006648 | 5.4e-05 | 0.00116 | 0.30663 | 1.3336 | 2.5476 | 98.45 | 19.93 | 48.79 | 226.0474 | 151.2128 | 189.3882 | 0.011549 | 0.76328 | 0.017243 | 22.5905 | 0.93403 | 0.15143 | 0.12002 | 6.1175 | 6.1175 | 0.000998
2019-04-09 15:29 | 1554838127.7167 | 169 | 0.18526 | 0.021892 | 0.000183 | 100.9963 | 90.987 | 85.0546 | 20.6929 | 0.008331 | 0.006897 | 137.9031 | 0.2555 | 8.404 | 0.00535 | 0.007738 | 0.007553 | 5.5e-05 | 0.001161 | 0.29374 | 1.3358 | 2.5477 | 98.45 | 19.93 | 48.792 | 225.3741 | 150.9589 | 189.593 | 0.011549 | 0.76294 | 0.017278 | 22.6461 | 0.93352 | 0.152 | 0.1188 | 6.0679 | 6.0679 | 0.001009
2019-04-09 15:29 | 1554838128.7178 | 170 | 0.18584 | 0.021499 | 0.000185 | 100.9695 | 91.3549 | 85.3635 | 20.6936 | 0.008285 | 0.006866 | 137.7067 | 0.25473 | 8.3996 | 0.005283 | 0.007034 | 0.006971 | 5.4e-05 | 0.001165 | 0.27848 | 1.3406 | 2.5477 | 98.45 | 19.93 | 48.795 | 224.2312 | 149.307 | 188.8297 | 0.01155 | 0.7642 | 0.017323 | 22.6685 | 0.89505 | 0.15154 | 0.11586 | 5.9347 | 5.9347 | 0.001029
2019-04-09 15:29 | 1554838129.7168 | 171 | 0.18677 | 0.021128 | 0.000187 | 100.898 | 91.7458 | 85.9113 | 20.6928 | 0.008255 | 0.006818 | 137.4864 | 0.25436 | 8.3997 | 0.005299 | 0.007043 | 0.006999 | 5.4e-05 | 0.001164 | 0.26017 | 1.3563 | 2.5477 | 98.45 | 19.93 | 48.802 | 223.4896 | 149.7043 | 187.6274 | 0.011551 | 0.76619 | 0.017401 | 22.7115 | 0.85052 | 0.15132 | 0.1108 | 5.7039 | 5.7039 | 0.001069
2019-04-09 15:29 | 1554838130.7171 | 172 | 0.18718 | 0.020754 | 0.000191 | 100.8188 | 92.048 | 86.631 | 20.6932 | 0.008222 | 0.006796 | 137.2862 | 0.25424 | 8.4014 | 0.005229 | 0.006642 | 0.006599 | 5.8e-05 | 0.001164 | 0.28351 | 1.336 | 2.5476 | 98.45 | 19.93 | 48.804 | 222.6755 | 147.9625 | 187.0747 | 0.011552 | 0.76711 | 0.017485 | 22.7927 | 0.81409 | 0.15125 | 0.10878 | 5.6276 | 5.6276 | 0.001083
2019-04-09 15:29 | 1554838131.7162 | 173 | 0.18776 | 0.020385 | 0.000194 | 100.8406 | 91.9909 | 87.3692 | 20.6928 | 0.008189 | 0.006816 | 137.0722 | 0.25457 | 8.403 | 0.005251 | 0.006372 | 0.006267 | 5.8e-05 | 0.00116 | 0.29407 | 1.3605 | 2.5477 | 98.45 | 19.93 | 48.81 | 221.8436 | 148.5039 | 187.5656 | 0.011553 | 0.7663 | 0.01755 | 22.9018 | 0.82157 | 0.15144 | 0.10579 | 5.4944 | 5.4944 | 0.001111
2019-04-09 15:29 | 1554838132.7159 | 174 | 0.18805 | 0.020018 | 0.000199 | 100.6672 | 91.4375 | 88.0377 | 20.6929 | 0.008178 | 0.006791 | 136.8462 | 0.254 | 8.4014 | 0.005208 | 0.006375 | 0.006273 | 5.7e-05 | 0.001154 | 0.28351 | 1.3853 | 2.5477 | 98.45 | 19.93 | 48.811 | 221.5605 | 147.4446 | 186.9358 | 0.011553 | 0.76735 | 0.017629 | 22.9736 | 0.86015 | 0.1511 | 0.10448 | 5.4514 | 5.4514 | 0.001117
2019-04-09 15:29 | 1554838133.7169 | 175 | 0.18836 | 0.019659 | 0.000203 | 100.597 | 91.0665 | 88.4182 | 20.693 | 0.008116 | 0.006735 | 136.6511 | 0.2541 | 8.4041 | 0.00515 | 0.006511 | 0.006438 | 5.5e-05 | 0.001153 | 0.26609 | 1.3775 | 2.5477 | 98.45 | 19.93 | 48.815 | 220.034 | 146.0011 | 185.5443 | 0.011554 | 0.76967 | 0.017694 | 22.9883 | 0.89027 | 0.15116 | 0.10307 | 5.3978 | 5.3978 | 0.001129
2019-04-09 15:29 | 1554838134.7166 | 176 | 0.18896 | 0.019305 | 0.000207 | 100.5383 | 91.0444 | 88.7759 | 20.6928 | 0.008095 | 0.006694 | 136.463 | 0.25447 | 8.4011 | 0.005073 | 0.005756 | 0.00561 | 5.6e-05 | 0.001147 | 0.28639 | 1.4001 | 2.5477 | 98.45 | 19.93 | 48.822 | 219.5146 | 144.0945 | 184.5245 | 0.011556 | 0.77139 | 0.017749 | 23.0091 | 0.88722 | 0.15139 | 0.099938 | 5.2517 | 5.2517 | 0.001162
2019-04-09 15:29 | 1554838135.7159 | 177 | 0.18947 | 0.018952 | 0.00021 | 100.415 | 91.3593 | 89.1357 | 20.6932 | 0.007991 | 0.006682 | 136.2222 | 0.25342 | 8.3988 | 0.005097 | 0.005791 | 0.005752 | 5.6e-05 | 0.00114 | 0.29767 | 1.3496 | 2.5477 | 98.45 | 19.93 | 48.825 | 216.9268 | 144.6883 | 184.2048 | 0.011557 | 0.77193 | 0.017791 | 23.0477 | 0.84536 | 0.15076 | 0.09731 | 5.1268 | 5.1268 | 0.001185
2019-04-09 15:29 | 1554838136.717 | 178 | 0.18998 | 0.018595 | 0.000213 | 100.3293 | 91.6694 | 89.738 | 20.693 | 0.008016 | 0.006652 | 136.0462 | 0.25378 | 8.4023 | 0.00504 | 0.006617 | 0.006508 | 5.7e-05 | 0.001137 | 0.28655 | 1.346 | 2.5477 | 98.45 | 19.93 | 48.829 | 217.5349 | 143.2696 | 183.4715 | 0.011558 | 0.77317 | 0.017866 | 23.1069 | 0.80741 | 0.15097 | 0.094715 | 5.0119 | 5.0119 | 0.001214
2019-04-09 15:29 | 1554838137.7166 | 179 | 0.19019 | 0.018223 | 0.000215 | 100.2634 | 91.751 | 90.1709 | 20.6936 | 0.008001 | 0.00654 | 135.8757 | 0.25422 | 8.4025 | 0.005061 | 0.006036 | 0.005942 | 5.8e-05 | 0.001128 | 0.28567 | 1.3411 | 2.5475 | 98.45 | 19.93 | 48.834 | 217.159 | 143.807 | 180.6739 | 0.011559 | 0.77794 | 0.017964 | 23.0915 | 0.79358 | 0.15124 | 0.093858 | 4.9941 | 4.9941 | 0.00122
2019-04-09 15:29 | 1554838138.7176 | 180 | 0.1904 | 0.017855 | 0.000217 | 100.0124 | 91.3452 | 90.3007 | 20.6937 | 0.007912 | 0.006557 | 135.6623 | 0.25488 | 8.4008 | 0.004979 | 0.005808 | 0.005761 | 5.7e-05 | 0.001131 | 0.28816 | 1.3693 | 2.5475 | 98.45 | 19.93 | 48.837 | 214.9393 | 141.7741 | 181.0822 | 0.01156 | 0.77724 | 0.017968 | 23.1185 | 0.81081 | 0.15163 | 0.093066 | 4.9535 | 4.9535 | 0.001234
2019-04-09 15:29 | 1554838139.7176 | 181 | 0.1908 | 0.017496 | 0.000218 | 99.8325 | 90.8596 | 90.3934 | 20.6923 | 0.007931 | 0.006487 | 135.4858 | 0.25471 | 8.4016 | 0.00497 | 0.006668 | 0.006493 | 5.7e-05 | 0.001134 | 0.30215 | 1.322 | 2.5474 | 98.45 | 19.93 | 48.843 | 215.431 | 141.5456 | 179.3348 | 0.011561 | 0.78024 | 0.018012 | 23.0858 | 0.84238 | 0.15153 | 0.091085 | 4.8606 | 4.8606 | 0.001256
2019-04-09 15:29 | 1554838140.7171 | 182 | 0.19127 | 0.017155 | 0.000219 | 99.5895 | 90.7251 | 90.3687 | 20.6932 | 0.007821 | 0.006556 | 135.324 | 0.25438 | 8.4013 | 0.004936 | 0.006131 | 0.006038 | 5.8e-05 | 0.001139 | 0.26952 | 1.3779 | 2.5476 | 98.45 | 19.93 | 48.845 | 212.6714 | 140.6967 | 181.0721 | 0.011562 | 0.77725 | 0.017975 | 23.1269 | 0.83384 | 0.15133 | 0.088714 | 4.7253 | 4.7253 | 0.001291
2019-04-09 15:29 | 1554838141.7161 | 183 | 0.19189 | 0.016816 | 0.00022 | 99.435 | 90.669 | 90.4603 | 20.6934 | 0.00784 | 0.006423 | 135.1561 | 0.25501 | 8.4012 | 0.004868 | 0.005857 | 0.005764 | 5.7e-05 | 0.001141 | 0.27489 | 1.3564 | 2.5476 | 98.45 | 19.93 | 48.847 | 213.1472 | 139.031 | 177.7282 | 0.011562 | 0.78302 | 0.018051 | 23.0533 | 0.82548 | 0.15171 | 0.085447 | 4.5715 | 4.5715 | 0.001337
2019-04-09 15:29 | 1554838142.7176 | 184 | 0.1926 | 0.016482 | 0.000223 | 99.1994 | 90.6903 | 90.6772 | 20.6942 | 0.007712 | 0.006416 | 135.0107 | 0.25462 | 8.4027 | 0.004833 | 0.005984 | 0.005881 | 5.8e-05 | 0.001144 | 0.28279 | 1.3645 | 2.5476 | 98.45 | 19.93 | 48.851 | 209.9699 | 138.1494 | 177.5479 | 0.011563 | 0.78333 | 0.018076 | 23.0763 | 0.80216 | 0.15147 | 0.081599 | 4.373 | 4.373 | 0.001396
2019-04-09 15:29 | 1554838143.7152 | 185 | 0.1932 | 0.015988 | 0.000227 | 99.0282 | 91.1061 | 91.0767 | 20.693 | 0.007723 | 0.00643 | 134.8829 | 0.25451 | 8.402 | 0.004843 | 0.006471 | 0.006403 | 5.7e-05 | 0.001146 | 0.27959 | 1.3902 | 2.5476 | 98.45 | 19.93 | 48.853 | 210.2504 | 138.3957 | 177.9086 | 0.011563 | 0.7827 | 0.018109 | 23.1363 | 0.7458 | 0.15141 | 0.078632 | 4.2223 | 4.2223 | 0.001445
2019-04-09 15:29 | 1554838144.7162 | 186 | 0.19398 | 0.015591 | 0.000233 | 98.7203 | 91.3566 | 91.5095 | 20.6935 | 0.007622 | 0.006312 | 134.7133 | 0.25353 | 8.4038 | 0.004866 | 0.006347 | 0.006256 | 5.6e-05 | 0.001147 | 0.29574 | 1.3499 | 2.5476 | 98.45 | 19.93 | 48.856 | 207.7151 | 138.9718 | 174.9626 | 0.011564 | 0.78785 | 0.018211 | 23.1154 | 0.69338 | 0.15082 | 0.074457 | 4.0217 | 4.0217 | 0.001511
2019-04-09 15:29 | 1554838145.7153 | 187 | 0.19434 | 0.015063 | 0.00024 | 98.4827 | 91.3521 | 91.8348 | 20.6928 | 0.007611 | 0.006259 | 134.5707 | 0.25366 | 8.4001 | 0.004718 | 0.006276 | 0.006229 | 5.9e-05 | 0.001147 | 0.27793 | 1.4069 | 2.5476 | 98.45 | 19.93 | 48.857 | 207.4543 | 135.3078 | 173.6438 | 0.011564 | 0.79017 | 0.018271 | 23.1223 | 0.67226 | 0.1509 | 0.072897 | 3.9503 | 3.9503 | 0.001539
2019-04-09 15:29 | 1554838146.7164 | 188 | 0.19469 | 0.014666 | 0.000246 | 98.1774 | 91.9716 | 92.0165 | 20.6922 | 0.007509 | 0.006257 | 134.4303 | 0.2551 | 8.4005 | 0.004775 | 0.005432 | 0.005272 | 6.1e-05 | 0.001145 | 0.26217 | 1.4064 | 2.5477 | 98.45 | 19.93 | 48.857 | 204.8848 | 136.7223 | 173.5694 | 0.011564 | 0.79031 | 0.01829 | 23.1433 | 0.58405 | 0.15176 | 0.071291 | 3.8676 | 3.8676 | 0.001581
2019-04-09 15:29 | 1554838147.7164 | 189 | 0.19482 | 0.014208 | 0.000253 | 98.0631 | 92.2252 | 92.5512 | 20.6928 | 0.007575 | 0.006196 | 134.2632 | 0.25435 | 8.4036 | 0.004652 | 0.006087 | 0.006011 | 5.9e-05 | 0.00115 | 0.27616 | 1.3532 | 2.5476 | 98.45 | 19.93 | 48.856 | 206.5456 | 133.6929 | 172.0556 | 0.011564 | 0.79299 | 0.018375 | 23.1711 | 0.54899 | 0.15132 | 0.071042 | 3.8715 | 3.8715 | 0.001575
2019-04-09 15:29 | 1554838148.7154 | 190 | 0.19505 | 0.013814 | 0.000261 | 97.7301 | 92.3646 | 93.0981 | 20.6934 | 0.007463 | 0.006185 | 134.1228 | 0.25343 | 8.4021 | 0.004648 | 0.006124 | 0.006063 | 6.2e-05 | 0.001149 | 0.2792 | 1.3098 | 2.5475 | 98.45 | 19.93 | 48.857 | 203.754 | 133.5888 | 171.7707 | 0.011564 | 0.7935 | 0.018435 | 23.232 | 0.50506 | 0.15076 | 0.070161 | 3.8359 | 3.8359 | 0.001584
2019-04-09 15:29 | 1554838149.7168 | 191 | 0.19539 | 0.013394 | 0.000267 | 97.4448 | 92.3261 | 93.6438 | 20.693 | 0.007447 | 0.006149 | 134.0077 | 0.25392 | 8.4015 | 0.004583 | 0.005512 | 0.005422 | 6.3e-05 | 0.001153 | 0.28479 | 1.3379 | 2.5476 | 98.45 | 19.93 | 48.856 | 203.35 | 131.9896 | 170.8809 | 0.011564 | 0.79509 | 0.018507 | 23.2767 | 0.48265 | 0.15106 | 0.068627 | 3.7668 | 3.7668 | 0.001616
2019-04-09 15:29 | 1554838150.7165 | 192 | 0.19589 | 0.013103 | 0.000274 | 97.1176 | 91.9399 | 94.2912 | 20.6922 | 0.007437 | 0.006058 | 133.8792 | 0.25324 | 8.4014 | 0.004497 | 0.00583 | 0.005778 | 6.3e-05 | 0.001158 | 0.25498 | 1.3368 | 2.5477 | 98.45 | 19.94 | 48.855 | 203.105 | 129.8734 | 168.6136 | 0.011571 | 0.79917 | 0.018619 | 23.2973 | 0.49004 | 0.15065 | 0.065946 | 3.6417 | 3.6417 | 0.001667
2019-04-09 15:29 | 1554838151.7155 | 193 | 0.19658 | 0.012798 | 0.00028 | 96.9329 | 91.9485 | 95.0102 | 20.6928 | 0.007327 | 0.005993 | 133.739 | 0.25311 | 8.3997 | 0.004552 | 0.005561 | 0.005432 | 6.3e-05 | 0.001155 | 0.2891 | 1.3699 | 2.5477 | 98.45 | 19.94 | 48.854 | 200.3507 | 131.2362 | 166.9766 | 0.011571 | 0.80214 | 0.018724 | 23.3426 | 0.47219 | 0.15058 | 0.062232 | 3.4564 | 3.4564 | 0.001756
2019-04-09 15:29 | 1554838152.7159 | 194 | 0.19722 | 0.012494 | 0.000287 | 96.536 | 91.7207 | 95.6996 | 20.6927 | 0.007309 | 0.006069 | 133.6269 | 0.25376 | 8.4029 | 0.004455 | 0.004856 | 0.00474 | 6e-05 | 0.001156 | 0.27887 | 1.3641 | 2.5476 | 98.45 | 19.94 | 48.853 | 199.8915 | 128.8506 | 168.8962 | 0.011571 | 0.79866 | 0.018751 | 23.4781 | 0.45767 | 0.15096 | 0.058777 | 3.2693 | 3.2693 | 0.001861
2019-04-09 15:29 | 1554838153.7157 | 195 | 0.1977 | 0.012029 | 0.000296 | 96.2349 | 91.8832 | 96.3293 | 20.6929 | 0.00725 | 0.005958 | 133.4904 | 0.25462 | 8.4003 | 0.004394 | 0.005337 | 0.005293 | 6.2e-05 | 0.001152 | 0.29319 | 1.3928 | 2.5476 | 98.45 | 19.94 | 48.854 | 198.4334 | 127.3527 | 166.1116 | 0.011571 | 0.80373 | 0.018872 | 23.4809 | 0.41391 | 0.15147 | 0.056449 | 3.1599 | 3.1599 | 0.001932
2019-04-09 15:29 | 1554838154.7155 | 196 | 0.19797 | 0.011712 | 0.000306 | 95.868 | 91.5851 | 96.8321 | 20.6931 | 0.007141 | 0.005924 | 133.3798 | 0.25354 | 8.4029 | 0.004386 | 0.006101 | 0.005923 | 6.2e-05 | 0.00115 | 0.29551 | 1.3723 | 2.5475 | 98.45 | 19.94 | 48.855 | 195.6911 | 127.1569 | 165.2678 | 0.011571 | 0.80527 | 0.01894 | 23.5195 | 0.4088 | 0.15083 | 0.055173 | 3.0992 | 3.0992 | 0.001961
2019-04-09 15:29 | 1554838155.7166 | 197 | 0.19837 | 0.011382 | 0.000316 | 95.5032 | 91.6037 | 97.5135 | 20.6923 | 0.007166 | 0.005918 | 133.3141 | 0.25403 | 8.4014 | 0.004329 | 0.005817 | 0.005767 | 6.2e-05 | 0.001139 | 0.27599 | 1.3592 | 2.5476 | 98.45 | 19.94 | 48.853 | 196.3135 | 125.7443 | 165.1139 | 0.011571 | 0.80555 | 0.019009 | 23.598 | 0.37289 | 0.15112 | 0.053157 | 2.9966 | 2.9966 | 0.002032
2019-04-09 15:29 | 1554838156.7161 | 198 | 0.19871 | 0.010954 | 0.000325 | 95.1224 | 91.2764 | 98.1001 | 20.6927 | 0.00714 | 0.005877 | 133.1987 | 0.25352 | 8.4028 | 0.004357 | 0.005551 | 0.005516 | 6e-05 | 0.001139 | 0.2891 | 1.3636 | 2.5476 | 98.45 | 19.94 | 48.854 | 195.6794 | 126.443 | 164.0851 | 0.011571 | 0.80745 | 0.019089 | 23.641 | 0.36916 | 0.15082 | 0.051629 | 2.9222 | 2.9222 | 0.00208
2019-04-09 15:29 | 1554838157.716 | 199 | 0.19903 | 0.010687 | 0.000336 | 94.7165 | 91.1105 | 98.5935 | 20.6925 | 0.007025 | 0.005835 | 133.1017 | 0.25399 | 8.4 | 0.004261 | 0.005509 | 0.005487 | 6.5e-05 | 0.001135 | 0.27577 | 1.3293 | 2.5476 | 98.45 | 19.94 | 48.854 | 192.8046 | 124.0893 | 163.0512 | 0.011571 | 0.80936 | 0.01916 | 23.6724 | 0.34718 | 0.1511 | 0.050019 | 2.841 | 2.841 | 0.002143
2019-04-09 15:29 | 1554838158.7157 | 200 | 0.19947 | 0.010408 | 0.000344 | 94.1682 | 91.2582 | 99.0993 | 20.6932 | 0.006932 | 0.005781 | 133.0059 | 0.25398 | 8.4011 | 0.004204 | 0.005455 | 0.005274 | 6.1e-05 | 0.001132 | 0.30525 | 1.3509 | 2.5477 | 98.45 | 19.94 | 48.855 | 190.4748 | 122.6838 | 161.6938 | 0.011571 | 0.81189 | 0.019239 | 23.6961 | 0.28077 | 0.15109 | 0.047769 | 2.7239 | 2.7239 | 0.002235
2019-04-09 15:29 | 1554838159.7147 | 201 | 0.19991 | 0.010123 | 0.000354 | 93.8827 | 91.0658 | 99.6988 | 20.6927 | 0.006902 | 0.005792 | 132.9073 | 0.25341 | 8.4014 | 0.004162 | 0.006022 | 0.005874 | 6.4e-05 | 0.001125 | 0.28622 | 1.3938 | 2.5477 | 98.45 | 19.94 | 48.855 | 189.7259 | 121.6653 | 161.9687 | 0.011571 | 0.81138 | 0.019291 | 23.7752 | 0.27248 | 0.15076 | 0.04543 | 2.5969 | 2.5969 | 0.002339
2019-04-09 15:29 | 1554838160.7151 | 202 | 0.2002 | 0.009697 | 0.000362 | 93.4616 | 90.9832 | 99.9734 | 20.6934 | 0.006895 | 0.005704 | 132.7834 | 0.25341 | 8.401 | 0.004152 | 0.005818 | 0.005741 | 6.2e-05 | 0.001128 | 0.28479 | 1.3651 | 2.5476 | 98.45 | 19.94 | 48.855 | 189.5447 | 121.4118 | 159.774 | 0.011571 | 0.81549 | 0.019366 | 23.7478 | 0.24039 | 0.15076 | 0.044186 | 2.5352 | 2.5352 | 0.002396
2019-04-09 15:29 | 1554838161.7157 | 203 | 0.20042 | 0.009421 | 0.000371 | 92.9708 | 90.4229 | 100.3245 | 20.6923 | 0.006743 | 0.005691 | 132.7136 | 0.254 | 8.4001 | 0.004147 | 0.005773 | 0.005692 | 6.1e-05 | 0.001125 | 0.274 | 1.3593 | 2.5475 | 98.45 | 19.94 | 48.854 | 185.7361 | 121.2949 | 159.4576 | 0.011571 | 0.81609 | 0.019407 | 23.7807 | 0.24855 | 0.1511 | 0.043229 | 2.4849 | 2.4849 | 0.002451
2019-04-09 15:29 | 1554838162.7159 | 204 | 0.20074 | 0.009153 | 0.000381 | 92.6025 | 90.0151 | 100.8402 | 20.6929 | 0.006873 | 0.005638 | 132.6256 | 0.25443 | 8.402 | 0.004144 | 0.005166 | 0.005133 | 6.2e-05 | 0.001127 | 0.27887 | 1.3612 | 2.5476 | 98.44 | 19.94 | 48.855 | 188.9825 | 121.2074 | 158.1364 | 0.011572 | 0.81859 | 0.019487 | 23.8053 | 0.25348 | 0.15136 | 0.04162 | 2.4015 | 2.4015 | 0.00254
2019-04-09 15:29 | 1554838163.7153 | 205 | 0.20109 | 0.008879 | 0.000389 | 91.9804 | 89.6207 | 101.2275 | 20.6927 | 0.006762 | 0.005584 | 132.5256 | 0.25347 | 8.4025 | 0.004189 | 0.005227 | 0.005047 | 6.4e-05 | 0.001125 | 0.27704 | 1.378 | 2.5477 | 98.44 | 19.94 | 48.854 | 186.2187 | 122.3227 | 156.7851 | 0.011572 | 0.82116 | 0.019555 | 23.8136 | 0.23246 | 0.15079 | 0.039849 | 2.3066 | 2.3066 | 0.002635
2019-04-09 15:29 | 1554838164.7149 | 206 | 0.20123 | 0.00852 | 0.000398 | 91.489 | 89.2972 | 101.7824 | 20.6922 | 0.006716 | 0.005571 | 132.453 | 0.25434 | 8.4002 | 0.004085 | 0.005951 | 0.005911 | 6.3e-05 | 0.001123 | 0.27151 | 1.349 | 2.5476 | 98.44 | 19.94 | 48.853 | 185.0741 | 119.769 | 156.4503 | 0.011572 | 0.8218 | 0.019616 | 23.8695 | 0.2169 | 0.1513 | 0.039451 | 2.2902 | 2.2902 | 0.002663
2019-04-09 15:29 | 1554838165.7163 | 207 | 0.20146 | 0.008207 | 0.000403 | 91.0006 | 89.0222 | 101.9923 | 20.6927 | 0.006687 | 0.005523 | 132.3534 | 0.25483 | 8.4045 | 0.004104 | 0.005607 | 0.00552 | 6.6e-05 | 0.001123 | 0.28766 | 1.3575 | 2.5475 | 98.44 | 19.94 | 48.851 | 184.3496 | 120.2311 | 155.2627 | 0.011571 | 0.82408 | 0.019663 | 23.861 | 0.19661 | 0.1516 | 0.038422 | 2.2353 | 2.2353 | 0.002733
2019-04-09 15:29 | 1554838166.716 | 208 | 0.20179 | 0.007972 | 0.000407 | 90.498 | 88.4526 | 102.257 | 20.6925 | 0.006634 | 0.005433 | 132.256 | 0.25356 | 8.3999 | 0.004135 | 0.005693 | 0.005648 | 6.8e-05 | 0.001123 | 0.29391 | 1.3336 | 2.5476 | 98.44 | 19.94 | 48.851 | 183.0057 | 120.9903 | 153.0384 | 0.011571 | 0.82838 | 0.01974 | 23.8299 | 0.20447 | 0.15084 | 0.036745 | 2.1454 | 2.1454 | 0.002834
2019-04-09 15:29 | 1554838167.716 | 209 | 0.20194 | 0.007743 | 0.000409 | 90.2273 | 88.0761 | 102.4751 | 20.6917 | 0.006596 | 0.005423 | 132.1955 | 0.25397 | 8.4035 | 0.004019 | 0.005592 | 0.005528 | 6.3e-05 | 0.001132 | 0.25984 | 1.331 | 2.5474 | 98.44 | 19.94 | 48.85 | 182.0666 | 118.1735 | 152.7785 | 0.011571 | 0.82888 | 0.019767 | 23.848 | 0.21583 | 0.15109 | 0.036092 | 2.1098 | 2.1098 | 0.002886
2019-04-09 15:29 | 1554838168.716 | 210 | 0.20225 | 0.007522 | 0.000412 | 90.0438 | 88.0803 | 102.7629 | 20.6927 | 0.006502 | 0.005402 | 132.1406 | 0.25403 | 8.4042 | 0.003998 | 0.004998 | 0.004844 | 6.6e-05 | 0.001135 | 0.27577 | 1.3567 | 2.5477 | 98.44 | 19.94 | 48.85 | 179.7045 | 117.65 | 152.2682 | 0.011571 | 0.82988 | 0.019807 | 23.8672 | 0.19721 | 0.15112 | 0.034494 | 2.0198 | 2.0198 | 0.003015
2019-04-09 15:29 | 1554838169.716 | 211 | 0.20251 | 0.00731 | 0.000413 | 90.1114 | 88.5025 | 102.6673 | 20.6923 | 0.006487 | 0.005366 | 132.0483 | 0.25216 | 8.4001 | 0.003994 | 0.005166 | 0.005075 | 6.5e-05 | 0.00114 | 0.29302 | 1.3522 | 2.5475 | 98.44 | 19.94 | 48.847 | 179.3265 | 117.5631 | 151.3541 | 0.01157 | 0.83166 | 0.019819 | 23.8304 | 0.16114 | 0.15001 | 0.033229 | 1.9464 | 1.9464 | 0.003106
2019-04-09 15:30 | 1554838170.715 | 212 | 0.20277 | 0.007101 | 0.000415 | 90.4214 | 88.5569 | 102.9577 | 20.6917 | 0.006385 | 0.005287 | 131.9964 | 0.25336 | 8.4008 | 0.004025 | 0.005943 | 0.005839 | 6.2e-05 | 0.001144 | 0.29231 | 1.3152 | 2.5475 | 98.44 | 19.94 | 48.845 | 176.7865 | 118.3129 | 149.3937 | 0.01157 | 0.83552 | 0.019893 | 23.8089 | 0.18637 | 0.15072 | 0.031944 | 1.8775 | 1.8775 | 0.003235
2019-04-09 15:30 | 1554838171.7161 | 213 | 0.20293 | 0.006896 | 0.000416 | 90.7213 | 88.9661 | 102.8499 | 20.6931 | 0.006306 | 0.005254 | 131.9297 | 0.25387 | 8.4017 | 0.00398 | 0.005097 | 0.005034 | 6.4e-05 | 0.001148 | 0.31277 | 1.3311 | 2.5476 | 98.44 | 19.94 | 48.844 | 174.7977 | 117.2178 | 148.5856 | 0.01157 | 0.83713 | 0.019902 | 23.7737 | 0.17475 | 0.15103 | 0.031197 | 1.8339 | 1.8339 | 0.003319
2019-04-09 15:30 | 1554838172.7162 | 214 | 0.20308 | 0.006701 | 0.000418 | 91.1195 | 89.5467 | 102.626 | 20.6925 | 0.006286 | 0.005233 | 131.843 | 0.25419 | 8.4004 | 0.003994 | 0.005113 | 0.005068 | 6.3e-05 | 0.001154 | 0.29966 | 1.3602 | 2.5477 | 98.44 | 19.94 | 48.841 | 174.2958 | 117.5517 | 148.0673 | 0.011569 | 0.83815 | 0.019892 | 23.7332 | 0.15574 | 0.15122 | 0.030556 | 1.795 | 1.795 | 0.003395
2019-04-09 15:30 | 1554838173.7156 | 215 | 0.2032 | 0.00651 | 0.00042 | 91.5956 | 90.1345 | 102.5225 | 20.6926 | 0.006269 | 0.00525 | 131.7796 | 0.25344 | 8.4008 | 0.00398 | 0.005882 | 0.00584 | 6.3e-05 | 0.001159 | 0.29767 | 1.3633 | 2.5476 | 98.44 | 19.94 | 48.839 | 173.8806 | 117.211 | 148.4941 | 0.011568 | 0.83731 | 0.019872 | 23.7332 | 0.14383 | 0.15077 | 0.030046 | 1.7629 | 1.7629 | 0.003447
2019-04-09 15:30 | 1554838174.7142 | 216 | 0.20344 | 0.006321 | 0.000422 | 92.0006 | 90.6511 | 102.3772 | 20.6929 | 0.006169 | 0.005154 | 131.7156 | 0.25368 | 8.4012 | 0.003954 | 0.005326 | 0.005243 | 6.4e-05 | 0.001165 | 0.28584 | 1.3526 | 2.5474 | 98.44 | 19.95 | 48.835 | 171.3868 | 116.5731 | 146.0911 | 0.011575 | 0.84211 | 0.019915 | 23.6487 | 0.13218 | 0.15091 | 0.02884 | 1.695 | 1.695 | 0.003588
2019-04-09 15:30 | 1554838175.7152 | 217 | 0.2037 | 0.006137 | 0.000424 | 92.4047 | 91.2021 | 102.4704 | 20.6932 | 0.006123 | 0.005162 | 131.6211 | 0.25335 | 8.4008 | 0.003865 | 0.005511 | 0.005429 | 6.4e-05 | 0.00117 | 0.2896 | 1.3526 | 2.5476 | 98.44 | 19.95 | 48.834 | 170.2431 | 114.4098 | 146.3046 | 0.011575 | 0.84168 | 0.019919 | 23.6655 | 0.11718 | 0.15072 | 0.027485 | 1.6149 | 1.6149 | 0.003761
2019-04-09 15:30 | 1554838176.7143 | 218 | 0.20378 | 0.005958 | 0.000425 | 92.8258 | 91.5693 | 102.7051 | 20.6932 | 0.006164 | 0.005125 | 131.5768 | 0.25399 | 8.4021 | 0.003791 | 0.005298 | 0.005257 | 5.9e-05 | 0.001167 | 0.29535 | 1.3706 | 2.5477 | 98.44 | 19.95 | 48.833 | 171.2589 | 112.6039 | 145.3764 | 0.011574 | 0.84354 | 0.019964 | 23.6664 | 0.1219 | 0.1511 | 0.027225 | 1.603 | 1.603 | 0.003799
2019-04-09 15:30 | 1554838177.7155 | 219 | 0.20397 | 0.00578 | 0.000426 | 93.2913 | 92.0728 | 102.9018 | 20.6912 | 0.006063 | 0.0051 | 131.5169 | 0.25364 | 8.4022 | 0.003858 | 0.005142 | 0.005058 | 6.2e-05 | 0.001165 | 0.28528 | 1.3951 | 2.5477 | 98.44 | 19.95 | 48.83 | 168.7438 | 114.2306 | 144.7568 | 0.011574 | 0.84479 | 0.019998 | 23.6715 | 0.11759 | 0.15089 | 0.026352 | 1.5536 | 1.5536 | 0.003914
2019-04-09 15:30 | 1554838178.7154 | 220 | 0.2041 | 0.005616 | 0.000427 | 93.7239 | 92.4617 | 103.175 | 20.6918 | 0.00599 | 0.005039 | 131.4346 | 0.25388 | 8.4029 | 0.00378 | 0.005429 | 0.005354 | 6.2e-05 | 0.001161 | 0.29231 | 1.4053 | 2.5477 | 98.44 | 19.95 | 48.828 | 166.9059 | 112.3253 | 143.2539 | 0.011573 | 0.84784 | 0.02006 | 23.6602 | 0.12128 | 0.15103 | 0.025765 | 1.5234 | 1.5234 | 0.003995
2019-04-09 15:30 | 1554838179.7144 | 221 | 0.20424 | 0.005457 | 0.000427 | 94.0125 | 92.9934 | 103.5919 | 20.6923 | 0.005942 | 0.004995 | 131.3852 | 0.25332 | 8.4009 | 0.003735 | 0.005151 | 0.004991 | 6.6e-05 | 0.001155 | 0.29192 | 1.3775 | 2.5475 | 98.44 | 19.95 | 48.823 | 165.709 | 111.2389 | 142.1717 | 0.011572 | 0.85005 | 0.020127 | 23.6772 | 0.09749 | 0.1507 | 0.025101 | 1.4886 | 1.4886 | 0.00408
2019-04-09 15:30 | 1554838180.7147 | 222 | 0.20437 | 0.005297 | 0.000426 | 94.3785 | 93.394 | 103.6912 | 20.6917 | 0.005941 | 0.004967 | 131.3277 | 0.25363 | 8.4031 | 0.003716 | 0.005019 | 0.004935 | 6.6e-05 | 0.001153 | 0.28727 | 1.3421 | 2.5476 | 98.44 | 19.95 | 48.82 | 165.681 | 110.7844 | 141.48 | 0.011571 | 0.85147 | 0.020153 | 23.6688 | 0.093794 | 0.15089 | 0.02451 | 1.4551 | 1.4551 | 0.004179
2019-04-09 15:30 | 1554838181.7144 | 223 | 0.20457 | 0.005145 | 0.000423 | 94.7013 | 93.9099 | 104.0318 | 20.6914 | 0.005912 | 0.004985 | 131.2555 | 0.25405 | 8.4014 | 0.003795 | 0.004356 | 0.0043 | 6.6e-05 | 0.001143 | 0.28208 | 1.341 | 2.5476 | 98.44 | 19.95 | 48.82 | 164.9663 | 112.7025 | 141.9124 | 0.011571 | 0.85058 | 0.020176 | 23.72 | 0.075064 | 0.15113 | 0.023518 | 1.3972 | 1.3972 | 0.004359
2019-04-09 15:30 | 1554838182.7155 | 224 | 0.20472 | 0.005001 | 0.00042 | 94.9836 | 94.2489 | 104.3071 | 20.6918 | 0.005894 | 0.004953 | 131.2088 | 0.25353 | 8.4009 | 0.00372 | 0.004926 | 0.004789 | 6.5e-05 | 0.001145 | 0.2724 | 1.3319 | 2.5476 | 98.44 | 19.96 | 48.815 | 164.5125 | 110.8707 | 141.1286 | 0.011577 | 0.85219 | 0.020222 | 23.7289 | 0.069459 | 0.15082 | 0.022754 | 1.3544 | 1.3544 | 0.004488
2019-04-09 15:30 | 1554838183.7146 | 225 | 0.20486 | 0.004862 | 0.000415 | 95.2299 | 94.4843 | 104.4059 | 20.6922 | 0.005787 | 0.004852 | 131.1749 | 0.25348 | 8.4014 | 0.003633 | 0.005583 | 0.005451 | 5.9e-05 | 0.001146 | 0.29264 | 1.3661 | 2.5477 | 98.44 | 19.96 | 48.812 | 161.8514 | 108.7544 | 138.6201 | 0.011576 | 0.85738 | 0.020293 | 23.6682 | 0.070308 | 0.15079 | 0.022107 | 1.3202 | 1.3202 | 0.004603
2019-04-09 15:30 | 1554838184.7145 | 226 | 0.205 | 0.004729 | 0.000411 | 95.6015 | 94.7689 | 104.731 | 20.6923 | 0.005756 | 0.004819 | 131.0945 | 0.25431 | 8.4029 | 0.003671 | 0.00515 | 0.005007 | 6.4e-05 | 0.001144 | 0.28313 | 1.35 | 2.5476 | 98.44 | 19.96 | 48.806 | 161.0693 | 109.6823 | 137.819 | 0.011575 | 0.85906 | 0.020344 | 23.6819 | 0.078243 | 0.15129 | 0.021469 | 1.2851 | 1.2851 | 0.004744
2019-04-09 15:30 | 1554838185.7145 | 227 | 0.20511 | 0.004593 | 0.000405 | 95.9282 | 95.1578 | 105.0746 | 20.6924 | 0.005732 | 0.00484 | 131.08 | 0.25378 | 8.4012 | 0.003665 | 0.004401 | 0.004438 | 6e-05 | 0.001147 | 0.2907 | 1.331 | 2.5475 | 98.44 | 19.96 | 48.804 | 160.4668 | 109.5414 | 138.3407 | 0.011575 | 0.85797 | 0.020365 | 23.7358 | 0.072126 | 0.15098 | 0.020945 | 1.2548 | 1.2548 | 0.004849
2019-04-09 15:30 | 1554838186.7168 | 228 | 0.20521 | 0.00446 | 0.0004 | 96.0723 | 95.5257 | 105.2669 | 20.6916 | 0.005741 | 0.00479 | 131.0475 | 0.25376 | 8.4027 | 0.003736 | 0.005494 | 0.005319 | 6.1e-05 | 0.001148 | 0.24779 | 1.3371 | 2.5477 | 98.44 | 19.96 | 48.802 | 160.6982 | 111.2587 | 137.0949 | 0.011574 | 0.86057 | 0.020414 | 23.7215 | 0.05104 | 0.15096 | 0.020552 | 1.2342 | 1.2342 | 0.004929
2019-04-09 15:30 | 1554838187.7146 | 229 | 0.20533 | 0.004333 | 0.000395 | 96.3552 | 95.8537 | 105.4185 | 20.6912 | 0.005599 | 0.004762 | 130.9699 | 0.25354 | 8.4029 | 0.003637 | 0.005418 | 0.005294 | 6.2e-05 | 0.001145 | 0.28174 | 1.3376 | 2.5476 | 98.44 | 19.96 | 48.799 | 157.1515 | 108.8663 | 136.4018 | 0.011573 | 0.86203 | 0.020446 | 23.7185 | 0.046683 | 0.15083 | 0.019962 | 1.2004 | 1.2004 | 0.005064
2019-04-09 15:30 | 1554838188.7153 | 230 | 0.20533 | 0.00421 | 0.000388 | 96.6006 | 96.1858 | 105.5713 | 20.6921 | 0.005584 | 0.00477 | 130.9259 | 0.25387 | 8.4013 | 0.003585 | 0.005645 | 0.005517 | 6.2e-05 | 0.001141 | 0.28871 | 1.3205 | 2.5476 | 98.44 | 19.96 | 48.799 | 156.795 | 107.5929 | 136.6045 | 0.011573 | 0.8616 | 0.020456 | 23.7416 | 0.038495 | 0.15103 | 0.020152 | 1.213 | 1.213 | 0.005018
2019-04-09 15:30 | 1554838189.7141 | 231 | 0.20553 | 0.004096 | 0.00038 | 96.7972 | 96.3059 | 105.7018 | 20.6922 | 0.00559 | 0.004741 | 130.878 | 0.25409 | 8.3999 | 0.003604 | 0.005528 | 0.005419 | 6.1e-05 | 0.001137 | 0.26609 | 1.3577 | 2.5476 | 98.44 | 19.96 | 48.8 | 156.9246 | 108.0538 | 135.8937 | 0.011574 | 0.8631 | 0.020486 | 23.7356 | 0.045513 | 0.15116 | 0.019093 | 1.1504 | 1.1504 | 0.005295
2019-04-09 15:30 | 1554838190.7138 | 232 | 0.20557 | 0.003989 | 0.000372 | 97.0068 | 96.5987 | 105.6382 | 20.6913 | 0.005554 | 0.004665 | 130.8314 | 0.25344 | 8.4022 | 0.003635 | 0.005486 | 0.005388 | 6.4e-05 | 0.001137 | 0.28766 | 1.3811 | 2.5477 | 98.44 | 19.97 | 48.801 | 156.0503 | 108.8119 | 134.0041 | 0.011581 | 0.86711 | 0.020528 | 23.6736 | 0.037704 | 0.15077 | 0.018966 | 1.1454 | 1.1454 | 0.005305
2019-04-09 15:30 | 1554838191.7148 | 233 | 0.20574 | 0.003888 | 0.000366 | 97.1885 | 96.8121 | 105.7518 | 20.6922 | 0.005505 | 0.004666 | 130.8084 | 0.25403 | 8.4016 | 0.003636 | 0.004642 | 0.004568 | 5.9e-05 | 0.001135 | 0.27367 | 1.3643 | 2.5475 | 98.44 | 19.97 | 48.803 | 154.8191 | 108.8238 | 134.0265 | 0.011582 | 0.86706 | 0.020538 | 23.687 | 0.034702 | 0.15112 | 0.018074 | 1.0916 | 1.0916 | 0.005579
2019-04-09 15:30 | 1554838192.7148 | 234 | 0.20586 | 0.003781 | 0.00036 | 97.4079 | 97.0754 | 106.1039 | 20.692 | 0.005448 | 0.004552 | 130.7818 | 0.25369 | 8.4011 | 0.003651 | 0.005014 | 0.004954 | 5.8e-05 | 0.001141 | 0.27096 | 1.3769 | 2.5476 | 98.44 | 19.97 | 48.806 | 153.409 | 109.2081 | 131.2244 | 0.011582 | 0.87307 | 0.020643 | 23.6446 | 0.030587 | 0.15092 | 0.017535 | 1.0643 | 1.0643 | 0.005715
2019-04-09 15:30 | 1554838193.7142 | 235 | 0.20596 | 0.003688 | 0.000354 | 97.5577 | 97.1762 | 106.4811 | 20.691 | 0.005413 | 0.004613 | 130.7236 | 0.25341 | 8.4037 | 0.00362 | 0.005109 | 0.004999 | 5.8e-05 | 0.00114 | 0.27688 | 1.3919 | 2.5476 | 98.44 | 19.97 | 48.808 | 152.5255 | 108.4553 | 132.7193 | 0.011583 | 0.86985 | 0.020642 | 23.7303 | 0.035044 | 0.15076 | 0.017056 | 1.0349 | 1.0349 | 0.00587
2019-04-09 15:30 | 1554838194.7139 | 236 | 0.20607 | 0.003592 | 0.000347 | 97.7429 | 97.1404 | 106.8299 | 20.6917 | 0.005429 | 0.004578 | 130.6971 | 0.25305 | 8.4014 | 0.003627 | 0.004508 | 0.004482 | 5.7e-05 | 0.001148 | 0.28174 | 1.3318 | 2.5476 | 98.44 | 19.97 | 48.811 | 152.9334 | 108.6121 | 131.8708 | 0.011583 | 0.87167 | 0.020697 | 23.7443 | 0.055305 | 0.15054 | 0.016504 | 1.0039 | 1.0039 | 0.006043
2019-04-09 15:30 | 1554838195.7149 | 237 | 0.20614 | 0.003499 | 0.000341 | 97.8114 | 97.4595 | 107.0139 | 20.6913 | 0.005402 | 0.004529 | 130.681 | 0.25363 | 8.4018 | 0.003554 | 0.00426 | 0.004164 | 5.6e-05 | 0.001154 | 0.30525 | 1.3517 | 2.5475 | 98.44 | 19.97 | 48.818 | 152.2516 | 106.8322 | 130.66 | 0.011585 | 0.87429 | 0.020746 | 23.7292 | 0.032243 | 0.15088 | 0.016239 | 0.99021 | 0.99021 | 0.006141
2019-04-09 15:30 | 1554838196.7139 | 238 | 0.20624 | 0.003411 | 0.000333 | 97.9961 | 97.6584 | 106.9599 | 20.6912 | 0.005329 | 0.004485 | 130.6124 | 0.25337 | 8.4024 | 0.00361 | 0.004485 | 0.004456 | 5.4e-05 | 0.001158 | 0.28008 | 1.3418 | 2.5477 | 98.44 | 19.97 | 48.822 | 150.4462 | 108.1956 | 129.5742 | 0.011586 | 0.87664 | 0.020769 | 23.6913 | 0.030874 | 0.15073 | 0.015748 | 0.96126 | 0.96126 | 0.006319
2019-04-09 15:30 | 1554838197.7139 | 239 | 0.20629 | 0.003316 | 0.000326 | 98.2221 | 97.8109 | 106.7355 | 20.6906 | 0.005233 | 0.004452 | 130.5954 | 0.2534 | 8.4028 | 0.003536 | 0.004884 | 0.00478 | 5.4e-05 | 0.001157 | 0.29192 | 1.3088 | 2.5476 | 98.44 | 19.97 | 48.828 | 148.0468 | 106.3994 | 128.7674 | 0.011587 | 0.8784 | 0.020768 | 23.6427 | 0.037527 | 0.15075 | 0.015565 | 0.95018 | 0.95018 | 0.006394
2019-04-09 15:30 | 1554838198.714 | 240 | 0.20638 | 0.003225 | 0.000319 | 98.3049 | 98.1368 | 106.8787 | 20.6913 | 0.005253 | 0.004457 | 130.5952 | 0.25426 | 8.3999 | 0.003552 | 0.005396 | 0.005306 | 5.4e-05 | 0.001158 | 0.28296 | 1.3386 | 2.5476 | 98.44 | 19.97 | 48.829 | 148.5476 | 106.7827 | 128.8908 | 0.011588 | 0.87813 | 0.020779 | 23.6622 | 0.015308 | 0.15126 | 0.015122 | 0.92355 | 0.92355 | 0.0066
2019-04-09 15:30 | 1554838199.714 | 241 | 0.2065 | 0.003142 | 0.000314 | 98.4365 | 98.0576 | 107.2611 | 20.6904 | 0.005243 | 0.004416 | 130.5667 | 0.25408 | 8.4018 | 0.003478 | 0.005435 | 0.00535 | 5.4e-05 | 0.001167 | 0.30879 | 1.329 | 2.5475 | 98.44 | 19.97 | 48.83 | 148.3171 | 105.0003 | 127.8855 | 0.011588 | 0.88033 | 0.020842 | 23.6748 | 0.034494 | 0.15115 | 0.014544 | 0.8907 | 0.8907 | 0.006839
2019-04-09 15:30 | 1554838200.7154 | 242 | 0.2066 | 0.003055 | 0.000308 | 98.5109 | 98.2469 | 107.6106 | 20.6917 | 0.005189 | 0.004385 | 130.5279 | 0.25339 | 8.4018 | 0.003426 | 0.005626 | 0.005548 | 5.5e-05 | 0.001169 | 0.3027 | 1.3151 | 2.5475 | 98.44 | 19.97 | 48.828 | 146.9759 | 103.7307 | 127.114 | 0.011587 | 0.88203 | 0.020896 | 23.6906 | 0.024007 | 0.15074 | 0.014009 | 0.85989 | 0.85989 | 0.007065
2019-04-09 15:30 | 1554838201.7151 | 243 | 0.20675 | 0.002968 | 0.000303 | 98.6712 | 98.6038 | 107.7275 | 20.6914 | 0.00509 | 0.004366 | 130.4944 | 0.25392 | 8.4042 | 0.003485 | 0.005034 | 0.004981 | 5.4e-05 | 0.001176 | 0.28711 | 1.3714 | 2.5475 | 98.44 | 19.97 | 48.826 | 144.5167 | 105.164 | 126.6653 | 0.011587 | 0.88302 | 0.020919 | 23.6901 | 0.006109 | 0.15106 | 0.013263 | 0.81441 | 0.81441 | 0.007475
2019-04-09 15:30 | 1554838202.7151 | 244 | 0.20684 | 0.002891 | 0.000298 | 98.8147 | 98.7339 | 107.7974 | 20.6908 | 0.005116 | 0.004363 | 130.4415 | 0.25309 | 8.4019 | 0.003522 | 0.005364 | 0.005261 | 5.8e-05 | 0.001176 | 0.29607 | 1.3336 | 2.5476 | 98.44 | 19.97 | 48.824 | 145.1626 | 106.0706 | 126.5884 | 0.011587 | 0.88319 | 0.020928 | 23.6955 | 0.007318 | 0.15057 | 0.012807 | 0.78652 | 0.78652 | 0.007715
2019-04-09 15:30 | 1554838203.7145 | 245 | 0.20689 | 0.00281 | 0.000294 | 98.9112 | 98.8812 | 107.947 | 20.6907 | 0.005097 | 0.00431 | 130.4239 | 0.25372 | 8.4 | 0.003477 | 0.005116 | 0.004987 | 5.3e-05 | 0.001177 | 0.29319 | 1.3402 | 2.5476 | 98.44 | 19.97 | 48.82 | 144.6754 | 104.9745 | 125.2925 | 0.011586 | 0.88606 | 0.020976 | 23.6735 | 0.002714 | 0.15094 | 0.012594 | 0.77524 | 0.77524 | 0.007846
2019-04-09 15:30 | 1554838204.7155 | 246 | 0.20702 | 0.00273 | 0.000289 | 99.0356 | 98.8323 | 108.158 | 20.6909 | 0.005021 | 0.004296 | 130.4054 | 0.25435 | 8.4013 | 0.003513 | 0.004746 | 0.004656 | 5.3e-05 | 0.00118 | 0.27743 | 1.3357 | 2.5476 | 98.44 | 19.97 | 48.818 | 142.8143 | 105.8463 | 124.9425 | 0.011585 | 0.88684 | 0.021006 | 23.6862 | 0.018385 | 0.15132 | 0.011907 | 0.73345 | 0.73345 | 0.008314
2019-04-09 15:30 | 1554838205.7142 | 247 | 0.20703 | 0.002653 | 0.000285 | 99.1018 | 98.9242 | 108.2028 | 20.691 | 0.004946 | 0.004291 | 130.3421 | 0.25331 | 8.4013 | 0.003455 | 0.005005 | 0.004894 | 5.1e-05 | 0.001181 | 0.28528 | 1.3734 | 2.5476 | 98.44 | 19.98 | 48.814 | 140.9519 | 104.4463 | 124.8099 | 0.011591 | 0.88714 | 0.021014 | 23.6872 | 0.016041 | 0.15069 | 0.011991 | 0.73924 | 0.73924 | 0.008215
2019-04-09 15:30 | 1554838206.7132 | 248 | 0.2071 | 0.002582 | 0.000281 | 99.2786 | 99.1644 | 108.3096 | 20.6907 | 0.004957 | 0.004223 | 130.3182 | 0.25394 | 8.4048 | 0.003412 | 0.005159 | 0.005093 | 5.1e-05 | 0.00117 | 0.27959 | 1.3726 | 2.5476 | 98.44 | 19.98 | 48.812 | 141.2237 | 103.3879 | 123.1462 | 0.011591 | 0.89086 | 0.021068 | 23.6493 | 0.010298 | 0.15107 | 0.011627 | 0.7184 | 0.7184 | 0.008474
2019-04-09 15:30 | 1554838207.7146 | 249 | 0.20716 | 0.002518 | 0.000275 | 99.4052 | 99.1634 | 108.6734 | 20.691 | 0.004902 | 0.004209 | 130.2892 | 0.25348 | 8.4014 | 0.003453 | 0.005024 | 0.004899 | 5.3e-05 | 0.001169 | 0.29734 | 1.3576 | 2.5477 | 98.44 | 19.98 | 48.809 | 139.8561 | 104.3899 | 122.8049 | 0.01159 | 0.89163 | 0.021113 | 23.6788 | 0.021787 | 0.15079 | 0.011364 | 0.70359 | 0.70359 | 0.008637
2019-04-09 15:30 | 1554838208.7137 | 250 | 0.20728 | 0.002453 | 0.000269 | 99.4415 | 99.2357 | 108.8502 | 20.6912 | 0.004899 | 0.00423 | 130.2353 | 0.25391 | 8.4031 | 0.003439 | 0.005056 | 0.005023 | 5.3e-05 | 0.001164 | 0.26936 | 1.3596 | 2.5475 | 98.44 | 19.98 | 48.803 | 139.7922 | 104.0559 | 123.3153 | 0.011589 | 0.89048 | 0.021116 | 23.7133 | 0.018539 | 0.15105 | 0.010717 | 0.66322 | 0.66322 | 0.009178
2019-04-09 15:30 | 1554838209.7147 | 251 | 0.20735 | 0.002394 | 0.000264 | 99.5094 | 99.1703 | 108.7673 | 20.6906 | 0.004844 | 0.004186 | 130.1924 | 0.25374 | 8.4023 | 0.003405 | 0.0054 | 0.005223 | 5e-05 | 0.00116 | 0.28008 | 1.3343 | 2.5475 | 98.44 | 19.98 | 48.798 | 138.4325 | 103.2408 | 122.2364 | 0.011588 | 0.89291 | 0.021137 | 23.672 | 0.03054 | 0.15095 | 0.01041 | 0.64469 | 0.64469 | 0.009436
2019-04-09 15:30 | 1554838210.7145 | 252 | 0.20742 | 0.002331 | 0.000259 | 99.5496 | 99.3268 | 109.0293 | 20.6916 | 0.004834 | 0.004097 | 130.194 | 0.25389 | 8.4016 | 0.003399 | 0.005126 | 0.005152 | 5.1e-05 | 0.00116 | 0.26737 | 1.3544 | 2.5475 | 98.44 | 19.98 | 48.792 | 138.1764 | 103.0801 | 120.0618 | 0.011586 | 0.89785 | 0.021221 | 23.6352 | 0.02004 | 0.15104 | 0.010069 | 0.62588 | 0.62588 | 0.009725
2019-04-09 15:30 | 1554838211.7137 | 253 | 0.20747 | 0.00227 | 0.000255 | 99.6086 | 99.4958 | 109.5776 | 20.6905 | 0.004813 | 0.004134 | 130.1675 | 0.25378 | 8.4019 | 0.003366 | 0.004815 | 0.004728 | 5.2e-05 | 0.001165 | 0.27417 | 1.384 | 2.5477 | 98.44 | 19.98 | 48.788 | 137.6688 | 102.2966 | 120.9646 | 0.011585 | 0.89579 | 0.02125 | 23.7217 | 0.010132 | 0.15097 | 0.009841 | 0.61251 | 0.61251 | 0.009933
2019-04-09 15:30 | 1554838212.7144 | 254 | 0.2075 | 0.002218 | 0.00025 | 99.6967 | 99.5065 | 109.774 | 20.6908 | 0.00478 | 0.004079 | 130.15 | 0.2533 | 8.4022 | 0.003279 | 0.00481 | 0.004763 | 4.9e-05 | 0.00116 | 0.274 | 1.37 | 2.5476 | 98.44 | 19.98 | 48.784 | 136.8354 | 100.1686 | 119.6303 | 0.011584 | 0.89884 | 0.021305 | 23.7028 | 0.017085 | 0.15069 | 0.009724 | 0.60687 | 0.60687 | 0.010007
2019-04-09 15:30 | 1554838213.7134 | 255 | 0.20761 | 0.002164 | 0.000245 | 99.7751 | 99.5051 | 110.0266 | 20.6906 | 0.004708 | 0.004105 | 130.1447 | 0.25304 | 8.4029 | 0.003317 | 0.004471 | 0.004416 | 4.9e-05 | 0.001167 | 0.28296 | 1.3609 | 2.5477 | 98.44 | 19.98 | 48.778 | 135.0718 | 101.0975 | 120.2693 | 0.011583 | 0.89738 | 0.021312 | 23.7493 | 0.024233 | 0.15053 | 0.009158 | 0.5713 | 0.5713 | 0.010619
2019-04-09 15:30 | 1554838214.7138 | 256 | 0.2077 | 0.002106 | 0.00024 | 99.8954 | 99.7777 | 110.0463 | 20.6907 | 0.004752 | 0.004044 | 130.1063 | 0.25344 | 8.4034 | 0.003351 | 0.005078 | 0.005041 | 4.9e-05 | 0.001171 | 0.2656 | 1.3797 | 2.5476 | 98.44 | 19.98 | 48.775 | 136.1515 | 101.9127 | 118.782 | 0.011582 | 0.90078 | 0.021354 | 23.7065 | 0.010543 | 0.15077 | 0.008696 | 0.5432 | 0.5432 | 0.011186
2019-04-09 15:30 | 1554838215.7146 | 257 | 0.20779 | 0.002051 | 0.000235 | 99.9195 | 99.7474 | 110.4331 | 20.6916 | 0.004651 | 0.003998 | 130.0794 | 0.25413 | 8.4023 | 0.003302 | 0.004717 | 0.004631 | 4.9e-05 | 0.001175 | 0.29214 | 1.3505 | 2.5476 | 98.44 | 19.98 | 48.771 | 133.6597 | 100.7445 | 117.6465 | 0.011581 | 0.9034 | 0.021423 | 23.7137 | 0.015418 | 0.15118 | 0.008233 | 0.51557 | 0.51557 | 0.011817
2019-04-09 15:30 | 1554838216.7135 | 258 | 0.20783 | 0.001994 | 0.00023 | 100.0157 | 99.789 | 110.8887 | 20.6905 | 0.004658 | 0.003983 | 130.0625 | 0.25409 | 8.4005 | 0.003405 | 0.004705 | 0.004651 | 4.9e-05 | 0.001175 | 0.28639 | 1.3929 | 2.5476 | 98.44 | 19.98 | 48.761 | 133.8369 | 103.2368 | 117.2723 | 0.011579 | 0.90426 | 0.021477 | 23.7512 | 0.020291 | 0.15116 | 0.008058 | 0.50586 | 0.50586 | 0.012042
2019-04-09 15:30 | 1554838217.7136 | 259 | 0.20787 | 0.001939 | 0.000226 | 100.066 | 99.8503 | 111.1192 | 20.6904 | 0.004581 | 0.00395 | 130.0566 | 0.25393 | 8.4039 | 0.003386 | 0.004669 | 0.004545 | 5.1e-05 | 0.001167 | 0.27721 | 1.4162 | 2.5476 | 98.44 | 19.98 | 48.757 | 131.9534 | 102.7747 | 116.4764 | 0.011578 | 0.90611 | 0.021522 | 23.7516 | 0.019301 | 0.15107 | 0.00787 | 0.49501 | 0.49501 | 0.012298
2019-04-09 15:30 | 1554838218.7146 | 260 | 0.20792 | 0.001889 | 0.000223 | 100.0338 | 100.0444 | 110.9763 | 20.6907 | 0.004649 | 0.003949 | 130.0339 | 0.25329 | 8.3981 | 0.003308 | 0.005061 | 0.005038 | 5.3e-05 | 0.001167 | 0.29048 | 1.3899 | 2.5476 | 98.44 | 19.98 | 48.747 | 133.627 | 100.8787 | 116.4499 | 0.011575 | 0.90617 | 0.021509 | 23.7355 | -0.000944 | 0.15069 | 0.007633 | 0.47965 | 0.47965 | 0.01266
2019-04-09 15:30 | 1554838219.7146 | 261 | 0.20793 | 0.00183 | 0.000218 | 100.1875 | 100.073 | 111.1077 | 20.6904 | 0.004602 | 0.003968 | 130.0433 | 0.25422 | 8.4029 | 0.003326 | 0.005087 | 0.004999 | 5.1e-05 | 0.001166 | 0.29358 | 1.3947 | 2.5477 | 98.44 | 19.98 | 48.742 | 132.4479 | 101.3274 | 116.9113 | 0.011574 | 0.9051 | 0.021509 | 23.7636 | 0.010231 | 0.15124 | 0.007695 | 0.48383 | 0.48383 | 0.012597
2019-04-09 15:30 | 1554838220.7146 | 262 | 0.20798 | 0.001786 | 0.000214 | 100.2069 | 99.9384 | 111.3777 | 20.6907 | 0.00454 | 0.003934 | 130.0067 | 0.25324 | 8.4039 | 0.003367 | 0.004854 | 0.004793 | 5.2e-05 | 0.001168 | 0.28694 | 1.3334 | 2.5478 | 98.44 | 19.98 | 48.735 | 130.9403 | 102.3077 | 116.0873 | 0.011573 | 0.90702 | 0.021557 | 23.7674 | 0.023995 | 0.15065 | 0.007431 | 0.46813 | 0.46813 | 0.01297
2019-04-09 15:30 | 1554838221.7131 | 263 | 0.20801 | 0.001734 | 0.000211 | 100.2039 | 99.9701 | 111.1587 | 20.6918 | 0.00448 | 0.003935 | 129.9606 | 0.25326 | 8.4013 | 0.003381 | 0.004876 | 0.004792 | 5.2e-05 | 0.001177 | 0.27505 | 1.3092 | 2.5476 | 98.44 | 19.98 | 48.729 | 129.4518 | 102.6567 | 116.1082 | 0.011571 | 0.90697 | 0.021536 | 23.7446 | 0.020895 | 0.15066 | 0.00731 | 0.4601 | 0.4601 | 0.013196
2019-04-09 15:30 | 1554838222.7137 | 264 | 0.20804 | 0.001693 | 0.000207 | 100.3191 | 100.1618 | 111.4234 | 20.6906 | 0.00452 | 0.003922 | 129.9431 | 0.2537 | 8.4054 | 0.003324 | 0.004755 | 0.004713 | 5.2e-05 | 0.001176 | 0.29551 | 1.3156 | 2.5476 | 98.44 | 19.98 | 48.724 | 130.4335 | 101.2805 | 115.7834 | 0.01157 | 0.90773 | 0.02157 | 23.763 | 0.014042 | 0.15093 | 0.007168 | 0.45188 | 0.45188 | 0.01346
2019-04-09 15:30 | 1554838223.7138 | 265 | 0.20813 | 0.001647 | 0.000203 | 100.3986 | 100.2447 | 111.4225 | 20.6906 | 0.00442 | 0.003914 | 129.9138 | 0.25366 | 8.4028 | 0.003286 | 0.005062 | 0.004948 | 4.9e-05 | 0.001171 | 0.29336 | 1.3352 | 2.5476 | 98.44 | 19.98 | 48.717 | 127.995 | 100.3519 | 115.5895 | 0.011568 | 0.90818 | 0.021576 | 23.7569 | 0.013724 | 0.1509 | 0.006736 | 0.42426 | 0.42426 | 0.014334
2019-04-09 15:30 | 1554838224.7128 | 266 | 0.20819 | 0.001607 | 0.0002 | 100.3966 | 100.2065 | 111.4997 | 20.6902 | 0.004409 | 0.003831 | 129.8991 | 0.25273 | 8.4043 | 0.00334 | 0.005095 | 0.00507 | 5.1e-05 | 0.00117 | 0.26913 | 1.3396 | 2.5477 | 98.44 | 19.98 | 48.713 | 127.7109 | 101.6648 | 113.5711 | 0.011567 | 0.91292 | 0.021639 | 23.7034 | 0.016948 | 0.15035 | 0.00638 | 0.40261 | 0.40261 | 0.015049
2019-04-09 15:30 | 1554838225.7129 | 267 | 0.2082 | 0.001569 | 0.000197 | 100.4424 | 100.1573 | 111.6018 | 20.6892 | 0.00444 | 0.003884 | 129.8924 | 0.25319 | 8.4011 | 0.003263 | 0.00478 | 0.004615 | 5.1e-05 | 0.001167 | 0.28385 | 1.3837 | 2.5477 | 98.44 | 19.98 | 48.709 | 128.4773 | 99.7989 | 114.8613 | 0.011566 | 0.90988 | 0.021613 | 23.7538 | 0.025427 | 0.15062 | 0.006374 | 0.40192 | 0.40192 | 0.015103
2019-04-09 15:30 | 1554838226.7142 | 268 | 0.20826 | 0.001532 | 0.000194 | 100.4972 | 100.2337 | 111.5638 | 20.6911 | 0.004377 | 0.00384 | 129.8869 | 0.2541 | 8.4017 | 0.003287 | 0.004845 | 0.004769 | 5.3e-05 | 0.001172 | 0.27903 | 1.3307 | 2.5476 | 98.44 | 19.99 | 48.702 | 126.9168 | 100.3633 | 113.7959 | 0.011572 | 0.91239 | 0.021639 | 23.7171 | 0.023484 | 0.15116 | 0.006075 | 0.3832 | 0.3832 | 0.015897
2019-04-09 15:30 | 1554838227.7139 | 269 | 0.20824 | 0.001492 | 0.000192 | 100.5505 | 100.4161 | 111.3739 | 20.6911 | 0.004328 | 0.003745 | 129.8547 | 0.25456 | 8.4007 | 0.003314 | 0.004875 | 0.004826 | 5.1e-05 | 0.001168 | 0.28313 | 1.3488 | 2.5477 | 98.44 | 19.99 | 48.7 | 125.7289 | 101.021 | 111.4831 | 0.011571 | 0.91788 | 0.021686 | 23.626 | 0.011965 | 0.15144 | 0.006264 | 0.39641 | 0.39641 | 0.015396
2019-04-09 15:30 | 1554838228.7139 | 270 | 0.20832 | 0.001456 | 0.00019 | 100.5833 | 100.5118 | 111.3901 | 20.6913 | 0.004327 | 0.003725 | 129.846 | 0.25366 | 8.4026 | 0.003275 | 0.005061 | 0.004936 | 5e-05 | 0.001166 | 0.29264 | 1.3356 | 2.5477 | 98.44 | 19.99 | 48.694 | 125.6933 | 100.089 | 111.009 | 0.01157 | 0.91901 | 0.021701 | 23.6131 | 0.006361 | 0.1509 | 0.005853 | 0.37012 | 0.37012 | 0.016431
2019-04-09 15:30 | 1554838229.713 | 271 | 0.20836 | 0.001422 | 0.000187 | 100.5868 | 100.5218 | 111.7349 | 20.6906 | 0.004287 | 0.003775 | 129.8254 | 0.25326 | 8.3989 | 0.003291 | 0.004826 | 0.004875 | 5e-05 | 0.001167 | 0.26648 | 1.3787 | 2.5477 | 98.44 | 19.98 | 48.693 | 124.7243 | 100.4751 | 112.2027 | 0.011563 | 0.91616 | 0.021701 | 23.6864 | 0.00579 | 0.15067 | 0.005612 | 0.35458 | 0.35458 | 0.017124
2019-04-09 15:31 | 1554838230.7144 | 272 | 0.20841 | 0.001387 | 0.000184 | 100.5901 | 100.3246 | 111.6354 | 20.69 | 0.004341 | 0.003691 | 129.801 | 0.25412 | 8.4006 | 0.003226 | 0.004726 | 0.004641 | 5e-05 | 0.001159 | 0.24674 | 1.3663 | 2.5477 | 98.44 | 19.99 | 48.691 | 126.0453 | 98.9055 | 110.1793 | 0.011569 | 0.921 | 0.021748 | 23.6136 | 0.023646 | 0.15118 | 0.005406 | 0.3421 | 0.3421 | 0.017809
2019-04-09 15:31 | 1554838231.7129 | 273 | 0.20848 | 0.001351 | 0.000181 | 100.6235 | 100.3246 | 111.742 | 20.6908 | 0.004283 | 0.003736 | 129.8138 | 0.25313 | 8.4024 | 0.003232 | 0.004299 | 0.004205 | 5.2e-05 | 0.001162 | 0.27721 | 1.3454 | 2.5475 | 98.44 | 19.99 | 48.693 | 124.6307 | 99.0498 | 111.2751 | 0.01157 | 0.91837 | 0.021727 | 23.6586 | 0.026611 | 0.15059 | 0.004987 | 0.31475 | 0.31475 | 0.019281
2019-04-09 15:31 | 1554838232.7131 | 274 | 0.20853 | 0.001324 | 0.000178 | 100.6544 | 100.287 | 111.7859 | 20.6903 | 0.004271 | 0.003714 | 129.8118 | 0.25383 | 8.4024 | 0.00325 | 0.004131 | 0.004066 | 5.1e-05 | 0.001155 | 0.28119 | 1.3385 | 2.5475 | 98.44 | 19.99 | 48.693 | 124.332 | 99.4736 | 110.724 | 0.01157 | 0.91969 | 0.021747 | 23.6462 | 0.032709 | 0.151 | 0.004763 | 0.30053 | 0.30053 | 0.020249
2019-04-09 15:31 | 1554838233.7141 | 275 | 0.20854 | 0.00129 | 0.000175 | 100.7417 | 100.6298 | 111.9957 | 20.6915 | 0.004203 | 0.003703 | 129.7806 | 0.25339 | 8.4021 | 0.003177 | 0.004908 | 0.004836 | 5.2e-05 | 0.001151 | 0.26521 | 1.3287 | 2.5476 | 98.44 | 19.98 | 48.697 | 122.6631 | 97.7116 | 110.456 | 0.011564 | 0.92033 | 0.021775 | 23.6602 | 0.009939 | 0.15074 | 0.004743 | 0.29983 | 0.29983 | 0.020261
2019-04-09 15:31 | 1554838234.7131 | 276 | 0.20856 | 0.001263 | 0.000173 | 100.7471 | 100.5324 | 112.0337 | 20.6907 | 0.004241 | 0.003665 | 129.7839 | 0.2533 | 8.4026 | 0.003126 | 0.004588 | 0.0045 | 4.9e-05 | 0.00115 | 0.3027 | 1.3569 | 2.5476 | 98.44 | 19.98 | 48.699 | 123.597 | 96.4799 | 109.542 | 0.011564 | 0.92253 | 0.021805 | 23.636 | 0.019076 | 0.15069 | 0.004639 | 0.29359 | 0.29359 | 0.020685
2019-04-09 15:31 | 1554838235.7145 | 277 | 0.2086 | 0.001232 | 0.00017 | 100.8233 | 100.7424 | 111.7272 | 20.6898 | 0.004143 | 0.003677 | 129.7886 | 0.25348 | 8.4041 | 0.003223 | 0.005355 | 0.005268 | 4.9e-05 | 0.001147 | 0.29214 | 1.346 | 2.5476 | 98.44 | 19.98 | 48.702 | 121.2002 | 98.82 | 109.8259 | 0.011565 | 0.92185 | 0.021767 | 23.6124 | 0.007183 | 0.1508 | 0.004423 | 0.27912 | 0.27912 | 0.021772
2019-04-09 15:31 | 1554838236.7122 | 278 | 0.20861 | 0.001198 | 0.000168 | 100.7995 | 100.6105 | 111.7359 | 20.6902 | 0.004201 | 0.003615 | 129.7845 | 0.25322 | 8.4036 | 0.003137 | 0.005157 | 0.005071 | 4.9e-05 | 0.001141 | 0.29767 | 1.3507 | 2.5477 | 98.44 | 19.99 | 48.71 | 122.6006 | 96.7377 | 108.3134 | 0.011574 | 0.9255 | 0.021811 | 23.5666 | 0.01678 | 0.15064 | 0.004448 | 0.28149 | 0.28149 | 0.021567
2019-04-09 15:31 | 1554838237.7122 | 279 | 0.20864 | 0.00117 | 0.000166 | 100.8024 | 100.5245 | 112.3537 | 20.6901 | 0.004171 | 0.003635 | 129.7811 | 0.25361 | 8.4007 | 0.003236 | 0.004316 | 0.004292 | 5.2e-05 | 0.00114 | 0.27184 | 1.3613 | 2.5477 | 98.44 | 19.99 | 48.713 | 121.8754 | 99.1512 | 108.7964 | 0.011575 | 0.92433 | 0.021857 | 23.6466 | 0.024697 | 0.15087 | 0.004291 | 0.27191 | 0.27191 | 0.02236
2019-04-09 15:31 | 1554838238.7132 | 280 | 0.20864 | 0.001148 | 0.000163 | 100.8084 | 100.7551 | 112.4886 | 20.6911 | 0.004089 | 0.003571 | 129.7589 | 0.25402 | 8.4022 | 0.003233 | 0.005266 | 0.005157 | 5.1e-05 | 0.001139 | 0.2912 | 1.3488 | 2.5476 | 98.44 | 19.99 | 48.716 | 119.8733 | 99.0746 | 107.2533 | 0.011575 | 0.92808 | 0.021915 | 23.613 | 0.004732 | 0.15112 | 0.004319 | 0.27461 | 0.27461 | 0.022177
2019-04-09 15:31 | 1554838239.7133 | 281 | 0.20869 | 0.001127 | 0.00016 | 100.8432 | 100.466 | 112.141 | 20.6896 | 0.004056 | 0.003587 | 129.7308 | 0.25369 | 8.4033 | 0.003226 | 0.005302 | 0.005259 | 5.3e-05 | 0.001143 | 0.27721 | 1.327 | 2.5475 | 98.44 | 19.99 | 48.721 | 119.0758 | 98.8959 | 107.6528 | 0.011576 | 0.92711 | 0.021869 | 23.5889 | 0.033524 | 0.15092 | 0.004075 | 0.25815 | 0.25815 | 0.023561
2019-04-09 15:31 | 1554838240.7127 | 282 | 0.20872 | 0.001101 | 0.000159 | 100.8397 | 100.4732 | 112.3086 | 20.6905 | 0.004019 | 0.003591 | 129.7327 | 0.25333 | 8.4027 | 0.003207 | 0.004796 | 0.004707 | 4.9e-05 | 0.001146 | 0.26792 | 1.3698 | 2.5475 | 98.44 | 19.99 | 48.724 | 118.1723 | 98.4301 | 107.7372 | 0.011577 | 0.9269 | 0.021883 | 23.6091 | 0.032563 | 0.15071 | 0.003925 | 0.2486 | 0.2486 | 0.024431
2019-04-09 15:31 | 1554838241.7134 | 283 | 0.20875 | 0.001075 | 0.000157 | 100.856 | 100.7437 | 112.5428 | 20.6908 | 0.00399 | 0.003539 | 129.7279 | 0.25304 | 8.4011 | 0.003178 | 0.004242 | 0.004167 | 5.5e-05 | 0.001147 | 0.27632 | 1.3711 | 2.5476 | 98.44 | 19.99 | 48.731 | 117.4445 | 97.7484 | 106.4838 | 0.011579 | 0.92996 | 0.021942 | 23.5948 | 0.009967 | 0.15053 | 0.003755 | 0.23822 | 0.23822 | 0.025466
2019-04-09 15:31 | 1554838242.7124 | 284 | 0.20875 | 0.001057 | 0.000155 | 100.8403 | 100.7169 | 112.5668 | 20.6918 | 0.00402 | 0.003565 | 129.6855 | 0.25287 | 8.3987 | 0.003182 | 0.00439 | 0.004307 | 5e-05 | 0.001143 | 0.28512 | 1.3768 | 2.5477 | 98.44 | 19.99 | 48.733 | 118.1782 | 97.845 | 107.0982 | 0.011579 | 0.92846 | 0.021927 | 23.6164 | 0.010951 | 0.15043 | 0.003804 | 0.24136 | 0.24136 | 0.025118
2019-04-09 15:31 | 1554838243.7124 | 285 | 0.20875 | 0.001037 | 0.000152 | 100.77 | 100.5582 | 112.6114 | 20.6901 | 0.004001 | 0.003582 | 129.6863 | 0.2541 | 8.4021 | 0.003175 | 0.004336 | 0.004314 | 5.4e-05 | 0.001146 | 0.29645 | 1.3494 | 2.5475 | 98.44 | 19.99 | 48.734 | 117.7226 | 97.6718 | 107.5295 | 0.01158 | 0.92741 | 0.021919 | 23.6344 | 0.018816 | 0.15116 | 0.003813 | 0.24207 | 0.24207 | 0.025167
2019-04-09 15:31 | 1554838244.7131 | 286 | 0.20877 | 0.001018 | 0.000149 | 100.795 | 100.5741 | 112.6844 | 20.6911 | 0.004089 | 0.003514 | 129.677 | 0.25374 | 8.4014 | 0.003165 | 0.003989 | 0.003821 | 5.2e-05 | 0.001151 | 0.3027 | 1.3157 | 2.5476 | 98.44 | 19.99 | 48.738 | 119.8764 | 97.4298 | 105.8671 | 0.011581 | 0.93148 | 0.021974 | 23.5904 | 0.019628 | 0.15095 | 0.003732 | 0.23746 | 0.23746 | 0.025618
2019-04-09 15:31 | 1554838245.7125 | 287 | 0.2088 | 0.000998 | 0.000148 | 100.7617 | 100.497 | 113.1196 | 20.6903 | 0.003945 | 0.003549 | 129.6291 | 0.25339 | 8.403 | 0.003122 | 0.004735 | 0.004673 | 5.4e-05 | 0.00115 | 0.28241 | 1.3212 | 2.5475 | 98.44 | 19.99 | 48.74 | 116.3505 | 96.3807 | 106.7179 | 0.011581 | 0.92939 | 0.021992 | 23.6624 | 0.023526 | 0.15074 | 0.003567 | 0.22686 | 0.22686 | 0.026779
2019-04-09 15:31 | 1554838246.7125 | 288 | 0.20882 | 0.000976 | 0.000145 | 100.7905 | 100.6425 | 113.2261 | 20.6914 | 0.003887 | 0.003439 | 129.609 | 0.25328 | 8.4019 | 0.003198 | 0.004533 | 0.004482 | 5.1e-05 | 0.001152 | 0.26537 | 1.3184 | 2.5475 | 98.44 | 19.99 | 48.74 | 114.9337 | 98.2263 | 104.0498 | 0.011581 | 0.93596 | 0.02208 | 23.5903 | 0.013143 | 0.15068 | 0.003503 | 0.22371 | 0.22371 | 0.027144
2019-04-09 15:31 | 1554838247.7135 | 289 | 0.20883 | 0.000962 | 0.000144 | 100.7569 | 100.584 | 112.947 | 20.6893 | 0.003903 | 0.003465 | 129.5923 | 0.25384 | 8.4037 | 0.003204 | 0.004822 | 0.004682 | 5e-05 | 0.001153 | 0.24563 | 1.3836 | 2.5475 | 98.44 | 19.99 | 48.74 | 115.3255 | 98.3779 | 104.6846 | 0.011581 | 0.93439 | 0.022034 | 23.581 | 0.015357 | 0.15101 | 0.003398 | 0.21632 | 0.21632 | 0.028133
2019-04-09 15:31 | 1554838248.7126 | 290 | 0.20886 | 0.000945 | 0.000142 | 100.7899 | 100.5152 | 113.2745 | 20.6905 | 0.003864 | 0.003457 | 129.5816 | 0.25382 | 8.4026 | 0.003112 | 0.004302 | 0.004253 | 5.1e-05 | 0.001156 | 0.2756 | 1.3607 | 2.5476 | 98.44 | 19.99 | 48.738 | 114.3683 | 96.1506 | 104.4924 | 0.011581 | 0.93487 | 0.022071 | 23.6092 | 0.024414 | 0.15099 | 0.003265 | 0.20803 | 0.20803 | 0.029251
2019-04-09 15:31 | 1554838249.7126 | 291 | 0.20888 | 0.000927 | 0.00014 | 100.7351 | 100.4977 | 113.4423 | 20.69 | 0.003859 | 0.003518 | 129.5891 | 0.25375 | 8.4017 | 0.003117 | 0.003996 | 0.003903 | 4.9e-05 | 0.001156 | 0.28694 | 1.3684 | 2.5476 | 98.44 | 19.99 | 48.737 | 114.2516 | 96.2534 | 105.9621 | 0.01158 | 0.93124 | 0.022045 | 23.6726 | 0.021112 | 0.15096 | 0.003178 | 0.20217 | 0.20217 | 0.030091
2019-04-09 15:31 | 1554838250.7127 | 292 | 0.20888 | 0.000917 | 0.000138 | 100.7262 | 100.4932 | 113.4634 | 20.6911 | 0.00385 | 0.003444 | 129.5956 | 0.25382 | 8.4056 | 0.003162 | 0.004907 | 0.0048 | 5.1e-05 | 0.001157 | 0.26073 | 1.3515 | 2.5476 | 98.44 | 19.99 | 48.736 | 114.047 | 97.3544 | 104.176 | 0.01158 | 0.93565 | 0.022099 | 23.619 | 0.020717 | 0.151 | 0.00321 | 0.20494 | 0.20494 | 0.029693
2019-04-09 15:31 | 1554838251.7129 | 293 | 0.20899 | 0.000902 | 0.000136 | 100.6756 | 100.5156 | 113.4278 | 20.6904 | 0.003843 | 0.003485 | 129.5751 | 0.25444 | 8.4034 | 0.003196 | 0.005318 | 0.005138 | 4.9e-05 | 0.001154 | 0.29192 | 1.3473 | 2.5475 | 98.44 | 19.99 | 48.736 | 113.8575 | 98.1768 | 105.17 | 0.01158 | 0.93319 | 0.022067 | 23.6463 | 0.014229 | 0.15137 | 0.002529 | 0.15968 | 0.15968 | 0.038202
2019-04-09 15:31 | 1554838252.7127 | 294 | 0.209 | 0.000889 | 0.000133 | 100.6062 | 100.5162 | 113.293 | 20.6902 | 0.003863 | 0.003465 | 129.5948 | 0.25386 | 8.4017 | 0.00314 | 0.004573 | 0.004541 | 5.2e-05 | 0.001153 | 0.24762 | 1.3388 | 2.5475 | 98.44 | 19.99 | 48.736 | 114.3578 | 96.8216 | 104.6902 | 0.01158 | 0.93438 | 0.022067 | 23.6173 | 0.008001 | 0.15102 | 0.002506 | 0.15831 | 0.15831 | 0.038446
2019-04-09 15:31 | 1554838253.7133 | 295 | 0.20909 | 0.000878 | 0.000133 | 100.5645 | 100.3208 | 113.4686 | 20.6894 | 0.003788 | 0.003413 | 129.5995 | 0.25361 | 8.4002 | 0.003171 | 0.004038 | 0.004008 | 5e-05 | 0.001152 | 0.28495 | 1.384 | 2.5476 | 98.44 | 19.99 | 48.735 | 112.5292 | 97.5598 | 103.4165 | 0.01158 | 0.93754 | 0.022122 | 23.5957 | 0.021706 | 0.15087 | 0.001979 | 0.12383 | 0.12383 | 0.049102
2019-04-09 15:31 | 1554838254.7128 | 296 | 0.20892 | 0.000863 | 0.000131 | 100.6059 | 100.5004 | 113.5444 | 20.6905 | 0.003769 | 0.003332 | 129.5708 | 0.25354 | 8.4014 | 0.003175 | 0.004229 | 0.004176 | 4.7e-05 | 0.001148 | 0.25857 | 1.3664 | 2.5476 | 98.44 | 20 | 48.734 | 112.0636 | 97.6705 | 101.457 | 0.011587 | 0.94244 | 0.022187 | 23.5421 | 0.009384 | 0.15083 | 0.003049 | 0.19539 | 0.19539 | 0.031108
2019-04-09 15:31 | 1554838255.7128 | 297 | 0.20897 | 0.000856 | 0.00013 | 100.5377 | 100.3428 | 112.9792 | 20.6896 | 0.003846 | 0.003365 | 129.5876 | 0.25339 | 8.4029 | 0.003124 | 0.004504 | 0.004408 | 4.9e-05 | 0.001143 | 0.26178 | 1.3947 | 2.5477 | 98.44 | 20 | 48.734 | 113.942 | 96.4341 | 102.251 | 0.011587 | 0.94045 | 0.022108 | 23.5083 | 0.017358 | 0.15074 | 0.002769 | 0.17622 | 0.17622 | 0.034474
2019-04-09 15:31 | 1554838256.7127 | 298 | 0.20893 | 0.00084 | 0.000128 | 100.5178 | 100.2781 | 112.9975 | 20.6895 | 0.003802 | 0.003402 | 129.6031 | 0.25404 | 8.4015 | 0.003114 | 0.004238 | 0.004237 | 5.2e-05 | 0.001143 | 0.27544 | 1.3556 | 2.5476 | 98.44 | 20 | 48.731 | 112.8688 | 96.1904 | 103.1557 | 0.011586 | 0.93819 | 0.022084 | 23.5385 | 0.021352 | 0.15113 | 0.002999 | 0.19132 | 0.19132 | 0.031835
2019-04-09 15:31 | 1554838257.7128 | 299 | 0.20899 | 0.000833 | 0.000127 | 100.5005 | 100.4626 | 113.2198 | 20.6898 | 0.00374 | 0.003402 | 129.5923 | 0.25408 | 8.4009 | 0.003087 | 0.004072 | 0.003953 | 4.6e-05 | 0.001141 | 0.27975 | 1.3351 | 2.5476 | 98.44 | 20 | 48.73 | 111.3564 | 95.5438 | 103.1609 | 0.011586 | 0.93817 | 0.022105 | 23.5618 | 0.003376 | 0.15115 | 0.002677 | 0.17026 | 0.17026 | 0.035777
2019-04-09 15:31 | 1554838258.7129 | 300 | 0.20896 | 0.00082 | 0.000126 | 100.4317 | 100.1542 | 113.2048 | 20.6901 | 0.003716 | 0.003336 | 129.616 | 0.25354 | 8.4019 | 0.003073 | 0.004648 | 0.004576 | 5.2e-05 | 0.001142 | 0.27992 | 1.3528 | 2.5477 | 98.44 | 20 | 48.729 | 110.7877 | 95.2128 | 101.55 | 0.011586 | 0.94221 | 0.022151 | 23.5098 | 0.024749 | 0.15083 | 0.002849 | 0.18207 | 0.18207 | 0.033384
2019-04-09 15:31 | 1554838259.712 | 301 | 0.209 | 0.000812 | 0.000124 | 100.331 | 100.3862 | 113.1081 | 20.6895 | 0.003711 | 0.003404 | 129.6101 | 0.25389 | 8.4031 | 0.003098 | 0.004254 | 0.004152 | 4.9e-05 | 0.001143 | 0.27417 | 1.3883 | 2.5476 | 98.44 | 20 | 48.727 | 110.6536 | 95.8042 | 103.2012 | 0.011585 | 0.93807 | 0.022093 | 23.5515 | -0.004912 | 0.15104 | 0.002636 | 0.16763 | 0.16763 | 0.036313
2019-04-09 15:31 | 1554838260.7119 | 302 | 0.20902 | 0.00081 | 0.000123 | 100.2899 | 100.223 | 113.1679 | 20.6903 | 0.003723 | 0.003401 | 129.5827 | 0.25482 | 8.4017 | 0.003033 | 0.004731 | 0.004579 | 4.8e-05 | 0.001142 | 0.26178 | 1.3923 | 2.5476 | 98.44 | 20 | 48.725 | 110.9395 | 94.2337 | 103.1212 | 0.011585 | 0.93827 | 0.022101 | 23.5552 | 0.005971 | 0.1516 | 0.002529 | 0.1607 | 0.1607 | 0.038017
2019-04-09 15:31 | 1554838261.713 | 303 | 0.209 | 0.000796 | 0.000121 | 100.2301 | 100.0978 | 113.5019 | 20.6893 | 0.003682 | 0.003337 | 129.5908 | 0.25352 | 8.4005 | 0.003099 | 0.004016 | 0.004057 | 4.7e-05 | 0.001147 | 0.28799 | 1.3867 | 2.5477 | 98.44 | 20 | 48.722 | 109.9477 | 95.8306 | 101.5909 | 0.011584 | 0.9421 | 0.022179 | 23.5419 | 0.011818 | 0.15082 | 0.002616 | 0.1671 | 0.1671 | 0.036373
2019-04-09 15:31 | 1554838262.713 | 304 | 0.20901 | 0.000789 | 0.00012 | 100.1923 | 100.0368 | 113.5775 | 20.6897 | 0.003709 | 0.003307 | 129.5928 | 0.25329 | 8.4024 | 0.003122 | 0.004516 | 0.00436 | 5e-05 | 0.001148 | 0.26073 | 1.3956 | 2.5477 | 98.44 | 20 | 48.72 | 110.6035 | 96.3794 | 100.865 | 0.011583 | 0.94393 | 0.022208 | 23.5269 | 0.013887 | 0.15068 | 0.002605 | 0.16665 | 0.16665 | 0.036438
2019-04-09 15:31 | 1554838263.712 | 305 | 0.20902 | 0.000782 | 0.000118 | 100.1661 | 99.9012 | 113.4101 | 20.6905 | 0.003644 | 0.003331 | 129.5661 | 0.25387 | 8.4042 | 0.003058 | 0.004445 | 0.00435 | 5.1e-05 | 0.001146 | 0.2557 | 1.3484 | 2.5475 | 98.44 | 20 | 48.712 | 109.0317 | 94.8336 | 101.4318 | 0.011582 | 0.94251 | 0.022175 | 23.5274 | 0.023679 | 0.15103 | 0.002544 | 0.16243 | 0.16243 | 0.037469
2019-04-09 15:31 | 1554838264.7121 | 306 | 0.20904 | 0.000773 | 0.000117 | 100.0612 | 99.8217 | 113.4794 | 20.6896 | 0.003695 | 0.003328 | 129.5785 | 0.25325 | 8.401 | 0.003072 | 0.004162 | 0.004088 | 5.3e-05 | 0.001145 | 0.27505 | 1.3946 | 2.5476 | 98.44 | 20 | 48.706 | 110.2734 | 95.1888 | 101.3588 | 0.01158 | 0.94269 | 0.022184 | 23.5323 | 0.021435 | 0.15066 | 0.002412 | 0.15382 | 0.15382 | 0.039471
2019-04-09 15:31 | 1554838265.7121 | 307 | 0.20905 | 0.000767 | 0.000115 | 100.0097 | 99.9343 | 113.8474 | 20.6888 | 0.0036 | 0.003274 | 129.5729 | 0.25432 | 8.4031 | 0.00304 | 0.004579 | 0.004548 | 4.8e-05 | 0.001148 | 0.27759 | 1.3572 | 2.5477 | 98.44 | 20 | 48.701 | 107.9568 | 94.4163 | 100.059 | 0.011579 | 0.94597 | 0.022258 | 23.5295 | 0.00675 | 0.15129 | 0.002359 | 0.15091 | 0.15091 | 0.040402
2019-04-09 15:31 | 1554838266.7134 | 308 | 0.20905 | 0.000758 | 0.000114 | 99.943 | 99.7553 | 113.9641 | 20.6894 | 0.003594 | 0.003318 | 129.5348 | 0.25435 | 8.4012 | 0.003037 | 0.004315 | 0.00432 | 5.2e-05 | 0.001147 | 0.27704 | 1.3886 | 2.5476 | 98.44 | 20 | 48.69 | 107.8071 | 94.3245 | 101.1239 | 0.011576 | 0.94328 | 0.022238 | 23.5751 | 0.016814 | 0.15132 | 0.002428 | 0.15545 | 0.15545 | 0.039229
2019-04-09 15:31 | 1554838267.7122 | 309 | 0.20903 | 0.000758 | 0.000113 | 99.9427 | 99.6263 | 113.7388 | 20.6895 | 0.003607 | 0.003304 | 129.5353 | 0.25467 | 8.3996 | 0.003079 | 0.004574 | 0.004421 | 4.9e-05 | 0.001144 | 0.2855 | 1.3538 | 2.5476 | 98.44 | 20 | 48.685 | 108.1333 | 95.3486 | 100.7817 | 0.011575 | 0.94414 | 0.022226 | 23.541 | 0.028367 | 0.1515 | 0.002541 | 0.16293 | 0.16293 | 0.037474
2019-04-09 15:31 | 1554838268.7114 | 310 | 0.20903 | 0.00075 | 0.000112 | 99.8406 | 99.6985 | 113.8723 | 20.6896 | 0.003604 | 0.003284 | 129.5164 | 0.25388 | 8.4032 | 0.003053 | 0.004529 | 0.004446 | 5e-05 | 0.001142 | 0.28495 | 1.3599 | 2.5476 | 98.44 | 20 | 48.679 | 108.0435 | 94.7302 | 100.2891 | 0.011574 | 0.94539 | 0.022254 | 23.5393 | 0.012735 | 0.15103 | 0.002538 | 0.16301 | 0.16301 | 0.037338
2019-04-09 15:31 | 1554838269.7123 | 311 | 0.20911 | 0.000745 | 0.00011 | 99.773 | 99.5275 | 114.1366 | 20.6894 | 0.003569 | 0.003285 | 129.4805 | 0.25371 | 8.4031 | 0.003077 | 0.004328 | 0.004208 | 5e-05 | 0.001142 | 0.30613 | 1.378 | 2.5476 | 98.44 | 20 | 48.667 | 107.2085 | 95.3101 | 100.3201 | 0.011571 | 0.94531 | 0.022279 | 23.5676 | 0.022035 | 0.15093 | 0.002042 | 0.13014 | 0.13014 | 0.04674
2019-04-09 15:31 | 1554838270.7113 | 312 | 0.20905 | 0.000736 | 0.00011 | 99.753 | 99.6128 | 114.0167 | 20.6893 | 0.003577 | 0.00324 | 129.4869 | 0.25357 | 8.4009 | 0.003041 | 0.004819 | 0.004728 | 4.9e-05 | 0.001141 | 0.26792 | 1.3497 | 2.5476 | 98.44 | 20 | 48.661 | 107.3989 | 94.4192 | 99.228 | 0.011569 | 0.94808 | 0.0223 | 23.5207 | 0.012579 | 0.15085 | 0.002425 | 0.1559 | 0.1559 | 0.038994
2019-04-09 15:31 | 1554838271.7127 | 313 | 0.20908 | 0.000726 | 0.000109 | 99.6446 | 99.521 | 113.6327 | 20.6899 | 0.003571 | 0.003234 | 129.4671 | 0.25391 | 8.3999 | 0.003032 | 0.004504 | 0.004489 | 5.4e-05 | 0.001137 | 0.26145 | 1.3472 | 2.5474 | 98.44 | 20.01 | 48.652 | 107.245 | 94.2209 | 99.0944 | 0.011574 | 0.94842 | 0.022266 | 23.4769 | 0.011109 | 0.15105 | 0.002269 | 0.1453 | 0.1453 | 0.041894
2019-04-09 15:31 | 1554838272.7118 | 314 | 0.20911 | 0.00073 | 0.000108 | 99.6685 | 99.5316 | 113.7071 | 20.689 | 0.003548 | 0.003212 | 129.4426 | 0.25428 | 8.4021 | 0.003016 | 0.004502 | 0.004418 | 5.4e-05 | 0.001134 | 0.26665 | 1.3677 | 2.5477 | 98.44 | 20.01 | 48.646 | 106.6843 | 93.8365 | 98.5564 | 0.011573 | 0.9498 | 0.022289 | 23.4676 | 0.012288 | 0.15127 | 0.002061 | 0.13163 | 0.13163 | 0.046313
2019-04-09 15:31 | 1554838273.7128 | 315 | 0.20906 | 0.000724 | 0.000107 | 99.63 | 99.5681 | 113.87 | 20.6902 | 0.003565 | 0.003275 | 129.4371 | 0.25434 | 8.401 | 0.003056 | 0.003981 | 0.003885 | 5e-05 | 0.001134 | 0.27383 | 1.3276 | 2.5476 | 98.44 | 20.01 | 48.641 | 107.1124 | 94.7838 | 100.0746 | 0.011572 | 0.94593 | 0.02226 | 23.5323 | 0.005563 | 0.1513 | 0.002396 | 0.15386 | 0.15386 | 0.03963
2019-04-09 15:31 | 1554838274.7118 | 316 | 0.20909 | 0.000715 | 0.000106 | 99.5467 | 99.3475 | 113.8038 | 20.69 | 0.003529 | 0.003211 | 129.4166 | 0.25436 | 8.4014 | 0.002962 | 0.004329 | 0.004227 | 5.1e-05 | 0.001134 | 0.27129 | 1.3449 | 2.5476 | 98.44 | 20.01 | 48.63 | 106.2335 | 92.5145 | 98.5268 | 0.011569 | 0.94987 | 0.0223 | 23.4766 | 0.017913 | 0.15132 | 0.002241 | 0.14377 | 0.14377 | 0.042417
2019-04-09 15:31 | 1554838275.7124 | 317 | 0.2091 | 0.000715 | 0.000105 | 99.4951 | 99.3785 | 113.819 | 20.6902 | 0.003501 | 0.003274 | 129.3776 | 0.25482 | 8.4034 | 0.002987 | 0.004752 | 0.004634 | 4.8e-05 | 0.001131 | 0.29463 | 1.4083 | 2.5476 | 98.44 | 20.01 | 48.625 | 105.5554 | 93.1203 | 100.0495 | 0.011568 | 0.946 | 0.022256 | 23.5262 | 0.010495 | 0.15159 | 0.002155 | 0.13781 | 0.13781 | 0.044329
2019-04-09 15:31 | 1554838276.7125 | 318 | 0.20903 | 0.000711 | 0.000104 | 99.5562 | 99.4834 | 113.8291 | 20.6898 | 0.003508 | 0.003224 | 129.3843 | 0.25431 | 8.4017 | 0.003089 | 0.004528 | 0.00445 | 5.1e-05 | 0.001127 | 0.26344 | 1.3627 | 2.5476 | 98.44 | 20.01 | 48.612 | 105.7345 | 95.5829 | 98.8403 | 0.011565 | 0.94907 | 0.022293 | 23.4891 | 0.006538 | 0.15129 | 0.002613 | 0.16873 | 0.16873 | 0.036135
2019-04-09 15:31 | 1554838277.7115 | 319 | 0.2091 | 0.000706 | 0.000103 | 99.4651 | 99.193 | 114.0211 | 20.6888 | 0.00346 | 0.003198 | 129.3807 | 0.25374 | 8.4041 | 0.00309 | 0.004921 | 0.004818 | 4.9e-05 | 0.001126 | 0.26792 | 1.3449 | 2.5476 | 98.44 | 20.01 | 48.604 | 104.5661 | 95.6235 | 98.2238 | 0.011563 | 0.95065 | 0.02233 | 23.4894 | 0.024502 | 0.15095 | 0.002207 | 0.14192 | 0.14192 | 0.042865
2019-04-09 15:31 | 1554838278.7129 | 320 | 0.20913 | 0.000701 | 0.000101 | 99.4183 | 99.2966 | 113.919 | 20.6897 | 0.003441 | 0.00319 | 129.3581 | 0.25371 | 8.4023 | 0.003054 | 0.00451 | 0.004456 | 5.2e-05 | 0.001122 | 0.26488 | 1.4007 | 2.5476 | 98.44 | 20.01 | 48.599 | 104.0976 | 94.7522 | 98.0224 | 0.011562 | 0.95116 | 0.022326 | 23.4726 | 0.01096 | 0.15093 | 0.002032 | 0.13027 | 0.13027 | 0.046693
2019-04-09 15:31 | 1554838279.713 | 321 | 0.20915 | 0.000696 | 0.0001 | 99.3308 | 99.2036 | 113.7611 | 20.6896 | 0.003465 | 0.003184 | 129.3587 | 0.25393 | 8.3984 | 0.003066 | 0.004709 | 0.004599 | 5.2e-05 | 0.001123 | 0.28423 | 1.3753 | 2.5476 | 98.44 | 20.01 | 48.586 | 104.6888 | 95.026 | 97.8789 | 0.011559 | 0.95153 | 0.022315 | 23.4517 | 0.011454 | 0.15107 | 0.001906 | 0.12185 | 0.12185 | 0.049962
2019-04-09 15:31 | 1554838280.7122 | 322 | 0.2091 | 0.000697 | 9.9e-05 | 99.3873 | 99.0257 | 113.8066 | 20.6894 | 0.003419 | 0.003179 | 129.3453 | 0.25382 | 8.3996 | 0.003053 | 0.004545 | 0.004463 | 4.6e-05 | 0.001118 | 0.26322 | 1.355 | 2.5476 | 98.44 | 20.01 | 48.579 | 103.5737 | 94.7283 | 97.7675 | 0.011557 | 0.95182 | 0.022323 | 23.4529 | 0.032605 | 0.15099 | 0.002172 | 0.13971 | 0.13971 | 0.043554
2019-04-09 15:31 | 1554838281.7126 | 323 | 0.20911 | 0.000697 | 9.8e-05 | 99.3501 | 99.2804 | 114.0974 | 20.6896 | 0.003429 | 0.003187 | 129.3256 | 0.25352 | 8.4021 | 0.003051 | 0.004186 | 0.004145 | 5e-05 | 0.001121 | 0.24851 | 1.3578 | 2.5476 | 98.43 | 20.01 | 48.571 | 103.8215 | 94.6619 | 97.9619 | 0.011556 | 0.95132 | 0.022346 | 23.489 | 0.00628 | 0.15082 | 0.002156 | 0.13885 | 0.13885 | 0.043775
2019-04-09 15:31 | 1554838282.712 | 324 | 0.20914 | 0.000687 | 9.7e-05 | 99.2468 | 99.0209 | 114.2702 | 20.6896 | 0.003456 | 0.003173 | 129.3172 | 0.25365 | 8.4013 | 0.00305 | 0.004504 | 0.004446 | 5.1e-05 | 0.001125 | 0.28385 | 1.3114 | 2.5476 | 98.43 | 20.01 | 48.557 | 104.4632 | 94.6417 | 97.6141 | 0.011553 | 0.95221 | 0.022373 | 23.4958 | 0.020381 | 0.15089 | 0.001953 | 0.12549 | 0.12549 | 0.048459
2019-04-09 15:31 | 1554838283.7117 | 325 | 0.20908 | 0.000677 | 9.7e-05 | 99.2956 | 99.0381 | 113.9615 | 20.6901 | 0.003459 | 0.003199 | 129.3252 | 0.25413 | 8.4002 | 0.003041 | 0.004506 | 0.00452 | 5.1e-05 | 0.001126 | 0.27959 | 1.3107 | 2.5476 | 98.43 | 20.01 | 48.549 | 104.5264 | 94.4385 | 98.2446 | 0.011551 | 0.95059 | 0.022324 | 23.484 | 0.023227 | 0.15118 | 0.002317 | 0.14961 | 0.14961 | 0.040724
2019-04-09 15:31 | 1554838284.7121 | 326 | 0.20914 | 0.000679 | 9.5e-05 | 99.2748 | 99.0323 | 113.9507 | 20.6897 | 0.003399 | 0.003141 | 129.332 | 0.25378 | 8.4019 | 0.003134 | 0.00488 | 0.004659 | 5.1e-05 | 0.001129 | 0.2724 | 1.3485 | 2.5478 | 98.43 | 20.01 | 48.533 | 103.0844 | 96.6747 | 96.8328 | 0.011547 | 0.95422 | 0.022365 | 23.4382 | 0.021878 | 0.15098 | 0.001966 | 0.12639 | 0.12639 | 0.04814
2019-04-09 15:31 | 1554838285.7108 | 327 | 0.20913 | 0.000674 | 9.5e-05 | 99.2379 | 99.1296 | 114.0216 | 20.6894 | 0.003426 | 0.003139 | 129.3527 | 0.25423 | 8.3997 | 0.003095 | 0.004577 | 0.00448 | 4.9e-05 | 0.00113 | 0.25984 | 1.3533 | 2.5475 | 98.43 | 20.01 | 48.526 | 103.7365 | 95.7242 | 96.7886 | 0.011546 | 0.95434 | 0.022373 | 23.4441 | 0.009758 | 0.15124 | 0.002017 | 0.12993 | 0.12993 | 0.046912
2019-04-09 15:31 | 1554838286.7118 | 328 | 0.20917 | 0.000674 | 9.4e-05 | 99.2718 | 99.0212 | 113.9167 | 20.6891 | 0.003423 | 0.0032 | 129.3555 | 0.25337 | 8.4027 | 0.00307 | 0.005187 | 0.005039 | 4.9e-05 | 0.001137 | 0.28119 | 1.3321 | 2.5476 | 98.43 | 20.01 | 48.513 | 103.6616 | 95.1218 | 98.265 | 0.011543 | 0.95054 | 0.022319 | 23.48 | 0.022604 | 0.15073 | 0.001791 | 0.11456 | 0.11456 | 0.053025
2019-04-09 15:31 | 1554838287.7118 | 329 | 0.20917 | 0.000672 | 9.2e-05 | 99.2563 | 99.1324 | 113.778 | 20.6898 | 0.003425 | 0.003133 | 129.3464 | 0.25447 | 8.4017 | 0.003054 | 0.00471 | 0.004651 | 5e-05 | 0.001138 | 0.29839 | 1.3294 | 2.5477 | 98.43 | 20.02 | 48.507 | 103.7243 | 94.7366 | 96.641 | 0.011548 | 0.95472 | 0.022354 | 23.4143 | 0.011173 | 0.15138 | 0.001805 | 0.11572 | 0.11572 | 0.05272
2019-04-09 15:31 | 1554838288.7122 | 330 | 0.20917 | 0.000664 | 9.1e-05 | 99.1667 | 99.1727 | 113.5892 | 20.6901 | 0.00335 | 0.003134 | 129.3447 | 0.25362 | 8.4047 | 0.003072 | 0.00478 | 0.004668 | 4.8e-05 | 0.001141 | 0.2891 | 1.3837 | 2.5476 | 98.43 | 20.02 | 48.502 | 101.898 | 95.18 | 96.671 | 0.011547 | 0.95464 | 0.022335 | 23.3959 | -0.00054 | 0.15088 | 0.001798 | 0.11521 | 0.11521 | 0.052776
2019-04-09 15:31 | 1554838289.712 | 331 | 0.2092 | 0.000667 | 9.1e-05 | 99.1941 | 99.1272 | 113.8514 | 20.6894 | 0.003404 | 0.003166 | 129.3524 | 0.2538 | 8.4033 | 0.003106 | 0.00458 | 0.004431 | 4.8e-05 | 0.00114 | 0.27721 | 1.3496 | 2.5477 | 98.43 | 20.02 | 48.493 | 103.2021 | 96.0068 | 97.4578 | 0.011545 | 0.95261 | 0.022337 | 23.4477 | 0.006029 | 0.15098 | 0.001641 | 0.10479 | 0.10479 | 0.058067
2019-04-09 15:32 | 1554838290.7118 | 332 | 0.20921 | 0.000658 | 9e-05 | 99.2352 | 99.1231 | 113.878 | 20.6882 | 0.003317 | 0.003134 | 129.3236 | 0.25352 | 8.4022 | 0.002988 | 0.004843 | 0.004795 | 5e-05 | 0.001135 | 0.30165 | 1.3344 | 2.5477 | 98.43 | 20.02 | 48.486 | 101.092 | 93.1449 | 96.6669 | 0.011543 | 0.95465 | 0.022363 | 23.4254 | 0.010106 | 0.15082 | 0.001575 | 0.10049 | 0.10049 | 0.060486
2019-04-09 15:32 | 1554838291.711 | 333 | 0.2092 | 0.000662 | 9e-05 | 99.2209 | 99.0539 | 113.5336 | 20.6894 | 0.003379 | 0.003177 | 129.3221 | 0.2532 | 8.4008 | 0.003082 | 0.004515 | 0.004519 | 5.1e-05 | 0.001135 | 0.25769 | 1.3225 | 2.5475 | 98.43 | 20.02 | 48.482 | 102.5892 | 95.4266 | 97.7252 | 0.011542 | 0.95193 | 0.022297 | 23.4234 | 0.015062 | 0.15063 | 0.001665 | 0.10625 | 0.10625 | 0.057131
2019-04-09 15:32 | 1554838292.7113 | 334 | 0.20919 | 0.000658 | 8.9e-05 | 99.1857 | 99.2866 | 113.2482 | 20.6895 | 0.003395 | 0.003125 | 129.3182 | 0.25316 | 8.4019 | 0.003077 | 0.004357 | 0.004271 | 4.8e-05 | 0.001132 | 0.28335 | 1.3111 | 2.5476 | 98.43 | 20.02 | 48.476 | 102.9964 | 95.2937 | 96.4661 | 0.011541 | 0.95517 | 0.022307 | 23.3542 | -0.00909 | 0.15061 | 0.001749 | 0.11193 | 0.11193 | 0.054228
2019-04-09 15:32 | 1554838293.7111 | 335 | 0.20919 | 0.000652 | 8.9e-05 | 99.2906 | 99.1173 | 113.3972 | 20.6882 | 0.003287 | 0.003051 | 129.2901 | 0.25323 | 8.4011 | 0.003027 | 0.004576 | 0.004532 | 5.1e-05 | 0.001134 | 0.27007 | 1.3193 | 2.5475 | 98.43 | 20.02 | 48.472 | 100.3684 | 94.1025 | 94.6803 | 0.01154 | 0.95981 | 0.022376 | 23.3131 | 0.015625 | 0.15065 | 0.001736 | 0.1114 | 0.1114 | 0.054497
2019-04-09 15:32 | 1554838294.7114 | 336 | 0.20919 | 0.00065 | 8.8e-05 | 99.2971 | 98.9882 | 113.5031 | 20.6885 | 0.003341 | 0.003064 | 129.2881 | 0.25374 | 8.4004 | 0.003064 | 0.004295 | 0.004184 | 5e-05 | 0.001138 | 0.29214 | 1.342 | 2.5475 | 98.43 | 20.02 | 48.466 | 101.6677 | 94.9928 | 94.9952 | 0.011539 | 0.95898 | 0.022377 | 23.3339 | 0.027869 | 0.15095 | 0.001705 | 0.10941 | 0.10941 | 0.055599
2019-04-09 15:32 | 1554838295.7121 | 337 | 0.2092 | 0.000644 | 8.8e-05 | 99.3031 | 99.0904 | 113.416 | 20.6897 | 0.003304 | 0.003078 | 129.3011 | 0.25349 | 8.4029 | 0.003035 | 0.004757 | 0.004685 | 5.3e-05 | 0.001144 | 0.27831 | 1.3065 | 2.5475 | 98.43 | 20.01 | 48.463 | 100.7818 | 94.2933 | 95.3212 | 0.011531 | 0.95814 | 0.022358 | 23.3353 | 0.019175 | 0.1508 | 0.00168 | 0.10764 | 0.10764 | 0.05646
2019-04-09 15:32 | 1554838296.7121 | 338 | 0.20919 | 0.000642 | 8.7e-05 | 99.3135 | 99.1407 | 113.4571 | 20.689 | 0.003337 | 0.003133 | 129.2502 | 0.25384 | 8.3989 | 0.003019 | 0.0043 | 0.004279 | 5.3e-05 | 0.001148 | 0.28335 | 1.3161 | 2.5474 | 98.43 | 20.01 | 48.461 | 101.5797 | 93.8889 | 96.6395 | 0.01153 | 0.95472 | 0.022323 | 23.3812 | 0.015579 | 0.15101 | 0.001755 | 0.11249 | 0.11249 | 0.054098
2019-04-09 15:32 | 1554838297.7115 | 339 | 0.20919 | 0.00064 | 8.7e-05 | 99.2941 | 99.1142 | 113.7721 | 20.6889 | 0.003292 | 0.003051 | 129.2487 | 0.25405 | 8.4021 | 0.002983 | 0.004483 | 0.004437 | 4.9e-05 | 0.001152 | 0.27151 | 1.3593 | 2.5476 | 98.43 | 20.01 | 48.456 | 100.4829 | 93.0345 | 94.6692 | 0.011529 | 0.95983 | 0.022413 | 23.3512 | 0.016227 | 0.15113 | 0.001722 | 0.11079 | 0.11079 | 0.054973
2019-04-09 15:32 | 1554838298.7116 | 340 | 0.2092 | 0.000637 | 8.7e-05 | 99.3924 | 99.2198 | 113.8582 | 20.6892 | 0.003263 | 0.003086 | 129.257 | 0.25418 | 8.4027 | 0.002955 | 0.004845 | 0.004672 | 4.7e-05 | 0.001153 | 0.27041 | 1.4061 | 2.5478 | 98.43 | 20.01 | 48.451 | 99.7864 | 92.352 | 95.5102 | 0.011528 | 0.95764 | 0.022396 | 23.3867 | 0.015545 | 0.15121 | 0.001672 | 0.10731 | 0.10731 | 0.056785
2019-04-09 15:32 | 1554838299.7102 | 341 | 0.2092 | 0.000632 | 8.6e-05 | 99.4472 | 99.2683 | 113.5385 | 20.6888 | 0.003316 | 0.003093 | 129.2429 | 0.25264 | 8.4029 | 0.003098 | 0.004489 | 0.004408 | 4.8e-05 | 0.001154 | 0.27273 | 1.3822 | 2.5477 | 98.43 | 20.01 | 48.446 | 101.0667 | 95.8011 | 95.6758 | 0.011527 | 0.95721 | 0.02236 | 23.3591 | 0.016101 | 0.15029 | 0.001704 | 0.10937 | 0.10937 | 0.055382
2019-04-09 15:32 | 1554838300.7113 | 342 | 0.20925 | 0.00063 | 8.6e-05 | 99.4246 | 99.2745 | 113.6507 | 20.6888 | 0.003285 | 0.003108 | 129.2371 | 0.25382 | 8.4016 | 0.00291 | 0.00455 | 0.004412 | 4.9e-05 | 0.001153 | 0.26753 | 1.3515 | 2.5477 | 98.43 | 20.01 | 48.445 | 100.3274 | 91.2828 | 96.0446 | 0.011526 | 0.95626 | 0.02236 | 23.3824 | 0.013507 | 0.15099 | 0.00137 | 0.086997 | 0.086997 | 0.069946
2019-04-09 15:32 | 1554838301.7113 | 343 | 0.20925 | 0.000627 | 8.5e-05 | 99.4847 | 99.5003 | 113.8838 | 20.6888 | 0.003287 | 0.003107 | 129.2633 | 0.25384 | 8.4021 | 0.002999 | 0.004623 | 0.00449 | 4.9e-05 | 0.001156 | 0.26825 | 1.3786 | 2.5476 | 98.43 | 20.01 | 48.443 | 100.3761 | 93.4178 | 96.0261 | 0.011526 | 0.95631 | 0.022383 | 23.4057 | -0.001401 | 0.15101 | 0.001401 | 0.089188 | 0.089188 | 0.068235
2019-04-09 15:32 | 1554838302.7104 | 344 | 0.20922 | 0.000626 | 8.5e-05 | 99.4811 | 99.3781 | 113.7949 | 20.6893 | 0.003274 | 0.003083 | 129.2571 | 0.25341 | 8.4043 | 0.003058 | 0.004645 | 0.004554 | 4.7e-05 | 0.001162 | 0.26056 | 1.3849 | 2.5476 | 98.43 | 20.01 | 48.442 | 100.0573 | 94.8426 | 95.4439 | 0.011526 | 0.95782 | 0.022392 | 23.3781 | 0.009267 | 0.15076 | 0.001594 | 0.10223 | 0.10223 | 0.05943
2019-04-09 15:32 | 1554838303.7107 | 345 | 0.20924 | 0.000628 | 8.4e-05 | 99.4889 | 99.2921 | 114.1451 | 20.6898 | 0.003269 | 0.003095 | 129.2517 | 0.25326 | 8.4028 | 0.003055 | 0.004562 | 0.004538 | 4.7e-05 | 0.001162 | 0.26792 | 1.3488 | 2.5476 | 98.43 | 20.02 | 48.44 | 99.94 | 94.76 | 95.7449 | 0.011532 | 0.95703 | 0.022417 | 23.4236 | 0.017711 | 0.15067 | 0.001468 | 0.093896 | 0.093896 | 0.064665
2019-04-09 15:32 | 1554838304.7108 | 346 | 0.20924 | 0.000621 | 8.4e-05 | 99.5252 | 99.4232 | 114.1695 | 20.6897 | 0.00331 | 0.003063 | 129.2387 | 0.25351 | 8.4036 | 0.003099 | 0.004019 | 0.004024 | 5e-05 | 0.00117 | 0.29717 | 1.3456 | 2.5476 | 98.43 | 20.02 | 48.439 | 100.9395 | 95.835 | 94.9556 | 0.011532 | 0.95909 | 0.022444 | 23.4011 | 0.009172 | 0.15081 | 0.00148 | 0.094802 | 0.094802 | 0.06411
2019-04-09 15:32 | 1554838305.7117 | 347 | 0.20922 | 0.000619 | 8.3e-05 | 99.5496 | 99.5075 | 114.1615 | 20.6887 | 0.003244 | 0.003115 | 129.2099 | 0.25373 | 8.4015 | 0.003029 | 0.004598 | 0.004473 | 4.8e-05 | 0.001166 | 0.29286 | 1.3456 | 2.5476 | 98.43 | 20.02 | 48.438 | 99.3412 | 94.1375 | 96.2113 | 0.011532 | 0.95583 | 0.022405 | 23.4401 | 0.003784 | 0.15094 | 0.001621 | 0.10417 | 0.10417 | 0.058398
2019-04-09 15:32 | 1554838306.7114 | 348 | 0.20923 | 0.000621 | 8.2e-05 | 99.5976 | 99.5182 | 114.073 | 20.6896 | 0.003268 | 0.002986 | 129.2217 | 0.2534 | 8.4016 | 0.002982 | 0.003983 | 0.00391 | 5.1e-05 | 0.001161 | 0.29662 | 1.346 | 2.5476 | 98.43 | 20.02 | 48.438 | 99.9241 | 93.0125 | 93.1094 | 0.011532 | 0.96392 | 0.022491 | 23.3325 | 0.007132 | 0.15075 | 0.001538 | 0.099024 | 0.099024 | 0.06135
2019-04-09 15:32 | 1554838307.7104 | 349 | 0.20922 | 0.000615 | 8.3e-05 | 99.6003 | 99.3891 | 114.1195 | 20.689 | 0.003276 | 0.002988 | 129.217 | 0.25352 | 8.4005 | 0.003074 | 0.003933 | 0.003818 | 4.9e-05 | 0.001159 | 0.26432 | 1.3694 | 2.5475 | 98.43 | 20.02 | 48.441 | 100.1119 | 95.2243 | 93.1481 | 0.011533 | 0.96382 | 0.022494 | 23.3384 | 0.01898 | 0.15082 | 0.001574 | 0.10141 | 0.10141 | 0.059932
2019-04-09 15:32 | 1554838308.7118 | 350 | 0.20924 | 0.000617 | 8.2e-05 | 99.6015 | 99.7182 | 114.2641 | 20.6881 | 0.003247 | 0.003023 | 129.2127 | 0.2541 | 8.4024 | 0.003056 | 0.004412 | 0.004254 | 5.3e-05 | 0.001164 | 0.2568 | 1.3559 | 2.5477 | 98.43 | 20.02 | 48.442 | 99.4088 | 94.7919 | 93.9968 | 0.011533 | 0.96159 | 0.022482 | 23.3803 | -0.010474 | 0.15116 | 0.001481 | 0.095167 | 0.095167 | 0.064014
2019-04-09 15:32 | 1554838309.7115 | 351 | 0.20925 | 0.000616 | 8.1e-05 | 99.6696 | 99.7202 | 114.188 | 20.6894 | 0.00326 | 0.003071 | 129.2136 | 0.25413 | 8.4019 | 0.003007 | 0.004141 | 0.004057 | 5.1e-05 | 0.001167 | 0.27527 | 1.3632 | 2.5475 | 98.43 | 20.02 | 48.445 | 99.7215 | 93.6093 | 95.158 | 0.011534 | 0.95856 | 0.022439 | 23.4094 | -0.004537 | 0.15119 | 0.001407 | 0.090062 | 0.090062 | 0.06765
2019-04-09 15:32 | 1554838310.7109 | 352 | 0.20922 | 0.000616 | 8e-05 | 99.7113 | 99.6125 | 113.9991 | 20.6891 | 0.003304 | 0.003033 | 129.2357 | 0.25292 | 8.3996 | 0.003057 | 0.004365 | 0.004221 | 5.1e-05 | 0.001169 | 0.28528 | 1.3862 | 2.5475 | 98.43 | 20.02 | 48.45 | 100.7719 | 94.8048 | 94.2419 | 0.011535 | 0.96095 | 0.022449 | 23.3609 | 0.008871 | 0.15046 | 0.0016 | 0.1031 | 0.1031 | 0.058814
2019-04-09 15:32 | 1554838311.7106 | 353 | 0.20921 | 0.000615 | 8e-05 | 99.6959 | 99.6775 | 113.969 | 20.6892 | 0.00323 | 0.003031 | 129.2106 | 0.25418 | 8.4011 | 0.003097 | 0.004321 | 0.004254 | 4.8e-05 | 0.001169 | 0.28943 | 1.3389 | 2.5476 | 98.43 | 20.02 | 48.454 | 98.9972 | 95.7804 | 94.1935 | 0.011536 | 0.96108 | 0.022447 | 23.3563 | 0.001643 | 0.15121 | 0.001672 | 0.10794 | 0.10794 | 0.056456
2019-04-09 15:32 | 1554838312.7116 | 354 | 0.20925 | 0.000617 | 8e-05 | 99.7879 | 99.6569 | 114.4887 | 20.689 | 0.003213 | 0.003028 | 129.1954 | 0.25339 | 8.4009 | 0.003106 | 0.004536 | 0.004504 | 4.8e-05 | 0.001166 | 0.29319 | 1.343 | 2.5477 | 98.43 | 20.02 | 48.458 | 98.587 | 96.0019 | 94.119 | 0.011537 | 0.96127 | 0.022501 | 23.4071 | 0.011748 | 0.15074 | 0.001438 | 0.092456 | 0.092456 | 0.065705
2019-04-09 15:32 | 1554838313.7106 | 355 | 0.20922 | 0.00061 | 7.8e-05 | 99.8013 | 99.6896 | 114.3305 | 20.6883 | 0.00321 | 0.003084 | 129.2027 | 0.2532 | 8.4037 | 0.003029 | 0.004667 | 0.004635 | 4.9e-05 | 0.001155 | 0.27903 | 1.3411 | 2.5477 | 98.43 | 20.02 | 48.461 | 98.5167 | 94.1285 | 95.475 | 0.011537 | 0.95774 | 0.022444 | 23.4341 | 0.010015 | 0.15063 | 0.001616 | 0.10428 | 0.10428 | 0.058214
2019-04-09 15:32 | 1554838314.711 | 356 | 0.20924 | 0.000607 | 7.8e-05 | 99.8587 | 99.8303 | 113.9146 | 20.6895 | 0.003204 | 0.002995 | 129.1921 | 0.25385 | 8.3998 | 0.003007 | 0.004462 | 0.004405 | 4.7e-05 | 0.001157 | 0.30038 | 1.3482 | 2.5476 | 98.43 | 20.02 | 48.462 | 98.3743 | 93.6063 | 93.3215 | 0.011538 | 0.96336 | 0.022469 | 23.323 | 0.002544 | 0.15101 | 0.001514 | 0.097551 | 0.097551 | 0.062387
2019-04-09 15:32 | 1554838315.7107 | 357 | 0.20918 | 0.000611 | 7.8e-05 | 99.9326 | 99.9587 | 113.85 | 20.6887 | 0.003193 | 0.003012 | 129.1654 | 0.25326 | 8.4019 | 0.002993 | 0.004509 | 0.004478 | 4.7e-05 | 0.001162 | 0.27472 | 1.3738 | 2.5477 | 98.43 | 20.03 | 48.467 | 98.099 | 93.2666 | 93.7219 | 0.011546 | 0.96231 | 0.02245 | 23.3291 | -0.002337 | 0.15066 | 0.00186 | 0.12073 | 0.12073 | 0.05029
2019-04-09 15:32 | 1554838316.7111 | 358 | 0.20925 | 0.000606 | 7.6e-05 | 99.8957 | 99.8985 | 113.8991 | 20.6889 | 0.003219 | 0.003035 | 129.189 | 0.25311 | 8.4015 | 0.002956 | 0.004624 | 0.004488 | 4.6e-05 | 0.001171 | 0.27367 | 1.3781 | 2.5476 | 98.43 | 20.03 | 48.469 | 98.7217 | 92.3922 | 94.2919 | 0.011546 | 0.96082 | 0.022437 | 23.3523 | -0.000252 | 0.15058 | 0.001462 | 0.093985 | 0.093985 | 0.064566
2019-04-09 15:32 | 1554838317.7101 | 359 | 0.20927 | 0.000607 | 7.6e-05 | 99.9609 | 99.8193 | 113.7256 | 20.6885 | 0.00315 | 0.003014 | 129.1973 | 0.25422 | 8.403 | 0.003 | 0.004653 | 0.00462 | 4.7e-05 | 0.001178 | 0.28368 | 1.3863 | 2.5477 | 98.43 | 20.03 | 48.47 | 97.0652 | 93.4527 | 93.7688 | 0.011547 | 0.96219 | 0.022436 | 23.3179 | 0.012675 | 0.15124 | 0.001342 | 0.085955 | 0.085955 | 0.070906
2019-04-09 15:32 | 1554838318.7108 | 360 | 0.20928 | 0.000608 | 7.5e-05 | 99.9642 | 100.0627 | 113.6737 | 20.6896 | 0.003136 | 0.002952 | 129.1632 | 0.25421 | 8.4036 | 0.003051 | 0.004337 | 0.004307 | 4.4e-05 | 0.001178 | 0.26537 | 1.4382 | 2.5477 | 98.43 | 20.03 | 48.473 | 96.7302 | 94.6675 | 92.2967 | 0.011547 | 0.96606 | 0.022476 | 23.2657 | -0.008808 | 0.15123 | 0.001277 | 0.081814 | 0.081814 | 0.074491
2019-04-09 15:32 | 1554838319.7108 | 361 | 0.20928 | 0.000604 | 7.4e-05 | 100.0362 | 100.1824 | 113.8768 | 20.6892 | 0.003223 | 0.003045 | 129.1568 | 0.25359 | 8.4022 | 0.002994 | 0.004524 | 0.004399 | 4.6e-05 | 0.001173 | 0.26698 | 1.4318 | 2.5476 | 98.43 | 20.03 | 48.473 | 98.8278 | 93.2905 | 94.516 | 0.011547 | 0.96023 | 0.022428 | 23.3571 | -0.013061 | 0.15086 | 0.001244 | 0.079496 | 0.079496 | 0.076479
2019-04-09 15:32 | 1554838320.7113 | 362 | 0.20924 | 0.000602 | 7.4e-05 | 100.035 | 100.1769 | 114.0188 | 20.6885 | 0.003239 | 0.003025 | 129.1469 | 0.25355 | 8.4036 | 0.003011 | 0.004476 | 0.004432 | 4.4e-05 | 0.001175 | 0.27759 | 1.3645 | 2.5476 | 98.43 | 20.03 | 48.473 | 99.221 | 93.6984 | 94.0387 | 0.011547 | 0.96148 | 0.022457 | 23.3565 | -0.012676 | 0.15084 | 0.001524 | 0.098381 | 0.098381 | 0.061789
2019-04-09 15:32 | 1554838321.7102 | 363 | 0.20927 | 0.000608 | 7.3e-05 | 100.0377 | 100.2609 | 114.0509 | 20.6904 | 0.003176 | 0.002952 | 129.1156 | 0.25399 | 8.4008 | 0.003016 | 0.004657 | 0.004546 | 4.8e-05 | 0.00118 | 0.27848 | 1.3662 | 2.5476 | 98.43 | 20.03 | 48.474 | 97.6861 | 93.8192 | 92.2937 | 0.011548 | 0.96607 | 0.022514 | 23.3042 | -0.019931 | 0.1511 | 0.001313 | 0.084418 | 0.084418 | 0.072131
2019-04-09 15:32 | 1554838322.7099 | 364 | 0.20935 | 0.0006 | 7.2e-05 | 100.1006 | 100.1604 | 114.0763 | 20.6895 | 0.003138 | 0.002995 | 129.1399 | 0.25395 | 8.4028 | 0.003048 | 0.004948 | 0.004827 | 4.6e-05 | 0.001175 | 0.27992 | 1.335 | 2.5475 | 98.43 | 20.03 | 48.475 | 96.7747 | 94.5955 | 93.3316 | 0.011548 | 0.96334 | 0.022484 | 23.3399 | -0.005345 | 0.15107 | 0.000872 | 0.054688 | 0.054688 | 0.11132
2019-04-09 15:32 | 1554838323.7114 | 365 | 0.20925 | 0.000598 | 7.2e-05 | 100.1687 | 100.1139 | 114.1549 | 20.6899 | 0.003093 | 0.002989 | 129.1397 | 0.25339 | 8.4031 | 0.003021 | 0.00434 | 0.004248 | 4.9e-05 | 0.001182 | 0.28456 | 1.3612 | 2.5475 | 98.43 | 20.03 | 48.475 | 95.6856 | 93.9525 | 93.1786 | 0.011548 | 0.96374 | 0.022497 | 23.343 | 0.004896 | 0.15074 | 0.001481 | 0.095776 | 0.095776 | 0.063427
2019-04-09 15:32 | 1554838324.7104 | 366 | 0.20927 | 0.000601 | 7.1e-05 | 100.1726 | 100.2719 | 114.6612 | 20.6896 | 0.003126 | 0.003 | 129.1402 | 0.25399 | 8.4006 | 0.002999 | 0.004155 | 0.004003 | 5e-05 | 0.001179 | 0.25658 | 1.3307 | 2.5476 | 98.43 | 20.03 | 48.475 | 96.4889 | 93.4115 | 93.4481 | 0.011548 | 0.96303 | 0.022538 | 23.4033 | -0.00886 | 0.1511 | 0.001313 | 0.084639 | 0.084639 | 0.071944
2019-04-09 15:32 | 1554838325.711 | 367 | 0.20927 | 0.000597 | 7.1e-05 | 100.1929 | 100.4192 | 115.0534 | 20.6884 | 0.003197 | 0.002996 | 129.1681 | 0.25394 | 8.4013 | 0.003025 | 0.004489 | 0.004412 | 4.8e-05 | 0.00118 | 0.26161 | 1.305 | 2.5474 | 98.43 | 20.02 | 48.476 | 98.2055 | 94.0518 | 93.3484 | 0.011541 | 0.96329 | 0.02258 | 23.4401 | -0.020182 | 0.15107 | 0.001321 | 0.085362 | 0.085362 | 0.07132
2019-04-09 15:32 | 1554838326.7104 | 368 | 0.20929 | 0.000597 | 7e-05 | 100.2238 | 100.2509 | 115.0532 | 20.6896 | 0.003066 | 0.00298 | 129.1536 | 0.25304 | 8.3996 | 0.003044 | 0.004695 | 0.004609 | 5e-05 | 0.001182 | 0.29031 | 1.3927 | 2.5476 | 98.43 | 20.02 | 48.476 | 95.0354 | 94.4962 | 92.9491 | 0.011541 | 0.96434 | 0.022592 | 23.4274 | -0.002415 | 0.15053 | 0.001197 | 0.077064 | 0.077064 | 0.078719
2019-04-09 15:32 | 1554838327.7115 | 369 | 0.20929 | 0.000598 | 7.1e-05 | 100.2745 | 100.5279 | 114.8253 | 20.6895 | 0.003102 | 0.002978 | 129.1097 | 0.25404 | 8.4016 | 0.00307 | 0.004703 | 0.004601 | 5e-05 | 0.001176 | 0.25608 | 1.3667 | 2.5475 | 98.43 | 20.02 | 48.477 | 95.8949 | 95.1323 | 92.9091 | 0.011541 | 0.96445 | 0.022571 | 23.4029 | -0.022582 | 0.15113 | 0.001237 | 0.0797 | 0.0797 | 0.076419
2019-04-09 15:32 | 1554838328.7111 | 370 | 0.20926 | 0.0006 | 7e-05 | 100.368 | 100.5754 | 114.7335 | 20.6896 | 0.003191 | 0.002941 | 129.1242 | 0.25313 | 8.4024 | 0.003094 | 0.004583 | 0.004598 | 4.5e-05 | 0.00117 | 0.29031 | 1.3855 | 2.5477 | 98.43 | 20.03 | 48.476 | 98.0444 | 95.7167 | 92.0309 | 0.011548 | 0.96677 | 0.022589 | 23.3654 | -0.01847 | 0.15058 | 0.001378 | 0.089279 | 0.089279 | 0.067973
2019-04-09 15:32 | 1554838329.7102 | 371 | 0.20924 | 0.000596 | 7e-05 | 100.4505 | 100.5668 | 114.6483 | 20.6889 | 0.003103 | 0.002952 | 129.1231 | 0.25379 | 8.4025 | 0.003105 | 0.004702 | 0.004528 | 4.5e-05 | 0.001166 | 0.29479 | 1.3757 | 2.5477 | 98.43 | 20.03 | 48.475 | 95.9383 | 95.9799 | 92.2852 | 0.011548 | 0.9661 | 0.022573 | 23.3649 | -0.010357 | 0.15098 | 0.001517 | 0.098614 | 0.098614 | 0.0617
2019-04-09 15:32 | 1554838330.7101 | 372 | 0.20927 | 0.000594 | 7e-05 | 100.4582 | 100.625 | 114.7776 | 20.6888 | 0.003135 | 0.002978 | 129.12 | 0.25449 | 8.4042 | 0.003071 | 0.00443 | 0.004362 | 4.9e-05 | 0.001158 | 0.25713 | 1.3811 | 2.5476 | 98.43 | 20.03 | 48.473 | 96.7105 | 95.1574 | 92.9186 | 0.011547 | 0.96442 | 0.022566 | 23.3983 | -0.014839 | 0.1514 | 0.001361 | 0.08806 | 0.08806 | 0.069286
2019-04-09 15:32 | 1554838331.7102 | 373 | 0.20926 | 0.000594 | 6.9e-05 | 100.4382 | 100.3631 | 115.1227 | 20.689 | 0.003083 | 0.002937 | 129.1397 | 0.25434 | 8.3999 | 0.003023 | 0.003653 | 0.003601 | 4.4e-05 | 0.00115 | 0.2677 | 1.3762 | 2.5476 | 98.43 | 20.03 | 48.472 | 95.4357 | 94.0047 | 91.9334 | 0.011547 | 0.96703 | 0.02263 | 23.4019 | 0.006695 | 0.1513 | 0.001378 | 0.08953 | 0.08953 | 0.068107
2019-04-09 15:32 | 1554838332.7092 | 374 | 0.2093 | 0.000596 | 6.8e-05 | 100.4927 | 100.6308 | 114.9546 | 20.6894 | 0.00313 | 0.002942 | 129.1475 | 0.25408 | 8.4021 | 0.002992 | 0.004847 | 0.004707 | 4.7e-05 | 0.001145 | 0.27417 | 1.391 | 2.5476 | 98.43 | 20.03 | 48.472 | 96.5733 | 93.2399 | 92.0355 | 0.011547 | 0.96676 | 0.022611 | 23.3881 | -0.012284 | 0.15115 | 0.001181 | 0.076202 | 0.076202 | 0.079939
2019-04-09 15:32 | 1554838333.7093 | 375 | 0.20927 | 0.000595 | 6.8e-05 | 100.4969 | 100.6074 | 114.8159 | 20.6895 | 0.003083 | 0.002953 | 129.1458 | 0.25325 | 8.4008 | 0.00305 | 0.004591 | 0.00454 | 4.5e-05 | 0.001139 | 0.28047 | 1.3293 | 2.5476 | 98.43 | 20.03 | 48.467 | 95.4487 | 94.6384 | 92.3053 | 0.011546 | 0.96604 | 0.022589 | 23.3826 | -0.009832 | 0.15066 | 0.00132 | 0.085503 | 0.085503 | 0.071009
2019-04-09 15:32 | 1554838334.7094 | 376 | 0.20929 | 0.000591 | 6.7e-05 | 100.5544 | 100.562 | 114.6981 | 20.6885 | 0.003102 | 0.002869 | 129.1363 | 0.25369 | 8.4023 | 0.003057 | 0.00437 | 0.004288 | 4.6e-05 | 0.001144 | 0.29247 | 1.3465 | 2.5476 | 98.43 | 20.03 | 48.466 | 95.9128 | 94.8184 | 90.2757 | 0.011546 | 0.97144 | 0.02264 | 23.3056 | -0.000679 | 0.15092 | 0.001253 | 0.081176 | 0.081176 | 0.074925
2019-04-09 15:32 | 1554838335.7097 | 377 | 0.20931 | 0.000591 | 6.6e-05 | 100.6062 | 100.881 | 114.5814 | 20.6895 | 0.003077 | 0.002915 | 129.1514 | 0.25477 | 8.4033 | 0.003037 | 0.005217 | 0.005167 | 4.4e-05 | 0.001139 | 0.27201 | 1.3425 | 2.5476 | 98.43 | 20.03 | 48.463 | 95.3109 | 94.3277 | 91.3982 | 0.011545 | 0.96845 | 0.022594 | 23.3297 | -0.024402 | 0.15156 | 0.001106 | 0.071112 | 0.071112 | 0.08589
2019-04-09 15:32 | 1554838336.7098 | 378 | 0.20927 | 0.000586 | 6.6e-05 | 100.6387 | 100.8542 | 114.5375 | 20.69 | 0.003138 | 0.002968 | 129.1476 | 0.2537 | 8.401 | 0.00304 | 0.004085 | 0.003964 | 4.6e-05 | 0.001142 | 0.28694 | 1.3483 | 2.5477 | 98.43 | 20.04 | 48.462 | 96.7627 | 94.4189 | 92.6647 | 0.011552 | 0.96509 | 0.02255 | 23.3657 | -0.019136 | 0.15093 | 0.001328 | 0.086037 | 0.086037 | 0.070694
2019-04-09 15:32 | 1554838337.7104 | 379 | 0.20929 | 0.000586 | 6.5e-05 | 100.6899 | 101.0359 | 114.6248 | 20.6898 | 0.003047 | 0.002984 | 129.1572 | 0.25453 | 8.3999 | 0.002999 | 0.004665 | 0.00457 | 4.6e-05 | 0.00113 | 0.26698 | 1.3039 | 2.5476 | 98.43 | 20.04 | 48.46 | 94.5683 | 93.4282 | 93.0542 | 0.011551 | 0.96407 | 0.022547 | 23.3871 | -0.030686 | 0.15142 | 0.001257 | 0.081207 | 0.081207 | 0.075144
2019-04-09 15:32 | 1554838338.7105 | 380 | 0.2093 | 0.000588 | 6.5e-05 | 100.7197 | 100.9884 | 114.8612 | 20.6897 | 0.00309 | 0.002943 | 129.1488 | 0.25394 | 8.4002 | 0.003016 | 0.004449 | 0.004368 | 4.7e-05 | 0.001127 | 0.30143 | 1.3425 | 2.5476 | 98.43 | 20.04 | 48.458 | 95.6132 | 93.816 | 92.0724 | 0.011551 | 0.96666 | 0.0226 | 23.3798 | -0.023836 | 0.15107 | 0.001194 | 0.077214 | 0.077214 | 0.078846
2019-04-09 15:32 | 1554838339.7095 | 381 | 0.20927 | 0.000583 | 6.5e-05 | 100.7268 | 100.9017 | 114.952 | 20.689 | 0.003049 | 0.002914 | 129.1382 | 0.25382 | 8.4013 | 0.003083 | 0.004558 | 0.004494 | 4.7e-05 | 0.001119 | 0.2453 | 1.3253 | 2.5474 | 98.43 | 20.04 | 48.457 | 94.6177 | 95.4444 | 91.3705 | 0.011551 | 0.96852 | 0.022631 | 23.3665 | -0.015517 | 0.151 | 0.001361 | 0.088615 | 0.088615 | 0.06867
2019-04-09 15:32 | 1554838340.7105 | 382 | 0.20929 | 0.000586 | 6.4e-05 | 100.8075 | 100.9237 | 114.9359 | 20.6895 | 0.003003 | 0.002907 | 129.1219 | 0.2543 | 8.4032 | 0.003037 | 0.004346 | 0.00427 | 4.9e-05 | 0.001124 | 0.25984 | 1.3389 | 2.5475 | 98.43 | 20.04 | 48.457 | 93.5248 | 94.3402 | 91.2121 | 0.011551 | 0.96894 | 0.022634 | 23.3598 | -0.010306 | 0.15129 | 0.001241 | 0.08053 | 0.08053 | 0.075709
2019-04-09 15:32 | 1554838341.7099 | 383 | 0.20929 | 0.000584 | 6.4e-05 | 100.8164 | 101.3226 | 114.7025 | 20.6905 | 0.003027 | 0.002938 | 129.1184 | 0.25392 | 8.3999 | 0.003024 | 0.003627 | 0.003561 | 4.6e-05 | 0.001126 | 0.30829 | 1.3777 | 2.5475 | 98.43 | 20.04 | 48.457 | 94.0908 | 94.0211 | 91.9425 | 0.011551 | 0.967 | 0.022589 | 23.3594 | -0.044794 | 0.15106 | 0.001266 | 0.082045 | 0.082045 | 0.074199
2019-04-09 15:32 | 1554838342.71 | 384 | 0.20927 | 0.000578 | 6.4e-05 | 100.8188 | 101.1863 | 114.8544 | 20.6895 | 0.003032 | 0.002918 | 129.0826 | 0.25387 | 8.4002 | 0.003025 | 0.004226 | 0.004099 | 5.1e-05 | 0.001133 | 0.29319 | 1.3409 | 2.5475 | 98.43 | 20.04 | 48.457 | 94.2095 | 94.0365 | 91.4632 | 0.011551 | 0.96827 | 0.022618 | 23.3596 | -0.032544 | 0.15103 | 0.001401 | 0.091365 | 0.091365 | 0.066617
2019-04-09 15:32 | 1554838343.7106 | 385 | 0.20928 | 0.000585 | 6.3e-05 | 100.8745 | 101.0184 | 114.8898 | 20.689 | 0.00309 | 0.002903 | 129.0662 | 0.25363 | 8.4018 | 0.003048 | 0.003998 | 0.003955 | 4.5e-05 | 0.001136 | 0.31078 | 1.3355 | 2.5475 | 98.43 | 20.04 | 48.458 | 95.6009 | 94.609 | 91.1039 | 0.011551 | 0.96923 | 0.022633 | 23.3516 | -0.012746 | 0.15089 | 0.001301 | 0.084647 | 0.084647 | 0.071836
2019-04-09 15:32 | 1554838344.7106 | 386 | 0.2093 | 0.000581 | 6.2e-05 | 100.9311 | 101.1385 | 115.1565 | 20.6892 | 0.003019 | 0.002914 | 129.0692 | 0.25404 | 8.4029 | 0.003064 | 0.004346 | 0.004373 | 4.4e-05 | 0.001138 | 0.25536 | 1.3689 | 2.5475 | 98.43 | 20.04 | 48.459 | 93.8922 | 94.9927 | 91.3699 | 0.011551 | 0.96852 | 0.022651 | 23.3873 | -0.018359 | 0.15113 | 0.001215 | 0.078939 | 0.078939 | 0.077155
2019-04-09 15:32 | 1554838345.7087 | 387 | 0.20932 | 0.000575 | 6.1e-05 | 100.9186 | 101.1904 | 115.1699 | 20.6895 | 0.003057 | 0.002874 | 129.0395 | 0.25353 | 8.404 | 0.003042 | 0.004137 | 0.003996 | 4.5e-05 | 0.001144 | 0.27527 | 1.3212 | 2.5476 | 98.43 | 20.04 | 48.46 | 94.8131 | 94.4662 | 90.4017 | 0.011551 | 0.9711 | 0.022683 | 23.3575 | -0.024062 | 0.15082 | 0.001081 | 0.06997 | 0.06997 | 0.086869
2019-04-09 15:32 | 1554838346.7091 | 388 | 0.20931 | 0.000579 | 6.1e-05 | 100.9612 | 101.484 | 115.3788 | 20.6885 | 0.003063 | 0.002942 | 129.0038 | 0.25458 | 8.4012 | 0.002996 | 0.004529 | 0.004487 | 4.6e-05 | 0.001152 | 0.28584 | 1.3207 | 2.5476 | 98.43 | 20.04 | 48.461 | 94.9631 | 93.3439 | 92.0524 | 0.011552 | 0.96671 | 0.022652 | 23.4317 | -0.046201 | 0.15145 | 0.001141 | 0.073967 | 0.073967 | 0.082514
2019-04-09 15:32 | 1554838347.7087 | 389 | 0.20936 | 0.000577 | 6.1e-05 | 100.9731 | 101.0462 | 115.6353 | 20.689 | 0.003041 | 0.00289 | 129.0025 | 0.25334 | 8.4031 | 0.003019 | 0.004885 | 0.004764 | 4.7e-05 | 0.001155 | 0.30502 | 1.3275 | 2.5476 | 98.43 | 20.04 | 48.464 | 94.4321 | 93.9038 | 90.7818 | 0.011552 | 0.97009 | 0.022716 | 23.4169 | -0.006478 | 0.15071 | 0.000825 | 0.052716 | 0.052716 | 0.11522
2019-04-09 15:32 | 1554838348.7097 | 390 | 0.20919 | 0.000572 | 6e-05 | 101.0032 | 101.3068 | 115.556 | 20.6892 | 0.003069 | 0.002874 | 129.0088 | 0.25357 | 8.4007 | 0.002975 | 0.00402 | 0.003951 | 4.6e-05 | 0.001154 | 0.30038 | 1.3253 | 2.5475 | 98.43 | 20.04 | 48.466 | 95.0992 | 92.8361 | 90.4139 | 0.011553 | 0.97107 | 0.02272 | 23.3971 | -0.026846 | 0.15085 | 0.001862 | 0.12335 | 0.12335 | 0.049285
2019-04-09 15:32 | 1554838349.7099 | 391 | 0.20934 | 0.000575 | 5.9e-05 | 101.0925 | 101.3377 | 115.5297 | 20.6892 | 0.003079 | 0.00287 | 129.0008 | 0.25328 | 8.4004 | 0.003013 | 0.00471 | 0.004678 | 4.5e-05 | 0.001155 | 0.28191 | 1.3524 | 2.5477 | 98.43 | 20.04 | 48.47 | 95.3465 | 93.7654 | 90.3123 | 0.011554 | 0.97134 | 0.022721 | 23.3911 | -0.021673 | 0.15068 | 0.000932 | 0.060052 | 0.060052 | 0.10112
2019-04-09 15:32 | 1554838350.7098 | 392 | 0.20936 | 0.000579 | 5.9e-05 | 101.0496 | 101.5346 | 115.2551 | 20.6888 | 0.003011 | 0.002879 | 128.9996 | 0.25465 | 8.404 | 0.002977 | 0.004834 | 0.004657 | 4.4e-05 | 0.001153 | 0.2672 | 1.372 | 2.5476 | 98.43 | 20.04 | 48.475 | 93.7113 | 92.8984 | 90.5143 | 0.011555 | 0.9708 | 0.022687 | 23.3698 | -0.042826 | 0.15149 | 0.000834 | 0.05337 | 0.05337 | 0.11439
2019-04-09 15:33 | 1554838351.7092 | 393 | 0.20936 | 0.00057 | 5.8e-05 | 101.1083 | 101.3797 | 115.2508 | 20.6892 | 0.003059 | 0.00294 | 129.0069 | 0.25441 | 8.4032 | 0.002967 | 0.00458 | 0.004587 | 4.3e-05 | 0.001161 | 0.27903 | 1.3559 | 2.5476 | 98.43 | 20.04 | 48.476 | 94.8697 | 92.6461 | 91.996 | 0.011555 | 0.96686 | 0.022641 | 23.4169 | -0.023981 | 0.15135 | 0.000873 | 0.055932 | 0.055932 | 0.10905
2019-04-09 15:33 | 1554838352.7089 | 394 | 0.20932 | 0.000569 | 5.7e-05 | 101.099 | 101.5226 | 114.8704 | 20.6898 | 0.003031 | 0.002836 | 129.008 | 0.25432 | 8.4047 | 0.002927 | 0.004451 | 0.004357 | 4.4e-05 | 0.00117 | 0.27936 | 1.3197 | 2.5476 | 98.43 | 20.04 | 48.48 | 94.1886 | 91.6792 | 89.5017 | 0.011556 | 0.97351 | 0.022681 | 23.2983 | -0.03739 | 0.15129 | 0.001099 | 0.071432 | 0.071432 | 0.085355
2019-04-09 15:33 | 1554838353.7103 | 395 | 0.20935 | 0.000568 | 5.7e-05 | 101.0943 | 101.4716 | 114.8577 | 20.6897 | 0.003077 | 0.002853 | 129.0099 | 0.25409 | 8.402 | 0.00297 | 0.004486 | 0.004382 | 4e-05 | 0.001176 | 0.28224 | 1.3358 | 2.5476 | 98.43 | 20.04 | 48.483 | 95.308 | 92.7177 | 89.9072 | 0.011557 | 0.97242 | 0.022667 | 23.31 | -0.033323 | 0.15116 | 0.000884 | 0.056791 | 0.056791 | 0.10726
2019-04-09 15:33 | 1554838354.7103 | 396 | 0.20937 | 0.000573 | 5.6e-05 | 101.0705 | 101.2741 | 114.6295 | 20.6894 | 0.003042 | 0.002864 | 129.0173 | 0.25458 | 8.4005 | 0.003001 | 0.004722 | 0.004674 | 4.2e-05 | 0.001181 | 0.29031 | 1.3776 | 2.5477 | 98.43 | 20.04 | 48.485 | 94.4535 | 93.4581 | 90.1593 | 0.011557 | 0.97175 | 0.022637 | 23.2949 | -0.018001 | 0.15145 | 0.00079 | 0.050417 | 0.050417 | 0.12106
2019-04-09 15:33 | 1554838355.709 | 397 | 0.20936 | 0.000568 | 5.6e-05 | 101.1026 | 101.473 | 114.5425 | 20.6902 | 0.002995 | 0.002934 | 129.0057 | 0.25415 | 8.4018 | 0.002977 | 0.004011 | 0.00397 | 4.2e-05 | 0.001175 | 0.28816 | 1.3484 | 2.5476 | 98.43 | 20.04 | 48.486 | 93.3236 | 92.8966 | 91.8531 | 0.011558 | 0.96724 | 0.022576 | 23.3403 | -0.032707 | 0.15119 | 0.000834 | 0.053284 | 0.053284 | 0.11435
2019-04-09 15:33 | 1554838356.71 | 398 | 0.20934 | 0.000562 | 5.6e-05 | 101.0839 | 101.229 | 114.0906 | 20.69 | 0.003019 | 0.002902 | 129.0292 | 0.25419 | 8.402 | 0.00293 | 0.004881 | 0.004736 | 4.2e-05 | 0.001177 | 0.27959 | 1.3447 | 2.5474 | 98.43 | 20.04 | 48.488 | 93.9038 | 91.7598 | 91.0805 | 0.011558 | 0.96929 | 0.022555 | 23.2695 | -0.012833 | 0.15121 | 0.000975 | 0.062779 | 0.062779 | 0.097071
2019-04-09 15:33 | 1554838357.71 | 399 | 0.20934 | 0.000569 | 5.6e-05 | 101.1297 | 101.3574 | 113.6467 | 20.6896 | 0.003034 | 0.002887 | 129.0476 | 0.25329 | 8.4022 | 0.002939 | 0.003824 | 0.003764 | 4.5e-05 | 0.001181 | 0.27721 | 1.341 | 2.5476 | 98.43 | 20.04 | 48.49 | 94.2518 | 91.9747 | 90.7082 | 0.011559 | 0.97028 | 0.022523 | 23.2124 | -0.020111 | 0.15069 | 0.000978 | 0.062873 | 0.062873 | 0.096585
2019-04-09 15:33 | 1554838358.7091 | 400 | 0.20933 | 0.000569 | 5.5e-05 | 101.1247 | 101.3801 | 113.4848 | 20.6894 | 0.003052 | 0.002901 | 129.0482 | 0.25433 | 8.4019 | 0.002884 | 0.00422 | 0.00404 | 4.2e-05 | 0.001185 | 0.27455 | 1.3822 | 2.5476 | 98.43 | 20.04 | 48.491 | 94.7041 | 90.6461 | 91.0494 | 0.011559 | 0.96937 | 0.022496 | 23.2067 | -0.022563 | 0.1513 | 0.001043 | 0.067212 | 0.067212 | 0.090719
2019-04-09 15:33 | 1554838359.7092 | 401 | 0.20937 | 0.000563 | 5.5e-05 | 101.1646 | 101.2717 | 113.7655 | 20.6904 | 0.003047 | 0.002908 | 129.0359 | 0.25453 | 8.4013 | 0.002961 | 0.003917 | 0.003913 | 4.6e-05 | 0.001189 | 0.28512 | 1.3517 | 2.5475 | 98.43 | 20.04 | 48.491 | 94.5706 | 92.4954 | 91.2224 | 0.011559 | 0.96891 | 0.022518 | 23.2409 | -0.009464 | 0.15142 | 0.000824 | 0.052535 | 0.052535 | 0.11616
2019-04-09 15:33 | 1554838360.7095 | 402 | 0.20931 | 0.000564 | 5.5e-05 | 101.1256 | 101.3102 | 114.0411 | 20.6903 | 0.003099 | 0.002872 | 129.0564 | 0.25489 | 8.4011 | 0.00293 | 0.003915 | 0.003856 | 4.2e-05 | 0.00119 | 0.2844 | 1.3385 | 2.5476 | 98.43 | 20.04 | 48.489 | 95.8419 | 91.7618 | 90.3572 | 0.011558 | 0.97122 | 0.022572 | 23.2414 | -0.016318 | 0.15163 | 0.001178 | 0.076544 | 0.076544 | 0.079835
2019-04-09 15:33 | 1554838361.7086 | 403 | 0.20934 | 0.000564 | 5.5e-05 | 101.2167 | 101.4458 | 114.0406 | 20.6898 | 0.003041 | 0.002914 | 129.0466 | 0.25433 | 8.402 | 0.002905 | 0.004569 | 0.00443 | 4.2e-05 | 0.001191 | 0.25697 | 1.3527 | 2.5476 | 98.43 | 20.04 | 48.489 | 94.4274 | 91.1414 | 91.379 | 0.011558 | 0.9685 | 0.022541 | 23.274 | -0.020225 | 0.1513 | 0.000984 | 0.063387 | 0.063387 | 0.096193
2019-04-09 15:33 | 1554838362.7082 | 404 | 0.20935 | 0.000561 | 5.5e-05 | 101.1729 | 101.4561 | 114.3441 | 20.6896 | 0.003072 | 0.002867 | 129.0415 | 0.25445 | 8.4012 | 0.002889 | 0.004197 | 0.004166 | 4.2e-05 | 0.00119 | 0.26913 | 1.3079 | 2.5475 | 98.43 | 20.04 | 48.49 | 95.1809 | 90.7567 | 90.2313 | 0.011559 | 0.97156 | 0.022606 | 23.2682 | -0.025005 | 0.15137 | 0.000898 | 0.057748 | 0.057748 | 0.10564
2019-04-09 15:33 | 1554838363.7096 | 405 | 0.20934 | 0.000564 | 5.4e-05 | 101.2307 | 101.5456 | 114.5683 | 20.6895 | 0.003024 | 0.002823 | 129.0455 | 0.25402 | 8.4027 | 0.00293 | 0.004678 | 0.004589 | 4.3e-05 | 0.001186 | 0.27527 | 1.3285 | 2.5477 | 98.43 | 20.04 | 48.49 | 94.0272 | 91.7556 | 89.1788 | 0.011559 | 0.97438 | 0.022661 | 23.2572 | -0.027784 | 0.15111 | 0.000972 | 0.062958 | 0.062958 | 0.09673
2019-04-09 15:33 | 1554838364.7093 | 406 | 0.20928 | 0.000563 | 5.4e-05 | 101.2036 | 101.4592 | 114.6305 | 20.6901 | 0.003001 | 0.002889 | 129.0241 | 0.25418 | 8.4005 | 0.002848 | 0.004397 | 0.004373 | 4.2e-05 | 0.001183 | 0.28888 | 1.3167 | 2.5475 | 98.43 | 20.04 | 48.49 | 93.4675 | 89.7739 | 90.7579 | 0.011559 | 0.97015 | 0.022618 | 23.3142 | -0.022567 | 0.15121 | 0.001355 | 0.088804 | 0.088804 | 0.068621
2019-04-09 15:33 | 1554838365.7089 | 407 | 0.20931 | 0.000562 | 5.3e-05 | 101.2953 | 101.5955 | 114.6833 | 20.69 | 0.003008 | 0.002848 | 129.0017 | 0.25482 | 8.4015 | 0.002912 | 0.004368 | 0.004287 | 4.3e-05 | 0.001182 | 0.28479 | 1.3218 | 2.5477 | 98.43 | 20.03 | 48.492 | 93.6451 | 91.3324 | 89.7817 | 0.011552 | 0.97276 | 0.022654 | 23.2883 | -0.026471 | 0.15159 | 0.001178 | 0.076916 | 0.076916 | 0.079426
2019-04-09 15:33 | 1554838366.709 | 408 | 0.20931 | 0.000567 | 5.3e-05 | 101.2741 | 101.483 | 115.2274 | 20.6902 | 0.003003 | 0.002836 | 129.0146 | 0.25475 | 8.402 | 0.002923 | 0.004585 | 0.004455 | 4.5e-05 | 0.00118 | 0.28174 | 1.337 | 2.5476 | 98.43 | 20.03 | 48.493 | 93.5254 | 91.5935 | 89.4833 | 0.011552 | 0.97356 | 0.022717 | 23.3338 | -0.018424 | 0.15155 | 0.001147 | 0.075056 | 0.075056 | 0.081371
2019-04-09 15:33 | 1554838367.7098 | 409 | 0.2093 | 0.000568 | 5.3e-05 | 101.2795 | 101.6482 | 115.6149 | 20.6901 | 0.003041 | 0.00283 | 129.0141 | 0.2547 | 8.4032 | 0.002909 | 0.004481 | 0.004409 | 4.5e-05 | 0.001181 | 0.28639 | 1.3681 | 2.5477 | 98.43 | 20.03 | 48.497 | 94.4206 | 91.2495 | 89.3536 | 0.011553 | 0.97391 | 0.022759 | 23.3689 | -0.032499 | 0.15152 | 0.001232 | 0.081008 | 0.081008 | 0.075379
2019-04-09 15:33 | 1554838368.7084 | 410 | 0.20934 | 0.000567 | 5.3e-05 | 101.1842 | 101.4541 | 115.9311 | 20.6904 | 0.003013 | 0.002911 | 129.0058 | 0.25478 | 8.4002 | 0.002891 | 0.004197 | 0.004187 | 4.6e-05 | 0.001175 | 0.28224 | 1.3518 | 2.5476 | 98.43 | 20.03 | 48.503 | 93.747 | 90.8176 | 91.299 | 0.011555 | 0.96871 | 0.022729 | 23.4635 | -0.023822 | 0.15157 | 0.000996 | 0.064789 | 0.064789 | 0.094277
2019-04-09 15:33 | 1554838369.7098 | 411 | 0.20934 | 0.000565 | 5.3e-05 | 101.1506 | 101.5287 | 115.848 | 20.6911 | 0.00299 | 0.00289 | 129.0291 | 0.25462 | 8.4031 | 0.002852 | 0.004225 | 0.00414 | 4.2e-05 | 0.001179 | 0.27831 | 1.3386 | 2.5475 | 98.43 | 20.03 | 48.506 | 93.2126 | 89.874 | 90.7838 | 0.011555 | 0.97008 | 0.022737 | 23.4385 | -0.033379 | 0.15147 | 0.000988 | 0.064319 | 0.064319 | 0.094908
2019-04-09 15:33 | 1554838370.7085 | 412 | 0.20933 | 0.000565 | 5.2e-05 | 101.2345 | 101.5824 | 115.8233 | 20.6893 | 0.002932 | 0.002825 | 129.0196 | 0.25411 | 8.401 | 0.002942 | 0.004019 | 0.003917 | 4.3e-05 | 0.001182 | 0.27345 | 1.3762 | 2.5477 | 98.43 | 20.03 | 48.51 | 91.8093 | 92.0448 | 89.2192 | 0.011556 | 0.97427 | 0.022784 | 23.3856 | -0.030685 | 0.15117 | 0.001036 | 0.067725 | 0.067725 | 0.089956
2019-04-09 15:33 | 1554838371.7086 | 413 | 0.20932 | 0.000564 | 5.1e-05 | 101.2819 | 101.6609 | 115.898 | 20.6895 | 0.002921 | 0.002874 | 128.996 | 0.25404 | 8.4028 | 0.002922 | 0.004877 | 0.004815 | 4.4e-05 | 0.001179 | 0.28136 | 1.355 | 2.5476 | 98.43 | 20.03 | 48.512 | 91.5288 | 91.5555 | 90.3947 | 0.011557 | 0.97112 | 0.022754 | 23.431 | -0.033409 | 0.15113 | 0.001136 | 0.074497 | 0.074497 | 0.081756
2019-04-09 15:33 | 1554838372.7085 | 414 | 0.20933 | 0.000569 | 5.2e-05 | 101.2435 | 101.6458 | 115.7989 | 20.6903 | 0.002997 | 0.002865 | 128.9969 | 0.25459 | 8.4006 | 0.002953 | 0.003727 | 0.003697 | 4.3e-05 | 0.001174 | 0.29358 | 1.3311 | 2.5476 | 98.43 | 20.03 | 48.515 | 93.3675 | 92.3092 | 90.1977 | 0.011557 | 0.97165 | 0.022751 | 23.4147 | -0.03547 | 0.15146 | 0.001009 | 0.065843 | 0.065843 | 0.092699
2019-04-09 15:33 | 1554838373.7096 | 415 | 0.20933 | 0.000565 | 5.1e-05 | 101.2092 | 101.4541 | 115.5452 | 20.6894 | 0.002915 | 0.002854 | 128.9819 | 0.25396 | 8.4022 | 0.002938 | 0.004184 | 0.004076 | 4.5e-05 | 0.001169 | 0.2912 | 1.347 | 2.5475 | 98.43 | 20.03 | 48.518 | 91.4005 | 91.9478 | 89.9281 | 0.011558 | 0.97237 | 0.022734 | 23.3803 | -0.021611 | 0.15108 | 0.001041 | 0.067995 | 0.067995 | 0.089544
2019-04-09 15:33 | 1554838374.7096 | 416 | 0.20932 | 0.000563 | 5.1e-05 | 101.2688 | 101.6489 | 115.3851 | 20.6906 | 0.00299 | 0.002809 | 128.9743 | 0.25436 | 8.4017 | 0.002926 | 0.004149 | 0.004082 | 4.4e-05 | 0.001159 | 0.27024 | 1.3335 | 2.5476 | 98.43 | 20.03 | 48.52 | 93.1918 | 91.6541 | 88.8424 | 0.011559 | 0.97528 | 0.022753 | 23.3292 | -0.033506 | 0.15132 | 0.001102 | 0.072277 | 0.072277 | 0.084372
2019-04-09 15:33 | 1554838375.7086 | 417 | 0.20935 | 0.000565 | 5e-05 | 101.2718 | 101.5081 | 115.2267 | 20.6899 | 0.002961 | 0.002788 | 128.9989 | 0.25513 | 8.3999 | 0.002956 | 0.00405 | 0.00399 | 4.6e-05 | 0.001159 | 0.2896 | 1.3554 | 2.5476 | 98.43 | 20.03 | 48.522 | 92.5091 | 92.3768 | 88.3293 | 0.011559 | 0.97667 | 0.022753 | 23.2966 | -0.020849 | 0.15178 | 0.000899 | 0.058467 | 0.058467 | 0.10462
2019-04-09 15:33 | 1554838376.7077 | 418 | 0.20933 | 0.000562 | 5e-05 | 101.2521 | 101.5246 | 115.0154 | 20.6897 | 0.002994 | 0.002832 | 129.003 | 0.25391 | 8.4012 | 0.00295 | 0.004428 | 0.004369 | 4.5e-05 | 0.001156 | 0.28385 | 1.3437 | 2.5476 | 98.43 | 20.03 | 48.522 | 93.2998 | 92.2469 | 89.3978 | 0.011559 | 0.97379 | 0.022699 | 23.3096 | -0.02404 | 0.15105 | 0.001032 | 0.067339 | 0.067339 | 0.090397
2019-04-09 15:33 | 1554838377.7088 | 419 | 0.20931 | 0.000558 | 5e-05 | 101.281 | 101.7142 | 114.8025 | 20.6896 | 0.002961 | 0.002818 | 128.984 | 0.25426 | 8.4028 | 0.002848 | 0.004628 | 0.004448 | 4.2e-05 | 0.001157 | 0.27057 | 1.3392 | 2.5476 | 98.43 | 20.03 | 48.523 | 92.4946 | 89.7709 | 89.0678 | 0.011559 | 0.97468 | 0.022688 | 23.2774 | -0.03818 | 0.15126 | 0.00115 | 0.075379 | 0.075379 | 0.080867
2019-04-09 15:33 | 1554838378.7077 | 420 | 0.20934 | 0.000556 | 4.9e-05 | 101.2506 | 101.6464 | 114.664 | 20.6897 | 0.00295 | 0.002809 | 128.986 | 0.25469 | 8.4013 | 0.002895 | 0.00429 | 0.004274 | 4.6e-05 | 0.001161 | 0.26753 | 1.3168 | 2.5476 | 98.43 | 20.03 | 48.522 | 92.2351 | 90.9177 | 88.8479 | 0.011559 | 0.97527 | 0.022681 | 23.2563 | -0.034899 | 0.15152 | 0.001011 | 0.065907 | 0.065907 | 0.092647
2019-04-09 15:33 | 1554838379.7077 | 421 | 0.20934 | 0.000559 | 4.9e-05 | 101.2646 | 101.6447 | 114.7117 | 20.6892 | 0.002936 | 0.002864 | 128.981 | 0.25438 | 8.4003 | 0.002859 | 0.004276 | 0.00423 | 4.1e-05 | 0.001158 | 0.28102 | 1.3174 | 2.5476 | 98.43 | 20.03 | 48.519 | 91.9046 | 90.0528 | 90.173 | 0.011558 | 0.97171 | 0.022644 | 23.3037 | -0.033511 | 0.15133 | 0.001009 | 0.065662 | 0.065662 | 0.092879
2019-04-09 15:33 | 1554838380.7092 | 422 | 0.20935 | 0.000559 | 4.9e-05 | 101.2634 | 101.8171 | 114.7053 | 20.6895 | 0.002988 | 0.002796 | 128.9653 | 0.25445 | 8.4017 | 0.002901 | 0.004066 | 0.004028 | 4.3e-05 | 0.001156 | 0.27743 | 1.3348 | 2.5475 | 98.43 | 20.03 | 48.517 | 93.1641 | 91.0568 | 88.5328 | 0.011558 | 0.97612 | 0.022695 | 23.2504 | -0.048776 | 0.15137 | 0.000914 | 0.059381 | 0.059381 | 0.10273
2019-04-09 15:33 | 1554838381.7078 | 423 | 0.20935 | 0.000554 | 4.8e-05 | 101.2607 | 101.685 | 114.7194 | 20.6889 | 0.003011 | 0.002797 | 128.9872 | 0.25423 | 8.4026 | 0.002942 | 0.004534 | 0.004419 | 4.4e-05 | 0.001146 | 0.28943 | 1.3486 | 2.5477 | 98.43 | 20.03 | 48.515 | 93.7145 | 92.0476 | 88.5479 | 0.011557 | 0.97608 | 0.022696 | 23.2523 | -0.037396 | 0.15124 | 0.00093 | 0.060502 | 0.060502 | 0.10074
2019-04-09 15:33 | 1554838382.7082 | 424 | 0.20935 | 0.000554 | 4.8e-05 | 101.2816 | 101.6623 | 114.7302 | 20.6889 | 0.002945 | 0.002794 | 129.005 | 0.25458 | 8.4006 | 0.002922 | 0.004962 | 0.004855 | 4.4e-05 | 0.00114 | 0.29015 | 1.334 | 2.5476 | 98.43 | 20.03 | 48.512 | 92.1106 | 91.5602 | 88.4845 | 0.011557 | 0.97625 | 0.022699 | 23.2514 | -0.033557 | 0.15145 | 0.00096 | 0.062547 | 0.062547 | 0.09758
2019-04-09 15:33 | 1554838383.7078 | 425 | 0.20934 | 0.000552 | 4.8e-05 | 101.284 | 101.6926 | 114.8213 | 20.688 | 0.002949 | 0.002801 | 129.01 | 0.25348 | 8.4036 | 0.002956 | 0.004834 | 0.00474 | 4.3e-05 | 0.001141 | 0.28727 | 1.3415 | 2.5476 | 98.43 | 20.03 | 48.506 | 92.1999 | 92.3843 | 88.6468 | 0.011555 | 0.97581 | 0.022703 | 23.2658 | -0.03601 | 0.15079 | 0.000986 | 0.0644 | 0.0644 | 0.094363
2019-04-09 15:33 | 1554838384.7089 | 426 | 0.20932 | 0.000554 | 4.7e-05 | 101.2706 | 101.7986 | 114.9347 | 20.6886 | 0.002952 | 0.00281 | 129.0021 | 0.25508 | 8.4032 | 0.002866 | 0.004459 | 0.004404 | 4.5e-05 | 0.001145 | 0.28766 | 1.3822 | 2.5476 | 98.43 | 20.03 | 48.503 | 92.2912 | 90.2163 | 88.8631 | 0.011555 | 0.97523 | 0.022707 | 23.2842 | -0.046512 | 0.15175 | 0.001113 | 0.073048 | 0.073048 | 0.083717
2019-04-09 15:33 | 1554838385.7083 | 427 | 0.20935 | 0.000553 | 4.7e-05 | 101.2887 | 101.7373 | 115.3051 | 20.6878 | 0.003002 | 0.002824 | 129.0111 | 0.25536 | 8.4023 | 0.002887 | 0.004071 | 0.003997 | 4.4e-05 | 0.001147 | 0.27721 | 1.3765 | 2.5476 | 98.43 | 20.03 | 48.5 | 93.4917 | 90.7103 | 89.1927 | 0.011554 | 0.97434 | 0.022734 | 23.3323 | -0.039524 | 0.15192 | 0.000924 | 0.060243 | 0.060243 | 0.10163
2019-04-09 15:33 | 1554838386.7084 | 428 | 0.20935 | 0.000552 | 4.6e-05 | 101.2968 | 101.7129 | 115.4058 | 20.6884 | 0.002902 | 0.002776 | 128.987 | 0.25524 | 8.4004 | 0.002875 | 0.004305 | 0.004265 | 4.3e-05 | 0.001148 | 0.27975 | 1.3827 | 2.5477 | 98.43 | 20.03 | 48.494 | 91.0867 | 90.4225 | 88.0536 | 0.011552 | 0.97741 | 0.022779 | 23.3058 | -0.036665 | 0.15184 | 0.000937 | 0.061343 | 0.061343 | 0.099755
2019-04-09 15:33 | 1554838387.7081 | 429 | 0.20934 | 0.000548 | 4.6e-05 | 101.2989 | 101.6344 | 115.4229 | 20.6888 | 0.002928 | 0.002777 | 128.9682 | 0.25462 | 8.4018 | 0.002922 | 0.00517 | 0.005102 | 4.3e-05 | 0.001153 | 0.27743 | 1.3361 | 2.5475 | 98.43 | 20.03 | 48.492 | 91.697 | 91.5594 | 88.0612 | 0.011552 | 0.97739 | 0.022781 | 23.3078 | -0.029578 | 0.15147 | 0.001035 | 0.068052 | 0.068052 | 0.089702
2019-04-09 15:33 | 1554838388.709 | 430 | 0.20936 | 0.000547 | 4.6e-05 | 101.2554 | 101.642 | 115.2954 | 20.6883 | 0.002945 | 0.002726 | 128.97 | 0.25493 | 8.4017 | 0.002879 | 0.003536 | 0.003426 | 4.2e-05 | 0.001148 | 0.27184 | 1.3372 | 2.5477 | 98.43 | 20.03 | 48.489 | 92.1142 | 90.5311 | 86.8511 | 0.011551 | 0.98068 | 0.022807 | 23.2559 | -0.034084 | 0.15166 | 0.00089 | 0.05815 | 0.05815 | 0.10511
2019-04-09 15:33 | 1554838389.7081 | 431 | 0.20938 | 0.000548 | 4.6e-05 | 101.2524 | 101.5601 | 115.2912 | 20.6889 | 0.003015 | 0.002833 | 128.9523 | 0.25488 | 8.4013 | 0.00289 | 0.004353 | 0.00435 | 4.2e-05 | 0.001155 | 0.27256 | 1.3719 | 2.5476 | 98.43 | 20.03 | 48.488 | 93.8013 | 90.7786 | 89.4291 | 0.011551 | 0.97371 | 0.022725 | 23.3386 | -0.027137 | 0.15163 | 0.000803 | 0.05209 | 0.05209 | 0.11731
2019-04-09 15:33 | 1554838390.7081 | 432 | 0.20936 | 0.000545 | 4.5e-05 | 101.2599 | 101.8144 | 115.202 | 20.6885 | 0.002954 | 0.002811 | 128.9403 | 0.25456 | 8.4023 | 0.002934 | 0.004242 | 0.004204 | 4.2e-05 | 0.001154 | 0.2995 | 1.3133 | 2.5476 | 98.43 | 20.03 | 48.486 | 92.3277 | 91.8524 | 88.8763 | 0.01155 | 0.97519 | 0.022733 | 23.3117 | -0.04885 | 0.15144 | 0.000881 | 0.057462 | 0.057462 | 0.10621
2019-04-09 15:33 | 1554838391.7088 | 433 | 0.20937 | 0.000549 | 4.5e-05 | 101.236 | 101.5556 | 115.0917 | 20.688 | 0.002912 | 0.002804 | 128.9306 | 0.25454 | 8.4015 | 0.002933 | 0.003996 | 0.003891 | 4.1e-05 | 0.001153 | 0.27671 | 1.3396 | 2.5476 | 98.43 | 20.03 | 48.486 | 91.3311 | 91.8261 | 88.7188 | 0.01155 | 0.97562 | 0.022727 | 23.2955 | -0.02819 | 0.15143 | 0.000804 | 0.052188 | 0.052188 | 0.11693
2019-04-09 15:33 | 1554838392.7076 | 434 | 0.20935 | 0.000546 | 4.5e-05 | 101.3173 | 101.4348 | 115.2257 | 20.6874 | 0.002993 | 0.002772 | 128.9128 | 0.25523 | 8.4047 | 0.00299 | 0.004623 | 0.00453 | 4.3e-05 | 0.001154 | 0.28119 | 1.3423 | 2.5477 | 98.43 | 20.03 | 48.485 | 93.2748 | 93.2082 | 87.9534 | 0.01155 | 0.97769 | 0.022765 | 23.2844 | -0.010364 | 0.15184 | 0.000934 | 0.061189 | 0.061189 | 0.1
2019-04-09 15:33 | 1554838393.7072 | 435 | 0.20939 | 0.000546 | 4.4e-05 | 101.2268 | 101.7053 | 115.1668 | 20.6889 | 0.002936 | 0.002744 | 128.8921 | 0.25533 | 8.4018 | 0.002953 | 0.004575 | 0.004521 | 4.2e-05 | 0.001147 | 0.27295 | 1.3281 | 2.5476 | 98.43 | 20.03 | 48.483 | 91.9058 | 92.3148 | 87.2725 | 0.01155 | 0.97953 | 0.02278 | 23.2565 | -0.042181 | 0.15189 | 0.000738 | 0.04788 | 0.04788 | 0.12785
2019-04-09 15:33 | 1554838394.7082 | 436 | 0.2094 | 0.000541 | 4.3e-05 | 101.217 | 101.5263 | 115.0929 | 20.6887 | 0.002986 | 0.0028 | 128.8895 | 0.25477 | 8.4001 | 0.00289 | 0.00379 | 0.003673 | 4.4e-05 | 0.001136 | 0.28838 | 1.3534 | 2.5475 | 98.43 | 20.03 | 48.483 | 93.1147 | 90.7947 | 88.6155 | 0.01155 | 0.9759 | 0.022731 | 23.2923 | -0.027297 | 0.15156 | 0.000638 | 0.041013 | 0.041013 | 0.14892
2019-04-09 15:33 | 1554838395.7083 | 437 | 0.20937 | 0.000539 | 4.4e-05 | 101.2828 | 101.7497 | 114.8143 | 20.6892 | 0.003 | 0.002789 | 128.8774 | 0.25464 | 8.4044 | 0.002956 | 0.004931 | 0.004862 | 4.1e-05 | 0.001129 | 0.27151 | 1.3445 | 2.5477 | 98.43 | 20.03 | 48.481 | 93.4408 | 92.3817 | 88.3466 | 0.011549 | 0.97662 | 0.022712 | 23.2554 | -0.041139 | 0.15148 | 0.000819 | 0.053258 | 0.053258 | 0.11463
2019-04-09 15:33 | 1554838396.708 | 438 | 0.20941 | 0.000543 | 4.4e-05 | 101.301 | 101.6038 | 114.9366 | 20.6887 | 0.002921 | 0.002809 | 128.839 | 0.25443 | 8.4023 | 0.002958 | 0.004613 | 0.004447 | 4.1e-05 | 0.001129 | 0.28816 | 1.3355 | 2.5476 | 98.43 | 20.03 | 48.481 | 91.5248 | 92.4344 | 88.8334 | 0.011549 | 0.97531 | 0.022709 | 23.2835 | -0.026698 | 0.15136 | 0.000632 | 0.040527 | 0.040527 | 0.15051
2019-04-09 15:33 | 1554838397.7077 | 439 | 0.2094 | 0.00054 | 4.3e-05 | 101.2116 | 101.6396 | 114.957 | 20.6882 | 0.002937 | 0.00281 | 128.8467 | 0.25383 | 8.4015 | 0.002831 | 0.004457 | 0.004447 | 4e-05 | 0.001132 | 0.2792 | 1.383 | 2.5477 | 98.43 | 20.03 | 48.482 | 91.9186 | 89.3759 | 88.8681 | 0.01155 | 0.97522 | 0.022709 | 23.2867 | -0.03774 | 0.15101 | 0.000693 | 0.044733 | 0.044733 | 0.13604
2019-04-09 15:33 | 1554838398.7744 | 440 | 0.20938 | 0.000547 | 4.3e-05 | 101.2435 | 101.5663 | 115.0499 | 20.6893 | 0.002879 | 0.002781 | 119.9371 | 0.25402 | 8.4019 | 0.002859 | 0.004356 | 0.004285 | 4.1e-05 | 0.001127 | 0.28368 | 1.3449 | 2.5476 | 98.43 | 20.03 | 48.483 | 90.5373 | 90.0397 | 88.1755 | 0.01155 | 0.97708 | 0.02274 | 23.2738 | -0.028471 | 0.15112 | 0.000796 | 0.051773 | 0.051773 | 0.11763
2019-04-09 15:33 | 1554838399.7207 | 441 | 0.20938 | 0.000544 | 4.3e-05 | 101.2527 | 101.5549 | 115.4515 | 20.6886 | 0.002904 | 0.002724 | 91.6239 | 0.25445 | 8.4035 | 0.002874 | 0.00415 | 0.004107 | 4.1e-05 | 0.001128 | 0.28224 | 1.3267 | 2.5476 | 98.43 | 20.03 | 48.482 | 91.1356 | 90.4116 | 86.8037 | 0.01155 | 0.98081 | 0.022823 | 23.2701 | -0.026656 | 0.15137 | 0.000799 | 0.052193 | 0.052193 | 0.11688
2019-04-09 15:33 | 1554838400.7198 | 442 | 0.20938 | 0.000538 | 4.3e-05 | 101.2295 | 101.6705 | 115.9351 | 20.689 | 0.002975 | 0.00277 | 63.2923 | 0.25516 | 8.3993 | 0.002844 | 0.003767 | 0.00367 | 4.2e-05 | 0.001135 | 0.27704 | 1.3661 | 2.5477 | 98.43 | 20.03 | 48.481 | 92.8453 | 89.6721 | 87.8951 | 0.011549 | 0.97784 | 0.022837 | 23.3541 | -0.038887 | 0.15179 | 0.000761 | 0.049681 | 0.049681 | 0.12313
2019-04-09 15:33 | 1554838401.7202 | 443 | 0.2094 | 0.000541 | 4.2e-05 | 101.2837 | 101.6382 | 116.2339 | 20.6894 | 0.002891 | 0.002758 | 34.9575 | 0.2545 | 8.4002 | 0.00289 | 0.003949 | 0.003856 | 4.6e-05 | 0.001135 | 0.28512 | 1.3178 | 2.5476 | 98.43 | 20.02 | 48.478 | 90.8031 | 90.7923 | 87.6182 | 0.011541 | 0.97859 | 0.022875 | 23.3752 | -0.031253 | 0.1514 | 0.000678 | 0.04406 | 0.04406 | 0.13848
2019-04-09 15:33 | 1554838402.7198 | 444 | 0.20944 | 0.000538 | 4e-05 | 101.2599 | 101.6389 | 116.4491 | 20.6885 | 0.002961 | 0.002769 | 6.6141 | 0.25515 | 8.4023 | 0.002887 | 0.004285 | 0.00423 | 4.7e-05 | 0.001133 | 0.27577 | 1.3175 | 2.5475 | 98.43 | 20.02 | 48.476 | 92.5117 | 90.7204 | 87.8749 | 0.011541 | 0.9779 | 0.022888 | 23.4051 | -0.033417 | 0.15179 | 0.000453 | 0.028747 | 0.028747 | 0.21279
2019-04-09 15:33 | 1554838403.7209 | 445 | 0.2094 | 0.00054 | 3.9e-05 | 101.197 | 101.7493 | 116.6672 | 20.6886 | 0.00289 | 0.002694 | -21.7249 | 0.24758 | 8.4026 | 0.002842 | 0.00488 | 0.004784 | 4.3e-05 | 0.001131 | 0.25962 | 1.3408 | 2.5476 | 98.43 | 20.02 | 48.474 | 90.787 | 89.6443 | 86.0818 | 0.01154 | 0.98278 | 0.022966 | 23.3688 | -0.048686 | 0.14729 | 0.000669 | 0.043843 | 0.043843 | 0.13538
2019-04-09 15:33 | 1554838404.7208 | 446 | 0.20943 | 0.00054 | 3.6e-05 | 101.1783 | 101.6206 | 117.1343 | 20.6897 | 0.002871 | 0.002756 | -50.0469 | 0.23254 | 8.4025 | 0.002873 | 0.004619 | 0.004442 | 4.5e-05 | 0.00113 | 0.28639 | 1.3888 | 2.5476 | 98.43 | 20.02 | 48.468 | 90.3214 | 90.3792 | 87.5621 | 0.011539 | 0.97874 | 0.022965 | 23.4637 | -0.039022 | 0.13834 | 0.000511 | 0.033106 | 0.033106 | 0.1684
2019-04-09 15:33 | 1554838405.7209 | 447 | 0.20947 | 0.000539 | 3.3e-05 | 101.1491 | 101.7638 | 117.4814 | 20.6884 | 0.002907 | 0.002722 | -78.3889 | 0.23213 | 8.4037 | 0.002842 | 0.004569 | 0.004355 | 4.2e-05 | 0.001134 | 0.29479 | 1.3809 | 2.5477 | 98.43 | 20.02 | 48.465 | 91.2103 | 89.6417 | 86.7395 | 0.011538 | 0.98098 | 0.023025 | 23.4717 | -0.054195 | 0.13809 | 0.000283 | 0.017592 | 0.017592 | 0.31634
2019-04-09 15:33 | 1554838406.7209 | 448 | 0.20946 | 0.000531 | 3e-05 | 101.1434 | 101.4978 | 118.0187 | 20.6884 | 0.002805 | 0.00268 | -106.7077 | 0.23174 | 8.4023 | 0.002862 | 0.003935 | 0.00371 | 4.5e-05 | 0.001134 | 0.29142 | 1.3599 | 2.5477 | 98.43 | 20.02 | 48.462 | 88.7339 | 90.113 | 85.7365 | 0.011538 | 0.98372 | 0.02311 | 23.4925 | -0.031282 | 0.13786 | 0.000338 | 0.021652 | 0.021652 | 0.2566
2019-04-09 15:33 | 1554838407.7201 | 449 | 0.20946 | 0.000532 | 2.7e-05 | 101.1119 | 101.5656 | 118.7546 | 20.688 | 0.002802 | 0.002677 | -135.0323 | 0.23258 | 8.4016 | 0.002881 | 0.004183 | 0.004073 | 4.4e-05 | 0.001126 | 0.26001 | 1.3163 | 2.5474 | 98.43 | 20.02 | 48.46 | 88.6642 | 90.5699 | 85.6558 | 0.011537 | 0.98394 | 0.023185 | 23.5629 | -0.040047 | 0.13836 | 0.000331 | 0.021387 | 0.021387 | 0.26072
2019-04-09 15:33 | 1554838408.7204 | 450 | 0.20947 | 0.000529 | 2.5e-05 | 101.1095 | 101.5676 | 119.2548 | 20.6879 | 0.002839 | 0.002652 | -154.481 | 0.23267 | 8.398 | 0.002897 | 0.00483 | 0.004606 | 4.6e-05 | 0.00113 | 0.29286 | 1.3336 | 2.5476 | 98.43 | 20.02 | 48.457 | 89.5625 | 90.955 | 85.0681 | 0.011536 | 0.98556 | 0.023252 | 23.5932 | -0.040439 | 0.13841 | 0.000279 | 0.018013 | 0.018013 | 0.30967
2019-04-09 15:33 | 1554838409.7191 | 451 | 0.20946 | 0.000523 | 2.2e-05 | 101.1678 | 101.5232 | 119.1761 | 20.689 | 0.002872 | 0.002649 | -154.5059 | 0.23179 | 8.3994 | 0.002899 | 0.00424 | 0.004078 | 4.4e-05 | 0.00114 | 0.2677 | 1.3454 | 2.5476 | 98.43 | 20.02 | 48.452 | 90.3591 | 91.0148 | 84.9872 | 0.011535 | 0.98578 | 0.023247 | 23.5827 | -0.031367 | 0.13789 | 0.000348 | 0.022917 | 0.022917 | 0.24249
2019-04-09 15:34 | 1554838410.7205 | 452 | 0.2095 | 0.000523 | 1.9e-05 | 101.1494 | 101.4496 | 119.1552 | 20.6888 | 0.002852 | 0.002646 | -154.4968 | 0.23208 | 8.3996 | 0.002885 | 0.004746 | 0.004553 | 4.7e-05 | 0.001143 | 0.2896 | 1.393 | 2.5477 | 98.43 | 20.02 | 48.45 | 89.8757 | 90.6675 | 84.9244 | 0.011535 | 0.98596 | 0.023247 | 23.5786 | -0.026508 | 0.13807 | 0.000141 | 0.008678 | 0.008678 | 0.64119
2019-04-09 15:34 | 1554838411.7201 | 453 | 0.20952 | 0.000515 | 1.7e-05 | 101.1631 | 101.4365 | 119.3085 | 20.6879 | 0.002848 | 0.002647 | -154.52 | 0.23193 | 8.4027 | 0.002936 | 0.004418 | 0.004279 | 4.6e-05 | 0.001134 | 0.28639 | 1.362 | 2.5476 | 98.43 | 20.02 | 48.448 | 89.7735 | 91.908 | 84.9361 | 0.011534 | 0.98592 | 0.023262 | 23.5941 | -0.024144 | 0.13798 | 6e-06 | -0.000563 | -0.000563 | -9.8817
2019-04-09 15:34 | 1554838412.7192 | 454 | 0.2095 | 0.000506 | 1.5e-05 | 101.0717 | 101.6323 | 119.4897 | 20.6896 | 0.002791 | 0.002638 | -154.5187 | 0.23175 | 8.4031 | 0.00292 | 0.003642 | 0.003518 | 4.4e-05 | 0.001137 | 0.27599 | 1.3689 | 2.5477 | 98.43 | 20.02 | 48.445 | 88.3947 | 91.5031 | 84.7294 | 0.011534 | 0.98649 | 0.023286 | 23.6052 | -0.049482 | 0.13787 | 0.000111 | 0.006914 | 0.006914 | 0.80355
2019-04-09 15:34 | 1554838413.7201 | 455 | 0.20952 | 0.000504 | 1.2e-05 | 101.1571 | 101.6279 | 119.7294 | 20.689 | 0.002796 | 0.002654 | -154.4999 | 0.23239 | 8.4004 | 0.002915 | 0.004233 | 0.004069 | 4.5e-05 | 0.001134 | 0.28926 | 1.3398 | 2.5477 | 98.43 | 20.02 | 48.443 | 88.5246 | 91.3993 | 85.1095 | 0.011533 | 0.98545 | 0.023297 | 23.6414 | -0.041528 | 0.13825 | 5e-06 | -0.000375 | -0.000375 | -14.867
2019-04-09 15:34 | 1554838414.7206 | 456 | 0.20956 | 0.000498 | 1.1e-05 | 101.1279 | 101.569 | 119.8923 | 20.6892 | 0.002738 | 0.002648 | -154.5079 | 0.23275 | 8.3999 | 0.002821 | 0.004295 | 0.004158 | 4.4e-05 | 0.001133 | 0.28031 | 1.3101 | 2.5476 | 98.43 | 20.02 | 48.437 | 87.1407 | 89.1338 | 84.9709 | 0.011532 | 0.98583 | 0.023318 | 23.6529 | -0.038928 | 0.13846 | -0.000221 | -0.016101 | -0.016101 | -0.34656
2019-04-09 15:34 | 1554838415.7206 | 457 | 0.20955 | 0.000496 | 9e-06 | 101.1154 | 101.7507 | 120.1535 | 20.6892 | 0.002762 | 0.002588 | -154.4997 | 0.23382 | 8.402 | 0.002897 | 0.004159 | 0.003983 | 4.4e-05 | 0.001137 | 0.28766 | 1.3417 | 2.5476 | 98.43 | 20.02 | 48.434 | 87.7029 | 90.9704 | 83.5141 | 0.011531 | 0.98985 | 0.023391 | 23.6305 | -0.056022 | 0.1391 | -0.000124 | -0.009283 | -0.009283 | -0.6039
2019-04-09 15:34 | 1554838416.7203 | 458 | 0.20949 | 0.000491 | 9e-06 | 101.0353 | 101.8667 | 119.7709 | 20.6878 | 0.002753 | 0.002559 | -154.4775 | 0.23236 | 8.4023 | 0.002921 | 0.00451 | 0.004319 | 4.8e-05 | 0.001139 | 0.29479 | 1.3463 | 2.5476 | 98.43 | 20.02 | 48.43 | 87.4792 | 91.5257 | 82.8187 | 0.01153 | 0.99179 | 0.023376 | 23.5698 | -0.0733 | 0.13823 | 0.000208 | 0.014047 | 0.014047 | 0.39659
2019-04-09 15:34 | 1554838417.7198 | 459 | 0.20959 | 0.000486 | 7e-06 | 101.0886 | 101.706 | 119.6648 | 20.6881 | 0.002803 | 0.00259 | -154.5015 | 0.23215 | 8.4022 | 0.002985 | 0.004144 | 0.003972 | 4.7e-05 | 0.001141 | 0.29607 | 1.3473 | 2.5476 | 98.43 | 20.02 | 48.42 | 88.6828 | 93.0727 | 83.5811 | 0.011528 | 0.98967 | 0.023341 | 23.5846 | -0.054459 | 0.13811 | -0.000353 | -0.025107 | -0.025107 | -0.22168
2019-04-09 15:34 | 1554838418.7205 | 460 | 0.20934 | 0.000481 | 6e-06 | 101.0332 | 101.5473 | 119.6984 | 20.6888 | 0.002696 | 0.002535 | -154.4887 | 0.23221 | 8.4006 | 0.002924 | 0.003988 | 0.003884 | 4.8e-05 | 0.001142 | 0.27367 | 1.3419 | 2.5475 | 98.43 | 20.02 | 48.413 | 86.1288 | 91.6211 | 82.2557 | 0.011526 | 0.99336 | 0.023388 | 23.544 | -0.045398 | 0.13814 | 0.001162 | 0.081052 | 0.081052 | 0.068686
2019-04-09 15:34 | 1554838419.7194 | 461 | 0.20957 | 0.000479 | 5e-06 | 101.0621 | 101.5814 | 119.7491 | 20.6888 | 0.002722 | 0.002512 | -154.4943 | 0.23308 | 8.3993 | 0.002915 | 0.003921 | 0.003732 | 4.7e-05 | 0.001144 | 0.27417 | 1.3362 | 2.5474 | 98.43 | 20.02 | 48.402 | 86.7556 | 91.3984 | 81.6847 | 0.011523 | 0.99496 | 0.023411 | 23.5301 | -0.045841 | 0.13866 | -0.000236 | -0.016883 | -0.016883 | -0.33099
2019-04-09 15:34 | 1554838420.7204 | 462 | 0.20954 | 0.000473 | 5e-06 | 100.9969 | 101.6244 | 119.7944 | 20.6877 | 0.002736 | 0.002528 | -154.4936 | 0.23211 | 8.4034 | 0.002842 | 0.004678 | 0.004544 | 4.4e-05 | 0.001146 | 0.29645 | 1.3418 | 2.5475 | 98.43 | 20.03 | 48.397 | 87.0885 | 89.6246 | 82.0735 | 0.011529 | 0.99387 | 0.023403 | 23.5474 | -0.055402 | 0.13808 | -4.6e-05 | -0.003525 | -0.003525 | -1.5786
2019-04-09 15:34 | 1554838421.7194 | 463 | 0.20955 | 0.000472 | 4e-06 | 100.9058 | 101.6654 | 119.7348 | 20.6881 | 0.002688 | 0.002559 | -154.4662 | 0.23308 | 8.4027 | 0.002876 | 0.004363 | 0.00422 | 4.5e-05 | 0.001147 | 0.26305 | 1.3628 | 2.5475 | 98.43 | 20.02 | 48.391 | 85.9258 | 90.448 | 82.8189 | 0.011521 | 0.99179 | 0.023373 | 23.5663 | -0.06708 | 0.13866 | -0.000112 | -0.00804 | -0.00804 | -0.69501
2019-04-09 15:34 | 1554838422.7204 | 464 | 0.20955 | 0.00047 | 4e-06 | 100.9189 | 101.6103 | 119.7662 | 20.6881 | 0.002706 | 0.002558 | -154.4705 | 0.23306 | 8.402 | 0.002818 | 0.003868 | 0.003606 | 4.6e-05 | 0.00114 | 0.29822 | 1.3443 | 2.5476 | 98.43 | 20.02 | 48.378 | 86.366 | 89.0499 | 82.7922 | 0.011518 | 0.99186 | 0.023377 | 23.5685 | -0.061073 | 0.13865 | -0.000101 | -0.007295 | -0.007295 | -0.76591
2019-04-09 15:34 | 1554838423.7188 | 465 | 0.20955 | 0.00047 | 3e-06 | 100.8977 | 101.538 | 119.751 | 20.688 | 0.002625 | 0.002478 | -154.4694 | 0.23241 | 8.4022 | 0.002806 | 0.003996 | 0.003891 | 4.5e-05 | 0.001142 | 0.29358 | 1.3625 | 2.5476 | 98.43 | 20.02 | 48.374 | 84.4032 | 88.7586 | 80.8776 | 0.011517 | 0.99723 | 0.023438 | 23.5035 | -0.056583 | 0.13826 | -7.1e-05 | -0.00517 | -0.00517 | -1.0778
2019-04-09 15:34 | 1554838424.7195 | 466 | 0.20954 | 0.000463 | 3e-06 | 100.8802 | 101.5435 | 119.5791 | 20.6885 | 0.002664 | 0.002504 | -154.465 | 0.23308 | 8.3993 | 0.002858 | 0.004148 | 0.003938 | 4.8e-05 | 0.001137 | 0.28351 | 1.3767 | 2.5476 | 98.43 | 20.02 | 48.369 | 85.3448 | 90.0264 | 81.5121 | 0.011515 | 0.99544 | 0.023401 | 23.5077 | -0.058625 | 0.13866 | -4.7e-05 | -0.003439 | -0.003439 | -1.6247
2019-04-09 15:34 | 1554838425.7196 | 467 | 0.20953 | 0.000464 | 3e-06 | 100.8543 | 101.5969 | 119.4552 | 20.6878 | 0.002671 | 0.002455 | -154.4669 | 0.23248 | 8.4041 | 0.002859 | 0.00399 | 0.003869 | 4.7e-05 | 0.001132 | 0.27793 | 1.3251 | 2.5475 | 98.43 | 20.02 | 48.37 | 85.5245 | 90.0484 | 80.3167 | 0.011516 | 0.99881 | 0.023428 | 23.4559 | -0.06562 | 0.1383 | 1.6e-05 | 0.000973 | 0.000973 | 5.7268
2019-04-09 15:34 | 1554838426.7191 | 468 | 0.20955 | 0.000462 | 2e-06 | 100.8733 | 101.5721 | 119.5543 | 20.6882 | 0.002605 | 0.002477 | -154.4707 | 0.23146 | 8.4012 | 0.002775 | 0.003947 | 0.00377 | 4.6e-05 | 0.00113 | 0.28528 | 1.3941 | 2.5476 | 98.43 | 20.02 | 48.372 | 83.9386 | 88.0238 | 80.8545 | 0.011516 | 0.99729 | 0.02342 | 23.4834 | -0.061749 | 0.13769 | -9.4e-05 | -0.00674 | -0.00674 | -0.82326
2019-04-09 15:34 | 1554838427.7193 | 469 | 0.20956 | 0.00046 | 2e-06 | 100.7908 | 101.4434 | 119.616 | 20.6887 | 0.002661 | 0.002473 | -154.463 | 0.23202 | 8.4013 | 0.002803 | 0.004143 | 0.003897 | 4.8e-05 | 0.001128 | 0.28512 | 1.4001 | 2.5474 | 98.43 | 20.02 | 48.378 | 85.2873 | 88.6875 | 80.7591 | 0.011518 | 0.99756 | 0.023429 | 23.4863 | -0.057724 | 0.13803 | -0.000125 | -0.008904 | -0.008904 | -0.62468
2019-04-09 15:34 | 1554838428.7186 | 470 | 0.20954 | 0.000455 | 2e-06 | 100.7185 | 101.262 | 119.6451 | 20.6895 | 0.002578 | 0.002481 | -154.4692 | 0.23288 | 8.4017 | 0.002816 | 0.003996 | 0.003909 | 4.5e-05 | 0.001125 | 0.28838 | 1.3567 | 2.5477 | 98.43 | 20.02 | 48.382 | 83.2836 | 89.0131 | 80.9554 | 0.011519 | 0.99701 | 0.023425 | 23.4957 | -0.048142 | 0.13854 | -1.6e-05 | -0.001224 | -0.001224 | -4.5625
2019-04-09 15:34 | 1554838429.7186 | 471 | 0.20957 | 0.000464 | 2e-06 | 100.6753 | 101.2486 | 119.5064 | 20.6877 | 0.002635 | 0.002481 | -154.4649 | 0.23198 | 8.4008 | 0.002779 | 0.004321 | 0.00416 | 4.7e-05 | 0.001124 | 0.29679 | 1.3751 | 2.5476 | 98.43 | 20.02 | 48.387 | 84.6535 | 88.1273 | 80.9381 | 0.01152 | 0.99706 | 0.023412 | 23.4815 | -0.050792 | 0.13801 | -0.000223 | -0.015728 | -0.015728 | -0.35362
2019-04-09 15:34 | 1554838430.7188 | 472 | 0.20957 | 0.000458 | 2e-06 | 100.6961 | 101.3343 | 119.5282 | 20.6883 | 0.002633 | 0.002429 | -154.4794 | 0.2313 | 8.4025 | 0.002792 | 0.00372 | 0.003598 | 4.6e-05 | 0.001124 | 0.28296 | 1.3584 | 2.5476 | 98.43 | 20.01 | 48.399 | 84.5963 | 88.4317 | 79.688 | 0.011515 | 1.0006 | 0.023456 | 23.4421 | -0.056507 | 0.1376 | -0.000188 | -0.013344 | -0.013344 | -0.41555
2019-04-09 15:34 | 1554838431.7197 | 473 | 0.20957 | 0.000451 | 1e-06 | 100.6407 | 101.1577 | 119.4681 | 20.6886 | 0.002574 | 0.002387 | -154.4954 | 0.23225 | 8.4024 | 0.002764 | 0.004377 | 0.004139 | 4.7e-05 | 0.001126 | 0.31454 | 1.3329 | 2.5475 | 98.43 | 20.01 | 48.405 | 83.1781 | 87.7437 | 78.6841 | 0.011517 | 1.0034 | 0.023483 | 23.4029 | -0.045831 | 0.13817 | -0.000182 | -0.01284 | -0.01284 | -0.43364
2019-04-09 15:34 | 1554838432.7187 | 474 | 0.20955 | 0.000453 | 1e-06 | 100.5696 | 101.1065 | 119.1059 | 20.6883 | 0.002506 | 0.002436 | -154.5039 | 0.23243 | 8.4019 | 0.002767 | 0.003835 | 0.003717 | 4.6e-05 | 0.001135 | 0.26808 | 1.2975 | 2.5474 | 98.43 | 20.01 | 48.411 | 81.5564 | 87.8179 | 79.8777 | 0.011518 | 1.0001 | 0.023408 | 23.407 | -0.047624 | 0.13827 | -0.000101 | -0.007168 | -0.007168 | -0.77736
2019-04-09 15:34 | 1554838433.7187 | 475 | 0.20956 | 0.000455 | 1e-06 | 100.5041 | 101.1939 | 118.8125 | 20.6883 | 0.002552 | 0.002467 | -154.5049 | 0.23189 | 8.4014 | 0.002875 | 0.003883 | 0.00364 | 4.5e-05 | 0.001132 | 0.29502 | 1.3308 | 2.5476 | 98.43 | 20.01 | 48.417 | 82.6561 | 90.42 | 80.6238 | 0.01152 | 0.99794 | 0.023355 | 23.4028 | -0.061182 | 0.13795 | -0.000129 | -0.009077 | -0.009077 | -0.61251
2019-04-09 15:34 | 1554838434.7202 | 476 | 0.20953 | 0.000451 | 2e-06 | 100.5559 | 101.0865 | 118.7201 | 20.6878 | 0.002547 | 0.002409 | -154.4962 | 0.23233 | 8.4011 | 0.00287 | 0.003901 | 0.003763 | 4.5e-05 | 0.00114 | 0.2557 | 1.3523 | 2.5475 | 98.43 | 20.01 | 48.427 | 82.5424 | 90.2953 | 79.2151 | 0.011522 | 1.0019 | 0.023392 | 23.3471 | -0.047076 | 0.13821 | 4e-05 | 0.002685 | 0.002685 | 2.0742
2019-04-09 15:34 | 1554838435.7191 | 477 | 0.20955 | 0.000456 | 2e-06 | 100.3671 | 101.0776 | 118.8992 | 20.6884 | 0.002564 | 0.002356 | -154.5171 | 0.23275 | 8.4021 | 0.002855 | 0.003689 | 0.00356 | 4.2e-05 | 0.001133 | 0.2896 | 1.3782 | 2.5476 | 98.43 | 20.01 | 48.43 | 82.9526 | 89.9511 | 77.9489 | 0.011523 | 1.0055 | 0.023452 | 23.3227 | -0.063094 | 0.13846 | -5.1e-05 | -0.003678 | -0.003678 | -1.5173
2019-04-09 15:34 | 1554838436.719 | 478 | 0.20954 | 0.000453 | 2e-06 | 100.357 | 100.954 | 119.047 | 20.6884 | 0.00249 | 0.002425 | -154.5235 | 0.23212 | 8.4023 | 0.002941 | 0.004566 | 0.004345 | 4.1e-05 | 0.001135 | 0.274 | 1.3498 | 2.5475 | 98.43 | 20.01 | 48.434 | 81.1556 | 92.0102 | 79.6103 | 0.011524 | 1.0008 | 0.023411 | 23.3923 | -0.053056 | 0.13809 | -2.9e-05 | -0.002114 | -0.002114 | -2.632
2019-04-09 15:34 | 1554838437.719 | 479 | 0.20956 | 0.000455 | 1e-06 | 100.2989 | 100.9423 | 118.929 | 20.6889 | 0.002576 | 0.002421 | -154.5112 | 0.23228 | 8.4011 | 0.00295 | 0.003577 | 0.003443 | 4.9e-05 | 0.001141 | 0.2808 | 1.3819 | 2.5477 | 98.43 | 20.01 | 48.434 | 83.244 | 92.2413 | 79.5132 | 0.011524 | 1.0011 | 0.023403 | 23.3775 | -0.057197 | 0.13819 | -0.00016 | -0.011264 | -0.011264 | -0.49441
2019-04-09 15:34 | 1554838438.72 | 480 | 0.20955 | 0.000453 | 2e-06 | 100.3564 | 100.9055 | 119.0444 | 20.6884 | 0.002494 | 0.002379 | -154.5186 | 0.23161 | 8.4011 | 0.00303 | 0.004485 | 0.004318 | 4.6e-05 | 0.001137 | 0.27041 | 1.3901 | 2.5476 | 98.43 | 20.01 | 48.434 | 81.2647 | 94.1645 | 78.497 | 0.011524 | 1.004 | 0.023448 | 23.3551 | -0.048809 | 0.13779 | -5.3e-05 | -0.003847 | -0.003847 | -1.4435
2019-04-09 15:34 | 1554838439.719 | 481 | 0.20957 | 0.000457 | 2e-06 | 100.245 | 100.7926 | 119.6418 | 20.6887 | 0.002561 | 0.002352 | -154.5202 | 0.23156 | 8.4025 | 0.002954 | 0.004332 | 0.004155 | 4.2e-05 | 0.001135 | 0.2891 | 1.3346 | 2.5476 | 98.43 | 20.01 | 48.435 | 82.8779 | 92.3201 | 77.8415 | 0.011524 | 1.0059 | 0.023529 | 23.3918 | -0.048729 | 0.13775 | -0.00018 | -0.01283 | -0.01283 | -0.43269
2019-04-09 15:34 | 1554838440.718 | 482 | 0.20957 | 0.000455 | 2e-06 | 100.1804 | 100.8391 | 119.893 | 20.6888 | 0.002509 | 0.002373 | -154.5182 | 0.23208 | 8.4017 | 0.00297 | 0.00462 | 0.004463 | 4.5e-05 | 0.001128 | 0.2792 | 1.3399 | 2.5475 | 98.43 | 20.01 | 48.436 | 81.6131 | 92.713 | 78.3571 | 0.011524 | 1.0044 | 0.023536 | 23.4336 | -0.058619 | 0.13806 | -0.000194 | -0.013805 | -0.013805 | -0.40303
2019-04-09 15:34 | 1554838441.7181 | 483 | 0.20958 | 0.00046 | 1e-06 | 100.1798 | 100.7902 | 120.2537 | 20.6883 | 0.002555 | 0.002393 | -154.5238 | 0.23211 | 8.4021 | 0.002939 | 0.004202 | 0.004031 | 4.4e-05 | 0.001128 | 0.27433 | 1.3734 | 2.5477 | 98.43 | 20.01 | 48.438 | 82.7395 | 91.9671 | 78.8366 | 0.011525 | 1.003 | 0.023555 | 23.4848 | -0.054337 | 0.13809 | -0.000286 | -0.020282 | -0.020282 | -0.27437
2019-04-09 15:34 | 1554838442.7181 | 484 | 0.20955 | 0.000449 | 2e-06 | 100.0726 | 100.7892 | 120.8389 | 20.6881 | 0.002513 | 0.002327 | -154.5194 | 0.23184 | 8.402 | 0.002927 | 0.003737 | 0.003634 | 4.6e-05 | 0.001132 | 0.27096 | 1.3844 | 2.5477 | 98.43 | 20.01 | 48.441 | 81.709 | 91.6722 | 77.2432 | 0.011525 | 1.0076 | 0.023666 | 23.4885 | -0.063822 | 0.13792 | -8.3e-05 | -0.006008 | -0.006008 | -0.92509
2019-04-09 15:34 | 1554838443.7185 | 485 | 0.20956 | 0.000454 | 1e-06 | 100.0481 | 100.604 | 121.2173 | 20.6884 | 0.002514 | 0.002368 | -154.5326 | 0.23235 | 8.402 | 0.002984 | 0.004136 | 0.003943 | 4.6e-05 | 0.001135 | 0.29175 | 1.3834 | 2.5477 | 98.44 | 20.01 | 48.442 | 81.7437 | 93.0552 | 78.224 | 0.011524 | 1.0048 | 0.02367 | 23.5582 | -0.049558 | 0.13822 | -0.000151 | -0.01079 | -0.01079 | -0.51624
2019-04-09 15:34 | 1554838444.7193 | 486 | 0.20957 | 0.000454 | 2e-06 | 99.9773 | 100.4209 | 121.6637 | 20.6878 | 0.00245 | 0.002432 | -154.5447 | 0.23193 | 8.4004 | 0.002999 | 0.004326 | 0.004146 | 4.7e-05 | 0.001137 | 0.27345 | 1.3817 | 2.5476 | 98.44 | 20.01 | 48.444 | 80.2069 | 93.405 | 79.7702 | 0.011525 | 1.0004 | 0.023662 | 23.6534 | -0.039603 | 0.13797 | -0.000203 | -0.014499 | -0.014499 | -0.38349
2019-04-09 15:34 | 1554838445.7182 | 487 | 0.20955 | 0.000451 | 1e-06 | 99.9642 | 100.34 | 121.7203 | 20.6884 | 0.002506 | 0.00238 | -154.5299 | 0.23165 | 8.3971 | 0.002899 | 0.004743 | 0.00459 | 4.5e-05 | 0.001137 | 0.28528 | 1.3575 | 2.5475 | 98.44 | 20.01 | 48.446 | 81.5557 | 90.9978 | 78.5186 | 0.011525 | 1.0039 | 0.023709 | 23.6169 | -0.033569 | 0.13781 | -7e-05 | -0.005049 | -0.005049 | -1.1
2019-04-09 15:34 | 1554838446.7182 | 488 | 0.20953 | 0.000451 | 1e-06 | 99.8477 | 100.4753 | 121.8989 | 20.6877 | 0.002493 | 0.002365 | -154.5216 | 0.23248 | 8.4039 | 0.002909 | 0.00423 | 0.003995 | 4.4e-05 | 0.001131 | 0.286 | 1.3334 | 2.5474 | 98.44 | 20.01 | 48.449 | 81.247 | 91.2574 | 78.1501 | 0.011526 | 1.005 | 0.023739 | 23.6218 | -0.056042 | 0.1383 | 1.8e-05 | 0.001219 | 0.001219 | 4.5736
2019-04-09 15:34 | 1554838447.7187 | 489 | 0.20954 | 0.000451 | 1e-06 | 99.8742 | 100.2846 | 122.1552 | 20.6882 | 0.002464 | 0.002366 | -154.524 | 0.23317 | 8.4031 | 0.002915 | 0.004537 | 0.004434 | 4.5e-05 | 0.001127 | 0.28854 | 1.3797 | 2.5475 | 98.44 | 20.01 | 48.454 | 80.5511 | 91.3798 | 78.177 | 0.011527 | 1.0049 | 0.023763 | 23.6476 | -0.036679 | 0.13871 | -2.4e-05 | -0.001777 | -0.001777 | -3.1465
2019-04-09 15:34 | 1554838448.7197 | 490 | 0.20954 | 0.000449 | 1e-06 | 99.6914 | 100.3514 | 122.1428 | 20.6877 | 0.002486 | 0.002308 | -154.5452 | 0.23234 | 8.3992 | 0.00286 | 0.004319 | 0.004055 | 4.9e-05 | 0.001125 | 0.28368 | 1.3228 | 2.5476 | 98.44 | 20.01 | 48.456 | 81.0575 | 90.0584 | 76.8007 | 0.011528 | 1.0088 | 0.023809 | 23.6 | -0.059022 | 0.13822 | 7e-06 | 0.000443 | 0.000443 | 12.5714
2019-04-09 15:34 | 1554838449.7183 | 491 | 0.20954 | 0.000449 | 1e-06 | 99.7697 | 100.2309 | 121.7403 | 20.6884 | 0.002476 | 0.002275 | -154.5334 | 0.23329 | 8.399 | 0.002856 | 0.004097 | 0.003969 | 4.7e-05 | 0.001133 | 0.28982 | 1.3223 | 2.5475 | 98.44 | 20.01 | 48.461 | 80.8345 | 89.9762 | 76.0002 | 0.011529 | 1.0112 | 0.023797 | 23.5341 | -0.041255 | 0.13878 | -2.9e-05 | -0.002165 | -0.002165 | -2.5834
2019-04-09 15:34 | 1554838450.7194 | 492 | 0.20955 | 0.000453 | 0 | 99.6539 | 100.3886 | 121.7954 | 20.6887 | 0.00243 | 0.0023 | -154.542 | 0.23289 | 8.4008 | 0.002933 | 0.003169 | 0.003072 | 4.7e-05 | 0.00114 | 0.26648 | 1.3168 | 2.5474 | 98.44 | 20.01 | 48.462 | 79.7283 | 91.829 | 76.6122 | 0.011529 | 1.0094 | 0.023781 | 23.5601 | -0.065702 | 0.13855 | -6.5e-05 | -0.004684 | -0.004684 | -1.192
2019-04-09 15:34 | 1554838451.7194 | 493 | 0.20957 | 0.000448 | 1e-06 | 99.6196 | 100.1869 | 121.7813 | 20.6874 | 0.002487 | 0.002393 | -154.5433 | 0.23341 | 8.402 | 0.002896 | 0.004022 | 0.003815 | 4.6e-05 | 0.001143 | 0.28031 | 1.3319 | 2.5475 | 98.44 | 20.01 | 48.463 | 81.0837 | 90.9444 | 78.8243 | 0.011529 | 1.003 | 0.023705 | 23.6331 | -0.050789 | 0.13886 | -0.000162 | -0.01156 | -0.01156 | -0.48406
2019-04-09 15:34 | 1554838452.7188 | 494 | 0.20951 | 0.000453 | 1e-06 | 99.5613 | 100.1683 | 121.6252 | 20.6882 | 0.002438 | 0.002344 | -154.5399 | 0.23369 | 8.3995 | 0.002901 | 0.004062 | 0.003905 | 4.8e-05 | 0.001143 | 0.2756 | 1.3115 | 2.5475 | 98.44 | 20.01 | 48.465 | 79.9265 | 91.0492 | 77.6548 | 0.01153 | 1.0064 | 0.023729 | 23.5787 | -0.054372 | 0.13902 | 0.000147 | 0.010411 | 0.010411 | 0.53814
2019-04-09 15:34 | 1554838453.7195 | 495 | 0.20946 | 0.000448 | 2e-06 | 99.4922 | 100.0055 | 121.8048 | 20.6877 | 0.00239 | 0.002335 | -154.5407 | 0.23313 | 8.4011 | 0.002902 | 0.003919 | 0.003742 | 4.5e-05 | 0.001137 | 0.28584 | 1.3663 | 2.5476 | 98.44 | 20.01 | 48.465 | 78.7676 | 91.087 | 77.4509 | 0.01153 | 1.007 | 0.023754 | 23.5892 | -0.046033 | 0.13869 | 0.000494 | 0.03505 | 0.03505 | 0.15946
2019-04-09 15:34 | 1554838454.7175 | 496 | 0.20939 | 0.000448 | 1e-06 | 99.5252 | 99.9697 | 121.4879 | 20.6874 | 0.002399 | 0.002291 | -154.5327 | 0.2326 | 8.401 | 0.002964 | 0.004082 | 0.003876 | 4.5e-05 | 0.001135 | 0.28512 | 1.3401 | 2.5474 | 98.44 | 20.01 | 48.465 | 78.9699 | 92.5785 | 76.392 | 0.01153 | 1.01 | 0.023759 | 23.5229 | -0.03986 | 0.13837 | 0.000902 | 0.064185 | 0.064185 | 0.086883
2019-04-09 15:34 | 1554838455.7188 | 497 | 0.20955 | 0.000453 | 1e-06 | 99.4716 | 100.1649 | 121.4302 | 20.6883 | 0.002375 | 0.002282 | -154.5409 | 0.23188 | 8.404 | 0.002891 | 0.004541 | 0.004335 | 4.5e-05 | 0.001133 | 0.27992 | 1.3403 | 2.5475 | 98.44 | 20.01 | 48.464 | 78.3917 | 90.8062 | 76.1793 | 0.01153 | 1.0106 | 0.02376 | 23.5101 | -0.062121 | 0.13795 | -4.4e-05 | -0.003182 | -0.003182 | -1.7472
2019-04-09 15:34 | 1554838456.7185 | 498 | 0.20954 | 0.000451 | 1e-06 | 99.4466 | 99.958 | 121.179 | 20.6883 | 0.00237 | 0.002294 | -154.5504 | 0.23296 | 8.4021 | 0.002869 | 0.003943 | 0.003809 | 4.3e-05 | 0.00113 | 0.28854 | 1.3468 | 2.5476 | 98.44 | 20.01 | 48.462 | 78.2816 | 90.2808 | 76.4592 | 0.011529 | 1.0098 | 0.023726 | 23.4952 | -0.045881 | 0.13859 | -1e-06 | -8.9e-05 | -8.9e-05 | -62.6884
2019-04-09 15:34 | 1554838457.7176 | 499 | 0.20954 | 0.000447 | 1e-06 | 99.41 | 99.9776 | 121.0891 | 20.687 | 0.00245 | 0.002295 | -154.5371 | 0.23224 | 8.4025 | 0.002917 | 0.003893 | 0.003687 | 4.8e-05 | 0.001134 | 0.30879 | 1.3411 | 2.5475 | 98.44 | 20.01 | 48.459 | 80.2106 | 91.4381 | 76.4899 | 0.011529 | 1.0097 | 0.023716 | 23.4875 | -0.050932 | 0.13816 | -1e-05 | -0.00078 | -0.00078 | -7.1405
2019-04-09 15:34 | 1554838458.718 | 500 | 0.20954 | 0.000445 | 1e-06 | 99.4389 | 99.8882 | 121.0027 | 20.6877 | 0.002409 | 0.002285 | -154.5194 | 0.2331 | 8.4007 | 0.002825 | 0.004044 | 0.003895 | 4.7e-05 | 0.001143 | 0.27096 | 1.3465 | 2.5476 | 98.44 | 20.01 | 48.458 | 79.2299 | 89.2232 | 76.2441 | 0.011528 | 1.0105 | 0.023716 | 23.4709 | -0.040324 | 0.13867 | -2.4e-05 | -0.001769 | -0.001769 | -3.1586
2019-04-09 15:34 | 1554838459.7176 | 501 | 0.20952 | 0.000449 | 1e-06 | 99.438 | 99.9628 | 121.1544 | 20.6883 | 0.002407 | 0.002271 | -154.5336 | 0.23346 | 8.4023 | 0.002827 | 0.003905 | 0.003741 | 4.6e-05 | 0.001144 | 0.29463 | 1.3903 | 2.5476 | 98.44 | 20.01 | 48.456 | 79.1704 | 89.2816 | 75.9072 | 0.011528 | 1.0114 | 0.023742 | 23.4743 | -0.047089 | 0.13889 | 8.2e-05 | 0.005752 | 0.005752 | 0.97312
2019-04-09 15:34 | 1554838460.7187 | 502 | 0.20952 | 0.000448 | 1e-06 | 99.4261 | 99.7939 | 121.1882 | 20.6879 | 0.002359 | 0.002257 | -154.5241 | 0.2333 | 8.4006 | 0.00289 | 0.00428 | 0.004106 | 4.4e-05 | 0.001136 | 0.28313 | 1.3303 | 2.5476 | 98.44 | 20.01 | 48.454 | 78.017 | 90.7909 | 75.5757 | 0.011527 | 1.0124 | 0.023757 | 23.4664 | -0.033028 | 0.13879 | 0.000101 | 0.007176 | 0.007176 | 0.77946
2019-04-09 15:34 | 1554838461.718 | 503 | 0.20955 | 0.000449 | 1e-06 | 99.3814 | 99.8923 | 121.0891 | 20.6873 | 0.002394 | 0.002276 | -154.5131 | 0.23273 | 8.4033 | 0.002872 | 0.004141 | 0.00403 | 4.2e-05 | 0.00113 | 0.29192 | 1.3727 | 2.5476 | 98.44 | 20.01 | 48.453 | 78.8694 | 90.3557 | 76.0342 | 0.011527 | 1.0111 | 0.023732 | 23.4722 | -0.045865 | 0.13845 | -4.5e-05 | -0.003208 | -0.003208 | -1.7394
2019-04-09 15:34 | 1554838462.7187 | 504 | 0.20953 | 0.00045 | 1e-06 | 99.3293 | 100.0203 | 121.0023 | 20.6876 | 0.002423 | 0.002258 | -154.5197 | 0.23245 | 8.401 | 0.002822 | 0.004254 | 0.004087 | 4.2e-05 | 0.001128 | 0.30215 | 1.3568 | 2.5476 | 98.43 | 20.01 | 48.45 | 79.5595 | 89.1621 | 75.5799 | 0.011528 | 1.0124 | 0.023739 | 23.4485 | -0.062013 | 0.13828 | 4.5e-05 | 0.003137 | 0.003137 | 1.7764
2019-04-09 15:34 | 1554838463.7192 | 505 | 0.20954 | 0.000452 | 1e-06 | 99.343 | 99.7367 | 120.8532 | 20.6882 | 0.002411 | 0.002296 | -154.5046 | 0.23284 | 8.402 | 0.00292 | 0.00423 | 0.004058 | 4.2e-05 | 0.00113 | 0.26178 | 1.3412 | 2.5477 | 98.43 | 20.01 | 48.449 | 79.2569 | 91.5237 | 76.5106 | 0.011527 | 1.0097 | 0.023692 | 23.4653 | -0.035383 | 0.13852 | 1.4e-05 | 0.000913 | 0.000913 | 6.1133
2019-04-09 15:34 | 1554838464.7178 | 506 | 0.20955 | 0.00045 | 1e-06 | 99.3224 | 99.8087 | 120.8452 | 20.6878 | 0.002411 | 0.0023 | -154.5096 | 0.23192 | 8.403 | 0.002878 | 0.004455 | 0.004201 | 4.5e-05 | 0.001138 | 0.27865 | 1.3666 | 2.5476 | 98.43 | 20.01 | 48.449 | 79.2696 | 90.5101 | 76.5935 | 0.011527 | 1.0094 | 0.023689 | 23.4674 | -0.043682 | 0.13797 | -8.3e-05 | -0.005912 | -0.005912 | -0.94047
2019-04-09 15:34 | 1554838465.7193 | 507 | 0.20957 | 0.000447 | 1e-06 | 99.301 | 99.894 | 120.5986 | 20.6879 | 0.002395 | 0.002265 | -154.4946 | 0.23218 | 8.4068 | 0.002838 | 0.004057 | 0.00399 | 3.9e-05 | 0.001138 | 0.26504 | 1.3344 | 2.5475 | 98.43 | 20 | 48.447 | 78.8783 | 89.5367 | 75.7511 | 0.01152 | 1.0119 | 0.023693 | 23.4151 | -0.053259 | 0.13812 | -0.000156 | -0.01115 | -0.01115 | -0.49921
2019-04-09 15:34 | 1554838466.7189 | 508 | 0.20951 | 0.00045 | 1e-06 | 99.248 | 99.7801 | 120.687 | 20.6884 | 0.002304 | 0.002284 | -154.4907 | 0.23213 | 8.4029 | 0.002883 | 0.004068 | 0.003914 | 4.5e-05 | 0.001131 | 0.26576 | 1.3396 | 2.5475 | 98.43 | 20 | 48.445 | 76.7008 | 90.6311 | 76.22 | 0.011519 | 1.0105 | 0.023686 | 23.4395 | -0.047829 | 0.1381 | 0.000197 | 0.013886 | 0.013886 | 0.40078
2019-04-09 15:34 | 1554838467.7189 | 509 | 0.20952 | 0.000449 | 1e-06 | 99.1979 | 99.7632 | 120.769 | 20.6879 | 0.002392 | 0.002265 | -154.527 | 0.23216 | 8.4023 | 0.002833 | 0.004187 | 0.004074 | 4.1e-05 | 0.001139 | 0.28799 | 1.3728 | 2.5477 | 98.43 | 20 | 48.445 | 78.8059 | 89.424 | 75.7556 | 0.011519 | 1.0119 | 0.02371 | 23.4318 | -0.050827 | 0.13811 | 9.9e-05 | 0.00699 | 0.00699 | 0.79632
2019-04-09 15:34 | 1554838468.719 | 510 | 0.20953 | 0.000451 | 2e-06 | 99.223 | 99.7832 | 120.9478 | 20.6875 | 0.002353 | 0.002265 | -154.5203 | 0.23298 | 8.4002 | 0.002868 | 0.004046 | 0.003833 | 4.3e-05 | 0.001136 | 0.28479 | 1.3671 | 2.5476 | 98.43 | 20 | 48.444 | 77.8811 | 90.2567 | 75.7608 | 0.011519 | 1.0118 | 0.023727 | 23.4493 | -0.050361 | 0.1386 | 3.7e-05 | 0.002541 | 0.002541 | 2.1986
2019-04-09 15:34 | 1554838469.7174 | 511 | 0.2095 | 0.000445 | 1e-06 | 99.2661 | 99.5505 | 120.9645 | 20.6871 | 0.00238 | 0.002212 | -154.5051 | 0.23193 | 8.4026 | 0.002781 | 0.004096 | 0.003916 | 4.2e-05 | 0.001134 | 0.28102 | 1.3823 | 2.5477 | 98.43 | 20 | 48.444 | 78.5284 | 88.1539 | 74.4796 | 0.011519 | 1.0156 | 0.023773 | 23.4079 | -0.02559 | 0.13798 | 0.000208 | 0.014766 | 0.014766 | 0.37658
2019-04-09 15:34 | 1554838470.717 | 512 | 0.20953 | 0.00045 | 1e-06 | 99.2352 | 99.6593 | 120.6619 | 20.6883 | 0.002375 | 0.002263 | -154.5082 | 0.23252 | 8.4019 | 0.002854 | 0.003936 | 0.003815 | 4.4e-05 | 0.001129 | 0.27992 | 1.385 | 2.5477 | 98.43 | 20 | 48.442 | 78.3948 | 89.9121 | 75.7015 | 0.011519 | 1.012 | 0.023701 | 23.4196 | -0.038148 | 0.13833 | 6.2e-05 | 0.004359 | 0.004359 | 1.2789
2019-04-09 15:35 | 1554838471.718 | 513 | 0.20954 | 0.000449 | 1e-06 | 99.1768 | 99.5378 | 120.5281 | 20.6875 | 0.002337 | 0.00227 | -154.5056 | 0.23161 | 8.4039 | 0.002887 | 0.004539 | 0.004382 | 4.5e-05 | 0.001126 | 0.27776 | 1.3872 | 2.5475 | 98.43 | 20 | 48.442 | 77.499 | 90.7187 | 75.8667 | 0.011519 | 1.0115 | 0.023682 | 23.4122 | -0.0325 | 0.13778 | -2.7e-05 | -0.001954 | -0.001954 | -2.8418
2019-04-09 15:35 | 1554838472.717 | 514 | 0.20954 | 0.000449 | 1e-06 | 99.2271 | 99.7047 | 120.376 | 20.6863 | 0.002404 | 0.002261 | -154.516 | 0.23325 | 8.4019 | 0.002817 | 0.003652 | 0.00349 | 4.3e-05 | 0.001129 | 0.28174 | 1.4064 | 2.5476 | 98.43 | 20 | 48.443 | 79.107 | 89.0319 | 75.6558 | 0.011519 | 1.0122 | 0.023675 | 23.3903 | -0.042949 | 0.13876 | -3e-06 | -0.000234 | -0.000234 | -23.8613
2019-04-09 15:35 | 1554838473.7181 | 515 | 0.20952 | 0.000452 | 1e-06 | 99.2331 | 99.574 | 120.3892 | 20.6879 | 0.002331 | 0.002242 | -154.5201 | 0.23239 | 8.4008 | 0.00295 | 0.004391 | 0.00425 | 4.6e-05 | 0.001135 | 0.26985 | 1.4049 | 2.5478 | 98.43 | 20 | 48.443 | 77.3543 | 92.2303 | 75.1992 | 0.011519 | 1.0135 | 0.023691 | 23.3763 | -0.030672 | 0.13825 | 9.3e-05 | 0.006529 | 0.006529 | 0.85329
2019-04-09 15:35 | 1554838474.7181 | 516 | 0.20957 | 0.000449 | 1e-06 | 99.2316 | 99.8675 | 120.6058 | 20.6873 | 0.002294 | 0.002288 | -154.522 | 0.23281 | 8.403 | 0.0028 | 0.0041 | 0.003915 | 4.4e-05 | 0.001141 | 0.28368 | 1.3831 | 2.5478 | 98.43 | 20 | 48.444 | 76.452 | 88.6243 | 76.3218 | 0.011519 | 1.0102 | 0.023675 | 23.435 | -0.057138 | 0.1385 | -0.000158 | -0.011326 | -0.011326 | -0.49281
2019-04-09 15:35 | 1554838475.7186 | 517 | 0.20947 | 0.000446 | 1e-06 | 99.2316 | 99.5189 | 120.9671 | 20.6872 | 0.002303 | 0.002261 | -154.525 | 0.23291 | 8.4022 | 0.002829 | 0.003802 | 0.003649 | 4.4e-05 | 0.001143 | 0.27151 | 1.3861 | 2.5476 | 98.43 | 20 | 48.441 | 76.6662 | 89.3272 | 75.656 | 0.011518 | 1.0122 | 0.023733 | 23.4477 | -0.025858 | 0.13856 | 0.000431 | 0.030583 | 0.030583 | 0.18258
2019-04-09 15:35 | 1554838476.7183 | 518 | 0.20954 | 0.000446 | 1e-06 | 99.2337 | 99.638 | 121.1417 | 20.6869 | 0.002339 | 0.002243 | -154.5211 | 0.23267 | 8.4028 | 0.002846 | 0.004377 | 0.004199 | 4.5e-05 | 0.001148 | 0.27007 | 1.3586 | 2.5476 | 98.43 | 20 | 48.439 | 77.5318 | 89.7286 | 75.2333 | 0.011518 | 1.0134 | 0.023764 | 23.4504 | -0.036367 | 0.13842 | -4e-06 | -0.000323 | -0.000323 | -17.2681
2019-04-09 15:35 | 1554838477.7182 | 519 | 0.20951 | 0.000445 | 1e-06 | 99.2179 | 99.8207 | 121.0044 | 20.6875 | 0.002344 | 0.002201 | -154.5037 | 0.23271 | 8.4022 | 0.002817 | 0.004011 | 0.003839 | 4.3e-05 | 0.001155 | 0.29175 | 1.3546 | 2.5477 | 98.43 | 20 | 48.437 | 77.6494 | 89.0323 | 74.2306 | 0.011517 | 1.0163 | 0.023785 | 23.4033 | -0.054179 | 0.13844 | 0.000172 | 0.012207 | 0.012207 | 0.45703
2019-04-09 15:35 | 1554838478.7183 | 520 | 0.20955 | 0.000449 | 1e-06 | 99.1726 | 99.8118 | 120.8269 | 20.687 | 0.002345 | 0.002193 | -154.515 | 0.23317 | 8.4027 | 0.002843 | 0.004522 | 0.004367 | 4.5e-05 | 0.001151 | 0.26753 | 1.3681 | 2.5477 | 98.44 | 20 | 48.435 | 77.6778 | 89.6647 | 74.0235 | 0.011516 | 1.0169 | 0.023775 | 23.3792 | -0.057459 | 0.13871 | -4.5e-05 | -0.003306 | -0.003306 | -1.6907
2019-04-09 15:35 | 1554838479.7173 | 521 | 0.20953 | 0.000448 | 1e-06 | 99.2575 | 99.7581 | 120.7575 | 20.6878 | 0.002291 | 0.002277 | -154.5162 | 0.2325 | 8.3993 | 0.002867 | 0.003932 | 0.003787 | 4.4e-05 | 0.001154 | 0.30215 | 1.3549 | 2.5477 | 98.44 | 20 | 48.433 | 76.3855 | 90.2389 | 76.0362 | 0.011515 | 1.0111 | 0.023699 | 23.4401 | -0.044995 | 0.13831 | 8.3e-05 | 0.005826 | 0.005826 | 0.95676
2019-04-09 15:35 | 1554838480.718 | 522 | 0.20951 | 0.000451 | 1e-06 | 99.3043 | 99.9295 | 120.6042 | 20.6874 | 0.002304 | 0.002218 | -154.5237 | 0.23363 | 8.4033 | 0.00286 | 0.003905 | 0.003726 | 4.3e-05 | 0.001155 | 0.25586 | 1.3934 | 2.5476 | 98.44 | 20 | 48.429 | 76.6998 | 90.065 | 74.6178 | 0.011514 | 1.0152 | 0.023732 | 23.3776 | -0.056137 | 0.13899 | 0.000162 | 0.011437 | 0.011437 | 0.48974
2019-04-09 15:35 | 1554838481.7183 | 523 | 0.20952 | 0.000447 | 1e-06 | 99.2801 | 99.808 | 120.5413 | 20.6878 | 0.002282 | 0.002189 | -154.5191 | 0.23248 | 8.4017 | 0.002876 | 0.004749 | 0.004625 | 4.6e-05 | 0.00115 | 0.25785 | 1.3624 | 2.5475 | 98.44 | 20 | 48.428 | 76.1678 | 90.4617 | 73.9402 | 0.011514 | 1.0172 | 0.023749 | 23.3487 | -0.047429 | 0.13831 | 0.000144 | 0.010182 | 0.010182 | 0.54738
2019-04-09 15:35 | 1554838482.7176 | 524 | 0.20949 | 0.000449 | 1e-06 | 99.332 | 99.7133 | 120.734 | 20.6887 | 0.002279 | 0.002208 | -154.511 | 0.23332 | 8.4004 | 0.002825 | 0.004122 | 0.003986 | 4.2e-05 | 0.001151 | 0.27129 | 1.3447 | 2.5477 | 98.44 | 20 | 48.427 | 76.0846 | 89.2232 | 74.4002 | 0.011514 | 1.0158 | 0.023753 | 23.3829 | -0.034277 | 0.13881 | 0.000286 | 0.020332 | 0.020332 | 0.27513
2019-04-09 15:35 | 1554838483.7185 | 525 | 0.20952 | 0.000454 | 1e-06 | 99.3245 | 99.7777 | 120.5241 | 20.6878 | 0.002346 | 0.00219 | -154.5138 | 0.23328 | 8.4014 | 0.002827 | 0.004481 | 0.00437 | 4.1e-05 | 0.00115 | 0.27312 | 1.3448 | 2.5476 | 98.44 | 20 | 48.424 | 77.7011 | 89.2606 | 73.9464 | 0.011513 | 1.0171 | 0.023747 | 23.3473 | -0.040718 | 0.13878 | 9.6e-05 | 0.006768 | 0.006768 | 0.82637
2019-04-09 15:35 | 1554838484.7184 | 526 | 0.20951 | 0.000451 | 1e-06 | 99.3153 | 99.8427 | 120.2336 | 20.6878 | 0.002319 | 0.002182 | -154.5082 | 0.23249 | 8.4028 | 0.002833 | 0.004409 | 0.004257 | 4.3e-05 | 0.001148 | 0.28064 | 1.3247 | 2.5476 | 98.44 | 20 | 48.422 | 77.0447 | 89.4044 | 73.7732 | 0.011513 | 1.0176 | 0.023725 | 23.3133 | -0.047378 | 0.13831 | 0.000175 | 0.012407 | 0.012407 | 0.44926
2019-04-09 15:35 | 1554838485.7186 | 527 | 0.2095 | 0.000449 | 1e-06 | 99.3754 | 99.8892 | 120.0209 | 20.6874 | 0.002333 | 0.002185 | -154.532 | 0.23329 | 8.403 | 0.002823 | 0.004391 | 0.004189 | 4.2e-05 | 0.001138 | 0.2959 | 1.322 | 2.5476 | 98.44 | 20 | 48.419 | 77.3811 | 89.1791 | 73.8273 | 0.011512 | 1.0175 | 0.023702 | 23.2945 | -0.046124 | 0.13878 | 0.000245 | 0.017351 | 0.017351 | 0.32235
2019-04-09 15:35 | 1554838486.7165 | 528 | 0.20949 | 0.000451 | 1e-06 | 99.3239 | 99.8913 | 119.655 | 20.6883 | 0.002302 | 0.002222 | -154.5448 | 0.23316 | 8.4023 | 0.002873 | 0.004188 | 0.004037 | 4.6e-05 | 0.001134 | 0.27992 | 1.3428 | 2.5475 | 98.44 | 20 | 48.417 | 76.6377 | 90.3716 | 74.7207 | 0.011511 | 1.0149 | 0.023635 | 23.2889 | -0.050947 | 0.1387 | 0.000264 | 0.018679 | 0.018679 | 0.29926
2019-04-09 15:35 | 1554838487.7175 | 529 | 0.20954 | 0.000449 | 1e-06 | 99.3614 | 99.7856 | 119.6101 | 20.6878 | 0.002316 | 0.00221 | -154.5371 | 0.23246 | 8.402 | 0.002823 | 0.004245 | 0.004138 | 5e-05 | 0.001126 | 0.25465 | 1.3407 | 2.5476 | 98.44 | 20 | 48.415 | 76.9856 | 89.1865 | 74.439 | 0.011511 | 1.0157 | 0.02364 | 23.2751 | -0.038103 | 0.13829 | -2.5e-05 | -0.001805 | -0.001805 | -3.0868
2019-04-09 15:35 | 1554838488.7176 | 530 | 0.20949 | 0.000448 | 1e-06 | 99.3355 | 99.8462 | 119.724 | 20.6884 | 0.002345 | 0.00223 | -154.5378 | 0.23315 | 8.4001 | 0.002767 | 0.003998 | 0.003813 | 4.7e-05 | 0.001129 | 0.28655 | 1.3109 | 2.5476 | 98.44 | 20 | 48.414 | 77.676 | 87.8242 | 74.9251 | 0.011511 | 1.0143 | 0.023635 | 23.3024 | -0.045863 | 0.1387 | 0.000306 | 0.021644 | 0.021644 | 0.25826
2019-04-09 15:35 | 1554838489.7165 | 531 | 0.20952 | 0.000453 | 1e-06 | 99.4013 | 100.0692 | 119.9737 | 20.6878 | 0.00236 | 0.002239 | -154.5548 | 0.23233 | 8.4043 | 0.002773 | 0.004227 | 0.004041 | 4.7e-05 | 0.001128 | 0.26377 | 1.31 | 2.5475 | 98.44 | 20 | 48.414 | 78.0332 | 87.9618 | 75.1242 | 0.011511 | 1.0137 | 0.023653 | 23.3334 | -0.059895 | 0.13821 | 0.00012 | 0.008478 | 0.008478 | 0.657
2019-04-09 15:35 | 1554838490.7179 | 532 | 0.20952 | 0.000448 | 1e-06 | 99.4558 | 99.8445 | 120.3763 | 20.6874 | 0.00223 | 0.002249 | -154.539 | 0.23252 | 8.4034 | 0.002857 | 0.004294 | 0.004155 | 4.3e-05 | 0.001131 | 0.27776 | 1.3393 | 2.5476 | 98.44 | 20 | 48.415 | 74.9226 | 89.9839 | 75.3849 | 0.011511 | 1.0129 | 0.023684 | 23.3813 | -0.034883 | 0.13833 | 0.000103 | 0.007267 | 0.007267 | 0.76708
2019-04-09 15:35 | 1554838491.7183 | 533 | 0.20952 | 0.00045 | 2e-06 | 99.4633 | 99.9966 | 120.8556 | 20.6876 | 0.002306 | 0.002152 | -154.5486 | 0.23288 | 8.4004 | 0.002759 | 0.004133 | 0.003961 | 5e-05 | 0.001134 | 0.27815 | 1.33 | 2.5475 | 98.44 | 20 | 48.417 | 76.7375 | 87.643 | 73.0333 | 0.011511 | 1.0198 | 0.023811 | 23.3486 | -0.04783 | 0.13854 | 8.7e-05 | 0.006077 | 0.006077 | 0.91872
2019-04-09 15:35 | 1554838492.7167 | 534 | 0.20952 | 0.000449 | 1e-06 | 99.5279 | 100.0496 | 121.0161 | 20.6874 | 0.002299 | 0.00222 | -154.5382 | 0.23329 | 8.3995 | 0.00278 | 0.004902 | 0.004697 | 4.9e-05 | 0.00113 | 0.27345 | 1.3414 | 2.5475 | 98.44 | 20 | 48.417 | 76.5744 | 88.1333 | 74.6748 | 0.011511 | 1.015 | 0.023771 | 23.4194 | -0.04676 | 0.13878 | 0.000108 | 0.00764 | 0.00764 | 0.73211
2019-04-09 15:35 | 1554838493.7177 | 535 | 0.20953 | 0.000448 | 1e-06 | 99.5467 | 99.8104 | 121.375 | 20.6876 | 0.002311 | 0.00217 | -154.5539 | 0.23249 | 8.4029 | 0.002734 | 0.00449 | 0.004331 | 5e-05 | 0.001123 | 0.2688 | 1.3305 | 2.5476 | 98.44 | 20 | 48.417 | 76.8598 | 87.0425 | 73.4684 | 0.011511 | 1.0185 | 0.023848 | 23.4134 | -0.023665 | 0.13831 | 7.9e-05 | 0.005572 | 0.005572 | 1.0003
2019-04-09 15:35 | 1554838494.7174 | 536 | 0.20951 | 0.000448 | 2e-06 | 99.5487 | 100.0441 | 121.4133 | 20.6879 | 0.002253 | 0.002203 | -154.5637 | 0.23385 | 8.4003 | 0.002721 | 0.004995 | 0.004857 | 4.9e-05 | 0.001126 | 0.30127 | 1.3551 | 2.5476 | 98.44 | 20 | 48.418 | 75.4687 | 86.7095 | 74.2685 | 0.011512 | 1.0162 | 0.023824 | 23.4441 | -0.044395 | 0.13912 | 0.000155 | 0.010997 | 0.010997 | 0.50983
2019-04-09 15:35 | 1554838495.7168 | 537 | 0.20952 | 0.000449 | 1e-06 | 99.5976 | 100.0809 | 121.4551 | 20.6875 | 0.002293 | 0.002184 | -154.5588 | 0.23318 | 8.4005 | 0.002714 | 0.003495 | 0.003361 | 5e-05 | 0.001127 | 0.29087 | 1.3396 | 2.5476 | 98.44 | 20 | 48.419 | 76.4287 | 86.5537 | 73.8183 | 0.011512 | 1.0175 | 0.023843 | 23.433 | -0.043299 | 0.13872 | 0.000109 | 0.007702 | 0.007702 | 0.72582
2019-04-09 15:35 | 1554838496.7168 | 538 | 0.20952 | 0.000453 | 1e-06 | 99.6402 | 100.2678 | 121.2893 | 20.6881 | 0.002268 | 0.002162 | -154.5653 | 0.23356 | 8.4024 | 0.002743 | 0.004215 | 0.003984 | 5e-05 | 0.001122 | 0.27312 | 1.3615 | 2.5475 | 98.44 | 20 | 48.419 | 75.8221 | 87.2482 | 73.2872 | 0.011512 | 1.0191 | 0.023845 | 23.399 | -0.056162 | 0.13894 | 8.4e-05 | 0.005887 | 0.005887 | 0.95115
2019-04-09 15:35 | 1554838497.7172 | 539 | 0.20952 | 0.000451 | 1e-06 | 99.7125 | 100.2781 | 120.999 | 20.6882 | 0.002268 | 0.002143 | -154.5758 | 0.23313 | 8.4026 | 0.00274 | 0.004319 | 0.00415 | 4.8e-05 | 0.001115 | 0.28423 | 1.3592 | 2.5476 | 98.44 | 20 | 48.42 | 75.84 | 87.1672 | 72.8344 | 0.011512 | 1.0204 | 0.023832 | 23.3557 | -0.050589 | 0.13869 | 0.000126 | 0.008968 | 0.008968 | 0.62323
2019-04-09 15:35 | 1554838498.7169 | 540 | 0.20949 | 0.000445 | 1e-06 | 99.7429 | 100.3493 | 120.8291 | 20.6887 | 0.002284 | 0.002182 | -154.577 | 0.23342 | 8.4025 | 0.002738 | 0.004224 | 0.004059 | 4.8e-05 | 0.001112 | 0.27743 | 1.3875 | 2.5478 | 98.43 | 20 | 48.423 | 76.2193 | 87.1277 | 73.7695 | 0.011514 | 1.0177 | 0.023784 | 23.3709 | -0.054217 | 0.13886 | 0.000283 | 0.020093 | 0.020093 | 0.27851
2019-04-09 15:35 | 1554838499.718 | 541 | 0.20949 | 0.00045 | 1e-06 | 99.8236 | 100.4237 | 120.665 | 20.6877 | 0.002308 | 0.002119 | -154.5691 | 0.23328 | 8.3996 | 0.002749 | 0.004435 | 0.004261 | 5e-05 | 0.001115 | 0.29574 | 1.3516 | 2.5476 | 98.43 | 20 | 48.424 | 76.8001 | 87.4009 | 72.2438 | 0.011514 | 1.0222 | 0.02382 | 23.3036 | -0.053607 | 0.13878 | 0.000272 | 0.019369 | 0.019369 | 0.28875
2019-04-09 15:35 | 1554838500.717 | 542 | 0.20947 | 0.000448 | 1e-06 | 99.7944 | 100.3586 | 120.515 | 20.6882 | 0.002289 | 0.002176 | -154.5687 | 0.23243 | 8.4012 | 0.002782 | 0.004312 | 0.004123 | 4.7e-05 | 0.001112 | 0.27815 | 1.3554 | 2.5476 | 98.43 | 20 | 48.426 | 76.3265 | 88.1933 | 73.6208 | 0.011515 | 1.0181 | 0.023758 | 23.3355 | -0.050427 | 0.13828 | 0.000389 | 0.027659 | 0.027659 | 0.20147
2019-04-09 15:35 | 1554838501.7164 | 543 | 0.20952 | 0.000446 | 2e-06 | 99.8156 | 100.5063 | 120.2526 | 20.6876 | 0.002353 | 0.002193 | -154.572 | 0.23383 | 8.4025 | 0.002733 | 0.00408 | 0.003973 | 4.9e-05 | 0.001104 | 0.25409 | 1.3332 | 2.5475 | 98.43 | 20 | 48.428 | 77.8842 | 87.0049 | 74.0178 | 0.011515 | 1.0169 | 0.023718 | 23.3234 | -0.061681 | 0.13911 | 0.000101 | 0.007065 | 0.007065 | 0.7935
2019-04-09 15:35 | 1554838502.716 | 544 | 0.2095 | 0.000448 | 1e-06 | 99.8099 | 100.4374 | 120.1613 | 20.6877 | 0.002267 | 0.002204 | -154.5813 | 0.23295 | 8.4026 | 0.002742 | 0.004123 | 0.00397 | 4.7e-05 | 0.001104 | 0.27433 | 1.3337 | 2.5477 | 98.43 | 20 | 48.43 | 75.8037 | 87.2314 | 74.2961 | 0.011516 | 1.0161 | 0.0237 | 23.3239 | -0.05606 | 0.13858 | 0.00023 | 0.016281 | 0.016281 | 0.34303
2019-04-09 15:35 | 1554838503.7174 | 545 | 0.20952 | 0.000449 | 1e-06 | 99.8662 | 100.4175 | 120.2183 | 20.6881 | 0.002281 | 0.002175 | -154.5546 | 0.23357 | 8.4005 | 0.002668 | 0.004218 | 0.004036 | 4.5e-05 | 0.001102 | 0.28102 | 1.3351 | 2.5476 | 98.43 | 20 | 48.433 | 76.1361 | 85.4545 | 73.6036 | 0.011516 | 1.0181 | 0.023729 | 23.3061 | -0.04924 | 0.13895 | 0.000119 | 0.008442 | 0.008442 | 0.66335
2019-04-09 15:35 | 1554838504.7164 | 546 | 0.20952 | 0.00045 | 1e-06 | 99.8736 | 100.5004 | 120.4732 | 20.6877 | 0.00225 | 0.002176 | -154.5493 | 0.23271 | 8.4025 | 0.00271 | 0.003875 | 0.003654 | 4.9e-05 | 0.001111 | 0.28136 | 1.3528 | 2.5476 | 98.43 | 20 | 48.434 | 75.389 | 86.4617 | 73.6157 | 0.011517 | 1.0181 | 0.023754 | 23.3312 | -0.055958 | 0.13844 | 0.000106 | 0.007471 | 0.007471 | 0.74675
2019-04-09 15:35 | 1554838505.7175 | 547 | 0.20952 | 0.000448 | 1e-06 | 99.9463 | 100.6398 | 120.6833 | 20.6871 | 0.002247 | 0.002156 | -154.5353 | 0.23277 | 8.4027 | 0.00261 | 0.003993 | 0.003838 | 4.7e-05 | 0.00112 | 0.2975 | 1.2994 | 2.5476 | 98.43 | 20 | 48.437 | 75.3263 | 84.0428 | 73.1381 | 0.011517 | 1.0195 | 0.023791 | 23.3355 | -0.061849 | 0.13848 | 9.6e-05 | 0.006786 | 0.006786 | 0.82239
2019-04-09 15:35 | 1554838506.7174 | 548 | 0.20957 | 0.00045 | 1e-06 | 99.9594 | 100.507 | 121.1945 | 20.6887 | 0.002202 | 0.002193 | -154.5248 | 0.23316 | 8.4027 | 0.002712 | 0.004514 | 0.004363 | 5.3e-05 | 0.001124 | 0.28335 | 1.307 | 2.5475 | 98.43 | 20 | 48.44 | 74.2356 | 86.511 | 74.0391 | 0.011518 | 1.0169 | 0.02381 | 23.4153 | -0.048862 | 0.1387 | -0.000199 | -0.014283 | -0.014283 | -0.39136
2019-04-09 15:35 | 1554838507.7166 | 549 | 0.20956 | 0.000452 | 1e-06 | 100.0285 | 100.6529 | 121.817 | 20.6884 | 0.002215 | 0.002158 | -154.5288 | 0.23263 | 8.4014 | 0.002707 | 0.003773 | 0.003589 | 5e-05 | 0.001128 | 0.28512 | 1.3924 | 2.5477 | 98.43 | 20 | 48.443 | 74.5652 | 86.3928 | 73.1947 | 0.011519 | 1.0193 | 0.0239 | 23.4468 | -0.055659 | 0.13839 | -0.000127 | -0.009191 | -0.009191 | -0.60679
2019-04-09 15:35 | 1554838508.7176 | 550 | 0.20953 | 0.000446 | 1e-06 | 99.9487 | 100.6109 | 121.9879 | 20.6887 | 0.002226 | 0.002173 | -154.5244 | 0.23307 | 8.4004 | 0.002694 | 0.004118 | 0.003939 | 5.1e-05 | 0.001131 | 0.28512 | 1.396 | 2.5476 | 98.43 | 20 | 48.445 | 74.8194 | 86.0614 | 73.5473 | 0.011519 | 1.0183 | 0.023905 | 23.4751 | -0.059066 | 0.13865 | 9e-05 | 0.006371 | 0.006371 | 0.87711
2019-04-09 15:35 | 1554838509.7172 | 551 | 0.20954 | 0.000448 | 1e-06 | 100.0079 | 100.7892 | 121.9799 | 20.6888 | 0.002212 | 0.002184 | -154.5199 | 0.23299 | 8.4008 | 0.002768 | 0.004205 | 0.004057 | 5.2e-05 | 0.001135 | 0.29983 | 1.3621 | 2.5477 | 98.43 | 20 | 48.445 | 74.4844 | 87.8424 | 73.8187 | 0.011519 | 1.0175 | 0.023895 | 23.4836 | -0.069599 | 0.1386 | -8e-06 | -0.000629 | -0.000629 | -8.8766
2019-04-09 15:35 | 1554838510.7163 | 552 | 0.20959 | 0.000449 | 1e-06 | 100.0133 | 100.6178 | 121.6964 | 20.6883 | 0.002213 | 0.002136 | -154.5255 | 0.23318 | 8.4003 | 0.002745 | 0.003932 | 0.003784 | 5.2e-05 | 0.001141 | 0.27815 | 1.3675 | 2.5477 | 98.43 | 20 | 48.447 | 74.505 | 87.3079 | 72.663 | 0.01152 | 1.0209 | 0.023907 | 23.4171 | -0.053897 | 0.13872 | -0.000287 | -0.020627 | -0.020627 | -0.27102
2019-04-09 15:35 | 1554838511.7163 | 553 | 0.20955 | 0.000446 | 1e-06 | 100.0446 | 100.5882 | 121.4494 | 20.6879 | 0.002227 | 0.002127 | -154.515 | 0.23252 | 8.4029 | 0.002713 | 0.003339 | 0.003153 | 5.3e-05 | 0.001143 | 0.28423 | 1.3599 | 2.5475 | 98.43 | 20 | 48.449 | 74.8513 | 86.5275 | 72.4372 | 0.01152 | 1.0216 | 0.02389 | 23.3857 | -0.04847 | 0.13833 | -5.9e-05 | -0.004277 | -0.004277 | -1.3035
2019-04-09 15:35 | 1554838512.7173 | 554 | 0.20958 | 0.000452 | 1e-06 | 100.1059 | 100.6405 | 121.4389 | 20.687 | 0.002256 | 0.002134 | -154.5178 | 0.23324 | 8.3998 | 0.002623 | 0.004722 | 0.004531 | 5.2e-05 | 0.001148 | 0.27295 | 1.3701 | 2.5474 | 98.43 | 20 | 48.449 | 75.5405 | 84.3515 | 72.6166 | 0.01152 | 1.0211 | 0.023883 | 23.3908 | -0.047637 | 0.13875 | -0.000224 | -0.0161 | -0.0161 | -0.34732
2019-04-09 15:35 | 1554838513.7173 | 555 | 0.20956 | 0.000451 | 1e-06 | 100.0857 | 100.7568 | 120.9706 | 20.6877 | 0.0022 | 0.002157 | -154.5272 | 0.23297 | 8.4025 | 0.002637 | 0.004376 | 0.004286 | 5.1e-05 | 0.001147 | 0.2656 | 1.3886 | 2.5477 | 98.43 | 20 | 48.448 | 74.1902 | 84.7088 | 73.152 | 0.01152 | 1.0195 | 0.023819 | 23.3637 | -0.059779 | 0.1386 | -0.000116 | -0.008317 | -0.008317 | -0.67158
2019-04-09 15:35 | 1554838514.7178 | 556 | 0.20954 | 0.00045 | 1e-06 | 100.1333 | 100.7623 | 120.6401 | 20.6891 | 0.002193 | 0.002123 | -154.5249 | 0.23319 | 8.4016 | 0.002701 | 0.003721 | 0.003507 | 5.1e-05 | 0.001142 | 0.28385 | 1.3624 | 2.5476 | 98.43 | 20 | 48.446 | 74.0279 | 86.2374 | 72.3543 | 0.011519 | 1.0218 | 0.023814 | 23.3049 | -0.056011 | 0.13873 | 2.4e-05 | 0.001671 | 0.001671 | 3.3454
2019-04-09 15:35 | 1554838515.7167 | 557 | 0.20955 | 0.000449 | 1e-06 | 100.1408 | 100.8232 | 120.0927 | 20.6881 | 0.002253 | 0.002124 | -154.5242 | 0.23406 | 8.3994 | 0.002611 | 0.004986 | 0.004889 | 5.2e-05 | 0.001141 | 0.25393 | 1.3505 | 2.5476 | 98.43 | 20 | 48.444 | 75.4738 | 84.082 | 72.3814 | 0.011519 | 1.0217 | 0.023759 | 23.2529 | -0.060752 | 0.13924 | -3.7e-05 | -0.002685 | -0.002685 | -2.0897
2019-04-09 15:35 | 1554838516.7164 | 558 | 0.20954 | 0.00045 | 2e-06 | 100.1122 | 100.8601 | 119.8003 | 20.6893 | 0.002227 | 0.00214 | -154.5216 | 0.23229 | 8.4033 | 0.002699 | 0.003835 | 0.003647 | 5e-05 | 0.001146 | 0.2891 | 1.3339 | 2.5476 | 98.43 | 20 | 48.439 | 74.838 | 86.1847 | 72.7553 | 0.011518 | 1.0206 | 0.023717 | 23.2371 | -0.066572 | 0.13819 | -2.4e-05 | -0.001773 | -0.001773 | -3.1415
2019-04-09 15:35 | 1554838517.7158 | 559 | 0.20954 | 0.000453 | 1e-06 | 100.0693 | 100.5957 | 119.4245 | 20.6886 | 0.002243 | 0.002145 | -154.5112 | 0.23394 | 8.4011 | 0.002634 | 0.004659 | 0.00448 | 5.3e-05 | 0.00115 | 0.2672 | 1.3115 | 2.5474 | 98.43 | 20 | 48.435 | 75.241 | 84.6176 | 72.8838 | 0.011517 | 1.0203 | 0.023675 | 23.2049 | -0.046933 | 0.13917 | 1.5e-05 | 0.000973 | 0.000973 | 5.7628
2019-04-09 15:35 | 1554838518.7155 | 560 | 0.20953 | 0.000451 | 1e-06 | 100.0723 | 100.7245 | 119.1754 | 20.6886 | 0.002212 | 0.002071 | -154.5161 | 0.23273 | 8.4038 | 0.002752 | 0.004212 | 0.004031 | 4.8e-05 | 0.001148 | 0.28208 | 1.3339 | 2.5475 | 98.43 | 20 | 48.429 | 74.4943 | 87.4682 | 71.1026 | 0.011515 | 1.0255 | 0.023712 | 23.121 | -0.058105 | 0.13845 | 6.9e-05 | 0.004874 | 0.004874 | 1.1447
2019-04-09 15:35 | 1554838519.7166 | 561 | 0.20954 | 0.000448 | 1e-06 | 100.1083 | 100.9678 | 119.0754 | 20.6886 | 0.00226 | 0.002115 | -154.5333 | 0.23335 | 8.402 | 0.002691 | 0.004157 | 0.003954 | 5.2e-05 | 0.001147 | 0.2912 | 1.3469 | 2.5475 | 98.43 | 20 | 48.418 | 75.6448 | 86.0098 | 72.1659 | 0.011513 | 1.0224 | 0.023665 | 23.147 | -0.076465 | 0.13882 | -1.7e-05 | -0.00127 | -0.00127 | -4.405
2019-04-09 15:35 | 1554838520.7166 | 562 | 0.20959 | 0.00045 | 1e-06 | 100.1157 | 100.8022 | 119.4508 | 20.6887 | 0.002198 | 0.002137 | -154.532 | 0.23338 | 8.4036 | 0.002715 | 0.00386 | 0.003702 | 4.9e-05 | 0.001147 | 0.28766 | 1.3697 | 2.5477 | 98.43 | 20 | 48.413 | 74.1498 | 86.5837 | 72.6883 | 0.011512 | 1.0208 | 0.023684 | 23.2009 | -0.061124 | 0.13884 | -0.000329 | -0.023438 | -0.023438 | -0.23872
2019-04-09 15:35 | 1554838521.7169 | 563 | 0.20958 | 0.000447 | 1e-06 | 100.206 | 100.9196 | 119.6388 | 20.6894 | 0.00225 | 0.002096 | -154.5355 | 0.23367 | 8.4044 | 0.002694 | 0.004048 | 0.003954 | 4.8e-05 | 0.001143 | 0.28528 | 1.369 | 2.5476 | 98.43 | 20 | 48.407 | 75.4039 | 86.0668 | 71.7067 | 0.01151 | 1.0237 | 0.023737 | 23.1862 | -0.063473 | 0.13901 | -0.000227 | -0.016222 | -0.016222 | -0.34534
2019-04-09 15:35 | 1554838522.7167 | 564 | 0.20955 | 0.00045 | 1e-06 | 100.189 | 100.7437 | 119.5669 | 20.6896 | 0.002238 | 0.002141 | -154.5336 | 0.23334 | 8.4001 | 0.002634 | 0.004499 | 0.004402 | 5.2e-05 | 0.001138 | 0.28639 | 1.3622 | 2.5476 | 98.43 | 20 | 48.399 | 75.1151 | 84.6209 | 72.7827 | 0.011508 | 1.0206 | 0.023693 | 23.2154 | -0.049388 | 0.13881 | -8.7e-05 | -0.006248 | -0.006248 | -0.89541
2019-04-09 15:35 | 1554838523.7167 | 565 | 0.2095 | 0.000449 | 1e-06 | 100.2182 | 100.8714 | 119.8383 | 20.6884 | 0.002266 | 0.00217 | -154.5384 | 0.23301 | 8.4025 | 0.002799 | 0.004168 | 0.003965 | 5e-05 | 0.001138 | 0.26416 | 1.3859 | 2.5477 | 98.43 | 20 | 48.383 | 75.7897 | 88.5887 | 73.4761 | 0.011504 | 1.0185 | 0.023696 | 23.265 | -0.058112 | 0.13862 | 0.000213 | 0.015081 | 0.015081 | 0.37042
2019-04-09 15:35 | 1554838524.7171 | 566 | 0.20954 | 0.000448 | 1e-06 | 100.2072 | 100.964 | 120.3547 | 20.6879 | 0.002148 | 0.00213 | -154.5325 | 0.23264 | 8.4019 | 0.002777 | 0.003671 | 0.003445 | 4.9e-05 | 0.001132 | 0.28854 | 1.4065 | 2.5476 | 98.43 | 20 | 48.376 | 72.9525 | 88.0562 | 72.5215 | 0.011503 | 1.0213 | 0.02378 | 23.2829 | -0.067301 | 0.1384 | 1.6e-05 | 0.001103 | 0.001103 | 5.0578
2019-04-09 15:35 | 1554838525.7157 | 567 | 0.20955 | 0.000454 | 1e-06 | 100.3379 | 100.9337 | 120.9072 | 20.6892 | 0.002197 | 0.002111 | -154.5523 | 0.23323 | 8.3995 | 0.002791 | 0.003312 | 0.003151 | 5.4e-05 | 0.00113 | 0.27129 | 1.3884 | 2.5477 | 98.43 | 20 | 48.359 | 74.1169 | 88.4099 | 72.0698 | 0.011499 | 1.0227 | 0.02385 | 23.3211 | -0.052956 | 0.13875 | -5e-05 | -0.003616 | -0.003616 | -1.5462
2019-04-09 15:35 | 1554838526.7168 | 568 | 0.20954 | 0.000447 | 1e-06 | 100.2938 | 100.8552 | 120.9985 | 20.6891 | 0.002228 | 0.002107 | -154.5566 | 0.23316 | 8.4019 | 0.002752 | 0.004216 | 0.003997 | 4.7e-05 | 0.001129 | 0.2756 | 1.336 | 2.5476 | 98.43 | 20 | 48.351 | 74.8642 | 87.4679 | 71.9544 | 0.011497 | 1.023 | 0.023863 | 23.326 | -0.049929 | 0.13871 | 2e-05 | 0.001343 | 0.001343 | 4.1635
2019-04-09 15:35 | 1554838527.7158 | 569 | 0.20955 | 0.00045 | 2e-06 | 100.3051 | 100.9802 | 120.8967 | 20.6892 | 0.002249 | 0.00209 | -154.5542 | 0.23315 | 8.3999 | 0.002836 | 0.004081 | 0.003928 | 5.2e-05 | 0.001132 | 0.24818 | 1.3132 | 2.5475 | 98.43 | 20 | 48.343 | 75.3665 | 89.491 | 71.5458 | 0.011495 | 1.0242 | 0.023867 | 23.3023 | -0.059993 | 0.1387 | -7e-05 | -0.005074 | -0.005074 | -1.1016
2019-04-09 15:35 | 1554838528.7169 | 570 | 0.20952 | 0.000453 | 1e-06 | 100.407 | 101.0696 | 120.9715 | 20.6891 | 0.002236 | 0.002137 | -154.5535 | 0.23286 | 8.4026 | 0.002763 | 0.00401 | 0.003908 | 4.9e-05 | 0.001136 | 0.28639 | 1.3306 | 2.5476 | 98.43 | 20 | 48.326 | 75.0529 | 87.7335 | 72.688 | 0.011491 | 1.0208 | 0.023835 | 23.3482 | -0.058838 | 0.13853 | 0.000138 | 0.009775 | 0.009775 | 0.5711
2019-04-09 15:35 | 1554838529.7158 | 571 | 0.20958 | 0.00045 | 2e-06 | 100.4144 | 101.1984 | 120.8455 | 20.6897 | 0.002236 | 0.002138 | -154.536 | 0.23336 | 8.4023 | 0.002771 | 0.003747 | 0.00363 | 4.7e-05 | 0.001134 | 0.25841 | 1.325 | 2.5475 | 98.43 | 20 | 48.317 | 75.0565 | 87.9347 | 72.7032 | 0.011489 | 1.0208 | 0.023822 | 23.3365 | -0.069559 | 0.13883 | -0.00027 | -0.019341 | -0.019341 | -0.28927
2019-04-09 15:36 | 1554838530.717 | 572 | 0.20954 | 0.000447 | 1e-06 | 100.4144 | 101.0074 | 120.7129 | 20.688 | 0.002228 | 0.00203 | -154.535 | 0.23259 | 8.399 | 0.002766 | 0.003635 | 0.003397 | 5.1e-05 | 0.00113 | 0.26897 | 1.3278 | 2.5476 | 98.43 | 20 | 48.301 | 74.8631 | 87.812 | 70.1102 | 0.011485 | 1.0285 | 0.023898 | 23.2361 | -0.052661 | 0.13837 | -2.2e-05 | -0.001649 | -0.001649 | -3.3824
2019-04-09 15:36 | 1554838531.716 | 573 | 0.20954 | 0.000452 | 2e-06 | 100.4397 | 101.2101 | 120.8715 | 20.6887 | 0.002201 | 0.002134 | -154.5499 | 0.23218 | 8.4018 | 0.002778 | 0.003391 | 0.00325 | 5.1e-05 | 0.001133 | 0.25625 | 1.3174 | 2.5475 | 98.43 | 20 | 48.293 | 74.2137 | 88.082 | 72.6033 | 0.011483 | 1.0211 | 0.023828 | 23.3356 | -0.068338 | 0.13812 | 2.1e-05 | 0.00143 | 0.00143 | 3.8918
2019-04-09 15:36 | 1554838532.716 | 574 | 0.20955 | 0.000447 | 1e-06 | 100.4811 | 101.1213 | 120.9628 | 20.6888 | 0.002228 | 0.002167 | -154.5442 | 0.23271 | 8.4021 | 0.00281 | 0.003907 | 0.003745 | 5.2e-05 | 0.001129 | 0.27184 | 1.3492 | 2.5476 | 98.43 | 20 | 48.285 | 74.8741 | 88.861 | 73.3963 | 0.011481 | 1.0188 | 0.02381 | 23.3712 | -0.056802 | 0.13844 | -3.6e-05 | -0.002667 | -0.002667 | -2.0922
2019-04-09 15:36 | 1554838533.716 | 575 | 0.20957 | 0.000454 | 2e-06 | 100.4811 | 101.0996 | 120.8417 | 20.6903 | 0.002186 | 0.002133 | -154.552 | 0.23322 | 8.4013 | 0.002939 | 0.00341 | 0.003253 | 5.1e-05 | 0.001125 | 0.27273 | 1.3509 | 2.5477 | 98.43 | 20 | 48.272 | 73.852 | 91.9824 | 72.5808 | 0.011478 | 1.0212 | 0.023826 | 23.332 | -0.054884 | 0.13874 | -0.000204 | -0.014652 | -0.014652 | -0.38161
2019-04-09 15:36 | 1554838534.7163 | 576 | 0.20953 | 0.000453 | 1e-06 | 100.5621 | 101.2541 | 120.384 | 20.69 | 0.002187 | 0.00207 | -154.5599 | 0.23273 | 8.4047 | 0.002915 | 0.004906 | 0.004801 | 5.1e-05 | 0.001124 | 0.26344 | 1.3625 | 2.5476 | 98.43 | 20 | 48.267 | 73.8971 | 91.4041 | 71.0679 | 0.011477 | 1.0256 | 0.023833 | 23.2368 | -0.061338 | 0.13845 | 7.3e-05 | 0.005113 | 0.005113 | 1.0913
2019-04-09 15:36 | 1554838535.716 | 577 | 0.20962 | 0.00045 | 1e-06 | 100.4793 | 101.2741 | 120.2314 | 20.6896 | 0.002152 | 0.002126 | -154.5445 | 0.23301 | 8.4028 | 0.00289 | 0.004752 | 0.004562 | 5.1e-05 | 0.00112 | 0.29607 | 1.4138 | 2.5477 | 98.43 | 20.01 | 48.261 | 73.047 | 90.7988 | 72.4131 | 0.011483 | 1.0217 | 0.023771 | 23.2674 | -0.070467 | 0.13861 | -0.000484 | -0.034533 | -0.034533 | -0.16176
2019-04-09 15:36 | 1554838536.7161 | 578 | 0.20954 | 0.000449 | 2e-06 | 100.5166 | 101.3119 | 120.2268 | 20.689 | 0.002171 | 0.002095 | -154.5547 | 0.23269 | 8.4001 | 0.002929 | 0.004503 | 0.00432 | 5.3e-05 | 0.001119 | 0.28639 | 1.3965 | 2.5476 | 98.43 | 20.01 | 48.252 | 73.5086 | 91.7418 | 71.6713 | 0.01148 | 1.0238 | 0.023796 | 23.2419 | -0.070497 | 0.13843 | 2.2e-05 | 0.001476 | 0.001476 | 3.78
2019-04-09 15:36 | 1554838537.7165 | 579 | 0.20956 | 0.000449 | 1e-06 | 100.5597 | 101.2407 | 120.3479 | 20.6894 | 0.002188 | 0.002126 | -154.5719 | 0.23277 | 8.4014 | 0.002791 | 0.004165 | 0.004052 | 5.2e-05 | 0.001116 | 0.28871 | 1.4006 | 2.5476 | 98.43 | 20.01 | 48.248 | 73.916 | 88.4007 | 72.4285 | 0.01148 | 1.0216 | 0.023782 | 23.2791 | -0.060364 | 0.13848 | -0.000105 | -0.007526 | -0.007526 | -0.74149
2019-04-09 15:36 | 1554838538.7165 | 580 | 0.20955 | 0.000447 | 1e-06 | 100.5845 | 101.3057 | 120.6077 | 20.6889 | 0.002189 | 0.002116 | -154.5673 | 0.2321 | 8.4022 | 0.002866 | 0.004285 | 0.004096 | 5e-05 | 0.001119 | 0.28385 | 1.3734 | 2.5476 | 98.43 | 20.01 | 48.24 | 73.9295 | 90.2214 | 72.1911 | 0.011478 | 1.0223 | 0.023816 | 23.2963 | -0.06391 | 0.13808 | -4.2e-05 | -0.003022 | -0.003022 | -1.8416
2019-04-09 15:36 | 1554838539.7163 | 581 | 0.20955 | 0.000451 | 1e-06 | 100.5714 | 101.1853 | 120.8424 | 20.6891 | 0.002195 | 0.002121 | -154.5747 | 0.23268 | 8.4017 | 0.002857 | 0.003962 | 0.003769 | 5.1e-05 | 0.001116 | 0.26864 | 1.3695 | 2.5477 | 98.43 | 20.01 | 48.238 | 74.0717 | 89.9894 | 72.2997 | 0.011477 | 1.022 | 0.023835 | 23.3226 | -0.054434 | 0.13842 | -7.2e-05 | -0.005222 | -0.005222 | -1.0682
2019-04-09 15:36 | 1554838540.7152 | 582 | 0.20954 | 0.000448 | 1e-06 | 100.5377 | 101.1729 | 120.876 | 20.6901 | 0.002136 | 0.002117 | -154.5627 | 0.23258 | 8.4029 | 0.00285 | 0.004155 | 0.004011 | 5e-05 | 0.001121 | 0.29336 | 1.3755 | 2.5477 | 98.43 | 20.01 | 48.235 | 72.6654 | 89.8302 | 72.2142 | 0.011476 | 1.0222 | 0.023842 | 23.3229 | -0.056332 | 0.13836 | -1.9e-05 | -0.001445 | -0.001445 | -3.8598
2019-04-09 15:36 | 1554838541.7166 | 583 | 0.20954 | 0.000446 | 1e-06 | 100.6801 | 101.4444 | 120.8586 | 20.6904 | 0.002198 | 0.002149 | -154.5613 | 0.23306 | 8.4018 | 0.002853 | 0.004291 | 0.004161 | 5.2e-05 | 0.001127 | 0.30829 | 1.3202 | 2.5476 | 98.43 | 20.01 | 48.23 | 74.1553 | 89.9085 | 72.98 | 0.011475 | 1.02 | 0.023814 | 23.3471 | -0.067651 | 0.13865 | 3e-06 | 0.000137 | 0.000137 | 40.7849
2019-04-09 15:36 | 1554838542.7162 | 584 | 0.20953 | 0.000446 | 1e-06 | 100.7039 | 101.3677 | 120.7042 | 20.6895 | 0.002201 | 0.00215 | -154.5645 | 0.23258 | 8.4026 | 0.002815 | 0.004895 | 0.004614 | 5.2e-05 | 0.001125 | 0.30359 | 1.3276 | 2.5476 | 98.43 | 20.01 | 48.226 | 74.2154 | 88.9882 | 72.9969 | 0.011474 | 1.0199 | 0.023798 | 23.3327 | -0.058766 | 0.13836 | 8.1e-05 | 0.00566 | 0.00566 | 0.98522
2019-04-09 15:36 | 1554838543.7157 | 585 | 0.20954 | 0.000449 | 1e-06 | 100.7655 | 101.4572 | 120.7842 | 20.6891 | 0.002182 | 0.002093 | -154.5541 | 0.23243 | 8.4021 | 0.002832 | 0.003954 | 0.003901 | 5.3e-05 | 0.001122 | 0.26465 | 1.3263 | 2.5475 | 98.43 | 20.01 | 48.218 | 73.7692 | 89.3963 | 71.6308 | 0.011472 | 1.024 | 0.023853 | 23.2944 | -0.061184 | 0.13827 | 2e-05 | 0.001367 | 0.001367 | 4.0758
2019-04-09 15:36 | 1554838544.7164 | 586 | 0.20954 | 0.000452 | 1e-06 | 100.7092 | 101.4537 | 120.8854 | 20.6895 | 0.002164 | 0.00209 | -154.5551 | 0.23314 | 8.4039 | 0.002904 | 0.004127 | 0.003977 | 5.2e-05 | 0.001119 | 0.29247 | 1.3492 | 2.5476 | 98.43 | 20.01 | 48.214 | 73.3251 | 91.1365 | 71.5635 | 0.011471 | 1.0242 | 0.023865 | 23.3019 | -0.065878 | 0.13869 | -2.1e-05 | -0.001534 | -0.001534 | -3.6441
2019-04-09 15:36 | 1554838545.7164 | 587 | 0.20954 | 0.00045 | 1e-06 | 100.6824 | 101.3467 | 120.837 | 20.6898 | 0.002239 | 0.002092 | -154.5614 | 0.23267 | 8.401 | 0.002812 | 0.004011 | 0.003826 | 5e-05 | 0.001118 | 0.30237 | 1.3145 | 2.5477 | 98.43 | 20.02 | 48.21 | 75.1309 | 88.9211 | 71.6151 | 0.011478 | 1.024 | 0.023859 | 23.2989 | -0.058818 | 0.13842 | -1.7e-05 | -0.001288 | -0.001288 | -4.3311
2019-04-09 15:36 | 1554838546.7158 | 588 | 0.20955 | 0.000457 | 1e-06 | 100.7319 | 101.2496 | 120.8222 | 20.69 | 0.002226 | 0.002061 | -154.5543 | 0.23402 | 8.4016 | 0.002842 | 0.004036 | 0.003939 | 5.1e-05 | 0.001118 | 0.25658 | 1.346 | 2.5476 | 98.43 | 20.02 | 48.198 | 74.8339 | 89.6419 | 70.8481 | 0.011475 | 1.0263 | 0.023884 | 23.2716 | -0.045858 | 0.13922 | -5.6e-05 | -0.004095 | -0.004095 | -1.3701
2019-04-09 15:36 | 1554838547.7158 | 589 | 0.20952 | 0.000453 | 1e-06 | 100.7774 | 101.4943 | 120.8549 | 20.69 | 0.002225 | 0.002091 | -154.544 | 0.23351 | 8.4039 | 0.002787 | 0.004274 | 0.004097 | 4.7e-05 | 0.001116 | 0.28854 | 1.3645 | 2.5475 | 98.43 | 20.02 | 48.192 | 74.7957 | 88.3022 | 71.5848 | 0.011473 | 1.0241 | 0.023861 | 23.2996 | -0.063403 | 0.13892 | 0.000139 | 0.009865 | 0.009865 | 0.56749
2019-04-09 15:36 | 1554838548.7151 | 590 | 0.20958 | 0.000449 | 1e-06 | 100.6815 | 101.3357 | 120.6004 | 20.6902 | 0.002187 | 0.002072 | -154.5384 | 0.23333 | 8.3983 | 0.002771 | 0.004335 | 0.004137 | 4.9e-05 | 0.001113 | 0.27057 | 1.3657 | 2.5475 | 98.43 | 20.02 | 48.186 | 73.8987 | 87.9359 | 71.1298 | 0.011472 | 1.0255 | 0.023852 | 23.2597 | -0.057925 | 0.13881 | -0.000258 | -0.018549 | -0.018549 | -0.30158
2019-04-09 15:36 | 1554838549.715 | 591 | 0.20954 | 0.000448 | 1e-06 | 100.8078 | 101.5773 | 120.2399 | 20.6895 | 0.002172 | 0.002096 | -154.5401 | 0.23328 | 8.4016 | 0.002853 | 0.003948 | 0.003778 | 4.6e-05 | 0.001115 | 0.28174 | 1.3832 | 2.5476 | 98.43 | 20.02 | 48.174 | 73.5147 | 89.8906 | 71.6927 | 0.011469 | 1.0238 | 0.023797 | 23.2439 | -0.068014 | 0.13878 | -1.9e-05 | -0.001421 | -0.001421 | -3.9369
2019-04-09 15:36 | 1554838550.7152 | 592 | 0.20956 | 0.00045 | 1e-06 | 100.8013 | 101.5401 | 120.4368 | 20.6895 | 0.002167 | 0.002106 | -154.5585 | 0.23327 | 8.4016 | 0.002883 | 0.004277 | 0.004211 | 5e-05 | 0.001118 | 0.26864 | 1.3755 | 2.5477 | 98.43 | 20.02 | 48.168 | 73.3945 | 90.6315 | 71.9361 | 0.011468 | 1.0231 | 0.023808 | 23.2712 | -0.065322 | 0.13877 | -9.7e-05 | -0.007033 | -0.007033 | -0.79514
2019-04-09 15:36 | 1554838551.7145 | 593 | 0.20954 | 0.00045 | 1e-06 | 100.7745 | 101.42 | 120.6258 | 20.689 | 0.002182 | 0.002101 | -154.5465 | 0.23368 | 8.403 | 0.002811 | 0.003943 | 0.003769 | 4.4e-05 | 0.001115 | 0.30525 | 1.3721 | 2.5477 | 98.43 | 20.02 | 48.155 | 73.7753 | 88.8782 | 71.8248 | 0.011465 | 1.0234 | 0.02383 | 23.2856 | -0.057114 | 0.13902 | 5e-06 | 0.000257 | 0.000257 | 21.792
2019-04-09 15:36 | 1554838552.7146 | 594 | 0.20953 | 0.000447 | 1e-06 | 100.7298 | 101.4018 | 120.8814 | 20.6891 | 0.002159 | 0.002121 | -154.5555 | 0.23347 | 8.4018 | 0.00288 | 0.004304 | 0.004156 | 4.6e-05 | 0.001115 | 0.27527 | 1.3806 | 2.5475 | 98.43 | 20.02 | 48.149 | 73.2142 | 90.5459 | 72.3084 | 0.011463 | 1.022 | 0.023839 | 23.3266 | -0.059471 | 0.13889 | 4.2e-05 | 0.002916 | 0.002916 | 1.9195
2019-04-09 15:36 | 1554838553.7148 | 595 | 0.20955 | 0.000448 | 1e-06 | 100.6955 | 101.5019 | 121.2391 | 20.6905 | 0.002203 | 0.002102 | -154.5594 | 0.23236 | 8.4054 | 0.002926 | 0.003759 | 0.003665 | 4.9e-05 | 0.001113 | 0.28152 | 1.3295 | 2.5476 | 98.43 | 20.02 | 48.143 | 74.2693 | 91.6673 | 71.8551 | 0.011462 | 1.0233 | 0.02389 | 23.3458 | -0.071342 | 0.13823 | -6.3e-05 | -0.004532 | -0.004532 | -1.2292
2019-04-09 15:36 | 1554838554.7161 | 596 | 0.20958 | 0.000452 | 1e-06 | 100.778 | 101.4076 | 121.3675 | 20.6907 | 0.002181 | 0.002057 | -154.554 | 0.23223 | 8.4035 | 0.002838 | 0.004226 | 0.003959 | 4.8e-05 | 0.001114 | 0.26897 | 1.3544 | 2.5476 | 98.43 | 20.02 | 48.132 | 73.7364 | 89.5395 | 70.7661 | 0.011459 | 1.0265 | 0.02394 | 23.3213 | -0.055704 | 0.13816 | -0.000243 | -0.0175 | -0.0175 | -0.31815
2019-04-09 15:36 | 1554838555.7153 | 597 | 0.20959 | 0.000454 | 2e-06 | 100.6869 | 101.5449 | 121.3199 | 20.6888 | 0.002142 | 0.002117 | -154.5592 | 0.23344 | 8.4028 | 0.002807 | 0.004663 | 0.00449 | 4.6e-05 | 0.00112 | 0.26609 | 1.3449 | 2.5477 | 98.43 | 20.02 | 48.126 | 72.8 | 88.8008 | 72.212 | 0.011458 | 1.0222 | 0.023885 | 23.3656 | -0.075899 | 0.13887 | -0.000311 | -0.022366 | -0.022366 | -0.25022
2019-04-09 15:36 | 1554838556.7161 | 598 | 0.20959 | 0.000451 | 1e-06 | 100.7354 | 101.5026 | 121.2347 | 20.6896 | 0.002233 | 0.002163 | -154.5562 | 0.23275 | 8.405 | 0.002833 | 0.004476 | 0.00433 | 4.5e-05 | 0.001123 | 0.27848 | 1.3287 | 2.5477 | 98.43 | 20.02 | 48.116 | 75.0004 | 89.4266 | 73.3182 | 0.011455 | 1.019 | 0.023839 | 23.3948 | -0.067859 | 0.13846 | -0.000278 | -0.019968 | -0.019968 | -0.27946
2019-04-09 15:36 | 1554838557.7158 | 599 | 0.20957 | 0.000458 | 1e-06 | 100.7357 | 101.4269 | 121.2365 | 20.6906 | 0.002161 | 0.002033 | -154.565 | 0.2336 | 8.4037 | 0.002822 | 0.003968 | 0.003859 | 4.4e-05 | 0.00113 | 0.28727 | 1.321 | 2.5476 | 98.43 | 20.02 | 48.11 | 73.2608 | 89.1501 | 70.1968 | 0.011454 | 1.0282 | 0.023947 | 23.2894 | -0.061158 | 0.13897 | -0.000171 | -0.012297 | -0.012297 | -0.45542
2019-04-09 15:36 | 1554838558.7148 | 600 | 0.20952 | 0.000451 | 1e-06 | 100.6652 | 101.3694 | 121.267 | 20.6906 | 0.002148 | 0.002068 | -154.5781 | 0.23258 | 8.4034 | 0.002732 | 0.004466 | 0.004305 | 4.3e-05 | 0.001134 | 0.24099 | 1.3464 | 2.5475 | 98.43 | 20.02 | 48.106 | 72.9622 | 86.9954 | 71.0281 | 0.011453 | 1.0258 | 0.023921 | 23.3205 | -0.062357 | 0.13836 | 0.000107 | 0.007611 | 0.007611 | 0.73263
2019-04-09 15:36 | 1554838559.7158 | 601 | 0.20954 | 0.000454 | 1e-06 | 100.7682 | 101.4795 | 121.5295 | 20.6893 | 0.00215 | 0.002076 | -154.5716 | 0.23308 | 8.4019 | 0.002724 | 0.004128 | 0.003994 | 4.3e-05 | 0.001135 | 0.26969 | 1.3584 | 2.5477 | 98.43 | 20.02 | 48.096 | 73.0071 | 86.7912 | 71.2093 | 0.01145 | 1.0252 | 0.023941 | 23.3518 | -0.062918 | 0.13866 | -5e-06 | -0.000421 | -0.000421 | -13.2577
2019-04-09 15:36 | 1554838560.7149 | 602 | 0.2095 | 0.000452 | 1e-06 | 100.7328 | 101.3993 | 121.2736 | 20.69 | 0.002162 | 0.002072 | -154.5691 | 0.23222 | 8.404 | 0.002807 | 0.004215 | 0.003986 | 4.3e-05 | 0.001133 | 0.26537 | 1.3498 | 2.5477 | 98.43 | 20.02 | 48.091 | 73.2968 | 88.7918 | 71.1234 | 0.011449 | 1.0255 | 0.023919 | 23.3243 | -0.058994 | 0.13815 | 0.000225 | 0.016122 | 0.016122 | 0.34533
2019-04-09 15:36 | 1554838561.7153 | 603 | 0.20955 | 0.00045 | 1e-06 | 100.7667 | 101.3116 | 120.9544 | 20.69 | 0.002129 | 0.002109 | -154.5682 | 0.2332 | 8.4024 | 0.002823 | 0.003935 | 0.003897 | 4.7e-05 | 0.001132 | 0.27367 | 1.3359 | 2.5476 | 98.43 | 20.02 | 48.08 | 72.4955 | 89.1879 | 72.0207 | 0.011447 | 1.0228 | 0.023856 | 23.324 | -0.048235 | 0.13873 | -7.2e-05 | -0.005184 | -0.005184 | -1.0784
2019-04-09 15:36 | 1554838562.7149 | 604 | 0.20952 | 0.000452 | 1e-06 | 100.7637 | 101.6131 | 120.6894 | 20.6903 | 0.002161 | 0.002043 | -154.5532 | 0.23299 | 8.4023 | 0.002744 | 0.004151 | 0.003875 | 4.4e-05 | 0.001137 | 0.27041 | 1.3355 | 2.5476 | 98.43 | 20.02 | 48.075 | 73.2675 | 87.2723 | 70.4384 | 0.011445 | 1.0275 | 0.023885 | 23.2449 | -0.075077 | 0.1386 | 0.000133 | 0.009428 | 0.009428 | 0.59244
2019-04-09 15:36 | 1554838563.7153 | 605 | 0.20959 | 0.000449 | 1e-06 | 100.767 | 101.4382 | 120.6037 | 20.6903 | 0.002143 | 0.002082 | -154.5505 | 0.23317 | 8.4019 | 0.002754 | 0.003946 | 0.003754 | 4.5e-05 | 0.001137 | 0.27184 | 1.39 | 2.5476 | 98.43 | 20.02 | 48.072 | 72.824 | 87.5199 | 71.3763 | 0.011445 | 1.0247 | 0.023844 | 23.2684 | -0.059383 | 0.13871 | -0.000277 | -0.019862 | -0.019862 | -0.28145
2019-04-09 15:36 | 1554838564.714 | 606 | 0.20952 | 0.000451 | 1e-06 | 100.7503 | 101.4561 | 120.6284 | 20.6904 | 0.002169 | 0.002029 | -154.5736 | 0.23356 | 8.401 | 0.002713 | 0.004723 | 0.004569 | 4.4e-05 | 0.001142 | 0.2672 | 1.4015 | 2.5476 | 98.43 | 20.02 | 48.066 | 73.4461 | 86.5353 | 70.0883 | 0.011443 | 1.0286 | 0.023891 | 23.2272 | -0.062441 | 0.13894 | 9.9e-05 | 0.007047 | 0.007047 | 0.79464
2019-04-09 15:36 | 1554838565.715 | 607 | 0.20956 | 0.000457 | 1e-06 | 100.7816 | 101.578 | 120.8758 | 20.6905 | 0.00214 | 0.002075 | -154.5677 | 0.23328 | 8.4021 | 0.002751 | 0.004373 | 0.00423 | 4.4e-05 | 0.001149 | 0.28871 | 1.3671 | 2.5477 | 98.43 | 20.02 | 48.063 | 72.7509 | 87.4335 | 71.201 | 0.011443 | 1.0252 | 0.023877 | 23.2887 | -0.0704 | 0.13878 | -0.000122 | -0.008818 | -0.008818 | -0.63427
2019-04-09 15:36 | 1554838566.7151 | 608 | 0.20957 | 0.00045 | 2e-06 | 100.8111 | 101.377 | 121.3384 | 20.69 | 0.002168 | 0.002088 | -154.5709 | 0.23288 | 8.4032 | 0.002791 | 0.004001 | 0.003833 | 4.6e-05 | 0.001145 | 0.26034 | 1.3821 | 2.5477 | 98.43 | 20.03 | 48.061 | 73.4381 | 88.3972 | 71.5132 | 0.011449 | 1.0243 | 0.023911 | 23.3438 | -0.050069 | 0.13854 | -0.000176 | -0.012711 | -0.012711 | -0.43923
2019-04-09 15:36 | 1554838567.7151 | 609 | 0.20957 | 0.00045 | 1e-06 | 100.8108 | 101.4465 | 121.5241 | 20.6906 | 0.002197 | 0.002101 | -154.5681 | 0.23309 | 8.4013 | 0.002808 | 0.004629 | 0.004489 | 4.6e-05 | 0.001149 | 0.26737 | 1.3798 | 2.5477 | 98.43 | 20.03 | 48.059 | 74.1245 | 88.8196 | 71.8233 | 0.011449 | 1.0234 | 0.023919 | 23.3721 | -0.056227 | 0.13867 | -0.000172 | -0.012382 | -0.012382 | -0.45131
2019-04-09 15:36 | 1554838568.7157 | 610 | 0.20953 | 0.00045 | 1e-06 | 100.8272 | 101.6351 | 121.6076 | 20.6905 | 0.002165 | 0.002114 | -154.5533 | 0.23344 | 8.4017 | 0.002809 | 0.004121 | 0.003939 | 4.5e-05 | 0.001145 | 0.27007 | 1.3669 | 2.5476 | 98.43 | 20.02 | 48.059 | 73.3534 | 88.8382 | 72.1326 | 0.011442 | 1.0225 | 0.023917 | 23.3906 | -0.071387 | 0.13887 | 5.3e-05 | 0.003776 | 0.003776 | 1.4819
2019-04-09 15:36 | 1554838569.7152 | 611 | 0.20961 | 0.000456 | 1e-06 | 100.8007 | 101.6224 | 121.7083 | 20.69 | 0.002213 | 0.002028 | -154.5576 | 0.23321 | 8.402 | 0.002743 | 0.00418 | 0.004017 | 4.6e-05 | 0.001145 | 0.26609 | 1.3226 | 2.5476 | 98.43 | 20.02 | 48.059 | 74.5151 | 87.2525 | 70.0584 | 0.011442 | 1.0287 | 0.023999 | 23.3299 | -0.072617 | 0.13874 | -0.000396 | -0.028548 | -0.028548 | -0.19585
2019-04-09 15:36 | 1554838570.7151 | 612 | 0.20952 | 0.000452 | 1e-06 | 100.8346 | 101.6058 | 121.79 | 20.6906 | 0.002106 | 0.002086 | -154.5637 | 0.23301 | 8.4014 | 0.002839 | 0.004433 | 0.004208 | 4.8e-05 | 0.001141 | 0.25033 | 1.3353 | 2.5475 | 98.43 | 20.02 | 48.058 | 71.9511 | 89.5578 | 71.4522 | 0.011441 | 1.0245 | 0.023958 | 23.3851 | -0.068152 | 0.13862 | 0.000109 | 0.007779 | 0.007779 | 0.71811
2019-04-09 15:36 | 1554838571.7143 | 613 | 0.20958 | 0.00045 | 1e-06 | 100.8882 | 101.7297 | 121.7703 | 20.6909 | 0.002157 | 0.002106 | -154.5741 | 0.2337 | 8.4012 | 0.002827 | 0.003519 | 0.003434 | 4.6e-05 | 0.001138 | 0.27113 | 1.3818 | 2.5477 | 98.43 | 20.02 | 48.058 | 73.1698 | 89.2654 | 71.9396 | 0.011441 | 1.0231 | 0.023939 | 23.3997 | -0.074298 | 0.13903 | -0.000207 | -0.014881 | -0.014881 | -0.37651
2019-04-09 15:36 | 1554838572.7152 | 614 | 0.20954 | 0.000446 | 1e-06 | 100.7804 | 101.5404 | 121.9276 | 20.6907 | 0.002196 | 0.002083 | -154.5742 | 0.23362 | 8.4005 | 0.002807 | 0.004009 | 0.003767 | 4.4e-05 | 0.001132 | 0.25785 | 1.3378 | 2.5477 | 98.43 | 20.03 | 48.059 | 74.1083 | 88.7932 | 71.3899 | 0.011449 | 1.0247 | 0.023974 | 23.3962 | -0.067202 | 0.13898 | 1e-05 | 0.000692 | 0.000692 | 8.0914
2019-04-09 15:36 | 1554838573.7142 | 615 | 0.20956 | 0.00045 | 1e-06 | 100.9052 | 101.7111 | 121.6844 | 20.6907 | 0.002192 | 0.002094 | -154.5909 | 0.23267 | 8.4005 | 0.002745 | 0.00402 | 0.003861 | 4.7e-05 | 0.001128 | 0.27848 | 1.3433 | 2.5474 | 98.43 | 20.03 | 48.059 | 74.0031 | 87.3054 | 71.6612 | 0.011449 | 1.0239 | 0.02394 | 23.3821 | -0.071159 | 0.13841 | -0.000127 | -0.009202 | -0.009202 | -0.60616
2019-04-09 15:36 | 1554838574.7147 | 616 | 0.20954 | 0.00045 | 1e-06 | 100.8504 | 101.504 | 121.6114 | 20.6903 | 0.002159 | 0.002039 | -154.5851 | 0.23278 | 8.4026 | 0.002787 | 0.004095 | 0.003963 | 4.4e-05 | 0.001124 | 0.26305 | 1.3347 | 2.5476 | 98.43 | 20.03 | 48.054 | 73.2184 | 88.2993 | 70.3407 | 0.011448 | 1.0278 | 0.023979 | 23.3302 | -0.057779 | 0.13848 | 1.3e-05 | 0.000924 | 0.000924 | 6.0393
2019-04-09 15:36 | 1554838575.7147 | 617 | 0.20955 | 0.000453 | 2e-06 | 100.8093 | 101.7063 | 121.4189 | 20.6909 | 0.002165 | 0.00209 | -154.5824 | 0.23354 | 8.4013 | 0.002858 | 0.004074 | 0.003938 | 4.6e-05 | 0.001124 | 0.29551 | 1.3202 | 2.5477 | 98.43 | 20.03 | 48.052 | 73.3521 | 90.0309 | 71.5467 | 0.011447 | 1.0242 | 0.023918 | 23.3527 | -0.079239 | 0.13893 | -7.8e-05 | -0.005719 | -0.005719 | -0.97902
2019-04-09 15:36 | 1554838576.7143 | 618 | 0.2096 | 0.00045 | 1e-06 | 100.8331 | 101.7008 | 121.0382 | 20.6907 | 0.002173 | 0.002083 | -154.6059 | 0.23366 | 8.4022 | 0.002907 | 0.003667 | 0.00351 | 4.6e-05 | 0.001121 | 0.286 | 1.3303 | 2.5476 | 98.43 | 20.03 | 48.05 | 73.5429 | 91.2048 | 71.3985 | 0.011447 | 1.0247 | 0.023886 | 23.311 | -0.076642 | 0.13901 | -0.000337 | -0.024227 | -0.024227 | -0.23123
2019-04-09 15:36 | 1554838577.7154 | 619 | 0.20948 | 0.000452 | 1e-06 | 100.9111 | 101.6705 | 120.8218 | 20.6898 | 0.002111 | 0.002008 | -154.5873 | 0.23374 | 8.4029 | 0.00293 | 0.003967 | 0.003708 | 4.5e-05 | 0.001117 | 0.26416 | 1.3309 | 2.5476 | 98.43 | 20.03 | 48.043 | 72.0548 | 91.7589 | 69.5832 | 0.011445 | 1.0301 | 0.023928 | 23.2287 | -0.067059 | 0.13905 | 0.000349 | 0.025024 | 0.025024 | 0.22394
2019-04-09 15:36 | 1554838578.7135 | 620 | 0.20963 | 0.000451 | 1e-06 | 100.8307 | 101.6189 | 120.595 | 20.6906 | 0.002144 | 0.002092 | -154.6139 | 0.2335 | 8.4023 | 0.002871 | 0.003522 | 0.003357 | 4.7e-05 | 0.001114 | 0.26288 | 1.3468 | 2.5476 | 98.43 | 20.03 | 48.038 | 72.8437 | 90.3449 | 71.6159 | 0.011444 | 1.024 | 0.023835 | 23.2756 | -0.069646 | 0.13891 | -0.000527 | -0.037735 | -0.037735 | -0.14835
2019-04-09 15:36 | 1554838579.7134 | 621 | 0.20953 | 0.000452 | 1e-06 | 100.8775 | 101.6843 | 120.6917 | 20.6918 | 0.002136 | 0.002095 | -154.6032 | 0.2333 | 8.4015 | 0.002895 | 0.004405 | 0.004222 | 4.8e-05 | 0.00111 | 0.26737 | 1.3545 | 2.5476 | 98.43 | 20.03 | 48.027 | 72.6555 | 90.9151 | 71.6723 | 0.011441 | 1.0238 | 0.023842 | 23.2869 | -0.071253 | 0.13879 | 8.9e-05 | 0.006292 | 0.006292 | 0.8889
2019-04-09 15:36 | 1554838580.7135 | 622 | 0.20947 | 0.000448 | 1e-06 | 100.8805 | 101.6392 | 120.7502 | 20.6918 | 0.002116 | 0.002062 | -154.5919 | 0.23395 | 8.4042 | 0.002809 | 0.004164 | 0.004044 | 4.6e-05 | 0.00111 | 0.26233 | 1.33 | 2.5476 | 98.43 | 20.03 | 48.021 | 72.1822 | 88.8513 | 70.8966 | 0.01144 | 1.0262 | 0.023875 | 23.2663 | -0.067023 | 0.13918 | 0.000417 | 0.029803 | 0.029803 | 0.1882
2019-04-09 15:36 | 1554838581.7148 | 623 | 0.20956 | 0.000452 | 0 | 100.9198 | 101.5659 | 120.8943 | 20.6911 | 0.002157 | 0.002023 | -154.5974 | 0.23379 | 8.3994 | 0.002817 | 0.00388 | 0.003697 | 5e-05 | 0.001115 | 0.27848 | 1.3517 | 2.5476 | 98.43 | 20.04 | 48.014 | 73.1624 | 89.0257 | 69.9361 | 0.011445 | 1.029 | 0.023922 | 23.2476 | -0.057085 | 0.13908 | -0.000113 | -0.008104 | -0.008104 | -0.69166
2019-04-09 15:36 | 1554838582.7136 | 624 | 0.20957 | 0.00045 | 1e-06 | 100.8525 | 101.4709 | 120.9333 | 20.6907 | 0.002097 | 0.002072 | -154.6031 | 0.23332 | 8.4007 | 0.002801 | 0.004278 | 0.004094 | 4.7e-05 | 0.001119 | 0.24945 | 1.3576 | 2.5478 | 98.43 | 20.04 | 47.998 | 71.7313 | 88.6588 | 71.1318 | 0.011441 | 1.0255 | 0.023885 | 23.2919 | -0.054683 | 0.1388 | -0.000169 | -0.012156 | -0.012156 | -0.46015
2019-04-09 15:36 | 1554838583.7136 | 625 | 0.20955 | 0.000451 | 1e-06 | 100.8301 | 101.4186 | 121.2436 | 20.6915 | 0.002178 | 0.002048 | -154.5857 | 0.23349 | 8.4042 | 0.002821 | 0.004072 | 0.003863 | 4.7e-05 | 0.001121 | 0.26952 | 1.3352 | 2.5476 | 98.43 | 20.04 | 47.99 | 73.6677 | 89.126 | 70.5604 | 0.011439 | 1.0272 | 0.023935 | 23.3024 | -0.052051 | 0.1389 | -8.2e-05 | -0.005962 | -0.005962 | -0.93887
2019-04-09 15:36 | 1554838584.7146 | 626 | 0.20951 | 0.000451 | 1e-06 | 100.8331 | 101.6926 | 121.3952 | 20.6897 | 0.002146 | 0.002091 | -154.5794 | 0.23288 | 8.4014 | 0.002828 | 0.004364 | 0.004163 | 4.7e-05 | 0.00112 | 0.26626 | 1.3726 | 2.5476 | 98.43 | 20.04 | 47.981 | 72.9068 | 89.3087 | 71.5876 | 0.011437 | 1.0241 | 0.023915 | 23.3518 | -0.075915 | 0.13854 | 0.000163 | 0.011636 | 0.011636 | 0.47981
2019-04-09 15:36 | 1554838585.7136 | 627 | 0.20948 | 0.000448 | 1e-06 | 100.7887 | 101.4448 | 121.4576 | 20.6908 | 0.002094 | 0.002038 | -154.5787 | 0.23264 | 8.401 | 0.00279 | 0.003613 | 0.00351 | 4.7e-05 | 0.001119 | 0.26089 | 1.344 | 2.5477 | 98.43 | 20.04 | 47.966 | 71.6619 | 88.3948 | 70.3131 | 0.011434 | 1.0279 | 0.023965 | 23.3145 | -0.05803 | 0.13839 | 0.000348 | 0.024933 | 0.024933 | 0.22369
2019-04-09 15:36 | 1554838586.7137 | 628 | 0.2096 | 0.000445 | 1e-06 | 100.8671 | 101.59 | 121.6067 | 20.6916 | 0.002105 | 0.002107 | -154.5561 | 0.23304 | 8.4019 | 0.002814 | 0.004606 | 0.00444 | 5.2e-05 | 0.001126 | 0.26145 | 1.3162 | 2.5475 | 98.43 | 20.04 | 47.958 | 71.9257 | 88.9653 | 71.9608 | 0.011432 | 1.023 | 0.023922 | 23.3847 | -0.063879 | 0.13863 | -0.000351 | -0.025234 | -0.025234 | -0.22141
2019-04-09 15:36 | 1554838587.7141 | 629 | 0.20943 | 0.000446 | 1e-06 | 100.7334 | 101.4596 | 121.6788 | 20.6901 | 0.002168 | 0.002028 | -154.5709 | 0.23273 | 8.4021 | 0.002892 | 0.004431 | 0.004196 | 4.7e-05 | 0.00113 | 0.24851 | 1.3326 | 2.5476 | 98.43 | 20.04 | 47.941 | 73.4233 | 90.8432 | 70.0763 | 0.011428 | 1.0286 | 0.023995 | 23.3277 | -0.064252 | 0.13845 | 0.000692 | 0.049723 | 0.049723 | 0.11221
2019-04-09 15:36 | 1554838588.7147 | 630 | 0.20956 | 0.000449 | 1e-06 | 100.7664 | 101.3646 | 121.8839 | 20.6907 | 0.002124 | 0.002028 | -154.5382 | 0.23293 | 8.4022 | 0.002807 | 0.003746 | 0.003632 | 4.4e-05 | 0.00113 | 0.26626 | 1.398 | 2.5477 | 98.43 | 20.04 | 47.933 | 72.3839 | 88.8042 | 70.0746 | 0.011426 | 1.0286 | 0.024015 | 23.3473 | -0.05294 | 0.13857 | -0.00014 | -0.010171 | -0.010171 | -0.54907
2019-04-09 15:36 | 1554838589.7148 | 631 | 0.20957 | 0.000444 | 1e-06 | 100.7137 | 101.3109 | 121.5208 | 20.6908 | 0.002163 | 0.002109 | -154.553 | 0.23259 | 8.4016 | 0.002767 | 0.003548 | 0.003428 | 4.6e-05 | 0.001132 | 0.26913 | 1.3261 | 2.5475 | 98.43 | 20.05 | 47.928 | 73.321 | 87.8241 | 72.0183 | 0.011432 | 1.0228 | 0.023912 | 23.3784 | -0.052881 | 0.13837 | -0.000177 | -0.012716 | -0.012716 | -0.43853
2019-04-09 15:36 | 1554838590.7138 | 632 | 0.2097 | 0.000444 | 1e-06 | 100.7256 | 101.4152 | 121.4067 | 20.6908 | 0.002146 | 0.002048 | -154.564 | 0.23349 | 8.4028 | 0.002829 | 0.003954 | 0.003719 | 4.6e-05 | 0.001139 | 0.28744 | 1.3656 | 2.5476 | 98.43 | 20.05 | 47.917 | 72.9016 | 89.3307 | 70.5441 | 0.011429 | 1.0272 | 0.023952 | 23.3175 | -0.061025 | 0.13891 | -0.000945 | -0.067916 | -0.067916 | -0.082423
2019-04-09 15:37 | 1554838591.7139 | 633 | 0.20963 | 0.000446 | 1e-06 | 100.6583 | 101.4152 | 121.2654 | 20.6909 | 0.002119 | 0.002046 | -154.546 | 0.23321 | 8.4041 | 0.002805 | 0.004784 | 0.004705 | 4.7e-05 | 0.00115 | 0.28064 | 1.3238 | 2.5475 | 98.43 | 20.04 | 47.914 | 72.2508 | 88.7342 | 70.4959 | 0.011421 | 1.0274 | 0.02394 | 23.3023 | -0.067003 | 0.13873 | -0.000549 | -0.039463 | -0.039463 | -0.14168
2019-04-09 15:37 | 1554838592.7149 | 634 | 0.20935 | 0.000443 | 1e-06 | 100.6178 | 101.3732 | 120.9387 | 20.6914 | 0.002194 | 0.002082 | -154.5384 | 0.2328 | 8.4021 | 0.002786 | 0.004076 | 0.003956 | 4.6e-05 | 0.001147 | 0.28351 | 1.3241 | 2.5475 | 98.43 | 20.04 | 47.912 | 74.0573 | 88.2986 | 71.3624 | 0.011421 | 1.0248 | 0.023877 | 23.3002 | -0.066899 | 0.1385 | 0.001154 | 0.082539 | 0.082539 | 0.067621
2019-04-09 15:37 | 1554838593.7153 | 635 | 0.2098 | 0.000445 | 1e-06 | 100.5743 | 101.2548 | 120.5772 | 20.691 | 0.00216 | 0.00204 | -154.5389 | 0.23269 | 8.4022 | 0.002829 | 0.003801 | 0.003657 | 4.5e-05 | 0.001146 | 0.28102 | 1.354 | 2.5477 | 98.43 | 20.04 | 47.913 | 73.2355 | 89.3177 | 70.3484 | 0.011421 | 1.0278 | 0.023877 | 23.2311 | -0.060313 | 0.13843 | -0.001542 | -0.11045 | -0.11045 | -0.050505
2019-04-09 15:37 | 1554838594.7143 | 636 | 0.20924 | 0.000447 | 1e-06 | 100.5106 | 101.3312 | 120.4068 | 20.6908 | 0.002044 | 0.002018 | -154.5446 | 0.23365 | 8.3989 | 0.002743 | 0.003537 | 0.003387 | 4.4e-05 | 0.00115 | 0.26432 | 1.3665 | 2.5476 | 98.43 | 20.04 | 47.916 | 70.4552 | 87.2456 | 69.8219 | 0.011422 | 1.0294 | 0.023878 | 23.1969 | -0.072729 | 0.139 | 0.001844 | 0.13192 | 0.13192 | 0.042463
2019-04-09 15:37 | 1554838595.7129 | 637 | 0.20982 | 0.000452 | 1e-06 | 100.4945 | 101.1719 | 120.3901 | 20.6911 | 0.002158 | 0.0021 | -154.5465 | 0.23332 | 8.3996 | 0.002736 | 0.004247 | 0.004097 | 4.6e-05 | 0.00115 | 0.28926 | 1.3819 | 2.5476 | 98.43 | 20.05 | 47.921 | 73.1842 | 87.0918 | 71.787 | 0.01143 | 1.0235 | 0.023808 | 23.2616 | -0.060083 | 0.1388 | -0.00171 | -0.12214 | -0.12214 | -0.045799
2019-04-09 15:37 | 1554838596.7143 | 638 | 0.20947 | 0.000446 | 1e-06 | 100.4332 | 101.2558 | 120.526 | 20.6903 | 0.002145 | 0.002008 | -154.5299 | 0.23368 | 8.4002 | 0.002728 | 0.003582 | 0.003354 | 4.6e-05 | 0.001148 | 0.28208 | 1.3881 | 2.5477 | 98.43 | 20.05 | 47.925 | 72.8893 | 86.8802 | 69.5914 | 0.011431 | 1.0301 | 0.023898 | 23.2005 | -0.072966 | 0.13901 | 0.000437 | 0.031229 | 0.031229 | 0.17939
2019-04-09 15:37 | 1554838597.7135 | 639 | 0.20841 | 0.000453 | 1e-06 | 100.5198 | 101.0824 | 120.7549 | 20.6897 | 0.002068 | 0.001979 | -154.5209 | 0.23297 | 8.4023 | 0.002686 | 0.004227 | 0.004037 | 4.7e-05 | 0.001149 | 0.2724 | 1.3556 | 2.5475 | 98.43 | 20.05 | 47.929 | 71.0154 | 85.8845 | 68.8868 | 0.011432 | 1.0322 | 0.023945 | 23.1987 | -0.049916 | 0.13859 | 0.006831 | 0.4899 | 0.4899 | 0.011401
2019-04-09 15:37 | 1554838598.7131 | 640 | 0.20949 | 0.00045 | 1e-06 | 100.4308 | 101.0831 | 120.7957 | 20.6909 | 0.002149 | 0.002069 | -154.5425 | 0.23345 | 8.401 | 0.002686 | 0.004596 | 0.004504 | 4.6e-05 | 0.001147 | 0.26128 | 1.2928 | 2.5476 | 98.43 | 20.05 | 47.933 | 72.9718 | 85.8725 | 71.0625 | 0.011433 | 1.0257 | 0.023874 | 23.2763 | -0.057904 | 0.13888 | 0.000311 | 0.02215 | 0.02215 | 0.25268
2019-04-09 15:37 | 1554838599.7141 | 641 | 0.20938 | 0.000453 | 1e-06 | 100.4692 | 101.0191 | 121.056 | 20.6914 | 0.002169 | 0.002026 | -154.5318 | 0.23299 | 8.4024 | 0.002775 | 0.003893 | 0.003728 | 4.7e-05 | 0.001145 | 0.27544 | 1.3159 | 2.5476 | 98.43 | 20.05 | 47.938 | 73.4666 | 88.0183 | 70.0191 | 0.011434 | 1.0288 | 0.023936 | 23.266 | -0.048817 | 0.1386 | 0.000969 | 0.069444 | 0.069444 | 0.080435
2019-04-09 15:37 | 1554838600.7132 | 642 | 0.20916 | 0.000451 | 1e-06 | 100.4117 | 101.1581 | 121.1009 | 20.6905 | 0.002141 | 0.002038 | -154.5383 | 0.23325 | 8.4029 | 0.002734 | 0.004187 | 0.004046 | 4.6e-05 | 0.001137 | 0.26432 | 1.3139 | 2.5476 | 98.43 | 20.05 | 47.944 | 72.7734 | 87.0351 | 70.3121 | 0.011436 | 1.0279 | 0.02393 | 23.2802 | -0.066238 | 0.13876 | 0.002289 | 0.16404 | 0.16404 | 0.034089
2019-04-09 15:37 | 1554838601.7136 | 643 | 0.20987 | 0.000449 | 1e-06 | 100.3921 | 100.9681 | 121.0286 | 20.6912 | 0.002172 | 0.002018 | -154.566 | 0.2339 | 8.4043 | 0.002675 | 0.003956 | 0.003811 | 4.4e-05 | 0.001134 | 0.24945 | 1.3293 | 2.5476 | 98.43 | 20.05 | 47.947 | 73.5226 | 85.6113 | 69.817 | 0.011436 | 1.0294 | 0.02394 | 23.2565 | -0.051175 | 0.13915 | -0.001985 | -0.1425 | -0.1425 | -0.039351
2019-04-09 15:37 | 1554838602.7132 | 644 | 0.20947 | 0.000445 | 1e-06 | 100.2956 | 100.9416 | 121.3185 | 20.6908 | 0.002136 | 0.002006 | -154.5649 | 0.23248 | 8.3999 | 0.002729 | 0.004309 | 0.004112 | 4.6e-05 | 0.001136 | 0.27312 | 1.3393 | 2.5476 | 98.43 | 20.05 | 47.947 | 72.6546 | 86.9199 | 69.5474 | 0.011436 | 1.0302 | 0.023978 | 23.2752 | -0.057428 | 0.1383 | 0.000432 | 0.031027 | 0.031027 | 0.17964
2019-04-09 15:37 | 1554838603.7132 | 645 | 0.2097 | 0.000448 | 2e-06 | 100.359 | 100.8556 | 121.5938 | 20.6903 | 0.002171 | 0.002023 | -154.5803 | 0.23314 | 8.4039 | 0.002708 | 0.004354 | 0.004302 | 4.5e-05 | 0.001139 | 0.2568 | 1.332 | 2.5475 | 98.43 | 20.05 | 47.947 | 73.5054 | 86.4152 | 69.9495 | 0.011436 | 1.029 | 0.023991 | 23.3153 | -0.044146 | 0.1387 | -0.000966 | -0.069585 | -0.069585 | -0.080326
2019-04-09 15:37 | 1554838604.7143 | 646 | 0.20947 | 0.000445 | 1e-06 | 100.3051 | 100.9726 | 121.9583 | 20.6917 | 0.002127 | 0.002026 | -154.5693 | 0.2332 | 8.401 | 0.002683 | 0.003846 | 0.003555 | 4.6e-05 | 0.001147 | 0.26985 | 1.3625 | 2.5476 | 98.43 | 20.05 | 47.946 | 72.437 | 85.8138 | 70.0225 | 0.011436 | 1.0288 | 0.024025 | 23.3527 | -0.059322 | 0.13873 | 0.000451 | 0.032443 | 0.032443 | 0.17233
2019-04-09 15:37 | 1554838605.7147 | 647 | 0.20953 | 0.000444 | 1e-06 | 100.2557 | 100.8824 | 121.8989 | 20.6908 | 0.002096 | 0.002038 | -154.5821 | 0.23261 | 8.3998 | 0.002636 | 0.004285 | 0.004155 | 4.6e-05 | 0.001147 | 0.26698 | 1.3512 | 2.5476 | 98.43 | 20.05 | 47.941 | 71.7002 | 84.6868 | 70.2977 | 0.011435 | 1.0279 | 0.024009 | 23.3563 | -0.055741 | 0.13838 | 7.8e-05 | 0.005556 | 0.005556 | 1.0037
2019-04-09 15:37 | 1554838606.7145 | 648 | 0.2096 | 0.000442 | 1e-06 | 100.2334 | 100.8422 | 121.6626 | 20.6918 | 0.002092 | 0.002072 | -154.5995 | 0.23352 | 8.3998 | 0.002586 | 0.003836 | 0.00361 | 4.7e-05 | 0.001152 | 0.29015 | 1.3255 | 2.5476 | 98.43 | 20.05 | 47.939 | 71.6147 | 83.4669 | 71.1243 | 0.011434 | 1.0255 | 0.023957 | 23.3617 | -0.054163 | 0.13892 | -0.000318 | -0.022895 | -0.022895 | -0.24453
2019-04-09 15:37 | 1554838607.7147 | 649 | 0.20954 | 0.000443 | 1e-06 | 100.1735 | 100.8793 | 121.4565 | 20.6913 | 0.002137 | 0.002003 | -154.6158 | 0.23353 | 8.4009 | 0.002681 | 0.003989 | 0.003825 | 4.9e-05 | 0.00115 | 0.27273 | 1.38 | 2.5477 | 98.43 | 20.05 | 47.935 | 72.6971 | 85.772 | 69.4652 | 0.011433 | 1.0304 | 0.023995 | 23.2856 | -0.062802 | 0.13892 | 8e-06 | 0.000513 | 0.000513 | 10.921
2019-04-09 15:37 | 1554838608.7138 | 650 | 0.20946 | 0.000444 | 1e-06 | 100.144 | 100.6742 | 121.5255 | 20.6908 | 0.002057 | 0.002016 | -154.6208 | 0.2336 | 8.4011 | 0.002637 | 0.004324 | 0.004168 | 4.6e-05 | 0.001147 | 0.27151 | 1.3737 | 2.5476 | 98.43 | 20.05 | 47.928 | 70.7529 | 84.701 | 69.7774 | 0.011432 | 1.0295 | 0.02399 | 23.3029 | -0.047229 | 0.13897 | 0.000491 | 0.035232 | 0.035232 | 0.15896
2019-04-09 15:37 | 1554838609.7144 | 651 | 0.20965 | 0.000443 | 1e-06 | 100.1387 | 100.8339 | 121.2891 | 20.6908 | 0.002137 | 0.001989 | -154.644 | 0.23326 | 8.4016 | 0.002648 | 0.003546 | 0.003332 | 4.5e-05 | 0.001151 | 0.27312 | 1.3879 | 2.5475 | 98.43 | 20.05 | 47.925 | 72.6904 | 84.9663 | 69.1311 | 0.011431 | 1.0314 | 0.02399 | 23.2582 | -0.061884 | 0.13877 | -0.000674 | -0.048506 | -0.048506 | -0.11529
2019-04-09 15:37 | 1554838610.7135 | 652 | 0.20948 | 0.000448 | 1e-06 | 100.1381 | 100.7809 | 121.4116 | 20.691 | 0.002079 | 0.00202 | -154.6327 | 0.23269 | 8.4035 | 0.002669 | 0.004032 | 0.003768 | 4.6e-05 | 0.001147 | 0.2589 | 1.3731 | 2.5475 | 98.43 | 20.05 | 47.918 | 71.2981 | 85.4759 | 69.8745 | 0.011429 | 1.0292 | 0.023976 | 23.2952 | -0.057235 | 0.13843 | 0.00038 | 0.027277 | 0.027277 | 0.20452
2019-04-09 15:37 | 1554838611.7134 | 653 | 0.20966 | 0.000445 | 1e-06 | 100.0815 | 100.7434 | 121.7405 | 20.6904 | 0.002141 | 0.002023 | -154.623 | 0.233 | 8.3994 | 0.002619 | 0.00425 | 0.00413 | 4.7e-05 | 0.001145 | 0.2568 | 1.3524 | 2.5476 | 98.43 | 20.04 | 47.913 | 72.7703 | 84.2769 | 69.9472 | 0.011421 | 1.029 | 0.024006 | 23.3292 | -0.05896 | 0.13861 | -0.000703 | -0.050679 | -0.050679 | -0.11023
2019-04-09 15:37 | 1554838612.713 | 654 | 0.20947 | 0.000443 | 1e-06 | 100.1238 | 100.625 | 121.7433 | 20.69 | 0.002074 | 0.001944 | -154.6316 | 0.2336 | 8.4006 | 0.002735 | 0.00378 | 0.003632 | 4.7e-05 | 0.001145 | 0.2589 | 1.3499 | 2.5476 | 98.43 | 20.04 | 47.91 | 71.1648 | 87.0535 | 68.0525 | 0.01142 | 1.0347 | 0.024073 | 23.265 | -0.044664 | 0.13897 | 0.000458 | 0.033045 | 0.033045 | 0.16948
2019-04-09 15:37 | 1554838613.7135 | 655 | 0.20955 | 0.000443 | 1e-06 | 100.0523 | 100.6088 | 121.4337 | 20.6901 | 0.002085 | 0.001971 | -154.6449 | 0.23364 | 8.4019 | 0.002661 | 0.004251 | 0.004121 | 4.7e-05 | 0.001138 | 0.27527 | 1.2953 | 2.5477 | 98.43 | 20.04 | 47.907 | 71.4431 | 85.2886 | 68.6951 | 0.01142 | 1.0328 | 0.024019 | 23.2573 | -0.049614 | 0.13899 | -3.7e-05 | -0.002749 | -0.002749 | -2.0376
2019-04-09 15:37 | 1554838614.7136 | 656 | 0.20962 | 0.000445 | 1e-06 | 100.0228 | 100.6398 | 120.9753 | 20.6905 | 0.002181 | 0.002044 | -154.6509 | 0.23327 | 8.4032 | 0.002621 | 0.00425 | 0.004036 | 4.6e-05 | 0.001134 | 0.29192 | 1.3192 | 2.5475 | 98.43 | 20.04 | 47.903 | 73.7439 | 84.3058 | 70.4524 | 0.011419 | 1.0275 | 0.023913 | 23.2729 | -0.055003 | 0.13877 | -0.000481 | -0.034511 | -0.034511 | -0.16205
2019-04-09 15:37 | 1554838615.7127 | 657 | 0.20956 | 0.000443 | 1e-06 | 99.9499 | 100.5613 | 121.079 | 20.6911 | 0.002119 | 0.002035 | -154.6316 | 0.23372 | 8.4016 | 0.002767 | 0.00432 | 0.004226 | 4.5e-05 | 0.001128 | 0.28119 | 1.3691 | 2.5477 | 98.43 | 20.04 | 47.894 | 72.2649 | 87.8391 | 70.2457 | 0.011417 | 1.0281 | 0.02393 | 23.2759 | -0.054553 | 0.13904 | -9.2e-05 | -0.006706 | -0.006706 | -0.83553
2019-04-09 15:37 | 1554838616.7127 | 658 | 0.20954 | 0.000448 | 1e-06 | 99.9064 | 100.466 | 121.4534 | 20.6897 | 0.002064 | 0.002009 | -154.6402 | 0.23334 | 8.4039 | 0.002644 | 0.004096 | 0.00393 | 4.3e-05 | 0.001124 | 0.30414 | 1.3248 | 2.5475 | 98.43 | 20.04 | 47.89 | 70.9283 | 84.862 | 69.6033 | 0.011416 | 1.03 | 0.023989 | 23.29 | -0.049961 | 0.13881 | 3.6e-05 | 0.002541 | 0.002541 | 2.2011
2019-04-09 15:37 | 1554838617.7127 | 659 | 0.20957 | 0.000445 | 1e-06 | 99.8671 | 100.4299 | 121.7126 | 20.6907 | 0.002103 | 0.002036 | -154.6272 | 0.23267 | 8.4044 | 0.002716 | 0.004526 | 0.004405 | 4.3e-05 | 0.001121 | 0.2724 | 1.3526 | 2.5477 | 98.43 | 20.04 | 47.88 | 71.8733 | 86.6068 | 70.2484 | 0.011413 | 1.0281 | 0.023992 | 23.3368 | -0.050264 | 0.13842 | -0.000148 | -0.010649 | -0.010649 | -0.52384
2019-04-09 15:37 | 1554838618.7131 | 660 | 0.20951 | 0.000441 | 1e-06 | 99.8385 | 100.37 | 121.7604 | 20.6895 | 0.00209 | 0.002016 | -154.6033 | 0.2336 | 8.3995 | 0.002719 | 0.003655 | 0.003412 | 4.7e-05 | 0.001125 | 0.28926 | 1.362 | 2.5476 | 98.43 | 20.04 | 47.876 | 71.5555 | 86.6829 | 69.7841 | 0.011412 | 1.0295 | 0.024013 | 23.3256 | -0.047489 | 0.13897 | 0.000221 | 0.015863 | 0.015863 | 0.35305
2019-04-09 15:37 | 1554838619.7128 | 661 | 0.20955 | 0.000446 | 1e-06 | 99.8546 | 100.3992 | 121.7881 | 20.6902 | 0.002105 | 0.001982 | -154.589 | 0.23329 | 8.4008 | 0.002762 | 0.004052 | 0.003887 | 4.6e-05 | 0.001128 | 0.27887 | 1.3896 | 2.5476 | 98.43 | 20.04 | 47.871 | 71.9087 | 87.7133 | 68.9676 | 0.011411 | 1.0319 | 0.024045 | 23.3005 | -0.048655 | 0.13878 | -7.8e-05 | -0.005702 | -0.005702 | -0.98081
2019-04-09 15:37 | 1554838620.7138 | 662 | 0.20962 | 0.000444 | 1e-06 | 99.8209 | 100.3242 | 121.8252 | 20.6901 | 0.002107 | 0.002076 | -154.5908 | 0.23323 | 8.4004 | 0.002789 | 0.004453 | 0.004256 | 5e-05 | 0.001124 | 0.26681 | 1.3735 | 2.5474 | 98.43 | 20.04 | 47.864 | 71.9763 | 88.3591 | 71.2108 | 0.011409 | 1.0252 | 0.02397 | 23.3803 | -0.044984 | 0.13875 | -0.000463 | -0.033381 | -0.033381 | -0.16751
2019-04-09 15:37 | 1554838621.7138 | 663 | 0.2095 | 0.000445 | 1e-06 | 99.7965 | 100.3077 | 121.7393 | 20.6899 | 0.002082 | 0.002069 | -154.5808 | 0.23316 | 8.4012 | 0.002774 | 0.004146 | 0.00403 | 4.8e-05 | 0.001131 | 0.26681 | 1.3803 | 2.5476 | 98.43 | 20.05 | 47.859 | 71.3533 | 88.0004 | 71.0483 | 0.011415 | 1.0257 | 0.023967 | 23.3665 | -0.045699 | 0.13871 | 0.000277 | 0.019835 | 0.019835 | 0.28182
2019-04-09 15:37 | 1554838622.7132 | 664 | 0.20954 | 0.000442 | 1e-06 | 99.7283 | 100.4195 | 121.7018 | 20.6896 | 0.002101 | 0.001939 | -154.588 | 0.23371 | 8.4003 | 0.002702 | 0.003403 | 0.00326 | 4.5e-05 | 0.001128 | 0.28783 | 1.416 | 2.5477 | 98.43 | 20.05 | 47.85 | 71.8122 | 86.2677 | 67.9311 | 0.011413 | 1.0351 | 0.024073 | 23.2569 | -0.061781 | 0.13903 | -6e-06 | -0.000548 | -0.000548 | -10.2322
2019-04-09 15:37 | 1554838623.7142 | 665 | 0.20957 | 0.000442 | 1e-06 | 99.6848 | 100.3338 | 121.8879 | 20.6901 | 0.002025 | 0.00202 | -154.5734 | 0.23349 | 8.4042 | 0.002857 | 0.003757 | 0.00361 | 4.6e-05 | 0.001128 | 0.27223 | 1.393 | 2.5476 | 98.43 | 20.05 | 47.846 | 70.0015 | 89.996 | 69.865 | 0.011412 | 1.0292 | 0.024023 | 23.3406 | -0.058045 | 0.13891 | -0.000187 | -0.013537 | -0.013537 | -0.41352
2019-04-09 15:37 | 1554838624.7138 | 666 | 0.20955 | 0.000446 | 1e-06 | 99.7069 | 100.3036 | 121.9175 | 20.6899 | 0.002085 | 0.001957 | -154.5789 | 0.23288 | 8.4028 | 0.002776 | 0.003273 | 0.003154 | 4.5e-05 | 0.001127 | 0.25929 | 1.3839 | 2.5476 | 98.43 | 20.05 | 47.84 | 71.4266 | 88.045 | 68.3697 | 0.011411 | 1.0337 | 0.024079 | 23.2925 | -0.053368 | 0.13854 | -2.9e-05 | -0.00211 | -0.00211 | -2.6459
2019-04-09 15:37 | 1554838625.7133 | 667 | 0.20956 | 0.000441 | 1e-06 | 99.6053 | 100.1335 | 121.8046 | 20.6902 | 0.002152 | 0.002004 | -154.5824 | 0.23258 | 8.3998 | 0.002847 | 0.004097 | 0.003846 | 4.8e-05 | 0.001129 | 0.26609 | 1.3589 | 2.5476 | 98.43 | 20.05 | 47.831 | 73.0444 | 89.7624 | 69.4978 | 0.011409 | 1.0303 | 0.024028 | 23.3201 | -0.047309 | 0.13836 | -0.000114 | -0.008277 | -0.008277 | -0.67363
2019-04-09 15:37 | 1554838626.712 | 668 | 0.20957 | 0.000448 | 1e-06 | 99.6369 | 100.1177 | 121.6708 | 20.6909 | 0.002086 | 0.002024 | -154.57 | 0.2325 | 8.4018 | 0.002782 | 0.004606 | 0.004451 | 5e-05 | 0.001128 | 0.25785 | 1.3523 | 2.5476 | 98.43 | 20.05 | 47.828 | 71.4711 | 88.1952 | 69.9633 | 0.011408 | 1.0289 | 0.023998 | 23.3231 | -0.04306 | 0.13832 | -0.000147 | -0.010621 | -0.010621 | -0.52484
2019-04-09 15:37 | 1554838627.7131 | 669 | 0.2095 | 0.000447 | 1e-06 | 99.5705 | 100.2808 | 121.5715 | 20.6893 | 0.00214 | 0.002032 | -154.5739 | 0.23344 | 8.401 | 0.002718 | 0.004532 | 0.004316 | 4.5e-05 | 0.001129 | 0.28008 | 1.3612 | 2.5477 | 98.43 | 20.05 | 47.827 | 72.7497 | 86.6507 | 70.155 | 0.011408 | 1.0284 | 0.023982 | 23.3201 | -0.063586 | 0.13887 | 0.000225 | 0.01611 | 0.01611 | 0.34739
2019-04-09 15:37 | 1554838628.713 | 670 | 0.20954 | 0.00044 | 1e-06 | 99.505 | 100.2729 | 121.2408 | 20.6897 | 0.002127 | 0.002017 | -154.5809 | 0.23288 | 8.4024 | 0.002762 | 0.003554 | 0.003407 | 4.6e-05 | 0.001124 | 0.25658 | 1.3567 | 2.5476 | 98.43 | 20.05 | 47.823 | 72.457 | 87.7052 | 69.8086 | 0.011407 | 1.0294 | 0.023961 | 23.2766 | -0.068767 | 0.13854 | 1.9e-05 | 0.001339 | 0.001339 | 4.1703
2019-04-09 15:37 | 1554838629.7131 | 671 | 0.20962 | 0.000446 | 1e-06 | 99.5169 | 99.9332 | 121.2365 | 20.69 | 0.002059 | 0.002018 | -154.5788 | 0.23268 | 8.4005 | 0.002727 | 0.004038 | 0.003769 | 4.9e-05 | 0.001127 | 0.28313 | 1.3925 | 2.5476 | 98.43 | 20.05 | 47.822 | 70.8208 | 86.8614 | 69.8172 | 0.011406 | 1.0294 | 0.02396 | 23.2765 | -0.037345 | 0.13842 | -0.000466 | -0.033529 | -0.033529 | -0.16637
2019-04-09 15:37 | 1554838630.7131 | 672 | 0.20955 | 0.000445 | 1e-06 | 99.4016 | 99.9825 | 121.052 | 20.6903 | 0.002098 | 0.002029 | -154.5828 | 0.23394 | 8.4041 | 0.002817 | 0.004326 | 0.004125 | 4.4e-05 | 0.001126 | 0.27959 | 1.3616 | 2.5476 | 98.43 | 20.05 | 47.818 | 71.7372 | 89.0242 | 70.093 | 0.011406 | 1.0286 | 0.023933 | 23.2681 | -0.052113 | 0.13917 | -5.7e-05 | -0.004142 | -0.004142 | -1.354
2019-04-09 15:37 | 1554838631.7135 | 673 | 0.2096 | 0.000448 | 1e-06 | 99.4484 | 99.9288 | 121.2224 | 20.6908 | 0.00213 | 0.002032 | -154.5881 | 0.23273 | 8.4025 | 0.002776 | 0.004343 | 0.004147 | 4.9e-05 | 0.001125 | 0.28838 | 1.3742 | 2.5477 | 98.43 | 20.05 | 47.815 | 72.5243 | 88.0451 | 70.1665 | 0.011405 | 1.0283 | 0.023947 | 23.287 | -0.043102 | 0.13845 | -0.000349 | -0.025111 | -0.025111 | -0.2222
2019-04-09 15:37 | 1554838632.7125 | 674 | 0.20953 | 0.000444 | 1e-06 | 99.3349 | 99.9012 | 121.5201 | 20.691 | 0.002099 | 0.002002 | -154.5832 | 0.23328 | 8.4024 | 0.002663 | 0.004125 | 0.003999 | 4.6e-05 | 0.001132 | 0.28528 | 1.3635 | 2.5476 | 98.43 | 20.05 | 47.811 | 71.784 | 85.327 | 69.4455 | 0.011404 | 1.0305 | 0.024001 | 23.2911 | -0.050848 | 0.13878 | 8.3e-05 | 0.005915 | 0.005915 | 0.94552
2019-04-09 15:37 | 1554838633.7132 | 675 | 0.20956 | 0.000446 | 1e-06 | 99.3459 | 99.7178 | 121.625 | 20.6916 | 0.002147 | 0.002016 | -154.5706 | 0.23298 | 8.4011 | 0.002778 | 0.004016 | 0.003848 | 4.8e-05 | 0.001129 | 0.27367 | 1.3747 | 2.5477 | 98.43 | 20.05 | 47.808 | 72.9336 | 88.0928 | 69.7713 | 0.011403 | 1.0295 | 0.024 | 23.3122 | -0.033418 | 0.1386 | -0.000127 | -0.009215 | -0.009215 | -0.60615
2019-04-09 15:37 | 1554838634.7122 | 676 | 0.20955 | 0.000445 | 1e-06 | 99.3162 | 99.8579 | 121.5203 | 20.6902 | 0.002119 | 0.002012 | -154.567 | 0.23276 | 8.3988 | 0.002834 | 0.004458 | 0.00427 | 4.7e-05 | 0.001131 | 0.262 | 1.3485 | 2.5476 | 98.43 | 20.05 | 47.806 | 72.2563 | 89.4481 | 69.6864 | 0.011403 | 1.0298 | 0.023993 | 23.2993 | -0.048654 | 0.13847 | -2.6e-05 | -0.001903 | -0.001903 | -2.9326
2019-04-09 15:37 | 1554838635.7132 | 677 | 0.20953 | 0.000447 | 1e-06 | 99.2947 | 99.9501 | 121.6201 | 20.6899 | 0.002096 | 0.002035 | -154.5758 | 0.23333 | 8.4007 | 0.002852 | 0.004391 | 0.004197 | 4.6e-05 | 0.001132 | 0.25874 | 1.3306 | 2.5476 | 98.43 | 20.05 | 47.807 | 71.7017 | 89.8676 | 70.2382 | 0.011403 | 1.0281 | 0.023984 | 23.3276 | -0.058844 | 0.13881 | 4.2e-05 | 0.002998 | 0.002998 | 1.866
2019-04-09 15:37 | 1554838636.7123 | 678 | 0.20952 | 0.000451 | 1e-06 | 99.2733 | 99.6951 | 121.3776 | 20.6903 | 0.002089 | 0.001971 | -154.5735 | 0.23347 | 8.4027 | 0.002796 | 0.004297 | 0.00419 | 4.7e-05 | 0.001134 | 0.26128 | 1.3744 | 2.5476 | 98.43 | 20.05 | 47.806 | 71.5298 | 88.5336 | 68.6891 | 0.011403 | 1.0328 | 0.024014 | 23.2517 | -0.037924 | 0.13889 | 0.000114 | 0.008156 | 0.008156 | 0.68629
2019-04-09 15:37 | 1554838637.7123 | 679 | 0.20957 | 0.000454 | 1e-06 | 99.2236 | 99.8279 | 121.5935 | 20.6894 | 0.002085 | 0.002015 | -154.5793 | 0.23273 | 8.4029 | 0.002811 | 0.004287 | 0.0041 | 4.6e-05 | 0.001135 | 0.28351 | 1.3962 | 2.5476 | 98.43 | 20.05 | 47.804 | 71.4362 | 88.8866 | 69.7513 | 0.011402 | 1.0296 | 0.023998 | 23.3085 | -0.054317 | 0.13845 | -0.000186 | -0.013393 | -0.013393 | -0.41662
2019-04-09 15:37 | 1554838638.7136 | 680 | 0.20948 | 0.000453 | 1e-06 | 99.1911 | 99.8128 | 121.8043 | 20.6919 | 0.002119 | 0.001971 | -154.5827 | 0.23355 | 8.3999 | 0.00283 | 0.004006 | 0.003896 | 4.6e-05 | 0.001137 | 0.25929 | 1.3301 | 2.5475 | 98.43 | 20.05 | 47.798 | 72.2626 | 89.3474 | 68.6954 | 0.011401 | 1.0328 | 0.024056 | 23.2927 | -0.055886 | 0.13894 | 0.000359 | 0.025801 | 0.025801 | 0.21701
2019-04-09 15:37 | 1554838639.7124 | 681 | 0.20954 | 0.000443 | 1e-06 | 99.2167 | 99.6366 | 121.9686 | 20.6906 | 0.002117 | 0.002026 | -154.5819 | 0.23339 | 8.4021 | 0.002796 | 0.004092 | 0.003931 | 4.7e-05 | 0.001141 | 0.26272 | 1.3286 | 2.5475 | 98.43 | 20.05 | 47.795 | 72.2037 | 88.5189 | 70.02 | 0.0114 | 1.0288 | 0.024026 | 23.3536 | -0.037773 | 0.13884 | -2e-06 | -0.000239 | -0.000239 | -23.4179
2019-04-09 15:37 | 1554838640.7134 | 682 | 0.20955 | 0.000447 | 1e-06 | 99.2688 | 99.672 | 121.7884 | 20.6906 | 0.002119 | 0.002057 | -154.5858 | 0.23375 | 8.4026 | 0.002744 | 0.004527 | 0.004258 | 4.6e-05 | 0.001144 | 0.25912 | 1.3273 | 2.5476 | 98.43 | 20.05 | 47.792 | 72.2529 | 87.2812 | 70.7614 | 0.011399 | 1.0266 | 0.023982 | 23.3615 | -0.036258 | 0.13906 | -2.5e-05 | -0.001837 | -0.001837 | -3.05
2019-04-09 15:37 | 1554838641.7125 | 683 | 0.20952 | 0.000449 | 1e-06 | 99.2504 | 99.7216 | 121.5438 | 20.6899 | 0.002056 | 0.00195 | -154.5748 | 0.23272 | 8.404 | 0.002761 | 0.004621 | 0.004543 | 4.6e-05 | 0.001144 | 0.25586 | 1.3345 | 2.5477 | 98.43 | 20.05 | 47.785 | 70.735 | 87.6796 | 68.2029 | 0.011398 | 1.0343 | 0.024047 | 23.2511 | -0.042368 | 0.13845 | 0.000149 | 0.010714 | 0.010714 | 0.52074
2019-04-09 15:37 | 1554838642.7115 | 684 | 0.20957 | 0.000449 | 1e-06 | 99.2447 | 99.8947 | 121.4382 | 20.6903 | 0.002119 | 0.001991 | -154.5812 | 0.2336 | 8.4 | 0.002719 | 0.00426 | 0.004118 | 4.5e-05 | 0.001139 | 0.26344 | 1.3272 | 2.5475 | 98.43 | 20.05 | 47.782 | 72.261 | 86.6829 | 69.1765 | 0.011397 | 1.0313 | 0.024003 | 23.2741 | -0.058391 | 0.13897 | -0.000186 | -0.013425 | -0.013425 | -0.41716
2019-04-09 15:37 | 1554838643.7125 | 685 | 0.2096 | 0.00045 | 1e-06 | 99.2328 | 99.7742 | 121.4673 | 20.6907 | 0.002073 | 0.002031 | -154.588 | 0.23235 | 8.4008 | 0.002685 | 0.00377 | 0.003628 | 4.4e-05 | 0.001141 | 0.29192 | 1.2969 | 2.5476 | 98.43 | 20.05 | 47.778 | 71.1575 | 85.8556 | 70.1313 | 0.011396 | 1.0284 | 0.023972 | 23.3093 | -0.048673 | 0.13822 | -0.000337 | -0.024293 | -0.024293 | -0.2293
2019-04-09 15:37 | 1554838644.7125 | 686 | 0.20954 | 0.000453 | 1e-06 | 99.2322 | 99.6201 | 121.7745 | 20.6905 | 0.002077 | 0.001914 | -154.5883 | 0.23311 | 8.402 | 0.002782 | 0.003956 | 0.003777 | 4.8e-05 | 0.001145 | 0.28854 | 1.3656 | 2.5476 | 98.43 | 20.05 | 47.776 | 71.2439 | 88.2011 | 67.3314 | 0.011396 | 1.0369 | 0.024101 | 23.2434 | -0.034894 | 0.13868 | -1.2e-05 | -0.000904 | -0.000904 | -6.1852
2019-04-09 15:37 | 1554838645.7117 | 687 | 0.20958 | 0.000453 | 1e-06 | 99.1449 | 99.64 | 121.7891 | 20.6907 | 0.002096 | 0.001986 | -154.5787 | 0.23419 | 8.4017 | 0.002828 | 0.004273 | 0.004096 | 4.5e-05 | 0.001146 | 0.28208 | 1.3564 | 2.5477 | 98.42 | 20.05 | 47.775 | 71.69 | 89.2955 | 69.0686 | 0.011396 | 1.0316 | 0.024041 | 23.304 | -0.044554 | 0.13932 | -0.000244 | -0.017606 | -0.017606 | -0.3189
2019-04-09 15:37 | 1554838646.7116 | 688 | 0.20954 | 0.000446 | 1e-06 | 99.1851 | 99.6896 | 122.1266 | 20.6908 | 0.002118 | 0.001976 | -154.5964 | 0.23314 | 8.4019 | 0.002706 | 0.004094 | 0.003982 | 4.8e-05 | 0.001152 | 0.274 | 1.3712 | 2.5475 | 98.42 | 20.05 | 47.769 | 72.2305 | 86.3547 | 68.8215 | 0.011395 | 1.0324 | 0.024083 | 23.3278 | -0.045376 | 0.13869 | 2.6e-05 | 0.001807 | 0.001807 | 3.0936
2019-04-09 15:37 | 1554838647.712 | 689 | 0.20955 | 0.00045 | 1e-06 | 99.1101 | 99.6225 | 121.8285 | 20.6907 | 0.002065 | 0.002026 | -154.5961 | 0.23372 | 8.4015 | 0.002755 | 0.003848 | 0.003689 | 5e-05 | 0.001151 | 0.27688 | 1.3428 | 2.5477 | 98.42 | 20.05 | 47.767 | 70.9532 | 87.5391 | 70.0285 | 0.011395 | 1.0288 | 0.024012 | 23.3404 | -0.046122 | 0.13904 | -3.5e-05 | -0.002528 | -0.002528 | -2.2168
2019-04-09 15:37 | 1554838648.7121 | 690 | 0.20955 | 0.00045 | 2e-06 | 99.1473 | 99.7006 | 121.7128 | 20.6913 | 0.002071 | 0.001998 | -154.604 | 0.23316 | 8.4019 | 0.00273 | 0.004508 | 0.004267 | 4.5e-05 | 0.001151 | 0.28456 | 1.3467 | 2.5475 | 98.42 | 20.05 | 47.763 | 71.1011 | 86.9438 | 69.3352 | 0.011394 | 1.0308 | 0.024024 | 23.3058 | -0.049775 | 0.1387 | -7.7e-05 | -0.005623 | -0.005623 | -0.99409
2019-04-09 15:37 | 1554838649.7128 | 691 | 0.20955 | 0.000451 | 1e-06 | 99.1455 | 99.6851 | 121.7266 | 20.6906 | 0.002133 | 0.001998 | -154.5908 | 0.23307 | 8.4031 | 0.002762 | 0.00437 | 0.004177 | 4.5e-05 | 0.001145 | 0.29015 | 1.3609 | 2.5476 | 98.42 | 20.04 | 47.761 | 72.5843 | 87.7169 | 69.3509 | 0.011386 | 1.0308 | 0.024025 | 23.3076 | -0.048547 | 0.13865 | -4.6e-05 | -0.003391 | -0.003391 | -1.6477
2019-04-09 15:38 | 1554838650.7116 | 692 | 0.20953 | 0.000451 | 1e-06 | 99.1256 | 99.6507 | 121.86 | 20.6911 | 0.002105 | 0.002026 | -154.5799 | 0.23318 | 8.4008 | 0.002764 | 0.004308 | 0.00415 | 4.7e-05 | 0.001148 | 0.28528 | 1.3805 | 2.5476 | 98.42 | 20.04 | 47.761 | 71.912 | 87.7674 | 70.0133 | 0.011386 | 1.0288 | 0.024015 | 23.3429 | -0.047258 | 0.13872 | 4e-05 | 0.002802 | 0.002802 | 1.995
2019-04-09 15:38 | 1554838651.7122 | 693 | 0.20958 | 0.000449 | 1e-06 | 99.08 | 99.8018 | 122.1306 | 20.6905 | 0.002027 | 0.001972 | -154.609 | 0.23336 | 8.4024 | 0.002732 | 0.00431 | 0.004185 | 4.9e-05 | 0.001148 | 0.26488 | 1.355 | 2.5476 | 98.42 | 20.04 | 47.76 | 70.0408 | 86.9875 | 68.7316 | 0.011386 | 1.0327 | 0.024087 | 23.3252 | -0.064921 | 0.13883 | -0.000199 | -0.014447 | -0.014447 | -0.38727
2019-04-09 15:38 | 1554838652.7122 | 694 | 0.2095 | 0.000451 | 2e-06 | 99.0827 | 99.6641 | 122.0649 | 20.6903 | 0.002083 | 0.002016 | -154.5898 | 0.23375 | 8.4012 | 0.002826 | 0.004036 | 0.003901 | 4.5e-05 | 0.001147 | 0.28727 | 1.3107 | 2.5475 | 98.42 | 20.04 | 47.761 | 71.377 | 89.2505 | 69.7863 | 0.011386 | 1.0295 | 0.024043 | 23.3548 | -0.052333 | 0.13906 | 0.000229 | 0.016426 | 0.016426 | 0.34117
2019-04-09 15:38 | 1554838653.7118 | 695 | 0.20954 | 0.00045 | 1e-06 | 99.1369 | 99.5636 | 121.8362 | 20.69 | 0.002076 | 0.002028 | -154.5827 | 0.23388 | 8.4024 | 0.002774 | 0.003903 | 0.003775 | 4.7e-05 | 0.001148 | 0.27704 | 1.3896 | 2.5475 | 98.43 | 20.04 | 47.761 | 71.2193 | 87.9921 | 70.0715 | 0.011385 | 1.0286 | 0.024011 | 23.3426 | -0.038419 | 0.13914 | -2e-06 | -0.000214 | -0.000214 | -26.2583
2019-04-09 15:38 | 1554838654.7118 | 696 | 0.20954 | 0.000452 | 1e-06 | 99.0675 | 99.6555 | 121.7156 | 20.6905 | 0.002075 | 0.002016 | -154.5909 | 0.23239 | 8.4034 | 0.002866 | 0.004412 | 0.00423 | 4.5e-05 | 0.001147 | 0.2891 | 1.3878 | 2.5476 | 98.43 | 20.04 | 47.761 | 71.201 | 90.2167 | 69.7781 | 0.011385 | 1.0295 | 0.024009 | 23.3211 | -0.052932 | 0.13825 | 2e-05 | 0.001347 | 0.001347 | 4.1365
2019-04-09 15:38 | 1554838655.7133 | 697 | 0.20954 | 0.000451 | 1e-06 | 99.0437 | 99.3946 | 121.7485 | 20.6897 | 0.002 | 0.001945 | -154.5904 | 0.23304 | 8.401 | 0.002911 | 0.00402 | 0.003925 | 4.6e-05 | 0.001157 | 0.28174 | 1.3835 | 2.5475 | 98.43 | 20.04 | 47.761 | 69.4014 | 91.2944 | 68.0785 | 0.011385 | 1.0346 | 0.024072 | 23.2664 | -0.031639 | 0.13863 | 1.8e-05 | 0.001241 | 0.001241 | 4.5014
2019-04-09 15:38 | 1554838656.7119 | 698 | 0.20955 | 0.00045 | 1e-06 | 99.0416 | 99.6218 | 121.3403 | 20.6909 | 0.002114 | 0.001992 | -154.5887 | 0.23222 | 8.4031 | 0.002926 | 0.004778 | 0.004669 | 4.9e-05 | 0.001155 | 0.28744 | 1.3469 | 2.5476 | 98.43 | 20.04 | 47.762 | 72.1237 | 91.6545 | 69.199 | 0.011385 | 1.0312 | 0.023992 | 23.2655 | -0.052244 | 0.13815 | -4.5e-05 | -0.003274 | -0.003274 | -1.7004
2019-04-09 15:38 | 1554838657.712 | 699 | 0.20952 | 0.000452 | 1e-06 | 99.0553 | 99.6445 | 121.2755 | 20.6908 | 0.002123 | 0.001993 | -154.5919 | 0.23312 | 8.4007 | 0.002829 | 0.004683 | 0.004508 | 4.6e-05 | 0.001154 | 0.27256 | 1.3555 | 2.5477 | 98.42 | 20.04 | 47.763 | 72.3395 | 89.3244 | 69.2214 | 0.011386 | 1.0312 | 0.023985 | 23.26 | -0.053046 | 0.13868 | 0.00016 | 0.011464 | 0.011464 | 0.48753
2019-04-09 15:38 | 1554838658.712 | 700 | 0.20954 | 0.000449 | 1e-06 | 99.0258 | 99.7336 | 121.5612 | 20.691 | 0.002078 | 0.001959 | -154.5808 | 0.23289 | 8.4037 | 0.002892 | 0.004162 | 0.003972 | 4.9e-05 | 0.00115 | 0.28456 | 1.3278 | 2.5476 | 98.42 | 20.04 | 47.763 | 71.2794 | 90.851 | 68.4152 | 0.011386 | 1.0336 | 0.024042 | 23.2599 | -0.063706 | 0.13855 | 1e-05 | 0.000682 | 0.000682 | 8.1827
2019-04-09 15:38 | 1554838659.7122 | 701 | 0.20953 | 0.000451 | 1e-06 | 99.03 | 99.5196 | 121.4011 | 20.6908 | 0.002022 | 0.001984 | -154.5814 | 0.23332 | 8.3994 | 0.002847 | 0.004327 | 0.004241 | 4.8e-05 | 0.001152 | 0.28263 | 1.3728 | 2.5477 | 98.42 | 20.04 | 47.764 | 69.9327 | 89.763 | 69.0189 | 0.011387 | 1.0318 | 0.024005 | 23.2652 | -0.044111 | 0.13881 | 4.6e-05 | 0.003264 | 0.003264 | 1.7139
2019-04-09 15:38 | 1554838660.7111 | 702 | 0.20955 | 0.000448 | 1e-06 | 98.9782 | 99.4645 | 121.517 | 20.6905 | 0.002052 | 0.001975 | -154.5762 | 0.23305 | 8.4023 | 0.002791 | 0.004734 | 0.004494 | 4.7e-05 | 0.001148 | 0.29767 | 1.3772 | 2.5477 | 98.42 | 20.04 | 47.763 | 70.6317 | 88.3994 | 68.8008 | 0.011386 | 1.0324 | 0.024024 | 23.2688 | -0.043842 | 0.13864 | -3.3e-05 | -0.002403 | -0.002403 | -2.325
2019-04-09 15:38 | 1554838661.7124 | 703 | 0.20955 | 0.000448 | 0 | 98.9585 | 99.4363 | 121.1861 | 20.691 | 0.002088 | 0.001955 | -154.5743 | 0.23242 | 8.404 | 0.002905 | 0.004398 | 0.00428 | 4.6e-05 | 0.001152 | 0.27527 | 1.3774 | 2.5476 | 98.42 | 20.04 | 47.764 | 71.4993 | 91.1488 | 68.3109 | 0.011387 | 1.0339 | 0.024008 | 23.2205 | -0.04308 | 0.13827 | -1.1e-05 | -0.000844 | -0.000844 | -6.6015
2019-04-09 15:38 | 1554838662.7115 | 704 | 0.20956 | 0.000449 | 1e-06 | 98.9835 | 99.4239 | 121.1469 | 20.6905 | 0.002139 | 0.00199 | -154.597 | 0.23362 | 8.4028 | 0.00281 | 0.004186 | 0.004043 | 5e-05 | 0.001147 | 0.28838 | 1.3473 | 2.5477 | 98.42 | 20.04 | 47.763 | 72.7429 | 88.8587 | 69.1479 | 0.011386 | 1.0314 | 0.023975 | 23.2452 | -0.039705 | 0.13898 | -0.000121 | -0.008767 | -0.008767 | -0.63887
2019-04-09 15:38 | 1554838663.7111 | 705 | 0.20955 | 0.000453 | 1e-06 | 98.854 | 99.5475 | 120.861 | 20.6897 | 0.002108 | 0.001961 | -154.606 | 0.23331 | 8.4026 | 0.002778 | 0.004092 | 0.00389 | 4.7e-05 | 0.00115 | 0.27472 | 1.3745 | 2.5477 | 98.42 | 20.04 | 47.762 | 71.9915 | 88.1042 | 68.4521 | 0.011386 | 1.0335 | 0.023971 | 23.1941 | -0.062526 | 0.13879 | -2e-05 | -0.001481 | -0.001481 | -3.7766
2019-04-09 15:38 | 1554838664.7122 | 706 | 0.2096 | 0.00045 | 1e-06 | 98.8686 | 99.2326 | 120.6932 | 20.691 | 0.002051 | 0.001984 | -154.6143 | 0.23326 | 8.4023 | 0.002807 | 0.004486 | 0.004366 | 4.4e-05 | 0.00115 | 0.27223 | 1.3893 | 2.5477 | 98.42 | 20.04 | 47.763 | 70.6145 | 88.7873 | 69.0191 | 0.011386 | 1.0318 | 0.023935 | 23.1972 | -0.032867 | 0.13877 | -0.000371 | -0.026673 | -0.026673 | -0.20966
2019-04-09 15:38 | 1554838665.7113 | 707 | 0.20955 | 0.000451 | 1e-06 | 98.7995 | 99.3272 | 120.5119 | 20.6906 | 0.002046 | 0.001983 | -154.6301 | 0.2335 | 8.4022 | 0.002786 | 0.004159 | 0.003933 | 4.5e-05 | 0.00115 | 0.28982 | 1.3971 | 2.5477 | 98.42 | 20.04 | 47.764 | 70.4908 | 88.2955 | 68.9836 | 0.011387 | 1.0319 | 0.023918 | 23.1786 | -0.047646 | 0.13891 | -8.1e-05 | -0.005836 | -0.005836 | -0.95916
2019-04-09 15:38 | 1554838666.7122 | 708 | 0.20959 | 0.000449 | 1e-06 | 98.8421 | 99.3534 | 120.5666 | 20.6909 | 0.002065 | 0.001989 | -154.6099 | 0.23344 | 8.4039 | 0.002798 | 0.004095 | 0.003909 | 4.6e-05 | 0.001153 | 0.26593 | 1.3796 | 2.5476 | 98.42 | 20.04 | 47.765 | 70.9534 | 88.5669 | 69.1233 | 0.011387 | 1.0315 | 0.023918 | 23.1886 | -0.046147 | 0.13887 | -0.000299 | -0.021496 | -0.021496 | -0.26036
2019-04-09 15:38 | 1554838667.7117 | 709 | 0.20961 | 0.000446 | 1e-06 | 98.7504 | 99.2828 | 120.7065 | 20.691 | 0.002113 | 0.002028 | -154.6073 | 0.23347 | 8.4018 | 0.002832 | 0.003998 | 0.003904 | 4.9e-05 | 0.001153 | 0.27759 | 1.3707 | 2.5475 | 98.42 | 20.04 | 47.765 | 72.1176 | 89.4061 | 70.0577 | 0.011387 | 1.0287 | 0.0239 | 23.2337 | -0.048097 | 0.13889 | -0.000429 | -0.030766 | -0.030766 | -0.18193
2019-04-09 15:38 | 1554838668.7117 | 710 | 0.20959 | 0.000449 | 1e-06 | 98.8084 | 99.2955 | 120.6694 | 20.6905 | 0.002124 | 0.001962 | -154.5993 | 0.23264 | 8.4007 | 0.002766 | 0.004399 | 0.004218 | 4.5e-05 | 0.001151 | 0.27455 | 1.3579 | 2.5476 | 98.42 | 20.03 | 47.765 | 72.3682 | 87.8075 | 68.4703 | 0.01138 | 1.0334 | 0.023951 | 23.1763 | -0.043986 | 0.13839 | -0.000286 | -0.020553 | -0.020553 | -0.27136
2019-04-09 15:38 | 1554838669.7113 | 711 | 0.20959 | 0.000446 | 1e-06 | 98.737 | 99.151 | 121.0168 | 20.6907 | 0.002071 | 0.001951 | -154.6085 | 0.23377 | 8.4037 | 0.002874 | 0.004115 | 0.00399 | 4.8e-05 | 0.001143 | 0.28694 | 1.3337 | 2.5477 | 98.42 | 20.03 | 47.767 | 71.0921 | 90.4004 | 68.2098 | 0.01138 | 1.0342 | 0.023995 | 23.2008 | -0.037428 | 0.13907 | -0.000264 | -0.019073 | -0.019073 | -0.29386
2019-04-09 15:38 | 1554838670.7128 | 712 | 0.20958 | 0.000447 | 1e-06 | 98.7822 | 99.3451 | 121.1318 | 20.6903 | 0.002102 | 0.002002 | -154.6261 | 0.23337 | 8.4017 | 0.002768 | 0.004492 | 0.004263 | 4.7e-05 | 0.001144 | 0.28224 | 1.3249 | 2.5475 | 98.42 | 20.03 | 47.769 | 71.8379 | 87.8456 | 69.4332 | 0.011381 | 1.0305 | 0.023964 | 23.2534 | -0.050821 | 0.13883 | -0.000238 | -0.017191 | -0.017191 | -0.32546
2019-04-09 15:38 | 1554838671.7114 | 713 | 0.20957 | 0.000442 | 1e-06 | 98.7524 | 99.1266 | 121.2417 | 20.6906 | 0.002127 | 0.002012 | -154.6022 | 0.23285 | 8.4039 | 0.002837 | 0.004417 | 0.004306 | 5e-05 | 0.001142 | 0.26969 | 1.3574 | 2.5477 | 98.42 | 20.03 | 47.772 | 72.4533 | 89.5042 | 69.6782 | 0.011382 | 1.0298 | 0.023966 | 23.2723 | -0.03382 | 0.13852 | -0.000173 | -0.012435 | -0.012435 | -0.44893
2019-04-09 15:38 | 1554838672.7108 | 714 | 0.20959 | 0.000449 | 1e-06 | 98.6911 | 99.0804 | 120.9187 | 20.6914 | 0.002089 | 0.002003 | -154.5995 | 0.23288 | 8.4036 | 0.002855 | 0.004342 | 0.004159 | 4.8e-05 | 0.001146 | 0.2677 | 1.3561 | 2.5475 | 98.42 | 20.03 | 47.778 | 71.5357 | 89.9534 | 69.46 | 0.011383 | 1.0305 | 0.023942 | 23.2339 | -0.035216 | 0.13854 | -0.000308 | -0.022174 | -0.022174 | -0.25179
2019-04-09 15:38 | 1554838673.7114 | 715 | 0.20957 | 0.000449 | 1e-06 | 98.6575 | 99.0732 | 121.0673 | 20.6913 | 0.002079 | 0.00201 | -154.6117 | 0.23275 | 8.4021 | 0.002864 | 0.004933 | 0.004727 | 4.9e-05 | 0.001151 | 0.28639 | 1.3877 | 2.5476 | 98.42 | 20.03 | 47.78 | 71.2991 | 90.1751 | 69.626 | 0.011383 | 1.03 | 0.02395 | 23.2538 | -0.037614 | 0.13846 | -0.000146 | -0.010525 | -0.010525 | -0.53016
2019-04-09 15:38 | 1554838674.7115 | 716 | 0.20958 | 0.000446 | 1e-06 | 98.6697 | 99.1534 | 121.2837 | 20.6918 | 0.002093 | 0.001951 | -154.578 | 0.23356 | 8.4033 | 0.002775 | 0.00414 | 0.003976 | 5.2e-05 | 0.001148 | 0.28191 | 1.3459 | 2.5476 | 98.43 | 20.03 | 47.788 | 71.6319 | 88.0321 | 68.2116 | 0.011384 | 1.0342 | 0.024021 | 23.2265 | -0.043744 | 0.13894 | -0.000251 | -0.018128 | -0.018128 | -0.30889
2019-04-09 15:38 | 1554838675.7109 | 717 | 0.20956 | 0.000441 | 1e-06 | 98.5931 | 99.2033 | 121.6104 | 20.6912 | 0.002132 | 0.001961 | -154.5733 | 0.2336 | 8.4016 | 0.00284 | 0.004408 | 0.004218 | 4.7e-05 | 0.001152 | 0.26936 | 1.3661 | 2.5477 | 98.43 | 20.03 | 47.791 | 72.5692 | 89.5936 | 68.4452 | 0.011385 | 1.0335 | 0.024046 | 23.2657 | -0.055185 | 0.13897 | -9.5e-05 | -0.00692 | -0.00692 | -0.80933
2019-04-09 15:38 | 1554838676.7105 | 718 | 0.20955 | 0.000445 | 1e-06 | 98.6155 | 98.9655 | 121.7755 | 20.691 | 0.002097 | 0.001972 | -154.5802 | 0.23333 | 8.4018 | 0.002802 | 0.004336 | 0.004153 | 5.1e-05 | 0.001145 | 0.28926 | 1.357 | 2.5476 | 98.43 | 20.03 | 47.794 | 71.7348 | 88.6734 | 68.714 | 0.011386 | 1.0327 | 0.024052 | 23.2906 | -0.031691 | 0.13881 | -3.6e-05 | -0.002659 | -0.002659 | -2.1035
2019-04-09 15:38 | 1554838677.711 | 719 | 0.20953 | 0.000445 | 1e-06 | 98.5514 | 99.0529 | 121.8452 | 20.6919 | 0.002074 | 0.002006 | -154.5827 | 0.2333 | 8.4018 | 0.002844 | 0.004246 | 0.004121 | 5e-05 | 0.001139 | 0.26504 | 1.3336 | 2.5476 | 98.43 | 20.03 | 47.797 | 71.168 | 89.6834 | 69.536 | 0.011386 | 1.0302 | 0.02403 | 23.3253 | -0.045397 | 0.13879 | 5.1e-05 | 0.003595 | 0.003595 | 1.5559
2019-04-09 15:38 | 1554838678.7117 | 720 | 0.20951 | 0.000448 | 1e-06 | 98.5243 | 99.0436 | 121.8992 | 20.6923 | 0.002089 | 0.001984 | -154.6041 | 0.23334 | 8.4033 | 0.00279 | 0.004188 | 0.004078 | 4.9e-05 | 0.001137 | 0.27113 | 1.3409 | 2.5476 | 98.43 | 20.03 | 47.8 | 71.5286 | 88.3761 | 69.0141 | 0.011387 | 1.0318 | 0.024054 | 23.3127 | -0.047018 | 0.13881 | 0.000182 | 0.013052 | 0.013052 | 0.4286
2019-04-09 15:38 | 1554838679.7116 | 721 | 0.20954 | 0.000448 | 1e-06 | 98.4606 | 98.9063 | 121.8515 | 20.6914 | 0.002052 | 0.001952 | -154.6161 | 0.23248 | 8.4024 | 0.0028 | 0.004589 | 0.004469 | 5.2e-05 | 0.001134 | 0.29142 | 1.3894 | 2.5477 | 98.43 | 20.03 | 47.804 | 70.6492 | 88.6275 | 68.2331 | 0.011388 | 1.0342 | 0.024077 | 23.2815 | -0.040396 | 0.13831 | 8e-06 | 0.000536 | 0.000536 | 10.3998
2019-04-09 15:38 | 1554838680.7116 | 722 | 0.20959 | 0.000446 | 1e-06 | 98.5047 | 98.8894 | 121.9022 | 20.6912 | 0.00209 | 0.00193 | -154.6099 | 0.23278 | 8.403 | 0.002736 | 0.00411 | 0.003888 | 5.1e-05 | 0.001131 | 0.27936 | 1.3299 | 2.5477 | 98.43 | 20.03 | 47.806 | 71.5487 | 87.0859 | 67.7018 | 0.011388 | 1.0358 | 0.024101 | 23.2682 | -0.034868 | 0.13848 | -0.000292 | -0.021168 | -0.021168 | -0.26365
2019-04-09 15:38 | 1554838681.7117 | 723 | 0.20957 | 0.000446 | 1e-06 | 98.5062 | 99.0082 | 121.8341 | 20.692 | 0.002128 | 0.001971 | -154.6232 | 0.23353 | 8.4034 | 0.002754 | 0.003855 | 0.00378 | 5.1e-05 | 0.001137 | 0.27599 | 1.3193 | 2.5476 | 98.43 | 20.03 | 47.808 | 72.4669 | 87.5091 | 68.6931 | 0.011389 | 1.0328 | 0.024059 | 23.2955 | -0.045466 | 0.13892 | -0.000157 | -0.011343 | -0.011343 | -0.49358
2019-04-09 15:38 | 1554838682.7122 | 724 | 0.20957 | 0.000446 | 1e-06 | 98.3805 | 98.7635 | 121.8445 | 20.6913 | 0.002085 | 0.001948 | -154.6017 | 0.23305 | 8.398 | 0.002699 | 0.004224 | 0.004083 | 5.2e-05 | 0.001144 | 0.28296 | 1.3375 | 2.5475 | 98.42 | 20.03 | 47.811 | 71.4412 | 86.1874 | 68.1513 | 0.011391 | 1.0344 | 0.024079 | 23.278 | -0.03475 | 0.13864 | -0.000165 | -0.011958 | -0.011958 | -0.46722
2019-04-09 15:38 | 1554838683.7118 | 725 | 0.20956 | 0.000444 | 1e-06 | 98.4651 | 98.7704 | 121.8935 | 20.6912 | 0.002091 | 0.002013 | -154.5782 | 0.23335 | 8.401 | 0.002741 | 0.003621 | 0.003464 | 5.3e-05 | 0.001149 | 0.26128 | 1.3428 | 2.5475 | 98.42 | 20.03 | 47.812 | 71.5797 | 87.1959 | 69.7179 | 0.011391 | 1.0297 | 0.024029 | 23.3361 | -0.027688 | 0.13882 | -7.6e-05 | -0.00556 | -0.00556 | -1.0063
2019-04-09 15:38 | 1554838684.7108 | 726 | 0.20956 | 0.000451 | 1e-06 | 98.3757 | 98.8123 | 122.0158 | 20.6918 | 0.002022 | 0.001924 | -154.5977 | 0.23241 | 8.4001 | 0.00271 | 0.003866 | 0.003736 | 4.4e-05 | 0.00115 | 0.23258 | 1.3626 | 2.5477 | 98.42 | 20.04 | 47.813 | 69.9236 | 86.4494 | 67.5784 | 0.011398 | 1.0361 | 0.024116 | 23.2748 | -0.039609 | 0.13826 | -0.000106 | -0.007733 | -0.007733 | -0.72055
2019-04-09 15:38 | 1554838685.7109 | 727 | 0.20953 | 0.000446 | 1e-06 | 98.3168 | 98.9097 | 122.148 | 20.692 | 0.002073 | 0.001968 | -154.5873 | 0.23301 | 8.4033 | 0.002684 | 0.003816 | 0.003646 | 4.9e-05 | 0.00115 | 0.27367 | 1.3701 | 2.5477 | 98.42 | 20.03 | 47.812 | 71.1479 | 85.836 | 68.6299 | 0.011391 | 1.033 | 0.024092 | 23.3233 | -0.053783 | 0.13862 | 5.4e-05 | 0.00386 | 0.00386 | 1.4473
2019-04-09 15:38 | 1554838686.7119 | 728 | 0.20955 | 0.000446 | 1e-06 | 98.2933 | 98.6702 | 122.2994 | 20.6917 | 0.00209 | 0.001931 | -154.5989 | 0.23247 | 8.4028 | 0.002715 | 0.003966 | 0.003817 | 5.2e-05 | 0.001151 | 0.30021 | 1.327 | 2.5475 | 98.42 | 20.03 | 47.811 | 71.5476 | 86.5734 | 67.7305 | 0.011391 | 1.0357 | 0.024139 | 23.3071 | -0.034236 | 0.13829 | -1.5e-05 | -0.001115 | -0.001115 | -5.0003
2019-04-09 15:38 | 1554838687.71 | 729 | 0.20955 | 0.000446 | 1e-06 | 98.3132 | 98.811 | 122.2299 | 20.6922 | 0.00208 | 0.001981 | -154.6048 | 0.23231 | 8.4006 | 0.002638 | 0.004248 | 0.004214 | 4.9e-05 | 0.001154 | 0.29551 | 1.3041 | 2.5477 | 98.42 | 20.03 | 47.807 | 71.3079 | 84.7396 | 68.9375 | 0.01139 | 1.032 | 0.024089 | 23.3417 | -0.045172 | 0.1382 | -4.6e-05 | -0.003343 | -0.003343 | -1.6663
2019-04-09 15:38 | 1554838688.71 | 730 | 0.20954 | 0.000449 | 1e-06 | 98.2921 | 98.5171 | 122.0494 | 20.6917 | 0.002087 | 0.001985 | -154.5989 | 0.23263 | 8.4019 | 0.002676 | 0.003417 | 0.003274 | 5.3e-05 | 0.001154 | 0.29103 | 1.3068 | 2.5476 | 98.42 | 20.04 | 47.802 | 71.4766 | 85.642 | 69.0305 | 0.011396 | 1.0318 | 0.024068 | 23.3276 | -0.020451 | 0.13839 | 3e-06 | 0.000117 | 0.000117 | 47.4949
2019-04-09 15:38 | 1554838689.71 | 731 | 0.20952 | 0.00045 | 1e-06 | 98.2787 | 98.7514 | 122.013 | 20.6914 | 0.002101 | 0.001944 | -154.5924 | 0.23241 | 8.4018 | 0.002689 | 0.004189 | 0.004038 | 4.9e-05 | 0.001151 | 0.30071 | 1.3609 | 2.5477 | 98.42 | 20.04 | 47.791 | 71.8317 | 85.9634 | 68.0538 | 0.011393 | 1.0347 | 0.024099 | 23.2908 | -0.042924 | 0.13826 | 0.000141 | 0.010154 | 0.010154 | 0.54875
2019-04-09 15:38 | 1554838690.7111 | 732 | 0.20957 | 0.000447 | 1e-06 | 98.2501 | 98.6231 | 121.6445 | 20.6918 | 0.002011 | 0.001996 | -154.6021 | 0.23251 | 8.402 | 0.00276 | 0.004065 | 0.003905 | 5.2e-05 | 0.001159 | 0.26952 | 1.3601 | 2.5477 | 98.42 | 20.04 | 47.786 | 69.671 | 87.6557 | 69.3059 | 0.011392 | 1.0309 | 0.024019 | 23.2982 | -0.033891 | 0.13832 | -0.000165 | -0.011949 | -0.011949 | -0.46651
2019-04-09 15:38 | 1554838691.7101 | 733 | 0.20956 | 0.000449 | 1e-06 | 98.2185 | 98.7084 | 121.9266 | 20.6911 | 0.002042 | 0.001993 | -154.599 | 0.23232 | 8.4034 | 0.002753 | 0.004346 | 0.004173 | 5.1e-05 | 0.001163 | 0.28711 | 1.383 | 2.5477 | 98.43 | 20.04 | 47.779 | 70.397 | 87.4838 | 69.2234 | 0.011389 | 1.0312 | 0.024049 | 23.3224 | -0.044503 | 0.13821 | -0.000106 | -0.007719 | -0.007719 | -0.72158
2019-04-09 15:38 | 1554838692.71 | 734 | 0.20958 | 0.000447 | 1e-06 | 98.0919 | 98.4551 | 122.3411 | 20.6915 | 0.001967 | 0.002009 | -154.633 | 0.2328 | 8.4006 | 0.002753 | 0.004313 | 0.004118 | 5.1e-05 | 0.001158 | 0.27975 | 1.3956 | 2.5478 | 98.43 | 20.04 | 47.766 | 68.5897 | 87.4818 | 69.612 | 0.011386 | 1.03 | 0.024077 | 23.3753 | -0.033054 | 0.1385 | -0.000203 | -0.014715 | -0.014715 | -0.37929
2019-04-09 15:38 | 1554838693.7101 | 735 | 0.20956 | 0.000447 | 1e-06 | 98.1128 | 98.5811 | 122.3099 | 20.6914 | 0.002052 | 0.001943 | -154.6321 | 0.23199 | 8.4039 | 0.002688 | 0.00447 | 0.004316 | 5.1e-05 | 0.001161 | 0.28296 | 1.345 | 2.5477 | 98.43 | 20.04 | 47.757 | 70.65 | 85.9353 | 68.0202 | 0.011384 | 1.0348 | 0.02413 | 23.318 | -0.042589 | 0.13801 | -7.9e-05 | -0.005807 | -0.005807 | -0.95784
2019-04-09 15:38 | 1554838694.7112 | 736 | 0.20953 | 0.000451 | 1e-06 | 98.1015 | 98.3123 | 122.202 | 20.6909 | 0.00202 | 0.002006 | -154.6321 | 0.2331 | 8.4011 | 0.002726 | 0.004143 | 0.003963 | 5.2e-05 | 0.001164 | 0.30165 | 1.3349 | 2.5476 | 98.43 | 20.04 | 47.748 | 69.8657 | 86.8481 | 69.5406 | 0.011382 | 1.0302 | 0.024065 | 23.3596 | -0.0192 | 0.13867 | 4.7e-05 | 0.003336 | 0.003336 | 1.6754
2019-04-09 15:38 | 1554838695.7102 | 737 | 0.20955 | 0.000444 | 2e-06 | 98.1071 | 98.5016 | 121.81 | 20.6916 | 0.002064 | 0.001969 | -154.631 | 0.23316 | 8.4 | 0.002762 | 0.003793 | 0.003664 | 5.3e-05 | 0.001166 | 0.29103 | 1.3263 | 2.5477 | 98.43 | 20.04 | 47.732 | 70.9363 | 87.7209 | 68.6561 | 0.011378 | 1.0329 | 0.024058 | 23.2919 | -0.035889 | 0.13871 | -5.4e-05 | -0.003981 | -0.003981 | -1.4043
2019-04-09 15:38 | 1554838696.7113 | 738 | 0.20957 | 0.000447 | 1e-06 | 98.0777 | 98.5288 | 121.8708 | 20.691 | 0.002051 | 0.002015 | -154.6084 | 0.23264 | 8.4028 | 0.002742 | 0.004515 | 0.004379 | 5.1e-05 | 0.001164 | 0.27312 | 1.3551 | 2.5476 | 98.43 | 20.04 | 47.723 | 70.6159 | 87.2323 | 69.7606 | 0.011376 | 1.0296 | 0.024025 | 23.3354 | -0.041046 | 0.1384 | -0.000171 | -0.012366 | -0.012366 | -0.45102
2019-04-09 15:38 | 1554838697.7102 | 739 | 0.20957 | 0.000451 | 1e-06 | 98.0785 | 98.5253 | 121.8675 | 20.6906 | 0.002078 | 0.001965 | -154.5994 | 0.23321 | 8.4047 | 0.002775 | 0.004627 | 0.004507 | 5.2e-05 | 0.001164 | 0.28584 | 1.3977 | 2.5476 | 98.43 | 20.04 | 47.709 | 71.2737 | 88.0206 | 68.5624 | 0.011372 | 1.0332 | 0.024067 | 23.2943 | -0.040652 | 0.13874 | -0.000157 | -0.011432 | -0.011432 | -0.48909
2019-04-09 15:38 | 1554838698.7113 | 740 | 0.20959 | 0.000445 | 1e-06 | 98.0109 | 98.315 | 121.9705 | 20.6909 | 0.002063 | 0.001909 | -154.5979 | 0.23354 | 8.4034 | 0.002749 | 0.003824 | 0.003722 | 5e-05 | 0.001165 | 0.27007 | 1.3556 | 2.5475 | 98.43 | 20.05 | 47.701 | 70.9075 | 87.3931 | 67.1959 | 0.011378 | 1.0373 | 0.024125 | 23.2575 | -0.027708 | 0.13893 | -0.000294 | -0.021307 | -0.021307 | -0.26277
2019-04-09 15:38 | 1554838699.7701 | 741 | 0.20957 | 0.000449 | 1e-06 | 97.9508 | 98.3546 | 121.682 | 20.6914 | 0.002118 | 0.001967 | -154.5967 | 0.23323 | 8.4008 | 0.00275 | 0.003857 | 0.003603 | 5.1e-05 | 0.001163 | 0.29767 | 1.3189 | 2.5474 | 98.43 | 20.05 | 47.694 | 72.2294 | 87.4187 | 68.6067 | 0.011376 | 1.033 | 0.024047 | 23.278 | -0.0368 | 0.13875 | -0.000179 | -0.012971 | -0.012971 | -0.43108
2019-04-09 15:38 | 1554838700.7231 | 742 | 0.2096 | 0.000448 | 1e-06 | 97.9439 | 98.2737 | 121.7391 | 20.6907 | 0.002096 | 0.001971 | -154.5874 | 0.23346 | 8.404 | 0.002759 | 0.003642 | 0.003491 | 5.1e-05 | 0.001163 | 0.28584 | 1.3244 | 2.5475 | 98.43 | 20.04 | 47.684 | 71.7057 | 87.6414 | 68.7005 | 0.011366 | 1.0327 | 0.024049 | 23.2867 | -0.030067 | 0.13889 | -0.00033 | -0.023877 | -0.023877 | -0.23442
2019-04-09 15:38 | 1554838701.7231 | 743 | 0.20957 | 0.000443 | 1e-06 | 97.9865 | 98.3291 | 121.6126 | 20.6905 | 0.002082 | 0.001963 | -154.5778 | 0.23287 | 8.4013 | 0.002786 | 0.004297 | 0.00415 | 5.3e-05 | 0.001159 | 0.27632 | 1.3669 | 2.5477 | 98.43 | 20.04 | 47.679 | 71.3601 | 88.2869 | 68.5015 | 0.011365 | 1.0333 | 0.024044 | 23.2678 | -0.031221 | 0.13853 | -0.000172 | -0.012476 | -0.012476 | -0.44749
2019-04-09 15:38 | 1554838702.7208 | 744 | 0.20958 | 0.00045 | 1e-06 | 97.8978 | 98.3219 | 121.6773 | 20.6906 | 0.002015 | 0.001983 | -154.5545 | 0.2334 | 8.3994 | 0.002727 | 0.003865 | 0.003735 | 5.1e-05 | 0.001157 | 0.274 | 1.3359 | 2.5476 | 98.43 | 20.04 | 47.674 | 69.7602 | 86.8666 | 68.9749 | 0.011364 | 1.0319 | 0.024034 | 23.2901 | -0.038668 | 0.13885 | -0.000215 | -0.01556 | -0.01556 | -0.35961
2019-04-09 15:38 | 1554838703.7211 | 745 | 0.20959 | 0.000448 | 1e-06 | 97.9645 | 98.3443 | 121.8632 | 20.6911 | 0.002089 | 0.001912 | -154.5721 | 0.23376 | 8.4012 | 0.002789 | 0.003915 | 0.003742 | 4.8e-05 | 0.001154 | 0.29734 | 1.3698 | 2.5476 | 98.43 | 20.05 | 47.67 | 71.5369 | 88.3711 | 67.2836 | 0.01137 | 1.037 | 0.024112 | 23.2502 | -0.03461 | 0.13906 | -0.000287 | -0.020791 | -0.020791 | -0.26955
2019-04-09 15:38 | 1554838704.7218 | 746 | 0.20957 | 0.000447 | 1e-06 | 97.9484 | 98.2569 | 121.752 | 20.6908 | 0.002095 | 0.001985 | -154.5521 | 0.23348 | 8.4018 | 0.002779 | 0.004372 | 0.00421 | 5.2e-05 | 0.001148 | 0.28102 | 1.3243 | 2.5476 | 98.42 | 20.05 | 47.665 | 71.6769 | 88.1116 | 69.0328 | 0.01137 | 1.0317 | 0.024039 | 23.2992 | -0.028124 | 0.1389 | -0.000162 | -0.011757 | -0.011757 | -0.4761
2019-04-09 15:38 | 1554838705.7218 | 747 | 0.20958 | 0.000444 | 2e-06 | 97.9413 | 98.2225 | 121.6558 | 20.6917 | 0.001991 | 0.001936 | -154.5597 | 0.23356 | 8.4011 | 0.002841 | 0.003949 | 0.003773 | 5.3e-05 | 0.001148 | 0.27455 | 1.3869 | 2.5477 | 98.42 | 20.05 | 47.663 | 69.1765 | 89.624 | 67.8449 | 0.01137 | 1.0353 | 0.024071 | 23.2496 | -0.025644 | 0.13894 | -0.00021 | -0.015222 | -0.015222 | -0.36784
2019-04-09 15:38 | 1554838706.7218 | 748 | 0.20958 | 0.000447 | 1e-06 | 97.8945 | 98.1216 | 121.3339 | 20.691 | 0.002039 | 0.001976 | -154.5865 | 0.23317 | 8.4006 | 0.002724 | 0.00437 | 0.004239 | 4.9e-05 | 0.001143 | 0.28943 | 1.3611 | 2.5475 | 98.42 | 20.05 | 47.66 | 70.3407 | 86.8037 | 68.8241 | 0.011369 | 1.0324 | 0.024005 | 23.2521 | -0.020727 | 0.13871 | -0.000228 | -0.01647 | -0.01647 | -0.33941
2019-04-09 15:38 | 1554838707.7215 | 749 | 0.20957 | 0.000445 | 1e-06 | 97.8763 | 98.2813 | 121.2471 | 20.6912 | 0.002102 | 0.00196 | -154.5858 | 0.23295 | 8.4017 | 0.002712 | 0.004047 | 0.003897 | 5e-05 | 0.001142 | 0.26969 | 1.3566 | 2.5476 | 98.42 | 20.05 | 47.659 | 71.8465 | 86.5128 | 68.4328 | 0.011369 | 1.0336 | 0.02401 | 23.2305 | -0.036932 | 0.13858 | -0.000151 | -0.010928 | -0.010928 | -0.51106
2019-04-09 15:38 | 1554838708.7208 | 750 | 0.20957 | 0.000446 | 1e-06 | 97.8921 | 98.2479 | 121.5004 | 20.6908 | 0.002121 | 0.001947 | -154.5927 | 0.23355 | 8.4007 | 0.002783 | 0.00375 | 0.003634 | 5.4e-05 | 0.001148 | 0.28008 | 1.3791 | 2.5476 | 98.42 | 20.05 | 47.658 | 72.3139 | 88.2201 | 68.1273 | 0.011369 | 1.0345 | 0.024046 | 23.2443 | -0.032451 | 0.13894 | -0.000153 | -0.011096 | -0.011096 | -0.50464
2019-04-09 15:38 | 1554838709.7209 | 751 | 0.20957 | 0.000446 | 1e-06 | 97.9928 | 98.3742 | 121.6783 | 20.6915 | 0.002049 | 0.002005 | -154.5939 | 0.23275 | 8.4016 | 0.00276 | 0.003821 | 0.00364 | 4.8e-05 | 0.001147 | 0.27201 | 1.3291 | 2.5475 | 98.42 | 20.05 | 47.656 | 70.5781 | 87.663 | 69.5046 | 0.011368 | 1.0303 | 0.024015 | 23.3082 | -0.03475 | 0.13846 | -0.000145 | -0.010515 | -0.010515 | -0.5307
2019-04-09 15:38 | 1554838710.7219 | 752 | 0.20965 | 0.000447 | 2e-06 | 97.8713 | 98.0146 | 121.6994 | 20.6927 | 0.002052 | 0.00199 | -154.5929 | 0.23366 | 8.4032 | 0.0028 | 0.004214 | 0.004104 | 5.2e-05 | 0.001149 | 0.26322 | 1.412 | 2.5475 | 98.42 | 20.05 | 47.655 | 70.6356 | 88.6242 | 69.1445 | 0.011368 | 1.0314 | 0.02403 | 23.298 | -0.013087 | 0.139 | -0.000651 | -0.046957 | -0.046957 | -0.11929
2019-04-09 15:39 | 1554838711.722 | 753 | 0.20952 | 0.000449 | 1e-06 | 97.8963 | 98.3157 | 121.4454 | 20.691 | 0.002028 | 0.001957 | -154.6007 | 0.23317 | 8.4004 | 0.002775 | 0.004213 | 0.004093 | 5.1e-05 | 0.001148 | 0.28512 | 1.3606 | 2.5477 | 98.42 | 20.05 | 47.655 | 70.0701 | 88.0282 | 68.3584 | 0.011368 | 1.0338 | 0.024032 | 23.2469 | -0.03824 | 0.13872 | 0.000155 | 0.011127 | 0.011127 | 0.50238
2019-04-09 15:39 | 1554838712.721 | 754 | 0.20962 | 0.000444 | 1e-06 | 97.8632 | 98.1712 | 121.4851 | 20.6911 | 0.002045 | 0.002001 | -154.5976 | 0.23412 | 8.3999 | 0.002751 | 0.004311 | 0.004156 | 5.2e-05 | 0.001144 | 0.28512 | 1.3507 | 2.5476 | 98.42 | 20.05 | 47.655 | 70.4867 | 87.4449 | 69.4128 | 0.011368 | 1.0306 | 0.023999 | 23.2866 | -0.028101 | 0.13928 | -0.000432 | -0.031162 | -0.031162 | -0.18013
2019-04-09 15:39 | 1554838713.7221 | 755 | 0.20956 | 0.000448 | 1e-06 | 97.8659 | 98.1922 | 121.7689 | 20.6922 | 0.002058 | 0.001969 | -154.5877 | 0.23298 | 8.4015 | 0.002718 | 0.00442 | 0.00427 | 5e-05 | 0.001137 | 0.26521 | 1.3636 | 2.5476 | 98.42 | 20.05 | 47.656 | 70.781 | 86.6636 | 68.6544 | 0.011368 | 1.0329 | 0.024054 | 23.288 | -0.029768 | 0.1386 | -8.6e-05 | -0.006304 | -0.006304 | -0.88604
2019-04-09 15:39 | 1554838714.7225 | 756 | 0.20955 | 0.000453 | 1e-06 | 97.8421 | 98.1798 | 122.1797 | 20.6911 | 0.00205 | 0.001955 | -154.6185 | 0.23375 | 8.4037 | 0.002719 | 0.004423 | 0.004254 | 4.9e-05 | 0.001134 | 0.27527 | 1.3657 | 2.5476 | 98.42 | 20.05 | 47.651 | 70.6027 | 86.6663 | 68.314 | 0.011367 | 1.0339 | 0.024106 | 23.3156 | -0.030817 | 0.13906 | -7.1e-05 | -0.005188 | -0.005188 | -1.0801
2019-04-09 15:39 | 1554838715.7221 | 757 | 0.20957 | 0.000447 | 1e-06 | 97.8338 | 98.283 | 122.2447 | 20.6906 | 0.002043 | 0.001967 | -154.6139 | 0.23199 | 8.4005 | 0.002739 | 0.004147 | 0.004033 | 5.1e-05 | 0.001138 | 0.27168 | 1.3442 | 2.5476 | 98.42 | 20.05 | 47.649 | 70.4379 | 87.1657 | 68.5903 | 0.011366 | 1.0331 | 0.024103 | 23.3312 | -0.040981 | 0.13801 | -0.000171 | -0.012446 | -0.012446 | -0.44687
2019-04-09 15:39 | 1554838716.7225 | 758 | 0.20952 | 0.000442 | 1e-06 | 97.7724 | 98.0352 | 122.1304 | 20.6917 | 0.002037 | 0.001903 | -154.598 | 0.23224 | 8.4007 | 0.002769 | 0.004846 | 0.004686 | 5e-05 | 0.001147 | 0.28672 | 1.3653 | 2.5477 | 98.42 | 20.05 | 47.647 | 70.2737 | 87.8749 | 67.0726 | 0.011366 | 1.0377 | 0.024145 | 23.2685 | -0.024012 | 0.13816 | 0.000162 | 0.011675 | 0.011675 | 0.4769
2019-04-09 15:39 | 1554838717.7223 | 759 | 0.2096 | 0.000446 | 1e-06 | 97.7968 | 98.177 | 121.8611 | 20.6918 | 0.002062 | 0.001973 | -154.6135 | 0.23232 | 8.4015 | 0.002857 | 0.003354 | 0.003186 | 4.7e-05 | 0.001155 | 0.274 | 1.3905 | 2.5476 | 98.42 | 20.05 | 47.64 | 70.8764 | 89.9865 | 68.753 | 0.011364 | 1.0326 | 0.024059 | 23.3001 | -0.034705 | 0.13821 | -0.000334 | -0.024161 | -0.024161 | -0.23053
2019-04-09 15:39 | 1554838718.7216 | 760 | 0.20958 | 0.000445 | 1e-06 | 97.8358 | 98.2352 | 121.609 | 20.6911 | 0.002072 | 0.001897 | -154.6036 | 0.23294 | 8.4025 | 0.002736 | 0.004746 | 0.004604 | 4.3e-05 | 0.001158 | 0.27815 | 1.3563 | 2.5476 | 98.42 | 20.06 | 47.638 | 71.136 | 87.074 | 66.9284 | 0.011371 | 1.0381 | 0.024099 | 23.2138 | -0.036435 | 0.13858 | -0.000189 | -0.013731 | -0.013731 | -0.40671
2019-04-09 15:39 | 1554838719.7202 | 761 | 0.2096 | 0.000449 | 2e-06 | 97.8263 | 98.1041 | 121.7165 | 20.6907 | 0.002035 | 0.001941 | -154.6231 | 0.23275 | 8.4039 | 0.002854 | 0.004177 | 0.004001 | 4.3e-05 | 0.001163 | 0.27577 | 1.3251 | 2.5476 | 98.42 | 20.05 | 47.635 | 70.2282 | 89.923 | 67.9874 | 0.011363 | 1.0349 | 0.024072 | 23.2602 | -0.02536 | 0.13846 | -0.000312 | -0.02263 | -0.02263 | -0.24657
2019-04-09 15:39 | 1554838720.7213 | 762 | 0.20952 | 0.000442 | 1e-06 | 97.9046 | 98.175 | 122.1287 | 20.6912 | 0.002055 | 0.001955 | -154.6213 | 0.2332 | 8.4025 | 0.002774 | 0.003971 | 0.003802 | 4.5e-05 | 0.00117 | 0.2808 | 1.3335 | 2.5476 | 98.42 | 20.05 | 47.634 | 70.7175 | 88.0009 | 68.3132 | 0.011363 | 1.0339 | 0.024101 | 23.3107 | -0.024663 | 0.13873 | 0.000148 | 0.010634 | 0.010634 | 0.52577
2019-04-09 15:39 | 1554838721.7213 | 763 | 0.20956 | 0.00045 | 1e-06 | 97.8537 | 98.0934 | 122.301 | 20.6902 | 0.00205 | 0.001962 | -154.6174 | 0.23269 | 8.4032 | 0.002789 | 0.003606 | 0.003454 | 4.4e-05 | 0.001171 | 0.26272 | 1.33 | 2.5476 | 98.42 | 20.06 | 47.634 | 70.6044 | 88.3647 | 68.4731 | 0.01137 | 1.0334 | 0.024113 | 23.3326 | -0.021882 | 0.13843 | -8.4e-05 | -0.006082 | -0.006082 | -0.91721
2019-04-09 15:39 | 1554838722.7207 | 764 | 0.20952 | 0.00045 | 1e-06 | 97.8796 | 98.2486 | 121.9527 | 20.6906 | 0.002076 | 0.001975 | -154.6336 | 0.23429 | 8.4037 | 0.002786 | 0.004171 | 0.004011 | 5e-05 | 0.001172 | 0.30829 | 1.3184 | 2.5476 | 98.42 | 20.06 | 47.634 | 71.2213 | 88.2753 | 68.7956 | 0.01137 | 1.0325 | 0.024067 | 23.3103 | -0.033656 | 0.13938 | 0.000131 | 0.009378 | 0.009378 | 0.59892
2019-04-09 15:39 | 1554838723.7207 | 765 | 0.20955 | 0.000448 | 1e-06 | 97.9585 | 98.0662 | 121.5443 | 20.6912 | 0.002065 | 0.001951 | -154.6415 | 0.23247 | 8.4023 | 0.002759 | 0.004269 | 0.004106 | 4.6e-05 | 0.001173 | 0.29927 | 1.327 | 2.5476 | 98.42 | 20.06 | 47.632 | 70.9475 | 87.645 | 68.2274 | 0.011369 | 1.0342 | 0.024047 | 23.2519 | -0.009826 | 0.1383 | -6.6e-05 | -0.004798 | -0.004798 | -1.1616
2019-04-09 15:39 | 1554838724.7204 | 766 | 0.20953 | 0.000446 | 2e-06 | 97.893 | 98.3443 | 121.3325 | 20.6903 | 0.002027 | 0.001931 | -154.6215 | 0.23366 | 8.4012 | 0.002801 | 0.003961 | 0.003739 | 4.4e-05 | 0.00118 | 0.30469 | 1.3652 | 2.5476 | 98.42 | 20.06 | 47.632 | 70.0329 | 88.6459 | 67.736 | 0.011369 | 1.0357 | 0.024043 | 23.215 | -0.041138 | 0.139 | 8.4e-05 | 0.005982 | 0.005982 | 0.93644
2019-04-09 15:39 | 1554838725.7218 | 767 | 0.20953 | 0.000448 | 2e-06 | 97.8165 | 98.2386 | 121.6546 | 20.6914 | 0.002004 | 0.001939 | -154.6266 | 0.2333 | 8.4 | 0.0028 | 0.00423 | 0.004057 | 4.4e-05 | 0.001177 | 0.29391 | 1.3654 | 2.5476 | 98.42 | 20.06 | 47.634 | 69.4983 | 88.6181 | 67.9321 | 0.01137 | 1.0351 | 0.024068 | 23.2524 | -0.038519 | 0.13879 | 7e-05 | 0.004952 | 0.004952 | 1.1294
2019-04-09 15:39 | 1554838726.7215 | 768 | 0.20956 | 0.000446 | 1e-06 | 97.9115 | 98.0807 | 121.836 | 20.6918 | 0.002048 | 0.001942 | -154.6381 | 0.2337 | 8.4027 | 0.002777 | 0.004104 | 0.003967 | 4.5e-05 | 0.00118 | 0.28622 | 1.3658 | 2.5476 | 98.42 | 20.05 | 47.636 | 70.5402 | 88.0836 | 68.0066 | 0.011363 | 1.0348 | 0.024083 | 23.2723 | -0.015442 | 0.13903 | -9e-05 | -0.006535 | -0.006535 | -0.85735
2019-04-09 15:39 | 1554838727.7205 | 769 | 0.20953 | 0.000447 | 1e-06 | 97.9499 | 98.3068 | 122.1259 | 20.6913 | 0.002051 | 0.001969 | -154.6276 | 0.23398 | 8.4024 | 0.002831 | 0.004073 | 0.003942 | 4.3e-05 | 0.001185 | 0.2844 | 1.359 | 2.5477 | 98.42 | 20.06 | 47.638 | 70.6217 | 89.3743 | 68.6458 | 0.011371 | 1.0329 | 0.024089 | 23.3218 | -0.03253 | 0.1392 | 8.8e-05 | 0.006329 | 0.006329 | 0.88637
2019-04-09 15:39 | 1554838728.7209 | 770 | 0.20954 | 0.000448 | 1e-06 | 98.0047 | 98.3071 | 122.1412 | 20.6905 | 0.00204 | 0.001952 | -154.6421 | 0.23311 | 8.3999 | 0.002859 | 0.003994 | 0.003797 | 4.3e-05 | 0.001187 | 0.29407 | 1.3628 | 2.5477 | 98.42 | 20.05 | 47.64 | 70.3614 | 90.0563 | 68.2486 | 0.011364 | 1.0341 | 0.024105 | 23.3097 | -0.027559 | 0.13868 | 2.2e-05 | 0.001527 | 0.001527 | 3.6596
2019-04-09 15:39 | 1554838729.722 | 771 | 0.20953 | 0.000449 | 1e-06 | 97.9824 | 98.5164 | 122.186 | 20.6914 | 0.001994 | 0.001977 | -154.6164 | 0.23237 | 8.4016 | 0.002911 | 0.004132 | 0.004021 | 4.5e-05 | 0.001189 | 0.28871 | 1.3193 | 2.5474 | 98.42 | 20.05 | 47.641 | 69.2463 | 91.2937 | 68.8454 | 0.011364 | 1.0323 | 0.024088 | 23.3343 | -0.048616 | 0.13824 | 7.6e-05 | 0.005431 | 0.005431 | 1.0258
2019-04-09 15:39 | 1554838730.7216 | 772 | 0.20957 | 0.00045 | 1e-06 | 98.0235 | 98.4214 | 122.1055 | 20.6913 | 0.001992 | 0.001936 | -154.6067 | 0.23317 | 8.404 | 0.002908 | 0.004595 | 0.004454 | 4.1e-05 | 0.00118 | 0.28982 | 1.34 | 2.5476 | 98.42 | 20.05 | 47.643 | 69.1938 | 91.2344 | 67.8525 | 0.011365 | 1.0353 | 0.024115 | 23.2928 | -0.036237 | 0.13871 | -0.000171 | -0.012428 | -0.012428 | -0.44979
2019-04-09 15:39 | 1554838731.7211 | 773 | 0.20956 | 0.000447 | 1e-06 | 98.0943 | 98.3577 | 122.2822 | 20.6912 | 0.001932 | 0.001951 | -154.6136 | 0.23303 | 8.4035 | 0.00282 | 0.004241 | 0.004102 | 4.3e-05 | 0.001166 | 0.29048 | 1.383 | 2.5477 | 98.42 | 20.05 | 47.642 | 67.7634 | 89.1149 | 68.2079 | 0.011365 | 1.0342 | 0.02412 | 23.3218 | -0.023984 | 0.13863 | -7.6e-05 | -0.005563 | -0.005563 | -1.0042
2019-04-09 15:39 | 1554838732.7207 | 774 | 0.20956 | 0.000448 | 1e-06 | 98.0351 | 98.5346 | 122.5247 | 20.6912 | 0.001992 | 0.001945 | -154.5888 | 0.23359 | 8.4012 | 0.002842 | 0.004342 | 0.004181 | 4.5e-05 | 0.001166 | 0.27295 | 1.3574 | 2.5476 | 98.42 | 20.05 | 47.641 | 69.1954 | 89.6273 | 68.0762 | 0.011364 | 1.0346 | 0.024149 | 23.3404 | -0.045461 | 0.13896 | -9.9e-05 | -0.007231 | -0.007231 | -0.77442
2019-04-09 15:39 | 1554838733.7217 | 775 | 0.20952 | 0.000452 | 1e-06 | 98.1196 | 98.6028 | 122.9331 | 20.6911 | 0.002051 | 0.00194 | -154.597 | 0.2335 | 8.4031 | 0.002876 | 0.003976 | 0.003693 | 4.6e-05 | 0.00118 | 0.27417 | 1.3587 | 2.5476 | 98.42 | 20.05 | 47.639 | 70.6316 | 90.4517 | 67.96 | 0.011364 | 1.035 | 0.024193 | 23.3752 | -0.043932 | 0.13891 | 0.000112 | 0.008012 | 0.008012 | 0.69874
2019-04-09 15:39 | 1554838734.7207 | 776 | 0.20957 | 0.000454 | 1e-06 | 98.0976 | 98.5993 | 122.9319 | 20.6913 | 0.002029 | 0.001938 | -154.6086 | 0.23275 | 8.4045 | 0.002902 | 0.004524 | 0.004427 | 4.4e-05 | 0.00118 | 0.27328 | 1.3603 | 2.5476 | 98.42 | 20.05 | 47.639 | 70.0822 | 91.085 | 67.9015 | 0.011364 | 1.0352 | 0.024195 | 23.3731 | -0.045629 | 0.13846 | -0.000154 | -0.011261 | -0.011261 | -0.49553
2019-04-09 15:39 | 1554838735.7207 | 777 | 0.20955 | 0.000451 | 1e-06 | 98.1417 | 98.4813 | 122.7167 | 20.6915 | 0.00204 | 0.001969 | -154.6056 | 0.2332 | 8.4018 | 0.00283 | 0.003925 | 0.00381 | 4.4e-05 | 0.001184 | 0.28816 | 1.3527 | 2.5476 | 98.42 | 20.05 | 47.637 | 70.3444 | 89.3488 | 68.6614 | 0.011364 | 1.0329 | 0.024147 | 23.3787 | -0.030896 | 0.13873 | -4.6e-05 | -0.003365 | -0.003365 | -1.6613
2019-04-09 15:39 | 1554838736.7208 | 778 | 0.20954 | 0.000449 | 1e-06 | 98.1825 | 98.5009 | 122.6439 | 20.6909 | 0.002031 | 0.001904 | -154.6133 | 0.23398 | 8.3999 | 0.002796 | 0.004742 | 0.004498 | 4.3e-05 | 0.001178 | 0.29264 | 1.3521 | 2.5474 | 98.42 | 20.05 | 47.636 | 70.1395 | 88.5272 | 67.0906 | 0.011363 | 1.0376 | 0.024196 | 23.318 | -0.02896 | 0.1392 | 1.3e-05 | 0.00082 | 0.00082 | 6.8427
2019-04-09 15:39 | 1554838737.7202 | 779 | 0.20953 | 0.000448 | 1e-06 | 98.2221 | 98.7587 | 122.3813 | 20.6917 | 0.002061 | 0.001899 | -154.6099 | 0.23356 | 8.4028 | 0.002845 | 0.004165 | 0.004058 | 4.1e-05 | 0.001176 | 0.27671 | 1.3483 | 2.5475 | 98.42 | 20.05 | 47.634 | 70.8528 | 89.7082 | 66.9714 | 0.011363 | 1.038 | 0.024174 | 23.2889 | -0.048729 | 0.13895 | 8.4e-05 | 0.006051 | 0.006051 | 0.92535
2019-04-09 15:39 | 1554838738.7204 | 780 | 0.20954 | 0.000449 | 1e-06 | 98.2766 | 98.7439 | 122.2465 | 20.6918 | 0.002007 | 0.001935 | -154.609 | 0.23363 | 8.4018 | 0.002858 | 0.003895 | 0.003824 | 4.3e-05 | 0.001173 | 0.30071 | 1.3359 | 2.5477 | 98.42 | 20.05 | 47.632 | 69.5703 | 90.0123 | 67.8436 | 0.011362 | 1.0353 | 0.02413 | 23.3059 | -0.042428 | 0.13899 | 3.8e-05 | 0.002706 | 0.002706 | 2.0698
2019-04-09 15:39 | 1554838739.7213 | 781 | 0.20947 | 0.000448 | 1e-06 | 98.3504 | 98.6606 | 122.0261 | 20.6909 | 0.002054 | 0.001948 | -154.622 | 0.23383 | 8.4026 | 0.002814 | 0.00393 | 0.003724 | 4.5e-05 | 0.001175 | 0.27936 | 1.369 | 2.5476 | 98.42 | 20.05 | 47.631 | 70.6891 | 88.9496 | 68.1374 | 0.011362 | 1.0345 | 0.024097 | 23.2949 | -0.028161 | 0.1391 | 0.00047 | 0.033882 | 0.033882 | 0.16545
2019-04-09 15:39 | 1554838740.72 | 782 | 0.20955 | 0.00045 | 1e-06 | 98.3626 | 98.8251 | 121.6884 | 20.6917 | 0.00198 | 0.001892 | -154.627 | 0.23342 | 8.4006 | 0.002796 | 0.003299 | 0.003202 | 4.3e-05 | 0.00118 | 0.27201 | 1.3829 | 2.5477 | 98.42 | 20.05 | 47.626 | 68.9072 | 88.5248 | 66.8028 | 0.011361 | 1.0385 | 0.024111 | 23.2171 | -0.041953 | 0.13886 | -2.8e-05 | -0.002082 | -0.002082 | -2.6882
2019-04-09 15:39 | 1554838741.7209 | 783 | 0.20948 | 0.000455 | 1e-06 | 98.2825 | 98.801 | 121.3682 | 20.6924 | 0.002064 | 0.001956 | -154.6414 | 0.23459 | 8.4003 | 0.002778 | 0.003453 | 0.003192 | 4.1e-05 | 0.001192 | 0.28335 | 1.3739 | 2.5477 | 98.42 | 20.05 | 47.626 | 70.9299 | 88.095 | 68.3494 | 0.011361 | 1.0338 | 0.024025 | 23.2392 | -0.047059 | 0.13956 | 0.00036 | 0.02585 | 0.02585 | 0.21757
2019-04-09 15:39 | 1554838742.7209 | 784 | 0.20953 | 0.00045 | 1e-06 | 98.3784 | 98.6423 | 121.2002 | 20.6915 | 0.002068 | 0.001933 | -154.6714 | 0.23369 | 8.4011 | 0.002839 | 0.003423 | 0.003263 | 4.4e-05 | 0.001196 | 0.30165 | 1.3643 | 2.5478 | 98.42 | 20.05 | 47.625 | 71.0204 | 89.5643 | 67.7786 | 0.011361 | 1.0355 | 0.024028 | 23.2037 | -0.023962 | 0.13902 | 8.1e-05 | 0.005776 | 0.005776 | 0.97003
2019-04-09 15:39 | 1554838743.721 | 785 | 0.20953 | 0.00045 | 1e-06 | 98.3209 | 98.8192 | 120.7011 | 20.6917 | 0.002063 | 0.001954 | -154.6569 | 0.23397 | 8.4016 | 0.002785 | 0.00461 | 0.004469 | 4.6e-05 | 0.001193 | 0.2891 | 1.3313 | 2.5476 | 98.42 | 20.04 | 47.62 | 70.9014 | 88.2639 | 68.2818 | 0.011352 | 1.034 | 0.023961 | 23.173 | -0.045215 | 0.13919 | 7.3e-05 | 0.005207 | 0.005207 | 1.0772
2019-04-09 15:39 | 1554838744.721 | 786 | 0.20952 | 0.00045 | 1e-06 | 98.354 | 98.8185 | 120.591 | 20.6919 | 0.002018 | 0.001954 | -154.6527 | 0.23367 | 8.4001 | 0.002863 | 0.004875 | 0.004734 | 4.6e-05 | 0.001189 | 0.27223 | 1.3751 | 2.5476 | 98.42 | 20.04 | 47.617 | 69.8354 | 90.1467 | 68.2982 | 0.011352 | 1.034 | 0.02395 | 23.163 | -0.042146 | 0.13901 | 0.000144 | 0.010287 | 0.010287 | 0.54456
2019-04-09 15:39 | 1554838745.7203 | 787 | 0.2095 | 0.000451 | 1e-06 | 98.3885 | 98.7941 | 120.6103 | 20.6913 | 0.00204 | 0.00185 | -154.6505 | 0.23321 | 8.4018 | 0.002864 | 0.00382 | 0.003681 | 4.3e-05 | 0.00119 | 0.2907 | 1.3706 | 2.5478 | 98.42 | 20.04 | 47.612 | 70.3643 | 90.1625 | 65.7949 | 0.01135 | 1.0416 | 0.02404 | 23.0798 | -0.036793 | 0.13873 | 0.000231 | 0.016576 | 0.016576 | 0.3373
2019-04-09 15:39 | 1554838746.7211 | 788 | 0.20951 | 0.000454 | 1e-06 | 98.4097 | 98.9142 | 120.7021 | 20.6917 | 0.002106 | 0.001948 | -154.6439 | 0.23323 | 8.4009 | 0.002748 | 0.0047 | 0.004493 | 4.3e-05 | 0.001184 | 0.30359 | 1.3417 | 2.5476 | 98.42 | 20.04 | 47.608 | 71.9351 | 87.3752 | 68.1507 | 0.01135 | 1.0344 | 0.023966 | 23.1686 | -0.045739 | 0.13875 | 0.000219 | 0.015648 | 0.015648 | 0.35735
2019-04-09 15:39 | 1554838747.7211 | 789 | 0.20953 | 0.000453 | 0 | 98.4493 | 98.8622 | 120.7718 | 20.6918 | 0.002046 | 0.001937 | -154.66 | 0.23315 | 8.4023 | 0.002832 | 0.003986 | 0.003848 | 4.4e-05 | 0.001185 | 0.26825 | 1.3225 | 2.5476 | 98.42 | 20.04 | 47.606 | 70.4943 | 89.395 | 67.888 | 0.011349 | 1.0352 | 0.023982 | 23.1664 | -0.037439 | 0.1387 | 6e-05 | 0.004311 | 0.004311 | 1.2966
2019-04-09 15:39 | 1554838748.7212 | 790 | 0.20954 | 0.000448 | 1e-06 | 98.5127 | 98.9452 | 120.8309 | 20.6915 | 0.002023 | 0.001917 | -154.6383 | 0.23264 | 8.4028 | 0.002813 | 0.004188 | 0.004055 | 4.3e-05 | 0.00118 | 0.27848 | 1.3933 | 2.5477 | 98.42 | 20.04 | 47.602 | 69.9451 | 88.9269 | 67.3946 | 0.011348 | 1.0367 | 0.024005 | 23.1553 | -0.039179 | 0.13839 | 8e-06 | 0.000534 | 0.000534 | 10.4483
2019-04-09 15:39 | 1554838749.7196 | 791 | 0.20952 | 0.000453 | 1e-06 | 98.5574 | 99.1534 | 120.7737 | 20.6924 | 0.002011 | 0.001974 | -154.6392 | 0.23295 | 8.4012 | 0.002775 | 0.004616 | 0.004447 | 4.3e-05 | 0.001178 | 0.28584 | 1.3815 | 2.5477 | 98.42 | 20.04 | 47.602 | 69.666 | 88.0348 | 68.7737 | 0.011348 | 1.0325 | 0.023951 | 23.1966 | -0.053927 | 0.13858 | 0.000142 | 0.010163 | 0.010163 | 0.54953
2019-04-09 15:39 | 1554838750.7206 | 792 | 0.20951 | 0.000452 | 1e-06 | 98.6375 | 99.12 | 120.6253 | 20.6921 | 0.002004 | 0.001944 | -154.6566 | 0.23353 | 8.4014 | 0.002835 | 0.004038 | 0.003882 | 4.2e-05 | 0.001177 | 0.28102 | 1.379 | 2.5476 | 98.42 | 20.04 | 47.605 | 69.5011 | 89.4695 | 68.0482 | 0.011349 | 1.0347 | 0.023962 | 23.1578 | -0.043648 | 0.13893 | 0.000197 | 0.014085 | 0.014085 | 0.39751
2019-04-09 15:39 | 1554838751.7202 | 793 | 0.20956 | 0.000454 | 1e-06 | 98.6762 | 99.1403 | 120.8014 | 20.6925 | 0.00207 | 0.001946 | -154.6241 | 0.2337 | 8.4045 | 0.002745 | 0.003616 | 0.003476 | 4.9e-05 | 0.001181 | 0.2896 | 1.3781 | 2.5478 | 98.42 | 20.04 | 47.608 | 71.0845 | 87.3088 | 68.1032 | 0.01135 | 1.0346 | 0.023977 | 23.1765 | -0.04197 | 0.13903 | -0.00012 | -0.008713 | -0.008713 | -0.64308
2019-04-09 15:39 | 1554838752.7203 | 794 | 0.20953 | 0.000456 | 1e-06 | 98.6444 | 99.1424 | 120.9183 | 20.6925 | 0.001975 | 0.001909 | -154.6241 | 0.23374 | 8.4022 | 0.002719 | 0.004263 | 0.004095 | 4.4e-05 | 0.001176 | 0.30934 | 1.3474 | 2.5476 | 98.42 | 20.04 | 47.611 | 68.802 | 86.6764 | 67.2147 | 0.01135 | 1.0373 | 0.02402 | 23.1576 | -0.045045 | 0.13905 | 9.5e-05 | 0.00676 | 0.00676 | 0.82901
2019-04-09 15:39 | 1554838753.7194 | 795 | 0.20954 | 0.000455 | 1e-06 | 98.6637 | 99.1813 | 121.0811 | 20.692 | 0.002059 | 0.001897 | -154.6392 | 0.2338 | 8.4033 | 0.002773 | 0.004219 | 0.003942 | 4.3e-05 | 0.001175 | 0.29087 | 1.3325 | 2.5476 | 98.42 | 20.03 | 47.618 | 70.8118 | 87.9784 | 66.9294 | 0.011345 | 1.0381 | 0.024047 | 23.1634 | -0.046798 | 0.13909 | 9e-06 | 0.000562 | 0.000562 | 9.9696
2019-04-09 15:39 | 1554838754.7204 | 796 | 0.20959 | 0.000453 | 1e-06 | 98.7164 | 99.1727 | 121.1539 | 20.6919 | 0.00207 | 0.001982 | -154.6371 | 0.23342 | 8.4015 | 0.002849 | 0.004457 | 0.004366 | 4.4e-05 | 0.001176 | 0.27616 | 1.3468 | 2.5477 | 98.42 | 20.03 | 47.624 | 71.0738 | 89.8084 | 68.9692 | 0.011346 | 1.0319 | 0.023982 | 23.2398 | -0.041245 | 0.13886 | -0.000278 | -0.020028 | -0.020028 | -0.27941
2019-04-09 15:39 | 1554838755.7197 | 797 | 0.20951 | 0.000455 | 1e-06 | 98.7932 | 99.364 | 121.4492 | 20.6923 | 0.002045 | 0.001895 | -154.6409 | 0.23329 | 8.4028 | 0.002816 | 0.004486 | 0.004298 | 4.5e-05 | 0.001181 | 0.27383 | 1.3525 | 2.5477 | 98.42 | 20.03 | 47.631 | 70.4831 | 89.0088 | 66.8657 | 0.011348 | 1.0383 | 0.024085 | 23.1964 | -0.051528 | 0.13878 | 0.000192 | 0.013789 | 0.013789 | 0.4056
2019-04-09 15:39 | 1554838756.7198 | 798 | 0.20955 | 0.000451 | 1e-06 | 98.8352 | 99.417 | 121.5046 | 20.6912 | 0.002038 | 0.001869 | -154.6382 | 0.23398 | 8.4018 | 0.002823 | 0.004236 | 0.004044 | 4.6e-05 | 0.001183 | 0.29192 | 1.3194 | 2.5475 | 98.42 | 20.03 | 47.645 | 70.2965 | 89.1689 | 66.2499 | 0.011351 | 1.0402 | 0.024113 | 23.1807 | -0.052497 | 0.1392 | -7e-05 | -0.005154 | -0.005154 | -1.0884
2019-04-09 15:39 | 1554838757.7194 | 799 | 0.20954 | 0.000448 | 0 | 98.8647 | 99.4655 | 121.7341 | 20.6918 | 0.001985 | 0.001927 | -154.6302 | 0.23362 | 8.4041 | 0.002859 | 0.00436 | 0.004177 | 4.5e-05 | 0.001179 | 0.26665 | 1.3526 | 2.5477 | 98.42 | 20.03 | 47.654 | 69.0237 | 90.0421 | 67.651 | 0.011353 | 1.0359 | 0.024086 | 23.2504 | -0.054194 | 0.13898 | 3.8e-05 | 0.002743 | 0.002743 | 2.0418
2019-04-09 15:39 | 1554838758.7205 | 800 | 0.20955 | 0.00045 | 1e-06 | 98.9031 | 99.173 | 121.5785 | 20.6918 | 0.002055 | 0.001933 | -154.6428 | 0.23324 | 8.4027 | 0.002812 | 0.004563 | 0.004378 | 4.6e-05 | 0.001179 | 0.29407 | 1.3674 | 2.5476 | 98.42 | 20.03 | 47.674 | 70.7251 | 88.921 | 67.7765 | 0.011358 | 1.0355 | 0.024066 | 23.2399 | -0.024374 | 0.13875 | -6.8e-05 | -0.004982 | -0.004982 | -1.1224
2019-04-09 15:39 | 1554838759.7196 | 801 | 0.20955 | 0.000454 | 1e-06 | 98.9463 | 99.3657 | 121.8524 | 20.6926 | 0.002055 | 0.001915 | -154.6326 | 0.23293 | 8.4017 | 0.002822 | 0.004331 | 0.004272 | 4.6e-05 | 0.001175 | 0.27041 | 1.4242 | 2.5477 | 98.42 | 20.03 | 47.683 | 70.7251 | 89.1466 | 67.3584 | 0.01136 | 1.0368 | 0.024108 | 23.2517 | -0.037835 | 0.13857 | -3.7e-05 | -0.002733 | -0.002733 | -2.0434
2019-04-09 15:39 | 1554838760.7205 | 802 | 0.20954 | 0.000446 | 1e-06 | 98.946 | 99.5574 | 121.9787 | 20.6926 | 0.002078 | 0.001944 | -154.6113 | 0.23399 | 8.4023 | 0.002901 | 0.003884 | 0.003684 | 4.5e-05 | 0.001177 | 0.28031 | 1.3735 | 2.5474 | 98.42 | 20.03 | 47.7 | 71.2733 | 91.0667 | 68.0542 | 0.011364 | 1.0347 | 0.024096 | 23.2875 | -0.055101 | 0.1392 | 2.6e-05 | 0.001823 | 0.001823 | 3.0764
2019-04-09 15:39 | 1554838761.7206 | 803 | 0.20955 | 0.000448 | 1e-06 | 98.9651 | 99.4766 | 122.0585 | 20.6924 | 0.00204 | 0.001895 | -154.6509 | 0.23369 | 8.4006 | 0.002846 | 0.004608 | 0.004499 | 4.4e-05 | 0.001177 | 0.30757 | 1.3703 | 2.5476 | 98.42 | 20.03 | 47.707 | 70.3609 | 89.7427 | 66.8595 | 0.011366 | 1.0383 | 0.024146 | 23.2543 | -0.046109 | 0.13902 | -6.1e-05 | -0.004502 | -0.004502 | -1.2444
2019-04-09 15:39 | 1554838762.7196 | 804 | 0.20954 | 0.000448 | 1e-06 | 99.0839 | 99.427 | 121.9254 | 20.6922 | 0.002044 | 0.00189 | -154.6356 | 0.23388 | 8.4035 | 0.002901 | 0.004329 | 0.004179 | 4.4e-05 | 0.001176 | 0.27688 | 1.3656 | 2.5475 | 98.42 | 20.03 | 47.714 | 70.451 | 91.0546 | 66.7487 | 0.011368 | 1.0387 | 0.024137 | 23.2379 | -0.03092 | 0.13913 | 2.7e-05 | 0.001879 | 0.001879 | 2.9841
2019-04-09 15:39 | 1554838763.7206 | 805 | 0.20955 | 0.00045 | 1e-06 | 99.0163 | 99.6507 | 121.8816 | 20.6917 | 0.002013 | 0.001929 | -154.6412 | 0.23209 | 8.402 | 0.002842 | 0.003859 | 0.003686 | 4.3e-05 | 0.001176 | 0.27831 | 1.3669 | 2.5476 | 98.42 | 20.03 | 47.728 | 69.6997 | 89.631 | 67.6923 | 0.011371 | 1.0358 | 0.024099 | 23.2659 | -0.057126 | 0.13807 | -6.1e-05 | -0.00452 | -0.00452 | -1.231
2019-04-09 15:39 | 1554838764.7211 | 806 | 0.20956 | 0.000448 | 1e-06 | 99.1095 | 99.7085 | 121.9027 | 20.6928 | 0.002019 | 0.001922 | -154.6466 | 0.23277 | 8.4035 | 0.002854 | 0.00385 | 0.003694 | 4.1e-05 | 0.001174 | 0.2959 | 1.3843 | 2.5478 | 98.42 | 20.03 | 47.733 | 69.8621 | 89.9164 | 67.52 | 0.011372 | 1.0363 | 0.024107 | 23.2621 | -0.053898 | 0.13848 | -0.000115 | -0.00836 | -0.00836 | -0.66755
2019-04-09 15:39 | 1554838765.7201 | 807 | 0.20955 | 0.000447 | 1e-06 | 99.0613 | 99.7051 | 121.7384 | 20.6923 | 0.002026 | 0.001909 | -154.6513 | 0.23346 | 8.4024 | 0.002852 | 0.003825 | 0.00371 | 4.4e-05 | 0.001174 | 0.27345 | 1.3502 | 2.5476 | 98.42 | 20.03 | 47.737 | 70.0175 | 89.8807 | 67.202 | 0.011373 | 1.0373 | 0.024102 | 23.2355 | -0.057944 | 0.13889 | -6.2e-05 | -0.004577 | -0.004577 | -1.2228
2019-04-09 15:39 | 1554838766.7188 | 808 | 0.20956 | 0.000452 | 1e-06 | 99.0583 | 99.5832 | 121.6806 | 20.6923 | 0.002069 | 0.00193 | -154.6576 | 0.23324 | 8.4013 | 0.002755 | 0.0041 | 0.003954 | 4.3e-05 | 0.001169 | 0.28694 | 1.3431 | 2.5476 | 98.42 | 20.03 | 47.746 | 71.0461 | 87.5313 | 67.7145 | 0.011375 | 1.0357 | 0.024078 | 23.2475 | -0.047277 | 0.13875 | -0.000103 | -0.007518 | -0.007518 | -0.74378
2019-04-09 15:39 | 1554838767.7191 | 809 | 0.20955 | 0.000451 | 1e-06 | 99.1393 | 99.6369 | 121.6229 | 20.693 | 0.002048 | 0.001883 | -154.6607 | 0.23347 | 8.4019 | 0.002789 | 0.003881 | 0.003635 | 4.5e-05 | 0.001172 | 0.2907 | 1.3153 | 2.5477 | 98.42 | 20.03 | 47.752 | 70.5476 | 88.369 | 66.5777 | 0.011377 | 1.0392 | 0.024113 | 23.2032 | -0.044787 | 0.13889 | -6.5e-05 | -0.004739 | -0.004739 | -1.1811
2019-04-09 15:39 | 1554838768.7187 | 810 | 0.20953 | 0.000449 | 1e-06 | 99.0988 | 99.607 | 121.6691 | 20.6924 | 0.002012 | 0.001916 | -154.6575 | 0.23332 | 8.4004 | 0.00277 | 0.004424 | 0.004289 | 4.4e-05 | 0.001181 | 0.28479 | 1.3374 | 2.5475 | 98.42 | 20.02 | 47.763 | 69.6878 | 87.9053 | 67.3721 | 0.011372 | 1.0368 | 0.024089 | 23.2347 | -0.045753 | 0.1388 | 6.8e-05 | 0.004868 | 0.004868 | 1.1491
2019-04-09 15:39 | 1554838769.7188 | 811 | 0.20958 | 0.000448 | 1e-06 | 99.1935 | 99.7801 | 121.4459 | 20.6929 | 0.002033 | 0.00188 | -154.6781 | 0.23347 | 8.4031 | 0.00282 | 0.004233 | 0.004086 | 4.3e-05 | 0.001183 | 0.29159 | 1.3216 | 2.5475 | 98.42 | 20.02 | 47.766 | 70.1821 | 89.0971 | 66.5051 | 0.011373 | 1.0394 | 0.024098 | 23.1838 | -0.052741 | 0.13889 | -0.000211 | -0.015273 | -0.015273 | -0.36648
2019-04-09 15:40 | 1554838770.7191 | 812 | 0.20958 | 0.000456 | 1e-06 | 99.1577 | 99.8455 | 121.376 | 20.693 | 0.001973 | 0.001953 | -154.6868 | 0.23352 | 8.4061 | 0.002822 | 0.004658 | 0.004519 | 4.7e-05 | 0.001191 | 0.29175 | 1.3265 | 2.5476 | 98.42 | 20.02 | 47.77 | 68.7487 | 89.1533 | 68.2551 | 0.011374 | 1.0341 | 0.024029 | 23.2368 | -0.061824 | 0.13892 | -0.000227 | -0.016437 | -0.016437 | -0.34061
2019-04-09 15:40 | 1554838771.7203 | 813 | 0.20954 | 0.000452 | 1e-06 | 99.2301 | 99.8365 | 121.2309 | 20.6927 | 0.002075 | 0.001906 | -154.6683 | 0.23358 | 8.4027 | 0.002788 | 0.004469 | 0.004242 | 4.5e-05 | 0.001192 | 0.28423 | 1.3192 | 2.5476 | 98.42 | 20.02 | 47.779 | 71.1931 | 88.3462 | 67.1275 | 0.011376 | 1.0375 | 0.024054 | 23.1845 | -0.054497 | 0.13896 | 4.3e-05 | 0.003036 | 0.003036 | 1.8446
2019-04-09 15:40 | 1554838772.7189 | 814 | 0.20955 | 0.000457 | 1e-06 | 99.2626 | 99.9105 | 121.2555 | 20.6929 | 0.00206 | 0.001877 | -154.6904 | 0.23341 | 8.4019 | 0.002771 | 0.003999 | 0.003877 | 4e-05 | 0.001188 | 0.27632 | 1.325 | 2.5476 | 98.42 | 20.02 | 47.783 | 70.8299 | 87.9323 | 66.4396 | 0.011377 | 1.0396 | 0.024081 | 23.1634 | -0.058199 | 0.13886 | -6.8e-05 | -0.00498 | -0.00498 | -1.1237
2019-04-09 15:40 | 1554838773.719 | 815 | 0.20952 | 0.000451 | 1e-06 | 99.2897 | 99.8032 | 121.1612 | 20.6932 | 0.001949 | 0.00192 | -154.6819 | 0.23372 | 8.403 | 0.002766 | 0.004128 | 0.004 | 4.1e-05 | 0.001187 | 0.27544 | 1.386 | 2.5478 | 98.42 | 20.02 | 47.793 | 68.1789 | 87.8022 | 67.472 | 0.011379 | 1.0365 | 0.024035 | 23.1896 | -0.046139 | 0.13904 | 0.000156 | 0.011179 | 0.011179 | 0.50122
2019-04-09 15:40 | 1554838774.719 | 816 | 0.20952 | 0.000455 | 1e-06 | 99.2852 | 99.839 | 121.0877 | 20.6933 | 0.002044 | 0.001905 | -154.6599 | 0.2346 | 8.4011 | 0.002831 | 0.00408 | 0.003981 | 4e-05 | 0.001187 | 0.28351 | 1.3397 | 2.5476 | 98.42 | 20.02 | 47.796 | 70.4559 | 89.3777 | 67.1196 | 0.01138 | 1.0375 | 0.024041 | 23.1705 | -0.049749 | 0.13956 | 0.000164 | 0.011774 | 0.011774 | 0.47771
2019-04-09 15:40 | 1554838775.719 | 817 | 0.20955 | 0.000455 | 1e-06 | 99.3266 | 99.8868 | 120.8654 | 20.6927 | 0.002021 | 0.001912 | -154.6575 | 0.23316 | 8.4028 | 0.002915 | 0.00396 | 0.00373 | 4.5e-05 | 0.001184 | 0.29231 | 1.3771 | 2.5476 | 98.42 | 20.02 | 47.8 | 69.9112 | 91.3849 | 67.2727 | 0.011381 | 1.0371 | 0.024013 | 23.1545 | -0.050305 | 0.1387 | -7e-05 | -0.005149 | -0.005149 | -1.0856
2019-04-09 15:40 | 1554838776.7201 | 818 | 0.20955 | 0.000454 | 1e-06 | 99.4529 | 100.0203 | 120.7338 | 20.6933 | 0.002049 | 0.001939 | -154.6453 | 0.2331 | 8.4026 | 0.002866 | 0.004344 | 0.004222 | 4.1e-05 | 0.001185 | 0.27975 | 1.3743 | 2.5476 | 98.42 | 20.02 | 47.811 | 70.5768 | 90.2045 | 67.9285 | 0.011384 | 1.0351 | 0.023977 | 23.1641 | -0.050891 | 0.13867 | -2.4e-05 | -0.001815 | -0.001815 | -3.0786
2019-04-09 15:40 | 1554838777.7191 | 819 | 0.20953 | 0.000459 | 1e-06 | 99.4058 | 99.8944 | 120.9598 | 20.6932 | 0.00202 | 0.001934 | -154.6512 | 0.23371 | 8.4024 | 0.002908 | 0.004716 | 0.004473 | 4.1e-05 | 0.001181 | 0.29607 | 1.3415 | 2.5477 | 98.42 | 20.02 | 47.816 | 69.8687 | 91.2252 | 67.8043 | 0.011385 | 1.0355 | 0.024004 | 23.1816 | -0.043853 | 0.13903 | 6e-05 | 0.004283 | 0.004283 | 1.3083
2019-04-09 15:40 | 1554838778.7202 | 820 | 0.2096 | 0.000456 | 1e-06 | 99.4633 | 99.9563 | 120.9997 | 20.6932 | 0.002007 | 0.002031 | -154.6551 | 0.23467 | 8.4011 | 0.002849 | 0.003502 | 0.003346 | 4.2e-05 | 0.001177 | 0.30469 | 1.3718 | 2.5476 | 98.42 | 20.02 | 47.82 | 69.5741 | 89.8176 | 70.1442 | 0.011386 | 1.0284 | 0.023926 | 23.2648 | -0.044228 | 0.13961 | -0.000353 | -0.025352 | -0.025352 | -0.22192
2019-04-09 15:40 | 1554838779.7188 | 821 | 0.20958 | 0.000458 | 1e-06 | 99.4945 | 100.1501 | 121.2518 | 20.6937 | 0.001926 | 0.001926 | -154.6224 | 0.23368 | 8.4008 | 0.002789 | 0.004231 | 0.004075 | 4.3e-05 | 0.001179 | 0.29286 | 1.3294 | 2.5474 | 98.42 | 20.02 | 47.825 | 67.6285 | 88.3534 | 67.6209 | 0.011387 | 1.036 | 0.024039 | 23.2033 | -0.058738 | 0.13901 | -0.000241 | -0.017393 | -0.017393 | -0.3221
2019-04-09 15:40 | 1554838780.7186 | 822 | 0.20953 | 0.000459 | 1e-06 | 99.474 | 100.1218 | 121.6534 | 20.6937 | 0.002002 | 0.001933 | -154.6278 | 0.23377 | 8.4018 | 0.002789 | 0.004296 | 0.004114 | 4.4e-05 | 0.001176 | 0.29103 | 1.3135 | 2.5474 | 98.42 | 20.02 | 47.828 | 69.448 | 88.3524 | 67.7922 | 0.011388 | 1.0355 | 0.024073 | 23.2475 | -0.058064 | 0.13907 | 7.6e-05 | 0.005398 | 0.005398 | 1.0383
2019-04-09 15:40 | 1554838781.7193 | 823 | 0.20957 | 0.000458 | 1e-06 | 99.5312 | 99.9663 | 121.8102 | 20.694 | 0.00198 | 0.001931 | -154.6172 | 0.2337 | 8.403 | 0.002816 | 0.004636 | 0.004488 | 4.2e-05 | 0.001176 | 0.28385 | 1.3443 | 2.5477 | 98.42 | 20.02 | 47.835 | 68.9094 | 89.0222 | 67.7355 | 0.011389 | 1.0357 | 0.02409 | 23.2606 | -0.039017 | 0.13903 | -0.000149 | -0.010799 | -0.010799 | -0.51884
2019-04-09 15:40 | 1554838782.7196 | 824 | 0.20955 | 0.000451 | 1e-06 | 99.4615 | 100.0657 | 122.0491 | 20.6938 | 0.002066 | 0.001948 | -154.6011 | 0.23373 | 8.3995 | 0.002812 | 0.003794 | 0.003623 | 4.2e-05 | 0.001175 | 0.26521 | 1.3177 | 2.5475 | 98.42 | 20.02 | 47.838 | 70.9899 | 88.9251 | 68.1538 | 0.01139 | 1.0344 | 0.024099 | 23.2977 | -0.054176 | 0.13905 | -6.4e-05 | -0.004664 | -0.004664 | -1.2015
2019-04-09 15:40 | 1554838783.7196 | 825 | 0.20958 | 0.000454 | 1e-06 | 99.5333 | 99.9962 | 122.5388 | 20.6943 | 0.002042 | 0.0019 | -154.6059 | 0.23416 | 8.4015 | 0.002806 | 0.003952 | 0.003863 | 4.5e-05 | 0.001173 | 0.26034 | 1.3153 | 2.5475 | 98.42 | 20.02 | 47.841 | 70.4016 | 88.7733 | 67.0003 | 0.011391 | 1.0379 | 0.024188 | 23.3049 | -0.041508 | 0.1393 | -0.000237 | -0.017287 | -0.017287 | -0.32473
2019-04-09 15:40 | 1554838784.7183 | 826 | 0.20951 | 0.00045 | 1e-06 | 99.5607 | 100.2561 | 122.6827 | 20.6933 | 0.002014 | 0.001892 | -154.6241 | 0.2337 | 8.4027 | 0.002795 | 0.004258 | 0.004132 | 4.3e-05 | 0.001172 | 0.28368 | 1.3237 | 2.5476 | 98.42 | 20.02 | 47.848 | 69.7294 | 88.495 | 66.7988 | 0.011393 | 1.0385 | 0.02421 | 23.3116 | -0.062258 | 0.13903 | 0.000226 | 0.016366 | 0.016366 | 0.34233
2019-04-09 15:40 | 1554838785.7187 | 827 | 0.20954 | 0.000453 | 2e-06 | 99.5601 | 100.2815 | 122.8594 | 20.6933 | 0.002005 | 0.00192 | -154.6139 | 0.23326 | 8.404 | 0.002835 | 0.004307 | 0.004132 | 4.6e-05 | 0.001171 | 0.28351 | 1.321 | 2.5476 | 98.42 | 20.02 | 47.851 | 69.522 | 89.4676 | 67.476 | 0.011393 | 1.0365 | 0.024203 | 23.3517 | -0.064584 | 0.13877 | 4.5e-05 | 0.003147 | 0.003147 | 1.7772
2019-04-09 15:40 | 1554838786.7184 | 828 | 0.20956 | 0.000451 | 1e-06 | 99.6661 | 100.3287 | 123.1174 | 20.6936 | 0.002046 | 0.001913 | -154.6269 | 0.23442 | 8.4026 | 0.002843 | 0.00401 | 0.003869 | 4.6e-05 | 0.001164 | 0.27041 | 1.3259 | 2.5475 | 98.42 | 20.02 | 47.86 | 70.4993 | 89.6687 | 67.3171 | 0.011395 | 1.0369 | 0.024234 | 23.3707 | -0.059269 | 0.13945 | -9.2e-05 | -0.006758 | -0.006758 | -0.83163
2019-04-09 15:40 | 1554838787.7194 | 829 | 0.20956 | 0.000452 | 1e-06 | 99.5669 | 100.2519 | 123.1864 | 20.6928 | 0.002034 | 0.001926 | -154.6094 | 0.23269 | 8.4017 | 0.002753 | 0.003997 | 0.003817 | 4.5e-05 | 0.001169 | 0.27936 | 1.3404 | 2.5476 | 98.42 | 20.02 | 47.866 | 70.2098 | 87.4918 | 67.6168 | 0.011397 | 1.036 | 0.02423 | 23.3875 | -0.061328 | 0.13843 | -0.000122 | -0.008972 | -0.008972 | -0.62174
2019-04-09 15:40 | 1554838788.7196 | 830 | 0.20956 | 0.000457 | 1e-06 | 99.5434 | 100.2836 | 123.3573 | 20.6942 | 0.00206 | 0.001978 | -154.5984 | 0.23317 | 8.4007 | 0.002825 | 0.00415 | 0.003979 | 4.5e-05 | 0.00117 | 0.3011 | 1.3617 | 2.5475 | 98.42 | 20.02 | 47.87 | 70.8382 | 89.2279 | 68.8676 | 0.011398 | 1.0322 | 0.024203 | 23.4467 | -0.066266 | 0.13871 | -0.000121 | -0.008859 | -0.008859 | -0.63101
2019-04-09 15:40 | 1554838789.7185 | 831 | 0.20956 | 0.000455 | 2e-06 | 99.6911 | 100.1738 | 123.3242 | 20.6948 | 0.002032 | 0.001957 | -154.6125 | 0.23228 | 8.4049 | 0.002829 | 0.004294 | 0.004094 | 4.8e-05 | 0.001169 | 0.29839 | 1.3955 | 2.5477 | 98.42 | 20.02 | 47.88 | 70.1635 | 89.3339 | 68.3668 | 0.0114 | 1.0338 | 0.024217 | 23.4263 | -0.043207 | 0.13819 | -0.000128 | -0.009367 | -0.009367 | -0.59455
2019-04-09 15:40 | 1554838790.7201 | 832 | 0.20955 | 0.000455 | 1e-06 | 99.6369 | 100.2626 | 123.2216 | 20.6947 | 0.002036 | 0.001904 | -154.6013 | 0.2334 | 8.3992 | 0.002795 | 0.003817 | 0.00369 | 4.6e-05 | 0.00117 | 0.28136 | 1.4102 | 2.5477 | 98.42 | 20.02 | 47.886 | 70.2571 | 88.5012 | 67.0984 | 0.011402 | 1.0376 | 0.024252 | 23.3731 | -0.055995 | 0.13885 | -6.5e-05 | -0.004762 | -0.004762 | -1.175
2019-04-09 15:40 | 1554838791.7195 | 833 | 0.20956 | 0.000457 | 1e-06 | 99.6277 | 100.348 | 123.0272 | 20.6939 | 0.00204 | 0.001889 | -154.6139 | 0.23231 | 8.401 | 0.002844 | 0.004099 | 0.003963 | 4.8e-05 | 0.001167 | 0.28368 | 1.425 | 2.5476 | 98.42 | 20.02 | 47.891 | 70.3678 | 89.6839 | 66.7378 | 0.011403 | 1.0387 | 0.024246 | 23.3423 | -0.064435 | 0.1382 | -8.1e-05 | -0.005959 | -0.005959 | -0.93458
2019-04-09 15:40 | 1554838792.7195 | 834 |  | 0.000453 |  | 99.6759 | 100.2506 | 122.647 | 20.6934 | 0.002052 | 0.001901 | -154.6146 | 0.23225 | 8.4034 | 0.002861 | 0.004105 | 0.003941 | 4.4e-05 | 0.001171 | 0.26609 | 1.3812 | 2.5476 | 98.42 | 20.02 | 47.897 | 70.6581 | 90.1005 | 67.0127 | 0.011404 | 1.0379 | 0.024199 | 23.3156 | -0.051419 | 0.13817
2019-04-09 15:40 | 1554838793.7195 | 835 |  | 0.000456 |  | 99.6959 | 100.3641 | 122.4493 | 20.6939 | 0.001995 | 0.001934 | -154.6119 | 0.23379 | 8.4021 | 0.002789 | 0.004425 | 0.004257 | 4.7e-05 | 0.001167 | 0.29302 | 1.3589 | 2.5476 | 98.42 | 20.02 | 47.91 | 69.2839 | 88.3668 | 67.8105 | 0.011407 | 1.0354 | 0.024151 | 23.3241 | -0.059757 | 0.13908
2019-04-09 15:40 | 1554838794.7195 | 836 |  |  |  | 99.7268 | 100.5221 | 122.185 | 20.693 | 0.002029 | 0.001937 | -154.6167 | 0.23362 | 8.4025 | 0.002814 | 0.003417 | 0.003232 | 4.7e-05 | 0.001168 | 0.29679 | 1.3938 | 2.5477 | 98.42 | 20.02 | 47.916 | 70.087 | 88.9534 | 67.8844 | 0.011409 | 1.0352 | 0.024122 | 23.3014 | -0.071045 | 0.13898
2019-04-09 15:40 | 1554838795.7195 | 837 |  |  |  | 99.7572 | 100.2857 | 121.983 | 20.6944 | 0.002013 | 0.001932 | -154.635 | 0.23282 | 8.4014 | 0.002918 | 0.004528 | 0.004346 | 4.6e-05 | 0.001167 | 0.30773 | 1.4155 | 2.5478 | 98.42 | 20.02 | 47.922 | 69.7163 | 91.4805 | 67.772 | 0.01141 | 1.0356 | 0.024106 | 23.2783 | -0.047258 | 0.1385
2019-04-09 15:40 | 1554838796.7195 | 838 |  |  |  | 99.7134 | 100.2691 | 121.7508 | 20.6933 | 0.002062 | 0.001935 | -154.6145 | 0.23308 | 8.4042 | 0.00286 | 0.004116 | 0.003947 | 4.7e-05 | 0.001169 | 0.29031 | 1.3621 | 2.5476 | 98.42 | 20.01 | 47.934 | 70.8841 | 90.0638 | 67.8356 | 0.011406 | 1.0354 | 0.024081 | 23.2583 | -0.049711 | 0.13866
2019-04-09 15:40 | 1554838797.7195 | 839 |  |  |  | 99.7596 | 100.2189 |  | 20.6936 | 0.002042 | 0.001934 | -154.6149 | 0.23268 | 8.4008 | 0.002891 | 0.004899 | 0.004724 | 4.5e-05 | 0.001166 | 0.30735 | 1.3876 | 2.5477 | 98.42 | 20.02 | 47.939 | 70.3964 | 90.8121 | 67.822 | 0.011414 | 1.0354 |  |  | -0.041088 | 0.13842
2019-04-09 15:40 | 1554838798.7195 | 840 |  |  |  | 99.8188 | 100.5317 |  | 20.6937 | 0.002014 |  | -154.6207 | 0.23261 | 8.4038 | 0.00287 | 0.004331 | 0.004195 | 5.1e-05 | 0.001167 | 0.3115 | 1.3505 | 2.5477 | 98.42 | 20.02 | 47.949 | 69.7249 | 90.3024 |  | 0.011417 |  |  |  | -0.063653 | 0.13838
2019-04-09 15:40 | 1554838799.7195 | 841 |  |  |  | 99.7956 | 100.4202 |  | 20.694 | 0.002034 |  |  | 0.23319 | 8.403 | 0.002788 | 0.004347 | 0.004195 | 4.7e-05 | 0.001158 | 0.28816 | 1.3605 | 2.5476 | 98.42 | 20.02 | 47.951 | 70.2032 | 88.345 |  | 0.011417 |  |  |  | -0.055809 | 0.13873
2019-04-09 15:40 | 1554838800.7195 | 842 |  |  |  | 99.8316 | 100.3597 |  | 20.6936 |  |  |  | 0.23332 | 8.4 | 0.002785 | 0.004284 | 0.00415 | 4.7e-05 | 0.001151 | 0.26394 | 1.375 | 2.5476 | 98.42 | 20.02 | 47.954 |  | 88.2548 |  | 0.011418 |  |  |  | -0.047183 | 0.1388
2019-04-09 15:40 | 1554838801.7195 | 843 |  |  |  | 99.8507 | 100.4987 |  | 20.6932 |  |  |  | 0.23319 | 8.402 | 0.002832 | 0.004568 | 0.004388 | 5.1e-05 | 0.00115 | 0.28224 | 1.3254 | 2.5475 | 98.42 | 20.02 | 47.957 |  | 89.405 |  | 0.011419 |  |  |  | -0.057859 | 0.13872
2019-04-09 15:40 | 1554838802.7195 | 844 |  |  |  | 99.9094 | 100.4829 |  | 20.6941 |  |  |  | 0.23315 | 8.4016 | 0.002787 | 0.003952 | 0.00377 | 4.6e-05 | 0.001149 | 0.29286 | 1.3273 | 2.5475 | 98.42 | 20.02 | 47.959 |  | 88.3114 |  | 0.011419 |  |  |  | -0.051196 | 0.1387
2019-04-09 15:40 | 1554838803.7195 | 845 |  |  |  | 99.9103 | 100.6226 |  | 20.6942 |  |  |  | 0.23317 | 8.4041 | 0.002818 | 0.004036 | 0.003903 | 5e-05 | 0.001145 | 0.27649 | 1.3355 | 2.5477 | 98.42 | 20.02 | 47.961 |  | 89.0514 |  | 0.011419 |  |  |  | -0.063545 | 0.13872
2019-04-09 15:40 | 1554838804.7195

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-04-09 15:26 | 0 | 0 | 0 | 0 | 0 | 43.6 | -0.004237
2019-04-09 15:26 | 1 | 0 | 0 | 0 | 0 | 44.1131 | 0.017714
2019-04-09 15:26 | 2 | 0 | 0 | 0 | 0 | 43.9452 | 0.000274
2019-04-09 15:26 | 3 | 0 | 0 | 0 | 0 | 43.6013 | 0.068778
2019-04-09 15:26 | 4 | 0 | 0 | 0 | 0 | 43.9805 | 0.39025
2019-04-09 15:26 | 5 | 0 | 0 | 0 | 0 | 43.8013 | 0.9343
2019-04-09 15:26 | 6 | 0 | 0 | 0 | 0 | 43.5629 | 1.5774
2019-04-09 15:26 | 7 | 0 | 0 | 0 | 0 | 43.4513 | 1.8376
2019-04-09 15:26 | 8 | 0 | 0 | 0 | 0 | 43.3755 | 2.2113
2019-04-09 15:26 | 9 | 0 | 0 | 0 | 0 | 43.2574 | 2.4747
2019-04-09 15:26 | 10 | 0 | 0 | 0 | 0 | 43.1236 | 2.7049
2019-04-09 15:26 | 11 | 0 | 0 | 0 | 0 | 43.0151 | 2.8427
2019-04-09 15:26 | 12 | 0 | 0 | 0 | 0 | 42.9275 | 3.0317
2019-04-09 15:26 | 13 | 0.15096 | 0 | 0 | 0 | 42.77 | 3.1463
2019-04-09 15:26 | 14 | 0.12563 | 0 | 0 | 0 | 42.6384 | 3.1435
2019-04-09 15:26 | 15 | 0.12753 | 0 | 0 | 0 | 42.5135 | 3.1271
2019-04-09 15:26 | 16 | 0.12942 | 0 | 0 | 0 | 42.3831 | 3.1456
2019-04-09 15:26 | 17 | 0.11311 | 0 | 0 | 0 | 42.2582 | 3.1285
2019-04-09 15:26 | 18 | 0.13065 | 0 | 0 | 0 | 42.149 | 3.1111
2019-04-09 15:26 | 19 | 0.14406 | 0 | 0 | 0 | 41.998 | 3.0976
2019-04-09 15:26 | 20 | 0.12011 | 0 | 0 | 0 | 41.8693 | 3.0708
2019-04-09 15:26 | 21 | 0.13367 | 0 | 0 | 0 | 41.7496 | 3.0361
2019-04-09 15:26 | 22 | 0.14552 | 0 | 0 | 0 | 41.603 | 3.0482
2019-04-09 15:26 | 23 | 0.099253 | 0 | 0 | 0 | 41.4718 | 2.9762
2019-04-09 15:26 | 24 | 0.1025 | 0 | 0 | 0 | 41.3927 | 2.9491
2019-04-09 15:26 | 25 | 0.1435 | 0 | 0 | 0 | 41.2578 | 2.9058
2019-04-09 15:26 | 26 | 0.11443 | 0 | 0 | 0 | 41.1211 | 2.867
2019-04-09 15:26 | 27 | 0.12104 | 0 | 0 | 0 | 41.0215 | 2.8095
2019-04-09 15:26 | 28 | 0.1457 | 0 | 0 | 0 | 40.8754 | 2.7392
2019-04-09 15:26 | 29 | 0.1084 | 0 | 0 | 0 | 40.7442 | 2.6915
2019-04-09 15:26 | 30 | 0.10163 | 0 | 0 | 0 | 40.6514 | 2.6946
2019-04-09 15:26 | 31 | 0.13978 | 0 | 0 | 0 | 40.5302 | 2.6826
2019-04-09 15:26 | 32 | 0.12774 | 0 | 0 | 0 | 40.3828 | 2.7042
2019-04-09 15:26 | 33 | 0.12703 | 0 | 0 | 0 | 40.273 | 2.7315
2019-04-09 15:27 | 34 | 0.13844 | 0 | 0 | 0 | 40.1266 | 2.7333
2019-04-09 15:27 | 35 | 0.11016 | 0 | 0 | 0 | 40.0054 | 2.7341
2019-04-09 15:27 | 36 | 0.094879 | 0 | 0 | 0 | 39.9032 | 2.7402
2019-04-09 15:27 | 37 | 0.11202 | 0 | 0 | 0 | 39.8073 | 2.704
2019-04-09 15:27 | 38 | 0.12104 | 0 | 0 | 0 | 39.6808 | 2.6692
2019-04-09 15:27 | 39 | 0.10973 | 0 | 0 | 0 | 39.5701 | 2.616
2019-04-09 15:27 | 40 | 0.11672 | 0 | 0 | 0 | 39.4578 | 2.562
2019-04-09 15:27 | 41 | 0.10615 | 0 | 0 | 0 | 39.3403 | 2.5085
2019-04-09 15:27 | 42 | 0.10372 | 0.026308 | 2.2296 | 23.9924 | 39.2436 | 2.4885
2019-04-09 15:27 | 43 | 0.11185 | 0.023649 | 2.049 | 21.744 | 39.13 | 2.432
2019-04-09 15:27 | 44 | 0.12091 | 0.021357 | 1.8786 | 19.5094 | 39.0198 | 2.3588
2019-04-09 15:27 | 45 | 0.1097 | 0.022902 | 2.0433 | 21.2929 | 38.8928 | 2.3358
2019-04-09 15:27 | 46 | 0.11637 | 0.021266 | 1.9077 | 20.0475 | 38.7963 | 2.3328
2019-04-09 15:27 | 47 | 0.13348 | 0.018384 | 1.6445 | 17.4434 | 38.6581 | 2.3284
2019-04-09 15:27 | 48 | 0.10983 | 0.021846 | 1.9655 | 21.0253 | 38.5386 | 2.3092
2019-04-09 15:27 | 49 | 0.10994 | 0.021808 | 1.9426 | 21.1861 | 38.4331 | 2.3293
2019-04-09 15:27 | 50 | 0.12651 | 0.018941 | 1.6728 | 18.3784 | 38.3145 | 2.325
2019-04-09 15:27 | 51 | 0.12491 | 0.019137 | 1.6788 | 18.7763 | 38.1847 | 2.3453
2019-04-09 15:27 | 52 | 0.10315 | 0.023155 | 2.0218 | 22.9123 | 38.0696 | 2.3634
2019-04-09 15:27 | 53 | 0.1071 | 0.022092 | 1.9407 | 22.22 | 37.9725 | 2.3797
2019-04-09 15:27 | 54 | 0.11075 | 0.021316 | 1.8702 | 22.3523 | 37.8547 | 2.4754
2019-04-09 15:27 | 55 | 0.11266 | 0.020753 | 1.8398 | 22.5275 | 37.7508 | 2.5378
2019-04-09 15:27 | 56 | 0.13642 | 0.016999 | 1.5262 | 18.8664 | 37.6249 | 2.5737
2019-04-09 15:27 | 57 | 0.12864 | 0.017772 | 1.6203 | 20.811 | 37.4858 | 2.6771
2019-04-09 15:27 | 58 | 0.10253 | 0.022158 | 2.0578 | 26.2768 | 37.3721 | 2.6941
2019-04-09 15:27 | 59 | 0.10997 | 0.020604 | 1.9324 | 24.5249 | 37.2725 | 2.697
2019-04-09 15:27 | 60 | 0.13161 | 0.017203 | 1.6328 | 20.5035 | 37.1487 | 2.6985
2019-04-09 15:27 | 61 | 0.11399 | 0.019637 | 1.9027 | 23.8077 | 37.0182 | 2.7137
2019-04-09 15:27 | 62 | 0.11471 | 0.019516 | 1.911 | 24.162 | 36.9165 | 2.7716
2019-04-09 15:27 | 63 | 0.14327 | 0.015601 | 1.5436 | 19.6204 | 36.7827 | 2.8109
2019-04-09 15:27 | 64 | 0.12241 | 0.018049 | 1.8216 | 23.1268 | 36.6413 | 2.8309
2019-04-09 15:27 | 65 | 0.11263 | 0.019726 | 1.9938 | 24.9429 | 36.5359 | 2.8092
2019-04-09 15:27 | 66 | 0.12338 | 0.01786 | 1.8403 | 22.775 | 36.411 | 2.81
2019-04-09 15:27 | 67 | 0.11929 | 0.018429 | 1.9167 | 23.5182 | 36.2918 | 2.8055
2019-04-09 15:27 | 68 | 0.12635 | 0.017413 | 1.8259 | 22.3609 | 36.1696 | 2.8253
2019-04-09 15:27 | 69 | 0.1506 | 0.014451 | 1.5398 | 18.791 | 36.0359 | 2.83
2019-04-09 15:27 | 70 | 0.12751 | 0.017062 | 1.8327 | 22.1176 | 35.8797 | 2.8202
2019-04-09 15:27 | 71 | 0.10702 | 0.020339 | 2.199 | 26.4519 | 35.7802 | 2.831
2019-04-09 15:27 | 72 | 0.13885 | 0.015539 | 1.707 | 20.3313 | 35.6537 | 2.8231
2019-04-09 15:27 | 73 | 0.12663 | 0.016691 | 1.8768 | 22.1333 | 35.5122 | 2.8028
2019-04-09 15:27 | 74 | 0.10699 | 0.019625 | 2.2363 | 26.2059 | 35.4024 | 2.8037
2019-04-09 15:27 | 75 | 0.12917 | 0.016242 | 1.8528 | 21.5859 | 35.2886 | 2.7883
2019-04-09 15:27 | 76 | 0.13262 | 0.015658 | 1.8076 | 20.9557 | 35.1486 | 2.7791
2019-04-09 15:27 | 77 | 0.098112 | 0.020926 | 2.4419 | 27.9556 | 35.0322 | 2.7428
2019-04-09 15:27 | 78 | 0.10502 | 0.019194 | 2.2857 | 25.6081 | 34.943 | 2.6894
2019-04-09 15:27 | 79 | 0.1122 | 0.017794 | 2.138 | 23.958 | 34.8214 | 2.688
2019-04-09 15:27 | 80 | 0.10286 | 0.018932 | 2.3171 | 26.0533 | 34.722 | 2.6799
2019-04-09 15:27 | 81 | 0.1199 | 0.016089 | 1.9803 | 22.4193 | 34.6102 | 2.6881
2019-04-09 15:27 | 82 | 0.11272 | 0.016881 | 2.0925 | 23.7515 | 34.4878 | 2.6773
2019-04-09 15:27 | 83 | 0.09712 | 0.019402 | 2.4223 | 27.6573 | 34.3864 | 2.6861
2019-04-09 15:27 | 84 | 0.12128 | 0.015517 | 1.9311 | 22.1652 | 34.2844 | 2.6882
2019-04-09 15:27 | 85 | 0.12265 | 0.015514 | 1.9075 | 21.9949 | 34.149 | 2.6976
2019-04-09 15:27 | 86 | 0.10287 | 0.018882 | 2.2708 | 26.2235 | 34.044 | 2.6976
2019-04-09 15:27 | 87 | 0.1203 | 0.016377 | 1.9489 | 22.9034 | 33.935 | 2.7552
2019-04-09 15:27 | 88 | 0.12053 | 0.016332 | 1.9433 | 22.8575 | 33.807 | 2.755
2019-04-09 15:27 | 89 | 0.11085 | 0.017757 | 2.1172 | 24.8146 | 33.6962 | 2.7508
2019-04-09 15:27 | 90 | 0.12434 | 0.015861 | 1.8955 | 22.0951 | 33.5799 | 2.7473
2019-04-09 15:27 | 91 | 0.13041 | 0.015102 | 1.8089 | 20.5976 | 33.4488 | 2.6862
2019-04-09 15:27 | 92 | 0.13506 | 0.014459 | 1.7526 | 19.9409 | 33.32 | 2.6932
2019-04-09 15:27 | 93 | 0.12989 | 0.015098 | 1.8107 | 20.468 | 33.1813 | 2.6585
2019-04-09 15:28 | 94 | 0.12055 | 0.016248 | 1.9259 | 21.9654 | 33.0605 | 2.6479
2019-04-09 15:28 | 95 | 0.14353 | 0.013792 | 1.5813 | 17.9203 | 32.9319 | 2.5722
2019-04-09 15:28 | 96 | 0.17267 | 0.012149 | 1.2863 | 14.8274 | 32.7712 | 2.5603
2019-04-09 15:28 | 97 | 0.18891 | 0.012121 | 1.1548 | 13.6514 | 32.5891 | 2.5789
2019-04-09 15:28 | 98 | 0.21959 | 0.011679 | 0.97684 | 12.8593 | 32.3903 | 2.8237
2019-04-09 15:28 | 99 | 0.24482 | 0.011775 | 0.86813 | 17.8047 | 32.1517 | 4.359
2019-04-09 15:28 | 100 | 0.24221 | 0.015936 | 0.91874 | 25.6117 | 31.9069 | 6.2033
2019-04-09 15:28 | 101 | 0.27405 | 0.01951 | 0.95306 | 27.7935 | 31.6599 | 7.6169
2019-04-09 15:28 | 102 | 0.28615 | 0.024376 | 1.1221 | 27.6172 | 31.3627 | 7.9026
2019-04-09 15:28 | 103 | 0.2991 | 0.027917 | 1.3211 | 26.9814 | 31.0869 | 8.0702
2019-04-09 15:28 | 104 | 0.34827 | 0.026957 | 1.3584 | 23.7225 | 30.7565 | 8.2619
2019-04-09 15:28 | 105 | 0.36238 | 0.028395 | 1.5152 | 23.47 | 30.3993 | 8.505
2019-04-09 15:28 | 106 | 0.3412 | 0.032738 | 1.7784 | 25.299 | 30.0407 | 8.632
2019-04-09 15:28 | 107 | 0.35306 | 0.03411 | 1.8142 | 24.3667 | 29.7092 | 8.6029
2019-04-09 15:28 | 108 | 0.38178 | 0.033719 | 1.7003 | 22.4221 | 29.3306 | 8.5603
2019-04-09 15:28 | 109 | 0.36835 | 0.036929 | 1.7795 | 23.2433 | 28.9548 | 8.5618
2019-04-09 15:28 | 110 | 0.38598 | 0.036304 | 1.7107 | 22.3458 | 28.5868 | 8.625
2019-04-09 15:28 | 111 | 0.41423 | 0.034199 | 1.5932 | 21.5123 | 28.1797 | 8.9111
2019-04-09 15:28 | 112 | 0.43031 | 0.032853 | 1.5272 | 21.2725 | 27.761 | 9.1537
2019-04-09 15:28 | 113 | 0.44959 | 0.031215 | 1.4661 | 21.2392 | 27.3189 | 9.549
2019-04-09 15:28 | 114 | 0.4594 | 0.029906 | 1.4596 | 22.3072 | 26.8646 | 10.248
2019-04-09 15:28 | 115 | 0.45584 | 0.028924 | 1.5216 | 23.0942 | 26.4033 | 10.5273
2019-04-09 15:28 | 116 | 0.46874 | 0.026818 | 1.5463 | 23.6757 | 25.9491 | 11.0977
2019-04-09 15:28 | 117 | 0.48726 | 0.024507 | 1.5538 | 22.8447 | 25.4648 | 11.1312
2019-04-09 15:28 | 118 | 0.47538 | 0.024023 | 1.669 | 23.8769 | 24.9816 | 11.3506
2019-04-09 15:28 | 119 | 0.48031 | 0.022775 | 1.7313 | 24.2604 | 24.5104 | 11.6525
2019-04-09 15:28 | 120 | 0.49934 | 0.020691 | 1.7265 | 23.7955 | 24.018 | 11.8819
2019-04-09 15:28 | 121 | 0.48277 | 0.019955 | 1.849 | 25.1909 | 23.5204 | 12.1613
2019-04-09 15:28 | 122 | 0.46253 | 0.019391 | 1.9969 | 27.2289 | 23.0537 | 12.5942
2019-04-09 15:28 | 123 | 0.47078 | 0.017847 | 2.0097 | 26.6385 | 22.5889 | 12.5408
2019-04-09 15:28 | 124 | 0.46156 | 0.017057 | 2.0869 | 27.1913 | 22.1162 | 12.5504
2019-04-09 15:28 | 125 | 0.43311 | 0.017124 | 2.2722 | 28.8166 | 21.6698 | 12.4807
2019-04-09 15:28 | 126 | 0.44853 | 0.01536 | 2.2404 | 27.4069 | 21.2401 | 12.2927
2019-04-09 15:28 | 127 | 0.44579 | 0.014621 | 2.2864 | 27.1118 | 20.777 | 12.0862
2019-04-09 15:28 | 128 | 0.41347 | 0.014883 | 2.4938 | 28.021 | 20.3555 | 11.5857
2019-04-09 15:28 | 129 | 0.41473 | 0.014038 | 2.4977 | 26.4567 | 19.9425 | 10.9723
2019-04-09 15:28 | 130 | 0.41473 | 0.013281 | 2.4973 | 25.187 | 19.5261 | 10.4458
2019-04-09 15:28 | 131 | 0.38844 | 0.013487 | 2.6435 | 26.3016 | 19.1189 | 10.2165
2019-04-09 15:28 | 132 | 0.38726 | 0.012879 | 2.6033 | 25.8271 | 18.7434 | 10.0017
2019-04-09 15:28 | 133 | 0.39504 | 0.012046 | 2.4583 | 25.0491 | 18.3424 | 9.8955
2019-04-09 15:28 | 134 | 0.36801 | 0.012303 | 2.5629 | 26.681 | 17.9609 | 9.8188
2019-04-09 15:28 | 135 | 0.37423 | 0.011506 | 2.4569 | 25.6067 | 17.599 | 9.5828
2019-04-09 15:28 | 136 | 0.37987 | 0.01077 | 2.3742 | 24.7516 | 17.2124 | 9.4023
2019-04-09 15:28 | 137 | 0.36154 | 0.010676 | 2.4455 | 25.7464 | 16.8449 | 9.3084
2019-04-09 15:28 | 138 | 0.35683 | 0.010349 | 2.4395 | 25.7891 | 16.4864 | 9.2022
2019-04-09 15:28 | 139 | 0.35837 | 0.009831 | 2.3914 | 25.5015 | 16.1298 | 9.1389
2019-04-09 15:28 | 140 | 0.34145 | 0.0099132 | 2.4723 | 26.298 | 15.7739 | 8.9796
2019-04-09 15:28 | 141 | 0.33155 | 0.0098709 | 2.5179 | 27.0306 | 15.445 | 8.9619
2019-04-09 15:28 | 142 | 0.34683 | 0.0091159 | 2.3828 | 25.5413 | 15.105 | 8.8584
2019-04-09 15:28 | 143 | 0.33641 | 0.0091546 | 2.4299 | 26.018 | 14.7572 | 8.7527
2019-04-09 15:28 | 144 | 0.31901 | 0.0093901 | 2.5349 | 27.1883 | 14.4337 | 8.6732
2019-04-09 15:28 | 145 | 0.333 | 0.0087435 | 2.4037 | 25.9485 | 14.112 | 8.6408
2019-04-09 15:28 | 146 | 0.33142 | 0.0085861 | 2.3939 | 25.8671 | 13.7714 | 8.5729
2019-04-09 15:28 | 147 | 0.30772 | 0.0090578 | 2.5493 | 27.4802 | 13.4545 | 8.4562
2019-04-09 15:28 | 148 | 0.30839 | 0.008894 | 2.5133 | 26.981 | 13.1505 | 8.3206
2019-04-09 15:28 | 149 | 0.3076 | 0.0087883 | 2.4943 | 26.5632 | 12.8377 | 8.1709
2019-04-09 15:28 | 150 | 0.29336 | 0.0090745 | 2.5858 | 27.3085 | 12.5381 | 8.0111
2019-04-09 15:28 | 151 | 0.30028 | 0.0087843 | 2.4981 | 26.3378 | 12.2464 | 7.9087
2019-04-09 15:28 | 152 | 0.2966 | 0.0088758 | 2.502 | 26.4214 | 11.9402 | 7.8365
2019-04-09 15:28 | 153 | 0.27278 | 0.0096305 | 2.6908 | 28.3398 | 11.6577 | 7.7306
2019-04-09 15:29 | 154 | 0.27972 | 0.0094056 | 2.5876 | 26.9815 | 11.3878 | 7.5473
2019-04-09 15:29 | 155 | 0.2755 | 0.0095575 | 2.5883 | 26.9045 | 11.1009 | 7.4123
2019-04-09 15:29 | 156 | 0.24611 | 0.010656 | 2.8559 | 29.6246 | 10.8424 | 7.291
2019-04-09 15:29 | 157 | 0.25238 | 0.01048 | 2.7439 | 28.3702 | 10.6001 | 7.16
2019-04-09 15:29 | 158 | 0.27042 | 0.0098777 | 2.5262 | 26.147 | 10.3346 | 7.0707
2019-04-09 15:29 | 159 | 0.25416 | 0.010661 | 2.6464 | 27.5236 | 10.0674 | 6.9954
2019-04-09 15:29 | 160 | 0.2422 | 0.011423 | 2.7367 | 28.8983 | 9.8257 | 6.9993
2019-04-09 15:29 | 161 | 0.24955 | 0.01124 | 2.6162 | 27.7046 | 9.5787 | 6.9138
2019-04-09 15:29 | 162 | 0.23175 | 0.012239 | 2.7698 | 28.932 | 9.332 | 6.7049
2019-04-09 15:29 | 163 | 0.22384 | 0.01278 | 2.8143 | 28.998 | 9.1127 | 6.4908
2019-04-09 15:29 | 164 | 0.24385 | 0.011952 | 2.5442 | 26.0331 | 8.878 | 6.3482
2019-04-09 15:29 | 165 | 0.23377 | 0.012594 | 2.6205 | 26.9818 | 8.6321 | 6.3074
2019-04-09 15:29 | 166 | 0.21078 | 0.01408 | 2.8615 | 29.5198 | 8.4136 | 6.2223
2019-04-09 15:29 | 167 | 0.21981 | 0.013693 | 2.702 | 27.9275 | 8.2032 | 6.1387
2019-04-09 15:29 | 168 | 0.21614 | 0.014033 | 2.7029 | 28.3032 | 7.9767 | 6.1175
2019-04-09 15:29 | 169 | 0.19523 | 0.015654 | 2.9427 | 31.0815 | 7.7746 | 6.0679
2019-04-09 15:29 | 170 | 0.20954 | 0.014782 | 2.6994 | 28.3224 | 7.5782 | 5.9347
2019-04-09 15:29 | 171 | 0.21112 | 0.014897 | 2.6449 | 27.0176 | 7.3579 | 5.7039
2019-04-09 15:29 | 172 | 0.20446 | 0.015787 | 2.6956 | 27.5243 | 7.1577 | 5.6276
2019-04-09 15:29 | 173 | 0.22374 | 0.014708 | 2.4286 | 24.5571 | 6.9437 | 5.4944
2019-04-09 15:29 | 174 | 0.21208 | 0.015988 | 2.5272 | 25.7041 | 6.7177 | 5.4514
2019-04-09 15:29 | 175 | 0.18462 | 0.018804 | 2.8617 | 29.2379 | 6.5226 | 5.3978
2019-04-09 15:29 | 176 | 0.2193 | 0.016193 | 2.3731 | 23.9478 | 6.3345 | 5.2517
2019-04-09 15:29 | 177 | 0.21324 | 0.016934 | 2.4016 | 24.0424 | 6.0937 | 5.1268
2019-04-09 15:29 | 178 | 0.16429 | 0.022388 | 3.0713 | 30.5063 | 5.9177 | 5.0119
2019-04-09 15:29 | 179 | 0.19458 | 0.019184 | 2.5552 | 25.6654 | 5.7472 | 4.9941
2019-04-09 15:29 | 180 | 0.1997 | 0.018871 | 2.44 | 24.8049 | 5.5338 | 4.9535
2019-04-09 15:29 | 181 | 0.16557 | 0.022921 | 2.8908 | 29.3567 | 5.3573 | 4.8606
2019-04-09 15:29 | 182 | 0.16551 | 0.022988 | 2.8297 | 28.5502 | 5.1955 | 4.7253
2019-04-09 15:29 | 183 | 0.15862 | 0.024198 | 2.9065 | 28.8205 | 5.0276 | 4.5715
2019-04-09 15:29 | 184 | 0.1313 | 0.029672 | 3.4462 | 33.3047 | 4.8822 | 4.373
2019-04-09 15:29 | 185 | 0.14945 | 0.026585 | 2.9423 | 28.2517 | 4.7543 | 4.2223
2019-04-09 15:29 | 186 | 0.15975 | 0.025672 | 2.6994 | 25.1747 | 4.5848 | 4.0217
2019-04-09 15:29 | 187 | 0.13705 | 0.030924 | 3.0499 | 28.8234 | 4.4422 | 3.9503
2019-04-09 15:29 | 188 | 0.15581 | 0.027911 | 2.6149 | 24.8231 | 4.3018 | 3.8676
2019-04-09 15:29 | 189 | 0.15804 | 0.02843 | 2.5089 | 24.4967 | 4.1346 | 3.8715
2019-04-09 15:29 | 190 | 0.12441 | 0 | 0 | 0 | 3.9943 | 3.8359
2019-04-09 15:29 | 191 | 0.11877 | 0 | 0 | 0 | 3.8791 | 3.7668
2019-04-09 15:29 | 192 | 0.13778 | 0 | 0 | 0 | 3.7507 | 3.6417
2019-04-09 15:29 | 193 | 0.12507 | 0 | 0 | 0 | 3.6105 | 3.4564
2019-04-09 15:29 | 194 | 0.12409 | 0 | 0 | 0 | 3.4984 | 3.2693
2019-04-09 15:29 | 195 | 0.12931 | 0 | 0 | 0 | 3.3619 | 3.1599
2019-04-09 15:29 | 196 | 0.08189 | 0 | 0 | 0 | 3.2513 | 3.0992
2019-04-09 15:29 | 197 | 0.088332 | 0 | 0 | 0 | 3.1856 | 2.9966
2019-04-09 15:29 | 198 | 0.11044 | 0 | 0 | 0 | 3.0702 | 2.9222
2019-04-09 15:29 | 199 | 0.094623 | 0 | 0 | 0 | 2.9732 | 2.841
2019-04-09 15:29 | 200 | 0.094983 | 0 | 0 | 0 | 2.8774 | 2.7239
2019-04-09 15:29 | 201 | 0.11604 | 0 | 0 | 0 | 2.7788 | 2.5969
2019-04-09 15:29 | 202 | 0.097438 | 0 | 0 | 0 | 2.6549 | 2.5352
2019-04-09 15:29 | 203 | 0.073394 | 0 | 0 | 0 | 2.5851 | 2.4849
2019-04-09 15:29 | 204 | 0.097796 | 0 | 0 | 0 | 2.4971 | 2.4015
2019-04-09 15:29 | 205 | 0.085041 | 0 | 0 | 0 | 2.3971 | 2.3066
2019-04-09 15:29 | 206 | 0.083974 | 0 | 0 | 0 | 2.3245 | 2.2902
2019-04-09 15:29 | 207 | 0.10382 | 0 | 0 | 0 | 2.2249 | 2.2353
2019-04-09 15:29 | 208 | 0.079271 | 0 | 0 | 0 | 2.1275 | 2.1454
2019-04-09 15:29 | 209 | 0.051506 | 0 | 0 | 0 | 2.067 | 2.1098
2019-04-09 15:29 | 210 | 0.076514 | 0 | 0 | 0 | 2.0121 | 2.0198
2019-04-09 15:29 | 211 | 0.073995 | 0 | 0 | 0 | 1.9197 | 1.9464
2019-04-09 15:29 | 212 | 0.054235 | 0 | 0 | 0 | 1.8679 | 1.8775
2019-04-09 15:30 | 213 | 0.079846 | 0 | 0 | 0 | 1.8012 | 1.8339
2019-04-09 15:30 | 214 | 0.076704 | 0 | 0 | 0 | 1.7145 | 1.795
2019-04-09 15:30 | 215 | 0.059241 | 0 | 0 | 0 | 1.651 | 1.7629
2019-04-09 15:30 | 216 | 0.083478 | 0 | 0 | 0 | 1.5871 | 1.695
2019-04-09 15:30 | 217 | 0.070628 | 0 | 0 | 0 | 1.4925 | 1.6149
2019-04-09 15:30 | 218 | 0.046009 | 0 | 0 | 0 | 1.4483 | 1.603
2019-04-09 15:30 | 219 | 0.075146 | 0 | 0 | 0 | 1.3884 | 1.5536
2019-04-09 15:30 | 220 | 0.067052 | 0 | 0 | 0 | 1.3061 | 1.5234
2019-04-09 15:30 | 221 | 0.049491 | 0 | 0 | 0 | 1.2567 | 1.4886
2019-04-09 15:30 | 222 | 0.067654 | 0 | 0 | 0 | 1.1992 | 1.4551
2019-04-09 15:30 | 223 | 0.061722 | 0 | 0 | 0 | 1.127 | 1.3972
2019-04-09 15:30 | 224 | 0.034323 | 0 | 0 | 0 | 1.0803 | 1.3544
2019-04-09 15:30 | 225 | 0.061596 | 0 | 0 | 0 | 1.0464 | 1.3202
2019-04-09 15:30 | 226 | 0.049813 | 0 | 0 | 0 | 0.96597 | 1.2851
2019-04-09 15:30 | 227 | 0.014217 | 0 | 0 | 0 | 0.95148 | 1.2548
2019-04-09 15:30 | 228 | 0.059311 | 0 | 0 | 0 | 0.91903 | 1.2342
2019-04-09 15:30 | 229 | 0.06426 | 0 | 0 | 0 | 0.84144 | 1.2004
2019-04-09 15:30 | 230 | 0.043271 | 0 | 0 | 0 | 0.79739 | 1.213
2019-04-09 15:30 | 231 | 0.049544 | 0 | 0 | 0 | 0.74951 | 1.1504
2019-04-09 15:30 | 232 | 0.034426 | 0 | 0 | 0 | 0.70288 | 1.1454
2019-04-09 15:30 | 233 | 0.020202 | 0 | 0 | 0 | 0.67986 | 1.0916
2019-04-09 15:30 | 234 | 0.045313 | 0 | 0 | 0 | 0.65328 | 1.0643
2019-04-09 15:30 | 235 | 0.045854 | 0 | 0 | 0 | 0.5951 | 1.0349
2019-04-09 15:30 | 236 | 0.014278 | 0 | 0 | 0 | 0.56858 | 1.0039
2019-04-09 15:30 | 237 | 0.045768 | 0 | 0 | 0 | 0.55251 | 0.99021
2019-04-09 15:30 | 238 | 0.048602 | 0 | 0 | 0 | 0.4839 | 0.96126
2019-04-09 15:30 | 239 | 0.0019454 | 0 | 0 | 0 | 0.46687 | 0.95018
2019-04-09 15:30 | 240 | 0.012096 | 0 | 0 | 0 | 0.46668 | 0.92355
2019-04-09 15:30 | 241 | 0.036471 | 0 | 0 | 0 | 0.43816 | 0.8907
2019-04-09 15:30 | 242 | 0.035405 | 0 | 0 | 0 | 0.39935 | 0.85989
2019-04-09 15:30 | 243 | 0.045671 | 0 | 0 | 0 | 0.36584 | 0.81441
2019-04-09 15:30 | 244 | 0.036745 | 0 | 0 | 0 | 0.313 | 0.78652
2019-04-09 15:30 | 245 | 0.011363 | 0 | 0 | 0 | 0.29542 | 0.77524
2019-04-09 15:30 | 246 | 0.044265 | 0 | 0 | 0 | 0.27693 | 0.73345
2019-04-09 15:30 | 247 | 0.046902 | 0 | 0 | 0 | 0.21362 | 0.73924
2019-04-09 15:30 | 248 | 0.021086 | 0 | 0 | 0 | 0.18974 | 0.7184
2019-04-09 15:30 | 249 | 0.042851 | 0 | 0 | 0 | 0.16072 | 0.70359
2019-04-09 15:30 | 250 | 0.0542 | 0 | 0 | 0 | 0.10674 | 0.66322
2019-04-09 15:30 | 251 | 0.017367 | 0 | 0 | 0 | 0.063885 | 0.64469
2019-04-09 15:30 | 252 | 0.0094947 | 0 | 0 | 0 | 0.065475 | 0.62588
2019-04-09 15:30 | 253 | 0.025351 | 0 | 0 | 0 | 0.038986 | 0.61251
2019-04-09 15:30 | 254 | 0.007882 | 0 | 0 | 0 | 0.021487 | 0.60687
2019-04-09 15:30 | 255 | 0.021822 | 0 | 0 | 0 | 0.016197 | 0.5713
2019-04-09 15:30 | 256 | 0.03622 | 0 | 0 | 0 | -0.022253 | 0.5432
2019-04-09 15:30 | 257 | 0.021813 | 0 | 0 | 0 | -0.049068 | 0.51557
2019-04-09 15:30 | 258 | 0.0092302 | 0 | 0 | 0 | -0.065992 | 0.50586
2019-04-09 15:30 | 259 | 0.016043 | 0 | 0 | 0 | -0.071953 | 0.49501
2019-04-09 15:30 | 260 | 0.0041661 | 0 | 0 | 0 | -0.094571 | 0.47965
2019-04-09 15:30 | 261 | 0.010163 | 0 | 0 | 0 | -0.085179 | 0.48383
2019-04-09 15:30 | 262 | 0.047571 | 0 | 0 | 0 | -0.12181 | 0.46813
2019-04-09 15:30 | 263 | 0.031579 | 0 | 0 | 0 | -0.16788 | 0.4601
2019-04-09 15:30 | 264 | 0.022234 | 0 | 0 | 0 | -0.18536 | 0.45188
2019-04-09 15:30 | 265 | 0.023707 | 0 | 0 | 0 | -0.21469 | 0.42426
2019-04-09 15:30 | 266 | 0.0095695 | 0 | 0 | 0 | -0.22945 | 0.40261
2019-04-09 15:30 | 267 | 0.0032117 | 0 | 0 | 0 | -0.23608 | 0.40192
2019-04-09 15:30 | 268 | 0.020742 | 0 | 0 | 0 | -0.24166 | 0.3832
2019-04-09 15:30 | 269 | 0.021622 | 0 | 0 | 0 | -0.27382 | 0.39641
2019-04-09 15:30 | 270 | 0.012342 | 0 | 0 | 0 | -0.28247 | 0.37012

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062463 | 0.062684 | 0 | 0.2095 | 2019-04-09 11:25 | Group A | 0.1 | 0.22 | Do Nothing | 0.20953 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025185 | 0.025304 | 0 | 0.0261 | 2019-04-09 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044991 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002722 | 0.002734 | 0 | 0.0011 | 2019-04-09 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 1.4077e-06 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.5996 | 183.5796 | 0 | 100 | 2019-04-09 11:25 | None | 0 | 1000 | Do Nothing | 100.8778 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.5348 | 212.1604 | 0 | 100 | 2019-04-09 11:25 | None | 0 | 1000 | Do Nothing | 101.6174 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -14.6891 | 1447 | 0 |  | 2019-04-09 11:13 | None | -10 | 1000 | Do Nothing | 122.2378 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 20.6914 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001903 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017872 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -154.5287 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23226 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4017 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0027095 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0040191 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0038575 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 4.2063e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.001319 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.28738 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3591 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5476 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 98.46 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.017 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 48.3234 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 67.062 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 86.4472 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.2763 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.011499 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0463 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.024269 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.1951 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN