## Cone Summary
Date: | 2019-04-09
Time: | 14:27
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B6S-Test3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\04_Apr\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0403 |  | Start Test | 0 | 2019-04-09 14:35
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 7 | 2019-04-09 14:35
Sample Area | 1 | m^2 | Wetting | 224 | 2019-04-09 14:38
Cone Heat Flux | 75 | kW/m^2 | Dripping | 300 | 2019-04-09 14:40
Cone Orientation |  |  | Bottom Ignition | 641 | 2019-04-09 14:45
Sample Grid |  |  | Flame Out | 991 | 2019-04-09 14:51
Sample Thickness | 106 | mm
Sample Initial Mass | 52.1 | g
Cone Separation Distance | 25 | mm
Exhaust Flow Rate | 23.1932 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.3441 | kW/m2 | Average Heat Release Rate | 0.86742 | kW/m2
Total Heat Release | 0.96141 | MJ/m2 | Average Mass Loss Rate | 0.042621 | g/m2s
Peak Mass Loss Rate | 0.19091 | g/m2s | Effective Heat of Combustion | 20.4107 | MJ/kg
Total Mass Lost | 44.8889 | g | Average Specific Extinction Area | 45.7884 | m2/kg
Time to Ignition | 7 | s | Peak CO Yield | 6.1806 | kg/kg
Time to Peak Heat Release | 677 | s | Peak CO2 Yield | 118.1848 | kg/kg
Time to Flame Out | 991 | s | Average CO Yield | 0.056145 | kg/kg
Average Heat Release Rate | 0.97605 | kW/m2 | Average CO2 Yield | 2.222 | kg/kg
Average Mass Loss Rate | 0.044782 | g/m2s
Effective Heat of Combustion | 21.4176 | MJ/kg
Dripping | 300 | s | Time to dripping
HRRD | 0.44263 | kW | HRR at dripping time
THRD | 0.3281 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.1198 | kW | AHRR from Ignition to Dripping
Wetting | 224 | s | Time to wetting
HRRW | 0.57174 | kW | HRR at wetting time
THRW | 0.29105 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.305 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-04-09 14:27 | 1554834427.525 | -454 | 0.20761 | 0.000497 | 6e-06 | 91.4485 | 90.7877 | 123.7586 | 20.4093 | 0.001758 | 0.001677 | -154.0859 | 0.22891 | 8.4001 | 0.002502 | 0.003969 | 0.003781 | 4.7e-05 | 0.001115 | 0.26698 | 1.3435 | 2.5473 | 98.56 | 20.01 | 50.16 | 63.2961 | 81.1875 | 61.3455 | 0.011919 | 1.0555 | 0.024513 | 23.2251 | 0.064868 | 0.13618 | -7.3e-05 | -0.005682 | -0.005682 | -0.96577
2019-04-09 14:27 | 1554834428.5264 | -453 | 0.20761 | 0.000497 | 6e-06 | 91.4485 | 90.7877 | 123.7586 | 20.4093 | 0.001758 | 0.001677 | -154.0859 | 0.22891 | 8.4001 | 0.002502 | 0.003969 | 0.003781 | 4.7e-05 | 0.001115 | 0.26698 | 1.3435 | 2.5473 | 98.56 | 20.01 | 50.159 | 63.2961 | 81.1875 | 61.3455 | 0.011918 | 1.0555 | 0.024513 | 23.2251 | 0.064868 | 0.13618 | -7.3e-05 | -0.005682 | -0.005682 | -0.96577
2019-04-09 14:27 | 1554834429.525 | -452 | 0.20758 | 0.000501 | 6e-06 | 91.4339 | 90.6239 | 123.7586 | 20.4103 | 0.001809 | 0.001675 | -154.0859 | 0.22862 | 8.4034 | 0.002555 | 0.004001 | 0.003828 | 5.3e-05 | 0.001119 | 0.27544 | 1.3638 | 2.5475 | 98.56 | 20.01 | 50.158 | 64.5177 | 82.4472 | 61.3081 | 0.011918 | 1.0556 | 0.024514 | 23.2238 | 0.079595 | 0.136 | -7.9e-05 | -0.006072 | -0.006072 | -0.90261
2019-04-09 14:27 | 1554834430.5252 | -451 | 0.20761 | 0.000497 | 5e-06 | 91.4432 | 90.6318 | 123.7586 | 20.4098 | 0.00183 | 0.001684 | -154.0859 | 0.22947 | 8.4014 | 0.002585 | 0.003675 | 0.003489 | 4.9e-05 | 0.001126 | 0.25841 | 1.3426 | 2.5475 | 98.56 | 20.01 | 50.155 | 65.0362 | 83.1822 | 61.5218 | 0.011917 | 1.0549 | 0.024507 | 23.2312 | 0.079717 | 0.13651 | -7.3e-05 | -0.005609 | -0.005609 | -0.98076
2019-04-09 14:27 | 1554834431.5262 | -450 | 0.2076 | 0.000495 | 5e-06 | 91.4491 | 90.6019 | 123.7586 | 20.4103 | 0.001805 | 0.001746 | -154.0859 | 0.2293 | 8.4022 | 0.002497 | 0.00406 | 0.003859 | 5.1e-05 | 0.001128 | 0.26056 | 1.3865 | 2.5476 | 98.56 | 20.01 | 50.151 | 64.4339 | 81.0536 | 63.019 | 0.011917 | 1.0502 | 0.024452 | 23.2831 | 0.083254 | 0.13641 | -7e-05 | -0.005386 | -0.005386 | -1.0206
2019-04-09 14:27 | 1554834432.5262 | -449 | 0.20759 | 0.000499 | 6e-06 | 91.4625 | 90.6287 | 123.7586 | 20.4106 | 0.001783 | 0.001693 | -154.0859 | 0.22851 | 8.4025 | 0.002527 | 0.003611 | 0.003488 | 4.8e-05 | 0.001135 | 0.27328 | 1.3512 | 2.5477 | 98.56 | 20.01 | 50.14 | 63.8996 | 81.7916 | 61.7475 | 0.011914 | 1.0542 | 0.024498 | 23.239 | 0.081916 | 0.13594 | -7.6e-05 | -0.005849 | -0.005849 | -0.93659
2019-04-09 14:27 | 1554834433.5255 | -448 | 0.20765 | 0.000495 | 6e-06 | 91.428 | 90.6001 | 123.7586 | 20.4096 | 0.001778 | 0.001643 | -154.0859 | 0.22915 | 8.4017 | 0.002492 | 0.004304 | 0.00411 | 4.9e-05 | 0.00114 | 0.26897 | 1.3225 | 2.5477 | 98.56 | 20.01 | 50.132 | 63.7907 | 80.9531 | 60.528 | 0.011912 | 1.058 | 0.024543 | 23.1967 | 0.081356 | 0.13632 | -7.1e-05 | -0.005525 | -0.005525 | -0.99428
2019-04-09 14:27 | 1554834434.5252 | -447 | 0.20753 | 0.000496 | 5e-06 | 91.3455 | 90.5902 | 123.7586 | 20.4103 | 0.001839 | 0.001704 | -154.0859 | 0.22876 | 8.3986 | 0.002483 | 0.003336 | 0.003186 | 5e-05 | 0.001142 | 0.29679 | 1.3335 | 2.5475 | 98.56 | 20.01 | 50.117 | 65.2397 | 80.7291 | 62.0068 | 0.011908 | 1.0534 | 0.024489 | 23.248 | 0.074268 | 0.13609 | -7.2e-05 | -0.005531 | -0.005531 | -0.99151
2019-04-09 14:27 | 1554834435.5243 | -446 | 0.20759 | 0.000499 | 5e-06 | 91.3916 | 90.6026 | 123.7586 | 20.4108 | 0.001809 | 0.00164 | -154.0859 | 0.22862 | 8.4 | 0.002486 | 0.004055 | 0.00381 | 4.8e-05 | 0.001141 | 0.27671 | 1.3294 | 2.5475 | 98.56 | 20.01 | 50.111 | 64.5307 | 80.7952 | 60.4512 | 0.011907 | 1.0583 | 0.024546 | 23.194 | 0.077563 | 0.136 | -7.6e-05 | -0.005822 | -0.005822 | -0.94135
2019-04-09 14:27 | 1554834436.5257 | -445 | 0.2076 | 0.000498 | 5e-06 | 91.2969 | 90.2997 | 123.7586 | 20.4112 | 0.001768 | 0.001656 | -154.0859 | 0.22945 | 8.4012 | 0.002503 | 0.004273 | 0.00403 | 5.3e-05 | 0.001138 | 0.26394 | 1.3815 | 2.5476 | 98.56 | 20.01 | 50.1 | 63.5503 | 81.2175 | 60.8542 | 0.011904 | 1.057 | 0.024531 | 23.208 | 0.098239 | 0.1365 | -7.4e-05 | -0.005725 | -0.005725 | -0.96094
2019-04-09 14:27 | 1554834437.5253 | -444 | 0.2076 | 0.0005 | 5e-06 | 91.2797 | 90.3902 | 123.7586 | 20.4112 | 0.001766 | 0.001693 | -154.0859 | 0.22831 | 8.4036 | 0.002518 | 0.004443 | 0.00428 | 4.8e-05 | 0.001144 | 0.27992 | 1.3552 | 2.5476 | 98.56 | 20.01 | 50.096 | 63.4941 | 81.5636 | 61.7295 | 0.011903 | 1.0542 | 0.024499 | 23.2384 | 0.087586 | 0.13583 | -7.6e-05 | -0.005851 | -0.005851 | -0.93549
2019-04-09 14:27 | 1554834438.5258 | -443 | 0.20762 | 0.000498 | 5e-06 | 91.2627 | 90.5021 | 123.8408 | 20.4106 | 0.001839 | 0.00172 | -154.0856 | 0.22807 | 8.402 | 0.002544 | 0.003436 | 0.003264 | 4.9e-05 | 0.001145 | 0.2688 | 1.3143 | 2.5474 | 98.56 | 20.01 | 50.092 | 65.2391 | 82.2002 | 62.3922 | 0.011903 | 1.0522 | 0.024483 | 23.2691 | 0.074861 | 0.13568 | -7.4e-05 | -0.005713 | -0.005713 | -0.95709
2019-04-09 14:27 | 1554834439.5245 | -442 | 0.20759 | 0.000504 | 6e-06 | 91.246 | 90.4339 | 123.7535 | 20.4119 | 0.001774 | 0.001729 | -154.0761 | 0.2293 | 8.4057 | 0.002531 | 0.004326 | 0.004171 | 4.8e-05 | 0.001146 | 0.27455 | 1.3633 | 2.5476 | 98.56 | 20.01 | 50.088 | 63.69 | 81.8832 | 62.606 | 0.011902 | 1.0515 | 0.024467 | 23.2683 | 0.079964 | 0.13641 | -8.2e-05 | -0.006325 | -0.006325 | -0.86914
2019-04-09 14:27 | 1554834440.5244 | -441 | 0.2076 | 0.000502 | 6e-06 | 91.2189 | 90.3916 | 123.9441 | 20.4118 | 0.001783 | 0.001669 | -154.0845 | 0.22938 | 8.4018 | 0.00252 | 0.004129 | 0.003928 | 4.9e-05 | 0.001143 | 0.27704 | 1.344 | 2.5476 | 98.56 | 20.01 | 50.085 | 63.9147 | 81.622 | 61.1525 | 0.011901 | 1.0561 | 0.024539 | 23.2358 | 0.081495 | 0.13646 | -8e-05 | -0.006178 | -0.006178 | -0.89014
2019-04-09 14:27 | 1554834441.5255 | -440 | 0.20761 | 0.000498 | 5e-06 | 91.2684 | 90.4126 | 123.8692 | 20.4119 | 0.001777 | 0.001732 | -154.0644 | 0.22847 | 8.4003 | 0.002507 | 0.003702 | 0.003434 | 5.1e-05 | 0.001151 | 0.2688 | 1.339 | 2.5474 | 98.56 | 20.01 | 50.082 | 63.768 | 81.3032 | 62.6816 | 0.0119 | 1.0513 | 0.024475 | 23.2818 | 0.084264 | 0.13592 | -7.4e-05 | -0.005706 | -0.005706 | -0.95995
2019-04-09 14:27 | 1554834442.5245 | -439 | 0.20763 | 0.000499 | 5e-06 | 91.2097 | 90.5358 | 123.7457 | 20.4117 | 0.001803 | 0.001707 | -154.0766 | 0.22837 | 8.4025 | 0.002488 | 0.004046 | 0.003923 | 5e-05 | 0.001152 | 0.26017 | 1.3328 | 2.5476 | 98.56 | 20.01 | 50.08 | 64.3956 | 80.86 | 62.0712 | 0.0119 | 1.0532 | 0.024485 | 23.2491 | 0.066332 | 0.13586 | -7.5e-05 | -0.005799 | -0.005799 | -0.94414
2019-04-09 14:27 | 1554834443.5245 | -438 | 0.20759 | 0.0005 | 6e-06 | 91.2025 | 90.2474 | 123.5406 | 20.4116 | 0.001804 | 0.001704 | -154.0777 | 0.22926 | 8.4013 | 0.00246 | 0.003809 | 0.003612 | 5e-05 | 0.001154 | 0.2844 | 1.3785 | 2.5475 | 98.56 | 20.01 | 50.079 | 64.4001 | 80.1659 | 62.0134 | 0.011899 | 1.0534 | 0.024467 | 23.2278 | 0.094169 | 0.13639 | -7.7e-05 | -0.005959 | -0.005959 | -0.92238
2019-04-09 14:27 | 1554834444.525 | -437 | 0.2076 | 0.0005 | 6e-06 | 91.2005 | 90.3992 | 123.5659 | 20.4125 | 0.001779 | 0.001677 | -154.0876 | 0.22847 | 8.4019 | 0.002492 | 0.003706 | 0.003625 | 4.7e-05 | 0.001145 | 0.26969 | 1.3879 | 2.5475 | 98.56 | 20.01 | 50.077 | 63.8069 | 80.9551 | 61.3611 | 0.011899 | 1.0554 | 0.024493 | 23.2075 | 0.078935 | 0.13592 | -7.8e-05 | -0.005983 | -0.005983 | -0.91557
2019-04-09 14:27 | 1554834445.5256 | -436 | 0.2076 | 0.000495 | 6e-06 | 91.2055 | 90.5544 | 123.6798 | 20.4118 | 0.001806 | 0.001687 | -154.071 | 0.22887 | 8.4002 | 0.002481 | 0.00409 | 0.003864 | 4.6e-05 | 0.001143 | 0.26737 | 1.3702 | 2.5476 | 98.56 | 20.01 | 50.078 | 64.4465 | 80.6823 | 61.5958 | 0.011899 | 1.0547 | 0.024496 | 23.2264 | 0.064087 | 0.13615 | -7.1e-05 | -0.005505 | -0.005505 | -0.99675
2019-04-09 14:27 | 1554834446.5251 | -435 | 0.2076 | 0.000498 | 5e-06 | 91.2061 | 90.2612 | 123.8359 | 20.4124 | 0.001793 | 0.001661 | -154.0884 | 0.22952 | 8.4019 | 0.002485 | 0.00393 | 0.003694 | 4.9e-05 | 0.00114 | 0.27632 | 1.3571 | 2.5476 | 98.56 | 20.01 | 50.078 | 64.1386 | 80.7718 | 60.9693 | 0.011899 | 1.0566 | 0.024535 | 23.2193 | 0.093155 | 0.13654 | -7.4e-05 | -0.005729 | -0.005729 | -0.96045
2019-04-09 14:27 | 1554834447.5247 | -434 | 0.20758 | 0.000498 | 5e-06 | 91.2692 | 90.4005 | 123.718 | 20.4129 | 0.001805 | 0.001655 | -154.0878 | 0.22825 | 8.403 | 0.002507 | 0.003562 | 0.003454 | 5e-05 | 0.001144 | 0.27328 | 1.3413 | 2.5475 | 98.56 | 20 | 50.078 | 64.4437 | 81.3104 | 60.8127 | 0.011892 | 1.0571 | 0.024529 | 23.2028 | 0.085543 | 0.13579 | -7.3e-05 | -0.005649 | -0.005649 | -0.96863
2019-04-09 14:27 | 1554834448.5247 | -433 | 0.20764 | 0.0005 | 6e-06 | 91.3169 | 90.4298 | 123.6467 | 20.4137 | 0.001821 | 0.001727 | -154.0737 | 0.2299 | 8.4016 | 0.002541 | 0.004869 | 0.004623 | 4.9e-05 | 0.001147 | 0.27544 | 1.3186 | 2.5475 | 98.56 | 20 | 50.077 | 64.8167 | 82.1214 | 62.5562 | 0.011892 | 1.0516 | 0.024458 | 23.2566 | 0.087317 | 0.13677 | -7.7e-05 | -0.005946 | -0.005946 | -0.927
2019-04-09 14:27 | 1554834449.5247 | -432 | 0.20759 | 0.000499 | 5e-06 | 91.3208 | 90.4587 | 123.4688 | 20.414 | 0.001746 | 0.001661 | -154.0819 | 0.22866 | 8.4002 | 0.00246 | 0.004032 | 0.003856 | 4.8e-05 | 0.001152 | 0.26322 | 1.3346 | 2.5475 | 98.56 | 20 | 50.079 | 63.0243 | 80.1667 | 60.9617 | 0.011892 | 1.0567 | 0.024498 | 23.1846 | 0.084837 | 0.13603 | -7.5e-05 | -0.005755 | -0.005755 | -0.9526
2019-04-09 14:27 | 1554834450.5247 | -431 | 0.20759 | 0.0005 | 5e-06 | 91.2955 | 90.3727 | 123.0575 | 20.4137 | 0.001701 | 0.001689 | -154.0586 | 0.2287 | 8.4018 | 0.002579 | 0.0043 | 0.004123 | 4.8e-05 | 0.001154 | 0.2808 | 1.3264 | 2.5476 | 98.56 | 20 | 50.078 | 61.9429 | 83.0316 | 61.6512 | 0.011892 | 1.0545 | 0.024432 | 23.1698 | 0.090869 | 0.13605 | -7.7e-05 | -0.005892 | -0.005892 | -0.93064
2019-04-09 14:28 | 1554834451.5251 | -430 | 0.20762 | 0.000499 | 5e-06 | 91.2752 | 90.7278 | 123.3462 | 20.414 | 0.001807 | 0.001671 | -154.0747 | 0.22876 | 8.4005 | 0.002507 | 0.003991 | 0.003814 | 4.9e-05 | 0.001156 | 0.27256 | 1.3684 | 2.5476 | 98.56 | 20 | 50.077 | 64.49 | 81.2973 | 61.217 | 0.011892 | 1.0559 | 0.024477 | 23.1819 | 0.053802 | 0.13609 | -7.6e-05 | -0.005817 | -0.005817 | -0.94283
2019-04-09 14:28 | 1554834452.5248 | -429 | 0.20759 | 0.0005 | 5e-06 | 91.302 | 90.4057 | 123.3817 | 20.4135 | 0.001811 | 0.00164 | -154.0589 | 0.22925 | 8.4019 | 0.002489 | 0.003864 | 0.003663 | 4.5e-05 | 0.001158 | 0.26952 | 1.3625 | 2.5476 | 98.56 | 20 | 50.078 | 64.5879 | 80.882 | 60.4663 | 0.011892 | 1.0582 | 0.024508 | 23.1592 | 0.088242 | 0.13638 | -7.7e-05 | -0.005938 | -0.005938 | -0.92564
2019-04-09 14:28 | 1554834453.5238 | -428 | 0.20757 | 0.000503 | 6e-06 | 91.3693 | 90.4085 | 123.7793 | 20.4138 | 0.001756 | 0.001715 | -154.0484 | 0.22817 | 8.3999 | 0.002579 | 0.003436 | 0.003241 | 4.9e-05 | 0.001161 | 0.26936 | 1.3951 | 2.5477 | 98.56 | 20 | 50.077 | 63.2601 | 83.0356 | 62.278 | 0.011892 | 1.0525 | 0.024481 | 23.2594 | 0.09456 | 0.13574 | -8e-05 | -0.006182 | -0.006182 | -0.8848
2019-04-09 14:28 | 1554834454.5248 | -427 | 0.20755 | 0.0005 | 6e-06 | 91.4735 | 90.5998 | 123.8485 | 20.4138 | 0.001851 | 0.001778 | -154.036 | 0.2288 | 8.4005 | 0.002514 | 0.003463 | 0.003312 | 4.9e-05 | 0.001156 | 0.25818 | 1.3309 | 2.5475 | 98.56 | 20 | 50.079 | 65.5498 | 81.4773 | 63.7854 | 0.011892 | 1.0478 | 0.024433 | 23.3181 | 0.085847 | 0.13612 | -7.6e-05 | -0.005869 | -0.005869 | -0.93463
2019-04-09 14:28 | 1554834455.5251 | -426 | 0.20754 | 0.0005 | 5e-06 | 91.5328 | 90.7519 | 123.7906 | 20.4143 | 0.001824 | 0.001705 | -154.0427 | 0.22954 | 8.4 | 0.002516 | 0.003915 | 0.003642 | 4.9e-05 | 0.001159 | 0.27688 | 1.3373 | 2.5477 | 98.56 | 20 | 50.084 | 64.8915 | 81.5339 | 62.0293 | 0.011893 | 1.0533 | 0.024491 | 23.2518 | 0.076634 | 0.13655 | -7.6e-05 | -0.005876 | -0.005876 | -0.93656
2019-04-09 14:28 | 1554834456.525 | -425 | 0.20754 | 0.000502 | 5e-06 | 91.5751 | 90.6012 | 123.7976 | 20.4141 | 0.001788 | 0.001652 | -154.0315 | 0.22865 | 8.4033 | 0.002507 | 0.003743 | 0.003647 | 4.7e-05 | 0.001153 | 0.26593 | 1.3115 | 2.5476 | 98.56 | 20 | 50.086 | 64.0372 | 81.3123 | 60.7522 | 0.011894 | 1.0573 | 0.024539 | 23.2081 | 0.095635 | 0.13603 | -7.9e-05 | -0.006103 | -0.006103 | -0.89819
2019-04-09 14:28 | 1554834457.525 | -424 | 0.20761 | 0.000498 | 5e-06 | 91.5775 | 90.7726 | 123.8239 | 20.4141 | 0.001784 | 0.001624 | -154.0254 | 0.2293 | 8.4029 | 0.002526 | 0.003688 | 0.003435 | 4.7e-05 | 0.001153 | 0.28351 | 1.3506 | 2.5475 | 98.56 | 20 | 50.09 | 63.9319 | 81.763 | 60.0877 | 0.011895 | 1.0594 | 0.024566 | 23.1875 | 0.078963 | 0.13641 | -7.3e-05 | -0.005639 | -0.005639 | -0.97487
2019-04-09 14:28 | 1554834458.524 | -423 | 0.20763 | 0.000499 | 6e-06 | 91.6489 | 90.882 | 123.8502 | 20.4157 | 0.001776 | 0.001637 | -154.028 | 0.22869 | 8.4022 | 0.002512 | 0.003918 | 0.003753 | 4.7e-05 | 0.001156 | 0.26593 | 1.3816 | 2.5475 | 98.56 | 20 | 50.093 | 63.7382 | 81.4191 | 60.3824 | 0.011895 | 1.0585 | 0.024558 | 23.2002 | 0.075163 | 0.13605 | -7.7e-05 | -0.006003 | -0.006003 | -0.91339
2019-04-09 14:28 | 1554834459.5251 | -422 | 0.2076 | 0.000501 | 6e-06 | 91.7823 | 91.0159 | 124.1178 | 20.4154 | 0.001833 | 0.00168 | -154.0231 | 0.22915 | 8.4007 | 0.002447 | 0.003222 | 0.00311 | 5.1e-05 | 0.001159 | 0.29247 | 1.3635 | 2.5474 | 98.56 | 20 | 50.095 | 65.1004 | 79.8656 | 61.4256 | 0.011896 | 1.0552 | 0.024546 | 23.2615 | 0.075008 | 0.13632 | -7.8e-05 | -0.006049 | -0.006049 | -0.90818
2019-04-09 14:28 | 1554834460.5256 | -421 | 0.2076 | 0.000495 | 6e-06 | 91.8574 | 91.1656 | 124.2854 | 20.4155 | 0.001806 | 0.00171 | -154.0261 | 0.22876 | 8.4018 | 0.002432 | 0.004084 | 0.00387 | 4.7e-05 | 0.001157 | 0.27312 | 1.3611 | 2.5476 | 98.56 | 20 | 50.097 | 64.4619 | 79.5124 | 62.1626 | 0.011896 | 1.0529 | 0.024535 | 23.3029 | 0.067618 | 0.13609 | -7e-05 | -0.005437 | -0.005437 | -1.0087
2019-04-09 14:28 | 1554834461.5245 | -420 | 0.20759 | 0.000499 | 5e-06 | 91.9554 | 91.3377 | 124.0614 | 20.4162 | 0.001745 | 0.001699 | -154.0119 | 0.22856 | 8.4023 | 0.002501 | 0.003852 | 0.003648 | 4.7e-05 | 0.00115 | 0.28567 | 1.3629 | 2.5475 | 98.56 | 20 | 50.098 | 62.9968 | 81.1646 | 61.8844 | 0.011897 | 1.0538 | 0.024523 | 23.2722 | 0.060286 | 0.13597 | -7.6e-05 | -0.005846 | -0.005846 | -0.93733
2019-04-09 14:28 | 1554834462.5245 | -419 | 0.2076 | 0.000499 | 5e-06 | 92.0298 | 91.3356 | 123.9415 | 20.4157 | 0.001796 | 0.001697 | -154.0246 | 0.22897 | 8.4008 | 0.002457 | 0.003769 | 0.00363 | 4.9e-05 | 0.001151 | 0.25802 | 1.3511 | 2.5476 | 98.56 | 20 | 50.102 | 64.22 | 80.0966 | 61.83 | 0.011898 | 1.0539 | 0.024513 | 23.2591 | 0.067727 | 0.13621 | -7.5e-05 | -0.005769 | -0.005769 | -0.95157
2019-04-09 14:28 | 1554834463.5232 | -418 | 0.20761 | 0.0005 | 6e-06 | 92.0775 | 91.1866 | 123.6814 | 20.4154 | 0.001829 | 0.001745 | -154.0367 | 0.22871 | 8.4028 | 0.002553 | 0.003692 | 0.003491 | 5e-05 | 0.001143 | 0.24275 | 1.3769 | 2.5476 | 98.56 | 20 | 50.103 | 65.0113 | 82.4241 | 62.9818 | 0.011898 | 1.0503 | 0.024446 | 23.2746 | 0.086962 | 0.13606 | -7.8e-05 | -0.005975 | -0.005975 | -0.91767
2019-04-09 14:28 | 1554834464.5236 | -417 | 0.2076 | 0.0005 | 6e-06 | 92.2457 | 91.5572 | 123.6326 | 20.4159 | 0.001843 | 0.001714 | -154.028 | 0.22944 | 8.3998 | 0.002597 | 0.003441 | 0.003246 | 5.1e-05 | 0.001138 | 0.29214 | 1.3378 | 2.5476 | 98.56 | 20 | 50.105 | 65.3551 | 83.4637 | 62.2357 | 0.011898 | 1.0527 | 0.024468 | 23.2441 | 0.067007 | 0.13649 | -7.7e-05 | -0.005967 | -0.005967 | -0.92189
2019-04-09 14:28 | 1554834465.5242 | -416 | 0.20761 | 0.0005 | 6e-06 | 92.2999 | 91.5101 | 123.8222 | 20.417 | 0.001737 | 0.001693 | -154.0239 | 0.22763 | 8.4019 | 0.002608 | 0.004084 | 0.003871 | 4.7e-05 | 0.001137 | 0.28655 | 1.3305 | 2.5475 | 98.56 | 20 | 50.111 | 62.792 | 83.7314 | 61.7491 | 0.0119 | 1.0542 | 0.024505 | 23.2451 | 0.076868 | 0.13542 | -7.7e-05 | -0.00594 | -0.00594 | -0.9188
2019-04-09 14:28 | 1554834466.5243 | -415 | 0.20763 | 0.000498 | 5e-06 | 92.425 | 91.7844 | 123.7232 | 20.4162 | 0.001772 | 0.001729 | -154.0102 | 0.22889 | 8.4035 | 0.002501 | 0.004106 | 0.003892 | 5.1e-05 | 0.001138 | 0.26128 | 1.3902 | 2.5476 | 98.56 | 20 | 50.113 | 63.6458 | 81.1778 | 62.6109 | 0.0119 | 1.0515 | 0.024463 | 23.2656 | 0.062211 | 0.13617 | -7.5e-05 | -0.005742 | -0.005742 | -0.95569
2019-04-09 14:28 | 1554834467.5233 | -414 | 0.20761 | 0.000497 | 5e-06 | 92.5155 | 91.8587 | 123.5918 | 20.4162 | 0.001834 | 0.00168 | -154.0203 | 0.22942 | 8.4016 | 0.002536 | 0.004187 | 0.003961 | 5e-05 | 0.00114 | 0.27383 | 1.3916 | 2.5476 | 98.56 | 20 | 50.114 | 65.1464 | 82.003 | 61.4389 | 0.0119 | 1.0552 | 0.024493 | 23.2127 | 0.063726 | 0.13648 | -7.2e-05 | -0.005542 | -0.005542 | -0.99246
2019-04-09 14:28 | 1554834468.5243 | -413 | 0.20759 | 0.0005 | 6e-06 | 92.5923 | 91.995 | 123.6244 | 20.4163 | 0.001785 | 0.001656 | -154.0234 | 0.22881 | 8.4009 | 0.002463 | 0.003758 | 0.003615 | 4.9e-05 | 0.00114 | 0.262 | 1.3852 | 2.5476 | 98.56 | 20 | 50.121 | 63.9629 | 80.2517 | 60.8404 | 0.011902 | 1.0571 | 0.024518 | 23.195 | 0.05789 | 0.13612 | -7.7e-05 | -0.005956 | -0.005956 | -0.92105
2019-04-09 14:28 | 1554834469.524 | -412 | 0.20762 | 0.000498 | 6e-06 | 92.6721 | 92.0074 | 123.5054 | 20.4166 | 0.001787 | 0.001669 | -154.0476 | 0.22935 | 8.4013 | 0.002556 | 0.003651 | 0.003481 | 5.2e-05 | 0.00114 | 0.27256 | 1.3628 | 2.5477 | 98.56 | 20 | 50.124 | 64.0122 | 82.4947 | 61.1579 | 0.011903 | 1.056 | 0.024495 | 23.1948 | 0.064392 | 0.13644 | -7.5e-05 | -0.005761 | -0.005761 | -0.95436
2019-04-09 14:28 | 1554834470.5234 | -411 | 0.20761 | 0.0005 | 6e-06 | 92.6742 | 91.8622 | 123.8488 | 20.4185 | 0.001721 | 0.001681 | -154.0513 | 0.22894 | 8.4019 | 0.002519 | 0.003635 | 0.003525 | 5e-05 | 0.001138 | 0.30431 | 1.3282 | 2.5476 | 98.56 | 20 | 50.132 | 62.4232 | 81.6089 | 61.4599 | 0.011905 | 1.0551 | 0.024518 | 23.2375 | 0.078723 | 0.1362 | -7.8e-05 | -0.006053 | -0.006053 | -0.90675
2019-04-09 14:28 | 1554834471.5245 | -410 | 0.2076 | 0.0005 | 6e-06 | 92.7764 | 92.1898 | 123.7204 | 20.4165 | 0.001811 | 0.001693 | -154.0606 | 0.22882 | 8.4031 | 0.002525 | 0.003575 | 0.003227 | 5.2e-05 | 0.001134 | 0.24906 | 1.3161 | 2.5476 | 98.56 | 20 | 50.136 | 64.5767 | 81.7396 | 61.7487 | 0.011906 | 1.0542 | 0.024495 | 23.2355 | 0.056732 | 0.13613 | -7.8e-05 | -0.005995 | -0.005995 | -0.9151
2019-04-09 14:28 | 1554834472.5244 | -409 | 0.20763 | 0.000501 | 6e-06 | 92.7725 | 92.3798 | 123.922 | 20.4181 | 0.001778 | 0.001666 | -154.0543 | 0.22899 | 8.4025 | 0.002477 | 0.003742 | 0.003609 | 4.5e-05 | 0.001132 | 0.2792 | 1.3796 | 2.5476 | 98.56 | 20 | 50.139 | 63.7998 | 80.58 | 61.0821 | 0.011906 | 1.0563 | 0.024539 | 23.2313 | 0.037946 | 0.13622 | -7.8e-05 | -0.006049 | -0.006049 | -0.90754
2019-04-09 14:28 | 1554834473.5249 | -408 | 0.20765 | 0.000503 | 6e-06 | 92.8699 | 92.3268 | 124.3377 | 20.4177 | 0.00175 | 0.001688 | -154.0581 | 0.22919 | 8.4005 | 0.002558 | 0.003625 | 0.003498 | 5.1e-05 | 0.001138 | 0.29358 | 1.3706 | 2.5478 | 98.56 | 20 | 50.142 | 63.1169 | 82.534 | 61.6294 | 0.011907 | 1.0546 | 0.02456 | 23.2892 | 0.052462 | 0.13635 | -8e-05 | -0.006198 | -0.006198 | -0.88649
2019-04-09 14:28 | 1554834474.5246 | -407 | 0.20761 | 0.000504 | 6e-06 | 92.9786 | 92.3595 | 124.3497 | 20.4175 | 0.001753 | 0.001667 | -154.0778 | 0.22921 | 8.4019 | 0.002546 | 0.003955 | 0.003744 | 5.2e-05 | 0.001134 | 0.28208 | 1.3901 | 2.5475 | 98.56 | 20 | 50.149 | 63.1895 | 82.2608 | 61.1222 | 0.011909 | 1.0562 | 0.02458 | 23.2727 | 0.059758 | 0.13636 | -8.2e-05 | -0.006312 | -0.006312 | -0.87063
2019-04-09 14:28 | 1554834475.5245 | -406 | 0.20763 | 0.000497 | 6e-06 | 93.1432 | 92.5474 | 124.1884 | 20.4182 | 0.001785 | 0.001637 | -154.0881 | 0.22837 | 8.4005 | 0.002581 | 0.004423 | 0.004254 | 5.2e-05 | 0.00113 | 0.27776 | 1.3699 | 2.5476 | 98.56 | 20 | 50.154 | 63.9585 | 83.0937 | 60.3863 | 0.01191 | 1.0585 | 0.024591 | 23.232 | 0.057406 | 0.13586 | -7.3e-05 | -0.005676 | -0.005676 | -0.96458
2019-04-09 14:28 | 1554834476.5236 | -405 | 0.20762 | 0.000501 | 6e-06 | 93.1814 | 92.7642 | 124.471 | 20.4183 | 0.001838 | 0.001666 | -154.1014 | 0.22862 | 8.4023 | 0.002613 | 0.00415 | 0.003958 | 5e-05 | 0.001139 | 0.29551 | 1.3778 | 2.5477 | 98.56 | 20 | 50.158 | 65.2451 | 83.8529 | 61.0916 | 0.011911 | 1.0563 | 0.024593 | 23.283 | 0.040136 | 0.13601 | -7.8e-05 | -0.006053 | -0.006053 | -0.90555
2019-04-09 14:28 | 1554834477.524 | -404 | 0.20762 | 0.000503 | 5e-06 | 93.2746 | 92.7163 | 124.2616 | 20.419 | 0.001824 | 0.001695 | -154.0988 | 0.22846 | 8.4036 | 0.002543 | 0.004367 | 0.004144 | 5.2e-05 | 0.001136 | 0.27544 | 1.3754 | 2.5475 | 98.56 | 20 | 50.163 | 64.9001 | 82.1771 | 61.7949 | 0.011912 | 1.054 | 0.024546 | 23.2879 | 0.053692 | 0.13591 | -8.1e-05 | -0.006201 | -0.006201 | -0.88327
2019-04-09 14:28 | 1554834478.5237 | -403 | 0.20765 | 0.000503 | 6e-06 | 93.3368 | 92.7838 | 124.4656 | 20.4192 | 0.001768 | 0.001662 | -154.1134 | 0.22981 | 8.4019 | 0.002549 | 0.003652 | 0.003544 | 5e-05 | 0.001142 | 0.28008 | 1.3307 | 2.5476 | 98.56 | 20 | 50.166 | 63.5541 | 82.3225 | 60.9947 | 0.011913 | 1.0566 | 0.024596 | 23.2791 | 0.053153 | 0.13671 | -8.2e-05 | -0.006334 | -0.006334 | -0.86978
2019-04-09 14:28 | 1554834479.5237 | -402 | 0.20762 | 0.0005 | 6e-06 | 93.3633 | 92.8771 | 124.3386 | 20.4176 | 0.001781 | 0.001649 | -154.0965 | 0.22859 | 8.4029 | 0.002583 | 0.003866 | 0.003659 | 5.3e-05 | 0.001145 | 0.2975 | 1.346 | 2.5478 | 98.56 | 20 | 50.171 | 63.8705 | 83.1453 | 60.6797 | 0.011914 | 1.0576 | 0.024595 | 23.2563 | 0.046706 | 0.13599 | -7.8e-05 | -0.006028 | -0.006028 | -0.90924
2019-04-09 14:28 | 1554834480.5226 | -401 | 0.20888 | 0.0005 | 6e-06 | 93.4205 | 92.6885 | 123.7652 | 20.4189 | 0.001778 | 0.001662 | -154.1087 | 0.22863 | 8.3992 | 0.00265 | 0.003834 | 0.003662 | 5.3e-05 | 0.001154 | 0.30182 | 1.3398 | 2.5475 | 98.56 | 20 | 50.172 | 63.7967 | 84.7469 | 60.9937 | 0.011914 | 1.0566 | 0.024527 | 23.2135 | 0.070364 | 0.13601 | -7.7e-05 | -0.006014 | -0.006014 | -0.91138
2019-04-09 14:28 | 1554834481.5231 | -400 | 0.21062 | 0.000498 | 5e-06 | 93.4857 | 92.9517 | 124.0452 | 20.4185 | 0.001854 | 0.001692 | -154.1112 | 0.22919 | 8.4007 | 0.002627 | 0.004009 | 0.003811 | 5.2e-05 | 0.001159 | 0.27096 | 1.3455 | 2.5476 | 98.56 | 20 | 50.172 | 65.6254 | 84.2035 | 61.7099 | 0.011914 | 1.0543 | 0.024528 | 23.2646 | 0.051235 | 0.13634 | -7.5e-05 | -0.005837 | -0.005837 | -0.9414
2019-04-09 14:28 | 1554834482.5228 | -399 | 0.21093 | 0.000497 | 5e-06 | 93.4664 | 92.9834 | 123.8528 | 20.4191 | 0.001799 | 0.001682 | -154.1081 | 0.22937 | 8.4045 | 0.002638 | 0.004136 | 0.003958 | 5e-05 | 0.001163 | 0.27433 | 1.3725 | 2.5477 | 98.56 | 20 | 50.176 | 64.2943 | 84.4693 | 61.4692 | 0.011915 | 1.0551 | 0.024518 | 23.2382 | 0.046336 | 0.13645 | -7.3e-05 | -0.005753 | -0.005753 | -0.9559
2019-04-09 14:28 | 1554834483.5228 | -398 | 0.21111 | 0.000502 | 5e-06 | 93.6188 | 92.9745 | 123.2638 | 20.4198 | 0.001769 | 0.001722 | -154.1115 | 0.22897 | 8.4017 | 0.002562 | 0.003868 | 0.003682 | 4.9e-05 | 0.001168 | 0.29302 | 1.3549 | 2.5476 | 98.56 | 20 | 50.18 | 63.5672 | 82.6414 | 62.4456 | 0.011916 | 1.052 | 0.024424 | 23.2167 | 0.061777 | 0.13622 | -7.9e-05 | -0.006165 | -0.006165 | -0.89048
2019-04-09 14:28 | 1554834484.5229 | -397 | 0.21135 | 0.000499 | 5e-06 | 93.6965 | 93.2429 | 123.0648 | 20.4196 | 0.001772 | 0.001656 | -154.0968 | 0.2288 | 8.4034 | 0.002582 | 0.004027 | 0.003808 | 5.2e-05 | 0.001175 | 0.2896 | 1.3381 | 2.5476 | 98.56 | 20 | 50.192 | 63.6506 | 83.1208 | 60.8544 | 0.011919 | 1.057 | 0.024462 | 23.1429 | 0.043411 | 0.13611 | -7.6e-05 | -0.005918 | -0.005918 | -0.92688
2019-04-09 14:28 | 1554834485.5239 | -396 | 0.21158 | 0.000498 | 5e-06 | 93.7597 | 93.3145 | 123.0716 | 20.4196 | 0.001779 | 0.00169 | -154.0974 | 0.22865 | 8.4013 | 0.002533 | 0.003782 | 0.003646 | 4.9e-05 | 0.001181 | 0.26665 | 1.3459 | 2.5476 | 98.56 | 20 | 50.198 | 63.8106 | 81.9268 | 61.6852 | 0.01192 | 1.0544 | 0.024433 | 23.1723 | 0.042572 | 0.13603 | -7.4e-05 | -0.005751 | -0.005751 | -0.95325
2019-04-09 14:28 | 1554834486.5228 | -395 | 0.2118 | 0.00049 | 5e-06 | 93.8579 | 93.4401 | 122.9183 | 20.4195 | 0.001827 | 0.001668 | -154.0797 | 0.22828 | 8.4017 | 0.002514 | 0.003638 | 0.003523 | 4.9e-05 | 0.001174 | 0.28456 | 1.3547 | 2.5476 | 98.56 | 20 | 50.205 | 64.9747 | 81.4937 | 61.1461 | 0.011922 | 1.0561 | 0.024437 | 23.1392 | 0.039909 | 0.13581 | -6.3e-05 | -0.004947 | -0.004947 | -1.1063
2019-04-09 14:28 | 1554834487.523 | -394 | 0.21191 | 0.00049 | 5e-06 | 93.9235 | 93.4649 | 123.1939 | 20.4199 | 0.00175 | 0.00168 | -154.0786 | 0.22874 | 8.4002 | 0.002629 | 0.004134 | 0.003921 | 4.8e-05 | 0.001172 | 0.27671 | 1.3454 | 2.5476 | 98.56 | 20 | 50.216 | 63.1191 | 84.2414 | 61.4246 | 0.011925 | 1.0552 | 0.024454 | 23.1748 | 0.043779 | 0.13608 | -6.3e-05 | -0.004993 | -0.004993 | -1.0984
2019-04-09 14:28 | 1554834488.5234 | -393 | 0.21209 | 0.000488 | 5e-06 | 93.9824 | 93.6745 | 123.2462 | 20.4197 | 0.001796 | 0.001691 | -154.0793 | 0.22817 | 8.4035 | 0.002608 | 0.003629 | 0.003461 | 4.9e-05 | 0.001172 | 0.27417 | 1.3312 | 2.5477 | 98.56 | 20 | 50.221 | 64.2271 | 83.7473 | 61.6995 | 0.011926 | 1.0543 | 0.024449 | 23.1892 | 0.029357 | 0.13574 | -6.2e-05 | -0.004901 | -0.004901 | -1.1161
2019-04-09 14:28 | 1554834489.5234 | -392 | 0.21229 | 0.000487 | 4e-06 | 94.0363 | 93.5475 | 123.2317 | 20.4209 | 0.001766 | 0.001665 | -154.0851 | 0.22923 | 8.403 | 0.002608 | 0.004371 | 0.004159 | 5.1e-05 | 0.001171 | 0.25642 | 1.3258 | 2.5475 | 98.56 | 20 | 50.224 | 63.5113 | 83.7345 | 61.0684 | 0.011926 | 1.0563 | 0.024471 | 23.166 | 0.04662 | 0.13637 | -5.9e-05 | -0.00465 | -0.00465 | -1.1818
2019-04-09 14:28 | 1554834490.5231 | -391 | 0.21234 | 0.000482 | 4e-06 | 94.0893 | 93.5557 | 123.3469 | 20.4208 | 0.001755 | 0.001646 | -154.0748 | 0.22807 | 8.4033 | 0.002524 | 0.003998 | 0.003805 | 5e-05 | 0.001174 | 0.27113 | 1.3591 | 2.5475 | 98.56 | 20 | 50.231 | 63.2363 | 81.7164 | 60.6211 | 0.011928 | 1.0577 | 0.024499 | 23.1613 | 0.05087 | 0.13568 | -5.3e-05 | -0.004201 | -0.004201 | -1.3017
2019-04-09 14:28 | 1554834491.5234 | -390 | 0.21244 | 0.00048 | 4e-06 | 94.1828 | 93.6762 | 123.5788 | 20.4209 | 0.001801 | 0.001661 | -154.0682 | 0.22936 | 8.4026 | 0.0025 | 0.004403 | 0.00423 | 4.6e-05 | 0.001174 | 0.31116 | 1.3421 | 2.5475 | 98.56 | 20 | 50.234 | 64.3562 | 81.148 | 60.9855 | 0.011929 | 1.0566 | 0.024508 | 23.1957 | 0.048246 | 0.13645 | -5.1e-05 | -0.004058 | -0.004058 | -1.3552
2019-04-09 14:28 | 1554834492.5241 | -389 | 0.21246 | 0.000475 | 4e-06 | 94.1956 | 93.7381 | 123.4995 | 20.4203 | 0.001779 | 0.001664 | -154.0715 | 0.22876 | 8.3994 | 0.002467 | 0.003316 | 0.003176 | 4.7e-05 | 0.001179 | 0.27599 | 1.3281 | 2.5474 | 98.56 | 20 | 50.238 | 63.8081 | 80.3497 | 61.0466 | 0.01193 | 1.0564 | 0.024498 | 23.1904 | 0.043549 | 0.13609 | -4.4e-05 | -0.003525 | -0.003525 | -1.5557
2019-04-09 14:28 | 1554834493.5241 | -388 | 0.21249 | 0.000476 | 4e-06 | 94.2174 | 93.6184 | 123.7089 | 20.4215 | 0.001783 | 0.001688 | -154.0758 | 0.22821 | 8.3991 | 0.00251 | 0.004353 | 0.004176 | 5e-05 | 0.001185 | 0.28136 | 1.4129 | 2.5477 | 98.56 | 20 | 50.239 | 63.9116 | 81.3796 | 61.6362 | 0.01193 | 1.0545 | 0.024497 | 23.2305 | 0.057049 | 0.13576 | -4.5e-05 | -0.003569 | -0.003569 | -1.533
2019-04-09 14:28 | 1554834494.5231 | -387 | 0.21072 | 0.000472 | 4e-06 | 94.254 | 93.849 | 123.8016 | 20.421 | 0.001858 | 0.001632 | -154.1009 | 0.2297 | 8.4008 | 0.002536 | 0.003783 | 0.003619 | 4.7e-05 | 0.001191 | 0.28241 | 1.3905 | 2.5477 | 98.56 | 20 | 50.239 | 65.7206 | 82.0184 | 60.2758 | 0.01193 | 1.0588 | 0.024557 | 23.1919 | 0.038522 | 0.13665 | -3.9e-05 | -0.003129 | -0.003129 | -1.7598
2019-04-09 14:28 | 1554834495.5222 | -386 | 0.20922 | 0.000474 | 4e-06 | 94.3323 | 93.8159 | 123.4908 | 20.4214 | 0.001736 | 0.001694 | -154.0913 | 0.22806 | 8.4033 | 0.002479 | 0.003479 | 0.003287 | 4.4e-05 | 0.001193 | 0.28766 | 1.3931 | 2.5476 | 98.56 | 20.01 | 50.236 | 62.7924 | 80.6425 | 61.7715 | 0.011937 | 1.0541 | 0.024471 | 23.2147 | 0.049101 | 0.13567 | -4.2e-05 | -0.003336 | -0.003336 | -1.6388
2019-04-09 14:28 | 1554834496.5235 | -385 | 0.20919 | 0.000468 | 3e-06 | 94.4118 | 94.0947 | 123.2594 | 20.4215 | 0.001796 | 0.001632 | -154.113 | 0.2288 | 8.3998 | 0.002536 | 0.00346 | 0.003338 | 4.2e-05 | 0.001193 | 0.28622 | 1.3879 | 2.5476 | 98.56 | 20.01 | 50.233 | 64.2264 | 82.0009 | 60.2777 | 0.011936 | 1.0588 | 0.024503 | 23.1412 | 0.030098 | 0.13611 | -3.3e-05 | -0.002656 | -0.002656 | -2.0649
2019-04-09 14:28 | 1554834497.5223 | -384 | 0.20924 | 0.000466 | 4e-06 | 94.4086 | 94.0389 | 123.0007 | 20.4215 | 0.001787 | 0.001648 | -154.1034 | 0.22867 | 8.403 | 0.00257 | 0.003682 | 0.003484 | 4.3e-05 | 0.001196 | 0.28279 | 1.3811 | 2.5477 | 98.56 | 20.01 | 50.224 | 64.0062 | 82.8238 | 60.6594 | 0.011934 | 1.0576 | 0.024463 | 23.1301 | 0.035089 | 0.13604 | -3.2e-05 | -0.002596 | -0.002596 | -2.1122
2019-04-09 14:28 | 1554834498.5227 | -383 | 0.20926 | 0.000466 | 4e-06 | 94.4044 | 93.893 | 123.0885 | 20.4221 | 0.001782 | 0.001674 | -154.0846 | 0.2291 | 8.4013 | 0.002501 | 0.004177 | 0.004038 | 4.6e-05 | 0.001201 | 0.26842 | 1.3145 | 2.5476 | 98.56 | 20.01 | 50.221 | 63.885 | 81.161 | 61.2869 | 0.011933 | 1.0556 | 0.024449 | 23.1601 | 0.048584 | 0.13629 | -3.2e-05 | -0.002572 | -0.002572 | -2.1352
2019-04-09 14:28 | 1554834499.5223 | -382 | 0.20927 | 0.000465 | 4e-06 | 94.5092 | 94.0844 | 122.8681 | 20.4221 | 0.001809 | 0.001669 | -154.0791 | 0.22931 | 8.4013 | 0.002473 | 0.003851 | 0.003698 | 4.4e-05 | 0.001208 | 0.28313 | 1.3501 | 2.5476 | 98.56 | 20.01 | 50.217 | 64.5331 | 80.5106 | 61.1594 | 0.011932 | 1.056 | 0.024432 | 23.1349 | 0.0403 | 0.13642 | -3e-05 | -0.002465 | -0.002465 | -2.23
2019-04-09 14:28 | 1554834500.5227 | -381 | 0.20927 | 0.000462 | 4e-06 | 94.4878 | 94.1511 | 122.9101 | 20.4227 | 0.001773 | 0.001696 | -154.0906 | 0.22862 | 8.4012 | 0.002481 | 0.003908 | 0.003674 | 4.7e-05 | 0.00121 | 0.27865 | 1.3483 | 2.5477 | 98.56 | 20.01 | 50.21 | 63.6786 | 80.6905 | 61.825 | 0.011931 | 1.0539 | 0.024411 | 23.1619 | 0.031926 | 0.136 | -2.7e-05 | -0.002256 | -0.002256 | -2.43
2019-04-09 14:28 | 1554834501.5224 | -380 | 0.20927 | 0.000467 | 3e-06 | 94.5908 | 94.3273 | 122.3841 | 20.4219 | 0.001848 | 0.001699 | -154.113 | 0.22852 | 8.4033 | 0.002499 | 0.003793 | 0.003601 | 4.4e-05 | 0.001204 | 0.28263 | 1.3417 | 2.5477 | 98.56 | 20.01 | 50.207 | 65.4694 | 81.1319 | 61.9001 | 0.01193 | 1.0537 | 0.024356 | 23.1149 | 0.02495 | 0.13595 | -3.2e-05 | -0.00257 | -0.00257 | -2.1322
2019-04-09 14:28 | 1554834502.5229 | -379 | 0.2093 | 0.000463 | 4e-06 | 94.6364 | 94.4622 | 122.6186 | 20.4229 | 0.001807 | 0.001717 | -154.1092 | 0.22934 | 8.4026 | 0.002499 | 0.004086 | 0.003946 | 4.3e-05 | 0.001199 | 0.26792 | 1.3541 | 2.5477 | 98.56 | 20.01 | 50.204 | 64.4893 | 81.136 | 62.3261 | 0.011929 | 1.0524 | 0.024364 | 23.1517 | 0.016474 | 0.13643 | -2.8e-05 | -0.002241 | -0.002241 | -2.4533
2019-04-09 14:28 | 1554834503.5234 | -378 | 0.20934 | 0.00046 | 4e-06 | 94.6414 | 94.4147 | 122.8531 | 20.4227 | 0.001767 | 0.001669 | -154.1012 | 0.22805 | 8.4034 | 0.002487 | 0.003904 | 0.003692 | 4.3e-05 | 0.001198 | 0.27201 | 1.3407 | 2.5476 | 98.56 | 20.01 | 50.198 | 63.538 | 80.8356 | 61.1662 | 0.011928 | 1.056 | 0.02443 | 23.1338 | 0.02145 | 0.13567 | -2.4e-05 | -0.002047 | -0.002047 | -2.6705
2019-04-09 14:28 | 1554834504.5235 | -377 | 0.20931 | 0.000465 | 4e-06 | 94.8252 | 94.519 | 122.7733 | 20.4232 | 0.001772 | 0.001663 | -154.0772 | 0.22901 | 8.4031 | 0.002424 | 0.003844 | 0.003733 | 4.6e-05 | 0.001198 | 0.29878 | 1.3776 | 2.5477 | 98.56 | 20.01 | 50.196 | 63.6603 | 79.3156 | 61.0284 | 0.011927 | 1.0565 | 0.024427 | 23.1215 | 0.028925 | 0.13624 | -3.1e-05 | -0.002474 | -0.002474 | -2.2196
2019-04-09 14:28 | 1554834505.524 | -376 | 0.20938 | 0.000465 | 4e-06 | 94.8922 | 94.5483 | 123.0411 | 20.4228 | 0.001832 | 0.00169 | -154.0869 | 0.22874 | 8.4019 | 0.002505 | 0.003878 | 0.003687 | 4.2e-05 | 0.001196 | 0.27113 | 1.336 | 2.5476 | 98.56 | 20.02 | 50.19 | 65.0958 | 81.2708 | 61.6674 | 0.011933 | 1.0544 | 0.02443 | 23.1688 | 0.032475 | 0.13608 | -3e-05 | -0.002443 | -0.002443 | -2.2445
2019-04-09 14:28 | 1554834506.5226 | -375 | 0.20936 | 0.000463 | 4e-06 | 94.9711 | 94.5782 | 123.1157 | 20.424 | 0.001818 | 0.001691 | -154.0984 | 0.22889 | 8.4029 | 0.002494 | 0.003989 | 0.003829 | 4.3e-05 | 0.001193 | 0.27057 | 1.3273 | 2.5474 | 98.56 | 20.02 | 50.186 | 64.7613 | 81.0161 | 61.6902 | 0.011932 | 1.0544 | 0.024437 | 23.1766 | 0.037078 | 0.13617 | -2.8e-05 | -0.002261 | -0.002261 | -2.4267
2019-04-09 14:28 | 1554834507.523 | -374 | 0.20936 | 0.000464 | 4e-06 | 95.0455 | 94.7678 | 123.6819 | 20.424 | 0.001765 | 0.001656 | -154.1121 | 0.22908 | 8.4022 | 0.002437 | 0.003144 | 0.003021 | 4.3e-05 | 0.001201 | 0.30502 | 1.3281 | 2.5475 | 98.56 | 20.01 | 50.179 | 63.4738 | 79.6416 | 60.8642 | 0.011923 | 1.057 | 0.024523 | 23.2012 | 0.026172 | 0.13628 | -2.9e-05 | -0.002407 | -0.002407 | -2.2813
2019-04-09 14:28 | 1554834508.5226 | -373 | 0.20935 | 0.000462 | 4e-06 | 95.0941 | 94.932 | 123.4408 | 20.4234 | 0.001769 | 0.001613 | -154.119 | 0.22957 | 8.4013 | 0.002421 | 0.0036 | 0.0034 | 4.4e-05 | 0.0012 | 0.29839 | 1.3479 | 2.5475 | 98.56 | 20.01 | 50.176 | 63.57 | 79.2596 | 59.8098 | 0.011922 | 1.0603 | 0.024538 | 23.1419 | 0.015258 | 0.13657 | -2.7e-05 | -0.002212 | -0.002212 | -2.4886
2019-04-09 14:28 | 1554834509.5227 | -372 | 0.20936 | 0.00046 | 4e-06 | 95.1593 | 94.8467 | 123.4772 | 20.4237 | 0.00185 | 0.001694 | -154.1258 | 0.2281 | 8.4 | 0.002386 | 0.004014 | 0.003844 | 4.5e-05 | 0.001196 | 0.27959 | 1.3347 | 2.5476 | 98.56 | 20.02 | 50.172 | 65.5359 | 78.4189 | 61.7831 | 0.011929 | 1.0541 | 0.024469 | 23.2138 | 0.029435 | 0.1357 | -2.5e-05 | -0.002041 | -0.002041 | -2.6789
2019-04-09 14:28 | 1554834510.5221 | -371 | 0.2094 | 0.000457 | 4e-06 | 95.2731 | 95.0242 | 123.6554 | 20.4234 | 0.001837 | 0.001679 | -154.1096 | 0.22822 | 8.4037 | 0.002435 | 0.003553 | 0.003417 | 4.2e-05 | 0.0012 | 0.26344 | 1.3723 | 2.5476 | 98.56 | 20.02 | 50.162 | 65.2073 | 79.5888 | 61.4218 | 0.011927 | 1.0552 | 0.0245 | 23.218 | 0.02339 | 0.13577 | -2.1e-05 | -0.001769 | -0.001769 | -3.0933
2019-04-09 14:28 | 1554834511.5218 | -370 | 0.20939 | 0.000457 | 4e-06 | 95.3082 | 95.0494 | 123.8007 | 20.4242 | 0.001718 | 0.001678 | -154.0947 | 0.22929 | 8.4008 | 0.002455 | 0.003837 | 0.00367 | 4.5e-05 | 0.001198 | 0.28423 | 1.3275 | 2.5476 | 98.56 | 20.02 | 50.156 | 62.3633 | 80.0635 | 61.399 | 0.011925 | 1.0553 | 0.024515 | 23.2309 | 0.024324 | 0.1364 | -2e-05 | -0.001722 | -0.001722 | -3.1932
2019-04-09 14:29 | 1554834512.5218 | -369 | 0.20939 | 0.00046 | 4e-06 | 95.4309 | 95.3488 | 123.738 | 20.4246 | 0.001778 | 0.001681 | -154.1125 | 0.22859 | 8.3996 | 0.002477 | 0.004003 | 0.003771 | 4.1e-05 | 0.001198 | 0.28783 | 1.3217 | 2.5476 | 98.56 | 20.02 | 50.15 | 63.8001 | 80.6051 | 61.4588 | 0.011924 | 1.0551 | 0.024507 | 23.2271 | 0.007699 | 0.13599 | -2.5e-05 | -0.00203 | -0.00203 | -2.6997
2019-04-09 14:29 | 1554834513.5218 | -368 | 0.20937 | 0.000461 | 4e-06 | 95.561 | 95.3904 | 123.4864 | 20.4243 | 0.001812 | 0.001643 | -154.0997 | 0.2288 | 8.4005 | 0.002448 | 0.003876 | 0.003715 | 4.6e-05 | 0.001193 | 0.2855 | 1.3153 | 2.5476 | 98.56 | 20.02 | 50.141 | 64.6055 | 79.892 | 60.5528 | 0.011922 | 1.058 | 0.024515 | 23.172 | 0.015982 | 0.13612 | -2.5e-05 | -0.002077 | -0.002077 | -2.641
2019-04-09 14:29 | 1554834514.522 | -367 | 0.20939 | 0.00046 | 3e-06 | 95.5268 | 95.25 | 123.2038 | 20.4247 | 0.001801 | 0.001653 | -154.1111 | 0.22914 | 8.4009 | 0.002412 | 0.003651 | 0.003505 | 4.2e-05 | 0.001189 | 0.26432 | 1.3276 | 2.5475 | 98.56 | 20.02 | 50.136 | 64.3498 | 79.0403 | 60.7977 | 0.01192 | 1.0572 | 0.024478 | 23.154 | 0.025952 | 0.13631 | -2.4e-05 | -0.001972 | -0.001972 | -2.7857
2019-04-09 14:29 | 1554834515.5218 | -366 | 0.20938 | 0.000458 | 3e-06 | 95.6334 | 95.4111 | 123.1174 | 20.4248 | 0.001778 | 0.001682 | -154.1018 | 0.22877 | 8.4021 | 0.002469 | 0.003412 | 0.003309 | 4.3e-05 | 0.001187 | 0.27295 | 1.3713 | 2.5478 | 98.56 | 20.02 | 50.125 | 63.8072 | 80.3955 | 61.4966 | 0.011918 | 1.055 | 0.024444 | 23.1701 | 0.020817 | 0.1361 | -2.1e-05 | -0.001738 | -0.001738 | -3.1564
2019-04-09 14:29 | 1554834516.5219 | -365 | 0.20942 | 0.000459 | 3e-06 | 95.7257 | 95.5412 | 122.9838 | 20.4243 | 0.001817 | 0.001675 | -154.0868 | 0.22904 | 8.4022 | 0.002453 | 0.003489 | 0.003329 | 4.2e-05 | 0.001174 | 0.28838 | 1.4125 | 2.5478 | 98.56 | 20.02 | 50.119 | 64.7442 | 80.0128 | 61.3162 | 0.011916 | 1.0555 | 0.024437 | 23.1513 | 0.01726 | 0.13626 | -2.3e-05 | -0.001848 | -0.001848 | -2.9721
2019-04-09 14:29 | 1554834517.5223 | -364 | 0.20941 | 0.000461 | 4e-06 | 95.7945 | 95.5825 | 122.5564 | 20.4253 | 0.001782 | 0.001649 | -154.0832 | 0.22829 | 8.4034 | 0.002494 | 0.004104 | 0.003902 | 4.4e-05 | 0.001175 | 0.29159 | 1.3688 | 2.5477 | 98.56 | 20.02 | 50.113 | 63.8919 | 81.0056 | 60.6903 | 0.011915 | 1.0575 | 0.024418 | 23.0894 | 0.01982 | 0.13581 | -2.5e-05 | -0.002084 | -0.002084 | -2.6262
2019-04-09 14:29 | 1554834518.522 | -363 | 0.20939 | 0.000459 | 3e-06 | 95.8609 | 95.8017 | 122.6226 | 20.426 | 0.001795 | 0.001645 | -154.0733 | 0.22888 | 8.4018 | 0.002478 | 0.003729 | 0.003552 | 4.5e-05 | 0.001173 | 0.29551 | 1.3507 | 2.5476 | 98.56 | 20.02 | 50.102 | 64.2096 | 80.6345 | 60.6028 | 0.011912 | 1.0578 | 0.024427 | 23.0926 | 0.005527 | 0.13616 | -2.3e-05 | -0.001901 | -0.001901 | -2.886
2019-04-09 14:29 | 1554834519.5233 | -362 | 0.20938 | 0.000454 | 3e-06 | 95.9916 | 95.6823 | 123.2002 | 20.4259 | 0.001813 | 0.001667 | -154.0824 | 0.22947 | 8.4029 | 0.002397 | 0.003991 | 0.003832 | 4.8e-05 | 0.001167 | 0.30038 | 1.3121 | 2.5475 | 98.56 | 20.02 | 50.096 | 64.639 | 78.686 | 61.1286 | 0.011911 | 1.0561 | 0.024466 | 23.1651 | 0.028873 | 0.13651 | -1.6e-05 | -0.001388 | -0.001388 | -3.9622
2019-04-09 14:29 | 1554834520.523 | -361 | 0.20934 | 0.000455 | 3e-06 | 96.0625 | 95.8853 | 123.1845 | 20.4259 | 0.001814 | 0.00172 | -154.0908 | 0.22839 | 8.3992 | 0.002431 | 0.003074 | 0.002957 | 4.4e-05 | 0.001165 | 0.27975 | 1.3193 | 2.5476 | 98.56 | 20.02 | 50.086 | 64.6734 | 79.4968 | 62.3983 | 0.011909 | 1.0521 | 0.024418 | 23.2076 | 0.016513 | 0.13587 | -1.7e-05 | -0.001439 | -0.001439 | -3.8039
2019-04-09 14:29 | 1554834521.5234 | -360 | 0.20935 | 0.000456 | 3e-06 | 96.1066 | 95.8908 | 123.0761 | 20.4257 | 0.001807 | 0.001641 | -154.0973 | 0.22935 | 8.4029 | 0.002394 | 0.003403 | 0.003231 | 4.6e-05 | 0.00117 | 0.28335 | 1.3477 | 2.5475 | 98.56 | 20.03 | 50.082 | 64.4865 | 78.6118 | 60.4977 | 0.011915 | 1.0581 | 0.024476 | 23.1316 | 0.020103 | 0.13644 | -1.9e-05 | -0.001573 | -0.001573 | -3.4966
2019-04-09 14:29 | 1554834522.5231 | -359 | 0.20936 | 0.000454 | 3e-06 | 96.156 | 96.0354 | 122.896 | 20.4265 | 0.0018 | 0.001654 | -154.0788 | 0.22908 | 8.4027 | 0.002416 | 0.003939 | 0.003774 | 4.4e-05 | 0.001166 | 0.26626 | 1.3468 | 2.5475 | 98.56 | 20.03 | 50.077 | 64.3335 | 79.1365 | 60.8071 | 0.011914 | 1.0572 | 0.024447 | 23.1254 | 0.011229 | 0.13628 | -1.6e-05 | -0.001392 | -0.001392 | -3.9453
2019-04-09 14:29 | 1554834523.5231 | -358 | 0.20936 | 0.000456 | 4e-06 | 96.2034 | 96.1768 | 122.9204 | 20.4268 | 0.001809 | 0.001727 | -154.0928 | 0.22854 | 8.4028 | 0.002415 | 0.003559 | 0.0034 | 4.3e-05 | 0.001172 | 0.27721 | 1.3204 | 2.5476 | 98.56 | 20.03 | 50.069 | 64.5352 | 79.1141 | 62.5822 | 0.011912 | 1.0516 | 0.024385 | 23.1891 | 0.002468 | 0.13596 | -1.9e-05 | -0.00157 | -0.00157 | -3.4888
2019-04-09 14:29 | 1554834524.5231 | -357 | 0.20933 | 0.000454 | 3e-06 | 96.3031 | 96.1448 | 123.1143 | 20.4267 | 0.001827 | 0.001685 | -154.0918 | 0.22941 | 8.4012 | 0.002467 | 0.003587 | 0.003377 | 4.8e-05 | 0.001178 | 0.28982 | 1.3463 | 2.5476 | 98.56 | 20.02 | 50.064 | 64.984 | 80.3653 | 61.549 | 0.011903 | 1.0548 | 0.024442 | 23.1716 | 0.014716 | 0.13648 | -1.7e-05 | -0.001422 | -0.001422 | -3.8664
2019-04-09 14:29 | 1554834525.5231 | -356 | 0.20934 | 0.000453 | 3e-06 | 96.4198 | 96.331 | 123.0434 | 20.427 | 0.001768 | 0.001694 | -154.089 | 0.2293 | 8.4016 | 0.002423 | 0.004279 | 0.004156 | 4.5e-05 | 0.001178 | 0.28799 | 1.3307 | 2.5477 | 98.56 | 20.02 | 50.059 | 63.5579 | 79.2967 | 61.7854 | 0.011902 | 1.0541 | 0.024426 | 23.1731 | 0.008247 | 0.13641 | -1.5e-05 | -0.001293 | -0.001293 | -4.2509
2019-04-09 14:29 | 1554834526.5222 | -355 | 0.20936 | 0.000453 | 3e-06 | 96.4833 | 96.4053 | 123.5589 | 20.4268 | 0.001796 | 0.001659 | -154.0798 | 0.22919 | 8.401 | 0.002402 | 0.003582 | 0.003446 | 4.7e-05 | 0.001169 | 0.27848 | 1.3057 | 2.5478 | 98.56 | 20.03 | 50.051 | 64.2353 | 78.7936 | 60.9286 | 0.011908 | 1.0568 | 0.024509 | 23.1919 | 0.007229 | 0.13634 | -1.4e-05 | -0.001212 | -0.001212 | -4.5333
2019-04-09 14:29 | 1554834527.5212 | -354 | 0.20936 | 0.000453 | 3e-06 | 96.5202 | 96.4225 | 123.833 | 20.4279 | 0.001832 | 0.001682 | -154.0687 | 0.22919 | 8.4029 | 0.002442 | 0.003696 | 0.003471 | 4.2e-05 | 0.001155 | 0.28512 | 1.339 | 2.5476 | 98.56 | 20.03 | 50.047 | 65.099 | 79.7489 | 61.4885 | 0.011907 | 1.055 | 0.024515 | 23.237 | 0.009055 | 0.13634 | -1.4e-05 | -0.001223 | -0.001223 | -4.4931
2019-04-09 14:29 | 1554834528.5216 | -353 | 0.20934 | 0.000454 | 3e-06 | 96.5941 | 96.554 | 123.6763 | 20.4276 | 0.001831 | 0.001704 | -154.0858 | 0.22823 | 8.4022 | 0.00248 | 0.004042 | 0.003849 | 4.3e-05 | 0.001157 | 0.2943 | 1.3311 | 2.5475 | 98.56 | 20.03 | 50.044 | 65.0747 | 80.6806 | 62.0244 | 0.011906 | 1.0533 | 0.02448 | 23.2409 | 0.00371 | 0.13578 | -1.5e-05 | -0.001263 | -0.001263 | -4.3314
2019-04-09 14:29 | 1554834529.5213 | -352 | 0.20934 | 0.000455 | 3e-06 | 96.676 | 96.4298 | 123.1495 | 20.4271 | 0.001707 | 0.001665 | -154.0542 | 0.22941 | 8.4014 | 0.002429 | 0.003933 | 0.003754 | 4.7e-05 | 0.001159 | 0.27455 | 1.3215 | 2.5475 | 98.56 | 20.03 | 50.044 | 62.0879 | 79.4504 | 61.0884 | 0.011906 | 1.0563 | 0.024462 | 23.159 | 0.022806 | 0.13648 | -1.7e-05 | -0.001421 | -0.001421 | -3.8711
2019-04-09 14:29 | 1554834530.5212 | -351 | 0.20938 | 0.000458 | 3e-06 | 96.8314 | 96.7206 | 122.782 | 20.4277 | 0.0018 | 0.001678 | -154.0678 | 0.22889 | 8.4017 | 0.002403 | 0.003378 | 0.003205 | 4.5e-05 | 0.001165 | 0.28031 | 1.3778 | 2.5477 | 98.56 | 20.03 | 50.043 | 64.3229 | 78.8334 | 61.3915 | 0.011906 | 1.0553 | 0.024414 | 23.1349 | 0.010243 | 0.13617 | -2e-05 | -0.001649 | -0.001649 | -3.3275
2019-04-09 14:29 | 1554834531.5217 | -350 | 0.20935 | 0.000455 | 3e-06 | 96.8665 | 96.9154 | 122.8841 | 20.4285 | 0.001802 | 0.001631 | -154.0502 | 0.22886 | 8.4029 | 0.002438 | 0.004131 | 0.003973 | 4.3e-05 | 0.001165 | 0.27848 | 1.3631 | 2.5478 | 98.56 | 20.03 | 50.04 | 64.3838 | 79.6598 | 60.258 | 0.011905 | 1.0589 | 0.024466 | 23.1052 | -0.004508 | 0.13615 | -1.6e-05 | -0.001376 | -0.001376 | -3.9877
2019-04-09 14:29 | 1554834532.5214 | -349 | 0.20935 | 0.000453 | 4e-06 | 96.9645 | 97.0048 | 123.0343 | 20.4285 | 0.001804 | 0.001669 | -154.0685 | 0.22856 | 8.401 | 0.002417 | 0.003519 | 0.003275 | 4.4e-05 | 0.001173 | 0.28567 | 1.3231 | 2.5476 | 98.56 | 20.03 | 50.039 | 64.43 | 79.153 | 61.1687 | 0.011905 | 1.056 | 0.024448 | 23.1509 | -0.00372 | 0.13597 | -1.6e-05 | -0.001403 | -0.001403 | -3.9046
2019-04-09 14:29 | 1554834533.5218 | -348 | 0.20933 | 0.000453 | 3e-06 | 97.0616 | 96.8197 | 123.2671 | 20.4276 | 0.001775 | 0.001666 | -154.0635 | 0.22887 | 8.4031 | 0.002447 | 0.00414 | 0.004056 | 4.3e-05 | 0.001177 | 0.26073 | 1.3762 | 2.5477 | 98.56 | 20.03 | 50.036 | 63.7366 | 79.875 | 61.0944 | 0.011904 | 1.0562 | 0.024474 | 23.1702 | 0.02232 | 0.13616 | -1.4e-05 | -0.00119 | -0.00119 | -4.6104
2019-04-09 14:29 | 1554834534.5214 | -347 | 0.20936 | 0.000455 | 3e-06 | 97.1149 | 97.0874 | 123.554 | 20.4275 | 0.00182 | 0.001662 | -154.0672 | 0.22871 | 8.4027 | 0.002441 | 0.004053 | 0.003792 | 4.5e-05 | 0.001177 | 0.28783 | 1.3806 | 2.5477 | 98.56 | 20.03 | 50.035 | 64.8186 | 79.7322 | 61.0013 | 0.011904 | 1.0565 | 0.024505 | 23.1939 | 0.002529 | 0.13606 | -1.6e-05 | -0.001345 | -0.001345 | -4.0772
2019-04-09 14:29 | 1554834535.5215 | -346 | 0.20935 | 0.000453 | 3e-06 | 97.1518 | 97.0754 | 123.4528 | 20.4275 | 0.001838 | 0.001694 | -154.0769 | 0.22938 | 8.3993 | 0.002395 | 0.004023 | 0.003887 | 4.5e-05 | 0.001181 | 0.2808 | 1.3446 | 2.5476 | 98.56 | 20.03 | 50.034 | 65.2517 | 78.6213 | 61.7712 | 0.011904 | 1.0541 | 0.024467 | 23.2111 | 0.007039 | 0.13646 | -1.5e-05 | -0.00124 | -0.00124 | -4.4362
2019-04-09 14:29 | 1554834536.5225 | -345 | 0.20938 | 0.000455 | 3e-06 | 97.1905 | 97.233 | 122.9972 | 20.4284 | 0.001833 | 0.001721 | -154.0923 | 0.22808 | 8.4041 | 0.002432 | 0.003761 | 0.003636 | 4.5e-05 | 0.001175 | 0.30901 | 1.3251 | 2.5475 | 98.56 | 20.03 | 50.031 | 65.1137 | 79.5172 | 62.4352 | 0.011903 | 1.052 | 0.024398 | 23.1912 | -0.003909 | 0.13569 | -1.6e-05 | -0.001337 | -0.001337 | -4.0902
2019-04-09 14:29 | 1554834537.5215 | -344 | 0.20935 | 0.000456 | 3e-06 | 97.2602 | 97.169 | 122.9092 | 20.429 | 0.001799 | 0.001679 | -154.1174 | 0.22889 | 8.4026 | 0.002419 | 0.003085 | 0.002941 | 4.4e-05 | 0.001174 | 0.26897 | 1.3404 | 2.5475 | 98.56 | 20.03 | 50.029 | 64.2993 | 79.2121 | 61.4154 | 0.011902 | 1.0552 | 0.024426 | 23.1477 | 0.008392 | 0.13617 | -1.9e-05 | -0.001605 | -0.001605 | -3.4187
2019-04-09 14:29 | 1554834538.522 | -343 | 0.20937 | 0.000456 | 3e-06 | 97.3513 | 97.3249 | 123.1472 | 20.4291 | 0.001719 | 0.001737 | -154.0965 | 0.2287 | 8.4004 | 0.002487 | 0.003546 | 0.003284 | 4.4e-05 | 0.001166 | 0.27865 | 1.3388 | 2.5475 | 98.56 | 20.03 | 50.026 | 62.3829 | 80.8317 | 62.8211 | 0.011902 | 1.0508 | 0.024399 | 23.2187 | 0.002429 | 0.13605 | -1.8e-05 | -0.001493 | -0.001493 | -3.6728
2019-04-09 14:29 | 1554834539.523 | -342 | 0.20936 | 0.000455 | 3e-06 | 97.3793 | 97.3194 | 123.3014 | 20.4292 | 0.001773 | 0.001642 | -154.1284 | 0.22889 | 8.4006 | 0.002406 | 0.003537 | 0.003418 | 4.6e-05 | 0.001167 | 0.29048 | 1.3685 | 2.5476 | 98.56 | 20.03 | 50.021 | 63.6741 | 78.9043 | 60.5161 | 0.0119 | 1.0581 | 0.024498 | 23.1534 | 0.005507 | 0.13617 | -1.6e-05 | -0.001374 | -0.001374 | -3.9938
2019-04-09 14:29 | 1554834540.5227 | -341 | 0.2094 | 0.000455 | 3e-06 | 97.4559 | 97.6667 | 123.4141 | 20.4297 | 0.001735 | 0.001705 | -154.1074 | 0.22904 | 8.4016 | 0.00242 | 0.003751 | 0.003572 | 4.5e-05 | 0.001162 | 0.28495 | 1.3231 | 2.5476 | 98.56 | 20.03 | 50.017 | 62.7702 | 79.2347 | 62.0466 | 0.011899 | 1.0532 | 0.024453 | 23.217 | -0.019325 | 0.13626 | -1.7e-05 | -0.00147 | -0.00147 | -3.7356
2019-04-09 14:29 | 1554834541.521 | -340 | 0.20939 | 0.000456 | 3e-06 | 97.5917 | 97.6594 | 123.4993 | 20.4293 | 0.001816 | 0.001696 | -154.1291 | 0.22947 | 8.4039 | 0.002416 | 0.003889 | 0.003775 | 4.3e-05 | 0.00115 | 0.27041 | 1.3188 | 2.5475 | 98.56 | 20.03 | 50.008 | 64.7046 | 79.1406 | 61.8237 | 0.011897 | 1.0539 | 0.02447 | 23.2173 | -0.006209 | 0.13651 | -1.7e-05 | -0.001443 | -0.001443 | -3.8123
2019-04-09 14:29 | 1554834542.5206 | -339 | 0.20938 | 0.000454 | 3e-06 | 97.5735 | 97.7107 | 123.6831 | 20.43 | 0.001822 | 0.001655 | -154.12 | 0.22884 | 8.4013 | 0.002481 | 0.00391 | 0.003697 | 4.5e-05 | 0.001142 | 0.2756 | 1.3249 | 2.5475 | 98.56 | 20.03 | 50.005 | 64.8562 | 80.693 | 60.8506 | 0.011897 | 1.057 | 0.024524 | 23.2008 | -0.012569 | 0.13614 | -1.6e-05 | -0.001347 | -0.001347 | -4.0742
2019-04-09 14:29 | 1554834543.5211 | -338 | 0.20936 | 0.000456 | 3e-06 | 97.6262 | 97.7599 | 123.5441 | 20.4296 | 0.001752 | 0.001744 | -154.1327 | 0.22875 | 8.4021 | 0.00246 | 0.003812 | 0.003731 | 4.4e-05 | 0.001145 | 0.27544 | 1.3118 | 2.5477 | 98.56 | 20.03 | 49.997 | 63.1808 | 80.1872 | 62.9879 | 0.011895 | 1.0503 | 0.024432 | 23.2619 | -0.012243 | 0.13608 | -1.9e-05 | -0.001547 | -0.001547 | -3.5446
2019-04-09 14:29 | 1554834544.5217 | -337 | 0.20939 | 0.000453 | 3e-06 | 97.6506 | 97.7678 | 123.3103 | 20.4301 | 0.001848 | 0.001643 | -154.1403 | 0.22943 | 8.4022 | 0.00244 | 0.003217 | 0.00301 | 4.7e-05 | 0.001146 | 0.26952 | 1.3663 | 2.5477 | 98.56 | 20.03 | 49.995 | 65.4978 | 79.7128 | 60.5591 | 0.011894 | 1.0579 | 0.024497 | 23.1557 | -0.01073 | 0.13649 | -1.4e-05 | -0.001221 | -0.001221 | -4.5044
2019-04-09 14:29 | 1554834545.5208 | -336 | 0.20939 | 0.000455 | 3e-06 | 97.7048 | 97.9237 | 123.3514 | 20.4297 | 0.001711 | 0.001696 | -154.1379 | 0.22917 | 8.4023 | 0.002458 | 0.003852 | 0.003658 | 4.6e-05 | 0.001153 | 0.30287 | 1.3936 | 2.5477 | 98.56 | 20.03 | 49.992 | 62.196 | 80.1359 | 61.8287 | 0.011894 | 1.0539 | 0.024455 | 23.2036 | -0.020018 | 0.13634 | -1.8e-05 | -0.001483 | -0.001483 | -3.7042
2019-04-09 14:29 | 1554834546.521 | -335 | 0.20939 | 0.000452 | 2e-06 | 97.8704 | 98.0132 | 123.5638 | 20.4306 | 0.001805 | 0.001711 | -154.1252 | 0.22971 | 8.4031 | 0.002539 | 0.003526 | 0.00333 | 4.8e-05 | 0.001162 | 0.27024 | 1.3858 | 2.5477 | 98.56 | 20.03 | 49.987 | 64.4631 | 82.0949 | 62.1975 | 0.011892 | 1.0528 | 0.024463 | 23.2363 | -0.013043 | 0.13666 | -1.3e-05 | -0.001068 | -0.001068 | -5.1577
2019-04-09 14:29 | 1554834547.5208 | -334 | 0.20938 | 0.000456 | 3e-06 | 97.8093 | 97.932 | 123.8598 | 20.4312 | 0.00185 | 0.001706 | -154.1063 | 0.22843 | 8.4008 | 0.002491 | 0.00418 | 0.004035 | 5e-05 | 0.001167 | 0.2792 | 1.3779 | 2.5475 | 98.56 | 20.03 | 49.985 | 65.5331 | 80.9443 | 62.0724 | 0.011892 | 1.0532 | 0.024496 | 23.2598 | -0.011209 | 0.13589 | -1.8e-05 | -0.001527 | -0.001527 | -3.5864
2019-04-09 14:29 | 1554834548.5212 | -333 | 0.20938 | 0.000456 | 3e-06 | 97.9585 | 98.058 | 123.6094 | 20.4305 | 0.001785 | 0.001723 | -154.1407 | 0.22928 | 8.3995 | 0.002483 | 0.003059 | 0.002884 | 5.1e-05 | 0.00117 | 0.27704 | 1.3474 | 2.5476 | 98.56 | 20.03 | 49.983 | 63.9603 | 80.7418 | 62.4748 | 0.011891 | 1.0519 | 0.024457 | 23.2502 | -0.009073 | 0.1364 | -1.8e-05 | -0.00145 | -0.00145 | -3.7907
2019-04-09 14:29 | 1554834549.5209 | -332 | 0.20937 | 0.000459 | 3e-06 | 97.9023 | 98.0483 | 123.6603 | 20.4311 | 0.001773 | 0.001707 | -154.122 | 0.22862 | 8.4002 | 0.00249 | 0.003682 | 0.003611 | 5e-05 | 0.001172 | 0.28655 | 1.3396 | 2.5476 | 98.56 | 20.03 | 49.982 | 63.6822 | 80.913 | 62.1037 | 0.011891 | 1.0531 | 0.024476 | 23.2422 | -0.013335 | 0.136 | -2.2e-05 | -0.001791 | -0.001791 | -3.0606
2019-04-09 14:29 | 1554834550.521 | -331 | 0.2094 | 0.000454 | 3e-06 | 97.9237 | 98.0752 | 123.73 | 20.4318 | 0.001741 | 0.001652 | -154.1251 | 0.2288 | 8.3991 | 0.002466 | 0.003408 | 0.00331 | 4.7e-05 | 0.001173 | 0.29767 | 1.3654 | 2.5477 | 98.56 | 20.03 | 49.979 | 62.9087 | 80.3515 | 60.76 | 0.01189 | 1.0573 | 0.024532 | 23.2021 | -0.013825 | 0.13611 | -1.5e-05 | -0.001267 | -0.001267 | -4.328
2019-04-09 14:29 | 1554834551.5214 | -330 | 0.20941 | 0.000456 | 3e-06 | 98.0044 | 98.2218 | 123.745 | 20.4314 | 0.00179 | 0.001639 | -154.1137 | 0.22889 | 8.4005 | 0.002498 | 0.003684 | 0.003437 | 5.1e-05 | 0.001174 | 0.27367 | 1.3573 | 2.5476 | 98.56 | 20.03 | 49.975 | 64.1032 | 81.1029 | 60.4601 | 0.011889 | 1.0583 | 0.024544 | 23.1931 | -0.019817 | 0.13617 | -1.8e-05 | -0.001477 | -0.001477 | -3.7142
2019-04-09 14:29 | 1554834552.521 | -329 | 0.20937 | 0.000458 | 3e-06 | 98.0845 | 98.2596 | 123.4019 | 20.4314 | 0.001771 | 0.001675 | -154.0974 | 0.22988 | 8.4023 | 0.002457 | 0.003419 | 0.003285 | 5e-05 | 0.001177 | 0.27992 | 1.3992 | 2.5475 | 98.56 | 20.03 | 49.973 | 63.638 | 80.1229 | 61.3294 | 0.011889 | 1.0555 | 0.024478 | 23.191 | -0.015956 | 0.13675 | -2.1e-05 | -0.001737 | -0.001737 | -3.1725
2019-04-09 14:29 | 1554834553.5215 | -328 | 0.20939 | 0.000459 | 3e-06 | 98.1694 | 98.2149 | 123.276 | 20.431 | 0.001817 | 0.001696 | -154.0868 | 0.22857 | 8.4019 | 0.002452 | 0.003659 | 0.003457 | 5e-05 | 0.001182 | 0.29806 | 1.3494 | 2.5476 | 98.56 | 20.03 | 49.971 | 64.7361 | 80.003 | 61.8275 | 0.011889 | 1.0539 | 0.024448 | 23.1965 | -0.004146 | 0.13598 | -2.2e-05 | -0.001806 | -0.001806 | -3.0345
2019-04-09 14:29 | 1554834554.5211 | -327 | 0.20938 | 0.000458 | 2e-06 | 98.1539 | 98.2232 | 123.4514 | 20.4307 | 0.001789 | 0.001598 | -154.0594 | 0.22903 | 8.4036 | 0.002556 | 0.004153 | 0.004044 | 5e-05 | 0.001184 | 0.31454 | 1.3988 | 2.5477 | 98.56 | 20.03 | 49.965 | 64.058 | 82.5099 | 59.4742 | 0.011887 | 1.0614 | 0.024551 | 23.1313 | -0.006309 | 0.13625 | -2e-05 | -0.001619 | -0.001619 | -3.391
2019-04-09 14:29 | 1554834555.5212 | -326 | 0.20939 | 0.000458 | 3e-06 | 98.2081 | 98.3205 | 123.4972 | 20.4319 | 0.001795 | 0.00163 | -154.0696 | 0.22883 | 8.4012 | 0.002431 | 0.003584 | 0.00337 | 4.9e-05 | 0.001187 | 0.2896 | 1.3467 | 2.5476 | 98.56 | 20.03 | 49.963 | 64.2059 | 79.5041 | 60.2363 | 0.011887 | 1.059 | 0.024528 | 23.162 | -0.010236 | 0.13613 | -2e-05 | -0.001645 | -0.001645 | -3.3355
2019-04-09 14:29 | 1554834556.5202 | -325 | 0.20939 | 0.000459 | 3e-06 | 98.1593 | 98.2194 | 123.4584 | 20.4318 | 0.001853 | 0.001654 | -154.0741 | 0.22858 | 8.4029 | 0.002501 | 0.004012 | 0.003874 | 4.9e-05 | 0.001187 | 0.29374 | 1.322 | 2.5476 | 98.56 | 20.04 | 49.959 | 65.6119 | 81.186 | 60.8147 | 0.011893 | 1.0571 | 0.024503 | 23.1785 | -0.005476 | 0.13598 | -2.1e-05 | -0.001706 | -0.001706 | -3.2121
2019-04-09 14:29 | 1554834557.5221 | -324 | 0.20941 | 0.000458 | 3e-06 | 98.2551 | 98.4265 | 123.3474 | 20.4318 | 0.001775 | 0.001636 | -154.0797 | 0.2285 | 8.4013 | 0.002487 | 0.003754 | 0.00359 | 4.9e-05 | 0.001189 | 0.28385 | 1.3057 | 2.5475 | 98.56 | 20.04 | 49.956 | 63.7337 | 80.839 | 60.3793 | 0.011892 | 1.0585 | 0.024508 | 23.153 | -0.01559 | 0.13593 | -2.1e-05 | -0.001701 | -0.001701 | -3.2214
2019-04-09 14:29 | 1554834558.5203 | -323 | 0.20941 | 0.00046 | 3e-06 | 98.3156 | 98.4238 | 123.7751 | 20.4327 | 0.001765 | 0.001701 | -154.0621 | 0.22858 | 8.403 | 0.002497 | 0.003884 | 0.003774 | 4.5e-05 | 0.001188 | 0.28567 | 1.3514 | 2.5475 | 98.56 | 20.04 | 49.953 | 63.4931 | 81.0885 | 61.9433 | 0.011892 | 1.0536 | 0.024493 | 23.2474 | -0.009838 | 0.13598 | -2.3e-05 | -0.001859 | -0.001859 | -2.9471
2019-04-09 14:29 | 1554834559.5203 | -322 | 0.20938 | 0.000457 | 3e-06 | 98.3138 | 98.4513 | 123.6863 | 20.4334 | 0.00177 | 0.001672 | -154.0696 | 0.22878 | 8.4017 | 0.002473 | 0.003274 | 0.003133 | 4.7e-05 | 0.00119 | 0.29999 | 1.3368 | 2.5476 | 98.56 | 20.04 | 49.949 | 63.6218 | 80.5083 | 61.2651 | 0.011891 | 1.0557 | 0.024509 | 23.2155 | -0.012503 | 0.1361 | -1.9e-05 | -0.001586 | -0.001586 | -3.4581
2019-04-09 14:29 | 1554834560.5203 | -321 | 0.20942 | 0.000459 | 3e-06 | 98.3641 | 98.4847 | 123.6727 | 20.433 | 0.00176 | 0.001657 | -154.0849 | 0.22912 | 8.4003 | 0.002503 | 0.003986 | 0.003789 | 4.6e-05 | 0.001194 | 0.27472 | 1.3495 | 2.5476 | 98.56 | 20.04 | 49.945 | 63.3824 | 81.2338 | 60.8838 | 0.01189 | 1.0569 | 0.024521 | 23.201 | -0.010958 | 0.1363 | -2.3e-05 | -0.001865 | -0.001865 | -2.9457
2019-04-09 14:29 | 1554834561.5207 | -320 | 0.2094 | 0.00046 | 3e-06 | 98.4079 | 98.6458 | 123.7429 | 20.4327 | 0.001777 | 0.00159 | -154.0954 | 0.22937 | 8.4013 | 0.002465 | 0.00417 | 0.003962 | 4.8e-05 | 0.001193 | 0.27527 | 1.367 | 2.5476 | 98.56 | 20.04 | 49.942 | 63.7788 | 80.3222 | 59.2879 | 0.011889 | 1.062 | 0.024587 | 23.1521 | -0.021594 | 0.13645 | -2.3e-05 | -0.001877 | -0.001877 | -2.9299
2019-04-09 14:29 | 1554834562.5213 | -319 | 0.20941 | 0.000461 | 3e-06 | 98.4222 | 98.4961 | 124.0363 | 20.432 | 0.001789 | 0.001695 | -154.1141 | 0.22859 | 8.4022 | 0.002501 | 0.003868 | 0.003751 | 4.8e-05 | 0.001192 | 0.28694 | 1.3337 | 2.5476 | 98.56 | 20.04 | 49.94 | 64.0695 | 81.1768 | 61.8114 | 0.011889 | 1.054 | 0.024523 | 23.2673 | -0.00671 | 0.13599 | -2.5e-05 | -0.001991 | -0.001991 | -2.7531
2019-04-09 14:29 | 1554834563.5204 | -318 | 0.20941 | 0.000455 | 3e-06 | 98.4222 | 98.6909 | 124.11 | 20.4333 | 0.001816 | 0.001652 | -154.1278 | 0.22831 | 8.4032 | 0.002518 | 0.003327 | 0.003195 | 4.7e-05 | 0.001196 | 0.29999 | 1.3863 | 2.5476 | 98.56 | 20.04 | 49.936 | 64.7304 | 81.5881 | 60.7684 | 0.011888 | 1.0573 | 0.024569 | 23.238 | -0.024382 | 0.13583 | -1.7e-05 | -0.001439 | -0.001439 | -3.8042
2019-04-09 14:29 | 1554834564.5214 | -317 | 0.20938 | 0.000455 | 3e-06 | 98.3555 | 98.7628 | 123.8645 | 20.4318 | 0.001803 | 0.001642 | -154.1487 | 0.22918 | 8.4007 | 0.002475 | 0.003669 | 0.003442 | 4.8e-05 | 0.001198 | 0.28783 | 1.3924 | 2.5476 | 98.56 | 20.04 | 49.934 | 64.4096 | 80.5488 | 60.534 | 0.011887 | 1.058 | 0.024553 | 23.2068 | -0.036963 | 0.13634 | -1.7e-05 | -0.001407 | -0.001407 | -3.9053
2019-04-09 14:29 | 1554834565.5204 | -316 | 0.20943 | 0.00046 | 2e-06 | 98.3915 | 98.6379 | 123.9051 | 20.434 | 0.001819 | 0.001658 | -154.1318 | 0.22914 | 8.3993 | 0.002534 | 0.003811 | 0.003705 | 4.8e-05 | 0.001197 | 0.27295 | 1.3316 | 2.5476 | 98.56 | 20.04 | 49.931 | 64.7975 | 81.971 | 60.9262 | 0.011886 | 1.0568 | 0.024543 | 23.2243 | -0.022365 | 0.13632 | -2.2e-05 | -0.001799 | -0.001799 | -3.0533
2019-04-09 14:29 | 1554834566.5215 | -315 | 0.20943 | 0.000457 | 3e-06 | 98.4493 | 98.6613 | 123.9612 | 20.4334 | 0.001773 | 0.001684 | -154.1497 | 0.22958 | 8.4 | 0.002552 | 0.004465 | 0.004228 | 5e-05 | 0.001195 | 0.27793 | 1.3296 | 2.5476 | 98.56 | 20.04 | 49.928 | 63.6852 | 82.4128 | 61.5322 | 0.011886 | 1.0549 | 0.024526 | 23.2506 | -0.019237 | 0.13658 | -2e-05 | -0.001627 | -0.001627 | -3.3835
2019-04-09 14:29 | 1554834567.5218 | -314 | 0.20945 | 0.000454 | 3e-06 | 98.4392 | 98.6478 | 123.8952 | 20.434 | 0.001756 | 0.001659 | -154.1317 | 0.22926 | 8.4034 | 0.00254 | 0.004311 | 0.004172 | 5.2e-05 | 0.001198 | 0.27848 | 1.3419 | 2.5474 | 98.56 | 20.04 | 49.929 | 63.2749 | 82.1152 | 60.9513 | 0.011886 | 1.0567 | 0.024541 | 23.2242 | -0.01894 | 0.13639 | -1.6e-05 | -0.001375 | -0.001375 | -3.9977
2019-04-09 14:29 | 1554834568.5219 | -313 | 0.2094 | 0.000452 | 3e-06 | 98.5226 | 98.5938 | 123.8093 | 20.4334 | 0.001801 | 0.001686 | -154.1336 | 0.22889 | 8.4023 | 0.002534 | 0.004302 | 0.004176 | 4.4e-05 | 0.001205 | 0.27505 | 1.341 | 2.5474 | 98.56 | 20.04 | 49.928 | 64.3613 | 81.9759 | 61.6018 | 0.011886 | 1.0546 | 0.024509 | 23.2387 | -0.006466 | 0.13617 | -1.3e-05 | -0.001108 | -0.001108 | -4.953
2019-04-09 14:29 | 1554834569.5209 | -312 | 0.20942 | 0.000449 | 3e-06 | 98.4966 | 98.632 | 123.7347 | 20.434 | 0.001828 | 0.001719 | -154.1203 | 0.22923 | 8.4034 | 0.00254 | 0.003826 | 0.003643 | 4.7e-05 | 0.001202 | 0.26432 | 1.3338 | 2.5474 | 98.56 | 20.04 | 49.929 | 65.0027 | 82.1092 | 62.386 | 0.011886 | 1.0522 | 0.024473 | 23.2589 | -0.012283 | 0.13637 | -9e-06 | -0.000839 | -0.000839 | -6.5505
2019-04-09 14:29 | 1554834570.5215 | -311 | 0.20943 | 0.000454 | 2e-06 | 98.5023 | 98.5869 | 123.6917 | 20.4339 | 0.001768 | 0.001684 | -154.0964 | 0.22888 | 8.3983 | 0.002506 | 0.00381 | 0.003644 | 4.9e-05 | 0.001202 | 0.27721 | 1.3564 | 2.5476 | 98.56 | 20.04 | 49.929 | 63.5701 | 81.2947 | 61.5334 | 0.011886 | 1.0549 | 0.0245 | 23.2253 | -0.00768 | 0.13616 | -1.6e-05 | -0.001289 | -0.001289 | -4.2562
2019-04-09 14:29 | 1554834571.5206 | -310 | 0.20947 | 0.000452 | 3e-06 | 98.6628 | 98.8719 | 123.6727 | 20.4342 | 0.001778 | 0.00165 | -154.0858 | 0.22943 | 8.4011 | 0.002478 | 0.004023 | 0.003871 | 4.8e-05 | 0.001198 | 0.28567 | 1.3662 | 2.5476 | 98.56 | 20.04 | 49.932 | 63.8139 | 80.6199 | 60.7132 | 0.011887 | 1.0575 | 0.024528 | 23.1951 | -0.018933 | 0.13649 | -1.2e-05 | -0.001072 | -0.001072 | -5.1316
2019-04-09 14:30 | 1554834572.521 | -309 | 0.20946 | 0.000452 | 3e-06 | 98.6116 | 98.7642 | 123.6216 | 20.4346 | 0.001748 | 0.001661 | -154.0803 | 0.22934 | 8.4042 | 0.002552 | 0.003721 | 0.003578 | 4.5e-05 | 0.001196 | 0.27887 | 1.3195 | 2.5475 | 98.56 | 20.04 | 49.934 | 63.0944 | 82.4006 | 60.9964 | 0.011887 | 1.0566 | 0.024512 | 23.2001 | -0.013828 | 0.13643 | -1.3e-05 | -0.001094 | -0.001094 | -5.0248
2019-04-09 14:30 | 1554834573.5196 | -308 | 0.20945 | 0.000454 | 3e-06 | 98.6315 | 98.9352 | 123.8659 | 20.435 | 0.0018 | 0.001667 | -154.0873 | 0.2301 | 8.401 | 0.002478 | 0.004178 | 0.004021 | 4.9e-05 | 0.001195 | 0.29015 | 1.3233 | 2.5474 | 98.56 | 20.04 | 49.937 | 64.3456 | 80.6307 | 61.1406 | 0.011888 | 1.0561 | 0.024531 | 23.228 | -0.027495 | 0.13689 | -1.6e-05 | -0.00137 | -0.00137 | -4.0278
2019-04-09 14:30 | 1554834574.5207 | -307 | 0.20943 | 0.000456 | 3e-06 | 98.625 | 98.813 | 124.1661 | 20.4342 | 0.00178 | 0.001667 | -154.0776 | 0.22945 | 8.4035 | 0.002473 | 0.003418 | 0.003286 | 4.9e-05 | 0.001188 | 0.27759 | 1.3524 | 2.5476 | 98.56 | 20.04 | 49.941 | 63.8599 | 80.5053 | 61.1447 | 0.011889 | 1.0561 | 0.024561 | 23.2563 | -0.017037 | 0.1365 | -1.8e-05 | -0.001518 | -0.001518 | -3.6227
2019-04-09 14:30 | 1554834575.5207 | -306 | 0.20949 | 0.000448 | 2e-06 | 98.6902 | 98.747 | 124.3168 | 20.436 | 0.001755 | 0.001674 | -154.0975 | 0.22926 | 8.401 | 0.002488 | 0.004184 | 0.003947 | 4.4e-05 | 0.001186 | 0.27041 | 1.3126 | 2.5474 | 98.56 | 20.04 | 49.942 | 63.2643 | 80.8757 | 61.3035 | 0.011889 | 1.0556 | 0.02457 | 23.2759 | -0.005141 | 0.13639 | -7e-06 | -0.000693 | -0.000693 | -7.9344
2019-04-09 14:30 | 1554834576.5202 | -305 | 0.20942 | 0.000456 | 2e-06 | 98.639 | 98.8419 | 124.0929 | 20.4364 | 0.00187 | 0.001687 | -154.0687 | 0.22886 | 8.402 | 0.002483 | 0.003624 | 0.003466 | 4.5e-05 | 0.001194 | 0.29767 | 1.3153 | 2.5475 | 98.56 | 20.04 | 49.945 | 66.0103 | 80.7576 | 61.6117 | 0.01189 | 1.0546 | 0.024536 | 23.2657 | -0.018384 | 0.13615 | -1.6e-05 | -0.001339 | -0.001339 | -4.0977
2019-04-09 14:30 | 1554834577.5212 | -304 | 0.20948 | 0.000451 | 3e-06 | 98.6944 | 98.9349 | 123.7328 | 20.4358 | 0.001829 | 0.001659 | -154.0882 | 0.22958 | 8.4015 | 0.002449 | 0.003809 | 0.003722 | 4.7e-05 | 0.001199 | 0.27992 | 1.3307 | 2.5475 | 98.56 | 20.05 | 49.948 | 65.0232 | 79.9336 | 60.9383 | 0.011898 | 1.0567 | 0.024525 | 23.2085 | -0.021768 | 0.13658 | -1.2e-05 | -0.00107 | -0.00107 | -5.1462
2019-04-09 14:30 | 1554834578.52 | -303 | 0.20946 | 0.000449 | 3e-06 | 98.6539 | 98.8832 | 123.4145 | 20.4348 | 0.001744 | 0.00165 | -154.0761 | 0.22959 | 8.402 | 0.002509 | 0.004048 | 0.003846 | 4.8e-05 | 0.001204 | 0.27417 | 1.3944 | 2.5476 | 98.56 | 20.05 | 49.947 | 62.9798 | 81.3807 | 60.7159 | 0.011898 | 1.0574 | 0.024502 | 23.1709 | -0.020771 | 0.13658 | -8e-06 | -0.000764 | -0.000764 | -7.208
2019-04-09 14:30 | 1554834579.5212 | -302 | 0.20946 | 0.000455 | 2e-06 | 98.7623 | 99.1627 | 123.607 | 20.4359 | 0.001845 | 0.001659 | -154.0952 | 0.22915 | 8.4007 | 0.002519 | 0.003365 | 0.003241 | 4.7e-05 | 0.001209 | 0.30381 | 1.3418 | 2.5475 | 98.56 | 20.05 | 49.944 | 65.4145 | 81.6243 | 60.9447 | 0.011897 | 1.0567 | 0.024513 | 23.1969 | -0.03619 | 0.13632 | -1.6e-05 | -0.001343 | -0.001343 | -4.0917
2019-04-09 14:30 | 1554834580.5199 | -301 | 0.20948 | 0.000454 | 3e-06 | 98.7364 | 99.1135 | 123.6535 | 20.4364 | 0.001789 | 0.001698 | -154.1104 | 0.22917 | 8.4006 | 0.002479 | 0.004277 | 0.003973 | 4.4e-05 | 0.001209 | 0.28799 | 1.3772 | 2.5476 | 98.56 | 20.05 | 49.942 | 64.0801 | 80.6634 | 61.8906 | 0.011896 | 1.0537 | 0.024483 | 23.2341 | -0.034097 | 0.13633 | -1.5e-05 | -0.001279 | -0.001279 | -4.295
2019-04-09 14:30 | 1554834581.52 | -300 | 0.20947 | 0.000453 | 3e-06 | 98.7283 | 98.9713 | 123.5899 | 20.436 | 0.001762 | 0.001645 | -154.1176 | 0.22923 | 8.4015 | 0.002518 | 0.004287 | 0.004197 | 4.6e-05 | 0.001206 | 0.2891 | 1.3507 | 2.5476 | 98.56 | 20.05 | 49.94 | 63.4306 | 81.5994 | 60.6014 | 0.011896 | 1.0578 | 0.024524 | 23.1834 | -0.021989 | 0.13637 | -1.4e-05 | -0.001217 | -0.001217 | -4.5144
2019-04-09 14:30 | 1554834582.5209 | -299 | 0.20947 | 0.000457 | 3e-06 | 98.742 | 99.0629 | 123.4648 | 20.4361 | 0.001783 | 0.001624 | -154.1265 | 0.22915 | 8.402 | 0.00251 | 0.003456 | 0.003371 | 4.7e-05 | 0.001201 | 0.29535 | 1.3551 | 2.5476 | 98.56 | 20.05 | 49.938 | 63.9282 | 81.4115 | 60.0996 | 0.011895 | 1.0594 | 0.02453 | 23.1543 | -0.029017 | 0.13632 | -1.9e-05 | -0.00159 | -0.00159 | -3.4558
2019-04-09 14:30 | 1554834583.52 | -298 | 0.20945 | 0.000452 | 3e-06 | 98.7891 | 99.0347 | 123.1737 | 20.4366 | 0.001833 | 0.001652 | -154.0977 | 0.22941 | 8.4025 | 0.002453 | 0.003821 | 0.003655 | 4.4e-05 | 0.001201 | 0.2625 | 1.3458 | 2.5475 | 98.56 | 20.05 | 49.937 | 65.138 | 80.0216 | 60.7702 | 0.011895 | 1.0573 | 0.024476 | 23.1502 | -0.022208 | 0.13648 | -1.3e-05 | -0.001108 | -0.001108 | -4.9623
2019-04-09 14:30 | 1554834584.5201 | -297 | 0.20947 | 0.000453 | 3e-06 | 98.7563 | 99.0181 | 123.0631 | 20.4373 | 0.001794 | 0.001741 | -154.1084 | 0.22925 | 8.4024 | 0.002534 | 0.003347 | 0.003163 | 4.5e-05 | 0.001204 | 0.286 | 1.356 | 2.5473 | 98.56 | 20.04 | 49.932 | 64.2035 | 81.968 | 62.9196 | 0.011887 | 1.0505 | 0.024387 | 23.2142 | -0.023682 | 0.13638 | -1.4e-05 | -0.001153 | -0.001153 | -4.7687
2019-04-09 14:30 | 1554834585.521 | -296 | 0.20943 | 0.000453 | 3e-06 | 98.7918 | 99.0643 | 122.9244 | 20.4373 | 0.001832 | 0.001708 | -154.0857 | 0.22964 | 8.4014 | 0.002442 | 0.00404 | 0.003799 | 4.7e-05 | 0.001199 | 0.30917 | 1.3493 | 2.5474 | 98.56 | 20.04 | 49.931 | 65.1113 | 79.7635 | 62.1251 | 0.011886 | 1.053 | 0.024402 | 23.1736 | -0.024638 | 0.13661 | -1.4e-05 | -0.001209 | -0.001209 | -4.5525
2019-04-09 14:30 | 1554834586.5205 | -295 | 0.20946 | 0.000451 | 3e-06 | 98.8266 | 99.0071 | 123.1221 | 20.4373 | 0.001855 | 0.001653 | -154.1151 | 0.22905 | 8.4013 | 0.002407 | 0.003574 | 0.00342 | 4.7e-05 | 0.001191 | 0.27273 | 1.3226 | 2.5474 | 98.56 | 20.04 | 49.93 | 65.6639 | 78.9264 | 60.7967 | 0.011886 | 1.0572 | 0.02447 | 23.1463 | -0.016325 | 0.13626 | -1.2e-05 | -0.001044 | -0.001044 | -5.2617
2019-04-09 14:30 | 1554834587.5197 | -294 | 0.20942 | 0.000453 | 2e-06 | 98.8418 | 99.0033 | 123.3929 | 20.4373 | 0.001786 | 0.00166 | -154.1096 | 0.22926 | 8.403 | 0.002434 | 0.003971 | 0.003878 | 4.4e-05 | 0.001188 | 0.28854 | 1.3403 | 2.5476 | 98.56 | 20.05 | 49.929 | 63.9912 | 79.5835 | 60.9795 | 0.011893 | 1.0566 | 0.02449 | 23.1781 | -0.014608 | 0.13639 | -1.4e-05 | -0.001174 | -0.001174 | -4.681
2019-04-09 14:30 | 1554834588.5196 | -293 | 0.20941 | 0.000451 | 3e-06 | 98.8772 | 99.2047 | 123.6978 | 20.4375 | 0.0018 | 0.001645 | -154.1253 | 0.2297 | 8.4019 | 0.002403 | 0.00356 | 0.00335 | 4.7e-05 | 0.001187 | 0.2724 | 1.3121 | 2.5475 | 98.56 | 20.04 | 49.93 | 64.3466 | 78.8203 | 60.5966 | 0.011886 | 1.0578 | 0.024535 | 23.1934 | -0.029572 | 0.13664 | -1.1e-05 | -0.001037 | -0.001037 | -5.3089
2019-04-09 14:30 | 1554834589.5201 | -292 | 0.20943 | 0.000451 | 3e-06 | 98.8379 | 99.1179 | 123.6448 | 20.4382 | 0.00179 | 0.001632 | -154.1197 | 0.22936 | 8.4026 | 0.002506 | 0.003912 | 0.003748 | 4.4e-05 | 0.001185 | 0.28799 | 1.317 | 2.5477 | 98.56 | 20.04 | 49.931 | 64.0912 | 81.2967 | 60.3058 | 0.011886 | 1.0587 | 0.02454 | 23.1783 | -0.025306 | 0.13645 | -1.2e-05 | -0.001021 | -0.001021 | -5.3843
2019-04-09 14:30 | 1554834590.5196 | -291 | 0.20945 | 0.000449 | 3e-06 | 98.8212 | 99.0457 | 123.4209 | 20.4377 | 0.001811 | 0.001653 | -154.1159 | 0.2295 | 8.4026 | 0.002394 | 0.003219 | 0.003061 | 4.7e-05 | 0.001182 | 0.28368 | 1.3452 | 2.5477 | 98.56 | 20.04 | 49.932 | 64.6105 | 78.6194 | 60.7908 | 0.011887 | 1.0572 | 0.0245 | 23.1741 | -0.020291 | 0.13653 | -9e-06 | -0.000838 | -0.000838 | -6.5626
2019-04-09 14:30 | 1554834591.5193 | -290 | 0.20943 | 0.000455 | 3e-06 | 98.8993 | 99.0773 | 123.4504 | 20.4379 | 0.001804 | 0.001686 | -154.114 | 0.22925 | 8.4009 | 0.002371 | 0.00411 | 0.003967 | 4.4e-05 | 0.001178 | 0.29717 | 1.3423 | 2.5477 | 98.56 | 20.05 | 49.934 | 64.4344 | 78.0732 | 61.5961 | 0.011895 | 1.0547 | 0.024473 | 23.2048 | -0.016091 | 0.13638 | -1.7e-05 | -0.001406 | -0.001406 | -3.9097
2019-04-09 14:30 | 1554834592.5203 | -289 | 0.20947 | 0.000453 | 3e-06 | 98.8748 | 99.0945 | 123.3073 | 20.4388 | 0.001746 | 0.001667 | -154.1062 | 0.22942 | 8.4012 | 0.002437 | 0.003463 | 0.003265 | 4.5e-05 | 0.00117 | 0.29911 | 1.3577 | 2.5475 | 98.56 | 20.04 | 49.938 | 63.0515 | 79.6451 | 61.1311 | 0.011888 | 1.0561 | 0.024476 | 23.1753 | -0.019853 | 0.13648 | -1.4e-05 | -0.001172 | -0.001172 | -4.6911
2019-04-09 14:30 | 1554834593.5194 | -288 | 0.20941 | 0.000454 | 2e-06 | 98.8701 | 99.2436 | 123.5091 | 20.4383 | 0.001797 | 0.001653 | -154.1301 | 0.22957 | 8.4001 | 0.002435 | 0.003622 | 0.003473 | 4.5e-05 | 0.001171 | 0.29645 | 1.358 | 2.5476 | 98.56 | 20.04 | 49.94 | 64.2701 | 79.6085 | 60.8015 | 0.011889 | 1.0572 | 0.024508 | 23.1828 | -0.033725 | 0.13657 | -1.5e-05 | -0.001239 | -0.001239 | -4.4423
2019-04-09 14:30 | 1554834594.5194 | -287 | 0.20944 | 0.000453 | 3e-06 | 98.8912 | 98.9927 | 123.6103 | 20.439 | 0.001837 | 0.001655 | -154.1203 | 0.23 | 8.3982 | 0.002492 | 0.003512 | 0.003416 | 4.6e-05 | 0.00117 | 0.29607 | 1.324 | 2.5476 | 98.56 | 20.04 | 49.943 | 65.2301 | 80.9793 | 60.8499 | 0.011889 | 1.057 | 0.024517 | 23.194 | -0.009172 | 0.13683 | -1.4e-05 | -0.001192 | -0.001192 | -4.6258
2019-04-09 14:30 | 1554834595.5204 | -286 | 0.20944 | 0.000454 | 2e-06 | 98.9034 | 99.2061 | 123.4516 | 20.4388 | 0.001752 | 0.001651 | -154.1354 | 0.22907 | 8.4014 | 0.002471 | 0.003874 | 0.003657 | 4.6e-05 | 0.00117 | 0.32084 | 1.3562 | 2.5476 | 98.56 | 20.05 | 49.945 | 63.1953 | 80.4605 | 60.7488 | 0.011897 | 1.0573 | 0.024505 | 23.1756 | -0.027326 | 0.13627 | -1.5e-05 | -0.001224 | -0.001224 | -4.4867
2019-04-09 14:30 | 1554834596.5194 | -285 | 0.20947 | 0.000452 | 3e-06 | 98.9174 | 99.2628 | 123.3251 | 20.4385 | 0.001764 | 0.0017 | -154.1211 | 0.22925 | 8.4015 | 0.00242 | 0.003773 | 0.003608 | 4.3e-05 | 0.001174 | 0.27312 | 1.3486 | 2.5476 | 98.56 | 20.05 | 49.948 | 63.4848 | 79.2322 | 61.9319 | 0.011898 | 1.0536 | 0.024449 | 23.2047 | -0.031179 | 0.13638 | -1.3e-05 | -0.001153 | -0.001153 | -4.7664
2019-04-09 14:30 | 1554834597.5196 | -284 | 0.20943 | 0.000452 | 2e-06 | 98.9064 | 99.2219 | 123.2638 | 20.4396 | 0.00184 | 0.001654 | -154.1223 | 0.22885 | 8.4028 | 0.00248 | 0.003475 | 0.003351 | 4.9e-05 | 0.001178 | 0.28998 | 1.3272 | 2.5476 | 98.56 | 20.05 | 49.952 | 65.3086 | 80.696 | 60.8292 | 0.011899 | 1.0571 | 0.024483 | 23.1607 | -0.028484 | 0.13614 | -1.2e-05 | -0.001044 | -0.001044 | -5.2566
2019-04-09 14:30 | 1554834598.5207 | -283 | 0.20948 | 0.000449 | 2e-06 | 98.9797 | 99.2666 | 123.2246 | 20.4396 | 0.001801 | 0.001708 | -154.1164 | 0.22919 | 8.4017 | 0.002453 | 0.003783 | 0.0036 | 4.7e-05 | 0.001176 | 0.27505 | 1.3097 | 2.5475 | 98.56 | 20.05 | 49.954 | 64.3657 | 80.0309 | 62.115 | 0.011899 | 1.053 | 0.024432 | 23.2016 | -0.025894 | 0.13635 | -9e-06 | -0.000778 | -0.000778 | -7.0631
2019-04-09 14:30 | 1554834599.5195 | -282 | 0.20948 | 0.000453 | 3e-06 | 98.9829 | 99.2267 | 123.2521 | 20.4381 | 0.001771 | 0.00167 | -154.1161 | 0.22995 | 8.4 | 0.00246 | 0.003111 | 0.002964 | 4.4e-05 | 0.00118 | 0.29175 | 1.3318 | 2.5475 | 98.56 | 20.05 | 49.959 | 63.6375 | 80.201 | 61.2116 | 0.0119 | 1.0559 | 0.024468 | 23.1729 | -0.022 | 0.1368 | -1.4e-05 | -0.001207 | -0.001207 | -4.5666
2019-04-09 14:30 | 1554834600.5199 | -281 | 0.20942 | 0.000453 | 3e-06 | 98.9049 | 99.195 | 123.5446 | 20.4406 | 0.001753 | 0.001666 | -154.1293 | 0.22948 | 8.4032 | 0.00251 | 0.003796 | 0.00355 | 4.6e-05 | 0.001183 | 0.27743 | 1.3471 | 2.5475 | 98.56 | 20.05 | 49.961 | 63.1994 | 81.4105 | 61.1147 | 0.011901 | 1.0562 | 0.0245 | 23.197 | -0.026199 | 0.13652 | -1.4e-05 | -0.001167 | -0.001167 | -4.7136
2019-04-09 14:30 | 1554834601.5206 | -280 | 0.20946 | 0.000448 | 3e-06 | 99.0464 | 99.1981 | 123.2753 | 20.4391 | 0.001812 | 0.00165 | -154.1251 | 0.22919 | 8.4041 | 0.002446 | 0.003706 | 0.003548 | 4.4e-05 | 0.001175 | 0.29518 | 1.3125 | 2.5475 | 98.56 | 20.05 | 49.963 | 64.633 | 79.8549 | 60.7305 | 0.011901 | 1.0574 | 0.024488 | 23.1584 | -0.013695 | 0.13634 | -8e-06 | -0.000773 | -0.000773 | -7.1064
2019-04-09 14:30 | 1554834602.52 | -279 | 0.20945 | 0.000454 | 2e-06 | 99.0196 | 99.29 | 123.5051 | 20.44 | 0.001755 | 0.001677 | -154.1424 | 0.22904 | 8.3992 | 0.002466 | 0.003775 | 0.003608 | 4.4e-05 | 0.001171 | 0.28584 | 1.3503 | 2.5474 | 98.56 | 20.05 | 49.967 | 63.2535 | 80.3539 | 61.3847 | 0.011902 | 1.0553 | 0.024487 | 23.2026 | -0.024397 | 0.13626 | -1.5e-05 | -0.001273 | -0.001273 | -4.3133
2019-04-09 14:30 | 1554834603.5196 | -278 | 0.20944 | 0.000454 | 3e-06 | 99.0097 | 99.2591 | 123.1526 | 20.4403 | 0.001826 | 0.001655 | -154.1155 | 0.22999 | 8.4003 | 0.00239 | 0.003548 | 0.003436 | 4.4e-05 | 0.001163 | 0.30342 | 1.3339 | 2.5475 | 98.56 | 20.05 | 49.967 | 64.9612 | 78.5315 | 60.8605 | 0.011902 | 1.057 | 0.024471 | 23.1514 | -0.022494 | 0.13682 | -1.6e-05 | -0.001336 | -0.001336 | -4.1265
2019-04-09 14:30 | 1554834604.5194 | -277 | 0.20946 | 0.000454 | 3e-06 | 99.0297 | 99.2735 | 123.4209 | 20.4402 | 0.001778 | 0.001696 | -154.1306 | 0.22926 | 8.4008 | 0.002442 | 0.004267 | 0.004053 | 4.6e-05 | 0.001161 | 0.2907 | 1.3528 | 2.5476 | 98.56 | 20.05 | 49.971 | 63.8048 | 79.775 | 61.8288 | 0.011903 | 1.0539 | 0.024462 | 23.2101 | -0.021994 | 0.13639 | -1.6e-05 | -0.001351 | -0.001351 | -4.0693
2019-04-09 14:30 | 1554834605.5197 | -276 | 0.20947 | 0.00045 | 3e-06 | 99.0371 | 99.3423 | 123.4697 | 20.4402 | 0.00176 | 0.001671 | -154.1185 | 0.22924 | 8.4023 | 0.0025 | 0.00403 | 0.003927 | 4.7e-05 | 0.001156 | 0.28047 | 1.4152 | 2.5476 | 98.56 | 20.05 | 49.973 | 63.3724 | 81.1629 | 61.2334 | 0.011904 | 1.0558 | 0.024489 | 23.1941 | -0.027522 | 0.13637 | -1e-05 | -0.000909 | -0.000909 | -6.0482
2019-04-09 14:30 | 1554834606.5187 | -275 | 0.20949 | 0.00045 | 3e-06 | 98.9341 | 99.1923 | 123.3007 | 20.4415 | 0.001782 | 0.0017 | -154.1226 | 0.22966 | 8.4017 | 0.002479 | 0.002851 | 0.002659 | 4.8e-05 | 0.00116 | 0.28694 | 1.3626 | 2.5475 | 98.56 | 20.05 | 49.976 | 63.9132 | 80.651 | 61.93 | 0.011905 | 1.0536 | 0.024446 | 23.2023 | -0.023312 | 0.13662 | -1e-05 | -0.00094 | -0.00094 | -5.8577
2019-04-09 14:30 | 1554834607.52 | -274 | 0.20942 | 0.000448 | 3e-06 | 99.0249 | 99.4335 | 123.404 | 20.4409 | 0.001803 | 0.001637 | -154.1325 | 0.23018 | 8.4024 | 0.002432 | 0.004643 | 0.004392 | 4.6e-05 | 0.001154 | 0.286 | 1.3326 | 2.5476 | 98.56 | 20.05 | 49.98 | 64.4114 | 79.5356 | 60.4062 | 0.011905 | 1.0584 | 0.024512 | 23.1592 | -0.036832 | 0.13693 | -8e-06 | -0.000759 | -0.000759 | -7.2728
2019-04-09 14:30 | 1554834608.5188 | -273 | 0.20944 | 0.000448 | 3e-06 | 99.0502 | 99.3516 | 123.3819 | 20.4415 | 0.001836 | 0.001595 | -154.1242 | 0.22956 | 8.4052 | 0.002487 | 0.004197 | 0.004032 | 4.3e-05 | 0.001157 | 0.28241 | 1.4046 | 2.5476 | 98.56 | 20.05 | 49.982 | 65.22 | 80.8521 | 59.4104 | 0.011906 | 1.0616 | 0.024547 | 23.1225 | -0.027175 | 0.13656 | -7e-06 | -0.00067 | -0.00067 | -8.2177
2019-04-09 14:30 | 1554834609.5188 | -272 | 0.20944 | 0.000448 | 3e-06 | 99.0636 | 99.4387 | 123.0897 | 20.4409 | 0.001801 | 0.001688 | -154.12 | 0.22928 | 8.4023 | 0.0025 | 0.003263 | 0.003185 | 4.6e-05 | 0.001146 | 0.27345 | 1.3743 | 2.5477 | 98.56 | 20.05 | 49.986 | 64.3771 | 81.1597 | 61.6394 | 0.011907 | 1.0545 | 0.024436 | 23.1724 | -0.0338 | 0.1364 | -8e-06 | -0.000722 | -0.000722 | -7.6152
2019-04-09 14:30 | 1554834610.5203 | -271 | 0.20942 | 0.000446 | 3e-06 | 99.1134 | 99.4225 | 123.0401 | 20.4408 | 0.001845 | 0.001683 | -154.0869 | 0.22917 | 8.4001 | 0.002486 | 0.002998 | 0.002836 | 4.8e-05 | 0.001147 | 0.30342 | 1.4085 | 2.5477 | 98.56 | 20.05 | 49.99 | 65.4356 | 80.822 | 61.5348 | 0.011908 | 1.0549 | 0.024435 | 23.1641 | -0.027856 | 0.13634 | -5e-06 | -0.000546 | -0.000546 | -10.0622
2019-04-09 14:30 | 1554834611.52 | -270 | 0.20943 | 0.000448 | 3e-06 | 99.0988 | 99.4666 | 123.1282 | 20.441 | 0.001759 | 0.001693 | -154.0846 | 0.22968 | 8.4008 | 0.002474 | 0.004001 | 0.003808 | 4.8e-05 | 0.00114 | 0.28495 | 1.3655 | 2.5477 | 98.56 | 20.05 | 49.993 | 63.3606 | 80.5383 | 61.7717 | 0.011909 | 1.0541 | 0.024435 | 23.1806 | -0.033135 | 0.13664 | -7e-06 | -0.000681 | -0.000681 | -8.0909
2019-04-09 14:30 | 1554834612.5193 | -269 | 0.20945 | 0.000443 | 3e-06 | 99.0809 | 99.4917 | 123.1833 | 20.4427 | 0.001839 | 0.001689 | -154.0573 | 0.22947 | 8.4006 | 0.002485 | 0.004187 | 0.004008 | 4.7e-05 | 0.001143 | 0.30215 | 1.3189 | 2.5475 | 98.56 | 20.05 | 49.999 | 65.2859 | 80.8081 | 61.6725 | 0.01191 | 1.0544 | 0.024444 | 23.1824 | -0.037007 | 0.13651 | -2e-06 | -0.000287 | -0.000287 | -19.1458
2019-04-09 14:30 | 1554834613.519 | -268 | 0.20944 | 0.000449 | 3e-06 | 99.0744 | 99.1637 | 123.4526 | 20.4422 | 0.001755 | 0.001662 | -154.0743 | 0.22966 | 8.4 | 0.002487 | 0.003573 | 0.003438 | 4.9e-05 | 0.001144 | 0.27743 | 1.3749 | 2.5477 | 98.56 | 20.05 | 50.003 | 63.264 | 80.8466 | 61.0201 | 0.011911 | 1.0565 | 0.024495 | 23.1851 | -0.008064 | 0.13662 | -9e-06 | -0.000817 | -0.000817 | -6.7402
2019-04-09 14:30 | 1554834614.5194 | -267 | 0.20944 | 0.00045 | 3e-06 | 99.0821 | 99.1892 | 122.8909 | 20.4428 | 0.001753 | 0.001649 | -154.0518 | 0.22885 | 8.4024 | 0.002518 | 0.003778 | 0.003637 | 4.8e-05 | 0.001143 | 0.28495 | 1.3518 | 2.5476 | 98.56 | 20.05 | 50.005 | 63.2224 | 81.6035 | 60.7058 | 0.011911 | 1.0575 | 0.02445 | 23.1214 | -0.009662 | 0.13614 | -1.1e-05 | -0.000936 | -0.000936 | -5.8607
2019-04-09 14:30 | 1554834615.52 | -266 | 0.20945 | 0.000448 | 2e-06 | 99.0851 | 99.3867 | 123.0862 | 20.4421 | 0.001793 | 0.001688 | -154.0764 | 0.22999 | 8.4034 | 0.002455 | 0.004404 | 0.004341 | 4.5e-05 | 0.001155 | 0.28871 | 1.3677 | 2.5476 | 98.56 | 20.05 | 50.009 | 64.1768 | 80.087 | 61.6454 | 0.011912 | 1.0545 | 0.024435 | 23.1723 | -0.027189 | 0.13682 | -7e-06 | -0.000655 | -0.000655 | -8.4216
2019-04-09 14:30 | 1554834616.5201 | -265 | 0.20944 | 0.000449 | 2e-06 | 99.136 | 99.3792 | 123.2981 | 20.4423 | 0.001772 | 0.001687 | -154.0669 | 0.22941 | 8.4014 | 0.00243 | 0.003453 | 0.003292 | 4.5e-05 | 0.001156 | 0.30718 | 1.3634 | 2.5475 | 98.56 | 20.05 | 50.01 | 63.6587 | 79.4947 | 61.6191 | 0.011913 | 1.0546 | 0.024457 | 23.1913 | -0.021912 | 0.13648 | -8e-06 | -0.000737 | -0.000737 | -7.4626
2019-04-09 14:30 | 1554834617.5201 | -264 | 0.20945 | 0.000448 | 2e-06 | 99.0517 | 99.3289 | 123.3108 | 20.4427 | 0.001779 | 0.001689 | -154.0644 | 0.22884 | 8.4018 | 0.00253 | 0.00389 | 0.003693 | 4.6e-05 | 0.001153 | 0.30502 | 1.3612 | 2.5476 | 98.56 | 20.05 | 50.011 | 63.8323 | 81.8953 | 61.6631 | 0.011913 | 1.0545 | 0.024457 | 23.194 | -0.024995 | 0.13614 | -8e-06 | -0.000715 | -0.000715 | -7.6701
2019-04-09 14:30 | 1554834618.5204 | -263 | 0.20944 | 0.00045 | 2e-06 | 99.0756 | 99.1327 | 123.0704 | 20.4433 | 0.001819 | 0.001681 | -154.0678 | 0.22961 | 8.4023 | 0.002441 | 0.003862 | 0.003664 | 4.6e-05 | 0.001157 | 0.28191 | 1.3269 | 2.5476 | 98.56 | 20.05 | 50.012 | 64.7901 | 79.7561 | 61.4746 | 0.011913 | 1.055 | 0.02444 | 23.1649 | -0.005162 | 0.13659 | -1e-05 | -0.00088 | -0.00088 | -6.2581
2019-04-09 14:30 | 1554834619.5185 | -262 | 0.20944 | 0.000449 | 2e-06 | 99.0306 | 99.418 | 123.6037 | 20.4432 | 0.001789 | 0.001692 | -154.075 | 0.22927 | 8.401 | 0.002463 | 0.004242 | 0.004166 | 4.5e-05 | 0.00116 | 0.30646 | 1.3628 | 2.5476 | 98.56 | 20.05 | 50.011 | 64.0763 | 80.2861 | 61.7339 | 0.011913 | 1.0542 | 0.024484 | 23.224 | -0.034928 | 0.13639 | -8e-06 | -0.000737 | -0.000737 | -7.4616
2019-04-09 14:30 | 1554834620.5192 | -261 | 0.20945 | 0.000448 | 2e-06 | 99.013 | 99.3238 | 123.7575 | 20.444 | 0.001803 | 0.001704 | -154.1112 | 0.22957 | 8.4004 | 0.002486 | 0.003872 | 0.003727 | 4.5e-05 | 0.001152 | 0.29535 | 1.3791 | 2.5477 | 98.56 | 20.05 | 50.01 | 64.4138 | 80.8359 | 62.0302 | 0.011913 | 1.0533 | 0.024488 | 23.2487 | -0.028027 | 0.13657 | -7e-06 | -0.000653 | -0.000653 | -8.4297
2019-04-09 14:30 | 1554834621.5186 | -260 | 0.20943 | 0.000447 | 2e-06 | 99.0642 | 99.3499 | 123.9553 | 20.4438 | 0.001754 | 0.001671 | -154.1043 | 0.22975 | 8.4017 | 0.002412 | 0.00432 | 0.004153 | 4.5e-05 | 0.00115 | 0.28816 | 1.3461 | 2.5476 | 98.56 | 20.05 | 50.01 | 63.2391 | 79.0542 | 61.2414 | 0.011913 | 1.0558 | 0.024536 | 23.2399 | -0.025756 | 0.13668 | -6e-06 | -0.000546 | -0.000546 | -10.085
2019-04-09 14:30 | 1554834622.5182 | -259 | 0.20944 | 0.000447 | 2e-06 | 99.0589 | 99.4015 | 123.8812 | 20.444 | 0.001796 | 0.00171 | -154.1309 | 0.22975 | 8.4022 | 0.002459 | 0.003805 | 0.003634 | 4.4e-05 | 0.001152 | 0.25426 | 1.3543 | 2.5476 | 98.56 | 20.05 | 50.009 | 64.2405 | 80.1833 | 62.1896 | 0.011912 | 1.0528 | 0.024494 | 23.2659 | -0.030887 | 0.13668 | -5e-06 | -0.000525 | -0.000525 | -10.5006
2019-04-09 14:30 | 1554834623.5192 | -258 | 0.20941 | 0.000446 | 3e-06 | 99.1223 | 99.5117 | 123.7859 | 20.4432 | 0.001766 | 0.001657 | -154.1213 | 0.22956 | 8.3995 | 0.002457 | 0.003859 | 0.003723 | 4.6e-05 | 0.001152 | 0.29302 | 1.3458 | 2.5476 | 98.56 | 20.05 | 50.008 | 63.5314 | 80.1405 | 60.8961 | 0.011912 | 1.0569 | 0.024532 | 23.212 | -0.035065 | 0.13657 | -6e-06 | -0.000561 | -0.000561 | -9.816
2019-04-09 14:30 | 1554834624.5183 | -257 | 0.20941 | 0.00045 | 2e-06 | 98.9865 | 99.2604 | 124.0645 | 20.4436 | 0.001741 | 0.001658 | -154.1376 | 0.22927 | 8.4016 | 0.002505 | 0.003662 | 0.003464 | 4.7e-05 | 0.001154 | 0.30287 | 1.372 | 2.5476 | 98.56 | 20.05 | 50.005 | 62.9358 | 81.2933 | 60.9317 | 0.011911 | 1.0568 | 0.024559 | 23.2394 | -0.024717 | 0.13639 | -9e-06 | -0.000808 | -0.000808 | -6.8061
2019-04-09 14:30 | 1554834625.5194 | -256 | 0.20941 | 0.000447 | 2e-06 | 99.0354 | 99.3509 | 123.8307 | 20.4444 | 0.001783 | 0.001615 | -154.1281 | 0.2298 | 8.4025 | 0.002437 | 0.003672 | 0.003547 | 4.6e-05 | 0.001145 | 0.30973 | 1.4051 | 2.5477 | 98.56 | 20.05 | 50.003 | 63.9449 | 79.645 | 59.8966 | 0.011911 | 1.06 | 0.024573 | 23.1815 | -0.028458 | 0.13671 | -6e-06 | -0.000622 | -0.000622 | -8.8616
2019-04-09 14:30 | 1554834626.5196 | -255 | 0.20944 | 0.000448 | 3e-06 | 99.0541 | 99.3922 | 123.8722 | 20.4446 | 0.001783 | 0.001641 | -154.1313 | 0.22931 | 8.4028 | 0.002432 | 0.00355 | 0.003402 | 4.4e-05 | 0.001147 | 0.28047 | 1.3473 | 2.5476 | 98.56 | 20.05 | 50.001 | 63.9395 | 79.5391 | 60.5091 | 0.01191 | 1.0581 | 0.024555 | 23.2067 | -0.030481 | 0.13642 | -8e-06 | -0.000748 | -0.000748 | -7.347
2019-04-09 14:30 | 1554834627.5188 | -254 | 0.20944 | 0.000452 | 2e-06 | 98.9532 | 99.3355 | 123.6432 | 20.4448 | 0.001842 | 0.001634 | -154.1265 | 0.23013 | 8.402 | 0.002514 | 0.004213 | 0.004093 | 4.4e-05 | 0.00114 | 0.27793 | 1.3815 | 2.5476 | 98.56 | 20.05 | 50.002 | 65.3473 | 81.4949 | 60.3442 | 0.011911 | 1.0586 | 0.024538 | 23.1795 | -0.03449 | 0.1369 | -1.2e-05 | -0.001022 | -0.001022 | -5.3989
2019-04-09 14:30 | 1554834628.5198 | -253 | 0.20944 | 0.000446 | 3e-06 | 99.0225 | 99.278 | 123.6049 | 20.4433 | 0.001827 | 0.001673 | -154.1247 | 0.22899 | 8.4054 | 0.002519 | 0.003814 | 0.003621 | 4.5e-05 | 0.00114 | 0.31708 | 1.3406 | 2.5475 | 98.56 | 20.05 | 50.001 | 64.993 | 81.6183 | 61.2911 | 0.01191 | 1.0556 | 0.0245 | 23.2088 | -0.023043 | 0.13622 | -6e-06 | -0.000581 | -0.000581 | -9.4488
2019-04-09 14:30 | 1554834629.5199 | -252 | 0.20945 | 0.000449 | 2e-06 | 99.0029 | 99.299 | 123.392 | 20.4448 | 0.001884 | 0.001695 | -154.1316 | 0.22958 | 8.4017 | 0.002495 | 0.003819 | 0.00369 | 4.6e-05 | 0.001144 | 0.27455 | 1.3359 | 2.5474 | 98.56 | 20.05 | 50 | 66.3602 | 81.0527 | 61.8168 | 0.01191 | 1.054 | 0.02446 | 23.207 | -0.02671 | 0.13658 | -8e-06 | -0.000745 | -0.000745 | -7.3872
2019-04-09 14:30 | 1554834630.5185 | -251 | 0.20948 | 0.00045 | 2e-06 | 99.0243 | 99.3024 | 123.2437 | 20.445 | 0.001732 | 0.001672 | -154.1156 | 0.22957 | 8.4003 | 0.002438 | 0.003623 | 0.003501 | 4.6e-05 | 0.001153 | 0.28064 | 1.3529 | 2.5474 | 98.56 | 20.05 | 49.999 | 62.7076 | 79.6737 | 61.2559 | 0.01191 | 1.0557 | 0.024465 | 23.1736 | -0.025083 | 0.13657 | -1e-05 | -0.000844 | -0.000844 | -6.5206
2019-04-09 14:30 | 1554834631.5186 | -250 | 0.20947 | 0.000448 | 2e-06 | 99.0458 | 99.4026 | 123.1824 | 20.4452 | 0.001745 | 0.001679 | -154.1397 | 0.23012 | 8.4012 | 0.00252 | 0.003828 | 0.003606 | 4.1e-05 | 0.001158 | 0.26537 | 1.3158 | 2.5474 | 98.56 | 20.05 | 49.998 | 63.0303 | 81.642 | 61.435 | 0.01191 | 1.0552 | 0.024453 | 23.1741 | -0.032163 | 0.1369 | -7e-06 | -0.000667 | -0.000667 | -8.2668
2019-04-09 14:31 | 1554834632.5196 | -249 | 0.20947 | 0.000453 | 2e-06 | 99.0464 | 99.1771 | 123.33 | 20.4458 | 0.001799 | 0.001683 | -154.1204 | 0.22939 | 8.4024 | 0.002448 | 0.003964 | 0.003802 | 4.5e-05 | 0.001164 | 0.30237 | 1.3289 | 2.5476 | 98.56 | 20.06 | 49.996 | 64.332 | 79.927 | 61.5397 | 0.011917 | 1.0548 | 0.024463 | 23.1916 | -0.011802 | 0.13646 | -1.4e-05 | -0.001135 | -0.001135 | -4.844
2019-04-09 14:31 | 1554834633.5196 | -248 | 0.20948 | 0.00045 | 3e-06 | 99.0538 | 99.5626 | 123.5173 | 20.4451 | 0.001779 | 0.001682 | -154.1275 | 0.22995 | 8.4008 | 0.002495 | 0.003588 | 0.003456 | 4.5e-05 | 0.001161 | 0.2855 | 1.3174 | 2.5476 | 98.56 | 20.06 | 49.994 | 63.8464 | 81.0467 | 61.5132 | 0.011916 | 1.0549 | 0.024483 | 23.2083 | -0.045825 | 0.1368 | -1e-05 | -0.00087 | -0.00087 | -6.3384
2019-04-09 14:31 | 1554834634.5196 | -247 | 0.20947 | 0.00045 | 2e-06 | 99.0252 | 99.4432 | 123.304 | 20.4458 | 0.001767 | 0.001664 | -154.1087 | 0.22939 | 8.4023 | 0.002469 | 0.003844 | 0.003707 | 4.4e-05 | 0.001162 | 0.28385 | 1.3633 | 2.5476 | 98.56 | 20.06 | 49.993 | 63.5591 | 80.4158 | 61.0613 | 0.011916 | 1.0564 | 0.024478 | 23.1725 | -0.037672 | 0.13646 | -1e-05 | -0.000846 | -0.000846 | -6.498
2019-04-09 14:31 | 1554834635.5186 | -246 | 0.2095 | 0.000443 | 2e-06 | 99.1125 | 99.4167 | 123.5185 | 20.4461 | 0.001805 | 0.001662 | -154.1067 | 0.22854 | 8.4021 | 0.002481 | 0.004197 | 0.003975 | 4.8e-05 | 0.001158 | 0.27096 | 1.3681 | 2.5477 | 98.56 | 20.06 | 49.99 | 64.4666 | 80.7242 | 61.0327 | 0.011915 | 1.0564 | 0.024501 | 23.1917 | -0.02741 | 0.13596 | -1e-06 | -0.000214 | -0.000214 | -25.6432
2019-04-09 14:31 | 1554834636.5177 | -245 | 0.20947 | 0.000451 | 2e-06 | 99.0809 | 99.2067 | 123.6824 | 20.446 | 0.00182 | 0.001669 | -154.1172 | 0.22966 | 8.4002 | 0.002483 | 0.003815 | 0.003712 | 4.5e-05 | 0.001154 | 0.27472 | 1.3558 | 2.5475 | 98.56 | 20.06 | 49.988 | 64.8365 | 80.7582 | 61.1997 | 0.011915 | 1.0559 | 0.024511 | 23.2129 | -0.011352 | 0.13663 | -1.1e-05 | -0.000959 | -0.000959 | -5.7441
2019-04-09 14:31 | 1554834637.5191 | -244 | 0.20948 | 0.000447 | 2e-06 | 99.0362 | 99.4094 | 123.9089 | 20.4464 | 0.001768 | 0.00169 | -154.1195 | 0.22946 | 8.4011 | 0.002535 | 0.004439 | 0.004299 | 4.7e-05 | 0.001151 | 0.30038 | 1.3749 | 2.5477 | 98.56 | 20.06 | 49.985 | 63.5799 | 82.0181 | 61.7029 | 0.011914 | 1.0543 | 0.024515 | 23.2516 | -0.033643 | 0.13651 | -6e-06 | -0.000599 | -0.000599 | -9.1905
2019-04-09 14:31 | 1554834638.5177 | -243 | 0.20947 | 0.000448 | 3e-06 | 99.038 | 99.3472 | 123.956 | 20.4462 | 0.001766 | 0.001662 | -154.1213 | 0.22931 | 8.4013 | 0.002593 | 0.004386 | 0.004183 | 4.4e-05 | 0.001153 | 0.29623 | 1.3606 | 2.5477 | 98.56 | 20.06 | 49.982 | 63.5337 | 83.4181 | 61.0195 | 0.011913 | 1.0565 | 0.024545 | 23.2323 | -0.027875 | 0.13642 | -8e-06 | -0.000781 | -0.000781 | -7.0425
2019-04-09 14:31 | 1554834639.5182 | -242 | 0.20946 | 0.000451 | 2e-06 | 99.0529 | 99.3272 | 123.8159 | 20.4463 | 0.0018 | 0.001639 | -154.1135 | 0.2302 | 8.3995 | 0.002528 | 0.003689 | 0.003534 | 4.8e-05 | 0.001155 | 0.28047 | 1.3581 | 2.5475 | 98.56 | 20.06 | 49.981 | 64.3558 | 81.8401 | 60.4818 | 0.011913 | 1.0582 | 0.02455 | 23.2005 | -0.024733 | 0.13694 | -1.2e-05 | -0.000995 | -0.000995 | -5.5455
2019-04-09 14:31 | 1554834640.5179 | -241 | 0.20946 | 0.000449 | 3e-06 | 99.1435 | 99.4803 | 124.0436 | 20.446 | 0.001808 | 0.001649 | -154.1304 | 0.22919 | 8.4002 | 0.00253 | 0.004054 | 0.003904 | 4.3e-05 | 0.001159 | 0.29302 | 1.3832 | 2.5477 | 98.56 | 20.06 | 49.978 | 64.5462 | 81.8906 | 60.7122 | 0.011912 | 1.0575 | 0.024565 | 23.2298 | -0.030342 | 0.13634 | -9e-06 | -0.00081 | -0.00081 | -6.7849
2019-04-09 14:31 | 1554834641.5189 | -240 | 0.20947 | 0.000445 | 3e-06 | 99.063 | 99.3737 | 123.9253 | 20.4476 | 0.001787 | 0.001652 | -154.1192 | 0.22955 | 8.4035 | 0.002541 | 0.004246 | 0.004016 | 4.5e-05 | 0.001154 | 0.28799 | 1.3301 | 2.5476 | 98.56 | 20.06 | 49.976 | 64.0346 | 82.1697 | 60.7925 | 0.011912 | 1.0572 | 0.02455 | 23.2215 | -0.028001 | 0.13656 | -4e-06 | -0.000445 | -0.000445 | -12.3641
2019-04-09 14:31 | 1554834642.5179 | -239 | 0.20946 | 0.000451 | 3e-06 | 99.0455 | 99.1634 | 123.6528 | 20.4473 | 0.001814 | 0.001686 | -154.1206 | 0.22933 | 8.4037 | 0.002516 | 0.004178 | 0.003979 | 4.6e-05 | 0.001144 | 0.28385 | 1.3114 | 2.5476 | 98.56 | 20.06 | 49.975 | 64.6932 | 81.5613 | 61.6115 | 0.011912 | 1.0546 | 0.024493 | 23.2244 | -0.010641 | 0.13643 | -1.1e-05 | -0.000965 | -0.000965 | -5.6986
2019-04-09 14:31 | 1554834643.5189 | -238 | 0.20946 | 0.000452 | 2e-06 | 99.0848 | 99.4652 | 123.5669 | 20.4471 | 0.001784 | 0.00162 | -154.1146 | 0.22961 | 8.4014 | 0.002465 | 0.003772 | 0.003626 | 4.3e-05 | 0.001139 | 0.27168 | 1.36 | 2.5477 | 98.56 | 20.06 | 49.971 | 63.957 | 80.3267 | 60.0159 | 0.011911 | 1.0597 | 0.024543 | 23.1609 | -0.034275 | 0.13659 | -1.2e-05 | -0.001039 | -0.001039 | -5.3006
2019-04-09 14:31 | 1554834644.518 | -237 | 0.20946 | 0.000452 | 2e-06 | 99.0815 | 99.428 | 123.4957 | 20.4473 | 0.001735 | 0.001679 | -154.1326 | 0.22949 | 8.4026 | 0.002506 | 0.00427 | 0.004111 | 4.5e-05 | 0.001137 | 0.2995 | 1.3582 | 2.5476 | 98.56 | 20.06 | 49.968 | 62.7887 | 81.3077 | 61.4378 | 0.01191 | 1.0552 | 0.024484 | 23.2036 | -0.031227 | 0.13653 | -1.2e-05 | -0.001002 | -0.001002 | -5.4914
2019-04-09 14:31 | 1554834645.5182 | -236 | 0.20947 | 0.000446 | 3e-06 | 99.1655 | 99.5764 | 123.395 | 20.4477 | 0.001795 | 0.001695 | -154.1126 | 0.2293 | 8.402 | 0.002472 | 0.003657 | 0.003464 | 4.6e-05 | 0.001138 | 0.28727 | 1.3728 | 2.5476 | 98.56 | 20.06 | 49.966 | 64.2354 | 80.4989 | 61.8329 | 0.01191 | 1.0539 | 0.024459 | 23.2078 | -0.036983 | 0.13641 | -5e-06 | -0.0005 | -0.0005 | -11.0043
2019-04-09 14:31 | 1554834646.5196 | -235 | 0.20949 | 0.000447 | 3e-06 | 99.1312 | 99.4074 | 123.4803 | 20.4491 | 0.001773 | 0.001675 | -154.1142 | 0.22912 | 8.4013 | 0.002547 | 0.003936 | 0.003769 | 4.8e-05 | 0.001142 | 0.27848 | 1.3571 | 2.5475 | 98.56 | 20.06 | 49.959 | 63.7081 | 82.3008 | 61.335 | 0.011908 | 1.0555 | 0.024486 | 23.1986 | -0.024881 | 0.13631 | -6e-06 | -0.00062 | -0.00062 | -8.8645
2019-04-09 14:31 | 1554834647.5184 | -234 | 0.20949 | 0.000447 | 2e-06 | 99.1452 | 99.4215 | 123.3828 | 20.4482 | 0.001763 | 0.00164 | -154.1242 | 0.22959 | 8.4013 | 0.002493 | 0.003208 | 0.003048 | 4.7e-05 | 0.00115 | 0.30574 | 1.3705 | 2.5474 | 98.56 | 20.06 | 49.956 | 63.4566 | 81.0115 | 60.5008 | 0.011907 | 1.0581 | 0.024507 | 23.1605 | -0.024888 | 0.13658 | -6e-06 | -0.00061 | -0.00061 | -9.0285
2019-04-09 14:31 | 1554834648.5195 | -233 | 0.2095 | 0.000446 | 3e-06 | 99.1765 | 99.7288 | 123.3298 | 20.4487 | 0.001778 | 0.001682 | -154.124 | 0.22922 | 8.4033 | 0.002552 | 0.003744 | 0.003592 | 4.4e-05 | 0.001152 | 0.28224 | 1.4169 | 2.5477 | 98.56 | 20.06 | 49.949 | 63.8193 | 82.4157 | 61.5104 | 0.011905 | 1.0549 | 0.024465 | 23.1905 | -0.049674 | 0.13636 | -5e-06 | -0.000488 | -0.000488 | -11.2533
2019-04-09 14:31 | 1554834649.5191 | -232 | 0.20949 | 0.000456 | 3e-06 | 99.0946 | 99.628 | 123.327 | 20.4481 | 0.001801 | 0.001667 | -154.1242 | 0.22854 | 8.4025 | 0.002456 | 0.003928 | 0.003752 | 4.5e-05 | 0.001155 | 0.28982 | 1.3706 | 2.5475 | 98.56 | 20.06 | 49.945 | 64.3701 | 80.1256 | 61.1387 | 0.011905 | 1.0561 | 0.024478 | 23.1774 | -0.048014 | 0.13596 | -1.7e-05 | -0.001441 | -0.001441 | -3.8018
2019-04-09 14:31 | 1554834650.5185 | -231 | 0.20946 | 0.000448 | 2e-06 | 99.2384 | 99.4889 | 123.235 | 20.4488 | 0.001801 | 0.001611 | -154.1123 | 0.22912 | 8.4021 | 0.002517 | 0.003567 | 0.003377 | 4.1e-05 | 0.001153 | 0.2792 | 1.3542 | 2.5476 | 98.56 | 20.06 | 49.939 | 64.3691 | 81.5705 | 59.7869 | 0.011903 | 1.0604 | 0.024518 | 23.1218 | -0.02255 | 0.1363 | -8e-06 | -0.000706 | -0.000706 | -7.7858
2019-04-09 14:31 | 1554834651.5178 | -230 | 0.20947 | 0.000448 | 2e-06 | 99.2477 | 99.6211 | 123.5413 | 20.4484 | 0.00173 | 0.001717 | -154.1074 | 0.22853 | 8.402 | 0.002554 | 0.003389 | 0.003275 | 4.3e-05 | 0.001152 | 0.2891 | 1.3235 | 2.5476 | 98.56 | 20.06 | 49.933 | 62.6742 | 82.4763 | 62.3579 | 0.011902 | 1.0523 | 0.024455 | 23.2398 | -0.033591 | 0.13595 | -7e-06 | -0.00064 | -0.00064 | -8.5659
2019-04-09 14:31 | 1554834652.5182 | -229 | 0.20949 | 0.000452 | 2e-06 | 99.3138 | 99.649 | 123.596 | 20.4488 | 0.001782 | 0.001607 | -154.1002 | 0.22956 | 8.4039 | 0.002505 | 0.003677 | 0.003479 | 4.3e-05 | 0.001143 | 0.29535 | 1.3597 | 2.5476 | 98.56 | 20.06 | 49.923 | 63.9179 | 81.2841 | 59.6945 | 0.011899 | 1.0607 | 0.024558 | 23.1525 | -0.030138 | 0.13656 | -1.2e-05 | -0.001002 | -0.001002 | -5.4939
2019-04-09 14:31 | 1554834653.5178 | -228 | 0.2095 | 0.000452 | 2e-06 | 99.3028 | 99.6084 | 123.5472 | 20.4488 | 0.001751 | 0.001683 | -154.1139 | 0.22949 | 8.3999 | 0.002548 | 0.003891 | 0.003659 | 4.3e-05 | 0.001146 | 0.29319 | 1.3585 | 2.5477 | 98.56 | 20.06 | 49.917 | 63.1635 | 82.3238 | 61.5213 | 0.011898 | 1.0549 | 0.024486 | 23.2113 | -0.027484 | 0.13652 | -1.3e-05 | -0.001102 | -0.001102 | -4.9926
2019-04-09 14:31 | 1554834654.5173 | -227 | 0.20947 | 0.000451 | 3e-06 | 99.3394 | 99.7574 | 123.5791 | 20.4494 | 0.001829 | 0.001628 | -154.1137 | 0.22944 | 8.4016 | 0.00253 | 0.003425 | 0.003269 | 4e-05 | 0.00115 | 0.28191 | 1.3784 | 2.5477 | 98.56 | 20.06 | 49.912 | 65.0568 | 81.9032 | 60.2079 | 0.011897 | 1.0591 | 0.024537 | 23.1687 | -0.037557 | 0.13649 | -1.2e-05 | -0.001073 | -0.001073 | -5.1252
2019-04-09 14:31 | 1554834655.5177 | -226 | 0.20948 | 0.000451 | 2e-06 | 99.4076 | 99.7722 | 123.7323 | 20.4503 | 0.001759 | 0.001674 | -154.1331 | 0.22904 | 8.4022 | 0.002535 | 0.004073 | 0.00397 | 4.1e-05 | 0.001154 | 0.29695 | 1.3719 | 2.5477 | 98.56 | 20.06 | 49.901 | 63.3557 | 82.0104 | 61.3104 | 0.011894 | 1.0556 | 0.024512 | 23.2214 | -0.032746 | 0.13626 | -1.1e-05 | -0.000921 | -0.000921 | -5.9626
2019-04-09 14:31 | 1554834656.5183 | -225 | 0.20947 | 0.000453 | 2e-06 | 99.3793 | 99.7901 | 123.4896 | 20.4505 | 0.001805 | 0.001684 | -154.1399 | 0.22949 | 8.4027 | 0.002459 | 0.004016 | 0.00381 | 4.3e-05 | 0.001158 | 0.27007 | 1.3259 | 2.5476 | 98.56 | 20.06 | 49.896 | 64.474 | 80.1911 | 61.5558 | 0.011893 | 1.0548 | 0.024479 | 23.2071 | -0.036896 | 0.13652 | -1.3e-05 | -0.001126 | -0.001126 | -4.8844
2019-04-09 14:31 | 1554834657.5173 | -224 | 0.20946 | 0.000451 | 2e-06 | 99.3483 | 99.8042 | 123.0685 | 20.4501 | 0.001768 | 0.00168 | -154.1192 | 0.22879 | 8.4037 | 0.002469 | 0.004282 | 0.004113 | 4.2e-05 | 0.001154 | 0.27688 | 1.3378 | 2.5476 | 98.56 | 20.06 | 49.891 | 63.5904 | 80.4412 | 61.4673 | 0.011892 | 1.0551 | 0.02444 | 23.1645 | -0.040948 | 0.13611 | -1e-05 | -0.000883 | -0.000883 | -6.215
2019-04-09 14:31 | 1554834658.5173 | -223 | 0.20949 | 0.000451 | 2e-06 | 99.3787 | 99.7254 | 123.2969 | 20.4494 | 0.001759 | 0.001682 | -154.1259 | 0.22899 | 8.4015 | 0.00248 | 0.004273 | 0.004206 | 3.9e-05 | 0.001155 | 0.31996 | 1.3613 | 2.5476 | 98.56 | 20.07 | 49.883 | 63.375 | 80.6892 | 61.5142 | 0.011897 | 1.0549 | 0.024461 | 23.1876 | -0.031148 | 0.13623 | -1.1e-05 | -0.000946 | -0.000946 | -5.8004
2019-04-09 14:31 | 1554834659.5825 | -222 | 0.20945 | 0.000448 | 3e-06 | 99.3972 | 99.7209 | 123.5495 | 20.4502 | 0.001725 | 0.001632 | -154.1049 | 0.2297 | 8.3999 | 0.002448 | 0.004018 | 0.003871 | 4.5e-05 | 0.001152 | 0.27721 | 1.3333 | 2.5475 | 98.56 | 20.07 | 49.879 | 62.5493 | 79.9202 | 60.3056 | 0.011896 | 1.0587 | 0.024531 | 23.1694 | -0.029085 | 0.13664 | -8e-06 | -0.000717 | -0.000717 | -7.6849
2019-04-09 14:31 | 1554834660.5308 | -221 | 0.20944 | 0.000448 | 3e-06 | 99.3775 | 99.7753 | 123.4462 | 20.4497 | 0.001745 | 0.001656 | -154.1138 | 0.22981 | 8.3996 | 0.00244 | 0.0041 | 0.003916 | 4.3e-05 | 0.00115 | 0.26504 | 1.3325 | 2.5476 | 98.56 | 20.07 | 49.873 | 63.0183 | 79.7421 | 60.8846 | 0.011895 | 1.0569 | 0.024499 | 23.1798 | -0.03573 | 0.13672 | -8e-06 | -0.00073 | -0.00073 | -7.5516
2019-04-09 14:31 | 1554834661.5315 | -220 | 0.20946 | 0.000451 | 2e-06 | 99.3394 | 99.6424 | 123.1359 | 20.4511 | 0.001765 | 0.001651 | -154.1201 | 0.22941 | 8.4025 | 0.00251 | 0.004059 | 0.003966 | 4.4e-05 | 0.00115 | 0.25946 | 1.3127 | 2.5474 | 98.56 | 20.07 | 49.872 | 63.504 | 81.4261 | 60.76 | 0.011895 | 1.0573 | 0.024473 | 23.1463 | -0.027244 | 0.13648 | -1.1e-05 | -0.000926 | -0.000926 | -5.9364
2019-04-09 14:31 | 1554834662.5315 | -219 | 0.20947 | 0.000447 | 2e-06 | 99.438 | 99.6786 | 123.4399 | 20.4502 | 0.001799 | 0.001684 | -154.1368 | 0.22945 | 8.401 | 0.002539 | 0.003956 | 0.003725 | 4.2e-05 | 0.001148 | 0.29789 | 1.3306 | 2.5474 | 98.56 | 20.07 | 49.87 | 64.3175 | 82.1208 | 61.5579 | 0.011894 | 1.0548 | 0.024474 | 23.2025 | -0.021617 | 0.1365 | -5e-06 | -0.000526 | -0.000526 | -10.4632
2019-04-09 14:31 | 1554834663.5305 | -218 | 0.20945 | 0.000444 | 3e-06 | 99.5496 | 99.83 | 123.6882 | 20.4503 | 0.001787 | 0.001664 | -154.1334 | 0.2291 | 8.4038 | 0.002493 | 0.004244 | 0.004131 | 4.2e-05 | 0.001142 | 0.27865 | 1.3314 | 2.5476 | 98.56 | 20.07 | 49.866 | 64.0318 | 80.998 | 61.0765 | 0.011893 | 1.0563 | 0.024516 | 23.2091 | -0.025155 | 0.13629 | -3e-06 | -0.000403 | -0.000403 | -13.6369
2019-04-09 14:31 | 1554834664.5316 | -217 | 0.2095 | 0.000452 | 3e-06 | 99.5199 | 99.8386 | 123.8291 | 20.4514 | 0.001786 | 0.001653 | -154.1156 | 0.22957 | 8.402 | 0.002431 | 0.003664 | 0.003423 | 4.2e-05 | 0.001128 | 0.29031 | 1.3362 | 2.5477 | 98.56 | 20.07 | 49.863 | 64.015 | 79.509 | 60.8197 | 0.011892 | 1.0571 | 0.024539 | 23.2134 | -0.028602 | 0.13657 | -1.2e-05 | -0.001061 | -0.001061 | -5.1861
2019-04-09 14:31 | 1554834665.5308 | -216 | 0.20949 | 0.00045 | 2e-06 | 99.5487 | 99.6686 | 123.7685 | 20.4517 | 0.001784 | 0.001659 | -154.1191 | 0.22977 | 8.4037 | 0.002521 | 0.003498 | 0.003377 | 4.1e-05 | 0.001127 | 0.24187 | 1.3414 | 2.5476 | 98.56 | 20.07 | 49.858 | 63.9669 | 81.684 | 60.9614 | 0.011891 | 1.0567 | 0.024528 | 23.2127 | -0.010762 | 0.13669 | -1e-05 | -0.000869 | -0.000869 | -6.3372
2019-04-09 14:31 | 1554834666.5307 | -215 | 0.20948 | 0.000453 | 2e-06 | 99.5336 | 99.8892 | 123.9159 | 20.4512 | 0.001816 | 0.001678 | -154.1031 | 0.22867 | 8.402 | 0.002421 | 0.00399 | 0.003823 | 4.2e-05 | 0.001124 | 0.27688 | 1.3531 | 2.5476 | 98.56 | 20.07 | 49.856 | 64.7463 | 79.2796 | 61.4214 | 0.011891 | 1.0552 | 0.024526 | 23.2425 | -0.031902 | 0.13603 | -1.4e-05 | -0.001153 | -0.001153 | -4.7555
2019-04-09 14:31 | 1554834667.5311 | -214 | 0.20952 | 0.00045 | 3e-06 | 99.5589 | 99.9322 | 124.1673 | 20.4506 | 0.001807 | 0.001661 | -154.1146 | 0.22932 | 8.4019 | 0.002471 | 0.003574 | 0.003437 | 4.4e-05 | 0.001128 | 0.27887 | 1.345 | 2.5475 | 98.56 | 20.07 | 49.855 | 64.527 | 80.4719 | 61.012 | 0.01189 | 1.0565 | 0.024566 | 23.2518 | -0.033479 | 0.13642 | -1e-05 | -0.000899 | -0.000899 | -6.1168
2019-04-09 14:31 | 1554834668.5311 | -213 | 0.20946 | 0.000454 | 2e-06 | 99.5886 | 100.0885 | 124.2283 | 20.4518 | 0.001735 | 0.001683 | -154.0993 | 0.22942 | 8.4008 | 0.002515 | 0.003678 | 0.00344 | 4.2e-05 | 0.001133 | 0.27887 | 1.3301 | 2.5476 | 98.56 | 20.07 | 49.854 | 62.7774 | 81.5282 | 61.5361 | 0.01189 | 1.0549 | 0.024553 | 23.2757 | -0.044779 | 0.13648 | -1.5e-05 | -0.001221 | -0.001221 | -4.504
2019-04-09 14:31 | 1554834669.5307 | -212 | 0.20948 | 0.000452 | 2e-06 | 99.6 | 99.9924 | 124.1762 | 20.4514 | 0.001781 | 0.001676 | -154.1248 | 0.2285 | 8.4011 | 0.002483 | 0.003268 | 0.003119 | 4.2e-05 | 0.001142 | 0.2688 | 1.3379 | 2.5475 | 98.56 | 20.07 | 49.855 | 63.8891 | 80.7627 | 61.3762 | 0.01189 | 1.0554 | 0.024553 | 23.2653 | -0.035177 | 0.13593 | -1.3e-05 | -0.001107 | -0.001107 | -4.9472
2019-04-09 14:31 | 1554834670.5308 | -211 | 0.2095 | 0.00045 | 2e-06 | 99.5809 | 99.9914 | 124.1203 | 20.4522 | 0.001741 | 0.001692 | -154.1104 | 0.22934 | 8.3995 | 0.002404 | 0.003746 | 0.003607 | 4e-05 | 0.001149 | 0.26753 | 1.3402 | 2.5475 | 98.56 | 20.07 | 49.855 | 62.9341 | 78.8738 | 61.7584 | 0.01189 | 1.0542 | 0.024534 | 23.2734 | -0.036796 | 0.13643 | -1e-05 | -0.000842 | -0.000842 | -6.5295
2019-04-09 14:31 | 1554834671.5298 | -210 | 0.20943 | 0.000454 | 3e-06 | 99.6309 | 99.9708 | 124.2225 | 20.4534 | 0.001764 | 0.001713 | -154.1068 | 0.22984 | 8.4006 | 0.002443 | 0.004008 | 0.003847 | 4.5e-05 | 0.001147 | 0.28584 | 1.3629 | 2.5477 | 98.56 | 20.07 | 49.855 | 63.4983 | 79.8087 | 62.2524 | 0.01189 | 1.0526 | 0.024526 | 23.3001 | -0.030456 | 0.13673 | -1.5e-05 | -0.001257 | -0.001257 | -4.3837
2019-04-09 14:31 | 1554834672.5308 | -209 | 0.2095 | 0.000452 | 2e-06 | 99.6545 | 99.9856 | 123.9089 | 20.4523 | 0.001768 | 0.001694 | -154.0858 | 0.22958 | 8.4014 | 0.002493 | 0.003646 | 0.003432 | 4.4e-05 | 0.001145 | 0.28495 | 1.3458 | 2.5476 | 98.56 | 20.07 | 49.857 | 63.5747 | 81.0131 | 61.8053 | 0.011891 | 1.054 | 0.024511 | 23.2551 | -0.029669 | 0.13658 | -1.2e-05 | -0.001039 | -0.001039 | -5.2988
2019-04-09 14:31 | 1554834673.5305 | -208 | 0.20946 | 0.000452 | 2e-06 | 99.6822 | 100.1731 | 123.5556 | 20.4527 | 0.001761 | 0.001687 | -154.0734 | 0.22919 | 8.4025 | 0.002519 | 0.003383 | 0.003252 | 4.6e-05 | 0.001144 | 0.28528 | 1.3533 | 2.5476 | 98.56 | 20.07 | 49.857 | 63.4141 | 81.6274 | 61.6292 | 0.011891 | 1.0546 | 0.024483 | 23.2159 | -0.043947 | 0.13635 | -1.2e-05 | -0.001016 | -0.001016 | -5.4072
2019-04-09 14:31 | 1554834674.5298 | -207 | 0.20949 | 0.000451 | 2e-06 | 99.6571 | 99.9625 | 123.6009 | 20.4525 | 0.00173 | 0.00169 | -154.0751 | 0.22902 | 8.4008 | 0.002477 | 0.003815 | 0.003664 | 4.5e-05 | 0.001148 | 0.27776 | 1.3384 | 2.5476 | 98.56 | 20.07 | 49.858 | 62.6633 | 80.6221 | 61.6994 | 0.011891 | 1.0543 | 0.024484 | 23.2226 | -0.027365 | 0.13624 | -1.1e-05 | -0.000945 | -0.000945 | -5.8076
2019-04-09 14:31 | 1554834675.5313 | -206 | 0.20947 | 0.000447 | 3e-06 | 99.672 | 100.062 | 123.7481 | 20.4527 | 0.001764 | 0.00166 | -154.0697 | 0.22956 | 8.4042 | 0.002494 | 0.003876 | 0.003743 | 4.4e-05 | 0.00115 | 0.26864 | 1.3568 | 2.5477 | 98.56 | 20.07 | 49.859 | 63.4774 | 81.0277 | 60.989 | 0.011891 | 1.0566 | 0.024525 | 23.2117 | -0.034924 | 0.13656 | -6e-06 | -0.000642 | -0.000642 | -8.5755
2019-04-09 14:31 | 1554834676.5313 | -205 | 0.20949 | 0.00045 | 2e-06 | 99.7501 | 100.1786 | 123.4087 | 20.4529 | 0.001778 | 0.001688 | -154.0715 | 0.23034 | 8.3985 | 0.002495 | 0.003477 | 0.003306 | 4.7e-05 | 0.001141 | 0.27383 | 1.4043 | 2.5475 | 98.56 | 20.07 | 49.864 | 63.8151 | 81.0623 | 61.6511 | 0.011893 | 1.0545 | 0.024467 | 23.2028 | -0.038348 | 0.13703 | -1e-05 | -0.000865 | -0.000865 | -6.3846
2019-04-09 14:31 | 1554834677.53 | -204 | 0.20947 | 0.000448 | 3e-06 | 99.7054 | 100.2117 | 123.4925 | 20.4529 | 0.001789 | 0.001675 | -154.0749 | 0.22906 | 8.4027 | 0.00252 | 0.003862 | 0.00361 | 4.5e-05 | 0.00114 | 0.28335 | 1.3516 | 2.5475 | 98.56 | 20.07 | 49.866 | 64.0857 | 81.6696 | 61.3456 | 0.011893 | 1.0555 | 0.024487 | 23.2001 | -0.045304 | 0.13627 | -8e-06 | -0.000748 | -0.000748 | -7.3383
2019-04-09 14:31 | 1554834678.5304 | -203 | 0.20946 | 0.000452 | 2e-06 | 99.7134 | 100.0441 | 123.6312 | 20.4537 | 0.001736 | 0.001669 | -154.101 | 0.22863 | 8.4018 | 0.002454 | 0.003625 | 0.003467 | 4.1e-05 | 0.001147 | 0.29734 | 1.3651 | 2.5475 | 98.56 | 20.07 | 49.87 | 62.8151 | 80.0671 | 61.1964 | 0.011894 | 1.0559 | 0.024506 | 23.208 | -0.02961 | 0.13601 | -1.2e-05 | -0.001 | -0.001 | -5.4835
2019-04-09 14:31 | 1554834679.5311 | -202 | 0.20949 | 0.00045 | 2e-06 | 99.7798 | 100.1428 | 123.4105 | 20.4541 | 0.001784 | 0.001668 | -154.0853 | 0.22922 | 8.4006 | 0.002443 | 0.003877 | 0.00372 | 4.4e-05 | 0.00115 | 0.28152 | 1.3449 | 2.5476 | 98.56 | 20.07 | 49.871 | 63.962 | 79.8141 | 61.1748 | 0.011894 | 1.056 | 0.024485 | 23.1865 | -0.032482 | 0.13636 | -9e-06 | -0.000799 | -0.000799 | -6.8765
2019-04-09 14:31 | 1554834680.53 | -201 | 0.20947 | 0.000448 | 3e-06 | 99.8066 | 100.0602 | 123.5549 | 20.4532 | 0.001779 | 0.001695 | -154.0999 | 0.22979 | 8.4027 | 0.002501 | 0.003896 | 0.003806 | 4.4e-05 | 0.001157 | 0.29264 | 1.3492 | 2.5475 | 98.56 | 20.07 | 49.873 | 63.8401 | 81.212 | 61.8365 | 0.011895 | 1.0539 | 0.024475 | 23.223 | -0.022699 | 0.1367 | -8e-06 | -0.000725 | -0.000725 | -7.5962
2019-04-09 14:31 | 1554834681.5292 | -200 | 0.20948 | 0.000452 | 3e-06 | 99.6973 | 100.2 | 123.661 | 20.4537 | 0.001761 | 0.001697 | -154.1072 | 0.22944 | 8.4043 | 0.002525 | 0.003502 | 0.003299 | 4.4e-05 | 0.001163 | 0.28174 | 1.3174 | 2.5476 | 98.56 | 20.07 | 49.874 | 63.4252 | 81.7906 | 61.8669 | 0.011895 | 1.0538 | 0.024484 | 23.234 | -0.044981 | 0.13649 | -1.4e-05 | -0.001176 | -0.001176 | -4.6773
2019-04-09 14:31 | 1554834682.5302 | -199 | 0.20947 | 0.000452 | 3e-06 | 99.7128 | 100.1865 | 123.7835 | 20.4542 | 0.001787 | 0.001701 | -154.1298 | 0.22908 | 8.4015 | 0.002529 | 0.004444 | 0.004269 | 4.6e-05 | 0.001167 | 0.28456 | 1.3667 | 2.5475 | 98.56 | 20.07 | 49.874 | 64.045 | 81.8714 | 61.9698 | 0.011895 | 1.0535 | 0.024493 | 23.2491 | -0.042393 | 0.13628 | -1.3e-05 | -0.001092 | -0.001092 | -5.027
2019-04-09 14:31 | 1554834683.5291 | -198 | 0.20947 | 0.00045 | 2e-06 | 99.6976 | 100.1132 | 123.8298 | 20.4545 | 0.001715 | 0.001621 | -154.128 | 0.22939 | 8.401 | 0.002469 | 0.004587 | 0.004461 | 4.6e-05 | 0.001171 | 0.2896 | 1.3429 | 2.5477 | 98.56 | 20.07 | 49.875 | 62.3043 | 80.4353 | 60.0425 | 0.011895 | 1.0596 | 0.024568 | 23.1865 | -0.037209 | 0.13646 | -9e-06 | -0.000829 | -0.000829 | -6.6351
2019-04-09 14:31 | 1554834684.5292 | -197 | 0.20949 | 0.000455 | 3e-06 | 99.7369 | 100.0981 | 124.033 | 20.4539 | 0.001771 | 0.001633 | -154.1576 | 0.22926 | 8.4039 | 0.00257 | 0.003431 | 0.003279 | 4.4e-05 | 0.001171 | 0.2995 | 1.3482 | 2.5476 | 98.56 | 20.07 | 49.875 | 63.6449 | 82.8593 | 60.3305 | 0.011895 | 1.0587 | 0.024578 | 23.2155 | -0.03233 | 0.13639 | -1.7e-05 | -0.001388 | -0.001388 | -3.9586
2019-04-09 14:31 | 1554834685.5306 | -196 | 0.20947 | 0.00045 | 3e-06 | 99.7861 | 100.213 | 123.7438 | 20.4542 | 0.001873 | 0.001691 | -154.1698 | 0.23022 | 8.4013 | 0.00245 | 0.003735 | 0.003586 | 4.5e-05 | 0.001171 | 0.2573 | 1.3435 | 2.5476 | 98.56 | 20.07 | 49.875 | 66.1135 | 79.9779 | 61.7275 | 0.011895 | 1.0543 | 0.024498 | 23.237 | -0.03819 | 0.13695 | -1e-05 | -0.00086 | -0.00086 | -6.4181
2019-04-09 14:31 | 1554834686.5297 | -195 | 0.20949 | 0.000451 | 2e-06 | 99.7914 | 100.1893 | 123.8375 | 20.4552 | 0.001761 | 0.001663 | -154.1682 | 0.22962 | 8.403 | 0.002512 | 0.00371 | 0.003506 | 4.6e-05 | 0.001173 | 0.29103 | 1.3594 | 2.5476 | 98.56 | 20.07 | 49.873 | 63.4213 | 81.4644 | 61.068 | 0.011895 | 1.0563 | 0.024531 | 23.2229 | -0.03559 | 0.1366 | -1.1e-05 | -0.000965 | -0.000965 | -5.7024
2019-04-09 14:31 | 1554834687.5303 | -194 | 0.20945 | 0.000451 | 3e-06 | 99.7667 | 100.1821 | 123.83 | 20.4545 | 0.001763 | 0.001711 | -154.1782 | 0.23 | 8.4012 | 0.002489 | 0.003805 | 0.003747 | 4.5e-05 | 0.001183 | 0.2808 | 1.3515 | 2.5475 | 98.56 | 20.07 | 49.87 | 63.4721 | 80.9092 | 62.2176 | 0.011894 | 1.0527 | 0.024488 | 23.262 | -0.03716 | 0.13683 | -1.1e-05 | -0.000984 | -0.000984 | -5.602
2019-04-09 14:31 | 1554834688.5307 | -193 | 0.20949 | 0.000452 | 3e-06 | 99.7635 | 100.0441 | 123.7352 | 20.4537 | 0.001798 | 0.001642 | -154.1567 | 0.22893 | 8.4028 | 0.002507 | 0.003868 | 0.003659 | 4.6e-05 | 0.001185 | 0.29247 | 1.3331 | 2.5476 | 98.56 | 20.07 | 49.866 | 64.3122 | 81.3396 | 60.5562 | 0.011893 | 1.058 | 0.02454 | 23.1955 | -0.025124 | 0.13619 | -1.4e-05 | -0.001187 | -0.001187 | -4.6239
2019-04-09 14:31 | 1554834689.5293 | -192 | 0.20946 | 0.000449 | 2e-06 | 99.862 | 100.2915 | 123.5666 | 20.4552 | 0.001669 | 0.001684 | -154.1845 | 0.22941 | 8.4004 | 0.002477 | 0.003174 | 0.00303 | 4.6e-05 | 0.001186 | 0.26936 | 1.3088 | 2.5477 | 98.56 | 20.07 | 49.863 | 61.2079 | 80.6285 | 61.5666 | 0.011892 | 1.0548 | 0.024486 | 23.2147 | -0.038387 | 0.13648 | -9e-06 | -0.0008 | -0.0008 | -6.8773
2019-04-09 14:31 | 1554834690.5294 | -191 | 0.20945 | 0.000446 | 3e-06 | 99.8546 | 100.1463 | 123.691 | 20.4559 | 0.00184 | 0.001714 | -154.1725 | 0.2297 | 8.4 | 0.002518 | 0.003293 | 0.003099 | 4.5e-05 | 0.001197 | 0.29623 | 1.3764 | 2.5476 | 98.56 | 20.06 | 49.86 | 65.316 | 81.6093 | 62.2931 | 0.011884 | 1.0525 | 0.024472 | 23.2516 | -0.026092 | 0.13664 | -6e-06 | -0.000581 | -0.000581 | -9.4856
2019-04-09 14:32 | 1554834691.5294 | -190 | 0.20946 | 0.000449 | 2e-06 | 99.7706 | 100.2137 | 123.9891 | 20.4552 | 0.001807 | 0.001667 | -154.174 | 0.23005 | 8.3977 | 0.002419 | 0.003984 | 0.003869 | 4.6e-05 | 0.001198 | 0.29479 | 1.3759 | 2.5476 | 98.56 | 20.06 | 49.855 | 64.5339 | 79.2352 | 61.1501 | 0.011883 | 1.0561 | 0.024543 | 23.2399 | -0.039639 | 0.13686 | -8e-06 | -0.000761 | -0.000761 | -7.244
2019-04-09 14:32 | 1554834692.53 | -189 | 0.20946 | 0.000448 | 2e-06 | 99.7691 | 100.0912 | 123.7852 | 20.4548 | 0.001774 | 0.001671 | -154.1731 | 0.22939 | 8.4022 | 0.002437 | 0.003927 | 0.00371 | 4.1e-05 | 0.001206 | 0.26985 | 1.3452 | 2.5476 | 98.56 | 20.06 | 49.851 | 63.723 | 79.6665 | 61.242 | 0.011882 | 1.0558 | 0.024519 | 23.224 | -0.028831 | 0.13646 | -8e-06 | -0.000727 | -0.000727 | -7.5638
2019-04-09 14:32 | 1554834693.5299 | -188 | 0.20946 | 0.000453 | 3e-06 | 99.7837 | 100.0764 | 123.7455 | 20.4554 | 0.001791 | 0.001651 | -154.1783 | 0.23038 | 8.4017 | 0.002444 | 0.00416 | 0.004026 | 4.4e-05 | 0.001206 | 0.31725 | 1.3563 | 2.5475 | 98.56 | 20.06 | 49.847 | 64.1462 | 79.8318 | 60.7579 | 0.011881 | 1.0573 | 0.024533 | 23.2035 | -0.0262 | 0.13706 | -1.4e-05 | -0.00124 | -0.00124 | -4.4539
2019-04-09 14:32 | 1554834694.5295 | -187 | 0.2095 | 0.000452 | 2e-06 | 99.7965 | 100.2234 | 123.5986 | 20.4552 | 0.001851 | 0.001689 | -154.1762 | 0.23003 | 8.4011 | 0.002452 | 0.003734 | 0.003555 | 4.5e-05 | 0.001207 | 0.29336 | 1.3768 | 2.5476 | 98.56 | 20.06 | 49.844 | 65.5727 | 80.0382 | 61.6901 | 0.01188 | 1.0544 | 0.024485 | 23.222 | -0.038177 | 0.13685 | -1.2e-05 | -0.001061 | -0.001061 | -5.1973
2019-04-09 14:32 | 1554834695.5296 | -186 | 0.20946 | 0.000448 | 2e-06 | 99.7486 | 100.2106 | 123.8751 | 20.4567 | 0.001751 | 0.001672 | -154.1637 | 0.22899 | 8.4012 | 0.002465 | 0.003853 | 0.003647 | 4.6e-05 | 0.001211 | 0.28694 | 1.3386 | 2.5475 | 98.56 | 20.06 | 49.842 | 63.1775 | 80.3422 | 61.2804 | 0.01188 | 1.0557 | 0.024527 | 23.2338 | -0.041339 | 0.13623 | -7e-06 | -0.000664 | -0.000664 | -8.2682
2019-04-09 14:32 | 1554834696.5289 | -185 | 0.20943 | 0.000449 | 2e-06 | 99.8096 | 100.1463 | 123.7746 | 20.4571 | 0.001768 | 0.001634 | -154.1622 | 0.22964 | 8.4006 | 0.002419 | 0.004015 | 0.003843 | 4.4e-05 | 0.00121 | 0.30055 | 1.3703 | 2.5475 | 98.56 | 20.06 | 49.838 | 63.5979 | 79.2332 | 60.3663 | 0.011879 | 1.0586 | 0.024551 | 23.1926 | -0.03012 | 0.13661 | -8e-06 | -0.000699 | -0.000699 | -7.8815
2019-04-09 14:32 | 1554834697.5296 | -184 | 0.20951 | 0.000447 | 2e-06 | 99.6994 | 100.1927 | 123.5378 | 20.4565 | 0.001773 | 0.001686 | -154.1515 | 0.23009 | 8.4027 | 0.002338 | 0.003806 | 0.003704 | 4.4e-05 | 0.001211 | 0.25393 | 1.3536 | 2.5476 | 98.56 | 20.06 | 49.836 | 63.7013 | 77.2868 | 61.6173 | 0.011879 | 1.0546 | 0.024481 | 23.2138 | -0.044148 | 0.13688 | -6e-06 | -0.000575 | -0.000575 | -9.5989
2019-04-09 14:32 | 1554834698.5286 | -183 | 0.20951 | 0.00045 | 2e-06 | 99.6643 | 99.8496 | 123.5244 | 20.4566 | 0.001818 | 0.001634 | -154.1557 | 0.2296 | 8.4023 | 0.002457 | 0.004157 | 0.004014 | 4.4e-05 | 0.001211 | 0.27079 | 1.3852 | 2.5477 | 98.56 | 20.06 | 49.836 | 64.8013 | 80.1549 | 60.3542 | 0.011879 | 1.0586 | 0.024526 | 23.1687 | -0.016617 | 0.13659 | -9e-06 | -0.00083 | -0.00083 | -6.6288
2019-04-09 14:32 | 1554834699.5296 | -182 | 0.20947 | 0.000449 | 3e-06 | 99.6321 | 99.9195 | 123.7596 | 20.4568 | 0.001779 | 0.001685 | -154.1186 | 0.22902 | 8.403 | 0.002464 | 0.004554 | 0.004373 | 4.2e-05 | 0.00121 | 0.30973 | 1.3675 | 2.5476 | 98.56 | 20.07 | 49.835 | 63.8414 | 80.3073 | 61.5939 | 0.011886 | 1.0547 | 0.024504 | 23.2338 | -0.02576 | 0.13624 | -9e-06 | -0.000796 | -0.000796 | -6.9
2019-04-09 14:32 | 1554834700.5297 | -181 | 0.20946 | 0.000452 | 3e-06 | 99.7066 | 100.1308 | 123.2411 | 20.4574 | 0.00179 | 0.001724 | -154.1376 | 0.22952 | 8.4029 | 0.002513 | 0.004061 | 0.003882 | 4.7e-05 | 0.001203 | 0.2896 | 1.3519 | 2.5476 | 98.56 | 20.07 | 49.834 | 64.1245 | 81.5016 | 62.5361 | 0.011885 | 1.0517 | 0.024418 | 23.2177 | -0.037975 | 0.13654 | -1.2e-05 | -0.001065 | -0.001065 | -5.1672
2019-04-09 14:32 | 1554834701.5287 | -180 | 0.20945 | 0.000452 | 2e-06 | 99.6872 | 100.1187 | 122.8805 | 20.4568 | 0.001742 | 0.001674 | -154.1128 | 0.2304 | 8.4003 | 0.002441 | 0.004073 | 0.003878 | 4.5e-05 | 0.001202 | 0.28584 | 1.3542 | 2.5476 | 98.56 | 20.07 | 49.835 | 62.9682 | 79.7529 | 61.3136 | 0.011886 | 1.0556 | 0.024427 | 23.1414 | -0.038636 | 0.13707 | -1.3e-05 | -0.001056 | -0.001056 | -5.2317
2019-04-09 14:32 | 1554834702.5288 | -179 | 0.20947 | 0.000454 | 2e-06 | 99.6699 | 100.0499 | 123.0578 | 20.4566 | 0.001768 | 0.001692 | -154.1027 | 0.22986 | 8.4007 | 0.002526 | 0.003528 | 0.00344 | 4.7e-05 | 0.001205 | 0.2943 | 1.3789 | 2.5475 | 98.56 | 20.07 | 49.835 | 63.5859 | 81.7956 | 61.7595 | 0.011886 | 1.0542 | 0.024428 | 23.1736 | -0.034034 | 0.13675 | -1.6e-05 | -0.001285 | -0.001285 | -4.288
2019-04-09 14:32 | 1554834703.5293 | -178 | 0.20948 | 0.000447 | 2e-06 | 99.6229 | 100.0647 | 123.1451 | 20.4564 | 0.001832 | 0.001626 | -154.1131 | 0.22949 | 8.401 | 0.002483 | 0.003619 | 0.003482 | 4.8e-05 | 0.001202 | 0.28008 | 1.4045 | 2.5476 | 98.56 | 20.07 | 49.836 | 65.1283 | 80.7692 | 60.175 | 0.011886 | 1.0592 | 0.024495 | 23.1269 | -0.039581 | 0.13653 | -6e-06 | -0.000573 | -0.000573 | -9.6059
2019-04-09 14:32 | 1554834704.5289 | -177 | 0.20943 | 0.000455 | 2e-06 | 99.6297 | 99.9584 | 122.7421 | 20.4565 | 0.001745 | 0.001676 | -154.0923 | 0.22851 | 8.4012 | 0.002568 | 0.004246 | 0.004047 | 4.4e-05 | 0.001203 | 0.31078 | 1.3649 | 2.5476 | 98.56 | 20.07 | 49.839 | 63.0256 | 82.8164 | 61.3679 | 0.011887 | 1.0554 | 0.024411 | 23.1303 | -0.029455 | 0.13594 | -1.6e-05 | -0.001282 | -0.001282 | -4.2718
2019-04-09 14:32 | 1554834705.53 | -176 | 0.20948 | 0.000453 | 2e-06 | 99.6327 | 100.0554 | 122.8862 | 20.4575 | 0.001788 | 0.001669 | -154.0964 | 0.22977 | 8.4006 | 0.002495 | 0.003576 | 0.00342 | 4.4e-05 | 0.001199 | 0.30198 | 1.3667 | 2.5476 | 98.56 | 20.07 | 49.839 | 64.0676 | 81.0636 | 61.1938 | 0.011887 | 1.0559 | 0.024432 | 23.1378 | -0.037868 | 0.13669 | -1.4e-05 | -0.001158 | -0.001158 | -4.7588
2019-04-09 14:32 | 1554834706.5303 | -175 | 0.20952 | 0.000453 | 3e-06 | 99.63 | 100.0571 | 123.1812 | 20.4583 | 0.001782 | 0.001734 | -154.0873 | 0.23035 | 8.402 | 0.002532 | 0.003898 | 0.003721 | 4.7e-05 | 0.001186 | 0.30558 | 1.3338 | 2.5475 | 98.56 | 20.07 | 49.842 | 63.9177 | 81.951 | 62.7611 | 0.011887 | 1.051 | 0.024404 | 23.2198 | -0.038262 | 0.13704 | -1.4e-05 | -0.001173 | -0.001173 | -4.7076
2019-04-09 14:32 | 1554834707.5293 | -174 | 0.20949 | 0.000449 | 3e-06 | 99.5973 | 100.1759 | 123.3059 | 20.4574 | 0.001755 | 0.001659 | -154.0902 | 0.22992 | 8.4022 | 0.002611 | 0.003551 | 0.003394 | 4.5e-05 | 0.001184 | 0.28672 | 1.3517 | 2.5475 | 98.56 | 20.07 | 49.843 | 63.2725 | 83.8393 | 60.9646 | 0.011888 | 1.0567 | 0.024482 | 23.1694 | -0.051811 | 0.13678 | -9e-06 | -0.000793 | -0.000793 | -6.9523
2019-04-09 14:32 | 1554834708.5289 | -173 | 0.20951 | 0.00045 | 2e-06 | 99.6098 | 99.914 | 123.1338 | 20.4572 | 0.001749 | 0.001626 | -154.0769 | 0.22928 | 8.3992 | 0.002523 | 0.003685 | 0.003526 | 4.4e-05 | 0.001181 | 0.27793 | 1.3311 | 2.5474 | 98.56 | 20.07 | 49.843 | 63.131 | 81.7245 | 60.1615 | 0.011888 | 1.0592 | 0.024494 | 23.1254 | -0.027273 | 0.1364 | -1e-05 | -0.000833 | -0.000833 | -6.5965
2019-04-09 14:32 | 1554834709.53 | -172 | 0.2095 | 0.000448 | 2e-06 | 99.5422 | 99.9573 | 123.1831 | 20.4583 | 0.001816 | 0.001669 | -154.102 | 0.22928 | 8.4032 | 0.002596 | 0.003888 | 0.003635 | 4.5e-05 | 0.001187 | 0.29142 | 1.3424 | 2.5475 | 98.56 | 20.07 | 49.842 | 64.7351 | 83.4851 | 61.2008 | 0.011887 | 1.0559 | 0.024461 | 23.166 | -0.037226 | 0.1364 | -6e-06 | -0.00052 | -0.00052 | -10.5671
2019-04-09 14:32 | 1554834710.53 | -171 | 0.2095 | 0.000451 | 3e-06 | 99.5985 | 100.0134 | 123.5692 | 20.4578 | 0.001789 | 0.001664 | -154.0812 | 0.2295 | 8.3993 | 0.002566 | 0.004172 | 0.004037 | 4.8e-05 | 0.001198 | 0.26969 | 1.3306 | 2.5475 | 98.56 | 20.07 | 49.841 | 64.1054 | 82.7631 | 61.0948 | 0.011887 | 1.0562 | 0.024503 | 23.1986 | -0.037188 | 0.13653 | -1.2e-05 | -0.001038 | -0.001038 | -5.299
2019-04-09 14:32 | 1554834711.5301 | -170 | 0.20952 | 0.00045 | 2e-06 | 99.5702 | 99.7643 | 123.611 | 20.4581 | 0.001746 | 0.001672 | -154.1158 | 0.22886 | 8.4008 | 0.002593 | 0.003786 | 0.003653 | 4.2e-05 | 0.001197 | 0.28744 | 1.3389 | 2.5475 | 98.56 | 20.07 | 49.841 | 63.0484 | 83.4109 | 61.2718 | 0.011887 | 1.0557 | 0.024501 | 23.2087 | -0.017418 | 0.13615 | -9e-06 | -0.000808 | -0.000808 | -6.7941
2019-04-09 14:32 | 1554834712.5291 | -169 | 0.20948 | 0.000449 | 2e-06 | 99.5654 | 99.8596 | 123.5117 | 20.4586 | 0.001812 | 0.00166 | -154.1203 | 0.22945 | 8.402 | 0.002567 | 0.003701 | 0.003537 | 4.4e-05 | 0.001194 | 0.28047 | 1.337 | 2.5475 | 98.56 | 20.07 | 49.838 | 64.6496 | 82.7937 | 60.9964 | 0.011886 | 1.0566 | 0.024501 | 23.1898 | -0.026389 | 0.1365 | -9e-06 | -0.000774 | -0.000774 | -7.1065
2019-04-09 14:32 | 1554834713.5291 | -168 | 0.20951 | 0.00045 | 2e-06 | 99.5157 | 99.8035 | 123.7734 | 20.459 | 0.001794 | 0.001665 | -154.1226 | 0.22965 | 8.4018 | 0.002534 | 0.004231 | 0.004154 | 4.5e-05 | 0.001186 | 0.2943 | 1.351 | 2.5475 | 98.56 | 20.07 | 49.837 | 64.2089 | 81.9969 | 61.1061 | 0.011886 | 1.0562 | 0.024523 | 23.2182 | -0.025832 | 0.13662 | -1e-05 | -0.000848 | -0.000848 | -6.4943
2019-04-09 14:32 | 1554834714.5292 | -167 | 0.20949 | 0.000447 | 2e-06 | 99.5264 | 99.7798 | 124.1281 | 20.4589 | 0.001793 | 0.001677 | -154.147 | 0.22985 | 8.4 | 0.002554 | 0.004324 | 0.004106 | 4.5e-05 | 0.001186 | 0.27743 | 1.32 | 2.5475 | 98.56 | 20.07 | 49.838 | 64.1886 | 82.4865 | 61.41 | 0.011886 | 1.0553 | 0.024547 | 23.262 | -0.02274 | 0.13674 | -6e-06 | -0.000577 | -0.000577 | -9.5472
2019-04-09 14:32 | 1554834715.5282 | -166 | 0.20947 | 0.000447 | 2e-06 | 99.5303 | 99.7608 | 123.9173 | 20.4584 | 0.00179 | 0.001659 | -154.1459 | 0.23013 | 8.4014 | 0.002507 | 0.003533 | 0.003359 | 4.2e-05 | 0.001183 | 0.27328 | 1.3238 | 2.5475 | 98.56 | 20.08 | 49.838 | 64.1303 | 81.3615 | 60.9634 | 0.011894 | 1.0567 | 0.024543 | 23.2267 | -0.020695 | 0.1369 | -6e-06 | -0.000552 | -0.000552 | -9.9865
2019-04-09 14:32 | 1554834716.5286 | -165 | 0.20949 | 0.000447 | 2e-06 | 99.5109 | 99.7481 | 123.8657 | 20.4598 | 0.001809 | 0.001671 | -154.1689 | 0.2299 | 8.4018 | 0.002507 | 0.003851 | 0.003636 | 4.2e-05 | 0.00119 | 0.29806 | 1.3385 | 2.5476 | 98.56 | 20.07 | 49.838 | 64.5689 | 81.3605 | 61.25 | 0.011886 | 1.0558 | 0.024527 | 23.2318 | -0.021292 | 0.13676 | -6e-06 | -0.000606 | -0.000606 | -9.0982
2019-04-09 14:32 | 1554834717.5296 | -164 | 0.2095 | 0.000443 | 2e-06 | 99.4803 | 99.8978 | 124.1652 | 20.4588 | 0.001829 | 0.001712 | -154.1588 | 0.22931 | 8.4016 | 0.002568 | 0.004124 | 0.003975 | 4.7e-05 | 0.00119 | 0.26161 | 1.3317 | 2.5474 | 98.56 | 20.08 | 49.841 | 65.0598 | 82.8204 | 62.2524 | 0.011895 | 1.0526 | 0.02452 | 23.2947 | -0.037464 | 0.13642 | -1e-06 | -0.000221 | -0.000221 | -24.8747
2019-04-09 14:32 | 1554834718.5292 | -163 | 0.20949 | 0.000446 | 2e-06 | 99.5094 | 99.8981 | 124.1316 | 20.4591 | 0.001793 | 0.001709 | -154.1777 | 0.22919 | 8.4016 | 0.002537 | 0.003783 | 0.003591 | 4.5e-05 | 0.00119 | 0.29717 | 1.3559 | 2.5477 | 98.56 | 20.08 | 49.842 | 64.194 | 82.0835 | 62.1659 | 0.011895 | 1.0529 | 0.02452 | 23.2886 | -0.034871 | 0.13634 | -5e-06 | -0.000479 | -0.000479 | -11.4802
2019-04-09 14:32 | 1554834719.5284 | -162 | 0.20949 | 0.000438 | 2e-06 | 99.4817 | 99.842 | 123.9509 | 20.4599 | 0.001761 | 0.001674 | -154.1652 | 0.23033 | 8.4015 | 0.002501 | 0.003868 | 0.003716 | 4.1e-05 | 0.001198 | 0.28102 | 1.3194 | 2.5476 | 98.56 | 20.08 | 49.845 | 63.4305 | 81.2084 | 61.3376 | 0.011895 | 1.0555 | 0.024532 | 23.2428 | -0.032337 | 0.13702 | 5e-06 | 0.000261 | 0.000261 | 21.1667
2019-04-09 14:32 | 1554834720.5284 | -161 | 0.2095 | 0.000445 | 2e-06 | 99.5565 | 99.9432 | 124.0283 | 20.4598 | 0.001783 | 0.001684 | -154.178 | 0.22983 | 8.4011 | 0.002518 | 0.003946 | 0.003843 | 4.1e-05 | 0.001202 | 0.25752 | 1.3512 | 2.5475 | 98.56 | 20.08 | 49.847 | 63.955 | 81.6169 | 61.5783 | 0.011896 | 1.0547 | 0.024531 | 23.2585 | -0.034679 | 0.13673 | -3e-06 | -0.000355 | -0.000355 | -15.5061
2019-04-09 14:32 | 1554834721.5284 | -160 | 0.20948 | 0.000441 | 2e-06 | 99.5103 | 99.9553 | 124.0849 | 20.4602 | 0.001766 | 0.001607 | -154.163 | 0.22954 | 8.402 | 0.002458 | 0.004118 | 0.003885 | 4.7e-05 | 0.001205 | 0.28423 | 1.3748 | 2.5477 | 98.56 | 20.08 | 49.849 | 63.5456 | 80.1787 | 59.7157 | 0.011896 | 1.0606 | 0.024605 | 23.199 | -0.039905 | 0.13656 | 2e-06 | -2.5e-05 | -2.5e-05 | -222.5111
2019-04-09 14:32 | 1554834722.5296 | -159 | 0.2095 | 0.000444 | 2e-06 | 99.4707 | 99.7195 | 124.2236 | 20.4606 | 0.001845 | 0.001632 | -154.1729 | 0.22962 | 8.4004 | 0.002525 | 0.003728 | 0.003596 | 4.5e-05 | 0.001212 | 0.28495 | 1.3715 | 2.5476 | 98.56 | 20.08 | 49.851 | 65.4415 | 81.7778 | 60.3157 | 0.011897 | 1.0587 | 0.024597 | 23.2328 | -0.022345 | 0.1366 | -3e-06 | -0.000346 | -0.000346 | -15.8943
2019-04-09 14:32 | 1554834723.5298 | -158 | 0.20957 | 0.00044 | 2e-06 | 99.4549 | 99.6744 | 123.8758 | 20.4602 | 0.001832 | 0.001708 | -154.1635 | 0.22886 | 8.4009 | 0.002526 | 0.003852 | 0.003655 | 4.4e-05 | 0.001206 | 0.28208 | 1.3615 | 2.5476 | 98.56 | 20.08 | 49.852 | 65.122 | 81.8007 | 62.1435 | 0.011897 | 1.0529 | 0.024495 | 23.2638 | -0.01972 | 0.13615 | 3e-06 | 7.8e-05 | 7.8e-05 | 70.5995
2019-04-09 14:32 | 1554834724.5285 | -157 | 0.2095 | 0.000446 | 2e-06 | 99.4207 | 99.5915 | 123.772 | 20.4599 | 0.001805 | 0.001692 | -154.1602 | 0.22988 | 8.4022 | 0.002501 | 0.003523 | 0.003393 | 4.4e-05 | 0.001209 | 0.2756 | 1.3404 | 2.5475 | 98.56 | 20.08 | 49.853 | 64.4713 | 81.2164 | 61.7634 | 0.011897 | 1.0541 | 0.024499 | 23.2408 | -0.015354 | 0.13676 | -4e-06 | -0.000419 | -0.000419 | -13.169
2019-04-09 14:32 | 1554834725.5301 | -156 | 0.20946 | 0.000445 | 2e-06 | 99.3611 | 99.574 | 123.6957 | 20.4603 | 0.001734 | 0.001743 | -154.1468 | 0.22882 | 8.4014 | 0.002511 | 0.004025 | 0.003874 | 4.6e-05 | 0.001212 | 0.26681 | 1.3085 | 2.5476 | 98.56 | 20.08 | 49.852 | 62.7623 | 81.4506 | 62.9958 | 0.011897 | 1.0503 | 0.024447 | 23.2764 | -0.019137 | 0.13613 | -3e-06 | -0.00035 | -0.00035 | -15.681
2019-04-09 14:32 | 1554834726.5299 | -155 | 0.20948 | 0.000443 | 1e-06 | 99.3302 | 99.7061 | 123.518 | 20.4604 | 0.001764 | 0.001706 | -154.1268 | 0.2296 | 8.402 | 0.00251 | 0.003758 | 0.00362 | 4.4e-05 | 0.001201 | 0.28296 | 1.3545 | 2.5475 | 98.56 | 20.08 | 49.85 | 63.4863 | 81.4218 | 62.0952 | 0.011897 | 1.0531 | 0.024462 | 23.2285 | -0.033789 | 0.13659 | 0 | -5.2e-05 | -5.2e-05 | -106.1745
2019-04-09 14:32 | 1554834727.528 | -154 | 0.20951 | 0.000445 | 2e-06 | 99.4686 | 99.6724 | 123.3192 | 20.4606 | 0.001732 | 0.001629 | -154.1292 | 0.23004 | 8.3998 | 0.002485 | 0.003815 | 0.00362 | 4.2e-05 | 0.001198 | 0.28495 | 1.3235 | 2.5476 | 98.56 | 20.08 | 49.842 | 62.7352 | 80.8178 | 60.2415 | 0.011895 | 1.059 | 0.02451 | 23.1455 | -0.018302 | 0.13685 | -3e-06 | -0.000367 | -0.000367 | -15.0175
2019-04-09 14:32 | 1554834728.5286 | -153 | 0.20951 | 0.000447 | 2e-06 | 99.3805 | 99.7522 | 123.63 | 20.4605 | 0.001762 | 0.001648 | -154.1273 | 0.22942 | 8.4008 | 0.002536 | 0.003366 | 0.003203 | 4.6e-05 | 0.001195 | 0.27295 | 1.2982 | 2.5475 | 98.56 | 20.08 | 49.838 | 63.4421 | 82.0607 | 60.6935 | 0.011894 | 1.0575 | 0.024524 | 23.1904 | -0.033395 | 0.13648 | -6e-06 | -0.000562 | -0.000562 | -9.7863
2019-04-09 14:32 | 1554834729.5276 | -152 | 0.2095 | 0.000444 | 1e-06 | 99.3859 | 99.7367 | 123.7812 | 20.4619 | 0.0018 | 0.001614 | -154.1257 | 0.22917 | 8.4018 | 0.002499 | 0.003694 | 0.0036 | 4.4e-05 | 0.001199 | 0.29574 | 1.2998 | 2.5475 | 98.56 | 20.08 | 49.832 | 64.3592 | 81.1644 | 59.8825 | 0.011892 | 1.0601 | 0.024569 | 23.1764 | -0.031523 | 0.13634 | -1e-06 | -0.000189 | -0.000189 | -29.0061
2019-04-09 14:32 | 1554834730.5289 | -151 | 0.2095 | 0.000445 | 1e-06 | 99.4353 | 99.7199 | 123.4469 | 20.4623 | 0.001792 | 0.001673 | -154.1196 | 0.22976 | 8.4012 | 0.002474 | 0.004515 | 0.004308 | 4.6e-05 | 0.001199 | 0.2724 | 1.3488 | 2.5476 | 98.56 | 20.08 | 49.819 | 64.1589 | 80.5565 | 61.3085 | 0.011889 | 1.0556 | 0.024484 | 23.1945 | -0.025563 | 0.13668 | -3e-06 | -0.000292 | -0.000292 | -18.8818
2019-04-09 14:32 | 1554834731.5293 | -150 | 0.20952 | 0.000444 | 2e-06 | 99.429 | 99.7849 | 123.1592 | 20.4615 | 0.001772 | 0.001665 | -154.1319 | 0.23027 | 8.4021 | 0.002521 | 0.00394 | 0.003696 | 4.5e-05 | 0.001204 | 0.28727 | 1.2844 | 2.5475 | 98.56 | 20.08 | 49.814 | 63.6961 | 81.6905 | 61.1219 | 0.011888 | 1.0562 | 0.024462 | 23.161 | -0.031958 | 0.13699 | -3e-06 | -0.000332 | -0.000332 | -16.6353
2019-04-09 14:32 | 1554834732.5277 | -149 | 0.20954 | 0.000447 | 2e-06 | 99.4174 | 99.8606 | 122.9505 | 20.4626 | 0.001762 | 0.001675 | -154.1238 | 0.22955 | 8.4005 | 0.002459 | 0.004019 | 0.003922 | 4.2e-05 | 0.001204 | 0.30718 | 1.3421 | 2.5476 | 98.56 | 20.08 | 49.802 | 63.4533 | 80.2021 | 61.3433 | 0.011885 | 1.0555 | 0.024433 | 23.1491 | -0.039787 | 0.13656 | -5e-06 | -0.00045 | -0.00045 | -12.219
2019-04-09 14:32 | 1554834733.5281 | -148 | 0.20951 | 0.000444 | 2e-06 | 99.4013 | 99.7866 | 123.0195 | 20.461 | 0.001736 | 0.001641 | -154.1224 | 0.23007 | 8.401 | 0.002466 | 0.004569 | 0.004456 | 4.3e-05 | 0.001204 | 0.29806 | 1.3319 | 2.5476 | 98.56 | 20.08 | 49.797 | 62.8215 | 80.3589 | 60.5219 | 0.011884 | 1.0581 | 0.02447 | 23.1271 | -0.034604 | 0.13687 | -1e-06 | -0.000154 | -0.000154 | -35.8241
2019-04-09 14:32 | 1554834734.5287 | -147 | 0.20953 | 0.000441 | 2e-06 | 99.4106 | 99.8004 | 122.9798 | 20.4621 | 0.001768 | 0.001726 | -154.1006 | 0.22958 | 8.4028 | 0.002454 | 0.003683 | 0.003585 | 4.4e-05 | 0.001212 | 0.29894 | 1.3458 | 2.5475 | 98.56 | 20.08 | 49.793 | 63.5958 | 80.0877 | 62.5741 | 0.011883 | 1.0516 | 0.024391 | 23.1944 | -0.035007 | 0.13658 | 1e-06 | -4.2e-05 | -4.2e-05 | -130.2076
2019-04-09 14:32 | 1554834735.5288 | -146 | 0.2095 | 0.000445 | 2e-06 | 99.3314 | 99.8362 | 122.8073 | 20.4622 | 0.001787 | 0.001648 | -154.1161 | 0.23025 | 8.4018 | 0.002475 | 0.003738 | 0.003534 | 4.4e-05 | 0.001218 | 0.27113 | 1.3595 | 2.5476 | 98.56 | 20.08 | 49.785 | 64.0565 | 80.579 | 60.7083 | 0.011881 | 1.0575 | 0.024442 | 23.1136 | -0.045344 | 0.13698 | -4e-06 | -0.000417 | -0.000417 | -13.2502
2019-04-09 14:32 | 1554834736.528 | -145 | 0.20952 | 0.000443 | 2e-06 | 99.3608 | 99.7849 | 123.1782 | 20.4628 | 0.001745 | 0.001703 | -154.1275 | 0.22936 | 8.3982 | 0.002535 | 0.003317 | 0.003025 | 4.5e-05 | 0.001214 | 0.30397 | 1.3593 | 2.5475 | 98.56 | 20.08 | 49.781 | 63.0498 | 82.0309 | 62.0171 | 0.01188 | 1.0533 | 0.024431 | 23.1938 | -0.038095 | 0.13644 | 0 | -0.000145 | -0.000145 | -37.8938
2019-04-09 14:32 | 1554834737.5292 | -144 | 0.20951 | 0.000443 | 2e-06 | 99.3972 | 99.7973 | 123.6824 | 20.4628 | 0.001774 | 0.00164 | -154.1309 | 0.22917 | 8.4033 | 0.002519 | 0.004168 | 0.004039 | 4.3e-05 | 0.001209 | 0.28152 | 1.3848 | 2.5476 | 98.56 | 20.08 | 49.776 | 63.7294 | 81.6559 | 60.5018 | 0.011879 | 1.0581 | 0.024536 | 23.1886 | -0.035935 | 0.13634 | 0 | -0.000147 | -0.000147 | -37.4201
2019-04-09 14:32 | 1554834738.529 | -143 | 0.20953 | 0.000442 | 2e-06 | 99.3864 | 99.6318 | 123.4857 | 20.4627 | 0.001751 | 0.001703 | -154.1389 | 0.22903 | 8.4042 | 0.002451 | 0.00428 | 0.00419 | 4.5e-05 | 0.001209 | 0.29231 | 1.3665 | 2.5475 | 98.56 | 20.08 | 49.774 | 63.184 | 80.015 | 62.0341 | 0.011879 | 1.0533 | 0.024461 | 23.2233 | -0.022051 | 0.13625 | 1e-06 | -7.2e-05 | -7.2e-05 | -76.5231
2019-04-09 14:32 | 1554834739.528 | -142 | 0.20951 | 0.000448 | 2e-06 | 99.357 | 99.7078 | 123.2861 | 20.4633 | 0.001822 | 0.001658 | -154.1467 | 0.22928 | 8.4025 | 0.002444 | 0.003931 | 0.003762 | 4.5e-05 | 0.001212 | 0.28119 | 1.3653 | 2.5475 | 98.56 | 20.08 | 49.772 | 64.8884 | 79.8331 | 60.94 | 0.011878 | 1.0567 | 0.024481 | 23.1667 | -0.03153 | 0.1364 | -6e-06 | -0.000601 | -0.000601 | -9.1424
2019-04-09 14:32 | 1554834740.529 | -141 | 0.20948 | 0.000442 | 2e-06 | 99.4469 | 99.6359 | 123.3275 | 20.4631 | 0.001741 | 0.001637 | -154.1522 | 0.22963 | 8.4031 | 0.002514 | 0.003906 | 0.003772 | 4.6e-05 | 0.001213 | 0.29551 | 1.3682 | 2.5476 | 98.56 | 20.09 | 49.771 | 62.936 | 81.5204 | 60.4497 | 0.011885 | 1.0583 | 0.024503 | 23.1535 | -0.016983 | 0.13661 | 1e-06 | -6.7e-05 | -6.7e-05 | -81.5638
2019-04-09 14:32 | 1554834741.5281 | -140 | 0.2095 | 0.000443 | 2e-06 | 99.4272 | 99.9054 | 123.7788 | 20.4631 | 0.00182 | 0.00168 | -154.1307 | 0.22903 | 8.4013 | 0.002431 | 0.003851 | 0.003689 | 4.8e-05 | 0.001216 | 0.28064 | 1.3562 | 2.5476 | 98.56 | 20.09 | 49.771 | 64.849 | 79.5211 | 61.4799 | 0.011885 | 1.055 | 0.02451 | 23.2316 | -0.042909 | 0.13625 | -1e-06 | -0.000219 | -0.000219 | -25.0226
2019-04-09 14:32 | 1554834742.5286 | -139 | 0.20953 | 0.000441 | 2e-06 | 99.4127 | 99.9147 | 123.8765 | 20.4634 | 0.001793 | 0.001699 | -154.1298 | 0.22915 | 8.402 | 0.002414 | 0.003784 | 0.00357 | 4.4e-05 | 0.001211 | 0.29336 | 1.3333 | 2.5475 | 98.56 | 20.08 | 49.773 | 64.203 | 79.1104 | 61.9255 | 0.011878 | 1.0536 | 0.024504 | 23.2563 | -0.045054 | 0.13632 | 1e-06 | -4.2e-05 | -4.2e-05 | -129.7103
2019-04-09 14:32 | 1554834743.5278 | -138 | 0.20954 | 0.000445 | 2e-06 | 99.4415 | 99.8355 | 123.7145 | 20.463 | 0.001806 | 0.001613 | -154.1199 | 0.23006 | 8.4041 | 0.002414 | 0.003627 | 0.003511 | 4.8e-05 | 0.00121 | 0.28102 | 1.3505 | 2.5475 | 98.56 | 20.08 | 49.773 | 64.5102 | 79.1194 | 59.8484 | 0.011878 | 1.0602 | 0.024564 | 23.1689 | -0.035367 | 0.13686 | -2e-06 | -0.000299 | -0.000299 | -18.4617
2019-04-09 14:32 | 1554834744.5283 | -137 | 0.20951 | 0.000443 | 2e-06 | 99.4564 | 99.8379 | 123.7136 | 20.4642 | 0.001851 | 0.001719 | -154.1397 | 0.23034 | 8.4011 | 0.002421 | 0.004233 | 0.004003 | 4.6e-05 | 0.001212 | 0.28385 | 1.3459 | 2.5475 | 98.55 | 20.08 | 49.777 | 65.5807 | 79.2915 | 62.4045 | 0.01188 | 1.0521 | 0.02447 | 23.2576 | -0.034243 | 0.13703 | -1e-06 | -0.000212 | -0.000212 | -26.0117
2019-04-09 14:32 | 1554834745.5277 | -136 | 0.20953 | 0.000446 | 2e-06 | 99.4311 | 99.7075 | 123.8204 | 20.4638 | 0.001834 | 0.001694 | -154.1375 | 0.22983 | 8.3999 | 0.002459 | 0.003798 | 0.003665 | 4.9e-05 | 0.001205 | 0.29358 | 1.364 | 2.5475 | 98.55 | 20.08 | 49.78 | 65.1706 | 80.2133 | 61.822 | 0.011881 | 1.054 | 0.024502 | 23.2474 | -0.024827 | 0.13673 | -4e-06 | -0.000421 | -0.000421 | -13.0943
2019-04-09 14:32 | 1554834746.5283 | -135 | 0.20952 | 0.000442 | 2e-06 | 99.4975 | 99.7763 | 123.6934 | 20.4648 | 0.001757 | 0.001708 | -154.1272 | 0.22938 | 8.4008 | 0.002436 | 0.003614 | 0.003447 | 4.9e-05 | 0.001204 | 0.30901 | 1.3959 | 2.5477 | 98.55 | 20.08 | 49.782 | 63.3211 | 79.6491 | 62.1471 | 0.011882 | 1.0529 | 0.024477 | 23.2468 | -0.025027 | 0.13646 | 2e-06 | 2.2e-05 | 2.2e-05 | 251.9294
2019-04-09 14:32 | 1554834747.5287 | -134 | 0.2095 | 0.000443 | 2e-06 | 99.5342 | 99.958 | 123.2526 | 20.4645 | 0.001756 | 0.001645 | -154.114 | 0.23011 | 8.3995 | 0.002378 | 0.00382 | 0.00361 | 4.9e-05 | 0.001196 | 0.30215 | 1.3597 | 2.5476 | 98.55 | 20.08 | 49.786 | 63.3086 | 78.2634 | 60.6424 | 0.011883 | 1.0577 | 0.024489 | 23.1532 | -0.03801 | 0.13689 | 0 | -0.000106 | -0.000106 | -51.8266
2019-04-09 14:32 | 1554834748.5272 | -133 | 0.20965 | 0.000445 | 3e-06 | 99.5228 | 99.809 | 123.0225 | 20.4645 | 0.001789 | 0.001681 | -154.1313 | 0.22936 | 8.4011 | 0.002414 | 0.004173 | 0.004006 | 5e-05 | 0.001196 | 0.28456 | 1.3496 | 2.5475 | 98.56 | 20.08 | 49.788 | 64.0932 | 79.1334 | 61.5039 | 0.011882 | 1.055 | 0.024434 | 23.1614 | -0.025683 | 0.13645 | -4e-06 | -0.000447 | -0.000447 | -12.3124
2019-04-09 14:32 | 1554834749.5273 | -132 | 0.20953 | 0.000446 | 3e-06 | 99.5294 | 99.9212 | 123.2192 | 20.4656 | 0.001849 | 0.001663 | -154.1328 | 0.22973 | 8.4003 | 0.002468 | 0.003759 | 0.003488 | 5e-05 | 0.001196 | 0.28191 | 1.3242 | 2.5475 | 98.56 | 20.08 | 49.793 | 65.5455 | 80.4155 | 61.0571 | 0.011883 | 1.0564 | 0.02447 | 23.1644 | -0.035143 | 0.13667 | -5e-06 | -0.000498 | -0.000498 | -11.0558
2019-04-09 14:32 | 1554834750.5273 | -131 | 0.20951 | 0.000443 | 3e-06 | 99.6104 | 100.1139 | 123.5479 | 20.465 | 0.001791 | 0.001681 | -154.1214 | 0.22892 | 8.3996 | 0.002411 | 0.004017 | 0.003936 | 5e-05 | 0.001185 | 0.28783 | 1.3504 | 2.5477 | 98.56 | 20.08 | 49.796 | 64.1531 | 79.0505 | 61.4933 | 0.011884 | 1.055 | 0.024487 | 23.2104 | -0.045102 | 0.13619 | -1e-06 | -0.000231 | -0.000231 | -23.7782
2019-04-09 14:33 | 1554834751.5287 | -130 | 0.20955 | 0.000445 | 2e-06 | 99.6321 | 99.882 | 123.4439 | 20.4645 | 0.001829 | 0.001687 | -154.1367 | 0.22985 | 8.4016 | 0.002411 | 0.003178 | 0.003002 | 5.4e-05 | 0.001183 | 0.30646 | 1.3382 | 2.5475 | 98.56 | 20.08 | 49.798 | 65.0705 | 79.0406 | 61.6539 | 0.011884 | 1.0545 | 0.024471 | 23.2062 | -0.022402 | 0.13674 | -0.000262 | -0.019307 | -0.019307 | -0.28541
2019-04-09 14:33 | 1554834752.5287 | -129 | 0.20965 | 0.00044 | 2e-06 | 99.7191 | 100.2031 | 123.412 | 20.4658 | 0.001753 | 0.001691 | -154.1351 | 0.23053 | 8.4026 | 0.002415 | 0.00403 | 0.003809 | 5e-05 | 0.001174 | 0.28727 | 1.3593 | 2.5476 | 98.56 | 20.08 | 49.804 | 63.2364 | 79.1572 | 61.7419 | 0.011886 | 1.0542 | 0.024464 | 23.2063 | -0.043307 | 0.13714 | -0.000865 | -0.063566 | -0.063566 | -0.086948
2019-04-09 14:33 | 1554834753.5285 | -128 | 0.20944 | 0.000446 | 2e-06 | 99.6872 | 99.8685 | 123.5415 | 20.4645 | 0.001736 | 0.001665 | -154.1625 | 0.2294 | 8.4006 | 0.002471 | 0.003931 | 0.003736 | 4.8e-05 | 0.001185 | 0.28224 | 1.3595 | 2.5476 | 98.56 | 20.08 | 49.807 | 62.835 | 80.4986 | 61.1083 | 0.011886 | 1.0562 | 0.0245 | 23.1965 | -0.016255 | 0.13647 | 0.000422 | 0.030842 | 0.030842 | 0.17832
2019-04-09 14:33 | 1554834754.5286 | -127 | 0.20949 | 0.000445 | 2e-06 | 99.6428 | 99.8961 | 123.4716 | 20.4658 | 0.001785 | 0.001669 | -154.1447 | 0.22983 | 8.4015 | 0.002497 | 0.003664 | 0.00357 | 4.9e-05 | 0.001185 | 0.2844 | 1.363 | 2.5476 | 98.56 | 20.08 | 49.811 | 63.9991 | 81.1137 | 61.2238 | 0.011887 | 1.0558 | 0.024489 | 23.1939 | -0.022703 | 0.13672 | 0.000114 | 0.008247 | 0.008247 | 0.66813
2019-04-09 14:33 | 1554834755.529 | -126 | 0.2095 | 0.000446 | 3e-06 | 99.7256 | 100.0358 | 123.86 | 20.4656 | 0.001819 | 0.001676 | -154.1516 | 0.22931 | 8.4032 | 0.002481 | 0.004035 | 0.003879 | 4.8e-05 | 0.001191 | 0.29446 | 1.3622 | 2.5475 | 98.56 | 20.08 | 49.813 | 64.8195 | 80.7321 | 61.3884 | 0.011888 | 1.0553 | 0.024522 | 23.2361 | -0.027777 | 0.13642 | 4.3e-05 | 0.002957 | 0.002957 | 1.8594
2019-04-09 14:33 | 1554834756.5276 | -125 | 0.20946 | 0.000448 | 2e-06 | 99.7113 | 100.1015 | 123.8227 | 20.4657 | 0.001812 | 0.00169 | -154.1453 | 0.22942 | 8.4016 | 0.00246 | 0.003774 | 0.003591 | 5e-05 | 0.001194 | 0.26985 | 1.3399 | 2.5475 | 98.56 | 20.08 | 49.813 | 64.6666 | 80.227 | 61.7122 | 0.011888 | 1.0543 | 0.024506 | 23.2438 | -0.034934 | 0.13648 | 0.000254 | 0.018481 | 0.018481 | 0.29762
2019-04-09 14:33 | 1554834757.5276 | -124 | 0.20949 | 0.000442 | 2e-06 | 99.7581 | 100.1239 | 123.569 | 20.4656 | 0.001848 | 0.001684 | -154.1677 | 0.22923 | 8.4012 | 0.002476 | 0.004068 | 0.003945 | 4.8e-05 | 0.001197 | 0.28368 | 1.3006 | 2.5475 | 98.55 | 20.09 | 49.813 | 65.5261 | 80.6198 | 61.5791 | 0.011896 | 1.0547 | 0.024486 | 23.2154 | -0.03274 | 0.13637 | 0.000124 | 0.008977 | 0.008977 | 0.61218
2019-04-09 14:33 | 1554834758.5276 | -123 | 0.20952 | 0.000441 | 3e-06 | 99.7417 | 100.0909 | 123.945 | 20.4661 | 0.001778 | 0.001688 | -154.1343 | 0.2306 | 8.4004 | 0.002412 | 0.003952 | 0.003814 | 4.8e-05 | 0.001195 | 0.27671 | 1.3421 | 2.5475 | 98.55 | 20.08 | 49.814 | 63.8386 | 79.0703 | 61.6617 | 0.011889 | 1.0545 | 0.02452 | 23.2536 | -0.031257 | 0.13718 | -8e-05 | -0.006073 | -0.006073 | -0.91034
2019-04-09 14:33 | 1554834759.5277 | -122 | 0.2095 | 0.000445 | 3e-06 | 99.826 | 100.0055 | 123.8922 | 20.4658 | 0.001816 | 0.00173 | -154.1362 | 0.23022 | 8.4041 | 0.002413 | 0.003666 | 0.003523 | 4.4e-05 | 0.001204 | 0.29894 | 1.3143 | 2.5474 | 98.55 | 20.08 | 49.815 | 64.7417 | 79.1106 | 62.6919 | 0.01189 | 1.0512 | 0.024477 | 23.2843 | -0.016072 | 0.13696 | 5.6e-05 | 0.003979 | 0.003979 | 1.3872
2019-04-09 14:33 | 1554834760.5282 | -121 | 0.20951 | 0.000443 | 2e-06 | 99.7575 | 100.1686 | 124.1131 | 20.4654 | 0.001771 | 0.00166 | -154.144 | 0.22932 | 8.4015 | 0.00246 | 0.003039 | 0.003007 | 4.5e-05 | 0.001205 | 0.30093 | 1.3124 | 2.5476 | 98.55 | 20.09 | 49.814 | 63.6567 | 80.2219 | 61.0006 | 0.011897 | 1.0565 | 0.024561 | 23.2463 | -0.036789 | 0.13642 | 5e-06 | 0.000222 | 0.000222 | 24.7524
2019-04-09 14:33 | 1554834761.5278 | -120 | 0.20953 | 0.000447 | 2e-06 | 99.8004 | 100.3153 | 123.9161 | 20.4658 | 0.00175 | 0.001726 | -154.1381 | 0.22962 | 8.4031 | 0.002518 | 0.00436 | 0.004116 | 4.7e-05 | 0.001205 | 0.29983 | 1.3391 | 2.5475 | 98.55 | 20.09 | 49.812 | 63.1738 | 81.6268 | 62.5739 | 0.011896 | 1.0516 | 0.024484 | 23.2825 | -0.046027 | 0.1366 | -0.000134 | -0.009947 | -0.009947 | -0.55345
2019-04-09 14:33 | 1554834762.5277 | -119 | 0.20952 | 0.000449 | 2e-06 | 99.8194 | 100.2003 | 123.9469 | 20.4656 | 0.00181 | 0.001688 | -154.148 | 0.22865 | 8.4003 | 0.00249 | 0.003609 | 0.003487 | 4.2e-05 | 0.001211 | 0.28423 | 1.367 | 2.5478 | 98.55 | 20.09 | 49.81 | 64.5947 | 80.938 | 61.6689 | 0.011896 | 1.0544 | 0.02452 | 23.254 | -0.034064 | 0.13603 | -0.000105 | -0.007859 | -0.007859 | -0.69754
2019-04-09 14:33 | 1554834763.5268 | -118 | 0.20951 | 0.000443 | 3e-06 | 99.8534 | 100.3166 | 123.9427 | 20.4673 | 0.001792 | 0.001693 | -154.1501 | 0.22968 | 8.4021 | 0.002429 | 0.003765 | 0.003577 | 4.5e-05 | 0.001199 | 0.29391 | 1.375 | 2.5476 | 98.55 | 20.09 | 49.808 | 64.1668 | 79.4737 | 61.7855 | 0.011895 | 1.0541 | 0.024515 | 23.2576 | -0.0414 | 0.13664 | -3e-06 | -0.000366 | -0.000366 | -15.0561
2019-04-09 14:33 | 1554834764.5269 | -117 | 0.2095 | 0.000446 | 2e-06 | 99.8221 | 100.1804 | 123.9105 | 20.4664 | 0.001799 | 0.00168 | -154.1573 | 0.22938 | 8.4027 | 0.002498 | 0.00356 | 0.003361 | 4.3e-05 | 0.001197 | 0.28567 | 1.3612 | 2.5477 | 98.55 | 20.08 | 49.806 | 64.3489 | 81.147 | 61.4768 | 0.011887 | 1.055 | 0.024523 | 23.2439 | -0.032042 | 0.13646 | 6.2e-05 | 0.004464 | 0.004464 | 1.2319
2019-04-09 14:33 | 1554834765.5268 | -116 | 0.20949 | 0.00044 | 2e-06 | 99.8841 | 100.3717 | 123.6321 | 20.4671 | 0.001777 | 0.001734 | -154.1323 | 0.23028 | 8.4019 | 0.002378 | 0.003524 | 0.003344 | 4.6e-05 | 0.001198 | 0.29087 | 1.3424 | 2.5476 | 98.55 | 20.08 | 49.8 | 63.8045 | 78.2526 | 62.7807 | 0.011886 | 1.0509 | 0.024448 | 23.263 | -0.043562 | 0.13699 | 0.000124 | 0.008964 | 0.008964 | 0.61588
2019-04-09 14:33 | 1554834766.5269 | -115 | 0.20951 | 0.000448 | 2e-06 | 99.9004 | 100.3376 | 123.9659 | 20.4671 | 0.001806 | 0.001613 | -154.1519 | 0.22976 | 8.4023 | 0.002384 | 0.003274 | 0.003211 | 4.2e-05 | 0.001204 | 0.27577 | 1.3169 | 2.5477 | 98.55 | 20.08 | 49.798 | 64.5033 | 78.4117 | 59.8545 | 0.011885 | 1.0602 | 0.024588 | 23.1927 | -0.039058 | 0.13668 | -1.2e-05 | -0.001001 | -0.001001 | -5.5024
2019-04-09 14:33 | 1554834767.528 | -114 | 0.20949 | 0.000444 | 2e-06 | 99.8266 | 100.2629 | 124.0711 | 20.4672 | 0.001729 | 0.001651 | -154.1459 | 0.22938 | 8.401 | 0.002389 | 0.003904 | 0.003625 | 4.6e-05 | 0.001206 | 0.27671 | 1.3487 | 2.5474 | 98.56 | 20.08 | 49.8 | 62.658 | 78.5343 | 60.7741 | 0.011885 | 1.0573 | 0.024565 | 23.2345 | -0.039013 | 0.13646 | 8.8e-05 | 0.006297 | 0.006297 | 0.87328
2019-04-09 14:33 | 1554834768.527 | -113 | 0.20951 | 0.000447 | 2e-06 | 99.8453 | 100.3383 | 123.7765 | 20.4669 | 0.001787 | 0.001706 | -154.1585 | 0.23031 | 8.3989 | 0.002427 | 0.00423 | 0.004085 | 4.6e-05 | 0.001209 | 0.28926 | 1.3579 | 2.5476 | 98.56 | 20.08 | 49.802 | 64.0579 | 79.4445 | 62.0938 | 0.011885 | 1.0531 | 0.024487 | 23.2527 | -0.044054 | 0.13701 | -4.2e-05 | -0.003256 | -0.003256 | -1.6957
2019-04-09 14:33 | 1554834769.5285 | -112 | 0.20947 | 0.000444 | 2e-06 | 99.9019 | 100.1218 | 123.7232 | 20.467 | 0.001821 | 0.001684 | -154.1481 | 0.23038 | 8.401 | 0.002428 | 0.00424 | 0.004118 | 4.3e-05 | 0.001212 | 0.27743 | 1.3149 | 2.5475 | 98.56 | 20.08 | 49.804 | 64.8615 | 79.4588 | 61.5687 | 0.011886 | 1.0548 | 0.024501 | 23.2295 | -0.019668 | 0.13706 | 0.000241 | 0.017589 | 0.017589 | 0.31402
2019-04-09 14:33 | 1554834770.5285 | -111 | 0.20948 | 0.000444 | 3e-06 | 99.8602 | 100.2168 | 123.4023 | 20.4675 | 0.001762 | 0.001653 | -154.1555 | 0.22964 | 8.4017 | 0.002405 | 0.003989 | 0.003819 | 4.3e-05 | 0.001215 | 0.27649 | 1.3472 | 2.5475 | 98.55 | 20.08 | 49.806 | 63.4489 | 78.9196 | 60.8173 | 0.011887 | 1.0571 | 0.024497 | 23.1733 | -0.031884 | 0.13661 | 0.000171 | 0.012418 | 0.012418 | 0.44334
2019-04-09 14:33 | 1554834771.528 | -110 | 0.20953 | 0.000444 | 3e-06 | 99.8269 | 100.2537 | 123.3155 | 20.4682 | 0.001718 | 0.001654 | -154.1409 | 0.22943 | 8.4025 | 0.002432 | 0.00429 | 0.004084 | 4.6e-05 | 0.001203 | 0.28998 | 1.3494 | 2.5476 | 98.55 | 20.08 | 49.809 | 62.3858 | 79.5472 | 60.8556 | 0.011888 | 1.057 | 0.024487 | 23.1665 | -0.038158 | 0.13649 | -0.000133 | -0.009938 | -0.009938 | -0.55349
2019-04-09 14:33 | 1554834772.5285 | -109 | 0.20951 | 0.000444 | 3e-06 | 99.8299 | 100.3159 | 123.5486 | 20.4685 | 0.001803 | 0.001685 | -154.1473 | 0.22943 | 8.3991 | 0.002387 | 0.00366 | 0.00349 | 4.5e-05 | 0.001212 | 0.29407 | 1.3844 | 2.5477 | 98.55 | 20.08 | 49.816 | 64.4453 | 78.4662 | 61.6032 | 0.01189 | 1.0546 | 0.024483 | 23.2143 | -0.043447 | 0.13649 | 3e-06 | 3.5e-05 | 3.5e-05 | 157.9465
2019-04-09 14:33 | 1554834773.5281 | -108 | 0.20952 | 0.000444 | 2e-06 | 99.8981 | 100.1243 | 123.5018 | 20.4669 | 0.001798 | 0.00168 | -154.1224 | 0.23014 | 8.4011 | 0.002377 | 0.00348 | 0.003382 | 4.2e-05 | 0.001216 | 0.30287 | 1.3317 | 2.5475 | 98.55 | 20.08 | 49.818 | 64.3087 | 78.2354 | 61.4773 | 0.01189 | 1.055 | 0.024483 | 23.2056 | -0.02023 | 0.13691 | -7.3e-05 | -0.005484 | -0.005484 | -1.0062
2019-04-09 14:33 | 1554834774.5276 | -107 | 0.20953 | 0.000444 | 2e-06 | 99.8069 | 100.2017 | 123.5406 | 20.4683 | 0.001798 | 0.001686 | -154.1223 | 0.22944 | 8.4029 | 0.002413 | 0.00355 | 0.00342 | 4.3e-05 | 0.001216 | 0.26753 | 1.3615 | 2.5477 | 98.55 | 20.07 | 49.821 | 64.332 | 79.1131 | 61.635 | 0.011884 | 1.0545 | 0.024481 | 23.2147 | -0.035308 | 0.13649 | -0.000166 | -0.012306 | -0.012306 | -0.44699
2019-04-09 14:33 | 1554834775.5263 | -106 | 0.20948 | 0.000445 | 2e-06 | 99.8334 | 100.2247 | 123.4504 | 20.4681 | 0.001787 | 0.001657 | -154.1291 | 0.23019 | 8.4029 | 0.002413 | 0.004069 | 0.003836 | 4.3e-05 | 0.001219 | 0.27975 | 1.3225 | 2.5476 | 98.55 | 20.07 | 49.826 | 64.0451 | 79.0933 | 60.9172 | 0.011885 | 1.0568 | 0.024498 | 23.1813 | -0.034991 | 0.13694 | 0.000141 | 0.010226 | 0.010226 | 0.53967
2019-04-09 14:33 | 1554834776.5262 | -105 | 0.20951 | 0.000444 | 2e-06 | 99.8903 | 100.3418 | 123.0155 | 20.4697 | 0.001797 | 0.001651 | -154.1162 | 0.23043 | 8.4006 | 0.002391 | 0.003639 | 0.003462 | 4.3e-05 | 0.001216 | 0.31006 | 1.3437 | 2.5476 | 98.55 | 20.07 | 49.829 | 64.3006 | 78.5764 | 60.785 | 0.011885 | 1.0572 | 0.02446 | 23.1358 | -0.040333 | 0.13708 | 8e-06 | 0.000407 | 0.000407 | 13.5647
2019-04-09 14:33 | 1554834777.5262 | -104 | 0.2095 | 0.000448 | 2e-06 | 99.9168 | 100.414 | 123.4868 | 20.4685 | 0.001789 | 0.001634 | -154.1357 | 0.22976 | 8.4013 | 0.002384 | 0.003789 | 0.003616 | 4.7e-05 | 0.001204 | 0.29142 | 1.3526 | 2.5475 | 98.55 | 20.07 | 49.831 | 64.1072 | 78.3999 | 60.3791 | 0.011886 | 1.0585 | 0.024522 | 23.166 | -0.0444 | 0.13668 | 3.9e-05 | 0.002747 | 0.002747 | 2.0053
2019-04-09 14:33 | 1554834778.5277 | -103 | 0.20949 | 0.000449 | 3e-06 | 99.9493 | 100.2925 | 123.5676 | 20.4686 | 0.001858 | 0.001662 | -154.1197 | 0.22972 | 8.4028 | 0.002461 | 0.00419 | 0.004008 | 4.6e-05 | 0.001195 | 0.274 | 1.3724 | 2.5475 | 98.55 | 20.08 | 49.833 | 65.7567 | 80.2593 | 61.0374 | 0.011894 | 1.0564 | 0.024505 | 23.1965 | -0.030667 | 0.13666 | 0.000104 | 0.00749 | 0.00749 | 0.73534
2019-04-09 14:33 | 1554834779.5273 | -102 | 0.20951 | 0.000442 | 2e-06 | 99.9237 | 100.2172 | 123.4847 | 20.4684 | 0.001812 | 0.0017 | -154.1326 | 0.22987 | 8.4034 | 0.002503 | 0.003395 | 0.00328 | 4.6e-05 | 0.001182 | 0.29142 | 1.3399 | 2.5475 | 98.55 | 20.08 | 49.831 | 64.6685 | 81.2722 | 61.9663 | 0.011893 | 1.0535 | 0.024463 | 23.2209 | -0.026235 | 0.13675 | -1.6e-05 | -0.001328 | -0.001328 | -4.1514
2019-04-09 14:33 | 1554834780.5268 | -101 | 0.2095 | 0.000445 | 2e-06 | 99.938 | 100.3486 | 123.3432 | 20.4689 | 0.001847 | 0.001687 | -154.1091 | 0.23023 | 8.4008 | 0.00243 | 0.004318 | 0.004135 | 4.8e-05 | 0.001173 | 0.29551 | 1.3332 | 2.5476 | 98.55 | 20.08 | 49.826 | 65.4886 | 79.5173 | 61.6488 | 0.011892 | 1.0545 | 0.024461 | 23.1966 | -0.036681 | 0.13697 | 7.1e-05 | 0.005058 | 0.005058 | 1.0914
2019-04-09 14:33 | 1554834781.5268 | -100 | 0.20951 | 0.000441 | 2e-06 | 99.9222 | 100.2626 | 123.4565 | 20.4687 | 0.001793 | 0.00165 | -154.1262 | 0.23001 | 8.4015 | 0.00256 | 0.003394 | 0.003249 | 5e-05 | 0.001164 | 0.31238 | 1.3475 | 2.5476 | 98.55 | 20.08 | 49.824 | 64.2106 | 82.6412 | 60.7599 | 0.011892 | 1.0573 | 0.024505 | 23.1764 | -0.030422 | 0.13684 | -2.9e-05 | -0.002262 | -0.002262 | -2.4377
2019-04-09 14:33 | 1554834782.5275 | -99 | 0.20951 | 0.000443 | 2e-06 | 99.8567 | 100.2891 | 123.0847 | 20.4689 | 0.001828 | 0.001684 | -154.0991 | 0.22948 | 8.4011 | 0.002574 | 0.003431 | 0.003229 | 4.7e-05 | 0.00116 | 0.27041 | 1.371 | 2.5476 | 98.55 | 20.08 | 49.822 | 65.0383 | 82.9737 | 61.5737 | 0.011891 | 1.0547 | 0.024438 | 23.1697 | -0.038652 | 0.13652 | -1.1e-05 | -0.000913 | -0.000913 | -6.0249
2019-04-09 14:33 | 1554834783.5279 | -98 | 0.20948 | 0.000448 | 3e-06 | 99.8698 | 100.1951 | 123.3509 | 20.4686 | 0.001803 | 0.001657 | -154.0948 | 0.22936 | 8.4021 | 0.002585 | 0.003764 | 0.003687 | 5.2e-05 | 0.001157 | 0.30254 | 1.3263 | 2.5476 | 98.55 | 20.08 | 49.816 | 64.4384 | 83.2277 | 60.9216 | 0.01189 | 1.0568 | 0.024488 | 23.1721 | -0.029095 | 0.13644 | 0.000158 | 0.011404 | 0.011404 | 0.48217
2019-04-09 14:33 | 1554834784.5264 | -97 | 0.20945 | 0.000448 | 2e-06 | 99.9106 | 100.2196 | 123.3967 | 20.4685 | 0.001836 | 0.001682 | -154.0912 | 0.22918 | 8.4024 | 0.002531 | 0.003909 | 0.003737 | 4.9e-05 | 0.001157 | 0.26897 | 1.3074 | 2.5475 | 98.55 | 20.08 | 49.813 | 65.2448 | 81.9381 | 61.5194 | 0.011889 | 1.0549 | 0.024471 | 23.1971 | -0.027623 | 0.13634 | 0.000327 | 0.023834 | 0.023834 | 0.23053
2019-04-09 14:33 | 1554834785.5275 | -96 | 0.20951 | 0.000443 | 2e-06 | 99.926 | 100.4058 | 123.3988 | 20.4693 | 0.001859 | 0.001676 | -154.1073 | 0.22984 | 8.3999 | 0.002452 | 0.003265 | 0.003099 | 4.9e-05 | 0.001157 | 0.2943 | 1.3338 | 2.5475 | 98.55 | 20.09 | 49.808 | 65.7817 | 80.0308 | 61.3811 | 0.011895 | 1.0553 | 0.024476 | 23.1925 | -0.042838 | 0.13673 | -1.8e-05 | -0.001505 | -0.001505 | -3.6607
2019-04-09 14:33 | 1554834786.5276 | -95 | 0.2095 | 0.000445 | 2e-06 | 99.9126 | 100.3462 | 123.1146 | 20.4697 | 0.001786 | 0.001729 | -154.1136 | 0.23087 | 8.4028 | 0.002564 | 0.004112 | 0.003862 | 5e-05 | 0.001152 | 0.29734 | 1.3775 | 2.5477 | 98.55 | 20.08 | 49.804 | 64.0412 | 82.7311 | 62.6567 | 0.011887 | 1.0513 | 0.024401 | 23.2099 | -0.038732 | 0.13734 | 7.1e-05 | 0.005063 | 0.005063 | 1.0932
2019-04-09 14:33 | 1554834787.5266 | -94 | 0.20949 | 0.000448 | 2e-06 | 99.8841 | 100.2726 | 122.7611 | 20.4691 | 0.001814 | 0.001688 | -154.0951 | 0.22988 | 8.4025 | 0.002553 | 0.003792 | 0.00363 | 5.1e-05 | 0.001152 | 0.28136 | 1.3782 | 2.5476 | 98.55 | 20.09 | 49.803 | 64.7105 | 82.4567 | 61.6723 | 0.011894 | 1.0544 | 0.024402 | 23.1426 | -0.034725 | 0.13676 | 7.1e-05 | 0.005042 | 0.005042 | 1.0931
2019-04-09 14:33 | 1554834788.5265 | -93 | 0.20949 | 0.000444 | 2e-06 | 99.8587 | 100.3194 | 122.6805 | 20.4693 | 0.001775 | 0.001653 | -154.1162 | 0.2304 | 8.3999 | 0.002526 | 0.004745 | 0.004558 | 5.3e-05 | 0.001156 | 0.28584 | 1.3178 | 2.5475 | 98.55 | 20.09 | 49.796 | 63.7645 | 81.8119 | 60.8364 | 0.011892 | 1.0571 | 0.024425 | 23.1061 | -0.041166 | 0.13707 | 0.0001 | 0.007181 | 0.007181 | 0.76918
2019-04-09 14:33 | 1554834789.5277 | -92 | 0.20947 | 0.000444 | 3e-06 | 99.8838 | 100.2323 | 122.9362 | 20.47 | 0.001769 | 0.001631 | -154.1104 | 0.23046 | 8.4001 | 0.002541 | 0.003698 | 0.003511 | 5.4e-05 | 0.001152 | 0.27577 | 1.3768 | 2.5476 | 98.55 | 20.09 | 49.794 | 63.6198 | 82.1807 | 60.2914 | 0.011892 | 1.0588 | 0.02447 | 23.1113 | -0.031159 | 0.1371 | 0.00022 | 0.015997 | 0.015997 | 0.34538
2019-04-09 14:33 | 1554834790.5274 | -91 | 0.2095 | 0.000443 | 2e-06 | 99.812 | 100.1893 | 122.8329 | 20.4703 | 0.001811 | 0.00169 | -154.1437 | 0.22981 | 8.3997 | 0.002495 | 0.003999 | 0.003896 | 5.2e-05 | 0.00116 | 0.31542 | 1.4166 | 2.5476 | 98.55 | 20.09 | 49.79 | 64.6258 | 81.0649 | 61.7344 | 0.011891 | 1.0542 | 0.024407 | 23.1515 | -0.033748 | 0.13671 | 3.2e-05 | 0.002167 | 0.002167 | 2.5424
2019-04-09 14:33 | 1554834791.5256 | -90 | 0.20949 | 0.000446 | 2e-06 | 99.8718 | 100.1955 | 123.1577 | 20.4703 | 0.001805 | 0.001711 | -154.1386 | 0.23032 | 8.4031 | 0.002552 | 0.004439 | 0.004238 | 5e-05 | 0.001159 | 0.2844 | 1.3611 | 2.5475 | 98.55 | 20.09 | 49.788 | 64.4908 | 82.45 | 62.2274 | 0.01189 | 1.0527 | 0.024421 | 23.1992 | -0.028939 | 0.13702 | 7.9e-05 | 0.005658 | 0.005658 | 0.976
2019-04-09 14:33 | 1554834792.5256 | -89 | 0.20949 | 0.000445 | 2e-06 | 99.9195 | 100.4037 | 122.9941 | 20.471 | 0.001812 | 0.001695 | -154.1501 | 0.23014 | 8.4043 | 0.002551 | 0.003417 | 0.003303 | 5e-05 | 0.001163 | 0.28102 | 1.3787 | 2.5476 | 98.55 | 20.09 | 49.786 | 64.6703 | 82.4259 | 61.8399 | 0.01189 | 1.0539 | 0.024419 | 23.1703 | -0.043241 | 0.13691 | 0.00013 | 0.009408 | 0.009408 | 0.58646
2019-04-09 14:33 | 1554834793.5267 | -88 | 0.2095 | 0.000446 | 2e-06 | 99.8969 | 100.2533 | 122.6852 | 20.47 | 0.001784 | 0.001726 | -154.1579 | 0.22992 | 8.4008 | 0.00258 | 0.003706 | 0.003604 | 5e-05 | 0.001167 | 0.26305 | 1.3308 | 2.5476 | 98.55 | 20.09 | 49.782 | 63.9744 | 83.1291 | 62.5905 | 0.011889 | 1.0515 | 0.024361 | 23.1671 | -0.031859 | 0.13678 | 5.6e-05 | 0.003976 | 0.003976 | 1.3864
2019-04-09 14:33 | 1554834794.5271 | -87 | 0.2095 | 0.000448 | 3e-06 | 99.9275 | 100.3779 | 122.5355 | 20.4708 | 0.001822 | 0.001719 | -154.1421 | 0.22943 | 8.4012 | 0.002588 | 0.003842 | 0.003629 | 4.8e-05 | 0.00116 | 0.28727 | 1.3538 | 2.5475 | 98.55 | 20.09 | 49.78 | 64.8965 | 83.321 | 62.4193 | 0.011889 | 1.0521 | 0.024353 | 23.1471 | -0.040221 | 0.13649 | 5.3e-05 | 0.003652 | 0.003652 | 1.5063
2019-04-09 14:33 | 1554834795.5257 | -86 | 0.2095 | 0.000447 | 2e-06 | 99.8787 | 100.2636 | 122.5071 | 20.47 | 0.001751 | 0.001694 | -154.1351 | 0.23033 | 8.4006 | 0.00259 | 0.004261 | 0.004059 | 4.6e-05 | 0.001149 | 0.2656 | 1.317 | 2.5475 | 98.55 | 20.09 | 49.778 | 63.1803 | 83.3584 | 61.8225 | 0.011888 | 1.054 | 0.024371 | 23.1238 | -0.034406 | 0.13702 | 2.3e-05 | 0.001526 | 0.001526 | 3.6193
2019-04-09 14:33 | 1554834796.5268 | -85 | 0.2095 | 0.000443 | 3e-06 | 99.9377 | 100.3091 | 122.5292 | 20.4706 | 0.001749 | 0.001709 | -154.1551 | 0.23083 | 8.4039 | 0.002582 | 0.004033 | 0.003805 | 4.7e-05 | 0.001143 | 0.27489 | 1.3278 | 2.5475 | 98.55 | 20.09 | 49.779 | 63.1346 | 83.157 | 62.188 | 0.011888 | 1.0528 | 0.02436 | 23.1385 | -0.03318 | 0.13732 | 2.9e-05 | 0.001937 | 0.001937 | 2.8576
2019-04-09 14:33 | 1554834797.5268 | -84 | 0.20949 | 0.000445 | 2e-06 | 99.8984 | 100.105 | 122.3498 | 20.4715 | 0.001794 | 0.001708 | -154.158 | 0.2299 | 8.3996 | 0.00265 | 0.003761 | 0.003687 | 4.9e-05 | 0.00114 | 0.27367 | 1.368 | 2.5476 | 98.55 | 20.09 | 49.781 | 64.2304 | 84.8001 | 62.1646 | 0.011889 | 1.0529 | 0.024343 | 23.1208 | -0.018482 | 0.13676 | 8.5e-05 | 0.00605 | 0.00605 | 0.91095
2019-04-09 14:33 | 1554834798.5268 | -83 | 0.20949 | 0.000447 | 2e-06 | 99.8733 | 100.4842 | 122.666 | 20.4707 | 0.001781 | 0.00165 | -154.1312 | 0.23009 | 8.4034 | 0.002639 | 0.004109 | 0.003871 | 4.9e-05 | 0.001136 | 0.28008 | 1.3382 | 2.5476 | 98.55 | 20.09 | 49.784 | 63.9141 | 84.5287 | 60.7602 | 0.011889 | 1.0573 | 0.024426 | 23.1021 | -0.054545 | 0.13688 | 9.4e-05 | 0.006748 | 0.006748 | 0.81744
2019-04-09 14:33 | 1554834799.5268 | -82 | 0.20951 | 0.000449 | 2e-06 | 99.9746 | 100.4123 | 122.8322 | 20.4713 | 0.001775 | 0.001691 | -154.1491 | 0.22977 | 8.4041 | 0.002619 | 0.003846 | 0.003679 | 4.8e-05 | 0.001134 | 0.28407 | 1.3567 | 2.5476 | 98.55 | 20.09 | 49.789 | 63.7688 | 84.0618 | 61.7518 | 0.011891 | 1.0542 | 0.024406 | 23.152 | -0.039077 | 0.13669 | -1.7e-05 | -0.00141 | -0.00141 | -3.9075
2019-04-09 14:33 | 1554834800.526 | -81 | 0.2095 | 0.00045 | 3e-06 | 99.8915 | 100.402 | 122.9392 | 20.4726 | 0.001749 | 0.001748 | -154.1284 | 0.23012 | 8.4004 | 0.002635 | 0.004211 | 0.004046 | 5.2e-05 | 0.001137 | 0.28456 | 1.3583 | 2.5476 | 98.55 | 20.09 | 49.793 | 63.1508 | 84.4469 | 63.1147 | 0.011892 | 1.0499 | 0.024367 | 23.2092 | -0.045593 | 0.1369 | 6.7e-05 | 0.004716 | 0.004716 | 1.1697
2019-04-09 14:33 | 1554834801.5269 | -80 | 0.2095 | 0.000452 | 2e-06 | 99.9558 | 100.4278 | 122.6162 | 20.4714 | 0.001798 | 0.001689 | -154.1369 | 0.22999 | 8.4026 | 0.002713 | 0.003443 | 0.003273 | 5e-05 | 0.001137 | 0.26913 | 1.3459 | 2.5473 | 98.55 | 20.08 | 49.8 | 64.3326 | 86.3174 | 61.6903 | 0.011886 | 1.0544 | 0.024387 | 23.1296 | -0.042135 | 0.13682 | 1.1e-05 | 0.000659 | 0.000659 | 8.3681
2019-04-09 14:33 | 1554834802.5263 | -79 | 0.20949 | 0.000449 | 2e-06 | 100.0312 | 100.3855 | 122.9932 | 20.4723 | 0.001736 | 0.001695 | -154.1327 | 0.22987 | 8.4007 | 0.002758 | 0.004151 | 0.004056 | 5e-05 | 0.001141 | 0.3027 | 1.3211 | 2.5474 | 98.55 | 20.08 | 49.802 | 62.8436 | 87.4095 | 61.8349 | 0.011886 | 1.0539 | 0.024419 | 23.1701 | -0.031625 | 0.13675 | 9.9e-05 | 0.007114 | 0.007114 | 0.77472
2019-04-09 14:33 | 1554834803.526 | -78 | 0.2095 | 0.000452 | 2e-06 | 99.9877 | 100.3958 | 123.1566 | 20.4716 | 0.001777 | 0.001701 | -154.1392 | 0.22976 | 8.4003 | 0.002711 | 0.004422 | 0.00422 | 5.5e-05 | 0.00115 | 0.2855 | 1.3433 | 2.5474 | 98.55 | 20.08 | 49.81 | 63.8301 | 86.2753 | 61.9968 | 0.011888 | 1.0534 | 0.02443 | 23.1911 | -0.036433 | 0.13668 | 2.8e-05 | 0.001905 | 0.001905 | 2.8911
2019-04-09 14:33 | 1554834804.5274 | -77 | 0.20949 | 0.000447 | 2e-06 | 99.9829 | 100.3566 | 123.15 | 20.4724 | 0.001772 | 0.001706 | -154.1328 | 0.22942 | 8.4025 | 0.002735 | 0.004683 | 0.004503 | 5.2e-05 | 0.001154 | 0.28407 | 1.3614 | 2.5474 | 98.55 | 20.08 | 49.814 | 63.705 | 86.8547 | 62.1194 | 0.011889 | 1.053 | 0.024424 | 23.1947 | -0.033363 | 0.13648 | 7.9e-05 | 0.00562 | 0.00562 | 0.97861
2019-04-09 14:33 | 1554834805.5264 | -76 | 0.20949 | 0.000447 | 2e-06 | 100.1175 | 100.4612 | 123.1753 | 20.4721 | 0.001804 | 0.001687 | -154.1447 | 0.23029 | 8.4003 | 0.002704 | 0.003839 | 0.003702 | 5.3e-05 | 0.001154 | 0.27168 | 1.3672 | 2.5475 | 98.55 | 20.08 | 49.817 | 64.4651 | 86.1032 | 61.6418 | 0.01189 | 1.0545 | 0.024444 | 23.1806 | -0.03065 | 0.137 | 7.9e-05 | 0.005661 | 0.005661 | 0.97524
2019-04-09 14:33 | 1554834806.5265 | -75 | 0.20948 | 0.000447 | 2e-06 | 100.069 | 100.4237 | 123.4404 | 20.4727 | 0.001786 | 0.00171 | -154.1169 | 0.22925 | 8.4024 | 0.002757 | 0.003958 | 0.003831 | 5.2e-05 | 0.001147 | 0.29192 | 1.3199 | 2.5476 | 98.55 | 20.09 | 49.824 | 64.0246 | 87.3841 | 62.1963 | 0.011899 | 1.0528 | 0.02445 | 23.2247 | -0.031647 | 0.13638 | 0.000184 | 0.013315 | 0.013315 | 0.41278
2019-04-09 14:33 | 1554834807.5265 | -74 | 0.20952 | 0.000449 | 2e-06 | 100.1157 | 100.5837 | 123.5786 | 20.4733 | 0.001847 | 0.001687 | -154.1179 | 0.22996 | 8.4004 | 0.002709 | 0.004326 | 0.004104 | 5e-05 | 0.001142 | 0.29911 | 1.3713 | 2.5476 | 98.55 | 20.09 | 49.826 | 65.4943 | 86.2246 | 61.6646 | 0.0119 | 1.0545 | 0.024484 | 23.2193 | -0.041711 | 0.1368 | -9.8e-05 | -0.007342 | -0.007342 | -0.7509
2019-04-09 14:33 | 1554834808.5262 | -73 | 0.20949 | 0.000448 | 3e-06 | 100.1381 | 100.7699 | 123.3551 | 20.4736 | 0.0018 | 0.001724 | -154.1395 | 0.23013 | 8.4016 | 0.002746 | 0.004097 | 0.003971 | 5.2e-05 | 0.001137 | 0.27704 | 1.3523 | 2.5475 | 98.55 | 20.08 | 49.827 | 64.3698 | 87.1181 | 62.548 | 0.011892 | 1.0517 | 0.024429 | 23.2288 | -0.056257 | 0.1369 | 0.000132 | 0.009449 | 0.009449 | 0.5839
2019-04-09 14:33 | 1554834809.5252 | -72 | 0.20951 | 0.00045 | 2e-06 | 100.2092 | 100.7578 | 123.0857 | 20.4733 | 0.001786 | 0.001684 | -154.1268 | 0.23018 | 8.4021 | 0.002691 | 0.003549 | 0.00335 | 5.4e-05 | 0.001131 | 0.29391 | 1.3091 | 2.5476 | 98.55 | 20.09 | 49.828 | 64.0471 | 85.79 | 61.5821 | 0.0119 | 1.0547 | 0.024438 | 23.17 | -0.048833 | 0.13693 | 7e-06 | 0.000423 | 0.000423 | 13.0353
2019-04-09 14:33 | 1554834810.5252 | -71 | 0.20948 | 0.000449 | 2e-06 | 100.1649 | 100.551 | 122.9287 | 20.4731 | 0.001788 | 0.001697 | -154.1315 | 0.23012 | 8.3994 | 0.002723 | 0.004364 | 0.004212 | 5.5e-05 | 0.00113 | 0.29231 | 1.3217 | 2.5476 | 98.55 | 20.09 | 49.829 | 64.0891 | 86.5703 | 61.8967 | 0.0119 | 1.0537 | 0.024411 | 23.1661 | -0.034414 | 0.1369 | 0.000142 | 0.010261 | 0.010261 | 0.53768
2019-04-09 14:34 | 1554834811.5263 | -70 | 0.2095 | 0.00045 | 2e-06 | 100.144 | 100.5971 | 123.0331 | 20.4729 | 0.001803 | 0.001726 | -154.1358 | 0.22896 | 8.4025 | 0.002646 | 0.003951 | 0.003894 | 5.4e-05 | 0.001128 | 0.26825 | 1.3185 | 2.5476 | 98.55 | 20.09 | 49.828 | 64.4443 | 84.711 | 62.5863 | 0.0119 | 1.0516 | 0.024396 | 23.1998 | -0.040377 | 0.13621 | 6.7e-05 | 0.004784 | 0.004784 | 1.1474
2019-04-09 14:34 | 1554834812.5262 | -69 | 0.20952 | 0.000446 | 2e-06 | 100.1795 | 100.5913 | 123.0967 | 20.4736 | 0.00188 | 0.001703 | -154.1465 | 0.23055 | 8.4003 | 0.002699 | 0.004396 | 0.004215 | 5.3e-05 | 0.001128 | 0.26985 | 1.375 | 2.5476 | 98.55 | 20.09 | 49.826 | 66.3086 | 85.9791 | 62.0472 | 0.0119 | 1.0532 | 0.024422 | 23.1872 | -0.036692 | 0.13716 | -7.6e-05 | -0.005659 | -0.005659 | -0.97674
2019-04-09 14:34 | 1554834813.5253 | -68 | 0.20949 | 0.00045 | 2e-06 | 100.1786 | 100.6546 | 122.9836 | 20.4729 | 0.001768 | 0.001755 | -154.1429 | 0.22945 | 8.4005 | 0.0027 | 0.003911 | 0.003792 | 5.5e-05 | 0.001131 | 0.27168 | 1.3735 | 2.5477 | 98.55 | 20.09 | 49.826 | 63.5938 | 85.9979 | 63.2948 | 0.0119 | 1.0493 | 0.024365 | 23.2196 | -0.042401 | 0.1365 | 9.5e-05 | 0.006777 | 0.006777 | 0.81166
2019-04-09 14:34 | 1554834814.5268 | -67 | 0.2095 | 0.000453 | 2e-06 | 100.181 | 100.7589 | 122.9465 | 20.4727 | 0.001783 | 0.0017 | -154.1776 | 0.22996 | 8.4033 | 0.002771 | 0.00415 | 0.003998 | 5.3e-05 | 0.00114 | 0.29175 | 1.3264 | 2.5477 | 98.55 | 20.09 | 49.825 | 63.9575 | 87.7202 | 61.9548 | 0.011899 | 1.0535 | 0.02441 | 23.1698 | -0.05145 | 0.1368 | 5.8e-05 | 0.004131 | 0.004131 | 1.3345
2019-04-09 14:34 | 1554834815.5268 | -66 | 0.2095 | 0.000448 | 2e-06 | 100.2173 | 100.5926 | 123.0777 | 20.4729 | 0.001792 | 0.001696 | -154.1633 | 0.22939 | 8.4026 | 0.002679 | 0.004058 | 0.003844 | 5.5e-05 | 0.001137 | 0.26344 | 1.3515 | 2.5475 | 98.55 | 20.09 | 49.825 | 64.188 | 85.5062 | 61.8815 | 0.011899 | 1.0538 | 0.024426 | 23.1797 | -0.033439 | 0.13647 | 6.6e-05 | 0.004713 | 0.004713 | 1.1669
2019-04-09 14:34 | 1554834816.5264 | -65 | 0.20953 | 0.000452 | 3e-06 | 100.2962 | 100.9213 | 122.8857 | 20.4742 | 0.001823 | 0.001681 | -154.1779 | 0.23006 | 8.4025 | 0.002719 | 0.004329 | 0.004188 | 5.2e-05 | 0.001136 | 0.28047 | 1.3842 | 2.5476 | 98.55 | 20.09 | 49.829 | 64.9192 | 86.4714 | 61.5157 | 0.0119 | 1.0549 | 0.02442 | 23.1489 | -0.055575 | 0.13687 | -0.000124 | -0.009216 | -0.009216 | -0.59849
2019-04-09 14:34 | 1554834817.5265 | -64 | 0.20948 | 0.000452 | 3e-06 | 100.2819 | 100.7103 | 123.2279 | 20.4737 | 0.001774 | 0.001667 | -154.1565 | 0.22973 | 8.4035 | 0.002761 | 0.003754 | 0.003682 | 4.9e-05 | 0.001136 | 0.27273 | 1.3631 | 2.5476 | 98.55 | 20.09 | 49.832 | 63.7555 | 87.4758 | 61.1838 | 0.011901 | 1.056 | 0.024466 | 23.1696 | -0.038133 | 0.13667 | 0.000176 | 0.012749 | 0.012749 | 0.43201
2019-04-09 14:34 | 1554834818.5264 | -63 | 0.20945 | 0.000453 | 2e-06 | 100.2274 | 100.7613 | 123.3239 | 20.4748 | 0.001826 | 0.001706 | -154.1549 | 0.22979 | 8.4016 | 0.002722 | 0.003821 | 0.003632 | 5.1e-05 | 0.001136 | 0.26272 | 1.3522 | 2.5476 | 98.55 | 20.09 | 49.838 | 64.9988 | 86.5283 | 62.1124 | 0.011902 | 1.053 | 0.024442 | 23.2108 | -0.047518 | 0.1367 | 0.000356 | 0.025959 | 0.025959 | 0.21222
2019-04-09 14:34 | 1554834819.5265 | -62 | 0.20947 | 0.000448 | 2e-06 | 100.2679 | 100.4994 | 123.3676 | 20.4748 | 0.001822 | 0.001679 | -154.1348 | 0.22985 | 8.4006 | 0.002743 | 0.00414 | 0.003949 | 5.2e-05 | 0.001133 | 0.28655 | 1.3283 | 2.5476 | 98.55 | 20.09 | 49.851 | 64.8949 | 87.0373 | 61.4513 | 0.011905 | 1.0551 | 0.02447 | 23.192 | -0.020625 | 0.13674 | 0.000244 | 0.017728 | 0.017728 | 0.31084
2019-04-09 14:34 | 1554834820.5259 | -61 | 0.20953 | 0.000447 | 2e-06 | 100.2676 | 100.7609 | 123.5432 | 20.4747 | 0.001781 | 0.001736 | -154.1491 | 0.22966 | 8.4031 | 0.002775 | 0.004491 | 0.004407 | 4.7e-05 | 0.001134 | 0.26056 | 1.3559 | 2.5475 | 98.55 | 20.09 | 49.859 | 63.9221 | 87.8244 | 62.8348 | 0.011907 | 1.0508 | 0.024437 | 23.2565 | -0.0439 | 0.13662 | -0.000162 | -0.01195 | -0.01195 | -0.46075
2019-04-09 14:34 | 1554834821.5256 | -60 | 0.20953 | 0.000449 | 2e-06 | 100.3209 | 100.8597 | 123.1937 | 20.4748 | 0.001849 | 0.001674 | -154.1489 | 0.23042 | 8.4004 | 0.002686 | 0.004075 | 0.003897 | 5.4e-05 | 0.001139 | 0.26864 | 1.3356 | 2.5475 | 98.55 | 20.09 | 49.876 | 65.5444 | 85.6711 | 61.3506 | 0.011911 | 1.0554 | 0.024457 | 23.1722 | -0.04791 | 0.13708 | -0.000166 | -0.012334 | -0.012334 | -0.44789
2019-04-09 14:34 | 1554834822.525 | -59 | 0.20947 | 0.000451 | 3e-06 | 100.3278 | 100.9199 | 123.1178 | 20.4751 | 0.001804 | 0.001665 | -154.1465 | 0.22974 | 8.4035 | 0.002706 | 0.00443 | 0.004271 | 5e-05 | 0.00115 | 0.30127 | 1.3325 | 2.5475 | 98.55 | 20.09 | 49.884 | 64.4673 | 86.15 | 61.1266 | 0.011913 | 1.0561 | 0.024458 | 23.1573 | -0.052639 | 0.13667 | 0.000231 | 0.016761 | 0.016761 | 0.32861
2019-04-09 14:34 | 1554834823.5246 | -58 | 0.2095 | 0.000452 | 3e-06 | 100.2661 | 100.7737 | 123.3915 | 20.4755 | 0.001823 | 0.001743 | -154.1592 | 0.23032 | 8.4001 | 0.002661 | 0.003855 | 0.003657 | 4.9e-05 | 0.001157 | 0.27345 | 1.3186 | 2.5474 | 98.55 | 20.09 | 49.892 | 64.9282 | 85.0803 | 63.0044 | 0.011915 | 1.0502 | 0.024416 | 23.2481 | -0.045163 | 0.13702 | 6.1e-05 | 0.00426 | 0.00426 | 1.296
2019-04-09 14:34 | 1554834824.5246 | -57 | 0.2095 | 0.000451 | 2e-06 | 100.3078 | 100.7052 | 123.4617 | 20.4754 | 0.001814 | 0.001689 | -154.147 | 0.2301 | 8.4037 | 0.00268 | 0.003658 | 0.003471 | 4.7e-05 | 0.001157 | 0.31061 | 1.3309 | 2.5476 | 98.55 | 20.09 | 49.912 | 64.7198 | 85.5228 | 61.7098 | 0.01192 | 1.0543 | 0.02447 | 23.2098 | -0.035364 | 0.13689 | 2.2e-05 | 0.001464 | 0.001464 | 3.7675
2019-04-09 14:34 | 1554834825.5251 | -56 | 0.20952 | 0.000453 | 3e-06 | 100.3355 | 100.7926 | 123.1298 | 20.4759 | 0.001782 | 0.001719 | -154.1518 | 0.2302 | 8.4003 | 0.002654 | 0.003877 | 0.003781 | 4.9e-05 | 0.001153 | 0.2688 | 1.3242 | 2.5476 | 98.55 | 20.09 | 49.92 | 63.948 | 84.9113 | 62.4235 | 0.011922 | 1.0521 | 0.024411 | 23.2033 | -0.040655 | 0.13694 | -9.4e-05 | -0.007029 | -0.007029 | -0.7851
2019-04-09 14:34 | 1554834826.5259 | -55 | 0.2095 | 0.000449 | 2e-06 | 100.281 | 100.8088 | 123.3267 | 20.4752 | 0.001822 | 0.00173 | -154.1493 | 0.22965 | 8.4021 | 0.002623 | 0.004378 | 0.004256 | 4.7e-05 | 0.001149 | 0.27383 | 1.3234 | 2.5476 | 98.55 | 20.09 | 49.937 | 64.9164 | 84.1578 | 62.6839 | 0.011926 | 1.0513 | 0.024421 | 23.2309 | -0.04695 | 0.13662 | 7e-05 | 0.004973 | 0.004973 | 1.1071
2019-04-09 14:34 | 1554834827.5257 | -54 | 0.20947 | 0.000454 | 3e-06 | 100.3909 | 100.742 | 123.1012 | 20.4761 | 0.001796 | 0.001733 | -154.1643 | 0.22956 | 8.3992 | 0.002702 | 0.003394 | 0.003241 | 4.9e-05 | 0.001153 | 0.25658 | 1.3286 | 2.5477 | 98.55 | 20.09 | 49.944 | 64.2801 | 86.0597 | 62.7621 | 0.011928 | 1.051 | 0.024396 | 23.2123 | -0.031228 | 0.13657 | 0.000213 | 0.015445 | 0.015445 | 0.35635
2019-04-09 14:34 | 1554834828.5248 | -53 | 0.20952 | 0.000448 | 2e-06 | 100.44 | 100.7472 | 122.8075 | 20.4765 | 0.001761 | 0.0017 | -154.1562 | 0.22973 | 8.4021 | 0.002677 | 0.003723 | 0.003553 | 4.8e-05 | 0.001156 | 0.27295 | 1.3891 | 2.5477 | 98.55 | 20.09 | 49.949 | 63.4278 | 85.4551 | 61.9664 | 0.011929 | 1.0535 | 0.024396 | 23.1571 | -0.02731 | 0.13667 | -5.3e-05 | -0.00396 | -0.00396 | -1.3908
2019-04-09 14:34 | 1554834829.5252 | -52 | 0.20949 | 0.000446 | 2e-06 | 100.3629 | 100.7947 | 122.4855 | 20.4773 | 0.001747 | 0.001673 | -154.1598 | 0.22996 | 8.4025 | 0.002692 | 0.004218 | 0.004059 | 4.8e-05 | 0.00116 | 0.28296 | 1.3436 | 2.5476 | 98.55 | 20.09 | 49.958 | 63.0983 | 85.8063 | 61.3287 | 0.011931 | 1.0555 | 0.024387 | 23.1047 | -0.038397 | 0.1368 | 9e-05 | 0.006425 | 0.006425 | 0.85813
2019-04-09 14:34 | 1554834830.5249 | -51 | 0.20954 | 0.000451 | 2e-06 | 100.4475 | 100.995 | 122.3803 | 20.4766 | 0.001881 | 0.001678 | -154.1367 | 0.22986 | 8.4022 | 0.00261 | 0.003951 | 0.003736 | 4.8e-05 | 0.001168 | 0.27721 | 1.3597 | 2.5477 | 98.55 | 20.09 | 49.962 | 66.3278 | 83.8385 | 61.4414 | 0.011932 | 1.0552 | 0.024373 | 23.0987 | -0.048619 | 0.13675 | -0.00021 | -0.015446 | -0.015446 | -0.35678
2019-04-09 14:34 | 1554834831.5265 | -50 | 0.20953 | 0.000447 | 3e-06 | 100.4305 | 101.0005 | 122.6146 | 20.477 | 0.001858 | 0.001721 | -154.1289 | 0.23001 | 8.4015 | 0.002638 | 0.00371 | 0.003621 | 4.8e-05 | 0.001174 | 0.27201 | 1.3541 | 2.5476 | 98.55 | 20.09 | 49.964 | 65.7726 | 84.5293 | 62.4677 | 0.011932 | 1.0519 | 0.024359 | 23.1562 | -0.050619 | 0.13683 | -0.000138 | -0.010243 | -0.010243 | -0.53836
2019-04-09 14:34 | 1554834832.5259 | -49 | 0.20948 | 0.000451 | 2e-06 | 100.4314 | 100.6164 | 122.3658 | 20.4772 | 0.001817 | 0.001698 | -154.1283 | 0.23031 | 8.4002 | 0.002584 | 0.0037 | 0.003585 | 4.8e-05 | 0.001174 | 0.28385 | 1.4053 | 2.5476 | 98.55 | 20.09 | 49.964 | 64.7824 | 83.2125 | 61.9271 | 0.011932 | 1.0536 | 0.024354 | 23.1141 | -0.016461 | 0.13701 | 0.000167 | 0.012036 | 0.012036 | 0.45875
2019-04-09 14:34 | 1554834833.5264 | -48 | 0.20949 | 0.000451 | 2e-06 | 100.4904 | 100.9272 | 122.1069 | 20.4774 | 0.001778 | 0.00167 | -154.1082 | 0.23008 | 8.402 | 0.00259 | 0.004693 | 0.004525 | 4.6e-05 | 0.001168 | 0.28102 | 1.3581 | 2.5475 | 98.55 | 20.1 | 49.962 | 63.8535 | 83.3687 | 61.2466 | 0.011939 | 1.0558 | 0.024353 | 23.0662 | -0.038795 | 0.13687 | 7.7e-05 | 0.005501 | 0.005501 | 1.0028
2019-04-09 14:34 | 1554834834.5254 | -47 | 0.20947 | 0.000452 | 2e-06 | 100.4624 | 100.9444 | 122.4364 | 20.4765 | 0.001808 | 0.001753 | -154.1307 | 0.22988 | 8.4021 | 0.002534 | 0.003465 | 0.003317 | 4.7e-05 | 0.001167 | 0.29623 | 1.3084 | 2.5475 | 98.55 | 20.1 | 49.96 | 64.5796 | 82.0201 | 63.2468 | 0.011939 | 1.0495 | 0.024313 | 23.1663 | -0.042812 | 0.13676 | 0.00021 | 0.015197 | 0.015197 | 0.36265
2019-04-09 14:34 | 1554834835.525 | -46 | 0.20952 | 0.00045 | 2e-06 | 100.5258 | 100.7881 | 122.5233 | 20.4765 | 0.001849 | 0.00171 | -154.1209 | 0.23031 | 8.4016 | 0.00262 | 0.004354 | 0.004156 | 4.6e-05 | 0.001165 | 0.27113 | 1.361 | 2.5477 | 98.55 | 20.1 | 49.957 | 65.5607 | 84.0801 | 62.2069 | 0.011938 | 1.0527 | 0.024359 | 23.1386 | -0.023311 | 0.13701 | -4.2e-05 | -0.003199 | -0.003199 | -1.7261
2019-04-09 14:34 | 1554834836.5255 | -45 | 0.20955 | 0.000449 | 2e-06 | 100.5023 | 100.9981 | 122.5315 | 20.4778 | 0.00181 | 0.001694 | -154.1283 | 0.22942 | 8.4 | 0.002622 | 0.003729 | 0.003564 | 4.8e-05 | 0.001167 | 0.26864 | 1.3452 | 2.5475 | 98.55 | 20.1 | 49.949 | 64.6199 | 84.1244 | 61.8239 | 0.011936 | 1.0539 | 0.024374 | 23.1262 | -0.044016 | 0.13648 | -0.000253 | -0.018603 | -0.018603 | -0.29567
2019-04-09 14:34 | 1554834837.5253 | -44 | 0.20941 | 0.000453 | 2e-06 | 100.5112 | 100.9853 | 122.5343 | 20.4778 | 0.001794 | 0.001689 | -154.1359 | 0.22956 | 8.3991 | 0.002611 | 0.003809 | 0.00373 | 4.7e-05 | 0.001168 | 0.26488 | 1.3574 | 2.5477 | 98.56 | 20.1 | 49.945 | 64.2365 | 83.8738 | 61.7136 | 0.011934 | 1.0543 | 0.024378 | 23.1226 | -0.042093 | 0.13656 | 0.000567 | 0.041267 | 0.041267 | 0.13336
2019-04-09 14:34 | 1554834838.5252 | -43 | 0.20949 | 0.000449 | 2e-06 | 100.5121 | 100.9451 | 122.6698 | 20.4778 | 0.001788 | 0.001666 | -154.139 | 0.22975 | 8.4016 | 0.002558 | 0.003018 | 0.002888 | 4.7e-05 | 0.001158 | 0.28385 | 1.3361 | 2.5476 | 98.56 | 20.1 | 49.936 | 64.0811 | 82.607 | 61.1502 | 0.011932 | 1.0561 | 0.024412 | 23.1159 | -0.038446 | 0.13668 | 9.6e-05 | 0.006904 | 0.006904 | 0.79776
2019-04-09 14:34 | 1554834839.5242 | -42 | 0.20949 | 0.000447 | 2e-06 | 100.5041 | 100.9464 | 123.1406 | 20.4779 | 0.001796 | 0.001664 | -154.1416 | 0.22961 | 8.4049 | 0.002586 | 0.003861 | 0.003669 | 4.7e-05 | 0.001146 | 0.27273 | 1.3733 | 2.5476 | 98.56 | 20.1 | 49.932 | 64.2742 | 83.2577 | 61.1096 | 0.011931 | 1.0562 | 0.02446 | 23.1589 | -0.039284 | 0.13659 | 0.0001 | 0.007192 | 0.007192 | 0.76541
2019-04-09 14:34 | 1554834840.5241 | -41 | 0.20949 | 0.000452 | 2e-06 | 100.491 | 100.9038 | 123.2059 | 20.4776 | 0.001818 | 0.001712 | -154.1739 | 0.22964 | 8.4007 | 0.002516 | 0.004011 | 0.003915 | 5.1e-05 | 0.001142 | 0.28191 | 1.3661 | 2.5476 | 98.56 | 20.1 | 49.926 | 64.8117 | 81.5869 | 62.2613 | 0.01193 | 1.0526 | 0.024425 | 23.2049 | -0.036668 | 0.13661 | 0.000122 | 0.008788 | 0.008788 | 0.62647
2019-04-09 14:34 | 1554834841.5257 | -40 | 0.20949 | 0.000451 | 3e-06 | 100.4835 | 100.9082 | 123.3509 | 20.4781 | 0.001841 | 0.001645 | -154.1737 | 0.22988 | 8.4016 | 0.00257 | 0.003925 | 0.003714 | 5.2e-05 | 0.001138 | 0.26698 | 1.3554 | 2.5475 | 98.56 | 20.1 | 49.914 | 65.3547 | 82.8843 | 60.6542 | 0.011927 | 1.0576 | 0.024498 | 23.1628 | -0.037727 | 0.13676 | 0.00013 | 0.009376 | 0.009376 | 0.5878
2019-04-09 14:34 | 1554834842.5247 | -39 | 0.20946 | 0.000453 | 2e-06 | 100.4907 | 101.0631 | 123.7328 | 20.4785 | 0.001846 | 0.001696 | -154.1839 | 0.23041 | 8.3999 | 0.002527 | 0.00406 | 0.003879 | 5.2e-05 | 0.001133 | 0.27527 | 1.3514 | 2.5476 | 98.56 | 20.1 | 49.908 | 65.494 | 81.8417 | 61.8674 | 0.011925 | 1.0538 | 0.024491 | 23.2408 | -0.050808 | 0.13707 | 0.000284 | 0.020695 | 0.020695 | 0.26692
2019-04-09 14:34 | 1554834843.5253 | -38 | 0.20952 | 0.000448 | 2e-06 | 100.4567 | 100.8642 | 123.746 | 20.4789 | 0.001856 | 0.001758 | -154.1782 | 0.2293 | 8.4007 | 0.002546 | 0.003641 | 0.003458 | 5.1e-05 | 0.001132 | 0.28136 | 1.3648 | 2.5474 | 98.56 | 20.1 | 49.897 | 65.7324 | 82.3179 | 63.3637 | 0.011923 | 1.0491 | 0.024438 | 23.2939 | -0.036208 | 0.13641 | -8.3e-05 | -0.006196 | -0.006196 | -0.88733
2019-04-09 14:34 | 1554834844.5257 | -37 | 0.20947 | 0.000447 | 3e-06 | 100.4898 | 101.016 | 123.4615 | 20.4778 | 0.001864 | 0.001729 | -154.1506 | 0.22966 | 8.4057 | 0.002676 | 0.003964 | 0.00383 | 5e-05 | 0.001119 | 0.29031 | 1.3782 | 2.5475 | 98.56 | 20.1 | 49.893 | 65.9201 | 85.4292 | 62.6746 | 0.011922 | 1.0513 | 0.024435 | 23.2432 | -0.046714 | 0.13663 | 0.000238 | 0.017238 | 0.017238 | 0.31941
2019-04-09 14:34 | 1554834845.5243 | -36 | 0.20951 | 0.000449 | 2e-06 | 100.4832 | 101.0218 | 123.8467 | 20.4788 | 0.001788 | 0.001663 | -154.1547 | 0.23061 | 8.4004 | 0.002691 | 0.003731 | 0.003551 | 5.2e-05 | 0.001111 | 0.28241 | 1.3829 | 2.5476 | 98.56 | 20.1 | 49.889 | 64.0931 | 85.7977 | 61.0878 | 0.011921 | 1.0563 | 0.024531 | 23.2244 | -0.047815 | 0.13719 | -3.2e-05 | -0.002455 | -0.002455 | -2.2521
2019-04-09 14:34 | 1554834846.5244 | -35 | 0.20935 | 0.000449 | 3e-06 | 100.4662 | 100.8446 | 123.9293 | 20.4791 | 0.001767 | 0.001656 | -154.1554 | 0.23007 | 8.4014 | 0.002596 | 0.004335 | 0.004244 | 5.1e-05 | 0.001106 | 0.26034 | 1.3442 | 2.5475 | 98.56 | 20.1 | 49.882 | 63.5967 | 83.5188 | 60.9234 | 0.011919 | 1.0568 | 0.024545 | 23.2264 | -0.033619 | 0.13687 | 0.000963 | 0.070656 | 0.070656 | 0.078066
2019-04-09 14:34 | 1554834847.5255 | -34 | 0.20953 | 0.000446 | 3e-06 | 100.5026 | 100.8669 | 123.553 | 20.4798 | 0.001802 | 0.001732 | -154.168 | 0.22995 | 8.4029 | 0.002662 | 0.003478 | 0.0033 | 5e-05 | 0.001113 | 0.26626 | 1.324 | 2.5475 | 98.56 | 20.1 | 49.881 | 64.4249 | 85.0876 | 62.7413 | 0.011919 | 1.0511 | 0.024442 | 23.2542 | -0.032369 | 0.1368 | -0.000119 | -0.008862 | -0.008862 | -0.62207
2019-04-09 14:34 | 1554834848.5254 | -33 | 0.20953 | 0.00045 | 3e-06 | 100.5321 | 100.9007 | 123.5227 | 20.4785 | 0.001842 | 0.00169 | -154.1683 | 0.22939 | 8.4021 | 0.002645 | 0.003367 | 0.003205 | 4.7e-05 | 0.001119 | 0.2896 | 1.3509 | 2.5474 | 98.56 | 20.1 | 49.879 | 65.3996 | 84.6876 | 61.7237 | 0.011918 | 1.0543 | 0.024476 | 23.2161 | -0.032736 | 0.13646 | -0.00015 | -0.011188 | -0.011188 | -0.49157
2019-04-09 14:34 | 1554834849.5245 | -32 | 0.20957 | 0.000449 | 2e-06 | 100.5091 | 100.9162 | 123.5112 | 20.4796 | 0.001847 | 0.001648 | -154.1795 | 0.22979 | 8.4011 | 0.00266 | 0.003472 | 0.003392 | 4.6e-05 | 0.001116 | 0.29767 | 1.3527 | 2.5475 | 98.56 | 20.1 | 49.873 | 65.5178 | 85.0578 | 60.7111 | 0.011917 | 1.0575 | 0.024512 | 23.1799 | -0.036149 | 0.1367 | -0.000348 | -0.025671 | -0.025671 | -0.2146
2019-04-09 14:34 | 1554834850.5249 | -31 | 0.20945 | 0.00045 | 2e-06 | 100.5172 | 100.9984 | 123.7044 | 20.4797 | 0.001785 | 0.0017 | -154.1584 | 0.2302 | 8.4014 | 0.002579 | 0.004099 | 0.003867 | 4.8e-05 | 0.001111 | 0.26609 | 1.3345 | 2.5476 | 98.56 | 20.11 | 49.871 | 64.0073 | 83.0909 | 61.9709 | 0.011924 | 1.0535 | 0.024485 | 23.2417 | -0.042721 | 0.13694 | 0.00034 | 0.024809 | 0.024809 | 0.22246
2019-04-09 14:34 | 1554834851.5245 | -30 | 0.2095 | 0.000453 | 2e-06 | 100.4573 | 100.9248 | 123.3587 | 20.479 | 0.001821 | 0.001694 | -154.1636 | 0.2297 | 8.4014 | 0.002646 | 0.003466 | 0.003384 | 5.1e-05 | 0.001103 | 0.29391 | 1.3638 | 2.5476 | 98.56 | 20.11 | 49.867 | 64.8903 | 84.717 | 61.8383 | 0.011923 | 1.0539 | 0.024455 | 23.2046 | -0.041524 | 0.13664 | 3.6e-05 | 0.002572 | 0.002572 | 2.1408
2019-04-09 14:34 | 1554834852.525 | -29 | 0.20951 | 0.000451 | 2e-06 | 100.5809 | 101.1701 | 123.1608 | 20.4791 | 0.001779 | 0.001753 | -154.1463 | 0.23 | 8.4029 | 0.002717 | 0.003846 | 0.003654 | 4.7e-05 | 0.001101 | 0.27903 | 1.3735 | 2.5476 | 98.56 | 20.11 | 49.864 | 63.8857 | 86.4271 | 63.2515 | 0.011922 | 1.0495 | 0.024384 | 23.2348 | -0.052248 | 0.13683 | 1.2e-05 | 0.000752 | 0.000752 | 7.3363
2019-04-09 14:34 | 1554834853.525 | -28 | 0.20949 | 0.000448 | 2e-06 | 100.5294 | 100.9705 | 123.2929 | 20.4796 | 0.001779 | 0.001707 | -154.1776 | 0.2299 | 8.402 | 0.002571 | 0.003994 | 0.003805 | 4.7e-05 | 0.001098 | 0.28655 | 1.3656 | 2.5476 | 98.56 | 20.11 | 49.862 | 63.8751 | 82.9044 | 62.1515 | 0.011922 | 1.0529 | 0.024437 | 23.2093 | -0.039166 | 0.13677 | 7.7e-05 | 0.005526 | 0.005526 | 0.99739
2019-04-09 14:34 | 1554834854.5236 | -27 | 0.20937 | 0.000454 | 1e-06 | 100.5511 | 101.0741 | 123.4849 | 20.4798 | 0.001803 | 0.001741 | -154.1812 | 0.22951 | 8.4043 | 0.002585 | 0.004018 | 0.003884 | 4.5e-05 | 0.001096 | 0.27848 | 1.3862 | 2.5477 | 98.56 | 20.11 | 49.859 | 64.4485 | 83.258 | 62.957 | 0.011921 | 1.0504 | 0.024427 | 23.2552 | -0.046404 | 0.13654 | 0.000814 | 0.059478 | 0.059478 | 0.092512
2019-04-09 14:34 | 1554834855.5237 | -26 | 0.20951 | 0.00045 | 2e-06 | 100.5276 | 101.0886 | 123.5558 | 20.4808 | 0.001821 | 0.001687 | -154.193 | 0.22944 | 8.4018 | 0.002506 | 0.003051 | 0.002994 | 4.4e-05 | 0.001101 | 0.27776 | 1.3277 | 2.5474 | 98.56 | 20.11 | 49.857 | 64.885 | 81.3371 | 61.6565 | 0.011921 | 1.0545 | 0.024482 | 23.2168 | -0.049776 | 0.13649 | 1.2e-05 | 0.000797 | 0.000797 | 6.8984
2019-04-09 14:34 | 1554834856.5247 | -25 | 0.20949 | 0.000444 | 2e-06 | 100.4871 | 101.0232 | 123.1667 | 20.4805 | 0.001846 | 0.001697 | -154.1795 | 0.22971 | 8.4025 | 0.002537 | 0.003847 | 0.003718 | 4.3e-05 | 0.001099 | 0.28351 | 1.3518 | 2.5475 | 98.56 | 20.11 | 49.851 | 65.4982 | 82.0988 | 61.8961 | 0.011919 | 1.0537 | 0.024434 | 23.1885 | -0.047592 | 0.13666 | 0.000122 | 0.008816 | 0.008816 | 0.62472
2019-04-09 14:34 | 1554834857.5239 | -24 | 0.20939 | 0.000451 | 2e-06 | 100.4722 | 101.0555 | 123.4258 | 20.4811 | 0.001838 | 0.001705 | -154.1652 | 0.2295 | 8.4011 | 0.002628 | 0.003235 | 0.003019 | 4.7e-05 | 0.0011 | 0.2677 | 1.3482 | 2.5475 | 98.56 | 20.11 | 49.849 | 65.2967 | 84.2788 | 62.0958 | 0.011919 | 1.0531 | 0.024453 | 23.2198 | -0.051782 | 0.13653 | 0.000708 | 0.051723 | 0.051723 | 0.10638
2019-04-09 14:34 | 1554834858.5237 | -23 | 0.2096 | 0.00045 | 2e-06 | 100.482 | 101.0008 | 123.3941 | 20.4803 | 0.001863 | 0.00168 | -154.1621 | 0.22901 | 8.4023 | 0.0025 | 0.004487 | 0.004264 | 4.4e-05 | 0.001109 | 0.28479 | 1.3541 | 2.5476 | 98.56 | 20.11 | 49.847 | 65.8853 | 81.1941 | 61.4925 | 0.011918 | 1.055 | 0.024472 | 23.196 | -0.046062 | 0.13624 | -0.000548 | -0.040284 | -0.040284 | -0.13629
2019-04-09 14:34 | 1554834859.5238 | -22 | 0.20965 | 0.000449 | 2e-06 | 100.4609 | 100.8635 | 123.1401 | 20.4808 | 0.001813 | 0.001759 | -154.1612 | 0.22985 | 8.402 | 0.002533 | 0.003464 | 0.003445 | 4.2e-05 | 0.001106 | 0.28998 | 1.3921 | 2.5477 | 98.56 | 20.11 | 49.843 | 64.6954 | 82.0048 | 63.3874 | 0.011917 | 1.0491 | 0.024377 | 23.2376 | -0.035775 | 0.13674 | -0.000843 | -0.061684 | -0.061684 | -0.089334
2019-04-09 14:34 | 1554834860.5247 | -21 | 0.20939 | 0.000449 | 2e-06 | 100.5139 | 100.8318 | 122.9979 | 20.4809 | 0.001842 | 0.001715 | -154.179 | 0.22977 | 8.4021 | 0.002489 | 0.003978 | 0.003767 | 4.6e-05 | 0.001112 | 0.30381 | 1.3587 | 2.5476 | 98.56 | 20.11 | 49.842 | 65.3911 | 80.9505 | 62.3372 | 0.011917 | 1.0523 | 0.024401 | 23.1879 | -0.028249 | 0.13669 | 0.000723 | 0.052739 | 0.052739 | 0.10445
2019-04-09 14:34 | 1554834861.5242 | -20 | 0.20927 | 0.000447 | 2e-06 | 100.5282 | 100.8687 | 123.2643 | 20.4816 | 0.00178 | 0.001753 | -154.1861 | 0.2291 | 8.4008 | 0.002616 | 0.003569 | 0.003375 | 4.6e-05 | 0.001116 | 0.29734 | 1.3238 | 2.5475 | 98.56 | 20.11 | 49.841 | 63.9122 | 83.9877 | 63.2532 | 0.011917 | 1.0495 | 0.024395 | 23.2447 | -0.030243 | 0.13629 | 0.00141 | 0.10293 | 0.10293 | 0.053364
2019-04-09 14:34 | 1554834862.5249 | -19 | 0.20932 | 0.000452 | 2e-06 | 100.482 | 101.0514 | 123.449 | 20.4817 | 0.001844 | 0.001683 | -154.2102 | 0.22961 | 8.4015 | 0.002547 | 0.003615 | 0.003451 | 4.7e-05 | 0.001118 | 0.25962 | 1.3518 | 2.5474 | 98.56 | 20.11 | 49.837 | 65.4377 | 82.3424 | 61.562 | 0.011916 | 1.0548 | 0.024474 | 23.2035 | -0.050541 | 0.13659 | 0.001148 | 0.084028 | 0.084028 | 0.065511
2019-04-09 14:34 | 1554834863.5253 | -18 | 0.20937 | 0.000445 | 1e-06 | 100.4707 | 100.9778 | 123.7828 | 20.4815 | 0.001732 | 0.001676 | -154.2072 | 0.22903 | 8.4019 | 0.002555 | 0.003926 | 0.00377 | 4.4e-05 | 0.001121 | 0.28064 | 1.3175 | 2.5474 | 98.56 | 20.11 | 49.837 | 62.7564 | 82.5194 | 61.4048 | 0.011916 | 1.0553 | 0.024513 | 23.2294 | -0.045027 | 0.13625 | 0.000836 | 0.061279 | 0.061279 | 0.089605
2019-04-09 14:34 | 1554834864.5245 | -17 | 0.20947 | 0.000451 | 2e-06 | 100.4627 | 100.8635 | 123.7542 | 20.4824 | 0.001878 | 0.001646 | -154.2287 | 0.22899 | 8.4026 | 0.002503 | 0.004251 | 0.004064 | 4.5e-05 | 0.001118 | 0.2688 | 1.3449 | 2.5474 | 98.56 | 20.11 | 49.834 | 66.2548 | 81.2803 | 60.6702 | 0.011915 | 1.0576 | 0.024537 | 23.2012 | -0.035616 | 0.13622 | 0.000196 | 0.014287 | 0.014287 | 0.38425
2019-04-09 14:34 | 1554834865.525 | -16 | 0.20955 | 0.000449 | 2e-06 | 100.4097 | 101.2207 | 123.6251 | 20.4815 | 0.001851 | 0.00164 | -154.2259 | 0.23005 | 8.3984 | 0.002489 | 0.00419 | 0.004025 | 4.6e-05 | 0.00112 | 0.26681 | 1.3246 | 2.5476 | 98.56 | 20.11 | 49.833 | 65.5992 | 80.9378 | 60.5206 | 0.011915 | 1.0581 | 0.02453 | 23.1839 | -0.07196 | 0.13685 | -0.000266 | -0.019609 | -0.019609 | -0.28126
2019-04-09 14:34 | 1554834866.5241 | -15 | 0.20935 | 0.000448 | 2e-06 | 100.485 | 101.0349 | 124.21 | 20.4815 | 0.001856 | 0.00165 | -154.2383 | 0.23003 | 8.4003 | 0.00254 | 0.003924 | 0.003724 | 4.6e-05 | 0.001122 | 0.28528 | 1.34 | 2.5476 | 98.56 | 20.11 | 49.83 | 65.7186 | 82.1591 | 60.7794 | 0.011914 | 1.0572 | 0.024579 | 23.2477 | -0.048814 | 0.13685 | 0.000974 | 0.071613 | 0.071613 | 0.07701
2019-04-09 14:34 | 1554834867.524 | -14 | 0.20947 | 0.000452 | 2e-06 | 100.4564 | 100.8497 | 124.5102 | 20.4824 | 0.001813 | 0.001649 | -154.2203 | 0.22953 | 8.4014 | 0.002603 | 0.003913 | 0.003771 | 4.5e-05 | 0.001119 | 0.27096 | 1.3166 | 2.5475 | 98.55 | 20.12 | 49.825 | 64.6946 | 83.6871 | 60.7521 | 0.011921 | 1.0573 | 0.024609 | 23.2748 | -0.034952 | 0.13655 | 0.000236 | 0.017268 | 0.017268 | 0.31868
2019-04-09 14:34 | 1554834868.5235 | -13 | 0.20937 | 0.000448 | 2e-06 | 100.4793 | 101.0948 | 124.629 | 20.4814 | 0.001839 | 0.001683 | -154.2283 | 0.22959 | 8.4008 | 0.002663 | 0.004581 | 0.004369 | 4.3e-05 | 0.001117 | 0.28584 | 1.3099 | 2.5475 | 98.55 | 20.12 | 49.822 | 65.311 | 85.1352 | 61.5632 | 0.011921 | 1.0548 | 0.024591 | 23.3142 | -0.054617 | 0.13658 | 0.000818 | 0.06015 | 0.06015 | 0.09151
2019-04-09 14:34 | 1554834869.5235 | -12 | 0.20944 | 0.000444 | 2e-06 | 100.4537 | 100.8508 | 124.7715 | 20.4818 | 0.001838 | 0.001687 | -154.222 | 0.22904 | 8.4017 | 0.002655 | 0.004066 | 0.003969 | 4.6e-05 | 0.00111 | 0.26017 | 1.3305 | 2.5476 | 98.55 | 20.12 | 49.814 | 65.2856 | 84.9217 | 61.6606 | 0.011919 | 1.0545 | 0.024602 | 23.3309 | -0.035282 | 0.13625 | 0.00039 | 0.028576 | 0.028576 | 0.19216
2019-04-09 14:34 | 1554834870.5232 | -11 | 0.20936 | 0.000448 | 2e-06 | 100.4448 | 100.9055 | 125.1684 | 20.4818 | 0.001775 | 0.001665 | -154.2096 | 0.22974 | 8.4014 | 0.002618 | 0.004068 | 0.003875 | 4.4e-05 | 0.001114 | 0.28351 | 1.3072 | 2.5476 | 98.56 | 20.12 | 49.809 | 63.7837 | 84.0425 | 61.1346 | 0.011916 | 1.0561 | 0.02466 | 23.3496 | -0.04093 | 0.13667 | 0.000874 | 0.064492 | 0.064492 | 0.085404
2019-04-09 14:34 | 1554834871.5242 | -10 | 0.2096 | 0.00045 | 2e-06 | 100.4412 | 100.9423 | 125.3165 | 20.4824 | 0.00177 | 0.001636 | -134.7098 | 0.23027 | 8.4017 | 0.002752 | 0.00359 | 0.003397 | 4.1e-05 | 0.001114 | 0.2724 | 1.3103 | 2.5475 | 98.56 | 20.12 | 49.804 | 63.6562 | 87.2672 | 60.4302 | 0.011915 | 1.0584 | 0.024701 | 23.3388 | -0.044511 | 0.13699 | -0.000557 | -0.041294 | -0.041294 | -0.13369
2019-04-09 14:35 | 1554834872.5232 | -9 | 0.2096 | 0.000452 | 2e-06 | 100.4466 | 100.8504 | 125.0832 | 20.4834 | 0.001791 | 0.001655 | -92.6667 | 0.22983 | 8.4031 | 0.002771 | 0.003979 | 0.003742 | 4.5e-05 | 0.001113 | 0.27671 | 1.3581 | 2.5475 | 98.56 | 20.11 | 49.793 | 64.1655 | 87.7188 | 60.8837 | 0.011905 | 1.0569 | 0.024661 | 23.3329 | -0.035888 | 0.13672 | -0.000563 | -0.041689 | -0.041689 | -0.13217
2019-04-09 14:35 | 1554834873.5232 | -8 | 0.20952 | 0.000454 | 2e-06 | 100.4594 | 100.912 | 124.814 | 20.4828 | 0.001788 | 0.001638 | -50.6032 | 0.22976 | 8.4012 | 0.00283 | 0.004321 | 0.004128 | 4.1e-05 | 0.00111 | 0.27256 | 1.3774 | 2.5477 | 98.56 | 20.11 | 49.79 | 64.0869 | 89.1537 | 60.4711 | 0.011904 | 1.0582 | 0.02465 | 23.2934 | -0.04021 | 0.13668 | -8.8e-05 | -0.006627 | -0.006627 | -0.8312
2019-04-09 14:35 | 1554834874.5233 | -7 | 0.20953 | 0.00045 | 2e-06 | 100.4031 | 101.1436 | 125.0423 | 20.4828 | 0.001779 | 0.001691 | -8.4995 | 0.2299 | 8.4018 | 0.00273 | 0.004706 | 0.004569 | 4.3e-05 | 0.001107 | 0.26178 | 1.3803 | 2.5476 | 98.56 | 20.11 | 49.783 | 63.8657 | 86.7476 | 61.7692 | 0.011903 | 1.0541 | 0.024624 | 23.36 | -0.065729 | 0.13677 | -0.000108 | -0.008059 | -0.008059 | -0.68397
2019-04-09 14:35 | 1554834875.5243 | -6 | 0.20936 | 0.000445 | 2e-06 | 100.4156 | 100.7658 | 125.0099 | 20.4828 | 0.001775 | 0.00159 | 33.5592 | 0.22977 | 8.402 | 0.002783 | 0.004445 | 0.004334 | 4.7e-05 | 0.001109 | 0.26825 | 1.3726 | 2.5476 | 98.56 | 20.11 | 49.78 | 63.7909 | 88.0171 | 59.3224 | 0.011902 | 1.0619 | 0.024712 | 23.2715 | -0.031135 | 0.13669 | 0.000926 | 0.068411 | 0.068411 | 0.080521
2019-04-09 14:35 | 1554834876.5237 | -5 | 0.20945 | 0.000453 | 2e-06 | 100.4192 | 100.868 | 124.5917 | 20.4832 | 0.001778 | 0.001637 | 40.7362 | 0.24434 | 8.402 | 0.002791 | 0.004038 | 0.003892 | 4.3e-05 | 0.001118 | 0.27201 | 1.3406 | 2.5476 | 98.56 | 20.11 | 49.778 | 63.8428 | 88.1988 | 60.4638 | 0.011902 | 1.0582 | 0.024628 | 23.2724 | -0.039885 | 0.14536 | 0.00034 | 0.02499 | 0.02499 | 0.2344
2019-04-09 14:35 | 1554834877.5244 | -4 | 0.20953 | 0.000448 | 2e-06 | 100.4311 | 101.0872 | 124.1506 | 20.4835 | 0.001766 | 0.001629 | 266.0999 | 0.26246 | 8.4016 | 0.002816 | 0.004858 | 0.004794 | 4.4e-05 | 0.001123 | 0.27417 | 1.349 | 2.5476 | 98.56 | 20.11 | 49.777 | 63.5699 | 88.7966 | 60.2725 | 0.011901 | 1.0589 | 0.024591 | 23.2245 | -0.058243 | 0.15614 | -0.000124 | -0.009311 | -0.009311 | -0.67581
2019-04-09 14:35 | 1554834878.5239 | -3 | 0.20954 | 0.00045 | 2e-06 | 100.4228 | 100.9471 | 124.0802 | 20.4835 | 0.001787 | 0.001659 | 266.3976 | 0.26148 | 8.4039 | 0.002826 | 0.004083 | 0.003955 | 4.2e-05 | 0.001121 | 0.28783 | 1.2998 | 2.5475 | 98.56 | 20.1 | 49.777 | 64.0594 | 89.0505 | 60.9986 | 0.011894 | 1.0566 | 0.024558 | 23.2432 | -0.046583 | 0.15555 | -0.000191 | -0.014193 | -0.014193 | -0.44167
2019-04-09 14:35 | 1554834879.5235 | -2 | 0.20949 | 0.000452 | 2e-06 | 100.4603 | 100.9544 | 124.0664 | 20.4834 | 0.001786 | 0.001594 | 266.7342 | 0.26102 | 8.4038 | 0.002856 | 0.004593 | 0.004586 | 4e-05 | 0.001117 | 0.2896 | 1.3463 | 2.5474 | 98.56 | 20.1 | 49.778 | 64.0366 | 89.7769 | 59.4152 | 0.011894 | 1.0616 | 0.024615 | 23.1868 | -0.043882 | 0.15528 | 0.000126 | 0.009168 | 0.009168 | 0.68256
2019-04-09 14:35 | 1554834880.5234 | -1 | 0.20956 | 0.000447 | 2e-06 | 100.3876 | 100.7675 | 123.5866 | 20.484 | 0.00172 | 0.001649 | 266.3899 | 0.26153 | 8.402 | 0.002898 | 0.003925 | 0.003864 | 4.3e-05 | 0.00112 | 0.2891 | 1.3612 | 2.5476 | 98.56 | 20.1 | 49.784 | 62.4673 | 90.788 | 60.7404 | 0.011896 | 1.0574 | 0.024518 | 23.1879 | -0.033781 | 0.15559 | -0.000282 | -0.02083 | -0.02083 | -0.30101
2019-04-09 14:35 | 1554834881.5235 | 0 | 0.2095 | 0.000443 | 2e-06 | 100.4231 | 100.8635 | 123.1883 | 20.4836 | 0.001799 | 0.001589 | 266.7721 | 0.26212 | 8.4012 | 0.002908 | 0.004316 | 0.004271 | 4.2e-05 | 0.001119 | 0.27599 | 1.3527 | 2.6619 | 98.56 | 20.1 | 49.79 | 64.3656 | 91.0346 | 59.2935 | 0.011897 | 1.062 | 0.024532 | 23.1003 | -0.039143 | 0.15594 | 5.3e-05 | 0.00378 | 0.00378 | 1.6626
2019-04-09 14:35 | 1554834882.5239 | 1 | 0.2095 | 0.000446 | 2e-06 | 98.7072 | 100.4188 | 122.4109 | 20.4758 | 0.001693 | 0.001617 | 266.4923 | 0.26199 | 8.4003 | 0.002866 | 0.005239 | 0.005168 | 4e-05 | 0.00112 | 0.26537 | 1.3474 | 2.6113 | 98.56 | 20.1 | 49.802 | 61.8008 | 90.0151 | 59.9756 | 0.0119 | 1.0598 | 0.024429 | 23.0509 | -0.15378 | 0.15586 | 7.2e-05 | 0.005139 | 0.005139 | 1.2222
2019-04-09 14:35 | 1554834883.5231 | 2 | 0.20951 | 0.000452 | 2e-06 | 87.1782 | 96.0391 | 121.1842 | 20.4744 | 0.001705 | 0.001648 | 266.922 | 0.26173 | 8.4024 | 0.002823 | 0.004553 | 0.004704 | 4.2e-05 | 0.001116 | 0.22948 | 1.3486 | 2.5479 | 98.56 | 20.1 | 49.809 | 62.0901 | 88.9572 | 60.7255 | 0.011902 | 1.0574 | 0.024279 | 22.9609 | -0.86585 | 0.15571 | -2.6e-05 | -0.002035 | -0.002035 | -3.0832
2019-04-09 14:35 | 1554834884.524 | 3 | 0.20949 | 0.000448 | 4e-06 | 79.4139 | 92.6052 | 119.5709 | 20.46 | 0.001733 | 0.001722 | 266.4131 | 0.2632 | 8.4019 | 0.002778 | 0.005347 | 0.005171 | 4e-05 | 0.001118 | 0.28031 | 1.3173 | 2.5475 | 98.56 | 20.1 | 49.816 | 62.7448 | 87.8819 | 62.4715 | 0.011903 | 1.0519 | 0.024054 | 22.8672 | -1.3745 | 0.15658 | 0.000134 | 0.009405 | 0.009405 | 0.6709
2019-04-09 14:35 | 1554834885.5235 | 4 | 0.20955 | 0.000451 | 5e-06 | 81.7897 | 95.3422 | 117.6819 | 20.4697 | 0.001783 | 0.001706 | 266.8764 | 0.26213 | 8.3989 | 0.002726 | 0.003738 | 0.003619 | 3.9e-05 | 0.001118 | 0.32372 | 1.3396 | 2.5478 | 98.56 | 20.1 | 49.828 | 63.9524 | 86.6288 | 62.1029 | 0.011906 | 1.0531 | 0.023877 | 22.6733 | -1.3714 | 0.15594 | -0.000281 | -0.020431 | -0.020431 | -0.30759
2019-04-09 14:35 | 1554834886.5231 | 5 | 0.20933 | 0.000448 | 7e-06 | 86.8983 | 97.5823 | 115.8879 | 20.4703 | 0.001804 | 0.001705 | 266.706 | 0.2624 | 8.4023 | 0.002793 | 0.004641 | 0.004498 | 4.3e-05 | 0.001125 | 0.27632 | 1.3385 | 2.5475 | 98.56 | 20.1 | 49.832 | 64.4737 | 88.2426 | 62.0791 | 0.011907 | 1.0531 | 0.023695 | 22.4991 | -1.0372 | 0.1561 | 0.001071 | 0.075599 | 0.075599 | 0.083213
2019-04-09 14:35 | 1554834887.5242 | 6 | 0.20882 | 0.000454 | 1.1e-05 | 92.3044 | 98.7993 | 113.8282 | 20.4747 | 0.001882 | 0.001826 | 266.5182 | 0.26229 | 8.4028 | 0.002806 | 0.004286 | 0.004148 | 4.4e-05 | 0.001127 | 0.32045 | 1.3612 | 2.5476 | 98.55 | 20.1 | 49.839 | 66.344 | 88.5638 | 64.9892 | 0.01191 | 1.0441 | 0.023382 | 22.3948 | -0.60822 | 0.15604 | 0.004118 | 0.28766 | 0.28766 | 0.02186
2019-04-09 14:35 | 1554834888.5231 | 7 | 0.20792 | 0.000478 | 1.7e-05 | 95.9541 | 97.753 | 111.4101 | 20.4784 | 0.001928 | 0.001861 | 266.4038 | 0.26234 | 8.4007 | 0.002765 | 0.004763 | 0.004741 | 3.6e-05 | 0.001118 | 0.25011 | 1.3422 | 1.7587 | 98.55 | 20.1 | 49.842 | 67.4572 | 87.5695 | 65.8373 | 0.011911 | 1.0415 | 0.023103 | 22.1834 | -0.16614 | 0.15606 | 0.009508 | 0.65639 | 0.65639 | 0.009582
2019-04-09 14:35 | 1554834889.5227 | 8 | 0.20664 | 0.000554 | 2.6e-05 | 98.1473 | 98.4534 | 109.2314 | 20.4844 | 0.001979 | 0.001968 | 266.3844 | 0.26202 | 8.4037 | 0.002726 | 0.004645 | 0.004503 | 3.7e-05 | 0.001112 | 0.27129 | 1.3492 | 2.5475 | 98.55 | 20.1 | 49.843 | 68.6814 | 86.6483 | 68.4222 | 0.011911 | 1.0336 | 0.02279 | 22.049 | -0.027848 | 0.15588 | 0.017072 | 1.1619 | 1.1619 | 0.005406
2019-04-09 14:35 | 1554834890.524 | 9 | 0.2054 | 0.00098 | 3.9e-05 | 99.2906 | 97.7806 | 107.085 | 20.4837 | 0.002072 | 0.002084 | 266.2495 | 0.26272 | 8.403 | 0.002815 | 0.004702 | 0.00469 | 4e-05 | 0.001109 | 0.24889 | 1.3544 | 2.5476 | 98.55 | 20.1 | 49.844 | 70.9296 | 88.7922 | 71.2173 | 0.011911 | 1.0252 | 0.022473 | 21.9205 | 0.13707 | 0.15629 | 0.02397 | 1.6073 | 1.6073 | 0.003919
2019-04-09 14:35 | 1554834891.5232 | 10 | 0.20359 | 0.001718 | 5.3e-05 | 99.8745 | 94.2523 | 105.4509 | 20.4849 | 0.002063 | 0.002129 | 266.1599 | 0.26199 | 8.4007 | 0.002745 | 0.004238 | 0.004163 | 4e-05 | 0.00111 | 0.2417 | 1.3522 | 2.5476 | 98.55 | 20.1 | 49.843 | 70.7047 | 87.1017 | 72.3036 | 0.011911 | 1.022 | 0.022266 | 21.7869 | 0.51824 | 0.15586 | 0.033853 | 2.247 | 2.247 | 0.002795
2019-04-09 14:35 | 1554834892.5229 | 11 | 0.20243 | 0.002608 | 6.8e-05 | 100.1032 | 88.6705 | 104.508 | 20.4851 | 0.002123 | 0.002175 | 266.029 | 0.26243 | 8.4029 | 0.002799 | 0.004897 | 0.004866 | 4e-05 | 0.001117 | 0.23275 | 1.3884 | 2.5475 | 98.55 | 20.1 | 49.842 | 72.1464 | 88.4101 | 73.392 | 0.011911 | 1.0188 | 0.022131 | 21.7234 | 1.0848 | 0.15612 | 0.039705 | 2.6175 | 2.6175 | 0.002404
2019-04-09 14:35 | 1554834893.5239 | 12 | 0.20125 | 0.003629 | 8.1e-05 | 100.2426 | 82.2062 | 103.8588 | 20.4855 | 0.002274 | 0.002238 | 265.9115 | 0.26154 | 8.4013 | 0.002869 | 0.004002 | 0.003985 | 4.2e-05 | 0.001109 | 0.25785 | 1.3382 | 2.5475 | 98.55 | 20.1 | 49.841 | 75.7715 | 90.0972 | 74.9148 | 0.011911 | 1.0143 | 0.022014 | 21.7033 | 1.7743 | 0.15559 | 0.04543 | 2.9772 | 2.9772 | 0.002106
2019-04-09 14:35 | 1554834894.5229 | 13 | 0.20017 | 0.004737 | 9.1e-05 | 100.3581 | 76.972 | 103.0894 | 20.4853 | 0.002286 | 0.002252 | 265.8246 | 0.26197 | 8.4004 | 0.002921 | 0.004929 | 0.004747 | 3.9e-05 | 0.001104 | 0.25393 | 1.3113 | 2.5473 | 98.55 | 20.1 | 49.841 | 76.0752 | 91.3389 | 75.2411 | 0.011911 | 1.0134 | 0.021922 | 21.6329 | 2.373 | 0.15585 | 0.050548 | 3.297 | 3.297 | 0.001905
2019-04-09 14:35 | 1554834895.5239 | 14 | 0.2 | 0.005817 | 9.9e-05 | 100.3212 | 74.8837 | 102.5762 | 20.4856 | 0.00235 | 0.00236 | 265.6837 | 0.26216 | 8.403 | 0.002929 | 0.004071 | 0.004137 | 4.1e-05 | 0.001097 | 0.25337 | 1.3116 | 2.5475 | 98.55 | 20.09 | 49.843 | 77.5964 | 91.5394 | 77.8559 | 0.011904 | 1.0058 | 0.021786 | 21.6598 | 2.6158 | 0.15596 | 0.050246 | 3.2565 | 3.2565 | 0.00193
2019-04-09 14:35 | 1554834896.5233 | 15 | 0.20003 | 0.006606 | 0.000105 | 100.3564 | 73.5549 | 102.5345 | 20.4865 | 0.00238 | 0.00236 | 265.5633 | 0.26284 | 8.4018 | 0.002994 | 0.004805 | 0.004642 | 4.4e-05 | 0.001092 | 0.26753 | 1.3508 | 2.5475 | 98.55 | 20.09 | 49.846 | 78.3214 | 93.1013 | 77.8421 | 0.011904 | 1.0058 | 0.021782 | 21.655 | 2.779 | 0.15636 | 0.049108 | 3.1821 | 3.1821 | 0.00198
2019-04-09 14:35 | 1554834897.524 | 16 | 0.20005 | 0.006919 | 0.000109 | 100.3114 | 73.6588 | 102.4561 | 20.4858 | 0.002424 | 0.002395 | 265.3856 | 0.26183 | 8.4011 | 0.003013 | 0.003985 | 0.003985 | 4.5e-05 | 0.001094 | 0.25642 | 1.3393 | 2.5474 | 98.55 | 20.09 | 49.85 | 79.3732 | 93.5552 | 78.6934 | 0.011905 | 1.0034 | 0.021747 | 21.673 | 2.7624 | 0.15576 | 0.048637 | 3.1464 | 3.1464 | 0.001995
2019-04-09 14:35 | 1554834898.5234 | 17 | 0.20008 | 0.007117 | 0.000113 | 100.3668 | 74.883 | 102.3824 | 20.4865 | 0.002503 | 0.002464 | 265.294 | 0.2626 | 8.403 | 0.003007 | 0.004556 | 0.004634 | 4.5e-05 | 0.001097 | 0.27433 | 1.3197 | 2.5475 | 98.55 | 20.09 | 49.861 | 81.2935 | 93.4133 | 80.3524 | 0.011908 | 0.99871 | 0.021688 | 21.7162 | 2.6199 | 0.15622 | 0.048195 | 3.1093 | 3.1093 | 0.002025
2019-04-09 14:35 | 1554834899.5231 | 18 | 0.20006 | 0.007264 | 0.000115 | 100.3558 | 75.0348 | 101.6787 | 20.486 | 0.002593 | 0.00254 | 265.1161 | 0.26231 | 8.4022 | 0.002952 | 0.00381 | 0.003774 | 4.6e-05 | 0.001107 | 0.28008 | 1.3465 | 2.5475 | 98.55 | 20.09 | 49.868 | 83.456 | 92.091 | 82.1774 | 0.01191 | 0.99358 | 0.021558 | 21.6972 | 2.6008 | 0.15605 | 0.048149 | 3.0876 | 3.0876 | 0.002037
2019-04-09 14:35 | 1554834900.5234 | 19 | 0.20006 | 0.007363 | 0.000116 | 100.4111 | 76.0951 | 101.2752 | 20.4873 | 0.002669 | 0.00267 | 265.0338 | 0.26248 | 8.4018 | 0.003031 | 0.004382 | 0.004295 | 4.4e-05 | 0.001108 | 0.28528 | 1.358 | 2.5475 | 98.55 | 20.09 | 49.882 | 85.2673 | 93.99 | 85.2907 | 0.011913 | 0.98495 | 0.021421 | 21.7488 | 2.4802 | 0.15615 | 0.048042 | 3.0611 | 3.0611 | 0.002056
2019-04-09 14:35 | 1554834901.5234 | 20 | 0.1999 | 0.007423 | 0.000118 | 100.3552 | 76.8929 | 100.3569 | 20.4867 | 0.002695 | 0.002682 | 264.8532 | 0.26228 | 8.4003 | 0.003039 | 0.004192 | 0.004218 | 4.6e-05 | 0.001108 | 0.27184 | 1.3787 | 2.5476 | 98.55 | 20.09 | 49.891 | 85.9052 | 94.1773 | 85.5898 | 0.011915 | 0.98413 | 0.021315 | 21.659 | 2.382 | 0.15603 | 0.048891 | 3.0995 | 3.0995 | 0.002029
2019-04-09 14:35 | 1554834902.5221 | 21 | 0.19949 | 0.007483 | 0.000121 | 100.331 | 77.9697 | 99.5507 | 20.4873 | 0.002797 | 0.00278 | 264.764 | 0.26155 | 8.4033 | 0.00308 | 0.004463 | 0.0044 | 4.8e-05 | 0.001107 | 0.2625 | 1.3599 | 2.5476 | 98.55 | 20.09 | 49.9 | 88.3489 | 95.1758 | 87.9505 | 0.011917 | 0.97769 | 0.02116 | 21.6427 | 2.2554 | 0.1556 | 0.051285 | 3.2269 | 3.2269 | 0.001943
2019-04-09 14:35 | 1554834903.5222 | 22 | 0.19825 | 0.007555 | 0.000123 | 100.3766 | 77.6214 | 98.9308 | 20.487 | 0.00287 | 0.002909 | 264.6009 | 0.26286 | 8.3971 | 0.003173 | 0.004348 | 0.004364 | 4.9e-05 | 0.001108 | 0.27383 | 1.3548 | 2.5476 | 98.55 | 20.09 | 49.917 | 90.1088 | 97.4222 | 91.0545 | 0.011921 | 0.96936 | 0.021004 | 21.6677 | 2.2995 | 0.15637 | 0.058556 | 3.6553 | 3.6553 | 0.001724
2019-04-09 14:35 | 1554834904.5233 | 23 | 0.19726 | 0.00767 | 0.000125 | 100.3016 | 77.7722 | 98.6775 | 20.4873 | 0.002947 | 0.002878 | 264.4715 | 0.26206 | 8.4005 | 0.003133 | 0.00448 | 0.004393 | 4.6e-05 | 0.001108 | 0.2808 | 1.3728 | 2.5476 | 98.55 | 20.09 | 49.926 | 91.9633 | 96.4507 | 90.3145 | 0.011923 | 0.97133 | 0.020998 | 21.618 | 2.2755 | 0.1559 | 0.064302 | 4.011 | 4.011 | 0.001566
2019-04-09 14:35 | 1554834905.5236 | 24 | 0.19662 | 0.00785 | 0.000127 | 100.3156 | 76.9259 | 98.5038 | 20.4876 | 0.003029 | 0.002927 | 264.3312 | 0.26213 | 8.4026 | 0.003206 | 0.005059 | 0.004979 | 4.6e-05 | 0.001105 | 0.2573 | 1.349 | 2.5476 | 98.55 | 20.09 | 49.94 | 93.9543 | 98.2173 | 91.4787 | 0.011927 | 0.96823 | 0.020946 | 21.6335 | 2.3746 | 0.15594 | 0.06787 | 4.2217 | 4.2217 | 0.001489
2019-04-09 14:35 | 1554834906.5223 | 25 | 0.19636 | 0.008232 | 0.000129 | 100.2363 | 76.3821 | 98.105 | 20.4875 | 0.003028 | 0.002947 | 264.1658 | 0.26236 | 8.4017 | 0.003247 | 0.00514 | 0.005089 | 4.7e-05 | 0.001104 | 0.26106 | 1.3029 | 2.5475 | 98.55 | 20.09 | 49.946 | 93.9089 | 99.2177 | 91.9719 | 0.011928 | 0.96693 | 0.02089 | 21.6042 | 2.431 | 0.15608 | 0.068957 | 4.2773 | 4.2773 | 0.001471
2019-04-09 14:35 | 1554834907.5223 | 26 | 0.19632 | 0.008728 | 0.000131 | 100.2867 | 76.8113 | 98.2925 | 20.4871 | 0.003159 | 0.002954 | 264.0618 | 0.26161 | 8.4019 | 0.003256 | 0.004942 | 0.0049 | 4.8e-05 | 0.001097 | 0.25625 | 1.3163 | 2.5475 | 98.55 | 20.09 | 49.951 | 97.0735 | 99.4323 | 92.1398 | 0.011929 | 0.96648 | 0.020905 | 21.6299 | 2.3853 | 0.15563 | 0.068587 | 4.2575 | 4.2575 | 0.001473
2019-04-09 14:35 | 1554834908.5223 | 27 | 0.19633 | 0.009114 | 0.000132 | 100.2932 | 76.3298 | 98.6876 | 20.4881 | 0.00315 | 0.00307 | 263.8701 | 0.26171 | 8.4012 | 0.003273 | 0.00513 | 0.005109 | 4.8e-05 | 0.001098 | 0.25354 | 1.3282 | 2.5476 | 98.55 | 20.09 | 49.957 | 96.8683 | 99.8475 | 94.9379 | 0.011931 | 0.95913 | 0.020867 | 21.7561 | 2.4422 | 0.15569 | 0.068061 | 4.2173 | 4.2173 | 0.001488
2019-04-09 14:35 | 1554834909.5224 | 28 | 0.19651 | 0.009356 | 0.000135 | 100.2634 | 75.7255 | 99.2307 | 20.4875 | 0.003224 | 0.00312 | 263.7705 | 0.26243 | 8.4028 | 0.003303 | 0.004282 | 0.004217 | 5.2e-05 | 0.001093 | 0.25802 | 1.3695 | 2.5475 | 98.55 | 20.09 | 49.959 | 98.643 | 100.5617 | 96.1431 | 0.011931 | 0.956 | 0.02089 | 21.8516 | 2.5106 | 0.15612 | 0.066698 | 4.1377 | 4.1377 | 0.001521
2019-04-09 14:35 | 1554834910.5224 | 29 | 0.19661 | 0.009566 | 0.000136 | 100.2506 | 75.0255 | 99.688 | 20.4874 | 0.003191 | 0.003143 | 263.5771 | 0.26182 | 8.4019 | 0.00333 | 0.004953 | 0.004951 | 5e-05 | 0.001094 | 0.25194 | 1.3435 | 2.5475 | 98.55 | 20.09 | 49.959 | 97.8612 | 101.2069 | 96.6961 | 0.011931 | 0.95457 | 0.020923 | 21.9183 | 2.5925 | 0.15576 | 0.065884 | 4.0937 | 4.0937 | 0.001533
2019-04-09 14:35 | 1554834911.5223 | 30 | 0.19685 | 0.009716 | 0.000139 | 100.1979 | 74.2793 | 100.5154 | 20.4881 | 0.003298 | 0.003133 | 263.4867 | 0.26214 | 8.4028 | 0.003398 | 0.004936 | 0.004821 | 5e-05 | 0.001094 | 0.26913 | 1.3281 | 2.5475 | 98.55 | 20.08 | 49.959 | 100.4469 | 102.8608 | 96.4476 | 0.011924 | 0.95522 | 0.021016 | 22.0017 | 2.6772 | 0.15595 | 0.064287 | 4.0128 | 4.0128 | 0.001566
2019-04-09 14:35 | 1554834912.5224 | 31 | 0.19713 | 0.009815 | 0.000139 | 100.211 | 73.0924 | 101.1698 | 20.4877 | 0.003314 | 0.003233 | 263.3028 | 0.26161 | 8.4027 | 0.003446 | 0.004576 | 0.004594 | 5.3e-05 | 0.00109 | 0.2521 | 1.3425 | 2.5475 | 98.55 | 20.08 | 49.958 | 100.8336 | 104.0266 | 98.8636 | 0.011924 | 0.94901 | 0.021016 | 22.1452 | 2.8225 | 0.15563 | 0.062509 | 3.9022 | 3.9022 | 0.001607
2019-04-09 14:35 | 1554834913.5236 | 32 | 0.19742 | 0.009849 | 0.000139 | 100.1988 | 72.5545 | 101.5613 | 20.4876 | 0.003424 | 0.003195 | 263.2051 | 0.26205 | 8.4006 | 0.003518 | 0.004858 | 0.004767 | 4.8e-05 | 0.001085 | 0.262 | 1.3578 | 2.5476 | 98.55 | 20.08 | 49.956 | 103.4898 | 105.7788 | 97.9588 | 0.011923 | 0.95133 | 0.021082 | 22.161 | 2.8875 | 0.15589 | 0.060754 | 3.8052 | 3.8052 | 0.001651
2019-04-09 14:35 | 1554834914.5216 | 33 | 0.19776 | 0.009834 | 0.000138 | 100.1324 | 72.3827 | 102.0343 | 20.4883 | 0.003405 | 0.003268 | 263.0548 | 0.26087 | 8.4026 | 0.003443 | 0.005291 | 0.005184 | 4.5e-05 | 0.001092 | 0.27079 | 1.3477 | 2.5476 | 98.55 | 20.08 | 49.954 | 103.0385 | 103.9479 | 99.7144 | 0.011923 | 0.94685 | 0.021082 | 22.265 | 2.9027 | 0.15519 | 0.058714 | 3.6778 | 3.6778 | 0.001701
2019-04-09 14:35 | 1554834915.522 | 34 | 0.19806 | 0.009776 | 0.000135 | 100.1458 | 72.9564 | 102.7539 | 20.4885 | 0.0034 | 0.003194 | 262.9389 | 0.2616 | 8.4028 | 0.003513 | 0.004527 | 0.004617 | 4.9e-05 | 0.001088 | 0.23955 | 1.3542 | 2.5476 | 98.55 | 20.08 | 49.951 | 102.9004 | 105.6437 | 97.9251 | 0.011922 | 0.95141 | 0.021207 | 22.2897 | 2.8333 | 0.15563 | 0.056999 | 3.5922 | 3.5922 | 0.001746
2019-04-09 14:35 | 1554834916.522 | 35 | 0.19853 | 0.00968 | 0.000133 | 100.225 | 72.8222 | 103.5018 | 20.4875 | 0.003402 | 0.003236 | 262.8503 | 0.26271 | 8.4024 | 0.003551 | 0.004709 | 0.004602 | 4.6e-05 | 0.001086 | 0.26753 | 1.3321 | 2.5475 | 98.55 | 20.08 | 49.949 | 102.9537 | 106.5611 | 98.9308 | 0.011921 | 0.94884 | 0.021255 | 22.401 | 2.8569 | 0.15629 | 0.054316 | 3.4317 | 3.4317 | 0.001835
2019-04-09 14:35 | 1554834917.5226 | 36 | 0.19878 | 0.009547 | 0.000128 | 100.1595 | 73.1522 | 104.1705 | 20.4878 | 0.003564 | 0.003201 | 262.7019 | 0.26168 | 8.4015 | 0.003507 | 0.005007 | 0.004961 | 5.1e-05 | 0.001074 | 0.26161 | 1.3123 | 2.5477 | 98.55 | 20.08 | 49.947 | 106.8795 | 105.5133 | 98.0931 | 0.011921 | 0.95098 | 0.021348 | 22.4479 | 2.8106 | 0.15567 | 0.053018 | 3.3648 | 3.3648 | 0.001865
2019-04-09 14:35 | 1554834918.5227 | 37 | 0.19926 | 0.009375 | 0.000123 | 100.1717 | 74.2284 | 104.9906 | 20.4892 | 0.003458 | 0.003217 | 262.6401 | 0.26232 | 8.4031 | 0.00353 | 0.005615 | 0.005534 | 4.8e-05 | 0.001066 | 0.26504 | 1.3091 | 2.5476 | 98.55 | 20.08 | 49.943 | 104.3229 | 106.055 | 98.491 | 0.01192 | 0.94996 | 0.02142 | 22.5482 | 2.681 | 0.15606 | 0.050367 | 3.2081 | 3.2081 | 0.00196
2019-04-09 14:35 | 1554834919.5217 | 38 | 0.19994 | 0.009189 | 0.000119 | 100.0946 | 75.6419 | 105.6556 | 20.4889 | 0.003509 | 0.003231 | 262.4813 | 0.26198 | 8.4013 | 0.003498 | 0.005292 | 0.00523 | 5.1e-05 | 0.001066 | 0.2589 | 1.32 | 2.5476 | 98.55 | 20.08 | 49.942 | 105.5488 | 105.294 | 98.8278 | 0.01192 | 0.9491 | 0.021478 | 22.6297 | 2.5054 | 0.15585 | 0.046534 | 2.973 | 2.973 | 0.002113
2019-04-09 14:35 | 1554834920.5221 | 39 | 0.20039 | 0.008953 | 0.000115 | 100.0994 | 76.6702 | 106.2379 | 20.4897 | 0.003511 | 0.003234 | 262.4428 | 0.2621 | 8.4016 | 0.003489 | 0.00519 | 0.005192 | 4.9e-05 | 0.001064 | 0.27472 | 1.3393 | 2.5476 | 98.55 | 20.08 | 49.941 | 105.6126 | 105.0627 | 98.897 | 0.01192 | 0.94893 | 0.021535 | 22.6941 | 2.3851 | 0.15592 | 0.04414 | 2.8281 | 2.8281 | 0.002222
2019-04-09 14:35 | 1554834921.5234 | 40 | 0.20083 | 0.008601 | 0.00011 | 100.125 | 77.8262 | 106.4269 | 20.4893 | 0.003485 | 0.00321 | 262.3081 | 0.26295 | 8.3995 | 0.003493 | 0.004603 | 0.004545 | 5.3e-05 | 0.001057 | 0.26161 | 1.3238 | 2.5475 | 98.55 | 20.08 | 49.939 | 104.9765 | 105.1551 | 98.3052 | 0.011919 | 0.95044 | 0.021571 | 22.6962 | 2.2535 | 0.15643 | 0.04194 | 2.6922 | 2.6922 | 0.002342
2019-04-09 14:35 | 1554834922.5218 | 41 | 0.20129 | 0.008177 | 0.000106 | 100.0946 | 79.4565 | 106.9099 | 20.4898 | 0.003498 | 0.003202 | 262.258 | 0.26206 | 8.4 | 0.003405 | 0.004845 | 0.004778 | 5.4e-05 | 0.00105 | 0.25066 | 1.3595 | 2.5475 | 98.55 | 20.08 | 49.94 | 105.2776 | 103.0242 | 98.1192 | 0.011919 | 0.95091 | 0.021626 | 22.742 | 2.0654 | 0.1559 | 0.039677 | 2.5539 | 2.5539 | 0.00246
2019-04-09 14:35 | 1554834923.5229 | 42 | 0.20183 | 0.007834 | 0.000101 | 100.0678 | 80.7119 | 107.258 | 20.4903 | 0.003544 | 0.003217 | 262.1427 | 0.26195 | 8.4007 | 0.00349 | 0.004771 | 0.00465 | 5.3e-05 | 0.001049 | 0.2792 | 1.3187 | 2.5475 | 98.55 | 20.08 | 49.94 | 106.3954 | 105.0948 | 98.492 | 0.011919 | 0.94996 | 0.02165 | 22.7904 | 1.9227 | 0.15583 | 0.036906 | 2.3788 | 2.3788 | 0.00264
2019-04-09 14:35 | 1554834924.5219 | 43 | 0.20234 | 0.007452 | 9.7e-05 | 99.9814 | 82.3711 | 107.4348 | 20.4906 | 0.003521 | 0.003258 | 262.0767 | 0.26166 | 8.4029 | 0.003572 | 0.005488 | 0.005475 | 5.7e-05 | 0.001048 | 0.27151 | 1.3334 | 2.5476 | 98.55 | 20.08 | 49.94 | 105.8417 | 107.0751 | 99.4814 | 0.011919 | 0.94744 | 0.021639 | 22.8395 | 1.733 | 0.15566 | 0.034294 | 2.2097 | 2.2097 | 0.002839
2019-04-09 14:35 | 1554834925.5233 | 44 | 0.20253 | 0.007209 | 9.2e-05 | 99.9624 | 83.5082 | 107.4205 | 20.4903 | 0.003517 | 0.003221 | 261.9903 | 0.26191 | 8.4013 | 0.003423 | 0.00481 | 0.004706 | 5.7e-05 | 0.001052 | 0.26576 | 1.3151 | 2.5474 | 98.55 | 20.08 | 49.94 | 105.7374 | 103.482 | 98.5846 | 0.011919 | 0.94972 | 0.021664 | 22.8105 | 1.6087 | 0.15581 | 0.033468 | 2.1592 | 2.1592 | 0.002908
2019-04-09 14:35 | 1554834926.523 | 45 | 0.20283 | 0.006965 | 8.9e-05 | 99.9567 | 84.4092 | 107.2026 | 20.4904 | 0.003479 | 0.003218 | 261.9055 | 0.26177 | 8.3995 | 0.003481 | 0.004826 | 0.004818 | 5.4e-05 | 0.001048 | 0.26537 | 1.3243 | 2.5475 | 98.55 | 20.08 | 49.941 | 104.8384 | 104.8699 | 98.5054 | 0.01192 | 0.94993 | 0.021644 | 22.7849 | 1.5122 | 0.15573 | 0.031965 | 2.0607 | 2.0607 | 0.003045
2019-04-09 14:35 | 1554834927.5219 | 46 | 0.20296 | 0.006728 | 8.5e-05 | 99.9883 | 85.5679 | 107.2785 | 20.49 | 0.003578 | 0.003269 | 261.8575 | 0.26155 | 8.4004 | 0.003446 | 0.005313 | 0.005231 | 5.2e-05 | 0.001051 | 0.26504 | 1.2997 | 2.5475 | 98.55 | 20.08 | 49.941 | 107.2177 | 104.0343 | 99.7537 | 0.01192 | 0.94675 | 0.021615 | 22.8312 | 1.393 | 0.1556 | 0.031503 | 2.0284 | 2.0284 | 0.003091
2019-04-09 14:35 | 1554834928.5237 | 47 | 0.20302 | 0.0065 | 8.2e-05 | 99.9278 | 86.5563 | 106.8752 | 20.4901 | 0.003573 | 0.003233 | 261.7393 | 0.2616 | 8.4023 | 0.003353 | 0.005151 | 0.005091 | 5.6e-05 | 0.00105 | 0.23955 | 1.3174 | 2.5475 | 98.55 | 20.08 | 49.94 | 107.1013 | 101.7828 | 98.8674 | 0.011919 | 0.949 | 0.0216 | 22.7612 | 1.2849 | 0.15562 | 0.031389 | 2.0199 | 2.0199 | 0.003105
2019-04-09 14:35 | 1554834929.5223 | 48 | 0.20302 | 0.006308 | 8e-05 | 99.8978 | 86.9029 | 106.7792 | 20.4902 | 0.003568 | 0.003267 | 261.7133 | 0.26175 | 8.3997 | 0.003445 | 0.005048 | 0.004972 | 5.4e-05 | 0.001052 | 0.25138 | 1.347 | 2.5475 | 98.55 | 20.07 | 49.94 | 106.9843 | 104.0094 | 99.6856 | 0.011912 | 0.94692 | 0.021567 | 22.776 | 1.2465 | 0.15571 | 0.031646 | 2.0334 | 2.0334 | 0.003086
2019-04-09 14:35 | 1554834930.522 | 49 | 0.20303 | 0.006124 | 7.8e-05 | 99.9016 | 86.429 | 106.632 | 20.4902 | 0.003544 | 0.003248 | 261.571 | 0.2613 | 8.4014 | 0.003443 | 0.004935 | 0.004887 | 5.3e-05 | 0.001047 | 0.26842 | 1.3556 | 2.5476 | 98.55 | 20.07 | 49.941 | 106.3978 | 103.9561 | 99.2347 | 0.011912 | 0.94807 | 0.021565 | 22.7465 | 1.2957 | 0.15545 | 0.031813 | 2.0441 | 2.0441 | 0.003065
2019-04-09 14:36 | 1554834931.5211 | 50 | 0.20299 | 0.005971 | 7.7e-05 | 99.9287 | 85.9076 | 106.7189 | 20.4908 | 0.003518 | 0.003264 | 261.5276 | 0.26158 | 8.4034 | 0.003391 | 0.004514 | 0.004429 | 5.3e-05 | 0.001048 | 0.26106 | 1.3418 | 2.5475 | 98.55 | 20.07 | 49.941 | 105.7662 | 102.7 | 99.6124 | 0.011912 | 0.94711 | 0.021563 | 22.7673 | 1.3523 | 0.15562 | 0.032256 | 2.0724 | 2.0724 | 0.003026
2019-04-09 14:36 | 1554834932.5212 | 51 | 0.20296 | 0.005844 | 7.5e-05 | 99.9115 | 86.0047 | 106.3868 | 20.4912 | 0.003588 | 0.00329 | 261.3986 | 0.26183 | 8.4022 | 0.003432 | 0.00502 | 0.004886 | 5.2e-05 | 0.00105 | 0.26665 | 1.3437 | 2.5476 | 98.55 | 20.07 | 49.94 | 107.4687 | 103.6794 | 100.25 | 0.011912 | 0.94549 | 0.021511 | 22.7513 | 1.3406 | 0.15576 | 0.032576 | 2.0879 | 2.0879 | 0.003006
2019-04-09 14:36 | 1554834933.5211 | 52 | 0.20296 | 0.005755 | 7.4e-05 | 99.8686 | 86.1217 | 106.3307 | 20.4907 | 0.003515 | 0.003283 | 261.3244 | 0.26227 | 8.4003 | 0.003427 | 0.003923 | 0.003936 | 5.3e-05 | 0.001049 | 0.26377 | 1.3398 | 2.5476 | 98.55 | 20.07 | 49.94 | 105.689 | 103.5594 | 100.0736 | 0.011912 | 0.94593 | 0.02151 | 22.7399 | 1.3246 | 0.15603 | 0.032695 | 2.0956 | 2.0956 | 0.003001
2019-04-09 14:36 | 1554834934.5222 | 53 | 0.20296 | 0.005695 | 7.3e-05 | 99.8334 | 86.7842 | 106.6731 | 20.4913 | 0.003606 | 0.003329 | 261.2213 | 0.26177 | 8.4015 | 0.003502 | 0.005024 | 0.004965 | 5.5e-05 | 0.00105 | 0.26272 | 1.3262 | 2.5476 | 98.55 | 20.07 | 49.94 | 107.9067 | 105.3828 | 101.1879 | 0.011912 | 0.94312 | 0.021513 | 22.8105 | 1.2529 | 0.15573 | 0.032771 | 2.1007 | 2.1007 | 0.002988
2019-04-09 14:36 | 1554834935.5212 | 54 | 0.20294 | 0.005647 | 7.2e-05 | 99.8486 | 87.7058 | 107.0853 | 20.4919 | 0.003552 | 0.00331 | 261.1421 | 0.26281 | 8.4029 | 0.003415 | 0.005376 | 0.005321 | 5.3e-05 | 0.001046 | 0.26521 | 1.3022 | 2.5475 | 98.55 | 20.07 | 49.941 | 106.6088 | 103.2675 | 100.7475 | 0.011912 | 0.94423 | 0.021567 | 22.841 | 1.1598 | 0.15635 | 0.03296 | 2.1182 | 2.1182 | 0.002975
2019-04-09 14:36 | 1554834936.5222 | 55 | 0.20298 | 0.005614 | 7.3e-05 | 99.7804 | 88.116 | 107.673 | 20.4915 | 0.003596 | 0.003277 | 261.0553 | 0.26094 | 8.4017 | 0.00351 | 0.004917 | 0.004841 | 5.1e-05 | 0.001045 | 0.27295 | 1.3192 | 2.5475 | 98.55 | 20.07 | 49.942 | 107.6564 | 105.5846 | 99.9458 | 0.011912 | 0.94626 | 0.02165 | 22.8791 | 1.112 | 0.15523 | 0.032765 | 2.1137 | 2.1137 | 0.00296
2019-04-09 14:36 | 1554834937.5212 | 56 | 0.20298 | 0.005592 | 7.2e-05 | 99.7384 | 87.5888 | 107.6038 | 20.4922 | 0.003562 | 0.003285 | 260.9526 | 0.26256 | 8.4011 | 0.003457 | 0.004538 | 0.004518 | 5.1e-05 | 0.001043 | 0.2557 | 1.3053 | 2.5476 | 98.55 | 20.07 | 49.942 | 106.8459 | 104.2953 | 100.1401 | 0.011912 | 0.94577 | 0.021637 | 22.8777 | 1.1619 | 0.1562 | 0.032801 | 2.1148 | 2.1148 | 0.002977
2019-04-09 14:36 | 1554834938.5223 | 57 | 0.20313 | 0.00558 | 7.1e-05 | 99.7346 | 87.8762 | 107.9596 | 20.491 | 0.003612 | 0.003287 | 260.9165 | 0.26156 | 8.4038 | 0.003481 | 0.005213 | 0.00511 | 4.8e-05 | 0.001042 | 0.24873 | 1.3637 | 2.5475 | 98.55 | 20.07 | 49.942 | 108.0482 | 104.8736 | 100.1711 | 0.011912 | 0.94569 | 0.021672 | 22.9164 | 1.1322 | 0.1556 | 0.031916 | 2.0612 | 2.0612 | 0.003042
2019-04-09 14:36 | 1554834939.5217 | 58 | 0.20331 | 0.005558 | 7.1e-05 | 99.7045 | 88.1869 | 108.4305 | 20.4922 | 0.003629 | 0.003291 | 260.7824 | 0.26137 | 8.4021 | 0.00349 | 0.005076 | 0.005056 | 5.5e-05 | 0.001039 | 0.27688 | 1.3348 | 2.5475 | 98.55 | 20.07 | 49.944 | 108.4756 | 105.0998 | 100.2705 | 0.011913 | 0.94544 | 0.021716 | 22.9694 | 1.098 | 0.15549 | 0.030858 | 1.9971 | 1.9971 | 0.003138
2019-04-09 14:36 | 1554834940.5228 | 59 | 0.20345 | 0.005532 | 6.9e-05 | 99.6628 | 88.9224 | 108.9309 | 20.492 | 0.003586 | 0.003241 | 260.7535 | 0.26225 | 8.3997 | 0.003442 | 0.004926 | 0.004863 | 5.5e-05 | 0.001041 | 0.26576 | 1.3021 | 2.5474 | 98.55 | 20.07 | 49.946 | 107.4284 | 103.9319 | 99.0582 | 0.011913 | 0.94852 | 0.021802 | 22.9849 | 1.0199 | 0.15601 | 0.03006 | 1.9532 | 1.9532 | 0.003219
2019-04-09 14:36 | 1554834941.5225 | 60 | 0.20351 | 0.005501 | 6.8e-05 | 99.5979 | 88.5948 | 109.2168 | 20.493 | 0.00366 | 0.003268 | 260.6381 | 0.26132 | 8.4041 | 0.003497 | 0.005389 | 0.005382 | 5.3e-05 | 0.00104 | 0.27743 | 1.3426 | 2.5475 | 98.55 | 20.07 | 49.947 | 109.2118 | 105.2548 | 99.7183 | 0.011914 | 0.94684 | 0.021811 | 23.0355 | 1.0471 | 0.15546 | 0.029736 | 1.933 | 1.933 | 0.003241
2019-04-09 14:36 | 1554834942.5214 | 61 | 0.20368 | 0.005456 | 6.7e-05 | 99.5898 | 88.75 | 109.7987 | 20.4922 | 0.003603 | 0.003266 | 260.597 | 0.26161 | 8.3986 | 0.003419 | 0.004938 | 0.00486 | 5.1e-05 | 0.001045 | 0.2756 | 1.3283 | 2.5474 | 98.55 | 20.07 | 49.948 | 107.8273 | 103.3645 | 99.6759 | 0.011914 | 0.94694 | 0.02187 | 23.0954 | 1.0307 | 0.15563 | 0.028764 | 1.8751 | 1.8751 | 0.003345
2019-04-09 14:36 | 1554834943.5215 | 62 | 0.20372 | 0.00539 | 6.6e-05 | 99.5821 | 89.1217 | 110.4294 | 20.4929 | 0.003618 | 0.003159 | 260.5101 | 0.26133 | 8.4023 | 0.003497 | 0.004898 | 0.004893 | 5.2e-05 | 0.001042 | 0.26969 | 1.3311 | 2.5477 | 98.55 | 20.07 | 49.949 | 108.1909 | 105.2765 | 97.083 | 0.011914 | 0.95358 | 0.022009 | 23.081 | 0.99267 | 0.15546 | 0.028618 | 1.8775 | 1.8775 | 0.003337
2019-04-09 14:36 | 1554834944.5215 | 63 | 0.20388 | 0.005315 | 6.4e-05 | 99.5693 | 89.3309 | 110.673 | 20.4928 | 0.003515 | 0.003198 | 260.4597 | 0.26195 | 8.4012 | 0.003393 | 0.00575 | 0.005666 | 5.6e-05 | 0.001036 | 0.2688 | 1.3528 | 2.5475 | 98.55 | 20.07 | 49.949 | 105.7093 | 102.7492 | 98.0335 | 0.011914 | 0.95113 | 0.022006 | 23.1361 | 0.97054 | 0.15584 | 0.027715 | 1.8181 | 1.8181 | 0.003454
2019-04-09 14:36 | 1554834945.5218 | 64 | 0.20405 | 0.005222 | 6.4e-05 | 99.5479 | 89.9728 | 110.6514 | 20.4927 | 0.003625 | 0.003175 | 260.3875 | 0.26169 | 8.4025 | 0.003435 | 0.00521 | 0.005234 | 5.2e-05 | 0.001033 | 0.26344 | 1.3288 | 2.5475 | 98.55 | 20.07 | 49.948 | 108.3603 | 103.7589 | 97.4816 | 0.011914 | 0.95255 | 0.02202 | 23.1166 | 0.90458 | 0.15568 | 0.026843 | 1.7621 | 1.7621 | 0.003561
2019-04-09 14:36 | 1554834946.5206 | 65 | 0.20439 | 0.005137 | 6.3e-05 | 99.5074 | 90.1975 | 111.1174 | 20.4934 | 0.003569 | 0.003172 | 260.3148 | 0.26117 | 8.402 | 0.003429 | 0.005373 | 0.005274 | 5.3e-05 | 0.001033 | 0.2521 | 1.3531 | 2.5476 | 98.55 | 20.07 | 49.947 | 107.0133 | 103.6086 | 97.411 | 0.011914 | 0.95273 | 0.022068 | 23.163 | 0.87862 | 0.15537 | 0.024871 | 1.6364 | 1.6364 | 0.003826
2019-04-09 14:36 | 1554834947.5217 | 66 | 0.20472 | 0.00503 | 6.3e-05 | 99.5199 | 90.4033 | 111.4394 | 20.4923 | 0.003512 | 0.003189 | 260.2757 | 0.26193 | 8.4024 | 0.003438 | 0.00488 | 0.004848 | 5.2e-05 | 0.001038 | 0.24386 | 1.3272 | 2.5476 | 98.55 | 20.07 | 49.946 | 105.6285 | 103.8471 | 97.8097 | 0.011913 | 0.95171 | 0.022088 | 23.209 | 0.85936 | 0.15582 | 0.023063 | 1.5189 | 1.5189 | 0.004134
2019-04-09 14:36 | 1554834948.5206 | 67 | 0.20496 | 0.004926 | 6.1e-05 | 99.4329 | 90.1218 | 111.5666 | 20.4926 | 0.003545 | 0.003169 | 260.1829 | 0.26181 | 8.4001 | 0.003365 | 0.004541 | 0.004441 | 5.5e-05 | 0.001037 | 0.25227 | 1.3122 | 2.5475 | 98.55 | 20.07 | 49.945 | 106.4261 | 102.0764 | 97.3247 | 0.011913 | 0.95295 | 0.022115 | 23.2071 | 0.87944 | 0.15575 | 0.021759 | 1.4348 | 1.4348 | 0.004375
2019-04-09 14:36 | 1554834949.521 | 68 | 0.20519 | 0.004808 | 6.1e-05 | 99.4183 | 90.459 | 111.9485 | 20.493 | 0.003533 | 0.003186 | 260.1592 | 0.26102 | 8.4022 | 0.003311 | 0.004625 | 0.00459 | 5.4e-05 | 0.001038 | 0.26737 | 1.3189 | 2.5475 | 98.55 | 20.07 | 49.945 | 106.1327 | 100.7587 | 97.7406 | 0.011913 | 0.95189 | 0.022141 | 23.2598 | 0.84471 | 0.15528 | 0.020481 | 1.3521 | 1.3521 | 0.004628
2019-04-09 14:36 | 1554834950.5221 | 69 | 0.20529 | 0.004687 | 5.9e-05 | 99.3912 | 90.6359 | 112.628 | 20.4932 | 0.003521 | 0.003194 | 260.0457 | 0.26163 | 8.4022 | 0.003296 | 0.005531 | 0.005443 | 5.3e-05 | 0.001032 | 0.27793 | 1.3279 | 2.5476 | 98.55 | 20.07 | 49.943 | 105.8464 | 100.3893 | 97.943 | 0.011913 | 0.95137 | 0.022202 | 23.3367 | 0.8248 | 0.15564 | 0.020035 | 1.3264 | 1.3264 | 0.004729
2019-04-09 14:36 | 1554834951.5212 | 70 | 0.20545 | 0.004557 | 5.7e-05 | 99.4076 | 91.3098 | 112.9365 | 20.4938 | 0.003497 | 0.003158 | 260.0395 | 0.26168 | 8.4023 | 0.003364 | 0.004764 | 0.004727 | 5.2e-05 | 0.001026 | 0.24657 | 1.291 | 2.5475 | 98.55 | 20.07 | 49.944 | 105.2578 | 102.0423 | 97.0668 | 0.011913 | 0.95362 | 0.022258 | 23.341 | 0.76002 | 0.15567 | 0.019263 | 1.2786 | 1.2786 | 0.004907
2019-04-09 14:36 | 1554834952.5218 | 71 | 0.20553 | 0.004418 | 5.6e-05 | 99.3325 | 91.3848 | 113.2111 | 20.493 | 0.003471 | 0.003071 | 259.9523 | 0.26102 | 8.4014 | 0.003302 | 0.005154 | 0.005112 | 5.4e-05 | 0.001019 | 0.26001 | 1.3191 | 2.5475 | 98.55 | 20.07 | 49.944 | 104.6342 | 100.5356 | 94.9646 | 0.011913 | 0.95906 | 0.022349 | 23.3029 | 0.74592 | 0.15528 | 0.018955 | 1.2634 | 1.2634 | 0.004953
2019-04-09 14:36 | 1554834953.5218 | 72 | 0.20574 | 0.004279 | 5.4e-05 | 99.3605 | 91.3903 | 113.3599 | 20.4936 | 0.003472 | 0.003149 | 259.9427 | 0.26182 | 8.4012 | 0.003296 | 0.004908 | 0.004854 | 5.4e-05 | 0.001016 | 0.27079 | 1.3526 | 2.5475 | 98.55 | 20.07 | 49.943 | 104.6583 | 100.3887 | 96.8564 | 0.011913 | 0.95416 | 0.022306 | 23.3781 | 0.7479 | 0.15576 | 0.017862 | 1.1883 | 1.1883 | 0.005282
2019-04-09 14:36 | 1554834954.5208 | 73 | 0.20599 | 0.004144 | 5.3e-05 | 99.2879 | 91.6126 | 113.7876 | 20.4946 | 0.003463 | 0.00304 | 259.8624 | 0.26226 | 8.4012 | 0.003261 | 0.004611 | 0.004529 | 5.4e-05 | 0.001017 | 0.25426 | 1.3591 | 2.5475 | 98.55 | 20.07 | 49.942 | 104.4391 | 99.5573 | 94.2146 | 0.011912 | 0.96102 | 0.022429 | 23.3384 | 0.71963 | 0.15602 | 0.016512 | 1.1045 | 1.1045 | 0.005693
2019-04-09 14:36 | 1554834955.5211 | 74 | 0.20614 | 0.003995 | 5.1e-05 | 99.3662 | 91.8745 | 113.8207 | 20.4936 | 0.003396 | 0.003079 | 259.8186 | 0.26107 | 8.4043 | 0.003313 | 0.004942 | 0.004937 | 5.5e-05 | 0.001012 | 0.26609 | 1.3503 | 2.5475 | 98.55 | 20.07 | 49.942 | 102.8245 | 100.8212 | 95.1577 | 0.011912 | 0.95856 | 0.022403 | 23.3717 | 0.70115 | 0.15531 | 0.015772 | 1.0539 | 1.0539 | 0.005939
2019-04-09 14:36 | 1554834956.522 | 75 | 0.20628 | 0.003855 | 4.9e-05 | 99.3337 | 92.3894 | 113.7021 | 20.4941 | 0.003404 | 0.003008 | 259.7694 | 0.26231 | 8.4015 | 0.003291 | 0.00479 | 0.004687 | 5.4e-05 | 0.001006 | 0.25818 | 1.3244 | 2.5474 | 98.55 | 20.07 | 49.942 | 103.0221 | 100.2707 | 93.4552 | 0.011912 | 0.96301 | 0.022443 | 23.3055 | 0.64824 | 0.15604 | 0.01513 | 1.0128 | 1.0128 | 0.006209
2019-04-09 14:36 | 1554834957.5223 | 76 | 0.20637 | 0.003715 | 4.8e-05 | 99.2349 | 92.9428 | 114.0986 | 20.4932 | 0.003356 | 0.002983 | 259.6965 | 0.26121 | 8.4009 | 0.003274 | 0.004755 | 0.004737 | 5.3e-05 | 0.00101 | 0.26864 | 1.3546 | 2.5475 | 98.55 | 20.07 | 49.942 | 101.8579 | 99.8587 | 92.8479 | 0.011912 | 0.96461 | 0.022501 | 23.3267 | 0.58591 | 0.1554 | 0.014751 | 0.99 | 0.99 | 0.006326
2019-04-09 14:36 | 1554834958.5212 | 77 | 0.20648 | 0.003584 | 4.7e-05 | 99.1851 | 93.4687 | 114.326 | 20.4945 | 0.003419 | 0.002971 | 259.6714 | 0.26081 | 8.4018 | 0.003149 | 0.004659 | 0.00458 | 5e-05 | 0.001008 | 0.24403 | 1.3331 | 2.5475 | 98.55 | 20.07 | 49.944 | 103.3778 | 96.8503 | 92.5473 | 0.011913 | 0.9654 | 0.022533 | 23.3404 | 0.53097 | 0.15515 | 0.014255 | 0.9581 | 0.9581 | 0.006526
2019-04-09 14:36 | 1554834959.521 | 78 | 0.20653 | 0.003459 | 4.6e-05 | 99.1703 | 93.4738 | 114.1951 | 20.494 | 0.003347 | 0.003021 | 259.5806 | 0.26089 | 8.4017 | 0.003138 | 0.004605 | 0.00454 | 5.4e-05 | 0.001016 | 0.25514 | 1.3584 | 2.5475 | 98.55 | 20.07 | 49.947 | 101.638 | 96.5765 | 93.7626 | 0.011914 | 0.96221 | 0.022483 | 23.3657 | 0.52913 | 0.15521 | 0.014154 | 0.94924 | 0.94924 | 0.006589
2019-04-09 14:36 | 1554834960.5451 | 79 | 0.20657 | 0.003345 | 4.5e-05 | 99.1524 | 93.8589 | 113.9662 | 20.494 | 0.00343 | 0.002995 | 259.5801 | 0.26063 | 8.4029 | 0.003175 | 0.004663 | 0.00465 | 5.2e-05 | 0.001015 | 0.27743 | 1.3405 | 2.5475 | 98.55 | 20.07 | 49.948 | 103.6435 | 97.4664 | 93.1264 | 0.011914 | 0.96388 | 0.02248 | 23.3221 | 0.49074 | 0.15505 | 0.014012 | 0.93958 | 0.93958 | 0.00665
2019-04-09 14:36 | 1554834961.5368 | 80 | 0.20662 | 0.003239 | 4.4e-05 | 99.1807 | 94.0596 | 114.0289 | 20.495 | 0.003328 | 0.002935 | 259.4724 | 0.26177 | 8.4016 | 0.003141 | 0.004583 | 0.004504 | 5.6e-05 | 0.001009 | 0.26449 | 1.3149 | 2.5476 | 98.55 | 20.07 | 49.949 | 101.1821 | 96.647 | 91.6782 | 0.011914 | 0.9677 | 0.02253 | 23.2823 | 0.47419 | 0.15572 | 0.013852 | 0.93106 | 0.93106 | 0.00674
2019-04-09 14:36 | 1554834962.536 | 81 | 0.20666 | 0.003144 | 4.2e-05 | 99.1047 | 94.2712 | 114.0524 | 20.4946 | 0.00337 | 0.002986 | 259.4784 | 0.2613 | 8.4018 | 0.003037 | 0.00473 | 0.004698 | 5.6e-05 | 0.001012 | 0.27848 | 1.3344 | 2.5476 | 98.55 | 20.07 | 49.95 | 102.1851 | 94.1341 | 92.9057 | 0.011914 | 0.96446 | 0.022495 | 23.3238 | 0.44724 | 0.15545 | 0.013732 | 0.92157 | 0.92157 | 0.006798
2019-04-09 14:36 | 1554834963.5359 | 82 | 0.20668 | 0.003061 | 4.2e-05 | 99.1024 | 94.2244 | 114.2169 | 20.4953 | 0.003356 | 0.002938 | 259.3958 | 0.26097 | 8.4013 | 0.003078 | 0.004399 | 0.00432 | 5.4e-05 | 0.00101 | 0.27793 | 1.2978 | 2.5476 | 98.55 | 20.07 | 49.95 | 101.8491 | 95.1249 | 91.7684 | 0.011914 | 0.96746 | 0.022546 | 23.3044 | 0.45146 | 0.15525 | 0.013722 | 0.92306 | 0.92306 | 0.006778
2019-04-09 14:36 | 1554834964.536 | 83 | 0.20672 | 0.002994 | 4.1e-05 | 99.0607 | 94.6522 | 114.1441 | 20.4948 | 0.003327 | 0.002965 | 259.3741 | 0.2604 | 8.4008 | 0.003092 | 0.005266 | 0.005263 | 5.4e-05 | 0.001013 | 0.26808 | 1.3145 | 2.5476 | 98.55 | 20.07 | 49.95 | 101.149 | 95.4649 | 92.4114 | 0.011914 | 0.96576 | 0.022519 | 23.3175 | 0.40718 | 0.15491 | 0.013585 | 0.91283 | 0.91283 | 0.006839
2019-04-09 14:36 | 1554834965.536 | 84 | 0.20669 | 0.002933 | 4.1e-05 | 99.0717 | 95.1296 | 114.4415 | 20.4951 | 0.003341 | 0.002985 | 259.3166 | 0.26043 | 8.4022 | 0.003141 | 0.00454 | 0.004524 | 5.2e-05 | 0.00101 | 0.25841 | 1.3662 | 2.5475 | 98.55 | 20.07 | 49.952 | 101.49 | 96.659 | 92.8828 | 0.011915 | 0.96452 | 0.022534 | 23.3629 | 0.36318 | 0.15493 | 0.013846 | 0.93098 | 0.93098 | 0.006707
2019-04-09 14:36 | 1554834966.5374 | 85 | 0.20674 | 0.002875 | 4.1e-05 | 99.0648 | 95.0707 | 114.5408 | 20.4955 | 0.003301 | 0.002887 | 259.2846 | 0.26159 | 8.4025 | 0.003123 | 0.004892 | 0.004849 | 4.8e-05 | 0.001001 | 0.25194 | 1.3487 | 2.5476 | 98.55 | 20.07 | 49.953 | 100.5335 | 96.2199 | 90.5398 | 0.011915 | 0.97073 | 0.022616 | 23.2981 | 0.3681 | 0.15562 | 0.013595 | 0.91738 | 0.91738 | 0.006836
2019-04-09 14:36 | 1554834967.5371 | 86 | 0.20675 | 0.00283 | 4.1e-05 | 98.98 | 95.1313 | 114.4866 | 20.4963 | 0.003262 | 0.00294 | 259.2548 | 0.26157 | 8.3983 | 0.003128 | 0.004825 | 0.004761 | 5.3e-05 | 0.001 | 0.274 | 1.3186 | 2.5477 | 98.55 | 20.07 | 49.954 | 99.5786 | 96.3436 | 91.8 | 0.011915 | 0.96738 | 0.022572 | 23.3329 | 0.35474 | 0.15561 | 0.013614 | 0.9169 | 0.9169 | 0.006839
2019-04-09 14:36 | 1554834968.5361 | 87 | 0.20677 | 0.002792 | 4.1e-05 | 99.0285 | 95.0831 | 114.6166 | 20.4952 | 0.003379 | 0.002829 | 259.1667 | 0.26105 | 8.404 | 0.003005 | 0.003925 | 0.003942 | 5e-05 | 0.000997 | 0.26913 | 1.32 | 2.5475 | 98.55 | 20.07 | 49.955 | 102.4198 | 93.3795 | 89.125 | 0.011916 | 0.97452 | 0.022668 | 23.2604 | 0.36365 | 0.1553 | 0.013524 | 0.91471 | 0.91471 | 0.006842
2019-04-09 14:36 | 1554834969.5371 | 88 | 0.20673 | 0.002755 | 4e-05 | 98.9642 | 95.0022 | 114.845 | 20.496 | 0.003245 | 0.002855 | 259.1813 | 0.2607 | 8.4037 | 0.00297 | 0.004837 | 0.004771 | 5.2e-05 | 0.00099 | 0.26034 | 1.3382 | 2.5475 | 98.55 | 20.07 | 49.954 | 99.1579 | 92.53 | 89.768 | 0.011915 | 0.9728 | 0.02267 | 23.3042 | 0.36545 | 0.15509 | 0.013821 | 0.93495 | 0.93495 | 0.006685
2019-04-09 14:36 | 1554834970.5371 | 89 | 0.20671 | 0.002723 | 4e-05 | 98.9418 | 95.1908 | 114.9978 | 20.4961 | 0.003318 | 0.002845 | 259.0985 | 0.2607 | 8.4026 | 0.003079 | 0.004175 | 0.004128 | 5.5e-05 | 0.000986 | 0.28982 | 1.3454 | 2.5476 | 98.55 | 20.07 | 49.953 | 100.9436 | 95.1609 | 89.5282 | 0.011915 | 0.97344 | 0.022693 | 23.312 | 0.3457 | 0.15509 | 0.013972 | 0.94611 | 0.94611 | 0.006606
2019-04-09 14:36 | 1554834971.5376 | 90 | 0.20669 | 0.002704 | 3.9e-05 | 98.9359 | 95.1378 | 114.979 | 20.4957 | 0.00329 | 0.002914 | 259.1071 | 0.26077 | 8.3998 | 0.003059 | 0.004547 | 0.004492 | 5.3e-05 | 0.000985 | 0.26609 | 1.3557 | 2.5475 | 98.55 | 20.07 | 49.951 | 100.2456 | 94.6767 | 91.1863 | 0.011915 | 0.96901 | 0.022639 | 23.3633 | 0.35014 | 0.15513 | 0.014102 | 0.95273 | 0.95273 | 0.006562
2019-04-09 14:36 | 1554834972.5376 | 91 | 0.20673 | 0.002677 | 3.8e-05 | 98.9856 | 95.1144 | 114.8708 | 20.4954 | 0.003277 | 0.002883 | 259.0253 | 0.26116 | 8.3992 | 0.003017 | 0.004835 | 0.004812 | 5.5e-05 | 0.000984 | 0.27367 | 1.3258 | 2.5474 | 98.55 | 20.07 | 49.95 | 99.9363 | 93.6614 | 90.431 | 0.011914 | 0.97102 | 0.022652 | 23.3281 | 0.35684 | 0.15536 | 0.013939 | 0.94232 | 0.94232 | 0.006644
2019-04-09 14:36 | 1554834973.5372 | 92 | 0.20679 | 0.002656 | 3.8e-05 | 98.9106 | 94.9062 | 114.7758 | 20.4955 | 0.003265 | 0.002861 | 259.0221 | 0.26136 | 8.4022 | 0.003036 | 0.004792 | 0.00476 | 5.3e-05 | 0.000987 | 0.27184 | 1.3569 | 2.5476 | 98.55 | 20.07 | 49.949 | 99.6453 | 94.1208 | 89.9123 | 0.011914 | 0.97241 | 0.022659 | 23.3018 | 0.36965 | 0.15548 | 0.013596 | 0.91943 | 0.91943 | 0.006815
2019-04-09 14:36 | 1554834974.5363 | 93 | 0.20675 | 0.002643 | 3.7e-05 | 98.9115 | 94.9279 | 114.5713 | 20.4958 | 0.003236 | 0.002778 | 258.968 | 0.261 | 8.4027 | 0.003001 | 0.004175 | 0.00411 | 5.5e-05 | 0.000983 | 0.24873 | 1.4034 | 2.5475 | 98.55 | 20.07 | 49.947 | 98.9546 | 93.2738 | 87.8954 | 0.011914 | 0.97784 | 0.022702 | 23.2163 | 0.36769 | 0.15527 | 0.013869 | 0.93973 | 0.93973 | 0.006659
2019-04-09 14:36 | 1554834975.5373 | 94 | 0.20681 | 0.002623 | 3.6e-05 | 98.8972 | 94.9764 | 114.5413 | 20.4956 | 0.003265 | 0.002837 | 258.9268 | 0.26098 | 8.4029 | 0.003012 | 0.004245 | 0.004191 | 5.4e-05 | 0.00099 | 0.26106 | 1.3923 | 2.5476 | 98.55 | 20.07 | 49.946 | 99.6408 | 93.5428 | 89.3332 | 0.011913 | 0.97396 | 0.022654 | 23.2595 | 0.36183 | 0.15526 | 0.013518 | 0.91405 | 0.91405 | 0.006845
2019-04-09 14:36 | 1554834976.5354 | 95 | 0.20685 | 0.002609 | 3.5e-05 | 98.9377 | 95.177 | 114.2786 | 20.497 | 0.003261 | 0.002817 | 258.8762 | 0.26151 | 8.404 | 0.00306 | 0.003474 | 0.003378 | 6.1e-05 | 0.000988 | 0.27007 | 1.3846 | 2.5476 | 98.55 | 20.07 | 49.946 | 99.5636 | 94.7094 | 88.8472 | 0.011913 | 0.97527 | 0.022643 | 23.2172 | 0.34661 | 0.15557 | 0.013309 | 0.89947 | 0.89947 | 0.00697
2019-04-09 14:36 | 1554834977.5354 | 96 | 0.20682 | 0.002603 | 3.6e-05 | 98.8671 | 95.2032 | 114.2303 | 20.4974 | 0.003212 | 0.002842 | 258.8173 | 0.26123 | 8.4007 | 0.002972 | 0.004974 | 0.004931 | 5.6e-05 | 0.00099 | 0.26753 | 1.3861 | 2.5477 | 98.55 | 20.07 | 49.945 | 98.3672 | 92.5885 | 89.4546 | 0.011913 | 0.97364 | 0.022619 | 23.2317 | 0.33777 | 0.15541 | 0.013479 | 0.91002 | 0.91002 | 0.006882
2019-04-09 14:36 | 1554834978.5354 | 97 | 0.20683 | 0.002592 | 3.7e-05 | 98.8963 | 95.2104 | 114.3195 | 20.496 | 0.003202 | 0.002859 | 258.8236 | 0.26141 | 8.401 | 0.002965 | 0.004512 | 0.004505 | 5.7e-05 | 0.000986 | 0.26842 | 1.3411 | 2.5476 | 98.55 | 20.07 | 49.944 | 98.1252 | 92.4126 | 89.8596 | 0.011913 | 0.97255 | 0.022615 | 23.2538 | 0.33973 | 0.15551 | 0.013417 | 0.90557 | 0.90557 | 0.006921
2019-04-09 14:36 | 1554834979.5354 | 98 | 0.2068 | 0.002574 | 3.8e-05 | 98.8263 | 95.2517 | 114.1047 | 20.4967 | 0.003174 | 0.002708 | 258.7357 | 0.26104 | 8.4008 | 0.002873 | 0.004626 | 0.004629 | 5.7e-05 | 0.00099 | 0.25608 | 1.3065 | 2.5476 | 98.55 | 20.07 | 49.942 | 97.4638 | 90.1854 | 86.2178 | 0.011912 | 0.98241 | 0.022708 | 23.1151 | 0.32952 | 0.15529 | 0.013591 | 0.92105 | 0.92105 | 0.006795
2019-04-09 14:36 | 1554834980.5371 | 99 | 0.20682 | 0.002565 | 3.9e-05 | 98.885 | 95.1633 | 114.0216 | 20.4969 | 0.003173 | 0.002797 | 258.745 | 0.26105 | 8.403 | 0.002887 | 0.004616 | 0.004555 | 5.6e-05 | 0.000993 | 0.2636 | 1.3132 | 2.5474 | 98.55 | 20.07 | 49.942 | 97.42 | 90.5199 | 88.3738 | 0.011912 | 0.97655 | 0.022632 | 23.1759 | 0.34314 | 0.1553 | 0.013519 | 0.91304 | 0.91304 | 0.006855
2019-04-09 14:36 | 1554834981.5355 | 100 | 0.20684 | 0.002558 | 3.9e-05 | 98.8835 | 95.2679 | 114.1221 | 20.4969 | 0.003173 | 0.002807 | 258.6546 | 0.26063 | 8.4008 | 0.002938 | 0.004074 | 0.004037 | 5.6e-05 | 0.000993 | 0.26488 | 1.3197 | 2.5476 | 98.55 | 20.07 | 49.945 | 97.4183 | 91.7668 | 88.5948 | 0.011913 | 0.97595 | 0.022635 | 23.1932 | 0.33318 | 0.15505 | 0.013425 | 0.90681 | 0.90681 | 0.00689
2019-04-09 14:36 | 1554834982.5359 | 101 | 0.20685 | 0.002542 | 4e-05 | 98.9767 | 95.2872 | 114.4131 | 20.4978 | 0.003186 | 0.0028 | 258.6446 | 0.26108 | 8.3979 | 0.002825 | 0.004088 | 0.004015 | 5.6e-05 | 0.000989 | 0.28296 | 1.3376 | 2.5476 | 98.55 | 20.07 | 49.946 | 97.7554 | 89.0331 | 88.4442 | 0.011913 | 0.97636 | 0.022669 | 23.2179 | 0.3398 | 0.15532 | 0.013387 | 0.9055 | 0.9055 | 0.006913
2019-04-09 14:36 | 1554834983.5365 | 102 | 0.20687 | 0.002542 | 4e-05 | 98.9401 | 95.4434 | 114.2685 | 20.4976 | 0.003217 | 0.002792 | 258.5511 | 0.26171 | 8.4007 | 0.002927 | 0.004683 | 0.004558 | 5.5e-05 | 0.000985 | 0.26985 | 1.3261 | 2.5476 | 98.55 | 20.07 | 49.947 | 98.4914 | 91.4885 | 88.2363 | 0.011914 | 0.97692 | 0.022661 | 23.1966 | 0.32183 | 0.15569 | 0.013228 | 0.89443 | 0.89443 | 0.007015
2019-04-09 14:36 | 1554834984.536 | 103 | 0.20684 | 0.002536 | 4e-05 | 98.949 | 95.3894 | 114.3495 | 20.4979 | 0.003158 | 0.002753 | 258.5355 | 0.26126 | 8.4024 | 0.002863 | 0.004631 | 0.004656 | 5.5e-05 | 0.000985 | 0.28943 | 1.3517 | 2.5475 | 98.55 | 20.07 | 49.949 | 97.0634 | 89.9637 | 87.3 | 0.011914 | 0.97946 | 0.022699 | 23.1747 | 0.3277 | 0.15543 | 0.013438 | 0.91018 | 0.91018 | 0.006882
2019-04-09 14:36 | 1554834985.5367 | 104 | 0.20685 | 0.002533 | 4e-05 | 98.9082 | 95.409 | 114.5716 | 20.4975 | 0.003136 | 0.002784 | 258.5006 | 0.26087 | 8.4001 | 0.002876 | 0.004598 | 0.004551 | 5.9e-05 | 0.000986 | 0.24082 | 1.3276 | 2.5475 | 98.55 | 20.07 | 49.949 | 96.5321 | 90.2666 | 88.0511 | 0.011914 | 0.97742 | 0.022697 | 23.2214 | 0.32217 | 0.15519 | 0.01339 | 0.9068 | 0.9068 | 0.006897
2019-04-09 14:36 | 1554834986.5356 | 105 | 0.20689 | 0.002526 | 4.2e-05 | 98.9978 | 95.6809 | 114.722 | 20.4985 | 0.003177 | 0.002794 | 258.4644 | 0.26137 | 8.4023 | 0.002928 | 0.004338 | 0.004249 | 5.6e-05 | 0.000988 | 0.27776 | 1.3228 | 2.5475 | 98.55 | 20.07 | 49.954 | 97.5179 | 91.512 | 88.2995 | 0.011915 | 0.97675 | 0.022704 | 23.2446 | 0.30482 | 0.15549 | 0.01314 | 0.89006 | 0.89006 | 0.00704
2019-04-09 14:36 | 1554834987.5367 | 106 | 0.20687 | 0.002524 | 4.3e-05 | 98.9946 | 95.7267 | 114.6009 | 20.4985 | 0.003111 | 0.002723 | 258.4544 | 0.26122 | 8.3993 | 0.002837 | 0.003988 | 0.003982 | 5.6e-05 | 0.000994 | 0.28711 | 1.3361 | 2.5474 | 98.55 | 20.07 | 49.953 | 95.9385 | 89.3314 | 86.5855 | 0.011915 | 0.9814 | 0.022746 | 23.1772 | 0.30025 | 0.1554 | 0.013251 | 0.89917 | 0.89917 | 0.006965
2019-04-09 14:36 | 1554834988.5361 | 107 | 0.20684 | 0.002515 | 4.3e-05 | 98.9451 | 95.7408 | 114.541 | 20.4984 | 0.003155 | 0.002744 | 258.3779 | 0.26165 | 8.4018 | 0.002819 | 0.00428 | 0.004173 | 5.7e-05 | 0.001003 | 0.28672 | 1.3412 | 2.5475 | 98.55 | 20.07 | 49.953 | 96.9888 | 88.8852 | 87.0897 | 0.011915 | 0.98003 | 0.022724 | 23.1873 | 0.29446 | 0.15565 | 0.013433 | 0.91065 | 0.91065 | 0.006888
2019-04-09 14:36 | 1554834989.5368 | 108 | 0.20688 | 0.002507 | 4.3e-05 | 99.0017 | 95.759 | 114.6525 | 20.4978 | 0.003112 | 0.002767 | 258.353 | 0.26054 | 8.4022 | 0.002882 | 0.004431 | 0.004399 | 5.6e-05 | 0.001007 | 0.27743 | 1.3152 | 2.5475 | 98.55 | 20.07 | 49.957 | 95.959 | 90.4201 | 87.6523 | 0.011916 | 0.9785 | 0.022718 | 23.2167 | 0.29787 | 0.15499 | 0.013229 | 0.89658 | 0.89658 | 0.006967
2019-04-09 14:36 | 1554834990.5362 | 109 | 0.20694 | 0.002503 | 4.2e-05 | 99.0791 | 95.8361 | 114.5497 | 20.4981 | 0.003107 | 0.002764 | 258.2684 | 0.26126 | 8.4025 | 0.002861 | 0.004235 | 0.0042 | 5.6e-05 | 0.001008 | 0.26272 | 1.3324 | 2.5475 | 98.55 | 20.07 | 49.96 | 95.831 | 89.8992 | 87.5675 | 0.011917 | 0.97873 | 0.02271 | 23.2036 | 0.29767 | 0.15543 | 0.012886 | 0.87303 | 0.87303 | 0.007175
2019-04-09 14:36 | 1554834991.5368 | 110 | 0.20695 | 0.002498 | 4.2e-05 | 98.9424 | 95.8702 | 114.6957 | 20.4988 | 0.003125 | 0.002768 | 258.2704 | 0.26019 | 8.4026 | 0.002818 | 0.003878 | 0.003802 | 5.4e-05 | 0.001006 | 0.26792 | 1.3465 | 2.5476 | 98.55 | 20.07 | 49.963 | 96.2772 | 88.8771 | 87.6564 | 0.011917 | 0.97849 | 0.022722 | 23.2212 | 0.28214 | 0.15479 | 0.012831 | 0.8698 | 0.8698 | 0.007172
2019-04-09 14:37 | 1554834992.5353 | 111 | 0.2069 | 0.002497 | 4.1e-05 | 99.0663 | 96.2326 | 114.7192 | 20.4981 | 0.003085 | 0.00275 | 258.1909 | 0.26067 | 8.4035 | 0.002849 | 0.004285 | 0.004258 | 5.3e-05 | 0.001006 | 0.26128 | 1.3232 | 2.5475 | 98.55 | 20.07 | 49.971 | 95.2976 | 89.6256 | 87.2373 | 0.011919 | 0.97963 | 0.022737 | 23.2101 | 0.25959 | 0.15507 | 0.013107 | 0.88919 | 0.88919 | 0.007028
2019-04-09 14:37 | 1554834993.5352 | 112 | 0.20684 | 0.002487 | 4e-05 | 98.9788 | 95.9146 | 114.6859 | 20.4989 | 0.003043 | 0.00277 | 258.194 | 0.26057 | 8.4021 | 0.002793 | 0.004556 | 0.004471 | 5.5e-05 | 0.001005 | 0.28296 | 1.3387 | 2.5476 | 98.55 | 20.07 | 49.974 | 94.298 | 88.2759 | 87.7134 | 0.01192 | 0.97833 | 0.022719 | 23.2221 | 0.28128 | 0.15501 | 0.013493 | 0.91467 | 0.91467 | 0.00683
2019-04-09 14:37 | 1554834994.5359 | 113 | 0.20697 | 0.002487 | 4e-05 | 99.0333 | 96.012 | 114.8621 | 20.499 | 0.003107 | 0.00269 | 258.1166 | 0.26124 | 8.3983 | 0.002789 | 0.003715 | 0.003725 | 5.4e-05 | 0.001007 | 0.27312 | 1.3388 | 2.5475 | 98.55 | 20.07 | 49.981 | 95.8412 | 88.1797 | 85.793 | 0.011922 | 0.98357 | 0.022797 | 23.178 | 0.27713 | 0.15541 | 0.012693 | 0.86333 | 0.86333 | 0.007255
2019-04-09 14:37 | 1554834995.5359 | 114 | 0.20689 | 0.002486 | 4.1e-05 | 99.1 | 96.0426 | 114.872 | 20.4989 | 0.003026 | 0.002781 | 258.0932 | 0.26131 | 8.4039 | 0.002878 | 0.004205 | 0.004034 | 5.4e-05 | 0.000996 | 0.26017 | 1.3283 | 2.5476 | 98.55 | 20.06 | 49.985 | 93.8792 | 90.318 | 87.979 | 0.011915 | 0.97762 | 0.022729 | 23.2495 | 0.2803 | 0.15545 | 0.013196 | 0.89492 | 0.89492 | 0.007
2019-04-09 14:37 | 1554834996.535 | 115 | 0.20687 | 0.002492 | 4.2e-05 | 99.1485 | 96.1964 | 114.7751 | 20.4982 | 0.003095 | 0.002756 | 258.0456 | 0.26147 | 8.4015 | 0.002825 | 0.00392 | 0.003898 | 5.8e-05 | 0.000991 | 0.25514 | 1.3484 | 2.5476 | 98.55 | 20.06 | 49.988 | 95.5554 | 89.028 | 87.3692 | 0.011916 | 0.97927 | 0.022739 | 23.22 | 0.27036 | 0.15555 | 0.013311 | 0.90306 | 0.90306 | 0.006942
2019-04-09 14:37 | 1554834997.536 | 116 | 0.20683 | 0.002493 | 4.2e-05 | 99.0657 | 96.2463 | 114.6246 | 20.4997 | 0.003072 | 0.002744 | 257.9829 | 0.26141 | 8.4031 | 0.002876 | 0.005097 | 0.005089 | 5.3e-05 | 0.000982 | 0.27743 | 1.3333 | 2.5476 | 98.55 | 20.06 | 49.996 | 94.9985 | 90.2608 | 87.0911 | 0.011918 | 0.98003 | 0.022733 | 23.1958 | 0.25825 | 0.15551 | 0.013532 | 0.91778 | 0.91778 | 0.006829
2019-04-09 14:37 | 1554834998.535 | 117 | 0.20679 | 0.002507 | 4.4e-05 | 99.1455 | 96.6576 | 114.5626 | 20.499 | 0.003051 | 0.002691 | 257.9724 | 0.26063 | 8.4006 | 0.002831 | 0.004226 | 0.004153 | 5.3e-05 | 0.000982 | 0.28871 | 1.3971 | 2.5475 | 98.55 | 20.06 | 49.999 | 94.492 | 89.1795 | 85.8188 | 0.011919 | 0.9835 | 0.022767 | 23.1486 | 0.22732 | 0.15505 | 0.013785 | 0.93626 | 0.93626 | 0.006674
2019-04-09 14:37 | 1554834999.5351 | 118 | 0.20676 | 0.002511 | 4.4e-05 | 99.0934 | 96.4036 | 114.9889 | 20.499 | 0.003089 | 0.002736 | 257.928 | 0.26099 | 8.4034 | 0.002873 | 0.004298 | 0.004296 | 5.5e-05 | 0.00098 | 0.28622 | 1.3542 | 2.5475 | 98.55 | 20.06 | 50.008 | 95.3915 | 90.1876 | 86.9022 | 0.011921 | 0.98054 | 0.022775 | 23.2266 | 0.24615 | 0.15526 | 0.013947 | 0.94761 | 0.94761 | 0.006603
2019-04-09 14:37 | 1554835000.5355 | 119 | 0.20679 | 0.002521 | 4.4e-05 | 99.1601 | 96.4793 | 115.014 | 20.4997 | 0.003056 | 0.002709 | 257.9271 | 0.26186 | 8.4034 | 0.002858 | 0.004393 | 0.004375 | 5.5e-05 | 0.000981 | 0.28279 | 1.3313 | 2.5476 | 98.55 | 20.06 | 50.011 | 94.6001 | 89.8406 | 86.259 | 0.011921 | 0.98229 | 0.022797 | 23.2084 | 0.24515 | 0.15578 | 0.013738 | 0.93434 | 0.93434 | 0.006719
2019-04-09 14:37 | 1554835001.5352 | 120 | 0.20678 | 0.002538 | 4.4e-05 | 99.2301 | 96.5922 | 114.7903 | 20.5001 | 0.003072 | 0.00266 | 257.836 | 0.26138 | 8.4005 | 0.002761 | 0.00442 | 0.004385 | 5.6e-05 | 0.000976 | 0.25536 | 1.3134 | 2.5476 | 98.55 | 20.06 | 50.014 | 94.9926 | 87.4921 | 85.066 | 0.011922 | 0.98557 | 0.022813 | 23.1473 | 0.241 | 0.1555 | 0.01381 | 0.93987 | 0.93987 | 0.006667
2019-04-09 14:37 | 1554835002.5352 | 121 | 0.20676 | 0.00255 | 4.4e-05 | 99.2075 | 96.7984 | 114.5713 | 20.4998 | 0.00299 | 0.002717 | 257.8311 | 0.2606 | 8.4036 | 0.002735 | 0.004585 | 0.004448 | 5.6e-05 | 0.000974 | 0.27041 | 1.3244 | 2.5475 | 98.55 | 20.06 | 50.016 | 93.0069 | 86.8858 | 86.444 | 0.011923 | 0.98179 | 0.022748 | 23.1696 | 0.21989 | 0.15503 | 0.013884 | 0.94221 | 0.94221 | 0.006631
2019-04-09 14:37 | 1554835003.5363 | 122 | 0.20681 | 0.002563 | 4.4e-05 | 99.2096 | 96.8221 | 114.6103 | 20.4997 | 0.003122 | 0.00277 | 257.7482 | 0.26023 | 8.404 | 0.002743 | 0.004974 | 0.00493 | 5.5e-05 | 0.000973 | 0.28407 | 1.3175 | 2.5474 | 98.55 | 20.06 | 50.016 | 96.2083 | 87.0747 | 87.7252 | 0.011923 | 0.9783 | 0.022711 | 23.2148 | 0.21788 | 0.15481 | 0.013551 | 0.9181 | 0.9181 | 0.006795
2019-04-09 14:37 | 1554835004.5353 | 123 | 0.20684 | 0.002568 | 4.3e-05 | 99.2069 | 96.9226 | 114.3627 | 20.4998 | 0.00301 | 0.002682 | 257.7487 | 0.26104 | 8.4014 | 0.002752 | 0.004356 | 0.004432 | 5.3e-05 | 0.000966 | 0.27184 | 1.3267 | 2.5475 | 98.55 | 20.06 | 50.016 | 93.4889 | 87.2894 | 85.5955 | 0.011923 | 0.98411 | 0.022754 | 23.1212 | 0.20836 | 0.15529 | 0.013364 | 0.90716 | 0.90716 | 0.006899
2019-04-09 14:37 | 1554835005.5363 | 124 | 0.20685 | 0.002567 | 4.2e-05 | 99.1694 | 97.0286 | 114.3242 | 20.4999 | 0.002993 | 0.002692 | 257.6675 | 0.26102 | 8.4027 | 0.002728 | 0.004528 | 0.004428 | 5.4e-05 | 0.000964 | 0.27383 | 1.3525 | 2.5474 | 98.55 | 20.06 | 50.014 | 93.083 | 86.6985 | 85.8492 | 0.011922 | 0.98342 | 0.022742 | 23.1255 | 0.1952 | 0.15528 | 0.013314 | 0.90332 | 0.90332 | 0.006928
2019-04-09 14:37 | 1554835006.5353 | 125 | 0.20689 | 0.002566 | 4.2e-05 | 99.1595 | 97.1886 | 114.6084 | 20.5003 | 0.003009 | 0.002759 | 257.6437 | 0.26043 | 8.4013 | 0.002791 | 0.004705 | 0.004607 | 5.4e-05 | 0.000963 | 0.26106 | 1.3698 | 2.5477 | 98.55 | 20.06 | 50.013 | 93.476 | 88.2209 | 87.4555 | 0.011922 | 0.97903 | 0.022719 | 23.2059 | 0.17957 | 0.15493 | 0.013083 | 0.88674 | 0.88674 | 0.007041
2019-04-09 14:37 | 1554835007.5353 | 126 | 0.20684 | 0.002562 | 4.2e-05 | 99.234 | 97.1191 | 114.9262 | 20.5008 | 0.003068 | 0.002667 | 257.6028 | 0.26086 | 8.401 | 0.00277 | 0.004136 | 0.004176 | 5.3e-05 | 0.000963 | 0.2625 | 1.429 | 2.5477 | 98.55 | 20.06 | 50.013 | 94.8938 | 87.713 | 85.2377 | 0.011922 | 0.98509 | 0.022821 | 23.1665 | 0.19269 | 0.15519 | 0.013389 | 0.91161 | 0.91161 | 0.00686
2019-04-09 14:37 | 1554835008.5344 | 127 | 0.20689 | 0.002553 | 4.2e-05 | 99.228 | 97.202 | 115.0891 | 20.5003 | 0.003029 | 0.002647 | 257.5407 | 0.26139 | 8.3991 | 0.00271 | 0.004028 | 0.003919 | 5.4e-05 | 0.000964 | 0.25265 | 1.3689 | 2.5475 | 98.55 | 20.06 | 50.013 | 93.9549 | 86.2805 | 84.7508 | 0.011922 | 0.98643 | 0.022853 | 23.1672 | 0.18452 | 0.1555 | 0.013119 | 0.89447 | 0.89447 | 0.007006
2019-04-09 14:37 | 1554835009.5354 | 128 | 0.20691 | 0.002543 | 4.1e-05 | 99.1828 | 97.1401 | 114.8239 | 20.5005 | 0.003048 | 0.00269 | 257.5178 | 0.26126 | 8.4008 | 0.002784 | 0.004653 | 0.004616 | 5.5e-05 | 0.000967 | 0.27649 | 1.3669 | 2.5475 | 98.55 | 20.06 | 50.012 | 94.4098 | 88.0548 | 85.7837 | 0.011922 | 0.98359 | 0.022794 | 23.1738 | 0.18613 | 0.15543 | 0.013015 | 0.88508 | 0.88508 | 0.007077
2019-04-09 14:37 | 1554835010.5347 | 129 | 0.20691 | 0.002535 | 4.1e-05 | 99.2331 | 96.9997 | 114.7394 | 20.5004 | 0.003014 | 0.002652 | 257.4435 | 0.26043 | 8.4033 | 0.002717 | 0.004017 | 0.003974 | 5.2e-05 | 0.000966 | 0.2589 | 1.3271 | 2.5475 | 98.55 | 20.06 | 50.012 | 93.6037 | 86.4478 | 84.8789 | 0.011922 | 0.98608 | 0.022814 | 23.1361 | 0.20361 | 0.15493 | 0.013033 | 0.8871 | 0.8871 | 0.007038
2019-04-09 14:37 | 1554835011.5356 | 130 | 0.20696 | 0.002532 | 4e-05 | 99.3278 | 97.1669 | 114.7227 | 20.5005 | 0.002997 | 0.002631 | 257.4594 | 0.26106 | 8.4009 | 0.002748 | 0.004111 | 0.004032 | 5.3e-05 | 0.000966 | 0.274 | 1.3344 | 2.5476 | 98.55 | 20.06 | 50.012 | 93.1906 | 87.1941 | 84.3804 | 0.011922 | 0.98746 | 0.022828 | 23.1183 | 0.19673 | 0.15531 | 0.012731 | 0.8671 | 0.8671 | 0.007218
2019-04-09 14:37 | 1554835012.5355 | 131 | 0.20695 | 0.002519 | 4.1e-05 | 99.3459 | 97.3166 | 114.9356 | 20.5011 | 0.003006 | 0.002629 | 257.3705 | 0.26082 | 8.4011 | 0.002699 | 0.004239 | 0.00427 | 5.6e-05 | 0.000967 | 0.27024 | 1.3337 | 2.5475 | 98.55 | 20.06 | 50.013 | 93.4025 | 86.0065 | 84.3146 | 0.011922 | 0.98764 | 0.022852 | 23.1376 | 0.1846 | 0.15516 | 0.012766 | 0.87039 | 0.87039 | 0.007184
2019-04-09 14:37 | 1554835013.5345 | 132 | 0.20701 | 0.002511 | 4e-05 | 99.2191 | 97.3659 | 115.2994 | 20.5017 | 0.002998 | 0.002719 | 257.3805 | 0.26079 | 8.4018 | 0.002711 | 0.004236 | 0.004191 | 5.7e-05 | 0.00097 | 0.27079 | 1.3561 | 2.5475 | 98.55 | 20.06 | 50.014 | 93.2193 | 86.2954 | 86.4975 | 0.011922 | 0.98164 | 0.022818 | 23.2448 | 0.16865 | 0.15515 | 0.012452 | 0.84774 | 0.84774 | 0.007375
2019-04-09 14:37 | 1554835014.5356 | 133 | 0.20702 | 0.002503 | 4e-05 | 99.2953 | 97.5338 | 115.4055 | 20.5012 | 0.003005 | 0.002661 | 257.3095 | 0.26082 | 8.401 | 0.002736 | 0.004426 | 0.004349 | 5.7e-05 | 0.000969 | 0.26864 | 1.3241 | 2.5475 | 98.55 | 20.06 | 50.016 | 93.3736 | 86.9106 | 85.1029 | 0.011923 | 0.98546 | 0.022873 | 23.2104 | 0.1601 | 0.15516 | 0.01242 | 0.84758 | 0.84758 | 0.007377
2019-04-09 14:37 | 1554835015.5357 | 134 | 0.20698 | 0.002493 | 4e-05 | 99.3016 | 97.5603 | 115.3105 | 20.5009 | 0.002936 | 0.002664 | 257.316 | 0.26039 | 8.4024 | 0.002715 | 0.004245 | 0.004188 | 5.6e-05 | 0.000966 | 0.27383 | 1.3284 | 2.5474 | 98.55 | 20.06 | 50.017 | 91.7177 | 86.3951 | 85.1588 | 0.011923 | 0.98531 | 0.022862 | 23.2027 | 0.15823 | 0.15491 | 0.012663 | 0.86382 | 0.86382 | 0.007227
2019-04-09 14:37 | 1554835016.536 | 135 | 0.207 | 0.002486 | 4e-05 | 99.3081 | 97.8453 | 115.2112 | 20.5006 | 0.002992 | 0.002649 | 257.2492 | 0.26076 | 8.402 | 0.002724 | 0.00422 | 0.004131 | 5.4e-05 | 0.00097 | 0.27848 | 1.321 | 2.5475 | 98.55 | 20.06 | 50.018 | 93.0555 | 86.6174 | 84.8161 | 0.011923 | 0.98625 | 0.022863 | 23.1816 | 0.13274 | 0.15513 | 0.012537 | 0.85522 | 0.85522 | 0.00731
2019-04-09 14:37 | 1554835017.5346 | 136 | 0.20702 | 0.002472 | 4e-05 | 99.3257 | 97.8439 | 114.964 | 20.5017 | 0.002961 | 0.002675 | 257.2338 | 0.26117 | 8.4018 | 0.002712 | 0.004406 | 0.004356 | 5.8e-05 | 0.000967 | 0.27256 | 1.3199 | 2.5475 | 98.54 | 20.06 | 50.021 | 92.3275 | 86.3264 | 85.4367 | 0.011925 | 0.98455 | 0.022819 | 23.1768 | 0.13445 | 0.15537 | 0.012459 | 0.84819 | 0.84819 | 0.007382
2019-04-09 14:37 | 1554835018.5357 | 137 | 0.20696 | 0.002463 | 4e-05 | 99.3057 | 97.7585 | 114.6201 | 20.5004 | 0.002973 | 0.002633 | 257.198 | 0.26008 | 8.4006 | 0.002669 | 0.004154 | 0.004091 | 5.6e-05 | 0.000967 | 0.26034 | 1.3137 | 2.5475 | 98.54 | 20.06 | 50.023 | 92.6172 | 85.2794 | 84.4193 | 0.011926 | 0.98735 | 0.022817 | 23.1092 | 0.14045 | 0.15472 | 0.012811 | 0.87215 | 0.87215 | 0.007149
2019-04-09 14:37 | 1554835019.535 | 138 | 0.20701 | 0.002459 | 4e-05 | 99.3233 | 97.8621 | 114.7469 | 20.5015 | 0.00291 | 0.002667 | 257.1383 | 0.2613 | 8.4018 | 0.002628 | 0.004032 | 0.004053 | 5.5e-05 | 0.000968 | 0.27057 | 1.3084 | 2.5475 | 98.54 | 20.06 | 50.026 | 91.1014 | 84.3119 | 85.2468 | 0.011926 | 0.98507 | 0.022803 | 23.1487 | 0.13256 | 0.15545 | 0.012528 | 0.85233 | 0.85233 | 0.00735
2019-04-09 14:37 | 1554835020.5357 | 139 | 0.207 | 0.00245 | 4e-05 | 99.3793 | 97.9592 | 115.0647 | 20.5022 | 0.002982 | 0.002657 | 257.1213 | 0.26073 | 8.4021 | 0.002632 | 0.003949 | 0.003887 | 5.6e-05 | 0.000971 | 0.27936 | 1.3376 | 2.5477 | 98.54 | 20.06 | 50.028 | 92.8215 | 84.3977 | 85.0079 | 0.011927 | 0.98573 | 0.022842 | 23.173 | 0.12874 | 0.15511 | 0.012581 | 0.85745 | 0.85745 | 0.00729
2019-04-09 14:37 | 1554835021.5361 | 140 | 0.207 | 0.002453 | 3.9e-05 | 99.3862 | 97.827 | 115.2623 | 20.5022 | 0.002937 | 0.002681 | 257.0457 | 0.26078 | 8.4036 | 0.002688 | 0.004595 | 0.004594 | 5.4e-05 | 0.000969 | 0.27041 | 1.3249 | 2.5476 | 98.54 | 20.06 | 50.031 | 91.7485 | 85.7391 | 85.5877 | 0.011927 | 0.98413 | 0.022843 | 23.2117 | 0.14143 | 0.15514 | 0.012595 | 0.85851 | 0.85851 | 0.007283
2019-04-09 14:37 | 1554835022.5353 | 141 | 0.20703 | 0.002455 | 3.8e-05 | 99.3864 | 98.1175 | 115.1271 | 20.502 | 0.002951 | 0.002685 | 257.0864 | 0.261 | 8.4028 | 0.002713 | 0.004501 | 0.004398 | 6e-05 | 0.000966 | 0.29159 | 1.3598 | 2.5477 | 98.54 | 20.06 | 50.039 | 92.0769 | 86.3572 | 85.6742 | 0.011929 | 0.98389 | 0.022827 | 23.2009 | 0.11494 | 0.15527 | 0.012365 | 0.84221 | 0.84221 | 0.00743
2019-04-09 14:37 | 1554835023.5342 | 142 | 0.20705 | 0.002447 | 3.8e-05 | 99.3987 | 97.996 | 115.0401 | 20.502 | 0.002949 | 0.002629 | 257.0089 | 0.26048 | 8.4025 | 0.002665 | 0.003871 | 0.003789 | 5.8e-05 | 0.000966 | 0.27721 | 1.3362 | 2.5476 | 98.54 | 20.06 | 50.044 | 92.026 | 85.1919 | 84.3227 | 0.011931 | 0.98762 | 0.022862 | 23.1484 | 0.12712 | 0.15496 | 0.012298 | 0.83895 | 0.83895 | 0.007444
2019-04-09 14:37 | 1554835024.5338 | 143 | 0.20709 | 0.00244 | 3.8e-05 | 99.3492 | 98.0817 | 114.9189 | 20.5019 | 0.00294 | 0.002652 | 257.0157 | 0.2601 | 8.4019 | 0.002656 | 0.004532 | 0.004507 | 5.8e-05 | 0.000964 | 0.27671 | 1.3421 | 2.5475 | 98.54 | 20.06 | 50.053 | 91.8109 | 84.9686 | 84.8865 | 0.011933 | 0.98606 | 0.022832 | 23.1544 | 0.11485 | 0.15473 | 0.012068 | 0.82211 | 0.82211 | 0.007585
2019-04-09 14:37 | 1554835025.5343 | 144 | 0.20707 | 0.002437 | 3.8e-05 | 99.3561 | 98.2397 | 114.988 | 20.5019 | 0.002911 | 0.002604 | 256.9206 | 0.26102 | 8.403 | 0.002655 | 0.00444 | 0.004356 | 5.8e-05 | 0.000962 | 0.25929 | 1.3268 | 2.5475 | 98.54 | 20.06 | 50.059 | 91.1189 | 84.9484 | 83.7195 | 0.011934 | 0.98928 | 0.022876 | 23.1236 | 0.10107 | 0.15528 | 0.012203 | 0.833 | 0.833 | 0.007512
2019-04-09 14:37 | 1554835026.5343 | 145 | 0.2071 | 0.002422 | 3.8e-05 | 99.3504 | 98.3026 | 115.1875 | 20.5021 | 0.002962 | 0.002617 | 256.9114 | 0.26115 | 8.4018 | 0.002634 | 0.003459 | 0.003492 | 5.8e-05 | 0.000962 | 0.27345 | 1.3312 | 2.5474 | 98.54 | 20.06 | 50.065 | 92.345 | 84.4428 | 84.0263 | 0.011936 | 0.98843 | 0.022886 | 23.1536 | 0.094832 | 0.15536 | 0.012017 | 0.8206 | 0.8206 | 0.00763
2019-04-09 14:37 | 1554835027.535 | 146 | 0.20707 | 0.002411 | 3.8e-05 | 99.323 | 98.2276 | 115.6438 | 20.5021 | 0.002914 | 0.00256 | 256.8376 | 0.26011 | 8.4002 | 0.002638 | 0.004001 | 0.00391 | 5.5e-05 | 0.000961 | 0.26161 | 1.3436 | 2.5477 | 98.54 | 20.06 | 50.077 | 91.1747 | 84.5544 | 82.6735 | 0.011938 | 0.99219 | 0.022975 | 23.1554 | 0.099194 | 0.15474 | 0.01222 | 0.83773 | 0.83773 | 0.007444
2019-04-09 14:37 | 1554835028.534 | 147 | 0.2071 | 0.002407 | 3.8e-05 | 99.3909 | 98.3415 | 115.6895 | 20.5029 | 0.00298 | 0.002631 | 256.8129 | 0.26016 | 8.4028 | 0.002681 | 0.004312 | 0.004317 | 5.8e-05 | 0.000958 | 0.26537 | 1.3264 | 2.5476 | 98.54 | 20.06 | 50.084 | 92.7799 | 85.5846 | 84.3772 | 0.01194 | 0.98746 | 0.022924 | 23.2154 | 0.094939 | 0.15477 | 0.01204 | 0.82353 | 0.82353 | 0.007574
2019-04-09 14:37 | 1554835029.534 | 148 | 0.2071 | 0.002398 | 3.9e-05 | 99.413 | 98.5212 | 115.602 | 20.5022 | 0.003016 | 0.002621 | 256.7952 | 0.26142 | 8.4027 | 0.002611 | 0.003499 | 0.003494 | 5.7e-05 | 0.000962 | 0.26056 | 1.3343 | 2.5476 | 98.54 | 20.06 | 50.091 | 93.6437 | 83.9007 | 84.1301 | 0.011942 | 0.98815 | 0.022924 | 23.1986 | 0.080597 | 0.15552 | 0.012035 | 0.82315 | 0.82315 | 0.007614
2019-04-09 14:37 | 1554835030.5354 | 149 | 0.20712 | 0.002395 | 3.8e-05 | 99.3576 | 98.6626 | 115.6494 | 20.5024 | 0.002931 | 0.002627 | 256.7353 | 0.26141 | 8.4023 | 0.002602 | 0.004497 | 0.004471 | 5.7e-05 | 0.000964 | 0.27295 | 1.3124 | 2.5476 | 98.54 | 20.06 | 50.104 | 91.6042 | 83.6762 | 84.273 | 0.011945 | 0.98775 | 0.022924 | 23.208 | 0.06278 | 0.15551 | 0.011953 | 0.81753 | 0.81753 | 0.007666
2019-04-09 14:37 | 1554835031.5357 | 150 | 0.20713 | 0.002385 | 3.8e-05 | 99.343 | 98.7057 | 115.8614 | 20.5026 | 0.002919 | 0.002635 | 256.7357 | 0.26072 | 8.4024 | 0.002604 | 0.004829 | 0.004806 | 5.7e-05 | 0.000967 | 0.27776 | 1.354 | 2.5476 | 98.54 | 20.06 | 50.111 | 91.2993 | 83.7373 | 84.4637 | 0.011947 | 0.98723 | 0.022939 | 23.2354 | 0.057567 | 0.1551 | 0.01187 | 0.81244 | 0.81244 | 0.007694
2019-04-09 14:37 | 1554835032.5341 | 151 | 0.20717 | 0.002383 | 3.7e-05 | 99.3519 | 98.7187 | 115.6187 | 20.5037 | 0.002912 | 0.002625 | 256.6304 | 0.26122 | 8.4026 | 0.002627 | 0.003785 | 0.003757 | 5.6e-05 | 0.000966 | 0.28655 | 1.3132 | 2.5475 | 98.54 | 20.06 | 50.124 | 91.1401 | 84.2892 | 84.2399 | 0.01195 | 0.98784 | 0.022922 | 23.2038 | 0.057186 | 0.1554 | 0.011671 | 0.79827 | 0.79827 | 0.007845
2019-04-09 14:37 | 1554835033.5353 | 152 | 0.2072 | 0.002381 | 3.7e-05 | 99.3915 | 98.8027 | 115.3201 | 20.5023 | 0.002965 | 0.002605 | 256.646 | 0.26027 | 8.4012 | 0.002609 | 0.003982 | 0.003969 | 5.7e-05 | 0.000964 | 0.30431 | 1.3478 | 2.5476 | 98.54 | 20.06 | 50.13 | 92.411 | 83.8548 | 83.7425 | 0.011951 | 0.98922 | 0.022908 | 23.1577 | 0.053146 | 0.15483 | 0.011485 | 0.78506 | 0.78506 | 0.007948
2019-04-09 14:37 | 1554835034.5346 | 153 | 0.2072 | 0.002376 | 3.7e-05 | 99.4034 | 98.7962 | 115.7461 | 20.5031 | 0.002917 | 0.002588 | 256.5712 | 0.26095 | 8.4031 | 0.002644 | 0.004877 | 0.004814 | 5.9e-05 | 0.000963 | 0.28639 | 1.2987 | 2.5473 | 98.54 | 20.05 | 50.137 | 91.2562 | 84.6934 | 83.3516 | 0.011945 | 0.9903 | 0.022963 | 23.1877 | 0.05481 | 0.15524 | 0.011476 | 0.78627 | 0.78627 | 0.007957
2019-04-09 14:37 | 1554835035.5347 | 154 | 0.20722 | 0.002365 | 3.7e-05 | 99.3772 | 98.5721 | 116.0006 | 20.5026 | 0.00291 | 0.002624 | 256.5904 | 0.26102 | 8.402 | 0.002589 | 0.003964 | 0.00391 | 5.9e-05 | 0.000958 | 0.25155 | 1.3089 | 2.5475 | 98.54 | 20.05 | 50.149 | 91.0875 | 83.3542 | 84.21 | 0.011948 | 0.98793 | 0.022961 | 23.2411 | 0.072759 | 0.15528 | 0.011338 | 0.77676 | 0.77676 | 0.008056
2019-04-09 14:37 | 1554835036.5343 | 155 | 0.20724 | 0.002356 | 3.7e-05 | 99.2894 | 98.8395 | 115.8402 | 20.5024 | 0.002942 | 0.00261 | 256.5414 | 0.26106 | 8.4014 | 0.002637 | 0.004787 | 0.004661 | 5.8e-05 | 0.000952 | 0.26128 | 1.3523 | 2.5475 | 98.54 | 20.05 | 50.155 | 91.8606 | 84.5119 | 83.8768 | 0.01195 | 0.98885 | 0.022955 | 23.2142 | 0.040616 | 0.1553 | 0.01128 | 0.77258 | 0.77258 | 0.008101
2019-04-09 14:37 | 1554835037.5357 | 156 | 0.20721 | 0.002337 | 3.7e-05 | 99.3108 | 98.7518 | 115.5968 | 20.5033 | 0.002943 | 0.002628 | 256.524 | 0.2599 | 8.4038 | 0.002616 | 0.004059 | 0.003999 | 5.9e-05 | 0.000952 | 0.27256 | 1.3339 | 2.5475 | 98.54 | 20.05 | 50.164 | 91.881 | 84.0132 | 84.3146 | 0.011952 | 0.98764 | 0.022917 | 23.2041 | 0.050493 | 0.15461 | 0.011467 | 0.78407 | 0.78407 | 0.007947
2019-04-09 14:37 | 1554835038.5337 | 157 | 0.20725 | 0.002327 | 3.8e-05 | 99.3159 | 98.853 | 115.9388 | 20.5037 | 0.002869 | 0.002681 | 256.4843 | 0.26033 | 8.4011 | 0.002638 | 0.003643 | 0.003625 | 6e-05 | 0.000951 | 0.27113 | 1.3496 | 2.5476 | 98.54 | 20.05 | 50.168 | 90.1069 | 84.5403 | 85.5729 | 0.011953 | 0.98417 | 0.022911 | 23.2792 | 0.041789 | 0.15487 | 0.011234 | 0.76791 | 0.76791 | 0.008128
2019-04-09 14:37 | 1554835039.534 | 158 | 0.20725 | 0.002314 | 3.8e-05 | 99.3638 | 98.8664 | 115.9203 | 20.5033 | 0.002857 | 0.002615 | 256.4114 | 0.26043 | 8.4026 | 0.002636 | 0.004315 | 0.004198 | 6e-05 | 0.000956 | 0.27776 | 1.3484 | 2.5476 | 98.54 | 20.05 | 50.172 | 89.8103 | 84.487 | 83.9986 | 0.011954 | 0.98851 | 0.022959 | 23.2262 | 0.044891 | 0.15493 | 0.011233 | 0.76945 | 0.76945 | 0.008114
2019-04-09 14:37 | 1554835040.5334 | 159 | 0.20725 | 0.002303 | 3.8e-05 | 99.2986 | 98.9641 | 115.8792 | 20.5043 | 0.002863 | 0.002635 | 256.4134 | 0.26043 | 8.4014 | 0.002632 | 0.00442 | 0.004365 | 5.9e-05 | 0.000961 | 0.28224 | 1.3395 | 2.5476 | 98.54 | 20.05 | 50.177 | 89.9661 | 84.3904 | 84.4676 | 0.011955 | 0.98721 | 0.02294 | 23.2374 | 0.030181 | 0.15493 | 0.011292 | 0.77284 | 0.77284 | 0.008079
2019-04-09 14:37 | 1554835041.5338 | 160 | 0.20728 | 0.00229 | 3.8e-05 | 99.3885 | 98.9232 | 115.5532 | 20.5035 | 0.0029 | 0.00259 | 256.3318 | 0.25998 | 8.4038 | 0.002654 | 0.004245 | 0.004325 | 5.9e-05 | 0.000966 | 0.274 | 1.3433 | 2.5476 | 98.54 | 20.05 | 50.18 | 90.8564 | 84.9187 | 83.394 | 0.011956 | 0.99019 | 0.022942 | 23.1698 | 0.04198 | 0.15466 | 0.011103 | 0.75996 | 0.75996 | 0.008202
2019-04-09 14:37 | 1554835042.5335 | 161 | 0.20728 | 0.002282 | 3.7e-05 | 99.3191 | 98.7803 | 115.544 | 20.5034 | 0.002858 | 0.002627 | 256.352 | 0.26061 | 8.4019 | 0.002597 | 0.003948 | 0.003824 | 6e-05 | 0.000963 | 0.27721 | 1.2822 | 2.5475 | 98.54 | 20.05 | 50.183 | 89.8333 | 83.5531 | 84.2849 | 0.011956 | 0.98772 | 0.022913 | 23.1978 | 0.048657 | 0.15504 | 0.011122 | 0.76033 | 0.76033 | 0.008218
2019-04-09 14:37 | 1554835043.5334 | 162 | 0.20732 | 0.002277 | 3.7e-05 | 99.3111 | 99.023 | 115.9391 | 20.5052 | 0.002887 | 0.002588 | 256.2556 | 0.26107 | 8.3983 | 0.002579 | 0.004023 | 0.003919 | 5.8e-05 | 0.000956 | 0.26449 | 1.3265 | 2.5476 | 98.54 | 20.05 | 50.188 | 90.5415 | 83.1351 | 83.3466 | 0.011957 | 0.99032 | 0.022982 | 23.2069 | 0.02599 | 0.15531 | 0.01086 | 0.74464 | 0.74464 | 0.008405
2019-04-09 14:37 | 1554835044.5345 | 163 | 0.20735 | 0.002265 | 3.7e-05 | 99.2897 | 98.9727 | 115.786 | 20.5044 | 0.002863 | 0.0026 | 256.2621 | 0.26133 | 8.4012 | 0.002627 | 0.004134 | 0.004126 | 5.7e-05 | 0.000954 | 0.25625 | 1.3175 | 2.5475 | 98.54 | 20.05 | 50.189 | 89.9694 | 84.2695 | 83.633 | 0.011958 | 0.98952 | 0.022958 | 23.2009 | 0.028598 | 0.15546 | 0.010733 | 0.73511 | 0.73511 | 0.008523
2019-04-09 14:37 | 1554835045.5345 | 164 | 0.20737 | 0.002258 | 3.6e-05 | 99.2974 | 99.0051 | 115.5541 | 20.5038 | 0.002891 | 0.002625 | 256.1861 | 0.26056 | 8.399 | 0.002575 | 0.003696 | 0.00367 | 5.9e-05 | 0.00095 | 0.29142 | 1.3392 | 2.5476 | 98.54 | 20.05 | 50.192 | 90.6384 | 83.0388 | 84.2252 | 0.011958 | 0.98788 | 0.022916 | 23.1969 | 0.026373 | 0.15501 | 0.010612 | 0.72552 | 0.72552 | 0.00861
2019-04-09 14:37 | 1554835046.535 | 165 | 0.20741 | 0.002244 | 3.7e-05 | 99.3281 | 99.14 | 115.3039 | 20.5043 | 0.002852 | 0.002568 | 256.1575 | 0.26141 | 8.4006 | 0.002591 | 0.004295 | 0.004163 | 5.5e-05 | 0.000955 | 0.26056 | 1.3475 | 2.5475 | 98.54 | 20.04 | 50.192 | 89.6848 | 83.4198 | 82.8738 | 0.011951 | 0.99163 | 0.022934 | 23.1279 | 0.016955 | 0.15551 | 0.010344 | 0.70775 | 0.70775 | 0.008855
2019-04-09 14:37 | 1554835047.5336 | 166 | 0.20736 | 0.002229 | 3.6e-05 | 99.2766 | 98.9779 | 115.3037 | 20.5047 | 0.002862 | 0.002566 | 256.0958 | 0.26086 | 8.399 | 0.002642 | 0.00388 | 0.003893 | 5.5e-05 | 0.00096 | 0.3027 | 1.3242 | 2.5476 | 98.54 | 20.04 | 50.195 | 89.9297 | 84.6351 | 82.8156 | 0.011952 | 0.9918 | 0.022936 | 23.126 | 0.026951 | 0.15519 | 0.010718 | 0.73349 | 0.73349 | 0.008526
2019-04-09 14:37 | 1554835048.535 | 167 | 0.20744 | 0.002216 | 3.6e-05 | 99.2676 | 99.1259 | 115.2342 | 20.5044 | 0.002849 | 0.002606 | 256.0399 | 0.26063 | 8.401 | 0.002627 | 0.00443 | 0.004335 | 5.6e-05 | 0.00096 | 0.27759 | 1.3469 | 2.5476 | 98.54 | 20.04 | 50.196 | 89.6146 | 84.2805 | 83.7779 | 0.011952 | 0.98912 | 0.022898 | 23.1502 | 0.012783 | 0.15505 | 0.010207 | 0.69726 | 0.69726 | 0.008962
2019-04-09 14:37 | 1554835049.5347 | 168 | 0.20735 | 0.002208 | 3.6e-05 | 99.3043 | 99.0071 | 115.2703 | 20.505 | 0.002836 | 0.002557 | 256.039 | 0.26121 | 8.4009 | 0.002623 | 0.004257 | 0.00415 | 5.4e-05 | 0.000958 | 0.274 | 1.3294 | 2.5475 | 98.54 | 20.04 | 50.197 | 89.3221 | 84.1883 | 82.6014 | 0.011952 | 0.99239 | 0.02294 | 23.1157 | 0.026803 | 0.15539 | 0.010792 | 0.73861 | 0.73861 | 0.008478
2019-04-09 14:37 | 1554835050.5341 | 169 | 0.20749 | 0.002194 | 3.6e-05 | 99.2986 | 98.9768 | 114.9427 | 20.5054 | 0.002823 | 0.002577 | 255.96 | 0.26037 | 8.4016 | 0.002647 | 0.003934 | 0.003928 | 5.5e-05 | 0.000958 | 0.27599 | 1.3421 | 2.5475 | 98.54 | 20.04 | 50.199 | 88.9897 | 84.7675 | 83.08 | 0.011953 | 0.99106 | 0.022892 | 23.0983 | 0.02903 | 0.1549 | 0.009977 | 0.68139 | 0.68139 | 0.009161
2019-04-09 14:38 | 1554835051.5343 | 170 | 0.20741 | 0.002189 | 3.5e-05 | 99.3382 | 99.1207 | 115.0121 | 20.5054 | 0.00286 | 0.002648 | 255.9736 | 0.26149 | 8.4008 | 0.002596 | 0.004155 | 0.004103 | 5.5e-05 | 0.00096 | 0.24779 | 1.3421 | 2.5476 | 98.54 | 20.04 | 50.2 | 89.8796 | 83.5286 | 84.794 | 0.011953 | 0.98631 | 0.022844 | 23.1608 | 0.019606 | 0.15556 | 0.010413 | 0.70974 | 0.70974 | 0.008833
2019-04-09 14:38 | 1554835052.5338 | 171 | 0.20746 | 0.002181 | 3.4e-05 | 99.2795 | 99.1231 | 115.5266 | 20.5046 | 0.002873 | 0.002645 | 255.882 | 0.26131 | 8.4023 | 0.002597 | 0.004867 | 0.004756 | 5.3e-05 | 0.000963 | 0.27312 | 1.3602 | 2.5475 | 98.54 | 20.04 | 50.201 | 90.1983 | 83.5574 | 84.7129 | 0.011953 | 0.98654 | 0.022897 | 23.2099 | 0.014105 | 0.15545 | 0.010179 | 0.69547 | 0.69547 | 0.009008
2019-04-09 14:38 | 1554835053.5348 | 172 | 0.20736 | 0.002173 | 3.5e-05 | 99.2757 | 99.1345 | 115.2011 | 20.5061 | 0.002788 | 0.002587 | 255.8756 | 0.26076 | 8.3988 | 0.002616 | 0.003977 | 0.003977 | 5.5e-05 | 0.000965 | 0.29302 | 1.3843 | 2.5477 | 98.54 | 20.04 | 50.203 | 88.1514 | 84.016 | 83.3146 | 0.011954 | 0.99041 | 0.02291 | 23.1319 | 0.012731 | 0.15513 | 0.010798 | 0.73819 | 0.73819 | 0.008469
2019-04-09 14:38 | 1554835054.5328 | 173 | 0.20749 | 0.002171 | 3.4e-05 | 99.242 | 99.2831 | 115.0769 | 20.506 | 0.002846 | 0.002629 | 255.787 | 0.2606 | 8.4004 | 0.002586 | 0.004223 | 0.004198 | 5.3e-05 | 0.000966 | 0.26465 | 1.3792 | 2.5476 | 98.54 | 20.03 | 50.205 | 89.5606 | 83.2967 | 84.342 | 0.011947 | 0.98756 | 0.022865 | 23.1527 | -0.003705 | 0.15503 | 0.009969 | 0.68013 | 0.68013 | 0.009186
2019-04-09 14:38 | 1554835055.5339 | 174 | 0.20754 | 0.002161 | 3.5e-05 | 99.2292 | 98.9645 | 115.1492 | 20.5054 | 0.002905 | 0.00255 | 255.7871 | 0.25966 | 8.401 | 0.002692 | 0.004315 | 0.004243 | 5.5e-05 | 0.000963 | 0.24635 | 1.3511 | 2.5476 | 98.54 | 20.03 | 50.206 | 90.9645 | 85.8336 | 82.4217 | 0.011947 | 0.99289 | 0.022934 | 23.0977 | 0.023897 | 0.15447 | 0.009693 | 0.66319 | 0.66319 | 0.009387
2019-04-09 14:38 | 1554835056.5329 | 175 | 0.2074 | 0.002158 | 3.5e-05 | 99.2748 | 99.0801 | 115.5424 | 20.5057 | 0.002788 | 0.002569 | 255.7453 | 0.26 | 8.4022 | 0.002582 | 0.003927 | 0.00394 | 5.3e-05 | 0.00096 | 0.274 | 1.3385 | 2.5474 | 98.54 | 20.03 | 50.205 | 88.1542 | 83.2068 | 82.8767 | 0.011947 | 0.99163 | 0.022958 | 23.1519 | 0.017558 | 0.15468 | 0.010539 | 0.72197 | 0.72197 | 0.008634
2019-04-09 14:38 | 1554835057.5343 | 176 | 0.20745 | 0.002154 | 3.5e-05 | 99.2533 | 99.1314 | 115.7029 | 20.5061 | 0.002891 | 0.002564 | 255.7316 | 0.26094 | 8.4005 | 0.002575 | 0.004267 | 0.004234 | 5.6e-05 | 0.000962 | 0.27489 | 1.3514 | 2.5476 | 98.54 | 20.03 | 50.205 | 90.6329 | 83.0205 | 82.7717 | 0.011947 | 0.99192 | 0.022977 | 23.1646 | 0.010998 | 0.15523 | 0.010225 | 0.70105 | 0.70105 | 0.008924
2019-04-09 14:38 | 1554835058.534 | 177 | 0.20744 | 0.002146 | 3.5e-05 | 99.298 | 99.1909 | 115.7435 | 20.506 | 0.002816 | 0.002574 | 255.712 | 0.26048 | 8.4016 | 0.002592 | 0.003842 | 0.003757 | 5.3e-05 | 0.000966 | 0.2677 | 1.3642 | 2.5476 | 98.54 | 20.03 | 50.206 | 88.8309 | 83.4479 | 83.0065 | 0.011947 | 0.99126 | 0.022974 | 23.1763 | 0.009652 | 0.15496 | 0.010333 | 0.7083 | 0.7083 | 0.008817
2019-04-09 14:38 | 1554835059.5343 | 178 | 0.20739 | 0.002145 | 3.5e-05 | 99.3135 | 99.1469 | 115.4137 | 20.5051 | 0.002829 | 0.002583 | 255.6535 | 0.26046 | 8.4032 | 0.002583 | 0.004769 | 0.004774 | 5.5e-05 | 0.000975 | 0.26897 | 1.3804 | 2.5477 | 98.54 | 20.03 | 50.205 | 89.1464 | 83.2302 | 83.2249 | 0.011947 | 0.99066 | 0.022934 | 23.1503 | 0.01502 | 0.15495 | 0.010623 | 0.72696 | 0.72696 | 0.00859
2019-04-09 14:38 | 1554835060.5334 | 179 | 0.20743 | 0.00214 | 3.5e-05 | 99.2789 | 99.2233 | 115.5022 | 20.5066 | 0.002828 | 0.002581 | 255.6567 | 0.26166 | 8.4011 | 0.002506 | 0.00412 | 0.004058 | 5.4e-05 | 0.000976 | 0.28727 | 1.3675 | 2.5475 | 98.54 | 20.04 | 50.203 | 89.1291 | 81.384 | 83.1747 | 0.011954 | 0.9908 | 0.022944 | 23.1576 | 0.005018 | 0.15566 | 0.010404 | 0.71231 | 0.71231 | 0.008807
2019-04-09 14:38 | 1554835061.5331 | 180 | 0.20741 | 0.002138 | 3.5e-05 | 99.2209 | 98.9135 | 115.5412 | 20.506 | 0.002828 | 0.002543 | 255.5646 | 0.26046 | 8.4007 | 0.002522 | 0.004355 | 0.004354 | 5.4e-05 | 0.000986 | 0.2995 | 1.3606 | 2.5476 | 98.54 | 20.03 | 50.201 | 89.1201 | 81.768 | 82.2531 | 0.011946 | 0.99337 | 0.022978 | 23.1315 | 0.02775 | 0.15495 | 0.010522 | 0.7214 | 0.7214 | 0.008656
2019-04-09 14:38 | 1554835062.533 | 181 | 0.20741 | 0.002135 | 3.5e-05 | 99.1488 | 99.1809 | 115.5985 | 20.5056 | 0.00287 | 0.00255 | 255.5702 | 0.26152 | 8.4024 | 0.002522 | 0.003857 | 0.003754 | 5.7e-05 | 0.000986 | 0.25426 | 1.3996 | 2.5477 | 98.54 | 20.03 | 50.199 | 90.1213 | 81.7464 | 82.4313 | 0.011945 | 0.99287 | 0.022978 | 23.143 | -0.002897 | 0.15558 | 0.010531 | 0.72207 | 0.72207 | 0.008683
2019-04-09 14:38 | 1554835063.5335 | 182 | 0.20745 | 0.002136 | 3.4e-05 | 99.2459 | 99.1909 | 115.5273 | 20.5065 | 0.002856 | 0.002548 | 255.5057 | 0.26044 | 8.404 | 0.00258 | 0.004269 | 0.004238 | 5.3e-05 | 0.000988 | 0.26626 | 1.347 | 2.5476 | 98.54 | 20.03 | 50.195 | 89.7918 | 83.1575 | 82.3877 | 0.011944 | 0.99299 | 0.022972 | 23.1345 | 0.004957 | 0.15494 | 0.010258 | 0.7032 | 0.7032 | 0.008879
2019-04-09 14:38 | 1554835064.5346 | 183 | 0.2075 | 0.002136 | 3.4e-05 | 99.2024 | 99.1833 | 115.5534 | 20.5075 | 0.002835 | 0.002574 | 255.4862 | 0.26127 | 8.4029 | 0.002542 | 0.004305 | 0.004224 | 5.4e-05 | 0.000992 | 0.25083 | 1.3267 | 2.5475 | 98.54 | 20.03 | 50.193 | 89.3005 | 82.2483 | 83.0125 | 0.011944 | 0.99125 | 0.022955 | 23.1574 | 0.00172 | 0.15543 | 0.009971 | 0.68291 | 0.68291 | 0.009172
2019-04-09 14:38 | 1554835065.5346 | 184 | 0.20744 | 0.002127 | 3.4e-05 | 99.2268 | 99.3637 | 115.5102 | 20.5077 | 0.00282 | 0.002575 | 255.4172 | 0.26044 | 8.402 | 0.002579 | 0.004216 | 0.004105 | 5.3e-05 | 0.00099 | 0.27129 | 1.3331 | 2.5473 | 98.54 | 20.03 | 50.19 | 88.9318 | 83.1337 | 83.037 | 0.011943 | 0.99118 | 0.02295 | 23.1539 | -0.012327 | 0.15494 | 0.010336 | 0.70779 | 0.70779 | 0.008822
2019-04-09 14:38 | 1554835066.5332 | 185 | 0.2075 | 0.00212 | 3.5e-05 | 99.2572 | 99.1272 | 115.4595 | 20.5064 | 0.002836 | 0.002589 | 255.3731 | 0.26059 | 8.4027 | 0.00253 | 0.005143 | 0.005166 | 5.4e-05 | 0.000991 | 0.28528 | 1.3478 | 2.5475 | 98.54 | 20.03 | 50.184 | 89.3116 | 81.946 | 83.368 | 0.011942 | 0.99026 | 0.022934 | 23.1595 | 0.01172 | 0.15503 | 0.009996 | 0.68401 | 0.68401 | 0.009134
2019-04-09 14:38 | 1554835067.5343 | 186 | 0.20749 | 0.002114 | 3.5e-05 | 99.2316 | 99.2243 | 115.3964 | 20.5074 | 0.002779 | 0.00261 | 255.3601 | 0.26022 | 8.4008 | 0.002587 | 0.003899 | 0.003841 | 5.2e-05 | 0.000995 | 0.28622 | 1.3316 | 2.5476 | 98.54 | 20.03 | 50.18 | 87.9419 | 83.3294 | 83.8668 | 0.011941 | 0.98888 | 0.022912 | 23.1694 | 0.000658 | 0.15481 | 0.010058 | 0.68756 | 0.68756 | 0.009074
2019-04-09 14:38 | 1554835068.5347 | 187 | 0.20749 | 0.002101 | 3.5e-05 | 99.2569 | 99.4108 | 115.2619 | 20.5086 | 0.00281 | 0.002515 | 255.3107 | 0.26081 | 8.4023 | 0.002606 | 0.00448 | 0.004459 | 5.5e-05 | 0.000994 | 0.29407 | 1.3749 | 2.5477 | 98.54 | 20.03 | 50.172 | 88.6945 | 83.7796 | 81.5864 | 0.011939 | 0.99523 | 0.022972 | 23.0818 | -0.013859 | 0.15515 | 0.010072 | 0.69036 | 0.69036 | 0.009057
2019-04-09 14:38 | 1554835069.5334 | 188 | 0.20751 | 0.002095 | 3.4e-05 | 99.2203 | 99.2078 | 115.4982 | 20.5071 | 0.002808 | 0.002574 | 255.3019 | 0.2607 | 8.4022 | 0.002595 | 0.004195 | 0.004102 | 5.5e-05 | 0.000996 | 0.27903 | 1.3141 | 2.5476 | 98.54 | 20.03 | 50.166 | 88.6413 | 83.5192 | 83.0037 | 0.011937 | 0.99127 | 0.02295 | 23.1516 | 0.001128 | 0.15509 | 0.009959 | 0.68195 | 0.68195 | 0.009165
2019-04-09 14:38 | 1554835070.5323 | 189 | 0.20756 | 0.002088 | 3.4e-05 | 99.2009 | 99.3516 | 115.6694 | 20.5087 | 0.002795 | 0.002539 | 255.2136 | 0.26074 | 8.3996 | 0.002542 | 0.00423 | 0.004188 | 5.4e-05 | 0.000992 | 0.27489 | 1.2769 | 2.5474 | 98.54 | 20.03 | 50.161 | 88.3282 | 82.237 | 82.1598 | 0.011936 | 0.99363 | 0.022994 | 23.1413 | -0.013578 | 0.15511 | 0.009669 | 0.66336 | 0.66336 | 0.009423
2019-04-09 14:38 | 1554835071.5324 | 190 | 0.20755 | 0.00208 | 3.3e-05 | 99.2033 | 99.3114 | 115.6001 | 20.5081 | 0.002789 | 0.002604 | 255.2313 | 0.26093 | 8.401 | 0.002513 | 0.004052 | 0.004089 | 5.4e-05 | 0.000992 | 0.2792 | 1.3143 | 2.5476 | 98.54 | 20.03 | 50.151 | 88.1854 | 81.5539 | 83.7419 | 0.011934 | 0.98922 | 0.022936 | 23.1858 | -0.009737 | 0.15523 | 0.009774 | 0.66892 | 0.66892 | 0.009352
2019-04-09 14:38 | 1554835072.5324 | 191 | 0.20755 | 0.002072 | 3.2e-05 | 99.189 | 99.288 | 115.4964 | 20.508 | 0.002804 | 0.002543 | 255.1419 | 0.26085 | 8.4019 | 0.002523 | 0.003413 | 0.003345 | 5.1e-05 | 0.00099 | 0.31205 | 1.3559 | 2.5477 | 98.54 | 20.03 | 50.146 | 88.5353 | 81.7745 | 82.2707 | 0.011933 | 0.99332 | 0.022973 | 23.1276 | -0.008918 | 0.15518 | 0.009787 | 0.67095 | 0.67095 | 0.009321
2019-04-09 14:38 | 1554835073.5334 | 192 | 0.20754 | 0.002059 | 3.2e-05 | 99.2426 | 99.2274 | 115.5374 | 20.5085 | 0.002773 | 0.002592 | 255.1531 | 0.2605 | 8.3993 | 0.002493 | 0.004154 | 0.004045 | 5.5e-05 | 0.000988 | 0.28102 | 1.3257 | 2.5475 | 98.54 | 20.03 | 50.135 | 87.8016 | 81.0724 | 83.4503 | 0.01193 | 0.99003 | 0.022939 | 23.17 | 0.001372 | 0.15497 | 0.009851 | 0.6744 | 0.6744 | 0.009261
2019-04-09 14:38 | 1554835074.5329 | 193 | 0.20754 | 0.002049 | 3.2e-05 | 99.2289 | 99.2783 | 115.46 | 20.5092 | 0.002779 | 0.002583 | 255.0947 | 0.26076 | 8.4017 | 0.002603 | 0.003738 | 0.003669 | 5.1e-05 | 0.000993 | 0.30055 | 1.3415 | 2.5475 | 98.54 | 20.03 | 50.13 | 87.9358 | 83.7172 | 83.2163 | 0.011929 | 0.99068 | 0.022939 | 23.1547 | -0.004452 | 0.15513 | 0.009861 | 0.67507 | 0.67507 | 0.009261
2019-04-09 14:38 | 1554835075.5339 | 194 | 0.20755 | 0.002045 | 3.2e-05 | 99.2626 | 99.3678 | 115.6137 | 20.5076 | 0.002835 | 0.002622 | 255.0907 | 0.26153 | 8.402 | 0.002651 | 0.003838 | 0.003843 | 5.3e-05 | 0.000987 | 0.29911 | 1.3074 | 2.5475 | 98.54 | 20.03 | 50.124 | 89.2887 | 84.8663 | 84.1648 | 0.011927 | 0.98805 | 0.022924 | 23.2009 | -0.009478 | 0.15559 | 0.009811 | 0.67119 | 0.67119 | 0.009342
2019-04-09 14:38 | 1554835076.5329 | 195 | 0.20753 | 0.00204 | 3.2e-05 | 99.2295 | 99.29 | 115.8245 | 20.5083 | 0.002831 | 0.002551 | 255.0637 | 0.26085 | 8.4032 | 0.002606 | 0.003891 | 0.003842 | 5.4e-05 | 0.000986 | 0.27776 | 1.3214 | 2.5476 | 98.54 | 20.02 | 50.112 | 89.2049 | 83.7772 | 82.4541 | 0.011917 | 0.99281 | 0.023 | 23.1664 | -0.005453 | 0.15518 | 0.009884 | 0.67845 | 0.67845 | 0.009218
2019-04-09 14:38 | 1554835077.5326 | 196 | 0.20755 | 0.002035 | 3.3e-05 | 99.2447 | 99.3737 | 116.1304 | 20.5091 | 0.002762 | 0.002563 | 254.9889 | 0.26074 | 8.4023 | 0.002585 | 0.003849 | 0.00381 | 4.7e-05 | 0.000982 | 0.27168 | 1.3415 | 2.5477 | 98.54 | 20.02 | 50.108 | 87.5312 | 83.2733 | 82.7436 | 0.011916 | 0.992 | 0.023021 | 23.2064 | -0.011615 | 0.15511 | 0.00978 | 0.67186 | 0.67186 | 0.009304
2019-04-09 14:38 | 1554835078.5326 | 197 | 0.20752 | 0.002031 | 3.5e-05 | 99.2182 | 99.2835 | 116.3092 | 20.5091 | 0.002838 | 0.00255 | 254.9846 | 0.26059 | 8.4043 | 0.002612 | 0.004347 | 0.004311 | 5e-05 | 0.00098 | 0.25465 | 1.3842 | 2.5476 | 98.54 | 20.02 | 50.103 | 89.362 | 83.9254 | 82.4431 | 0.011915 | 0.99284 | 0.023048 | 23.2144 | -0.005884 | 0.15502 | 0.009965 | 0.68528 | 0.68528 | 0.009117
2019-04-09 14:38 | 1554835079.5326 | 198 | 0.20754 | 0.002032 | 3.6e-05 | 99.2849 | 99.1586 | 116.3045 | 20.5091 | 0.00276 | 0.002574 | 254.9123 | 0.26045 | 8.4012 | 0.002513 | 0.004234 | 0.00423 | 5e-05 | 0.000987 | 0.27992 | 1.3793 | 2.5477 | 98.54 | 20.02 | 50.092 | 87.4898 | 81.5378 | 83.004 | 0.011912 | 0.99127 | 0.023029 | 23.2323 | 0.011389 | 0.15494 | 0.009845 | 0.6764 | 0.6764 | 0.009231
2019-04-09 14:38 | 1554835080.5336 | 199 | 0.20753 | 0.002033 | 3.7e-05 | 99.2536 | 99.4064 | 116.1437 | 20.5095 | 0.002799 | 0.002562 | 254.9232 | 0.26141 | 8.4006 | 0.00254 | 0.004706 | 0.004643 | 5.2e-05 | 0.000986 | 0.25818 | 1.3362 | 2.5475 | 98.54 | 20.02 | 50.089 | 88.4212 | 82.1978 | 82.7322 | 0.011912 | 0.99203 | 0.023022 | 23.2074 | -0.013752 | 0.15551 | 0.009918 | 0.68117 | 0.68117 | 0.009201
2019-04-09 14:38 | 1554835081.5327 | 200 | 0.20754 | 0.002031 | 3.6e-05 | 99.2569 | 99.364 | 116.162 | 20.5092 | 0.002755 | 0.002572 | 254.84 | 0.26082 | 8.3996 | 0.002607 | 0.003894 | 0.003867 | 4.9e-05 | 0.000985 | 0.25984 | 1.3738 | 2.5476 | 98.54 | 20.02 | 50.084 | 87.3705 | 83.8163 | 82.9651 | 0.01191 | 0.99138 | 0.023017 | 23.2168 | -0.009647 | 0.15516 | 0.009837 | 0.67543 | 0.67543 | 0.009258
2019-04-09 14:38 | 1554835082.5337 | 201 | 0.20758 | 0.002036 | 3.6e-05 | 99.2224 | 99.4528 | 116.2688 | 20.5095 | 0.0028 | 0.002519 | 254.847 | 0.26122 | 8.398 | 0.002596 | 0.004299 | 0.004201 | 4.9e-05 | 0.000985 | 0.25105 | 1.3421 | 2.5475 | 98.54 | 20.02 | 50.082 | 88.4525 | 83.5504 | 81.6786 | 0.01191 | 0.99498 | 0.023069 | 23.1854 | -0.02075 | 0.1554 | 0.009629 | 0.66267 | 0.66267 | 0.00945
2019-04-09 14:38 | 1554835083.5341 | 202 | 0.20759 | 0.002035 | 3.6e-05 | 99.2066 | 99.4621 | 116.5982 | 20.5088 | 0.002794 | 0.002538 | 254.7502 | 0.26001 | 8.4024 | 0.002609 | 0.004445 | 0.00436 | 4.9e-05 | 0.000976 | 0.25874 | 1.3403 | 2.5477 | 98.54 | 20.02 | 50.081 | 88.3113 | 83.8587 | 82.1496 | 0.01191 | 0.99366 | 0.023086 | 23.2337 | -0.023008 | 0.15468 | 0.009549 | 0.65765 | 0.65765 | 0.009479
2019-04-09 14:38 | 1554835084.5338 | 203 | 0.20763 | 0.002036 | 3.6e-05 | 99.2221 | 99.4387 | 116.6475 | 20.5103 | 0.002793 | 0.002492 | 254.7229 | 0.26094 | 8.4038 | 0.002519 | 0.004541 | 0.004465 | 4.8e-05 | 0.000974 | 0.28191 | 1.3273 | 2.5475 | 98.54 | 20.02 | 50.078 | 88.2689 | 81.6878 | 81.0509 | 0.011909 | 0.99674 | 0.023127 | 23.2026 | -0.019507 | 0.15524 | 0.009343 | 0.6446 | 0.6446 | 0.009705
2019-04-09 14:38 | 1554835085.5318 | 204 | 0.20766 | 0.002027 | 3.7e-05 | 99.2393 | 99.4081 | 116.3444 | 20.5105 | 0.002716 | 0.002545 | 254.6618 | 0.26055 | 8.4011 | 0.002567 | 0.004388 | 0.004239 | 5.5e-05 | 0.000975 | 0.2625 | 1.2846 | 2.5476 | 98.54 | 20.02 | 50.078 | 86.4201 | 82.8541 | 82.3113 | 0.011909 | 0.9932 | 0.023056 | 23.2137 | -0.015195 | 0.155 | 0.009118 | 0.627 | 0.627 | 0.009963
2019-04-09 14:38 | 1554835086.5319 | 205 | 0.20769 | 0.002019 | 3.7e-05 | 99.1908 | 99.3427 | 116.0139 | 20.5102 | 0.002797 | 0.002526 | 254.5987 | 0.26142 | 8.4019 | 0.002528 | 0.004035 | 0.004047 | 4.8e-05 | 0.000971 | 0.26521 | 1.3259 | 2.5476 | 98.54 | 20.02 | 50.078 | 88.3749 | 81.9031 | 81.861 | 0.011909 | 0.99446 | 0.023038 | 23.166 | -0.013685 | 0.15552 | 0.008993 | 0.61794 | 0.61794 | 0.010142
2019-04-09 14:38 | 1554835087.533 | 206 | 0.2077 | 0.002008 | 3.7e-05 | 99.2849 | 99.3379 | 115.7898 | 20.5089 | 0.002772 | 0.002549 | 254.5772 | 0.26127 | 8.4027 | 0.002501 | 0.003675 | 0.00366 | 4.5e-05 | 0.000974 | 0.28567 | 1.3202 | 2.5475 | 98.54 | 20.02 | 50.077 | 87.7748 | 81.2506 | 82.4213 | 0.011909 | 0.9929 | 0.022997 | 23.1618 | -0.00477 | 0.15543 | 0.008932 | 0.61261 | 0.61261 | 0.010225
2019-04-09 14:38 | 1554835088.5329 | 207 | 0.20767 | 0.001999 | 3.8e-05 | 99.1554 | 99.4373 | 115.7459 | 20.5103 | 0.002708 | 0.002517 | 254.4805 | 0.26195 | 8.4025 | 0.002549 | 0.004425 | 0.004376 | 5e-05 | 0.000975 | 0.28495 | 1.3315 | 2.5476 | 98.54 | 20.02 | 50.076 | 86.2442 | 82.4171 | 81.6301 | 0.011909 | 0.99511 | 0.023019 | 23.1317 | -0.025399 | 0.15583 | 0.009101 | 0.62474 | 0.62474 | 0.010052
2019-04-09 14:38 | 1554835089.532 | 208 | 0.20769 | 0.001987 | 3.8e-05 | 99.2265 | 99.341 | 115.9499 | 20.5099 | 0.002817 | 0.002521 | 254.4946 | 0.26153 | 8.4019 | 0.002514 | 0.004114 | 0.0041 | 4.9e-05 | 0.00098 | 0.27544 | 1.3414 | 2.5476 | 98.54 | 20.01 | 50.076 | 88.8653 | 81.58 | 81.7295 | 0.011901 | 0.99483 | 0.023036 | 23.1553 | -0.010309 | 0.15559 | 0.008999 | 0.6182 | 0.6182 | 0.010143
2019-04-09 14:38 | 1554835090.5329 | 209 | 0.2077 | 0.001979 | 3.8e-05 | 99.2247 | 99.5241 | 116.0771 | 20.5112 | 0.00284 | 0.002541 | 254.4061 | 0.2604 | 8.403 | 0.00259 | 0.004369 | 0.004324 | 4.9e-05 | 0.000978 | 0.27168 | 1.3683 | 2.5475 | 98.54 | 20.01 | 50.076 | 89.4172 | 83.3921 | 82.2098 | 0.011901 | 0.99349 | 0.023033 | 23.1837 | -0.02694 | 0.15491 | 0.008945 | 0.61442 | 0.61442 | 0.010161
2019-04-09 14:38 | 1554835091.5319 | 210 | 0.20772 | 0.001972 | 3.8e-05 | 99.2027 | 99.4772 | 116.132 | 20.5103 | 0.002741 | 0.002515 | 254.4032 | 0.26093 | 8.3999 | 0.00249 | 0.004024 | 0.004024 | 5.1e-05 | 0.000978 | 0.30702 | 1.3404 | 2.5477 | 98.54 | 20.01 | 50.076 | 87.0248 | 81.0023 | 81.5849 | 0.011901 | 0.99524 | 0.023058 | 23.1687 | -0.024719 | 0.15523 | 0.00884 | 0.60788 | 0.60788 | 0.010291
2019-04-09 14:38 | 1554835092.533 | 211 | 0.20771 | 0.001959 | 3.8e-05 | 99.2492 | 99.3516 | 115.9959 | 20.511 | 0.002779 | 0.002497 | 254.3243 | 0.26117 | 8.3996 | 0.002539 | 0.003575 | 0.003497 | 5.2e-05 | 0.000975 | 0.27007 | 1.3357 | 2.5475 | 98.54 | 20.02 | 50.074 | 87.9501 | 82.1721 | 81.1721 | 0.011908 | 0.9964 | 0.023058 | 23.1417 | -0.009229 | 0.15537 | 0.008949 | 0.61537 | 0.61537 | 0.010175
2019-04-09 14:38 | 1554835093.5321 | 212 | 0.20776 | 0.001952 | 3.8e-05 | 99.2009 | 99.3472 | 115.9205 | 20.5104 | 0.002756 | 0.002487 | 254.2922 | 0.26097 | 8.4004 | 0.002564 | 0.004127 | 0.004134 | 5.4e-05 | 0.000976 | 0.27848 | 1.3442 | 2.5475 | 98.54 | 20.02 | 50.071 | 87.3943 | 82.7675 | 80.9137 | 0.011907 | 0.99712 | 0.023059 | 23.1257 | -0.013175 | 0.15525 | 0.00865 | 0.59478 | 0.59478 | 0.010519
2019-04-09 14:38 | 1554835094.5322 | 213 | 0.20775 | 0.001946 | 3.8e-05 | 99.2298 | 99.3076 | 116.0858 | 20.5106 | 0.00274 | 0.002472 | 254.2436 | 0.2612 | 8.4026 | 0.002508 | 0.003584 | 0.003515 | 5.1e-05 | 0.000981 | 0.28727 | 1.3319 | 2.5476 | 98.54 | 20.02 | 50.062 | 86.9972 | 81.4143 | 80.5722 | 0.011905 | 0.99809 | 0.023087 | 23.131 | -0.007007 | 0.15539 | 0.008712 | 0.59976 | 0.59976 | 0.010441
2019-04-09 14:38 | 1554835095.5322 | 214 | 0.20776 | 0.001933 | 3.8e-05 | 99.1971 | 99.4969 | 116.4111 | 20.5108 | 0.00277 | 0.002505 | 254.1868 | 0.26052 | 8.4026 | 0.0025 | 0.00379 | 0.003771 | 5.4e-05 | 0.000978 | 0.29015 | 1.3597 | 2.5476 | 98.54 | 20.02 | 50.058 | 87.7335 | 81.2313 | 81.3568 | 0.011904 | 0.99588 | 0.023094 | 23.1891 | -0.026993 | 0.15498 | 0.008651 | 0.59571 | 0.59571 | 0.010485
2019-04-09 14:38 | 1554835096.5332 | 215 | 0.20778 | 0.001919 | 3.8e-05 | 99.2277 | 99.5894 | 116.5003 | 20.5115 | 0.002785 | 0.002497 | 254.1778 | 0.26124 | 8.401 | 0.002481 | 0.004069 | 0.003957 | 5.5e-05 | 0.000969 | 0.27455 | 1.3873 | 2.5475 | 98.54 | 20.02 | 50.054 | 88.0936 | 80.7884 | 81.1696 | 0.011903 | 0.9964 | 0.023108 | 23.1919 | -0.032546 | 0.15541 | 0.008556 | 0.58952 | 0.58952 | 0.010624
2019-04-09 14:38 | 1554835097.5324 | 216 | 0.20779 | 0.001909 | 3.8e-05 | 99.2435 | 99.2632 | 116.4766 | 20.5116 | 0.002776 | 0.002503 | 254.0821 | 0.26074 | 8.4 | 0.002459 | 0.003257 | 0.003313 | 5.1e-05 | 0.000971 | 0.27489 | 1.3319 | 2.5475 | 98.54 | 20.02 | 50.045 | 87.8836 | 80.2551 | 81.3054 | 0.011901 | 0.99602 | 0.023102 | 23.1939 | -0.001773 | 0.15512 | 0.008489 | 0.58478 | 0.58478 | 0.01069
2019-04-09 14:38 | 1554835098.5335 | 217 | 0.20781 | 0.001905 | 3.8e-05 | 99.1473 | 99.2831 | 116.2616 | 20.5126 | 0.002763 | 0.002487 | 254.0994 | 0.26138 | 8.4014 | 0.002522 | 0.00421 | 0.004158 | 5.3e-05 | 0.000974 | 0.28799 | 1.3032 | 2.5474 | 98.54 | 20.02 | 50.04 | 87.5487 | 81.7695 | 80.9127 | 0.0119 | 0.99713 | 0.023093 | 23.1597 | -0.012244 | 0.1555 | 0.008362 | 0.57578 | 0.57578 | 0.010883
2019-04-09 14:38 | 1554835099.5333 | 218 | 0.20783 | 0.001896 | 3.8e-05 | 99.1705 | 99.3447 | 116.1132 | 20.5113 | 0.00275 | 0.002508 | 253.9953 | 0.2618 | 8.4007 | 0.002466 | 0.003214 | 0.003162 | 4.7e-05 | 0.000973 | 0.286 | 1.314 | 2.5473 | 98.54 | 20.02 | 50.035 | 87.2553 | 80.4279 | 81.4267 | 0.011899 | 0.99568 | 0.023062 | 23.1617 | -0.015698 | 0.15574 | 0.008295 | 0.57035 | 0.57035 | 0.011005
2019-04-09 14:38 | 1554835100.5323 | 219 | 0.20781 | 0.001891 | 3.7e-05 | 99.1309 | 99.5117 | 116.2494 | 20.5129 | 0.002751 | 0.002503 | 253.9911 | 0.26102 | 8.3994 | 0.00257 | 0.004212 | 0.004183 | 5.2e-05 | 0.000963 | 0.28512 | 1.2965 | 2.5473 | 98.54 | 20.02 | 50.03 | 87.2781 | 82.9187 | 81.3067 | 0.011898 | 0.99602 | 0.023079 | 23.1714 | -0.034286 | 0.15528 | 0.008381 | 0.57682 | 0.57682 | 0.010849
2019-04-09 14:38 | 1554835101.5327 | 220 | 0.20785 | 0.001886 | 3.7e-05 | 99.1711 | 99.501 | 116.2933 | 20.5119 | 0.002737 | 0.002497 | 253.9098 | 0.26069 | 8.4018 | 0.002511 | 0.004322 | 0.004254 | 5.2e-05 | 0.000968 | 0.30215 | 1.2935 | 2.5474 | 98.54 | 20.02 | 50.024 | 86.9234 | 81.51 | 81.1615 | 0.011896 | 0.99643 | 0.023088 | 23.171 | -0.0297 | 0.15508 | 0.00819 | 0.56387 | 0.56387 | 0.011084
2019-04-09 14:38 | 1554835102.5324 | 221 | 0.20783 | 0.00188 | 3.6e-05 | 99.1533 | 99.4132 | 116.0947 | 20.5124 | 0.002771 | 0.002495 | 253.8988 | 0.26072 | 8.4045 | 0.002519 | 0.0035 | 0.003451 | 5.1e-05 | 0.000962 | 0.27472 | 1.2906 | 2.5474 | 98.54 | 20.02 | 50.02 | 87.7539 | 81.696 | 81.1269 | 0.011895 | 0.99652 | 0.02307 | 23.1501 | -0.02342 | 0.1551 | 0.008293 | 0.57055 | 0.57055 | 0.010955
2019-04-09 14:38 | 1554835103.5325 | 222 | 0.20784 | 0.001874 | 3.7e-05 | 99.1321 | 99.3554 | 116.2536 | 20.5123 | 0.002737 | 0.002516 | 253.8682 | 0.2599 | 8.4022 | 0.002532 | 0.004844 | 0.004812 | 5e-05 | 0.000962 | 0.30254 | 1.3044 | 2.5474 | 98.54 | 20.02 | 50.011 | 86.9394 | 81.9948 | 81.6196 | 0.011893 | 0.99514 | 0.023069 | 23.182 | -0.020123 | 0.15462 | 0.008232 | 0.56629 | 0.56629 | 0.011003
2019-04-09 14:38 | 1554835104.5325 | 223 | 0.20783 | 0.001863 | 3.7e-05 | 99.1563 | 99.3279 | 116.3454 | 20.5127 | 0.002687 | 0.002458 | 253.8269 | 0.2611 | 8.4037 | 0.002537 | 0.003606 | 0.003605 | 5e-05 | 0.000957 | 0.30143 | 1.3305 | 2.5475 | 98.54 | 20.01 | 50.008 | 85.7283 | 82.1168 | 80.2288 | 0.011885 | 0.99906 | 0.023124 | 23.1456 | -0.015468 | 0.15533 | 0.008325 | 0.574 | 0.574 | 0.010906
2019-04-09 14:38 | 1554835105.5316 | 224 | 0.20783 | 0.001864 | 3.8e-05 | 99.2185 | 99.3564 | 116.3482 | 20.5124 | 0.002656 | 0.002522 | 253.8155 | 0.26155 | 8.4008 | 0.002534 | 0.004433 | 0.004254 | 5.5e-05 | 0.000956 | 0.2891 | 1.3416 | 2.7686 | 98.54 | 20.01 | 50.006 | 84.9843 | 82.054 | 81.7742 | 0.011885 | 0.99471 | 0.023074 | 23.1965 | -0.012428 | 0.1556 | 0.00831 | 0.57174 | 0.57174 | 0.010967
2019-04-09 14:38 | 1554835106.5315 | 225 | 0.20782 | 0.001854 | 3.7e-05 | 99.1551 | 99.3454 | 116.2038 | 20.5128 | 0.002738 | 0.002486 | 253.7517 | 0.26072 | 8.4035 | 0.002584 | 0.004481 | 0.004428 | 5.1e-05 | 0.000949 | 0.26808 | 1.3639 | 2.5474 | 98.54 | 20.01 | 50.003 | 86.9625 | 83.2478 | 80.8879 | 0.011884 | 0.9972 | 0.023088 | 23.1531 | -0.017156 | 0.1551 | 0.008394 | 0.57789 | 0.57789 | 0.010816
2019-04-09 14:38 | 1554835107.5319 | 226 | 0.20786 | 0.00185 | 3.7e-05 | 99.119 | 99.2508 | 116.182 | 20.5134 | 0.002665 | 0.002514 | 253.7725 | 0.26119 | 8.4023 | 0.00253 | 0.004127 | 0.004138 | 5e-05 | 0.000947 | 0.28102 | 1.3539 | 2.5475 | 98.54 | 20.01 | 50.001 | 85.1959 | 81.9506 | 81.5778 | 0.011883 | 0.99526 | 0.023064 | 23.1735 | -0.011882 | 0.15538 | 0.008164 | 0.56147 | 0.56147 | 0.011153
2019-04-09 14:38 | 1554835108.532 | 227 | 0.20785 | 0.001846 | 3.7e-05 | 99.1295 | 99.4886 | 116.6301 | 20.5133 | 0.002715 | 0.002535 | 253.6823 | 0.26193 | 8.3996 | 0.002555 | 0.004426 | 0.004348 | 5.2e-05 | 0.000943 | 0.26106 | 1.3404 | 2.5475 | 98.54 | 20.01 | 49.999 | 86.4182 | 82.5697 | 82.0721 | 0.011883 | 0.99387 | 0.023092 | 23.2343 | -0.032348 | 0.15582 | 0.008203 | 0.56482 | 0.56482 | 0.011118
2019-04-09 14:38 | 1554835109.5317 | 228 | 0.20787 | 0.00185 | 3.7e-05 | 99.1872 | 99.4446 | 116.7366 | 20.5138 | 0.002705 | 0.002495 | 253.6956 | 0.26245 | 8.4009 | 0.002628 | 0.004132 | 0.004059 | 5.1e-05 | 0.000941 | 0.27992 | 1.3454 | 2.5475 | 98.54 | 20.01 | 49.997 | 86.1757 | 84.3238 | 81.1181 | 0.011882 | 0.99655 | 0.023134 | 23.2137 | -0.023175 | 0.15613 | 0.008126 | 0.56049 | 0.56049 | 0.011226
2019-04-09 14:38 | 1554835110.5326 | 229 | 0.20789 | 0.001839 | 3.8e-05 | 99.1643 | 99.4518 | 116.5749 | 20.5131 | 0.00267 | 0.002476 | 253.6247 | 0.26139 | 8.4 | 0.002602 | 0.003535 | 0.003525 | 5.5e-05 | 0.000942 | 0.28495 | 1.3015 | 2.5476 | 98.54 | 20.01 | 49.996 | 85.3314 | 83.6819 | 80.6573 | 0.011882 | 0.99785 | 0.023133 | 23.1825 | -0.025893 | 0.1555 | 0.007996 | 0.55145 | 0.55145 | 0.011364
2019-04-09 14:38 | 1554835111.5326 | 230 | 0.20789 | 0.001839 | 3.8e-05 | 99.0958 | 99.3076 | 116.6432 | 20.5137 | 0.002656 | 0.002491 | 253.6217 | 0.26082 | 8.4021 | 0.002544 | 0.004256 | 0.004136 | 5.2e-05 | 0.000949 | 0.26842 | 1.3347 | 2.5474 | 98.54 | 20.01 | 49.992 | 84.9846 | 82.2838 | 81.0096 | 0.011881 | 0.99685 | 0.023128 | 23.2008 | -0.019094 | 0.15516 | 0.008016 | 0.55266 | 0.55266 | 0.011314
2019-04-09 14:39 | 1554835112.5319 | 231 | 0.20786 | 0.001833 | 3.9e-05 | 99.128 | 99.4287 | 116.9554 | 20.5131 | 0.002678 | 0.002444 | 253.5549 | 0.26044 | 8.402 | 0.002625 | 0.003975 | 0.003892 | 5.3e-05 | 0.000951 | 0.28495 | 1.3583 | 2.5474 | 98.54 | 20.01 | 49.992 | 85.5226 | 84.2397 | 79.8885 | 0.011881 | 1 | 0.023196 | 23.1951 | -0.027097 | 0.15493 | 0.008195 | 0.56669 | 0.56669 | 0.011018
2019-04-09 14:39 | 1554835113.5328 | 232 | 0.20789 | 0.001824 | 3.9e-05 | 99.1211 | 99.3981 | 116.774 | 20.5136 | 0.002714 | 0.002474 | 253.4966 | 0.26035 | 8.4024 | 0.002616 | 0.004399 | 0.004382 | 5.1e-05 | 0.000957 | 0.28711 | 1.3341 | 2.5476 | 98.54 | 20.01 | 49.991 | 86.3723 | 84.0344 | 80.6174 | 0.011881 | 0.99796 | 0.023154 | 23.201 | -0.024959 | 0.15488 | 0.008037 | 0.55476 | 0.55476 | 0.011251
2019-04-09 14:39 | 1554835114.5322 | 233 | 0.20791 | 0.00182 | 3.9e-05 | 99.111 | 99.417 | 116.6707 | 20.5142 | 0.002773 | 0.002413 | 253.4532 | 0.26169 | 8.4023 | 0.002559 | 0.004295 | 0.004223 | 5.3e-05 | 0.000962 | 0.28584 | 1.3421 | 2.5476 | 98.54 | 20.01 | 49.987 | 87.7915 | 82.6471 | 79.1458 | 0.01188 | 1.0021 | 0.023192 | 23.1424 | -0.027576 | 0.15568 | 0.007885 | 0.54512 | 0.54512 | 0.011509
2019-04-09 14:39 | 1554835115.5332 | 234 | 0.20796 | 0.001813 | 3.8e-05 | 99.2134 | 99.237 | 116.9188 | 20.5141 | 0.002693 | 0.002495 | 253.3649 | 0.26103 | 8.4035 | 0.002678 | 0.004187 | 0.004131 | 5.3e-05 | 0.000959 | 0.27168 | 1.3046 | 2.5475 | 98.54 | 20.01 | 49.986 | 85.8796 | 85.5087 | 81.1098 | 0.01188 | 0.99657 | 0.023152 | 23.2315 | -0.002127 | 0.15529 | 0.007634 | 0.5268 | 0.5268 | 0.011879
2019-04-09 14:39 | 1554835116.5328 | 235 | 0.20795 | 0.0018 | 3.8e-05 | 99.1345 | 99.3265 | 116.9223 | 20.5143 | 0.002735 | 0.002479 | 253.3579 | 0.26007 | 8.4041 | 0.002586 | 0.003529 | 0.003463 | 4.8e-05 | 0.000965 | 0.28313 | 1.3103 | 2.5476 | 98.54 | 20.01 | 49.983 | 86.8773 | 83.3103 | 80.7296 | 0.011879 | 0.99764 | 0.023165 | 23.2194 | -0.017306 | 0.15471 | 0.007661 | 0.529 | 0.529 | 0.011786
2019-04-09 14:39 | 1554835117.5308 | 236 | 0.20801 | 0.001798 | 3.7e-05 | 99.1363 | 99.3358 | 116.8242 | 20.5137 | 0.002667 | 0.002411 | 253.2767 | 0.26091 | 8.3993 | 0.002579 | 0.004272 | 0.004275 | 5.2e-05 | 0.000964 | 0.26609 | 1.3222 | 2.5475 | 98.54 | 20.01 | 49.982 | 85.2492 | 83.1294 | 79.0891 | 0.011879 | 1.0023 | 0.023209 | 23.1558 | -0.017981 | 0.15521 | 0.007328 | 0.5069 | 0.5069 | 0.01234
2019-04-09 14:39 | 1554835118.5319 | 237 | 0.208 | 0.001777 | 3.7e-05 | 99.1232 | 99.4232 | 116.893 | 20.5143 | 0.00271 | 0.002432 | 253.2733 | 0.26155 | 8.4002 | 0.00258 | 0.003605 | 0.003579 | 5.1e-05 | 0.000971 | 0.27223 | 1.3565 | 2.5476 | 98.54 | 20.01 | 49.98 | 86.2756 | 83.1581 | 79.5974 | 0.011878 | 1.0008 | 0.023199 | 23.1793 | -0.027031 | 0.1556 | 0.007389 | 0.51097 | 0.51097 | 0.012272
2019-04-09 14:39 | 1554835119.5309 | 238 | 0.20802 | 0.001764 | 3.7e-05 | 99.2274 | 99.3365 | 116.8301 | 20.5143 | 0.002751 | 0.002469 | 253.1792 | 0.26165 | 8.399 | 0.002617 | 0.004491 | 0.004362 | 5.1e-05 | 0.000967 | 0.29463 | 1.3729 | 2.5476 | 98.54 | 20 | 49.975 | 87.2701 | 84.0475 | 80.4909 | 0.01187 | 0.99832 | 0.023163 | 23.2024 | -0.009825 | 0.15565 | 0.007323 | 0.50562 | 0.50562 | 0.012406
2019-04-09 14:39 | 1554835120.5309 | 239 | 0.20803 | 0.001742 | 3.6e-05 | 99.1988 | 99.3977 | 116.9704 | 20.5149 | 0.002746 | 0.002465 | 253.1655 | 0.26106 | 8.4006 | 0.002635 | 0.004017 | 0.003957 | 5e-05 | 0.000973 | 0.27887 | 1.4112 | 2.5475 | 98.54 | 20 | 49.973 | 87.1597 | 84.4817 | 80.3946 | 0.011869 | 0.99859 | 0.02318 | 23.2132 | -0.017917 | 0.1553 | 0.007274 | 0.50263 | 0.50263 | 0.012452
2019-04-09 14:39 | 1554835121.532 | 240 | 0.20802 | 0.001734 | 3.7e-05 | 99.1539 | 99.5354 | 117.1406 | 20.5152 | 0.002684 | 0.002465 | 253.1067 | 0.26122 | 8.4004 | 0.002593 | 0.004372 | 0.004429 | 5.2e-05 | 0.000966 | 0.2724 | 1.3836 | 2.5476 | 98.54 | 20 | 49.972 | 85.6574 | 83.4655 | 80.3904 | 0.011869 | 0.9986 | 0.023197 | 23.2299 | -0.034352 | 0.1554 | 0.007362 | 0.50908 | 0.50908 | 0.012302
2019-04-09 14:39 | 1554835122.5314 | 241 | 0.20803 | 0.001723 | 3.7e-05 | 99.2376 | 99.4339 | 117.0054 | 20.5147 | 0.002659 | 0.002445 | 253.0581 | 0.26108 | 8.3998 | 0.002633 | 0.003318 | 0.003212 | 5.5e-05 | 0.000965 | 0.26808 | 1.3538 | 2.5476 | 98.54 | 20 | 49.969 | 85.0536 | 84.4293 | 79.9188 | 0.011868 | 0.99993 | 0.023199 | 23.201 | -0.017678 | 0.15531 | 0.007272 | 0.50286 | 0.50286 | 0.012447
2019-04-09 14:39 | 1554835123.5324 | 242 | 0.20807 | 0.001706 | 3.8e-05 | 99.1685 | 99.5533 | 116.9308 | 20.5149 | 0.002695 | 0.002488 | 253.0631 | 0.26052 | 8.402 | 0.002616 | 0.004182 | 0.004128 | 5.1e-05 | 0.000962 | 0.26521 | 1.3932 | 2.5476 | 98.54 | 20 | 49.968 | 85.9187 | 84.0356 | 80.9477 | 0.011868 | 0.99703 | 0.023158 | 23.2274 | -0.034643 | 0.15498 | 0.007057 | 0.48702 | 0.48702 | 0.012824
2019-04-09 14:39 | 1554835124.5311 | 243 | 0.20804 | 0.001696 | 3.8e-05 | 99.1542 | 99.373 | 117.1322 | 20.5146 | 0.002678 | 0.002468 | 253.0051 | 0.26057 | 8.4 | 0.002598 | 0.004157 | 0.004107 | 5.3e-05 | 0.000968 | 0.27903 | 1.3768 | 2.5477 | 98.54 | 20 | 49.964 | 85.5116 | 83.596 | 80.4669 | 0.011867 | 0.99838 | 0.023194 | 23.2316 | -0.019716 | 0.15501 | 0.007293 | 0.50414 | 0.50414 | 0.012391
2019-04-09 14:39 | 1554835125.5325 | 244 | 0.20804 | 0.001689 | 3.7e-05 | 99.1554 | 99.4074 | 117.0315 | 20.5157 | 0.002681 | 0.002436 | 253.0148 | 0.26156 | 8.4027 | 0.002549 | 0.004212 | 0.004151 | 5e-05 | 0.000967 | 0.2589 | 1.3047 | 2.5475 | 98.54 | 20 | 49.964 | 85.5867 | 82.4168 | 79.711 | 0.011867 | 1.0005 | 0.023209 | 23.1968 | -0.022705 | 0.1556 | 0.0073 | 0.50499 | 0.50499 | 0.012417
2019-04-09 14:39 | 1554835126.5311 | 245 | 0.20808 | 0.001684 | 3.8e-05 | 99.2432 | 99.5808 | 117.1538 | 20.516 | 0.002722 | 0.002401 | 252.9188 | 0.26068 | 8.4014 | 0.002601 | 0.004007 | 0.003964 | 5e-05 | 0.000968 | 0.26913 | 1.3028 | 2.5475 | 98.54 | 20 | 49.963 | 86.5729 | 83.6731 | 78.8537 | 0.011867 | 1.003 | 0.023249 | 23.1807 | -0.030378 | 0.15508 | 0.007059 | 0.48909 | 0.48909 | 0.012778
2019-04-09 14:39 | 1554835127.5311 | 246 | 0.20804 | 0.00168 | 3.8e-05 | 99.211 | 99.4163 | 117.1911 | 20.5158 | 0.002673 | 0.002488 | 252.9297 | 0.26091 | 8.4008 | 0.002633 | 0.00415 | 0.004053 | 5.4e-05 | 0.000965 | 0.27312 | 1.3366 | 2.5475 | 98.54 | 20 | 49.963 | 85.408 | 84.4248 | 80.946 | 0.011867 | 0.99703 | 0.023184 | 23.2532 | -0.018488 | 0.15522 | 0.007279 | 0.50296 | 0.50296 | 0.012437
2019-04-09 14:39 | 1554835128.5311 | 247 | 0.20805 | 0.001671 | 3.8e-05 | 99.1968 | 99.438 | 116.7817 | 20.5154 | 0.002709 | 0.002469 | 252.8759 | 0.2616 | 8.4002 | 0.002633 | 0.003522 | 0.00351 | 5e-05 | 0.000963 | 0.29286 | 1.3463 | 2.5474 | 98.54 | 20 | 49.964 | 86.2694 | 84.4286 | 80.5052 | 0.011867 | 0.99828 | 0.023158 | 23.1981 | -0.021728 | 0.15562 | 0.007274 | 0.50203 | 0.50203 | 0.012493
2019-04-09 14:39 | 1554835129.5322 | 248 | 0.20804 | 0.001669 | 3.8e-05 | 99.2584 | 99.4724 | 116.7451 | 20.5153 | 0.002686 | 0.002445 | 252.8729 | 0.26164 | 8.4001 | 0.002592 | 0.005017 | 0.004894 | 5.1e-05 | 0.000958 | 0.27936 | 1.3193 | 2.5475 | 98.54 | 20 | 49.97 | 85.7136 | 83.4548 | 79.9204 | 0.011869 | 0.99993 | 0.023174 | 23.1753 | -0.019266 | 0.15565 | 0.0073 | 0.50421 | 0.50421 | 0.012441
2019-04-09 14:39 | 1554835130.5317 | 249 | 0.20807 | 0.001671 | 3.8e-05 | 99.2923 | 99.5003 | 116.7256 | 20.5159 | 0.00267 | 0.002458 | 252.8325 | 0.26128 | 8.4013 | 0.002613 | 0.004026 | 0.004018 | 5.1e-05 | 0.000958 | 0.27007 | 1.3422 | 2.5475 | 98.54 | 20 | 49.974 | 85.3302 | 83.9507 | 80.2201 | 0.01187 | 0.99908 | 0.023162 | 23.1832 | -0.018714 | 0.15543 | 0.007102 | 0.4902 | 0.4902 | 0.012778
2019-04-09 14:39 | 1554835131.5313 | 250 | 0.20803 | 0.001666 | 3.8e-05 | 99.1524 | 99.4704 | 116.6357 | 20.516 | 0.002617 | 0.002424 | 252.7851 | 0.2612 | 8.4 | 0.002594 | 0.004799 | 0.00479 | 5e-05 | 0.000952 | 0.27345 | 1.3855 | 2.5476 | 98.54 | 20 | 49.979 | 84.0583 | 83.5038 | 79.4056 | 0.011871 | 1.0014 | 0.02318 | 23.1475 | -0.028639 | 0.15539 | 0.007366 | 0.50886 | 0.50886 | 0.012306
2019-04-09 14:39 | 1554835132.5317 | 251 | 0.20802 | 0.001665 | 3.8e-05 | 99.2712 | 99.5299 | 116.9148 | 20.5155 | 0.002677 | 0.002422 | 252.7895 | 0.26098 | 8.401 | 0.002617 | 0.003884 | 0.003776 | 5e-05 | 0.000952 | 0.29479 | 1.3429 | 2.5476 | 98.54 | 20 | 49.99 | 85.4922 | 84.054 | 79.3669 | 0.011873 | 1.0015 | 0.023209 | 23.1739 | -0.023279 | 0.15526 | 0.007403 | 0.51208 | 0.51208 | 0.012219
2019-04-09 14:39 | 1554835133.5308 | 252 | 0.20803 | 0.001658 | 3.8e-05 | 99.1926 | 99.5808 | 117.0418 | 20.5158 | 0.002659 | 0.002422 | 252.7112 | 0.26072 | 8.4027 | 0.00262 | 0.003578 | 0.003518 | 5e-05 | 0.000943 | 0.28064 | 1.3594 | 2.5476 | 98.54 | 20 | 49.995 | 85.0694 | 84.1238 | 79.3689 | 0.011875 | 1.0015 | 0.023221 | 23.1866 | -0.034941 | 0.1551 | 0.007351 | 0.50875 | 0.50875 | 0.012286
2019-04-09 14:39 | 1554835134.5305 | 253 | 0.20809 | 0.001657 | 3.8e-05 | 99.2137 | 99.309 | 116.7829 | 20.5146 | 0.002687 | 0.002466 | 252.7459 | 0.26147 | 8.4034 | 0.00257 | 0.003624 | 0.003693 | 4.7e-05 | 0.000947 | 0.27417 | 1.3196 | 2.5476 | 98.54 | 20 | 50.007 | 85.7223 | 82.9158 | 80.4147 | 0.011877 | 0.99853 | 0.023161 | 23.1952 | -0.008581 | 0.15555 | 0.007049 | 0.48652 | 0.48652 | 0.012885
2019-04-09 14:39 | 1554835135.5304 | 254 | 0.20809 | 0.001645 | 3.9e-05 | 99.27 | 99.6035 | 116.5057 | 20.5158 | 0.002702 | 0.002406 | 252.6919 | 0.2615 | 8.4056 | 0.002514 | 0.004544 | 0.004458 | 4.8e-05 | 0.000949 | 0.28423 | 1.3392 | 2.5474 | 98.54 | 20 | 50.014 | 86.0973 | 81.5645 | 78.968 | 0.011879 | 1.0026 | 0.023181 | 23.1202 | -0.030002 | 0.15557 | 0.007054 | 0.48727 | 0.48727 | 0.012866
2019-04-09 14:39 | 1554835136.5314 | 255 | 0.20808 | 0.001649 | 3.8e-05 | 99.2438 | 99.4463 | 116.612 | 20.5159 | 0.002663 | 0.002424 | 252.6913 | 0.26086 | 8.4033 | 0.002508 | 0.003648 | 0.003582 | 4.9e-05 | 0.000946 | 0.28102 | 1.3301 | 2.5475 | 98.54 | 20 | 50.021 | 85.1623 | 81.4273 | 79.4204 | 0.011881 | 1.0013 | 0.023177 | 23.1456 | -0.018229 | 0.15519 | 0.007125 | 0.49215 | 0.49215 | 0.012708
2019-04-09 14:39 | 1554835137.5306 | 256 | 0.20811 | 0.001644 | 3.8e-05 | 99.2879 | 99.3334 | 116.7242 | 20.516 | 0.002683 | 0.002393 | 252.6119 | 0.26122 | 8.4031 | 0.00253 | 0.004066 | 0.004047 | 4.9e-05 | 0.00095 | 0.2656 | 1.3421 | 2.5475 | 98.54 | 20 | 50.033 | 85.6495 | 81.957 | 78.6758 | 0.011884 | 1.0035 | 0.023212 | 23.1323 | -0.004099 | 0.1554 | 0.006888 | 0.47641 | 0.47641 | 0.013146
2019-04-09 14:39 | 1554835138.5305 | 257 | 0.20811 | 0.001639 | 3.9e-05 | 99.3019 | 99.4903 | 116.7911 | 20.5172 | 0.002634 | 0.002393 | 252.5703 | 0.26116 | 8.4006 | 0.002559 | 0.00396 | 0.00389 | 5.2e-05 | 0.000944 | 0.2724 | 1.3364 | 2.5473 | 98.54 | 20 | 50.04 | 84.4731 | 82.6539 | 78.6756 | 0.011885 | 1.0035 | 0.023219 | 23.1389 | -0.016958 | 0.15536 | 0.006934 | 0.47972 | 0.47972 | 0.013052
2019-04-09 14:39 | 1554835139.5309 | 258 | 0.20828 | 0.00163 | 3.9e-05 | 99.3373 | 99.534 | 116.7876 | 20.5164 | 0.002626 | 0.002463 | 252.5431 | 0.26121 | 8.4023 | 0.002584 | 0.003977 | 0.003889 | 5e-05 | 0.000948 | 0.2943 | 1.404 | 2.5477 | 98.54 | 20 | 50.052 | 84.2633 | 83.2513 | 80.3603 | 0.011888 | 0.99869 | 0.023163 | 23.1939 | -0.017695 | 0.1554 | 0.005925 | 0.40863 | 0.40863 | 0.015325
2019-04-09 14:39 | 1554835140.531 | 259 | 0.20802 | 0.001626 | 3.9e-05 | 99.3069 | 99.7533 | 117.0789 | 20.517 | 0.002656 | 0.002411 | 252.5076 | 0.26163 | 8.4017 | 0.002558 | 0.004007 | 0.003937 | 5.1e-05 | 0.000949 | 0.2844 | 1.3456 | 2.5476 | 98.54 | 20 | 50.056 | 84.9794 | 82.6428 | 79.0891 | 0.011889 | 1.0023 | 0.023234 | 23.181 | -0.04011 | 0.15564 | 0.007502 | 0.51947 | 0.51947 | 0.012075
2019-04-09 14:39 | 1554835141.5306 | 260 | 0.20827 | 0.001619 | 3.9e-05 | 99.354 | 99.65 | 117.1516 | 20.5162 | 0.002633 | 0.002463 | 252.5097 | 0.26133 | 8.4011 | 0.002514 | 0.004222 | 0.004088 | 5.2e-05 | 0.00095 | 0.27599 | 1.3467 | 2.5476 | 98.54 | 20 | 50.059 | 84.4331 | 81.5757 | 80.3499 | 0.01189 | 0.99871 | 0.0232 | 23.2297 | -0.02661 | 0.15547 | 0.006016 | 0.41557 | 0.41557 | 0.015077
2019-04-09 14:39 | 1554835142.5318 | 261 | 0.20808 | 0.001614 | 4e-05 | 99.3454 | 99.5857 | 117.1225 | 20.517 | 0.002654 | 0.002459 | 252.4419 | 0.26064 | 8.4013 | 0.002605 | 0.004224 | 0.004151 | 5.2e-05 | 0.000953 | 0.27776 | 1.3584 | 2.5474 | 98.54 | 20 | 50.063 | 84.9349 | 83.7544 | 80.2573 | 0.011891 | 0.99898 | 0.0232 | 23.2238 | -0.021609 | 0.15505 | 0.007139 | 0.49347 | 0.49347 | 0.012663
2019-04-09 14:39 | 1554835143.5307 | 262 | 0.20817 | 0.00161 | 4e-05 | 99.3483 | 99.8035 | 117.1871 | 20.5163 | 0.002611 | 0.002433 | 252.4719 | 0.26144 | 8.4022 | 0.002655 | 0.004474 | 0.00446 | 5e-05 | 0.000948 | 0.28208 | 1.4384 | 2.5475 | 98.54 | 20 | 50.064 | 83.9097 | 84.9638 | 79.6359 | 0.011891 | 1.0007 | 0.023227 | 23.2097 | -0.040887 | 0.15553 | 0.006577 | 0.45505 | 0.45505 | 0.013774
2019-04-09 14:39 | 1554835144.5317 | 263 | 0.20821 | 0.001607 | 4e-05 | 99.3608 | 99.6183 | 117.1946 | 20.5165 | 0.002611 | 0.00242 | 252.3962 | 0.2605 | 8.4021 | 0.002563 | 0.003821 | 0.003805 | 4.9e-05 | 0.000956 | 0.28512 | 1.3809 | 2.5475 | 98.54 | 19.99 | 50.067 | 83.9099 | 82.7529 | 79.322 | 0.011884 | 1.0016 | 0.023238 | 23.2001 | -0.023151 | 0.15497 | 0.006393 | 0.44241 | 0.44241 | 0.014117
2019-04-09 14:39 | 1554835145.5307 | 264 | 0.20814 | 0.001602 | 3.9e-05 | 99.3864 | 99.6875 | 117.2136 | 20.5159 | 0.002652 | 0.002463 | 252.3826 | 0.26128 | 8.4003 | 0.002568 | 0.003918 | 0.003867 | 5e-05 | 0.000953 | 0.28766 | 1.3459 | 2.5476 | 98.54 | 19.99 | 50.07 | 84.905 | 82.8688 | 80.3552 | 0.011885 | 0.9987 | 0.023206 | 23.236 | -0.027055 | 0.15544 | 0.006771 | 0.4681 | 0.4681 | 0.013382
2019-04-09 14:39 | 1554835146.5309 | 265 | 0.2081 | 0.001596 | 4e-05 | 99.4415 | 99.7508 | 117.0636 | 20.5169 | 0.002647 | 0.002405 | 252.3324 | 0.26063 | 8.4027 | 0.002695 | 0.004224 | 0.004211 | 5.3e-05 | 0.00095 | 0.26344 | 1.348 | 2.5477 | 98.54 | 19.99 | 50.071 | 84.7676 | 85.9244 | 78.9612 | 0.011885 | 1.0027 | 0.023237 | 23.1753 | -0.027778 | 0.15505 | 0.007014 | 0.48554 | 0.48554 | 0.012869
2019-04-09 14:39 | 1554835147.5308 | 266 | 0.20816 | 0.001592 | 4e-05 | 99.413 | 99.6627 | 117.1338 | 20.5173 | 0.002637 | 0.002409 | 252.3259 | 0.26152 | 8.4023 | 0.00259 | 0.00471 | 0.00453 | 4.9e-05 | 0.000958 | 0.27599 | 1.3391 | 2.5473 | 98.54 | 19.99 | 50.075 | 84.5396 | 83.3988 | 79.0498 | 0.011886 | 1.0024 | 0.023241 | 23.1852 | -0.022447 | 0.15558 | 0.00667 | 0.46176 | 0.46176 | 0.013578
2019-04-09 14:39 | 1554835148.5312 | 267 | 0.20816 | 0.001587 | 4e-05 | 99.4302 | 99.575 | 117.2568 | 20.5175 | 0.002658 | 0.002407 | 252.3224 | 0.26139 | 8.402 | 0.002572 | 0.003512 | 0.003534 | 5.1e-05 | 0.000957 | 0.27599 | 1.3502 | 2.5475 | 98.54 | 19.99 | 50.077 | 85.0453 | 82.977 | 79.0048 | 0.011887 | 1.0025 | 0.023255 | 23.1959 | -0.013013 | 0.1555 | 0.006704 | 0.46441 | 0.46441 | 0.013494
2019-04-09 14:39 | 1554835149.5302 | 268 | 0.20806 | 0.001582 | 4e-05 | 99.4535 | 99.6338 | 117.2969 | 20.5174 | 0.002574 | 0.002483 | 252.2645 | 0.26074 | 8.4015 | 0.002603 | 0.003586 | 0.003635 | 4.8e-05 | 0.000952 | 0.274 | 1.3542 | 2.5477 | 98.54 | 20 | 50.078 | 83.0308 | 83.7242 | 80.833 | 0.011894 | 0.99735 | 0.023198 | 23.26 | -0.016208 | 0.15511 | 0.007323 | 0.50621 | 0.50621 | 0.012349
2019-04-09 14:39 | 1554835150.5309 | 269 | 0.2081 | 0.001583 | 4.1e-05 | 99.4445 | 99.7188 | 117.2152 | 20.5173 | 0.002615 | 0.002413 | 252.2478 | 0.26113 | 8.4044 | 0.002591 | 0.004443 | 0.004359 | 4.7e-05 | 0.000948 | 0.25818 | 1.3222 | 2.5475 | 98.54 | 19.99 | 50.084 | 84.0004 | 83.4216 | 79.1473 | 0.011888 | 1.0021 | 0.023246 | 23.1964 | -0.02464 | 0.15535 | 0.007051 | 0.48831 | 0.48831 | 0.01282
2019-04-09 14:39 | 1554835151.5309 | 270 | 0.208 | 0.00158 | 4.1e-05 | 99.399 | 99.7326 | 117.2704 | 20.5174 | 0.002592 | 0.00243 | 252.1553 | 0.26095 | 8.4 | 0.002579 | 0.004012 | 0.003974 | 5.1e-05 | 0.000952 | 0.27151 | 1.3402 | 2.5473 | 98.54 | 19.99 | 50.086 | 83.4597 | 83.1368 | 79.5692 | 0.011889 | 1.0009 | 0.023237 | 23.2158 | -0.029974 | 0.15524 | 0.007656 | 0.53015 | 0.53015 | 0.011801
2019-04-09 14:39 | 1554835152.5302 | 271 | 0.2082 | 0.001576 | 4.1e-05 | 99.4258 | 99.4807 | 117.2127 | 20.5174 | 0.002595 | 0.002414 | 252.1652 | 0.26108 | 8.4054 | 0.002649 | 0.004742 | 0.004676 | 4.8e-05 | 0.000952 | 0.28888 | 1.3594 | 2.5476 | 98.54 | 20 | 50.09 | 83.5324 | 84.8269 | 79.1722 | 0.011897 | 1.0021 | 0.023245 | 23.197 | -0.00494 | 0.15532 | 0.006483 | 0.44877 | 0.44877 | 0.013948
2019-04-09 14:39 | 1554835153.5313 | 272 | 0.20812 | 0.001578 | 4.1e-05 | 99.4725 | 99.8001 | 117.3368 | 20.5181 | 0.002618 | 0.002375 | 252.087 | 0.26103 | 8.402 | 0.002581 | 0.004564 | 0.004493 | 4.7e-05 | 0.000951 | 0.26017 | 1.3221 | 2.5475 | 98.54 | 19.99 | 50.093 | 84.0682 | 83.1882 | 78.2297 | 0.01189 | 1.0047 | 0.023288 | 23.1782 | -0.029405 | 0.15529 | 0.00692 | 0.48004 | 0.48004 | 0.013036
2019-04-09 14:39 | 1554835154.5313 | 273 | 0.20809 | 0.001578 | 4.1e-05 | 99.4472 | 99.8087 | 117.3857 | 20.5178 | 0.002646 | 0.002393 | 252.0614 | 0.26061 | 8.4008 | 0.00259 | 0.003815 | 0.003883 | 4.9e-05 | 0.000961 | 0.27345 | 1.3444 | 2.5476 | 98.54 | 19.99 | 50.096 | 84.7406 | 83.3954 | 78.6668 | 0.011891 | 1.0035 | 0.023278 | 23.1975 | -0.032452 | 0.15504 | 0.007098 | 0.49221 | 0.49221 | 0.012694
2019-04-09 14:39 | 1554835155.5311 | 274 | 0.20813 | 0.001575 | 4.2e-05 | 99.4127 | 99.7567 | 117.4502 | 20.5172 | 0.002587 | 0.002419 | 252.0135 | 0.26124 | 8.4027 | 0.002576 | 0.003486 | 0.003427 | 5e-05 | 0.000956 | 0.27688 | 1.3394 | 2.5475 | 98.54 | 19.99 | 50.101 | 83.3309 | 83.0611 | 79.2914 | 0.011892 | 1.0017 | 0.023264 | 23.2244 | -0.030901 | 0.15541 | 0.006862 | 0.47543 | 0.47543 | 0.013174
2019-04-09 14:39 | 1554835156.5315 | 275 | 0.2081 | 0.00157 | 4.1e-05 | 99.4579 | 99.6242 | 117.5051 | 20.5178 | 0.002621 | 0.002402 | 251.999 | 0.26141 | 8.3994 | 0.002587 | 0.004242 | 0.004061 | 4.9e-05 | 0.000959 | 0.29231 | 1.3539 | 2.5475 | 98.54 | 19.99 | 50.104 | 84.1396 | 83.3269 | 78.8897 | 0.011893 | 1.0029 | 0.023283 | 23.2166 | -0.014941 | 0.15551 | 0.00707 | 0.49033 | 0.49033 | 0.012781
2019-04-09 14:39 | 1554835157.5302 | 276 | 0.20803 | 0.001572 | 4.1e-05 | 99.5208 | 99.8517 | 117.3859 | 20.518 | 0.00267 | 0.002412 | 251.9727 | 0.26158 | 8.4041 | 0.002581 | 0.003711 | 0.003599 | 5.1e-05 | 0.000947 | 0.28783 | 1.3318 | 2.5474 | 98.54 | 19.99 | 50.108 | 85.3368 | 83.1967 | 79.1275 | 0.011894 | 1.0022 | 0.023263 | 23.2127 | -0.029693 | 0.15562 | 0.007459 | 0.51702 | 0.51702 | 0.01213
2019-04-09 14:39 | 1554835158.5301 | 277 | 0.20816 | 0.001568 | 4.1e-05 | 99.4749 | 99.6617 | 117.4838 | 20.5186 | 0.002642 | 0.002369 | 251.9045 | 0.261 | 8.3992 | 0.002551 | 0.004253 | 0.004293 | 5e-05 | 0.000947 | 0.27417 | 1.37 | 2.5474 | 98.54 | 20 | 50.109 | 84.6649 | 82.4766 | 78.0945 | 0.011902 | 1.0051 | 0.023307 | 23.1883 | -0.016782 | 0.15527 | 0.006682 | 0.46383 | 0.46383 | 0.01349
2019-04-09 14:39 | 1554835159.5302 | 278 | 0.2081 | 0.001562 | 4.1e-05 | 99.4409 | 99.7749 | 117.5239 | 20.518 | 0.002591 | 0.002415 | 251.9042 | 0.26094 | 8.4023 | 0.002554 | 0.003916 | 0.003802 | 4.9e-05 | 0.000944 | 0.28241 | 1.3569 | 2.5475 | 98.54 | 20 | 50.11 | 83.4206 | 82.5298 | 79.2064 | 0.011902 | 1.002 | 0.023274 | 23.2289 | -0.029992 | 0.15523 | 0.007082 | 0.49102 | 0.49102 | 0.012741
2019-04-09 14:39 | 1554835160.5306 | 279 | 0.20816 | 0.001562 | 4e-05 | 99.4663 | 99.8782 | 117.5122 | 20.5198 | 0.002618 | 0.002446 | 251.8185 | 0.26048 | 8.4026 | 0.002551 | 0.004265 | 0.004287 | 5.2e-05 | 0.000946 | 0.27743 | 1.3671 | 2.5475 | 98.54 | 19.99 | 50.11 | 84.0863 | 82.4684 | 79.9387 | 0.011894 | 0.99988 | 0.023249 | 23.2519 | -0.036966 | 0.15496 | 0.006729 | 0.466 | 0.466 | 0.013401
2019-04-09 14:39 | 1554835161.5317 | 280 | 0.20814 | 0.001567 | 4e-05 | 99.4427 | 99.5836 | 117.5516 | 20.519 | 0.002565 | 0.00242 | 251.8522 | 0.26122 | 8.403 | 0.00253 | 0.004062 | 0.003975 | 5e-05 | 0.000948 | 0.27848 | 1.3677 | 2.5476 | 98.54 | 20 | 50.109 | 82.8109 | 81.9592 | 79.3235 | 0.011902 | 1.0016 | 0.023273 | 23.2355 | -0.012661 | 0.1554 | 0.006854 | 0.47514 | 0.47514 | 0.013181
2019-04-09 14:39 | 1554835162.5305 | 281 | 0.20812 | 0.001558 | 4e-05 | 99.4654 | 99.8211 | 117.5152 | 20.5194 | 0.002571 | 0.002401 | 251.7892 | 0.26183 | 8.4016 | 0.002502 | 0.004497 | 0.004489 | 5.3e-05 | 0.000949 | 0.29999 | 1.3375 | 2.5475 | 98.54 | 19.99 | 50.107 | 82.9501 | 81.2787 | 78.8639 | 0.011894 | 1.0029 | 0.023285 | 23.2168 | -0.031929 | 0.15576 | 0.006983 | 0.48441 | 0.48441 | 0.012959
2019-04-09 14:39 | 1554835163.5293 | 282 | 0.20823 | 0.001559 | 4e-05 | 99.4814 | 99.7175 | 117.3084 | 20.5195 | 0.002618 | 0.002402 | 251.7729 | 0.26059 | 8.4015 | 0.002519 | 0.003986 | 0.003878 | 5e-05 | 0.000947 | 0.28296 | 1.3091 | 2.5475 | 98.54 | 19.99 | 50.105 | 84.0797 | 81.7046 | 78.8952 | 0.011893 | 1.0028 | 0.023263 | 23.1973 | -0.021196 | 0.15503 | 0.006301 | 0.43652 | 0.43652 | 0.014312
2019-04-09 14:39 | 1554835164.5303 | 283 | 0.20814 | 0.001553 | 4e-05 | 99.4782 | 99.7402 | 117.4279 | 20.5195 | 0.002624 | 0.002441 | 251.7335 | 0.26097 | 8.4022 | 0.002515 | 0.003704 | 0.003648 | 4.9e-05 | 0.000952 | 0.29031 | 1.3178 | 2.5475 | 98.54 | 19.99 | 50.103 | 84.235 | 81.6032 | 79.8324 | 0.011893 | 1.0002 | 0.023244 | 23.24 | -0.023528 | 0.15525 | 0.006854 | 0.47454 | 0.47454 | 0.013185
2019-04-09 14:39 | 1554835165.5304 | 284 | 0.2082 | 0.001546 | 4.1e-05 | 99.4567 | 100.0189 | 117.3598 | 20.5196 | 0.002605 | 0.002402 | 251.6971 | 0.26069 | 8.4019 | 0.002504 | 0.004028 | 0.004011 | 5e-05 | 0.000944 | 0.27433 | 1.295 | 2.5475 | 98.54 | 19.99 | 50.1 | 83.7733 | 81.3413 | 78.8978 | 0.011892 | 1.0028 | 0.023268 | 23.2025 | -0.05042 | 0.15508 | 0.006473 | 0.44852 | 0.44852 | 0.013934
2019-04-09 14:39 | 1554835166.5308 | 285 | 0.20815 | 0.001539 | 4.1e-05 | 99.5458 | 99.7712 | 117.3359 | 20.5198 | 0.002566 | 0.002408 | 251.684 | 0.26147 | 8.4038 | 0.002517 | 0.004112 | 0.004032 | 5.1e-05 | 0.000948 | 0.2557 | 1.2965 | 2.5475 | 98.54 | 19.99 | 50.099 | 82.8202 | 81.648 | 79.0304 | 0.011892 | 1.0025 | 0.023262 | 23.2045 | -0.020229 | 0.15555 | 0.006788 | 0.47031 | 0.47031 | 0.013329
2019-04-09 14:39 | 1554835167.5295 | 286 | 0.20817 | 0.001541 | 4.1e-05 | 99.5071 | 99.7003 | 117.4535 | 20.5207 | 0.002599 | 0.002433 | 251.6234 | 0.26054 | 8.4028 | 0.002551 | 0.004304 | 0.004162 | 4.9e-05 | 0.000946 | 0.27151 | 1.361 | 2.5476 | 98.54 | 19.99 | 50.096 | 83.6262 | 82.4693 | 79.6284 | 0.011891 | 1.0008 | 0.023253 | 23.2358 | -0.01735 | 0.15499 | 0.006702 | 0.46414 | 0.46414 | 0.013458
2019-04-09 14:39 | 1554835168.5308 | 287 | 0.20821 | 0.001531 | 4.1e-05 | 99.5219 | 99.8345 | 117.3411 | 20.5211 | 0.002625 | 0.002391 | 251.6434 | 0.26165 | 8.4041 | 0.002512 | 0.004648 | 0.004695 | 5.1e-05 | 0.000944 | 0.27936 | 1.3266 | 2.5475 | 98.54 | 19.99 | 50.094 | 84.239 | 81.5408 | 78.6161 | 0.011891 | 1.0036 | 0.023276 | 23.1914 | -0.028045 | 0.15565 | 0.006483 | 0.44933 | 0.44933 | 0.013961
2019-04-09 14:39 | 1554835169.5295 | 288 | 0.20819 | 0.00152 | 4.1e-05 | 99.558 | 99.9763 | 117.4066 | 20.5217 | 0.002623 | 0.002388 | 251.5657 | 0.26091 | 8.4014 | 0.002492 | 0.003868 | 0.003928 | 4.9e-05 | 0.000954 | 0.27345 | 1.339 | 2.5475 | 98.54 | 19.99 | 50.093 | 84.2105 | 81.0619 | 78.549 | 0.01189 | 1.0038 | 0.023284 | 23.1956 | -0.037502 | 0.15522 | 0.006561 | 0.45497 | 0.45497 | 0.013749
2019-04-09 14:39 | 1554835170.5309 | 289 | 0.20818 | 0.00152 | 4.2e-05 | 99.5297 | 99.9381 | 117.4692 | 20.5209 | 0.002649 | 0.002388 | 251.588 | 0.2622 | 8.4013 | 0.002562 | 0.003948 | 0.00381 | 5e-05 | 0.000961 | 0.27312 | 1.3379 | 2.5476 | 98.53 | 19.99 | 50.089 | 84.8278 | 82.7409 | 78.5504 | 0.011891 | 1.0038 | 0.023291 | 23.2019 | -0.036625 | 0.15598 | 0.006659 | 0.46185 | 0.46185 | 0.013611
2019-04-09 14:40 | 1554835171.5296 | 290 | 0.2082 | 0.001518 | 4.1e-05 | 99.5997 | 99.831 | 117.3704 | 20.5206 | 0.002614 | 0.002386 | 251.5214 | 0.26178 | 8.4035 | 0.002479 | 0.004067 | 0.004021 | 5.1e-05 | 0.000962 | 0.28799 | 1.3789 | 2.5476 | 98.53 | 19.99 | 50.088 | 83.9935 | 80.7409 | 78.494 | 0.01189 | 1.004 | 0.023283 | 23.1902 | -0.020754 | 0.15574 | 0.006526 | 0.45243 | 0.45243 | 0.013872
2019-04-09 14:40 | 1554835172.5297 | 291 | 0.20821 | 0.001514 | 4.1e-05 | 99.5193 | 99.755 | 117.1866 | 20.5217 | 0.002628 | 0.002397 | 251.4988 | 0.26096 | 8.3984 | 0.002454 | 0.004034 | 0.00403 | 5.3e-05 | 0.000966 | 0.28136 | 1.3368 | 2.5477 | 98.53 | 19.99 | 50.087 | 84.3332 | 80.1402 | 78.7765 | 0.01189 | 1.0032 | 0.023255 | 23.1814 | -0.02116 | 0.15524 | 0.0065 | 0.4501 | 0.4501 | 0.0139
2019-04-09 14:40 | 1554835173.5296 | 292 | 0.20819 | 0.00151 | 4.1e-05 | 99.5937 | 99.9291 | 117.2692 | 20.5205 | 0.002582 | 0.002368 | 251.4691 | 0.26094 | 8.4022 | 0.002488 | 0.004392 | 0.004352 | 5e-05 | 0.000967 | 0.2808 | 1.3543 | 2.5476 | 98.53 | 19.99 | 50.085 | 83.2079 | 80.9586 | 78.0769 | 0.01189 | 1.0052 | 0.023286 | 23.1665 | -0.030073 | 0.15523 | 0.006585 | 0.45664 | 0.45664 | 0.0137
2019-04-09 14:40 | 1554835174.531 | 293 | 0.20821 | 0.001501 | 4.1e-05 | 99.5514 | 99.9085 | 117.0775 | 20.5212 | 0.002585 | 0.002387 | 251.4488 | 0.26078 | 8.4018 | 0.002495 | 0.004084 | 0.004039 | 5.1e-05 | 0.000957 | 0.27831 | 1.3626 | 2.5475 | 98.53 | 19.99 | 50.083 | 83.2857 | 81.1136 | 78.5266 | 0.011889 | 1.0039 | 0.023253 | 23.1624 | -0.032023 | 0.15514 | 0.006506 | 0.45049 | 0.45049 | 0.013879
2019-04-09 14:40 | 1554835175.5311 | 294 | 0.20821 | 0.001496 | 4.1e-05 | 99.5481 | 99.9185 | 117.4697 | 20.5212 | 0.002546 | 0.002396 | 251.4572 | 0.26128 | 8.404 | 0.002625 | 0.004428 | 0.00434 | 5.1e-05 | 0.000959 | 0.29878 | 1.3338 | 2.5476 | 98.53 | 19.99 | 50.078 | 82.3537 | 84.2514 | 78.7503 | 0.011888 | 1.0033 | 0.023284 | 23.2085 | -0.03321 | 0.15543 | 0.006481 | 0.4494 | 0.4494 | 0.013939
2019-04-09 14:40 | 1554835176.5307 | 295 | 0.20821 | 0.001491 | 4.2e-05 | 99.5604 | 99.765 | 117.7021 | 20.5211 | 0.002588 | 0.00235 | 251.3761 | 0.26074 | 8.4014 | 0.002545 | 0.003845 | 0.003893 | 5.5e-05 | 0.000951 | 0.28208 | 1.3361 | 2.5476 | 98.53 | 19.99 | 50.074 | 83.3534 | 82.3355 | 77.6308 | 0.011887 | 1.0065 | 0.023344 | 23.1945 | -0.018362 | 0.15512 | 0.006486 | 0.45081 | 0.45081 | 0.013866
2019-04-09 14:40 | 1554835177.5289 | 296 | 0.20823 | 0.001484 | 4.2e-05 | 99.5487 | 99.839 | 117.6091 | 20.5215 | 0.002518 | 0.002355 | 251.393 | 0.261 | 8.4026 | 0.002574 | 0.004052 | 0.003914 | 4.9e-05 | 0.000954 | 0.28567 | 1.3228 | 2.5475 | 98.53 | 19.99 | 50.072 | 81.6715 | 83.0114 | 77.7707 | 0.011887 | 1.0061 | 0.02333 | 23.1899 | -0.026037 | 0.15527 | 0.006399 | 0.44452 | 0.44452 | 0.014077
2019-04-09 14:40 | 1554835178.5301 | 297 | 0.20826 | 0.001479 | 4.2e-05 | 99.4853 | 99.6892 | 117.7417 | 20.5213 | 0.002602 | 0.002407 | 251.3195 | 0.26057 | 8.402 | 0.002567 | 0.003711 | 0.00368 | 5.3e-05 | 0.000949 | 0.28368 | 1.3405 | 2.5477 | 98.53 | 19.98 | 50.069 | 83.7001 | 82.8528 | 79.0054 | 0.011879 | 1.0025 | 0.023303 | 23.2438 | -0.018316 | 0.15502 | 0.006234 | 0.4325 | 0.4325 | 0.014444
2019-04-09 14:40 | 1554835179.5302 | 298 | 0.20823 | 0.001479 | 4.2e-05 | 99.5949 | 99.9904 | 117.7063 | 20.5221 | 0.002589 | 0.002395 | 251.3275 | 0.26131 | 8.4016 | 0.002474 | 0.004389 | 0.004393 | 5.1e-05 | 0.000949 | 0.2844 | 1.3342 | 2.5476 | 98.53 | 19.98 | 50.067 | 83.3941 | 80.6162 | 78.713 | 0.011878 | 1.0034 | 0.023309 | 23.2306 | -0.035447 | 0.15545 | 0.006379 | 0.44266 | 0.44266 | 0.014153
2019-04-09 14:40 | 1554835180.5288 | 299 | 0.20827 | 0.001466 | 4.2e-05 | 99.4934 | 99.7687 | 117.3824 | 20.5216 | 0.002572 | 0.002462 | 251.2775 | 0.26083 | 8.4013 | 0.002463 | 0.00411 | 0.004061 | 5.1e-05 | 0.00095 | 0.30143 | 1.3194 | 2.5475 | 98.53 | 19.98 | 50.065 | 82.9755 | 80.3629 | 80.3326 | 0.011878 | 0.99876 | 0.023223 | 23.252 | -0.024723 | 0.15517 | 0.006163 | 0.426 | 0.426 | 0.014679
2019-04-09 14:40 | 1554835181.5299 | 300 | 0.20823 | 0.001473 | 4.2e-05 | 99.5142 | 99.9171 | 117.1998 | 20.5219 | 0.002598 | 0.002364 | 251.2563 | 0.26087 | 8.4006 | 0.002534 | 0.004218 | 0.004229 | 5.4e-05 | 0.000948 | 0.27345 | 1.3291 | 2.7264 | 98.53 | 19.98 | 50.066 | 83.6031 | 82.0594 | 77.9713 | 0.011878 | 1.0055 | 0.023283 | 23.1562 | -0.036138 | 0.15519 | 0.006386 | 0.44263 | 0.44263 | 0.01413
2019-04-09 14:40 | 1554835182.5289 | 301 | 0.20824 | 0.001466 | 4.2e-05 | 99.5613 | 99.8155 | 117.1038 | 20.5218 | 0.002533 | 0.00238 | 251.2446 | 0.26132 | 8.4029 | 0.002541 | 0.004312 | 0.004259 | 5.1e-05 | 0.000951 | 0.27704 | 1.3598 | 2.573 | 98.53 | 19.98 | 50.065 | 82.0266 | 82.234 | 78.3643 | 0.011878 | 1.0044 | 0.02326 | 23.1596 | -0.022817 | 0.15546 | 0.006369 | 0.44102 | 0.44102 | 0.014206
2019-04-09 14:40 | 1554835183.5291 | 302 | 0.20825 | 0.001464 | 4.2e-05 | 99.5892 | 99.6813 | 117.1368 | 20.5228 | 0.002625 | 0.002394 | 251.1835 | 0.26043 | 8.4019 | 0.002589 | 0.004038 | 0.003939 | 4.9e-05 | 0.00095 | 0.27007 | 1.3495 | 2.5475 | 98.53 | 19.98 | 50.063 | 84.2406 | 83.3793 | 78.6949 | 0.011877 | 1.0034 | 0.023253 | 23.1738 | -0.008267 | 0.15493 | 0.00627 | 0.43403 | 0.43403 | 0.014385
2019-04-09 14:40 | 1554835184.5294 | 303 | 0.20825 | 0.001466 | 4.1e-05 | 99.5791 | 99.9443 | 117.5889 | 20.5218 | 0.002517 | 0.002416 | 251.1783 | 0.26156 | 8.4017 | 0.002577 | 0.004866 | 0.004842 | 5.2e-05 | 0.000947 | 0.29856 | 1.3331 | 2.5476 | 98.53 | 19.98 | 50.062 | 81.6624 | 83.0858 | 79.2334 | 0.011877 | 1.0019 | 0.02328 | 23.2362 | -0.032738 | 0.1556 | 0.006292 | 0.43615 | 0.43615 | 0.014377
2019-04-09 14:40 | 1554835185.5294 | 304 | 0.20826 | 0.00147 | 4.2e-05 | 99.5922 | 99.7763 | 117.3608 | 20.5223 | 0.002597 | 0.002391 | 251.0993 | 0.26079 | 8.4007 | 0.002536 | 0.003908 | 0.003893 | 5.1e-05 | 0.000948 | 0.28313 | 1.3091 | 2.5476 | 98.53 | 19.98 | 50.062 | 83.5684 | 82.1223 | 78.6243 | 0.011877 | 1.0036 | 0.023277 | 23.1936 | -0.016518 | 0.15515 | 0.006221 | 0.43108 | 0.43108 | 0.014504
2019-04-09 14:40 | 1554835186.5301 | 305 | 0.20825 | 0.001466 | 4.2e-05 | 99.6119 | 99.9449 | 117.3397 | 20.5233 | 0.002476 | 0.002439 | 251.1088 | 0.26128 | 8.4016 | 0.002484 | 0.004034 | 0.00394 | 5.1e-05 | 0.000961 | 0.2844 | 1.3334 | 2.5476 | 98.53 | 19.98 | 50.062 | 80.6804 | 80.8589 | 79.7753 | 0.011877 | 1.0003 | 0.023237 | 23.2294 | -0.029858 | 0.15544 | 0.00632 | 0.43724 | 0.43724 | 0.014327
2019-04-09 14:40 | 1554835187.5297 | 306 | 0.20825 | 0.00146 | 4.2e-05 | 99.5842 | 99.7316 | 117.2547 | 20.5222 | 0.002573 | 0.002411 | 251.0367 | 0.26085 | 8.4011 | 0.002535 | 0.004359 | 0.004277 | 5e-05 | 0.000965 | 0.29159 | 1.3465 | 2.5474 | 98.53 | 19.98 | 50.062 | 83.0052 | 82.0815 | 79.0997 | 0.011877 | 1.0023 | 0.023251 | 23.1988 | -0.013229 | 0.15518 | 0.006309 | 0.43668 | 0.43668 | 0.014321
2019-04-09 14:40 | 1554835188.5292 | 307 | 0.20825 | 0.001455 | 4.2e-05 | 99.5922 | 99.8792 | 117.0725 | 20.5235 | 0.002494 | 0.002379 | 251.0312 | 0.26131 | 8.4021 | 0.0026 | 0.004123 | 0.004122 | 5.5e-05 | 0.00097 | 0.30093 | 1.3189 | 2.5476 | 98.53 | 19.98 | 50.062 | 81.1119 | 83.6436 | 78.3308 | 0.011877 | 1.0045 | 0.023259 | 23.1554 | -0.025738 | 0.15545 | 0.006335 | 0.4387 | 0.4387 | 0.01428
2019-04-09 14:40 | 1554835189.5302 | 308 | 0.20824 | 0.001465 | 4.2e-05 | 99.5901 | 99.9088 | 117.1545 | 20.5231 | 0.002546 | 0.002365 | 250.9923 | 0.26084 | 8.4028 | 0.0025 | 0.003816 | 0.003792 | 5.1e-05 | 0.000979 | 0.28655 | 1.339 | 2.5476 | 98.53 | 19.98 | 50.061 | 82.3625 | 81.2387 | 77.9975 | 0.011877 | 1.0054 | 0.023278 | 23.1525 | -0.028576 | 0.15517 | 0.006383 | 0.44239 | 0.44239 | 0.014136
2019-04-09 14:40 | 1554835190.5292 | 309 | 0.20825 | 0.001461 | 4.2e-05 | 99.6229 | 100.0286 | 117.2174 | 20.5229 | 0.002588 | 0.002346 | 250.9542 | 0.26109 | 8.4035 | 0.002508 | 0.004604 | 0.004547 | 5.1e-05 | 0.000984 | 0.29247 | 1.2963 | 2.5475 | 98.53 | 19.98 | 50.06 | 83.37 | 81.4356 | 77.5443 | 0.011876 | 1.0067 | 0.023299 | 23.1438 | -0.03635 | 0.15532 | 0.006332 | 0.43921 | 0.43921 | 0.014252
2019-04-09 14:40 | 1554835191.5292 | 310 | 0.20827 | 0.001455 | 4.1e-05 | 99.6357 | 99.7884 | 117.4455 | 20.5229 | 0.002548 | 0.002396 | 250.9259 | 0.261 | 8.403 | 0.00256 | 0.003778 | 0.003785 | 5.1e-05 | 0.000991 | 0.28799 | 1.3302 | 2.5476 | 98.53 | 19.98 | 50.06 | 82.3911 | 82.6844 | 78.737 | 0.011876 | 1.0033 | 0.023282 | 23.2057 | -0.013694 | 0.15527 | 0.00619 | 0.42907 | 0.42907 | 0.014583
2019-04-09 14:40 | 1554835192.5297 | 311 | 0.20829 | 0.001453 | 4.1e-05 | 99.6059 | 99.8902 | 117.8952 | 20.5238 | 0.002563 | 0.002371 | 250.8578 | 0.26081 | 8.4015 | 0.002503 | 0.00355 | 0.003468 | 5.1e-05 | 0.000993 | 0.26825 | 1.3375 | 2.5476 | 98.53 | 19.98 | 50.056 | 82.7632 | 81.3156 | 78.1363 | 0.011875 | 1.005 | 0.023347 | 23.2302 | -0.025494 | 0.15515 | 0.006088 | 0.42313 | 0.42313 | 0.014777
2019-04-09 14:40 | 1554835193.5296 | 312 | 0.20825 | 0.001449 | 4.1e-05 | 99.6318 | 99.9673 | 117.8161 | 20.5227 | 0.002537 | 0.002341 | 250.871 | 0.26085 | 8.4001 | 0.002564 | 0.004046 | 0.003942 | 5.2e-05 | 0.000991 | 0.27831 | 1.3202 | 2.5476 | 98.53 | 19.98 | 50.054 | 82.1467 | 82.7753 | 77.4188 | 0.011875 | 1.0071 | 0.023363 | 23.1987 | -0.030068 | 0.15518 | 0.006313 | 0.43915 | 0.43915 | 0.014241
2019-04-09 14:40 | 1554835194.5293 | 313 | 0.20826 | 0.001449 | 4.1e-05 | 99.6485 | 100.0568 | 117.4702 | 20.5228 | 0.00257 | 0.002377 | 250.8086 | 0.26154 | 8.399 | 0.00255 | 0.004287 | 0.004236 | 5.1e-05 | 0.00099 | 0.28136 | 1.3257 | 2.5476 | 98.53 | 19.98 | 50.052 | 82.9237 | 82.4487 | 78.2989 | 0.011874 | 1.0045 | 0.023299 | 23.1937 | -0.036574 | 0.15559 | 0.006278 | 0.43554 | 0.43554 | 0.014397
2019-04-09 14:40 | 1554835195.5293 | 314 | 0.20827 | 0.001441 | 4.1e-05 | 99.6321 | 99.8679 | 117.5603 | 20.5223 | 0.002557 | 0.00235 | 250.8204 | 0.26072 | 8.4032 | 0.002531 | 0.004773 | 0.004656 | 5.3e-05 | 0.000985 | 0.27544 | 1.3411 | 2.5476 | 98.53 | 19.98 | 50.051 | 82.605 | 82.0021 | 77.6459 | 0.011874 | 1.0064 | 0.02333 | 23.181 | -0.021138 | 0.1551 | 0.006233 | 0.43292 | 0.43292 | 0.014439
2019-04-09 14:40 | 1554835196.5284 | 315 | 0.2083 | 0.001443 | 4.2e-05 | 99.6938 | 100.0702 | 117.5807 | 20.5231 | 0.002568 | 0.002356 | 250.7659 | 0.26082 | 8.4018 | 0.002512 | 0.004176 | 0.004154 | 5.1e-05 | 0.000981 | 0.27489 | 1.3375 | 2.5475 | 98.53 | 19.98 | 50.051 | 82.8858 | 81.5317 | 77.7769 | 0.011874 | 1.006 | 0.023327 | 23.1873 | -0.033712 | 0.15516 | 0.006049 | 0.42001 | 0.42001 | 0.014887
2019-04-09 14:40 | 1554835197.5284 | 316 | 0.20827 | 0.001435 | 4.2e-05 | 99.7623 | 100.0575 | 117.7676 | 20.5238 | 0.002572 | 0.002389 | 250.7608 | 0.2613 | 8.4026 | 0.002585 | 0.004582 | 0.004437 | 5.2e-05 | 0.000984 | 0.28982 | 1.3383 | 2.5476 | 98.53 | 19.98 | 50.05 | 82.988 | 83.2904 | 78.5851 | 0.011874 | 1.0037 | 0.023319 | 23.2325 | -0.02643 | 0.15544 | 0.006208 | 0.43098 | 0.43098 | 0.014535
2019-04-09 14:40 | 1554835198.5284 | 317 | 0.20825 | 0.001434 | 4.2e-05 | 99.79 | 100.3314 | 118.0945 | 20.5226 | 0.00261 | 0.002451 | 250.7311 | 0.26158 | 8.4015 | 0.002522 | 0.00377 | 0.003748 | 5.2e-05 | 0.000976 | 0.29286 | 1.3736 | 2.5476 | 98.53 | 19.98 | 50.048 | 83.8947 | 81.7677 | 80.0748 | 0.011874 | 0.99949 | 0.023302 | 23.3139 | -0.048404 | 0.15561 | 0.006327 | 0.43896 | 0.43896 | 0.014286
2019-04-09 14:40 | 1554835199.5294 | 318 | 0.20825 | 0.001436 | 4.1e-05 | 99.7947 | 100.2041 | 118.0887 | 20.522 | 0.002536 | 0.002408 | 250.6825 | 0.26044 | 8.4023 | 0.002523 | 0.004076 | 0.003934 | 5e-05 | 0.000968 | 0.2855 | 1.3665 | 2.5476 | 98.53 | 19.98 | 50.048 | 82.1008 | 81.7919 | 79.0342 | 0.011874 | 1.0024 | 0.023336 | 23.279 | -0.036617 | 0.15494 | 0.006365 | 0.44227 | 0.44227 | 0.014118
2019-04-09 14:40 | 1554835200.5295 | 319 | 0.20828 | 0.00143 | 4.1e-05 | 99.7646 | 100.1559 | 118.0002 | 20.5238 | 0.002524 | 0.002408 | 250.7042 | 0.26125 | 8.4019 | 0.002515 | 0.003829 | 0.00379 | 5.2e-05 | 0.000962 | 0.27041 | 1.3541 | 2.5475 | 98.53 | 19.98 | 50.049 | 81.8256 | 81.6073 | 79.0272 | 0.011874 | 1.0025 | 0.023327 | 23.27 | -0.035012 | 0.15542 | 0.006165 | 0.42817 | 0.42817 | 0.014628
2019-04-09 14:40 | 1554835201.5286 | 320 | 0.20826 | 0.001433 | 4e-05 | 99.8221 | 99.9553 | 117.7124 | 20.5238 | 0.002557 | 0.002375 | 250.6698 | 0.26087 | 8.4027 | 0.002503 | 0.003593 | 0.003623 | 4.9e-05 | 0.000952 | 0.27472 | 1.4038 | 2.5475 | 98.53 | 19.98 | 50.05 | 82.615 | 81.3259 | 78.2485 | 0.011874 | 1.0047 | 0.023325 | 23.2159 | -0.011924 | 0.15519 | 0.006307 | 0.43806 | 0.43806 | 0.014277
2019-04-09 14:40 | 1554835202.5286 | 321 | 0.20826 | 0.001431 | 4e-05 | 99.8156 | 100.0661 | 117.8941 | 20.5241 | 0.002547 | 0.002382 | 250.6987 | 0.26091 | 8.402 | 0.002492 | 0.003992 | 0.003896 | 5.3e-05 | 0.000952 | 0.29286 | 1.3612 | 2.5476 | 98.53 | 19.98 | 50.051 | 82.3822 | 81.0493 | 78.4117 | 0.011874 | 1.0042 | 0.023337 | 23.2392 | -0.022422 | 0.15521 | 0.006308 | 0.43834 | 0.43834 | 0.01427
2019-04-09 14:40 | 1554835203.5286 | 322 | 0.20827 | 0.001426 | 4e-05 | 99.8564 | 100.2103 | 117.9246 | 20.5239 | 0.002492 | 0.002334 | 250.6198 | 0.26082 | 8.4009 | 0.002483 | 0.003594 | 0.003502 | 5e-05 | 0.000951 | 0.29031 | 1.3208 | 2.5474 | 98.53 | 19.98 | 50.055 | 81.0561 | 80.8324 | 77.264 | 0.011875 | 1.0075 | 0.023378 | 23.2042 | -0.031647 | 0.15516 | 0.006206 | 0.43203 | 0.43203 | 0.014474
2019-04-09 14:40 | 1554835204.5296 | 323 | 0.20829 | 0.001425 | 4e-05 | 99.8629 | 100.2075 | 118.0232 | 20.524 | 0.002554 | 0.002372 | 250.6187 | 0.26093 | 8.4008 | 0.002587 | 0.004053 | 0.003963 | 5.1e-05 | 0.00095 | 0.27383 | 1.312 | 2.5474 | 98.53 | 19.98 | 50.058 | 82.5535 | 83.35 | 78.1747 | 0.011876 | 1.0049 | 0.023358 | 23.2441 | -0.030814 | 0.15523 | 0.006094 | 0.4238 | 0.4238 | 0.014761
2019-04-09 14:40 | 1554835205.5286 | 324 | 0.20829 | 0.001423 | 4e-05 | 99.9073 | 100.2179 | 117.9678 | 20.5235 | 0.002604 | 0.002446 | 250.5722 | 0.26081 | 8.4023 | 0.002521 | 0.004038 | 0.003998 | 5.1e-05 | 0.000951 | 0.26449 | 1.329 | 2.5476 | 98.53 | 19.98 | 50.064 | 83.7395 | 81.7443 | 79.9437 | 0.011877 | 0.99986 | 0.023294 | 23.2971 | -0.027762 | 0.15515 | 0.006107 | 0.42355 | 0.42355 | 0.014763
2019-04-09 14:40 | 1554835206.5297 | 325 | 0.20829 | 0.001418 | 4e-05 | 99.9356 | 100.2199 | 117.9502 | 20.5242 | 0.002443 | 0.002377 | 250.5661 | 0.26005 | 8.4015 | 0.002545 | 0.004238 | 0.004308 | 4.7e-05 | 0.000956 | 0.29479 | 1.3052 | 2.5475 | 98.53 | 19.98 | 50.066 | 79.8696 | 82.3205 | 78.2885 | 0.011878 | 1.0046 | 0.023347 | 23.2407 | -0.025414 | 0.15471 | 0.006146 | 0.42722 | 0.42722 | 0.014594
2019-04-09 14:40 | 1554835207.5288 | 326 | 0.20828 | 0.001416 | 4.1e-05 | 99.9257 | 100.0926 | 118.1011 | 20.5247 | 0.002524 | 0.002379 | 250.54 | 0.26215 | 8.4016 | 0.002482 | 0.004124 | 0.004041 | 5.2e-05 | 0.000959 | 0.29463 | 1.3078 | 2.5474 | 98.53 | 19.97 | 50.068 | 81.8306 | 80.8213 | 78.3389 | 0.011871 | 1.0044 | 0.02336 | 23.2572 | -0.014922 | 0.15595 | 0.006158 | 0.42824 | 0.42824 | 0.014676
2019-04-09 14:40 | 1554835208.5302 | 327 | 0.20834 | 0.001411 | 4.1e-05 | 99.932 | 100.2963 | 118.0534 | 20.5238 | 0.00254 | 0.002359 | 250.4797 | 0.26117 | 8.4023 | 0.00256 | 0.004292 | 0.004197 | 4.5e-05 | 0.000959 | 0.29192 | 1.3062 | 2.5475 | 98.53 | 19.97 | 50.074 | 82.2109 | 82.6968 | 77.847 | 0.011872 | 1.0058 | 0.023372 | 23.2362 | -0.03255 | 0.15537 | 0.005819 | 0.40478 | 0.40478 | 0.015468
2019-04-09 14:40 | 1554835209.5288 | 328 | 0.20832 | 0.001406 | 4.2e-05 | 99.91 | 100.2309 | 117.9253 | 20.5232 | 0.002538 | 0.002367 | 250.5075 | 0.26132 | 8.4024 | 0.002417 | 0.004212 | 0.004214 | 4.9e-05 | 0.000949 | 0.26737 | 1.3865 | 2.5475 | 98.53 | 19.97 | 50.078 | 82.1666 | 79.2473 | 78.0382 | 0.011873 | 1.0053 | 0.023353 | 23.2299 | -0.02869 | 0.15546 | 0.005958 | 0.41411 | 0.41411 | 0.015129
2019-04-09 14:40 | 1554835210.5278 | 329 | 0.20834 | 0.0014 | 4.3e-05 | 99.9573 | 100.2055 | 118.1192 | 20.5239 | 0.002523 | 0.002386 | 250.4182 | 0.26137 | 8.4 | 0.002462 | 0.00425 | 0.004208 | 4.9e-05 | 0.000945 | 0.28871 | 1.34 | 2.5476 | 98.53 | 19.97 | 50.087 | 81.7969 | 80.3247 | 78.5136 | 0.011875 | 1.0039 | 0.023356 | 23.2648 | -0.022179 | 0.15549 | 0.005834 | 0.40542 | 0.40542 | 0.015456
2019-04-09 14:40 | 1554835211.5278 | 330 | 0.20835 | 0.001397 | 4.3e-05 | 99.9409 | 100.1539 | 118.4691 | 20.5234 | 0.002563 | 0.00235 | 250.4199 | 0.26158 | 8.3987 | 0.00248 | 0.004035 | 0.003977 | 4.8e-05 | 0.000952 | 0.2656 | 1.3574 | 2.5476 | 98.53 | 19.97 | 50.091 | 82.7628 | 80.7735 | 77.6455 | 0.011876 | 1.0064 | 0.02342 | 23.2704 | -0.019035 | 0.15562 | 0.005772 | 0.40218 | 0.40218 | 0.015593
2019-04-09 14:40 | 1554835212.5279 | 331 | 0.20833 | 0.001391 | 4.3e-05 | 99.9073 | 100.2867 | 117.9241 | 20.5244 | 0.002551 | 0.002348 | 250.354 | 0.26076 | 8.4019 | 0.002468 | 0.003769 | 0.003802 | 5.1e-05 | 0.000961 | 0.26842 | 1.3528 | 2.5476 | 98.53 | 19.97 | 50.096 | 82.4628 | 80.4688 | 77.592 | 0.011878 | 1.0066 | 0.023368 | 23.2151 | -0.033904 | 0.15512 | 0.005887 | 0.40931 | 0.40931 | 0.015273
2019-04-09 14:40 | 1554835213.5279 | 332 | 0.20834 | 0.001396 | 4.4e-05 | 99.9966 | 100.3418 | 117.7753 | 20.5244 | 0.002618 | 0.00237 | 250.3502 | 0.26052 | 8.4045 | 0.002598 | 0.004347 | 0.004236 | 5.1e-05 | 0.000966 | 0.26952 | 1.3157 | 2.5476 | 98.53 | 19.97 | 50.106 | 84.0831 | 83.6149 | 78.1297 | 0.01188 | 1.005 | 0.023335 | 23.2182 | -0.030819 | 0.15498 | 0.005856 | 0.40658 | 0.40658 | 0.015362
2019-04-09 14:40 | 1554835214.528 | 333 | 0.20833 | 0.001387 | 4.3e-05 | 100.0264 | 100.4333 | 117.6854 | 20.5245 | 0.002542 | 0.002335 | 250.3237 | 0.26162 | 8.3999 | 0.002605 | 0.004229 | 0.004163 | 5.1e-05 | 0.000974 | 0.2855 | 1.326 | 2.5476 | 98.53 | 19.97 | 50.112 | 82.2592 | 83.7655 | 77.2795 | 0.011881 | 1.0075 | 0.023354 | 23.1812 | -0.036312 | 0.15564 | 0.00589 | 0.40927 | 0.40927 | 0.015325
2019-04-09 14:40 | 1554835215.5283 | 334 | 0.20835 | 0.001386 | 4.4e-05 | 100.0121 | 100.3576 | 117.9328 | 20.5243 | 0.002529 | 0.002339 | 250.2708 | 0.26128 | 8.4024 | 0.002624 | 0.004008 | 0.00397 | 4.9e-05 | 0.000974 | 0.29374 | 1.3307 | 2.5474 | 98.53 | 19.97 | 50.117 | 81.9471 | 84.2249 | 77.3766 | 0.011883 | 1.0072 | 0.023376 | 23.2088 | -0.030845 | 0.15543 | 0.005823 | 0.40495 | 0.40495 | 0.015468
2019-04-09 14:40 | 1554835216.5281 | 335 | 0.20835 | 0.001388 | 4.3e-05 | 100.0133 | 100.2853 | 117.7298 | 20.5242 | 0.002495 | 0.002309 | 250.2576 | 0.26065 | 8.3984 | 0.002589 | 0.004127 | 0.004075 | 5.1e-05 | 0.000975 | 0.27151 | 1.3076 | 2.5475 | 98.53 | 19.97 | 50.129 | 81.1359 | 83.3939 | 76.6672 | 0.011885 | 1.0092 | 0.023379 | 23.1653 | -0.024295 | 0.15506 | 0.00577 | 0.40135 | 0.40135 | 0.01557
2019-04-09 14:40 | 1554835217.529 | 336 | 0.20836 | 0.001385 | 4.3e-05 | 100.0204 | 100.2881 | 117.6565 | 20.5239 | 0.002575 | 0.002344 | 250.2032 | 0.26097 | 8.4039 | 0.002543 | 0.003443 | 0.003481 | 4.9e-05 | 0.000971 | 0.27992 | 1.2689 | 2.5475 | 98.53 | 19.97 | 50.136 | 83.0573 | 82.2861 | 77.5009 | 0.011887 | 1.0068 | 0.023344 | 23.1857 | -0.023901 | 0.15525 | 0.005722 | 0.39739 | 0.39739 | 0.015744
2019-04-09 14:40 | 1554835218.5291 | 337 | 0.20839 | 0.001381 | 4.3e-05 | 100.0243 | 100.2833 | 117.4838 | 20.5244 | 0.00258 | 0.002406 | 250.2135 | 0.26062 | 8.4 | 0.002567 | 0.003513 | 0.003385 | 4.8e-05 | 0.000971 | 0.28047 | 1.3093 | 2.5474 | 98.53 | 19.97 | 50.148 | 83.1585 | 82.8575 | 78.9841 | 0.01189 | 1.0026 | 0.023278 | 23.2176 | -0.023125 | 0.15504 | 0.005587 | 0.38686 | 0.38686 | 0.016151
2019-04-09 14:40 | 1554835219.5291 | 338 | 0.20836 | 0.001371 | 4.3e-05 | 100.0464 | 100.4567 | 117.5408 | 20.5238 | 0.002502 | 0.00235 | 250.1329 | 0.26068 | 8.4027 | 0.002559 | 0.00404 | 0.00394 | 5.2e-05 | 0.000977 | 0.30021 | 1.3202 | 2.5474 | 98.53 | 19.97 | 50.153 | 81.3023 | 82.6696 | 77.631 | 0.011891 | 1.0065 | 0.023328 | 23.1786 | -0.036612 | 0.15508 | 0.00575 | 0.39908 | 0.39908 | 0.01566
2019-04-09 14:40 | 1554835220.5281 | 339 | 0.20836 | 0.001368 | 4.4e-05 | 100.0389 | 100.3459 | 117.3669 | 20.5244 | 0.002509 | 0.002354 | 250.1435 | 0.26031 | 8.3992 | 0.002546 | 0.00392 | 0.003932 | 5.3e-05 | 0.00099 | 0.29319 | 1.3661 | 2.5474 | 98.53 | 19.97 | 50.157 | 81.4738 | 82.3488 | 77.7337 | 0.011892 | 1.0062 | 0.023308 | 23.1648 | -0.027405 | 0.15486 | 0.005767 | 0.3999 | 0.3999 | 0.015606
2019-04-09 14:40 | 1554835221.5282 | 340 | 0.20835 | 0.001366 | 4.3e-05 | 100.0931 | 100.4997 | 117.2631 | 20.5244 | 0.00254 | 0.002365 | 250.092 | 0.26094 | 8.4037 | 0.002541 | 0.004264 | 0.004214 | 5e-05 | 0.000993 | 0.28313 | 1.3719 | 2.5475 | 98.53 | 19.97 | 50.167 | 82.2115 | 82.2247 | 78.0011 | 0.011894 | 1.0054 | 0.023288 | 23.1634 | -0.036263 | 0.15523 | 0.00584 | 0.40469 | 0.40469 | 0.015458
2019-04-09 14:40 | 1554835222.5291 | 341 | 0.20836 | 0.001365 | 4.3e-05 | 100.1455 | 100.6553 | 117.6063 | 20.525 | 0.002466 | 0.002322 | 250.0796 | 0.26123 | 8.4019 | 0.002558 | 0.003994 | 0.003917 | 5.2e-05 | 0.000998 | 0.28512 | 1.3844 | 2.5476 | 98.53 | 19.97 | 50.173 | 80.4404 | 82.6364 | 76.9695 | 0.011896 | 1.0084 | 0.023357 | 23.1632 | -0.045414 | 0.15541 | 0.005787 | 0.40216 | 0.40216 | 0.015573
2019-04-09 14:40 | 1554835223.5281 | 342 | 0.20835 | 0.00136 | 4.4e-05 | 100.0484 | 100.4247 | 117.8873 | 20.5248 | 0.002557 | 0.002365 | 250.1023 | 0.26078 | 8.4036 | 0.002552 | 0.00377 | 0.003776 | 5.2e-05 | 0.001001 | 0.26178 | 1.4029 | 2.5478 | 98.53 | 19.96 | 50.181 | 82.6185 | 82.5038 | 77.997 | 0.01189 | 1.0054 | 0.02335 | 23.2248 | -0.033575 | 0.15514 | 0.005832 | 0.40516 | 0.40516 | 0.015431
2019-04-09 14:40 | 1554835224.5282 | 343 | 0.20832 | 0.001358 | 4.4e-05 | 100.0937 | 100.4608 | 117.8075 | 20.5246 | 0.002526 | 0.002325 | 250.0336 | 0.261 | 8.4001 | 0.002617 | 0.004201 | 0.004172 | 5.3e-05 | 0.001 | 0.29878 | 1.3489 | 2.5477 | 98.53 | 19.96 | 50.184 | 81.8625 | 84.0544 | 77.0331 | 0.011891 | 1.0082 | 0.023375 | 23.1851 | -0.032748 | 0.15527 | 0.006028 | 0.41924 | 0.41924 | 0.014925
2019-04-09 14:40 | 1554835225.5292 | 344 | 0.20837 | 0.001358 | 4.4e-05 | 100.1172 | 100.4257 | 117.4129 | 20.5241 | 0.002572 | 0.002367 | 250.058 | 0.26138 | 8.4007 | 0.002564 | 0.004113 | 0.00411 | 5.3e-05 | 0.001006 | 0.28241 | 1.331 | 2.5475 | 98.53 | 19.97 | 50.185 | 82.9779 | 82.783 | 78.0468 | 0.011899 | 1.0053 | 0.023302 | 23.1797 | -0.02752 | 0.1555 | 0.005738 | 0.39775 | 0.39775 | 0.015755
2019-04-09 14:40 | 1554835226.5283 | 345 | 0.20838 | 0.001359 | 4.4e-05 | 100.0553 | 100.4106 | 117.7039 | 20.5251 | 0.002533 | 0.002344 | 250.0015 | 0.26085 | 8.402 | 0.00261 | 0.004076 | 0.004003 | 5.1e-05 | 0.001004 | 0.28799 | 1.3548 | 2.5476 | 98.53 | 19.96 | 50.187 | 82.0405 | 83.8839 | 77.5061 | 0.011892 | 1.0068 | 0.023349 | 23.1905 | -0.031706 | 0.15518 | 0.005663 | 0.39326 | 0.39326 | 0.015903
2019-04-09 14:40 | 1554835227.5288 | 346 | 0.20834 | 0.001362 | 4.4e-05 | 100.0779 | 100.3562 | 117.7546 | 20.5243 | 0.002572 | 0.002357 | 250.0272 | 0.26108 | 8.4026 | 0.002636 | 0.004208 | 0.004106 | 5.5e-05 | 0.001008 | 0.30165 | 1.3196 | 2.5475 | 98.53 | 19.96 | 50.188 | 82.9673 | 84.5057 | 77.8179 | 0.011892 | 1.0059 | 0.023343 | 23.2058 | -0.024838 | 0.15532 | 0.005876 | 0.40807 | 0.40807 | 0.015339
2019-04-09 14:40 | 1554835228.5284 | 347 | 0.20836 | 0.001363 | 4.4e-05 | 100.1279 | 100.5744 | 117.6173 | 20.5245 | 0.002489 | 0.002351 | 249.9822 | 0.26049 | 8.4002 | 0.002661 | 0.003711 | 0.003714 | 5.1e-05 | 0.00101 | 0.27649 | 1.3324 | 2.5476 | 98.53 | 19.96 | 50.187 | 80.9733 | 85.1247 | 77.6629 | 0.011892 | 1.0064 | 0.023335 | 23.1872 | -0.039794 | 0.15496 | 0.005739 | 0.39838 | 0.39838 | 0.015676
2019-04-09 14:40 | 1554835229.5274 | 348 | 0.20835 | 0.001359 | 4.4e-05 | 100.1482 | 100.5737 | 117.4774 | 20.5255 | 0.002517 | 0.002333 | 249.9819 | 0.26122 | 8.4034 | 0.002694 | 0.004284 | 0.004209 | 5.3e-05 | 0.001004 | 0.29015 | 1.3631 | 2.5475 | 98.53 | 19.96 | 50.187 | 81.659 | 85.9077 | 77.2235 | 0.011892 | 1.0076 | 0.023335 | 23.1589 | -0.037924 | 0.1554 | 0.005844 | 0.40571 | 0.40571 | 0.015436
2019-04-09 14:40 | 1554835230.5274 | 349 | 0.2084 | 0.001361 | 4.5e-05 | 100.094 | 100.5699 | 117.3746 | 20.5257 | 0.002518 | 0.002319 | 249.9557 | 0.26112 | 8.404 | 0.002704 | 0.004596 | 0.004564 | 5.2e-05 | 0.001003 | 0.29999 | 1.3433 | 2.5474 | 98.53 | 19.96 | 50.185 | 81.6693 | 86.1646 | 76.8938 | 0.011891 | 1.0086 | 0.023336 | 23.1378 | -0.042429 | 0.15534 | 0.005553 | 0.38539 | 0.38539 | 0.016244
2019-04-09 14:40 | 1554835231.5274 | 350 | 0.20841 | 0.001356 | 4.6e-05 | 100.1134 | 100.5589 | 117.4204 | 20.5259 | 0.002513 | 0.002394 | 249.9265 | 0.2608 | 8.4027 | 0.002663 | 0.004136 | 0.004062 | 5.2e-05 | 0.001001 | 0.29192 | 1.385 | 2.5477 | 98.53 | 19.96 | 50.183 | 81.5541 | 85.1682 | 78.709 | 0.011891 | 1.0034 | 0.023281 | 23.2023 | -0.03972 | 0.15515 | 0.005489 | 0.37991 | 0.37991 | 0.016458
2019-04-09 14:41 | 1554835232.5289 | 351 | 0.20834 | 0.001358 | 4.6e-05 | 100.1503 | 100.3538 | 117.0887 | 20.5255 | 0.002588 | 0.002373 | 249.9353 | 0.26113 | 8.402 | 0.002606 | 0.004355 | 0.004291 | 5.2e-05 | 0.000995 | 0.28047 | 1.3613 | 2.5476 | 98.53 | 19.96 | 50.182 | 83.3538 | 83.8039 | 78.197 | 0.011891 | 1.0048 | 0.023265 | 23.1526 | -0.018158 | 0.15535 | 0.005905 | 0.40864 | 0.40864 | 0.015321
2019-04-09 14:41 | 1554835233.5275 | 352 | 0.2083 | 0.001359 | 4.6e-05 | 100.181 | 100.5462 | 116.9125 | 20.5266 | 0.002505 | 0.002382 | 249.8569 | 0.26108 | 8.4008 | 0.002712 | 0.004353 | 0.004359 | 5.2e-05 | 0.000996 | 0.26416 | 1.3515 | 2.5475 | 98.53 | 19.96 | 50.18 | 81.3637 | 86.3522 | 78.4218 | 0.01189 | 1.0042 | 0.02324 | 23.1426 | -0.03255 | 0.15532 | 0.006136 | 0.42423 | 0.42423 | 0.014754
2019-04-09 14:41 | 1554835234.5289 | 353 | 0.20834 | 0.001359 | 4.6e-05 | 100.1705 | 100.5544 | 117.2676 | 20.5257 | 0.002515 | 0.002337 | 249.8826 | 0.26091 | 8.401 | 0.00265 | 0.00422 | 0.004135 | 5.3e-05 | 0.000994 | 0.27903 | 1.3465 | 2.5477 | 98.53 | 19.96 | 50.178 | 81.6193 | 84.8505 | 77.3196 | 0.01189 | 1.0073 | 0.023311 | 23.1413 | -0.034215 | 0.15522 | 0.005897 | 0.40885 | 0.40885 | 0.0153
2019-04-09 14:41 | 1554835235.5279 | 354 | 0.20835 | 0.001359 | 4.6e-05 | 100.1714 | 100.4915 | 117.4065 | 20.5256 | 0.00249 | 0.002369 | 249.8192 | 0.26047 | 8.4032 | 0.002712 | 0.004135 | 0.004145 | 5.1e-05 | 0.000987 | 0.28008 | 1.3532 | 2.5475 | 98.53 | 19.96 | 50.176 | 80.9969 | 86.3488 | 78.0943 | 0.011889 | 1.0051 | 0.023299 | 23.1806 | -0.028531 | 0.15495 | 0.005849 | 0.40529 | 0.40529 | 0.015408
2019-04-09 14:41 | 1554835236.5276 | 355 | 0.20834 | 0.00136 | 4.6e-05 | 100.2218 | 100.531 | 117.3373 | 20.526 | 0.002511 | 0.002403 | 249.8195 | 0.26063 | 8.4007 | 0.002687 | 0.004541 | 0.0045 | 5.1e-05 | 0.000984 | 0.2896 | 1.3339 | 2.5475 | 98.53 | 19.96 | 50.175 | 81.505 | 85.7507 | 78.9165 | 0.011889 | 1.0028 | 0.023265 | 23.2009 | -0.027561 | 0.15505 | 0.005896 | 0.408 | 0.408 | 0.015315
2019-04-09 14:41 | 1554835237.5286 | 356 | 0.20835 | 0.00136 | 4.7e-05 | 100.2328 | 100.6343 | 117.0209 | 20.5257 | 0.002474 | 0.002342 | 249.7745 | 0.26079 | 8.4012 | 0.002678 | 0.004106 | 0.004035 | 5e-05 | 0.000991 | 0.28584 | 1.3218 | 2.5475 | 98.53 | 19.96 | 50.175 | 80.6125 | 85.5356 | 77.4397 | 0.011889 | 1.007 | 0.023283 | 23.121 | -0.035759 | 0.15515 | 0.005832 | 0.40378 | 0.40378 | 0.015485
2019-04-09 14:41 | 1554835238.5276 | 357 | 0.20836 | 0.001359 | 4.7e-05 | 100.2688 | 100.6377 | 117.4044 | 20.5267 | 0.002452 | 0.002329 | 249.7454 | 0.26082 | 8.4003 | 0.002699 | 0.004325 | 0.004254 | 5.3e-05 | 0.000991 | 0.25713 | 1.3414 | 2.5477 | 98.53 | 19.96 | 50.174 | 80.0934 | 86.0292 | 77.1302 | 0.011889 | 1.0079 | 0.023331 | 23.1486 | -0.03285 | 0.15516 | 0.005781 | 0.40106 | 0.40106 | 0.015591
2019-04-09 14:41 | 1554835239.5278 | 358 | 0.20836 | 0.001362 | 4.7e-05 | 100.189 | 100.5496 | 117.4577 | 20.5262 | 0.002538 | 0.002354 | 249.7293 | 0.26139 | 8.4016 | 0.002683 | 0.003908 | 0.003921 | 4.9e-05 | 0.000996 | 0.27007 | 1.3543 | 2.5476 | 98.53 | 19.96 | 50.17 | 82.1518 | 85.6438 | 77.7448 | 0.011888 | 1.0061 | 0.023316 | 23.1741 | -0.032138 | 0.1555 | 0.005796 | 0.40186 | 0.40186 | 0.015594
2019-04-09 14:41 | 1554835240.5277 | 359 | 0.20838 | 0.00136 | 4.7e-05 | 100.2521 | 100.7155 | 117.5619 | 20.5258 | 0.002541 | 0.002331 | 249.6866 | 0.25979 | 8.3988 | 0.00262 | 0.004681 | 0.004567 | 5.5e-05 | 0.000998 | 0.27776 | 1.3519 | 2.5476 | 98.53 | 19.96 | 50.168 | 82.234 | 84.1342 | 77.1925 | 0.011887 | 1.0077 | 0.023345 | 23.1662 | -0.041248 | 0.15455 | 0.005653 | 0.39238 | 0.39238 | 0.015873
2019-04-09 14:41 | 1554835241.5272 | 360 | 0.20838 | 0.001358 | 4.7e-05 | 100.3156 | 100.6415 | 117.6659 | 20.5269 | 0.002548 | 0.002321 | 249.6984 | 0.26116 | 8.4034 | 0.002613 | 0.004487 | 0.004463 | 5.1e-05 | 0.000997 | 0.27367 | 1.3653 | 2.5475 | 98.53 | 19.96 | 50.166 | 82.4043 | 83.9753 | 76.9466 | 0.011887 | 1.0084 | 0.023363 | 23.1683 | -0.029017 | 0.15536 | 0.005659 | 0.39306 | 0.39306 | 0.015929
2019-04-09 14:41 | 1554835242.5282 | 361 | 0.20839 | 0.001356 | 4.8e-05 | 100.3013 | 100.7155 | 117.8535 | 20.5266 | 0.002599 | 0.002331 | 249.6305 | 0.26099 | 8.4005 | 0.002621 | 0.004277 | 0.00427 | 5.1e-05 | 0.000994 | 0.26161 | 1.3288 | 2.5476 | 98.53 | 19.96 | 50.166 | 83.6209 | 84.1482 | 77.1996 | 0.011887 | 1.0077 | 0.023374 | 23.1951 | -0.036865 | 0.15526 | 0.005595 | 0.38877 | 0.38877 | 0.016095
2019-04-09 14:41 | 1554835243.5269 | 362 | 0.20836 | 0.001349 | 4.8e-05 | 100.3671 | 100.7355 | 118.018 | 20.5265 | 0.002543 | 0.002426 | 249.6277 | 0.26074 | 8.4013 | 0.002593 | 0.004485 | 0.004379 | 5e-05 | 0.000997 | 0.28783 | 1.3464 | 2.5476 | 98.53 | 19.96 | 50.165 | 82.2782 | 83.4945 | 79.4821 | 0.011886 | 1.0012 | 0.023314 | 23.2868 | -0.032771 | 0.15511 | 0.005787 | 0.40112 | 0.40112 | 0.015584
2019-04-09 14:41 | 1554835244.5279 | 363 | 0.20839 | 0.001345 | 4.8e-05 | 100.3254 | 100.7885 | 117.9542 | 20.5275 | 0.002556 | 0.002287 | 249.5811 | 0.26101 | 8.4021 | 0.002615 | 0.003823 | 0.003827 | 4.8e-05 | 0.000995 | 0.30359 | 1.377 | 2.5476 | 98.53 | 19.96 | 50.162 | 82.5944 | 84.0033 | 76.1232 | 0.011886 | 1.0108 | 0.02342 | 23.1693 | -0.041192 | 0.15527 | 0.005581 | 0.38848 | 0.38848 | 0.016108
2019-04-09 14:41 | 1554835245.5282 | 364 | 0.20837 | 0.001343 | 4.8e-05 | 100.3087 | 100.719 | 118.3544 | 20.5268 | 0.002432 | 0.002343 | 249.562 | 0.26066 | 8.4013 | 0.002672 | 0.004367 | 0.004308 | 5.2e-05 | 0.000993 | 0.27649 | 1.3336 | 2.5476 | 98.53 | 19.96 | 50.161 | 79.6252 | 85.3954 | 77.4648 | 0.011886 | 1.0069 | 0.023414 | 23.2531 | -0.036507 | 0.15507 | 0.005753 | 0.40049 | 0.40049 | 0.015604
2019-04-09 14:41 | 1554835246.5273 | 365 | 0.2084 | 0.001341 | 4.7e-05 | 100.3986 | 100.6701 | 118.422 | 20.5269 | 0.002517 | 0.002315 | 249.5484 | 0.26117 | 8.4026 | 0.002668 | 0.004394 | 0.004346 | 5.1e-05 | 0.000985 | 0.27599 | 1.3181 | 2.5474 | 98.53 | 19.96 | 50.159 | 81.6563 | 85.2856 | 76.8032 | 0.011885 | 1.0088 | 0.023443 | 23.2378 | -0.02415 | 0.15537 | 0.005569 | 0.38807 | 0.38807 | 0.016134
2019-04-09 14:41 | 1554835247.528 | 366 | 0.20839 | 0.001335 | 4.8e-05 | 100.4016 | 100.5885 | 118.6574 | 20.527 | 0.002518 | 0.002331 | 249.5019 | 0.26139 | 8.4007 | 0.002606 | 0.004243 | 0.004257 | 5.2e-05 | 0.000989 | 0.30055 | 1.2984 | 2.5474 | 98.53 | 19.96 | 50.157 | 81.6845 | 83.7937 | 77.1913 | 0.011885 | 1.0077 | 0.023453 | 23.2738 | -0.016635 | 0.1555 | 0.005632 | 0.39264 | 0.39264 | 0.01596
2019-04-09 14:41 | 1554835248.528 | 367 | 0.20842 | 0.001341 | 4.8e-05 | 100.4248 | 100.8494 | 118.567 | 20.5267 | 0.002478 | 0.002339 | 249.5168 | 0.26139 | 8.3998 | 0.002527 | 0.004749 | 0.00474 | 4.6e-05 | 0.000991 | 0.28816 | 1.3324 | 2.5477 | 98.53 | 19.96 | 50.154 | 80.725 | 81.8941 | 77.3712 | 0.011884 | 1.0072 | 0.023439 | 23.2709 | -0.037733 | 0.1555 | 0.005458 | 0.38024 | 0.38024 | 0.016481
2019-04-09 14:41 | 1554835249.5284 | 368 | 0.2084 | 0.001337 | 4.8e-05 | 100.3843 | 100.6233 | 118.6513 | 20.5274 | 0.00248 | 0.002322 | 249.4722 | 0.26062 | 8.4029 | 0.002558 | 0.00405 | 0.003989 | 5.2e-05 | 0.000994 | 0.28871 | 1.3258 | 2.5474 | 98.53 | 19.96 | 50.149 | 80.768 | 82.6349 | 76.9791 | 0.011883 | 1.0083 | 0.02346 | 23.2662 | -0.021264 | 0.15504 | 0.005543 | 0.38656 | 0.38656 | 0.016164
2019-04-09 14:41 | 1554835250.5271 | 369 | 0.20844 | 0.001331 | 4.8e-05 | 100.3483 | 100.8415 | 118.7797 | 20.5276 | 0.002543 | 0.002359 | 249.4757 | 0.26094 | 8.4012 | 0.002631 | 0.003531 | 0.003606 | 5.2e-05 | 0.000998 | 0.27721 | 1.3318 | 2.5475 | 98.53 | 19.96 | 50.146 | 82.2911 | 84.3911 | 77.8546 | 0.011882 | 1.0058 | 0.023443 | 23.3078 | -0.04385 | 0.15524 | 0.005314 | 0.3702 | 0.3702 | 0.016899
2019-04-09 14:41 | 1554835251.5282 | 370 | 0.20844 | 0.001327 | 4.8e-05 | 100.3778 | 100.794 | 118.4713 | 20.5276 | 0.002612 | 0.002301 | 249.4243 | 0.26029 | 8.4024 | 0.002674 | 0.004403 | 0.004284 | 5e-05 | 0.001001 | 0.28191 | 1.3464 | 2.5474 | 98.53 | 19.97 | 50.143 | 83.9354 | 85.4289 | 76.4702 | 0.011889 | 1.0098 | 0.023459 | 23.2316 | -0.037009 | 0.15485 | 0.005352 | 0.3731 | 0.3731 | 0.016726
2019-04-09 14:41 | 1554835252.5271 | 371 | 0.20843 | 0.001322 | 4.8e-05 | 100.4555 | 100.8043 | 118.2102 | 20.5284 | 0.00253 | 0.002342 | 249.3932 | 0.26106 | 8.4019 | 0.002595 | 0.004465 | 0.00447 | 5.1e-05 | 0.001004 | 0.28982 | 1.329 | 2.5475 | 98.53 | 19.97 | 50.134 | 81.9673 | 83.5447 | 77.4514 | 0.011887 | 1.007 | 0.023401 | 23.2385 | -0.031002 | 0.15531 | 0.005374 | 0.37367 | 0.37367 | 0.01675
2019-04-09 14:41 | 1554835253.5282 | 372 | 0.20846 | 0.001316 | 4.9e-05 | 100.4442 | 100.7723 | 118.3314 | 20.5277 | 0.002549 | 0.00236 | 249.3903 | 0.26126 | 8.4008 | 0.00266 | 0.003987 | 0.003946 | 5.1e-05 | 0.001002 | 0.29048 | 1.3575 | 2.5476 | 98.53 | 19.97 | 50.129 | 82.4316 | 85.0935 | 77.8967 | 0.011885 | 1.0057 | 0.023398 | 23.2652 | -0.029169 | 0.15542 | 0.005243 | 0.36447 | 0.36447 | 0.017185
2019-04-09 14:41 | 1554835254.5282 | 373 | 0.20845 | 0.001313 | 4.9e-05 | 100.4305 | 100.953 | 118.2485 | 20.5276 | 0.002554 | 0.002346 | 249.356 | 0.2612 | 8.4028 | 0.00261 | 0.004279 | 0.004249 | 5.2e-05 | 0.001003 | 0.28102 | 1.3542 | 2.5476 | 98.53 | 19.96 | 50.12 | 82.5416 | 83.9034 | 77.5559 | 0.011876 | 1.0067 | 0.023401 | 23.2458 | -0.046412 | 0.15539 | 0.005292 | 0.36789 | 0.36789 | 0.017022
2019-04-09 14:41 | 1554835255.5275 | 374 | 0.20844 | 0.001307 | 5e-05 | 100.4192 | 100.8707 | 117.9586 | 20.5279 | 0.002493 | 0.002324 | 249.3654 | 0.26027 | 8.3986 | 0.002567 | 0.004036 | 0.004001 | 5.3e-05 | 0.000997 | 0.28031 | 1.3254 | 2.5475 | 98.53 | 19.97 | 50.115 | 81.0819 | 82.8631 | 77.0293 | 0.011882 | 1.0082 | 0.02339 | 23.1998 | -0.040129 | 0.15483 | 0.005386 | 0.37424 | 0.37424 | 0.016673
2019-04-09 14:41 | 1554835256.5277 | 375 | 0.20845 | 0.001308 | 5e-05 | 100.5398 | 100.9124 | 117.8713 | 20.5283 | 0.002459 | 0.002341 | 249.3088 | 0.26045 | 8.4022 | 0.00259 | 0.004203 | 0.004023 | 5.2e-05 | 0.000999 | 0.29767 | 1.3323 | 2.5475 | 98.53 | 19.97 | 50.11 | 80.2612 | 83.4094 | 77.4318 | 0.011881 | 1.007 | 0.023368 | 23.2046 | -0.033085 | 0.15494 | 0.005273 | 0.366 | 0.366 | 0.017061
2019-04-09 14:41 | 1554835257.5263 | 376 | 0.20836 | 0.001307 | 5e-05 | 100.4889 | 101.0187 | 117.7558 | 20.5283 | 0.002522 | 0.002365 | 249.3148 | 0.26077 | 8.4017 | 0.002611 | 0.003749 | 0.003815 | 5.1e-05 | 0.001003 | 0.2896 | 1.359 | 2.5474 | 98.53 | 19.97 | 50.098 | 81.7841 | 83.9158 | 78.0167 | 0.011878 | 1.0053 | 0.023337 | 23.2125 | -0.047037 | 0.15513 | 0.005848 | 0.40565 | 0.40565 | 0.015412
2019-04-09 14:41 | 1554835258.5264 | 377 | 0.20849 | 0.001303 | 5.1e-05 | 100.5061 | 100.7892 | 117.6293 | 20.5273 | 0.002528 | 0.002368 | 249.2541 | 0.26076 | 8.4028 | 0.002633 | 0.003962 | 0.003925 | 5e-05 | 0.001 | 0.27527 | 1.313 | 2.5476 | 98.53 | 19.97 | 50.092 | 81.9242 | 84.4415 | 78.0796 | 0.011877 | 1.0052 | 0.023322 | 23.2021 | -0.025151 | 0.15513 | 0.005082 | 0.35196 | 0.35196 | 0.017763
2019-04-09 14:41 | 1554835259.5264 | 378 | 0.20844 | 0.001301 | 5.1e-05 | 100.493 | 100.9571 | 117.6521 | 20.5277 | 0.002491 | 0.002295 | 249.246 | 0.26059 | 8.4002 | 0.002631 | 0.004437 | 0.004377 | 5.1e-05 | 0.000998 | 0.29031 | 1.3545 | 2.5475 | 98.53 | 19.97 | 50.08 | 81.0458 | 84.388 | 76.315 | 0.011874 | 1.0102 | 0.023383 | 23.146 | -0.04121 | 0.15502 | 0.005336 | 0.37063 | 0.37063 | 0.016856
2019-04-09 14:41 | 1554835260.5274 | 379 | 0.20854 | 0.001306 | 5.1e-05 | 100.5484 | 100.9867 | 117.5058 | 20.5289 | 0.00252 | 0.002264 | 249.214 | 0.26103 | 8.3992 | 0.002627 | 0.004273 | 0.004218 | 5.1e-05 | 0.001003 | 0.2959 | 1.3077 | 2.5474 | 98.53 | 19.97 | 50.074 | 81.7263 | 84.295 | 75.5897 | 0.011872 | 1.0123 | 0.023393 | 23.1076 | -0.038903 | 0.15529 | 0.004758 | 0.3303 | 0.3303 | 0.018947
2019-04-09 14:41 | 1554835261.5881 | 380 | 0.20847 | 0.001303 | 5.1e-05 | 100.5976 | 100.8755 | 117.7701 | 20.5284 | 0.002507 | 0.002376 | 249.1796 | 0.26089 | 8.4015 | 0.002649 | 0.003985 | 0.003899 | 5e-05 | 0.001001 | 0.29015 | 1.3235 | 2.5475 | 98.53 | 19.97 | 50.068 | 81.4075 | 84.8313 | 78.2669 | 0.011871 | 1.0046 | 0.02333 | 23.2222 | -0.024682 | 0.15521 | 0.005186 | 0.35929 | 0.35929 | 0.017408
2019-04-09 14:41 | 1554835262.54 | 381 | 0.20843 | 0.001303 | 5.1e-05 | 100.5672 | 101.0376 | 117.979 | 20.5278 | 0.002592 | 0.002302 | 249.2055 | 0.26073 | 8.4024 | 0.00266 | 0.004038 | 0.004057 | 4.9e-05 | 0.001006 | 0.28982 | 1.3144 | 2.5474 | 98.53 | 19.97 | 50.056 | 83.4715 | 85.1067 | 76.4898 | 0.011868 | 1.0097 | 0.02341 | 23.1839 | -0.041744 | 0.15511 | 0.005422 | 0.37709 | 0.37709 | 0.016577
2019-04-09 14:41 | 1554835263.541 | 382 | 0.2085 | 0.001307 | 5.2e-05 | 100.5088 | 101.1047 | 117.7234 | 20.5281 | 0.002464 | 0.002369 | 249.1437 | 0.26093 | 8.403 | 0.002641 | 0.00339 | 0.003281 | 4.7e-05 | 0.001002 | 0.28783 | 1.3469 | 2.5475 | 98.53 | 19.97 | 50.051 | 80.3951 | 84.6506 | 78.0991 | 0.011867 | 1.0051 | 0.023331 | 23.2121 | -0.052876 | 0.15523 | 0.005022 | 0.3478 | 0.3478 | 0.017986
2019-04-09 14:41 | 1554835264.5396 | 383 | 0.20843 | 0.001301 | 5.2e-05 | 100.5279 | 101.0473 | 117.7828 | 20.5274 | 0.002552 | 0.002323 | 249.145 | 0.26077 | 8.4 | 0.00263 | 0.004167 | 0.004185 | 4.8e-05 | 0.001005 | 0.27776 | 1.3322 | 2.5475 | 98.53 | 19.98 | 50.041 | 82.5094 | 84.37 | 77.0038 | 0.011872 | 1.0083 | 0.023373 | 23.1817 | -0.046094 | 0.15513 | 0.005403 | 0.37506 | 0.37506 | 0.016669
2019-04-09 14:41 | 1554835265.5406 | 384 | 0.2085 | 0.0013 | 5.2e-05 | 100.6131 | 100.9716 | 117.4645 | 20.5299 | 0.002521 | 0.002309 | 249.0867 | 0.26086 | 8.4026 | 0.002652 | 0.003973 | 0.00392 | 4.8e-05 | 0.001007 | 0.30574 | 1.3343 | 2.5477 | 98.53 | 19.98 | 50.036 | 81.7499 | 84.9138 | 76.6509 | 0.011871 | 1.0093 | 0.023353 | 23.1387 | -0.031815 | 0.15519 | 0.004987 | 0.34567 | 0.34567 | 0.018092
2019-04-09 14:41 | 1554835266.5401 | 385 | 0.20847 | 0.001299 | 5.1e-05 | 100.6446 | 100.921 | 117.6108 | 20.5283 | 0.002551 | 0.002339 | 249.104 | 0.26077 | 8.4021 | 0.002603 | 0.004471 | 0.004466 | 4.8e-05 | 0.001007 | 0.28672 | 1.3342 | 2.5475 | 98.53 | 19.98 | 50.031 | 82.4828 | 83.7247 | 77.3852 | 0.011869 | 1.0072 | 0.023343 | 23.1774 | -0.024526 | 0.15513 | 0.005168 | 0.35824 | 0.35824 | 0.017452
2019-04-09 14:41 | 1554835267.5408 | 386 | 0.20843 | 0.001296 | 5.2e-05 | 100.6422 | 100.9069 | 117.6795 | 20.5299 | 0.002511 | 0.002298 | 249.0845 | 0.26103 | 8.4026 | 0.002659 | 0.004346 | 0.004354 | 5.1e-05 | 0.001009 | 0.286 | 1.3356 | 2.5475 | 98.53 | 19.98 | 50.021 | 81.5237 | 85.0653 | 76.403 | 0.011867 | 1.01 | 0.023383 | 23.1516 | -0.023487 | 0.15529 | 0.005451 | 0.37856 | 0.37856 | 0.016531
2019-04-09 14:41 | 1554835268.5408 | 387 | 0.20847 | 0.001303 | 5.2e-05 | 100.6187 | 101.165 | 117.7793 | 20.5282 | 0.002498 | 0.00232 | 249.0622 | 0.2613 | 8.402 | 0.002603 | 0.004111 | 0.00401 | 4.5e-05 | 0.001017 | 0.30326 | 1.3184 | 2.5475 | 98.53 | 19.98 | 50.017 | 81.2061 | 83.7356 | 76.9274 | 0.011866 | 1.0085 | 0.023375 | 23.1788 | -0.048429 | 0.15544 | 0.005146 | 0.35712 | 0.35712 | 0.017542
2019-04-09 14:41 | 1554835269.5398 | 388 | 0.20844 | 0.001296 | 5.3e-05 | 100.6345 | 101.0228 | 117.7516 | 20.5287 | 0.002536 | 0.002344 | 249.0674 | 0.26172 | 8.3988 | 0.002675 | 0.00417 | 0.004086 | 4.9e-05 | 0.001019 | 0.29192 | 1.3689 | 2.5474 | 98.53 | 19.98 | 50.013 | 82.1222 | 85.4495 | 77.5048 | 0.011865 | 1.0068 | 0.023353 | 23.1952 | -0.03445 | 0.1557 | 0.00535 | 0.37098 | 0.37098 | 0.016913
2019-04-09 14:41 | 1554835270.5409 | 389 | 0.20847 | 0.001302 | 5.3e-05 | 100.6336 | 100.9709 | 117.7945 | 20.5294 | 0.002538 | 0.002298 | 249.0221 | 0.26094 | 8.399 | 0.002654 | 0.00413 | 0.004106 | 4.8e-05 | 0.001022 | 0.27632 | 1.3503 | 2.5476 | 98.53 | 19.98 | 50.006 | 82.1727 | 84.9633 | 76.4049 | 0.011864 | 1.01 | 0.023394 | 23.163 | -0.029927 | 0.15523 | 0.005207 | 0.36161 | 0.36161 | 0.0173
2019-04-09 14:41 | 1554835271.5403 | 390 | 0.20847 | 0.001295 | 5.4e-05 | 100.6774 | 101.0466 | 117.5638 | 20.5293 | 0.002546 | 0.002354 | 249.036 | 0.26102 | 8.4007 | 0.002614 | 0.003498 | 0.003443 | 4.6e-05 | 0.001025 | 0.27417 | 1.3403 | 2.5475 | 98.53 | 19.98 | 50.002 | 82.3521 | 83.9832 | 77.7385 | 0.011863 | 1.0061 | 0.023327 | 23.1844 | -0.032742 | 0.15528 | 0.005207 | 0.36053 | 0.36053 | 0.017357
2019-04-09 14:41 | 1554835272.5399 | 391 | 0.20848 | 0.001295 | 5.5e-05 | 100.7488 | 101.0387 | 117.5608 | 20.5301 | 0.002464 | 0.0023 | 248.9828 | 0.26056 | 8.4017 | 0.002677 | 0.004232 | 0.004182 | 5e-05 | 0.001027 | 0.32012 | 1.3356 | 2.5476 | 98.53 | 19.98 | 49.995 | 80.3971 | 85.5204 | 76.4476 | 0.011861 | 1.0099 | 0.02337 | 23.1414 | -0.025694 | 0.15501 | 0.005138 | 0.3563 | 0.3563 | 0.017532
2019-04-09 14:41 | 1554835273.5399 | 392 | 0.20845 | 0.001293 | 5.5e-05 | 100.7176 | 101.0886 | 117.7481 | 20.5293 | 0.002506 | 0.002295 | 248.9967 | 0.26085 | 8.406 | 0.002706 | 0.004222 | 0.004226 | 4.8e-05 | 0.001026 | 0.31454 | 1.3397 | 2.5476 | 98.53 | 19.98 | 49.992 | 81.4045 | 86.1979 | 76.3197 | 0.01186 | 1.0102 | 0.023393 | 23.1556 | -0.032887 | 0.15518 | 0.005318 | 0.36928 | 0.36928 | 0.016935
2019-04-09 14:41 | 1554835274.5404 | 393 | 0.20843 | 0.001289 | 5.6e-05 | 100.7313 | 100.9698 | 117.8382 | 20.5287 | 0.002547 | 0.002278 | 248.9596 | 0.26078 | 8.4022 | 0.00272 | 0.004395 | 0.004484 | 4.9e-05 | 0.001021 | 0.27776 | 1.3541 | 2.5476 | 98.53 | 19.98 | 49.989 | 82.3857 | 86.5314 | 75.9084 | 0.01186 | 1.0114 | 0.023415 | 23.1508 | -0.021158 | 0.15514 | 0.005454 | 0.37907 | 0.37907 | 0.016493
2019-04-09 14:41 | 1554835275.5394 | 394 | 0.20845 | 0.001299 | 5.6e-05 | 100.6881 | 101.1509 | 118.0901 | 20.5287 | 0.002534 | 0.002321 | 248.9735 | 0.25997 | 8.4035 | 0.002709 | 0.003741 | 0.003564 | 4.8e-05 | 0.001021 | 0.29358 | 1.345 | 2.5475 | 98.53 | 19.98 | 49.983 | 82.076 | 86.2833 | 76.9561 | 0.011858 | 1.0084 | 0.023405 | 23.2103 | -0.041016 | 0.15466 | 0.005272 | 0.36618 | 0.36618 | 0.017021
2019-04-09 14:41 | 1554835276.5391 | 395 | 0.20844 | 0.00129 | 5.6e-05 | 100.7077 | 101.1581 | 118.1652 | 20.5296 | 0.002509 | 0.002336 | 248.9248 | 0.26089 | 8.4006 | 0.002705 | 0.004007 | 0.004008 | 5.1e-05 | 0.001024 | 0.28639 | 1.3335 | 2.5475 | 98.53 | 19.99 | 49.98 | 81.4625 | 86.1763 | 77.3132 | 0.011865 | 1.0074 | 0.023401 | 23.2295 | -0.039909 | 0.15521 | 0.005345 | 0.37124 | 0.37124 | 0.016848
2019-04-09 14:41 | 1554835277.5401 | 396 | 0.20844 | 0.00129 | 5.6e-05 | 100.7173 | 101.1519 | 117.9755 | 20.5298 | 0.002484 | 0.002296 | 248.8567 | 0.26079 | 8.4022 | 0.002646 | 0.004331 | 0.004291 | 4.8e-05 | 0.001014 | 0.28136 | 1.382 | 2.5475 | 98.53 | 19.99 | 49.977 | 80.8676 | 84.7593 | 76.3529 | 0.011864 | 1.0101 | 0.023414 | 23.1791 | -0.038515 | 0.15515 | 0.005391 | 0.37465 | 0.37465 | 0.016689
2019-04-09 14:41 | 1554835278.5391 | 397 | 0.20843 | 0.001289 | 5.7e-05 | 100.8114 | 101.1419 | 117.7246 | 20.5291 | 0.002445 | 0.002329 | 248.8542 | 0.26036 | 8.4006 | 0.002715 | 0.003769 | 0.003725 | 5e-05 | 0.001011 | 0.26432 | 1.3941 | 2.5476 | 98.53 | 19.99 | 49.976 | 79.9307 | 86.4163 | 77.1357 | 0.011864 | 1.0079 | 0.023363 | 23.1803 | -0.029279 | 0.15489 | 0.005432 | 0.37665 | 0.37665 | 0.016572
2019-04-09 14:41 | 1554835279.5405 | 398 | 0.20847 | 0.001296 | 5.7e-05 | 100.7596 | 101.1075 | 117.7518 | 20.5299 | 0.002501 | 0.002339 | 248.7848 | 0.26091 | 8.4021 | 0.002706 | 0.003863 | 0.003739 | 5e-05 | 0.00101 | 0.29734 | 1.3361 | 2.5476 | 98.53 | 19.99 | 49.976 | 81.2835 | 86.2111 | 77.3907 | 0.011864 | 1.0071 | 0.023357 | 23.1914 | -0.030834 | 0.15522 | 0.005209 | 0.361 | 0.361 | 0.017327
2019-04-09 14:41 | 1554835280.5392 | 399 | 0.20844 | 0.001292 | 5.7e-05 | 100.8093 | 101.1464 | 118.1394 | 20.5293 | 0.002439 | 0.002313 | 248.8028 | 0.26011 | 8.4028 | 0.002733 | 0.003702 | 0.003754 | 5e-05 | 0.001007 | 0.31758 | 1.362 | 2.5475 | 98.53 | 19.99 | 49.975 | 79.7843 | 86.8669 | 76.7508 | 0.011864 | 1.009 | 0.023417 | 23.2083 | -0.02986 | 0.15474 | 0.005344 | 0.37135 | 0.37135 | 0.016793
2019-04-09 14:41 | 1554835281.5396 | 400 | 0.20847 | 0.001284 | 5.7e-05 | 100.8054 | 101.2335 | 117.9938 | 20.5299 | 0.002491 | 0.002335 | 248.7453 | 0.26066 | 8.4017 | 0.00266 | 0.004491 | 0.004388 | 5.2e-05 | 0.001005 | 0.323 | 1.3274 | 2.5475 | 98.53 | 19.99 | 49.975 | 81.0292 | 85.1112 | 77.2815 | 0.011864 | 1.0075 | 0.023385 | 23.2116 | -0.0379 | 0.15507 | 0.005193 | 0.36027 | 0.36027 | 0.017346
2019-04-09 14:41 | 1554835282.5403 | 401 | 0.20845 | 0.001286 | 5.8e-05 | 100.8227 | 101.0869 | 117.8988 | 20.5303 | 0.002475 | 0.00228 | 248.7341 | 0.26126 | 8.4039 | 0.002739 | 0.004078 | 0.004069 | 4.9e-05 | 0.00101 | 0.26737 | 1.3346 | 2.5474 | 98.53 | 19.99 | 49.975 | 80.6614 | 87.0084 | 75.9675 | 0.011864 | 1.0113 | 0.023419 | 23.1587 | -0.023404 | 0.15543 | 0.005339 | 0.37094 | 0.37094 | 0.016886
2019-04-09 14:41 | 1554835283.5402 | 402 | 0.20847 | 0.001279 | 5.8e-05 | 100.7885 | 101.2703 | 117.9056 | 20.5296 | 0.002491 | 0.00233 | 248.714 | 0.26141 | 8.4008 | 0.002734 | 0.003963 | 0.003883 | 4.9e-05 | 0.001009 | 0.29192 | 1.3179 | 2.5473 | 98.53 | 19.99 | 49.974 | 81.0428 | 86.8835 | 77.1782 | 0.011863 | 1.0078 | 0.023379 | 23.1995 | -0.042659 | 0.15551 | 0.005222 | 0.36215 | 0.36215 | 0.017305
2019-04-09 14:41 | 1554835284.5397 | 403 | 0.20846 | 0.001283 | 5.8e-05 | 100.8385 | 101.1271 | 118.0377 | 20.5303 | 0.002513 | 0.002317 | 248.6929 | 0.26116 | 8.4048 | 0.002783 | 0.0047 | 0.004697 | 5e-05 | 0.001008 | 0.30198 | 1.3352 | 2.5476 | 98.53 | 20 | 49.975 | 81.5668 | 88.0541 | 76.8586 | 0.011871 | 1.0087 | 0.023403 | 23.2019 | -0.025566 | 0.15536 | 0.005242 | 0.36394 | 0.36394 | 0.017204
2019-04-09 14:41 | 1554835285.5393 | 404 | 0.20843 | 0.001273 | 5.8e-05 | 100.8471 | 101.1571 | 118.2722 | 20.5303 | 0.00255 | 0.002313 | 248.6671 | 0.26085 | 8.4019 | 0.002737 | 0.00415 | 0.004052 | 5e-05 | 0.00101 | 0.30055 | 1.3161 | 2.5476 | 98.53 | 20 | 49.976 | 82.4598 | 86.9562 | 76.7479 | 0.011871 | 1.009 | 0.02343 | 23.2213 | -0.027448 | 0.15518 | 0.005426 | 0.37722 | 0.37722 | 0.016579
2019-04-09 14:41 | 1554835286.5408 | 405 | 0.20846 | 0.001273 | 5.9e-05 | 100.8486 | 101.2734 | 118.088 | 20.5305 | 0.002486 | 0.002331 | 248.6298 | 0.26095 | 8.4009 | 0.002779 | 0.004024 | 0.004069 | 5.2e-05 | 0.001 | 0.2792 | 1.3313 | 2.5474 | 98.53 | 20 | 49.976 | 80.9238 | 87.9691 | 77.2009 | 0.011871 | 1.0077 | 0.023397 | 23.2182 | -0.037596 | 0.15524 | 0.005252 | 0.36449 | 0.36449 | 0.017164
2019-04-09 14:41 | 1554835287.5394 | 406 | 0.20847 | 0.001269 | 5.9e-05 | 100.8837 | 101.2737 | 118.44 | 20.5305 | 0.002484 | 0.002347 | 248.6661 | 0.2615 | 8.3996 | 0.002781 | 0.003779 | 0.003662 | 4.8e-05 | 0.000999 | 0.29446 | 1.3536 | 2.5476 | 98.53 | 20 | 49.976 | 80.8777 | 88.0004 | 77.5799 | 0.011871 | 1.0066 | 0.023419 | 23.2654 | -0.03451 | 0.15557 | 0.005227 | 0.36303 | 0.36303 | 0.017269
2019-04-09 14:41 | 1554835288.5395 | 407 | 0.20853 | 0.00127 | 6e-05 | 100.8677 | 101.3746 | 118.7705 | 20.5301 | 0.002464 | 0.002269 | 248.6412 | 0.26102 | 8.4001 | 0.002697 | 0.004344 | 0.004351 | 5e-05 | 0.000989 | 0.31686 | 1.3733 | 2.5476 | 98.53 | 20 | 49.975 | 80.3954 | 85.9942 | 75.6922 | 0.011871 | 1.012 | 0.023515 | 23.235 | -0.044837 | 0.15528 | 0.004847 | 0.33776 | 0.33776 | 0.018527
2019-04-09 14:41 | 1554835289.5385 | 408 | 0.20852 | 0.001269 | 6e-05 | 100.9019 | 101.2335 | 118.9013 | 20.5303 | 0.002453 | 0.002305 | 248.6508 | 0.2619 | 8.3999 | 0.002707 | 0.004455 | 0.004389 | 5.1e-05 | 0.000988 | 0.29048 | 1.3473 | 2.5475 | 98.53 | 20 | 49.973 | 80.1288 | 86.2362 | 76.5763 | 0.01187 | 1.0095 | 0.023498 | 23.2773 | -0.029342 | 0.1558 | 0.004879 | 0.33979 | 0.33979 | 0.018479
2019-04-09 14:41 | 1554835290.5384 | 409 | 0.20852 | 0.00126 | 6e-05 | 100.9195 | 101.3663 | 118.7337 | 20.5308 | 0.002483 | 0.002255 | 248.6153 | 0.2611 | 8.4005 | 0.002676 | 0.004335 | 0.004252 | 5.3e-05 | 0.000987 | 0.27671 | 1.3682 | 2.5474 | 98.53 | 20 | 49.971 | 80.8436 | 85.4921 | 75.3625 | 0.01187 | 1.013 | 0.023522 | 23.2205 | -0.039515 | 0.15533 | 0.004945 | 0.34472 | 0.34472 | 0.018159
2019-04-09 14:42 | 1554835291.5385 | 410 | 0.20853 | 0.001254 | 6e-05 | 100.9088 | 101.2937 | 118.9029 | 20.5302 | 0.002496 | 0.002302 | 248.6041 | 0.26157 | 8.4023 | 0.002668 | 0.004671 | 0.004636 | 5.2e-05 | 0.000992 | 0.28495 | 1.3674 | 2.5475 | 98.53 | 20 | 49.969 | 81.1566 | 85.2815 | 76.4967 | 0.01187 | 1.0097 | 0.023501 | 23.2748 | -0.034055 | 0.15561 | 0.004854 | 0.33798 | 0.33798 | 0.018554
2019-04-09 14:42 | 1554835292.5392 | 411 | 0.20854 | 0.001244 | 6e-05 | 100.8957 | 101.5126 | 118.7902 | 20.5306 | 0.00253 | 0.002309 | 248.5834 | 0.26093 | 8.4031 | 0.00264 | 0.004335 | 0.004273 | 5.1e-05 | 0.000991 | 0.29358 | 1.3887 | 2.5475 | 98.53 | 20 | 49.964 | 81.9851 | 84.6304 | 76.6741 | 0.011868 | 1.0092 | 0.023484 | 23.2696 | -0.054524 | 0.15523 | 0.004806 | 0.33436 | 0.33436 | 0.018709
2019-04-09 14:42 | 1554835293.5396 | 412 | 0.20855 | 0.001241 | 6.1e-05 | 100.9799 | 101.3849 | 118.8017 | 20.5303 | 0.002436 | 0.002301 | 248.5074 | 0.2616 | 8.402 | 0.00268 | 0.004322 | 0.004268 | 5.1e-05 | 0.000997 | 0.28672 | 1.36 | 2.5476 | 98.53 | 20 | 49.961 | 79.704 | 85.5901 | 76.4822 | 0.011868 | 1.0098 | 0.023492 | 23.2644 | -0.035798 | 0.15562 | 0.004771 | 0.33199 | 0.33199 | 0.018891
2019-04-09 14:42 | 1554835294.5396 | 413 | 0.20852 | 0.001235 | 6.1e-05 | 100.9138 | 101.2751 | 118.5734 | 20.5307 | 0.002453 | 0.002328 | 248.5103 | 0.26149 | 8.4022 | 0.002761 | 0.004124 | 0.00411 | 5e-05 | 0.001 | 0.30917 | 1.3328 | 2.5476 | 98.53 | 20 | 49.955 | 80.1153 | 87.5303 | 77.1248 | 0.011866 | 1.0079 | 0.023447 | 23.2634 | -0.031965 | 0.15556 | 0.004937 | 0.343 | 0.343 | 0.018277
2019-04-09 14:42 | 1554835295.5401 | 414 | 0.20851 | 0.001227 | 6.2e-05 | 100.9683 | 101.2982 | 118.4975 | 20.5303 | 0.002469 | 0.002279 | 248.4416 | 0.26058 | 8.403 | 0.002747 | 0.004133 | 0.004036 | 5.1e-05 | 0.000998 | 0.3011 | 1.3436 | 2.5476 | 98.53 | 20 | 49.952 | 80.5047 | 87.1888 | 75.9372 | 0.011866 | 1.0113 | 0.02348 | 23.2165 | -0.029173 | 0.15502 | 0.005014 | 0.34885 | 0.34885 | 0.017908
2019-04-09 14:42 | 1554835296.5387 | 415 | 0.20853 | 0.001222 | 6.2e-05 | 100.9841 | 101.314 | 118.8689 | 20.5309 | 0.002484 | 0.00233 | 248.4743 | 0.26098 | 8.4016 | 0.0027 | 0.003915 | 0.003893 | 5.1e-05 | 0.000997 | 0.27455 | 1.3249 | 2.5474 | 98.53 | 20 | 49.949 | 80.8634 | 86.0581 | 77.1701 | 0.011865 | 1.0078 | 0.023475 | 23.2938 | -0.029172 | 0.15526 | 0.004933 | 0.34313 | 0.34313 | 0.018235
2019-04-09 14:42 | 1554835297.5397 | 416 | 0.20851 | 0.001216 | 6.2e-05 | 100.9966 | 101.3845 | 118.9761 | 20.5309 | 0.002532 | 0.002271 | 248.4216 | 0.26068 | 8.4045 | 0.002708 | 0.003784 | 0.003744 | 4.8e-05 | 0.000997 | 0.27936 | 1.3161 | 2.5473 | 98.53 | 20 | 49.944 | 82.0112 | 86.251 | 75.7531 | 0.011864 | 1.0119 | 0.023533 | 23.2572 | -0.034291 | 0.15508 | 0.005034 | 0.35107 | 0.35107 | 0.017802
2019-04-09 14:42 | 1554835298.5388 | 417 | 0.20851 | 0.001219 | 6.4e-05 | 100.9558 | 101.3969 | 118.9484 | 20.5309 | 0.002507 | 0.002326 | 248.414 | 0.26139 | 8.4028 | 0.002759 | 0.004212 | 0.00407 | 5.5e-05 | 0.000998 | 0.25177 | 1.3484 | 2.5474 | 98.53 | 20 | 49.94 | 81.4326 | 87.4915 | 77.0798 | 0.011863 | 1.008 | 0.023486 | 23.2986 | -0.038998 | 0.1555 | 0.005045 | 0.35104 | 0.35104 | 0.017852
2019-04-09 14:42 | 1554835299.5398 | 418 | 0.20852 | 0.00122 | 6.4e-05 | 100.9567 | 101.3632 | 118.9956 | 20.531 | 0.002511 | 0.0023 | 248.3935 | 0.26116 | 8.4014 | 0.002772 | 0.004613 | 0.004547 | 4.8e-05 | 0.000996 | 0.28351 | 1.3363 | 2.5475 | 98.53 | 20 | 49.935 | 81.5129 | 87.8039 | 76.4578 | 0.011861 | 1.0098 | 0.023511 | 23.2826 | -0.035943 | 0.15536 | 0.005001 | 0.3483 | 0.3483 | 0.017976
2019-04-09 14:42 | 1554835300.5388 | 419 | 0.20849 | 0.001218 | 6.4e-05 | 100.9856 | 101.3467 | 118.6715 | 20.531 | 0.002511 | 0.002331 | 248.3705 | 0.26083 | 8.4033 | 0.002736 | 0.004319 | 0.004308 | 4.9e-05 | 0.000995 | 0.30198 | 1.3567 | 2.5476 | 98.53 | 20 | 49.931 | 81.5255 | 86.9179 | 77.1858 | 0.011861 | 1.0077 | 0.023455 | 23.275 | -0.031926 | 0.15517 | 0.005148 | 0.35775 | 0.35775 | 0.017479
2019-04-09 14:42 | 1554835301.5401 | 420 | 0.20849 | 0.001219 | 6.5e-05 | 100.9579 | 101.2345 | 118.6607 | 20.5308 | 0.002444 | 0.002311 | 248.3806 | 0.2604 | 8.4022 | 0.002675 | 0.004502 | 0.004479 | 4.8e-05 | 0.000998 | 0.28313 | 1.3813 | 2.5476 | 98.53 | 20 | 49.93 | 79.8972 | 85.45 | 76.7188 | 0.01186 | 1.0091 | 0.02347 | 23.2584 | -0.024472 | 0.15491 | 0.005165 | 0.35913 | 0.35913 | 0.017384
2019-04-09 14:42 | 1554835302.5389 | 421 | 0.20851 | 0.001215 | 6.5e-05 | 100.9269 | 101.3305 | 118.8339 | 20.5319 | 0.002466 | 0.00227 | 248.3192 | 0.26009 | 8.4004 | 0.002735 | 0.004579 | 0.004541 | 5.1e-05 | 0.000997 | 0.29463 | 1.3978 | 2.5473 | 98.53 | 20 | 49.924 | 80.4267 | 86.9045 | 75.7245 | 0.011859 | 1.012 | 0.02352 | 23.2423 | -0.035696 | 0.15473 | 0.005033 | 0.35064 | 0.35064 | 0.017783
2019-04-09 14:42 | 1554835303.5379 | 422 | 0.20851 | 0.001215 | 6.5e-05 | 100.9493 | 101.3333 | 118.4703 | 20.5313 | 0.002447 | 0.002298 | 248.3373 | 0.26171 | 8.3989 | 0.002678 | 0.003988 | 0.003917 | 5.4e-05 | 0.001001 | 0.27295 | 1.3782 | 2.5475 | 98.53 | 20 | 49.923 | 79.9709 | 85.5364 | 76.4091 | 0.011859 | 1.01 | 0.023461 | 23.2295 | -0.03396 | 0.15569 | 0.005031 | 0.34957 | 0.34957 | 0.017949
2019-04-09 14:42 | 1554835304.5389 | 423 | 0.20851 | 0.00121 | 6.5e-05 | 100.9281 | 101.3856 | 118.1429 | 20.5317 | 0.002479 | 0.002245 | 248.2887 | 0.26153 | 8.4019 | 0.002761 | 0.004615 | 0.004506 | 5e-05 | 0.001011 | 0.27759 | 1.3003 | 2.5475 | 98.53 | 20 | 49.922 | 80.7534 | 87.5375 | 75.1179 | 0.011858 | 1.0137 | 0.023472 | 23.1545 | -0.04045 | 0.15559 | 0.005022 | 0.34914 | 0.34914 | 0.017958
2019-04-09 14:42 | 1554835305.5392 | 424 | 0.20851 | 0.001208 | 6.5e-05 | 100.8656 | 101.3463 | 117.7845 | 20.5314 | 0.002504 | 0.00231 | 248.3086 | 0.26091 | 8.4011 | 0.002691 | 0.004459 | 0.004392 | 5e-05 | 0.001007 | 0.28407 | 1.2949 | 2.5475 | 98.53 | 20.01 | 49.919 | 81.358 | 85.855 | 76.6905 | 0.011865 | 1.0092 | 0.023384 | 23.1715 | -0.042532 | 0.15521 | 0.00503 | 0.34832 | 0.34832 | 0.017958
2019-04-09 14:42 | 1554835306.5381 | 425 | 0.20853 | 0.001206 | 6.6e-05 | 100.9433 | 101.2138 | 117.3415 | 20.5316 | 0.002464 | 0.002334 | 248.2461 | 0.26159 | 8.4035 | 0.002666 | 0.004583 | 0.004435 | 5.1e-05 | 0.001009 | 0.262 | 1.3034 | 2.5475 | 98.53 | 20.01 | 49.918 | 80.3836 | 85.2432 | 77.2652 | 0.011865 | 1.0075 | 0.023321 | 23.1468 | -0.023941 | 0.15562 | 0.004937 | 0.34088 | 0.34088 | 0.018398
2019-04-09 14:42 | 1554835307.5381 | 426 | 0.20854 | 0.001201 | 6.6e-05 | 100.9239 | 101.4076 | 117.2526 | 20.5311 | 0.002444 | 0.002339 | 248.2341 | 0.2603 | 8.4021 | 0.002719 | 0.004534 | 0.004534 | 4.9e-05 | 0.001014 | 0.28495 | 1.3362 | 2.5476 | 98.53 | 20.01 | 49.914 | 79.9037 | 86.5134 | 77.3855 | 0.011864 | 1.0072 | 0.023308 | 23.142 | -0.042762 | 0.15485 | 0.004894 | 0.33771 | 0.33771 | 0.018479
2019-04-09 14:42 | 1554835308.5385 | 427 | 0.2085 | 0.001203 | 6.7e-05 | 100.8992 | 101.3092 | 117.2777 | 20.5312 | 0.002476 | 0.002289 | 248.2105 | 0.26091 | 8.401 | 0.002694 | 0.004881 | 0.004784 | 5e-05 | 0.001011 | 0.30237 | 1.333 | 2.5475 | 98.53 | 20.01 | 49.912 | 80.6873 | 85.9209 | 76.175 | 0.011863 | 1.0107 | 0.023351 | 23.1045 | -0.036267 | 0.15521 | 0.005077 | 0.35105 | 0.35105 | 0.017818
2019-04-09 14:42 | 1554835309.5381 | 428 | 0.20856 | 0.001203 | 6.7e-05 | 100.9409 | 101.4028 | 117.3857 | 20.5306 | 0.002456 | 0.002284 | 248.1553 | 0.26094 | 8.4032 | 0.002738 | 0.004296 | 0.004218 | 4.7e-05 | 0.00101 | 0.28335 | 1.3215 | 2.5475 | 98.53 | 20.01 | 49.91 | 80.1854 | 86.9846 | 76.0598 | 0.011863 | 1.011 | 0.023365 | 23.1113 | -0.040833 | 0.15523 | 0.004735 | 0.32735 | 0.32735 | 0.019111
2019-04-09 14:42 | 1554835310.5393 | 429 | 0.20853 | 0.001199 | 6.7e-05 | 100.9231 | 101.2286 | 117.5662 | 20.5319 | 0.002444 | 0.00226 | 248.1447 | 0.2607 | 8.401 | 0.002728 | 0.00445 | 0.004493 | 4.6e-05 | 0.001014 | 0.28854 | 1.3337 | 2.5476 | 98.53 | 20.02 | 49.906 | 79.9005 | 86.7462 | 75.4821 | 0.011869 | 1.0127 | 0.023403 | 23.11 | -0.027043 | 0.15509 | 0.004921 | 0.34087 | 0.34087 | 0.018336
2019-04-09 14:42 | 1554835311.5392 | 430 | 0.20853 | 0.001199 | 6.8e-05 | 100.9147 | 101.2104 | 117.7213 | 20.5318 | 0.002408 | 0.002309 | 248.0815 | 0.2616 | 8.4006 | 0.002721 | 0.004207 | 0.004188 | 4.6e-05 | 0.001015 | 0.28407 | 1.3308 | 2.5475 | 98.53 | 20.02 | 49.903 | 79.0482 | 86.5789 | 76.6698 | 0.011869 | 1.0092 | 0.023378 | 23.1646 | -0.02617 | 0.15563 | 0.004934 | 0.34141 | 0.34141 | 0.01837
2019-04-09 14:42 | 1554835312.5396 | 431 | 0.20853 | 0.0012 | 6.8e-05 | 100.8823 | 101.4155 | 117.4864 | 20.5318 | 0.002481 | 0.002266 | 248.1079 | 0.26132 | 8.4028 | 0.00275 | 0.003937 | 0.003877 | 4.7e-05 | 0.001013 | 0.29374 | 1.3424 | 2.5476 | 98.53 | 20.02 | 49.896 | 80.7963 | 87.2582 | 75.6418 | 0.011867 | 1.0122 | 0.023389 | 23.1074 | -0.047156 | 0.15546 | 0.004929 | 0.34122 | 0.34122 | 0.018361
2019-04-09 14:42 | 1554835313.5387 | 432 | 0.20854 | 0.001193 | 6.8e-05 | 100.9359 | 101.3202 | 117.6554 | 20.5324 | 0.002433 | 0.002276 | 248.0813 | 0.26119 | 8.4034 | 0.002701 | 0.00403 | 0.003905 | 4.4e-05 | 0.001017 | 0.27671 | 1.3591 | 2.5475 | 98.53 | 20.02 | 49.893 | 79.6553 | 86.0915 | 75.8761 | 0.011866 | 1.0115 | 0.023398 | 23.1318 | -0.033993 | 0.15538 | 0.004859 | 0.33644 | 0.33644 | 0.018612
2019-04-09 14:42 | 1554835314.5393 | 433 | 0.20853 | 0.001193 | 6.9e-05 | 100.8727 | 101.3711 | 117.9847 | 20.5325 | 0.00242 | 0.002285 | 248.065 | 0.26085 | 8.4004 | 0.00264 | 0.004248 | 0.004344 | 4.6e-05 | 0.001014 | 0.30989 | 1.3265 | 2.5475 | 98.53 | 20.02 | 49.889 | 79.3272 | 84.6159 | 76.0852 | 0.011865 | 1.0109 | 0.023424 | 23.1711 | -0.044085 | 0.15518 | 0.004955 | 0.34347 | 0.34347 | 0.018208
2019-04-09 14:42 | 1554835315.5394 | 434 | 0.20855 | 0.001187 | 6.9e-05 | 100.9236 | 101.2854 | 118.1661 | 20.5325 | 0.002471 | 0.002354 | 248.0614 | 0.26167 | 8.4012 | 0.002697 | 0.004189 | 0.004105 | 4.7e-05 | 0.001018 | 0.31166 | 1.3634 | 2.5476 | 98.53 | 20.02 | 49.875 | 80.5623 | 85.9934 | 77.7432 | 0.011862 | 1.0061 | 0.023386 | 23.2439 | -0.032006 | 0.15567 | 0.004853 | 0.3358 | 0.3358 | 0.018682
2019-04-09 14:42 | 1554835316.5384 | 435 | 0.20857 | 0.001185 | 7e-05 | 100.8638 | 101.3636 | 118.4504 | 20.5321 | 0.00246 | 0.002279 | 248.0231 | 0.26111 | 8.4014 | 0.002783 | 0.003823 | 0.003763 | 4.4e-05 | 0.001021 | 0.2943 | 1.3176 | 2.5476 | 98.53 | 20.02 | 49.871 | 80.3046 | 88.0611 | 75.9508 | 0.011861 | 1.0113 | 0.023475 | 23.2123 | -0.044209 | 0.15534 | 0.004683 | 0.32503 | 0.32503 | 0.01926
2019-04-09 14:42 | 1554835317.5384 | 436 | 0.20854 | 0.001181 | 7.1e-05 | 100.8754 | 101.1226 | 118.2581 | 20.532 | 0.002502 | 0.002308 | 248.0287 | 0.26111 | 8.4027 | 0.002751 | 0.004158 | 0.00412 | 4.6e-05 | 0.001027 | 0.30071 | 1.3303 | 2.5475 | 98.53 | 20.02 | 49.865 | 81.3081 | 87.2886 | 76.6411 | 0.01186 | 1.0093 | 0.023432 | 23.2164 | -0.021896 | 0.15534 | 0.004863 | 0.33709 | 0.33709 | 0.018571
2019-04-09 14:42 | 1554835318.5384 | 437 | 0.20858 | 0.001186 | 7.1e-05 | 100.815 | 101.1075 | 117.9969 | 20.5317 | 0.002533 | 0.002315 | 247.9636 | 0.26057 | 8.4011 | 0.002739 | 0.004651 | 0.004594 | 4.7e-05 | 0.00103 | 0.28102 | 1.3616 | 2.5475 | 98.53 | 20.02 | 49.851 | 82.0386 | 86.9926 | 76.811 | 0.011856 | 1.0088 | 0.023401 | 23.1963 | -0.025918 | 0.15502 | 0.004675 | 0.32344 | 0.32344 | 0.019315
2019-04-09 14:42 | 1554835319.5375 | 438 | 0.20854 | 0.001177 | 7.2e-05 | 100.789 | 100.9158 | 117.5678 | 20.5317 | 0.002467 | 0.002288 | 247.9744 | 0.2614 | 8.3998 | 0.002634 | 0.003993 | 0.003976 | 4.4e-05 | 0.001029 | 0.28064 | 1.3162 | 2.5476 | 98.53 | 20.02 | 49.846 | 80.4647 | 84.4777 | 76.163 | 0.011855 | 1.0107 | 0.02338 | 23.1327 | -0.011242 | 0.15551 | 0.004892 | 0.33825 | 0.33825 | 0.018528
2019-04-09 14:42 | 1554835320.5378 | 439 | 0.20856 | 0.001175 | 7.2e-05 | 100.8114 | 101.2923 | 118.007 | 20.5325 | 0.002459 | 0.002283 | 247.9201 | 0.2611 | 8.4009 | 0.002713 | 0.003687 | 0.003589 | 4.4e-05 | 0.001025 | 0.28998 | 1.3121 | 2.5476 | 98.53 | 20.02 | 49.834 | 80.266 | 86.3837 | 76.0492 | 0.011852 | 1.011 | 0.023427 | 23.1721 | -0.042569 | 0.15532 | 0.004759 | 0.32958 | 0.32958 | 0.018993
2019-04-09 14:42 | 1554835321.5379 | 440 | 0.20854 | 0.001173 | 7.1e-05 | 100.753 | 101.1918 | 118.0565 | 20.5315 | 0.002473 | 0.002339 | 247.9299 | 0.26106 | 8.3988 | 0.002612 | 0.004032 | 0.004055 | 4.5e-05 | 0.001024 | 0.2625 | 1.3608 | 2.5476 | 98.53 | 20.02 | 49.828 | 80.6059 | 83.9346 | 77.3752 | 0.011851 | 1.0072 | 0.023388 | 23.2209 | -0.038871 | 0.15531 | 0.004897 | 0.33879 | 0.33879 | 0.018474
2019-04-09 14:42 | 1554835322.5385 | 441 | 0.20854 | 0.001172 | 7.2e-05 | 100.7164 | 101.1161 | 118.2438 | 20.5319 | 0.002475 | 0.002283 | 247.8966 | 0.26159 | 8.401 | 0.002615 | 0.004324 | 0.004176 | 4.5e-05 | 0.001022 | 0.28727 | 1.3394 | 2.5477 | 98.53 | 20.02 | 49.821 | 80.6646 | 84.0122 | 76.0486 | 0.011849 | 1.011 | 0.023451 | 23.1953 | -0.035429 | 0.15562 | 0.004901 | 0.3399 | 0.3399 | 0.018451
2019-04-09 14:42 | 1554835323.5376 | 442 | 0.20855 | 0.001171 | 7.2e-05 | 100.7569 | 101.0996 | 118.3288 | 20.5328 | 0.002449 | 0.002284 | 247.8576 | 0.26135 | 8.4007 | 0.002611 | 0.00407 | 0.004004 | 4.8e-05 | 0.001021 | 0.274 | 1.3266 | 2.5475 | 98.53 | 20.02 | 49.809 | 80.037 | 83.9125 | 76.0588 | 0.011846 | 1.011 | 0.023459 | 23.204 | -0.030371 | 0.15548 | 0.004871 | 0.33794 | 0.33794 | 0.018541
2019-04-09 14:42 | 1554835324.5376 | 443 | 0.20857 | 0.001167 | 7.2e-05 | 100.7137 | 101.1316 | 118.299 | 20.5319 | 0.002494 | 0.002284 | 247.8645 | 0.26086 | 8.4026 | 0.002633 | 0.003949 | 0.004001 | 4.4e-05 | 0.001022 | 0.28296 | 1.3393 | 2.5476 | 98.53 | 20.02 | 49.803 | 81.1013 | 84.4623 | 76.0642 | 0.011845 | 1.011 | 0.023456 | 23.2012 | -0.037036 | 0.15519 | 0.004701 | 0.32591 | 0.32591 | 0.019189
2019-04-09 14:42 | 1554835325.5389 | 444 | 0.20858 | 0.001171 | 7.2e-05 | 100.6926 | 100.9781 | 118.4849 | 20.5321 | 0.002399 | 0.002359 | 247.8045 | 0.2608 | 8.4059 | 0.002657 | 0.004198 | 0.004145 | 4.5e-05 | 0.00102 | 0.29446 | 1.2935 | 2.5475 | 98.53 | 20.02 | 49.789 | 78.8255 | 85.0229 | 77.8721 | 0.011841 | 1.0058 | 0.023414 | 23.2795 | -0.025328 | 0.15515 | 0.004643 | 0.32128 | 0.32128 | 0.019461
2019-04-09 14:42 | 1554835326.5387 | 445 | 0.20863 | 0.001165 | 7.2e-05 | 100.6812 | 101.2369 | 118.2257 | 20.5331 | 0.002526 | 0.002258 | 247.8376 | 0.26092 | 8.4009 | 0.002606 | 0.004344 | 0.004287 | 4.9e-05 | 0.001021 | 0.29159 | 1.2856 | 2.5476 | 98.53 | 20.02 | 49.784 | 81.8815 | 83.7972 | 75.4333 | 0.01184 | 1.0128 | 0.02347 | 23.1731 | -0.049229 | 0.15522 | 0.004393 | 0.30448 | 0.30448 | 0.020544
2019-04-09 14:42 | 1554835327.5387 | 446 | 0.2086 | 0.001165 | 7.3e-05 | 100.5949 | 101.1836 | 118.1506 | 20.5322 | 0.002423 | 0.00234 | 247.7942 | 0.26096 | 8.4034 | 0.002595 | 0.00405 | 0.00402 | 5e-05 | 0.00102 | 0.27223 | 1.3276 | 2.5476 | 98.53 | 20.03 | 49.78 | 79.4074 | 83.5297 | 77.4193 | 0.011847 | 1.0071 | 0.023396 | 23.2316 | -0.05219 | 0.15524 | 0.004559 | 0.31517 | 0.31517 | 0.019851
2019-04-09 14:42 | 1554835328.5378 | 447 | 0.20862 | 0.001165 | 7.3e-05 | 100.6541 | 100.8611 | 118.4281 | 20.5323 | 0.002504 | 0.002277 | 247.8102 | 0.26156 | 8.3995 | 0.002537 | 0.003568 | 0.003483 | 4.7e-05 | 0.001017 | 0.27831 | 1.3244 | 2.5474 | 98.53 | 20.03 | 49.766 | 81.3457 | 82.1475 | 75.8986 | 0.011843 | 1.0115 | 0.023474 | 23.2084 | -0.018371 | 0.1556 | 0.004451 | 0.30854 | 0.30854 | 0.020324
2019-04-09 14:42 | 1554835329.5378 | 448 | 0.20864 | 0.001157 | 7.3e-05 | 100.5955 | 100.9829 | 118.659 | 20.5332 | 0.002415 | 0.00227 | 247.7827 | 0.26138 | 8.403 | 0.002588 | 0.004121 | 0.004148 | 4.5e-05 | 0.001014 | 0.27975 | 1.2929 | 2.5473 | 98.53 | 20.03 | 49.762 | 79.2211 | 83.3769 | 75.7226 | 0.011842 | 1.012 | 0.023503 | 23.2251 | -0.034383 | 0.1555 | 0.00431 | 0.299 | 0.299 | 0.020958
2019-04-09 14:42 | 1554835330.5388 | 449 | 0.20863 | 0.001154 | 7.4e-05 | 100.5386 | 100.9743 | 118.3689 | 20.5328 | 0.002517 | 0.002312 | 247.7579 | 0.26132 | 8.3999 | 0.002588 | 0.003733 | 0.003646 | 4.4e-05 | 0.001011 | 0.27328 | 1.3097 | 2.5474 | 98.53 | 20.02 | 49.751 | 81.6526 | 83.3724 | 76.7418 | 0.011832 | 1.009 | 0.02344 | 23.2306 | -0.03868 | 0.15546 | 0.004361 | 0.30177 | 0.30177 | 0.020761
2019-04-09 14:42 | 1554835331.5388 | 450 | 0.20859 | 0.001149 | 7.3e-05 | 100.5285 | 100.7262 | 118.2715 | 20.5333 | 0.002443 | 0.002301 | 247.7461 | 0.26171 | 8.4034 | 0.002531 | 0.003565 | 0.003451 | 4.3e-05 | 0.001009 | 0.27704 | 1.3393 | 2.5474 | 98.53 | 20.02 | 49.745 | 79.8862 | 81.9995 | 76.4733 | 0.011831 | 1.0098 | 0.023439 | 23.2121 | -0.017573 | 0.15569 | 0.004628 | 0.32049 | 0.32049 | 0.019578
2019-04-09 14:42 | 1554835332.5378 | 451 | 0.20866 | 0.001146 | 7.4e-05 | 100.4511 | 100.8645 | 118.1126 | 20.5328 | 0.002417 | 0.002261 | 247.7073 | 0.26096 | 8.401 | 0.002614 | 0.004196 | 0.004164 | 4.7e-05 | 0.001004 | 0.27959 | 1.3236 | 2.5474 | 98.53 | 20.02 | 49.74 | 79.2707 | 83.989 | 75.5195 | 0.01183 | 1.0126 | 0.023456 | 23.1649 | -0.036742 | 0.15524 | 0.004198 | 0.29053 | 0.29053 | 0.021534
2019-04-09 14:42 | 1554835333.5379 | 452 | 0.20862 | 0.001144 | 7.4e-05 | 100.4945 | 100.7183 | 118.1142 | 20.5332 | 0.002477 | 0.002297 | 247.7202 | 0.26119 | 8.4019 | 0.002503 | 0.004592 | 0.004489 | 4.7e-05 | 0.001001 | 0.27544 | 1.3065 | 2.5473 | 98.53 | 20.02 | 49.732 | 80.7048 | 81.3187 | 76.3828 | 0.011828 | 1.01 | 0.023427 | 23.1937 | -0.019891 | 0.15538 | 0.004446 | 0.30759 | 0.30759 | 0.020358
2019-04-09 14:42 | 1554835334.5371 | 453 | 0.20865 | 0.001142 | 7.4e-05 | 100.4192 | 100.732 | 118.1206 | 20.533 | 0.002425 | 0.002281 | 247.6591 | 0.26078 | 8.4023 | 0.002583 | 0.004314 | 0.004223 | 4.2e-05 | 0.000999 | 0.28622 | 1.2948 | 2.5474 | 98.53 | 20.03 | 49.727 | 79.4641 | 83.2621 | 75.9875 | 0.011834 | 1.0112 | 0.023441 | 23.1812 | -0.027822 | 0.15514 | 0.004256 | 0.29439 | 0.29439 | 0.021237
2019-04-09 14:42 | 1554835335.5369 | 454 | 0.20863 | 0.00114 | 7.5e-05 | 100.3724 | 100.6422 | 117.8683 | 20.5334 | 0.002401 | 0.002271 | 247.6552 | 0.26175 | 8.4036 | 0.002586 | 0.004709 | 0.004692 | 4.3e-05 | 0.001 | 0.26576 | 1.3269 | 2.5476 | 98.53 | 20.03 | 49.725 | 78.8806 | 83.3202 | 75.757 | 0.011834 | 1.0119 | 0.023423 | 23.1488 | -0.024008 | 0.15572 | 0.00439 | 0.30351 | 0.30351 | 0.020676
2019-04-09 14:42 | 1554835336.5376 | 455 | 0.2086 | 0.001138 | 7.5e-05 | 100.3882 | 100.613 | 117.8593 | 20.5331 | 0.002403 | 0.002265 | 247.6275 | 0.26147 | 8.403 | 0.002595 | 0.004197 | 0.004171 | 4.5e-05 | 0.001003 | 0.26842 | 1.3684 | 2.5477 | 98.53 | 20.03 | 49.72 | 78.9334 | 83.5375 | 75.6204 | 0.011832 | 1.0123 | 0.023427 | 23.1434 | -0.020001 | 0.15555 | 0.004605 | 0.31864 | 0.31864 | 0.019673
2019-04-09 14:42 | 1554835337.5374 | 456 | 0.20858 | 0.001134 | 7.6e-05 | 100.3221 | 100.5606 | 118.0283 | 20.5342 | 0.002456 | 0.002302 | 247.5729 | 0.2602 | 8.4032 | 0.002581 | 0.004388 | 0.004285 | 4.3e-05 | 0.001007 | 0.29623 | 1.3323 | 2.5476 | 98.53 | 20.03 | 49.717 | 80.2028 | 83.2108 | 76.4991 | 0.011832 | 1.0097 | 0.023414 | 23.1891 | -0.021242 | 0.15479 | 0.004734 | 0.32748 | 0.32748 | 0.019049
2019-04-09 14:42 | 1554835338.538 | 457 | 0.20857 | 0.001133 | 7.6e-05 | 100.2378 | 100.5328 | 117.9563 | 20.5336 | 0.002418 | 0.002323 | 247.5846 | 0.26126 | 8.4043 | 0.002627 | 0.003752 | 0.003749 | 4.4e-05 | 0.001006 | 0.27936 | 1.345 | 2.5476 | 98.53 | 20.02 | 49.711 | 79.2847 | 84.3057 | 76.9973 | 0.011823 | 1.0083 | 0.023391 | 23.1985 | -0.026279 | 0.15543 | 0.004794 | 0.33133 | 0.33133 | 0.018905
2019-04-09 14:42 | 1554835339.5371 | 458 | 0.20857 | 0.00113 | 7.7e-05 | 100.2304 | 100.4509 | 117.8837 | 20.5328 | 0.002474 | 0.002278 | 247.5417 | 0.26145 | 8.4028 | 0.002637 | 0.004234 | 0.004224 | 4.4e-05 | 0.001005 | 0.27417 | 1.3041 | 2.5474 | 98.53 | 20.02 | 49.709 | 80.6318 | 84.5483 | 75.9108 | 0.011822 | 1.0114 | 0.02342 | 23.1554 | -0.019653 | 0.15554 | 0.004756 | 0.329 | 0.329 | 0.019052
2019-04-09 14:42 | 1554835340.5381 | 459 | 0.20863 | 0.001129 | 7.8e-05 | 100.1786 | 100.467 | 117.8114 | 20.5337 | 0.002412 | 0.002269 | 247.5518 | 0.26114 | 8.4026 | 0.002632 | 0.004174 | 0.004055 | 4.3e-05 | 0.001003 | 0.27024 | 1.3638 | 2.5476 | 98.53 | 20.03 | 49.707 | 79.1522 | 84.4345 | 75.695 | 0.011829 | 1.012 | 0.02342 | 23.1411 | -0.025718 | 0.15535 | 0.004379 | 0.30252 | 0.30252 | 0.020695
2019-04-09 14:42 | 1554835341.538 | 460 | 0.20863 | 0.001131 | 7.8e-05 | 100.1312 | 100.3428 | 117.8086 | 20.5336 | 0.002463 | 0.002242 | 247.4921 | 0.26195 | 8.3999 | 0.002604 | 0.004084 | 0.00402 | 4.1e-05 | 0.000994 | 0.29159 | 1.3464 | 2.5475 | 98.53 | 20.03 | 49.702 | 80.3731 | 83.7619 | 75.0525 | 0.011828 | 1.0139 | 0.023441 | 23.1195 | -0.018879 | 0.15583 | 0.004409 | 0.30489 | 0.30489 | 0.020598
2019-04-09 14:42 | 1554835342.5372 | 461 | 0.20863 | 0.001126 | 7.9e-05 | 100.1128 | 100.3263 | 117.9919 | 20.5334 | 0.002418 | 0.002265 | 247.4781 | 0.26068 | 8.4025 | 0.002554 | 0.004423 | 0.004347 | 4e-05 | 0.000989 | 0.28816 | 1.3586 | 2.5475 | 98.53 | 20.03 | 49.701 | 79.2828 | 82.5611 | 75.5996 | 0.011828 | 1.0123 | 0.023441 | 23.1557 | -0.019056 | 0.15508 | 0.004389 | 0.30341 | 0.30341 | 0.020598
2019-04-09 14:42 | 1554835343.5386 | 462 | 0.20864 | 0.001128 | 8e-05 | 100.0883 | 100.0602 | 118.0288 | 20.5329 | 0.002457 | 0.002273 | 247.4524 | 0.26178 | 8.4021 | 0.00262 | 0.004554 | 0.004525 | 4.2e-05 | 0.000987 | 0.28854 | 1.3366 | 2.5474 | 98.53 | 20.03 | 49.7 | 80.2234 | 84.1341 | 75.805 | 0.011828 | 1.0117 | 0.023438 | 23.1661 | 0.002512 | 0.15573 | 0.004337 | 0.29966 | 0.29966 | 0.020944
2019-04-09 14:42 | 1554835344.5386 | 463 | 0.20861 | 0.001125 | 8e-05 | 99.9794 | 100.1508 | 117.9821 | 20.5337 | 0.002503 | 0.002298 | 247.4174 | 0.2616 | 8.4033 | 0.002591 | 0.00401 | 0.004016 | 3.9e-05 | 0.000986 | 0.28174 | 1.335 | 2.5474 | 98.53 | 20.03 | 49.699 | 81.3345 | 83.4452 | 76.4052 | 0.011827 | 1.01 | 0.023413 | 23.1814 | -0.015321 | 0.15563 | 0.00453 | 0.31291 | 0.31291 | 0.020044
2019-04-09 14:42 | 1554835345.5382 | 464 | 0.20862 | 0.001124 | 8.1e-05 | 99.9969 | 100.3232 | 118.2736 | 20.5334 | 0.002373 | 0.002276 | 247.4114 | 0.26208 | 8.4029 | 0.002603 | 0.003671 | 0.003604 | 4.1e-05 | 0.000986 | 0.28567 | 1.3378 | 2.5475 | 98.53 | 20.03 | 49.698 | 78.2148 | 83.7258 | 75.8716 | 0.011827 | 1.0115 | 0.02346 | 23.1923 | -0.029135 | 0.15591 | 0.004473 | 0.3095 | 0.3095 | 0.020301
2019-04-09 14:42 | 1554835346.5373 | 465 | 0.20862 | 0.001123 | 8.1e-05 | 99.9225 | 100.1638 | 118.3907 | 20.5329 | 0.002412 | 0.002252 | 247.3526 | 0.26245 | 8.3994 | 0.002621 | 0.004115 | 0.004057 | 4.1e-05 | 0.000986 | 0.29407 | 1.3358 | 2.5474 | 98.53 | 20.03 | 49.694 | 79.1374 | 84.1615 | 75.2882 | 0.011826 | 1.0132 | 0.023491 | 23.1844 | -0.021579 | 0.15613 | 0.004486 | 0.31079 | 0.31079 | 0.020246
2019-04-09 14:42 | 1554835347.5383 | 466 | 0.20863 | 0.001122 | 8.1e-05 | 99.8581 | 100.1483 | 118.4783 | 20.534 | 0.002443 | 0.002326 | 247.3796 | 0.26121 | 8.4031 | 0.002686 | 0.004413 | 0.004299 | 4.2e-05 | 0.000986 | 0.27632 | 1.339 | 2.5475 | 98.53 | 20.03 | 49.692 | 79.89 | 85.7272 | 77.0648 | 0.011826 | 1.0081 | 0.02344 | 23.252 | -0.025956 | 0.1554 | 0.004433 | 0.30641 | 0.30641 | 0.020438
2019-04-09 14:42 | 1554835348.5384 | 467 | 0.20859 | 0.001121 | 8.1e-05 | 99.8331 | 100.0148 | 118.3717 | 20.533 | 0.002429 | 0.002272 | 247.325 | 0.26164 | 8.4021 | 0.002673 | 0.004615 | 0.004503 | 4.4e-05 | 0.000988 | 0.27201 | 1.3225 | 2.5475 | 98.53 | 20.03 | 49.689 | 79.5492 | 85.4202 | 75.7659 | 0.011825 | 1.0118 | 0.023473 | 23.1984 | -0.016263 | 0.15565 | 0.004665 | 0.32308 | 0.32308 | 0.019415
2019-04-09 14:42 | 1554835349.5384 | 468 | 0.20861 | 0.001114 | 8.1e-05 | 99.7381 | 100.0957 | 118.3936 | 20.5334 | 0.002447 | 0.002305 | 247.3131 | 0.26184 | 8.4007 | 0.002666 | 0.004166 | 0.004231 | 4.1e-05 | 0.000994 | 0.28838 | 1.3557 | 2.5475 | 98.53 | 20.03 | 49.687 | 79.9751 | 85.2481 | 76.5615 | 0.011825 | 1.0095 | 0.023448 | 23.227 | -0.032008 | 0.15577 | 0.004554 | 0.31499 | 0.31499 | 0.01993
2019-04-09 14:42 | 1554835350.5378 | 469 | 0.20862 | 0.001119 | 8.2e-05 | 99.6643 | 100.0127 | 118.3302 | 20.5334 | 0.002459 | 0.002247 | 247.2998 | 0.26184 | 8.3998 | 0.002584 | 0.004507 | 0.004427 | 4e-05 | 0.000999 | 0.24801 | 1.3677 | 2.5476 | 98.53 | 20.03 | 49.685 | 80.2638 | 83.2834 | 75.1763 | 0.011824 | 1.0135 | 0.023489 | 23.1748 | -0.031219 | 0.15577 | 0.004478 | 0.31018 | 0.31018 | 0.020238
2019-04-09 14:42 | 1554835351.5365 | 470 | 0.20864 | 0.001117 | 8.2e-05 | 99.6053 | 99.8228 | 118.2959 | 20.5339 | 0.002446 | 0.002262 | 247.2768 | 0.26079 | 8.4031 | 0.002627 | 0.004307 | 0.004295 | 4e-05 | 0.001003 | 0.27096 | 1.3177 | 2.5476 | 98.53 | 20.03 | 49.685 | 79.9485 | 84.3131 | 75.5456 | 0.011824 | 1.0125 | 0.023473 | 23.1837 | -0.019506 | 0.15514 | 0.004381 | 0.30309 | 0.30309 | 0.020629
2019-04-09 14:43 | 1554835352.5365 | 471 | 0.20865 | 0.001115 | 8.2e-05 | 99.6172 | 99.8355 | 118.3612 | 20.5338 | 0.002467 | 0.002272 | 247.3001 | 0.26154 | 8.4018 | 0.002638 | 0.004557 | 0.004497 | 4.3e-05 | 0.001001 | 0.27312 | 1.339 | 2.5474 | 98.53 | 20.03 | 49.684 | 80.4521 | 84.5719 | 75.7777 | 0.011824 | 1.0118 | 0.023472 | 23.1978 | -0.019577 | 0.15559 | 0.004321 | 0.29886 | 0.29886 | 0.02098
2019-04-09 14:43 | 1554835353.5375 | 472 | 0.20864 | 0.001115 | 8.3e-05 | 99.4957 | 99.7598 | 118.3006 | 20.5347 | 0.002495 | 0.002293 | 247.2576 | 0.26183 | 8.4012 | 0.002617 | 0.003968 | 0.003905 | 4.2e-05 | 0.001004 | 0.30735 | 1.2981 | 2.5474 | 98.53 | 20.03 | 49.683 | 81.1468 | 84.0781 | 76.2721 | 0.011824 | 1.0104 | 0.023449 | 23.2083 | -0.023707 | 0.15576 | 0.004376 | 0.30246 | 0.30246 | 0.020754
2019-04-09 14:43 | 1554835354.5375 | 473 | 0.20863 | 0.00111 | 8.2e-05 | 99.3888 | 99.5709 | 118.2431 | 20.5331 | 0.002465 | 0.002312 | 247.2861 | 0.26228 | 8.4028 | 0.002642 | 0.003674 | 0.00363 | 4e-05 | 0.00101 | 0.2891 | 1.3126 | 2.5474 | 98.53 | 20.03 | 49.681 | 80.4122 | 84.6773 | 76.7296 | 0.011823 | 1.009 | 0.023428 | 23.2178 | -0.016366 | 0.15603 | 0.004453 | 0.30758 | 0.30758 | 0.020444
2019-04-09 14:43 | 1554835355.5366 | 474 | 0.20863 | 0.001111 | 8.3e-05 | 99.4234 | 99.5034 | 118.1647 | 20.5339 | 0.00238 | 0.002311 | 247.2659 | 0.26159 | 8.4016 | 0.002648 | 0.003852 | 0.003798 | 3.9e-05 | 0.001009 | 0.2724 | 1.3044 | 2.5476 | 98.53 | 20.03 | 49.68 | 78.3681 | 84.8046 | 76.7187 | 0.011823 | 1.0091 | 0.023421 | 23.2098 | -0.007197 | 0.15562 | 0.004399 | 0.30362 | 0.30362 | 0.020656
2019-04-09 14:43 | 1554835356.5376 | 475 | 0.20864 | 0.001107 | 8.4e-05 | 99.2748 | 99.1709 | 118.3558 | 20.534 | 0.002418 | 0.002281 | 247.26 | 0.26217 | 8.403 | 0.002651 | 0.004489 | 0.00446 | 3.9e-05 | 0.001011 | 0.25105 | 1.298 | 2.5475 | 98.53 | 20.03 | 49.679 | 79.2914 | 84.8867 | 75.9892 | 0.011823 | 1.0112 | 0.023464 | 23.2043 | 0.00936 | 0.15597 | 0.004345 | 0.30037 | 0.30037 | 0.020926
2019-04-09 14:43 | 1554835357.5377 | 476 | 0.20863 | 0.001106 | 8.4e-05 | 99.231 | 99.1431 | 118.0464 | 20.5344 | 0.002457 | 0.002358 | 247.2273 | 0.26144 | 8.3988 | 0.002614 | 0.003464 | 0.003543 | 4e-05 | 0.001008 | 0.26034 | 1.3349 | 2.5475 | 98.53 | 20.03 | 49.679 | 80.2304 | 83.9983 | 77.8424 | 0.011823 | 1.0058 | 0.023371 | 23.2354 | 0.007929 | 0.15553 | 0.004446 | 0.30623 | 0.30623 | 0.020468
2019-04-09 14:43 | 1554835358.5366 | 477 | 0.20863 | 0.00111 | 8.5e-05 | 99.0744 | 99.0591 | 118.2694 | 20.5341 | 0.002437 | 0.00227 | 247.1892 | 0.26206 | 8.4007 | 0.002589 | 0.003753 | 0.003637 | 4.4e-05 | 0.001008 | 0.2677 | 1.3375 | 2.5475 | 98.53 | 20.03 | 49.68 | 79.7444 | 83.4004 | 75.7398 | 0.011823 | 1.0119 | 0.023464 | 23.1875 | 0.001378 | 0.1559 | 0.004418 | 0.30542 | 0.30542 | 0.020571
2019-04-09 14:43 | 1554835359.5377 | 478 | 0.20864 | 0.001102 | 8.5e-05 | 99.0511 | 98.9307 | 118.4907 | 20.5345 | 0.00242 | 0.002243 | 247.2179 | 0.26174 | 8.4017 | 0.002647 | 0.00397 | 0.003905 | 4.3e-05 | 0.00101 | 0.27024 | 1.3402 | 2.5474 | 98.53 | 20.03 | 49.681 | 79.3228 | 84.8003 | 75.0915 | 0.011823 | 1.0138 | 0.023508 | 23.1877 | 0.01088 | 0.15571 | 0.0044 | 0.30469 | 0.30469 | 0.020595
2019-04-09 14:43 | 1554835360.5371 | 479 | 0.2086 | 0.001103 | 8.5e-05 | 99.0029 | 98.9824 | 118.5837 | 20.5339 | 0.002429 | 0.002266 | 247.1681 | 0.26138 | 8.4009 | 0.002593 | 0.00415 | 0.004074 | 3.7e-05 | 0.001009 | 0.25409 | 1.3828 | 2.5476 | 98.53 | 20.03 | 49.681 | 79.5463 | 83.4986 | 75.6244 | 0.011823 | 1.0122 | 0.023499 | 23.2145 | 0.001856 | 0.1555 | 0.00462 | 0.32009 | 0.32009 | 0.019577
2019-04-09 14:43 | 1554835361.5368 | 480 | 0.20868 | 0.001102 | 8.6e-05 | 98.8314 | 98.9053 | 118.545 | 20.534 | 0.002484 | 0.00228 | 247.1844 | 0.261 | 8.4036 | 0.00261 | 0.00382 | 0.003764 | 3.8e-05 | 0.001009 | 0.28241 | 1.3951 | 2.5476 | 98.53 | 20.03 | 49.683 | 80.8748 | 83.8964 | 75.9798 | 0.011824 | 1.0112 | 0.023483 | 23.2225 | -0.006685 | 0.15527 | 0.004144 | 0.28642 | 0.28642 | 0.021847
2019-04-09 14:43 | 1554835362.5368 | 481 | 0.20863 | 0.001102 | 8.5e-05 | 98.7912 | 98.8729 | 118.7614 | 20.5346 | 0.002393 | 0.002281 | 247.1124 | 0.26195 | 8.3983 | 0.002648 | 0.004592 | 0.004564 | 4.1e-05 | 0.001008 | 0.27201 | 1.3757 | 2.5476 | 98.53 | 20.03 | 49.684 | 78.69 | 84.8036 | 75.9953 | 0.011824 | 1.0112 | 0.023504 | 23.2442 | -0.007398 | 0.15584 | 0.004466 | 0.30932 | 0.30932 | 0.020303
2019-04-09 14:43 | 1554835363.5382 | 482 | 0.20865 | 0.0011 | 8.6e-05 | 98.6616 | 98.5305 | 118.8834 | 20.5339 | 0.002453 | 0.002268 | 247.1036 | 0.26188 | 8.4031 | 0.002603 | 0.003526 | 0.003482 | 3.6e-05 | 0.001011 | 0.30359 | 1.3145 | 2.5475 | 98.53 | 20.03 | 49.684 | 80.127 | 83.7332 | 75.6702 | 0.011824 | 1.0121 | 0.023527 | 23.2453 | 0.011897 | 0.15579 | 0.004306 | 0.29836 | 0.29836 | 0.021044
2019-04-09 14:43 | 1554835364.5379 | 483 | 0.20861 | 0.001102 | 8.6e-05 | 98.6298 | 98.5088 | 118.8755 | 20.5341 | 0.002461 | 0.002336 | 247.089 | 0.26193 | 8.4047 | 0.002722 | 0.004228 | 0.004139 | 4e-05 | 0.001017 | 0.26593 | 1.3111 | 2.5476 | 98.53 | 20.03 | 49.684 | 80.3214 | 86.6006 | 77.3183 | 0.011824 | 1.0074 | 0.023471 | 23.2994 | 0.010977 | 0.15582 | 0.004543 | 0.31421 | 0.31421 | 0.019985
2019-04-09 14:43 | 1554835365.5359 | 484 | 0.20866 | 0.001096 | 8.6e-05 | 98.5443 | 98.5091 | 118.8245 | 20.5344 | 0.002475 | 0.00232 | 247.0383 | 0.26214 | 8.402 | 0.002592 | 0.004099 | 0.004105 | 4.1e-05 | 0.001022 | 0.27831 | 1.3438 | 2.5475 | 98.53 | 20.03 | 49.685 | 80.6469 | 83.4755 | 76.9227 | 0.011824 | 1.0085 | 0.023479 | 23.2813 | 0.003191 | 0.15594 | 0.004249 | 0.29367 | 0.29367 | 0.0214
2019-04-09 14:43 | 1554835366.537 | 485 | 0.20868 | 0.001094 | 8.6e-05 | 98.4335 | 98.4482 | 118.4405 | 20.535 | 0.002431 | 0.002268 | 247.0284 | 0.26146 | 8.4018 | 0.002662 | 0.004189 | 0.004058 | 3.9e-05 | 0.001029 | 0.274 | 1.3425 | 2.5474 | 98.53 | 20.03 | 49.683 | 79.5948 | 85.1551 | 75.6845 | 0.011824 | 1.0121 | 0.023483 | 23.2025 | -0.001336 | 0.15554 | 0.004139 | 0.28597 | 0.28597 | 0.021919
2019-04-09 14:43 | 1554835367.5369 | 486 | 0.20868 | 0.001094 | 8.7e-05 | 98.346 | 98.124 | 118.7412 | 20.5345 | 0.002405 | 0.002249 | 247.0046 | 0.26164 | 8.4017 | 0.002676 | 0.004181 | 0.004186 | 3.9e-05 | 0.001028 | 0.26808 | 1.3275 | 2.5475 | 98.53 | 20.03 | 49.683 | 78.9636 | 85.4918 | 75.2172 | 0.011824 | 1.0134 | 0.023528 | 23.2163 | 0.020207 | 0.15565 | 0.004122 | 0.28534 | 0.28534 | 0.021983
2019-04-09 14:43 | 1554835368.5359 | 487 | 0.20866 | 0.001091 | 8.7e-05 | 98.1658 | 98.006 | 118.6623 | 20.535 | 0.002458 | 0.002267 | 247.0132 | 0.26113 | 8.4027 | 0.002641 | 0.003843 | 0.003742 | 3.9e-05 | 0.00103 | 0.28031 | 1.3317 | 2.5475 | 98.53 | 20.03 | 49.682 | 80.2386 | 84.6387 | 75.6487 | 0.011823 | 1.0122 | 0.023506 | 23.223 | 0.014573 | 0.15535 | 0.004246 | 0.29372 | 0.29372 | 0.021315
2019-04-09 14:43 | 1554835369.5364 | 488 | 0.20868 | 0.00109 | 8.8e-05 | 98.0425 | 98.07 | 118.3004 | 20.536 | 0.00244 | 0.002214 | 246.9429 | 0.26137 | 8.4024 | 0.002649 | 0.004208 | 0.004143 | 3.7e-05 | 0.001032 | 0.28456 | 1.3119 | 2.5476 | 98.53 | 20.03 | 49.681 | 79.8241 | 84.8292 | 74.39 | 0.011823 | 1.0158 | 0.023512 | 23.1457 | -0.002507 | 0.15549 | 0.004172 | 0.28855 | 0.28855 | 0.021716
2019-04-09 14:43 | 1554835370.5364 | 489 | 0.2087 | 0.001088 | 8.8e-05 | 97.9886 | 97.6488 | 118.2304 | 20.5355 | 0.002436 | 0.002267 | 246.9594 | 0.26079 | 8.4034 | 0.002701 | 0.003892 | 0.00388 | 3.5e-05 | 0.001027 | 0.27577 | 1.3017 | 2.5475 | 98.53 | 20.03 | 49.68 | 79.727 | 86.0861 | 75.664 | 0.011823 | 1.0121 | 0.023462 | 23.1812 | 0.031077 | 0.15514 | 0.004004 | 0.27616 | 0.27616 | 0.02264
2019-04-09 14:43 | 1554835371.5361 | 490 | 0.20867 | 0.001087 | 8.9e-05 | 97.793 | 97.6453 | 118.3781 | 20.5351 | 0.002421 | 0.002259 | 246.9229 | 0.262 | 8.3975 | 0.002562 | 0.004293 | 0.004267 | 4.2e-05 | 0.001021 | 0.27759 | 1.306 | 2.5474 | 98.53 | 20.03 | 49.68 | 79.3516 | 82.7477 | 75.4727 | 0.011823 | 1.0127 | 0.023484 | 23.1893 | 0.013515 | 0.15586 | 0.004181 | 0.28878 | 0.28878 | 0.021751
2019-04-09 14:43 | 1554835372.5375 | 491 | 0.20868 | 0.00109 | 8.9e-05 | 97.6846 | 97.3913 | 118.4734 | 20.5351 | 0.002437 | 0.002277 | 246.9104 | 0.26132 | 8.4025 | 0.002585 | 0.004085 | 0.004021 | 3.8e-05 | 0.001017 | 0.28512 | 1.3464 | 2.5475 | 98.53 | 20.03 | 49.68 | 79.7538 | 83.2961 | 75.8912 | 0.011823 | 1.0115 | 0.023479 | 23.2126 | 0.026891 | 0.15546 | 0.004132 | 0.28524 | 0.28524 | 0.021963
2019-04-09 14:43 | 1554835373.5375 | 492 | 0.20862 | 0.001087 | 9e-05 | 97.6015 | 97.2988 | 118.6107 | 20.5354 | 0.002449 | 0.002269 | 246.9373 | 0.26157 | 8.4025 | 0.002605 | 0.004013 | 0.003853 | 3.7e-05 | 0.001015 | 0.27704 | 1.3154 | 2.5474 | 98.53 | 20.03 | 49.681 | 80.0223 | 83.7934 | 75.7099 | 0.011823 | 1.012 | 0.023499 | 23.22 | 0.027787 | 0.15561 | 0.004515 | 0.31246 | 0.31246 | 0.02007
2019-04-09 14:43 | 1554835374.5361 | 493 | 0.20868 | 0.001086 | 8.9e-05 | 97.416 | 97.3411 | 118.8485 | 20.5356 | 0.002454 | 0.002173 | 246.8751 | 0.26145 | 8.4 | 0.002587 | 0.004501 | 0.004417 | 4e-05 | 0.001016 | 0.27721 | 1.3485 | 2.5476 | 98.53 | 20.03 | 49.68 | 80.1528 | 83.338 | 73.395 | 0.011823 | 1.0188 | 0.023601 | 23.166 | 0.006878 | 0.15553 | 0.004143 | 0.2875 | 0.2875 | 0.021802
2019-04-09 14:43 | 1554835375.5362 | 494 | 0.2087 | 0.001083 | 8.9e-05 | 97.2936 | 97.0943 | 118.9156 | 20.5357 | 0.002438 | 0.002266 | 246.8935 | 0.26213 | 8.3993 | 0.002587 | 0.004401 | 0.004345 | 4.2e-05 | 0.001012 | 0.28263 | 1.3122 | 2.5474 | 98.53 | 20.03 | 49.68 | 79.763 | 83.3599 | 75.633 | 0.011823 | 1.0122 | 0.023531 | 23.2473 | 0.018336 | 0.15594 | 0.004031 | 0.27879 | 0.27879 | 0.022542
2019-04-09 14:43 | 1554835376.5366 | 495 | 0.20861 | 0.001083 | 9e-05 | 97.1176 | 97.0165 | 119.1477 | 20.5353 | 0.002434 | 0.002296 | 246.8603 | 0.26186 | 8.4027 | 0.002602 | 0.003928 | 0.003883 | 4e-05 | 0.001011 | 0.27472 | 1.3248 | 2.5474 | 98.53 | 20.03 | 49.681 | 79.6832 | 83.7065 | 76.3513 | 0.011823 | 1.0101 | 0.02353 | 23.2939 | 0.00931 | 0.15578 | 0.004543 | 0.31479 | 0.31479 | 0.019943
2019-04-09 14:43 | 1554835377.5378 | 496 | 0.20868 | 0.001086 | 8.9e-05 | 96.9788 | 96.6562 | 119.1721 | 20.5368 | 0.002421 | 0.002293 | 246.823 | 0.26116 | 8.4034 | 0.002581 | 0.00433 | 0.00433 | 4.1e-05 | 0.001021 | 0.28584 | 1.3739 | 2.5476 | 98.53 | 20.03 | 49.68 | 79.3674 | 83.2143 | 76.2881 | 0.011823 | 1.0103 | 0.023535 | 23.2942 | 0.029802 | 0.15536 | 0.004137 | 0.28632 | 0.28632 | 0.021868
2019-04-09 14:43 | 1554835378.5366 | 497 | 0.20866 | 0.00108 | 9e-05 | 96.7963 | 96.565 | 119.07 | 20.5356 | 0.002403 | 0.002298 | 246.8028 | 0.26185 | 8.4012 | 0.002569 | 0.004452 | 0.004403 | 4.1e-05 | 0.00102 | 0.28241 | 1.4142 | 2.5476 | 98.53 | 20.03 | 49.68 | 78.9248 | 82.9135 | 76.4008 | 0.011823 | 1.01 | 0.023521 | 23.2879 | 0.021394 | 0.15577 | 0.004255 | 0.29443 | 0.29443 | 0.021322
2019-04-09 14:43 | 1554835379.5377 | 498 | 0.2087 | 0.001079 | 9e-05 | 96.6253 | 96.2312 | 119.139 | 20.535 | 0.002473 | 0.002242 | 246.7806 | 0.26126 | 8.4009 | 0.002619 | 0.003771 | 0.003801 | 3.9e-05 | 0.001019 | 0.2656 | 1.3841 | 2.5475 | 98.53 | 20.03 | 49.681 | 80.6068 | 84.1111 | 75.0524 | 0.011823 | 1.0139 | 0.023573 | 23.2497 | 0.03656 | 0.15542 | 0.004024 | 0.27872 | 0.27872 | 0.022472
2019-04-09 14:43 | 1554835380.5367 | 499 | 0.20864 | 0.001075 | 9.1e-05 | 96.3981 | 96.0839 | 119.5078 | 20.5359 | 0.002439 | 0.002292 | 246.7829 | 0.2618 | 8.399 | 0.002599 | 0.003915 | 0.00383 | 4.3e-05 | 0.001022 | 0.28224 | 1.3641 | 2.5476 | 98.53 | 20.03 | 49.681 | 79.7914 | 83.6292 | 76.2567 | 0.011823 | 1.0104 | 0.023569 | 23.3259 | 0.029203 | 0.15574 | 0.004409 | 0.30581 | 0.30581 | 0.020524
2019-04-09 14:43 | 1554835381.5354 | 500 | 0.20865 | 0.001075 | 9.2e-05 | 96.2102 | 96.0312 | 119.6744 | 20.5355 | 0.002382 | 0.002292 | 246.7537 | 0.26162 | 8.4024 | 0.002613 | 0.005452 | 0.005443 | 4.3e-05 | 0.001017 | 0.25874 | 1.3636 | 2.5475 | 98.53 | 20.03 | 49.679 | 78.4273 | 83.9836 | 76.2531 | 0.011823 | 1.0104 | 0.023585 | 23.342 | 0.016654 | 0.15564 | 0.00435 | 0.30182 | 0.30182 | 0.020782
2019-04-09 14:43 | 1554835382.5354 | 501 | 0.20865 | 0.001075 | 9.3e-05 | 96.0947 | 95.7528 | 119.4759 | 20.5355 | 0.002414 | 0.002275 | 246.7647 | 0.26178 | 8.4019 | 0.002618 | 0.004172 | 0.004126 | 4.2e-05 | 0.001016 | 0.28385 | 1.3283 | 2.5474 | 98.53 | 20.03 | 49.678 | 79.2022 | 84.1 | 75.8589 | 0.011822 | 1.0116 | 0.023579 | 23.3095 | 0.031876 | 0.15574 | 0.004341 | 0.30105 | 0.30105 | 0.020848
2019-04-09 14:43 | 1554835383.5366 | 502 | 0.20868 | 0.001075 | 9.3e-05 | 95.9327 | 95.5019 | 119.1735 | 20.5358 | 0.002418 | 0.002284 | 246.6907 | 0.26164 | 8.3993 | 0.002589 | 0.004012 | 0.003956 | 4.1e-05 | 0.001017 | 0.27024 | 1.309 | 2.5475 | 98.53 | 20.03 | 49.678 | 79.2977 | 83.3934 | 76.0641 | 0.011822 | 1.011 | 0.023542 | 23.2868 | 0.040251 | 0.15565 | 0.004151 | 0.2872 | 0.2872 | 0.021841
2019-04-09 14:43 | 1554835384.5365 | 503 | 0.2087 | 0.001076 | 9.3e-05 | 95.7564 | 95.1013 | 119.2611 | 20.5358 | 0.002385 | 0.002237 | 246.698 | 0.26156 | 8.4013 | 0.002628 | 0.00457 | 0.004517 | 4.3e-05 | 0.001019 | 0.24657 | 1.336 | 2.5475 | 98.53 | 20.03 | 49.673 | 78.5026 | 84.3235 | 74.9372 | 0.011821 | 1.0142 | 0.023589 | 23.2578 | 0.061397 | 0.1556 | 0.004058 | 0.28118 | 0.28118 | 0.022302
2019-04-09 14:43 | 1554835385.5369 | 504 | 0.2087 | 0.001078 | 9.5e-05 | 95.6471 | 95.0394 | 119.2435 | 20.5353 | 0.00242 | 0.002316 | 246.6591 | 0.26178 | 8.4021 | 0.002635 | 0.003279 | 0.003136 | 3.9e-05 | 0.001024 | 0.28495 | 1.304 | 2.5475 | 98.53 | 20.04 | 49.671 | 79.3455 | 84.4924 | 76.8409 | 0.011828 | 1.0087 | 0.023523 | 23.3196 | 0.057011 | 0.15574 | 0.004046 | 0.27943 | 0.27943 | 0.02246
2019-04-09 14:43 | 1554835386.536 | 505 | 0.2087 | 0.001076 | 9.5e-05 | 95.5098 | 94.8997 | 118.9705 | 20.5355 | 0.002376 | 0.002301 | 246.6403 | 0.26145 | 8.4031 | 0.002537 | 0.004212 | 0.004142 | 3.6e-05 | 0.001023 | 0.26792 | 1.3112 | 2.5475 | 98.53 | 20.04 | 49.666 | 78.287 | 82.1444 | 76.4684 | 0.011827 | 1.0098 | 0.023509 | 23.2804 | 0.057323 | 0.15553 | 0.004043 | 0.27905 | 0.27905 | 0.022462
2019-04-09 14:43 | 1554835387.5366 | 506 | 0.20872 | 0.001071 | 9.5e-05 | 95.2784 | 94.697 | 118.8839 | 20.5358 | 0.002372 | 0.002311 | 246.6734 | 0.26185 | 8.4041 | 0.002581 | 0.004319 | 0.004291 | 3.7e-05 | 0.001022 | 0.27472 | 1.325 | 2.5475 | 98.53 | 20.04 | 49.662 | 78.182 | 83.2023 | 76.7104 | 0.011826 | 1.0091 | 0.023492 | 23.28 | 0.054754 | 0.15577 | 0.003921 | 0.27026 | 0.27026 | 0.023229
2019-04-09 14:43 | 1554835388.5356 | 507 | 0.20867 | 0.001069 | 9.6e-05 | 95.1075 | 94.6047 | 119.1186 | 20.5363 | 0.002453 | 0.002315 | 246.6339 | 0.26199 | 8.4028 | 0.002565 | 0.004128 | 0.00406 | 3.9e-05 | 0.001021 | 0.25481 | 1.3465 | 2.5476 | 98.53 | 20.04 | 49.658 | 80.1246 | 82.8184 | 76.8145 | 0.011825 | 1.0088 | 0.023512 | 23.3065 | 0.047409 | 0.15585 | 0.004197 | 0.28983 | 0.28983 | 0.021671
2019-04-09 14:43 | 1554835389.5356 | 508 | 0.20871 | 0.001068 | 9.6e-05 | 95.011 | 94.432 | 119.3184 | 20.536 | 0.002393 | 0.002266 | 246.6576 | 0.26171 | 8.4015 | 0.002608 | 0.003736 | 0.00376 | 4.1e-05 | 0.001021 | 0.26842 | 1.3807 | 2.5477 | 98.53 | 20.04 | 49.651 | 78.6952 | 83.8485 | 75.6302 | 0.011823 | 1.0122 | 0.023571 | 23.2865 | 0.054679 | 0.15569 | 0.003979 | 0.27521 | 0.27521 | 0.022798
2019-04-09 14:43 | 1554835390.5367 | 509 | 0.20869 | 0.001068 | 9.7e-05 | 94.7519 | 93.9976 | 119.0398 | 20.5358 | 0.00236 | 0.002322 | 246.6085 | 0.26191 | 8.402 | 0.00253 | 0.0043 | 0.004188 | 3.8e-05 | 0.001017 | 0.27688 | 1.3696 | 2.5476 | 98.53 | 20.04 | 49.647 | 77.9045 | 81.9817 | 76.9723 | 0.011822 | 1.0083 | 0.023499 | 23.304 | 0.071489 | 0.15581 | 0.004079 | 0.28133 | 0.28133 | 0.02232
2019-04-09 14:43 | 1554835391.5357 | 510 | 0.20869 | 0.001065 | 9.7e-05 | 94.5587 | 93.8397 | 118.8036 | 20.5365 | 0.002414 | 0.002297 | 246.6109 | 0.26173 | 8.4013 | 0.00257 | 0.004339 | 0.004261 | 4.1e-05 | 0.001019 | 0.28119 | 1.3362 | 2.5474 | 98.53 | 20.04 | 49.637 | 79.1907 | 82.9517 | 76.375 | 0.01182 | 1.0101 | 0.023495 | 23.261 | 0.068271 | 0.15571 | 0.004116 | 0.28387 | 0.28387 | 0.022105
2019-04-09 14:43 | 1554835392.5357 | 511 | 0.20869 | 0.001067 | 9.7e-05 | 94.3699 | 93.7471 | 118.9229 | 20.5364 | 0.002432 | 0.002262 | 246.5913 | 0.26164 | 8.4006 | 0.002612 | 0.004228 | 0.004199 | 4.1e-05 | 0.001019 | 0.28031 | 1.3252 | 2.5475 | 98.53 | 20.04 | 49.633 | 79.6272 | 83.9448 | 75.5518 | 0.011819 | 1.0125 | 0.023535 | 23.2453 | 0.059223 | 0.15565 | 0.00408 | 0.28181 | 0.28181 | 0.022258
2019-04-09 14:43 | 1554835393.5361 | 512 | 0.20868 | 0.001061 | 9.7e-05 | 94.1831 | 93.4907 | 119.0348 | 20.5364 | 0.002475 | 0.002294 | 246.5555 | 0.26143 | 8.4014 | 0.00256 | 0.004387 | 0.00434 | 4e-05 | 0.001022 | 0.27831 | 1.3006 | 2.5475 | 98.53 | 20.04 | 49.628 | 80.662 | 82.7045 | 76.3134 | 0.011818 | 1.0103 | 0.02352 | 23.2816 | 0.066004 | 0.15552 | 0.004151 | 0.28668 | 0.28668 | 0.021862
2019-04-09 14:43 | 1554835394.5368 | 513 | 0.2087 | 0.001059 | 9.7e-05 | 94.0098 | 93.2054 | 118.5804 | 20.5355 | 0.002372 | 0.002222 | 246.562 | 0.26154 | 8.4024 | 0.002524 | 0.004127 | 0.004085 | 4.3e-05 | 0.001025 | 0.26609 | 1.3396 | 2.5476 | 98.53 | 20.04 | 49.618 | 78.1928 | 81.8324 | 74.5892 | 0.011815 | 1.0153 | 0.023533 | 23.1797 | 0.076867 | 0.15559 | 0.004078 | 0.28166 | 0.28166 | 0.022262
2019-04-09 14:43 | 1554835395.5359 | 514 | 0.20868 | 0.001057 | 9.7e-05 | 93.7841 | 93.1145 | 118.5518 | 20.5358 | 0.002418 | 0.002297 | 246.5362 | 0.26167 | 8.402 | 0.002545 | 0.004268 | 0.004158 | 4e-05 | 0.001023 | 0.28423 | 1.3193 | 2.5475 | 98.53 | 20.04 | 49.613 | 79.2953 | 82.349 | 76.383 | 0.011814 | 1.01 | 0.02347 | 23.2366 | 0.064088 | 0.15567 | 0.004161 | 0.28676 | 0.28676 | 0.021877
2019-04-09 14:43 | 1554835396.5363 | 515 | 0.20869 | 0.001058 | 9.7e-05 | 93.6218 | 92.7848 | 118.9806 | 20.5377 | 0.002426 | 0.002197 | 246.5457 | 0.26228 | 8.4008 | 0.002522 | 0.003733 | 0.003681 | 4.4e-05 | 0.001022 | 0.27616 | 1.3459 | 2.5474 | 98.53 | 20.04 | 49.606 | 79.4812 | 81.7932 | 73.9684 | 0.011813 | 1.0171 | 0.023594 | 23.198 | 0.080323 | 0.15603 | 0.004105 | 0.28429 | 0.28429 | 0.022118
2019-04-09 14:43 | 1554835397.5359 | 516 | 0.20871 | 0.001053 | 9.7e-05 | 93.3529 | 92.6193 | 118.9902 | 20.5365 | 0.002383 | 0.002239 | 246.5058 | 0.26139 | 8.4015 | 0.002541 | 0.004127 | 0.004093 | 3.9e-05 | 0.001025 | 0.28512 | 1.3132 | 2.5475 | 98.53 | 20.04 | 49.603 | 78.4478 | 82.2334 | 74.9823 | 0.011812 | 1.0141 | 0.023561 | 23.2328 | 0.070567 | 0.1555 | 0.003995 | 0.27616 | 0.27616 | 0.022692
2019-04-09 14:43 | 1554835398.536 | 517 | 0.20871 | 0.001048 | 9.6e-05 | 93.1823 | 92.2899 | 118.6231 | 20.5377 | 0.002446 | 0.002297 | 246.4929 | 0.26165 | 8.4017 | 0.002546 | 0.004299 | 0.00427 | 4.2e-05 | 0.001029 | 0.30414 | 1.3005 | 2.5475 | 98.53 | 20.04 | 49.6 | 79.9521 | 82.3718 | 76.3824 | 0.011811 | 1.0101 | 0.023477 | 23.2436 | 0.086066 | 0.15566 | 0.004031 | 0.27775 | 0.27775 | 0.022585
2019-04-09 14:43 | 1554835399.5359 | 518 | 0.20871 | 0.001055 | 9.6e-05 | 92.9333 | 92.2789 | 118.5769 | 20.5358 | 0.002406 | 0.002302 | 246.4791 | 0.26243 | 8.4016 | 0.002472 | 0.003821 | 0.003785 | 3.7e-05 | 0.001035 | 0.25177 | 1.3259 | 2.5475 | 98.53 | 20.04 | 49.597 | 78.9913 | 80.5935 | 76.5093 | 0.01181 | 1.0097 | 0.023468 | 23.2433 | 0.063205 | 0.15612 | 0.003977 | 0.27385 | 0.27385 | 0.022975
2019-04-09 14:43 | 1554835400.535 | 519 | 0.20872 | 0.001052 | 9.6e-05 | 92.7254 | 91.9217 | 118.6205 | 20.5365 | 0.002412 | 0.002353 | 246.4287 | 0.26254 | 8.3999 | 0.00252 | 0.004165 | 0.004129 | 4.8e-05 | 0.001031 | 0.26864 | 1.3405 | 2.5476 | 98.53 | 20.04 | 49.596 | 79.1542 | 81.7406 | 77.733 | 0.01181 | 1.0062 | 0.023432 | 23.2882 | 0.07787 | 0.15619 | 0.003961 | 0.27233 | 0.27233 | 0.023113
2019-04-09 14:43 | 1554835401.5354 | 520 | 0.20871 | 0.001052 | 9.7e-05 | 92.6337 | 91.8264 | 118.572 | 20.5373 | 0.002425 | 0.002242 | 246.4407 | 0.26102 | 8.4001 | 0.002506 | 0.003436 | 0.003382 | 4.3e-05 | 0.001037 | 0.25177 | 1.3238 | 2.5476 | 98.53 | 20.04 | 49.596 | 79.4536 | 81.3981 | 75.0561 | 0.01181 | 1.0139 | 0.023517 | 23.1944 | 0.078299 | 0.15528 | 0.00398 | 0.27455 | 0.27455 | 0.022794
2019-04-09 14:43 | 1554835402.5361 | 521 | 0.20871 | 0.001051 | 9.7e-05 | 92.4059 | 91.4918 | 118.4229 | 20.5362 | 0.00238 | 0.002203 | 246.3834 | 0.26191 | 8.4038 | 0.002556 | 0.003966 | 0.003936 | 4.2e-05 | 0.001036 | 0.26985 | 1.3108 | 2.5476 | 98.53 | 20.03 | 49.599 | 78.3866 | 82.6138 | 74.1329 | 0.011804 | 1.0166 | 0.023533 | 23.1491 | 0.088919 | 0.15581 | 0.003981 | 0.2748 | 0.2748 | 0.02285
2019-04-09 14:43 | 1554835403.535 | 522 | 0.20871 | 0.001051 | 9.8e-05 | 92.2585 | 91.2262 | 118.3102 | 20.5367 | 0.00244 | 0.002246 | 246.4149 | 0.26245 | 8.4012 | 0.00249 | 0.00446 | 0.004344 | 4.4e-05 | 0.001033 | 0.25321 | 1.2791 | 2.5474 | 98.53 | 20.03 | 49.599 | 79.814 | 81.0084 | 75.1506 | 0.011804 | 1.0136 | 0.023488 | 23.172 | 0.10065 | 0.15613 | 0.003996 | 0.27531 | 0.27531 | 0.022855
2019-04-09 14:43 | 1554835404.5351 | 523 | 0.20872 | 0.001051 | 9.8e-05 | 91.9831 | 90.9171 | 118.2856 | 20.5364 | 0.002411 | 0.002244 | 246.3704 | 0.26174 | 8.4012 | 0.002505 | 0.003522 | 0.003523 | 4.4e-05 | 0.001031 | 0.27688 | 1.3372 | 2.5474 | 98.53 | 20.03 | 49.603 | 79.1284 | 81.3716 | 75.1122 | 0.011805 | 1.0137 | 0.023486 | 23.1683 | 0.10426 | 0.15571 | 0.003949 | 0.27202 | 0.27202 | 0.023069
2019-04-09 14:43 | 1554835405.5362 | 524 | 0.20871 | 0.001049 | 9.9e-05 | 91.8085 | 90.766 | 118.2283 | 20.5369 | 0.002439 | 0.002266 | 246.3589 | 0.26212 | 8.4006 | 0.002499 | 0.00368 | 0.003555 | 4.2e-05 | 0.001031 | 0.28622 | 1.3537 | 2.5476 | 98.53 | 20.03 | 49.604 | 79.7892 | 81.2286 | 75.646 | 0.011805 | 1.0122 | 0.023463 | 23.1804 | 0.10215 | 0.15594 | 0.004011 | 0.27606 | 0.27606 | 0.022764
2019-04-09 14:43 | 1554835406.5353 | 525 | 0.20871 | 0.001055 | 9.8e-05 | 91.5778 | 90.4346 | 118.1584 | 20.5364 | 0.002405 | 0.002293 | 246.3558 | 0.26138 | 8.4036 | 0.002477 | 0.004479 | 0.004409 | 4.5e-05 | 0.001035 | 0.28351 | 1.3189 | 2.5475 | 98.53 | 20.03 | 49.607 | 78.9739 | 80.7065 | 76.2783 | 0.011806 | 1.0104 | 0.023435 | 23.1945 | 0.11236 | 0.1555 | 0.003978 | 0.27338 | 0.27338 | 0.022922
2019-04-09 14:43 | 1554835407.5362 | 526 | 0.20871 | 0.001047 | 0.0001 | 91.4381 | 90.4535 | 118.2178 | 20.5368 | 0.002489 | 0.002276 | 246.3037 | 0.26216 | 8.4017 | 0.002556 | 0.004343 | 0.004381 | 4.3e-05 | 0.001034 | 0.30541 | 1.3222 | 2.5475 | 98.53 | 20.03 | 49.611 | 81.0054 | 82.6152 | 75.8852 | 0.011806 | 1.0115 | 0.023454 | 23.1873 | 0.096833 | 0.15596 | 0.004037 | 0.27773 | 0.27773 | 0.022631
2019-04-09 14:43 | 1554835408.5352 | 527 | 0.20872 | 0.001045 | 0.0001 | 91.355 | 90.2684 | 118.4968 | 20.5372 | 0.002417 | 0.00228 | 246.3189 | 0.26189 | 8.4002 | 0.002579 | 0.004 | 0.003947 | 4.7e-05 | 0.001038 | 0.28982 | 1.3243 | 2.5475 | 98.53 | 20.03 | 49.614 | 79.2695 | 83.1529 | 75.9783 | 0.011807 | 1.0112 | 0.023478 | 23.2178 | 0.10703 | 0.1558 | 0.003947 | 0.2717 | 0.2717 | 0.023109
2019-04-09 14:43 | 1554835409.5353 | 528 | 0.20871 | 0.001049 | 0.0001 | 91.2103 | 90.0715 | 118.3163 | 20.537 | 0.002421 | 0.002241 | 246.263 | 0.26121 | 8.4022 | 0.002569 | 0.004038 | 0.003933 | 4.4e-05 | 0.001035 | 0.28313 | 1.3296 | 2.5475 | 98.53 | 20.03 | 49.619 | 79.3589 | 82.9233 | 75.0301 | 0.011808 | 1.014 | 0.023492 | 23.1685 | 0.11238 | 0.15539 | 0.004008 | 0.27607 | 0.27607 | 0.022684
2019-04-09 14:43 | 1554835410.5364 | 529 | 0.20871 | 0.001048 | 0.000101 | 90.9908 | 89.5897 | 118.152 | 20.5362 | 0.002406 | 0.002222 | 246.2767 | 0.26118 | 8.4024 | 0.002527 | 0.003904 | 0.003967 | 4.5e-05 | 0.001032 | 0.28119 | 1.3355 | 2.5475 | 98.53 | 20.03 | 49.619 | 78.9987 | 81.9134 | 74.5844 | 0.011808 | 1.0153 | 0.023491 | 23.1376 | 0.1388 | 0.15538 | 0.004009 | 0.2761 | 0.2761 | 0.022679
2019-04-09 14:44 | 1554835411.5353 | 530 | 0.20872 | 0.001045 | 0.000101 | 90.7666 | 89.5522 | 118.1349 | 20.5378 | 0.002416 | 0.002281 | 246.2146 | 0.26129 | 8.4007 | 0.002501 | 0.003619 | 0.003548 | 4.3e-05 | 0.00103 | 0.29806 | 1.3237 | 2.5475 | 98.53 | 20.03 | 49.621 | 79.2506 | 81.285 | 75.9883 | 0.011809 | 1.0112 | 0.023442 | 23.1826 | 0.12048 | 0.15544 | 0.003968 | 0.27259 | 0.27259 | 0.02298
2019-04-09 14:44 | 1554835412.5345 | 531 | 0.20871 | 0.001047 | 0.000102 | 90.5534 | 89.3399 | 118.122 | 20.537 | 0.00244 | 0.002231 | 246.2014 | 0.26133 | 8.4013 | 0.002553 | 0.003599 | 0.003594 | 4.7e-05 | 0.00103 | 0.30326 | 1.2909 | 2.5474 | 98.53 | 20.03 | 49.621 | 79.8174 | 82.5365 | 74.7886 | 0.011809 | 1.0147 | 0.023481 | 23.1415 | 0.12068 | 0.15546 | 0.003984 | 0.27417 | 0.27417 | 0.022852
2019-04-09 14:44 | 1554835413.5349 | 532 | 0.20873 | 0.001049 | 0.000102 | 90.4861 | 89.0012 | 118.371 | 20.5373 | 0.002416 | 0.002267 | 246.1608 | 0.26277 | 8.4008 | 0.002446 | 0.004096 | 0.003977 | 4.2e-05 | 0.001032 | 0.30143 | 1.3274 | 2.5476 | 98.53 | 20.03 | 49.621 | 79.2297 | 79.9511 | 75.6555 | 0.011809 | 1.0122 | 0.023477 | 23.1947 | 0.14799 | 0.15632 | 0.003888 | 0.26734 | 0.26734 | 0.023564
2019-04-09 14:44 | 1554835414.5344 | 533 | 0.20872 | 0.001045 | 0.000102 | 90.28 | 88.9482 | 118.8654 | 20.5382 | 0.002426 | 0.002246 | 246.1165 | 0.2614 | 8.4021 | 0.002549 | 0.003464 | 0.00344 | 4.5e-05 | 0.001025 | 0.25785 | 1.3519 | 2.5476 | 98.53 | 20.03 | 49.62 | 79.4762 | 82.4348 | 75.1552 | 0.011809 | 1.0136 | 0.023543 | 23.2264 | 0.13293 | 0.15551 | 0.003946 | 0.27218 | 0.27218 | 0.023025
2019-04-09 14:44 | 1554835415.5349 | 534 | 0.20875 | 0.001048 | 0.000102 | 90.0418 | 88.5548 | 118.84 | 20.5374 | 0.002405 | 0.002271 | 246.1242 | 0.26133 | 8.4 | 0.002589 | 0.00438 | 0.004316 | 4.7e-05 | 0.001023 | 0.28816 | 1.3587 | 2.5476 | 98.53 | 20.03 | 49.62 | 78.9734 | 83.3981 | 75.7493 | 0.011809 | 1.0119 | 0.02352 | 23.2437 | 0.14894 | 0.15546 | 0.003791 | 0.26104 | 0.26104 | 0.024001
2019-04-09 14:44 | 1554835416.5356 | 535 | 0.20872 | 0.001041 | 0.000102 | 89.7943 | 88.3786 | 118.5245 | 20.5376 | 0.002434 | 0.002211 | 246.0502 | 0.26238 | 8.4004 | 0.002533 | 0.003815 | 0.003785 | 4.3e-05 | 0.001024 | 0.26377 | 1.3359 | 2.5476 | 98.53 | 20.03 | 49.619 | 79.6784 | 82.063 | 74.3245 | 0.011808 | 1.016 | 0.023537 | 23.1654 | 0.14214 | 0.15609 | 0.00397 | 0.27378 | 0.27378 | 0.022976
2019-04-09 14:44 | 1554835417.5345 | 536 | 0.20872 | 0.001044 | 0.000102 | 89.6499 | 88.4895 | 118.3828 | 20.5381 | 0.002402 | 0.002228 | 246.0526 | 0.2618 | 8.403 | 0.002557 | 0.004671 | 0.004612 | 4.5e-05 | 0.001016 | 0.27776 | 1.3364 | 2.5476 | 98.53 | 20.03 | 49.618 | 78.9031 | 82.6404 | 74.7366 | 0.011808 | 1.0148 | 0.023509 | 23.1653 | 0.11653 | 0.15574 | 0.003951 | 0.27215 | 0.27215 | 0.023063
2019-04-09 14:44 | 1554835418.5346 | 537 | 0.20874 | 0.001039 | 0.000102 | 89.6472 | 88.41 | 118.4788 | 20.5391 | 0.002458 | 0.002279 | 245.9968 | 0.26196 | 8.4031 | 0.002492 | 0.004143 | 0.004149 | 4.3e-05 | 0.001013 | 0.27649 | 1.3112 | 2.5475 | 98.53 | 20.03 | 49.618 | 80.26 | 81.0709 | 75.9528 | 0.011808 | 1.0113 | 0.023477 | 23.2151 | 0.1243 | 0.15584 | 0.003842 | 0.26419 | 0.26419 | 0.023772
2019-04-09 14:44 | 1554835419.5357 | 538 | 0.20874 | 0.001035 | 0.000103 | 89.6064 | 88.263 | 118.6173 | 20.5386 | 0.002399 | 0.002264 | 245.9823 | 0.26195 | 8.4012 | 0.002502 | 0.004209 | 0.003998 | 4.5e-05 | 0.001013 | 0.29358 | 1.3133 | 2.5475 | 98.53 | 20.03 | 49.618 | 78.8335 | 81.2997 | 75.5946 | 0.011808 | 1.0123 | 0.023503 | 23.2168 | 0.13511 | 0.15584 | 0.003857 | 0.26544 | 0.26544 | 0.02366
2019-04-09 14:44 | 1554835420.5346 | 539 | 0.20872 | 0.001036 | 0.000103 | 89.6144 | 88.2055 | 118.5809 | 20.538 | 0.002417 | 0.002253 | 245.9702 | 0.26187 | 8.4024 | 0.002564 | 0.004243 | 0.004153 | 4.2e-05 | 0.001012 | 0.27831 | 1.3447 | 2.5475 | 98.53 | 20.03 | 49.619 | 79.2655 | 82.8043 | 75.3345 | 0.011808 | 1.0131 | 0.023508 | 23.2046 | 0.14174 | 0.15579 | 0.003956 | 0.27243 | 0.27243 | 0.023045
2019-04-09 14:44 | 1554835421.5346 | 540 | 0.20872 | 0.00103 | 0.000104 | 89.6269 | 88.2289 | 118.6388 | 20.5378 | 0.00248 | 0.002284 | 245.937 | 0.26167 | 8.4013 | 0.002537 | 0.003987 | 0.004028 | 4.4e-05 | 0.001013 | 0.30431 | 1.3817 | 2.5477 | 98.53 | 20.03 | 49.619 | 80.7739 | 82.1529 | 76.063 | 0.011808 | 1.011 | 0.023489 | 23.2345 | 0.14062 | 0.15567 | 0.003942 | 0.27124 | 0.27124 | 0.023129
2019-04-09 14:44 | 1554835422.5351 | 541 | 0.20871 | 0.001026 | 0.000105 | 89.6088 | 88.4272 | 118.8222 | 20.5382 | 0.002447 | 0.00228 | 245.9591 | 0.26171 | 8.4017 | 0.002545 | 0.003731 | 0.003631 | 4.3e-05 | 0.001011 | 0.27295 | 1.3639 | 2.5476 | 98.53 | 20.03 | 49.616 | 79.9923 | 82.3458 | 75.9699 | 0.011808 | 1.0112 | 0.023511 | 23.2493 | 0.11873 | 0.15569 | 0.004017 | 0.27667 | 0.27667 | 0.022679
2019-04-09 14:44 | 1554835423.5352 | 542 | 0.20876 | 0.001026 | 0.000105 | 89.5817 | 88.274 | 118.7363 | 20.538 | 0.002428 | 0.002269 | 245.9246 | 0.26231 | 8.4013 | 0.002507 | 0.003977 | 0.004025 | 4.1e-05 | 0.001004 | 0.29463 | 1.3897 | 2.5476 | 98.53 | 20.03 | 49.618 | 79.5221 | 81.4387 | 75.6984 | 0.011808 | 1.012 | 0.023511 | 23.2319 | 0.13153 | 0.15604 | 0.003754 | 0.25813 | 0.25813 | 0.024362
2019-04-09 14:44 | 1554835424.5357 | 543 | 0.20874 | 0.001024 | 0.000105 | 89.6412 | 88.0599 | 118.8757 | 20.5369 | 0.002411 | 0.002178 | 245.9495 | 0.26139 | 8.4022 | 0.00255 | 0.004299 | 0.004136 | 4.4e-05 | 0.001003 | 0.27671 | 1.3188 | 2.5475 | 98.53 | 20.03 | 49.617 | 79.1147 | 82.4617 | 73.5147 | 0.011808 | 1.0184 | 0.023599 | 23.1727 | 0.15919 | 0.1555 | 0.003863 | 0.26678 | 0.26678 | 0.02349
2019-04-09 14:44 | 1554835425.5358 | 544 | 0.20873 | 0.001033 | 0.000105 | 89.6555 | 88.1601 | 118.8069 | 20.538 | 0.00249 | 0.002228 | 245.8949 | 0.26242 | 8.4023 | 0.002602 | 0.004513 | 0.004496 | 4e-05 | 0.001001 | 0.30326 | 1.3313 | 2.5476 | 98.53 | 20.03 | 49.617 | 81.0239 | 83.7028 | 74.7195 | 0.011808 | 1.0149 | 0.023551 | 23.2062 | 0.15045 | 0.15611 | 0.00391 | 0.26959 | 0.26959 | 0.023337
2019-04-09 14:44 | 1554835426.5348 | 545 | 0.20871 | 0.001025 | 0.000105 | 89.6505 | 88.2585 | 118.7994 | 20.538 | 0.002477 | 0.002217 | 245.8999 | 0.26193 | 8.4037 | 0.002482 | 0.004283 | 0.004192 | 4.5e-05 | 0.001006 | 0.29247 | 1.3275 | 2.5476 | 98.53 | 20.03 | 49.617 | 80.7001 | 80.8328 | 74.4513 | 0.011808 | 1.0157 | 0.02356 | 23.1965 | 0.13996 | 0.15582 | 0.004006 | 0.27647 | 0.27647 | 0.022714
2019-04-09 14:44 | 1554835427.5358 | 546 | 0.20873 | 0.001025 | 0.000105 | 89.649 | 88.198 | 119.0745 | 20.538 | 0.002404 | 0.002272 | 245.8909 | 0.26238 | 8.3998 | 0.002542 | 0.003986 | 0.003993 | 4.4e-05 | 0.001004 | 0.28031 | 1.3024 | 2.5475 | 98.53 | 20.03 | 49.618 | 78.9497 | 82.2615 | 75.7749 | 0.011808 | 1.0118 | 0.023542 | 23.2675 | 0.14596 | 0.15609 | 0.003938 | 0.27145 | 0.27145 | 0.023173
2019-04-09 14:44 | 1554835428.5349 | 547 | 0.20877 | 0.001022 | 0.000105 | 89.7255 | 88.2706 | 118.9313 | 20.5381 | 0.002459 | 0.002263 | 245.8615 | 0.26211 | 8.4051 | 0.002597 | 0.004145 | 0.004039 | 4.4e-05 | 0.001004 | 0.27848 | 1.3154 | 2.5476 | 98.53 | 20.03 | 49.618 | 80.2823 | 83.5914 | 75.5746 | 0.011808 | 1.0124 | 0.023535 | 23.2468 | 0.14623 | 0.15593 | 0.003682 | 0.25339 | 0.25339 | 0.0248
2019-04-09 14:44 | 1554835429.535 | 548 | 0.20871 | 0.001024 | 0.000105 | 89.9471 | 88.7283 | 118.6128 | 20.5372 | 0.002469 | 0.002226 | 245.8711 | 0.26171 | 8.3991 | 0.002558 | 0.004039 | 0.004022 | 4.1e-05 | 0.001005 | 0.29623 | 1.3174 | 2.5475 | 98.53 | 20.03 | 49.619 | 80.5037 | 82.6656 | 74.6694 | 0.011808 | 1.015 | 0.023534 | 23.1855 | 0.12202 | 0.15569 | 0.004006 | 0.27617 | 0.27617 | 0.022719
2019-04-09 14:44 | 1554835430.5343 | 549 | 0.20875 | 0.001019 | 0.000105 | 90.0289 | 88.9417 | 118.8416 | 20.5383 | 0.002358 | 0.002298 | 245.8127 | 0.26148 | 8.4039 | 0.002537 | 0.004442 | 0.004377 | 4.3e-05 | 0.001004 | 0.29015 | 1.3442 | 2.5475 | 98.53 | 20.03 | 49.619 | 77.8524 | 82.1485 | 76.4034 | 0.011808 | 1.01 | 0.023498 | 23.2657 | 0.10868 | 0.15555 | 0.003767 | 0.25895 | 0.25895 | 0.024209
2019-04-09 14:44 | 1554835431.535 | 550 | 0.20874 | 0.001018 | 0.000106 | 90.2115 | 88.8326 | 118.8783 | 20.5387 | 0.002415 | 0.002291 | 245.8308 | 0.26217 | 8.4007 | 0.002496 | 0.004394 | 0.004378 | 4.1e-05 | 0.001004 | 0.28479 | 1.3625 | 2.5476 | 98.53 | 20.03 | 49.619 | 79.2272 | 81.1694 | 76.2392 | 0.011808 | 1.0105 | 0.023507 | 23.2638 | 0.13778 | 0.15596 | 0.003883 | 0.26712 | 0.26712 | 0.02353
2019-04-09 14:44 | 1554835432.5354 | 551 | 0.20872 | 0.001011 | 0.000107 | 90.3297 | 88.9448 | 118.7287 | 20.5385 | 0.002397 | 0.002232 | 245.7846 | 0.26179 | 8.4025 | 0.002567 | 0.004277 | 0.00421 | 4e-05 | 0.001007 | 0.29806 | 1.3716 | 2.5476 | 98.53 | 20.03 | 49.62 | 78.7914 | 82.8796 | 74.8165 | 0.011809 | 1.0146 | 0.02354 | 23.2018 | 0.1382 | 0.15574 | 0.004 | 0.27567 | 0.27567 | 0.022767
2019-04-09 14:44 | 1554835433.5351 | 552 | 0.20877 | 0.001013 | 0.000107 | 90.3958 | 88.9692 | 118.5288 | 20.5384 | 0.002327 | 0.00228 | 245.7844 | 0.2621 | 8.4006 | 0.002548 | 0.004171 | 0.004143 | 3.9e-05 | 0.001001 | 0.28639 | 1.3542 | 2.5476 | 98.53 | 20.03 | 49.621 | 77.0933 | 82.4139 | 75.9778 | 0.011809 | 1.0112 | 0.023481 | 23.2209 | 0.14229 | 0.15592 | 0.003678 | 0.25237 | 0.25237 | 0.024899
2019-04-09 14:44 | 1554835434.5351 | 553 | 0.20873 | 0.001014 | 0.000108 | 90.5057 | 89.133 | 118.2546 | 20.5382 | 0.002432 | 0.002191 | 245.761 | 0.26199 | 8.4003 | 0.002561 | 0.003394 | 0.003356 | 4.4e-05 | 0.000999 | 0.30663 | 1.3212 | 2.5476 | 98.53 | 20.03 | 49.623 | 79.6313 | 82.7304 | 73.8247 | 0.011809 | 1.0175 | 0.023527 | 23.1224 | 0.1367 | 0.15585 | 0.003925 | 0.27017 | 0.27017 | 0.023248
2019-04-09 14:44 | 1554835435.5345 | 554 | 0.20874 | 0.001015 | 0.000108 | 90.5352 | 89.1888 | 118.5499 | 20.5392 | 0.002353 | 0.002239 | 245.7314 | 0.26136 | 8.4034 | 0.002572 | 0.003723 | 0.003606 | 4e-05 | 0.000995 | 0.26432 | 1.3505 | 2.5476 | 98.53 | 20.03 | 49.622 | 77.7177 | 82.9824 | 74.9832 | 0.011809 | 1.0141 | 0.023517 | 23.1898 | 0.13402 | 0.15548 | 0.00388 | 0.26688 | 0.26688 | 0.023479
2019-04-09 14:44 | 1554835436.5356 | 555 | 0.20874 | 0.001018 | 0.000109 | 90.7344 | 89.4032 | 118.7912 | 20.5381 | 0.002504 | 0.002269 | 245.7445 | 0.26173 | 8.4046 | 0.002617 | 0.003781 | 0.003847 | 4.2e-05 | 0.000994 | 0.30287 | 1.3678 | 2.5476 | 98.53 | 20.03 | 49.622 | 81.3624 | 84.079 | 75.7024 | 0.011809 | 1.012 | 0.023517 | 23.2374 | 0.1322 | 0.15571 | 0.003874 | 0.26638 | 0.26638 | 0.023556
2019-04-09 14:44 | 1554835437.5342 | 556 | 0.20875 | 0.001017 | 0.00011 | 90.7508 | 89.534 | 118.5553 | 20.5379 | 0.002501 | 0.002256 | 245.703 | 0.26278 | 8.3998 | 0.002584 | 0.004047 | 0.003933 | 4.1e-05 | 0.000994 | 0.27887 | 1.3399 | 2.5476 | 98.53 | 20.03 | 49.622 | 81.2725 | 83.2815 | 75.3873 | 0.011809 | 1.0129 | 0.023504 | 23.2038 | 0.12075 | 0.15633 | 0.003827 | 0.2629 | 0.2629 | 0.023963
2019-04-09 14:44 | 1554835438.5343 | 557 | 0.20872 | 0.001009 | 0.00011 | 90.7996 | 89.4999 | 118.5212 | 20.5386 | 0.002402 | 0.002287 | 245.7044 | 0.2626 | 8.4012 | 0.002547 | 0.003831 | 0.003783 | 4e-05 | 0.000997 | 0.29048 | 1.3645 | 2.5475 | 98.53 | 20.03 | 49.62 | 78.9043 | 82.3945 | 76.1408 | 0.011809 | 1.0107 | 0.023475 | 23.2256 | 0.12896 | 0.15622 | 0.004001 | 0.27481 | 0.27481 | 0.022909
2019-04-09 14:44 | 1554835439.5353 | 558 | 0.20871 | 0.001014 | 0.000111 | 90.9595 | 89.6751 | 118.94 | 20.5392 | 0.002343 | 0.002229 | 245.6742 | 0.26183 | 8.4016 | 0.002543 | 0.003735 | 0.003717 | 4.1e-05 | 0.000994 | 0.30093 | 1.3505 | 2.5476 | 98.53 | 20.03 | 49.62 | 77.497 | 82.2937 | 74.7415 | 0.011809 | 1.0148 | 0.023564 | 23.2199 | 0.12721 | 0.15576 | 0.004024 | 0.27747 | 0.27747 | 0.022624
2019-04-09 14:44 | 1554835440.5343 | 559 | 0.20874 | 0.001016 | 0.000111 | 90.9131 | 89.7508 | 118.8482 | 20.5384 | 0.002407 | 0.002265 | 245.6524 | 0.26175 | 8.4003 | 0.002574 | 0.00387 | 0.003746 | 4.4e-05 | 0.000999 | 0.29142 | 1.3516 | 2.5476 | 98.53 | 20.03 | 49.618 | 79.0234 | 83.0435 | 75.6239 | 0.011808 | 1.0122 | 0.023525 | 23.2404 | 0.11509 | 0.15572 | 0.003868 | 0.26598 | 0.26598 | 0.023593
2019-04-09 14:44 | 1554835441.5357 | 560 | 0.20874 | 0.001012 | 0.00011 | 91.093 | 89.9593 | 118.6649 | 20.5389 | 0.002419 | 0.002273 | 245.6345 | 0.26175 | 8.4008 | 0.002553 | 0.004749 | 0.004757 | 4.3e-05 | 0.001001 | 0.29822 | 1.3375 | 2.5475 | 98.53 | 20.03 | 49.616 | 79.3108 | 82.5454 | 75.8099 | 0.011808 | 1.0117 | 0.023501 | 23.2286 | 0.11201 | 0.15571 | 0.003866 | 0.26562 | 0.26562 | 0.023625
2019-04-09 14:44 | 1554835442.5356 | 561 | 0.20872 | 0.001013 | 0.000112 | 91.2499 | 90.0271 | 118.6522 | 20.5382 | 0.002439 | 0.002269 | 245.5955 | 0.26177 | 8.4042 | 0.002598 | 0.004348 | 0.004272 | 4e-05 | 0.000999 | 0.29502 | 1.3139 | 2.5475 | 98.53 | 20.03 | 49.615 | 79.7866 | 83.6275 | 75.7032 | 0.011807 | 1.012 | 0.023503 | 23.2238 | 0.12067 | 0.15573 | 0.003949 | 0.27135 | 0.27135 | 0.023128
2019-04-09 14:44 | 1554835443.5333 | 562 | 0.20874 | 0.001011 | 0.000113 | 91.3306 | 90.1496 | 118.8292 | 20.5384 | 0.002363 | 0.002218 | 245.5969 | 0.2621 | 8.4035 | 0.002627 | 0.004451 | 0.00447 | 4.1e-05 | 0.001001 | 0.2672 | 1.3032 | 2.5476 | 98.53 | 20.03 | 49.612 | 77.972 | 84.3167 | 74.4952 | 0.011807 | 1.0155 | 0.023561 | 23.2009 | 0.11641 | 0.15593 | 0.003833 | 0.26379 | 0.26379 | 0.023821
2019-04-09 14:44 | 1554835444.5348 | 563 | 0.20869 | 0.001015 | 0.000114 | 91.4271 | 90.1607 | 118.6522 | 20.5394 | 0.002379 | 0.002269 | 245.5657 | 0.26217 | 8.3996 | 0.002634 | 0.004257 | 0.004223 | 4.2e-05 | 0.001004 | 0.28854 | 1.2893 | 2.5476 | 98.53 | 20.03 | 49.609 | 78.3657 | 84.4752 | 75.7106 | 0.011806 | 1.012 | 0.023503 | 23.2241 | 0.12476 | 0.15596 | 0.004186 | 0.28796 | 0.28796 | 0.021827
2019-04-09 14:44 | 1554835445.5345 | 564 | 0.20871 | 0.001016 | 0.000114 | 91.5629 | 90.3561 | 118.4227 | 20.5388 | 0.002418 | 0.002202 | 245.5721 | 0.262 | 8.4012 | 0.002612 | 0.003909 | 0.003828 | 4.2e-05 | 0.001003 | 0.28854 | 1.2857 | 2.5476 | 98.53 | 20.03 | 49.606 | 79.2819 | 83.9518 | 74.1122 | 0.011805 | 1.0167 | 0.023534 | 23.1484 | 0.11867 | 0.15587 | 0.00402 | 0.27661 | 0.27661 | 0.022709
2019-04-09 14:44 | 1554835446.5335 | 565 | 0.2087 | 0.001013 | 0.000114 | 91.7034 | 90.6201 | 118.2865 | 20.5399 | 0.002396 | 0.002264 | 245.5317 | 0.26276 | 8.4009 | 0.002569 | 0.003522 | 0.003605 | 3.9e-05 | 0.001002 | 0.2792 | 1.3439 | 2.5476 | 98.53 | 20.03 | 49.604 | 78.7685 | 82.9139 | 75.5864 | 0.011805 | 1.0124 | 0.023471 | 23.1841 | 0.10629 | 0.15632 | 0.004105 | 0.28182 | 0.28182 | 0.022353
2019-04-09 14:44 | 1554835447.5345 | 566 | 0.20869 | 0.001012 | 0.000113 | 91.7663 | 90.5189 | 118.6665 | 20.5392 | 0.002483 | 0.002297 | 245.5045 | 0.26168 | 8.4041 | 0.002625 | 0.004265 | 0.004172 | 4.2e-05 | 0.001 | 0.28031 | 1.3714 | 2.5476 | 98.53 | 20.03 | 49.602 | 80.8597 | 84.2783 | 76.3933 | 0.011804 | 1.01 | 0.023481 | 23.2482 | 0.12241 | 0.15567 | 0.004158 | 0.28574 | 0.28574 | 0.021956
2019-04-09 14:44 | 1554835448.5349 | 567 | 0.20871 | 0.00101 | 0.000113 | 91.9697 | 90.7956 | 118.806 | 20.5395 | 0.00241 | 0.002292 | 245.5174 | 0.26201 | 8.403 | 0.002633 | 0.003796 | 0.003751 | 4.1e-05 | 0.000999 | 0.30718 | 1.3469 | 2.5476 | 98.53 | 20.04 | 49.6 | 79.1033 | 84.4582 | 76.2737 | 0.011811 | 1.0104 | 0.023499 | 23.2579 | 0.11492 | 0.15587 | 0.004066 | 0.2795 | 0.2795 | 0.022474
2019-04-09 14:44 | 1554835449.535 | 568 | 0.20869 | 0.001018 | 0.000113 | 92.0215 | 90.8848 | 118.5424 | 20.5396 | 0.002395 | 0.002248 | 245.4674 | 0.26171 | 8.4023 | 0.002639 | 0.003818 | 0.003739 | 4.4e-05 | 0.001 | 0.29734 | 1.3675 | 2.5476 | 98.53 | 20.04 | 49.6 | 78.7439 | 84.6134 | 75.201 | 0.011811 | 1.0135 | 0.023509 | 23.1964 | 0.11118 | 0.15569 | 0.004166 | 0.28667 | 0.28667 | 0.021887
2019-04-09 14:44 | 1554835450.5336 | 569 | 0.20869 | 0.001013 | 0.000113 | 92.0763 | 91.3604 | 118.552 | 20.5394 | 0.002389 | 0.002247 | 245.4817 | 0.26178 | 8.4018 | 0.002657 | 0.004248 | 0.004137 | 4.4e-05 | 0.000999 | 0.29175 | 1.3591 | 2.5476 | 98.53 | 20.04 | 49.6 | 78.6057 | 85.0252 | 75.1893 | 0.011811 | 1.0135 | 0.02351 | 23.1969 | 0.069814 | 0.15574 | 0.00417 | 0.28694 | 0.28694 | 0.021873
2019-04-09 14:44 | 1554835451.5347 | 570 | 0.20874 | 0.001014 | 0.000112 | 92.1987 | 91.2499 | 118.771 | 20.5393 | 0.002421 | 0.002241 | 245.4445 | 0.26168 | 8.4008 | 0.00266 | 0.004986 | 0.004922 | 4.5e-05 | 0.000999 | 0.27992 | 1.3463 | 2.5478 | 98.53 | 20.04 | 49.6 | 79.3518 | 85.1172 | 75.034 | 0.011811 | 1.014 | 0.023537 | 23.2132 | 0.092519 | 0.15567 | 0.003879 | 0.26682 | 0.26682 | 0.023512
2019-04-09 14:44 | 1554835452.5337 | 571 | 0.20873 | 0.001005 | 0.000112 | 92.2615 | 91.3194 | 118.7471 | 20.5397 | 0.002428 | 0.00222 | 245.4475 | 0.26189 | 8.4005 | 0.002637 | 0.004785 | 0.00473 | 4.2e-05 | 0.000997 | 0.28136 | 1.3365 | 2.5476 | 98.53 | 20.04 | 49.599 | 79.5269 | 84.5527 | 74.5277 | 0.011811 | 1.0154 | 0.023552 | 23.1939 | 0.0918 | 0.1558 | 0.003952 | 0.27215 | 0.27215 | 0.023071
2019-04-09 14:44 | 1554835453.5347 | 572 | 0.20872 | 0.00101 | 0.000112 | 92.4006 | 91.6935 | 118.6299 | 20.5399 | 0.00246 | 0.002315 | 245.4082 | 0.26154 | 8.403 | 0.002699 | 0.004226 | 0.004121 | 4.4e-05 | 0.000994 | 0.30143 | 1.3536 | 2.5476 | 98.53 | 20.04 | 49.599 | 80.3083 | 86.0584 | 76.8211 | 0.011811 | 1.0088 | 0.023463 | 23.2588 | 0.068706 | 0.15559 | 0.003964 | 0.27195 | 0.27195 | 0.023057
2019-04-09 14:44 | 1554835454.5347 | 573 | 0.20875 | 0.001009 | 0.000113 | 92.4839 | 91.5366 | 118.6809 | 20.54 | 0.002386 | 0.00227 | 245.3808 | 0.26205 | 8.4013 | 0.002693 | 0.004525 | 0.004493 | 4.3e-05 | 0.000989 | 0.2943 | 1.357 | 2.5475 | 98.53 | 20.04 | 49.598 | 78.5292 | 85.8912 | 75.7424 | 0.011811 | 1.0119 | 0.023504 | 23.2279 | 0.092095 | 0.15589 | 0.003833 | 0.26322 | 0.26322 | 0.023868
2019-04-09 14:44 | 1554835455.534 | 574 | 0.20872 | 0.001006 | 0.000112 | 92.5712 | 91.6051 | 118.7391 | 20.5394 | 0.002367 | 0.002252 | 245.3907 | 0.26187 | 8.4029 | 0.002714 | 0.004166 | 0.004055 | 4.6e-05 | 0.000986 | 0.26897 | 1.3432 | 2.5476 | 98.53 | 20.04 | 49.597 | 78.0601 | 86.4137 | 75.3099 | 0.01181 | 1.0132 | 0.023525 | 23.2192 | 0.09384 | 0.15579 | 0.00396 | 0.2724 | 0.2724 | 0.023048
2019-04-09 14:44 | 1554835456.5352 | 575 | 0.20872 | 0.001008 | 0.000112 | 92.704 | 91.5861 | 118.6574 | 20.5392 | 0.002457 | 0.002298 | 245.341 | 0.26196 | 8.3994 | 0.002563 | 0.003674 | 0.003687 | 4.9e-05 | 0.000986 | 0.28584 | 1.3241 | 2.5475 | 98.53 | 20.04 | 49.597 | 80.2396 | 82.7805 | 76.4002 | 0.01181 | 1.01 | 0.02348 | 23.2475 | 0.10851 | 0.15584 | 0.003955 | 0.27152 | 0.27152 | 0.023131
2019-04-09 14:44 | 1554835457.5348 | 576 | 0.20872 | 0.001004 | 0.000113 | 92.7633 | 91.9004 | 118.2025 | 20.5395 | 0.002413 | 0.002241 | 245.3603 | 0.26253 | 8.4033 | 0.002674 | 0.004641 | 0.004543 | 4.4e-05 | 0.000986 | 0.28335 | 1.3109 | 2.5475 | 98.53 | 20.04 | 49.596 | 79.172 | 85.4539 | 75.0342 | 0.01181 | 1.014 | 0.023481 | 23.1575 | 0.083594 | 0.15618 | 0.003971 | 0.27265 | 0.27265 | 0.023085
2019-04-09 14:44 | 1554835458.5329 | 577 | 0.20874 | 0.001008 | 0.000113 | 92.7862 | 91.8174 | 117.8894 | 20.5406 | 0.002406 | 0.002305 | 245.3132 | 0.26162 | 8.4031 | 0.002615 | 0.004534 | 0.004498 | 4.6e-05 | 0.000996 | 0.32261 | 1.32 | 2.5475 | 98.52 | 20.04 | 49.596 | 78.9953 | 84.0269 | 76.5842 | 0.011811 | 1.0095 | 0.023398 | 23.1783 | 0.093881 | 0.15564 | 0.003878 | 0.26514 | 0.26514 | 0.023656
2019-04-09 14:44 | 1554835459.5329 | 578 | 0.20874 | 0.001003 | 0.000113 | 92.9955 | 91.9998 | 117.9783 | 20.5407 | 0.002369 | 0.002264 | 245.3102 | 0.26237 | 8.4007 | 0.002684 | 0.003984 | 0.003915 | 4.8e-05 | 0.000994 | 0.2756 | 1.3042 | 2.5476 | 98.52 | 20.04 | 49.596 | 78.1268 | 85.6923 | 75.5848 | 0.011811 | 1.0124 | 0.02344 | 23.1539 | 0.096288 | 0.15608 | 0.003863 | 0.26454 | 0.26454 | 0.023777
2019-04-09 14:44 | 1554835460.5332 | 579 | 0.20876 | 0.001004 | 0.000115 | 93.1552 | 92.3488 | 117.9192 | 20.5405 | 0.002414 | 0.002253 | 245.2999 | 0.26232 | 8.4027 | 0.00264 | 0.00417 | 0.004043 | 4.8e-05 | 0.000997 | 0.29789 | 1.278 | 2.5475 | 98.52 | 20.04 | 49.595 | 79.2008 | 84.6229 | 75.3383 | 0.011811 | 1.0131 | 0.023442 | 23.1399 | 0.077763 | 0.15605 | 0.003745 | 0.2562 | 0.2562 | 0.024547
2019-04-09 14:44 | 1554835461.534 | 580 | 0.20872 | 0.001004 | 0.000115 | 93.2862 | 92.186 | 117.9854 | 20.5402 | 0.002447 | 0.002229 | 245.2526 | 0.26186 | 8.4039 | 0.002639 | 0.003944 | 0.003879 | 4.5e-05 | 0.000992 | 0.28512 | 1.2915 | 2.5475 | 98.52 | 20.04 | 49.592 | 79.9866 | 84.6132 | 74.7425 | 0.011811 | 1.0148 | 0.023469 | 23.1266 | 0.10612 | 0.15578 | 0.004011 | 0.27517 | 0.27517 | 0.022815
2019-04-09 14:44 | 1554835462.5333 | 581 | 0.20875 | 0.001001 | 0.000116 | 93.313 | 92.3075 | 117.6617 | 20.5402 | 0.002418 | 0.002215 | 245.2674 | 0.2622 | 8.4029 | 0.002706 | 0.004761 | 0.004725 | 4.9e-05 | 0.000983 | 0.29264 | 1.3171 | 2.5476 | 98.52 | 20.04 | 49.59 | 79.2887 | 86.2131 | 74.4167 | 0.01181 | 1.0158 | 0.023448 | 23.084 | 0.096905 | 0.15598 | 0.003828 | 0.26201 | 0.26201 | 0.023992
2019-04-09 14:44 | 1554835463.5331 | 582 | 0.20873 | 0.001001 | 0.000116 | 93.4253 | 92.4414 | 117.6124 | 20.54 | 0.00239 | 0.002229 | 245.2261 | 0.26237 | 8.4039 | 0.002668 | 0.004296 | 0.004223 | 4.8e-05 | 0.000987 | 0.29518 | 1.3252 | 2.5476 | 98.52 | 20.04 | 49.589 | 78.6275 | 85.3108 | 74.7486 | 0.01181 | 1.0148 | 0.023432 | 23.0902 | 0.094697 | 0.15608 | 0.003905 | 0.26723 | 0.26723 | 0.023539
2019-04-09 14:44 | 1554835464.5331 | 583 | 0.20875 | 0.000999 | 0.000117 | 93.5113 | 92.4606 | 117.9884 | 20.5403 | 0.002448 | 0.002279 | 245.2376 | 0.26229 | 8.4023 | 0.002647 | 0.004753 | 0.004683 | 4.6e-05 | 0.00098 | 0.29695 | 1.3335 | 2.5477 | 98.52 | 20.04 | 49.587 | 80.0131 | 84.7921 | 75.9438 | 0.011809 | 1.0113 | 0.023429 | 23.1668 | 0.10107 | 0.15604 | 0.003798 | 0.25965 | 0.25965 | 0.024219
2019-04-09 14:44 | 1554835465.5331 | 584 | 0.20876 | 0.000998 | 0.000116 | 93.595 | 92.6613 | 117.7258 | 20.5391 | 0.002478 | 0.002243 | 245.1762 | 0.26262 | 8.402 | 0.002706 | 0.003967 | 0.003921 | 4.7e-05 | 0.00098 | 0.27345 | 1.3263 | 2.5476 | 98.52 | 20.04 | 49.584 | 80.7278 | 86.2245 | 75.0977 | 0.011809 | 1.0138 | 0.023431 | 23.1129 | 0.08968 | 0.15623 | 0.003781 | 0.25851 | 0.25851 | 0.024355
2019-04-09 14:44 | 1554835466.5342 | 585 | 0.20872 | 0.000995 | 0.000116 | 93.587 | 92.622 | 117.5913 | 20.5393 | 0.002425 | 0.002235 | 245.1671 | 0.26197 | 8.3989 | 0.002599 | 0.003982 | 0.003988 | 4.5e-05 | 0.000982 | 0.28224 | 1.336 | 2.5476 | 98.52 | 20.04 | 49.583 | 79.4688 | 83.6387 | 74.9055 | 0.011808 | 1.0143 | 0.023424 | 23.0933 | 0.0927 | 0.15584 | 0.003982 | 0.27255 | 0.27255 | 0.023043
2019-04-09 14:44 | 1554835467.5345 | 586 | 0.20876 | 0.000996 | 0.000115 | 93.5756 | 92.5532 | 117.5333 | 20.5404 | 0.002411 | 0.002258 | 245.139 | 0.26252 | 8.4038 | 0.002571 | 0.004257 | 0.004158 | 4.6e-05 | 0.000988 | 0.30198 | 1.3057 | 2.5476 | 98.52 | 20.04 | 49.581 | 79.1236 | 82.9728 | 75.4505 | 0.011808 | 1.0128 | 0.0234 | 23.1057 | 0.098268 | 0.15617 | 0.003738 | 0.25522 | 0.25522 | 0.02466
2019-04-09 14:44 | 1554835468.5331 | 587 | 0.20882 | 0.000993 | 0.000115 | 93.6864 | 92.6245 | 117.9748 | 20.5402 | 0.002462 | 0.002259 | 245.0978 | 0.26239 | 8.4045 | 0.002684 | 0.004657 | 0.00456 | 4.2e-05 | 0.000991 | 0.29911 | 1.3166 | 2.5475 | 98.52 | 20.04 | 49.579 | 80.3506 | 85.6949 | 75.4651 | 0.011807 | 1.0127 | 0.023444 | 23.1495 | 0.10197 | 0.1561 | 0.003396 | 0.2317 | 0.2317 | 0.027151
2019-04-09 14:44 | 1554835469.5336 | 588 | 0.20876 | 0.001 | 0.000116 | 93.6957 | 92.6785 | 117.9626 | 20.5415 | 0.00243 | 0.002278 | 245.0923 | 0.26275 | 8.4009 | 0.002673 | 0.004465 | 0.004455 | 4.5e-05 | 0.000987 | 0.27007 | 1.3248 | 2.5476 | 98.52 | 20.04 | 49.578 | 79.5721 | 85.4284 | 75.9387 | 0.011807 | 1.0113 | 0.023427 | 23.1641 | 0.097633 | 0.15631 | 0.003727 | 0.25474 | 0.25474 | 0.024728
2019-04-09 14:44 | 1554835470.5337 | 589 | 0.2088 | 0.000985 | 0.000115 | 93.7927 | 92.7002 | 118.0539 | 20.5404 | 0.002452 | 0.002251 | 245.0501 | 0.262 | 8.4018 | 0.002624 | 0.004 | 0.003936 | 4.5e-05 | 0.000992 | 0.28313 | 1.3546 | 2.5477 | 98.52 | 20.04 | 49.577 | 80.1029 | 84.2359 | 75.283 | 0.011807 | 1.0132 | 0.023458 | 23.1512 | 0.1048 | 0.15587 | 0.003503 | 0.23931 | 0.23931 | 0.026248
2019-04-09 14:44 | 1554835471.5333 | 590 | 0.20884 | 0.000994 | 0.000116 | 93.8627 | 92.8664 | 118.0025 | 20.5405 | 0.002464 | 0.002268 | 245.0703 | 0.26189 | 8.4026 | 0.002667 | 0.004294 | 0.004269 | 4.6e-05 | 0.000991 | 0.28694 | 1.3379 | 2.5474 | 98.52 | 20.04 | 49.577 | 80.4024 | 85.2808 | 75.6783 | 0.011807 | 1.0121 | 0.023439 | 23.1593 | 0.09545 | 0.1558 | 0.00326 | 0.22203 | 0.22203 | 0.028279
2019-04-09 14:45 | 1554835472.5334 | 591 | 0.20877 | 0.00099 | 0.000117 | 93.7841 | 92.6854 | 118.4299 | 20.5403 | 0.00242 | 0.002283 | 245.0152 | 0.26258 | 8.4008 | 0.002634 | 0.004179 | 0.004101 | 4.6e-05 | 0.000988 | 0.27113 | 1.358 | 2.5476 | 98.52 | 20.04 | 49.577 | 79.3419 | 84.4909 | 76.0421 | 0.011807 | 1.011 | 0.023469 | 23.2133 | 0.10541 | 0.15621 | 0.003723 | 0.25485 | 0.25485 | 0.024702
2019-04-09 14:45 | 1554835473.5344 | 592 | 0.2088 | 0.000987 | 0.000117 | 93.8928 | 92.6544 | 118.2771 | 20.5407 | 0.00241 | 0.002278 | 245.013 | 0.263 | 8.403 | 0.002574 | 0.004817 | 0.004728 | 4.3e-05 | 0.000994 | 0.31365 | 1.3583 | 2.5476 | 98.52 | 20.04 | 49.578 | 79.0999 | 83.0426 | 75.938 | 0.011807 | 1.0113 | 0.023458 | 23.1949 | 0.11876 | 0.15646 | 0.003511 | 0.2398 | 0.2398 | 0.026294
2019-04-09 14:45 | 1554835474.5324 | 593 | 0.20875 | 0.000996 | 0.000118 | 93.9416 | 92.7373 | 117.7873 | 20.5408 | 0.002435 | 0.002271 | 244.9945 | 0.26227 | 8.4025 | 0.002667 | 0.004408 | 0.004361 | 4.2e-05 | 0.000993 | 0.27649 | 1.3385 | 2.5476 | 98.52 | 20.04 | 49.579 | 79.7111 | 85.2693 | 75.7541 | 0.011807 | 1.0119 | 0.023415 | 23.1407 | 0.11541 | 0.15602 | 0.003791 | 0.25897 | 0.25897 | 0.02428
2019-04-09 14:45 | 1554835475.5334 | 594 | 0.2088 | 0.000994 | 0.000118 | 93.8847 | 92.8836 | 118.0757 | 20.5409 | 0.002466 | 0.002255 | 244.9705 | 0.26213 | 8.3993 | 0.002663 | 0.004197 | 0.004123 | 4.3e-05 | 0.000994 | 0.28031 | 1.3288 | 2.5476 | 98.52 | 20.04 | 49.582 | 80.4434 | 85.1828 | 75.3793 | 0.011808 | 1.013 | 0.023457 | 23.1566 | 0.095893 | 0.15594 | 0.003502 | 0.2391 | 0.2391 | 0.026284
2019-04-09 14:45 | 1554835476.5324 | 595 | 0.20881 | 0.000992 | 0.000118 | 93.9973 | 92.8382 | 118.472 | 20.5407 | 0.00241 | 0.00224 | 244.9811 | 0.26171 | 8.4013 | 0.002758 | 0.004526 | 0.004466 | 4.2e-05 | 0.000991 | 0.29856 | 1.3004 | 2.5475 | 98.52 | 20.04 | 49.583 | 79.088 | 87.4689 | 75.0127 | 0.011808 | 1.014 | 0.023508 | 23.1832 | 0.11099 | 0.15569 | 0.003466 | 0.23709 | 0.23709 | 0.026463
2019-04-09 14:45 | 1554835477.533 | 596 | 0.20875 | 0.000991 | 0.000118 | 94.075 | 93.0326 | 118.6076 | 20.5403 | 0.002428 | 0.002204 | 244.9373 | 0.26159 | 8.4007 | 0.002654 | 0.00447 | 0.00445 | 4.6e-05 | 0.000992 | 0.28528 | 1.3028 | 2.5475 | 98.52 | 20.04 | 49.583 | 79.5421 | 84.9544 | 74.1511 | 0.011808 | 1.0165 | 0.023551 | 23.1678 | 0.099665 | 0.15562 | 0.003805 | 0.26143 | 0.26143 | 0.023989
2019-04-09 14:45 | 1554835478.5334 | 597 | 0.20875 | 0.000989 | 0.000118 | 94.1179 | 93.0653 | 118.0725 | 20.5415 | 0.002403 | 0.002311 | 244.9364 | 0.26205 | 8.3994 | 0.002708 | 0.004112 | 0.004071 | 4.4e-05 | 0.000989 | 0.27903 | 1.2805 | 2.5476 | 98.52 | 20.04 | 49.584 | 78.9266 | 86.2565 | 76.7229 | 0.011809 | 1.0091 | 0.023411 | 23.2009 | 0.1006 | 0.15589 | 0.003819 | 0.26083 | 0.26083 | 0.024086
2019-04-09 14:45 | 1554835479.534 | 598 | 0.20879 | 0.000987 | 0.000118 | 94.1745 | 93.2484 | 118.2095 | 20.5404 | 0.002482 | 0.002246 | 244.8985 | 0.26195 | 8.4018 | 0.002697 | 0.004553 | 0.004475 | 4.3e-05 | 0.000993 | 0.274 | 1.3107 | 2.5475 | 98.52 | 20.04 | 49.583 | 80.8209 | 85.9925 | 75.1684 | 0.011808 | 1.0136 | 0.023477 | 23.1627 | 0.088394 | 0.15584 | 0.003604 | 0.24645 | 0.24645 | 0.025482
2019-04-09 14:45 | 1554835480.5337 | 599 | 0.20878 | 0.000986 | 0.000118 | 94.2233 | 93.2326 | 117.6917 | 20.5409 | 0.002422 | 0.002274 | 244.8632 | 0.26212 | 8.4036 | 0.002693 | 0.004492 | 0.004429 | 4.5e-05 | 0.00099 | 0.29966 | 1.3313 | 2.5475 | 98.52 | 20.04 | 49.581 | 79.386 | 85.9046 | 75.8322 | 0.011808 | 1.0116 | 0.023403 | 23.1339 | 0.09455 | 0.15594 | 0.003675 | 0.25066 | 0.25066 | 0.025071
2019-04-09 14:45 | 1554835481.5335 | 600 | 0.20883 | 0.000989 | 0.000119 | 94.2892 | 93.3819 | 117.5861 | 20.5406 | 0.002462 | 0.002317 | 244.8388 | 0.2624 | 8.4012 | 0.002664 | 0.003936 | 0.003908 | 4.1e-05 | 0.000987 | 0.30431 | 1.3121 | 2.5475 | 98.52 | 20.04 | 49.579 | 80.3525 | 85.2152 | 76.87 | 0.011807 | 1.0086 | 0.023358 | 23.1579 | 0.086477 | 0.1561 | 0.00337 | 0.22881 | 0.22881 | 0.027493
2019-04-09 14:45 | 1554835482.5326 | 601 | 0.20878 | 0.000991 | 0.00012 | 94.4267 | 93.4614 | 117.3298 | 20.541 | 0.002432 | 0.002251 | 244.7916 | 0.26283 | 8.4009 | 0.002701 | 0.004012 | 0.003972 | 4.2e-05 | 0.000983 | 0.28047 | 1.3174 | 2.5476 | 98.52 | 20.04 | 49.575 | 79.6201 | 86.1061 | 75.2773 | 0.011806 | 1.0133 | 0.023386 | 23.0799 | 0.091908 | 0.15636 | 0.003626 | 0.24691 | 0.24691 | 0.02552
2019-04-09 14:45 | 1554835483.5337 | 602 | 0.20874 | 0.000988 | 0.00012 | 94.5858 | 93.5895 | 117.6375 | 20.541 | 0.002435 | 0.002249 | 244.7995 | 0.26259 | 8.4026 | 0.002654 | 0.004341 | 0.004303 | 4.5e-05 | 0.000982 | 0.27959 | 1.309 | 2.5475 | 98.52 | 20.04 | 49.568 | 79.698 | 84.9769 | 75.2422 | 0.011805 | 1.0134 | 0.023418 | 23.109 | 0.094714 | 0.15621 | 0.003917 | 0.26764 | 0.26764 | 0.023522
2019-04-09 14:45 | 1554835484.5331 | 603 | 0.2088 | 0.000996 | 0.000121 | 94.6087 | 93.6142 | 117.5171 | 20.541 | 0.00243 | 0.002278 | 244.7262 | 0.26212 | 8.4001 | 0.002635 | 0.004805 | 0.004709 | 4.6e-05 | 0.00098 | 0.27632 | 1.3046 | 2.5474 | 98.52 | 20.04 | 49.565 | 79.5819 | 84.5173 | 75.9212 | 0.011804 | 1.0114 | 0.023383 | 23.1197 | 0.094514 | 0.15594 | 0.003493 | 0.23754 | 0.23754 | 0.026456
2019-04-09 14:45 | 1554835485.5328 | 604 | 0.20879 | 0.000991 | 0.000121 | 94.6534 | 93.6845 | 117.422 | 20.5416 | 0.002427 | 0.00225 | 244.7124 | 0.26233 | 8.3993 | 0.002597 | 0.00456 | 0.004463 | 4.1e-05 | 0.000982 | 0.28479 | 1.3415 | 2.5475 | 98.52 | 20.04 | 49.556 | 79.5145 | 83.601 | 75.2498 | 0.011802 | 1.0133 | 0.023396 | 23.0881 | 0.092031 | 0.15606 | 0.003555 | 0.24197 | 0.24197 | 0.025992
2019-04-09 14:45 | 1554835486.5328 | 605 | 0.20882 | 0.000997 | 0.000122 | 94.6831 | 93.8452 | 117.2789 | 20.54 | 0.002368 | 0.002284 | 244.6777 | 0.26238 | 8.4013 | 0.00263 | 0.004286 | 0.00428 | 4.5e-05 | 0.000978 | 0.27649 | 1.3703 | 2.5475 | 98.52 | 20.04 | 49.55 | 78.0951 | 84.3844 | 76.0822 | 0.011801 | 1.0109 | 0.023354 | 23.1016 | 0.079513 | 0.15609 | 0.003378 | 0.22913 | 0.22913 | 0.027453
2019-04-09 14:45 | 1554835487.5327 | 606 | 0.20874 | 0.000999 | 0.000122 | 94.8386 | 93.9181 | 117.3089 | 20.5409 | 0.00244 | 0.002251 | 244.6594 | 0.26242 | 8.4022 | 0.002657 | 0.004206 | 0.004112 | 4.5e-05 | 0.000983 | 0.27168 | 1.3507 | 2.5475 | 98.52 | 20.04 | 49.545 | 79.8164 | 85.0391 | 75.2686 | 0.011799 | 1.0133 | 0.023384 | 23.0776 | 0.087234 | 0.15611 | 0.003879 | 0.26452 | 0.26452 | 0.023784
2019-04-09 14:45 | 1554835488.5329 | 607 | 0.20879 | 0.001002 | 0.000123 | 94.9282 | 93.9815 | 117.2188 | 20.5414 | 0.0024 | 0.002303 | 244.6739 | 0.26255 | 8.4024 | 0.002616 | 0.004215 | 0.004162 | 4.3e-05 | 0.000985 | 0.27577 | 1.3294 | 2.5475 | 98.52 | 20.05 | 49.533 | 78.8588 | 84.0471 | 76.5304 | 0.011804 | 1.0096 | 0.023333 | 23.1105 | 0.089655 | 0.15619 | 0.003539 | 0.24017 | 0.24017 | 0.026208
2019-04-09 14:45 | 1554835489.5329 | 608 | 0.20885 | 0.001001 | 0.000124 | 95.0661 | 94.0255 | 116.9944 | 20.5405 | 0.002429 | 0.002222 | 244.6373 | 0.26206 | 8.4018 | 0.002659 | 0.004499 | 0.004464 | 4.5e-05 | 0.000992 | 0.25393 | 1.3718 | 2.5475 | 98.52 | 20.05 | 49.528 | 79.5682 | 85.0888 | 74.5784 | 0.011803 | 1.0153 | 0.023376 | 23.0238 | 0.098446 | 0.1559 | 0.003194 | 0.2164 | 0.2164 | 0.029033
2019-04-09 14:45 | 1554835490.5319 | 609 | 0.20869 | 0.000995 | 0.000124 | 95.2463 | 94.3796 | 117.5044 | 20.5408 | 0.002391 | 0.002243 | 244.6638 | 0.26186 | 8.4019 | 0.002628 | 0.004834 | 0.004802 | 4.4e-05 | 0.000988 | 0.28191 | 1.3259 | 2.5476 | 98.52 | 20.04 | 49.516 | 78.6414 | 84.3322 | 75.076 | 0.011792 | 1.0138 | 0.02341 | 23.0904 | 0.081756 | 0.15578 | 0.004141 | 0.2831 | 0.2831 | 0.022176
2019-04-09 14:45 | 1554835491.5329 | 610 | 0.2088 | 0.000994 | 0.000124 | 95.3844 | 94.3786 | 117.445 | 20.5413 | 0.002446 | 0.002287 | 244.6068 | 0.26157 | 8.4003 | 0.002669 | 0.003891 | 0.003826 | 4.7e-05 | 0.000993 | 0.28744 | 1.3256 | 2.5475 | 98.52 | 20.04 | 49.511 | 79.9759 | 85.3253 | 76.1567 | 0.011791 | 1.0107 | 0.023368 | 23.1204 | 0.09482 | 0.15561 | 0.003537 | 0.24027 | 0.24027 | 0.0261
2019-04-09 14:45 | 1554835492.5323 | 611 | 0.20878 | 0.000992 | 0.000124 | 95.5458 | 94.5393 | 117.3549 | 20.5408 | 0.002451 | 0.002273 | 244.6167 | 0.26233 | 8.4025 | 0.002645 | 0.004466 | 0.004376 | 4.6e-05 | 0.000994 | 0.28064 | 1.3152 | 2.5475 | 98.52 | 20.04 | 49.506 | 80.0931 | 84.7449 | 75.8102 | 0.01179 | 1.0117 | 0.02337 | 23.1001 | 0.094724 | 0.15606 | 0.00366 | 0.24891 | 0.24891 | 0.025267
2019-04-09 14:45 | 1554835493.533 | 612 | 0.2088 | 0.00099 | 0.000123 | 95.7454 | 94.53 | 117.2016 | 20.5412 | 0.002437 | 0.002258 | 244.566 | 0.26262 | 8.4016 | 0.002619 | 0.004269 | 0.004187 | 4.6e-05 | 0.000995 | 0.28456 | 1.3394 | 2.5475 | 98.52 | 20.04 | 49.498 | 79.7413 | 84.1297 | 75.4592 | 0.011788 | 1.0127 | 0.023367 | 23.0734 | 0.11426 | 0.15623 | 0.003524 | 0.23938 | 0.23938 | 0.026302
2019-04-09 14:45 | 1554835494.533 | 613 | 0.20873 | 0.000989 | 0.000123 | 95.8511 | 95.0604 | 117.4242 | 20.541 | 0.002388 | 0.002276 | 244.5836 | 0.26197 | 8.3996 | 0.002697 | 0.00407 | 0.004096 | 4.7e-05 | 0.000993 | 0.27505 | 1.3067 | 2.5476 | 98.52 | 20.04 | 49.494 | 78.5708 | 85.9914 | 75.8696 | 0.011787 | 1.0115 | 0.023375 | 23.1088 | 0.074091 | 0.15585 | 0.003941 | 0.26869 | 0.26869 | 0.023375
2019-04-09 14:45 | 1554835495.5334 | 614 | 0.20878 | 0.000993 | 0.000124 | 95.9839 | 95.124 | 117.6774 | 20.5413 | 0.002449 | 0.002251 | 244.5898 | 0.26266 | 8.4008 | 0.00262 | 0.004606 | 0.004503 | 4.9e-05 | 0.00099 | 0.274 | 1.3477 | 2.5477 | 98.52 | 20.05 | 49.491 | 80.0312 | 84.1502 | 75.2697 | 0.011794 | 1.0133 | 0.023421 | 23.1139 | 0.080488 | 0.15626 | 0.003656 | 0.24914 | 0.24914 | 0.025275
2019-04-09 14:45 | 1554835496.5331 | 615 | 0.2086 | 0.000989 | 0.000125 | 96.0973 | 95.0084 | 117.5577 | 20.541 | 0.002458 | 0.002212 | 244.5464 | 0.26173 | 8.4019 | 0.002592 | 0.00381 | 0.00377 | 5.1e-05 | 0.00099 | 0.30021 | 1.3328 | 2.5475 | 98.52 | 20.04 | 49.482 | 80.2525 | 83.4788 | 74.3404 | 0.011784 | 1.016 | 0.02344 | 23.0713 | 0.10193 | 0.1557 | 0.004745 | 0.32568 | 0.32568 | 0.019267
2019-04-09 14:45 | 1554835497.5331 | 616 | 0.20915 | 0.000989 | 0.000125 | 96.1521 | 95.1505 | 117.3425 | 20.5415 | 0.002402 | 0.002245 | 244.558 | 0.26252 | 8.4014 | 0.002644 | 0.004595 | 0.004502 | 4.7e-05 | 0.000991 | 0.28854 | 1.3278 | 2.5474 | 98.52 | 20.04 | 49.478 | 78.9065 | 84.7307 | 75.1312 | 0.011783 | 1.0137 | 0.023392 | 23.0764 | 0.093662 | 0.15617 | 0.001429 | 0.092797 | 0.092797 | 0.067823
2019-04-09 14:45 | 1554835498.5321 | 617 | 0.20913 | 0.000996 | 0.000125 | 96.2778 | 95.398 | 117.5868 | 20.5415 | 0.002482 | 0.00225 | 244.4896 | 0.26186 | 8.4016 | 0.002655 | 0.004613 | 0.004612 | 5.3e-05 | 0.000991 | 0.26698 | 1.3673 | 2.5475 | 98.52 | 20.04 | 49.472 | 80.8268 | 84.99 | 75.2574 | 0.011782 | 1.0133 | 0.023412 | 23.1046 | 0.082113 | 0.15578 | 0.001495 | 0.097517 | 0.097517 | 0.064379
2019-04-09 14:45 | 1554835499.5336 | 618 | 0.20892 | 0.000994 | 0.000124 | 96.3984 | 95.6802 | 118.0767 | 20.5413 | 0.002463 | 0.002269 | 244.4927 | 0.26198 | 8.4026 | 0.002675 | 0.004586 | 0.004526 | 5.1e-05 | 0.00099 | 0.27721 | 1.3371 | 2.5476 | 98.52 | 20.04 | 49.469 | 80.375 | 85.4672 | 75.7172 | 0.011781 | 1.012 | 0.023445 | 23.1679 | 0.06689 | 0.15585 | 0.002758 | 0.18636 | 0.18636 | 0.033703
2019-04-09 14:45 | 1554835500.5322 | 619 | 0.20895 | 0.000997 | 0.000123 | 96.4913 | 95.6716 | 118.2776 | 20.5407 | 0.002443 | 0.002294 | 244.4626 | 0.26234 | 8.4007 | 0.002578 | 0.004062 | 0.004018 | 4.8e-05 | 0.000996 | 0.27959 | 1.3423 | 2.5474 | 98.52 | 20.04 | 49.467 | 79.8888 | 83.1337 | 76.3245 | 0.011781 | 1.0102 | 0.023445 | 23.2078 | 0.076311 | 0.15607 | 0.002588 | 0.17446 | 0.17446 | 0.036051
2019-04-09 14:45 | 1554835501.5326 | 620 | 0.20852 | 0.001002 | 0.000124 | 96.5911 | 95.8949 | 118.1628 | 20.5416 | 0.002463 | 0.002317 | 244.4566 | 0.26259 | 8.3999 | 0.002565 | 0.003936 | 0.003832 | 4.6e-05 | 0.000994 | 0.27527 | 1.3164 | 2.5475 | 98.52 | 20.04 | 49.461 | 80.3747 | 82.835 | 76.8766 | 0.011779 | 1.0086 | 0.023415 | 23.2148 | 0.064697 | 0.15621 | 0.005225 | 0.35908 | 0.35908 | 0.017532
2019-04-09 14:45 | 1554835502.5323 | 621 | 0.20883 | 0.000997 | 0.000123 | 96.6221 | 95.7497 | 118.0955 | 20.5415 | 0.002381 | 0.002283 | 244.4586 | 0.2621 | 8.4012 | 0.002582 | 0.003969 | 0.003958 | 5.1e-05 | 0.000995 | 0.27527 | 1.3044 | 2.5475 | 98.52 | 20.04 | 49.458 | 78.4029 | 83.2335 | 76.0377 | 0.011779 | 1.011 | 0.023436 | 23.1804 | 0.081121 | 0.15592 | 0.003338 | 0.22712 | 0.22712 | 0.027667
2019-04-09 14:45 | 1554835503.5322 | 622 | 0.20874 | 0.000996 | 0.000123 | 96.6855 | 96.0471 | 118.1635 | 20.5417 | 0.002408 | 0.002265 | 244.4155 | 0.2621 | 8.4009 | 0.002589 | 0.004157 | 0.004136 | 4.6e-05 | 0.000998 | 0.27776 | 1.3426 | 2.5476 | 98.52 | 20.04 | 49.454 | 79.0634 | 83.3922 | 75.6139 | 0.011778 | 1.0123 | 0.023457 | 23.173 | 0.059257 | 0.15593 | 0.003876 | 0.26509 | 0.26509 | 0.023705
2019-04-09 14:45 | 1554835504.5323 | 623 | 0.20875 | 0.000999 | 0.000123 | 96.7864 | 96.2116 | 118.2997 | 20.5413 | 0.002363 | 0.002275 | 244.3967 | 0.26214 | 8.3999 | 0.002568 | 0.003752 | 0.0037 | 4.6e-05 | 0.001 | 0.28567 | 1.3604 | 2.5476 | 98.52 | 20.04 | 49.451 | 77.981 | 82.9003 | 75.8504 | 0.011777 | 1.0116 | 0.023463 | 23.1942 | 0.053284 | 0.15594 | 0.003777 | 0.25813 | 0.25813 | 0.024347
2019-04-09 14:45 | 1554835505.5313 | 624 | 0.20866 | 0.000998 | 0.000125 | 96.9008 | 96.0512 | 118.2586 | 20.5417 | 0.002386 | 0.002285 | 244.3276 | 0.26217 | 8.4026 | 0.002623 | 0.004774 | 0.004737 | 4.4e-05 | 0.000997 | 0.29031 | 1.3287 | 2.5475 | 98.52 | 20.04 | 49.449 | 78.5278 | 84.2323 | 76.098 | 0.011776 | 1.0109 | 0.023451 | 23.1984 | 0.078768 | 0.15597 | 0.004356 | 0.29861 | 0.29861 | 0.021049
2019-04-09 14:45 | 1554835506.5315 | 625 | 0.20878 | 0.000999 | 0.000126 | 96.9731 | 96.2697 | 118.3994 | 20.5425 | 0.00238 | 0.002281 | 244.3321 | 0.26284 | 8.4013 | 0.002632 | 0.004055 | 0.003979 | 4.5e-05 | 0.001 | 0.2844 | 1.3335 | 2.5476 | 98.52 | 20.04 | 49.442 | 78.3913 | 84.4337 | 76.0084 | 0.011775 | 1.0111 | 0.023468 | 23.2092 | 0.065117 | 0.15636 | 0.003641 | 0.2485 | 0.2485 | 0.025358
2019-04-09 14:45 | 1554835507.5314 | 626 | 0.20868 | 0.000994 | 0.000127 | 97.086 | 96.3207 | 118.5522 | 20.5415 | 0.002487 | 0.002269 | 244.2814 | 0.26246 | 8.403 | 0.002656 | 0.004034 | 0.003905 | 4.4e-05 | 0.000998 | 0.28744 | 1.323 | 2.5477 | 98.52 | 20.04 | 49.439 | 80.9453 | 85.0133 | 75.7072 | 0.011774 | 1.012 | 0.023493 | 23.2142 | 0.07079 | 0.15614 | 0.004241 | 0.29091 | 0.29091 | 0.02163
2019-04-09 14:45 | 1554835508.5318 | 627 | 0.20877 | 0.000997 | 0.000128 | 97.1351 | 96.3513 | 118.7205 | 20.5417 | 0.00243 | 0.002276 | 244.2514 | 0.26247 | 8.4004 | 0.002709 | 0.003883 | 0.003916 | 4.6e-05 | 0.000998 | 0.28512 | 1.3388 | 2.5477 | 98.52 | 20.05 | 49.435 | 79.574 | 86.2872 | 75.8882 | 0.01178 | 1.0115 | 0.023503 | 23.2367 | 0.072472 | 0.15615 | 0.003707 | 0.25342 | 0.25342 | 0.024831
2019-04-09 14:45 | 1554835509.5324 | 628 | 0.20864 | 0.000997 | 0.000129 | 97.3463 | 96.3861 | 118.537 | 20.5418 | 0.002412 | 0.002215 | 244.2476 | 0.26196 | 8.4036 | 0.002618 | 0.004008 | 0.003986 | 4.8e-05 | 0.000994 | 0.27151 | 1.3524 | 2.5476 | 98.52 | 20.05 | 49.429 | 79.1538 | 84.0905 | 74.4182 | 0.011779 | 1.0158 | 0.023535 | 23.1698 | 0.088667 | 0.15584 | 0.004487 | 0.30864 | 0.30864 | 0.020348
2019-04-09 14:45 | 1554835510.5316 | 629 | 0.2087 | 0.001003 | 0.00013 | 97.298 | 96.1558 | 118.0664 | 20.5411 | 0.002449 | 0.002289 | 244.1776 | 0.26156 | 8.4021 | 0.002623 | 0.003825 | 0.003724 | 4.6e-05 | 0.000994 | 0.32405 | 1.4049 | 2.5476 | 98.52 | 20.05 | 49.424 | 80.0444 | 84.214 | 76.1947 | 0.011778 | 1.0106 | 0.023428 | 23.1827 | 0.10562 | 0.1556 | 0.004075 | 0.27831 | 0.27831 | 0.022531
2019-04-09 14:45 | 1554835511.5325 | 630 | 0.20876 | 0.001002 | 0.000131 | 97.4121 | 96.4132 | 118.2459 | 20.5421 | 0.00246 | 0.002257 | 244.1784 | 0.262 | 8.4025 | 0.002683 | 0.004185 | 0.00417 | 4.5e-05 | 0.000989 | 0.29391 | 1.4012 | 2.5477 | 98.52 | 20.05 | 49.416 | 80.3049 | 85.664 | 75.435 | 0.011776 | 1.0128 | 0.023472 | 23.1751 | 0.092189 | 0.15586 | 0.003732 | 0.25468 | 0.25468 | 0.024664
2019-04-09 14:45 | 1554835512.5326 | 631 | 0.20867 | 0.001001 | 0.000131 | 97.4824 | 96.6149 | 118.6454 | 20.5413 | 0.002457 | 0.002302 | 244.1071 | 0.26206 | 8.4016 | 0.002629 | 0.00456 | 0.004549 | 4.4e-05 | 0.000989 | 0.30182 | 1.3731 | 2.5477 | 98.52 | 20.05 | 49.413 | 80.2309 | 84.3724 | 76.4974 | 0.011775 | 1.0097 | 0.023476 | 23.2496 | 0.07995 | 0.1559 | 0.004286 | 0.2936 | 0.2936 | 0.021399
2019-04-09 14:45 | 1554835513.5326 | 632 | 0.20872 | 0.001007 | 0.000132 | 97.678 | 96.5543 | 118.9484 | 20.5423 | 0.00244 | 0.002291 | 244.1128 | 0.26324 | 8.4015 | 0.002587 | 0.004121 | 0.004012 | 4.2e-05 | 0.000989 | 0.26842 | 1.3424 | 2.5475 | 98.52 | 20.05 | 49.41 | 79.8156 | 83.3553 | 76.2427 | 0.011774 | 1.0105 | 0.023514 | 23.2708 | 0.10349 | 0.1566 | 0.003949 | 0.27032 | 0.27032 | 0.023346
2019-04-09 14:45 | 1554835514.533 | 633 | 0.20874 | 0.001006 | 0.000132 | 97.6992 | 96.6493 | 118.7994 | 20.5421 | 0.002409 | 0.002258 | 244.0678 | 0.26148 | 8.3998 | 0.002658 | 0.004518 | 0.004502 | 4.5e-05 | 0.000988 | 0.28423 | 1.3043 | 2.5475 | 98.52 | 20.05 | 49.403 | 79.0661 | 85.0552 | 75.4412 | 0.011773 | 1.0128 | 0.023526 | 23.2295 | 0.096634 | 0.15555 | 0.00388 | 0.26563 | 0.26563 | 0.0236
2019-04-09 14:45 | 1554835515.5317 | 634 | 0.20873 | 0.001006 | 0.000132 | 97.798 | 96.7357 | 118.3264 | 20.5422 | 0.002421 | 0.002277 | 244.0508 | 0.26225 | 8.4032 | 0.002629 | 0.004243 | 0.004165 | 4.4e-05 | 0.000992 | 0.27992 | 1.2858 | 2.5475 | 98.52 | 20.05 | 49.4 | 79.3757 | 84.3593 | 75.8957 | 0.011772 | 1.0115 | 0.023464 | 23.1983 | 0.09769 | 0.15601 | 0.00392 | 0.26772 | 0.26772 | 0.023484
2019-04-09 14:45 | 1554835516.5317 | 635 | 0.20869 | 0.001004 | 0.000133 | 97.8743 | 96.8641 | 118.0776 | 20.5417 | 0.002453 | 0.00225 | 244.011 | 0.26241 | 8.4025 | 0.002569 | 0.003779 | 0.003713 | 4.5e-05 | 0.000983 | 0.26272 | 1.3108 | 2.5474 | 98.52 | 20.05 | 49.394 | 80.1291 | 82.9187 | 75.253 | 0.011771 | 1.0133 | 0.023461 | 23.1526 | 0.092798 | 0.15611 | 0.004182 | 0.28607 | 0.28607 | 0.021991
2019-04-09 14:45 | 1554835517.5317 | 636 | 0.20872 | 0.001009 | 0.000133 | 97.9898 | 96.9309 | 118.1987 | 20.5418 | 0.00239 | 0.002268 | 243.9445 | 0.2626 | 8.4017 | 0.002593 | 0.004171 | 0.004144 | 4.3e-05 | 0.000985 | 0.29286 | 1.3054 | 2.5476 | 98.52 | 20.05 | 49.392 | 78.6161 | 83.5101 | 75.6967 | 0.01177 | 1.012 | 0.023458 | 23.1792 | 0.097188 | 0.15622 | 0.003935 | 0.2686 | 0.2686 | 0.023438
2019-04-09 14:45 | 1554835518.5317 | 637 | 0.2087 | 0.001009 | 0.000136 | 98.0366 | 96.9732 | 117.7488 | 20.5418 | 0.002464 | 0.002249 | 243.9363 | 0.26184 | 8.4019 | 0.002589 | 0.00432 | 0.004189 | 4.4e-05 | 0.000981 | 0.27489 | 1.3007 | 2.5475 | 98.52 | 20.05 | 49.391 | 80.3981 | 83.4132 | 75.2414 | 0.01177 | 1.0134 | 0.023429 | 23.1199 | 0.097549 | 0.15577 | 0.004069 | 0.27757 | 0.27757 | 0.022616
2019-04-09 14:45 | 1554835519.5311 | 638 | 0.20867 | 0.001008 | 0.00014 | 98.1152 | 96.9825 | 117.4425 | 20.5416 | 0.002391 | 0.002311 | 243.8653 | 0.26246 | 8.3998 | 0.002539 | 0.004009 | 0.003997 | 3.9e-05 | 0.000983 | 0.25299 | 1.328 | 2.5475 | 98.52 | 20.05 | 49.389 | 78.6499 | 82.2104 | 76.7136 | 0.011769 | 1.0091 | 0.023349 | 23.1386 | 0.10386 | 0.15614 | 0.004252 | 0.28916 | 0.28916 | 0.021761
2019-04-09 14:45 | 1554835520.5318 | 639 | 0.20859 | 0.00101 | 0.000144 | 98.2221 | 97.0344 | 117.4979 | 20.5419 | 0.002438 | 0.00224 | 243.8405 | 0.26206 | 8.4006 | 0.002577 | 0.00414 | 0.004071 | 4.1e-05 | 0.000982 | 0.28047 | 1.3279 | 2.5475 | 98.52 | 20.05 | 49.389 | 79.7652 | 83.1084 | 75.0084 | 0.011769 | 1.014 | 0.023412 | 23.0876 | 0.10881 | 0.1559 | 0.004762 | 0.32544 | 0.32544 | 0.019305
2019-04-09 14:45 | 1554835521.5323 | 640 | 0.20826 | 0.001019 | 0.000149 | 98.2581 | 96.7416 | 117.5594 | 20.5422 | 0.002431 | 0.002287 | 243.7769 | 0.2627 | 8.4017 | 0.002651 | 0.004678 | 0.004683 | 4e-05 | 0.000979 | 0.27201 | 1.3689 | 2.5476 | 98.52 | 20.05 | 49.388 | 79.611 | 84.8875 | 76.1444 | 0.011769 | 1.0107 | 0.02338 | 23.1312 | 0.13913 | 0.15628 | 0.006725 | 0.46203 | 0.46203 | 0.013631
2019-04-09 14:45 | 1554835522.5322 | 641 | 0.20786 | 0.00104 | 0.000153 | 98.2635 | 97.1542 | 117.1655 | 20.5419 | 0.002439 | 0.002306 | 243.7228 | 0.26238 | 8.4005 | 0.002546 | 0.003738 | 0.003696 | 3.7e-05 | 0.000974 | 0.2385 | 1.3743 | 2.0636 | 98.52 | 20.05 | 49.388 | 79.8059 | 82.3658 | 76.5987 | 0.011769 | 1.0094 | 0.023325 | 23.1075 | 0.10155 | 0.15609 | 0.009112 | 0.62729 | 0.62729 | 0.010028
2019-04-09 14:45 | 1554835523.5325 | 642 | 0.20768 | 0.001087 | 0.000156 | 98.3743 | 97.5307 | 116.8892 | 20.5418 | 0.002454 | 0.002263 | 243.7012 | 0.26201 | 8.3997 | 0.002589 | 0.003752 | 0.003696 | 4e-05 | 0.000971 | 0.23706 | 1.3567 | 2.5476 | 98.52 | 20.05 | 49.388 | 80.1698 | 83.4125 | 75.5626 | 0.011769 | 1.0124 | 0.023332 | 23.046 | 0.077029 | 0.15587 | 0.01012 | 0.69759 | 0.69759 | 0.009005
2019-04-09 14:45 | 1554835524.5324 | 643 | 0.2075 | 0.001163 | 0.000159 | 98.4877 | 98.2001 | 116.9653 | 20.5423 | 0.002462 | 0.002305 | 243.6131 | 0.26187 | 8.4061 | 0.002671 | 0.004386 | 0.004247 | 3.6e-05 | 0.000974 | 0.24099 | 1.3811 | 2.5476 | 98.52 | 20.05 | 49.387 | 80.3441 | 85.3816 | 76.588 | 0.011769 | 1.0095 | 0.023306 | 23.0874 | 0.02616 | 0.15579 | 0.011067 | 0.76267 | 0.76267 | 0.008232
2019-04-09 14:45 | 1554835525.5324 | 644 | 0.20743 | 0.001258 | 0.000161 | 98.5479 | 98.3643 | 116.942 | 20.5427 | 0.002435 | 0.002318 | 243.581 | 0.26165 | 8.4004 | 0.002531 | 0.004362 | 0.004309 | 3.8e-05 | 0.000973 | 0.25066 | 1.3458 | 2.5475 | 98.52 | 20.05 | 49.385 | 79.7136 | 82.0218 | 76.8908 | 0.011769 | 1.0086 | 0.023293 | 23.0951 | 0.016681 | 0.15565 | 0.011367 | 0.78302 | 0.78302 | 0.008011
2019-04-09 14:45 | 1554835526.5324 | 645 | 0.20731 | 0.001363 | 0.000162 | 98.5613 | 98.4689 | 116.8693 | 20.5415 | 0.002457 | 0.002321 | 243.4799 | 0.26204 | 8.4017 | 0.002603 | 0.004233 | 0.004252 | 4e-05 | 0.000963 | 0.2385 | 1.3402 | 2.5477 | 98.52 | 20.05 | 49.383 | 80.2286 | 83.7336 | 76.9692 | 0.011768 | 1.0084 | 0.023283 | 23.0905 | 0.008389 | 0.15589 | 0.011991 | 0.82603 | 0.82603 | 0.007605
2019-04-09 14:45 | 1554835527.531 | 646 | 0.20719 | 0.001486 | 0.000162 | 98.5961 | 98.7163 | 116.5038 | 20.5418 | 0.002507 | 0.002338 | 243.4409 | 0.2622 | 8.4016 | 0.002623 | 0.004508 | 0.004467 | 3.9e-05 | 0.000962 | 0.24563 | 1.3104 | 2.5475 | 98.52 | 20.05 | 49.382 | 81.4408 | 84.2116 | 77.3679 | 0.011768 | 1.0072 | 0.023234 | 23.0675 | -0.010899 | 0.15598 | 0.012538 | 0.86222 | 0.86222 | 0.007291
2019-04-09 14:45 | 1554835528.5315 | 647 | 0.20697 | 0.001603 | 0.000163 | 98.6607 | 98.8323 | 116.5346 | 20.5422 | 0.002535 | 0.002339 | 243.3777 | 0.26215 | 8.4009 | 0.002638 | 0.004319 | 0.004204 | 4e-05 | 0.000959 | 0.25426 | 1.3007 | 2.5476 | 98.52 | 20.05 | 49.38 | 82.0964 | 84.5808 | 77.3915 | 0.011767 | 1.0071 | 0.023236 | 23.0713 | -0.015541 | 0.15595 | 0.013695 | 0.94258 | 0.94258 | 0.006668
2019-04-09 14:45 | 1554835529.5326 | 648 | 0.20671 | 0.001719 | 0.000164 | 98.7102 | 98.8351 | 115.9651 | 20.5427 | 0.002499 | 0.002333 | 243.3071 | 0.26184 | 8.4016 | 0.002631 | 0.004371 | 0.004379 | 4.2e-05 | 0.000959 | 0.2505 | 1.3006 | 2.5476 | 98.52 | 20.05 | 49.373 | 81.2421 | 84.4043 | 77.2579 | 0.011766 | 1.0075 | 0.023184 | 23.0105 | -0.01131 | 0.15577 | 0.015146 | 1.0409 | 1.0409 | 0.006031
2019-04-09 14:45 | 1554835530.5312 | 649 | 0.2066 | 0.001837 | 0.000164 | 98.8227 | 99.0687 | 115.5781 | 20.5422 | 0.002502 | 0.002355 | 243.2502 | 0.26181 | 8.4019 | 0.00256 | 0.003744 | 0.003627 | 4.3e-05 | 0.000964 | 0.26106 | 1.3086 | 2.5475 | 98.52 | 20.05 | 49.369 | 81.3211 | 82.6968 | 77.7742 | 0.011765 | 1.006 | 0.023128 | 22.9889 | -0.022239 | 0.15575 | 0.015659 | 1.0739 | 1.0739 | 0.005845
2019-04-09 14:46 | 1554835531.5316 | 650 | 0.20637 | 0.001964 | 0.000165 | 98.9043 | 99.0756 | 114.9556 | 20.5429 | 0.002517 | 0.002351 | 243.1627 | 0.26214 | 8.4015 | 0.002621 | 0.004465 | 0.004441 | 4.4e-05 | 0.000965 | 0.25769 | 1.3181 | 2.5475 | 98.52 | 20.05 | 49.366 | 81.679 | 84.1649 | 77.6877 | 0.011764 | 1.0063 | 0.023068 | 22.9241 | -0.015478 | 0.15594 | 0.016894 | 1.1561 | 1.1561 | 0.005436
2019-04-09 14:46 | 1554835532.5326 | 651 | 0.20622 | 0.002092 | 0.000165 | 99.0178 | 99.2893 | 114.7108 | 20.5422 | 0.002602 | 0.002396 | 243.1112 | 0.26143 | 8.4009 | 0.00261 | 0.004027 | 0.003975 | 4.1e-05 | 0.000962 | 0.28335 | 1.3506 | 2.5476 | 98.52 | 20.04 | 49.357 | 83.718 | 83.8974 | 78.7721 | 0.011755 | 1.0032 | 0.023008 | 22.9351 | -0.024497 | 0.15553 | 0.017644 | 1.2046 | 1.2046 | 0.005203
2019-04-09 14:46 | 1554835533.5322 | 652 | 0.20606 | 0.00222 | 0.000165 | 99.0642 | 99.1937 | 114.1385 | 20.5428 | 0.002579 | 0.002411 | 242.982 | 0.2626 | 8.4011 | 0.002615 | 0.003654 | 0.003615 | 4.5e-05 | 0.000958 | 0.25393 | 1.3639 | 2.5476 | 98.52 | 20.04 | 49.352 | 83.1538 | 84.0391 | 79.1315 | 0.011753 | 1.0022 | 0.022939 | 22.8895 | -0.011678 | 0.15622 | 0.018423 | 1.2544 | 1.2544 | 0.005019
2019-04-09 14:46 | 1554835534.5317 | 653 | 0.20589 | 0.002361 | 0.000165 | 99.1705 | 99.1857 | 113.2965 | 20.5422 | 0.002593 | 0.00248 | 242.9395 | 0.26231 | 8.3998 | 0.00256 | 0.003561 | 0.003464 | 4.3e-05 | 0.000957 | 0.26609 | 1.335 | 2.5476 | 98.52 | 20.05 | 49.346 | 83.49 | 82.7007 | 80.7792 | 0.011759 | 0.9975 | 0.022801 | 22.8581 | -0.001371 | 0.15604 | 0.019282 | 1.3052 | 1.3052 | 0.004818
2019-04-09 14:46 | 1554835535.5313 | 654 | 0.2059 | 0.002492 | 0.000164 | 99.1703 | 99.1565 | 112.8522 | 20.5428 | 0.002616 | 0.002454 | 242.8349 | 0.26243 | 8.4043 | 0.002639 | 0.004197 | 0.004164 | 4.5e-05 | 0.000951 | 0.25409 | 1.3417 | 2.5476 | 98.52 | 20.05 | 49.344 | 84.0516 | 84.6103 | 80.1614 | 0.011759 | 0.99925 | 0.022776 | 22.7934 | 0.001241 | 0.15612 | 0.019048 | 1.2879 | 1.2879 | 0.004885
2019-04-09 14:46 | 1554835536.5304 | 655 | 0.20549 | 0.002636 | 0.000164 | 99.3507 | 99.0794 | 112.4285 | 20.5425 | 0.002503 | 0.002407 | 242.7588 | 0.26239 | 8.403 | 0.002615 | 0.004253 | 0.004233 | 4.2e-05 | 0.000946 | 0.25818 | 1.3545 | 2.5477 | 98.52 | 20.05 | 49.344 | 81.3368 | 84.021 | 79.0305 | 0.011759 | 1.0025 | 0.02277 | 22.7141 | 0.024459 | 0.1561 | 0.021352 | 1.4441 | 1.4441 | 0.004356
2019-04-09 14:46 | 1554835537.5304 | 656 | 0.20517 | 0.00279 | 0.000165 | 99.3713 | 99.0312 | 111.8042 | 20.5421 | 0.002648 | 0.002516 | 242.7032 | 0.26173 | 8.4011 | 0.002599 | 0.004247 | 0.00419 | 4.2e-05 | 0.000945 | 0.2724 | 1.3092 | 2.5475 | 98.52 | 20.05 | 49.344 | 84.814 | 83.6446 | 81.6477 | 0.011759 | 0.99506 | 0.022623 | 22.7349 | 0.03066 | 0.15571 | 0.023028 | 1.5478 | 1.5478 | 0.004054
2019-04-09 14:46 | 1554835538.5305 | 657 | 0.20494 | 0.002944 | 0.000166 | 99.4645 | 99.0619 | 111.2566 | 20.5435 | 0.002682 | 0.002485 | 242.5619 | 0.26219 | 8.4019 | 0.002665 | 0.003837 | 0.003786 | 4.3e-05 | 0.000939 | 0.26432 | 1.2879 | 2.5475 | 98.52 | 20.05 | 49.344 | 85.6426 | 85.2305 | 80.9153 | 0.011759 | 0.99712 | 0.022591 | 22.6558 | 0.03628 | 0.15598 | 0.024253 | 1.6281 | 1.6281 | 0.003861
2019-04-09 14:46 | 1554835539.5308 | 658 | 0.20472 | 0.003112 | 0.000168 | 99.4856 | 99.0963 | 110.5991 | 20.5429 | 0.002588 | 0.002519 | 242.4596 | 0.26262 | 8.4002 | 0.002642 | 0.004076 | 0.004028 | 4.5e-05 | 0.000939 | 0.28639 | 1.321 | 2.5475 | 98.52 | 20.05 | 49.346 | 83.3822 | 84.6853 | 81.728 | 0.011759 | 0.99484 | 0.022498 | 22.6146 | 0.035074 | 0.15623 | 0.025361 | 1.6956 | 1.6956 | 0.003713
2019-04-09 14:46 | 1554835540.5315 | 659 | 0.20425 | 0.003296 | 0.00017 | 99.5654 | 99.0364 | 109.4745 | 20.5432 | 0.002693 | 0.002605 | 242.3319 | 0.26216 | 8.4039 | 0.002667 | 0.004557 | 0.004487 | 4.3e-05 | 0.000938 | 0.27273 | 1.2991 | 2.5475 | 98.52 | 20.05 | 49.349 | 85.91 | 85.2834 | 83.7954 | 0.01176 | 0.98907 | 0.022318 | 22.5648 | 0.047653 | 0.15596 | 0.027971 | 1.8555 | 1.8555 | 0.003387
2019-04-09 14:46 | 1554835541.5305 | 660 | 0.20371 | 0.003548 | 0.000173 | 99.6622 | 99.0371 | 108.6776 | 20.5426 | 0.002706 | 0.002621 | 242.23 | 0.26224 | 8.4001 | 0.00265 | 0.0044 | 0.004288 | 4.6e-05 | 0.000934 | 0.2677 | 1.3382 | 2.5476 | 98.52 | 20.05 | 49.351 | 86.221 | 84.8809 | 84.163 | 0.01176 | 0.98806 | 0.022225 | 22.4941 | 0.056281 | 0.15601 | 0.030887 | 2.0406 | 2.0406 | 0.003081
2019-04-09 14:46 | 1554835542.5315 | 661 | 0.20319 | 0.003856 | 0.000176 | 99.6536 | 99.1716 | 107.6303 | 20.5428 | 0.002694 | 0.002596 | 242.0924 | 0.26132 | 8.4021 | 0.002705 | 0.003913 | 0.003882 | 4.4e-05 | 0.000931 | 0.27577 | 1.4257 | 2.5477 | 98.52 | 20.05 | 49.357 | 85.9183 | 86.1906 | 83.5695 | 0.011762 | 0.9897 | 0.022136 | 22.3669 | 0.043361 | 0.15546 | 0.03358 | 2.2097 | 2.2097 | 0.002835
2019-04-09 14:46 | 1554835543.5306 | 662 | 0.20259 | 0.004168 | 0.00018 | 99.7396 | 99.2078 | 106.444 | 20.5429 | 0.002769 | 0.002706 | 241.9406 | 0.26226 | 8.4011 | 0.002706 | 0.004332 | 0.004271 | 4.3e-05 | 0.000929 | 0.26272 | 1.3622 | 2.5476 | 98.52 | 20.05 | 49.358 | 87.7233 | 86.2164 | 86.2244 | 0.011762 | 0.98239 | 0.021933 | 22.3259 | 0.047824 | 0.15602 | 0.036808 | 2.3997 | 2.3997 | 0.00262
2019-04-09 14:46 | 1554835544.5306 | 663 | 0.20187 | 0.004525 | 0.000184 | 99.7715 | 99.3186 | 105.3941 | 20.5431 | 0.002785 | 0.002671 | 241.8072 | 0.26151 | 8.4022 | 0.002728 | 0.00472 | 0.004632 | 4.1e-05 | 0.000931 | 0.27057 | 1.3433 | 2.5476 | 98.52 | 20.05 | 49.36 | 88.1233 | 86.7452 | 85.3809 | 0.011763 | 0.9847 | 0.02185 | 22.1894 | 0.040695 | 0.15557 | 0.040619 | 2.6379 | 2.6379 | 0.002377
2019-04-09 14:46 | 1554835545.5306 | 664 | 0.20094 | 0.004922 | 0.000188 | 99.9662 | 99.3241 | 104.14 | 20.5431 | 0.002857 | 0.002772 | 241.6378 | 0.26217 | 8.4018 | 0.002766 | 0.003759 | 0.003774 | 4.4e-05 | 0.000933 | 0.27433 | 1.3935 | 2.5476 | 98.52 | 20.05 | 49.366 | 89.8615 | 87.6643 | 87.8174 | 0.011764 | 0.97805 | 0.021646 | 22.1318 | 0.057642 | 0.15597 | 0.045696 | 2.9394 | 2.9394 | 0.002138
2019-04-09 14:46 | 1554835546.5306 | 665 | 0.20069 | 0.005344 | 0.000191 | 99.8733 | 99.1837 | 103.2553 | 20.5424 | 0.002926 | 0.00283 | 241.5198 | 0.26228 | 8.4008 | 0.002765 | 0.004008 | 0.00392 | 4.8e-05 | 0.000941 | 0.26985 | 1.3503 | 2.5475 | 98.52 | 20.05 | 49.367 | 91.5195 | 87.6432 | 89.2133 | 0.011764 | 0.97429 | 0.021512 | 22.0802 | 0.061979 | 0.15603 | 0.04666 | 2.9826 | 2.9826 | 0.002108
2019-04-09 14:46 | 1554835547.5316 | 666 | 0.19996 | 0.005901 | 0.000193 | 99.9591 | 99.0412 | 102.5342 | 20.5444 | 0.002963 | 0.002839 | 241.324 | 0.26147 | 8.4006 | 0.002798 | 0.00442 | 0.00429 | 4.5e-05 | 0.000945 | 0.25818 | 1.3615 | 2.5475 | 98.52 | 20.05 | 49.367 | 92.4153 | 88.4363 | 89.4175 | 0.011764 | 0.97374 | 0.021431 | 22.0092 | 0.082515 | 0.15555 | 0.05036 | 3.2065 | 3.2065 | 0.001955
2019-04-09 14:46 | 1554835548.5311 | 667 | 0.19923 | 0.006539 | 0.000194 | 99.9817 | 98.9486 | 101.3728 | 20.5435 | 0.00305 | 0.002918 | 241.1885 | 0.26188 | 8.4019 | 0.002819 | 0.004196 | 0.00415 | 4.3e-05 | 0.000949 | 0.26936 | 1.3547 | 2.5475 | 98.52 | 20.05 | 49.366 | 94.5112 | 88.9463 | 91.3258 | 0.011764 | 0.96864 | 0.021254 | 21.9417 | 0.092905 | 0.15579 | 0.053906 | 3.4032 | 3.4032 | 0.001845
2019-04-09 14:46 | 1554835549.5317 | 668 | 0.19868 | 0.007102 | 0.000195 | 100.0258 | 99.2315 | 100.755 | 20.5432 | 0.00312 | 0.002982 | 241.0365 | 0.26232 | 8.4041 | 0.002798 | 0.004122 | 0.0041 | 4.6e-05 | 0.000952 | 0.26648 | 1.3224 | 2.5476 | 98.52 | 20.05 | 49.367 | 96.1879 | 88.4212 | 92.856 | 0.011764 | 0.96459 | 0.021144 | 21.9206 | 0.071311 | 0.15606 | 0.056507 | 3.5485 | 3.5485 | 0.001772
2019-04-09 14:46 | 1554835550.5307 | 669 | 0.19806 | 0.007573 | 0.000199 | 100.077 | 99.5877 | 100.4358 | 20.5436 | 0.003104 | 0.003055 | 240.8585 | 0.26129 | 8.4003 | 0.002864 | 0.004886 | 0.00487 | 4.4e-05 | 0.000958 | 0.26985 | 1.3364 | 2.5476 | 98.52 | 20.05 | 49.366 | 95.8163 | 90.0152 | 94.6286 | 0.011764 | 0.95994 | 0.02106 | 21.9388 | 0.04384 | 0.15544 | 0.05962 | 3.7282 | 3.7282 | 0.00168
2019-04-09 14:46 | 1554835551.5308 | 670 | 0.19776 | 0.008023 | 0.000203 | 100.1536 | 99.5609 | 100.0671 | 20.5438 | 0.003139 | 0.003032 | 240.7033 | 0.26174 | 8.4019 | 0.002811 | 0.004708 | 0.004572 | 4.6e-05 | 0.000962 | 0.25752 | 1.3058 | 2.5476 | 98.52 | 20.05 | 49.365 | 96.6649 | 88.7483 | 94.081 | 0.011764 | 0.96137 | 0.021037 | 21.8821 | 0.053089 | 0.15571 | 0.060857 | 3.8009 | 3.8009 | 0.001651
2019-04-09 14:46 | 1554835552.5298 | 671 | 0.1973 | 0.008281 | 0.000207 | 100.2107 | 99.5994 | 99.4101 | 20.5435 | 0.003123 | 0.003095 | 240.5095 | 0.26176 | 8.4001 | 0.002765 | 0.004471 | 0.004443 | 4.3e-05 | 0.000963 | 0.2656 | 1.3271 | 2.5475 | 98.52 | 20.05 | 49.363 | 96.2674 | 87.6341 | 95.5901 | 0.011763 | 0.95744 | 0.020925 | 21.855 | 0.054731 | 0.15572 | 0.063318 | 3.9327 | 3.9327 | 0.001596
2019-04-09 14:46 | 1554835553.5299 | 672 | 0.19704 | 0.008544 | 0.000211 | 100.2119 | 99.8211 | 99.1871 | 20.5439 | 0.003308 | 0.003167 | 240.3895 | 0.26192 | 8.4015 | 0.002875 | 0.004378 | 0.004329 | 4.5e-05 | 0.000965 | 0.27079 | 1.3564 | 2.5475 | 98.52 | 20.05 | 49.363 | 100.7472 | 90.2821 | 97.3254 | 0.011763 | 0.95295 | 0.020852 | 21.8817 | 0.034956 | 0.15582 | 0.064507 | 3.9922 | 3.9922 | 0.001573
2019-04-09 14:46 | 1554835554.5303 | 673 | 0.19653 | 0.008813 | 0.000213 | 100.2215 | 99.5623 | 99.1096 | 20.5427 | 0.003328 | 0.003235 | 240.1839 | 0.26165 | 8.4019 | 0.00285 | 0.004729 | 0.004635 | 4.8e-05 | 0.000964 | 0.29679 | 1.3485 | 2.5475 | 98.52 | 20.05 | 49.364 | 101.2168 | 89.6738 | 98.9697 | 0.011764 | 0.94874 | 0.020798 | 21.9217 | 0.059027 | 0.15565 | 0.06724 | 4.1497 | 4.1497 | 0.001512
2019-04-09 14:46 | 1554835555.5309 | 674 | 0.19626 | 0.009079 | 0.000215 | 100.2703 | 99.4797 | 99.4788 | 20.544 | 0.003346 | 0.003229 | 240.0341 | 0.26087 | 8.4024 | 0.002946 | 0.003727 | 0.003581 | 4.5e-05 | 0.000971 | 0.274 | 1.3403 | 2.5476 | 98.52 | 20.05 | 49.366 | 101.6689 | 91.9948 | 98.8172 | 0.011764 | 0.94913 | 0.020841 | 21.958 | 0.070808 | 0.15519 | 0.068509 | 4.2363 | 4.2363 | 0.001476
2019-04-09 14:46 | 1554835556.5299 | 675 | 0.19601 | 0.009359 | 0.000216 | 100.3858 | 99.5812 | 99.4871 | 20.5438 | 0.003425 | 0.00324 | 239.8651 | 0.26232 | 8.4016 | 0.002844 | 0.004529 | 0.004508 | 4.4e-05 | 0.000969 | 0.29264 | 1.3535 | 2.5475 | 98.52 | 20.05 | 49.368 | 103.5801 | 89.5426 | 99.0923 | 0.011764 | 0.94843 | 0.020834 | 21.967 | 0.071985 | 0.15606 | 0.069643 | 4.3046 | 4.3046 | 0.001461
2019-04-09 14:46 | 1554835557.5303 | 676 | 0.19582 | 0.009631 | 0.000217 | 100.3686 | 99.6648 | 99.1798 | 20.5431 | 0.003507 | 0.003272 | 239.6659 | 0.26164 | 8.4031 | 0.002927 | 0.004756 | 0.00471 | 4.5e-05 | 0.000963 | 0.27793 | 1.3655 | 2.5476 | 98.52 | 20.05 | 49.368 | 105.5475 | 91.5476 | 99.8752 | 0.011764 | 0.94644 | 0.02078 | 21.9561 | 0.062938 | 0.15565 | 0.07045 | 4.3429 | 4.3429 | 0.001444
2019-04-09 14:46 | 1554835558.5314 | 677 | 0.19578 | 0.00989 | 0.000217 | 100.4192 | 99.6886 | 99.4173 | 20.5437 | 0.003422 | 0.003288 | 239.4916 | 0.26072 | 8.4017 | 0.002943 | 0.004499 | 0.004435 | 4.7e-05 | 0.000958 | 0.29319 | 1.3446 | 2.5477 | 98.52 | 20.05 | 49.369 | 103.4967 | 91.9171 | 100.2635 | 0.011765 | 0.94545 | 0.020794 | 21.9938 | 0.065317 | 0.1551 | 0.070421 | 4.3441 | 4.3441 | 0.001439
2019-04-09 14:46 | 1554835559.5301 | 678 | 0.19577 | 0.010137 | 0.000218 | 100.4784 | 99.8579 | 99.6328 | 20.5441 | 0.003535 | 0.003338 | 239.2901 | 0.26188 | 8.4032 | 0.002943 | 0.004385 | 0.004311 | 4.3e-05 | 0.00096 | 0.26808 | 1.3872 | 2.5477 | 98.52 | 20.05 | 49.371 | 106.2474 | 91.9175 | 101.4742 | 0.011765 | 0.9424 | 0.020783 | 22.0533 | 0.055414 | 0.15579 | 0.070169 | 4.3262 | 4.3262 | 0.001451
2019-04-09 14:46 | 1554835560.5301 | 679 | 0.19578 | 0.010367 | 0.000221 | 100.4987 | 99.5949 | 99.8553 | 20.5431 | 0.003602 | 0.003366 | 239.1499 | 0.26286 | 8.4018 | 0.003061 | 0.005177 | 0.005074 | 4.3e-05 | 0.00096 | 0.2792 | 1.3783 | 2.5476 | 98.52 | 20.05 | 49.373 | 107.8512 | 94.777 | 102.1408 | 0.011766 | 0.94073 | 0.020788 | 22.0975 | 0.080799 | 0.15637 | 0.069834 | 4.3065 | 4.3065 | 0.001463
2019-04-09 14:46 | 1554835561.5312 | 680 | 0.1958 | 0.010563 | 0.000224 | 100.546 | 99.6497 | 99.7579 | 20.5433 | 0.003607 | 0.003442 | 238.9512 | 0.26118 | 8.4006 | 0.003052 | 0.004319 | 0.004256 | 4.6e-05 | 0.000963 | 0.28407 | 1.34 | 2.5475 | 98.52 | 20.05 | 49.374 | 107.9893 | 94.5651 | 103.9898 | 0.011766 | 0.93611 | 0.020727 | 22.1411 | 0.080099 | 0.15538 | 0.069484 | 4.2722 | 4.2722 | 0.001466
2019-04-09 14:46 | 1554835562.5898 | 681 | 0.19585 | 0.010709 | 0.000228 | 100.4978 | 99.6115 | 100.3668 | 20.5436 | 0.003714 | 0.003432 | 238.8099 | 0.2618 | 8.4034 | 0.003056 | 0.004676 | 0.004544 | 4.6e-05 | 0.000961 | 0.26394 | 1.3303 | 2.5476 | 98.52 | 20.05 | 49.378 | 110.5776 | 94.6561 | 103.7401 | 0.011767 | 0.93673 | 0.020797 | 22.2012 | 0.079236 | 0.15574 | 0.069013 | 4.2575 | 4.2575 | 0.001474
2019-04-09 14:46 | 1554835563.5419 | 682 | 0.19599 | 0.01083 | 0.000233 | 100.5032 | 99.8056 | 101.1017 | 20.5442 | 0.003724 | 0.003445 | 238.6327 | 0.26187 | 8.4002 | 0.003184 | 0.004643 | 0.004508 | 4.4e-05 | 0.000963 | 0.25713 | 1.3343 | 2.5475 | 98.52 | 20.05 | 49.381 | 110.8284 | 97.7325 | 104.0544 | 0.011768 | 0.93595 | 0.020864 | 22.2916 | 0.062301 | 0.15579 | 0.068028 | 4.2103 | 4.2103 | 0.001491
2019-04-09 14:46 | 1554835564.5416 | 683 | 0.196 | 0.010904 | 0.000238 | 100.519 | 99.5901 | 101.8751 | 20.5438 | 0.003667 | 0.003459 | 238.4675 | 0.26137 | 8.403 | 0.003187 | 0.004878 | 0.004831 | 4.4e-05 | 0.000964 | 0.29192 | 1.3594 | 2.5475 | 98.52 | 20.05 | 49.385 | 109.4414 | 97.8239 | 104.3874 | 0.011769 | 0.93513 | 0.020934 | 22.3866 | 0.083034 | 0.15549 | 0.067862 | 4.214 | 4.214 | 0.001487
2019-04-09 14:46 | 1554835565.5424 | 684 | 0.19626 | 0.010959 | 0.000243 | 100.5061 | 99.4473 | 102.1779 | 20.5436 | 0.003653 | 0.003483 | 238.3402 | 0.26132 | 8.4026 | 0.003101 | 0.004918 | 0.004861 | 4.6e-05 | 0.000965 | 0.26737 | 1.3725 | 2.5476 | 98.52 | 20.05 | 49.386 | 109.1068 | 95.7248 | 104.9825 | 0.011769 | 0.93366 | 0.020949 | 22.4375 | 0.094731 | 0.15546 | 0.06625 | 4.1169 | 4.1169 | 0.001522
2019-04-09 14:46 | 1554835566.542 | 685 | 0.19676 | 0.010981 | 0.000249 | 100.5761 | 99.4466 | 102.4159 | 20.5443 | 0.003738 | 0.003454 | 238.162 | 0.26133 | 8.4008 | 0.003116 | 0.004016 | 0.003954 | 4.2e-05 | 0.000964 | 0.26322 | 1.3834 | 2.5476 | 98.52 | 20.05 | 49.388 | 111.1678 | 96.0873 | 104.2619 | 0.011769 | 0.93544 | 0.020993 | 22.4422 | 0.10102 | 0.15546 | 0.063239 | 3.9385 | 3.9385 | 0.001591
2019-04-09 14:46 | 1554835567.54 | 686 | 0.19709 | 0.010988 | 0.000251 | 100.5991 | 99.246 | 103.2907 | 20.5433 | 0.00376 | 0.00347 | 238.075 | 0.26165 | 8.3994 | 0.003104 | 0.004499 | 0.004516 | 4.5e-05 | 0.000963 | 0.26432 | 1.3754 | 2.5477 | 98.52 | 20.04 | 49.392 | 111.6996 | 95.8028 | 104.6652 | 0.011763 | 0.93444 | 0.021072 | 22.5499 | 0.12112 | 0.15565 | 0.061278 | 3.8309 | 3.8309 | 0.001637
2019-04-09 14:46 | 1554835568.5411 | 687 | 0.19748 | 0.010943 | 0.00025 | 100.5973 | 99.1754 | 104.2167 | 20.5444 | 0.003751 | 0.003519 | 237.9604 | 0.26083 | 8.4011 | 0.00312 | 0.005526 | 0.005388 | 4.3e-05 | 0.000965 | 0.26985 | 1.341 | 2.5476 | 98.52 | 20.04 | 49.395 | 111.4858 | 96.1915 | 105.8446 | 0.011764 | 0.93153 | 0.021133 | 22.6861 | 0.12732 | 0.15517 | 0.05906 | 3.7033 | 3.7033 | 0.001689
2019-04-09 14:46 | 1554835569.5403 | 688 | 0.19803 | 0.010847 | 0.000247 | 100.5472 | 98.8867 | 105.1216 | 20.5447 | 0.003774 | 0.003449 | 237.8575 | 0.26148 | 8.4029 | 0.003146 | 0.004952 | 0.004941 | 4.1e-05 | 0.00097 | 0.274 | 1.3216 | 2.5476 | 98.52 | 20.04 | 49.401 | 112.0451 | 96.8256 | 104.1593 | 0.011765 | 0.93569 | 0.021272 | 22.7337 | 0.14896 | 0.15555 | 0.055857 | 3.5261 | 3.5261 | 0.001778
2019-04-09 14:46 | 1554835570.5402 | 689 | 0.19871 | 0.010689 | 0.000242 | 100.6479 | 99.013 | 105.9213 | 20.5433 | 0.003809 | 0.003479 | 237.7472 | 0.26225 | 8.4019 | 0.003065 | 0.004182 | 0.004135 | 4.4e-05 | 0.000972 | 0.28567 | 1.3918 | 2.5477 | 98.52 | 20.04 | 49.404 | 112.8997 | 94.8608 | 104.8782 | 0.011766 | 0.93391 | 0.021332 | 22.8417 | 0.14649 | 0.15601 | 0.052006 | 3.293 | 3.293 | 0.001909
2019-04-09 14:46 | 1554835571.5412 | 690 | 0.19964 | 0.010472 | 0.000236 | 100.6244 | 98.9806 | 106.8975 | 20.5437 | 0.00376 | 0.003408 | 237.6429 | 0.26078 | 8.4022 | 0.003115 | 0.004899 | 0.00486 | 4.4e-05 | 0.000975 | 0.25409 | 1.3575 | 2.5477 | 98.52 | 20.04 | 49.408 | 111.6937 | 96.0774 | 103.1477 | 0.011767 | 0.93821 | 0.021479 | 22.8941 | 0.14732 | 0.15514 | 0.04674 | 2.9806 | 2.9806 | 0.002098
2019-04-09 14:46 | 1554835572.5403 | 691 | 0.20023 | 0.010048 | 0.000229 | 100.6553 | 99.1334 | 107.8838 | 20.5435 | 0.003717 | 0.003403 | 237.5884 | 0.26192 | 8.3997 | 0.003049 | 0.004778 | 0.004719 | 4.5e-05 | 0.000976 | 0.25769 | 1.3822 | 2.5476 | 98.52 | 20.04 | 49.413 | 110.6619 | 94.4743 | 103.0273 | 0.011768 | 0.93851 | 0.021582 | 22.9958 | 0.13627 | 0.15582 | 0.043703 | 2.8007 | 2.8007 | 0.002242
2019-04-09 14:46 | 1554835573.5412 | 692 | 0.2011 | 0.009447 | 0.000221 | 100.6175 | 99.0643 | 108.791 | 20.5438 | 0.003774 | 0.003379 | 237.4878 | 0.26179 | 8.4019 | 0.003124 | 0.004112 | 0.004108 | 4.3e-05 | 0.000978 | 0.29175 | 1.3671 | 2.5477 | 98.52 | 20.04 | 49.416 | 112.0459 | 96.2897 | 102.4683 | 0.011769 | 0.9399 | 0.021688 | 23.0751 | 0.13916 | 0.15574 | 0.039215 | 2.5259 | 2.5259 | 0.002485
2019-04-09 14:46 | 1554835574.5417 | 693 | 0.20164 | 0.008857 | 0.000212 | 100.6333 | 99.31 | 109.3058 | 20.5441 | 0.003778 | 0.003452 | 237.4381 | 0.26111 | 8.4017 | 0.003059 | 0.004897 | 0.004767 | 3.9e-05 | 0.00097 | 0.27256 | 1.3298 | 2.5476 | 98.52 | 20.04 | 49.42 | 112.1425 | 94.7308 | 104.221 | 0.01177 | 0.93554 | 0.021689 | 23.1836 | 0.1184 | 0.15533 | 0.03674 | 2.3668 | 2.3668 | 0.002645
2019-04-09 14:46 | 1554835575.5413 | 694 | 0.20216 | 0.008324 | 0.000204 | 100.5895 | 99.0818 | 109.898 | 20.5436 | 0.003767 | 0.0034 | 237.3297 | 0.26171 | 8.3999 | 0.003045 | 0.004455 | 0.004398 | 4.2e-05 | 0.000969 | 0.25857 | 1.3402 | 2.5477 | 98.52 | 20.04 | 49.422 | 111.8845 | 94.3788 | 102.9661 | 0.01177 | 0.93866 | 0.021784 | 23.2076 | 0.13508 | 0.15569 | 0.034278 | 2.2181 | 2.2181 | 0.002829
2019-04-09 14:46 | 1554835576.5407 | 695 | 0.20267 | 0.007818 | 0.000196 | 100.6399 | 99.4081 | 110.6449 | 20.5445 | 0.003712 | 0.003402 | 237.2914 | 0.26124 | 8.3994 | 0.003018 | 0.004919 | 0.004803 | 4.3e-05 | 0.000966 | 0.26648 | 1.3578 | 2.5476 | 98.52 | 20.04 | 49.425 | 110.5384 | 93.7382 | 103.0031 | 0.011771 | 0.93857 | 0.021857 | 23.2875 | 0.11015 | 0.15541 | 0.031855 | 2.0684 | 2.0684 | 0.003028
2019-04-09 14:46 | 1554835577.5414 | 696 | 0.20317 | 0.007334 | 0.000187 | 100.6658 | 99.5371 | 111.3429 | 20.544 | 0.003718 | 0.003338 | 237.2418 | 0.26122 | 8.403 | 0.003055 | 0.004506 | 0.004437 | 4.2e-05 | 0.000961 | 0.27632 | 1.3609 | 2.5476 | 98.52 | 20.04 | 49.427 | 110.675 | 94.6314 | 101.4545 | 0.011771 | 0.94245 | 0.021971 | 23.3127 | 0.10085 | 0.1554 | 0.029467 | 1.9235 | 1.9235 | 0.003256
2019-04-09 14:46 | 1554835578.5413 | 697 | 0.2036 | 0.006867 | 0.000179 | 100.7474 | 99.8038 | 111.8107 | 20.5441 | 0.003704 | 0.00332 | 237.1537 | 0.26199 | 8.4004 | 0.003067 | 0.004623 | 0.004643 | 3.9e-05 | 0.000959 | 0.26504 | 1.4223 | 2.5477 | 98.52 | 20.04 | 49.428 | 110.3359 | 94.9172 | 101.0244 | 0.011771 | 0.94353 | 0.02203 | 23.3482 | 0.084161 | 0.15586 | 0.027437 | 1.7959 | 1.7959 | 0.003498
2019-04-09 14:46 | 1554835579.5407 | 698 | 0.20415 | 0.00647 | 0.000173 | 100.748 | 99.9116 | 112.0053 | 20.5439 | 0.003674 | 0.003285 | 237.1252 | 0.262 | 8.4026 | 0.002985 | 0.005091 | 0.004973 | 4.2e-05 | 0.000959 | 0.27007 | 1.3672 | 2.5476 | 98.52 | 20.04 | 49.429 | 109.6133 | 92.9325 | 100.1812 | 0.011772 | 0.94566 | 0.022074 | 23.3422 | 0.074565 | 0.15586 | 0.024624 | 1.6148 | 1.6148 | 0.00389
2019-04-09 14:46 | 1554835580.5408 | 699 | 0.20446 | 0.006115 | 0.000167 | 100.6946 | 100.2674 | 112.2971 | 20.5447 | 0.003666 | 0.003296 | 237.035 | 0.26226 | 8.4032 | 0.002952 | 0.004616 | 0.004557 | 4.3e-05 | 0.000956 | 0.27328 | 1.3595 | 2.5477 | 98.52 | 20.05 | 49.435 | 109.4184 | 92.155 | 100.4489 | 0.01178 | 0.94498 | 0.022095 | 23.3809 | 0.03803 | 0.15602 | 0.023202 | 1.523 | 1.523 | 0.004128
2019-04-09 14:46 | 1554835581.5405 | 700 | 0.20472 | 0.005774 | 0.000161 | 100.6524 | 100.615 | 112.5461 | 20.5445 | 0.00369 | 0.003306 | 237.0155 | 0.26197 | 8.4005 | 0.002895 | 0.003788 | 0.00375 | 4.5e-05 | 0.000946 | 0.26936 | 1.391 | 2.5475 | 98.52 | 20.05 | 49.438 | 110.0122 | 90.771 | 100.6873 | 0.011781 | 0.94438 | 0.022112 | 23.4143 | 0.003319 | 0.15584 | 0.022113 | 1.4527 | 1.4527 | 0.004323
2019-04-09 14:46 | 1554835582.5415 | 701 | 0.20493 | 0.005458 | 0.000156 | 100.6949 | 100.5926 | 112.56 | 20.5448 | 0.003572 | 0.003268 | 236.9289 | 0.26195 | 8.4016 | 0.002927 | 0.004313 | 0.004229 | 4.3e-05 | 0.000941 | 0.27345 | 1.341 | 2.5477 | 98.52 | 20.04 | 49.444 | 107.129 | 91.5483 | 99.7609 | 0.011775 | 0.94673 | 0.022141 | 23.3867 | 0.009092 | 0.15583 | 0.02121 | 1.3953 | 1.3953 | 0.004501
2019-04-09 14:46 | 1554835583.5405 | 702 | 0.20506 | 0.005183 | 0.00015 | 100.6693 | 100.5527 | 112.1452 | 20.5444 | 0.003637 | 0.003313 | 236.8737 | 0.26057 | 8.4013 | 0.002823 | 0.003847 | 0.003794 | 4.2e-05 | 0.000943 | 0.27616 | 1.35 | 2.5476 | 98.52 | 20.05 | 49.447 | 108.7119 | 89.0447 | 100.8527 | 0.011783 | 0.94397 | 0.022068 | 23.3777 | 0.010367 | 0.15502 | 0.0208 | 1.364 | 1.364 | 0.00458
2019-04-09 14:46 | 1554835584.5398 | 703 | 0.20533 | 0.004974 | 0.000145 | 100.6488 | 100.6174 | 111.9574 | 20.544 | 0.003628 | 0.00328 | 236.8402 | 0.26228 | 8.3987 | 0.002905 | 0.004626 | 0.00448 | 4.6e-05 | 0.000947 | 0.27599 | 1.363 | 2.5476 | 98.52 | 20.04 | 49.451 | 108.5004 | 90.9989 | 100.0512 | 0.011777 | 0.94599 | 0.022073 | 23.3331 | 0.002788 | 0.15603 | 0.019414 | 1.2733 | 1.2733 | 0.004938
2019-04-09 14:46 | 1554835585.5406 | 704 | 0.20555 | 0.004761 | 0.00014 | 100.7012 | 100.519 | 112.0342 | 20.545 | 0.003564 | 0.00322 | 236.7808 | 0.26201 | 8.4038 | 0.002952 | 0.004732 | 0.004704 | 4.5e-05 | 0.000943 | 0.27151 | 1.3739 | 2.5477 | 98.52 | 20.04 | 49.457 | 106.9418 | 92.1508 | 98.6176 | 0.011778 | 0.94964 | 0.022123 | 23.2962 | 0.016198 | 0.15587 | 0.0184 | 1.2095 | 1.2095 | 0.005194
2019-04-09 14:46 | 1554835586.5406 | 705 | 0.2057 | 0.004563 | 0.000136 | 100.6875 | 100.5933 | 112.2931 | 20.5451 | 0.003563 | 0.003231 | 236.7355 | 0.26158 | 8.4014 | 0.002924 | 0.004361 | 0.00435 | 4.6e-05 | 0.000947 | 0.28694 | 1.3239 | 2.5476 | 98.52 | 20.04 | 49.46 | 106.9127 | 91.458 | 98.8686 | 0.011779 | 0.949 | 0.022141 | 23.331 | 0.008369 | 0.15561 | 0.017739 | 1.167 | 1.167 | 0.005374
2019-04-09 14:46 | 1554835587.5411 | 706 | 0.20583 | 0.00437 | 0.000133 | 100.6887 | 100.4966 | 112.2729 | 20.5447 | 0.003552 | 0.003185 | 236.6388 | 0.26233 | 8.4008 | 0.002905 | 0.004535 | 0.004416 | 4.3e-05 | 0.000948 | 0.26576 | 1.3032 | 2.5475 | 98.52 | 20.05 | 49.466 | 106.6542 | 91.0153 | 97.7722 | 0.011788 | 0.9518 | 0.022172 | 23.2945 | 0.017078 | 0.15606 | 0.017177 | 1.1316 | 1.1316 | 0.005558
2019-04-09 14:46 | 1554835588.5397 | 707 | 0.20581 | 0.004202 | 0.000131 | 100.6482 | 100.465 | 112.1466 | 20.5446 | 0.0035 | 0.003186 | 236.6109 | 0.26161 | 8.401 | 0.002897 | 0.004685 | 0.004567 | 4.3e-05 | 0.000947 | 0.25321 | 1.3185 | 2.5476 | 98.52 | 20.05 | 49.471 | 105.3792 | 90.8094 | 97.7863 | 0.011789 | 0.95177 | 0.022159 | 23.2818 | 0.016298 | 0.15563 | 0.017494 | 1.152 | 1.152 | 0.005444
2019-04-09 14:46 | 1554835589.5407 | 708 | 0.20587 | 0.00404 | 0.000131 | 100.5574 | 100.4942 | 112.3443 | 20.5456 | 0.003521 | 0.003207 | 236.5376 | 0.26177 | 8.4009 | 0.00286 | 0.004761 | 0.004728 | 4.3e-05 | 0.00095 | 0.25586 | 1.3122 | 2.5475 | 98.52 | 20.04 | 49.475 | 105.8987 | 89.9196 | 98.299 | 0.011783 | 0.95045 | 0.022163 | 23.3184 | 0.005618 | 0.15572 | 0.017365 | 1.1438 | 1.1438 | 0.005487
2019-04-09 14:46 | 1554835590.5407 | 709 | 0.20583 | 0.003904 | 0.00013 | 100.6419 | 100.3751 | 112.5924 | 20.5456 | 0.003539 | 0.003181 | 236.5067 | 0.2613 | 8.4017 | 0.002923 | 0.004088 | 0.004072 | 4.2e-05 | 0.000954 | 0.25857 | 1.3314 | 2.5475 | 98.52 | 20.05 | 49.483 | 106.3244 | 91.4404 | 97.6616 | 0.011792 | 0.95209 | 0.022207 | 23.3241 | 0.023743 | 0.15544 | 0.017743 | 1.1711 | 1.1711 | 0.005349
2019-04-09 14:46 | 1554835591.5398 | 710 | 0.20585 | 0.003787 | 0.000129 | 100.6449 | 100.3696 | 112.7367 | 20.5455 | 0.003521 | 0.003132 | 236.4579 | 0.26145 | 8.4008 | 0.002898 | 0.004259 | 0.004055 | 4.1e-05 | 0.000952 | 0.28064 | 1.3495 | 2.5476 | 98.52 | 20.04 | 49.486 | 105.8923 | 90.8485 | 96.4886 | 0.011785 | 0.95511 | 0.022256 | 23.3021 | 0.024498 | 0.15554 | 0.017811 | 1.1783 | 1.1783 | 0.00532
2019-04-09 14:47 | 1554835592.5398 | 711 | 0.20592 | 0.003698 | 0.000129 | 100.6339 | 100.2337 | 112.7703 | 20.5453 | 0.003527 | 0.003114 | 236.4021 | 0.26119 | 8.3998 | 0.00282 | 0.004172 | 0.004118 | 4.6e-05 | 0.000948 | 0.2677 | 1.3269 | 2.5475 | 98.52 | 20.04 | 49.487 | 106.0337 | 88.9683 | 96.0427 | 0.011785 | 0.95626 | 0.022273 | 23.2916 | 0.035642 | 0.15538 | 0.017509 | 1.1591 | 1.1591 | 0.005402
2019-04-09 14:47 | 1554835593.5398 | 712 | 0.20603 | 0.003618 | 0.000128 | 100.6148 | 100.3648 | 113.1024 | 20.5444 | 0.003473 | 0.003141 | 236.3778 | 0.26211 | 8.4027 | 0.002885 | 0.00448 | 0.004482 | 4.1e-05 | 0.000947 | 0.23435 | 1.3377 | 2.5476 | 98.52 | 20.05 | 49.491 | 104.7283 | 90.5239 | 96.707 | 0.011794 | 0.95455 | 0.022286 | 23.3468 | 0.022254 | 0.15593 | 0.016908 | 1.1198 | 1.1198 | 0.005612
2019-04-09 14:47 | 1554835594.5398 | 713 | 0.20598 | 0.003551 | 0.000127 | 100.5734 | 100.3586 | 113.4646 | 20.5451 | 0.003505 | 0.003118 | 236.2759 | 0.26124 | 8.4028 | 0.002908 | 0.004091 | 0.004015 | 4.5e-05 | 0.000948 | 0.30453 | 1.3235 | 2.5474 | 98.52 | 20.05 | 49.493 | 105.5227 | 91.0727 | 96.1545 | 0.011794 | 0.95597 | 0.022338 | 23.3667 | 0.019126 | 0.15541 | 0.017325 | 1.1503 | 1.1503 | 0.005445
2019-04-09 14:47 | 1554835595.5398 | 714 | 0.20607 | 0.003496 | 0.000124 | 100.5675 | 100.6508 | 113.5439 | 20.5456 | 0.0035 | 0.003134 | 236.2512 | 0.2616 | 8.4024 | 0.002932 | 0.004096 | 0.004033 | 4.3e-05 | 0.000947 | 0.2672 | 1.3242 | 2.5476 | 98.52 | 20.04 | 49.496 | 105.3935 | 91.6633 | 96.5452 | 0.011788 | 0.95496 | 0.022334 | 23.3872 | -0.007407 | 0.15563 | 0.016828 | 1.1171 | 1.1171 | 0.005614
2019-04-09 14:47 | 1554835596.5399 | 715 | 0.20606 | 0.003438 | 0.000123 | 100.5529 | 100.6794 | 113.6594 | 20.5456 | 0.003434 | 0.003087 | 236.1798 | 0.26106 | 8.4011 | 0.002895 | 0.00419 | 0.004055 | 4.4e-05 | 0.000946 | 0.27544 | 1.3499 | 2.5477 | 98.52 | 20.04 | 49.496 | 103.798 | 90.7658 | 95.4049 | 0.011788 | 0.95792 | 0.02238 | 23.363 | -0.011243 | 0.1553 | 0.016984 | 1.1299 | 1.1299 | 0.005539
2019-04-09 14:47 | 1554835597.54 | 716 | 0.20608 | 0.003394 | 0.000121 | 100.519 | 100.6838 | 113.6369 | 20.5455 | 0.003438 | 0.003032 | 236.1387 | 0.26143 | 8.4023 | 0.002876 | 0.00495 | 0.004869 | 4.1e-05 | 0.000951 | 0.25514 | 1.3729 | 2.5477 | 98.52 | 20.04 | 49.498 | 103.8929 | 90.3077 | 94.0795 | 0.011788 | 0.96137 | 0.022418 | 23.3186 | -0.014661 | 0.15552 | 0.016892 | 1.1258 | 1.1258 | 0.005567
2019-04-09 14:47 | 1554835598.5401 | 717 | 0.20627 | 0.003344 | 0.000119 | 100.5586 | 100.626 | 113.6686 | 20.545 | 0.0035 | 0.003071 | 236.0833 | 0.26197 | 8.402 | 0.00286 | 0.003843 | 0.003774 | 4.4e-05 | 0.00095 | 0.24513 | 1.3838 | 2.5477 | 98.52 | 20.04 | 49.5 | 105.3983 | 89.9274 | 95.0199 | 0.011789 | 0.95892 | 0.022392 | 23.3517 | -0.006 | 0.15585 | 0.015857 | 1.0554 | 1.0554 | 0.005951
2019-04-09 14:47 | 1554835599.5391 | 718 | 0.20625 | 0.003301 | 0.000119 | 100.5481 | 100.6949 | 113.3123 | 20.5443 | 0.003408 | 0.00309 | 236.0126 | 0.26121 | 8.3997 | 0.002911 | 0.003875 | 0.003747 | 4.5e-05 | 0.000947 | 0.25769 | 1.341 | 2.5475 | 98.52 | 20.04 | 49.501 | 103.1686 | 91.1476 | 95.4663 | 0.011789 | 0.95776 | 0.022344 | 23.3292 | -0.013043 | 0.15539 | 0.016024 | 1.0643 | 1.0643 | 0.005884
2019-04-09 14:47 | 1554835600.5391 | 719 | 0.20627 | 0.003255 | 0.000117 | 100.5145 | 100.7159 | 112.9663 | 20.5452 | 0.003439 | 0.003051 | 235.9832 | 0.26241 | 8.4001 | 0.002821 | 0.004476 | 0.004477 | 4.5e-05 | 0.000945 | 0.28456 | 1.3401 | 2.5476 | 98.52 | 20.04 | 49.503 | 103.9226 | 88.9793 | 94.5433 | 0.011789 | 0.96016 | 0.022338 | 23.2644 | -0.017902 | 0.15611 | 0.015929 | 1.0578 | 1.0578 | 0.005947
2019-04-09 14:47 | 1554835601.5392 | 720 | 0.20627 | 0.003207 | 0.000115 | 100.5446 | 100.5197 | 112.8405 | 20.546 | 0.003402 | 0.003037 | 235.8934 | 0.2612 | 8.4001 | 0.002818 | 0.004452 | 0.004352 | 4.6e-05 | 0.000947 | 0.27455 | 1.3625 | 2.5477 | 98.52 | 20.04 | 49.504 | 103.0133 | 88.916 | 94.1887 | 0.01179 | 0.96109 | 0.022336 | 23.2402 | 0.002213 | 0.15539 | 0.016019 | 1.0638 | 1.0638 | 0.005887
2019-04-09 14:47 | 1554835602.5404 | 721 | 0.2063 | 0.003165 | 0.000113 | 100.5365 | 100.5851 | 113.0773 | 20.5459 | 0.003438 | 0.003047 | 235.863 | 0.26108 | 8.4054 | 0.002865 | 0.004303 | 0.004239 | 4.6e-05 | 0.000949 | 0.27295 | 1.3594 | 2.5475 | 98.52 | 20.04 | 49.506 | 103.883 | 90.041 | 94.4281 | 0.01179 | 0.96046 | 0.022352 | 23.2722 | -0.004319 | 0.15532 | 0.01586 | 1.054 | 1.054 | 0.005939
2019-04-09 14:47 | 1554835603.5392 | 722 | 0.20633 | 0.003131 | 0.000112 | 100.4636 | 100.6147 | 113.0994 | 20.5458 | 0.003442 | 0.00302 | 235.7928 | 0.26094 | 8.4019 | 0.002796 | 0.004646 | 0.004581 | 4.3e-05 | 0.000952 | 0.28622 | 1.3555 | 2.5477 | 98.52 | 20.04 | 49.509 | 103.9946 | 88.3792 | 93.784 | 0.011791 | 0.96215 | 0.022374 | 23.254 | -0.013444 | 0.15524 | 0.015717 | 1.0456 | 1.0456 | 0.005983
2019-04-09 14:47 | 1554835604.5405 | 723 | 0.2063 | 0.003111 | 0.000111 | 100.4713 | 100.5531 | 112.9421 | 20.5456 | 0.003384 | 0.003036 | 235.7458 | 0.26147 | 8.4024 | 0.00286 | 0.003942 | 0.003854 | 4.4e-05 | 0.00095 | 0.25752 | 1.3557 | 2.5477 | 98.52 | 20.04 | 49.51 | 102.572 | 89.928 | 94.1768 | 0.011791 | 0.96112 | 0.022346 | 23.2503 | -0.007276 | 0.15555 | 0.015963 | 1.0608 | 1.0608 | 0.00591
2019-04-09 14:47 | 1554835605.5393 | 724 | 0.20596 | 0.003091 | 0.00011 | 100.4936 | 100.5482 | 112.8846 | 20.5463 | 0.003356 | 0.002999 | 235.7151 | 0.26119 | 8.4047 | 0.002918 | 0.003995 | 0.003924 | 4.6e-05 | 0.000946 | 0.27527 | 1.3576 | 2.5476 | 98.52 | 20.04 | 49.511 | 101.8916 | 91.317 | 93.2829 | 0.011791 | 0.96347 | 0.022368 | 23.2161 | -0.00486 | 0.15538 | 0.018011 | 1.1986 | 1.1986 | 0.005224
2019-04-09 14:47 | 1554835606.5402 | 725 | 0.20632 | 0.003061 | 0.00011 | 100.4656 | 100.5968 | 112.4372 | 20.5453 | 0.003395 | 0.003028 | 235.6482 | 0.26137 | 8.4005 | 0.002754 | 0.004542 | 0.004474 | 4.8e-05 | 0.000949 | 0.27544 | 1.3522 | 2.5474 | 98.52 | 20.04 | 49.517 | 102.8569 | 87.368 | 93.976 | 0.011793 | 0.96165 | 0.022302 | 23.1919 | -0.011667 | 0.15549 | 0.015879 | 1.0531 | 1.0531 | 0.00595
2019-04-09 14:47 | 1554835607.5393 | 726 | 0.20623 | 0.003046 | 0.00011 | 100.4555 | 100.3985 | 112.2403 | 20.5455 | 0.003363 | 0.002988 | 235.6213 | 0.26116 | 8.402 | 0.00278 | 0.004168 | 0.004087 | 4.6e-05 | 0.00095 | 0.27743 | 1.3258 | 2.5476 | 98.52 | 20.04 | 49.52 | 102.0654 | 87.9961 | 93.02 | 0.011793 | 0.96416 | 0.022312 | 23.1414 | 0.005075 | 0.15536 | 0.016431 | 1.0904 | 1.0904 | 0.005742
2019-04-09 14:47 | 1554835608.5393 | 727 | 0.20623 | 0.003033 | 0.00011 | 100.4665 | 100.4935 | 112.3339 | 20.5455 | 0.003336 | 0.00298 | 235.5469 | 0.26185 | 8.4017 | 0.002835 | 0.00459 | 0.004524 | 5.1e-05 | 0.000949 | 0.26145 | 1.3404 | 2.5477 | 98.52 | 20.05 | 49.529 | 101.4066 | 89.327 | 92.8151 | 0.011803 | 0.9647 | 0.022328 | 23.1446 | -0.002403 | 0.15577 | 0.016497 | 1.0955 | 1.0955 | 0.00573
2019-04-09 14:47 | 1554835609.5394 | 728 | 0.20629 | 0.003024 | 0.000109 | 100.4832 | 100.7485 | 112.8905 | 20.5454 | 0.003396 | 0.002976 | 235.5184 | 0.26132 | 8.4002 | 0.00274 | 0.004458 | 0.004391 | 4.8e-05 | 0.000945 | 0.25194 | 1.3486 | 2.5475 | 98.52 | 20.05 | 49.534 | 102.864 | 87.0473 | 92.7282 | 0.011804 | 0.96493 | 0.022385 | 23.1991 | -0.023588 | 0.15546 | 0.016143 | 1.0747 | 1.0747 | 0.005829
2019-04-09 14:47 | 1554835610.5394 | 729 | 0.20635 | 0.003026 | 0.000109 | 100.4829 | 100.6367 | 112.9949 | 20.5454 | 0.003404 | 0.002978 | 235.4427 | 0.26201 | 8.4008 | 0.002691 | 0.003664 | 0.003635 | 4.7e-05 | 0.000944 | 0.26825 | 1.3132 | 2.5475 | 98.52 | 20.05 | 49.545 | 103.0622 | 85.8568 | 92.7666 | 0.011807 | 0.96483 | 0.022395 | 23.2111 | -0.013678 | 0.15587 | 0.015745 | 1.0486 | 1.0486 | 0.005991
2019-04-09 14:47 | 1554835611.5398 | 730 | 0.20611 | 0.003028 | 0.000109 | 100.4651 | 100.6126 | 113.2036 | 20.5451 | 0.003337 | 0.003016 | 235.3951 | 0.26096 | 8.4034 | 0.002746 | 0.003743 | 0.003703 | 4.8e-05 | 0.000948 | 0.25033 | 1.3185 | 2.5475 | 98.52 | 20.05 | 49.55 | 101.4304 | 87.1712 | 93.6801 | 0.011808 | 0.96242 | 0.022387 | 23.2615 | -0.013128 | 0.15524 | 0.017229 | 1.1475 | 1.1475 | 0.005452
2019-04-09 14:47 | 1554835612.5399 | 731 | 0.20638 | 0.003023 | 0.000108 | 100.4823 | 100.5734 | 113.0541 | 20.5469 | 0.00336 | 0.002999 | 235.3441 | 0.26167 | 8.4011 | 0.00272 | 0.004269 | 0.004179 | 4.8e-05 | 0.000948 | 0.26842 | 1.3415 | 2.5476 | 98.52 | 20.05 | 49.556 | 101.9944 | 86.5507 | 93.2804 | 0.011809 | 0.96347 | 0.022385 | 23.2334 | -0.008102 | 0.15567 | 0.015615 | 1.0395 | 1.0395 | 0.006035
2019-04-09 14:47 | 1554835613.5405 | 732 | 0.2061 | 0.003027 | 0.000107 | 100.4933 | 100.6484 | 113.0752 | 20.5463 | 0.003356 | 0.002913 | 235.2706 | 0.26108 | 8.4035 | 0.002742 | 0.004637 | 0.004617 | 4.8e-05 | 0.000951 | 0.27273 | 1.348 | 2.5478 | 98.52 | 20.05 | 49.567 | 101.9091 | 87.0966 | 91.1973 | 0.011812 | 0.96898 | 0.022451 | 23.1695 | -0.013791 | 0.15532 | 0.01724 | 1.1515 | 1.1515 | 0.005436
2019-04-09 14:47 | 1554835614.5386 | 733 | 0.20619 | 0.003019 | 0.000107 | 100.4558 | 100.5452 | 113.1355 | 20.5456 | 0.003318 | 0.003001 | 235.2817 | 0.26126 | 8.4012 | 0.002792 | 0.004409 | 0.004331 | 4.6e-05 | 0.000954 | 0.26377 | 1.403 | 2.5477 | 98.52 | 20.05 | 49.571 | 100.9906 | 88.2907 | 93.3175 | 0.011813 | 0.96337 | 0.022392 | 23.243 | -0.007951 | 0.15543 | 0.016747 | 1.1156 | 1.1156 | 0.005615
2019-04-09 14:47 | 1554835615.5395 | 734 | 0.20644 | 0.003022 | 0.000106 | 100.4344 | 100.4158 | 113.2062 | 20.5463 | 0.003349 | 0.002968 | 235.1848 | 0.26121 | 8.4008 | 0.002791 | 0.004345 | 0.004322 | 5e-05 | 0.000952 | 0.28064 | 1.3494 | 2.5477 | 98.52 | 20.05 | 49.576 | 101.7348 | 88.2631 | 92.535 | 0.011814 | 0.96544 | 0.022423 | 23.2254 | 0.001659 | 0.15539 | 0.015242 | 1.0164 | 1.0164 | 0.006161
2019-04-09 14:47 | 1554835616.5389 | 735 | 0.20638 | 0.003016 | 0.000106 | 100.4177 | 100.604 | 113.3416 | 20.5456 | 0.003346 | 0.002993 | 235.1479 | 0.2605 | 8.4049 | 0.002719 | 0.003933 | 0.003918 | 4.9e-05 | 0.000955 | 0.26698 | 1.327 | 2.5477 | 98.52 | 20.05 | 49.584 | 101.6509 | 86.5252 | 93.1439 | 0.011816 | 0.96383 | 0.022417 | 23.2586 | -0.016579 | 0.15497 | 0.015619 | 1.0414 | 1.0414 | 0.005997
2019-04-09 14:47 | 1554835617.54 | 736 | 0.20638 | 0.003017 | 0.000105 | 100.3924 | 100.5713 | 113.5998 | 20.5468 | 0.003283 | 0.002898 | 235.0809 | 0.26143 | 8.4002 | 0.002772 | 0.004169 | 0.004072 | 4.7e-05 | 0.000954 | 0.27649 | 1.3232 | 2.5476 | 98.52 | 20.05 | 49.588 | 100.1395 | 87.7998 | 90.8327 | 0.011817 | 0.96995 | 0.022514 | 23.2115 | -0.015927 | 0.15552 | 0.015624 | 1.0463 | 1.0463 | 0.00599
2019-04-09 14:47 | 1554835618.5387 | 737 | 0.20597 | 0.003008 | 0.000104 | 100.3766 | 100.5434 | 113.266 | 20.5456 | 0.003275 | 0.002941 | 235.0358 | 0.26129 | 8.4041 | 0.002796 | 0.004498 | 0.004349 | 4.7e-05 | 0.000956 | 0.26488 | 1.3467 | 2.5475 | 98.52 | 20.05 | 49.594 | 99.9461 | 88.3875 | 91.8894 | 0.011818 | 0.96714 | 0.022448 | 23.211 | -0.014854 | 0.15544 | 0.018049 | 1.2058 | 1.2058 | 0.005195
2019-04-09 14:47 | 1554835619.5388 | 738 | 0.20595 | 0.003008 | 0.000104 | 100.4752 | 100.4777 | 113.1341 | 20.5456 | 0.00329 | 0.002923 | 234.9978 | 0.26046 | 8.4019 | 0.002783 | 0.004486 | 0.00439 | 5.1e-05 | 0.000953 | 0.26521 | 1.3135 | 2.5475 | 98.52 | 20.05 | 49.596 | 100.296 | 88.0829 | 91.4445 | 0.011819 | 0.96832 | 0.022449 | 23.1834 | -0.000225 | 0.15495 | 0.018211 | 1.2167 | 1.2167 | 0.005132
2019-04-09 14:47 | 1554835620.5397 | 739 | 0.20625 | 0.003015 | 0.000103 | 100.3596 | 100.679 | 113.5536 | 20.5459 | 0.003288 | 0.002938 | 234.93 | 0.26128 | 8.3996 | 0.002711 | 0.004658 | 0.004611 | 4.7e-05 | 0.000952 | 0.262 | 1.3106 | 2.5475 | 98.52 | 20.05 | 49.597 | 100.2547 | 86.331 | 91.8098 | 0.011819 | 0.96735 | 0.022479 | 23.2379 | -0.028421 | 0.15543 | 0.016413 | 1.0977 | 1.0977 | 0.005706
2019-04-09 14:47 | 1554835621.5398 | 740 | 0.20635 | 0.003016 | 0.000103 | 100.3129 | 100.4829 | 113.3456 | 20.5468 | 0.003255 | 0.002931 | 234.9047 | 0.26181 | 8.4001 | 0.002696 | 0.004143 | 0.004146 | 4.9e-05 | 0.000951 | 0.2568 | 1.2953 | 2.5475 | 98.52 | 20.05 | 49.6 | 99.4489 | 85.9924 | 91.6488 | 0.01182 | 0.96778 | 0.022464 | 23.2115 | -0.015144 | 0.15575 | 0.015756 | 1.0528 | 1.0528 | 0.005962
2019-04-09 14:47 | 1554835622.5394 | 741 | 0.20627 | 0.00302 | 0.000103 | 100.3379 | 100.551 | 113.1623 | 20.5457 | 0.003257 | 0.002959 | 234.8136 | 0.26166 | 8.404 | 0.002795 | 0.004348 | 0.004215 | 5e-05 | 0.000948 | 0.26842 | 1.3063 | 2.5476 | 98.52 | 20.05 | 49.6 | 99.5108 | 88.3697 | 92.3132 | 0.01182 | 0.96602 | 0.022425 | 23.2139 | -0.018977 | 0.15566 | 0.016267 | 1.0853 | 1.0853 | 0.00578
2019-04-09 14:47 | 1554835623.5392 | 742 | 0.20635 | 0.003022 | 0.000103 | 100.3668 | 100.3979 | 113.3017 | 20.5473 | 0.00326 | 0.002959 | 234.7855 | 0.26085 | 8.4024 | 0.00274 | 0.003983 | 0.003944 | 5.1e-05 | 0.000948 | 0.26394 | 1.3076 | 2.5475 | 98.52 | 20.05 | 49.604 | 99.5749 | 87.046 | 92.3084 | 0.011821 | 0.96603 | 0.022439 | 23.228 | -0.002769 | 0.15518 | 0.015785 | 1.0537 | 1.0537 | 0.005935
2019-04-09 14:47 | 1554835624.5392 | 743 | 0.20647 | 0.003026 | 0.000103 | 100.3436 | 100.5331 | 113.3433 | 20.5462 | 0.003211 | 0.002988 | 234.7032 | 0.26126 | 8.4002 | 0.002744 | 0.004224 | 0.004156 | 5e-05 | 0.000944 | 0.2636 | 1.3461 | 2.5475 | 98.52 | 20.05 | 49.604 | 98.4023 | 87.13 | 93.0169 | 0.011821 | 0.96417 | 0.022421 | 23.2548 | -0.016881 | 0.15543 | 0.015048 | 1.0035 | 1.0035 | 0.006242
2019-04-09 14:47 | 1554835625.5399 | 744 | 0.2063 | 0.003029 | 0.000102 | 100.3659 | 100.731 | 113.412 | 20.5455 | 0.003329 | 0.002954 | 234.6509 | 0.26069 | 8.4026 | 0.002704 | 0.004416 | 0.004399 | 4.7e-05 | 0.000944 | 0.27815 | 1.3969 | 2.5476 | 98.52 | 20.05 | 49.606 | 101.2433 | 86.1723 | 92.199 | 0.011821 | 0.96632 | 0.022453 | 23.2358 | -0.032479 | 0.15509 | 0.016096 | 1.0752 | 1.0752 | 0.005813
2019-04-09 14:47 | 1554835626.5389 | 745 | 0.20586 | 0.003035 | 0.000101 | 100.3805 | 100.6016 | 113.5665 | 20.5468 | 0.003263 | 0.00292 | 234.6178 | 0.26106 | 8.4019 | 0.002748 | 0.00377 | 0.003741 | 4.7e-05 | 0.000946 | 0.26288 | 1.3387 | 2.5476 | 98.52 | 20.05 | 49.61 | 99.6429 | 87.2369 | 91.3806 | 0.011822 | 0.96849 | 0.022494 | 23.2256 | -0.019681 | 0.1553 | 0.018677 | 1.2505 | 1.2505 | 0.005005
2019-04-09 14:47 | 1554835627.539 | 746 | 0.20608 | 0.003036 | 0.000101 | 100.4025 | 100.6608 | 113.4806 | 20.5466 | 0.003211 | 0.002916 | 234.5578 | 0.26192 | 8.4009 | 0.002867 | 0.004779 | 0.004636 | 5.3e-05 | 0.000941 | 0.25354 | 1.3431 | 2.5475 | 98.52 | 20.05 | 49.612 | 98.3901 | 90.0892 | 91.2836 | 0.011823 | 0.96875 | 0.022488 | 23.2137 | -0.022979 | 0.15582 | 0.01736 | 1.1618 | 1.1618 | 0.005405
2019-04-09 14:47 | 1554835628.539 | 747 | 0.20599 | 0.003038 | 9.9e-05 | 100.2911 | 100.6033 | 113.5277 | 20.5464 | 0.0032 | 0.002893 | 234.5319 | 0.26106 | 8.4028 | 0.002784 | 0.00447 | 0.004484 | 4.7e-05 | 0.000942 | 0.26665 | 1.353 | 2.5477 | 98.52 | 20.05 | 49.615 | 98.1181 | 88.0907 | 90.715 | 0.011823 | 0.97027 | 0.022511 | 23.2004 | -0.027799 | 0.1553 | 0.017901 | 1.1994 | 1.1994 | 0.005218
2019-04-09 14:47 | 1554835629.5381 | 748 | 0.20634 | 0.003037 | 9.9e-05 | 100.3337 | 100.6598 | 113.4266 | 20.5466 | 0.003244 | 0.002915 | 234.4385 | 0.26025 | 8.4001 | 0.002723 | 0.003704 | 0.003597 | 5.1e-05 | 0.000932 | 0.27975 | 1.3583 | 2.5477 | 98.52 | 20.05 | 49.619 | 99.186 | 86.6425 | 91.2428 | 0.011824 | 0.96886 | 0.022484 | 23.2069 | -0.029018 | 0.15482 | 0.01582 | 1.0583 | 1.0583 | 0.005896
2019-04-09 14:47 | 1554835630.538 | 749 | 0.20599 | 0.003041 | 9.8e-05 | 100.2843 | 100.636 | 113.2113 | 20.5462 | 0.003176 | 0.002843 | 234.4118 | 0.26106 | 8.3995 | 0.002761 | 0.004709 | 0.004584 | 5.1e-05 | 0.000934 | 0.2792 | 1.3266 | 2.5476 | 98.52 | 20.05 | 49.619 | 97.5557 | 87.5345 | 89.5254 | 0.011824 | 0.97345 | 0.022516 | 23.1302 | -0.031316 | 0.15531 | 0.017919 | 1.201 | 1.201 | 0.005211
2019-04-09 14:47 | 1554835631.5384 | 750 | 0.20636 | 0.003038 | 9.8e-05 | 100.2736 | 100.6711 | 113.1198 | 20.5468 | 0.003213 | 0.002908 | 234.3467 | 0.26112 | 8.4026 | 0.002747 | 0.004142 | 0.004136 | 5e-05 | 0.000931 | 0.26161 | 1.3302 | 2.5476 | 98.52 | 20.05 | 49.622 | 98.442 | 87.2008 | 91.0915 | 0.011825 | 0.96926 | 0.022458 | 23.1707 | -0.035392 | 0.15534 | 0.015672 | 1.0473 | 1.0473 | 0.005978
2019-04-09 14:47 | 1554835632.5392 | 751 | 0.20632 | 0.00304 | 9.7e-05 | 100.256 | 100.5304 | 113.3076 | 20.5463 | 0.003243 | 0.002887 | 234.2997 | 0.26188 | 8.4015 | 0.002716 | 0.004522 | 0.004396 | 4.7e-05 | 0.000938 | 0.28031 | 1.3141 | 2.5474 | 98.52 | 20.05 | 49.623 | 99.1671 | 86.4573 | 90.568 | 0.011825 | 0.97066 | 0.022493 | 23.1732 | -0.024444 | 0.15579 | 0.015959 | 1.0682 | 1.0682 | 0.005878
2019-04-09 14:47 | 1554835633.5392 | 752 | 0.20608 | 0.003037 | 9.8e-05 | 100.3266 | 100.6536 | 113.1527 | 20.5457 | 0.003179 | 0.002867 | 234.2626 | 0.26113 | 8.4005 | 0.002718 | 0.004168 | 0.004098 | 4.9e-05 | 0.000943 | 0.26432 | 1.3497 | 2.5476 | 98.52 | 20.05 | 49.624 | 97.6211 | 86.5042 | 90.0913 | 0.011825 | 0.97193 | 0.022493 | 23.1422 | -0.029105 | 0.15535 | 0.0174 | 1.1649 | 1.1649 | 0.005374
2019-04-09 14:47 | 1554835634.5395 | 753 | 0.20629 | 0.003036 | 9.7e-05 | 100.1991 | 100.3221 | 113.1041 | 20.5464 | 0.003199 | 0.00286 | 234.1772 | 0.26067 | 8.4036 | 0.002717 | 0.004461 | 0.004391 | 5e-05 | 0.000947 | 0.28152 | 1.3111 | 2.5477 | 98.52 | 20.05 | 49.626 | 98.1171 | 86.4768 | 89.9297 | 0.011826 | 0.97236 | 0.022493 | 23.1321 | -0.010975 | 0.15507 | 0.016128 | 1.0796 | 1.0796 | 0.005789
2019-04-09 14:47 | 1554835635.5391 | 754 | 0.20628 | 0.003044 | 9.6e-05 | 100.2691 | 100.4705 | 113.1649 | 20.5467 | 0.003213 | 0.002944 | 234.1516 | 0.26124 | 8.4035 | 0.002764 | 0.004158 | 0.004152 | 4.7e-05 | 0.000944 | 0.26842 | 1.3185 | 2.5475 | 98.52 | 20.05 | 49.626 | 98.4327 | 87.6209 | 91.956 | 0.011826 | 0.96697 | 0.022436 | 23.2028 | -0.017945 | 0.15541 | 0.016197 | 1.0816 | 1.0816 | 0.005791
2019-04-09 14:47 | 1554835636.5392 | 755 | 0.20604 | 0.003047 | 9.5e-05 | 100.2218 | 100.4574 | 113.0942 | 20.5466 | 0.003203 | 0.002854 | 234.0622 | 0.26089 | 8.4028 | 0.002677 | 0.004746 | 0.00463 | 5.1e-05 | 0.000946 | 0.26952 | 1.3318 | 2.5475 | 98.52 | 20.05 | 49.626 | 98.213 | 85.513 | 89.7883 | 0.011826 | 0.97274 | 0.022496 | 23.1266 | -0.021006 | 0.1552 | 0.017592 | 1.1782 | 1.1782 | 0.005309
2019-04-09 14:47 | 1554835637.5393 | 756 | 0.20603 | 0.003053 | 9.4e-05 | 100.1953 | 100.4533 | 112.8269 | 20.5466 | 0.003167 | 0.002843 | 234.0224 | 0.26117 | 8.4017 | 0.002802 | 0.003997 | 0.003956 | 5e-05 | 0.000948 | 0.27367 | 1.364 | 2.5476 | 98.52 | 20.05 | 49.626 | 97.3298 | 88.5283 | 89.5212 | 0.011826 | 0.97346 | 0.022478 | 23.0907 | -0.023004 | 0.15537 | 0.017659 | 1.1817 | 1.1817 | 0.005299
2019-04-09 14:47 | 1554835638.5383 | 757 | 0.20636 | 0.003056 | 9.4e-05 | 100.2122 | 100.3896 | 113.0348 | 20.5467 | 0.003209 | 0.002868 | 233.9601 | 0.26136 | 8.4025 | 0.002745 | 0.00358 | 0.003567 | 4.8e-05 | 0.000943 | 0.286 | 1.3243 | 2.5477 | 98.52 | 20.05 | 49.626 | 98.3542 | 87.161 | 90.1244 | 0.011826 | 0.97184 | 0.02248 | 23.1312 | -0.015818 | 0.15548 | 0.015685 | 1.0494 | 1.0494 | 0.005971
2019-04-09 14:47 | 1554835639.5393 | 758 | 0.20661 | 0.003063 | 9.3e-05 | 100.1613 | 100.4078 | 113.1841 | 20.5462 | 0.003231 | 0.002873 | 233.9211 | 0.26048 | 8.4041 | 0.002842 | 0.004031 | 0.003956 | 5e-05 | 0.000944 | 0.26985 | 1.3041 | 2.5474 | 98.52 | 20.05 | 49.623 | 98.8689 | 89.4909 | 90.2491 | 0.011825 | 0.97151 | 0.022491 | 23.1504 | -0.021989 | 0.15496 | 0.014159 | 0.94745 | 0.94745 | 0.006591
2019-04-09 14:47 | 1554835640.5384 | 759 | 0.20597 | 0.003067 | 9.3e-05 | 100.1821 | 100.2795 | 113.0393 | 20.5463 | 0.003287 | 0.002869 | 233.8826 | 0.2609 | 8.4027 | 0.002716 | 0.004062 | 0.003978 | 4.7e-05 | 0.000938 | 0.28241 | 1.2921 | 2.5474 | 98.52 | 20.05 | 49.622 | 100.223 | 86.4657 | 90.1445 | 0.011825 | 0.97179 | 0.02248 | 23.1323 | -0.008685 | 0.15521 | 0.018022 | 1.2063 | 1.2063 | 0.005185
2019-04-09 14:47 | 1554835641.5384 | 760 | 0.20587 | 0.003078 | 9.2e-05 | 100.1747 | 100.4106 | 113.1785 | 20.5466 | 0.003207 | 0.002858 | 233.8164 | 0.26066 | 8.4015 | 0.00276 | 0.004424 | 0.004386 | 4.7e-05 | 0.000938 | 0.29645 | 1.3311 | 2.5475 | 98.52 | 20.05 | 49.621 | 98.2937 | 87.5161 | 89.8907 | 0.011825 | 0.97247 | 0.022501 | 23.1384 | -0.021038 | 0.15507 | 0.01859 | 1.2456 | 1.2456 | 0.005017
2019-04-09 14:47 | 1554835642.5398 | 761 | 0.20627 | 0.003087 | 9.1e-05 | 100.0973 | 100.552 | 113.0881 | 20.5465 | 0.00318 | 0.002884 | 233.8004 | 0.26067 | 8.4012 | 0.002816 | 0.004157 | 0.004053 | 5e-05 | 0.000941 | 0.28694 | 1.3486 | 2.5477 | 98.52 | 20.05 | 49.622 | 97.6415 | 88.8657 | 90.5096 | 0.011825 | 0.97081 | 0.022473 | 23.1489 | -0.040545 | 0.15507 | 0.016176 | 1.0821 | 1.0821 | 0.005775
2019-04-09 14:47 | 1554835643.5385 | 762 | 0.20613 | 0.0031 | 9.1e-05 | 100.063 | 100.3424 | 113.055 | 20.546 | 0.00321 | 0.002861 | 233.7284 | 0.2613 | 8.4008 | 0.002817 | 0.003892 | 0.003837 | 5.1e-05 | 0.000938 | 0.3147 | 1.3685 | 2.5477 | 98.52 | 20.05 | 49.621 | 98.3762 | 88.9035 | 89.9636 | 0.011825 | 0.97227 | 0.022487 | 23.1281 | -0.024942 | 0.15544 | 0.017001 | 1.1382 | 1.1382 | 0.005504
2019-04-09 14:47 | 1554835644.5385 | 763 | 0.2058 | 0.003106 | 9e-05 | 100.1893 | 100.4615 | 113.1663 | 20.5467 | 0.003188 | 0.002892 | 233.6899 | 0.26197 | 8.4018 | 0.002725 | 0.004298 | 0.004253 | 4.9e-05 | 0.00094 | 0.27129 | 1.3558 | 2.5475 | 98.52 | 20.05 | 49.621 | 97.8446 | 86.6824 | 90.709 | 0.011825 | 0.97028 | 0.022475 | 23.1633 | -0.024271 | 0.15584 | 0.019014 | 1.2727 | 1.2727 | 0.004935
2019-04-09 14:47 | 1554835645.5388 | 764 | 0.20602 | 0.003121 | 9e-05 | 100.1527 | 100.4041 | 113.1538 | 20.547 | 0.003162 | 0.002943 | 233.6331 | 0.26161 | 8.4006 | 0.00276 | 0.004425 | 0.004366 | 4.9e-05 | 0.00094 | 0.26017 | 1.3766 | 2.5477 | 98.52 | 20.05 | 49.619 | 97.207 | 87.5307 | 91.9305 | 0.011824 | 0.96703 | 0.022436 | 23.2008 | -0.022423 | 0.15563 | 0.01765 | 1.1792 | 1.1792 | 0.005319
2019-04-09 14:47 | 1554835646.538 | 765 | 0.2061 | 0.003129 | 9e-05 | 100.1676 | 100.4109 | 112.94 | 20.5463 | 0.003126 | 0.002915 | 233.5781 | 0.26139 | 8.4022 | 0.002802 | 0.004328 | 0.004259 | 4.8e-05 | 0.000939 | 0.27041 | 1.3585 | 2.5478 | 98.52 | 20.05 | 49.618 | 96.3406 | 88.5405 | 91.2589 | 0.011824 | 0.96882 | 0.022435 | 23.1576 | -0.021706 | 0.1555 | 0.017194 | 1.1486 | 1.1486 | 0.005456
2019-04-09 14:47 | 1554835647.5376 | 766 | 0.20611 | 0.003127 | 9e-05 | 100.1175 | 100.275 | 113.2465 | 20.546 | 0.003125 | 0.002883 | 233.5484 | 0.26126 | 8.4015 | 0.002773 | 0.003513 | 0.003463 | 4.8e-05 | 0.000931 | 0.27831 | 1.3311 | 2.5475 | 98.52 | 20.05 | 49.617 | 96.3277 | 87.8334 | 90.4847 | 0.011824 | 0.97088 | 0.02249 | 23.1643 | -0.014057 | 0.15542 | 0.017126 | 1.1468 | 1.1468 | 0.005462
2019-04-09 14:47 | 1554835648.5377 | 767 | 0.20596 | 0.003131 | 9e-05 | 100.144 | 100.3053 | 113.3578 | 20.5469 | 0.003114 | 0.00288 | 233.4608 | 0.26089 | 8.4042 | 0.002759 | 0.004983 | 0.004915 | 4.7e-05 | 0.000928 | 0.29679 | 1.3027 | 2.5477 | 98.52 | 20.05 | 49.616 | 96.0609 | 87.5024 | 90.4114 | 0.011824 | 0.97108 | 0.022503 | 23.1734 | -0.014391 | 0.1552 | 0.018005 | 1.2065 | 1.2065 | 0.005184
2019-04-09 14:47 | 1554835649.5377 | 768 | 0.20599 | 0.003133 | 8.9e-05 | 100.0782 | 100.4192 | 113.1308 | 20.5478 | 0.003133 | 0.002885 | 233.4362 | 0.26137 | 8.4007 | 0.002846 | 0.005007 | 0.004956 | 5.3e-05 | 0.000928 | 0.27433 | 1.2954 | 2.5475 | 98.52 | 20.05 | 49.616 | 96.5207 | 89.5871 | 90.5226 | 0.011824 | 0.97078 | 0.022477 | 23.1537 | -0.030423 | 0.15549 | 0.017843 | 1.1943 | 1.1943 | 0.005247
2019-04-09 14:47 | 1554835650.539 | 769 | 0.20602 | 0.003134 | 8.9e-05 | 100.0755 | 100.4257 | 113.3123 | 20.5476 | 0.003157 | 0.002912 | 233.3567 | 0.2618 | 8.3997 | 0.002799 | 0.003791 | 0.003807 | 5e-05 | 0.000929 | 0.28456 | 1.3161 | 2.5476 | 98.52 | 20.05 | 49.616 | 97.0962 | 88.4503 | 91.1939 | 0.011824 | 0.96899 | 0.022474 | 23.1936 | -0.031245 | 0.15575 | 0.017667 | 1.1824 | 1.1824 | 0.005309
2019-04-09 14:48 | 1554835651.5387 | 770 | 0.20602 | 0.003146 | 8.9e-05 | 100.0511 | 100.138 | 113.3313 | 20.5479 | 0.003171 | 0.00277 | 233.3198 | 0.26075 | 8.4023 | 0.002756 | 0.004887 | 0.004804 | 5e-05 | 0.00093 | 0.29518 | 1.3103 | 2.5475 | 98.52 | 20.05 | 49.618 | 97.4195 | 87.4269 | 87.7557 | 0.011824 | 0.97822 | 0.022583 | 23.0859 | -0.007765 | 0.15512 | 0.017608 | 1.1841 | 1.1841 | 0.005279
2019-04-09 14:48 | 1554835652.5378 | 771 | 0.20595 | 0.003152 | 8.8e-05 | 100.0612 | 100.0898 | 113.2538 | 20.5471 | 0.003104 | 0.002886 | 233.2511 | 0.26149 | 8.4019 | 0.002809 | 0.00385 | 0.003764 | 4.8e-05 | 0.000924 | 0.28047 | 1.3807 | 2.5476 | 98.52 | 20.05 | 49.618 | 95.8166 | 88.6927 | 90.5516 | 0.011824 | 0.9707 | 0.022488 | 23.1672 | -0.002556 | 0.15556 | 0.018031 | 1.2076 | 1.2076 | 0.005191
2019-04-09 14:48 | 1554835653.5381 | 772 | 0.20597 | 0.003165 | 8.8e-05 | 99.971 | 100.1504 | 112.8353 | 20.5473 | 0.003205 | 0.002811 | 233.177 | 0.26071 | 8.4027 | 0.002741 | 0.004349 | 0.004268 | 4.8e-05 | 0.000922 | 0.28512 | 1.3485 | 2.5476 | 98.52 | 20.05 | 49.62 | 98.2513 | 87.0687 | 88.7617 | 0.011825 | 0.9755 | 0.022502 | 23.0674 | -0.016036 | 0.1551 | 0.017899 | 1.1995 | 1.1995 | 0.005211
2019-04-09 14:48 | 1554835654.5388 | 773 | 0.20596 | 0.003171 | 8.8e-05 | 100.0526 | 100.2495 | 112.9074 | 20.5476 | 0.003227 | 0.002881 | 233.1391 | 0.26203 | 8.401 | 0.002684 | 0.004547 | 0.004526 | 4.8e-05 | 0.000922 | 0.30359 | 1.3614 | 2.5476 | 98.52 | 20.05 | 49.621 | 98.7816 | 85.7028 | 90.4372 | 0.011825 | 0.97101 | 0.022458 | 23.1281 | -0.017587 | 0.15588 | 0.017953 | 1.2008 | 1.2008 | 0.005232
2019-04-09 14:48 | 1554835655.5378 | 774 | 0.20602 | 0.003182 | 8.7e-05 | 100.0032 | 100.2334 | 112.8923 | 20.5484 | 0.003207 | 0.002901 | 233.0441 | 0.2611 | 8.4004 | 0.002748 | 0.004168 | 0.004074 | 4.5e-05 | 0.000923 | 0.27151 | 1.385 | 2.5477 | 98.52 | 20.05 | 49.624 | 98.2971 | 87.2406 | 90.9202 | 0.011825 | 0.96972 | 0.022441 | 23.1419 | -0.020564 | 0.15532 | 0.017563 | 1.1738 | 1.1738 | 0.005333
2019-04-09 14:48 | 1554835656.5389 | 775 | 0.20597 | 0.003193 | 8.7e-05 | 99.9665 | 100.136 | 112.8067 | 20.5481 | 0.00317 | 0.002884 | 233.0112 | 0.26042 | 8.402 | 0.002696 | 0.004361 | 0.004369 | 4.8e-05 | 0.000921 | 0.25752 | 1.3479 | 2.5475 | 98.52 | 20.05 | 49.629 | 97.4087 | 85.9912 | 90.5205 | 0.011827 | 0.97078 | 0.022445 | 23.1204 | -0.015145 | 0.15492 | 0.017879 | 1.1952 | 1.1952 | 0.005224
2019-04-09 14:48 | 1554835657.5379 | 776 | 0.20576 | 0.003197 | 8.8e-05 | 99.9683 | 100.1022 | 112.6266 | 20.5476 | 0.003172 | 0.002812 | 232.9193 | 0.2614 | 8.4012 | 0.002717 | 0.003794 | 0.003704 | 4.8e-05 | 0.000926 | 0.29391 | 1.3112 | 2.5475 | 98.52 | 20.05 | 49.631 | 97.4481 | 86.4835 | 88.7801 | 0.011827 | 0.97545 | 0.022481 | 23.0466 | -0.011972 | 0.15551 | 0.01913 | 1.2809 | 1.2809 | 0.004893
2019-04-09 14:48 | 1554835658.539 | 777 | 0.20583 | 0.003213 | 8.8e-05 | 99.9597 | 100.1091 | 112.4532 | 20.5477 | 0.003139 | 0.00289 | 232.8752 | 0.26093 | 8.4026 | 0.002735 | 0.00485 | 0.004814 | 4.6e-05 | 0.000923 | 0.27865 | 1.3361 | 2.5476 | 98.52 | 20.05 | 49.636 | 96.6641 | 86.9144 | 90.6623 | 0.011828 | 0.9704 | 0.022405 | 23.0887 | -0.01336 | 0.15523 | 0.01868 | 1.2465 | 1.2465 | 0.005018
2019-04-09 14:48 | 1554835659.539 | 778 | 0.20582 | 0.003221 | 8.8e-05 | 99.9543 | 100.1834 | 112.5013 | 20.5474 | 0.003189 | 0.002911 | 232.8096 | 0.26106 | 8.4018 | 0.002763 | 0.004221 | 0.00422 | 4.3e-05 | 0.000928 | 0.26128 | 1.3365 | 2.5476 | 98.52 | 20.05 | 49.639 | 97.8736 | 87.5958 | 91.1679 | 0.011829 | 0.96906 | 0.022395 | 23.1097 | -0.020479 | 0.1553 | 0.018732 | 1.2494 | 1.2494 | 0.005009
2019-04-09 14:48 | 1554835660.537 | 779 | 0.20579 | 0.003234 | 8.8e-05 | 99.9278 | 100.1222 | 112.2215 | 20.5473 | 0.003173 | 0.002886 | 232.7503 | 0.2613 | 8.3994 | 0.002715 | 0.004142 | 0.004011 | 4.8e-05 | 0.000932 | 0.27273 | 1.319 | 2.5475 | 98.52 | 20.05 | 49.642 | 97.4809 | 86.4308 | 90.5494 | 0.01183 | 0.97071 | 0.022386 | 23.0613 | -0.01738 | 0.15544 | 0.018903 | 1.2603 | 1.2603 | 0.004971
2019-04-09 14:48 | 1554835661.5375 | 780 | 0.20585 | 0.003253 | 8.9e-05 | 99.8966 | 100.189 | 112.2734 | 20.5477 | 0.003155 | 0.002894 | 232.7018 | 0.26113 | 8.4015 | 0.002762 | 0.004806 | 0.004802 | 4.5e-05 | 0.000929 | 0.2855 | 1.3295 | 2.5475 | 98.52 | 20.05 | 49.65 | 97.0384 | 87.5774 | 90.757 | 0.011832 | 0.97015 | 0.022385 | 23.0732 | -0.026142 | 0.15535 | 0.018528 | 1.2352 | 1.2352 | 0.005069
2019-04-09 14:48 | 1554835662.5375 | 781 | 0.20576 | 0.003266 | 8.9e-05 | 99.8912 | 100.1711 | 112.0842 | 20.5482 | 0.003148 | 0.002898 | 232.6222 | 0.261 | 8.4013 | 0.002731 | 0.003877 | 0.003824 | 4.7e-05 | 0.000927 | 0.28119 | 1.3945 | 2.5477 | 98.52 | 20.05 | 49.654 | 96.8768 | 86.8322 | 90.8418 | 0.011833 | 0.96993 | 0.022363 | 23.0565 | -0.025024 | 0.15527 | 0.019044 | 1.2684 | 1.2684 | 0.004933
2019-04-09 14:48 | 1554835663.5381 | 782 | 0.20575 | 0.003283 | 8.9e-05 | 99.8233 | 99.9704 | 112.2584 | 20.5481 | 0.003195 | 0.002877 | 232.5879 | 0.26072 | 8.401 | 0.002752 | 0.004381 | 0.004291 | 4.7e-05 | 0.000928 | 0.27184 | 1.3481 | 2.5475 | 98.52 | 20.05 | 49.658 | 97.9997 | 87.3262 | 90.344 | 0.011834 | 0.97125 | 0.022396 | 23.0586 | -0.013172 | 0.1551 | 0.019099 | 1.274 | 1.274 | 0.004906
2019-04-09 14:48 | 1554835664.5386 | 783 | 0.2057 | 0.003303 | 8.9e-05 | 99.82 | 100.095 | 112.1523 | 20.5488 | 0.003163 | 0.002901 | 232.5147 | 0.26133 | 8.4008 | 0.002768 | 0.004507 | 0.004425 | 4.6e-05 | 0.000927 | 0.29142 | 1.303 | 2.5476 | 98.52 | 20.05 | 49.67 | 97.2496 | 87.7166 | 90.9137 | 0.011836 | 0.96974 | 0.022368 | 23.0657 | -0.024605 | 0.15547 | 0.019359 | 1.2897 | 1.2897 | 0.004858
2019-04-09 14:48 | 1554835665.5382 | 784 | 0.2057 | 0.00332 | 8.9e-05 | 99.8358 | 99.9102 | 111.7882 | 20.5487 | 0.003157 | 0.002907 | 232.4707 | 0.26085 | 8.4009 | 0.00274 | 0.00416 | 0.004146 | 4.4e-05 | 0.000928 | 0.29103 | 1.3128 | 2.5476 | 98.52 | 20.05 | 49.677 | 97.0849 | 87.0324 | 91.0738 | 0.011838 | 0.96931 | 0.022326 | 23.0333 | -0.006661 | 0.15518 | 0.019315 | 1.2844 | 1.2844 | 0.004869
2019-04-09 14:48 | 1554835666.5382 | 785 | 0.20574 | 0.003337 | 8.9e-05 | 99.7581 | 100.1122 | 111.945 | 20.5481 | 0.00316 | 0.002929 | 232.4164 | 0.26118 | 8.4009 | 0.002722 | 0.004892 | 0.004767 | 4.3e-05 | 0.00093 | 0.28031 | 1.3552 | 2.5476 | 98.52 | 20.05 | 49.687 | 97.1633 | 86.6026 | 91.5898 | 0.01184 | 0.96794 | 0.022326 | 23.0658 | -0.031695 | 0.15538 | 0.01907 | 1.2681 | 1.2681 | 0.004938
2019-04-09 14:48 | 1554835667.5382 | 786 | 0.2056 | 0.003353 | 8.9e-05 | 99.7471 | 99.9825 | 111.9652 | 20.5489 | 0.003204 | 0.002859 | 232.349 | 0.26201 | 8.3992 | 0.002775 | 0.004488 | 0.004449 | 4.7e-05 | 0.000929 | 0.26681 | 1.3316 | 2.5476 | 98.52 | 20.05 | 49.693 | 98.2302 | 87.8736 | 89.9127 | 0.011842 | 0.97241 | 0.02238 | 23.0148 | -0.021083 | 0.15587 | 0.019898 | 1.3264 | 1.3264 | 0.004736
2019-04-09 14:48 | 1554835668.5382 | 787 | 0.20568 | 0.003373 | 8.8e-05 | 99.6905 | 99.8895 | 111.637 | 20.5486 | 0.003181 | 0.002955 | 232.3244 | 0.26071 | 8.4009 | 0.002722 | 0.004456 | 0.004317 | 4.6e-05 | 0.000932 | 0.27688 | 1.3175 | 2.5476 | 98.52 | 20.05 | 49.7 | 97.6694 | 86.6115 | 92.2247 | 0.011844 | 0.96626 | 0.022276 | 23.0541 | -0.017841 | 0.1551 | 0.019374 | 1.2854 | 1.2854 | 0.004863
2019-04-09 14:48 | 1554835669.5373 | 788 | 0.20558 | 0.003387 | 8.8e-05 | 99.7009 | 99.882 | 111.4476 | 20.5481 | 0.003237 | 0.002887 | 232.2337 | 0.26161 | 8.4056 | 0.002734 | 0.004754 | 0.004686 | 4.3e-05 | 0.000933 | 0.28263 | 1.3457 | 2.5475 | 98.52 | 20.05 | 49.715 | 99.0291 | 86.8858 | 90.5871 | 0.011847 | 0.97061 | 0.022307 | 22.9828 | -0.016228 | 0.15563 | 0.019981 | 1.3277 | 1.3277 | 0.004724
2019-04-09 14:48 | 1554835670.5383 | 789 | 0.20558 | 0.003403 | 8.8e-05 | 99.6691 | 99.979 | 111.826 | 20.5489 | 0.0031 | 0.002902 | 232.1749 | 0.26164 | 8.403 | 0.002698 | 0.004932 | 0.004852 | 4.5e-05 | 0.000933 | 0.26665 | 1.3722 | 2.5476 | 98.52 | 20.05 | 49.724 | 95.7182 | 86.0281 | 90.9511 | 0.011849 | 0.96964 | 0.022334 | 23.0333 | -0.027774 | 0.15565 | 0.019931 | 1.3259 | 1.3259 | 0.004731
2019-04-09 14:48 | 1554835671.5383 | 790 | 0.20559 | 0.003413 | 8.7e-05 | 99.6673 | 99.8458 | 111.607 | 20.5481 | 0.003153 | 0.002904 | 232.0941 | 0.26156 | 8.402 | 0.002695 | 0.004761 | 0.004733 | 4.8e-05 | 0.000934 | 0.2792 | 1.3474 | 2.5475 | 98.52 | 20.05 | 49.739 | 96.9934 | 85.9694 | 90.999 | 0.011853 | 0.96951 | 0.022311 | 23.0123 | -0.016011 | 0.1556 | 0.019855 | 1.3195 | 1.3195 | 0.004752
2019-04-09 14:48 | 1554835672.5384 | 791 | 0.20555 | 0.003431 | 8.7e-05 | 99.7099 | 99.9484 | 111.4305 | 20.5487 | 0.003207 | 0.002913 | 232.0301 | 0.26119 | 8.4022 | 0.002839 | 0.004034 | 0.003974 | 4.7e-05 | 0.000937 | 0.27649 | 1.3336 | 2.5476 | 98.52 | 20.05 | 49.746 | 98.3051 | 89.4147 | 91.2058 | 0.011855 | 0.96896 | 0.022287 | 23.0006 | -0.021373 | 0.15538 | 0.020074 | 1.3327 | 1.3327 | 0.004699
2019-04-09 14:48 | 1554835673.5378 | 792 | 0.20559 | 0.003451 | 8.6e-05 | 99.5991 | 99.7398 | 111.5798 | 20.5483 | 0.003208 | 0.002904 | 231.9909 | 0.25994 | 8.4014 | 0.002766 | 0.003953 | 0.003883 | 4.6e-05 | 0.00094 | 0.2808 | 1.351 | 2.5475 | 98.52 | 20.05 | 49.753 | 98.3275 | 87.6638 | 90.9795 | 0.011856 | 0.96956 | 0.022308 | 23.0089 | -0.012632 | 0.15464 | 0.019802 | 1.3159 | 1.3159 | 0.004736
2019-04-09 14:48 | 1554835674.5387 | 793 | 0.20553 | 0.003463 | 8.7e-05 | 99.6205 | 99.8276 | 111.4436 | 20.5488 | 0.003135 | 0.002887 | 231.919 | 0.26126 | 8.3999 | 0.002766 | 0.003938 | 0.00389 | 4.7e-05 | 0.000943 | 0.29015 | 1.3035 | 2.5475 | 98.52 | 20.05 | 49.767 | 96.5695 | 87.6799 | 90.5807 | 0.01186 | 0.97062 | 0.022307 | 22.9822 | -0.018574 | 0.15543 | 0.020204 | 1.3425 | 1.3425 | 0.004666
2019-04-09 14:48 | 1554835675.5371 | 794 | 0.20553 | 0.003474 | 8.6e-05 | 99.5982 | 99.7866 | 111.596 | 20.5484 | 0.003198 | 0.002898 | 231.8956 | 0.26121 | 8.402 | 0.002757 | 0.004409 | 0.004297 | 4.6e-05 | 0.000946 | 0.27151 | 1.2951 | 2.5475 | 98.52 | 20.05 | 49.774 | 98.0768 | 87.4594 | 90.8569 | 0.011861 | 0.96989 | 0.022314 | 23.0067 | -0.016909 | 0.1554 | 0.020181 | 1.3415 | 1.3415 | 0.004668
2019-04-09 14:48 | 1554835676.5366 | 795 | 0.20558 | 0.003485 | 8.6e-05 | 99.6423 | 99.7512 | 111.692 | 20.5487 | 0.003194 | 0.002905 | 231.8169 | 0.26137 | 8.3974 | 0.002692 | 0.005111 | 0.005028 | 4.8e-05 | 0.000936 | 0.26432 | 1.3297 | 2.5476 | 98.52 | 20.05 | 49.78 | 97.9799 | 85.8927 | 91.0049 | 0.011863 | 0.96949 | 0.022319 | 23.0212 | -0.009774 | 0.15549 | 0.019838 | 1.319 | 1.319 | 0.004751
2019-04-09 14:48 | 1554835677.538 | 796 | 0.20553 | 0.003494 | 8.6e-05 | 99.5464 | 99.8076 | 111.7882 | 20.549 | 0.003211 | 0.002906 | 231.7806 | 0.26062 | 8.4031 | 0.002695 | 0.00389 | 0.003862 | 4.6e-05 | 0.000934 | 0.27489 | 1.2882 | 2.5475 | 98.52 | 20.06 | 49.794 | 98.3997 | 85.9556 | 91.0396 | 0.011873 | 0.9694 | 0.022327 | 23.0322 | -0.023445 | 0.15504 | 0.020164 | 1.3411 | 1.3411 | 0.004659
2019-04-09 14:48 | 1554835678.538 | 797 | 0.20548 | 0.003505 | 8.6e-05 | 99.5505 | 99.6896 | 111.9774 | 20.5488 | 0.003199 | 0.002944 | 231.717 | 0.2607 | 8.4018 | 0.00279 | 0.003861 | 0.003791 | 4.8e-05 | 0.000931 | 0.29336 | 1.3192 | 2.5477 | 98.52 | 20.05 | 49.8 | 98.1117 | 88.243 | 91.9667 | 0.011867 | 0.96694 | 0.022318 | 23.081 | -0.012485 | 0.15509 | 0.020452 | 1.3597 | 1.3597 | 0.004597
2019-04-09 14:48 | 1554835679.5376 | 798 | 0.20542 | 0.003514 | 8.6e-05 | 99.527 | 99.5337 | 112.0408 | 20.549 | 0.003187 | 0.002916 | 231.657 | 0.26112 | 8.4006 | 0.002778 | 0.004354 | 0.004367 | 4.4e-05 | 0.000931 | 0.30807 | 1.3013 | 2.5475 | 98.52 | 20.06 | 49.81 | 97.829 | 87.9495 | 91.2922 | 0.011877 | 0.96873 | 0.022345 | 23.0662 | -0.0006 | 0.15534 | 0.020788 | 1.3838 | 1.3838 | 0.004524
2019-04-09 14:48 | 1554835680.5377 | 799 | 0.20543 | 0.003521 | 8.6e-05 | 99.5428 | 99.566 | 111.6983 | 20.5493 | 0.003155 | 0.002884 | 231.635 | 0.26145 | 8.3997 | 0.002777 | 0.003954 | 0.003917 | 4.7e-05 | 0.000927 | 0.28639 | 1.3389 | 2.5476 | 98.52 | 20.06 | 49.815 | 97.0549 | 87.9241 | 90.5078 | 0.011878 | 0.97082 | 0.022335 | 23.0062 | -0.002089 | 0.15553 | 0.020692 | 1.3768 | 1.3768 | 0.004553
2019-04-09 14:48 | 1554835681.5367 | 800 | 0.20542 | 0.003534 | 8.6e-05 | 99.5151 | 99.5936 | 111.829 | 20.5495 | 0.003215 | 0.002978 | 231.549 | 0.26106 | 8.4003 | 0.002856 | 0.004381 | 0.004205 | 4.5e-05 | 0.000922 | 0.28224 | 1.3532 | 2.5477 | 98.52 | 20.05 | 49.819 | 98.5036 | 89.8458 | 92.7635 | 0.011872 | 0.96483 | 0.022279 | 23.0909 | -0.007051 | 0.1553 | 0.020739 | 1.3764 | 1.3764 | 0.004547
2019-04-09 14:48 | 1554835682.5371 | 801 | 0.20537 | 0.00355 | 8.6e-05 | 99.5139 | 99.555 | 112.0814 | 20.5492 | 0.003256 | 0.002979 | 231.5245 | 0.26067 | 8.4016 | 0.002863 | 0.004463 | 0.004516 | 4.7e-05 | 0.000925 | 0.31205 | 1.367 | 2.5476 | 98.52 | 20.05 | 49.824 | 99.496 | 90.0112 | 92.8093 | 0.011873 | 0.96471 | 0.022303 | 23.1184 | -0.003695 | 0.15507 | 0.021039 | 1.3979 | 1.3979 | 0.004471
2019-04-09 14:48 | 1554835683.5368 | 802 | 0.20534 | 0.003567 | 8.6e-05 | 99.505 | 99.4951 | 111.7509 | 20.5481 | 0.00321 | 0.002945 | 231.4433 | 0.26103 | 8.4038 | 0.002914 | 0.004229 | 0.004127 | 4.5e-05 | 0.000926 | 0.29966 | 1.3368 | 2.5477 | 98.52 | 20.06 | 49.827 | 98.3838 | 91.2429 | 91.9668 | 0.011881 | 0.96694 | 0.022295 | 23.0577 | 0.000883 | 0.15529 | 0.021184 | 1.407 | 1.407 | 0.004448
2019-04-09 14:48 | 1554835684.537 | 803 | 0.20538 | 0.003584 | 8.6e-05 | 99.4106 | 99.513 | 111.7685 | 20.5494 | 0.003203 | 0.002955 | 231.407 | 0.26133 | 8.4006 | 0.002759 | 0.004287 | 0.004294 | 4.7e-05 | 0.000929 | 0.30486 | 1.3579 | 2.5477 | 98.52 | 20.06 | 49.831 | 98.2111 | 87.501 | 92.2255 | 0.011882 | 0.96625 | 0.022289 | 23.0677 | -0.009215 | 0.15547 | 0.020957 | 1.3915 | 1.3915 | 0.004503
2019-04-09 14:48 | 1554835685.5368 | 804 | 0.20536 | 0.003593 | 8.7e-05 | 99.4365 | 99.4067 | 112.018 | 20.5492 | 0.003305 | 0.002911 | 231.3541 | 0.26045 | 8.4027 | 0.002852 | 0.00461 | 0.004532 | 4.8e-05 | 0.000935 | 0.29159 | 1.3443 | 2.5475 | 98.52 | 20.06 | 49.832 | 100.6744 | 89.7347 | 91.1523 | 0.011882 | 0.9691 | 0.022347 | 23.0595 | 0.00268 | 0.15494 | 0.021048 | 1.4012 | 1.4012 | 0.004456
2019-04-09 14:48 | 1554835686.5369 | 805 | 0.20533 | 0.003605 | 8.7e-05 | 99.4317 | 99.3048 | 111.6875 | 20.5486 | 0.003235 | 0.002941 | 231.274 | 0.26122 | 8.3987 | 0.002938 | 0.004095 | 0.003992 | 4.9e-05 | 0.00094 | 0.29103 | 1.3793 | 2.5476 | 98.52 | 20.06 | 49.833 | 98.9803 | 91.8046 | 91.8728 | 0.011883 | 0.96719 | 0.022292 | 23.0482 | 0.011422 | 0.1554 | 0.021212 | 1.4086 | 1.4086 | 0.004446
2019-04-09 14:48 | 1554835687.5369 | 806 | 0.2052 | 0.003622 | 8.7e-05 | 99.4362 | 99.4476 | 111.3603 | 20.5491 | 0.003233 | 0.002926 | 231.2349 | 0.2611 | 8.4007 | 0.00295 | 0.004113 | 0.003991 | 4.9e-05 | 0.000939 | 0.28495 | 1.3483 | 2.5476 | 98.52 | 20.06 | 49.832 | 98.9305 | 92.0883 | 91.5197 | 0.011882 | 0.96812 | 0.02227 | 23.0033 | -0.001031 | 0.15532 | 0.021972 | 1.4577 | 1.4577 | 0.004294
2019-04-09 14:48 | 1554835688.5379 | 807 | 0.20519 | 0.003634 | 8.8e-05 | 99.3331 | 99.3313 | 111.4112 | 20.5491 | 0.003168 | 0.002969 | 231.1427 | 0.26076 | 8.4026 | 0.002939 | 0.003702 | 0.003657 | 4.2e-05 | 0.00094 | 0.28622 | 1.3583 | 2.5476 | 98.52 | 20.06 | 49.831 | 97.3539 | 91.8277 | 92.5667 | 0.011882 | 0.96535 | 0.022243 | 23.0415 | 0.000164 | 0.15513 | 0.021999 | 1.4577 | 1.4577 | 0.004289
2019-04-09 14:48 | 1554835689.5379 | 808 | 0.20515 | 0.003653 | 8.7e-05 | 99.3763 | 99.3286 | 111.4774 | 20.5499 | 0.003186 | 0.00299 | 231.1174 | 0.26012 | 8.4035 | 0.002897 | 0.004575 | 0.004525 | 4.5e-05 | 0.000938 | 0.27544 | 1.3377 | 2.5476 | 98.52 | 20.05 | 49.83 | 97.7904 | 90.8137 | 93.0556 | 0.011875 | 0.96406 | 0.022235 | 23.0638 | 0.004299 | 0.15475 | 0.022213 | 1.4714 | 1.4714 | 0.004238
2019-04-09 14:48 | 1554835690.537 | 809 | 0.20517 | 0.003679 | 8.7e-05 | 99.312 | 99.3234 | 111.4866 | 20.5513 | 0.00325 | 0.002962 | 231.0248 | 0.26107 | 8.4009 | 0.002903 | 0.004594 | 0.004479 | 4.6e-05 | 0.000938 | 0.28838 | 1.3588 | 2.5475 | 98.52 | 20.05 | 49.83 | 99.3405 | 90.9802 | 92.3915 | 0.011875 | 0.96581 | 0.022256 | 23.0438 | -0.001027 | 0.15531 | 0.022067 | 1.4631 | 1.4631 | 0.004278
2019-04-09 14:48 | 1554835691.537 | 810 | 0.20514 | 0.003706 | 8.6e-05 | 99.2435 | 99.2742 | 111.1462 | 20.55 | 0.00324 | 0.002984 | 230.9795 | 0.26007 | 8.4024 | 0.002807 | 0.004846 | 0.004758 | 4.6e-05 | 0.00094 | 0.28495 | 1.3181 | 2.5475 | 98.52 | 20.05 | 49.83 | 99.1063 | 88.6666 | 92.9226 | 0.011875 | 0.96441 | 0.022206 | 23.0253 | -0.002765 | 0.15471 | 0.022227 | 1.4705 | 1.4705 | 0.00424
2019-04-09 14:48 | 1554835692.536 | 811 | 0.20515 | 0.003726 | 8.6e-05 | 99.2373 | 99.1992 | 111.3333 | 20.5498 | 0.003247 | 0.002928 | 230.9231 | 0.26097 | 8.402 | 0.002923 | 0.004281 | 0.004303 | 4.5e-05 | 0.000934 | 0.26576 | 1.346 | 2.5476 | 98.52 | 20.05 | 49.83 | 99.2659 | 91.457 | 91.5821 | 0.011875 | 0.96796 | 0.022265 | 23.0024 | 0.003433 | 0.15525 | 0.022141 | 1.4687 | 1.4687 | 0.00426
2019-04-09 14:48 | 1554835693.536 | 812 | 0.20515 | 0.003744 | 8.5e-05 | 99.2867 | 99.1699 | 111.3704 | 20.5502 | 0.003295 | 0.003016 | 230.8465 | 0.26024 | 8.403 | 0.002873 | 0.004098 | 0.004102 | 4.2e-05 | 0.000932 | 0.28008 | 1.3377 | 2.5476 | 98.52 | 20.06 | 49.831 | 100.4324 | 90.2482 | 93.6904 | 0.011882 | 0.9624 | 0.022205 | 23.0727 | 0.010526 | 0.15482 | 0.022105 | 1.4624 | 1.4624 | 0.004266
2019-04-09 14:48 | 1554835694.536 | 813 | 0.20511 | 0.003762 | 8.5e-05 | 99.2784 | 99.2205 | 111.2976 | 20.5493 | 0.003269 | 0.002999 | 230.8031 | 0.26066 | 8.4021 | 0.002888 | 0.004514 | 0.004377 | 4.3e-05 | 0.000934 | 0.26792 | 1.2849 | 2.5475 | 98.52 | 20.05 | 49.831 | 99.8068 | 90.6178 | 93.2865 | 0.011875 | 0.96346 | 0.02221 | 23.0524 | 0.005213 | 0.15507 | 0.022341 | 1.4784 | 1.4784 | 0.004227
2019-04-09 14:48 | 1554835695.5361 | 814 | 0.20511 | 0.003767 | 8.5e-05 | 99.2444 | 99.4129 | 111.4314 | 20.5501 | 0.003207 | 0.003044 | 230.7166 | 0.26108 | 8.4026 | 0.002926 | 0.004583 | 0.004525 | 4.5e-05 | 0.000936 | 0.28854 | 1.3554 | 2.5476 | 98.52 | 20.05 | 49.832 | 98.3104 | 91.512 | 94.3673 | 0.011875 | 0.96062 | 0.022191 | 23.1003 | -0.015172 | 0.15532 | 0.022339 | 1.4769 | 1.4769 | 0.004238
2019-04-09 14:48 | 1554835696.5377 | 815 | 0.20506 | 0.003786 | 8.6e-05 | 99.2518 | 99.0326 | 111.7859 | 20.5504 | 0.003259 | 0.002961 | 230.6913 | 0.26145 | 8.4026 | 0.002868 | 0.004259 | 0.00424 | 4.5e-05 | 0.000935 | 0.27793 | 1.3439 | 2.5477 | 98.52 | 20.05 | 49.833 | 99.5694 | 90.1258 | 92.3623 | 0.011875 | 0.96589 | 0.022287 | 23.0738 | 0.01978 | 0.15554 | 0.022617 | 1.5017 | 1.5017 | 0.004174
2019-04-09 14:48 | 1554835697.5372 | 816 | 0.20506 | 0.003799 | 8.5e-05 | 99.2724 | 99.1534 | 111.5812 | 20.5487 | 0.003187 | 0.002982 | 230.5858 | 0.2613 | 8.4019 | 0.002936 | 0.004415 | 0.004272 | 4.5e-05 | 0.000937 | 0.28335 | 1.3418 | 2.5476 | 98.52 | 20.05 | 49.837 | 97.8193 | 91.7705 | 92.8605 | 0.011876 | 0.96458 | 0.022251 | 23.0684 | 0.010728 | 0.15545 | 0.022588 | 1.4974 | 1.4974 | 0.004184
2019-04-09 14:48 | 1554835698.5362 | 817 | 0.2051 | 0.003819 | 8.5e-05 | 99.1586 | 99.2012 | 111.2901 | 20.5496 | 0.003278 | 0.002977 | 230.5167 | 0.26035 | 8.4 | 0.002982 | 0.004526 | 0.004494 | 4.6e-05 | 0.000933 | 0.29175 | 1.3411 | 2.5477 | 98.52 | 20.05 | 49.841 | 100.0102 | 92.8748 | 92.762 | 0.011877 | 0.96484 | 0.022225 | 23.0352 | -0.003841 | 0.15488 | 0.022344 | 1.4796 | 1.4796 | 0.004219
2019-04-09 14:48 | 1554835699.5375 | 818 | 0.20508 | 0.003835 | 8.5e-05 | 99.1619 | 99.2291 | 111.2444 | 20.5494 | 0.003295 | 0.003044 | 230.4562 | 0.26121 | 8.4002 | 0.002977 | 0.003973 | 0.003952 | 4.3e-05 | 0.000931 | 0.28567 | 1.3539 | 2.5477 | 98.52 | 20.05 | 49.842 | 100.4246 | 92.7604 | 94.3713 | 0.011877 | 0.96061 | 0.022172 | 23.081 | -0.006059 | 0.1554 | 0.022437 | 1.4822 | 1.4822 | 0.004225
2019-04-09 14:48 | 1554835700.5363 | 819 | 0.20511 | 0.003847 | 8.4e-05 | 99.1542 | 99.1207 | 111.2566 | 20.5495 | 0.003231 | 0.00302 | 230.3597 | 0.26065 | 8.4027 | 0.002888 | 0.00449 | 0.004412 | 4.2e-05 | 0.00093 | 0.28423 | 1.3205 | 2.5475 | 98.52 | 20.05 | 49.847 | 98.8816 | 90.5954 | 93.7922 | 0.011879 | 0.96213 | 0.022191 | 23.0641 | 0.00302 | 0.15506 | 0.022261 | 1.4718 | 1.4718 | 0.004246
2019-04-09 14:48 | 1554835701.5377 | 820 | 0.20505 | 0.003861 | 8.5e-05 | 99.1446 | 99.0763 | 111.4805 | 20.55 | 0.003205 | 0.002979 | 230.3216 | 0.26119 | 8.401 | 0.002933 | 0.004933 | 0.004802 | 4.6e-05 | 0.000928 | 0.24065 | 1.3124 | 2.5475 | 98.52 | 20.05 | 49.85 | 98.2627 | 91.6809 | 92.807 | 0.011879 | 0.96472 | 0.022243 | 23.0563 | 0.006167 | 0.15538 | 0.02259 | 1.4971 | 1.4971 | 0.004183
2019-04-09 14:48 | 1554835702.5378 | 821 | 0.20509 | 0.003873 | 8.5e-05 | 99.2048 | 99.1018 | 111.1401 | 20.5506 | 0.003255 | 0.00298 | 230.2178 | 0.26121 | 8.4018 | 0.002957 | 0.004307 | 0.004327 | 4.4e-05 | 0.000927 | 0.26001 | 1.3331 | 2.5475 | 98.52 | 20.05 | 49.855 | 99.4715 | 92.2687 | 92.8287 | 0.011881 | 0.96466 | 0.022208 | 23.0217 | 0.009294 | 0.1554 | 0.022325 | 1.4771 | 1.4771 | 0.00424
2019-04-09 14:48 | 1554835703.5374 | 822 | 0.205 | 0.003875 | 8.5e-05 | 99.1694 | 99.1971 | 111.1364 | 20.5504 | 0.003283 | 0.002991 | 230.1616 | 0.26132 | 8.4025 | 0.002908 | 0.004313 | 0.004225 | 4.5e-05 | 0.000929 | 0.27168 | 1.3486 | 2.5474 | 98.52 | 20.05 | 49.857 | 100.142 | 91.087 | 93.1007 | 0.011881 | 0.96395 | 0.0222 | 23.0299 | -0.002502 | 0.15546 | 0.022841 | 1.5107 | 1.5107 | 0.004147
2019-04-09 14:48 | 1554835704.5378 | 823 | 0.20501 | 0.003886 | 8.4e-05 | 99.125 | 99.142 | 111.1836 | 20.5497 | 0.003248 | 0.002989 | 230.1019 | 0.26052 | 8.4009 | 0.002878 | 0.004551 | 0.004509 | 4.9e-05 | 0.000935 | 0.27223 | 1.3163 | 2.5476 | 98.52 | 20.05 | 49.859 | 99.291 | 90.3696 | 93.0367 | 0.011881 | 0.96411 | 0.022206 | 23.0328 | -0.001539 | 0.15498 | 0.022773 | 1.5068 | 1.5068 | 0.004145
2019-04-09 14:48 | 1554835705.5378 | 824 | 0.20503 | 0.003897 | 8.4e-05 | 99.0717 | 99.1334 | 110.9777 | 20.5499 | 0.00324 | 0.00301 | 230.0309 | 0.26055 | 8.4031 | 0.002922 | 0.004644 | 0.004544 | 4.7e-05 | 0.000939 | 0.31078 | 1.3244 | 2.5476 | 98.52 | 20.05 | 49.861 | 99.099 | 91.4167 | 93.5416 | 0.011882 | 0.96279 | 0.02217 | 23.0273 | -0.005574 | 0.155 | 0.022645 | 1.4959 | 1.4959 | 0.004176
2019-04-09 14:48 | 1554835706.5364 | 825 | 0.20502 | 0.003909 | 8.3e-05 | 99.1178 | 99.1789 | 110.8413 | 20.5495 | 0.003296 | 0.003023 | 229.9938 | 0.26129 | 8.401 | 0.002868 | 0.00422 | 0.004148 | 4.7e-05 | 0.00094 | 0.28351 | 1.3459 | 2.5473 | 98.52 | 20.05 | 49.863 | 100.4582 | 90.1148 | 93.8657 | 0.011882 | 0.96193 | 0.022147 | 23.0233 | -0.005505 | 0.15544 | 0.022724 | 1.4995 | 1.4995 | 0.004178
2019-04-09 14:48 | 1554835707.5366 | 826 | 0.20503 | 0.003913 | 8.4e-05 | 99.0866 | 99.1823 | 110.9981 | 20.5494 | 0.003283 | 0.003012 | 229.9148 | 0.26009 | 8.4038 | 0.002894 | 0.004112 | 0.004103 | 4.5e-05 | 0.000941 | 0.28622 | 1.3192 | 2.5475 | 98.52 | 20.05 | 49.865 | 100.1466 | 90.7579 | 93.6052 | 0.011883 | 0.96262 | 0.02217 | 23.0314 | -0.008638 | 0.15473 | 0.022635 | 1.4952 | 1.4952 | 0.00417
2019-04-09 14:48 | 1554835708.5355 | 827 | 0.20498 | 0.003919 | 8.3e-05 | 99.0702 | 98.9404 | 111.1378 | 20.5496 | 0.003255 | 0.003033 | 229.8864 | 0.26018 | 8.4029 | 0.002889 | 0.004782 | 0.004671 | 4.9e-05 | 0.000945 | 0.27328 | 1.3276 | 2.5475 | 98.52 | 20.05 | 49.866 | 99.4691 | 90.622 | 94.0936 | 0.011883 | 0.96134 | 0.02217 | 23.0612 | 0.011729 | 0.15478 | 0.02293 | 1.5147 | 1.5147 | 0.004118
2019-04-09 14:48 | 1554835709.5365 | 828 | 0.20489 | 0.003924 | 8.3e-05 | 99.041 | 98.9349 | 111.1695 | 20.55 | 0.003306 | 0.003034 | 229.8048 | 0.26059 | 8.4039 | 0.002882 | 0.004562 | 0.004542 | 4.6e-05 | 0.000944 | 0.26609 | 1.3675 | 2.5476 | 98.52 | 20.05 | 49.866 | 100.7052 | 90.4544 | 94.1207 | 0.011883 | 0.96127 | 0.022172 | 23.0654 | 0.009592 | 0.15503 | 0.023467 | 1.5503 | 1.5503 | 0.00403
2019-04-09 14:48 | 1554835710.5366 | 829 | 0.20493 | 0.003938 | 8.3e-05 | 99.0708 | 99.0278 | 111.0946 | 20.5498 | 0.003291 | 0.002975 | 229.7825 | 0.25997 | 8.4034 | 0.002902 | 0.004745 | 0.004604 | 5.1e-05 | 0.00095 | 0.30254 | 1.3315 | 2.5476 | 98.52 | 20.05 | 49.867 | 100.3338 | 90.9411 | 92.713 | 0.011883 | 0.96497 | 0.022207 | 23.0134 | 0.003884 | 0.15466 | 0.023226 | 1.5368 | 1.5368 | 0.004056
2019-04-09 14:49 | 1554835711.5366 | 830 | 0.20491 | 0.003948 | 8.2e-05 | 99.1041 | 99.3351 | 110.8467 | 20.5507 | 0.003286 | 0.003005 | 229.7284 | 0.26089 | 8.4034 | 0.002892 | 0.005006 | 0.004974 | 4.6e-05 | 0.000952 | 0.27721 | 1.3809 | 2.5476 | 98.52 | 20.05 | 49.867 | 100.22 | 90.6934 | 93.4328 | 0.011883 | 0.96307 | 0.022161 | 23.0103 | -0.020821 | 0.1552 | 0.023325 | 1.5402 | 1.5402 | 0.004061
2019-04-09 14:49 | 1554835712.5356 | 831 | 0.2049 | 0.003964 | 8.1e-05 | 99.1009 | 98.9762 | 110.8085 | 20.5497 | 0.003285 | 0.002973 | 229.6717 | 0.26126 | 8.399 | 0.002905 | 0.004587 | 0.004485 | 4.9e-05 | 0.000957 | 0.28335 | 1.3406 | 2.5476 | 98.52 | 20.05 | 49.865 | 100.1941 | 91.0169 | 92.6432 | 0.011883 | 0.96515 | 0.022181 | 22.9815 | 0.011263 | 0.15542 | 0.023372 | 1.5447 | 1.5447 | 0.004055
2019-04-09 14:49 | 1554835713.5367 | 832 | 0.20491 | 0.003979 | 8.1e-05 | 99.0273 | 99.0206 | 110.8402 | 20.5503 | 0.003309 | 0.002978 | 229.6359 | 0.26118 | 8.4018 | 0.002852 | 0.004448 | 0.004389 | 4.8e-05 | 0.000952 | 0.27345 | 1.3231 | 2.5477 | 98.52 | 20.05 | 49.863 | 100.7714 | 89.731 | 92.7822 | 0.011882 | 0.96478 | 0.02218 | 22.9892 | 0.00061 | 0.15538 | 0.023264 | 1.5375 | 1.5375 | 0.004073
2019-04-09 14:49 | 1554835714.5363 | 833 | 0.20493 | 0.003992 | 8.1e-05 | 99.1298 | 98.9882 | 110.7852 | 20.5515 | 0.003292 | 0.003032 | 229.5383 | 0.26133 | 8.4034 | 0.002843 | 0.004154 | 0.004035 | 4.8e-05 | 0.000953 | 0.29717 | 1.3273 | 2.5476 | 98.52 | 20.05 | 49.864 | 100.3707 | 89.5305 | 94.0821 | 0.011883 | 0.96137 | 0.022135 | 23.0243 | 0.012781 | 0.15546 | 0.023139 | 1.5262 | 1.5262 | 0.004105
2019-04-09 14:49 | 1554835715.5367 | 834 | 0.20496 | 0.004001 | 8.1e-05 | 99.1131 | 98.9276 | 110.6761 | 20.5499 | 0.003239 | 0.003076 | 229.4928 | 0.26065 | 8.4025 | 0.002819 | 0.004353 | 0.004261 | 4.6e-05 | 0.000948 | 0.2792 | 1.3308 | 2.5475 | 98.52 | 20.05 | 49.862 | 99.0826 | 88.9487 | 95.1421 | 0.011882 | 0.9586 | 0.022092 | 23.0461 | 0.016751 | 0.15506 | 0.022964 | 1.5118 | 1.5118 | 0.004134
2019-04-09 14:49 | 1554835716.5367 | 835 | 0.2049 | 0.004006 | 8.1e-05 | 99.0541 | 99.0636 | 110.2857 | 20.5509 | 0.003305 | 0.003043 | 229.4144 | 0.2604 | 8.4048 | 0.002852 | 0.004318 | 0.0043 | 5.1e-05 | 0.000949 | 0.29983 | 1.3105 | 2.5476 | 98.52 | 20.05 | 49.861 | 100.6739 | 89.7364 | 94.3431 | 0.011882 | 0.96069 | 0.022077 | 22.9805 | -0.000854 | 0.15491 | 0.023329 | 1.5348 | 1.5348 | 0.004068
2019-04-09 14:49 | 1554835717.5371 | 836 | 0.2049 | 0.004009 | 8.1e-05 | 99.0508 | 99.0199 | 110.2827 | 20.5509 | 0.003319 | 0.002999 | 229.3641 | 0.26051 | 8.4032 | 0.002851 | 0.004314 | 0.004282 | 5.3e-05 | 0.000947 | 0.27865 | 1.3158 | 2.5475 | 98.52 | 20.05 | 49.861 | 101.0027 | 89.7245 | 93.2804 | 0.011882 | 0.96347 | 0.022109 | 22.9469 | 0.002797 | 0.15498 | 0.023282 | 1.5338 | 1.5338 | 0.004072
2019-04-09 14:49 | 1554835718.5373 | 837 | 0.20485 | 0.004013 | 8.2e-05 | 99.0809 | 99.0061 | 110.2083 | 20.5508 | 0.00328 | 0.003002 | 229.3052 | 0.25991 | 8.4013 | 0.002827 | 0.004555 | 0.004457 | 4.7e-05 | 0.000947 | 0.28528 | 1.3185 | 2.5474 | 98.52 | 20.06 | 49.858 | 100.0593 | 89.1494 | 93.3539 | 0.011889 | 0.96328 | 0.022099 | 22.9415 | 0.006757 | 0.15462 | 0.023586 | 1.5532 | 1.5532 | 0.004012
2019-04-09 14:49 | 1554835719.5359 | 838 | 0.20486 | 0.004019 | 8.2e-05 | 99.0053 | 98.9476 | 110.0306 | 20.5508 | 0.003319 | 0.002995 | 229.2402 | 0.26022 | 8.4021 | 0.002804 | 0.004185 | 0.004102 | 4.7e-05 | 0.000948 | 0.31045 | 1.3226 | 2.5476 | 98.52 | 20.06 | 49.856 | 101.0176 | 88.5846 | 93.1897 | 0.011888 | 0.96371 | 0.022086 | 22.9178 | 0.005213 | 0.1548 | 0.023506 | 1.547 | 1.547 | 0.004033
2019-04-09 14:49 | 1554835720.537 | 839 | 0.2048 | 0.00403 | 8.3e-05 | 99.0389 | 98.9448 | 110.0273 | 20.5509 | 0.00331 | 0.003069 | 229.1829 | 0.26076 | 8.3994 | 0.002808 | 0.003912 | 0.003878 | 4.7e-05 | 0.000941 | 0.28744 | 1.3323 | 2.5476 | 98.52 | 20.05 | 49.856 | 100.7883 | 88.6817 | 94.9754 | 0.011881 | 0.95904 | 0.022032 | 22.9733 | 0.008501 | 0.15513 | 0.023859 | 1.5664 | 1.5664 | 0.003991
2019-04-09 14:49 | 1554835721.5363 | 840 | 0.20474 | 0.004034 | 8.4e-05 | 99.077 | 99.1799 | 110.1397 | 20.5517 | 0.003291 | 0.002995 | 229.0993 | 0.26076 | 8.3994 | 0.002757 | 0.004282 | 0.004205 | 5.1e-05 | 0.000944 | 0.27649 | 1.3347 | 2.5476 | 98.52 | 20.05 | 49.856 | 100.3436 | 87.4431 | 93.185 | 0.011881 | 0.96372 | 0.022097 | 22.929 | -0.009281 | 0.15513 | 0.0242 | 1.5933 | 1.5933 | 0.003924
2019-04-09 14:49 | 1554835722.5369 | 841 | 0.20474 | 0.004049 | 8.5e-05 | 99.0788 | 99.0481 | 110.4207 | 20.5515 | 0.003323 | 0.003099 | 229.0459 | 0.26076 | 8.4019 | 0.002795 | 0.004086 | 0.004063 | 4.9e-05 | 0.000942 | 0.29717 | 1.3882 | 2.5476 | 98.52 | 20.05 | 49.854 | 101.1037 | 88.365 | 95.7004 | 0.01188 | 0.95715 | 0.02205 | 23.0369 | 0.002776 | 0.15513 | 0.024183 | 1.5888 | 1.5888 | 0.003935
2019-04-09 14:49 | 1554835723.536 | 842 | 0.20467 | 0.004061 | 8.5e-05 | 99.0666 | 99.184 | 110.5937 | 20.552 | 0.00336 | 0.003042 | 228.987 | 0.26026 | 8.4044 | 0.002761 | 0.004646 | 0.004603 | 4.9e-05 | 0.00094 | 0.28672 | 1.3793 | 2.5476 | 98.52 | 20.05 | 49.857 | 102.0074 | 87.5509 | 94.3183 | 0.011881 | 0.96075 | 0.022109 | 23.0117 | -0.010594 | 0.15483 | 0.024599 | 1.6204 | 1.6204 | 0.003851
2019-04-09 14:49 | 1554835724.536 | 843 | 0.2047 | 0.004074 | 8.5e-05 | 99.0848 | 99.0354 | 110.5054 | 20.5517 | 0.003342 | 0.003029 | 228.9085 | 0.25967 | 8.4029 | 0.002816 | 0.004595 | 0.00454 | 4.9e-05 | 0.000934 | 0.29159 | 1.3574 | 2.5476 | 98.52 | 20.05 | 49.857 | 101.5704 | 88.8628 | 93.9962 | 0.011881 | 0.96159 | 0.022109 | 22.9925 | 0.004464 | 0.15448 | 0.02441 | 1.608 | 1.608 | 0.003871
2019-04-09 14:49 | 1554835725.5351 | 844 | 0.20468 | 0.00409 | 8.4e-05 | 99.1378 | 98.9321 | 110.5043 | 20.5517 | 0.003354 | 0.003025 | 228.8517 | 0.26085 | 8.4016 | 0.00275 | 0.00436 | 0.004309 | 4.9e-05 | 0.000929 | 0.28263 | 1.3586 | 2.5476 | 98.52 | 20.05 | 49.857 | 101.8696 | 87.283 | 93.9063 | 0.011881 | 0.96183 | 0.022112 | 22.9895 | 0.018577 | 0.15518 | 0.024491 | 1.6136 | 1.6136 | 0.003876
2019-04-09 14:49 | 1554835726.5361 | 845 | 0.20469 | 0.004105 | 8.4e-05 | 99.052 | 98.9627 | 110.9319 | 20.5514 | 0.003283 | 0.003053 | 228.7401 | 0.2608 | 8.3981 | 0.00278 | 0.004505 | 0.004441 | 4.9e-05 | 0.000923 | 0.27959 | 1.3573 | 2.5476 | 98.52 | 20.05 | 49.857 | 100.1524 | 88.0081 | 94.5775 | 0.011881 | 0.96007 | 0.022135 | 23.055 | 0.008067 | 0.15515 | 0.024457 | 1.613 | 1.613 | 0.003876
2019-04-09 14:49 | 1554835727.5362 | 846 | 0.20471 | 0.004125 | 8.4e-05 | 99.1506 | 99.2195 | 110.9763 | 20.5521 | 0.00335 | 0.003068 | 228.7046 | 0.26093 | 8.4009 | 0.002767 | 0.004018 | 0.003852 | 5e-05 | 0.000917 | 0.26665 | 1.3157 | 2.5475 | 98.52 | 20.05 | 49.857 | 101.7695 | 87.7057 | 94.9423 | 0.011881 | 0.95912 | 0.022128 | 23.0711 | -0.006211 | 0.15523 | 0.024273 | 1.6004 | 1.6004 | 0.003909
2019-04-09 14:49 | 1554835728.5351 | 847 | 0.20465 | 0.004136 | 8.2e-05 | 99.1312 | 99.2346 | 111.0582 | 20.5518 | 0.00331 | 0.003034 | 228.6132 | 0.26104 | 8.4004 | 0.002733 | 0.003705 | 0.003641 | 4.7e-05 | 0.000919 | 0.29623 | 1.3005 | 2.5474 | 98.52 | 20.05 | 49.857 | 100.787 | 86.8684 | 94.1302 | 0.011881 | 0.96124 | 0.022161 | 23.0541 | -0.009323 | 0.15529 | 0.024665 | 1.6288 | 1.6288 | 0.003842
2019-04-09 14:49 | 1554835729.5352 | 848 | 0.20472 | 0.004147 | 8.3e-05 | 99.1637 | 99.0949 | 111.2462 | 20.5522 | 0.003268 | 0.003036 | 228.5527 | 0.26098 | 8.4029 | 0.002851 | 0.004216 | 0.004133 | 5.3e-05 | 0.000911 | 0.2975 | 1.3257 | 2.5476 | 98.52 | 20.05 | 49.858 | 99.7886 | 89.7119 | 94.1777 | 0.011881 | 0.96112 | 0.022178 | 23.0751 | 0.006209 | 0.15526 | 0.024199 | 1.5992 | 1.5992 | 0.003913
2019-04-09 14:49 | 1554835730.5363 | 849 | 0.20472 | 0.00415 | 8.3e-05 | 99.1155 | 99.1348 | 111.2535 | 20.5511 | 0.003345 | 0.003031 | 228.4885 | 0.2609 | 8.4022 | 0.0028 | 0.00452 | 0.004455 | 4.7e-05 | 0.000918 | 0.26698 | 1.3173 | 2.5477 | 98.52 | 20.05 | 49.859 | 101.6374 | 88.4857 | 94.0475 | 0.011881 | 0.96146 | 0.022183 | 23.0718 | -0.001746 | 0.15521 | 0.0242 | 1.5996 | 1.5996 | 0.00391
2019-04-09 14:49 | 1554835731.5363 | 850 | 0.20471 | 0.004149 | 8.2e-05 | 99.1881 | 99.1785 | 111.2798 | 20.5521 | 0.003376 | 0.00309 | 228.4119 | 0.2605 | 8.4006 | 0.002836 | 0.004627 | 0.004586 | 5.2e-05 | 0.000916 | 0.28655 | 1.309 | 2.5476 | 98.52 | 20.05 | 49.86 | 102.3899 | 89.3476 | 95.4707 | 0.011882 | 0.95775 | 0.022142 | 23.1192 | 0.000866 | 0.15497 | 0.024289 | 1.6026 | 1.6026 | 0.003897
2019-04-09 14:49 | 1554835732.5353 | 851 | 0.20473 | 0.004142 | 8.2e-05 | 99.2224 | 99.2074 | 111.1061 | 20.5511 | 0.003374 | 0.003068 | 228.365 | 0.26126 | 8.4012 | 0.002826 | 0.004335 | 0.004242 | 5.3e-05 | 0.000919 | 0.2912 | 1.2756 | 2.5475 | 98.52 | 20.05 | 49.86 | 102.3442 | 89.1114 | 94.9483 | 0.011882 | 0.95911 | 0.022141 | 23.0848 | 0.001347 | 0.15543 | 0.02413 | 1.592 | 1.592 | 0.003935
2019-04-09 14:49 | 1554835733.5364 | 852 | 0.20471 | 0.004135 | 8.2e-05 | 99.1848 | 99.0791 | 110.9256 | 20.5521 | 0.00333 | 0.003107 | 228.2594 | 0.26061 | 8.4022 | 0.002837 | 0.00452 | 0.004433 | 5e-05 | 0.000914 | 0.29966 | 1.334 | 2.5474 | 98.52 | 20.05 | 49.86 | 101.2676 | 89.3905 | 95.8786 | 0.011882 | 0.95669 | 0.022095 | 23.0951 | 0.009545 | 0.15504 | 0.024308 | 1.6004 | 1.6004 | 0.003904
2019-04-09 14:49 | 1554835734.5364 | 853 | 0.2047 | 0.004141 | 8.2e-05 | 99.262 | 99.3936 | 110.7303 | 20.5514 | 0.003396 | 0.00305 | 228.2296 | 0.2604 | 8.4023 | 0.002829 | 0.004368 | 0.004242 | 4.9e-05 | 0.000917 | 0.25857 | 1.3566 | 2.5475 | 98.52 | 20.05 | 49.861 | 102.8842 | 89.1813 | 94.5189 | 0.011882 | 0.96023 | 0.022116 | 23.0322 | -0.011855 | 0.15491 | 0.02431 | 1.6021 | 1.6021 | 0.003897
2019-04-09 14:49 | 1554835735.5355 | 854 | 0.2047 | 0.004143 | 8.2e-05 | 99.2417 | 99.236 | 110.5221 | 20.552 | 0.003331 | 0.003147 | 228.1621 | 0.26093 | 8.4019 | 0.002808 | 0.00378 | 0.003795 | 4.8e-05 | 0.000912 | 0.29518 | 1.3345 | 2.5476 | 98.52 | 20.05 | 49.858 | 101.2926 | 88.6757 | 96.8624 | 0.011881 | 0.95415 | 0.022025 | 23.0838 | 0.000516 | 0.15523 | 0.024337 | 1.5973 | 1.5973 | 0.003916
2019-04-09 14:49 | 1554835736.5364 | 855 | 0.20464 | 0.00415 | 8.2e-05 | 99.2501 | 99.4143 | 110.4806 | 20.5516 | 0.003365 | 0.003075 | 228.1251 | 0.26005 | 8.4006 | 0.002903 | 0.004404 | 0.004329 | 4.9e-05 | 0.000911 | 0.29856 | 1.3562 | 2.5476 | 98.52 | 20.05 | 49.858 | 102.1333 | 90.9812 | 95.1148 | 0.011881 | 0.95867 | 0.022073 | 23.0249 | -0.014786 | 0.15471 | 0.024656 | 1.6217 | 1.6217 | 0.003844
2019-04-09 14:49 | 1554835737.5345 | 856 | 0.20464 | 0.004161 | 8.2e-05 | 99.2709 | 99.269 | 110.3353 | 20.5525 | 0.003387 | 0.003127 | 228.0611 | 0.25986 | 8.4012 | 0.00294 | 0.004117 | 0.004129 | 5.1e-05 | 0.000907 | 0.28279 | 1.3059 | 2.5477 | 98.52 | 20.05 | 49.858 | 102.6486 | 91.8671 | 96.3647 | 0.011881 | 0.95543 | 0.022021 | 23.0488 | 0.000169 | 0.15459 | 0.024658 | 1.618 | 1.618 | 0.00385
2019-04-09 14:49 | 1554835738.5345 | 857 | 0.20466 | 0.004162 | 8.2e-05 | 99.3486 | 99.4363 | 110.3113 | 20.5528 | 0.003382 | 0.003061 | 227.9886 | 0.26033 | 8.4026 | 0.002832 | 0.004724 | 0.004619 | 5e-05 | 0.0009 | 0.26305 | 1.3295 | 2.5476 | 98.52 | 20.05 | 49.858 | 102.5477 | 89.2568 | 94.7707 | 0.011881 | 0.95957 | 0.022067 | 22.9965 | -0.007889 | 0.15487 | 0.024546 | 1.6141 | 1.6141 | 0.003867
2019-04-09 14:49 | 1554835739.5356 | 858 | 0.20466 | 0.004171 | 8.2e-05 | 99.3001 | 99.2838 | 110.2827 | 20.5521 | 0.003372 | 0.003059 | 227.9378 | 0.26117 | 8.4037 | 0.00285 | 0.004485 | 0.004392 | 5.1e-05 | 0.000902 | 0.28423 | 1.372 | 2.5477 | 98.52 | 20.05 | 49.859 | 102.2972 | 89.6945 | 94.7372 | 0.011881 | 0.95966 | 0.022065 | 22.9925 | 0.001464 | 0.15537 | 0.024553 | 1.6144 | 1.6144 | 0.003879
2019-04-09 14:49 | 1554835740.5359 | 859 | 0.20466 | 0.004179 | 8.1e-05 | 99.3742 | 99.4879 | 110.247 | 20.5527 | 0.003362 | 0.003148 | 227.8558 | 0.26001 | 8.4017 | 0.002947 | 0.003993 | 0.003941 | 5e-05 | 0.000902 | 0.28064 | 1.3309 | 2.5474 | 98.52 | 20.05 | 49.859 | 102.0536 | 92.0277 | 96.8866 | 0.011881 | 0.95408 | 0.021997 | 23.0558 | -0.010226 | 0.15468 | 0.02456 | 1.6099 | 1.6099 | 0.003872
2019-04-09 14:49 | 1554835741.535 | 860 | 0.20467 | 0.004186 | 8.1e-05 | 99.3996 | 99.3723 | 110.1055 | 20.5517 | 0.003369 | 0.003145 | 227.8146 | 0.26009 | 8.4027 | 0.002879 | 0.004162 | 0.004167 | 5.2e-05 | 0.000899 | 0.29264 | 1.3588 | 2.5476 | 98.52 | 20.05 | 49.858 | 102.215 | 90.3973 | 96.7972 | 0.011881 | 0.95431 | 0.021986 | 23.0382 | 0.002455 | 0.15473 | 0.024441 | 1.6013 | 1.6013 | 0.003894
2019-04-09 14:49 | 1554835742.5356 | 861 | 0.20471 | 0.004196 | 8.1e-05 | 99.3996 | 99.4504 | 110.0696 | 20.5518 | 0.003357 | 0.003089 | 227.7431 | 0.26089 | 8.4023 | 0.002834 | 0.004617 | 0.004557 | 4.9e-05 | 0.0009 | 0.30237 | 1.3264 | 2.5476 | 98.52 | 20.05 | 49.859 | 101.9207 | 89.298 | 95.4562 | 0.011881 | 0.95778 | 0.022022 | 22.9927 | -0.004574 | 0.1552 | 0.02423 | 1.59 | 1.59 | 0.003934
2019-04-09 14:49 | 1554835743.5356 | 862 | 0.20472 | 0.004199 | 8.2e-05 | 99.3978 | 99.587 | 110.3078 | 20.552 | 0.00338 | 0.003091 | 227.6872 | 0.26013 | 8.4018 | 0.002826 | 0.003671 | 0.003688 | 4.7e-05 | 0.000898 | 0.28152 | 1.3114 | 2.5475 | 98.52 | 20.05 | 49.859 | 102.4904 | 89.1054 | 95.4958 | 0.011881 | 0.95768 | 0.022045 | 23.0188 | -0.017015 | 0.15475 | 0.024143 | 1.586 | 1.586 | 0.003932
2019-04-09 14:49 | 1554835744.5357 | 863 | 0.20469 | 0.004193 | 8.2e-05 | 99.4249 | 99.5667 | 110.3386 | 20.5527 | 0.003349 | 0.003054 | 227.6542 | 0.2603 | 8.4014 | 0.002844 | 0.004178 | 0.00409 | 5e-05 | 0.0009 | 0.28655 | 1.2891 | 2.5476 | 98.52 | 20.05 | 49.858 | 101.7386 | 89.5428 | 94.6227 | 0.011881 | 0.95996 | 0.022074 | 22.9947 | -0.012753 | 0.15485 | 0.024346 | 1.6014 | 1.6014 | 0.003897
2019-04-09 14:49 | 1554835745.5358 | 864 | 0.20469 | 0.004179 | 8.3e-05 | 99.538 | 99.4143 | 110.2552 | 20.5529 | 0.003418 | 0.003067 | 227.5648 | 0.26035 | 8.4027 | 0.002827 | 0.004087 | 0.003973 | 4.8e-05 | 0.000896 | 0.27544 | 1.3508 | 2.5477 | 98.52 | 20.05 | 49.857 | 103.42 | 89.1522 | 94.9368 | 0.011881 | 0.95914 | 0.022056 | 22.9958 | 0.011128 | 0.15488 | 0.024343 | 1.5999 | 1.5999 | 0.003902
2019-04-09 14:49 | 1554835746.5357 | 865 | 0.20468 | 0.004174 | 8.3e-05 | 99.5282 | 99.6221 | 110.3709 | 20.5525 | 0.003487 | 0.003088 | 227.5104 | 0.26175 | 8.4004 | 0.002822 | 0.004389 | 0.004374 | 4.8e-05 | 0.0009 | 0.30127 | 1.3455 | 2.5476 | 98.52 | 20.05 | 49.856 | 105.088 | 89.03 | 95.4398 | 0.011881 | 0.95783 | 0.022053 | 23.0236 | -0.008438 | 0.15572 | 0.024441 | 1.606 | 1.606 | 0.003907
2019-04-09 14:49 | 1554835747.5347 | 866 | 0.20461 | 0.00417 | 8.3e-05 | 99.5467 | 99.4711 | 110.4881 | 20.5533 | 0.003375 | 0.003159 | 227.4444 | 0.26147 | 8.402 | 0.002857 | 0.00422 | 0.004148 | 5e-05 | 0.0009 | 0.28423 | 1.3671 | 2.5476 | 98.52 | 20.05 | 49.855 | 102.3674 | 89.8724 | 97.1419 | 0.011881 | 0.95342 | 0.022014 | 23.089 | 0.006796 | 0.15555 | 0.024839 | 1.6293 | 1.6293 | 0.003847
2019-04-09 14:49 | 1554835748.5358 | 867 | 0.20471 | 0.004169 | 8.4e-05 | 99.5812 | 99.6786 | 110.4545 | 20.5525 | 0.003435 | 0.0031 | 227.3922 | 0.26019 | 8.4037 | 0.002848 | 0.004284 | 0.004175 | 5.2e-05 | 0.000911 | 0.26665 | 1.3981 | 2.5476 | 98.52 | 20.05 | 49.853 | 103.8244 | 89.6409 | 95.7263 | 0.01188 | 0.95708 | 0.022052 | 23.0413 | -0.008742 | 0.15479 | 0.024225 | 1.5918 | 1.5918 | 0.003919
2019-04-09 14:49 | 1554835749.5359 | 868 | 0.20465 | 0.004176 | 8.4e-05 | 99.6253 | 99.6328 | 110.3235 | 20.5537 | 0.003393 | 0.003058 | 227.3474 | 0.26031 | 8.4034 | 0.002824 | 0.00504 | 0.004986 | 4.8e-05 | 0.000909 | 0.30773 | 1.4121 | 2.5475 | 98.52 | 20.05 | 49.851 | 102.8157 | 89.0656 | 94.6973 | 0.01188 | 0.95976 | 0.02207 | 22.9955 | -0.000675 | 0.15486 | 0.024585 | 1.6167 | 1.6167 | 0.00386
2019-04-09 14:49 | 1554835750.5359 | 869 | 0.20464 | 0.00419 | 8.3e-05 | 99.6408 | 99.7495 | 110.285 | 20.5529 | 0.003376 | 0.003103 | 227.2757 | 0.26061 | 8.4014 | 0.002867 | 0.004754 | 0.004686 | 5.2e-05 | 0.00091 | 0.31669 | 1.372 | 2.5475 | 98.52 | 20.05 | 49.847 | 102.3984 | 90.1063 | 95.7855 | 0.011879 | 0.95693 | 0.022034 | 23.0254 | -0.009753 | 0.15504 | 0.024652 | 1.6185 | 1.6185 | 0.00386
2019-04-09 14:49 | 1554835751.5351 | 870 | 0.20466 | 0.004193 | 8.3e-05 | 99.7238 | 99.7371 | 110.3306 | 20.5532 | 0.003431 | 0.003096 | 227.2365 | 0.26077 | 8.4018 | 0.002881 | 0.004741 | 0.004626 | 4.9e-05 | 0.000909 | 0.30021 | 1.3296 | 2.5476 | 98.52 | 20.05 | 49.845 | 103.7322 | 90.4364 | 95.626 | 0.011878 | 0.95734 | 0.022043 | 23.0252 | -0.001187 | 0.15513 | 0.024516 | 1.6103 | 1.6103 | 0.003882
2019-04-09 14:49 | 1554835752.5359 | 871 | 0.20476 | 0.004201 | 8.2e-05 | 99.7176 | 99.8448 | 110.5139 | 20.553 | 0.003411 | 0.003154 | 227.1708 | 0.26051 | 8.4012 | 0.002846 | 0.004744 | 0.004679 | 4.6e-05 | 0.000915 | 0.29662 | 1.3358 | 2.5474 | 98.52 | 20.05 | 49.843 | 103.2425 | 89.5923 | 97.0264 | 0.011878 | 0.95372 | 0.02202 | 23.0881 | -0.011403 | 0.15498 | 0.023922 | 1.5696 | 1.5696 | 0.003979
2019-04-09 14:49 | 1554835753.5349 | 872 | 0.20473 | 0.004203 | 8.2e-05 | 99.7989 | 99.7202 | 110.208 | 20.553 | 0.003425 | 0.003098 | 227.11 | 0.26007 | 8.4011 | 0.00287 | 0.004157 | 0.004053 | 5.4e-05 | 0.00091 | 0.28064 | 1.3485 | 2.5475 | 98.52 | 20.05 | 49.838 | 103.5867 | 90.1705 | 95.6635 | 0.011876 | 0.95725 | 0.02203 | 23.0136 | 0.007054 | 0.15472 | 0.024105 | 1.5824 | 1.5824 | 0.00394
2019-04-09 14:49 | 1554835754.5354 | 873 | 0.2048 | 0.004199 | 8.2e-05 | 99.7712 | 99.9525 | 110.1752 | 20.553 | 0.003404 | 0.003088 | 227.0241 | 0.26046 | 8.4016 | 0.002866 | 0.005052 | 0.004929 | 4.9e-05 | 0.000913 | 0.2724 | 1.3498 | 2.5475 | 98.52 | 20.04 | 49.835 | 103.0804 | 90.0821 | 95.4284 | 0.011868 | 0.95786 | 0.022033 | 23.0028 | -0.016241 | 0.15495 | 0.023693 | 1.5556 | 1.5556 | 0.004014
2019-04-09 14:49 | 1554835755.5354 | 874 | 0.20478 | 0.004191 | 8.3e-05 | 99.812 | 100.0991 | 110.5299 | 20.5526 | 0.003426 | 0.003078 | 226.9827 | 0.26029 | 8.4019 | 0.002875 | 0.00421 | 0.004206 | 4.9e-05 | 0.000911 | 0.27472 | 1.3442 | 2.5475 | 98.52 | 20.04 | 49.83 | 103.6147 | 90.2952 | 95.1972 | 0.011867 | 0.95846 | 0.022076 | 23.0326 | -0.025695 | 0.15485 | 0.023801 | 1.5657 | 1.5657 | 0.003986
2019-04-09 14:49 | 1554835756.535 | 875 | 0.20479 | 0.004178 | 8.2e-05 | 99.8596 | 100.0685 | 110.2165 | 20.5516 | 0.003374 | 0.003108 | 226.9457 | 0.26069 | 8.3998 | 0.002811 | 0.004934 | 0.00485 | 5e-05 | 0.000908 | 0.28639 | 1.3035 | 2.5476 | 98.52 | 20.04 | 49.828 | 102.3447 | 88.7635 | 95.9142 | 0.011867 | 0.9566 | 0.022023 | 23.0223 | -0.018688 | 0.15509 | 0.023776 | 1.5603 | 1.5603 | 0.004006
2019-04-09 14:49 | 1554835757.5345 | 876 | 0.20477 | 0.004171 | 8.3e-05 | 99.9171 | 99.9615 | 110.0895 | 20.553 | 0.003387 | 0.003065 | 226.8805 | 0.26015 | 8.4017 | 0.002944 | 0.004731 | 0.00467 | 4.9e-05 | 0.000901 | 0.27688 | 1.3159 | 2.5476 | 98.52 | 20.04 | 49.826 | 102.6594 | 91.9708 | 94.8672 | 0.011866 | 0.95932 | 0.022042 | 22.9764 | -0.00397 | 0.15476 | 0.023899 | 1.5697 | 1.5697 | 0.003973
2019-04-09 14:49 | 1554835758.5351 | 877 | 0.20475 | 0.004162 | 8.3e-05 | 99.9287 | 100.0331 | 110.3925 | 20.5529 | 0.003481 | 0.003106 | 226.8337 | 0.26089 | 8.4037 | 0.002942 | 0.004546 | 0.004394 | 5.1e-05 | 0.000906 | 0.29103 | 1.3599 | 2.5475 | 98.52 | 20.04 | 49.821 | 104.9405 | 91.9191 | 95.8676 | 0.011865 | 0.95672 | 0.022042 | 23.0392 | -0.009333 | 0.1552 | 0.024001 | 1.5764 | 1.5764 | 0.003968
2019-04-09 14:49 | 1554835759.5341 | 878 | 0.20478 | 0.004155 | 8.4e-05 | 99.9761 | 100.211 | 110.3519 | 20.5532 | 0.003407 | 0.003058 | 226.7632 | 0.26017 | 8.405 | 0.002935 | 0.004785 | 0.004734 | 5e-05 | 0.000908 | 0.29645 | 1.3684 | 2.5475 | 98.52 | 20.04 | 49.819 | 103.1436 | 91.7466 | 94.7077 | 0.011865 | 0.95973 | 0.022073 | 22.9988 | -0.020992 | 0.15478 | 0.023857 | 1.569 | 1.569 | 0.003975
2019-04-09 14:49 | 1554835760.5352 | 879 | 0.20482 | 0.004147 | 8.4e-05 | 100.0475 | 100.3648 | 110.2771 | 20.5524 | 0.003426 | 0.003058 | 226.7114 | 0.26032 | 8.4007 | 0.002912 | 0.004599 | 0.004492 | 4.9e-05 | 0.000905 | 0.286 | 1.3528 | 2.5476 | 98.52 | 20.04 | 49.818 | 103.604 | 91.1844 | 94.7013 | 0.011864 | 0.95975 | 0.022065 | 22.9908 | -0.02832 | 0.15487 | 0.023611 | 1.5524 | 1.5524 | 0.00402
2019-04-09 14:49 | 1554835761.5342 | 880 | 0.20475 | 0.004142 | 8.4e-05 | 100.0821 | 100.2389 | 110.3921 | 20.5537 | 0.003423 | 0.003148 | 226.6531 | 0.26042 | 8.4038 | 0.002913 | 0.004614 | 0.004543 | 4.9e-05 | 0.000907 | 0.28351 | 1.3447 | 2.5475 | 98.52 | 20.04 | 49.815 | 103.544 | 91.2232 | 96.8791 | 0.011864 | 0.9541 | 0.022012 | 23.0707 | -0.014 | 0.15492 | 0.024014 | 1.575 | 1.575 | 0.003964
2019-04-09 14:49 | 1554835762.5348 | 881 | 0.20476 | 0.004137 | 8.3e-05 | 100.0892 | 100.3338 | 110.1477 | 20.5526 | 0.003383 | 0.00308 | 226.5624 | 0.26052 | 8.4021 | 0.002988 | 0.004851 | 0.004759 | 5.1e-05 | 0.000913 | 0.29927 | 1.3396 | 2.5475 | 98.52 | 20.04 | 49.815 | 102.5679 | 93.0301 | 95.2355 | 0.011864 | 0.95836 | 0.022036 | 22.9939 | -0.021833 | 0.15498 | 0.023967 | 1.5737 | 1.5737 | 0.003969
2019-04-09 14:49 | 1554835763.5342 | 882 | 0.20488 | 0.004131 | 8.3e-05 | 100.1181 | 100.3806 | 110.3397 | 20.5529 | 0.003371 | 0.003151 | 226.5428 | 0.25955 | 8.3999 | 0.002924 | 0.004164 | 0.004128 | 5.1e-05 | 0.000911 | 0.27865 | 1.3055 | 2.5475 | 98.52 | 20.04 | 49.813 | 102.2805 | 91.4704 | 96.9463 | 0.011863 | 0.95393 | 0.022005 | 23.0674 | -0.023423 | 0.1544 | 0.023248 | 1.5243 | 1.5243 | 0.004082
2019-04-09 14:49 | 1554835764.5357 | 883 | 0.20488 | 0.004131 | 8.3e-05 | 100.1804 | 100.4529 | 110.535 | 20.5533 | 0.003422 | 0.003078 | 226.4591 | 0.26012 | 8.406 | 0.002986 | 0.004673 | 0.004561 | 4.9e-05 | 0.000915 | 0.30917 | 1.3078 | 2.5475 | 98.52 | 20.04 | 49.811 | 103.5005 | 92.9777 | 95.1939 | 0.011863 | 0.95847 | 0.022076 | 23.033 | -0.024302 | 0.15475 | 0.023298 | 1.5325 | 1.5325 | 0.004069
2019-04-09 14:49 | 1554835765.5351 | 884 | 0.2049 | 0.00412 | 8.2e-05 | 100.15 | 100.3717 | 110.4064 | 20.5533 | 0.003422 | 0.003135 | 226.3992 | 0.2603 | 8.4002 | 0.003049 | 0.004867 | 0.004848 | 4.9e-05 | 0.000914 | 0.28783 | 1.3119 | 2.5474 | 98.52 | 20.04 | 49.81 | 103.505 | 94.4971 | 96.555 | 0.011862 | 0.95494 | 0.022023 | 23.0621 | -0.01978 | 0.15485 | 0.023149 | 1.5191 | 1.5191 | 0.004108
2019-04-09 14:49 | 1554835766.5358 | 885 | 0.20493 | 0.004106 | 8.2e-05 | 100.2224 | 100.3782 | 110.5601 | 20.5527 | 0.003429 | 0.003076 | 226.3217 | 0.26065 | 8.4029 | 0.00301 | 0.00474 | 0.004663 | 4.8e-05 | 0.000923 | 0.28672 | 1.3908 | 2.5476 | 98.52 | 20.04 | 49.806 | 103.6831 | 93.5475 | 95.1449 | 0.011861 | 0.95859 | 0.02208 | 23.0341 | -0.013902 | 0.15506 | 0.023014 | 1.5142 | 1.5142 | 0.004127
2019-04-09 14:49 | 1554835767.5354 | 886 | 0.20493 | 0.004099 | 8.2e-05 | 100.2372 | 100.2633 | 110.7688 | 20.5531 | 0.003352 | 0.003129 | 226.2512 | 0.25978 | 8.4007 | 0.002986 | 0.004499 | 0.004416 | 4.9e-05 | 0.000925 | 0.29336 | 1.3655 | 2.5475 | 98.52 | 20.04 | 49.803 | 101.816 | 92.9609 | 96.4272 | 0.011861 | 0.95527 | 0.022063 | 23.0959 | -0.002324 | 0.15454 | 0.023014 | 1.513 | 1.513 | 0.004116
2019-04-09 14:49 | 1554835768.5355 | 887 | 0.20491 | 0.004081 | 8.2e-05 | 100.2777 | 100.5544 | 110.3721 | 20.5531 | 0.0034 | 0.003065 | 226.2033 | 0.25994 | 8.4011 | 0.002955 | 0.004005 | 0.003976 | 4.5e-05 | 0.000927 | 0.27632 | 1.3841 | 2.5477 | 98.52 | 20.04 | 49.798 | 102.9635 | 92.2321 | 94.8718 | 0.01186 | 0.95931 | 0.02207 | 23.006 | -0.024648 | 0.15464 | 0.023139 | 1.5217 | 1.5217 | 0.004095
2019-04-09 14:49 | 1554835769.5345 | 888 | 0.20494 | 0.004068 | 8.3e-05 | 100.3266 | 100.6866 | 110.4458 | 20.5532 | 0.003454 | 0.003107 | 226.1016 | 0.26005 | 8.4003 | 0.003032 | 0.004417 | 0.00434 | 4.8e-05 | 0.000928 | 0.26593 | 1.3762 | 2.5476 | 98.52 | 20.04 | 49.796 | 104.2731 | 94.088 | 95.8815 | 0.011859 | 0.95668 | 0.022047 | 23.0452 | -0.03204 | 0.15471 | 0.022973 | 1.5092 | 1.5092 | 0.004131
2019-04-09 14:49 | 1554835770.5334 | 889 | 0.2049 | 0.004056 | 8.3e-05 | 100.3686 | 100.3741 | 110.6237 | 20.5535 | 0.003437 | 0.003095 | 226.0549 | 0.26094 | 8.3983 | 0.002947 | 0.004644 | 0.004568 | 4.9e-05 | 0.000926 | 0.28711 | 1.3433 | 2.5476 | 98.52 | 20.04 | 49.795 | 103.8827 | 92.0204 | 95.6029 | 0.011859 | 0.9574 | 0.022073 | 23.0551 | -0.000494 | 0.15524 | 0.023251 | 1.5293 | 1.5293 | 0.004091
2019-04-09 14:50 | 1554835771.5335 | 890 | 0.20495 | 0.004046 | 8.3e-05 | 100.421 | 100.5107 | 110.4738 | 20.553 | 0.003403 | 0.003092 | 225.986 | 0.26045 | 8.4022 | 0.002952 | 0.004883 | 0.00485 | 4.6e-05 | 0.000923 | 0.28456 | 1.3525 | 2.5477 | 98.52 | 20.04 | 49.79 | 103.0429 | 92.1554 | 95.5216 | 0.011858 | 0.95761 | 0.02206 | 23.0369 | -0.007992 | 0.15494 | 0.02298 | 1.5106 | 1.5106 | 0.004134
2019-04-09 14:50 | 1554835772.5345 | 891 | 0.20484 | 0.004038 | 8.2e-05 | 100.4651 | 100.7224 | 110.7643 | 20.553 | 0.003524 | 0.003107 | 225.9451 | 0.26069 | 8.4018 | 0.00295 | 0.004588 | 0.004502 | 4.9e-05 | 0.000921 | 0.26322 | 1.333 | 2.5475 | 98.52 | 20.03 | 49.788 | 105.9877 | 92.1147 | 95.8858 | 0.01185 | 0.95667 | 0.022079 | 23.0786 | -0.022882 | 0.15509 | 0.023625 | 1.5543 | 1.5543 | 0.004021
2019-04-09 14:50 | 1554835773.5346 | 892 | 0.20511 | 0.004033 | 8.3e-05 | 100.5076 | 100.9767 | 110.7136 | 20.5531 | 0.003465 | 0.003127 | 225.8688 | 0.26009 | 8.401 | 0.002992 | 0.00473 | 0.004724 | 4.5e-05 | 0.000915 | 0.28456 | 1.3508 | 2.5476 | 98.52 | 20.03 | 49.782 | 104.5546 | 93.1261 | 96.3787 | 0.011848 | 0.95539 | 0.022059 | 23.0887 | -0.041649 | 0.15473 | 0.022022 | 1.4475 | 1.4475 | 0.004308
2019-04-09 14:50 | 1554835774.5336 | 893 | 0.20481 | 0.004028 | 8.3e-05 | 100.5041 | 100.8291 | 110.7129 | 20.5532 | 0.003408 | 0.003158 | 225.7998 | 0.26043 | 8.3993 | 0.002984 | 0.004165 | 0.004136 | 4.8e-05 | 0.000918 | 0.29264 | 1.3229 | 2.5476 | 98.52 | 20.04 | 49.78 | 103.1688 | 92.9299 | 97.1121 | 0.011855 | 0.9535 | 0.022037 | 23.1115 | -0.028879 | 0.15493 | 0.023802 | 1.563 | 1.563 | 0.003995
2019-04-09 14:50 | 1554835775.5341 | 894 | 0.2052 | 0.00402 | 8.3e-05 | 100.5347 | 100.8917 | 110.6963 | 20.5535 | 0.003404 | 0.003124 | 225.7634 | 0.26031 | 8.4031 | 0.002986 | 0.004828 | 0.004753 | 4.3e-05 | 0.000916 | 0.30917 | 1.2761 | 2.5475 | 98.52 | 20.04 | 49.779 | 103.0637 | 92.9606 | 96.3037 | 0.011855 | 0.95559 | 0.022059 | 23.0845 | -0.031705 | 0.15486 | 0.021505 | 1.4135 | 1.4135 | 0.004415
2019-04-09 14:50 | 1554835776.5337 | 895 | 0.20472 | 0.004008 | 8.3e-05 | 100.5335 | 100.7678 | 110.5409 | 20.5536 | 0.003398 | 0.003128 | 225.6783 | 0.25992 | 8.3993 | 0.002969 | 0.004476 | 0.004388 | 4.7e-05 | 0.000915 | 0.28031 | 1.2917 | 2.5476 | 98.52 | 20.04 | 49.78 | 102.9263 | 92.5533 | 96.397 | 0.011855 | 0.95535 | 0.022041 | 23.0712 | -0.02082 | 0.15463 | 0.024353 | 1.5995 | 1.5995 | 0.003896
2019-04-09 14:50 | 1554835777.5347 | 896 | 0.20493 | 0.004003 | 8.2e-05 | 100.6547 | 100.9753 | 110.4066 | 20.554 | 0.003385 | 0.003116 | 225.6303 | 0.26039 | 8.4055 | 0.003008 | 0.004703 | 0.004642 | 4.8e-05 | 0.000914 | 0.27433 | 1.2994 | 2.5476 | 98.52 | 20.04 | 49.78 | 102.6218 | 93.5041 | 96.1014 | 0.011855 | 0.95611 | 0.022036 | 23.048 | -0.028445 | 0.1549 | 0.023108 | 1.5174 | 1.5174 | 0.004114
2019-04-09 14:50 | 1554835778.5347 | 897 | 0.20472 | 0.003993 | 8.2e-05 | 100.6244 | 100.8938 | 110.8441 | 20.5544 | 0.003425 | 0.003149 | 225.5721 | 0.26063 | 8.4009 | 0.002998 | 0.004564 | 0.004522 | 4.9e-05 | 0.000911 | 0.29302 | 1.3027 | 2.5475 | 98.52 | 20.04 | 49.781 | 103.5927 | 93.2605 | 96.9007 | 0.011856 | 0.95405 | 0.022056 | 23.1186 | -0.023916 | 0.15505 | 0.024385 | 1.6027 | 1.6027 | 0.003899
2019-04-09 14:50 | 1554835779.5348 | 898 | 0.20476 | 0.003984 | 8.2e-05 | 100.6646 | 101.039 | 110.7763 | 20.5541 | 0.003375 | 0.0031 | 225.5181 | 0.26031 | 8.4016 | 0.003055 | 0.003843 | 0.00381 | 5.1e-05 | 0.000913 | 0.29214 | 1.3324 | 2.5476 | 98.52 | 20.03 | 49.786 | 102.3587 | 94.628 | 95.7184 | 0.011849 | 0.9571 | 0.022085 | 23.0746 | -0.033209 | 0.15486 | 0.024177 | 1.5911 | 1.5911 | 0.003922
2019-04-09 14:50 | 1554835780.5342 | 899 | 0.20529 | 0.003972 | 8.1e-05 | 100.7301 | 101.1092 | 110.7303 | 20.554 | 0.003415 | 0.00313 | 225.4819 | 0.26012 | 8.401 | 0.002931 | 0.004637 | 0.004507 | 4.5e-05 | 0.000913 | 0.29645 | 1.335 | 2.5476 | 98.52 | 20.03 | 49.79 | 103.3314 | 91.6496 | 96.4578 | 0.01185 | 0.95519 | 0.022058 | 23.0929 | -0.033603 | 0.15475 | 0.021011 | 1.381 | 1.381 | 0.004516
2019-04-09 14:50 | 1554835781.5349 | 900 | 0.20507 | 0.003961 | 8.2e-05 | 100.6726 | 101.1863 | 110.9207 | 20.5534 | 0.003422 | 0.003086 | 225.3892 | 0.26084 | 8.4001 | 0.003029 | 0.004255 | 0.004148 | 4.9e-05 | 0.000915 | 0.29502 | 1.3389 | 2.5475 | 98.52 | 20.04 | 49.794 | 103.5187 | 94.0073 | 95.3827 | 0.011859 | 0.95798 | 0.022109 | 23.0791 | -0.045525 | 0.15517 | 0.022366 | 1.4735 | 1.4735 | 0.004244
2019-04-09 14:50 | 1554835782.5352 | 901 | 0.20503 | 0.003935 | 8.1e-05 | 100.7223 | 100.9998 | 110.9448 | 20.5541 | 0.003494 | 0.003181 | 225.3708 | 0.26001 | 8.4015 | 0.003013 | 0.004443 | 0.004437 | 5e-05 | 0.00092 | 0.27671 | 1.3427 | 2.5475 | 98.52 | 20.04 | 49.798 | 105.2564 | 93.6325 | 97.6685 | 0.01186 | 0.95207 | 0.022043 | 23.1531 | -0.024604 | 0.15468 | 0.022589 | 1.4838 | 1.4838 | 0.004201
2019-04-09 14:50 | 1554835783.5343 | 902 | 0.20478 | 0.003912 | 8.2e-05 | 100.7593 | 100.9272 | 110.8122 | 20.5538 | 0.003391 | 0.003128 | 225.3084 | 0.26124 | 8.4004 | 0.002965 | 0.004201 | 0.00414 | 4.7e-05 | 0.000921 | 0.286 | 1.3367 | 2.5477 | 98.52 | 20.04 | 49.801 | 102.7567 | 92.4611 | 96.3918 | 0.01186 | 0.95536 | 0.022068 | 23.0994 | -0.014892 | 0.15541 | 0.02413 | 1.5868 | 1.5868 | 0.003947
2019-04-09 14:50 | 1554835784.534 | 903 | 0.20475 | 0.003884 | 8.2e-05 | 100.823 | 101.229 | 111.0812 | 20.5535 | 0.003387 | 0.003143 | 225.2645 | 0.25959 | 8.4039 | 0.002986 | 0.004327 | 0.004272 | 4.4e-05 | 0.000921 | 0.29231 | 1.3564 | 2.5474 | 98.52 | 20.04 | 49.81 | 102.6545 | 92.9637 | 96.762 | 0.011862 | 0.9544 | 0.022084 | 23.139 | -0.035945 | 0.15443 | 0.024392 | 1.6052 | 1.6052 | 0.003877
2019-04-09 14:50 | 1554835785.534 | 904 | 0.20534 | 0.003862 | 8.2e-05 | 100.7658 | 101.2689 | 111.1493 | 20.5543 | 0.003452 | 0.003127 | 225.2085 | 0.26061 | 8.4007 | 0.002998 | 0.004241 | 0.004211 | 4.7e-05 | 0.00092 | 0.25769 | 1.4155 | 2.5476 | 98.52 | 20.04 | 49.814 | 104.2254 | 93.2736 | 96.3843 | 0.011863 | 0.95538 | 0.022102 | 23.1342 | -0.044546 | 0.15504 | 0.020844 | 1.3726 | 1.3726 | 0.004552
2019-04-09 14:50 | 1554835786.5347 | 905 | 0.20516 | 0.003834 | 8.2e-05 | 100.8918 | 101.387 | 111.23 | 20.554 | 0.003378 | 0.003129 | 225.1635 | 0.25968 | 8.401 | 0.002979 | 0.004739 | 0.004708 | 4.6e-05 | 0.000923 | 0.27959 | 1.3383 | 2.5477 | 98.52 | 20.04 | 49.821 | 102.4475 | 92.8133 | 96.4312 | 0.011865 | 0.95526 | 0.022109 | 23.1441 | -0.043792 | 0.15449 | 0.021949 | 1.446 | 1.446 | 0.004306
2019-04-09 14:50 | 1554835787.533 | 906 | 0.20578 | 0.003822 | 8.1e-05 | 100.9388 | 101.3026 | 111.1812 | 20.5544 | 0.003413 | 0.003095 | 225.1222 | 0.26046 | 8.4006 | 0.002989 | 0.003936 | 0.003856 | 4.7e-05 | 0.000923 | 0.28279 | 1.3155 | 2.5474 | 98.52 | 20.03 | 49.823 | 103.2918 | 93.0398 | 95.6077 | 0.011858 | 0.95739 | 0.022128 | 23.1132 | -0.03218 | 0.15495 | 0.018246 | 1.2028 | 1.2028 | 0.005191
2019-04-09 14:50 | 1554835788.5331 | 907 | 0.20538 | 0.003805 | 8e-05 | 100.9454 | 101.2644 | 110.7115 | 20.5543 | 0.003414 | 0.003121 | 225.0565 | 0.26102 | 8.4021 | 0.002926 | 0.004575 | 0.004552 | 4.6e-05 | 0.000923 | 0.28528 | 1.3099 | 2.5475 | 98.52 | 20.03 | 49.825 | 103.3198 | 91.5216 | 96.2262 | 0.011859 | 0.95579 | 0.022063 | 23.0837 | -0.028225 | 0.15528 | 0.020696 | 1.3606 | 1.3606 | 0.004599
2019-04-09 14:50 | 1554835789.5331 | 908 | 0.20525 | 0.003791 | 8.1e-05 | 100.8689 | 101.1849 | 110.7777 | 20.5542 | 0.00343 | 0.003106 | 225.0319 | 0.26076 | 8.4012 | 0.002969 | 0.004049 | 0.003931 | 5.1e-05 | 0.000929 | 0.28495 | 1.3312 | 2.5474 | 98.52 | 20.04 | 49.831 | 103.7048 | 92.5608 | 95.8768 | 0.011867 | 0.95669 | 0.02208 | 23.0797 | -0.027985 | 0.15513 | 0.021474 | 1.4129 | 1.4129 | 0.004425
2019-04-09 14:50 | 1554835790.5332 | 909 | 0.2057 | 0.003782 | 8.1e-05 | 100.9335 | 101.2469 | 110.5322 | 20.555 | 0.003364 | 0.003067 | 224.9741 | 0.26037 | 8.4021 | 0.002946 | 0.004442 | 0.004349 | 4.6e-05 | 0.00094 | 0.28263 | 1.3152 | 2.5475 | 98.52 | 20.04 | 49.832 | 102.1139 | 91.9985 | 94.9176 | 0.011868 | 0.95919 | 0.022084 | 23.0241 | -0.027731 | 0.15489 | 0.018767 | 1.2348 | 1.2348 | 0.005055
2019-04-09 14:50 | 1554835791.5332 | 910 | 0.20525 | 0.003759 | 8.1e-05 | 100.9951 | 101.4245 | 110.6254 | 20.5541 | 0.003394 | 0.003109 | 224.9175 | 0.25968 | 8.4036 | 0.002987 | 0.004909 | 0.004838 | 4.5e-05 | 0.000933 | 0.28854 | 1.329 | 2.5476 | 98.52 | 20.04 | 49.832 | 102.8214 | 93.0075 | 95.9362 | 0.011868 | 0.95654 | 0.022063 | 23.0657 | -0.037944 | 0.15449 | 0.021486 | 1.4125 | 1.4125 | 0.004407
2019-04-09 14:50 | 1554835792.5336 | 911 | 0.20541 | 0.003745 | 8.1e-05 | 100.9493 | 101.2734 | 110.6751 | 20.5533 | 0.003445 | 0.003071 | 224.8656 | 0.25993 | 8.4032 | 0.0029 | 0.003583 | 0.003517 | 4.9e-05 | 0.000923 | 0.27671 | 1.3622 | 2.5477 | 98.52 | 20.04 | 49.833 | 104.072 | 90.9063 | 95.0115 | 0.011868 | 0.95894 | 0.022096 | 23.0419 | -0.028673 | 0.15463 | 0.020566 | 1.3541 | 1.3541 | 0.004602
2019-04-09 14:50 | 1554835793.5338 | 912 | 0.20514 | 0.003729 | 8e-05 | 101.0183 | 101.6871 | 110.5627 | 20.5544 | 0.003355 | 0.003093 | 224.7922 | 0.26044 | 8.4026 | 0.002932 | 0.004107 | 0.004017 | 4.7e-05 | 0.000923 | 0.27992 | 1.3578 | 2.5477 | 98.52 | 20.04 | 49.833 | 101.8949 | 91.6798 | 95.5606 | 0.011868 | 0.95751 | 0.022068 | 23.0474 | -0.059015 | 0.15493 | 0.022218 | 1.4611 | 1.4611 | 0.004273
2019-04-09 14:50 | 1554835794.5332 | 913 | 0.20565 | 0.003706 | 8e-05 | 100.9799 | 101.3443 | 110.5899 | 20.5552 | 0.003474 | 0.003106 | 224.7607 | 0.25997 | 8.4009 | 0.002928 | 0.00437 | 0.004292 | 5e-05 | 0.000921 | 0.29463 | 1.3497 | 2.5476 | 98.52 | 20.04 | 49.832 | 104.7728 | 91.5835 | 95.8684 | 0.011868 | 0.95671 | 0.022062 | 23.0598 | -0.032215 | 0.15466 | 0.019176 | 1.2605 | 1.2605 | 0.004945
2019-04-09 14:50 | 1554835795.5333 | 914 | 0.20543 | 0.003687 | 7.9e-05 | 101.0365 | 101.4424 | 110.6505 | 20.5547 | 0.003383 | 0.003077 | 224.6858 | 0.25927 | 8.4016 | 0.002977 | 0.003995 | 0.003951 | 4.7e-05 | 0.000921 | 0.27776 | 1.4001 | 2.5475 | 98.52 | 20.03 | 49.829 | 102.5612 | 92.7522 | 95.1741 | 0.01186 | 0.95852 | 0.022089 | 23.0444 | -0.035857 | 0.15424 | 0.020518 | 1.3505 | 1.3505 | 0.004603
2019-04-09 14:50 | 1554835796.5333 | 915 | 0.20561 | 0.003673 | 7.9e-05 | 101.0204 | 101.4692 | 110.6855 | 20.5546 | 0.003404 | 0.003126 | 224.6546 | 0.26019 | 8.4033 | 0.002921 | 0.004544 | 0.004515 | 4.6e-05 | 0.000916 | 0.28008 | 1.355 | 2.5475 | 98.52 | 20.03 | 49.824 | 103.0725 | 91.4101 | 96.3504 | 0.011858 | 0.95547 | 0.022057 | 23.0849 | -0.039647 | 0.15479 | 0.019435 | 1.2774 | 1.2774 | 0.004883
2019-04-09 14:50 | 1554835797.5343 | 916 | 0.20538 | 0.003656 | 7.8e-05 | 100.9966 | 101.4031 | 110.9953 | 20.5546 | 0.003414 | 0.00304 | 224.6167 | 0.25971 | 8.4026 | 0.002889 | 0.004511 | 0.004495 | 4.8e-05 | 0.000917 | 0.2808 | 1.3258 | 2.5476 | 98.52 | 20.03 | 49.82 | 103.3201 | 90.6362 | 94.2825 | 0.011857 | 0.96084 | 0.02215 | 23.0524 | -0.03593 | 0.1545 | 0.020854 | 1.3765 | 1.3765 | 0.004523
2019-04-09 14:50 | 1554835798.5334 | 917 | 0.20534 | 0.003637 | 7.8e-05 | 101.1002 | 101.4141 | 111.1768 | 20.5546 | 0.003358 | 0.003093 | 224.5436 | 0.26063 | 8.4023 | 0.002953 | 0.00457 | 0.004461 | 4.8e-05 | 0.000921 | 0.28241 | 1.3091 | 2.5475 | 98.52 | 20.03 | 49.816 | 101.9533 | 92.1695 | 95.5469 | 0.011856 | 0.95755 | 0.02213 | 23.1109 | -0.027729 | 0.15505 | 0.021101 | 1.3916 | 1.3916 | 0.00449
2019-04-09 14:50 | 1554835799.5335 | 918 | 0.20583 | 0.003628 | 7.8e-05 | 101.0374 | 101.5852 | 111.2807 | 20.5553 | 0.003399 | 0.003092 | 224.5034 | 0.25949 | 8.4018 | 0.002923 | 0.004689 | 0.004631 | 5.3e-05 | 0.000918 | 0.27113 | 1.3126 | 2.5476 | 98.52 | 20.03 | 49.807 | 102.9468 | 91.4602 | 95.5356 | 0.011854 | 0.95758 | 0.02214 | 23.1213 | -0.048363 | 0.15437 | 0.018192 | 1.2002 | 1.2002 | 0.005184
2019-04-09 14:50 | 1554835800.5324 | 919 | 0.20546 | 0.003611 | 7.7e-05 | 101.0812 | 101.5005 | 111.6938 | 20.5563 | 0.003363 | 0.003062 | 224.4216 | 0.26066 | 8.4009 | 0.002861 | 0.004407 | 0.004287 | 4.8e-05 | 0.000922 | 0.29911 | 1.302 | 2.5475 | 98.52 | 20.03 | 49.8 | 102.0876 | 89.9673 | 94.7992 | 0.011853 | 0.95949 | 0.022204 | 23.1411 | -0.037031 | 0.15507 | 0.020446 | 1.3529 | 1.3529 | 0.004619
2019-04-09 14:50 | 1554835801.5335 | 920 | 0.20563 | 0.00359 | 7.7e-05 | 100.9731 | 101.3825 | 111.6086 | 20.5547 | 0.003361 | 0.00308 | 224.4183 | 0.25994 | 8.4013 | 0.002898 | 0.00388 | 0.003904 | 5e-05 | 0.000918 | 0.29048 | 1.3065 | 2.5476 | 98.52 | 20.03 | 49.796 | 102.0381 | 90.8568 | 95.2361 | 0.011852 | 0.95836 | 0.022182 | 23.146 | -0.036193 | 0.15464 | 0.019462 | 1.2865 | 1.2865 | 0.004844
2019-04-09 14:50 | 1554835802.5335 | 921 | 0.20566 | 0.003577 | 7.7e-05 | 101.0407 | 101.4406 | 111.4145 | 20.5555 | 0.003377 | 0.00305 | 224.3423 | 0.26012 | 8.4042 | 0.002885 | 0.004785 | 0.004663 | 4.7e-05 | 0.000918 | 0.28567 | 1.3402 | 2.5476 | 98.52 | 20.03 | 49.786 | 102.4175 | 90.5442 | 94.508 | 0.011849 | 0.96025 | 0.022185 | 23.1029 | -0.035336 | 0.15474 | 0.01928 | 1.2746 | 1.2746 | 0.004893
2019-04-09 14:50 | 1554835803.5329 | 922 | 0.20567 | 0.003554 | 7.7e-05 | 101.0207 | 101.4923 | 111.6753 | 20.5557 | 0.003365 | 0.003104 | 224.3151 | 0.26123 | 8.3996 | 0.00285 | 0.004547 | 0.004505 | 4.9e-05 | 0.000919 | 0.28047 | 1.3587 | 2.5476 | 98.52 | 20.03 | 49.781 | 102.1233 | 89.6958 | 95.8267 | 0.011848 | 0.95682 | 0.022171 | 23.1714 | -0.041653 | 0.15541 | 0.01925 | 1.2718 | 1.2718 | 0.004924
2019-04-09 14:50 | 1554835804.533 | 923 | 0.20575 | 0.00353 | 7.7e-05 | 101.0508 | 101.3381 | 111.5152 | 20.5557 | 0.003396 | 0.003064 | 224.2775 | 0.26054 | 8.3999 | 0.002908 | 0.004466 | 0.004324 | 5e-05 | 0.00092 | 0.28456 | 1.3216 | 2.5475 | 98.52 | 20.03 | 49.773 | 102.8813 | 91.0946 | 94.8642 | 0.011846 | 0.95932 | 0.022184 | 23.1246 | -0.025393 | 0.155 | 0.018822 | 1.2443 | 1.2443 | 0.00502
2019-04-09 14:50 | 1554835805.5326 | 924 | 0.20579 | 0.003505 | 7.7e-05 | 101.035 | 101.494 | 111.2868 | 20.5559 | 0.003358 | 0.003128 | 224.2291 | 0.26074 | 8.4021 | 0.00298 | 0.004537 | 0.004464 | 5.1e-05 | 0.000921 | 0.27704 | 1.341 | 2.5476 | 98.52 | 20.03 | 49.769 | 101.971 | 92.8363 | 96.3939 | 0.011845 | 0.95535 | 0.022115 | 23.1488 | -0.04054 | 0.15512 | 0.018588 | 1.225 | 1.225 | 0.005103
2019-04-09 14:50 | 1554835806.5337 | 925 | 0.20581 | 0.00348 | 7.7e-05 | 101.0359 | 101.431 | 111.5117 | 20.5552 | 0.00335 | 0.003094 | 224.1955 | 0.26059 | 8.4026 | 0.002888 | 0.004227 | 0.004208 | 5.1e-05 | 0.000917 | 0.30718 | 1.3559 | 2.5475 | 98.52 | 20.03 | 49.766 | 101.7725 | 90.6236 | 95.5865 | 0.011845 | 0.95745 | 0.022162 | 23.1469 | -0.034908 | 0.15502 | 0.018526 | 1.2234 | 1.2234 | 0.005107
2019-04-09 14:50 | 1554835807.5328 | 926 | 0.20586 | 0.003451 | 7.7e-05 | 101.0797 | 101.5986 | 111.4145 | 20.5566 | 0.003402 | 0.003082 | 224.1219 | 0.26098 | 8.3983 | 0.002857 | 0.004604 | 0.004483 | 5.2e-05 | 0.000918 | 0.30469 | 1.3248 | 2.5476 | 98.51 | 20.03 | 49.76 | 103.0329 | 89.8575 | 95.2857 | 0.011844 | 0.95823 | 0.022161 | 23.1274 | -0.045802 | 0.15526 | 0.018222 | 1.2033 | 1.2033 | 0.0052
2019-04-09 14:50 | 1554835808.5328 | 927 | 0.20589 | 0.003422 | 7.7e-05 | 101.038 | 101.6041 | 111.5244 | 20.5553 | 0.003349 | 0.003057 | 224.1082 | 0.25996 | 8.4014 | 0.002887 | 0.003641 | 0.003574 | 5e-05 | 0.000918 | 0.30254 | 1.3419 | 2.5475 | 98.51 | 20.03 | 49.759 | 101.7373 | 90.5766 | 94.6783 | 0.011844 | 0.95981 | 0.022191 | 23.1197 | -0.049976 | 0.15465 | 0.018094 | 1.1964 | 1.1964 | 0.005209
2019-04-09 14:50 | 1554835809.5337 | 928 | 0.20579 | 0.003393 | 7.6e-05 | 101.0237 | 101.5387 | 111.5859 | 20.5567 | 0.003381 | 0.003088 | 224.0514 | 0.26076 | 8.4028 | 0.002965 | 0.004979 | 0.004944 | 5.5e-05 | 0.000919 | 0.30326 | 1.33 | 2.5475 | 98.51 | 20.03 | 49.758 | 102.5184 | 92.4616 | 95.4471 | 0.011844 | 0.95781 | 0.022173 | 23.1502 | -0.045483 | 0.15512 | 0.018727 | 1.2374 | 1.2374 | 0.005052
2019-04-09 14:50 | 1554835810.5338 | 929 | 0.20594 | 0.00337 | 7.7e-05 | 101.0761 | 101.7204 | 111.3873 | 20.5569 | 0.00334 | 0.003047 | 224.0182 | 0.26046 | 8.4045 | 0.00291 | 0.004016 | 0.003926 | 5.3e-05 | 0.000916 | 0.29767 | 1.3376 | 2.5476 | 98.51 | 20.03 | 49.759 | 101.5332 | 91.1348 | 94.442 | 0.011844 | 0.96043 | 0.022184 | 23.0981 | -0.056837 | 0.15495 | 0.017831 | 1.1786 | 1.1786 | 0.005298
2019-04-09 14:50 | 1554835811.5332 | 930 | 0.20576 | 0.00334 | 7.6e-05 | 101.032 | 101.4974 | 111.3347 | 20.5569 | 0.003303 | 0.003032 | 223.9755 | 0.26058 | 8.4006 | 0.002873 | 0.0041 | 0.00406 | 5.5e-05 | 0.000916 | 0.30215 | 1.3459 | 2.5474 | 98.51 | 20.03 | 49.761 | 100.6186 | 90.2578 | 94.0741 | 0.011845 | 0.96139 | 0.02219 | 23.081 | -0.041106 | 0.15502 | 0.018985 | 1.2554 | 1.2554 | 0.004976
2019-04-09 14:50 | 1554835812.534 | 931 | 0.20597 | 0.003325 | 7.6e-05 | 100.9764 | 101.5263 | 111.4868 | 20.5571 | 0.003311 | 0.003022 | 223.9308 | 0.26001 | 8.4023 | 0.002907 | 0.003895 | 0.003769 | 4.9e-05 | 0.000916 | 0.2995 | 1.3859 | 2.5477 | 98.51 | 20.03 | 49.766 | 100.8244 | 91.0625 | 93.8513 | 0.011846 | 0.96197 | 0.022212 | 23.0898 | -0.048585 | 0.15468 | 0.017761 | 1.1755 | 1.1755 | 0.005303
2019-04-09 14:50 | 1554835813.5339 | 932 | 0.20599 | 0.0033 | 7.6e-05 | 100.999 | 101.5256 | 111.499 | 20.5569 | 0.003403 | 0.003116 | 223.9038 | 0.26031 | 8.4032 | 0.002832 | 0.004324 | 0.004262 | 4.9e-05 | 0.000921 | 0.28479 | 1.3707 | 2.5478 | 98.51 | 20.03 | 49.769 | 103.0426 | 89.2719 | 96.1035 | 0.011846 | 0.95611 | 0.022145 | 23.1618 | -0.04652 | 0.15486 | 0.017633 | 1.1635 | 1.1635 | 0.005364
2019-04-09 14:50 | 1554835814.533 | 933 | 0.20601 | 0.003279 | 7.6e-05 | 101.0493 | 101.5201 | 111.5317 | 20.5567 | 0.003381 | 0.003025 | 223.8472 | 0.25964 | 8.402 | 0.002875 | 0.004651 | 0.004665 | 5.2e-05 | 0.00092 | 0.28584 | 1.3722 | 2.5477 | 98.51 | 20.03 | 49.771 | 102.5162 | 90.296 | 93.9058 | 0.011847 | 0.96183 | 0.022215 | 23.0962 | -0.041579 | 0.15446 | 0.017546 | 1.1614 | 1.1614 | 0.00536
2019-04-09 14:50 | 1554835815.5333 | 934 | 0.20598 | 0.003263 | 7.6e-05 | 100.9469 | 101.3766 | 111.799 | 20.557 | 0.003347 | 0.003029 | 223.8113 | 0.25996 | 8.4026 | 0.002882 | 0.004275 | 0.004203 | 4.9e-05 | 0.000929 | 0.3027 | 1.3531 | 2.5475 | 98.51 | 20.03 | 49.777 | 101.7019 | 90.4634 | 94.0219 | 0.011848 | 0.96153 | 0.022238 | 23.1275 | -0.037998 | 0.15465 | 0.017747 | 1.176 | 1.176 | 0.0053
2019-04-09 14:50 | 1554835816.5331 | 935 | 0.20611 | 0.003238 | 7.6e-05 | 100.9773 | 101.2152 | 111.6495 | 20.5558 | 0.003335 | 0.003015 | 223.7718 | 0.2606 | 8.3996 | 0.002913 | 0.004775 | 0.004694 | 5.5e-05 | 0.000931 | 0.27151 | 1.3473 | 2.5475 | 98.51 | 20.03 | 49.779 | 101.3929 | 91.2221 | 93.6676 | 0.011849 | 0.96245 | 0.022234 | 23.1009 | -0.021055 | 0.15503 | 0.017013 | 1.127 | 1.127 | 0.005543
2019-04-09 14:50 | 1554835817.533 | 936 | 0.20608 | 0.003219 | 7.6e-05 | 100.9692 | 101.5862 | 111.6687 | 20.5561 | 0.003324 | 0.003062 | 223.7177 | 0.25992 | 8.4024 | 0.002943 | 0.004125 | 0.004066 | 5.2e-05 | 0.000932 | 0.31454 | 1.3327 | 2.5477 | 98.52 | 20.03 | 49.784 | 101.128 | 91.9316 | 94.8011 | 0.011849 | 0.95949 | 0.022201 | 23.1385 | -0.054492 | 0.15463 | 0.017215 | 1.1387 | 1.1387 | 0.005472
2019-04-09 14:50 | 1554835818.533 | 937 | 0.20614 | 0.003197 | 7.5e-05 | 100.884 | 101.4107 | 111.6615 | 20.558 | 0.003362 | 0.003066 | 223.6977 | 0.26063 | 8.4027 | 0.002896 | 0.004318 | 0.004206 | 5.3e-05 | 0.000938 | 0.28926 | 1.3659 | 2.5478 | 98.52 | 20.03 | 49.786 | 102.057 | 90.7976 | 94.894 | 0.011849 | 0.95925 | 0.022198 | 23.1407 | -0.046573 | 0.15505 | 0.016892 | 1.1172 | 1.1172 | 0.005593
2019-04-09 14:50 | 1554835819.533 | 938 | 0.20613 | 0.003176 | 7.4e-05 | 100.8477 | 101.4141 | 111.7943 | 20.5562 | 0.003326 | 0.003003 | 223.6288 | 0.25996 | 8.3979 | 0.00286 | 0.004652 | 0.004557 | 4.9e-05 | 0.000937 | 0.28152 | 1.3245 | 2.5476 | 98.52 | 20.03 | 49.787 | 101.1807 | 89.9268 | 93.392 | 0.01185 | 0.96318 | 0.022256 | 23.1071 | -0.050098 | 0.15465 | 0.016954 | 1.1244 | 1.1244 | 0.005543
2019-04-09 14:50 | 1554835820.532 | 939 | 0.20615 | 0.003157 | 7.4e-05 | 100.9001 | 101.4548 | 111.7021 | 20.5569 | 0.003266 | 0.002998 | 223.605 | 0.25927 | 8.4009 | 0.002892 | 0.004028 | 0.003921 | 5.7e-05 | 0.000939 | 0.30198 | 1.3388 | 2.5476 | 98.52 | 20.03 | 49.788 | 99.7366 | 90.7057 | 93.2761 | 0.01185 | 0.96348 | 0.022251 | 23.094 | -0.049032 | 0.15424 | 0.016857 | 1.1176 | 1.1176 | 0.005562
2019-04-09 14:50 | 1554835821.5331 | 940 | 0.20614 | 0.003134 | 7.3e-05 | 100.8066 | 101.305 | 112.0584 | 20.5573 | 0.003265 | 0.003052 | 223.5421 | 0.25955 | 8.4034 | 0.002931 | 0.004627 | 0.004555 | 4.9e-05 | 0.000941 | 0.29878 | 1.3552 | 2.5476 | 98.52 | 20.03 | 49.788 | 99.7104 | 91.6566 | 94.5676 | 0.01185 | 0.9601 | 0.022247 | 23.1715 | -0.044117 | 0.1544 | 0.016937 | 1.1228 | 1.1228 | 0.005542
2019-04-09 14:50 | 1554835822.5323 | 941 | 0.20627 | 0.003106 | 7.3e-05 | 100.7524 | 101.2059 | 112.0825 | 20.5572 | 0.003348 | 0.003018 | 223.491 | 0.25991 | 8.4019 | 0.002807 | 0.004383 | 0.004406 | 4.9e-05 | 0.000946 | 0.30397 | 1.3429 | 2.5476 | 98.52 | 20.03 | 49.787 | 101.7101 | 88.6526 | 93.7466 | 0.01185 | 0.96225 | 0.022274 | 23.1481 | -0.040171 | 0.15462 | 0.01623 | 1.0772 | 1.0772 | 0.005785
2019-04-09 14:50 | 1554835823.5332 | 942 | 0.2063 | 0.003084 | 7.3e-05 | 100.7587 | 101.2861 | 112.0567 | 20.5579 | 0.003314 | 0.003023 | 223.4676 | 0.26022 | 8.401 | 0.002857 | 0.004067 | 0.003947 | 5.3e-05 | 0.000945 | 0.29839 | 1.3632 | 2.5477 | 98.52 | 20.03 | 49.787 | 100.898 | 89.8643 | 93.8788 | 0.01185 | 0.9619 | 0.022268 | 23.1496 | -0.0467 | 0.1548 | 0.016079 | 1.0668 | 1.0668 | 0.005848
2019-04-09 14:50 | 1554835824.5322 | 943 | 0.20629 | 0.003064 | 7.3e-05 | 100.7348 | 101.1698 | 112.4741 | 20.5572 | 0.003297 | 0.003008 | 223.409 | 0.26044 | 8.4037 | 0.002904 | 0.005183 | 0.005068 | 5e-05 | 0.000945 | 0.28296 | 1.3197 | 2.5476 | 98.52 | 20.03 | 49.783 | 100.4791 | 90.9929 | 93.5142 | 0.011849 | 0.96286 | 0.02232 | 23.1811 | -0.038537 | 0.15494 | 0.016122 | 1.0722 | 1.0722 | 0.005823
2019-04-09 14:50 | 1554835825.5333 | 944 | 0.20628 | 0.003033 | 7.3e-05 | 100.6675 | 101.1929 | 112.4177 | 20.5566 | 0.003312 | 0.003007 | 223.386 | 0.26031 | 8.4013 | 0.002837 | 0.004434 | 0.004448 | 5.3e-05 | 0.000945 | 0.30807 | 1.3491 | 2.5475 | 98.52 | 20.03 | 49.78 | 100.8553 | 89.3949 | 93.4822 | 0.011848 | 0.96294 | 0.022316 | 23.1743 | -0.046553 | 0.15486 | 0.016244 | 1.0801 | 1.0801 | 0.005778
2019-04-09 14:50 | 1554835826.5333 | 945 | 0.20634 | 0.003009 | 7.3e-05 | 100.6723 | 101.2603 | 112.5928 | 20.5574 | 0.003258 | 0.003018 | 223.3159 | 0.26007 | 8.4041 | 0.00282 | 0.004181 | 0.004107 | 5.5e-05 | 0.000951 | 0.33395 | 1.3309 | 2.5475 | 98.51 | 20.03 | 49.775 | 99.541 | 88.9825 | 93.7459 | 0.011848 | 0.96225 | 0.022325 | 23.2007 | -0.05209 | 0.15472 | 0.015934 | 1.0599 | 1.0599 | 0.005882
2019-04-09 14:50 | 1554835827.5323 | 946 | 0.20631 | 0.002989 | 7.3e-05 | 100.6515 | 101.2266 | 112.9844 | 20.5581 | 0.003306 | 0.002974 | 223.2935 | 0.25989 | 8.4034 | 0.002895 | 0.003973 | 0.003879 | 5.5e-05 | 0.000948 | 0.32781 | 1.3293 | 2.5475 | 98.51 | 20.03 | 49.772 | 100.6974 | 90.7858 | 92.681 | 0.011847 | 0.96505 | 0.022396 | 23.2073 | -0.050963 | 0.15461 | 0.01614 | 1.0771 | 1.0771 | 0.005785
2019-04-09 14:50 | 1554835828.5333 | 947 | 0.20637 | 0.002968 | 7.3e-05 | 100.5678 | 101.2658 | 113.2116 | 20.5577 | 0.003277 | 0.003058 | 223.2426 | 0.26014 | 8.4029 | 0.002863 | 0.003884 | 0.003783 | 5.5e-05 | 0.000946 | 0.30879 | 1.3644 | 2.5476 | 98.51 | 20.03 | 49.768 | 99.9947 | 90.0212 | 94.7169 | 0.011846 | 0.95971 | 0.022357 | 23.2951 | -0.061868 | 0.15476 | 0.01577 | 1.0505 | 1.0505 | 0.005937
2019-04-09 14:50 | 1554835829.5338 | 948 | 0.20642 | 0.002948 | 7.2e-05 | 100.5669 | 101.2593 | 113.04 | 20.5582 | 0.003277 | 0.003023 | 223.2079 | 0.26059 | 8.4027 | 0.002894 | 0.004687 | 0.004602 | 5.6e-05 | 0.00095 | 0.28136 | 1.3637 | 2.5476 | 98.51 | 20.03 | 49.759 | 100.014 | 90.7546 | 93.868 | 0.011844 | 0.96193 | 0.022365 | 23.2506 | -0.06137 | 0.15503 | 0.015528 | 1.0348 | 1.0348 | 0.006038
2019-04-09 14:50 | 1554835830.5335 | 949 | 0.20642 | 0.002925 | 7.2e-05 | 100.5386 | 100.8711 | 113.4329 | 20.5568 | 0.003271 | 0.002979 | 223.1798 | 0.26003 | 8.3994 | 0.0028 | 0.004626 | 0.004483 | 5.5e-05 | 0.00095 | 0.29894 | 1.3469 | 2.5475 | 98.51 | 20.03 | 49.755 | 99.8581 | 88.4881 | 92.7996 | 0.011843 | 0.96474 | 0.022437 | 23.2571 | -0.029529 | 0.15469 | 0.015561 | 1.0403 | 1.0403 | 0.005993
2019-04-09 14:50 | 1554835831.5324 | 950 | 0.20649 | 0.002915 | 7.1e-05 | 100.4266 | 101.0166 | 113.2881 | 20.5579 | 0.003247 | 0.002976 | 223.1244 | 0.25968 | 8.402 | 0.002892 | 0.004644 | 0.004568 | 4.9e-05 | 0.000945 | 0.27831 | 1.3496 | 2.5475 | 98.51 | 20.03 | 49.752 | 99.2833 | 90.7168 | 92.7358 | 0.011842 | 0.96491 | 0.022425 | 23.2402 | -0.052396 | 0.15448 | 0.015138 | 1.0114 | 1.0114 | 0.006155
2019-04-09 14:51 | 1554835832.5325 | 951 | 0.20647 | 0.002895 | 7.2e-05 | 100.3736 | 100.9275 | 113.3273 | 20.5568 | 0.003269 | 0.002935 | 223.1045 | 0.25996 | 8.4016 | 0.002832 | 0.004897 | 0.004871 | 5e-05 | 0.00095 | 0.29286 | 1.3071 | 2.5474 | 98.51 | 20.03 | 49.742 | 99.7984 | 89.2576 | 91.7556 | 0.01184 | 0.9675 | 0.022459 | 23.213 | -0.049222 | 0.15465 | 0.015257 | 1.0209 | 1.0209 | 0.006105
2019-04-09 14:51 | 1554835833.5315 | 952 | 0.20652 | 0.002874 | 7.1e-05 | 100.3001 | 100.8587 | 113.4716 | 20.5579 | 0.003233 | 0.003016 | 223.0455 | 0.26047 | 8.402 | 0.002893 | 0.004112 | 0.004043 | 5.3e-05 | 0.000944 | 0.28495 | 1.3343 | 2.5476 | 98.51 | 20.03 | 49.738 | 98.9321 | 90.745 | 93.702 | 0.011839 | 0.96237 | 0.022413 | 23.2897 | -0.049677 | 0.15495 | 0.015009 | 1.0023 | 1.0023 | 0.006231
2019-04-09 14:51 | 1554835834.5319 | 953 | 0.20651 | 0.002856 | 7.2e-05 | 100.3272 | 100.7265 | 113.2156 | 20.5577 | 0.003252 | 0.002959 | 223.0149 | 0.26002 | 8.4018 | 0.00286 | 0.004488 | 0.004378 | 5.1e-05 | 0.000941 | 0.31116 | 1.302 | 2.5475 | 98.52 | 20.03 | 49.73 | 99.3931 | 89.9366 | 92.3241 | 0.011836 | 0.96599 | 0.02243 | 23.2197 | -0.035533 | 0.15469 | 0.015064 | 1.0067 | 1.0067 | 0.006192
2019-04-09 14:51 | 1554835835.5328 | 954 | 0.20658 | 0.002835 | 7.1e-05 | 100.2143 | 100.6364 | 113.1243 | 20.5582 | 0.003256 | 0.002947 | 222.9812 | 0.25992 | 8.4025 | 0.002822 | 0.003971 | 0.004004 | 4.8e-05 | 0.000941 | 0.29966 | 1.3225 | 2.5476 | 98.52 | 20.03 | 49.727 | 99.4877 | 89.031 | 92.026 | 0.011835 | 0.96678 | 0.02243 | 23.2008 | -0.03759 | 0.15463 | 0.014696 | 0.9821 | 0.9821 | 0.006345
2019-04-09 14:51 | 1554835836.5317 | 955 | 0.20658 | 0.002809 | 7.1e-05 | 100.1396 | 100.7251 | 113.0264 | 20.5586 | 0.003177 | 0.002958 | 222.9467 | 0.26055 | 8.4057 | 0.002875 | 0.004448 | 0.00434 | 5.1e-05 | 0.000938 | 0.27599 | 1.2927 | 2.5476 | 98.52 | 20.03 | 49.723 | 97.5811 | 90.2978 | 92.2968 | 0.011834 | 0.96606 | 0.022412 | 23.1994 | -0.052152 | 0.155 | 0.01469 | 0.98091 | 0.98091 | 0.006368
2019-04-09 14:51 | 1554835837.533 | 956 | 0.20662 | 0.002795 | 7.2e-05 | 100.1 | 100.6009 | 112.9459 | 20.5582 | 0.003253 | 0.002947 | 222.9471 | 0.26036 | 8.4033 | 0.002815 | 0.004539 | 0.004443 | 5e-05 | 0.000936 | 0.28064 | 1.3577 | 2.5476 | 98.52 | 20.03 | 49.719 | 99.432 | 88.866 | 92.0306 | 0.011833 | 0.96677 | 0.022412 | 23.1827 | -0.044651 | 0.15489 | 0.01449 | 0.96751 | 0.96751 | 0.006452
2019-04-09 14:51 | 1554835838.5321 | 957 | 0.20668 | 0.002769 | 7.1e-05 | 100.136 | 100.4443 | 112.909 | 20.5591 | 0.003317 | 0.002954 | 222.8926 | 0.25964 | 8.4033 | 0.002874 | 0.004205 | 0.00414 | 4.9e-05 | 0.000927 | 0.30431 | 1.3333 | 2.5477 | 98.52 | 20.03 | 49.717 | 100.9761 | 90.2781 | 92.2129 | 0.011833 | 0.96629 | 0.022403 | 23.1847 | -0.027499 | 0.15446 | 0.014164 | 0.94528 | 0.94528 | 0.006585
2019-04-09 14:51 | 1554835839.5327 | 958 | 0.20665 | 0.002749 | 7.2e-05 | 99.9898 | 100.3177 | 113.067 | 20.5586 | 0.003269 | 0.002951 | 222.8948 | 0.26063 | 8.4012 | 0.002825 | 0.003991 | 0.003905 | 5.3e-05 | 0.000922 | 0.28174 | 1.3289 | 2.5474 | 98.52 | 20.03 | 49.712 | 99.8071 | 89.0898 | 92.142 | 0.011832 | 0.96647 | 0.022421 | 23.1986 | -0.029284 | 0.15505 | 0.014372 | 0.95998 | 0.95998 | 0.006509
2019-04-09 14:51 | 1554835840.5317 | 959 | 0.20664 | 0.002728 | 7.2e-05 | 99.9496 | 100.2987 | 113.3801 | 20.5582 | 0.003218 | 0.002931 | 222.8549 | 0.2602 | 8.4016 | 0.002888 | 0.004326 | 0.004231 | 5e-05 | 0.000919 | 0.29574 | 1.3122 | 2.5475 | 98.51 | 20.03 | 49.711 | 98.5771 | 90.6178 | 91.656 | 0.011833 | 0.96776 | 0.022467 | 23.2153 | -0.031193 | 0.15479 | 0.014458 | 0.96768 | 0.96768 | 0.006447
2019-04-09 14:51 | 1554835841.5317 | 960 | 0.20671 | 0.002708 | 7.3e-05 | 99.8864 | 100.34 | 113.708 | 20.5591 | 0.003174 | 0.002942 | 222.8359 | 0.26001 | 8.4026 | 0.002824 | 0.004177 | 0.004126 | 5.6e-05 | 0.000917 | 0.29336 | 1.329 | 2.5475 | 98.51 | 20.03 | 49.71 | 97.504 | 89.0725 | 91.9119 | 0.011832 | 0.96708 | 0.022491 | 23.257 | -0.040526 | 0.15468 | 0.014082 | 0.94345 | 0.94345 | 0.006607
2019-04-09 14:51 | 1554835842.5318 | 961 | 0.20681 | 0.002689 | 7.3e-05 | 99.7926 | 100.095 | 113.8521 | 20.5587 | 0.003217 | 0.002989 | 222.8223 | 0.25973 | 8.4007 | 0.002938 | 0.004276 | 0.004217 | 5.5e-05 | 0.000916 | 0.28528 | 1.3481 | 2.5476 | 98.51 | 20.03 | 49.707 | 98.5534 | 91.8112 | 93.0548 | 0.011832 | 0.96406 | 0.022471 | 23.3081 | -0.02706 | 0.15451 | 0.013495 | 0.90315 | 0.90315 | 0.006895
2019-04-09 14:51 | 1554835843.5319 | 962 | 0.20677 | 0.002667 | 7.4e-05 | 99.7792 | 100.3029 | 114.0528 | 20.5587 | 0.003236 | 0.00294 | 222.7783 | 0.25991 | 8.4024 | 0.002884 | 0.004691 | 0.004637 | 5.4e-05 | 0.000917 | 0.29966 | 1.3423 | 2.5474 | 98.51 | 20.03 | 49.709 | 99.0129 | 90.527 | 91.8645 | 0.011832 | 0.96721 | 0.022527 | 23.2907 | -0.046818 | 0.15462 | 0.013736 | 0.92155 | 0.92155 | 0.006762
2019-04-09 14:51 | 1554835844.5333 | 963 | 0.20682 | 0.002653 | 7.5e-05 | 99.7626 | 100.0227 | 114.1183 | 20.5591 | 0.003212 | 0.00293 | 222.7741 | 0.25962 | 8.4025 | 0.002961 | 0.004412 | 0.004323 | 5.4e-05 | 0.000916 | 0.30862 | 1.3182 | 2.5477 | 98.51 | 20.03 | 49.709 | 98.4416 | 92.3684 | 91.6162 | 0.011832 | 0.96787 | 0.022541 | 23.2895 | -0.023296 | 0.15445 | 0.013461 | 0.90358 | 0.90358 | 0.006889
2019-04-09 14:51 | 1554835845.5319 | 964 | 0.20677 | 0.002633 | 7.6e-05 | 99.7182 | 99.9171 | 114.5145 | 20.5591 | 0.003167 | 0.002961 | 222.7169 | 0.26058 | 8.4009 | 0.002954 | 0.004952 | 0.004928 | 5.5e-05 | 0.000913 | 0.2688 | 1.3317 | 2.5475 | 98.51 | 20.03 | 49.708 | 97.3466 | 92.2074 | 92.3865 | 0.011832 | 0.96583 | 0.022556 | 23.3545 | -0.017822 | 0.15502 | 0.013767 | 0.92478 | 0.92478 | 0.006755
2019-04-09 14:51 | 1554835846.533 | 965 | 0.20681 | 0.002604 | 7.7e-05 | 99.5735 | 100.0437 | 114.5432 | 20.5594 | 0.003224 | 0.002908 | 222.6922 | 0.26048 | 8.4007 | 0.002792 | 0.004379 | 0.004366 | 5.1e-05 | 0.000919 | 0.30182 | 1.3967 | 2.5476 | 98.51 | 20.03 | 49.707 | 98.7218 | 88.2932 | 91.088 | 0.011832 | 0.96927 | 0.022599 | 23.3159 | -0.042143 | 0.15496 | 0.013575 | 0.91347 | 0.91347 | 0.006836
2019-04-09 14:51 | 1554835847.5309 | 966 | 0.20687 | 0.002588 | 7.7e-05 | 99.5467 | 99.8789 | 114.6176 | 20.5594 | 0.003187 | 0.002902 | 222.656 | 0.25963 | 8.4003 | 0.002886 | 0.004452 | 0.004358 | 5.2e-05 | 0.000923 | 0.29623 | 1.3622 | 2.5476 | 98.51 | 20.03 | 49.706 | 97.8367 | 90.5558 | 90.956 | 0.011832 | 0.96962 | 0.022611 | 23.3192 | -0.0298 | 0.15446 | 0.013223 | 0.89018 | 0.89018 | 0.006992
2019-04-09 14:51 | 1554835848.5321 | 967 | 0.20683 | 0.002567 | 7.8e-05 | 99.4847 | 99.7787 | 114.6734 | 20.5599 | 0.003221 | 0.002902 | 222.6291 | 0.26031 | 8.4006 | 0.002909 | 0.004449 | 0.004248 | 5.1e-05 | 0.000923 | 0.29336 | 1.3908 | 2.5477 | 98.51 | 20.03 | 49.705 | 98.6626 | 91.1292 | 90.9627 | 0.011831 | 0.96961 | 0.022616 | 23.3251 | -0.026394 | 0.15486 | 0.013501 | 0.90911 | 0.90911 | 0.006865
2019-04-09 14:51 | 1554835849.531 | 968 | 0.20689 | 0.002547 | 7.9e-05 | 99.4329 | 99.6583 | 114.591 | 20.5601 | 0.003239 | 0.002909 | 222.599 | 0.26103 | 8.4023 | 0.002909 | 0.004862 | 0.004898 | 5.3e-05 | 0.000928 | 0.31293 | 1.3565 | 2.5476 | 98.51 | 20.03 | 49.703 | 99.0777 | 91.1284 | 91.1145 | 0.011831 | 0.9692 | 0.022603 | 23.3216 | -0.02025 | 0.15529 | 0.013174 | 0.88646 | 0.88646 | 0.00706
2019-04-09 14:51 | 1554835850.5321 | 969 | 0.20688 | 0.00253 | 7.9e-05 | 99.3665 | 99.5726 | 114.5028 | 20.5603 | 0.003259 | 0.002911 | 222.5429 | 0.2602 | 8.4041 | 0.002878 | 0.004131 | 0.004086 | 5.1e-05 | 0.000927 | 0.29031 | 1.3362 | 2.5477 | 98.51 | 20.03 | 49.701 | 99.5666 | 90.3809 | 91.1663 | 0.01183 | 0.96906 | 0.022593 | 23.3143 | -0.018531 | 0.15479 | 0.013261 | 0.89195 | 0.89195 | 0.006994
2019-04-09 14:51 | 1554835851.5315 | 970 | 0.20693 | 0.002513 | 7.9e-05 | 99.2885 | 99.6662 | 114.2432 | 20.5599 | 0.003207 | 0.002933 | 222.5275 | 0.25934 | 8.4018 | 0.002874 | 0.004289 | 0.004275 | 5e-05 | 0.000931 | 0.31581 | 1.3716 | 2.5476 | 98.51 | 20.03 | 49.7 | 98.3083 | 90.2719 | 91.6905 | 0.01183 | 0.96767 | 0.022551 | 23.3046 | -0.033962 | 0.15428 | 0.012998 | 0.87253 | 0.87253 | 0.007126
2019-04-09 14:51 | 1554835852.5325 | 971 | 0.20698 | 0.002501 | 8e-05 | 99.1697 | 99.5757 | 114.085 | 20.5594 | 0.003098 | 0.002834 | 222.4849 | 0.26044 | 8.4001 | 0.002914 | 0.004184 | 0.004083 | 4.9e-05 | 0.00093 | 0.28816 | 1.3385 | 2.5476 | 98.51 | 20.03 | 49.7 | 95.6869 | 91.2389 | 89.3165 | 0.01183 | 0.97401 | 0.022609 | 23.2125 | -0.036546 | 0.15493 | 0.012664 | 0.85221 | 0.85221 | 0.007327
2019-04-09 14:51 | 1554835853.5321 | 972 | 0.20697 | 0.002485 | 8.1e-05 | 99.1494 | 99.4618 | 114.2436 | 20.5598 | 0.003175 | 0.002907 | 222.4353 | 0.26021 | 8.3991 | 0.002868 | 0.004354 | 0.004278 | 4.9e-05 | 0.000929 | 0.29983 | 1.3361 | 2.5476 | 98.51 | 20.04 | 49.7 | 97.5284 | 90.1357 | 91.0682 | 0.011837 | 0.96932 | 0.02257 | 23.2847 | -0.028134 | 0.1548 | 0.012747 | 0.85625 | 0.85625 | 0.007286
2019-04-09 14:51 | 1554835854.5316 | 973 | 0.20701 | 0.002468 | 8.1e-05 | 99.0485 | 99.3072 | 114.1009 | 20.5595 | 0.003112 | 0.002896 | 222.4125 | 0.25978 | 8.4017 | 0.002935 | 0.004727 | 0.004613 | 5.2e-05 | 0.000928 | 0.30397 | 1.3192 | 2.5476 | 98.51 | 20.03 | 49.701 | 96.008 | 91.752 | 90.8145 | 0.01183 | 0.97 | 0.022564 | 23.2621 | -0.023338 | 0.15454 | 0.012532 | 0.84153 | 0.84153 | 0.007401
2019-04-09 14:51 | 1554835855.5312 | 974 | 0.20705 | 0.002449 | 8.2e-05 | 98.9681 | 99.2539 | 113.762 | 20.5595 | 0.003182 | 0.002867 | 222.3693 | 0.25943 | 8.4008 | 0.002937 | 0.004002 | 0.003988 | 5.1e-05 | 0.000921 | 0.31061 | 1.3395 | 2.5476 | 98.51 | 20.03 | 49.7 | 97.7035 | 91.8043 | 90.1183 | 0.01183 | 0.97186 | 0.022552 | 23.2053 | -0.025796 | 0.15433 | 0.012354 | 0.82907 | 0.82907 | 0.007502
2019-04-09 14:51 | 1554835856.5323 | 975 | 0.20707 | 0.002434 | 8.3e-05 | 98.9713 | 99.2353 | 113.9331 | 20.5602 | 0.003128 | 0.002963 | 222.3485 | 0.25973 | 8.4029 | 0.0029 | 0.00389 | 0.003817 | 5.2e-05 | 0.000919 | 0.3027 | 1.3502 | 2.5476 | 98.51 | 20.04 | 49.701 | 96.4097 | 90.913 | 92.4259 | 0.011838 | 0.96572 | 0.022498 | 23.2964 | -0.023825 | 0.15452 | 0.012218 | 0.81782 | 0.81782 | 0.007614
2019-04-09 14:51 | 1554835857.5316 | 976 | 0.20709 | 0.002413 | 8.4e-05 | 98.7694 | 99.2009 | 113.6019 | 20.5605 | 0.00319 | 0.002839 | 222.2946 | 0.26035 | 8.401 | 0.002848 | 0.004153 | 0.004047 | 5.1e-05 | 0.000915 | 0.26792 | 1.3387 | 2.5476 | 98.51 | 20.04 | 49.702 | 97.9139 | 89.6549 | 89.4355 | 0.011838 | 0.97369 | 0.022558 | 23.1671 | -0.038988 | 0.15488 | 0.012122 | 0.81347 | 0.81347 | 0.007673
2019-04-09 14:51 | 1554835858.5314 | 977 | 0.20713 | 0.002395 | 8.5e-05 | 98.6661 | 98.8887 | 113.6355 | 20.5603 | 0.003134 | 0.002899 | 222.2631 | 0.26035 | 8.4014 | 0.002982 | 0.004076 | 0.004023 | 4.7e-05 | 0.000913 | 0.28838 | 1.3426 | 2.5475 | 98.51 | 20.04 | 49.701 | 96.5588 | 92.893 | 90.8928 | 0.011838 | 0.96979 | 0.022516 | 23.2171 | -0.020162 | 0.15488 | 0.01193 | 0.79897 | 0.79897 | 0.007812
2019-04-09 14:51 | 1554835859.5313 | 978 | 0.20714 | 0.002372 | 8.6e-05 | 98.6378 | 98.9431 | 114.0124 | 20.5608 | 0.003101 | 0.002867 | 222.2335 | 0.26063 | 8.4007 | 0.002846 | 0.004636 | 0.004582 | 5.1e-05 | 0.000915 | 0.28816 | 1.3692 | 2.5476 | 98.51 | 20.04 | 49.701 | 95.7423 | 89.6099 | 90.1217 | 0.011838 | 0.97185 | 0.022577 | 23.2309 | -0.027644 | 0.15505 | 0.011892 | 0.79851 | 0.79851 | 0.007825
2019-04-09 14:51 | 1554835860.5327 | 979 | 0.20719 | 0.002361 | 8.8e-05 | 98.4886 | 98.6227 | 113.9275 | 20.5605 | 0.003168 | 0.002865 | 222.2004 | 0.25906 | 8.4013 | 0.002788 | 0.005351 | 0.00529 | 5.1e-05 | 0.000923 | 0.30901 | 1.383 | 2.5474 | 98.51 | 20.04 | 49.7 | 97.3681 | 88.2097 | 90.0573 | 0.011837 | 0.97202 | 0.022571 | 23.2202 | -0.012171 | 0.15411 | 0.01161 | 0.77919 | 0.77919 | 0.007971
2019-04-09 14:51 | 1554835861.5329 | 980 | 0.20721 | 0.002337 | 8.9e-05 | 98.4704 | 98.6564 | 113.842 | 20.5604 | 0.003116 | 0.002908 | 222.1816 | 0.26057 | 8.401 | 0.002863 | 0.004098 | 0.004048 | 5.1e-05 | 0.000914 | 0.28854 | 1.3572 | 2.5476 | 98.51 | 20.04 | 49.699 | 96.1226 | 90.0096 | 91.0891 | 0.011837 | 0.96927 | 0.02253 | 23.2444 | -0.016879 | 0.15502 | 0.011525 | 0.77201 | 0.77201 | 0.008092
2019-04-09 14:51 | 1554835862.5325 | 981 | 0.20723 | 0.00232 | 9e-05 | 98.4043 | 98.7617 | 114.3636 | 20.5619 | 0.003099 | 0.002861 | 222.1537 | 0.26105 | 8.4017 | 0.00291 | 0.00468 | 0.004601 | 5.3e-05 | 0.000917 | 0.2912 | 1.3391 | 2.5475 | 98.51 | 20.04 | 49.695 | 95.7144 | 91.1444 | 89.9775 | 0.011836 | 0.97223 | 0.022616 | 23.262 | -0.032429 | 0.1553 | 0.011425 | 0.76816 | 0.76816 | 0.008147
2019-04-09 14:51 | 1554835863.6006 | 982 | 0.20722 | 0.002296 | 9.1e-05 | 98.2239 | 98.3247 | 114.0911 | 20.5598 | 0.003148 | 0.002843 | 222.1494 | 0.25981 | 8.4026 | 0.00287 | 0.004711 | 0.004639 | 5.1e-05 | 0.000922 | 0.30845 | 1.3321 | 2.5475 | 98.51 | 20.04 | 49.692 | 96.8818 | 90.1732 | 89.523 | 0.011836 | 0.97345 | 0.022603 | 23.2198 | -0.009175 | 0.15456 | 0.011487 | 0.7718 | 0.7718 | 0.00807
2019-04-09 14:51 | 1554835864.5491 | 983 | 0.20728 | 0.002277 | 9.3e-05 | 98.1884 | 98.3911 | 114.2847 | 20.5591 | 0.003126 | 0.002854 | 222.0953 | 0.26026 | 8.4013 | 0.002865 | 0.004506 | 0.004442 | 5.2e-05 | 0.000916 | 0.27743 | 1.3249 | 2.5474 | 98.51 | 20.04 | 49.689 | 96.35 | 90.0539 | 89.8015 | 0.011835 | 0.97271 | 0.022614 | 23.2484 | -0.018443 | 0.15483 | 0.011152 | 0.74948 | 0.74948 | 0.008325
2019-04-09 14:51 | 1554835865.5492 | 984 | 0.20727 | 0.002259 | 9.5e-05 | 98.1268 | 98.2879 | 114.3643 | 20.5612 | 0.003184 | 0.002898 | 222.0825 | 0.25945 | 8.3993 | 0.002908 | 0.003953 | 0.00388 | 5.2e-05 | 0.000915 | 0.27671 | 1.3288 | 2.5475 | 98.51 | 20.04 | 49.682 | 97.7466 | 91.0953 | 90.8481 | 0.011833 | 0.96991 | 0.022589 | 23.29 | -0.014669 | 0.15435 | 0.0112 | 0.75175 | 0.75175 | 0.008274
2019-04-09 14:51 | 1554835866.5493 | 985 | 0.20732 | 0.002237 | 9.6e-05 | 98.0276 | 98.249 | 114.6201 | 20.5606 | 0.003145 | 0.002849 | 222.0764 | 0.26024 | 8.4024 | 0.002847 | 0.00456 | 0.004484 | 4.7e-05 | 0.000921 | 0.29734 | 1.3465 | 2.5476 | 98.51 | 20.04 | 49.68 | 96.8114 | 89.6184 | 89.6661 | 0.011833 | 0.97307 | 0.022651 | 23.2781 | -0.020173 | 0.15482 | 0.01096 | 0.73745 | 0.73745 | 0.00846
2019-04-09 14:51 | 1554835867.5507 | 986 | 0.20735 | 0.002221 | 9.8e-05 | 97.9827 | 98.112 | 114.9652 | 20.5608 | 0.003149 | 0.002836 | 222.0241 | 0.26039 | 8.4017 | 0.00288 | 0.0046 | 0.004581 | 4.9e-05 | 0.000923 | 0.31238 | 1.3253 | 2.5475 | 98.51 | 20.04 | 49.679 | 96.9163 | 90.4148 | 89.3568 | 0.011832 | 0.9739 | 0.022695 | 23.3032 | -0.011798 | 0.1549 | 0.010795 | 0.72761 | 0.72761 | 0.00858
2019-04-09 14:51 | 1554835868.5497 | 987 | 0.20736 | 0.002201 | 0.0001 | 97.8963 | 98.1433 | 115.076 | 20.5606 | 0.003157 | 0.002823 | 222.0246 | 0.26016 | 8.405 | 0.002928 | 0.004619 | 0.004534 | 5e-05 | 0.000922 | 0.31078 | 1.3285 | 2.5476 | 98.51 | 20.04 | 49.672 | 97.1028 | 91.5894 | 89.0409 | 0.011831 | 0.97475 | 0.022716 | 23.3043 | -0.02254 | 0.15477 | 0.01077 | 0.72645 | 0.72645 | 0.008586
2019-04-09 14:51 | 1554835869.5493 | 988 | 0.20738 | 0.00218 | 0.000101 | 97.8841 | 98.0407 | 114.7791 | 20.5608 | 0.003136 | 0.002819 | 221.9911 | 0.26046 | 8.4011 | 0.002898 | 0.00432 | 0.004175 | 4.9e-05 | 0.000922 | 0.28031 | 1.3297 | 2.5475 | 98.51 | 20.05 | 49.671 | 96.5865 | 90.8495 | 88.9663 | 0.011838 | 0.97495 | 0.022689 | 23.2718 | -0.014304 | 0.15495 | 0.010659 | 0.71801 | 0.71801 | 0.008697
2019-04-09 14:51 | 1554835870.5493 | 989 | 0.20735 | 0.002162 | 0.000103 | 97.7078 | 97.7637 | 114.8168 | 20.5613 | 0.003088 | 0.002804 | 221.9774 | 0.26083 | 8.4022 | 0.002908 | 0.004962 | 0.004886 | 5.1e-05 | 0.000924 | 0.27865 | 1.3442 | 2.5476 | 98.51 | 20.04 | 49.668 | 95.4358 | 91.1055 | 88.5926 | 0.01183 | 0.97596 | 0.022704 | 23.2636 | -0.005116 | 0.15517 | 0.01084 | 0.73068 | 0.73068 | 0.008558
2019-04-09 14:51 | 1554835871.5499 | 990 | 0.20739 | 0.00215 | 0.000104 | 97.6071 | 97.6763 | 115.0103 | 20.5612 | 0.003131 | 0.002872 | 221.9492 | 0.26111 | 8.4009 | 0.002944 | 0.004658 | 0.004561 | 5.2e-05 | 0.000926 | 0.31293 | 1.3353 | 2.5475 | 98.51 | 20.04 | 49.667 | 96.4853 | 91.9638 | 90.2261 | 0.01183 | 0.97157 | 0.022672 | 23.3357 | -0.006334 | 0.15533 | 0.010656 | 0.71711 | 0.71711 | 0.008729
2019-04-09 14:51 | 1554835872.5495 | 991 | 0.2074 | 0.002138 | 0.000105 | 97.606 | 97.4467 | 114.6784 | 20.561 | 0.003159 | 0.002778 | 221.8832 | 0.26001 | 8.4001 | 0.002895 | 0.004411 | 0.004342 | 5.1e-05 | 0.000927 | 0.28639 | 1.2982 | 2.5474 | 98.51 | 20.05 | 49.668 | 97.1494 | 90.7957 | 87.9682 | 0.011837 | 0.97764 | 0.02271 | 23.2295 | 0.014602 | 0.15468 | 0.010579 | 0.71302 | 0.71302 | 0.008742
2019-04-09 14:51 | 1554835873.5504 | 992 | 0.20741 | 0.002125 | 0.000106 | 97.541 | 97.6629 | 114.5831 | 20.5609 | 0.002993 | 0.002854 | 221.8707 | 0.25973 | 8.4012 | 0.002934 | 0.004552 | 0.004555 | 4.9e-05 | 0.00093 | 0.33179 | 1.3144 | 2.5475 | 98.51 | 20.04 | 49.669 | 93.1438 | 91.736 | 89.8063 | 0.01183 | 0.97269 | 0.022643 | 23.2789 | -0.011165 | 0.15451 | 0.010545 | 0.70857 | 0.70857 | 0.008788
2019-04-09 14:51 | 1554835874.5495 | 993 | 0.20747 | 0.002105 | 0.000107 | 97.4821 | 97.5565 | 114.7342 | 20.5609 | 0.00308 | 0.002749 | 221.8297 | 0.26019 | 8.4046 | 0.002892 | 0.004209 | 0.004123 | 5.2e-05 | 0.000927 | 0.30326 | 1.3257 | 2.5475 | 98.51 | 20.04 | 49.67 | 95.2366 | 90.7065 | 87.277 | 0.01183 | 0.97952 | 0.022738 | 23.2129 | -0.006829 | 0.15479 | 0.010232 | 0.69022 | 0.69022 | 0.009038
2019-04-09 14:51 | 1554835875.5495 | 994 | 0.2075 | 0.002093 | 0.000108 | 97.4481 | 97.4502 | 115.2229 | 20.5604 | 0.003101 | 0.00277 | 221.8283 | 0.26061 | 8.4015 | 0.002864 | 0.004711 | 0.004648 | 5e-05 | 0.000928 | 0.28495 | 1.3331 | 2.5477 | 98.51 | 20.04 | 49.673 | 95.7443 | 90.0276 | 87.7834 | 0.011831 | 0.97815 | 0.02277 | 23.2786 | -0.000189 | 0.15504 | 0.010066 | 0.67983 | 0.67983 | 0.009191
2019-04-09 14:51 | 1554835876.5496 | 995 | 0.20754 | 0.002079 | 0.000109 | 97.3874 | 97.5335 | 114.9966 | 20.5621 | 0.003071 | 0.00283 | 221.785 | 0.25901 | 8.403 | 0.002854 | 0.004315 | 0.004191 | 4.7e-05 | 0.000933 | 0.29015 | 1.3119 | 2.5476 | 98.51 | 20.04 | 49.675 | 95.0396 | 89.8065 | 89.2276 | 0.011831 | 0.97425 | 0.022702 | 23.3022 | -0.013408 | 0.15408 | 0.00979 | 0.65903 | 0.65903 | 0.009422
2019-04-09 14:51 | 1554835877.5496 | 996 | 0.20757 | 0.002059 | 0.000111 | 97.296 | 97.4399 | 114.942 | 20.5611 | 0.003101 | 0.002782 | 221.7841 | 0.26033 | 8.4033 | 0.002923 | 0.004492 | 0.004448 | 5e-05 | 0.000935 | 0.28888 | 1.342 | 2.5475 | 98.51 | 20.04 | 49.676 | 95.7472 | 91.4548 | 88.0549 | 0.011832 | 0.97741 | 0.022734 | 23.259 | -0.01322 | 0.15487 | 0.009653 | 0.65048 | 0.65048 | 0.009595
2019-04-09 14:51 | 1554835878.5497 | 997 | 0.20756 | 0.002035 | 0.000114 | 97.2197 | 97.264 | 114.7687 | 20.5615 | 0.003025 | 0.002824 | 221.752 | 0.25955 | 8.4023 | 0.002938 | 0.004314 | 0.00433 | 5.1e-05 | 0.000935 | 0.28727 | 1.3722 | 2.5477 | 98.51 | 20.04 | 49.677 | 93.9088 | 91.8181 | 89.0669 | 0.011832 | 0.97468 | 0.022685 | 23.274 | -0.004073 | 0.1544 | 0.009734 | 0.65446 | 0.65446 | 0.009508
2019-04-09 14:51 | 1554835879.5503 | 998 | 0.20752 | 0.002013 | 0.000117 | 97.2102 | 97.3115 | 115.0783 | 20.562 | 0.003141 | 0.002832 | 221.697 | 0.26063 | 8.4012 | 0.002844 | 0.004474 | 0.004362 | 4.7e-05 | 0.000943 | 0.27936 | 1.3601 | 2.5476 | 98.51 | 20.04 | 49.679 | 96.7311 | 89.554 | 89.28 | 0.011832 | 0.97411 | 0.022709 | 23.3122 | -0.009316 | 0.15505 | 0.010016 | 0.67409 | 0.67409 | 0.00927
2019-04-09 14:51 | 1554835880.5497 | 999 | 0.20764 | 0.001989 | 0.00012 | 97.0979 | 97.148 | 115.2677 | 20.5622 | 0.003012 | 0.00278 | 221.7164 | 0.25949 | 8.4025 | 0.002878 | 0.004563 | 0.004528 | 4.9e-05 | 0.000937 | 0.26864 | 1.3418 | 2.5476 | 98.51 | 20.04 | 49.683 | 93.6176 | 90.3897 | 88.0263 | 0.011833 | 0.97749 | 0.022767 | 23.291 | -0.004613 | 0.15437 | 0.009327 | 0.62875 | 0.62875 | 0.009894
2019-04-09 14:51 | 1554835881.5497 | 1000 | 0.20751 | 0.001975 | 0.000123 | 97.0789 | 97.202 | 115.3133 | 20.5619 | 0.003104 | 0.002827 | 221.6655 | 0.2603 | 8.402 | 0.002909 | 0.004872 | 0.004761 | 4.6e-05 | 0.000939 | 0.28407 | 1.3483 | 2.5475 | 98.51 | 20.04 | 49.686 | 95.8241 | 91.1296 | 89.1501 | 0.011834 | 0.97446 | 0.022736 | 23.3318 | -0.011343 | 0.15485 | 0.010118 | 0.68147 | 0.68147 | 0.009157
2019-04-09 14:51 | 1554835882.5492 | 1001 | 0.20765 | 0.001953 | 0.000126 | 97.0172 | 97.0413 | 115.1342 | 20.5618 | 0.003124 | 0.002803 | 221.6959 | 0.26089 | 8.4017 | 0.002837 | 0.00435 | 0.004298 | 4.8e-05 | 0.000941 | 0.29374 | 1.3291 | 2.5476 | 98.51 | 20.04 | 49.688 | 96.3096 | 89.3915 | 88.5664 | 0.011835 | 0.97603 | 0.022736 | 23.2949 | -0.002222 | 0.1552 | 0.0093 | 0.62574 | 0.62574 | 0.009996
2019-04-09 14:51 | 1554835883.5492 | 1002 | 0.20761 | 0.001938 | 0.000128 | 96.9234 | 96.9759 | 115.3295 | 20.562 | 0.003016 | 0.002767 | 221.681 | 0.25983 | 8.4004 | 0.00283 | 0.004917 | 0.004865 | 5e-05 | 0.000943 | 0.27168 | 1.3526 | 2.5476 | 98.51 | 20.04 | 49.692 | 93.6925 | 89.2261 | 87.6983 | 0.011836 | 0.97838 | 0.022783 | 23.2867 | -0.004846 | 0.15457 | 0.009511 | 0.64129 | 0.64129 | 0.009714
2019-04-09 14:51 | 1554835884.5499 | 1003 | 0.2074 | 0.001926 | 0.00013 | 96.8933 | 96.9491 | 115.375 | 20.5615 | 0.003092 | 0.002771 | 221.679 | 0.26098 | 8.4011 | 0.002834 | 0.004494 | 0.004412 | 5e-05 | 0.000938 | 0.27616 | 1.326 | 2.5476 | 98.51 | 20.04 | 49.695 | 95.5386 | 89.3286 | 87.7929 | 0.011836 | 0.97812 | 0.022785 | 23.2943 | -0.005145 | 0.15526 | 0.010811 | 0.72979 | 0.72979 | 0.008574
2019-04-09 14:51 | 1554835885.5489 | 1004 | 0.20765 | 0.001916 | 0.000132 | 96.8621 | 96.819 | 115.4076 | 20.5626 | 0.003095 | 0.00281 | 221.6578 | 0.26077 | 8.4021 | 0.002837 | 0.004575 | 0.00446 | 5.1e-05 | 0.00094 | 0.27024 | 1.3623 | 2.5475 | 98.51 | 20.04 | 49.698 | 95.6063 | 89.3829 | 88.7407 | 0.011837 | 0.97556 | 0.022758 | 23.3282 | 0.003978 | 0.15513 | 0.009351 | 0.62947 | 0.62947 | 0.009932
2019-04-09 14:51 | 1554835886.5489 | 1005 | 0.2077 | 0.001904 | 0.000134 | 96.7864 | 96.8768 | 115.1408 | 20.5624 | 0.003092 | 0.002729 | 221.6072 | 0.26035 | 8.4007 | 0.002905 | 0.004478 | 0.00449 | 4.9e-05 | 0.000937 | 0.31421 | 1.3313 | 2.5475 | 98.51 | 20.04 | 49.702 | 95.5339 | 91.0199 | 86.7991 | 0.011838 | 0.98082 | 0.022793 | 23.2386 | -0.008349 | 0.15488 | 0.00904 | 0.60911 | 0.60911 | 0.010247
2019-04-09 14:51 | 1554835887.5505 | 1006 | 0.20791 | 0.001893 | 0.000136 | 96.7483 | 96.7034 | 115.1105 | 20.5624 | 0.003061 | 0.002761 | 221.6001 | 0.26047 | 8.4004 | 0.002859 | 0.004212 | 0.004129 | 4.8e-05 | 0.000937 | 0.30038 | 1.3299 | 2.5476 | 98.51 | 20.04 | 49.705 | 94.7981 | 89.9115 | 87.5721 | 0.011839 | 0.97872 | 0.022765 | 23.2605 | 0.004156 | 0.15495 | 0.007808 | 0.52437 | 0.52437 | 0.011909
2019-04-09 14:51 | 1554835888.5505 | 1007 | 0.20768 | 0.00187 | 0.000138 | 96.7751 | 96.7089 | 115.0675 | 20.5627 | 0.003087 | 0.002781 | 221.5842 | 0.26076 | 8.4016 | 0.002823 | 0.004235 | 0.0042 | 5.4e-05 | 0.000936 | 0.30071 | 1.342 | 2.5476 | 98.51 | 20.04 | 49.714 | 95.4152 | 89.0492 | 88.0386 | 0.011841 | 0.97745 | 0.022746 | 23.2712 | 0.006124 | 0.15512 | 0.009221 | 0.61992 | 0.61992 | 0.010084
2019-04-09 14:51 | 1554835889.55 | 1008 | 0.20786 | 0.001849 | 0.000141 | 96.673 | 96.6442 | 115.0084 | 20.5637 | 0.003047 | 0.002767 | 221.5693 | 0.26077 | 8.4022 | 0.00289 | 0.004913 | 0.004813 | 5.2e-05 | 0.000929 | 0.30021 | 1.3552 | 2.5476 | 98.51 | 20.04 | 49.719 | 94.4549 | 90.6802 | 87.7045 | 0.011842 | 0.97836 | 0.022751 | 23.2544 | 0.002665 | 0.15513 | 0.008151 | 0.5471 | 0.5471 | 0.011427
2019-04-09 14:51 | 1554835890.5501 | 1009 | 0.2078 | 0.001832 | 0.000144 | 96.5774 | 96.7178 | 114.8711 | 20.5629 | 0.003038 | 0.002779 | 221.5424 | 0.26093 | 8.4011 | 0.002876 | 0.004087 | 0.003957 | 5.1e-05 | 0.000929 | 0.27671 | 1.3609 | 2.5476 | 98.51 | 20.04 | 49.724 | 94.2459 | 90.3222 | 87.9989 | 0.011843 | 0.97756 | 0.022728 | 23.25 | -0.012997 | 0.15523 | 0.008527 | 0.57194 | 0.57194 | 0.010938
2019-04-09 14:52 | 1554835891.5501 | 1010 | 0.20784 | 0.001815 | 0.000146 | 96.5798 | 96.2852 | 114.6309 | 20.5625 | 0.00302 | 0.002719 | 221.5161 | 0.26026 | 8.4035 | 0.00281 | 0.00505 | 0.004997 | 5.3e-05 | 0.00093 | 0.286 | 1.3279 | 2.5476 | 98.51 | 20.04 | 49.732 | 93.8042 | 88.7366 | 86.5447 | 0.011845 | 0.98151 | 0.02275 | 23.1789 | 0.02732 | 0.15483 | 0.008265 | 0.55455 | 0.55455 | 0.011251
2019-04-09 14:52 | 1554835892.5491 | 1011 | 0.20789 | 0.001792 | 0.000149 | 96.4595 | 96.2477 | 114.6652 | 20.5632 | 0.003047 | 0.002835 | 221.4976 | 0.26001 | 8.4021 | 0.002877 | 0.004354 | 0.004321 | 5.1e-05 | 0.000929 | 0.27024 | 1.3513 | 2.5476 | 98.51 | 20.04 | 49.737 | 94.4614 | 90.3608 | 89.3442 | 0.011846 | 0.97393 | 0.022666 | 23.2724 | 0.019657 | 0.15468 | 0.008019 | 0.53563 | 0.53563 | 0.011638
2019-04-09 14:52 | 1554835893.5491 | 1012 | 0.20783 | 0.001767 | 0.000152 | 96.4216 | 96.331 | 114.7298 | 20.5639 | 0.003081 | 0.002726 | 221.4481 | 0.26073 | 8.4012 | 0.002859 | 0.004246 | 0.004161 | 5.1e-05 | 0.000935 | 0.29966 | 1.3568 | 2.5477 | 98.51 | 20.04 | 49.748 | 95.2747 | 89.9106 | 86.7234 | 0.011849 | 0.98103 | 0.022755 | 23.1946 | 0.008412 | 0.15511 | 0.008379 | 0.56204 | 0.56204 | 0.011122
2019-04-09 14:52 | 1554835894.5485 | 1013 | 0.20788 | 0.001746 | 0.000155 | 96.3543 | 96.1479 | 114.6619 | 20.564 | 0.002924 | 0.002788 | 221.4463 | 0.26082 | 8.4012 | 0.002859 | 0.004243 | 0.004159 | 5.3e-05 | 0.000938 | 0.27848 | 1.3385 | 2.5476 | 98.51 | 20.04 | 49.753 | 91.4802 | 89.9319 | 88.2139 | 0.01185 | 0.97698 | 0.022701 | 23.2357 | 0.019184 | 0.15516 | 0.008138 | 0.54424 | 0.54424 | 0.011489
2019-04-09 14:52 | 1554835895.5486 | 1014 | 0.20795 | 0.001725 | 0.000159 | 96.3141 | 96.3448 | 114.4596 | 20.5634 | 0.003122 | 0.002734 | 221.414 | 0.25998 | 8.4028 | 0.002905 | 0.00482 | 0.004755 | 4.9e-05 | 0.000938 | 0.25426 | 1.3748 | 2.5476 | 98.51 | 20.04 | 49.759 | 96.2703 | 91.0411 | 86.9016 | 0.011851 | 0.98054 | 0.022722 | 23.173 | -0.002844 | 0.15466 | 0.00774 | 0.51741 | 0.51741 | 0.012046
2019-04-09 14:52 | 1554835896.5492 | 1015 | 0.20788 | 0.001705 | 0.000163 | 96.3076 | 96.3255 | 114.8187 | 20.563 | 0.002939 | 0.002757 | 221.3883 | 0.25998 | 8.4017 | 0.002794 | 0.003947 | 0.003888 | 5.4e-05 | 0.000936 | 0.29264 | 1.3911 | 2.5478 | 98.51 | 20.04 | 49.77 | 91.8436 | 88.3497 | 87.4572 | 0.011854 | 0.97903 | 0.02274 | 23.2273 | -0.001663 | 0.15466 | 0.008175 | 0.54722 | 0.54722 | 0.01139
2019-04-09 14:52 | 1554835897.5484 | 1016 | 0.20789 | 0.001686 | 0.000166 | 96.2275 | 96.2202 | 114.5582 | 20.5642 | 0.003001 | 0.002696 | 221.3901 | 0.25957 | 8.4018 | 0.002815 | 0.003982 | 0.003892 | 5.2e-05 | 0.00094 | 0.26217 | 1.391 | 2.5476 | 98.51 | 20.04 | 49.777 | 93.3529 | 88.8703 | 86.0054 | 0.011856 | 0.98299 | 0.02276 | 23.1541 | 0.000679 | 0.15442 | 0.008102 | 0.54258 | 0.54258 | 0.01147
2019-04-09 14:52 | 1554835898.5486 | 1017 | 0.20801 | 0.001674 | 0.000169 | 96.1584 | 95.9951 | 114.8112 | 20.5644 | 0.002975 | 0.002681 | 221.3403 | 0.26083 | 8.4023 | 0.002852 | 0.004297 | 0.004206 | 5.1e-05 | 0.00094 | 0.29192 | 1.3719 | 2.5476 | 98.51 | 20.04 | 49.791 | 92.7247 | 89.7533 | 85.6464 | 0.011859 | 0.98397 | 0.022797 | 23.1681 | 0.015202 | 0.15517 | 0.007419 | 0.49663 | 0.49663 | 0.012591
2019-04-09 14:52 | 1554835899.5495 | 1018 | 0.20809 | 0.001656 | 0.000173 | 96.0461 | 95.9566 | 114.91 | 20.5642 | 0.002996 | 0.002709 | 221.3675 | 0.26014 | 8.4033 | 0.002856 | 0.004635 | 0.004565 | 5e-05 | 0.000945 | 0.30486 | 1.3419 | 2.5476 | 98.51 | 20.04 | 49.797 | 93.2348 | 89.85 | 86.3021 | 0.011861 | 0.98218 | 0.022786 | 23.1993 | 0.008345 | 0.15476 | 0.006943 | 0.46374 | 0.46374 | 0.013449
2019-04-09 14:52 | 1554835900.5484 | 1019 | 0.20777 | 0.001642 | 0.000176 | 96.05 | 95.8774 | 114.6577 | 20.5649 | 0.003017 | 0.002744 | 221.3303 | 0.25987 | 8.4003 | 0.002727 | 0.00412 | 0.004023 | 5e-05 | 0.000942 | 0.28495 | 1.367 | 2.5476 | 98.51 | 20.04 | 49.803 | 93.7385 | 86.7501 | 87.1616 | 0.011862 | 0.97983 | 0.022734 | 23.2015 | 0.016087 | 0.1546 | 0.008863 | 0.59306 | 0.59306 | 0.010505
2019-04-09 14:52 | 1554835901.5483 | 1020 | 0.20831 | 0.00163 | 0.000179 | 95.8412 | 95.791 | 114.8093 | 20.5648 | 0.00298 | 0.002766 | 221.3309 | 0.26019 | 8.4017 | 0.002815 | 0.00458 | 0.004515 | 5e-05 | 0.000937 | 0.29623 | 1.3415 | 2.5476 | 98.51 | 20.04 | 49.813 | 92.827 | 88.8717 | 87.6907 | 0.011864 | 0.9784 | 0.022732 | 23.2338 | 0.004689 | 0.15479 | 0.005629 | 0.37297 | 0.37297 | 0.016725
2019-04-09 14:52 | 1554835902.5494 | 1021 | 0.20798 | 0.001607 | 0.000181 | 95.8159 | 95.5618 | 115.0959 | 20.5642 | 0.002959 | 0.002712 | 221.3263 | 0.26062 | 8.3997 | 0.002776 | 0.00523 | 0.005093 | 5.1e-05 | 0.00093 | 0.28241 | 1.3206 | 2.5475 | 98.51 | 20.04 | 49.818 | 92.3391 | 87.9277 | 86.3793 | 0.011866 | 0.98196 | 0.022802 | 23.2205 | 0.023754 | 0.15504 | 0.007683 | 0.51408 | 0.51408 | 0.012154
2019-04-09 14:52 | 1554835903.5484 | 1022 | 0.20784 | 0.001587 | 0.000184 | 95.6694 | 95.4455 | 115.0502 | 20.5642 | 0.002933 | 0.002721 | 221.301 | 0.2606 | 8.4006 | 0.002803 | 0.004528 | 0.004472 | 5.3e-05 | 0.000925 | 0.27383 | 1.3166 | 2.5476 | 98.51 | 20.04 | 49.824 | 91.7103 | 88.5687 | 86.5957 | 0.011867 | 0.98137 | 0.02279 | 23.2229 | 0.020961 | 0.15503 | 0.008527 | 0.57122 | 0.57122 | 0.010938
2019-04-09 14:52 | 1554835904.5486 | 1023 | 0.20807 | 0.00157 | 0.000186 | 95.5161 | 95.2156 | 114.9316 | 20.5656 | 0.00295 | 0.002704 | 221.292 | 0.26058 | 8.4021 | 0.002766 | 0.004528 | 0.004478 | 5.1e-05 | 0.000923 | 0.29717 | 1.3523 | 2.5476 | 98.51 | 20.04 | 49.83 | 92.1136 | 87.6907 | 86.2012 | 0.011868 | 0.98245 | 0.022791 | 23.1982 | 0.028181 | 0.15502 | 0.007145 | 0.47688 | 0.47688 | 0.0131
2019-04-09 14:52 | 1554835905.5484 | 1024 | 0.2081 | 0.00155 | 0.000187 | 95.4291 | 95.143 | 114.9443 | 20.5657 | 0.002948 | 0.002693 | 221.2687 | 0.26039 | 8.4018 | 0.002834 | 0.004227 | 0.004146 | 5.2e-05 | 0.000919 | 0.27616 | 1.3113 | 2.5476 | 98.51 | 20.03 | 49.833 | 92.0762 | 89.3303 | 85.9288 | 0.011862 | 0.9832 | 0.022801 | 23.1907 | 0.026858 | 0.15491 | 0.006994 | 0.46669 | 0.46669 | 0.013377
2019-04-09 14:52 | 1554835906.5499 | 1025 | 0.20828 | 0.001532 | 0.00019 | 95.2808 | 95.0101 | 114.9635 | 20.5646 | 0.002936 | 0.002713 | 221.2888 | 0.26017 | 8.4019 | 0.002775 | 0.004101 | 0.00407 | 5.1e-05 | 0.000921 | 0.30021 | 1.3441 | 2.5476 | 98.51 | 20.03 | 49.84 | 91.7685 | 87.896 | 86.4117 | 0.011863 | 0.98188 | 0.022788 | 23.2082 | 0.025445 | 0.15478 | 0.00593 | 0.3938 | 0.3938 | 0.015839
2019-04-09 14:52 | 1554835907.5485 | 1026 | 0.20812 | 0.001519 | 0.000191 | 95.1664 | 94.9299 | 115.1863 | 20.5642 | 0.002971 | 0.002726 | 221.2672 | 0.25968 | 8.4042 | 0.002836 | 0.004394 | 0.004341 | 4.6e-05 | 0.000916 | 0.26609 | 1.321 | 2.5476 | 98.51 | 20.03 | 49.841 | 92.6269 | 89.3741 | 86.7205 | 0.011864 | 0.98103 | 0.0228 | 23.2406 | 0.022256 | 0.15449 | 0.006918 | 0.46131 | 0.46131 | 0.013496
2019-04-09 14:52 | 1554835908.5485 | 1027 | 0.20813 | 0.001498 | 0.000194 | 95.0634 | 94.8171 | 115.2041 | 20.5649 | 0.00294 | 0.002664 | 221.2667 | 0.2602 | 8.4037 | 0.002865 | 0.004322 | 0.004288 | 5.1e-05 | 0.000916 | 0.24729 | 1.3299 | 2.5475 | 98.51 | 20.03 | 49.843 | 91.8768 | 90.0734 | 85.231 | 0.011864 | 0.98511 | 0.022849 | 23.1943 | 0.023209 | 0.15479 | 0.006907 | 0.46142 | 0.46142 | 0.01352
2019-04-09 14:52 | 1554835909.5487 | 1028 | 0.20829 | 0.001486 | 0.000195 | 95.0575 | 94.8993 | 115.0013 | 20.5643 | 0.002918 | 0.002705 | 221.2531 | 0.26003 | 8.4043 | 0.002825 | 0.004965 | 0.004855 | 4.8e-05 | 0.000922 | 0.27959 | 1.3724 | 2.5476 | 98.51 | 20.04 | 49.843 | 91.3407 | 89.1064 | 86.2084 | 0.011871 | 0.98243 | 0.022798 | 23.2054 | 0.014893 | 0.15469 | 0.005921 | 0.39304 | 0.39304 | 0.015861
2019-04-09 14:52 | 1554835910.548 | 1029 | 0.20823 | 0.001462 | 0.000199 | 94.9163 | 94.7823 | 115.2764 | 20.565 | 0.002925 | 0.002697 | 221.1957 | 0.26037 | 8.4002 | 0.00284 | 0.004601 | 0.00456 | 4.9e-05 | 0.000923 | 0.27312 | 1.3691 | 2.5475 | 98.51 | 20.04 | 49.843 | 91.5093 | 89.4709 | 86.0351 | 0.011871 | 0.98291 | 0.022831 | 23.2276 | 0.012637 | 0.15489 | 0.006299 | 0.41923 | 0.41923 | 0.01489
2019-04-09 14:52 | 1554835911.5477 | 1030 | 0.20821 | 0.00145 | 0.000201 | 94.8293 | 94.5514 | 115.1328 | 20.5643 | 0.002989 | 0.002669 | 221.2179 | 0.26026 | 8.4011 | 0.002776 | 0.004064 | 0.003993 | 4.7e-05 | 0.000922 | 0.286 | 1.3479 | 2.5474 | 98.51 | 20.04 | 49.843 | 93.0611 | 87.9215 | 85.357 | 0.011871 | 0.98477 | 0.022838 | 23.1912 | 0.026257 | 0.15483 | 0.006461 | 0.43023 | 0.43023 | 0.014503
2019-04-09 14:52 | 1554835912.548 | 1031 | 0.20821 | 0.001431 | 0.000203 | 94.6602 | 94.3834 | 114.9145 | 20.565 | 0.002995 | 0.002742 | 221.1739 | 0.25959 | 8.4011 | 0.002791 | 0.004576 | 0.004544 | 4.8e-05 | 0.000922 | 0.28838 | 1.3689 | 2.5476 | 98.51 | 20.04 | 49.843 | 93.1917 | 88.2932 | 87.1015 | 0.011871 | 0.98 | 0.022761 | 23.2255 | 0.026191 | 0.15443 | 0.006467 | 0.42914 | 0.42914 | 0.014502
2019-04-09 14:52 | 1554835913.5487 | 1032 | 0.20829 | 0.001414 | 0.000206 | 94.5173 | 94.0045 | 115.1213 | 20.5646 | 0.002915 | 0.002677 | 221.1776 | 0.25967 | 8.4041 | 0.00279 | 0.004069 | 0.00399 | 4.7e-05 | 0.000924 | 0.30342 | 1.3719 | 2.5476 | 98.51 | 20.04 | 49.843 | 91.2682 | 88.2583 | 85.536 | 0.011871 | 0.98427 | 0.022831 | 23.1958 | 0.048655 | 0.15448 | 0.006012 | 0.3992 | 0.3992 | 0.015595
2019-04-09 14:52 | 1554835914.5478 | 1033 | 0.20827 | 0.001396 | 0.000208 | 94.4514 | 94.0789 | 115.2607 | 20.5663 | 0.002891 | 0.002663 | 221.1555 | 0.26033 | 8.4017 | 0.00274 | 0.004102 | 0.004092 | 4.7e-05 | 0.000921 | 0.29391 | 1.3663 | 2.5477 | 98.51 | 20.04 | 49.844 | 90.7047 | 87.0645 | 85.206 | 0.011872 | 0.98518 | 0.022855 | 23.1992 | 0.035355 | 0.15487 | 0.006136 | 0.40796 | 0.40796 | 0.015299
2019-04-09 14:52 | 1554835915.5478 | 1034 | 0.20834 | 0.001377 | 0.00021 | 94.3299 | 93.8961 | 115.1992 | 20.5657 | 0.002958 | 0.002714 | 221.1427 | 0.25985 | 8.3993 | 0.002828 | 0.00446 | 0.004407 | 5e-05 | 0.000915 | 0.28512 | 1.3459 | 2.5476 | 98.51 | 20.03 | 49.843 | 92.317 | 89.1803 | 86.4277 | 0.011864 | 0.98183 | 0.02281 | 23.2325 | 0.041233 | 0.15458 | 0.00573 | 0.37939 | 0.37939 | 0.01642
2019-04-09 14:52 | 1554835916.5478 | 1035 | 0.20832 | 0.001359 | 0.000213 | 94.1614 | 93.6549 | 115.237 | 20.5649 | 0.002945 | 0.002668 | 221.1569 | 0.26048 | 8.4011 | 0.002856 | 0.004235 | 0.004153 | 4.8e-05 | 0.000912 | 0.28567 | 1.3098 | 2.5474 | 98.51 | 20.04 | 49.842 | 91.9855 | 89.8618 | 85.3256 | 0.011871 | 0.98485 | 0.022849 | 23.2007 | 0.048247 | 0.15496 | 0.005886 | 0.39051 | 0.39051 | 0.015992
2019-04-09 14:52 | 1554835917.5489 | 1036 | 0.20835 | 0.001339 | 0.000217 | 94.0277 | 93.511 | 115.3689 | 20.5659 | 0.002888 | 0.00267 | 221.108 | 0.25942 | 8.4017 | 0.00288 | 0.004264 | 0.004217 | 4.9e-05 | 0.000912 | 0.30917 | 1.3148 | 2.5473 | 98.51 | 20.04 | 49.842 | 90.6207 | 90.4323 | 85.3631 | 0.011871 | 0.98475 | 0.022861 | 23.2151 | 0.049287 | 0.15433 | 0.00576 | 0.38187 | 0.38187 | 0.016287
2019-04-09 14:52 | 1554835918.5502 | 1037 | 0.20836 | 0.001323 | 0.00022 | 93.7909 | 93.555 | 115.4518 | 20.5662 | 0.002923 | 0.00265 | 221.1173 | 0.25952 | 8.4006 | 0.002774 | 0.003924 | 0.003932 | 5.2e-05 | 0.000909 | 0.28982 | 1.3405 | 2.5474 | 98.51 | 20.04 | 49.841 | 91.4667 | 87.8694 | 84.8999 | 0.011871 | 0.98602 | 0.022884 | 23.2085 | 0.022525 | 0.15439 | 0.005689 | 0.37719 | 0.37719 | 0.016495
2019-04-09 14:52 | 1554835919.5481 | 1038 | 0.20838 | 0.001306 | 0.000225 | 93.6921 | 93.1527 | 115.5783 | 20.5666 | 0.00295 | 0.002695 | 221.0811 | 0.26048 | 8.4 | 0.002802 | 0.003699 | 0.003605 | 5.2e-05 | 0.000909 | 0.28296 | 1.3026 | 2.5474 | 98.51 | 20.04 | 49.841 | 92.121 | 88.5564 | 85.9842 | 0.011871 | 0.98305 | 0.022862 | 23.2563 | 0.051639 | 0.15496 | 0.005585 | 0.36943 | 0.36943 | 0.016904
2019-04-09 14:52 | 1554835920.549 | 1039 | 0.20847 | 0.001291 | 0.00023 | 93.6152 | 93.2429 | 115.4358 | 20.5659 | 0.002942 | 0.00264 | 221.0902 | 0.25984 | 8.401 | 0.002815 | 0.004345 | 0.004297 | 4.8e-05 | 0.000912 | 0.27455 | 1.307 | 2.5474 | 98.51 | 20.04 | 49.841 | 91.9293 | 88.8711 | 84.6476 | 0.011871 | 0.98672 | 0.022891 | 23.1987 | 0.035647 | 0.15458 | 0.005052 | 0.33306 | 0.33306 | 0.018704
2019-04-09 14:52 | 1554835921.548 | 1040 | 0.20848 | 0.001273 | 0.000234 | 93.4854 | 93.1641 | 115.9729 | 20.5658 | 0.002933 | 0.002704 | 221.072 | 0.26029 | 8.4042 | 0.00285 | 0.004672 | 0.00458 | 4.6e-05 | 0.00091 | 0.29662 | 1.2886 | 2.5474 | 98.51 | 20.04 | 49.84 | 91.7088 | 89.7102 | 86.2046 | 0.011871 | 0.98244 | 0.022894 | 23.3031 | 0.030797 | 0.15485 | 0.005027 | 0.33121 | 0.33121 | 0.018841
2019-04-09 14:52 | 1554835922.549 | 1041 | 0.20853 | 0.001251 | 0.000238 | 93.4604 | 92.9215 | 116.0975 | 20.567 | 0.002938 | 0.002675 | 221.049 | 0.25985 | 8.4008 | 0.002856 | 0.004647 | 0.004491 | 4.6e-05 | 0.000909 | 0.27201 | 1.3532 | 2.5474 | 98.51 | 20.04 | 49.839 | 91.8171 | 89.8644 | 85.4876 | 0.011871 | 0.98441 | 0.022929 | 23.2924 | 0.051728 | 0.15458 | 0.004717 | 0.31022 | 0.31022 | 0.020082
2019-04-09 14:52 | 1554835923.5472 | 1042 | 0.20858 | 0.001235 | 0.000242 | 93.2954 | 92.6726 | 116.0395 | 20.5659 | 0.002942 | 0.002693 | 221.0613 | 0.25986 | 8.4006 | 0.002795 | 0.004658 | 0.004558 | 4.7e-05 | 0.000911 | 0.29391 | 1.3131 | 2.5474 | 98.51 | 20.04 | 49.836 | 91.9185 | 88.3907 | 85.9375 | 0.01187 | 0.98317 | 0.022909 | 23.3012 | 0.059908 | 0.15459 | 0.004444 | 0.29098 | 0.29098 | 0.02141
2019-04-09 14:52 | 1554835924.5481 | 1043 | 0.20857 | 0.001202 | 0.000246 | 93.1963 | 92.5646 | 115.7898 | 20.567 | 0.002869 | 0.002653 | 221.0272 | 0.26003 | 8.4003 | 0.002781 | 0.004064 | 0.004089 | 4.9e-05 | 0.000906 | 0.30182 | 1.3223 | 2.5474 | 98.51 | 20.04 | 49.836 | 90.166 | 88.0473 | 84.9591 | 0.01187 | 0.98586 | 0.022916 | 23.2443 | 0.060832 | 0.15469 | 0.004521 | 0.29612 | 0.29612 | 0.021053
2019-04-09 14:52 | 1554835925.5471 | 1044 | 0.20861 | 0.001175 | 0.00025 | 93.0879 | 92.3612 | 115.9656 | 20.5673 | 0.002938 | 0.002696 | 221.0161 | 0.2605 | 8.4 | 0.002824 | 0.004603 | 0.004452 | 5.1e-05 | 0.000907 | 0.27865 | 1.3352 | 2.5476 | 98.51 | 20.04 | 49.835 | 91.8245 | 89.0889 | 85.9905 | 0.01187 | 0.98303 | 0.0229 | 23.2955 | 0.070099 | 0.15497 | 0.004349 | 0.28394 | 0.28394 | 0.021995
2019-04-09 14:52 | 1554835926.5482 | 1045 | 0.20861 | 0.001146 | 0.000252 | 92.9336 | 92.3058 | 115.7229 | 20.5666 | 0.002824 | 0.002676 | 220.9807 | 0.26009 | 8.4032 | 0.002853 | 0.004338 | 0.004289 | 5.1e-05 | 0.000909 | 0.28296 | 1.3357 | 2.5475 | 98.51 | 20.04 | 49.831 | 89.0726 | 89.7788 | 85.5286 | 0.011869 | 0.9843 | 0.022891 | 23.2561 | 0.060632 | 0.15473 | 0.004358 | 0.28431 | 0.28431 | 0.021932
2019-04-09 14:52 | 1554835927.5474 | 1046 | 0.20861 | 0.001125 | 0.000255 | 92.8958 | 92.3302 | 115.7952 | 20.567 | 0.002799 | 0.002639 | 220.9545 | 0.25985 | 8.4012 | 0.002829 | 0.004359 | 0.004273 | 4.8e-05 | 0.000912 | 0.29407 | 1.35 | 2.5476 | 98.51 | 20.04 | 49.829 | 88.4747 | 89.2136 | 84.6369 | 0.011868 | 0.98675 | 0.022927 | 23.2344 | 0.054623 | 0.15458 | 0.004368 | 0.28527 | 0.28527 | 0.021838
2019-04-09 14:52 | 1554835928.5482 | 1047 | 0.20864 | 0.001105 | 0.000257 | 92.7889 | 92.2831 | 115.6724 | 20.567 | 0.002815 | 0.002649 | 220.9465 | 0.26004 | 8.4014 | 0.002797 | 0.004561 | 0.004503 | 5.3e-05 | 0.000909 | 0.28711 | 1.3373 | 2.5476 | 98.51 | 20.04 | 49.828 | 88.8605 | 88.4286 | 84.8756 | 0.011868 | 0.98609 | 0.022907 | 23.2299 | 0.048893 | 0.1547 | 0.004214 | 0.27436 | 0.27436 | 0.022723
2019-04-09 14:52 | 1554835929.5486 | 1048 | 0.20867 | 0.001081 | 0.000259 | 92.771 | 92.2817 | 115.8522 | 20.5661 | 0.002891 | 0.002627 | 220.9178 | 0.25924 | 8.4029 | 0.00279 | 0.004783 | 0.004642 | 5e-05 | 0.00091 | 0.28727 | 1.3289 | 2.5475 | 98.51 | 20.04 | 49.826 | 90.6905 | 88.2528 | 84.3356 | 0.011867 | 0.98758 | 0.022942 | 23.2304 | 0.047304 | 0.15422 | 0.004039 | 0.26262 | 0.26262 | 0.023666
2019-04-09 14:52 | 1554835930.5473 | 1049 | 0.20868 | 0.001064 | 0.00026 | 92.7091 | 92.028 | 116.1412 | 20.5668 | 0.002843 | 0.002668 | 220.9367 | 0.26022 | 8.403 | 0.002712 | 0.004556 | 0.004639 | 4.8e-05 | 0.000914 | 0.27793 | 1.3123 | 2.5475 | 98.51 | 20.04 | 49.824 | 89.5502 | 86.3794 | 85.3171 | 0.011867 | 0.98487 | 0.022939 | 23.2912 | 0.065948 | 0.15481 | 0.004017 | 0.26104 | 0.26104 | 0.023899
2019-04-09 14:52 | 1554835931.5479 | 1050 | 0.2087 | 0.001045 | 0.000261 | 92.6903 | 91.8394 | 115.9576 | 20.5672 | 0.002886 | 0.002673 | 220.9084 | 0.26016 | 8.4014 | 0.002754 | 0.004369 | 0.004263 | 5.1e-05 | 0.000923 | 0.31094 | 1.3114 | 2.5475 | 98.51 | 20.04 | 49.822 | 90.5649 | 87.3956 | 85.4495 | 0.011866 | 0.98451 | 0.022917 | 23.2771 | 0.082487 | 0.15477 | 0.003912 | 0.25346 | 0.25346 | 0.024608
2019-04-09 14:52 | 1554835932.5474 | 1051 | 0.2087 | 0.001028 | 0.000262 | 92.5879 | 91.8061 | 115.7834 | 20.5675 | 0.002836 | 0.00264 | 220.9172 | 0.25984 | 8.4019 | 0.002777 | 0.004328 | 0.00428 | 5.1e-05 | 0.000919 | 0.27815 | 1.298 | 2.5475 | 98.51 | 20.04 | 49.821 | 89.372 | 87.9511 | 84.6453 | 0.011866 | 0.98672 | 0.022925 | 23.2335 | 0.075848 | 0.15458 | 0.003936 | 0.25518 | 0.25518 | 0.024412
2019-04-09 14:52 | 1554835933.5474 | 1052 | 0.20868 | 0.001009 | 0.000261 | 92.4878 | 91.6405 | 115.8057 | 20.5677 | 0.002859 | 0.0026 | 220.9078 | 0.25961 | 8.4008 | 0.002712 | 0.004512 | 0.004435 | 5.4e-05 | 0.000917 | 0.29192 | 1.3078 | 2.5475 | 98.51 | 20.04 | 49.821 | 89.9355 | 86.3921 | 83.6997 | 0.011866 | 0.98934 | 0.022958 | 23.205 | 0.082319 | 0.15444 | 0.004086 | 0.26591 | 0.26591 | 0.023407
2019-04-09 14:52 | 1554835934.5474 | 1053 | 0.2087 | 0.001 | 0.000259 | 92.3863 | 91.623 | 115.8855 | 20.5676 | 0.002879 | 0.002632 | 220.883 | 0.25949 | 8.4037 | 0.002788 | 0.004486 | 0.004434 | 5.3e-05 | 0.000916 | 0.28263 | 1.3463 | 2.5475 | 98.51 | 20.04 | 49.819 | 90.4034 | 88.2246 | 84.4549 | 0.011866 | 0.98725 | 0.022941 | 23.2376 | 0.074207 | 0.15437 | 0.004013 | 0.26083 | 0.26083 | 0.023852
2019-04-09 14:52 | 1554835935.5477 | 1054 | 0.20868 | 0.000984 | 0.000258 | 92.3127 | 91.5469 | 115.748 | 20.5694 | 0.002831 | 0.002661 | 220.9055 | 0.25962 | 8.4018 | 0.00288 | 0.004347 | 0.004228 | 5.1e-05 | 0.00092 | 0.27992 | 1.3253 | 2.5473 | 98.51 | 20.04 | 49.816 | 89.2565 | 90.428 | 85.1627 | 0.011865 | 0.9853 | 0.022905 | 23.2468 | 0.074511 | 0.15445 | 0.004109 | 0.26709 | 0.26709 | 0.023305
2019-04-09 14:52 | 1554835936.5475 | 1055 | 0.20869 | 0.000977 | 0.000257 | 92.1998 | 91.4925 | 116.1376 | 20.5678 | 0.002837 | 0.002653 | 220.866 | 0.26054 | 8.403 | 0.002791 | 0.004676 | 0.004572 | 5.2e-05 | 0.000923 | 0.30397 | 1.329 | 2.5474 | 98.51 | 20.04 | 49.812 | 89.4043 | 88.2864 | 84.9724 | 0.011864 | 0.98582 | 0.02295 | 23.2797 | 0.068883 | 0.155 | 0.004087 | 0.26611 | 0.26611 | 0.023472
2019-04-09 14:52 | 1554835937.5481 | 1056 | 0.20868 | 0.000971 | 0.000255 | 92.1695 | 91.3401 | 116.1308 | 20.5674 | 0.00283 | 0.002654 | 220.8803 | 0.25937 | 8.4021 | 0.00281 | 0.003846 | 0.003859 | 5.2e-05 | 0.000924 | 0.30254 | 1.3125 | 2.5475 | 98.51 | 20.04 | 49.811 | 89.2366 | 88.7468 | 84.9977 | 0.011864 | 0.98575 | 0.022948 | 23.2798 | 0.080853 | 0.1543 | 0.004155 | 0.27087 | 0.27087 | 0.022957
2019-04-09 14:52 | 1554835938.5481 | 1057 | 0.20868 | 0.00096 | 0.000253 | 92.0131 | 91.2472 | 116.1712 | 20.5679 | 0.002814 | 0.002671 | 220.8546 | 0.26001 | 8.3988 | 0.002862 | 0.00445 | 0.004362 | 4.9e-05 | 0.000917 | 0.30757 | 1.3647 | 2.5475 | 98.51 | 20.04 | 49.809 | 88.8441 | 90.0091 | 85.405 | 0.011863 | 0.98463 | 0.022939 | 23.2971 | 0.074772 | 0.15468 | 0.004172 | 0.27208 | 0.27208 | 0.022911
2019-04-09 14:52 | 1554835939.5476 | 1058 | 0.2087 | 0.000959 | 0.000251 | 91.9163 | 91.1591 | 115.9743 | 20.5685 | 0.002814 | 0.002638 | 220.8572 | 0.25961 | 8.4024 | 0.002794 | 0.004335 | 0.00427 | 5.1e-05 | 0.000919 | 0.30934 | 1.3607 | 2.5476 | 98.51 | 20.04 | 49.807 | 88.8377 | 88.3581 | 84.6056 | 0.011863 | 0.98683 | 0.022945 | 23.2514 | 0.073999 | 0.15444 | 0.004051 | 0.26394 | 0.26394 | 0.023581
2019-04-09 14:52 | 1554835940.5476 | 1059 | 0.20867 | 0.000955 | 0.000249 | 91.8437 | 91.0389 | 115.9013 | 20.5682 | 0.002866 | 0.002594 | 220.8661 | 0.26001 | 8.4013 | 0.002799 | 0.004438 | 0.00444 | 4.8e-05 | 0.000928 | 0.27007 | 1.3138 | 2.5475 | 98.51 | 20.04 | 49.805 | 90.0836 | 88.4703 | 83.5412 | 0.011862 | 0.98978 | 0.022972 | 23.2094 | 0.078719 | 0.15468 | 0.004263 | 0.27895 | 0.27895 | 0.022347
2019-04-09 14:52 | 1554835941.548 | 1060 | 0.20871 | 0.00095 | 0.000246 | 91.7156 | 90.8118 | 116.1259 | 20.5688 | 0.002764 | 0.002609 | 220.8366 | 0.25981 | 8.4 | 0.002719 | 0.004182 | 0.004138 | 5.3e-05 | 0.000928 | 0.30038 | 1.3372 | 2.5475 | 98.51 | 20.04 | 49.803 | 87.6385 | 86.5602 | 83.9186 | 0.011862 | 0.98873 | 0.022982 | 23.2442 | 0.088584 | 0.15456 | 0.004004 | 0.26143 | 0.26143 | 0.023827
2019-04-09 14:52 | 1554835942.5467 | 1061 | 0.20867 | 0.000947 | 0.000244 | 91.676 | 90.8562 | 116.2576 | 20.5687 | 0.002829 | 0.002637 | 220.8398 | 0.26009 | 8.4008 | 0.002786 | 0.004539 | 0.004464 | 4.9e-05 | 0.000929 | 0.30038 | 1.3167 | 2.5475 | 98.51 | 20.04 | 49.798 | 89.2089 | 88.1708 | 84.5856 | 0.011861 | 0.98689 | 0.022974 | 23.2791 | 0.080348 | 0.15473 | 0.004262 | 0.2792 | 0.2792 | 0.022334
2019-04-09 14:52 | 1554835943.5467 | 1062 | 0.20871 | 0.000946 | 0.000241 | 91.59 | 90.8569 | 116.2933 | 20.5689 | 0.002792 | 0.002611 | 220.8146 | 0.26092 | 8.4021 | 0.002811 | 0.004401 | 0.004313 | 5.1e-05 | 0.000928 | 0.28263 | 1.3124 | 2.5475 | 98.51 | 20.04 | 49.795 | 88.3116 | 88.7639 | 83.9677 | 0.01186 | 0.9886 | 0.022997 | 23.2626 | 0.07188 | 0.15522 | 0.004054 | 0.26529 | 0.26529 | 0.02358
2019-04-09 14:52 | 1554835944.5466 | 1063 | 0.20869 | 0.000939 | 0.000238 | 91.4375 | 90.5416 | 116.435 | 20.5686 | 0.00283 | 0.002547 | 220.8192 | 0.26026 | 8.4003 | 0.002726 | 0.003792 | 0.00374 | 5.4e-05 | 0.000931 | 0.27455 | 1.3409 | 2.5475 | 98.51 | 20.04 | 49.788 | 89.2176 | 86.7231 | 82.4292 | 0.011858 | 0.99287 | 0.023061 | 23.2265 | 0.088066 | 0.15483 | 0.004189 | 0.27557 | 0.27557 | 0.022643
2019-04-09 14:52 | 1554835945.5482 | 1064 | 0.20875 | 0.000935 | 0.000234 | 91.2731 | 90.6263 | 116.6597 | 20.5698 | 0.002796 | 0.00263 | 220.8038 | 0.26002 | 8.4023 | 0.002761 | 0.004607 | 0.004539 | 5.3e-05 | 0.000931 | 0.28423 | 1.2975 | 2.5476 | 98.51 | 20.04 | 49.784 | 88.4109 | 87.573 | 84.4173 | 0.011857 | 0.98735 | 0.023019 | 23.3138 | 0.063612 | 0.15469 | 0.003825 | 0.2502 | 0.2502 | 0.024915
2019-04-09 14:52 | 1554835946.5472 | 1065 | 0.20874 | 0.000933 | 0.00023 | 91.1391 | 90.3197 | 116.573 | 20.5694 | 0.002846 | 0.002621 | 220.7999 | 0.25978 | 8.4011 | 0.002796 | 0.004213 | 0.004097 | 5.1e-05 | 0.000931 | 0.30973 | 1.3049 | 2.5475 | 98.51 | 20.03 | 49.779 | 89.6221 | 88.4192 | 84.1928 | 0.011849 | 0.98797 | 0.023018 | 23.2979 | 0.080786 | 0.15454 | 0.003878 | 0.25409 | 0.25409 | 0.024511
2019-04-09 14:52 | 1554835947.5469 | 1066 | 0.20875 | 0.000934 | 0.000226 | 91.0263 | 90.3334 | 116.4862 | 20.569 | 0.002833 | 0.002581 | 220.8171 | 0.26026 | 8.4015 | 0.002784 | 0.003955 | 0.003924 | 5e-05 | 0.000927 | 0.30414 | 1.2826 | 2.5477 | 98.51 | 20.03 | 49.777 | 89.3101 | 88.1119 | 83.2424 | 0.011848 | 0.99061 | 0.02304 | 23.2582 | 0.068339 | 0.15483 | 0.003815 | 0.25024 | 0.25024 | 0.024934
2019-04-09 14:52 | 1554835948.5469 | 1067 | 0.20879 | 0.000928 | 0.000222 | 90.9673 | 90.1476 | 116.4432 | 20.5683 | 0.002808 | 0.002638 | 220.7902 | 0.26063 | 8.3997 | 0.002726 | 0.00372 | 0.003609 | 4.9e-05 | 0.000923 | 0.27113 | 1.3148 | 2.5476 | 98.51 | 20.03 | 49.773 | 88.7031 | 86.7353 | 84.6155 | 0.011847 | 0.98681 | 0.022991 | 23.2987 | 0.080964 | 0.15505 | 0.003565 | 0.23273 | 0.23273 | 0.026848
2019-04-09 14:52 | 1554835949.5479 | 1068 | 0.20877 | 0.000921 | 0.000218 | 90.8363 | 89.8795 | 116.5806 | 20.5693 | 0.002823 | 0.002561 | 220.7984 | 0.2602 | 8.4034 | 0.002759 | 0.004994 | 0.004947 | 5.1e-05 | 0.000924 | 0.29894 | 1.3396 | 2.5476 | 98.51 | 20.04 | 49.77 | 89.0605 | 87.5297 | 82.7645 | 0.011854 | 0.99194 | 0.023065 | 23.252 | 0.094712 | 0.15479 | 0.003724 | 0.24464 | 0.24464 | 0.0255
2019-04-09 14:52 | 1554835950.548 | 1069 | 0.20877 | 0.000919 | 0.000214 | 90.7255 | 89.8478 | 116.547 | 20.5699 | 0.002795 | 0.002552 | 220.7594 | 0.25998 | 8.4018 | 0.002713 | 0.004402 | 0.004313 | 5.1e-05 | 0.000926 | 0.26842 | 1.304 | 2.5475 | 98.51 | 20.03 | 49.767 | 88.386 | 86.4168 | 82.5465 | 0.011846 | 0.99255 | 0.023068 | 23.2415 | 0.086948 | 0.15466 | 0.003706 | 0.24366 | 0.24366 | 0.02558
2019-04-09 14:52 | 1554835951.5483 | 1070 | 0.20874 | 0.000916 | 0.00021 | 90.5948 | 89.7122 | 116.5374 | 20.5692 | 0.002874 | 0.002622 | 220.7602 | 0.26023 | 8.402 | 0.002722 | 0.005127 | 0.005067 | 5.1e-05 | 0.000919 | 0.26488 | 1.3356 | 2.5475 | 98.51 | 20.04 | 49.764 | 90.2994 | 86.6325 | 84.2239 | 0.011853 | 0.98789 | 0.023013 | 23.2953 | 0.08756 | 0.15481 | 0.003901 | 0.25672 | 0.25672 | 0.024302
2019-04-09 14:53 | 1554835952.5484 | 1071 | 0.20875 | 0.00091 | 0.000207 | 90.414 | 89.6042 | 116.7338 | 20.5698 | 0.002873 | 0.002542 | 220.7441 | 0.26052 | 8.4031 | 0.002799 | 0.004237 | 0.004192 | 5.2e-05 | 0.000915 | 0.27505 | 1.3192 | 2.5476 | 98.51 | 20.03 | 49.759 | 90.2666 | 88.4718 | 82.2962 | 0.011844 | 0.99325 | 0.023095 | 23.252 | 0.080476 | 0.15499 | 0.003889 | 0.25702 | 0.25702 | 0.024301
2019-04-09 14:53 | 1554835953.5481 | 1072 | 0.20876 | 0.000902 | 0.000202 | 90.2845 | 89.2579 | 116.6019 | 20.5692 | 0.002759 | 0.002559 | 220.7189 | 0.26039 | 8.4001 | 0.002753 | 0.00452 | 0.004452 | 4.9e-05 | 0.000912 | 0.28744 | 1.3518 | 2.5475 | 98.51 | 20.03 | 49.757 | 87.5217 | 87.3631 | 82.7199 | 0.011844 | 0.99206 | 0.023068 | 23.2527 | 0.10228 | 0.1549 | 0.003813 | 0.25173 | 0.25173 | 0.024798
2019-04-09 14:53 | 1554835954.548 | 1073 | 0.20876 | 0.0009 | 0.000198 | 90.1665 | 89.2074 | 116.5989 | 20.571 | 0.002836 | 0.002572 | 220.7357 | 0.2596 | 8.4008 | 0.002711 | 0.003992 | 0.003917 | 5.3e-05 | 0.000916 | 0.29087 | 1.3413 | 2.5475 | 98.51 | 20.04 | 49.753 | 89.3863 | 86.378 | 83.029 | 0.01185 | 0.9912 | 0.023058 | 23.2625 | 0.09566 | 0.15444 | 0.003813 | 0.25184 | 0.25184 | 0.024713
2019-04-09 14:53 | 1554835955.5461 | 1074 | 0.20881 | 0.000899 | 0.000194 | 89.9605 | 88.9675 | 116.5442 | 20.5693 | 0.00283 | 0.00257 | 220.7092 | 0.25997 | 8.4023 | 0.002667 | 0.004044 | 0.004077 | 5e-05 | 0.000916 | 0.29247 | 1.307 | 2.5476 | 98.51 | 20.04 | 49.752 | 89.2244 | 85.3072 | 82.9751 | 0.01185 | 0.99135 | 0.023054 | 23.2553 | 0.099278 | 0.15466 | 0.003559 | 0.23456 | 0.23456 | 0.026572
2019-04-09 14:53 | 1554835956.5461 | 1075 | 0.20882 | 0.000896 | 0.00019 | 89.8783 | 88.8112 | 116.3803 | 20.5698 | 0.002818 | 0.002548 | 220.7226 | 0.26022 | 8.4015 | 0.002682 | 0.004252 | 0.004163 | 4.9e-05 | 0.000913 | 0.27455 | 1.3058 | 2.5475 | 98.51 | 20.04 | 49.753 | 88.9476 | 85.6674 | 82.439 | 0.01185 | 0.99285 | 0.023055 | 23.2214 | 0.10682 | 0.1548 | 0.003508 | 0.23125 | 0.23125 | 0.026978
2019-04-09 14:53 | 1554835957.5474 | 1076 | 0.20881 | 0.000895 | 0.000186 | 89.7365 | 88.6013 | 116.6315 | 20.5688 | 0.002803 | 0.002524 | 220.7059 | 0.2599 | 8.4016 | 0.002682 | 0.004836 | 0.004744 | 5.3e-05 | 0.000911 | 0.27776 | 1.333 | 2.5476 | 98.51 | 20.04 | 49.752 | 88.5857 | 85.6639 | 81.8724 | 0.01185 | 0.99443 | 0.023099 | 23.2279 | 0.11388 | 0.15461 | 0.003582 | 0.23705 | 0.23705 | 0.026286
2019-04-09 14:53 | 1554835958.5461 | 1077 | 0.2088 | 0.000894 | 0.000182 | 89.7112 | 88.5538 | 116.5857 | 20.5702 | 0.002781 | 0.002567 | 220.6904 | 0.25977 | 8.4012 | 0.002727 | 0.004248 | 0.004218 | 5.1e-05 | 0.000913 | 0.27671 | 1.3149 | 2.5475 | 98.51 | 20.04 | 49.751 | 88.0567 | 86.7545 | 82.9065 | 0.01185 | 0.99154 | 0.02306 | 23.2572 | 0.11615 | 0.15454 | 0.003611 | 0.23885 | 0.23885 | 0.026074
2019-04-09 14:53 | 1554835959.5462 | 1078 | 0.20883 | 0.000889 | 0.000179 | 89.78 | 88.7235 | 116.5097 | 20.5703 | 0.002836 | 0.002593 | 220.6923 | 0.26008 | 8.4014 | 0.002733 | 0.004396 | 0.004291 | 5.2e-05 | 0.000913 | 0.30917 | 1.3651 | 2.5476 | 98.51 | 20.04 | 49.749 | 89.3786 | 86.8889 | 83.5183 | 0.011849 | 0.98984 | 0.023033 | 23.2695 | 0.10588 | 0.15472 | 0.00342 | 0.22561 | 0.22561 | 0.027638
2019-04-09 14:53 | 1554835960.5472 | 1079 | 0.20884 | 0.000885 | 0.000177 | 89.7038 | 88.7218 | 116.7871 | 20.5711 | 0.002764 | 0.00255 | 220.662 | 0.26003 | 8.4009 | 0.002764 | 0.004335 | 0.004346 | 5.4e-05 | 0.000914 | 0.30254 | 1.3406 | 2.5475 | 98.51 | 20.04 | 49.749 | 87.6359 | 87.6297 | 82.4949 | 0.011849 | 0.99269 | 0.023094 | 23.2638 | 0.098457 | 0.15469 | 0.003394 | 0.2245 | 0.2245 | 0.027769
2019-04-09 14:53 | 1554835961.5463 | 1080 | 0.20883 | 0.000883 | 0.000175 | 89.749 | 88.6312 | 116.3292 | 20.5703 | 0.002726 | 0.002549 | 220.6674 | 0.26044 | 8.3989 | 0.002739 | 0.003858 | 0.003847 | 5.3e-05 | 0.000914 | 0.30486 | 1.3491 | 2.5476 | 98.51 | 20.04 | 49.748 | 86.7178 | 87.0403 | 82.459 | 0.011849 | 0.99279 | 0.02305 | 23.217 | 0.1121 | 0.15494 | 0.003456 | 0.22851 | 0.22851 | 0.027324
2019-04-09 14:53 | 1554835962.5466 | 1081 | 0.20886 | 0.000882 | 0.000173 | 89.7181 | 88.5944 | 116.3559 | 20.5708 | 0.002768 | 0.002545 | 220.6503 | 0.26015 | 8.4013 | 0.002684 | 0.003284 | 0.003221 | 5.3e-05 | 0.00091 | 0.32316 | 1.3706 | 2.5476 | 98.51 | 20.03 | 49.746 | 87.7348 | 85.7239 | 82.3799 | 0.011841 | 0.99301 | 0.023055 | 23.217 | 0.11273 | 0.15476 | 0.003286 | 0.21694 | 0.21694 | 0.02875
2019-04-09 14:53 | 1554835963.5467 | 1082 | 0.20888 | 0.000885 | 0.00017 | 89.7157 | 88.868 | 116.3303 | 20.5698 | 0.002737 | 0.00258 | 220.6516 | 0.26015 | 8.3982 | 0.002778 | 0.004573 | 0.004483 | 5.2e-05 | 0.000912 | 0.28854 | 1.3617 | 2.5476 | 98.51 | 20.03 | 49.745 | 86.9788 | 87.9853 | 83.2057 | 0.011841 | 0.99071 | 0.023026 | 23.2414 | 0.084911 | 0.15476 | 0.003154 | 0.20772 | 0.20772 | 0.030026
2019-04-09 14:53 | 1554835964.5463 | 1083 | 0.20883 | 0.00088 | 0.000168 | 89.7728 | 88.708 | 116.2876 | 20.5696 | 0.002721 | 0.002562 | 220.6521 | 0.25979 | 8.4014 | 0.002737 | 0.004354 | 0.0043 | 5.1e-05 | 0.00091 | 0.2808 | 1.3597 | 2.5476 | 98.51 | 20.03 | 49.745 | 86.6071 | 86.9873 | 82.7812 | 0.011841 | 0.99189 | 0.023035 | 23.2233 | 0.10673 | 0.15455 | 0.0035 | 0.23178 | 0.23178 | 0.026872
2019-04-09 14:53 | 1554835965.5474 | 1084 | 0.20886 | 0.000875 | 0.000166 | 89.8884 | 88.9148 | 116.4045 | 20.571 | 0.002715 | 0.002582 | 220.613 | 0.26029 | 8.4022 | 0.002699 | 0.004631 | 0.004507 | 5.5e-05 | 0.000912 | 0.30973 | 1.3333 | 2.5476 | 98.51 | 20.03 | 49.744 | 86.4662 | 86.0774 | 83.2659 | 0.011841 | 0.99054 | 0.023031 | 23.2508 | 0.097405 | 0.15485 | 0.003312 | 0.21886 | 0.21886 | 0.028512
2019-04-09 14:53 | 1554835966.5465 | 1085 | 0.20886 | 0.000873 | 0.000165 | 89.9265 | 88.8663 | 116.5353 | 20.5709 | 0.002785 | 0.002534 | 220.626 | 0.26053 | 8.4025 | 0.002697 | 0.004783 | 0.004721 | 5e-05 | 0.000914 | 0.27903 | 1.3197 | 2.5476 | 98.51 | 20.03 | 49.744 | 88.1378 | 86.0266 | 82.11 | 0.011841 | 0.99377 | 0.023081 | 23.2261 | 0.10608 | 0.15499 | 0.003322 | 0.2201 | 0.2201 | 0.028378
2019-04-09 14:53 | 1554835967.5475 | 1086 | 0.20891 | 0.000871 | 0.000164 | 90.0394 | 88.9038 | 116.2522 | 20.5716 | 0.002741 | 0.002553 | 220.6053 | 0.25953 | 8.4015 | 0.002771 | 0.004218 | 0.00416 | 5.3e-05 | 0.000918 | 0.28263 | 1.2935 | 2.5476 | 98.51 | 20.03 | 49.741 | 87.0853 | 87.8138 | 82.5693 | 0.01184 | 0.99248 | 0.023038 | 23.2129 | 0.11352 | 0.15439 | 0.002991 | 0.19696 | 0.19696 | 0.031591
2019-04-09 14:53 | 1554835968.5465 | 1087 | 0.20889 | 0.00087 | 0.000163 | 90.1376 | 89.1368 | 116.2801 | 20.5719 | 0.002701 | 0.002553 | 220.6275 | 0.26038 | 8.4051 | 0.002729 | 0.004223 | 0.004143 | 5.1e-05 | 0.000918 | 0.28672 | 1.3034 | 2.5475 | 98.51 | 20.03 | 49.74 | 86.1378 | 86.8098 | 82.5672 | 0.01184 | 0.99249 | 0.023041 | 23.2156 | 0.099871 | 0.1549 | 0.003129 | 0.20657 | 0.20657 | 0.030219
2019-04-09 14:53 | 1554835969.5476 | 1088 | 0.20887 | 0.000865 | 0.000161 | 90.2839 | 89.3454 | 116.4364 | 20.5716 | 0.002754 | 0.002593 | 220.605 | 0.26015 | 8.4033 | 0.002821 | 0.004259 | 0.004165 | 5e-05 | 0.000919 | 0.30862 | 1.3084 | 2.5476 | 98.51 | 20.03 | 49.738 | 87.4015 | 89.0126 | 83.5217 | 0.011839 | 0.98983 | 0.023026 | 23.2623 | 0.093464 | 0.15476 | 0.003277 | 0.21667 | 0.21667 | 0.028785
2019-04-09 14:53 | 1554835970.5479 | 1089 | 0.20889 | 0.000859 | 0.00016 | 90.4384 | 89.3932 | 116.1008 | 20.5713 | 0.002764 | 0.002562 | 220.6026 | 0.2603 | 8.4022 | 0.002749 | 0.003935 | 0.003823 | 5.4e-05 | 0.000921 | 0.29551 | 1.295 | 2.5474 | 98.51 | 20.03 | 49.734 | 87.6361 | 87.2848 | 82.7929 | 0.011838 | 0.99186 | 0.023016 | 23.205 | 0.10397 | 0.15485 | 0.003163 | 0.20878 | 0.20878 | 0.02989
2019-04-09 14:53 | 1554835971.547 | 1090 | 0.2089 | 0.00086 | 0.000159 | 90.5819 | 89.7859 | 116.4346 | 20.5714 | 0.002671 | 0.002579 | 220.6124 | 0.26016 | 8.4028 | 0.002842 | 0.004538 | 0.004466 | 5.5e-05 | 0.000922 | 0.29966 | 1.322 | 2.5476 | 98.51 | 20.03 | 49.733 | 85.3973 | 89.5172 | 83.1993 | 0.011838 | 0.99073 | 0.023036 | 23.2516 | 0.078955 | 0.15477 | 0.003089 | 0.20396 | 0.20396 | 0.030581
2019-04-09 14:53 | 1554835972.5456 | 1091 | 0.20887 | 0.000856 | 0.000158 | 90.6177 | 89.8003 | 116.5446 | 20.5716 | 0.00274 | 0.002596 | 220.5974 | 0.26016 | 8.4036 | 0.002743 | 0.004472 | 0.004514 | 5.2e-05 | 0.000926 | 0.30287 | 1.3456 | 2.5476 | 98.51 | 20.03 | 49.73 | 87.068 | 87.1361 | 83.5912 | 0.011837 | 0.98964 | 0.023034 | 23.2754 | 0.081043 | 0.15477 | 0.003255 | 0.21543 | 0.21543 | 0.028952
2019-04-09 14:53 | 1554835973.5461 | 1092 | 0.2089 | 0.000852 | 0.000156 | 90.6364 | 89.5952 | 116.3005 | 20.5724 | 0.002724 | 0.002481 | 220.6172 | 0.25998 | 8.4022 | 0.002751 | 0.004324 | 0.004221 | 5.2e-05 | 0.000932 | 0.30502 | 1.3287 | 2.5475 | 98.51 | 20.03 | 49.732 | 86.6815 | 87.3257 | 80.8334 | 0.011838 | 0.99735 | 0.0231 | 23.161 | 0.10335 | 0.15466 | 0.003097 | 0.20522 | 0.20522 | 0.030372
2019-04-09 14:53 | 1554835974.5457 | 1093 | 0.20889 | 0.000846 | 0.000155 | 90.738 | 89.768 | 116.5212 | 20.5712 | 0.002677 | 0.002533 | 220.5859 | 0.26046 | 8.3995 | 0.002707 | 0.00377 | 0.003757 | 5.3e-05 | 0.000936 | 0.28191 | 1.319 | 2.5475 | 98.51 | 20.03 | 49.733 | 85.5369 | 86.2759 | 82.0769 | 0.011838 | 0.99386 | 0.023081 | 23.2236 | 0.096134 | 0.15495 | 0.003167 | 0.20999 | 0.20999 | 0.029737
2019-04-09 14:53 | 1554835975.5467 | 1094 | 0.20887 | 0.000845 | 0.000154 | 90.9167 | 89.9125 | 116.3822 | 20.571 | 0.002731 | 0.002531 | 220.6049 | 0.26009 | 8.4009 | 0.002759 | 0.004193 | 0.004101 | 5.2e-05 | 0.000941 | 0.29822 | 1.342 | 2.5475 | 98.51 | 20.03 | 49.735 | 86.8464 | 87.521 | 82.0468 | 0.011838 | 0.99394 | 0.023068 | 23.2088 | 0.099338 | 0.15473 | 0.003294 | 0.21873 | 0.21873 | 0.028508
2019-04-09 14:53 | 1554835976.5472 | 1095 | 0.20894 | 0.000837 | 0.000153 | 91.0456 | 90.3293 | 116.2592 | 20.5722 | 0.0027 | 0.002535 | 220.5927 | 0.26076 | 8.4001 | 0.002748 | 0.004041 | 0.004 | 5.3e-05 | 0.000936 | 0.28208 | 1.3312 | 2.5477 | 98.51 | 20.03 | 49.737 | 86.1068 | 87.2646 | 82.1418 | 0.011839 | 0.99368 | 0.023053 | 23.1996 | 0.07065 | 0.15513 | 0.002857 | 0.18842 | 0.18842 | 0.03318
2019-04-09 14:53 | 1554835977.5461 | 1096 | 0.20892 | 0.000833 | 0.000152 | 91.2639 | 90.4986 | 116.3355 | 20.5724 | 0.002751 | 0.002526 | 220.5663 | 0.26012 | 8.4037 | 0.002768 | 0.004174 | 0.004081 | 5.2e-05 | 0.000938 | 0.28926 | 1.3605 | 2.5477 | 98.51 | 20.03 | 49.738 | 87.3289 | 87.7387 | 81.9134 | 0.011839 | 0.99432 | 0.023068 | 23.1998 | 0.075318 | 0.15474 | 0.003003 | 0.19868 | 0.19868 | 0.031389
2019-04-09 14:53 | 1554835978.5468 | 1097 | 0.20882 | 0.000833 | 0.000151 | 91.3363 | 90.5761 | 116.6479 | 20.5718 | 0.002703 | 0.002523 | 220.5695 | 0.26008 | 8.4032 | 0.002772 | 0.00448 | 0.004457 | 5.6e-05 | 0.000942 | 0.28943 | 1.3791 | 2.5477 | 98.51 | 20.03 | 49.742 | 86.1798 | 87.8467 | 81.8465 | 0.01184 | 0.9945 | 0.023101 | 23.2287 | 0.074757 | 0.15472 | 0.003605 | 0.24069 | 0.24069 | 0.025906
2019-04-09 14:53 | 1554835979.5459 | 1098 | 0.20891 | 0.000829 | 0.00015 | 91.4393 | 90.4298 | 116.4768 | 20.5725 | 0.002708 | 0.002527 | 220.565 | 0.26006 | 8.4004 | 0.002747 | 0.00466 | 0.004505 | 5.5e-05 | 0.000942 | 0.27704 | 1.3783 | 2.5476 | 98.51 | 20.03 | 49.744 | 86.2991 | 87.2306 | 81.9493 | 0.011841 | 0.99422 | 0.023081 | 23.215 | 0.099297 | 0.15471 | 0.003073 | 0.20375 | 0.20375 | 0.030601
2019-04-09 14:53 | 1554835980.5473 | 1099 | 0.20886 | 0.00083 | 0.000149 | 91.5402 | 90.7946 | 116.4904 | 20.5722 | 0.002819 | 0.002528 | 220.5843 | 0.2604 | 8.4022 | 0.002764 | 0.004485 | 0.004452 | 5.6e-05 | 0.000944 | 0.27527 | 1.3418 | 2.5476 | 98.51 | 20.03 | 49.75 | 88.9649 | 87.6512 | 81.9584 | 0.011842 | 0.99419 | 0.023082 | 23.2167 | 0.073157 | 0.15491 | 0.003371 | 0.22441 | 0.22441 | 0.027819
2019-04-09 14:53 | 1554835981.5473 | 1100 | 0.20891 | 0.000834 | 0.000148 | 91.6623 | 90.8731 | 116.4965 | 20.5726 | 0.002711 | 0.002558 | 220.5699 | 0.26019 | 8.3993 | 0.002784 | 0.004172 | 0.004009 | 5.8e-05 | 0.000945 | 0.29015 | 1.3278 | 2.5475 | 98.51 | 20.03 | 49.754 | 86.3632 | 88.1262 | 82.6857 | 0.011843 | 0.99216 | 0.023059 | 23.2411 | 0.077352 | 0.15479 | 0.003095 | 0.20519 | 0.20519 | 0.030401
2019-04-09 14:53 | 1554835982.547 | 1101 | 0.2089 | 0.000834 | 0.000147 | 91.7797 | 91.0362 | 116.2968 | 20.573 | 0.002711 | 0.002521 | 220.5484 | 0.26035 | 8.4031 | 0.002736 | 0.004208 | 0.004185 | 5.7e-05 | 0.000942 | 0.27433 | 1.317 | 2.5475 | 98.51 | 20.03 | 49.761 | 86.3605 | 86.9695 | 81.7986 | 0.011845 | 0.99464 | 0.023068 | 23.1922 | 0.072751 | 0.15488 | 0.003139 | 0.20837 | 0.20837 | 0.029955
2019-04-09 14:53 | 1554835983.5464 | 1102 | 0.20886 | 0.00083 | 0.000146 | 91.847 | 91.1078 | 116.9357 | 20.5731 | 0.00274 | 0.002497 | 220.5494 | 0.25983 | 8.4023 | 0.002767 | 0.004704 | 0.00463 | 5.8e-05 | 0.000938 | 0.29767 | 1.3149 | 2.5475 | 98.51 | 20.03 | 49.765 | 87.0568 | 87.7182 | 81.2191 | 0.011846 | 0.99626 | 0.02315 | 23.2368 | 0.072278 | 0.15457 | 0.003363 | 0.22471 | 0.22471 | 0.027722
2019-04-09 14:53 | 1554835984.5464 | 1103 | 0.20888 | 0.000833 | 0.000145 | 92.0664 | 91.3019 | 116.9599 | 20.5733 | 0.002669 | 0.002527 | 220.5375 | 0.2605 | 8.4005 | 0.002752 | 0.004033 | 0.003974 | 6.1e-05 | 0.000933 | 0.29214 | 1.3209 | 2.5475 | 98.51 | 20.03 | 49.768 | 85.3665 | 87.3663 | 81.9354 | 0.011846 | 0.99425 | 0.023129 | 23.2627 | 0.07459 | 0.15497 | 0.003247 | 0.21651 | 0.21651 | 0.028845
2019-04-09 14:53 | 1554835985.546 | 1104 | 0.20891 | 0.000834 | 0.000143 | 92.2341 | 91.5235 | 116.8622 | 20.5726 | 0.002745 | 0.002485 | 220.5443 | 0.26034 | 8.3996 | 0.002776 | 0.004048 | 0.00402 | 5.6e-05 | 0.000931 | 0.2943 | 1.291 | 2.5475 | 98.51 | 20.03 | 49.771 | 87.1931 | 87.9209 | 80.9271 | 0.011847 | 0.99709 | 0.023152 | 23.2199 | 0.069177 | 0.15488 | 0.003083 | 0.2055 | 0.2055 | 0.030373
2019-04-09 14:53 | 1554835986.5452 | 1105 | 0.20888 | 0.000834 | 0.000143 | 92.4351 | 91.7651 | 116.9613 | 20.5725 | 0.002737 | 0.002503 | 220.5142 | 0.26098 | 8.4013 | 0.002737 | 0.004882 | 0.004755 | 5.7e-05 | 0.000928 | 0.28495 | 1.3435 | 2.5475 | 98.51 | 20.03 | 49.77 | 86.9829 | 86.9839 | 81.3777 | 0.011847 | 0.99582 | 0.023147 | 23.2445 | 0.065069 | 0.15526 | 0.003275 | 0.2188 | 0.2188 | 0.028596
2019-04-09 14:53 | 1554835987.5461 | 1106 | 0.20892 | 0.000828 | 0.000142 | 92.5783 | 92.0828 | 117.0831 | 20.573 | 0.00272 | 0.002484 | 220.5251 | 0.26052 | 8.4021 | 0.002793 | 0.004432 | 0.004402 | 5.6e-05 | 0.000931 | 0.31045 | 1.3505 | 2.5476 | 98.51 | 20.03 | 49.769 | 86.5962 | 88.336 | 80.9064 | 0.011846 | 0.99714 | 0.023175 | 23.2412 | 0.048009 | 0.15499 | 0.003044 | 0.20302 | 0.20302 | 0.030764
2019-04-09 14:53 | 1554835988.5451 | 1107 | 0.20892 | 0.000823 | 0.000142 | 92.8184 | 92.0976 | 116.7754 | 20.5736 | 0.00268 | 0.002413 | 220.5235 | 0.26078 | 8.3998 | 0.002778 | 0.004032 | 0.003981 | 5.5e-05 | 0.000927 | 0.28528 | 1.3874 | 2.5477 | 98.51 | 20.03 | 49.766 | 85.6282 | 87.9703 | 79.1988 | 0.011846 | 1.002 | 0.0232 | 23.1546 | 0.069732 | 0.15514 | 0.003025 | 0.20198 | 0.20198 | 0.030954
2019-04-09 14:53 | 1554835989.5465 | 1108 | 0.20896 | 0.000825 | 0.000141 | 92.9917 | 92.3836 | 116.7277 | 20.5737 | 0.002719 | 0.002523 | 220.4862 | 0.25959 | 8.4036 | 0.002755 | 0.004412 | 0.004305 | 5.4e-05 | 0.000932 | 0.26969 | 1.335 | 2.5475 | 98.51 | 20.03 | 49.764 | 86.5572 | 87.4387 | 81.8445 | 0.011845 | 0.99451 | 0.023109 | 23.2366 | 0.058685 | 0.15443 | 0.002776 | 0.18396 | 0.18396 | 0.033831
2019-04-09 14:53 | 1554835990.5466 | 1109 | 0.20895 | 0.000816 | 0.000142 | 93.1537 | 92.7015 | 116.7179 | 20.574 | 0.002727 | 0.00253 | 220.5009 | 0.26034 | 8.3989 | 0.002683 | 0.004271 | 0.004274 | 5.5e-05 | 0.000939 | 0.2557 | 1.3635 | 2.5475 | 98.51 | 20.03 | 49.756 | 86.7551 | 85.7042 | 82.0265 | 0.011843 | 0.994 | 0.023102 | 23.2416 | 0.043518 | 0.15488 | 0.002883 | 0.1913 | 0.1913 | 0.032627
2019-04-09 14:53 | 1554835991.5452 | 1110 | 0.20894 | 0.000816 | 0.000142 | 93.4142 | 92.5629 | 116.4376 | 20.5742 | 0.002725 | 0.002538 | 220.4863 | 0.2607 | 8.4018 | 0.002733 | 0.004162 | 0.004061 | 5.6e-05 | 0.000943 | 0.29159 | 1.3618 | 2.5476 | 98.51 | 20.03 | 49.752 | 86.6958 | 86.8895 | 82.2101 | 0.011842 | 0.99349 | 0.023068 | 23.2196 | 0.081895 | 0.15509 | 0.002913 | 0.19309 | 0.19309 | 0.03237
2019-04-09 14:53 | 1554835992.5463 | 1111 | 0.20896 | 0.000811 | 0.000142 | 93.5786 | 93.0688 | 116.2843 | 20.5732 | 0.002706 | 0.002484 | 220.5101 | 0.25953 | 8.4039 | 0.002727 | 0.004478 | 0.004368 | 5.4e-05 | 0.00094 | 0.28368 | 1.3701 | 2.5476 | 98.51 | 20.03 | 49.746 | 86.25 | 86.7627 | 80.9013 | 0.011841 | 0.99716 | 0.023096 | 23.1616 | 0.048867 | 0.1544 | 0.002816 | 0.18658 | 0.18658 | 0.033348
2019-04-09 14:53 | 1554835993.5463 | 1112 | 0.20896 | 0.000806 | 0.000141 | 93.7126 | 93.2605 | 116.125 | 20.5741 | 0.002705 | 0.002537 | 220.4853 | 0.25995 | 8.4002 | 0.002716 | 0.00455 | 0.00444 | 5.4e-05 | 0.000947 | 0.30055 | 1.3356 | 2.5475 | 98.51 | 20.03 | 49.735 | 86.2178 | 86.4779 | 82.1902 | 0.011838 | 0.99354 | 0.023038 | 23.1878 | 0.043263 | 0.15464 | 0.002844 | 0.18813 | 0.18813 | 0.033126
2019-04-09 14:53 | 1554835994.5464 | 1113 | 0.20896 | 0.000799 | 0.000141 | 93.8627 | 93.152 | 116.1536 | 20.5741 | 0.002703 | 0.002505 | 220.4806 | 0.26024 | 8.4043 | 0.002741 | 0.00446 | 0.004485 | 5.2e-05 | 0.000944 | 0.29479 | 1.3424 | 2.5475 | 98.51 | 20.04 | 49.729 | 86.1695 | 87.0846 | 81.4129 | 0.011844 | 0.99572 | 0.023066 | 23.1653 | 0.06798 | 0.15482 | 0.002815 | 0.18638 | 0.18638 | 0.033476
2019-04-09 14:53 | 1554835995.5467 | 1114 | 0.20896 | 0.000795 | 0.000141 | 94.0768 | 93.6631 | 116.3029 | 20.5749 | 0.002765 | 0.002502 | 220.4795 | 0.26012 | 8.4012 | 0.002664 | 0.004301 | 0.004197 | 5e-05 | 0.000948 | 0.27975 | 1.3152 | 2.5477 | 98.51 | 20.04 | 49.715 | 87.6716 | 85.2494 | 81.356 | 0.011841 | 0.99588 | 0.023083 | 23.1783 | 0.039421 | 0.15475 | 0.002866 | 0.19 | 0.19 | 0.032823
2019-04-09 14:53 | 1554835996.5464 | 1115 | 0.20894 | 0.000794 | 0.00014 | 94.1715 | 93.7261 | 116.2996 | 20.575 | 0.002695 | 0.002509 | 220.47 | 0.26051 | 8.4016 | 0.002705 | 0.004305 | 0.00426 | 5.3e-05 | 0.000951 | 0.28174 | 1.3589 | 2.5477 | 98.51 | 20.04 | 49.709 | 85.9823 | 86.2212 | 81.5042 | 0.01184 | 0.99546 | 0.023078 | 23.1828 | 0.042408 | 0.15498 | 0.002947 | 0.19557 | 0.19557 | 0.031935
2019-04-09 14:53 | 1554835997.5458 | 1116 | 0.20894 | 0.000786 | 0.00014 | 94.3511 | 93.9336 | 116.3294 | 20.575 | 0.002737 | 0.002538 | 220.4916 | 0.26106 | 8.3996 | 0.002719 | 0.00439 | 0.004343 | 5.8e-05 | 0.000948 | 0.27776 | 1.3264 | 2.5475 | 98.51 | 20.04 | 49.703 | 86.9964 | 86.5592 | 82.2099 | 0.011838 | 0.99349 | 0.023058 | 23.2088 | 0.039665 | 0.1553 | 0.002955 | 0.19601 | 0.19601 | 0.03193
2019-04-09 14:53 | 1554835998.5459 | 1117 | 0.20896 | 0.00079 | 0.000141 | 94.5217 | 94.2502 | 116.2149 | 20.5753 | 0.002727 | 0.002495 | 220.4829 | 0.26052 | 8.4007 | 0.00269 | 0.004555 | 0.004432 | 5.5e-05 | 0.000951 | 0.30055 | 1.3155 | 2.5475 | 98.51 | 20.04 | 49.693 | 86.7477 | 85.8639 | 81.1862 | 0.011836 | 0.99636 | 0.02308 | 23.164 | 0.025728 | 0.15498 | 0.002821 | 0.18682 | 0.18682 | 0.033433
2019-04-09 14:53 | 1554835999.5455 | 1118 | 0.20899 | 0.000783 | 0.000141 | 94.645 | 94.3951 | 116.1304 | 20.5751 | 0.002703 | 0.002503 | 220.4813 | 0.26082 | 8.3983 | 0.002715 | 0.004698 | 0.004682 | 5.4e-05 | 0.000948 | 0.29103 | 1.3124 | 2.5477 | 98.51 | 20.04 | 49.688 | 86.1722 | 86.4679 | 81.3747 | 0.011835 | 0.99583 | 0.023065 | 23.1617 | 0.023647 | 0.15516 | 0.002696 | 0.17809 | 0.17809 | 0.035112
2019-04-09 14:53 | 1554836000.5454 | 1119 | 0.20895 | 0.000786 | 0.000141 | 94.7638 | 94.6429 | 116.0318 | 20.5757 | 0.002667 | 0.002501 | 220.4777 | 0.26002 | 8.4013 | 0.002767 | 0.004646 | 0.004518 | 5.6e-05 | 0.000951 | 0.30807 | 1.337 | 2.5476 | 98.51 | 20.04 | 49.684 | 85.3016 | 87.7161 | 81.3185 | 0.011834 | 0.99599 | 0.023057 | 23.15 | 0.01142 | 0.15469 | 0.002915 | 0.19321 | 0.19321 | 0.032264
2019-04-09 14:53 | 1554836001.5446 | 1120 | 0.20901 | 0.000783 | 0.000139 | 94.8749 | 94.5056 | 116.3963 | 20.5752 | 0.002706 | 0.002518 | 220.446 | 0.26033 | 8.402 | 0.002754 | 0.004075 | 0.004039 | 5.7e-05 | 0.000951 | 0.29319 | 1.3368 | 2.5476 | 98.51 | 20.04 | 49.68 | 86.2471 | 87.4135 | 81.7205 | 0.011833 | 0.99486 | 0.02308 | 23.1995 | 0.034885 | 0.15487 | 0.002544 | 0.16778 | 0.16778 | 0.037199
2019-04-09 14:53 | 1554836002.5446 | 1121 | 0.20894 | 0.000775 | 0.000138 | 94.9595 | 94.541 | 116.6801 | 20.5754 | 0.002618 | 0.002522 | 220.4546 | 0.25992 | 8.4028 | 0.00269 | 0.003923 | 0.003892 | 5.8e-05 | 0.000955 | 0.30453 | 1.3341 | 2.5476 | 98.51 | 20.04 | 49.676 | 84.1395 | 85.8648 | 81.818 | 0.011832 | 0.99458 | 0.023105 | 23.231 | 0.0395 | 0.15463 | 0.002991 | 0.19902 | 0.19902 | 0.031311
2019-04-09 14:53 | 1554836003.5456 | 1122 | 0.20909 | 0.000771 | 0.000137 | 95.0536 | 94.8511 | 116.6763 | 20.5756 | 0.002712 | 0.002478 | 220.4406 | 0.26066 | 8.4021 | 0.002728 | 0.004499 | 0.004385 | 5.7e-05 | 0.000948 | 0.26753 | 1.3433 | 2.5476 | 98.51 | 20.04 | 49.675 | 86.3866 | 86.773 | 80.7741 | 0.011831 | 0.99752 | 0.023139 | 23.1964 | 0.019071 | 0.15507 | 0.002106 | 0.13803 | 0.13803 | 0.045272
2019-04-09 14:53 | 1554836004.5446 | 1123 | 0.20901 | 0.000773 | 0.000137 | 95.2123 | 94.9286 | 116.6355 | 20.5767 | 0.002738 | 0.002511 | 220.4331 | 0.26046 | 8.399 | 0.002721 | 0.00449 | 0.004392 | 5.7e-05 | 0.000955 | 0.29679 | 1.3778 | 2.5476 | 98.51 | 20.04 | 49.674 | 87.0332 | 86.6099 | 81.551 | 0.011831 | 0.99533 | 0.023109 | 23.2178 | 0.026695 | 0.15495 | 0.002597 | 0.17182 | 0.17182 | 0.036341
2019-04-09 14:53 | 1554836005.5447 | 1124 | 0.20902 | 0.000768 | 0.000136 | 95.3657 | 95.1423 | 116.5442 | 20.576 | 0.002706 | 0.002516 | 220.4249 | 0.26028 | 8.4028 | 0.002725 | 0.005063 | 0.004968 | 5.8e-05 | 0.000956 | 0.30596 | 1.3185 | 2.5475 | 98.51 | 20.04 | 49.674 | 86.2495 | 86.7194 | 81.6851 | 0.011831 | 0.99496 | 0.023096 | 23.2131 | 0.020976 | 0.15484 | 0.002516 | 0.16621 | 0.16621 | 0.037543
2019-04-09 14:53 | 1554836006.5461 | 1125 | 0.209 | 0.000767 | 0.000135 | 95.4884 | 95.2934 | 116.5686 | 20.5767 | 0.002688 | 0.002546 | 220.4146 | 0.26052 | 8.4029 | 0.002698 | 0.004475 | 0.004472 | 5.7e-05 | 0.000951 | 0.27345 | 1.3194 | 2.5475 | 98.51 | 20.04 | 49.675 | 85.8093 | 86.0542 | 82.4027 | 0.011831 | 0.99295 | 0.023075 | 23.239 | 0.018284 | 0.15498 | 0.002637 | 0.17443 | 0.17443 | 0.035806
2019-04-09 14:53 | 1554836007.5447 | 1126 | 0.20901 | 0.000766 | 0.000135 | 95.6122 | 95.4596 | 116.5975 | 20.5764 | 0.002647 | 0.002438 | 220.419 | 0.26075 | 8.4028 | 0.0027 | 0.00359 | 0.003441 | 5.6e-05 | 0.000951 | 0.29159 | 1.3576 | 2.5477 | 98.51 | 20.04 | 49.679 | 84.8403 | 86.1078 | 79.8081 | 0.011832 | 1.0002 | 0.023163 | 23.1569 | 0.014291 | 0.15512 | 0.002568 | 0.17031 | 0.17031 | 0.036706
2019-04-09 14:53 | 1554836008.5448 | 1127 | 0.20901 | 0.000765 | 0.000135 | 95.7325 | 95.5866 | 116.9667 | 20.576 | 0.002619 | 0.002495 | 220.3702 | 0.25973 | 8.403 | 0.002716 | 0.003868 | 0.003828 | 5.7e-05 | 0.00096 | 0.28672 | 1.3382 | 2.5476 | 98.51 | 20.04 | 49.681 | 84.1627 | 86.4976 | 81.1886 | 0.011833 | 0.99635 | 0.023154 | 23.2389 | 0.013648 | 0.15451 | 0.002596 | 0.1722 | 0.1722 | 0.03616
2019-04-09 14:53 | 1554836009.5448 | 1128 | 0.20901 | 0.000761 | 0.000135 | 95.8055 | 95.7869 | 117.2974 | 20.5763 | 0.002676 | 0.002513 | 220.3895 | 0.26042 | 8.4017 | 0.002741 | 0.004429 | 0.004336 | 5.5e-05 | 0.000956 | 0.30127 | 1.2959 | 2.5476 | 98.51 | 20.04 | 49.684 | 85.5207 | 87.1024 | 81.6201 | 0.011834 | 0.99514 | 0.023173 | 23.2859 | 0.001739 | 0.15492 | 0.002589 | 0.17187 | 0.17187 | 0.036326
2019-04-09 14:53 | 1554836010.5458 | 1129 | 0.20901 | 0.000759 | 0.000135 | 95.8624 | 95.738 | 117.4901 | 20.5758 | 0.002669 | 0.002531 | 220.3849 | 0.26029 | 8.4022 | 0.00269 | 0.004319 | 0.004241 | 5.6e-05 | 0.000955 | 0.27632 | 1.3093 | 2.5476 | 98.51 | 20.05 | 49.691 | 85.372 | 85.8693 | 82.0525 | 0.011843 | 0.99393 | 0.023178 | 23.3192 | 0.011612 | 0.15485 | 0.002595 | 0.17236 | 0.17236 | 0.036205
2019-04-09 14:54 | 1554836011.5455 | 1130 | 0.20902 | 0.000754 | 0.000134 | 95.9758 | 95.8791 | 117.1638 | 20.5776 | 0.002714 | 0.00243 | 220.3962 | 0.26063 | 8.4015 | 0.002765 | 0.004927 | 0.004877 | 6.1e-05 | 0.000944 | 0.27721 | 1.3393 | 2.5476 | 98.51 | 20.05 | 49.696 | 86.4539 | 87.6641 | 79.6158 | 0.011844 | 1.0008 | 0.023225 | 23.2068 | 0.009019 | 0.15505 | 0.002558 | 0.17013 | 0.17013 | 0.036727
2019-04-09 14:54 | 1554836012.5459 | 1131 | 0.20899 | 0.000755 | 0.000133 | 96.0327 | 95.832 | 116.8535 | 20.5765 | 0.00268 | 0.00251 | 220.3842 | 0.25978 | 8.4039 | 0.002804 | 0.004674 | 0.004632 | 5.7e-05 | 0.00094 | 0.29551 | 1.3365 | 2.5476 | 98.51 | 20.05 | 49.701 | 85.6224 | 88.6019 | 81.5362 | 0.011845 | 0.99537 | 0.023132 | 23.239 | 0.018718 | 0.15454 | 0.002704 | 0.17962 | 0.17962 | 0.034673
2019-04-09 14:54 | 1554836013.546 | 1132 | 0.20898 | 0.000758 | 0.000132 | 96.1542 | 96.0901 | 116.9402 | 20.577 | 0.002665 | 0.002433 | 220.373 | 0.26115 | 8.4019 | 0.002787 | 0.004309 | 0.004239 | 5.9e-05 | 0.000941 | 0.2589 | 1.3508 | 2.5476 | 98.51 | 20.04 | 49.708 | 85.2765 | 88.2039 | 79.6784 | 0.011839 | 1.0006 | 0.023201 | 23.1867 | 0.005968 | 0.15536 | 0.002798 | 0.18679 | 0.18679 | 0.033519
2019-04-09 14:54 | 1554836014.5461 | 1133 | 0.20903 | 0.000751 | 0.000131 | 96.2179 | 96.1644 | 116.9357 | 20.5771 | 0.002701 | 0.002476 | 220.3931 | 0.2607 | 8.4009 | 0.00272 | 0.004087 | 0.003975 | 5.7e-05 | 0.000944 | 0.29518 | 1.317 | 2.5475 | 98.51 | 20.04 | 49.713 | 86.1209 | 86.5769 | 80.7123 | 0.011841 | 0.99769 | 0.023167 | 23.2202 | 0.004977 | 0.15509 | 0.002468 | 0.16361 | 0.16361 | 0.038202
2019-04-09 14:54 | 1554836015.546 | 1134 | 0.209 | 0.000751 | 0.000131 | 96.2477 | 96.1909 | 116.8376 | 20.5779 | 0.002602 | 0.002485 | 220.3822 | 0.2605 | 8.4031 | 0.002674 | 0.004487 | 0.004481 | 5.6e-05 | 0.000949 | 0.27577 | 1.3406 | 2.5476 | 98.51 | 20.04 | 49.722 | 83.762 | 85.4895 | 80.9431 | 0.011843 | 0.99704 | 0.023149 | 23.218 | 0.00528 | 0.15497 | 0.002677 | 0.17799 | 0.17799 | 0.035088
2019-04-09 14:54 | 1554836016.5464 | 1135 | 0.209 | 0.000745 | 0.000131 | 96.3788 | 96.2508 | 116.4451 | 20.5778 | 0.002712 | 0.002474 | 220.3922 | 0.261 | 8.4017 | 0.002672 | 0.004373 | 0.004288 | 5.8e-05 | 0.000945 | 0.26698 | 1.3197 | 2.5476 | 98.51 | 20.04 | 49.726 | 86.3901 | 85.4254 | 80.6724 | 0.011844 | 0.9978 | 0.023119 | 23.1701 | 0.011882 | 0.15527 | 0.002659 | 0.17654 | 0.17654 | 0.035445
2019-04-09 14:54 | 1554836017.5461 | 1136 | 0.20902 | 0.000746 | 0.000132 | 96.4032 | 96.342 | 116.4611 | 20.5772 | 0.002667 | 0.002506 | 220.3839 | 0.26069 | 8.4011 | 0.002685 | 0.003686 | 0.003665 | 5.9e-05 | 0.000951 | 0.28694 | 1.3179 | 2.5476 | 98.51 | 20.04 | 49.731 | 85.3149 | 85.7549 | 81.4322 | 0.011845 | 0.99567 | 0.023096 | 23.1965 | 0.005677 | 0.15508 | 0.002578 | 0.1707 | 0.1707 | 0.036612
2019-04-09 14:54 | 1554836018.5451 | 1137 | 0.20902 | 0.000745 | 0.000133 | 96.4973 | 96.4766 | 116.243 | 20.5777 | 0.002713 | 0.002517 | 220.3938 | 0.26126 | 8.403 | 0.002784 | 0.004439 | 0.004331 | 5.9e-05 | 0.000953 | 0.28888 | 1.3439 | 2.5477 | 98.51 | 20.04 | 49.739 | 86.4112 | 88.1281 | 81.7191 | 0.011847 | 0.99486 | 0.023065 | 23.1842 | 0.00192 | 0.15542 | 0.002558 | 0.16903 | 0.16903 | 0.037055
2019-04-09 14:54 | 1554836019.5441 | 1138 | 0.20901 | 0.000743 | 0.000133 | 96.5011 | 96.5189 | 115.871 | 20.5783 | 0.002688 | 0.002499 | 220.4104 | 0.2605 | 8.4013 | 0.002709 | 0.004154 | 0.004121 | 5.7e-05 | 0.000949 | 0.28512 | 1.3561 | 2.5476 | 98.51 | 20.04 | 49.742 | 85.8136 | 86.322 | 81.2694 | 0.011847 | 0.99612 | 0.023043 | 23.1324 | -0.001646 | 0.15497 | 0.002616 | 0.17285 | 0.17285 | 0.036132
2019-04-09 14:54 | 1554836020.5452 | 1139 | 0.20903 | 0.000746 | 0.000133 | 96.6471 | 96.5454 | 115.7992 | 20.5775 | 0.002691 | 0.002532 | 220.3634 | 0.26076 | 8.4003 | 0.002657 | 0.004675 | 0.004618 | 5.7e-05 | 0.000947 | 0.31238 | 1.3745 | 2.5477 | 98.51 | 20.05 | 49.746 | 85.8942 | 85.0614 | 82.0625 | 0.011856 | 0.9939 | 0.02301 | 23.1511 | 0.009414 | 0.15513 | 0.002475 | 0.16289 | 0.16289 | 0.03838
2019-04-09 14:54 | 1554836021.5442 | 1140 | 0.20902 | 0.000745 | 0.000133 | 96.5815 | 96.6528 | 116.1501 | 20.5777 | 0.002665 | 0.002512 | 220.3787 | 0.26082 | 8.402 | 0.002685 | 0.004402 | 0.004338 | 5.5e-05 | 0.000952 | 0.26665 | 1.3345 | 2.5476 | 98.51 | 20.04 | 49.747 | 85.2748 | 85.7453 | 81.5878 | 0.011849 | 0.99523 | 0.02306 | 23.1706 | -0.006594 | 0.15516 | 0.002539 | 0.1677 | 0.1677 | 0.037287
2019-04-09 14:54 | 1554836022.5446 | 1141 | 0.20898 | 0.000742 | 0.000132 | 96.676 | 96.6373 | 116.4191 | 20.5784 | 0.002659 | 0.002464 | 220.377 | 0.26048 | 8.4032 | 0.002752 | 0.004321 | 0.004221 | 5.8e-05 | 0.000951 | 0.24889 | 1.3688 | 2.5477 | 98.51 | 20.04 | 49.749 | 85.1319 | 87.3612 | 80.4274 | 0.011849 | 0.9985 | 0.023125 | 23.1595 | 0.003578 | 0.15496 | 0.002782 | 0.18507 | 0.18507 | 0.033743
2019-04-09 14:54 | 1554836023.5453 | 1142 | 0.20897 | 0.000746 | 0.000131 | 96.7159 | 96.8162 | 116.2555 | 20.5783 | 0.002671 | 0.002534 | 220.357 | 0.26097 | 8.4023 | 0.002795 | 0.003884 | 0.003829 | 5.6e-05 | 0.000949 | 0.28064 | 1.3614 | 2.5475 | 98.51 | 20.04 | 49.749 | 85.415 | 88.3958 | 82.122 | 0.011849 | 0.99373 | 0.023053 | 23.1986 | -0.00928 | 0.15525 | 0.002884 | 0.19152 | 0.19152 | 0.032668
2019-04-09 14:54 | 1554836024.5453 | 1143 | 0.20894 | 0.000746 | 0.000132 | 96.7912 | 96.724 | 115.9801 | 20.5766 | 0.002647 | 0.00253 | 220.3547 | 0.26083 | 8.4032 | 0.002722 | 0.00392 | 0.003833 | 5.9e-05 | 0.000945 | 0.28208 | 1.3473 | 2.5476 | 98.51 | 20.04 | 49.749 | 84.8412 | 86.631 | 82.0071 | 0.011849 | 0.99405 | 0.02303 | 23.1674 | 0.00621 | 0.15517 | 0.003026 | 0.20114 | 0.20114 | 0.031089
2019-04-09 14:54 | 1554836025.5447 | 1144 | 0.20897 | 0.000746 | 0.000132 | 96.7903 | 96.9615 | 115.9367 | 20.5791 | 0.002614 | 0.002474 | 220.3329 | 0.26077 | 8.4015 | 0.00272 | 0.003905 | 0.003882 | 5.5e-05 | 0.00094 | 0.28296 | 1.332 | 2.5475 | 98.51 | 20.05 | 49.748 | 84.0347 | 86.6008 | 80.6761 | 0.011856 | 0.99779 | 0.023069 | 23.1196 | -0.015805 | 0.15513 | 0.002839 | 0.18856 | 0.18856 | 0.033156
2019-04-09 14:54 | 1554836026.5454 | 1145 | 0.20898 | 0.00074 | 0.000132 | 96.9455 | 96.7822 | 116.3472 | 20.5796 | 0.002663 | 0.002459 | 220.3401 | 0.26059 | 8.4022 | 0.00268 | 0.004564 | 0.004524 | 5.7e-05 | 0.000935 | 0.2808 | 1.338 | 2.5477 | 98.51 | 20.05 | 49.746 | 85.2265 | 85.6364 | 80.3232 | 0.011856 | 0.99879 | 0.023121 | 23.1489 | 0.015077 | 0.15502 | 0.002797 | 0.18606 | 0.18606 | 0.033578
2019-04-09 14:54 | 1554836027.545 | 1146 | 0.209 | 0.000746 | 0.000131 | 97.0074 | 97.0785 | 116.2555 | 20.579 | 0.002634 | 0.00248 | 220.2965 | 0.26048 | 8.4021 | 0.002709 | 0.003977 | 0.0039 | 5.5e-05 | 0.000936 | 0.27367 | 1.3328 | 2.5475 | 98.51 | 20.05 | 49.744 | 84.5253 | 86.3328 | 80.8108 | 0.011855 | 0.99741 | 0.023096 | 23.1557 | -0.006553 | 0.15496 | 0.002673 | 0.1773 | 0.1773 | 0.035221
2019-04-09 14:54 | 1554836028.5454 | 1147 | 0.20895 | 0.000745 | 0.00013 | 97.1075 | 97.1305 | 116.6017 | 20.578 | 0.002674 | 0.002529 | 220.3012 | 0.26055 | 8.4013 | 0.002748 | 0.004551 | 0.004408 | 5.4e-05 | 0.00093 | 0.2756 | 1.3389 | 2.5475 | 98.51 | 20.05 | 49.737 | 85.4771 | 87.2592 | 81.9845 | 0.011854 | 0.99412 | 0.023092 | 23.2286 | -0.002119 | 0.155 | 0.002967 | 0.1977 | 0.1977 | 0.031596
2019-04-09 14:54 | 1554836029.5451 | 1148 | 0.209 | 0.000747 | 0.00013 | 97.212 | 97.3583 | 116.7582 | 20.579 | 0.002675 | 0.002529 | 220.2847 | 0.2609 | 8.4002 | 0.002722 | 0.004004 | 0.00392 | 5.3e-05 | 0.000932 | 0.2808 | 1.3701 | 2.5476 | 98.51 | 20.05 | 49.733 | 85.5149 | 86.6372 | 82.0025 | 0.011853 | 0.99407 | 0.023107 | 23.2448 | -0.013452 | 0.15521 | 0.002665 | 0.17697 | 0.17697 | 0.035344
2019-04-09 14:54 | 1554836030.5455 | 1149 | 0.20902 | 0.000743 | 0.000129 | 97.3174 | 97.3294 | 116.6484 | 20.5792 | 0.002655 | 0.002517 | 220.2568 | 0.2606 | 8.4024 | 0.0028 | 0.004544 | 0.004491 | 5.6e-05 | 0.00093 | 0.3027 | 1.3046 | 2.5476 | 98.51 | 20.05 | 49.725 | 85.0224 | 88.517 | 81.7149 | 0.011851 | 0.99487 | 0.023105 | 23.2244 | -0.001102 | 0.15503 | 0.002553 | 0.16922 | 0.16922 | 0.036922
2019-04-09 14:54 | 1554836031.5459 | 1150 | 0.20903 | 0.000739 | 0.000129 | 97.3501 | 97.4994 | 116.435 | 20.5783 | 0.002593 | 0.002433 | 220.2588 | 0.26025 | 8.4011 | 0.002758 | 0.004306 | 0.00423 | 5.6e-05 | 0.000935 | 0.28727 | 1.309 | 2.5476 | 98.51 | 20.05 | 49.721 | 83.5462 | 87.4922 | 79.6998 | 0.01185 | 1.0006 | 0.02315 | 23.1372 | -0.013702 | 0.15482 | 0.002516 | 0.16703 | 0.16703 | 0.037354
2019-04-09 14:54 | 1554836032.5445 | 1151 | 0.20903 | 0.000743 | 0.00013 | 97.4401 | 97.3972 | 116.6174 | 20.5787 | 0.002649 | 0.002486 | 220.2444 | 0.26084 | 8.4017 | 0.002735 | 0.003678 | 0.003704 | 5.3e-05 | 0.000935 | 0.25265 | 1.2906 | 2.5476 | 98.51 | 20.05 | 49.718 | 84.8731 | 86.9423 | 80.9595 | 0.011849 | 0.99699 | 0.023127 | 23.1966 | 0.003939 | 0.15517 | 0.002523 | 0.16725 | 0.16725 | 0.037389
2019-04-09 14:54 | 1554836033.5446 | 1152 | 0.20904 | 0.000736 | 0.00013 | 97.5589 | 97.7076 | 116.7106 | 20.5796 | 0.0027 | 0.002462 | 220.2806 | 0.26004 | 8.4009 | 0.002741 | 0.00417 | 0.004038 | 5.7e-05 | 0.000938 | 0.27472 | 1.3277 | 2.5477 | 98.51 | 20.05 | 49.712 | 86.1178 | 87.0976 | 80.3833 | 0.011848 | 0.99862 | 0.023155 | 23.187 | -0.013623 | 0.1547 | 0.002467 | 0.16359 | 0.16359 | 0.038109
2019-04-09 14:54 | 1554836034.5436 | 1153 | 0.20902 | 0.000736 | 0.00013 | 97.5967 | 97.7971 | 116.8683 | 20.5789 | 0.002554 | 0.002547 | 220.281 | 0.26087 | 8.4015 | 0.002747 | 0.004651 | 0.004601 | 5.6e-05 | 0.000934 | 0.28174 | 1.3279 | 2.5475 | 98.51 | 20.06 | 49.71 | 82.5959 | 87.2391 | 82.4391 | 0.011855 | 0.99285 | 0.023104 | 23.27 | -0.018344 | 0.15519 | 0.00259 | 0.17171 | 0.17171 | 0.036423
2019-04-09 14:54 | 1554836035.5439 | 1154 | 0.20898 | 0.000732 | 0.00013 | 97.6539 | 97.816 | 116.6984 | 20.5793 | 0.002596 | 0.002477 | 220.2774 | 0.26068 | 8.4016 | 0.002722 | 0.004181 | 0.004112 | 5.4e-05 | 0.000938 | 0.27831 | 1.3357 | 2.5475 | 98.51 | 20.06 | 49.709 | 83.6042 | 86.6298 | 80.744 | 0.011854 | 0.9976 | 0.023142 | 23.1976 | -0.014837 | 0.15508 | 0.002797 | 0.18637 | 0.18637 | 0.033533
2019-04-09 14:54 | 1554836036.5438 | 1155 | 0.20904 | 0.000729 | 0.000129 | 97.7382 | 97.9279 | 116.5667 | 20.5793 | 0.002635 | 0.002487 | 220.2817 | 0.26028 | 8.4016 | 0.00271 | 0.004201 | 0.004105 | 5.2e-05 | 0.00094 | 0.29302 | 1.3585 | 2.5477 | 98.51 | 20.06 | 49.71 | 84.5434 | 86.3564 | 80.999 | 0.011855 | 0.99688 | 0.023121 | 23.1929 | -0.017343 | 0.15484 | 0.002448 | 0.16203 | 0.16203 | 0.038512
2019-04-09 14:54 | 1554836037.5447 | 1156 | 0.20904 | 0.000733 | 0.00013 | 97.7423 | 97.7599 | 116.8019 | 20.5793 | 0.002672 | 0.002514 | 220.2676 | 0.2605 | 8.4009 | 0.002726 | 0.003536 | 0.00348 | 5.3e-05 | 0.000941 | 0.3235 | 1.3312 | 2.5474 | 98.51 | 20.06 | 49.71 | 85.4308 | 86.7328 | 81.6425 | 0.011855 | 0.99508 | 0.023123 | 23.2374 | -0.001608 | 0.15497 | 0.002437 | 0.16129 | 0.16129 | 0.038723
2019-04-09 14:54 | 1554836038.5442 | 1157 | 0.20905 | 0.000726 | 0.000128 | 97.8394 | 97.9131 | 116.463 | 20.5801 | 0.002687 | 0.00245 | 220.2799 | 0.26054 | 8.401 | 0.002785 | 0.004724 | 0.004595 | 5.1e-05 | 0.000944 | 0.29192 | 1.3435 | 2.5477 | 98.51 | 20.06 | 49.711 | 85.7858 | 88.1602 | 80.0985 | 0.011855 | 0.99943 | 0.02314 | 23.1531 | -0.00673 | 0.155 | 0.002438 | 0.16152 | 0.16152 | 0.038673
2019-04-09 14:54 | 1554836039.5442 | 1158 | 0.20904 | 0.000729 | 0.000128 | 97.8963 | 97.9533 | 116.5613 | 20.5791 | 0.00255 | 0.002459 | 220.2568 | 0.26016 | 8.4021 | 0.0027 | 0.004431 | 0.004337 | 5.2e-05 | 0.000943 | 0.28584 | 1.3999 | 2.5477 | 98.51 | 20.06 | 49.715 | 82.4916 | 86.1167 | 80.3074 | 0.011856 | 0.99883 | 0.023143 | 23.1697 | -0.005209 | 0.15477 | 0.002487 | 0.16494 | 0.16494 | 0.037815
2019-04-09 14:54 | 1554836040.5438 | 1159 | 0.20903 | 0.00073 | 0.000128 | 97.8862 | 97.9447 | 116.6134 | 20.5799 | 0.002601 | 0.002495 | 220.2753 | 0.25954 | 8.4032 | 0.002728 | 0.004097 | 0.004143 | 4.7e-05 | 0.000941 | 0.28528 | 1.3807 | 2.5477 | 98.51 | 20.06 | 49.718 | 83.7334 | 86.7824 | 81.1854 | 0.011856 | 0.99636 | 0.023119 | 23.2036 | -0.005349 | 0.1544 | 0.002547 | 0.16895 | 0.16895 | 0.036829
2019-04-09 14:54 | 1554836041.5448 | 1160 | 0.20906 | 0.000729 | 0.000129 | 97.9002 | 98.0432 | 116.7303 | 20.5799 | 0.00266 | 0.002529 | 220.2625 | 0.26078 | 8.4031 | 0.002732 | 0.00399 | 0.003931 | 4.9e-05 | 0.000942 | 0.30287 | 1.3728 | 2.5476 | 98.51 | 20.06 | 49.724 | 85.1395 | 86.8721 | 81.9987 | 0.011858 | 0.99408 | 0.023104 | 23.2419 | -0.013054 | 0.15514 | 0.002359 | 0.15579 | 0.15579 | 0.040131
2019-04-09 14:54 | 1554836042.5439 | 1161 | 0.20906 | 0.00072 | 0.000129 | 97.9401 | 97.9461 | 116.5698 | 20.5797 | 0.002671 | 0.002444 | 220.2346 | 0.2602 | 8.4024 | 0.002698 | 0.004282 | 0.004154 | 4.8e-05 | 0.000944 | 0.28495 | 1.3644 | 2.5475 | 98.51 | 20.06 | 49.727 | 85.4155 | 86.0508 | 79.9485 | 0.011859 | 0.99985 | 0.023155 | 23.1588 | -0.000551 | 0.15479 | 0.002346 | 0.15521 | 0.15521 | 0.040193
2019-04-09 14:54 | 1554836043.5447 | 1162 | 0.20905 | 0.000724 | 0.000129 | 97.9937 | 98.2245 | 116.4238 | 20.5802 | 0.002645 | 0.002476 | 220.2335 | 0.26151 | 8.3999 | 0.002717 | 0.004555 | 0.004484 | 5.1e-05 | 0.000938 | 0.29856 | 1.3617 | 2.5477 | 98.51 | 20.06 | 49.732 | 84.7908 | 86.5298 | 80.7302 | 0.01186 | 0.99764 | 0.023115 | 23.1699 | -0.021046 | 0.15557 | 0.002405 | 0.15908 | 0.15908 | 0.039413
2019-04-09 14:54 | 1554836044.545 | 1163 | 0.20906 | 0.000721 | 0.000129 | 97.9833 | 98.1116 | 116.4536 | 20.5794 | 0.002681 | 0.002462 | 220.204 | 0.26108 | 8.402 | 0.002693 | 0.0039 | 0.0039 | 4.7e-05 | 0.000937 | 0.28838 | 1.2931 | 2.5474 | 98.51 | 20.06 | 49.735 | 85.6579 | 85.943 | 80.3949 | 0.01186 | 0.99859 | 0.023129 | 23.1618 | -0.011712 | 0.15532 | 0.002362 | 0.15613 | 0.15613 | 0.040091
2019-04-09 14:54 | 1554836045.5443 | 1164 | 0.20906 | 0.000724 | 0.00013 | 98.0535 | 98.156 | 116.8406 | 20.5807 | 0.002641 | 0.002469 | 220.2251 | 0.26042 | 8.4008 | 0.002751 | 0.004499 | 0.004407 | 4.8e-05 | 0.000942 | 0.29767 | 1.327 | 2.5476 | 98.51 | 20.06 | 49.735 | 84.6821 | 87.3359 | 80.5452 | 0.01186 | 0.99816 | 0.023163 | 23.2052 | -0.009346 | 0.15492 | 0.00234 | 0.15485 | 0.15485 | 0.040319
2019-04-09 14:54 | 1554836046.5454 | 1165 | 0.20907 | 0.000725 | 0.000129 | 98.1843 | 98.2514 | 117.112 | 20.5808 | 0.002599 | 0.002536 | 220.2146 | 0.2604 | 8.4024 | 0.002704 | 0.004032 | 0.004073 | 4.8e-05 | 0.000943 | 0.28208 | 1.3381 | 2.5475 | 98.51 | 20.06 | 49.737 | 83.6859 | 86.1996 | 82.179 | 0.011861 | 0.99357 | 0.023136 | 23.2858 | -0.006111 | 0.15491 | 0.00226 | 0.14917 | 0.14917 | 0.041853
2019-04-09 14:54 | 1554836047.5443 | 1166 | 0.20909 | 0.000723 | 0.000129 | 98.2715 | 98.4493 | 116.8106 | 20.5797 | 0.002589 | 0.002493 | 220.2278 | 0.26108 | 8.4025 | 0.002708 | 0.00455 | 0.00439 | 4.8e-05 | 0.000947 | 0.29623 | 1.3373 | 2.5476 | 98.51 | 20.06 | 49.736 | 83.4471 | 86.3053 | 81.1366 | 0.011861 | 0.9965 | 0.02314 | 23.2216 | -0.016163 | 0.15531 | 0.002172 | 0.14308 | 0.14308 | 0.043744
2019-04-09 14:54 | 1554836048.5451 | 1167 | 0.20911 | 0.000721 | 0.000129 | 98.2176 | 98.3729 | 116.7493 | 20.5809 | 0.002655 | 0.002398 | 220.2344 | 0.25988 | 8.4024 | 0.002732 | 0.004283 | 0.004206 | 4.6e-05 | 0.000948 | 0.3095 | 1.333 | 2.5478 | 98.51 | 20.06 | 49.734 | 85.021 | 86.8756 | 78.8488 | 0.01186 | 1.003 | 0.023209 | 23.1405 | -0.014126 | 0.1546 | 0.002042 | 0.13449 | 0.13449 | 0.046326
2019-04-09 14:54 | 1554836049.5441 | 1168 | 0.20909 | 0.000722 | 0.000127 | 98.2542 | 98.4389 | 116.6329 | 20.5809 | 0.002641 | 0.002479 | 220.2283 | 0.26102 | 8.4023 | 0.002748 | 0.004278 | 0.004281 | 4.4e-05 | 0.000943 | 0.3115 | 1.3731 | 2.5475 | 98.51 | 20.06 | 49.731 | 84.7018 | 87.2714 | 80.7925 | 0.01186 | 0.99747 | 0.023134 | 23.1927 | -0.016797 | 0.15528 | 0.002147 | 0.14139 | 0.14139 | 0.044261
2019-04-09 14:54 | 1554836050.5432 | 1169 | 0.20906 | 0.000723 | 0.000127 | 98.315 | 98.5831 | 116.8674 | 20.5818 | 0.002671 | 0.002472 | 220.2558 | 0.26088 | 8.3993 | 0.002713 | 0.005019 | 0.004989 | 4.5e-05 | 0.000939 | 0.28174 | 1.3336 | 2.5475 | 98.51 | 20.06 | 49.73 | 85.4248 | 86.4145 | 80.6318 | 0.011859 | 0.99792 | 0.023162 | 23.2107 | -0.024362 | 0.1552 | 0.00233 | 0.15428 | 0.15428 | 0.040541
2019-04-09 14:54 | 1554836051.5441 | 1170 | 0.20905 | 0.000726 | 0.000127 | 98.2715 | 98.3488 | 116.7329 | 20.5813 | 0.002649 | 0.002445 | 220.251 | 0.26049 | 8.4015 | 0.002728 | 0.003854 | 0.003736 | 4.6e-05 | 0.000941 | 0.2568 | 1.3419 | 2.5476 | 98.51 | 20.06 | 49.726 | 84.8965 | 86.791 | 79.9875 | 0.011858 | 0.99974 | 0.02317 | 23.1762 | -0.007028 | 0.15496 | 0.002427 | 0.16108 | 0.16108 | 0.03877
2019-04-09 14:54 | 1554836052.5431 | 1171 | 0.2091 | 0.000723 | 0.000127 | 98.3531 | 98.7256 | 116.5592 | 20.5804 | 0.002672 | 0.002496 | 220.261 | 0.2608 | 8.4019 | 0.002735 | 0.004952 | 0.004853 | 4.7e-05 | 0.000943 | 0.31492 | 1.3529 | 2.5475 | 98.51 | 20.06 | 49.725 | 85.4419 | 86.9603 | 81.1974 | 0.011858 | 0.99633 | 0.023113 | 23.1986 | -0.033813 | 0.15515 | 0.0021 | 0.138 | 0.138 | 0.045307
2019-04-09 14:54 | 1554836053.5446 | 1172 | 0.20908 | 0.000723 | 0.000127 | 98.4174 | 98.6021 | 116.6789 | 20.5815 | 0.002653 | 0.00247 | 220.2508 | 0.26122 | 8.4017 | 0.002758 | 0.004554 | 0.004541 | 4.4e-05 | 0.000942 | 0.28241 | 1.377 | 2.5476 | 98.51 | 20.06 | 49.725 | 84.9763 | 87.504 | 80.5832 | 0.011858 | 0.99806 | 0.023145 | 23.1904 | -0.016764 | 0.1554 | 0.002219 | 0.14649 | 0.14649 | 0.042751
2019-04-09 14:54 | 1554836054.5432 | 1173 | 0.20909 | 0.000721 | 0.000126 | 98.441 | 98.6396 | 116.9495 | 20.5809 | 0.00265 | 0.00248 | 220.2426 | 0.25996 | 8.403 | 0.002756 | 0.004644 | 0.004612 | 4.1e-05 | 0.000944 | 0.27759 | 1.3293 | 2.5476 | 98.51 | 20.06 | 49.724 | 84.904 | 87.454 | 80.8313 | 0.011858 | 0.99736 | 0.023164 | 23.2254 | -0.018028 | 0.15465 | 0.002192 | 0.14477 | 0.14477 | 0.043049
2019-04-09 14:54 | 1554836055.5443 | 1174 | 0.20909 | 0.000724 | 0.000126 | 98.541 | 98.8622 | 116.7543 | 20.5809 | 0.002595 | 0.002451 | 220.2478 | 0.26025 | 8.4038 | 0.002742 | 0.003971 | 0.003917 | 4.1e-05 | 0.000953 | 0.30055 | 1.381 | 2.5476 | 98.51 | 20.06 | 49.724 | 83.5925 | 87.1246 | 80.1337 | 0.011858 | 0.99933 | 0.023168 | 23.1832 | -0.02911 | 0.15482 | 0.002178 | 0.14388 | 0.14388 | 0.043365
2019-04-09 14:54 | 1554836056.5447 | 1175 | 0.20905 | 0.000724 | 0.000125 | 98.4981 | 98.6884 | 116.4954 | 20.5809 | 0.002648 | 0.00246 | 220.212 | 0.26012 | 8.402 | 0.002645 | 0.003742 | 0.003686 | 4.1e-05 | 0.000953 | 0.28528 | 1.3684 | 2.5476 | 98.51 | 20.06 | 49.725 | 84.8666 | 84.7947 | 80.345 | 0.011858 | 0.99873 | 0.023135 | 23.1644 | -0.017265 | 0.15474 | 0.002396 | 0.15875 | 0.15875 | 0.039284
2019-04-09 14:54 | 1554836057.5434 | 1176 | 0.20906 | 0.00072 | 0.000124 | 98.5607 | 98.6527 | 116.6435 | 20.5821 | 0.00261 | 0.002452 | 220.2107 | 0.25963 | 8.4014 | 0.002752 | 0.003342 | 0.003304 | 4.1e-05 | 0.000948 | 0.29856 | 1.3713 | 2.5476 | 98.51 | 20.06 | 49.724 | 83.9368 | 87.3578 | 80.138 | 0.011858 | 0.99931 | 0.023156 | 23.1723 | -0.008343 | 0.15446 | 0.002383 | 0.15805 | 0.15805 | 0.039383
2019-04-09 14:54 | 1554836058.5444 | 1177 | 0.20907 | 0.000723 | 0.000124 | 98.4984 | 98.6733 | 116.8251 | 20.5811 | 0.002563 | 0.002423 | 220.2015 | 0.26076 | 8.4027 | 0.002689 | 0.004091 | 0.004007 | 4.7e-05 | 0.000953 | 0.28241 | 1.3785 | 2.5475 | 98.51 | 20.06 | 49.723 | 82.811 | 85.8351 | 79.4617 | 0.011858 | 1.0012 | 0.023197 | 23.1681 | -0.015865 | 0.15513 | 0.00228 | 0.1512 | 0.1512 | 0.041347
2019-04-09 14:54 | 1554836059.5434 | 1178 | 0.20904 | 0.000721 | 0.000124 | 98.6006 | 98.718 | 116.7662 | 20.5824 | 0.002589 | 0.002461 | 220.1858 | 0.26119 | 8.4038 | 0.002774 | 0.004504 | 0.004416 | 4.2e-05 | 0.000948 | 0.28816 | 1.3391 | 2.5475 | 98.51 | 20.06 | 49.724 | 83.4347 | 87.9022 | 80.3728 | 0.011858 | 0.99865 | 0.023161 | 23.1922 | -0.010649 | 0.15538 | 0.00248 | 0.16482 | 0.16482 | 0.037991
2019-04-09 14:54 | 1554836060.5441 | 1179 | 0.20909 | 0.000726 | 0.000126 | 98.5952 | 98.8192 | 116.5822 | 20.5824 | 0.002653 | 0.00241 | 220.1887 | 0.26104 | 8.402 | 0.002754 | 0.003827 | 0.00377 | 4.3e-05 | 0.000949 | 0.28943 | 1.3366 | 2.5475 | 98.51 | 20.06 | 49.725 | 84.9789 | 87.411 | 79.1504 | 0.011858 | 1.0021 | 0.023183 | 23.1338 | -0.020299 | 0.15529 | 0.002191 | 0.14489 | 0.14489 | 0.043194
2019-04-09 14:54 | 1554836061.5435 | 1180 | 0.2091 | 0.000727 | 0.000125 | 98.6378 | 98.7132 | 116.6172 | 20.5816 | 0.002646 | 0.002435 | 220.1579 | 0.26059 | 8.4008 | 0.002731 | 0.00389 | 0.003809 | 4.5e-05 | 0.000949 | 0.27704 | 1.3312 | 2.5476 | 98.51 | 20.06 | 49.725 | 84.8103 | 86.8481 | 79.7374 | 0.011858 | 1.0004 | 0.023167 | 23.1565 | -0.006837 | 0.15502 | 0.002129 | 0.14043 | 0.14043 | 0.044489
2019-04-09 14:54 | 1554836062.5439 | 1181 | 0.20907 | 0.000721 | 0.000126 | 98.6667 | 98.844 | 116.7315 | 20.5829 | 0.002634 | 0.002502 | 220.1414 | 0.26055 | 8.4007 | 0.002756 | 0.00426 | 0.004161 | 4.5e-05 | 0.000949 | 0.29502 | 1.3394 | 2.5477 | 98.51 | 20.06 | 49.725 | 84.5368 | 87.4702 | 81.3503 | 0.011858 | 0.9959 | 0.023125 | 23.2208 | -0.01606 | 0.155 | 0.002319 | 0.15332 | 0.15332 | 0.040743
2019-04-09 14:54 | 1554836063.5449 | 1182 | 0.20908 | 0.000719 | 0.000125 | 98.6813 | 98.9435 | 116.6942 | 20.5831 | 0.002639 | 0.002505 | 220.1302 | 0.26054 | 8.4023 | 0.002676 | 0.004327 | 0.004208 | 4.4e-05 | 0.000953 | 0.28799 | 1.3404 | 2.5476 | 98.51 | 20.06 | 49.724 | 84.6525 | 85.5321 | 81.4318 | 0.011858 | 0.99567 | 0.023119 | 23.2197 | -0.023733 | 0.155 | 0.002236 | 0.14763 | 0.14763 | 0.042312
2019-04-09 14:54 | 1554836064.5426 | 1183 | 0.20908 | 0.000722 | 0.000124 | 98.6866 | 98.9782 | 116.4334 | 20.5823 | 0.002591 | 0.002457 | 220.127 | 0.26086 | 8.4001 | 0.00267 | 0.003964 | 0.003962 | 4.4e-05 | 0.000951 | 0.27472 | 1.3548 | 2.5476 | 98.51 | 20.06 | 49.72 | 83.4955 | 85.3797 | 80.2737 | 0.011857 | 0.99893 | 0.023131 | 23.1559 | -0.026389 | 0.15519 | 0.002241 | 0.14811 | 0.14811 | 0.042226
2019-04-09 14:54 | 1554836065.5426 | 1184 | 0.20907 | 0.000719 | 0.000123 | 98.6971 | 98.9053 | 116.4907 | 20.5828 | 0.002557 | 0.002447 | 220.1318 | 0.26111 | 8.4025 | 0.002721 | 0.003864 | 0.003754 | 4.8e-05 | 0.000948 | 0.27671 | 1.3608 | 2.5477 | 98.51 | 20.06 | 49.718 | 82.6754 | 86.619 | 80.0293 | 0.011856 | 0.99962 | 0.023145 | 23.1536 | -0.018849 | 0.15534 | 0.002321 | 0.15373 | 0.15373 | 0.040722
2019-04-09 14:54 | 1554836066.543 | 1185 | 0.20907 | 0.000726 | 0.000124 | 98.8183 | 99.2078 | 116.7099 | 20.583 | 0.002621 | 0.002382 | 220.1346 | 0.26086 | 8.398 | 0.002676 | 0.003352 | 0.003312 | 4.6e-05 | 0.000944 | 0.28528 | 1.3738 | 2.5476 | 98.51 | 20.06 | 49.713 | 84.2164 | 85.532 | 78.4664 | 0.011855 | 1.0041 | 0.023218 | 23.124 | -0.035187 | 0.15519 | 0.002286 | 0.15182 | 0.15182 | 0.041193
2019-04-09 14:54 | 1554836067.5426 | 1186 | 0.20907 | 0.000722 | 0.000124 | 98.8254 | 99.1235 | 116.9331 | 20.5828 | 0.002597 | 0.00245 | 220.1344 | 0.2608 | 8.4 | 0.002712 | 0.003739 | 0.00367 | 4.4e-05 | 0.000944 | 0.286 | 1.3734 | 2.5477 | 98.51 | 20.06 | 49.711 | 83.6402 | 86.3918 | 80.0942 | 0.011855 | 0.99944 | 0.023187 | 23.1996 | -0.026935 | 0.15515 | 0.002325 | 0.15426 | 0.15426 | 0.040532
2019-04-09 14:54 | 1554836068.5427 | 1187 | 0.20906 | 0.000722 | 0.000124 | 98.8218 | 98.9796 | 116.7122 | 20.5832 | 0.002654 | 0.002436 | 220.1136 | 0.26066 | 8.4007 | 0.002661 | 0.00355 | 0.003498 | 4.1e-05 | 0.000946 | 0.30127 | 1.3314 | 2.5476 | 98.51 | 20.06 | 49.707 | 85.0054 | 85.1709 | 79.7764 | 0.011854 | 1.0003 | 0.023175 | 23.1673 | -0.014268 | 0.15507 | 0.00237 | 0.1573 | 0.1573 | 0.039729
2019-04-09 14:54 | 1554836069.5428 | 1188 | 0.20908 | 0.000717 | 0.000123 | 99.0017 | 99.2924 | 116.7977 | 20.5844 | 0.002633 | 0.002476 | 220.1239 | 0.26061 | 8.4032 | 0.002704 | 0.004625 | 0.004484 | 4.4e-05 | 0.000948 | 0.30558 | 1.3045 | 2.5477 | 98.51 | 20.06 | 49.704 | 84.5118 | 86.199 | 80.7218 | 0.011853 | 0.99767 | 0.023153 | 23.2068 | -0.026228 | 0.15504 | 0.002231 | 0.14756 | 0.14756 | 0.04234
2019-04-09 14:54 | 1554836070.5427 | 1189 | 0.20904 | 0.000717 | 0.000123 | 98.8894 | 99.0708 | 117.0876 | 20.5836 | 0.002632 | 0.002442 | 220.1143 | 0.261 | 8.4003 | 0.002718 | 0.004103 | 0.004076 | 4.6e-05 | 0.000946 | 0.30502 | 1.3552 | 2.5476 | 98.51 | 20.06 | 49.701 | 84.4861 | 86.5411 | 79.9023 | 0.011852 | 0.99998 | 0.023208 | 23.2086 | -0.016389 | 0.15527 | 0.002473 | 0.16477 | 0.16477 | 0.037975
2019-04-09 14:54 | 1554836071.5432 | 1190 | 0.2091 | 0.00072 | 0.000123 | 98.9371 | 99.3248 | 117.3265 | 20.5846 | 0.002567 | 0.002446 | 220.103 | 0.26085 | 8.4018 | 0.002729 | 0.003803 | 0.003706 | 4.5e-05 | 0.000949 | 0.30381 | 1.3149 | 2.5475 | 98.51 | 20.07 | 49.698 | 82.9048 | 86.8214 | 80.006 | 0.011859 | 0.99969 | 0.023228 | 23.2357 | -0.034982 | 0.15518 | 0.002128 | 0.14094 | 0.14094 | 0.044372
2019-04-09 14:55 | 1554836072.5442 | 1191 | 0.20909 | 0.000721 | 0.000123 | 99.0339 | 99.3014 | 117.4866 | 20.583 | 0.0026 | 0.002439 | 220.1244 | 0.26024 | 8.4028 | 0.002715 | 0.004176 | 0.004126 | 4.6e-05 | 0.000951 | 0.30486 | 1.297 | 2.5475 | 98.51 | 20.07 | 49.693 | 83.696 | 86.4686 | 79.8354 | 0.011858 | 1.0002 | 0.02325 | 23.2459 | -0.024129 | 0.15481 | 0.002186 | 0.14506 | 0.14506 | 0.043011
2019-04-09 14:55 | 1554836073.5442 | 1192 | 0.20905 | 0.000716 | 0.000123 | 99.0869 | 99.3953 | 117.5047 | 20.5844 | 0.002582 | 0.002449 | 220.08 | 0.26078 | 8.4001 | 0.002687 | 0.003716 | 0.003645 | 4.4e-05 | 0.000951 | 0.29048 | 1.3659 | 2.5477 | 98.51 | 20.07 | 49.69 | 83.2713 | 85.7994 | 80.0823 | 0.011857 | 0.99947 | 0.023244 | 23.2559 | -0.027802 | 0.15514 | 0.002451 | 0.16345 | 0.16345 | 0.03825
2019-04-09 14:55 | 1554836074.5434 | 1193 | 0.20908 | 0.000715 | 0.000123 | 99.1208 | 99.4201 | 117.6007 | 20.5838 | 0.002656 | 0.0025 | 220.1001 | 0.26092 | 8.4016 | 0.002731 | 0.004569 | 0.004487 | 4.7e-05 | 0.000946 | 0.27632 | 1.3513 | 2.5475 | 98.51 | 20.07 | 49.684 | 85.063 | 86.8668 | 81.3108 | 0.011856 | 0.99601 | 0.023213 | 23.3058 | -0.026967 | 0.15522 | 0.002225 | 0.14758 | 0.14758 | 0.042386
2019-04-09 14:55 | 1554836075.5428 | 1194 | 0.2091 | 0.000716 | 0.000123 | 99.1309 | 99.6008 | 117.787 | 20.5837 | 0.002622 | 0.002402 | 220.0872 | 0.25988 | 8.4029 | 0.002731 | 0.004225 | 0.004108 | 4.5e-05 | 0.000943 | 0.31454 | 1.3318 | 2.5475 | 98.51 | 20.07 | 49.682 | 84.2396 | 86.8563 | 78.9598 | 0.011855 | 1.0027 | 0.023309 | 23.2468 | -0.042295 | 0.1546 | 0.002109 | 0.14006 | 0.14006 | 0.044486
2019-04-09 14:55 | 1554836076.5439 | 1195 | 0.20913 | 0.000714 | 0.000122 | 99.2423 | 99.4869 | 117.7068 | 20.5836 | 0.002599 | 0.002442 | 220.0838 | 0.26022 | 8.403 | 0.002688 | 0.003978 | 0.003877 | 4.5e-05 | 0.000936 | 0.28119 | 1.3515 | 2.5474 | 98.51 | 20.07 | 49.679 | 83.681 | 85.8212 | 79.9148 | 0.011855 | 0.99994 | 0.023269 | 23.2703 | -0.022015 | 0.1548 | 0.001939 | 0.12801 | 0.12801 | 0.048733
2019-04-09 14:55 | 1554836077.5429 | 1196 | 0.2091 | 0.000716 | 0.000121 | 99.2304 | 99.395 | 117.1638 | 20.5843 | 0.002639 | 0.002444 | 220.0962 | 0.26035 | 8.4016 | 0.002676 | 0.004055 | 0.004038 | 4.4e-05 | 0.000935 | 0.30989 | 1.3524 | 2.5476 | 98.51 | 20.07 | 49.674 | 84.6436 | 85.5386 | 79.9561 | 0.011853 | 0.99983 | 0.023214 | 23.218 | -0.014823 | 0.15488 | 0.002157 | 0.14291 | 0.14291 | 0.043675
2019-04-09 14:55 | 1554836078.5443 | 1197 | 0.20907 | 0.00071 | 0.00012 | 99.371 | 99.6211 | 117.303 | 20.5836 | 0.002577 | 0.002429 | 220.0667 | 0.26084 | 8.4026 | 0.002678 | 0.004489 | 0.004417 | 4.2e-05 | 0.000929 | 0.29607 | 1.3291 | 2.5475 | 98.51 | 20.07 | 49.672 | 83.1645 | 85.5895 | 79.5944 | 0.011853 | 1.0009 | 0.02324 | 23.2198 | -0.022487 | 0.15517 | 0.002319 | 0.15444 | 0.15444 | 0.040491
2019-04-09 14:55 | 1554836079.5426 | 1198 | 0.20907 | 0.000716 | 0.000119 | 99.3081 | 99.7226 | 117.4063 | 20.5846 | 0.002643 | 0.002374 | 220.0845 | 0.26078 | 8.4017 | 0.002738 | 0.00438 | 0.004291 | 4.4e-05 | 0.000927 | 0.28407 | 1.3655 | 2.5476 | 98.51 | 20.07 | 49.667 | 84.7547 | 87.0182 | 78.2793 | 0.011852 | 1.0046 | 0.023293 | 23.1867 | -0.037255 | 0.15514 | 0.002296 | 0.15324 | 0.15324 | 0.040799
2019-04-09 14:55 | 1554836080.5431 | 1199 | 0.20911 | 0.000715 | 0.000119 | 99.3308 | 99.6035 | 117.3014 | 20.5846 | 0.002656 | 0.002441 | 220.0463 | 0.26029 | 8.4021 | 0.002665 | 0.004454 | 0.004361 | 4.4e-05 | 0.000929 | 0.31349 | 1.3931 | 2.5477 | 98.51 | 20.07 | 49.664 | 85.0565 | 85.2672 | 79.8892 | 0.011851 | 1 | 0.02323 | 23.2294 | -0.02453 | 0.15485 | 0.002056 | 0.13618 | 0.13618 | 0.045825
2019-04-09 14:55 | 1554836081.5435 | 1200 | 0.20908 | 0.000709 | 0.000119 | 99.3891 | 99.7299 | 117.2089 | 20.5845 | 0.002598 | 0.002449 | 220.0636 | 0.2603 | 8.4018 | 0.002651 | 0.004062 | 0.004028 | 4.5e-05 | 0.000936 | 0.26272 | 1.3873 | 2.5476 | 98.51 | 20.07 | 49.662 | 83.6513 | 84.9334 | 80.0755 | 0.01185 | 0.99949 | 0.023215 | 23.2263 | -0.030611 | 0.15485 | 0.00227 | 0.15094 | 0.15094 | 0.041343
2019-04-09 14:55 | 1554836082.5435 | 1201 | 0.20909 | 0.000709 | 0.000119 | 99.4454 | 99.6294 | 117.0737 | 20.5848 | 0.002544 | 0.002434 | 220.0673 | 0.26084 | 8.4023 | 0.002659 | 0.004434 | 0.004339 | 4.3e-05 | 0.000927 | 0.29159 | 1.3894 | 2.5477 | 98.51 | 20.07 | 49.659 | 82.3528 | 85.1248 | 79.7224 | 0.01185 | 1.0005 | 0.023213 | 23.2013 | -0.01653 | 0.15517 | 0.002206 | 0.14647 | 0.14647 | 0.042694
2019-04-09 14:55 | 1554836083.5422 | 1202 | 0.2091 | 0.000708 | 0.000118 | 99.5011 | 99.7742 | 117.1873 | 20.585 | 0.002557 | 0.00245 | 220.0754 | 0.26069 | 8.4034 | 0.002733 | 0.003811 | 0.003832 | 4.5e-05 | 0.000925 | 0.27815 | 1.3878 | 2.5477 | 98.51 | 20.07 | 49.657 | 82.6675 | 86.9154 | 80.111 | 0.011849 | 0.99939 | 0.023211 | 23.2254 | -0.024521 | 0.15509 | 0.002126 | 0.14093 | 0.14093 | 0.044348
2019-04-09 14:55 | 1554836084.5432 | 1203 | 0.20908 | 0.000711 | 0.000118 | 99.533 | 99.8486 | 117.4124 | 20.5849 | 0.002653 | 0.002383 | 220.074 | 0.26132 | 8.4023 | 0.002596 | 0.003492 | 0.003413 | 4.5e-05 | 0.000924 | 0.28672 | 1.3832 | 2.5476 | 98.51 | 20.07 | 49.656 | 84.9876 | 83.6135 | 78.4824 | 0.011849 | 1.004 | 0.023287 | 23.194 | -0.028319 | 0.15546 | 0.002264 | 0.15104 | 0.15104 | 0.04148
2019-04-09 14:55 | 1554836085.5425 | 1204 | 0.20912 | 0.000706 | 0.000118 | 99.5246 | 99.8035 | 117.4699 | 20.5851 | 0.002648 | 0.002479 | 220.0439 | 0.26161 | 8.4034 | 0.002699 | 0.004519 | 0.004398 | 4.6e-05 | 0.000925 | 0.30342 | 1.3445 | 2.5476 | 98.51 | 20.07 | 49.657 | 84.8581 | 86.1013 | 80.8051 | 0.011849 | 0.99743 | 0.023216 | 23.2762 | -0.025029 | 0.15563 | 0.00204 | 0.135 | 0.135 | 0.04646
2019-04-09 14:55 | 1554836086.5427 | 1205 | 0.2091 | 0.000708 | 0.000118 | 99.5071 | 99.765 | 117.3566 | 20.5854 | 0.002627 | 0.002435 | 220.168 | 0.26079 | 8.4017 | 0.002692 | 0.003583 | 0.003569 | 4.4e-05 | 0.000925 | 0.29806 | 1.3233 | 2.5476 | 98.51 | 20.07 | 49.657 | 84.3599 | 85.9231 | 79.7466 | 0.011849 | 1.0004 | 0.02324 | 23.2301 | -0.023153 | 0.15515 | 0.002168 | 0.14411 | 0.14411 | 0.043388
2019-04-09 14:55 | 1554836087.5424 | 1206 | 0.20912 | 0.000703 | 0.000116 | 99.5818 | 100.043 | 117.5328 | 20.5854 | 0.002554 | 0.002444 | 192.6444 | 0.26076 | 8.4036 | 0.002624 | 0.003795 | 0.003737 | 4.6e-05 | 0.000924 | 0.29336 | 1.3619 | 2.5476 | 98.51 | 20.07 | 49.66 | 82.6008 | 84.2904 | 79.9679 | 0.01185 | 0.99979 | 0.02325 | 23.2549 | -0.041333 | 0.15513 | 0.00205 | 0.13601 | 0.13601 | 0.045963
2019-04-09 14:55 | 1554836088.5433 | 1207 | 0.20911 | 0.000702 | 0.000116 | 99.7018 | 99.9542 | 117.722 | 20.5849 | 0.002641 | 0.002403 | 155.2231 | 0.26095 | 8.4001 | 0.002678 | 0.004846 | 0.004714 | 4.8e-05 | 0.000925 | 0.30342 | 1.3521 | 2.5475 | 98.51 | 20.07 | 49.662 | 84.691 | 85.5759 | 78.9827 | 0.01185 | 1.0026 | 0.023301 | 23.2411 | -0.022618 | 0.15524 | 0.002125 | 0.1416 | 0.1416 | 0.044183
2019-04-09 14:55 | 1554836089.5427 | 1208 | 0.20911 | 0.000708 | 0.000114 | 99.6503 | 100.0399 | 117.9579 | 20.5853 | 0.002553 | 0.002462 | 117.7376 | 0.26095 | 8.3995 | 0.00266 | 0.00414 | 0.004131 | 4.2e-05 | 0.000932 | 0.2959 | 1.3462 | 2.5476 | 98.51 | 20.07 | 49.664 | 82.585 | 85.144 | 80.3958 | 0.011851 | 0.99858 | 0.023278 | 23.311 | -0.034906 | 0.15524 | 0.002076 | 0.13806 | 0.13806 | 0.045315
2019-04-09 14:55 | 1554836090.5426 | 1209 | 0.20911 | 0.000705 | 0.000113 | 99.7334 | 100.1477 | 118.1081 | 20.5854 | 0.002583 | 0.002434 | 80.2889 | 0.26055 | 8.4019 | 0.002593 | 0.004185 | 0.004088 | 4.2e-05 | 0.000929 | 0.30021 | 1.3098 | 2.5475 | 98.51 | 20.07 | 49.668 | 83.298 | 83.55 | 79.7255 | 0.011852 | 1.0005 | 0.023315 | 23.3037 | -0.037078 | 0.155 | 0.002114 | 0.14103 | 0.14103 | 0.044293
2019-04-09 14:55 | 1554836091.5435 | 1210 | 0.20911 | 0.000709 | 0.000112 | 99.8197 | 100.085 | 118.283 | 20.5859 | 0.002573 | 0.002389 | 42.8208 | 0.26128 | 8.4016 | 0.00265 | 0.004134 | 0.004058 | 4.3e-05 | 0.000932 | 0.27793 | 1.3067 | 2.5474 | 98.51 | 20.07 | 49.67 | 83.0607 | 84.9017 | 78.6509 | 0.011852 | 1.0035 | 0.023368 | 23.2854 | -0.02374 | 0.15543 | 0.002101 | 0.14049 | 0.14049 | 0.044586
2019-04-09 14:55 | 1554836092.5429 | 1211 | 0.20911 | 0.000704 | 0.000109 | 99.7271 | 100.062 | 118.8447 | 20.5858 | 0.002575 | 0.002361 | 5.3501 | 0.2391 | 8.404 | 0.002569 | 0.004154 | 0.004072 | 4.3e-05 | 0.000933 | 0.30646 | 1.3355 | 2.5476 | 98.51 | 20.07 | 49.672 | 83.1167 | 82.963 | 77.9766 | 0.011853 | 1.0055 | 0.023446 | 23.3183 | -0.029982 | 0.14224 | 0.002071 | 0.13907 | 0.13907 | 0.041218
2019-04-09 14:55 | 1554836093.5438 | 1212 | 0.20913 | 0.000702 | 0.000104 | 99.8817 | 100.3232 | 119.2214 | 20.5869 | 0.00261 | 0.00237 | -32.0939 | 0.23151 | 8.401 | 0.002639 | 0.003428 | 0.003264 | 4.4e-05 | 0.000928 | 0.30431 | 1.3878 | 2.5477 | 98.51 | 20.07 | 49.675 | 83.9578 | 84.6516 | 78.1919 | 0.011854 | 1.0048 | 0.023476 | 23.3624 | -0.03945 | 0.13772 | 0.001987 | 0.13365 | 0.13365 | 0.041529
2019-04-09 14:55 | 1554836094.5435 | 1213 | 0.20919 | 0.000698 | 9.7e-05 | 99.8465 | 100.3504 | 119.0578 | 20.5867 | 0.00256 | 0.00235 | -69.5768 | 0.23111 | 8.4008 | 0.00266 | 0.004268 | 0.004144 | 4.4e-05 | 0.000923 | 0.2943 | 1.3584 | 2.5476 | 98.51 | 20.07 | 49.678 | 82.7415 | 85.1625 | 77.6983 | 0.011854 | 1.0063 | 0.023476 | 23.3299 | -0.045021 | 0.13749 | 0.001648 | 0.11023 | 0.11023 | 0.050263
2019-04-09 14:55 | 1554836095.5416 | 1214 | 0.2092 | 0.000704 | 9e-05 | 100.0059 | 100.2777 | 119.2085 | 20.5859 | 0.002534 | 0.002343 | -107.0497 | 0.23234 | 8.3994 | 0.002603 | 0.004392 | 0.00421 | 4.4e-05 | 0.000922 | 0.27721 | 1.3343 | 2.5475 | 98.51 | 20.07 | 49.68 | 82.1222 | 83.7806 | 77.5238 | 0.011855 | 1.0068 | 0.023497 | 23.3389 | -0.024285 | 0.13822 | 0.001563 | 0.10472 | 0.10472 | 0.053189
2019-04-09 14:55 | 1554836096.542 | 1215 | 0.20928 | 0.000699 | 8.3e-05 | 99.8888 | 100.6085 | 119.6906 | 20.5859 | 0.002512 | 0.002366 | -144.6423 | 0.23128 | 8.4043 | 0.002625 | 0.003525 | 0.003294 | 4.2e-05 | 0.000917 | 0.29048 | 1.334 | 2.5478 | 98.51 | 20.06 | 49.684 | 81.5967 | 84.3048 | 78.0871 | 0.011848 | 1.0051 | 0.023525 | 23.4048 | -0.064211 | 0.13759 | 0.001116 | 0.073752 | 0.073752 | 0.075181
2019-04-09 14:55 | 1554836097.543 | 1216 | 0.20934 | 0.000685 | 7.7e-05 | 99.9829 | 100.2457 | 120.0446 | 20.5855 | 0.002598 | 0.002381 | -154.6199 | 0.23182 | 8.4043 | 0.00263 | 0.004384 | 0.004221 | 4.6e-05 | 0.000924 | 0.29103 | 1.3209 | 2.5476 | 98.51 | 20.06 | 49.684 | 83.671 | 84.4395 | 78.4556 | 0.011848 | 1.0041 | 0.023548 | 23.4517 | -0.023481 | 0.13791 | 0.000766 | 0.049491 | 0.049491 | 0.1123
2019-04-09 14:55 | 1554836098.542 | 1217 | 0.20932 | 0.000671 | 7.3e-05 | 99.9597 | 100.3775 | 120.3955 | 20.5868 | 0.002539 | 0.00237 | -154.6483 | 0.232 | 8.4003 | 0.002614 | 0.003902 | 0.00378 | 4.4e-05 | 0.00092 | 0.29264 | 1.3339 | 2.5475 | 98.51 | 20.07 | 49.685 | 82.237 | 84.0373 | 78.1786 | 0.011856 | 1.0049 | 0.023591 | 23.4767 | -0.037312 | 0.13801 | 0.000903 | 0.059503 | 0.059503 | 0.093475
2019-04-09 14:55 | 1554836099.5427 | 1218 | 0.20933 | 0.000657 | 6.9e-05 | 100.0196 | 100.3345 | 120.5969 | 20.5871 | 0.002481 | 0.002363 | -154.6528 | 0.23195 | 8.4028 | 0.002678 | 0.003968 | 0.003717 | 4.5e-05 | 0.000924 | 0.28871 | 1.3372 | 2.5477 | 98.51 | 20.07 | 49.685 | 80.8422 | 85.5961 | 78.0043 | 0.011856 | 1.0054 | 0.023617 | 23.4905 | -0.028124 | 0.13799 | 0.000908 | 0.060125 | 0.060125 | 0.09249
2019-04-09 14:55 | 1554836100.5421 | 1219 | 0.20937 | 0.00064 | 6.5e-05 | 100.0267 | 100.5631 | 120.5138 | 20.5866 | 0.002549 | 0.002346 | -154.6593 | 0.23142 | 8.4027 | 0.002644 | 0.003699 | 0.003562 | 4e-05 | 0.000921 | 0.27472 | 1.3059 | 2.5475 | 98.51 | 20.07 | 49.684 | 82.4904 | 84.7612 | 77.6016 | 0.011856 | 1.0065 | 0.023622 | 23.4689 | -0.047833 | 0.13767 | 0.000658 | 0.042723 | 0.042723 | 0.12986
2019-04-09 14:55 | 1554836101.5431 | 1220 | 0.20937 | 0.000624 | 6e-05 | 100.0934 | 100.4392 | 120.7145 | 20.587 | 0.002551 | 0.00234 | -154.6542 | 0.23172 | 8.4009 | 0.002658 | 0.00371 | 0.003496 | 4.2e-05 | 0.00092 | 0.31045 | 1.3388 | 2.5476 | 98.51 | 20.07 | 49.684 | 82.5398 | 85.1012 | 77.4515 | 0.011856 | 1.007 | 0.023647 | 23.4834 | -0.030844 | 0.13785 | 0.000716 | 0.047176 | 0.047176 | 0.11776
2019-04-09 14:55 | 1554836102.543 | 1221 | 0.20934 | 0.000613 | 5.6e-05 | 100.116 | 100.5431 | 120.5833 | 20.5874 | 0.002528 | 0.002382 | -154.6636 | 0.23248 | 8.4011 | 0.002618 | 0.003863 | 0.003706 | 4.2e-05 | 0.000924 | 0.28174 | 1.3604 | 2.5477 | 98.51 | 20.07 | 49.682 | 81.9904 | 84.1456 | 78.4762 | 0.011855 | 1.004 | 0.0236 | 23.5049 | -0.038072 | 0.1383 | 0.000924 | 0.062056 | 0.062056 | 0.089815
2019-04-09 14:55 | 1554836103.5418 | 1222 | 0.2094 | 0.000597 | 5e-05 | 100.0985 | 100.4901 | 120.7945 | 20.5863 | 0.002414 | 0.002318 | -154.6656 | 0.23226 | 8.4038 | 0.002705 | 0.003948 | 0.003758 | 4e-05 | 0.000927 | 0.28385 | 1.3233 | 2.5476 | 98.51 | 20.07 | 49.681 | 79.2508 | 86.2391 | 76.9449 | 0.011855 | 1.0084 | 0.023672 | 23.4742 | -0.034926 | 0.13817 | 0.000556 | 0.036455 | 0.036455 | 0.15275
2019-04-09 14:55 | 1554836104.5422 | 1223 | 0.20939 | 0.00059 | 4.5e-05 | 100.1017 | 100.5813 | 121.0856 | 20.5863 | 0.002462 | 0.002271 | -154.6741 | 0.23172 | 8.4037 | 0.002688 | 0.004304 | 0.004077 | 4.1e-05 | 0.000931 | 0.31166 | 1.3464 | 2.5476 | 98.51 | 20.07 | 49.679 | 80.3978 | 85.8239 | 75.8107 | 0.011855 | 1.0117 | 0.023739 | 23.4644 | -0.042747 | 0.13785 | 0.000616 | 0.041164 | 0.041164 | 0.13496
2019-04-09 14:55 | 1554836105.5432 | 1224 | 0.20945 | 0.000577 | 4e-05 | 100.2328 | 100.6481 | 121.3346 | 20.5878 | 0.002487 | 0.002353 | -154.6746 | 0.23169 | 8.4013 | 0.002719 | 0.00462 | 0.00448 | 4.4e-05 | 0.000931 | 0.29966 | 1.352 | 2.5476 | 98.51 | 20.07 | 49.677 | 80.9979 | 86.5786 | 77.7671 | 0.011854 | 1.0061 | 0.023697 | 23.5542 | -0.036982 | 0.13783 | 0.000274 | 0.017085 | 0.017085 | 0.32512
2019-04-09 14:55 | 1554836106.5429 | 1225 | 0.20944 | 0.000569 | 3.5e-05 | 100.1041 | 100.4509 | 121.5666 | 20.5875 | 0.00253 | 0.002288 | -154.6579 | 0.23187 | 8.4042 | 0.002744 | 0.004133 | 0.003961 | 4.3e-05 | 0.000935 | 0.28385 | 1.3606 | 2.5475 | 98.51 | 20.07 | 49.676 | 82.035 | 87.1811 | 76.2149 | 0.011854 | 1.0105 | 0.023772 | 23.5245 | -0.030928 | 0.13794 | 0.000376 | 0.024726 | 0.024726 | 0.22482
2019-04-09 14:55 | 1554836107.543 | 1226 | 0.20945 | 0.000554 | 3e-05 | 100.178 | 100.4873 | 121.7743 | 20.5879 | 0.002415 | 0.002251 | -154.6573 | 0.23159 | 8.4025 | 0.002724 | 0.003792 | 0.003654 | 4e-05 | 0.000944 | 0.29358 | 1.3445 | 2.5476 | 98.51 | 20.08 | 49.669 | 79.2773 | 86.6834 | 75.3156 | 0.011859 | 1.0131 | 0.023823 | 23.5143 | -0.027579 | 0.13778 | 0.000345 | 0.022848 | 0.022848 | 0.24301
2019-04-09 14:55 | 1554836108.5432 | 1227 | 0.20946 | 0.000543 | 2.6e-05 | 100.2703 | 100.8721 | 121.4961 | 20.5874 | 0.002432 | 0.002264 | -154.6514 | 0.2316 | 8.4031 | 0.00277 | 0.003662 | 0.003439 | 4.4e-05 | 0.000942 | 0.29231 | 1.3775 | 2.5476 | 98.51 | 20.08 | 49.667 | 79.662 | 87.7901 | 75.6481 | 0.011859 | 1.0122 | 0.023785 | 23.4986 | -0.053523 | 0.13778 | 0.000324 | 0.02152 | 0.02152 | 0.25802
2019-04-09 14:55 | 1554836109.542 | 1228 | 0.20948 | 0.000536 | 2.2e-05 | 100.2691 | 100.7516 | 121.2994 | 20.5889 | 0.002465 | 0.002239 | -154.6553 | 0.23217 | 8.4031 | 0.002747 | 0.004208 | 0.004023 | 3.8e-05 | 0.000944 | 0.2855 | 1.3733 | 2.5476 | 98.51 | 20.08 | 49.663 | 80.4649 | 87.2426 | 75.0299 | 0.011858 | 1.014 | 0.023787 | 23.4588 | -0.042942 | 0.13812 | 0.000186 | 0.011892 | 0.011892 | 0.46808
2019-04-09 14:55 | 1554836110.5411 | 1229 | 0.20948 | 0.000526 | 1.9e-05 | 100.259 | 100.8766 | 121.5799 | 20.5878 | 0.002441 | 0.002203 | -154.6544 | 0.23185 | 8.4008 | 0.002739 | 0.003926 | 0.003818 | 4.2e-05 | 0.000938 | 0.26305 | 1.3787 | 2.5476 | 98.51 | 20.08 | 49.66 | 79.89 | 87.0666 | 74.1673 | 0.011857 | 1.0165 | 0.023844 | 23.4568 | -0.054929 | 0.13793 | 0.000224 | 0.014893 | 0.014893 | 0.37324
2019-04-09 14:55 | 1554836111.5425 | 1230 | 0.20953 | 0.000514 | 1.6e-05 | 100.264 | 100.815 | 121.575 | 20.5884 | 0.002385 | 0.002304 | -154.6677 | 0.23215 | 8.4037 | 0.002695 | 0.004256 | 0.004057 | 4.1e-05 | 0.000933 | 0.26626 | 1.3388 | 2.5477 | 98.51 | 20.08 | 49.657 | 78.5355 | 86.0077 | 76.5898 | 0.011857 | 1.0095 | 0.02376 | 23.538 | -0.049014 | 0.13811 | -8.2e-05 | -0.006779 | -0.006779 | -0.82099
2019-04-09 14:55 | 1554836112.5427 | 1231 | 0.20952 | 0.000509 | 1.3e-05 | 100.245 | 100.6064 | 121.4752 | 20.5889 | 0.002394 | 0.002247 | -154.6745 | 0.23204 | 8.4032 | 0.002721 | 0.004566 | 0.004353 | 4.4e-05 | 0.000937 | 0.28711 | 1.3464 | 2.5477 | 98.51 | 20.08 | 49.649 | 78.7677 | 86.6322 | 75.2231 | 0.011855 | 1.0134 | 0.023797 | 23.4823 | -0.032191 | 0.13804 | 4e-06 | -0.000522 | -0.000522 | -10.6606
2019-04-09 14:55 | 1554836113.5421 | 1232 | 0.20952 | 0.000499 | 1.1e-05 | 100.3167 | 100.6608 | 121.2361 | 20.5886 | 0.002415 | 0.002271 | -154.6743 | 0.23157 | 8.4029 | 0.00269 | 0.003746 | 0.003611 | 4.1e-05 | 0.000936 | 0.30917 | 1.3537 | 2.5476 | 98.51 | 20.08 | 49.645 | 79.2763 | 85.8817 | 75.8059 | 0.011854 | 1.0117 | 0.023754 | 23.4788 | -0.030623 | 0.13776 | -2e-05 | -0.002116 | -0.002116 | -2.6232
2019-04-09 14:55 | 1554836114.5412 | 1233 | 0.20951 | 0.000494 | 1e-05 | 100.2813 | 100.7541 | 121.4426 | 20.5875 | 0.002411 | 0.002202 | -154.668 | 0.23247 | 8.3996 | 0.002747 | 0.003821 | 0.003609 | 4.4e-05 | 0.00094 | 0.29247 | 1.3591 | 2.5475 | 98.51 | 20.08 | 49.641 | 79.163 | 87.2528 | 74.1507 | 0.011853 | 1.0165 | 0.023831 | 23.443 | -0.042067 | 0.13829 | 9.7e-05 | 0.006362 | 0.006362 | 0.876
2019-04-09 14:55 | 1554836115.5412 | 1234 | 0.20953 | 0.000481 | 9e-06 | 100.2215 | 100.7028 | 121.2624 | 20.5879 | 0.002374 | 0.002223 | -154.668 | 0.23202 | 8.4013 | 0.002671 | 0.003982 | 0.003745 | 4.5e-05 | 0.000942 | 0.28766 | 1.3347 | 2.5476 | 98.51 | 20.08 | 49.634 | 78.2919 | 85.4314 | 74.647 | 0.011851 | 1.0151 | 0.023796 | 23.4423 | -0.042854 | 0.13803 | -1e-06 | -0.000614 | -0.000614 | -9.0553
2019-04-09 14:55 | 1554836116.5412 | 1235 | 0.20955 | 0.000482 | 8e-06 | 100.2343 | 100.7578 | 121.1675 | 20.5886 | 0.002391 | 0.002181 | -154.6781 | 0.23226 | 8.4015 | 0.002665 | 0.00383 | 0.003695 | 4.1e-05 | 0.000948 | 0.28672 | 1.3308 | 2.5476 | 98.51 | 20.08 | 49.631 | 78.6987 | 85.2822 | 73.6499 | 0.01185 | 1.018 | 0.023821 | 23.3995 | -0.046601 | 0.13817 | -0.00012 | -0.009011 | -0.009011 | -0.61793
2019-04-09 14:55 | 1554836117.5422 | 1236 | 0.20952 | 0.000475 | 7e-06 | 100.2718 | 100.8453 | 121.1006 | 20.5883 | 0.002353 | 0.002124 | -154.6583 | 0.23252 | 8.3987 | 0.002626 | 0.003223 | 0.003059 | 4.6e-05 | 0.000943 | 0.26626 | 1.3281 | 2.5476 | 98.51 | 20.08 | 49.624 | 77.7676 | 84.3502 | 72.2714 | 0.011849 | 1.0221 | 0.023862 | 23.3465 | -0.05101 | 0.13833 | 4.9e-05 | 0.003111 | 0.003111 | 1.7917
2019-04-09 14:55 | 1554836118.5422 | 1237 | 0.20953 | 0.000471 | 6e-06 | 100.2661 | 100.6904 | 121.5748 | 20.5889 | 0.002406 | 0.002221 | -154.6658 | 0.2322 | 8.4031 | 0.002628 | 0.004039 | 0.003912 | 4.1e-05 | 0.000949 | 0.29463 | 1.3308 | 2.5477 | 98.51 | 20.08 | 49.622 | 79.0615 | 84.382 | 74.5965 | 0.011848 | 1.0152 | 0.023828 | 23.4708 | -0.037768 | 0.13814 | -2.9e-05 | -0.002435 | -0.002435 | -2.2865
2019-04-09 14:55 | 1554836119.5414 | 1238 | 0.20955 | 0.00047 | 6e-06 | 100.2134 | 100.6962 | 121.7543 | 20.5886 | 0.002467 | 0.002154 | -154.6485 | 0.23239 | 8.4031 | 0.002675 | 0.004058 | 0.003807 | 4.4e-05 | 0.00095 | 0.29717 | 1.3175 | 2.5476 | 98.51 | 20.08 | 49.62 | 80.524 | 85.5242 | 73.0052 | 0.011848 | 1.0199 | 0.023901 | 23.4343 | -0.042991 | 0.13825 | -9.4e-05 | -0.007048 | -0.007048 | -0.79048
2019-04-09 14:55 | 1554836120.5413 | 1239 | 0.20952 | 0.000471 | 5e-06 | 100.2959 | 100.7231 | 121.6067 | 20.5894 | 0.002354 | 0.002137 | -154.6538 | 0.23259 | 8.3998 | 0.002604 | 0.004046 | 0.003919 | 4.3e-05 | 0.000953 | 0.29806 | 1.4034 | 2.5477 | 98.51 | 20.08 | 49.612 | 77.8005 | 83.8198 | 72.5896 | 0.011846 | 1.0211 | 0.023901 | 23.406 | -0.038016 | 0.13837 | 5.6e-05 | 0.003719 | 0.003719 | 1.4993
2019-04-09 14:55 | 1554836121.5424 | 1240 | 0.20956 | 0.000465 | 4e-06 | 100.2539 | 100.8301 | 121.4978 | 20.5887 | 0.002405 | 0.00218 | -154.6403 | 0.23234 | 8.4021 | 0.002658 | 0.004023 | 0.003879 | 4.2e-05 | 0.000956 | 0.28655 | 1.3793 | 2.5476 | 98.51 | 20.08 | 49.61 | 79.0318 | 85.1019 | 73.6181 | 0.011845 | 1.0181 | 0.023854 | 23.4303 | -0.051261 | 0.13822 | -0.000164 | -0.011968 | -0.011968 | -0.46543
2019-04-09 14:55 | 1554836122.5414 | 1241 | 0.20955 | 0.000461 | 4e-06 | 100.2896 | 100.7527 | 121.6151 | 20.5895 | 0.002411 | 0.00218 | -154.6411 | 0.2311 | 8.3997 | 0.002626 | 0.004669 | 0.004391 | 4e-05 | 0.00096 | 0.28694 | 1.3949 | 2.5476 | 98.51 | 20.08 | 49.605 | 79.1698 | 84.3458 | 73.6164 | 0.011844 | 1.0181 | 0.023866 | 23.4416 | -0.041201 | 0.13748 | -0.000131 | -0.00964 | -0.00964 | -0.57476
2019-04-09 14:55 | 1554836123.5424 | 1242 | 0.20955 | 0.000464 | 4e-06 | 100.2673 | 100.6374 | 121.9344 | 20.5892 | 0.002306 | 0.002124 | -154.6471 | 0.23259 | 8.4031 | 0.002623 | 0.004053 | 0.003944 | 4.1e-05 | 0.000963 | 0.2891 | 1.3749 | 2.5476 | 98.51 | 20.08 | 49.604 | 76.6476 | 84.2676 | 72.2761 | 0.011844 | 1.0221 | 0.023944 | 23.4269 | -0.032952 | 0.13837 | -0.00014 | -0.010228 | -0.010228 | -0.5452
2019-04-09 14:55 | 1554836124.5415 | 1243 | 0.20955 | 0.000463 | 3e-06 | 100.1908 | 100.7231 | 121.7839 | 20.5894 | 0.002362 | 0.002154 | -154.6548 | 0.2314 | 8.402 | 0.002613 | 0.00365 | 0.003566 | 4.4e-05 | 0.000962 | 0.27295 | 1.345 | 2.5475 | 98.51 | 20.08 | 49.603 | 78.0017 | 84.0242 | 72.9919 | 0.011844 | 1.0199 | 0.023904 | 23.4367 | -0.047395 | 0.13766 | -9.8e-05 | -0.007236 | -0.007236 | -0.7667
2019-04-09 14:55 | 1554836125.5419 | 1244 | 0.20952 | 0.000465 | 3e-06 | 100.2768 | 100.6412 | 121.6238 | 20.5891 | 0.002357 | 0.002161 | -154.6392 | 0.23163 | 8.4017 | 0.002641 | 0.004386 | 0.004108 | 4.4e-05 | 0.000961 | 0.3079 | 1.3555 | 2.5476 | 98.5 | 20.08 | 49.6 | 77.862 | 84.6973 | 73.1677 | 0.011844 | 1.0194 | 0.023882 | 23.4272 | -0.032438 | 0.1378 | 3.6e-05 | 0.002406 | 0.002406 | 2.308
2019-04-09 14:55 | 1554836126.5406 | 1245 | 0.2096 | 0.000456 | 3e-06 | 100.1938 | 100.6845 | 121.4781 | 20.5902 | 0.002245 | 0.002207 | -154.643 | 0.23106 | 8.4031 | 0.002638 | 0.004344 | 0.004253 | 4.5e-05 | 0.000961 | 0.31133 | 1.3625 | 2.5476 | 98.5 | 20.08 | 49.599 | 75.1909 | 84.6375 | 74.2797 | 0.011844 | 1.0162 | 0.02383 | 23.4508 | -0.043706 | 0.13746 | -0.000406 | -0.029183 | -0.029183 | -0.18982
2019-04-09 14:55 | 1554836127.5406 | 1246 | 0.20965 | 0.000457 | 3e-06 | 100.1571 | 100.5878 | 121.5828 | 20.5904 | 0.002298 | 0.002137 | -154.6437 | 0.23204 | 8.4022 | 0.002593 | 0.004293 | 0.004143 | 4.3e-05 | 0.000947 | 0.29894 | 1.3703 | 2.5477 | 98.5 | 20.08 | 49.598 | 76.4484 | 83.5556 | 72.5868 | 0.011844 | 1.0211 | 0.023898 | 23.4036 | -0.038381 | 0.13804 | -0.00071 | -0.051043 | -0.051043 | -0.10898
2019-04-09 14:55 | 1554836128.5423 | 1247 | 0.20958 | 0.000459 | 3e-06 | 100.2423 | 100.7565 | 121.429 | 20.5893 | 0.002298 | 0.00214 | -154.6485 | 0.23216 | 8.4009 | 0.002678 | 0.004015 | 0.003808 | 4.6e-05 | 0.000943 | 0.26344 | 1.2952 | 2.5477 | 98.5 | 20.08 | 49.597 | 76.4629 | 85.5927 | 72.6667 | 0.011843 | 1.0209 | 0.02388 | 23.3915 | -0.045761 | 0.13811 | -0.0003 | -0.021649 | -0.021649 | -0.2571
2019-04-09 14:55 | 1554836129.5416 | 1248 | 0.2095 | 0.000462 | 3e-06 | 100.2084 | 100.6918 | 121.575 | 20.5907 | 0.002297 | 0.002152 | -154.6464 | 0.23222 | 8.4005 | 0.002655 | 0.00458 | 0.004378 | 4.7e-05 | 0.000938 | 0.28766 | 1.3168 | 2.5475 | 98.5 | 20.08 | 49.596 | 76.4299 | 85.0461 | 72.9486 | 0.011843 | 1.0201 | 0.023885 | 23.4151 | -0.043045 | 0.13815 | 0.000179 | 0.01262 | 0.01262 | 0.44117
2019-04-09 14:55 | 1554836130.541 | 1249 | 0.20953 | 0.000455 | 2e-06 | 100.2298 | 100.8449 | 121.7682 | 20.5907 | 0.002265 | 0.002154 | -154.645 | 0.23164 | 8.4029 | 0.002645 | 0.004203 | 0.004003 | 4.6e-05 | 0.000938 | 0.2912 | 1.3233 | 2.5476 | 98.5 | 20.08 | 49.593 | 75.663 | 84.7989 | 73.0042 | 0.011843 | 1.0199 | 0.023902 | 23.4356 | -0.054726 | 0.1378 | 3.6e-05 | 0.002396 | 0.002396 | 2.3176
2019-04-09 14:55 | 1554836131.5417 | 1250 | 0.20955 | 0.000453 | 2e-06 | 100.2533 | 100.7967 | 121.6288 | 20.5911 | 0.002282 | 0.00215 | -154.6323 | 0.23192 | 8.4029 | 0.002689 | 0.004011 | 0.003833 | 4.9e-05 | 0.000934 | 0.28174 | 1.3751 | 2.5475 | 98.5 | 20.08 | 49.593 | 76.0781 | 85.8489 | 72.8976 | 0.011843 | 1.0202 | 0.023892 | 23.4186 | -0.048352 | 0.13797 | -6.9e-05 | -0.005059 | -0.005059 | -1.0991
2019-04-09 14:56 | 1554836132.5407 | 1251 | 0.20956 | 0.00045 | 3e-06 | 100.2733 | 100.7644 | 121.3011 | 20.5903 | 0.002262 | 0.002102 | -154.6318 | 0.23167 | 8.4016 | 0.002675 | 0.004073 | 0.003888 | 5e-05 | 0.000926 | 0.29822 | 1.384 | 2.5476 | 98.5 | 20.07 | 49.593 | 75.5995 | 85.5093 | 71.7455 | 0.011835 | 1.0236 | 0.0239 | 23.3481 | -0.043701 | 0.13782 | -0.00018 | -0.013031 | -0.013031 | -0.42624
2019-04-09 14:56 | 1554836133.5407 | 1252 | 0.20956 | 0.000453 | 2e-06 | 100.295 | 100.7389 | 121.0204 | 20.591 | 0.002246 | 0.002176 | -154.6202 | 0.23184 | 8.4043 | 0.002637 | 0.003752 | 0.003548 | 5.2e-05 | 0.000928 | 0.26233 | 1.3735 | 2.5477 | 98.5 | 20.07 | 49.594 | 75.2175 | 84.5993 | 73.5369 | 0.011835 | 1.0183 | 0.02381 | 23.3815 | -0.039501 | 0.13792 | -0.000148 | -0.01069 | -0.01069 | -0.51993
2019-04-09 14:56 | 1554836134.5412 | 1253 | 0.20955 | 0.000453 | 3e-06 | 100.256 | 100.7121 | 120.9647 | 20.5913 | 0.00227 | 0.002154 | -154.6207 | 0.23113 | 8.4001 | 0.002655 | 0.004191 | 0.00405 | 4.9e-05 | 0.000921 | 0.28528 | 1.3455 | 2.5475 | 98.5 | 20.07 | 49.594 | 75.7797 | 85.0383 | 73.0007 | 0.011835 | 1.0199 | 0.023823 | 23.358 | -0.040597 | 0.1375 | -0.000127 | -0.009196 | -0.009196 | -0.60258
2019-04-09 14:56 | 1554836135.5424 | 1254 | 0.20955 | 0.000455 | 2e-06 | 100.2938 | 100.816 | 121.1804 | 20.5913 | 0.002258 | 0.002119 | -154.6149 | 0.2314 | 8.4023 | 0.002663 | 0.00402 | 0.003816 | 4.7e-05 | 0.00092 | 0.29302 | 1.3075 | 2.5475 | 98.5 | 20.07 | 49.595 | 75.496 | 85.2341 | 72.1682 | 0.011836 | 1.0224 | 0.023873 | 23.3507 | -0.046449 | 0.13766 | -0.000107 | -0.007812 | -0.007812 | -0.71012
2019-04-09 14:56 | 1554836136.5409 | 1255 | 0.20955 | 0.000451 | 3e-06 | 100.3266 | 100.8095 | 121.4602 | 20.5908 | 0.002277 | 0.002083 | -154.6208 | 0.23225 | 8.4013 | 0.002574 | 0.004563 | 0.004396 | 4.8e-05 | 0.000914 | 0.28031 | 1.3801 | 2.5476 | 98.5 | 20.08 | 49.597 | 75.9428 | 83.0982 | 71.2923 | 0.011843 | 1.025 | 0.023931 | 23.348 | -0.042948 | 0.13817 | -0.000101 | -0.007393 | -0.007393 | -0.75312
2019-04-09 14:56 | 1554836137.5409 | 1256 | 0.20954 | 0.00045 | 2e-06 | 100.3575 | 100.8927 | 121.2595 | 20.5916 | 0.00229 | 0.002131 | -154.629 | 0.23157 | 8.4059 | 0.002612 | 0.003887 | 0.003685 | 5.2e-05 | 0.000911 | 0.27417 | 1.3574 | 2.5475 | 98.5 | 20.08 | 49.598 | 76.2664 | 84.0147 | 72.4562 | 0.011844 | 1.0215 | 0.023871 | 23.3681 | -0.047574 | 0.13776 | -5.6e-05 | -0.004099 | -0.004099 | -1.3543
2019-04-09 14:56 | 1554836138.5419 | 1257 | 0.20957 | 0.00045 | 2e-06 | 100.3501 | 100.8243 | 120.9178 | 20.5919 | 0.002247 | 0.002079 | -154.6307 | 0.23236 | 8.3995 | 0.002518 | 0.003114 | 0.002912 | 5.1e-05 | 0.00091 | 0.27759 | 1.3563 | 2.5474 | 98.5 | 20.08 | 49.601 | 75.2277 | 81.7374 | 71.1958 | 0.011844 | 1.0253 | 0.023881 | 23.2925 | -0.042164 | 0.13823 | -0.00024 | -0.01732 | -0.01732 | -0.32163
2019-04-09 14:56 | 1554836139.5415 | 1258 | 0.20957 | 0.000449 | 2e-06 | 100.3578 | 100.751 | 120.4807 | 20.5922 | 0.002191 | 0.002104 | -154.6457 | 0.23183 | 8.3997 | 0.00257 | 0.00349 | 0.003349 | 5.1e-05 | 0.000908 | 0.27223 | 1.3369 | 2.5474 | 98.5 | 20.08 | 49.603 | 73.8763 | 82.9918 | 71.7925 | 0.011845 | 1.0235 | 0.023817 | 23.2705 | -0.034968 | 0.13792 | -0.000186 | -0.013381 | -0.013381 | -0.41535
2019-04-09 14:56 | 1554836140.542 | 1259 | 0.20956 | 0.000447 | 2e-06 | 100.3885 | 100.8718 | 120.4225 | 20.5919 | 0.002194 | 0.002134 | -154.6478 | 0.23206 | 8.4002 | 0.002633 | 0.003974 | 0.003788 | 5.3e-05 | 0.000897 | 0.28567 | 1.3561 | 2.5476 | 98.5 | 20.08 | 49.608 | 73.9681 | 84.5167 | 72.5102 | 0.011846 | 1.0214 | 0.023787 | 23.2891 | -0.042953 | 0.13805 | -0.000139 | -0.01008 | -0.01008 | -0.55194
2019-04-09 14:56 | 1554836141.542 | 1260 | 0.20954 | 0.000453 | 2e-06 | 100.2983 | 101.0449 | 120.7011 | 20.5918 | 0.002227 | 0.002069 | -154.6767 | 0.23174 | 8.4 | 0.002616 | 0.00362 | 0.003418 | 5.1e-05 | 0.000894 | 0.29717 | 1.3768 | 2.5475 | 98.5 | 20.08 | 49.611 | 74.7598 | 84.0922 | 70.9563 | 0.011847 | 1.026 | 0.023868 | 23.2636 | -0.066333 | 0.13786 | -4.1e-05 | -0.00308 | -0.00308 | -1.8038
2019-04-09 14:56 | 1554836142.5405 | 1261 | 0.20953 | 0.000451 | 2e-06 | 100.3215 | 100.8425 | 121.2187 | 20.592 | 0.002275 | 0.002102 | -154.6719 | 0.23166 | 8.4045 | 0.00266 | 0.004225 | 0.004037 | 5.4e-05 | 0.000894 | 0.29087 | 1.3749 | 2.5475 | 98.5 | 20.08 | 49.615 | 75.9144 | 85.1512 | 71.761 | 0.011848 | 1.0236 | 0.023891 | 23.3406 | -0.04633 | 0.13781 | 3.4e-05 | 0.002281 | 0.002281 | 2.4347
2019-04-09 14:56 | 1554836143.5415 | 1262 | 0.20953 | 0.000452 | 2e-06 | 100.4222 | 100.8611 | 121.3415 | 20.5924 | 0.002229 | 0.00207 | -154.6832 | 0.23203 | 8.3999 | 0.002529 | 0.004184 | 0.004056 | 5.6e-05 | 0.000897 | 0.28031 | 1.3872 | 2.5476 | 98.5 | 20.08 | 49.622 | 74.8094 | 82.0029 | 70.9694 | 0.011849 | 1.0259 | 0.023931 | 23.3256 | -0.039009 | 0.13803 | 6e-06 | 0.000288 | 0.000288 | 19.3451
2019-04-09 14:56 | 1554836144.5401 | 1263 | 0.20952 | 0.000449 | 2e-06 | 100.3626 | 100.987 | 121.6201 | 20.5924 | 0.002217 | 0.002137 | -154.6915 | 0.23185 | 8.4036 | 0.002637 | 0.003891 | 0.003694 | 5.4e-05 | 0.0009 | 0.29391 | 1.3756 | 2.5476 | 98.5 | 20.08 | 49.625 | 74.5195 | 84.5952 | 72.5926 | 0.01185 | 1.0211 | 0.023902 | 23.4074 | -0.055479 | 0.13793 | 6.6e-05 | 0.004554 | 0.004554 | 1.2206
2019-04-09 14:56 | 1554836145.5412 | 1264 | 0.20954 | 0.00045 | 3e-06 | 100.4216 | 100.8515 | 121.5374 | 20.5919 | 0.002221 | 0.002072 | -154.7109 | 0.23151 | 8.4027 | 0.002677 | 0.003791 | 0.003588 | 5.3e-05 | 0.000897 | 0.2912 | 1.3517 | 2.5475 | 98.5 | 20.08 | 49.629 | 74.5978 | 85.5671 | 71.0189 | 0.011851 | 1.0258 | 0.023948 | 23.3462 | -0.038208 | 0.13772 | -4.5e-05 | -0.003362 | -0.003362 | -1.651
2019-04-09 14:56 | 1554836146.5412 | 1265 | 0.20954 | 0.000451 | 2e-06 | 100.3977 | 100.9069 | 121.0049 | 20.5922 | 0.002239 | 0.002021 | -154.7049 | 0.23227 | 8.4025 | 0.002559 | 0.004109 | 0.003944 | 5.2e-05 | 0.000897 | 0.27383 | 1.3183 | 2.5475 | 98.5 | 20.08 | 49.631 | 75.0433 | 82.7399 | 69.7974 | 0.011852 | 1.0294 | 0.023938 | 23.2536 | -0.045243 | 0.13818 | -3.9e-05 | -0.002903 | -0.002903 | -1.9182
2019-04-09 14:56 | 1554836147.5416 | 1266 | 0.20953 | 0.000448 | 3e-06 | 100.3447 | 100.9058 | 120.783 | 20.5921 | 0.002214 | 0.002092 | -154.7228 | 0.23188 | 8.4058 | 0.002614 | 0.003979 | 0.003775 | 5.2e-05 | 0.000898 | 0.28224 | 1.3455 | 2.5475 | 98.5 | 20.08 | 49.634 | 74.4343 | 84.0528 | 71.504 | 0.011852 | 1.0243 | 0.023857 | 23.29 | -0.049875 | 0.13795 | 1.9e-05 | 0.001225 | 0.001225 | 4.5397
2019-04-09 14:56 | 1554836148.5412 | 1267 | 0.20952 | 0.000448 | 3e-06 | 100.3078 | 100.8174 | 121.1973 | 20.5927 | 0.002276 | 0.002086 | -154.7129 | 0.23221 | 8.4023 | 0.002602 | 0.00363 | 0.003515 | 5.3e-05 | 0.0009 | 0.27975 | 1.3441 | 2.5476 | 98.5 | 20.08 | 49.639 | 75.929 | 83.7621 | 71.3681 | 0.011854 | 1.0248 | 0.023903 | 23.3253 | -0.045323 | 0.13814 | 7.9e-05 | 0.005475 | 0.005475 | 1.0168
2019-04-09 14:56 | 1554836149.5413 | 1268 | 0.20955 | 0.000454 | 2e-06 | 100.3805 | 101.093 | 121.51 | 20.5927 | 0.002245 | 0.002055 | -154.7087 | 0.23235 | 8.4001 | 0.002629 | 0.003832 | 0.003622 | 4.7e-05 | 0.000908 | 0.28263 | 1.3909 | 2.5477 | 98.5 | 20.08 | 49.641 | 75.1886 | 84.4254 | 70.6119 | 0.011854 | 1.027 | 0.02396 | 23.3297 | -0.06327 | 0.13822 | -8.7e-05 | -0.006383 | -0.006383 | -0.87263
2019-04-09 14:56 | 1554836150.5413 | 1269 | 0.20955 | 0.000452 | 2e-06 | 100.3513 | 100.7768 | 121.5858 | 20.5932 | 0.002217 | 0.002074 | -154.6831 | 0.23187 | 8.4009 | 0.002602 | 0.00312 | 0.002925 | 4.9e-05 | 0.000916 | 0.27417 | 1.3753 | 2.5476 | 98.5 | 20.08 | 49.643 | 74.5133 | 83.7644 | 71.0724 | 0.011854 | 1.0256 | 0.023951 | 23.3526 | -0.037843 | 0.13794 | -0.000112 | -0.008209 | -0.008209 | -0.67717
2019-04-09 14:56 | 1554836151.5402 | 1270 | 0.20955 | 0.000452 | 2e-06 | 100.3507 | 100.8786 | 121.602 | 20.5934 | 0.002183 | 0.002101 | -154.6841 | 0.23243 | 8.4001 | 0.002632 | 0.00371 | 0.003516 | 4.8e-05 | 0.000924 | 0.28672 | 1.3979 | 2.5477 | 98.5 | 20.08 | 49.645 | 73.6879 | 84.4816 | 71.7297 | 0.011855 | 1.0237 | 0.02393 | 23.3765 | -0.046934 | 0.13828 | -9.9e-05 | -0.007242 | -0.007242 | -0.76942
2019-04-09 14:56 | 1554836152.5403 | 1271 | 0.20953 | 0.000446 | 2e-06 | 100.3903 | 100.889 | 121.6384 | 20.5925 | 0.002221 | 0.002092 | -154.6882 | 0.23162 | 8.4005 | 0.002713 | 0.004107 | 0.003952 | 5.2e-05 | 0.000931 | 0.28694 | 1.3627 | 2.5476 | 98.5 | 20.08 | 49.644 | 74.6101 | 86.4419 | 71.4989 | 0.011855 | 1.0244 | 0.023942 | 23.3721 | -0.044319 | 0.13779 | 4e-05 | 0.002769 | 0.002769 | 2.0056
2019-04-09 14:56 | 1554836153.5418 | 1272 | 0.20956 | 0.000453 | 2e-06 | 100.4537 | 101.0173 | 121.6295 | 20.5932 | 0.00219 | 0.002059 | -154.7103 | 0.23237 | 8.4035 | 0.002672 | 0.004722 | 0.004545 | 5e-05 | 0.000941 | 0.29302 | 1.3708 | 2.5474 | 98.5 | 20.08 | 49.645 | 73.8672 | 85.4442 | 70.7277 | 0.011855 | 1.0267 | 0.023967 | 23.3451 | -0.050043 | 0.13824 | -0.000129 | -0.009348 | -0.009348 | -0.59593
2019-04-09 14:56 | 1554836154.5405 | 1273 | 0.20953 | 0.000452 | 2e-06 | 100.4135 | 100.9568 | 121.3365 | 20.5934 | 0.002169 | 0.002087 | -154.6969 | 0.23203 | 8.4011 | 0.002679 | 0.004 | 0.003819 | 4.8e-05 | 0.000949 | 0.28672 | 1.3982 | 2.5477 | 98.5 | 20.08 | 49.643 | 73.3646 | 85.6107 | 71.397 | 0.011854 | 1.0247 | 0.023915 | 23.3397 | -0.048259 | 0.13803 | 1.7e-05 | 0.001114 | 0.001114 | 4.9953
2019-04-09 14:56 | 1554836155.5415 | 1274 | 0.20952 | 0.000447 | 2e-06 | 100.401 | 100.881 | 120.9605 | 20.5934 | 0.00217 | 0.002013 | -154.6935 | 0.23226 | 8.4023 | 0.002676 | 0.003762 | 0.003579 | 4.9e-05 | 0.000957 | 0.28296 | 1.3947 | 2.5477 | 98.5 | 20.08 | 49.642 | 73.3941 | 85.5404 | 69.6191 | 0.011854 | 1.03 | 0.02394 | 23.2433 | -0.042662 | 0.13817 | 7.9e-05 | 0.005537 | 0.005537 | 1.0056
2019-04-09 14:56 | 1554836156.5405 | 1275 | 0.20955 | 0.000446 | 2e-06 | 100.424 | 101.007 | 121.1671 | 20.5931 | 0.002211 | 0.002054 | -154.6794 | 0.23171 | 8.4021 | 0.002647 | 0.004526 | 0.004328 | 4.7e-05 | 0.000963 | 0.28423 | 1.3815 | 2.5476 | 98.5 | 20.07 | 49.641 | 74.3635 | 84.8395 | 70.5991 | 0.011847 | 1.027 | 0.023926 | 23.2963 | -0.051781 | 0.13784 | -8.3e-05 | -0.006116 | -0.006116 | -0.90827
2019-04-09 14:56 | 1554836157.5405 | 1276 | 0.20951 | 0.000452 | 2e-06 | 100.4782 | 100.9389 | 121.5898 | 20.593 | 0.002246 | 0.002038 | -154.6808 | 0.23215 | 8.4025 | 0.002622 | 0.004074 | 0.003869 | 4.2e-05 | 0.000969 | 0.27793 | 1.3774 | 2.5477 | 98.5 | 20.08 | 49.64 | 75.2194 | 84.2477 | 70.2068 | 0.011854 | 1.0282 | 0.023982 | 23.3236 | -0.040919 | 0.13811 | 0.000124 | 0.008805 | 0.008805 | 0.63215
2019-04-09 14:56 | 1554836158.5395 | 1277 | 0.20954 | 0.000447 | 2e-06 | 100.46 | 100.9406 | 121.828 | 20.5943 | 0.002159 | 0.002055 | -154.6812 | 0.2323 | 8.3995 | 0.002663 | 0.003694 | 0.003603 | 4.8e-05 | 0.000974 | 0.26952 | 1.4228 | 2.5478 | 98.51 | 20.07 | 49.637 | 73.1206 | 85.2221 | 70.6134 | 0.011844 | 1.027 | 0.023991 | 23.3603 | -0.042689 | 0.13819 | -1.7e-05 | -0.001305 | -0.001305 | -4.2667
2019-04-09 14:56 | 1554836159.5399 | 1278 | 0.20951 | 0.00045 | 2e-06 | 100.4764 | 101.0163 | 122.251 | 20.5933 | 0.002195 | 0.002002 | -154.6996 | 0.23234 | 8.3978 | 0.00257 | 0.003625 | 0.003365 | 4.7e-05 | 0.000984 | 0.28871 | 1.3809 | 2.5475 | 98.51 | 20.08 | 49.636 | 73.9889 | 82.989 | 69.3414 | 0.011852 | 1.0308 | 0.024077 | 23.3575 | -0.047937 | 0.13822 | 0.000167 | 0.011964 | 0.011964 | 0.46557
2019-04-09 14:56 | 1554836160.5407 | 1279 | 0.20951 | 0.000455 | 2e-06 | 100.4326 | 101.1399 | 122.2498 | 20.5939 | 0.002161 | 0.002014 | -154.7292 | 0.23194 | 8.4013 | 0.002614 | 0.003757 | 0.003596 | 4.7e-05 | 0.000987 | 0.28838 | 1.3527 | 2.5476 | 98.51 | 20.08 | 49.635 | 73.1803 | 84.0577 | 69.6383 | 0.011851 | 1.0299 | 0.024067 | 23.3675 | -0.062768 | 0.13798 | 0.000144 | 0.010245 | 0.010245 | 0.54277
2019-04-09 14:56 | 1554836161.5406 | 1280 | 0.20952 | 0.000448 | 2e-06 | 100.4222 | 100.9674 | 122.0707 | 20.5931 | 0.002155 | 0.002063 | -154.7225 | 0.2325 | 8.3985 | 0.002687 | 0.003658 | 0.003468 | 4.5e-05 | 0.000993 | 0.28672 | 1.3683 | 2.5476 | 98.51 | 20.08 | 49.635 | 73.033 | 85.818 | 70.8026 | 0.011851 | 1.0264 | 0.024008 | 23.39 | -0.048435 | 0.13831 | 8.3e-05 | 0.005833 | 0.005833 | 0.95565
2019-04-09 14:56 | 1554836162.5407 | 1281 | 0.20952 | 0.00044 | 2e-06 | 100.4347 | 100.9747 | 121.6109 | 20.5939 | 0.002154 | 0.001997 | -154.7113 | 0.2318 | 8.3998 | 0.002626 | 0.003844 | 0.003691 | 4.7e-05 | 0.000998 | 0.30541 | 1.3704 | 2.5476 | 98.51 | 20.08 | 49.636 | 73.0032 | 84.3323 | 69.2345 | 0.011852 | 1.0311 | 0.024018 | 23.2926 | -0.047961 | 0.13789 | 0.000113 | 0.008002 | 0.008002 | 0.69445
2019-04-09 14:56 | 1554836163.5415 | 1282 | 0.20956 | 0.000447 | 2e-06 | 100.4198 | 100.9895 | 121.2774 | 20.594 | 0.00217 | 0.002006 | -154.7045 | 0.23181 | 8.4021 | 0.002645 | 0.003646 | 0.00349 | 4.9e-05 | 0.001 | 0.2844 | 1.3599 | 2.5475 | 98.51 | 20.08 | 49.637 | 73.3908 | 84.7973 | 69.435 | 0.011852 | 1.0305 | 0.023978 | 23.2674 | -0.050598 | 0.1379 | -0.000142 | -0.010363 | -0.010363 | -0.53631
2019-04-09 14:56 | 1554836164.5828 | 1283 | 0.20951 | 0.000448 | 2e-06 | 100.3927 | 100.9939 | 121.4199 | 20.5934 | 0.002203 | 0.002032 | -154.7176 | 0.23239 | 8.4031 | 0.002609 | 0.00357 | 0.003402 | 4.6e-05 | 0.000999 | 0.28119 | 1.3667 | 2.5476 | 98.51 | 20.08 | 49.636 | 74.1714 | 83.9432 | 70.063 | 0.011852 | 1.0286 | 0.02397 | 23.3024 | -0.053408 | 0.13825 | 0.00018 | 0.012841 | 0.012841 | 0.43389
2019-04-09 14:56 | 1554836165.5588 | 1284 | 0.20948 | 0.00044 | 2e-06 | 100.4394 | 100.8208 | 121.9862 | 20.5946 | 0.002187 | 0.002048 | -154.7033 | 0.23267 | 8.4013 | 0.002628 | 0.003915 | 0.003695 | 4.9e-05 | 0.000999 | 0.28368 | 1.3493 | 2.5476 | 98.51 | 20.08 | 49.634 | 73.7967 | 84.3977 | 70.4547 | 0.011851 | 1.0275 | 0.024012 | 23.37 | -0.033899 | 0.13842 | 0.000351 | 0.025135 | 0.025135 | 0.22193
2019-04-09 14:56 | 1554836166.5589 | 1285 | 0.2095 | 0.000448 | 1e-06 | 100.4305 | 100.9633 | 122.1703 | 20.595 | 0.002224 | 0.001999 | -154.7198 | 0.23197 | 8.4019 | 0.002584 | 0.003881 | 0.00368 | 4.8e-05 | 0.001001 | 0.28313 | 1.3565 | 2.5475 | 98.51 | 20.08 | 49.632 | 74.6858 | 83.3236 | 69.2831 | 0.011851 | 1.031 | 0.024071 | 23.3478 | -0.047326 | 0.138 | 0.000186 | 0.013316 | 0.013316 | 0.41764
2019-04-09 14:56 | 1554836167.5598 | 1286 | 0.20953 | 0.00045 | 2e-06 | 100.3995 | 100.9148 | 121.8755 | 20.595 | 0.002159 | 0.00204 | -154.7029 | 0.23187 | 8.4006 | 0.002629 | 0.004352 | 0.004163 | 4.6e-05 | 0.001009 | 0.286 | 1.337 | 2.5476 | 98.51 | 20.08 | 49.632 | 73.1261 | 84.4207 | 70.2719 | 0.011851 | 1.028 | 0.024008 | 23.3532 | -0.045785 | 0.13794 | 6e-06 | 0.000329 | 0.000329 | 16.8804
2019-04-09 14:56 | 1554836168.5601 | 1287 | 0.20953 | 0.000452 | 2e-06 | 100.3796 | 100.8181 | 121.6046 | 20.5949 | 0.002136 | 0.002061 | -154.7056 | 0.23201 | 8.4019 | 0.002608 | 0.004393 | 0.004096 | 4.7e-05 | 0.001009 | 0.29607 | 1.3135 | 2.5474 | 98.51 | 20.08 | 49.629 | 72.5724 | 83.9008 | 70.7724 | 0.01185 | 1.0265 | 0.023963 | 23.3442 | -0.038987 | 0.13802 | 4.6e-05 | 0.003145 | 0.003145 | 1.7684
2019-04-09 14:56 | 1554836169.5589 | 1288 | 0.20949 | 0.000452 | 2e-06 | 100.3656 | 100.9702 | 121.3656 | 20.5948 | 0.00211 | 0.002032 | -154.6846 | 0.23244 | 8.4011 | 0.00255 | 0.004292 | 0.004148 | 4.8e-05 | 0.001013 | 0.286 | 1.3628 | 2.5475 | 98.5 | 20.08 | 49.628 | 71.9447 | 82.5224 | 70.0744 | 0.011851 | 1.0286 | 0.023964 | 23.2976 | -0.05372 | 0.13828 | 0.000286 | 0.020422 | 0.020422 | 0.27287
2019-04-09 14:56 | 1554836170.56 | 1289 | 0.20953 | 0.000449 | 3e-06 | 100.3349 | 100.7492 | 121.5245 | 20.5944 | 0.002101 | 0.002039 | -154.6851 | 0.23105 | 8.4018 | 0.002628 | 0.003411 | 0.003289 | 4.4e-05 | 0.001016 | 0.29551 | 1.3764 | 2.5476 | 98.5 | 20.08 | 49.627 | 71.7245 | 84.3881 | 70.229 | 0.011851 | 1.0282 | 0.023975 | 23.3181 | -0.03686 | 0.13745 | 5.3e-05 | 0.003649 | 0.003649 | 1.5181
2019-04-09 14:56 | 1554836171.5603 | 1290 | 0.20956 | 0.000447 | 3e-06 | 100.359 | 100.8415 | 121.6525 | 20.5951 | 0.002131 | 0.002032 | -154.6936 | 0.23256 | 8.4002 | 0.002641 | 0.004337 | 0.004125 | 4.8e-05 | 0.001016 | 0.27815 | 1.3623 | 2.5476 | 98.5 | 20.08 | 49.62 | 72.4525 | 84.7103 | 70.0723 | 0.011849 | 1.0286 | 0.023993 | 23.3251 | -0.042894 | 0.13835 | -0.000152 | -0.011055 | -0.011055 | -0.50435
2019-04-09 14:56 | 1554836172.5581 | 1291 | 0.20959 | 0.000453 | 3e-06 | 100.3528 | 100.8921 | 122.0841 | 20.5958 | 0.00208 | 0.002029 | -154.7014 | 0.23182 | 8.4014 | 0.002558 | 0.004148 | 0.003998 | 4.4e-05 | 0.00102 | 0.30646 | 1.3494 | 2.5475 | 98.5 | 20.08 | 49.617 | 71.2252 | 82.72 | 69.9878 | 0.011848 | 1.0289 | 0.024038 | 23.3635 | -0.047938 | 0.13791 | -0.000347 | -0.02518 | -0.02518 | -0.22072
2019-04-09 14:56 | 1554836173.558 | 1292 | 0.20954 | 0.000451 | 3e-06 | 100.3441 | 100.9557 | 122.439 | 20.5949 | 0.002117 | 0.001994 | -154.6802 | 0.23129 | 8.4014 | 0.00262 | 0.004054 | 0.003889 | 4.4e-05 | 0.001021 | 0.29518 | 1.3324 | 2.5475 | 98.5 | 20.08 | 49.613 | 72.1195 | 84.1991 | 69.1504 | 0.011847 | 1.0314 | 0.024102 | 23.3689 | -0.05435 | 0.13759 | -1.8e-05 | -0.001429 | -0.001429 | -3.8799
2019-04-09 14:56 | 1554836174.5591 | 1293 | 0.20926 | 0.000451 | 2e-06 | 100.3117 | 100.8862 | 122.1698 | 20.5954 | 0.002081 | 0.002084 | -154.6832 | 0.23222 | 8.4019 | 0.002678 | 0.003812 | 0.003647 | 4.7e-05 | 0.001026 | 0.29806 | 1.333 | 2.5476 | 98.5 | 20.08 | 49.605 | 71.26 | 85.6063 | 71.3114 | 0.011845 | 1.0249 | 0.024 | 23.4168 | -0.051083 | 0.13815 | 0.001654 | 0.11876 | 0.11876 | 0.046881
2019-04-09 14:56 | 1554836175.5595 | 1294 | 0.20959 | 0.000451 | 3e-06 | 100.3593 | 100.7981 | 121.9186 | 20.5955 | 0.00213 | 0.00202 | -154.6903 | 0.2321 | 8.3998 | 0.002618 | 0.003717 | 0.00351 | 4.5e-05 | 0.001032 | 0.27903 | 1.3412 | 2.5476 | 98.5 | 20.08 | 49.6 | 72.4341 | 84.1496 | 69.7881 | 0.011844 | 1.0295 | 0.024029 | 23.3409 | -0.039021 | 0.13808 | -0.000334 | -0.024228 | -0.024228 | -0.22967
2019-04-09 14:56 | 1554836176.5595 | 1295 | 0.20949 | 0.000452 | 2e-06 | 100.3781 | 100.783 | 121.9208 | 20.5963 | 0.002115 | 0.001989 | -154.6803 | 0.23159 | 8.4022 | 0.00257 | 0.003989 | 0.003823 | 4.7e-05 | 0.001032 | 0.28351 | 1.3296 | 2.5476 | 98.5 | 20.08 | 49.595 | 72.0584 | 82.9934 | 69.0332 | 0.011843 | 1.0317 | 0.024056 | 23.3154 | -0.036005 | 0.13778 | 0.000299 | 0.021443 | 0.021443 | 0.25893
2019-04-09 14:56 | 1554836177.5581 | 1296 | 0.2095 | 0.000452 | 2e-06 | 100.3915 | 100.9347 | 122.1717 | 20.5957 | 0.002113 | 0.002015 | -154.6861 | 0.23306 | 8.4009 | 0.002591 | 0.004148 | 0.004021 | 4.7e-05 | 0.001027 | 0.28423 | 1.4054 | 2.5476 | 98.5 | 20.08 | 49.593 | 72.0158 | 83.5131 | 69.673 | 0.011843 | 1.0298 | 0.024058 | 23.3612 | -0.048271 | 0.13865 | 0.000232 | 0.016587 | 0.016587 | 0.33686
2019-04-09 14:56 | 1554836178.5582 | 1297 | 0.20954 | 0.000451 | 1e-06 | 100.3641 | 100.9344 | 121.9883 | 20.5959 | 0.002124 | 0.002018 | -154.6778 | 0.2321 | 8.4019 | 0.002604 | 0.003927 | 0.003756 | 4.6e-05 | 0.001035 | 0.26521 | 1.3577 | 2.5475 | 98.5 | 20.08 | 49.589 | 72.2716 | 83.8129 | 69.7416 | 0.011842 | 1.0296 | 0.024037 | 23.346 | -0.050681 | 0.13808 | -2e-05 | -0.001499 | -0.001499 | -3.7113
2019-04-09 14:56 | 1554836179.5582 | 1298 | 0.20957 | 0.000451 | 2e-06 | 100.3903 | 100.826 | 121.8726 | 20.5968 | 0.002155 | 0.001993 | -154.6861 | 0.2325 | 8.3985 | 0.002556 | 0.004421 | 0.004287 | 4.6e-05 | 0.001034 | 0.29015 | 1.357 | 2.5475 | 98.5 | 20.08 | 49.581 | 73.0216 | 82.657 | 69.124 | 0.01184 | 1.0315 | 0.024048 | 23.3139 | -0.038734 | 0.13832 | -0.000231 | -0.016765 | -0.016765 | -0.33248
2019-04-09 14:56 | 1554836180.5597 | 1299 | 0.20953 | 0.000453 | 3e-06 | 100.3838 | 100.7695 | 121.837 | 20.596 | 0.002055 | 0.00199 | -154.7055 | 0.23239 | 8.4028 | 0.002584 | 0.004092 | 0.003883 | 4.6e-05 | 0.001035 | 0.30735 | 1.3355 | 2.5475 | 98.5 | 20.08 | 49.576 | 70.6353 | 83.3473 | 69.0632 | 0.011838 | 1.0317 | 0.024046 | 23.3084 | -0.034309 | 0.13825 | 4.9e-05 | 0.00336 | 0.00336 | 1.658
2019-04-09 14:56 | 1554836181.5593 | 1300 | 0.20954 | 0.000452 | 2e-06 | 100.3852 | 100.9626 | 121.7158 | 20.5962 | 0.002147 | 0.002017 | -154.6988 | 0.23144 | 8.4023 | 0.002581 | 0.004449 | 0.004355 | 4.6e-05 | 0.001038 | 0.28766 | 1.3981 | 2.5477 | 98.5 | 20.08 | 49.571 | 72.8332 | 83.2688 | 69.7092 | 0.011837 | 1.0297 | 0.024012 | 23.3188 | -0.051298 | 0.13768 | -3.5e-05 | -0.002669 | -0.002669 | -2.0786
2019-04-09 14:56 | 1554836182.5584 | 1301 | 0.20952 | 0.000451 | 3e-06 | 100.3638 | 100.817 | 121.6173 | 20.5967 | 0.002123 | 0.001985 | -154.7052 | 0.2318 | 8.4002 | 0.002597 | 0.004506 | 0.004251 | 4.4e-05 | 0.001038 | 0.28711 | 1.382 | 2.5476 | 98.5 | 20.08 | 49.565 | 72.2491 | 83.6593 | 68.9485 | 0.011836 | 1.032 | 0.024029 | 23.2835 | -0.040302 | 0.1379 | 6.6e-05 | 0.004575 | 0.004575 | 1.2148
2019-04-09 14:56 | 1554836183.5588 | 1302 | 0.20958 | 0.000452 | 3e-06 | 100.3876 | 100.8284 | 121.4461 | 20.5968 | 0.002125 | 0.002003 | -154.7042 | 0.23156 | 8.4017 | 0.002531 | 0.003648 | 0.003572 | 4.7e-05 | 0.00104 | 0.29247 | 1.3248 | 2.5475 | 98.5 | 20.08 | 49.555 | 72.3107 | 82.0729 | 69.3704 | 0.011833 | 1.0307 | 0.023997 | 23.2814 | -0.039187 | 0.13776 | -0.000248 | -0.018001 | -0.018001 | -0.30841
2019-04-09 14:56 | 1554836184.5584 | 1303 | 0.20957 | 0.000452 | 2e-06 | 100.317 | 100.8377 | 121.5889 | 20.596 | 0.002109 | 0.001973 | -154.7026 | 0.23197 | 8.4025 | 0.00265 | 0.003826 | 0.003685 | 4.4e-05 | 0.001039 | 0.2808 | 1.3172 | 2.5475 | 98.5 | 20.08 | 49.548 | 71.9233 | 84.9334 | 68.6625 | 0.011832 | 1.0329 | 0.024036 | 23.271 | -0.046301 | 0.138 | -0.000196 | -0.014281 | -0.014281 | -0.38942
2019-04-09 14:56 | 1554836185.5594 | 1304 | 0.20962 | 0.000443 | 2e-06 | 100.4222 | 100.976 | 121.8884 | 20.5967 | 0.002143 | 0.001962 | -154.6888 | 0.2318 | 8.4015 | 0.002633 | 0.003868 | 0.003746 | 4.6e-05 | 0.001037 | 0.3126 | 1.3255 | 2.5476 | 98.5 | 20.08 | 49.541 | 72.7402 | 84.5253 | 68.38 | 0.01183 | 1.0337 | 0.024075 | 23.29 | -0.049197 | 0.1379 | -0.000509 | -0.036822 | -0.036822 | -0.15093
2019-04-09 14:56 | 1554836186.5588 | 1305 | 0.20959 | 0.000451 | 2e-06 | 100.3156 | 100.7699 | 122.0055 | 20.5967 | 0.002135 | 0.002028 | -154.7169 | 0.23178 | 8.3999 | 0.002612 | 0.003966 | 0.003732 | 4.5e-05 | 0.001039 | 0.30038 | 1.3433 | 2.5475 | 98.5 | 20.09 | 49.528 | 72.5422 | 84.001 | 69.971 | 0.011834 | 1.0289 | 0.024031 | 23.3554 | -0.040418 | 0.13789 | -0.000338 | -0.024427 | -0.024427 | -0.22749
2019-04-09 14:56 | 1554836187.5589 | 1306 | 0.20951 | 0.000452 | 2e-06 | 100.3438 | 100.7816 | 121.9367 | 20.5963 | 0.002083 | 0.001981 | -154.7124 | 0.2325 | 8.4012 | 0.002663 | 0.004215 | 0.004153 | 4.6e-05 | 0.00104 | 0.28241 | 1.3325 | 2.5475 | 98.5 | 20.09 | 49.521 | 71.3053 | 85.2358 | 68.8546 | 0.011833 | 1.0323 | 0.024063 | 23.3108 | -0.038935 | 0.13832 | 0.000138 | 0.009779 | 0.009779 | 0.57
2019-04-09 14:56 | 1554836188.5585 | 1307 | 0.20954 | 0.000451 | 2e-06 | 100.343 | 100.7021 | 122.0362 | 20.5974 | 0.002098 | 0.001981 | -154.7383 | 0.23222 | 8.4029 | 0.002622 | 0.003826 | 0.003503 | 4.5e-05 | 0.001037 | 0.27959 | 1.3081 | 2.5475 | 98.5 | 20.09 | 49.508 | 71.6651 | 84.2575 | 68.8539 | 0.01183 | 1.0323 | 0.024073 | 23.3203 | -0.031956 | 0.13815 | -3.5e-05 | -0.002659 | -0.002659 | -2.0938
2019-04-09 14:56 | 1554836189.558 | 1308 | 0.20952 | 0.000451 | 2e-06 | 100.3858 | 101.0008 | 122.0008 | 20.597 | 0.002116 | 0.001953 | -154.7403 | 0.23255 | 8.4018 | 0.002677 | 0.003602 | 0.00349 | 4.6e-05 | 0.001038 | 0.28296 | 1.3501 | 2.5476 | 98.5 | 20.09 | 49.503 | 72.0896 | 85.5752 | 68.1853 | 0.011828 | 1.0343 | 0.024093 | 23.2941 | -0.054628 | 0.13834 | 6.7e-05 | 0.004738 | 0.004738 | 1.1768
2019-04-09 14:56 | 1554836190.558 | 1309 | 0.20954 | 0.000445 | 2e-06 | 100.3969 | 100.6226 | 121.8877 | 20.5965 | 0.002074 | 0.001992 | -154.7211 | 0.232 | 8.4035 | 0.002617 | 0.004517 | 0.004345 | 4.6e-05 | 0.001032 | 0.2995 | 1.356 | 2.5476 | 98.5 | 20.09 | 49.497 | 71.0849 | 84.1405 | 69.1019 | 0.011827 | 1.0315 | 0.02405 | 23.3146 | -0.020088 | 0.13802 | -1.4e-05 | -0.001141 | -0.001141 | -4.8729
2019-04-09 14:56 | 1554836191.5579 | 1310 | 0.20955 | 0.000453 | 3e-06 | 100.3781 | 100.8415 | 122.0506 | 20.5984 | 0.002027 | 0.001995 | -154.735 | 0.23266 | 8.4019 | 0.002665 | 0.004059 | 0.003959 | 4.4e-05 | 0.001033 | 0.31492 | 1.3455 | 2.5475 | 98.5 | 20.09 | 49.488 | 69.949 | 85.2869 | 69.1756 | 0.011825 | 1.0313 | 0.024063 | 23.3327 | -0.041196 | 0.13841 | -8.8e-05 | -0.006477 | -0.006477 | -0.86124
2019-04-09 14:57 | 1554836192.5586 | 1311 | 0.20954 | 0.000448 | 2e-06 | 100.3677 | 101.1323 | 122.1787 | 20.5979 | 0.002066 | 0.001955 | -154.7391 | 0.23282 | 8.4022 | 0.002662 | 0.004435 | 0.004186 | 4.7e-05 | 0.001032 | 0.29407 | 1.3648 | 2.5475 | 98.5 | 20.09 | 49.484 | 70.8928 | 85.2061 | 68.2188 | 0.011824 | 1.0342 | 0.02411 | 23.3123 | -0.067882 | 0.1385 | -4.2e-05 | -0.003211 | -0.003211 | -1.7381
2019-04-09 14:57 | 1554836193.5581 | 1312 | 0.20954 | 0.000448 | 3e-06 | 100.4207 | 100.7296 | 122.1773 | 20.5981 | 0.0021 | 0.001972 | -154.7391 | 0.23195 | 8.399 | 0.002645 | 0.00456 | 0.004377 | 4.6e-05 | 0.00103 | 0.30702 | 1.393 | 2.5476 | 98.5 | 20.09 | 49.479 | 71.7117 | 84.8037 | 68.6301 | 0.011823 | 1.033 | 0.024095 | 23.3262 | -0.027475 | 0.13799 | 3e-06 | 2.3e-05 | 2.3e-05 | 240.6056
2019-04-09 14:57 | 1554836194.5581 | 1313 | 0.20953 | 0.000449 | 2e-06 | 100.4126 | 100.9674 | 122.4104 | 20.5987 | 0.002111 | 0.002005 | -154.7337 | 0.23273 | 8.402 | 0.002656 | 0.003086 | 0.003 | 4.4e-05 | 0.001027 | 0.29502 | 1.3698 | 2.5476 | 98.5 | 20.09 | 49.477 | 71.9697 | 85.0813 | 69.4199 | 0.011822 | 1.0306 | 0.02409 | 23.3754 | -0.049284 | 0.13845 | 4.4e-05 | 0.003056 | 0.003056 | 1.8257
2019-04-09 14:57 | 1554836195.5578 | 1314 | 0.20959 | 0.000448 | 2e-06 | 100.4585 | 101.0166 | 122.6104 | 20.5987 | 0.002064 | 0.002002 | -154.7618 | 0.2322 | 8.4028 | 0.002686 | 0.004379 | 0.004148 | 4.3e-05 | 0.00103 | 0.28711 | 1.3785 | 2.5475 | 98.5 | 20.09 | 49.475 | 70.8421 | 85.7899 | 69.342 | 0.011822 | 1.0308 | 0.024113 | 23.3918 | -0.049558 | 0.13814 | -0.000335 | -0.024298 | -0.024298 | -0.2291
2019-04-09 14:57 | 1554836196.5591 | 1315 | 0.20956 | 0.000448 | 2e-06 | 100.357 | 100.9072 | 122.381 | 20.5981 | 0.002028 | 0.001962 | -154.7376 | 0.23277 | 8.3999 | 0.002725 | 0.003854 | 0.003646 | 4.3e-05 | 0.001031 | 0.29479 | 1.3836 | 2.5477 | 98.5 | 20.09 | 49.473 | 69.9675 | 86.7339 | 68.3921 | 0.011821 | 1.0337 | 0.024123 | 23.3375 | -0.048908 | 0.13848 | -0.000166 | -0.012156 | -0.012156 | -0.45908
2019-04-09 14:57 | 1554836197.5578 | 1316 | 0.20958 | 0.000447 | 3e-06 | 100.3718 | 100.8835 | 122.3559 | 20.5982 | 0.002095 | 0.001966 | -154.7474 | 0.23209 | 8.4017 | 0.002726 | 0.003987 | 0.003819 | 4.5e-05 | 0.00103 | 0.28495 | 1.3388 | 2.5477 | 98.5 | 20.09 | 49.471 | 71.5988 | 86.7643 | 68.479 | 0.011821 | 1.0334 | 0.024118 | 23.338 | -0.045476 | 0.13807 | -0.000279 | -0.020313 | -0.020313 | -0.27393
2019-04-09 14:57 | 1554836198.5588 | 1317 | 0.20958 | 0.000451 | 2e-06 | 100.404 | 100.8852 | 122.3613 | 20.5988 | 0.002083 | 0.001926 | -154.7313 | 0.23242 | 8.4014 | 0.002667 | 0.004003 | 0.003845 | 4.6e-05 | 0.001033 | 0.29856 | 1.3904 | 2.5477 | 98.5 | 20.09 | 49.47 | 71.2962 | 85.3374 | 67.5358 | 0.011821 | 1.0363 | 0.024152 | 23.3063 | -0.042763 | 0.13827 | -0.00026 | -0.018972 | -0.018972 | -0.2937
2019-04-09 14:57 | 1554836199.5589 | 1318 | 0.2096 | 0.000446 | 2e-06 | 100.3108 | 100.9 | 122.6625 | 20.5987 | 0.002037 | 0.001964 | -154.741 | 0.23198 | 8.4024 | 0.002716 | 0.004239 | 0.003974 | 5e-05 | 0.001031 | 0.28672 | 1.3941 | 2.5474 | 98.5 | 20.09 | 49.467 | 70.1887 | 86.512 | 68.4326 | 0.01182 | 1.0336 | 0.02415 | 23.3657 | -0.052383 | 0.138 | -0.000352 | -0.025578 | -0.025578 | -0.21743
2019-04-09 14:57 | 1554836200.5583 | 1319 | 0.20958 | 0.000456 | 2e-06 | 100.2914 | 100.8876 | 122.6299 | 20.5995 | 0.002109 | 0.001991 | -154.7446 | 0.23207 | 8.4036 | 0.002726 | 0.003959 | 0.003747 | 4.7e-05 | 0.00103 | 0.29087 | 1.348 | 2.5475 | 98.5 | 20.09 | 49.465 | 71.9159 | 86.7649 | 69.0902 | 0.011819 | 1.0316 | 0.024123 | 23.385 | -0.053011 | 0.13806 | -0.000278 | -0.020219 | -0.020219 | -0.27517
2019-04-09 14:57 | 1554836201.558 | 1320 | 0.20956 | 0.000443 | 3e-06 | 100.3343 | 100.6873 | 122.5402 | 20.5984 | 0.00202 | 0.00197 | -154.7292 | 0.23267 | 8.4019 | 0.00275 | 0.004497 | 0.004359 | 4.5e-05 | 0.001027 | 0.28816 | 1.3428 | 2.5476 | 98.51 | 20.09 | 49.46 | 69.7784 | 87.3368 | 68.5766 | 0.011817 | 1.0331 | 0.024133 | 23.3589 | -0.031411 | 0.13842 | -0.000121 | -0.008869 | -0.008869 | -0.62894
2019-04-09 14:57 | 1554836202.5589 | 1321 | 0.20957 | 0.000445 | 2e-06 | 100.3096 | 100.7902 | 122.5104 | 20.5988 | 0.002058 | 0.00198 | -154.7293 | 0.23294 | 8.4039 | 0.002728 | 0.004034 | 0.003873 | 4.5e-05 | 0.001031 | 0.30143 | 1.3496 | 2.5477 | 98.51 | 20.09 | 49.458 | 70.7106 | 86.8025 | 68.8313 | 0.011816 | 1.0324 | 0.024121 | 23.3648 | -0.042752 | 0.13858 | -0.000179 | -0.013071 | -0.013071 | -0.42724
2019-04-09 14:57 | 1554836203.559 | 1322 | 0.20955 | 0.000444 | 2e-06 | 100.2694 | 100.8566 | 122.5648 | 20.5996 | 0.002065 | 0.001947 | -154.7335 | 0.23185 | 8.404 | 0.002738 | 0.004011 | 0.003821 | 4.4e-05 | 0.001033 | 0.2995 | 1.3319 | 2.5476 | 98.51 | 20.09 | 49.455 | 70.8687 | 87.0523 | 68.0279 | 0.011816 | 1.0348 | 0.024154 | 23.3425 | -0.05223 | 0.13793 | -8.7e-05 | -0.006445 | -0.006445 | -0.86242
2019-04-09 14:57 | 1554836204.558 | 1323 | 0.20955 | 0.000442 | 2e-06 | 100.2587 | 100.7141 | 122.3707 | 20.6 | 0.002059 | 0.001931 | -154.7111 | 0.23236 | 8.4037 | 0.002775 | 0.004434 | 0.004284 | 4.5e-05 | 0.001031 | 0.28816 | 1.3573 | 2.5476 | 98.51 | 20.09 | 49.447 | 70.7351 | 87.9336 | 67.6499 | 0.011814 | 1.0359 | 0.024149 | 23.3111 | -0.040541 | 0.13823 | -8.8e-05 | -0.006538 | -0.006538 | -0.85198
2019-04-09 14:57 | 1554836205.5574 | 1324 | 0.20955 | 0.000444 | 2e-06 | 100.273 | 100.8422 | 121.905 | 20.5999 | 0.002163 | 0.001972 | -154.6995 | 0.23247 | 8.4019 | 0.002747 | 0.00404 | 0.003831 | 4.6e-05 | 0.001027 | 0.28456 | 1.3598 | 2.5476 | 98.51 | 20.09 | 49.441 | 73.218 | 87.2696 | 68.6326 | 0.011812 | 1.033 | 0.024068 | 23.3002 | -0.050629 | 0.1383 | -0.000106 | -0.007801 | -0.007801 | -0.71443
2019-04-09 14:57 | 1554836206.5573 | 1325 | 0.20957 | 0.000441 | 3e-06 | 100.203 | 100.7097 | 121.8435 | 20.5996 | 0.002031 | 0.001983 | -154.7051 | 0.23221 | 8.402 | 0.002709 | 0.004589 | 0.004416 | 4.4e-05 | 0.001023 | 0.26394 | 1.3442 | 2.5476 | 98.51 | 20.09 | 49.431 | 70.0423 | 86.3521 | 68.8936 | 0.01181 | 1.0322 | 0.024053 | 23.3032 | -0.045113 | 0.13814 | -0.000188 | -0.013722 | -0.013722 | -0.40571
2019-04-09 14:57 | 1554836207.5575 | 1326 | 0.20955 | 0.000445 | 2e-06 | 100.2453 | 100.7616 | 122.1447 | 20.5998 | 0.002046 | 0.001967 | -154.6953 | 0.23296 | 8.4031 | 0.002791 | 0.004088 | 0.004073 | 4.2e-05 | 0.001026 | 0.27367 | 1.3524 | 2.5477 | 98.51 | 20.09 | 49.425 | 70.4145 | 88.3113 | 68.5094 | 0.011809 | 1.0333 | 0.024096 | 23.3189 | -0.045954 | 0.13859 | -5.8e-05 | -0.004287 | -0.004287 | -1.3028
2019-04-09 14:57 | 1554836208.5575 | 1327 | 0.20955 | 0.000442 | 2e-06 | 100.2149 | 100.6071 | 122.1721 | 20.6004 | 0.002063 | 0.001981 | -154.6853 | 0.23276 | 8.4005 | 0.002764 | 0.004553 | 0.004327 | 4.5e-05 | 0.001026 | 0.27328 | 1.3417 | 2.5477 | 98.51 | 20.09 | 49.42 | 70.8166 | 87.6754 | 68.8444 | 0.011807 | 1.0323 | 0.024087 | 23.333 | -0.034937 | 0.13847 | -7e-05 | -0.005158 | -0.005158 | -1.0819
2019-04-09 14:57 | 1554836209.5571 | 1328 | 0.20958 | 0.000447 | 2e-06 | 100.2179 | 100.7375 | 121.7902 | 20.5996 | 0.002065 | 0.001945 | -154.6771 | 0.23193 | 8.4008 | 0.00276 | 0.003639 | 0.003518 | 4.1e-05 | 0.001029 | 0.29264 | 1.4135 | 2.5477 | 98.51 | 20.09 | 49.41 | 70.8781 | 87.5745 | 67.9834 | 0.011805 | 1.0349 | 0.02408 | 23.2671 | -0.04626 | 0.13797 | -0.000235 | -0.017044 | -0.017044 | -0.32623
2019-04-09 14:57 | 1554836210.5572 | 1329 | 0.20953 | 0.000442 | 2e-06 | 100.214 | 100.5613 | 121.6527 | 20.5996 | 0.002072 | 0.001964 | -154.6627 | 0.2315 | 8.4002 | 0.002737 | 0.004192 | 0.004034 | 4.4e-05 | 0.001033 | 0.30901 | 1.3369 | 2.5475 | 98.51 | 20.08 | 49.405 | 71.0466 | 87.0102 | 68.4328 | 0.011796 | 1.0336 | 0.02405 | 23.2693 | -0.030946 | 0.13772 | 6e-05 | 0.004236 | 0.004236 | 1.3102
2019-04-09 14:57 | 1554836211.5582 | 1330 | 0.20956 | 0.000445 | 2e-06 | 100.2631 | 100.8456 | 121.8337 | 20.6001 | 0.002071 | 0.001945 | -154.6787 | 0.23164 | 8.4046 | 0.002805 | 0.004241 | 0.004061 | 4.5e-05 | 0.001035 | 0.28694 | 1.3804 | 2.5475 | 98.51 | 20.08 | 49.395 | 71.0158 | 88.6505 | 67.9916 | 0.011794 | 1.0349 | 0.024084 | 23.2716 | -0.051811 | 0.1378 | -0.000143 | -0.0104 | -0.0104 | -0.53396
2019-04-09 14:57 | 1554836212.5587 | 1331 | 0.2095 | 0.000445 | 2e-06 | 100.2435 | 100.5992 | 121.7339 | 20.6001 | 0.002009 | 0.001937 | -154.6743 | 0.23174 | 8.4016 | 0.002801 | 0.00373 | 0.00356 | 4.4e-05 | 0.001037 | 0.29856 | 1.3351 | 2.5476 | 98.51 | 20.09 | 49.39 | 69.5313 | 88.5513 | 67.7966 | 0.0118 | 1.0355 | 0.024081 | 23.2554 | -0.031682 | 0.13786 | 0.000253 | 0.018116 | 0.018116 | 0.30669
2019-04-09 14:57 | 1554836213.5583 | 1332 | 0.20955 | 0.000445 | 1e-06 | 100.1154 | 100.508 | 121.8724 | 20.6002 | 0.002077 | 0.001889 | -154.6916 | 0.23298 | 8.4014 | 0.002827 | 0.004213 | 0.004007 | 4.7e-05 | 0.001039 | 0.30469 | 1.3289 | 2.5475 | 98.51 | 20.09 | 49.386 | 71.168 | 89.1834 | 66.6351 | 0.011799 | 1.039 | 0.024135 | 23.2289 | -0.035002 | 0.1386 | -9.6e-05 | -0.007046 | -0.007046 | -0.79271
2019-04-09 14:57 | 1554836214.5584 | 1333 | 0.20959 | 0.00044 | 2e-06 | 100.1333 | 100.646 | 121.7698 | 20.6003 | 0.001995 | 0.002005 | -154.7101 | 0.23325 | 8.4011 | 0.002808 | 0.004367 | 0.004162 | 4.4e-05 | 0.001038 | 0.29214 | 1.3479 | 2.5475 | 98.51 | 20.09 | 49.377 | 69.1869 | 88.7297 | 69.4339 | 0.011797 | 1.0305 | 0.024027 | 23.3146 | -0.045679 | 0.13876 | -0.000293 | -0.021251 | -0.021251 | -0.26313
2019-04-09 14:57 | 1554836215.5588 | 1334 | 0.20952 | 0.000444 | 2e-06 | 100.0961 | 100.5971 | 121.8872 | 20.5993 | 0.002077 | 0.001921 | -154.7189 | 0.23267 | 8.4 | 0.002775 | 0.00428 | 0.004112 | 4.7e-05 | 0.001041 | 0.30541 | 1.3619 | 2.5475 | 98.51 | 20.09 | 49.373 | 71.1539 | 87.9281 | 67.4177 | 0.011796 | 1.0366 | 0.024109 | 23.2571 | -0.044661 | 0.13842 | 0.00012 | 0.008544 | 0.008544 | 0.65286
2019-04-09 14:57 | 1554836216.5585 | 1335 | 0.20957 | 0.000445 | 2e-06 | 100.1568 | 100.7265 | 121.7656 | 20.6005 | 0.002023 | 0.001972 | -154.712 | 0.23305 | 8.4 | 0.002777 | 0.004371 | 0.004253 | 4.6e-05 | 0.001039 | 0.27433 | 1.3506 | 2.5476 | 98.51 | 20.08 | 49.369 | 69.8538 | 87.9723 | 68.6231 | 0.011788 | 1.033 | 0.024055 | 23.2866 | -0.050732 | 0.13864 | -0.000206 | -0.014923 | -0.014923 | -0.3744
2019-04-09 14:57 | 1554836217.5574 | 1336 | 0.20955 | 0.000445 | 2e-06 | 100.1056 | 100.5238 | 121.4684 | 20.6003 | 0.002116 | 0.001925 | -154.7364 | 0.23241 | 8.4013 | 0.002775 | 0.003986 | 0.003835 | 4.5e-05 | 0.001043 | 0.27505 | 1.3875 | 2.5478 | 98.51 | 20.08 | 49.362 | 72.0897 | 87.9401 | 67.4973 | 0.011786 | 1.0364 | 0.024065 | 23.2198 | -0.037289 | 0.13826 | -7.4e-05 | -0.005455 | -0.005455 | -1.0215
2019-04-09 14:57 | 1554836218.5585 | 1337 | 0.20952 | 0.000443 | 2e-06 | 100.1217 | 100.636 | 121.4797 | 20.6011 | 0.00203 | 0.001983 | -154.7393 | 0.23133 | 8.4028 | 0.002839 | 0.003898 | 0.003698 | 4.5e-05 | 0.001041 | 0.30198 | 1.3802 | 2.5476 | 98.51 | 20.08 | 49.359 | 70.0375 | 89.4834 | 68.9076 | 0.011785 | 1.0321 | 0.024016 | 23.2689 | -0.04583 | 0.13762 | 0.000131 | 0.009315 | 0.009315 | 0.5954
2019-04-09 14:57 | 1554836219.5585 | 1338 | 0.20953 | 0.000444 | 3e-06 | 100.1178 | 100.6752 | 121.6297 | 20.601 | 0.00201 | 0.001896 | -154.7414 | 0.23201 | 8.4021 | 0.002833 | 0.004181 | 0.003973 | 4.3e-05 | 0.001043 | 0.30989 | 1.3913 | 2.5476 | 98.51 | 20.09 | 49.355 | 69.554 | 89.3231 | 66.817 | 0.011792 | 1.0385 | 0.024105 | 23.212 | -0.049662 | 0.13802 | 1.9e-05 | 0.001186 | 0.001186 | 4.6891
2019-04-09 14:57 | 1554836220.5567 | 1339 | 0.20955 | 0.000441 | 2e-06 | 100.1238 | 100.6061 | 121.8344 | 20.5993 | 0.002077 | 0.001901 | -154.75 | 0.23316 | 8.4021 | 0.002848 | 0.003902 | 0.003742 | 4.4e-05 | 0.001042 | 0.27793 | 1.3809 | 2.5477 | 98.51 | 20.08 | 49.354 | 71.1532 | 89.6842 | 66.9315 | 0.011784 | 1.0381 | 0.024121 | 23.2354 | -0.042982 | 0.1387 | -4.9e-05 | -0.00365 | -0.00365 | -1.5315
2019-04-09 14:57 | 1554836221.5576 | 1340 | 0.20956 | 0.000444 | 2e-06 | 100.1399 | 100.5183 | 121.7679 | 20.6009 | 0.002065 | 0.001935 | -154.7233 | 0.23302 | 8.4006 | 0.002838 | 0.003773 | 0.003669 | 4.5e-05 | 0.001047 | 0.29048 | 1.3754 | 2.5476 | 98.51 | 20.09 | 49.352 | 70.8606 | 89.463 | 67.7498 | 0.011791 | 1.0356 | 0.024086 | 23.257 | -0.03374 | 0.13862 | -0.000149 | -0.010895 | -0.010895 | -0.51278
2019-04-09 14:57 | 1554836222.5576 | 1341 | 0.20956 | 0.000442 | 2e-06 | 100.0735 | 100.6308 | 121.5562 | 20.6017 | 0.002094 | 0.001909 | -154.7337 | 0.23252 | 8.4024 | 0.00279 | 0.004059 | 0.003843 | 4.4e-05 | 0.001048 | 0.30414 | 1.3686 | 2.5475 | 98.51 | 20.09 | 49.35 | 71.5734 | 88.2888 | 67.1125 | 0.011791 | 1.0376 | 0.024087 | 23.2151 | -0.049682 | 0.13833 | -0.000123 | -0.009038 | -0.009038 | -0.61683
2019-04-09 14:57 | 1554836223.5576 | 1342 | 0.20955 | 0.000442 | 2e-06 | 100.1491 | 100.6322 | 121.3544 | 20.6009 | 0.002101 | 0.001919 | -154.7168 | 0.23265 | 8.3998 | 0.002786 | 0.004692 | 0.004477 | 4.8e-05 | 0.00105 | 0.28456 | 1.3849 | 2.5477 | 98.51 | 20.09 | 49.351 | 71.7286 | 88.2093 | 67.3615 | 0.011791 | 1.0368 | 0.024058 | 23.2043 | -0.043046 | 0.13841 | -0.000101 | -0.007421 | -0.007421 | -0.75163
2019-04-09 14:57 | 1554836224.558 | 1343 | 0.20953 | 0.000444 | 2e-06 | 100.1336 | 100.7596 | 121.3537 | 20.6015 | 0.001981 | 0.001968 | -154.7277 | 0.23256 | 8.402 | 0.002739 | 0.00409 | 0.003977 | 4.3e-05 | 0.001051 | 0.27776 | 1.3612 | 2.5476 | 98.51 | 20.08 | 49.353 | 68.8417 | 87.0605 | 68.5477 | 0.011784 | 1.0332 | 0.024017 | 23.2446 | -0.05574 | 0.13835 | 6.3e-05 | 0.004365 | 0.004365 | 1.2774
2019-04-09 14:57 | 1554836225.5571 | 1344 | 0.20954 | 0.000443 | 2e-06 | 100.1729 | 100.4123 | 121.198 | 20.6012 | 0.002056 | 0.001911 | -154.7225 | 0.23247 | 8.4011 | 0.002778 | 0.004138 | 0.003961 | 4.7e-05 | 0.00105 | 0.27936 | 1.3839 | 2.5476 | 98.51 | 20.08 | 49.354 | 70.6497 | 88.0026 | 67.1728 | 0.011784 | 1.0374 | 0.02405 | 23.1829 | -0.02135 | 0.1383 | 1.2e-05 | 0.000717 | 0.000717 | 7.7736
2019-04-09 14:57 | 1554836226.5567 | 1345 | 0.20955 | 0.000441 | 3e-06 | 100.13 | 100.4722 | 121.6724 | 20.6015 | 0.002068 | 0.001937 | -154.7343 | 0.23224 | 8.4016 | 0.002776 | 0.004435 | 0.004225 | 4.5e-05 | 0.001056 | 0.28888 | 1.3523 | 2.5475 | 98.51 | 20.08 | 49.355 | 70.9487 | 87.9653 | 67.7841 | 0.011785 | 1.0355 | 0.024075 | 23.2491 | -0.030513 | 0.13816 | -7.1e-05 | -0.005273 | -0.005273 | -1.056
2019-04-09 14:57 | 1554836227.5577 | 1346 | 0.20957 | 0.000446 | 2e-06 | 100.1881 | 100.4842 | 121.6799 | 20.6014 | 0.002081 | 0.001944 | -154.7061 | 0.23255 | 8.3987 | 0.002779 | 0.004339 | 0.004133 | 4.8e-05 | 0.001054 | 0.27505 | 1.3505 | 2.5476 | 98.51 | 20.08 | 49.36 | 71.2511 | 88.0223 | 67.9696 | 0.011786 | 1.035 | 0.024069 | 23.2561 | -0.026399 | 0.13834 | -0.000223 | -0.016229 | -0.016229 | -0.34354
2019-04-09 14:57 | 1554836228.5571 | 1347 | 0.20955 | 0.000445 | 2e-06 | 100.1616 | 100.6278 | 121.5964 | 20.6008 | 0.002019 | 0.001938 | -154.7304 | 0.23217 | 8.4004 | 0.002713 | 0.00412 | 0.003997 | 4.7e-05 | 0.001057 | 0.27759 | 1.3348 | 2.5476 | 98.51 | 20.08 | 49.364 | 69.7713 | 86.4336 | 67.8074 | 0.011787 | 1.0355 | 0.024067 | 23.2426 | -0.041531 | 0.13812 | -6.5e-05 | -0.004822 | -0.004822 | -1.1542
2019-04-09 14:57 | 1554836229.5578 | 1348 | 0.20958 | 0.000445 | 3e-06 | 100.1842 | 100.5706 | 121.4201 | 20.6012 | 0.002023 | 0.002023 | -154.711 | 0.2321 | 8.4007 | 0.002738 | 0.003639 | 0.003453 | 4.2e-05 | 0.001061 | 0.28136 | 1.348 | 2.5477 | 98.51 | 20.08 | 49.372 | 69.8568 | 87.0365 | 69.8513 | 0.011789 | 1.0293 | 0.023977 | 23.2953 | -0.034431 | 0.13808 | -0.000245 | -0.017781 | -0.017781 | -0.31295
2019-04-09 14:57 | 1554836230.5569 | 1349 | 0.20964 | 0.000446 | 2e-06 | 100.1658 | 100.7193 | 121.3321 | 20.6013 | 0.002063 | 0.001949 | -154.7126 | 0.23174 | 8.402 | 0.002743 | 0.004059 | 0.003931 | 4.4e-05 | 0.001065 | 0.26842 | 1.3483 | 2.5476 | 98.51 | 20.09 | 49.375 | 70.8113 | 87.1702 | 68.0849 | 0.011797 | 1.0346 | 0.024031 | 23.2268 | -0.049292 | 0.13786 | -0.000593 | -0.042796 | -0.042796 | -0.12982
2019-04-09 14:57 | 1554836231.5569 | 1350 | 0.2095 | 0.000446 | 2e-06 | 100.1274 | 100.6047 | 121.1452 | 20.6018 | 0.001991 | 0.001904 | -154.7224 | 0.23245 | 8.4011 | 0.002677 | 0.004275 | 0.004105 | 4.5e-05 | 0.001061 | 0.3126 | 1.313 | 2.5475 | 98.51 | 20.09 | 49.377 | 69.0914 | 85.5666 | 67.0104 | 0.011797 | 1.0379 | 0.02405 | 23.1723 | -0.04254 | 0.13828 | 0.000251 | 0.017923 | 0.017923 | 0.31093
2019-04-09 14:57 | 1554836232.5573 | 1351 | 0.2095 | 0.000443 | 2e-06 | 100.0886 | 100.4939 | 121.3389 | 20.6012 | 0.002058 | 0.001898 | -154.7104 | 0.23213 | 8.4013 | 0.002695 | 0.003633 | 0.003467 | 4.5e-05 | 0.001065 | 0.28064 | 1.3789 | 2.5477 | 98.51 | 20.09 | 49.384 | 70.7129 | 86.0125 | 66.8499 | 0.011799 | 1.0384 | 0.024075 | 23.1854 | -0.036141 | 0.13809 | 0.000212 | 0.015222 | 0.015222 | 0.36561
2019-04-09 14:57 | 1554836233.5584 | 1352 | 0.20954 | 0.000447 | 2e-06 | 100.0863 | 100.4822 | 121.4752 | 20.6016 | 0.002065 | 0.00192 | -154.7113 | 0.23255 | 8.4026 | 0.002703 | 0.004406 | 0.004261 | 4.3e-05 | 0.001066 | 0.28208 | 1.3584 | 2.5476 | 98.51 | 20.09 | 49.386 | 70.8793 | 86.2087 | 67.3956 | 0.011799 | 1.0367 | 0.024069 | 23.217 | -0.035312 | 0.13834 | -1.4e-05 | -0.001133 | -0.001133 | -4.9196
2019-04-09 14:57 | 1554836234.5574 | 1353 | 0.20955 | 0.000447 | 3e-06 | 100.1137 | 100.6525 | 121.5426 | 20.6016 | 0.002029 | 0.001957 | -154.7131 | 0.23236 | 8.4044 | 0.002773 | 0.003517 | 0.00335 | 4.3e-05 | 0.001063 | 0.30574 | 1.3868 | 2.5476 | 98.51 | 20.09 | 49.389 | 69.9976 | 87.8951 | 68.2857 | 0.0118 | 1.034 | 0.024044 | 23.2538 | -0.048016 | 0.13823 | -8.5e-05 | -0.006312 | -0.006312 | -0.88261
2019-04-09 14:57 | 1554836235.558 | 1354 | 0.20955 | 0.000446 | 2e-06 | 100.1101 | 100.508 | 121.9219 | 20.6034 | 0.002076 | 0.001945 | -154.715 | 0.23215 | 8.401 | 0.002755 | 0.00385 | 0.003724 | 4.2e-05 | 0.001061 | 0.2891 | 1.3244 | 2.5477 | 98.51 | 20.09 | 49.392 | 71.1252 | 87.4494 | 67.9965 | 0.011801 | 1.0349 | 0.024092 | 23.2802 | -0.035481 | 0.1381 | -4.5e-05 | -0.003415 | -0.003415 | -1.6295
2019-04-09 14:57 | 1554836236.556 | 1355 | 0.20954 | 0.000446 | 2e-06 | 100.1506 | 100.6611 | 121.6832 | 20.6022 | 0.001966 | 0.001916 | -154.7196 | 0.23297 | 8.4039 | 0.002751 | 0.003542 | 0.003324 | 4.6e-05 | 0.001062 | 0.28799 | 1.3701 | 2.5477 | 98.51 | 20.09 | 49.393 | 68.486 | 87.3578 | 67.294 | 0.011801 | 1.037 | 0.024093 | 23.2334 | -0.045482 | 0.1386 | 7e-06 | 0.00036 | 0.00036 | 15.5122
2019-04-09 14:57 | 1554836237.5573 | 1356 | 0.20954 | 0.000441 | 3e-06 | 100.1006 | 100.5052 | 121.567 | 20.602 | 0.002052 | 0.001899 | -154.7293 | 0.23272 | 8.4 | 0.002706 | 0.003856 | 0.003694 | 4.2e-05 | 0.001061 | 0.28982 | 1.3609 | 2.5477 | 98.51 | 20.09 | 49.394 | 70.55 | 86.2779 | 66.8874 | 0.011801 | 1.0383 | 0.024096 | 23.2084 | -0.036088 | 0.13845 | -2.9e-05 | -0.002227 | -0.002227 | -2.5059
2019-04-09 14:57 | 1554836238.5561 | 1357 | 0.20951 | 0.000445 | 2e-06 | 100.1041 | 100.6556 | 121.5321 | 20.6021 | 0.002055 | 0.00194 | -154.7315 | 0.23225 | 8.401 | 0.002693 | 0.004113 | 0.003987 | 4.4e-05 | 0.00106 | 0.31078 | 1.3868 | 2.5477 | 98.51 | 20.09 | 49.393 | 70.6225 | 85.9752 | 67.8626 | 0.011801 | 1.0353 | 0.024058 | 23.2383 | -0.049143 | 0.13816 | 0.000146 | 0.010385 | 0.010385 | 0.53617
2019-04-09 14:57 | 1554836239.5565 | 1358 | 0.20952 | 0.000442 | 3e-06 | 100.086 | 100.3855 | 121.3985 | 20.6025 | 0.002108 | 0.001909 | -154.7499 | 0.23257 | 8.4008 | 0.002695 | 0.003799 | 0.003596 | 4.4e-05 | 0.00105 | 0.30021 | 1.383 | 2.5476 | 98.51 | 20.09 | 49.392 | 71.9132 | 86.0215 | 67.1142 | 0.011801 | 1.0376 | 0.024072 | 23.2001 | -0.026726 | 0.13836 | 0.000117 | 0.008256 | 0.008256 | 0.67533
2019-04-09 14:57 | 1554836240.5561 | 1359 | 0.20953 | 0.000441 | 2e-06 | 100.1011 | 100.6405 | 121.6778 | 20.6022 | 0.002053 | 0.0019 | -154.7484 | 0.23272 | 8.4019 | 0.002739 | 0.003833 | 0.003625 | 4.7e-05 | 0.001046 | 0.28191 | 1.3887 | 2.5477 | 98.51 | 20.09 | 49.391 | 70.5774 | 87.0618 | 66.9033 | 0.0118 | 1.0382 | 0.024107 | 23.2195 | -0.048063 | 0.13845 | 3.5e-05 | 0.00236 | 0.00236 | 2.3641
2019-04-09 14:57 | 1554836241.5572 | 1360 | 0.20952 | 0.000447 | 2e-06 | 100.1324 | 100.774 | 122.2059 | 20.6027 | 0.002001 | 0.001942 | -154.7709 | 0.2319 | 8.4013 | 0.002746 | 0.003575 | 0.003387 | 4.6e-05 | 0.00105 | 0.27096 | 1.3629 | 2.5477 | 98.51 | 20.09 | 49.388 | 69.3264 | 87.2356 | 67.9132 | 0.0118 | 1.0351 | 0.024123 | 23.3044 | -0.057129 | 0.13796 | 8.2e-05 | 0.00578 | 0.00578 | 0.96192
2019-04-09 14:57 | 1554836242.5576 | 1361 | 0.20952 | 0.000442 | 2e-06 | 100.08 | 100.7103 | 122.5637 | 20.6027 | 0.002027 | 0.001889 | -154.7697 | 0.23258 | 8.4016 | 0.002766 | 0.003782 | 0.003584 | 4.4e-05 | 0.001052 | 0.31133 | 1.3839 | 2.5477 | 98.51 | 20.09 | 49.383 | 69.9549 | 87.7181 | 66.6413 | 0.011799 | 1.039 | 0.024204 | 23.2949 | -0.05616 | 0.13836 | 9.4e-05 | 0.006705 | 0.006705 | 0.8316
2019-04-09 14:57 | 1554836243.5567 | 1362 | 0.20953 | 0.000447 | 2e-06 | 100.0738 | 100.6236 | 122.5385 | 20.6034 | 0.001974 | 0.001959 | -154.7977 | 0.23249 | 8.4037 | 0.002809 | 0.00374 | 0.003554 | 4.5e-05 | 0.001053 | 0.28495 | 1.3717 | 2.5476 | 98.51 | 20.09 | 49.381 | 68.6947 | 88.7631 | 68.3163 | 0.011798 | 1.0339 | 0.024142 | 23.3499 | -0.049012 | 0.13831 | 6.2e-05 | 0.004346 | 0.004346 | 1.2826
2019-04-09 14:57 | 1554836244.5567 | 1363 | 0.20956 | 0.000442 | 2e-06 | 100.1083 | 100.486 | 122.4339 | 20.6028 | 0.002043 | 0.001924 | -154.7878 | 0.23226 | 8.4018 | 0.002732 | 0.004355 | 0.004179 | 4.3e-05 | 0.001056 | 0.27649 | 1.4155 | 2.5476 | 98.51 | 20.09 | 49.378 | 70.3393 | 86.8922 | 67.4833 | 0.011797 | 1.0364 | 0.024161 | 23.3114 | -0.03368 | 0.13817 | -0.000127 | -0.009331 | -0.009331 | -0.59675
2019-04-09 14:57 | 1554836245.5573 | 1364 | 0.20953 | 0.000445 | 3e-06 | 100.0985 | 100.3941 | 122.0686 | 20.6043 | 0.001979 | 0.001921 | -154.7862 | 0.23193 | 8.4013 | 0.002725 | 0.004748 | 0.004567 | 4.1e-05 | 0.001055 | 0.29999 | 1.4171 | 2.5476 | 98.51 | 20.09 | 49.374 | 68.8071 | 86.7258 | 67.4038 | 0.011796 | 1.0367 | 0.024128 | 23.2739 | -0.026375 | 0.13798 | 7.1e-05 | 0.004973 | 0.004973 | 1.1181
2019-04-09 14:57 | 1554836246.5577 | 1365 | 0.20952 | 0.000452 | 2e-06 | 100.0219 | 100.5039 | 121.9461 | 20.6038 | 0.002069 | 0.001913 | -154.762 | 0.2322 | 8.4028 | 0.002715 | 0.004287 | 0.004133 | 4.5e-05 | 0.00106 | 0.28102 | 1.3613 | 2.5476 | 98.51 | 20.09 | 49.372 | 70.9572 | 86.4874 | 67.2278 | 0.011796 | 1.0372 | 0.024122 | 23.2562 | -0.042993 | 0.13814 | 8e-05 | 0.005643 | 0.005643 | 0.9865
2019-04-09 14:57 | 1554836247.5578 | 1366 | 0.20954 | 0.000442 | 3e-06 | 100.1485 | 100.7695 | 122.0837 | 20.603 | 0.002009 | 0.001914 | -154.7597 | 0.23214 | 8.4016 | 0.002701 | 0.004222 | 0.004062 | 4.4e-05 | 0.001065 | 0.30254 | 1.3806 | 2.5477 | 98.51 | 20.09 | 49.37 | 69.5303 | 86.161 | 67.2371 | 0.011795 | 1.0372 | 0.024135 | 23.2696 | -0.055296 | 0.1381 | -3e-06 | -0.000384 | -0.000384 | -14.4866
2019-04-09 14:57 | 1554836248.5568 | 1367 | 0.20953 | 0.000446 | 2e-06 | 100.1431 | 100.5369 | 122.2247 | 20.6043 | 0.002004 | 0.001901 | -154.7441 | 0.23302 | 8.4033 | 0.002813 | 0.003911 | 0.00379 | 4.6e-05 | 0.001063 | 0.29717 | 1.4356 | 2.5477 | 98.51 | 20.09 | 49.366 | 69.4128 | 88.8537 | 66.9248 | 0.011794 | 1.0381 | 0.02416 | 23.2724 | -0.0351 | 0.13862 | 3.4e-05 | 0.002311 | 0.002311 | 2.4179
2019-04-09 14:57 | 1554836249.5574 | 1368 | 0.2095 | 0.000442 | 2e-06 | 100.0449 | 100.4643 | 122.6414 | 20.6037 | 0.002008 | 0.00188 | -154.7439 | 0.23239 | 8.4015 | 0.002705 | 0.004336 | 0.004042 | 4.5e-05 | 0.001064 | 0.29789 | 1.4032 | 2.5476 | 98.51 | 20.09 | 49.367 | 69.4973 | 86.2639 | 66.4129 | 0.011795 | 1.0397 | 0.024219 | 23.2945 | -0.037419 | 0.13825 | 0.000232 | 0.016708 | 0.016708 | 0.33346
2019-04-09 14:57 | 1554836250.5581 | 1369 | 0.20955 | 0.000448 | 2e-06 | 100.0883 | 100.6119 | 122.5139 | 20.6045 | 0.002028 | 0.001904 | -154.7477 | 0.23261 | 8.3996 | 0.002684 | 0.003997 | 0.003832 | 4.5e-05 | 0.001066 | 0.29518 | 1.3551 | 2.5476 | 98.51 | 20.09 | 49.368 | 69.9723 | 85.7494 | 67.006 | 0.011795 | 1.0379 | 0.024186 | 23.3027 | -0.046668 | 0.13838 | -0.000101 | -0.007447 | -0.007447 | -0.74892
2019-04-09 14:58 | 1554836251.5559 | 1370 | 0.20955 | 0.000447 | 3e-06 | 100.0726 | 100.5682 | 122.1937 | 20.6045 | 0.002004 | 0.001904 | -154.7471 | 0.23227 | 8.3988 | 0.002684 | 0.003763 | 0.003566 | 4.3e-05 | 0.001068 | 0.29822 | 1.3628 | 2.5477 | 98.51 | 20.09 | 49.368 | 69.4108 | 85.7424 | 66.999 | 0.011795 | 1.0379 | 0.024154 | 23.272 | -0.044192 | 0.13818 | -5.4e-05 | -0.004069 | -0.004069 | -1.3685
2019-04-09 14:58 | 1554836252.5569 | 1371 | 0.20956 | 0.000447 | 2e-06 | 100.1226 | 100.5056 | 121.7337 | 20.6031 | 0.002002 | 0.001907 | -154.7357 | 0.23157 | 8.4021 | 0.00275 | 0.004441 | 0.004184 | 4.3e-05 | 0.001064 | 0.2995 | 1.3461 | 2.5476 | 98.51 | 20.09 | 49.37 | 69.3595 | 87.3305 | 67.0624 | 0.011795 | 1.0377 | 0.024107 | 23.2303 | -0.034149 | 0.13776 | -0.000117 | -0.008535 | -0.008535 | -0.65047
2019-04-09 14:58 | 1554836253.5569 | 1372 | 0.20952 | 0.000441 | 2e-06 | 100.0499 | 100.5737 | 121.7189 | 20.6039 | 0.001987 | 0.001894 | -154.7206 | 0.2317 | 8.4032 | 0.002711 | 0.004099 | 0.003949 | 4.2e-05 | 0.001066 | 0.28335 | 1.3096 | 2.5476 | 98.51 | 20.09 | 49.372 | 68.9891 | 86.3928 | 66.7683 | 0.011796 | 1.0386 | 0.024116 | 23.2189 | -0.046705 | 0.13784 | 0.000125 | 0.008929 | 0.008929 | 0.62215
2019-04-09 14:58 | 1554836254.5565 | 1373 | 0.20955 | 0.000444 | 3e-06 | 100.105 | 100.5658 | 121.7738 | 20.6038 | 0.002065 | 0.001951 | -154.7175 | 0.23222 | 8.4013 | 0.002751 | 0.003536 | 0.003367 | 4.5e-05 | 0.001062 | 0.27273 | 1.304 | 2.5476 | 98.51 | 20.09 | 49.375 | 70.8832 | 87.3664 | 68.1304 | 0.011797 | 1.0345 | 0.024073 | 23.2706 | -0.041077 | 0.13815 | -8.5e-05 | -0.006272 | -0.006272 | -0.8877
2019-04-09 14:58 | 1554836255.5561 | 1374 | 0.20954 | 0.000446 | 3e-06 | 100.066 | 100.6085 | 121.9416 | 20.604 | 0.002042 | 0.001882 | -154.7278 | 0.23206 | 8.3997 | 0.002758 | 0.004152 | 0.004002 | 4.5e-05 | 0.001061 | 0.28888 | 1.2852 | 2.5475 | 98.51 | 20.09 | 49.379 | 70.3234 | 87.5203 | 66.4774 | 0.011798 | 1.0395 | 0.024148 | 23.2301 | -0.048359 | 0.13805 | -2.7e-05 | -0.002119 | -0.002119 | -2.6251
2019-04-09 14:58 | 1554836256.5567 | 1375 | 0.20954 | 0.000448 | 2e-06 | 100.0502 | 100.53 | 122.1559 | 20.6038 | 0.001945 | 0.001921 | -154.7629 | 0.23266 | 8.4023 | 0.002758 | 0.004655 | 0.004413 | 4.5e-05 | 0.001061 | 0.28711 | 1.3171 | 2.5475 | 98.51 | 20.09 | 49.38 | 67.9782 | 87.531 | 67.4165 | 0.011798 | 1.0366 | 0.024136 | 23.2827 | -0.042791 | 0.13841 | -6e-06 | -0.000568 | -0.000568 | -9.8205
2019-04-09 14:58 | 1554836257.5567 | 1376 | 0.20953 | 0.000448 | 2e-06 | 100.0609 | 100.5245 | 122.4491 | 20.6042 | 0.002023 | 0.001883 | -154.7705 | 0.2326 | 8.4019 | 0.002653 | 0.003385 | 0.003274 | 4.6e-05 | 0.001058 | 0.27887 | 1.3483 | 2.5476 | 98.5 | 20.09 | 49.382 | 69.8689 | 84.9921 | 66.4884 | 0.011799 | 1.0395 | 0.024198 | 23.2788 | -0.041342 | 0.13837 | 1.7e-05 | 0.001104 | 0.001104 | 5.0496
2019-04-09 14:58 | 1554836258.5567 | 1377 | 0.20952 | 0.000446 | 2e-06 | 100.058 | 100.4319 | 122.29 | 20.6034 | 0.002037 | 0.001867 | -154.7652 | 0.23221 | 8.4022 | 0.002702 | 0.004965 | 0.004735 | 4.4e-05 | 0.001059 | 0.26864 | 1.3666 | 2.5476 | 98.5 | 20.1 | 49.383 | 70.2089 | 86.1715 | 66.106 | 0.011807 | 1.0406 | 0.024196 | 23.2506 | -0.033367 | 0.13814 | 0.000101 | 0.007184 | 0.007184 | 0.77491
2019-04-09 14:58 | 1554836259.5561 | 1378 | 0.20951 | 0.000446 | 2e-06 | 100.0806 | 100.6814 | 121.9621 | 20.6048 | 0.002043 | 0.001909 | -154.7619 | 0.23187 | 8.4038 | 0.002661 | 0.004255 | 0.004104 | 4.4e-05 | 0.001062 | 0.26305 | 1.3345 | 2.5475 | 98.5 | 20.1 | 49.383 | 70.3538 | 85.1933 | 67.1193 | 0.011807 | 1.0375 | 0.024127 | 23.254 | -0.053538 | 0.13794 | 0.000146 | 0.010437 | 0.010437 | 0.53261
2019-04-09 14:58 | 1554836260.5562 | 1379 | 0.20952 | 0.000447 | 2e-06 | 100.1089 | 100.5988 | 121.9947 | 20.6051 | 0.002025 | 0.001837 | -154.764 | 0.23203 | 8.4015 | 0.00271 | 0.003727 | 0.003569 | 4.2e-05 | 0.001061 | 0.28136 | 1.3521 | 2.5475 | 98.5 | 20.1 | 49.383 | 69.9059 | 86.3871 | 65.3868 | 0.011807 | 1.0429 | 0.024192 | 23.1979 | -0.043669 | 0.13803 | 0.000134 | 0.009607 | 0.009607 | 0.57901
2019-04-09 14:58 | 1554836261.5569 | 1380 | 0.20953 | 0.000444 | 2e-06 | 100.1178 | 100.625 | 122.0733 | 20.6044 | 0.001993 | 0.001898 | -154.7399 | 0.23198 | 8.4013 | 0.002621 | 0.003264 | 0.003125 | 4.2e-05 | 0.001064 | 0.28313 | 1.3403 | 2.5476 | 98.5 | 20.1 | 49.381 | 69.1425 | 84.2237 | 66.8523 | 0.011807 | 1.0384 | 0.024148 | 23.2555 | -0.045197 | 0.138 | 7e-05 | 0.004949 | 0.004949 | 1.1239
2019-04-09 14:58 | 1554836262.5559 | 1381 | 0.20952 | 0.000446 | 2e-06 | 100.0717 | 100.4003 | 122.2855 | 20.6044 | 0.002037 | 0.001919 | -154.7688 | 0.23186 | 8.3994 | 0.002633 | 0.003502 | 0.003319 | 4.5e-05 | 0.001065 | 0.29247 | 1.3239 | 2.5476 | 98.5 | 20.1 | 49.378 | 70.2022 | 84.5305 | 67.3546 | 0.011806 | 1.0368 | 0.024151 | 23.2929 | -0.029322 | 0.13793 | 0.000142 | 0.010133 | 0.010133 | 0.54859
2019-04-09 14:58 | 1554836263.5573 | 1382 | 0.20953 | 0.000444 | 2e-06 | 100.1443 | 100.72 | 122.2395 | 20.6061 | 0.00198 | 0.001905 | -154.7518 | 0.23179 | 8.4008 | 0.002641 | 0.00441 | 0.004169 | 4.8e-05 | 0.001071 | 0.29159 | 1.3517 | 2.5476 | 98.5 | 20.1 | 49.374 | 68.8203 | 84.7169 | 67.027 | 0.011805 | 1.0378 | 0.024158 | 23.2773 | -0.051268 | 0.13789 | 3.5e-05 | 0.002396 | 0.002396 | 2.3197
2019-04-09 14:58 | 1554836264.5559 | 1383 | 0.20954 | 0.000441 | 2e-06 | 100.1014 | 100.5971 | 122.2426 | 20.6063 | 0.001995 | 0.001917 | -154.7585 | 0.23218 | 8.4019 | 0.002706 | 0.003871 | 0.00375 | 4.5e-05 | 0.001074 | 0.28351 | 1.3796 | 2.5475 | 98.5 | 20.1 | 49.372 | 69.2004 | 86.2749 | 67.3245 | 0.011804 | 1.0369 | 0.024148 | 23.2878 | -0.044182 | 0.13812 | -7e-06 | -0.000637 | -0.000637 | -8.7394
2019-04-09 14:58 | 1554836265.557 | 1384 | 0.20956 | 0.000444 | 2e-06 | 100.0994 | 100.6546 | 122.1951 | 20.6063 | 0.002009 | 0.001923 | -154.7436 | 0.23185 | 8.4023 | 0.002673 | 0.003386 | 0.003298 | 3.9e-05 | 0.001071 | 0.2891 | 1.346 | 2.5475 | 98.51 | 20.1 | 49.367 | 69.5373 | 85.4781 | 67.4579 | 0.011802 | 1.0365 | 0.024138 | 23.2878 | -0.049477 | 0.13793 | -0.000127 | -0.009333 | -0.009333 | -0.59558
2019-04-09 14:58 | 1554836266.5551 | 1385 | 0.20952 | 0.000443 | 2e-06 | 100.0931 | 100.5324 | 122.2315 | 20.6055 | 0.001935 | 0.001929 | -154.7256 | 0.23179 | 8.4007 | 0.002693 | 0.003925 | 0.003728 | 3.8e-05 | 0.001074 | 0.28926 | 1.391 | 2.5476 | 98.51 | 20.1 | 49.365 | 67.7466 | 85.9579 | 67.5975 | 0.011802 | 1.0361 | 0.024137 | 23.2961 | -0.039172 | 0.13789 | 8.5e-05 | 0.006027 | 0.006027 | 0.92202
2019-04-09 14:58 | 1554836267.5561 | 1386 | 0.20952 | 0.000447 | 2e-06 | 100.0699 | 100.6477 | 121.956 | 20.6057 | 0.002032 | 0.001884 | -154.7129 | 0.23158 | 8.4029 | 0.00263 | 0.004019 | 0.003855 | 4.1e-05 | 0.001073 | 0.29175 | 1.3535 | 2.5477 | 98.51 | 20.1 | 49.364 | 70.0849 | 84.456 | 66.5163 | 0.011801 | 1.0394 | 0.024148 | 23.2328 | -0.0515 | 0.13776 | 0.000122 | 0.008722 | 0.008722 | 0.63655
2019-04-09 14:58 | 1554836268.5563 | 1387 | 0.20953 | 0.000445 | 2e-06 | 100.1169 | 100.6119 | 122.1336 | 20.6065 | 0.002016 | 0.001946 | -154.7146 | 0.23203 | 8.4042 | 0.002647 | 0.003921 | 0.003701 | 4.2e-05 | 0.001074 | 0.2656 | 1.3506 | 2.5475 | 98.51 | 20.1 | 49.363 | 69.705 | 84.8587 | 68.0159 | 0.011801 | 1.0348 | 0.024112 | 23.301 | -0.044114 | 0.13803 | 5e-05 | 0.003515 | 0.003515 | 1.5826
2019-04-09 14:58 | 1554836269.5557 | 1388 | 0.20954 | 0.000447 | 2e-06 | 100.0696 | 100.5362 | 122.2792 | 20.6057 | 0.001947 | 0.001893 | -154.7139 | 0.23232 | 8.4024 | 0.002728 | 0.004212 | 0.004046 | 4.4e-05 | 0.001073 | 0.29607 | 1.3452 | 2.5477 | 98.51 | 20.1 | 49.363 | 68.0331 | 86.8053 | 66.728 | 0.011801 | 1.0387 | 0.024172 | 23.2709 | -0.041611 | 0.13821 | 8e-06 | 0.000502 | 0.000502 | 11.0934
2019-04-09 14:58 | 1554836270.5555 | 1389 | 0.20954 | 0.000444 | 3e-06 | 100.1595 | 100.4935 | 122.0604 | 20.6063 | 0.001945 | 0.00186 | -154.7177 | 0.23223 | 8.4029 | 0.002653 | 0.003931 | 0.003798 | 4.2e-05 | 0.001075 | 0.28926 | 1.366 | 2.5477 | 98.51 | 20.1 | 49.364 | 68 | 84.9965 | 65.9551 | 0.011801 | 1.0411 | 0.024178 | 23.2236 | -0.029779 | 0.13816 | -1.4e-05 | -0.001162 | -0.001162 | -4.79
2019-04-09 14:58 | 1554836271.5562 | 1390 | 0.20954 | 0.000449 | 2e-06 | 100.1029 | 100.636 | 121.897 | 20.607 | 0.002056 | 0.001888 | -154.7269 | 0.23187 | 8.4014 | 0.002653 | 0.004228 | 0.004105 | 3.9e-05 | 0.001079 | 0.29336 | 1.3164 | 2.5475 | 98.51 | 20.1 | 49.367 | 70.6469 | 85.0121 | 66.6239 | 0.011802 | 1.0391 | 0.024138 | 23.2309 | -0.047506 | 0.13794 | -1.2e-05 | -0.001047 | -0.001047 | -5.3086
2019-04-09 14:58 | 1554836272.5562 | 1391 | 0.20953 | 0.000445 | 2e-06 | 100.1262 | 100.7362 | 121.4884 | 20.6056 | 0.001998 | 0.001908 | -154.7306 | 0.23173 | 8.403 | 0.00261 | 0.00439 | 0.004154 | 4.2e-05 | 0.001081 | 0.30829 | 1.3796 | 2.5477 | 98.51 | 20.1 | 49.368 | 69.2622 | 83.9675 | 67.0913 | 0.011802 | 1.0376 | 0.024081 | 23.2079 | -0.054327 | 0.13786 | 5e-05 | 0.003485 | 0.003485 | 1.5943
2019-04-09 14:58 | 1554836273.5563 | 1392 | 0.20951 | 0.000444 | 3e-06 | 100.1684 | 100.7031 | 121.4121 | 20.6065 | 0.002022 | 0.001874 | -154.7444 | 0.23211 | 8.4033 | 0.002678 | 0.003804 | 0.003607 | 3.9e-05 | 0.001082 | 0.2792 | 1.3636 | 2.5476 | 98.51 | 20.1 | 49.369 | 69.8516 | 85.6046 | 66.2925 | 0.011802 | 1.0401 | 0.024102 | 23.1733 | -0.047617 | 0.13808 | 0.000178 | 0.012662 | 0.012662 | 0.43948
2019-04-09 14:58 | 1554836274.5553 | 1393 | 0.20955 | 0.000442 | 3e-06 | 100.0341 | 100.753 | 121.4236 | 20.6064 | 0.001959 | 0.001872 | -154.7397 | 0.23176 | 8.4014 | 0.002716 | 0.003979 | 0.00375 | 4.1e-05 | 0.001081 | 0.27959 | 1.3694 | 2.5476 | 98.51 | 20.1 | 49.372 | 68.3331 | 86.5146 | 66.2371 | 0.011803 | 1.0402 | 0.024105 | 23.1725 | -0.064049 | 0.13787 | -7.1e-05 | -0.005262 | -0.005262 | -1.0558
2019-04-09 14:58 | 1554836275.5564 | 1394 | 0.20955 | 0.000444 | 3e-06 | 100.0702 | 100.6632 | 121.7428 | 20.6066 | 0.002067 | 0.001854 | -154.7648 | 0.23206 | 8.4032 | 0.002671 | 0.004209 | 0.004004 | 4.2e-05 | 0.001082 | 0.29927 | 1.4261 | 2.5478 | 98.51 | 20.1 | 49.373 | 70.9294 | 85.4489 | 65.7993 | 0.011803 | 1.0416 | 0.024152 | 23.188 | -0.052849 | 0.13805 | -6.6e-05 | -0.004936 | -0.004936 | -1.1271
2019-04-09 14:58 | 1554836276.5553 | 1395 | 0.20955 | 0.000447 | 3e-06 | 100.0618 | 100.4842 | 121.775 | 20.6074 | 0.002012 | 0.001923 | -154.7726 | 0.23165 | 8.403 | 0.002724 | 0.004468 | 0.004282 | 4.2e-05 | 0.00109 | 0.29623 | 1.3561 | 2.5477 | 98.51 | 20.11 | 49.374 | 69.6069 | 86.7211 | 67.4652 | 0.011811 | 1.0365 | 0.024096 | 23.248 | -0.037679 | 0.13781 | -5.6e-05 | -0.00418 | -0.00418 | -1.3287
2019-04-09 14:58 | 1554836277.5553 | 1396 | 0.20955 | 0.000449 | 3e-06 | 100.0475 | 100.6591 | 122.0426 | 20.6065 | 0.002039 | 0.001888 | -154.7869 | 0.23204 | 8.4017 | 0.002603 | 0.003526 | 0.003404 | 3.8e-05 | 0.001089 | 0.2896 | 1.37 | 2.5475 | 98.51 | 20.11 | 49.375 | 70.2439 | 83.7999 | 66.6249 | 0.011811 | 1.0391 | 0.024153 | 23.2448 | -0.054505 | 0.13804 | -4.9e-05 | -0.003727 | -0.003727 | -1.4925
2019-04-09 14:58 | 1554836278.5556 | 1397 | 0.20955 | 0.000445 | 2e-06 | 99.9779 | 100.5765 | 121.8987 | 20.6075 | 0.001944 | 0.00191 | -154.7895 | 0.23175 | 8.4022 | 0.002636 | 0.004068 | 0.003897 | 4.2e-05 | 0.001091 | 0.29048 | 1.3563 | 2.5475 | 98.51 | 20.11 | 49.375 | 67.9599 | 84.6052 | 67.1431 | 0.011811 | 1.0375 | 0.02412 | 23.2488 | -0.053395 | 0.13787 | -9.9e-05 | -0.007279 | -0.007279 | -0.76334
2019-04-09 14:58 | 1554836279.5554 | 1398 | 0.20956 | 0.000446 | 2e-06 | 99.9451 | 100.4681 | 121.8567 | 20.607 | 0.002023 | 0.001875 | -154.8101 | 0.23217 | 8.4009 | 0.002665 | 0.003919 | 0.003701 | 4.4e-05 | 0.001091 | 0.28064 | 1.3833 | 2.5477 | 98.51 | 20.11 | 49.375 | 69.8709 | 85.3032 | 66.2985 | 0.011811 | 1.0401 | 0.024146 | 23.2159 | -0.04668 | 0.13812 | -0.000149 | -0.010923 | -0.010923 | -0.50959
2019-04-09 14:58 | 1554836280.5564 | 1399 | 0.20955 | 0.000448 | 2e-06 | 99.943 | 100.3053 | 121.9506 | 20.6071 | 0.001941 | 0.00191 | -154.7918 | 0.2316 | 8.4018 | 0.002653 | 0.003634 | 0.003584 | 3.8e-05 | 0.001088 | 0.26288 | 1.3462 | 2.5475 | 98.51 | 20.11 | 49.374 | 67.8986 | 85.0091 | 67.1427 | 0.011811 | 1.0375 | 0.024125 | 23.2537 | -0.032362 | 0.13778 | -5.4e-05 | -0.004017 | -0.004017 | -1.3822
2019-04-09 14:58 | 1554836281.5555 | 1400 | 0.20954 | 0.000442 | 2e-06 | 99.9409 | 100.2471 | 122.1097 | 20.6076 | 0.001936 | 0.001916 | -154.8031 | 0.23207 | 8.4017 | 0.002691 | 0.00424 | 0.003987 | 4e-05 | 0.001091 | 0.28296 | 1.306 | 2.5476 | 98.51 | 20.11 | 49.375 | 67.7806 | 85.9321 | 67.2937 | 0.011811 | 1.037 | 0.024136 | 23.2741 | -0.027361 | 0.13806 | 0 | -0.000139 | -0.000139 | -40.0122
2019-04-09 14:58 | 1554836282.5546 | 1401 | 0.20957 | 0.000446 | 2e-06 | 99.8781 | 100.4006 | 122.513 | 20.6082 | 0.002002 | 0.001906 | -154.7889 | 0.23242 | 8.4014 | 0.002685 | 0.004133 | 0.003998 | 3.8e-05 | 0.001089 | 0.262 | 1.3472 | 2.5476 | 98.51 | 20.11 | 49.375 | 69.3668 | 85.7787 | 67.0635 | 0.011811 | 1.0377 | 0.024184 | 23.3046 | -0.046671 | 0.13827 | -0.000199 | -0.014583 | -0.014583 | -0.3821
2019-04-09 14:58 | 1554836283.5556 | 1402 | 0.20955 | 0.000449 | 2e-06 | 99.8305 | 100.2691 | 122.6867 | 20.6077 | 0.001959 | 0.00188 | -154.8008 | 0.23219 | 8.3989 | 0.002661 | 0.004715 | 0.004574 | 4.3e-05 | 0.001093 | 0.2855 | 1.3544 | 2.5476 | 98.51 | 20.11 | 49.375 | 68.3337 | 85.2069 | 66.4192 | 0.011811 | 1.0397 | 0.024224 | 23.299 | -0.039219 | 0.13813 | -8.2e-05 | -0.006058 | -0.006058 | -0.91882
2019-04-09 14:58 | 1554836284.555 | 1403 | 0.20955 | 0.000448 | 2e-06 | 99.8328 | 100.3032 | 122.8909 | 20.6073 | 0.001919 | 0.001906 | -154.8022 | 0.23176 | 8.4011 | 0.002694 | 0.003691 | 0.003536 | 3.5e-05 | 0.001092 | 0.27992 | 1.3697 | 2.5476 | 98.51 | 20.11 | 49.375 | 67.3738 | 85.984 | 67.0431 | 0.011811 | 1.0378 | 0.024222 | 23.3398 | -0.042044 | 0.13787 | -4.7e-05 | -0.003531 | -0.003531 | -1.5735
2019-04-09 14:58 | 1554836285.5547 | 1404 | 0.20955 | 0.000445 | 2e-06 | 99.7789 | 100.2712 | 122.7742 | 20.6076 | 0.001972 | 0.001877 | -154.7837 | 0.23183 | 8.4023 | 0.002705 | 0.004152 | 0.004006 | 3.9e-05 | 0.001089 | 0.28567 | 1.3726 | 2.5477 | 98.51 | 20.11 | 49.379 | 68.6307 | 86.2614 | 66.3676 | 0.011812 | 1.0398 | 0.024234 | 23.3055 | -0.04402 | 0.13792 | -3.8e-05 | -0.002907 | -0.002907 | -1.9116
2019-04-09 14:58 | 1554836286.5556 | 1405 | 0.20957 | 0.000445 | 2e-06 | 99.7679 | 100.1738 | 122.623 | 20.608 | 0.00201 | 0.001859 | -154.7709 | 0.23126 | 8.4042 | 0.002819 | 0.003818 | 0.00363 | 4.1e-05 | 0.001094 | 0.27743 | 1.3807 | 2.5477 | 98.51 | 20.11 | 49.381 | 69.5435 | 89.0052 | 65.9171 | 0.011813 | 1.0412 | 0.024235 | 23.2757 | -0.036316 | 0.13758 | -0.000167 | -0.01228 | -0.01228 | -0.4515
2019-04-09 14:58 | 1554836287.5556 | 1406 | 0.20955 | 0.000444 | 2e-06 | 99.7545 | 100.3091 | 122.2132 | 20.608 | 0.002041 | 0.001824 | -154.7505 | 0.23154 | 8.4014 | 0.002805 | 0.003368 | 0.003257 | 3.9e-05 | 0.001092 | 0.28512 | 1.3753 | 2.5476 | 98.51 | 20.11 | 49.385 | 70.2898 | 88.6659 | 65.0898 | 0.011814 | 1.0438 | 0.024224 | 23.2084 | -0.049586 | 0.13774 | -5.4e-05 | -0.00405 | -0.00405 | -1.3705
2019-04-09 14:58 | 1554836288.5547 | 1407 | 0.20955 | 0.000445 | 2e-06 | 99.7292 | 100.1508 | 122.4355 | 20.6079 | 0.002009 | 0.001946 | -154.7668 | 0.23247 | 8.4018 | 0.002712 | 0.004584 | 0.004331 | 4.5e-05 | 0.001093 | 0.27688 | 1.3802 | 2.5477 | 98.51 | 20.11 | 49.389 | 69.5412 | 86.4169 | 68.0218 | 0.011815 | 1.0348 | 0.024142 | 23.33 | -0.037728 | 0.13829 | -6.6e-05 | -0.004907 | -0.004907 | -1.1357
2019-04-09 14:58 | 1554836289.5548 | 1408 | 0.20955 | 0.000445 | 3e-06 | 99.7003 | 100.371 | 122.5355 | 20.6079 | 0.001979 | 0.001902 | -154.7545 | 0.23138 | 8.4048 | 0.002748 | 0.004163 | 0.003965 | 4.3e-05 | 0.001095 | 0.29103 | 1.3558 | 2.5477 | 98.51 | 20.11 | 49.392 | 68.819 | 87.2813 | 66.9572 | 0.011815 | 1.038 | 0.02419 | 23.3031 | -0.059969 | 0.13765 | -4e-05 | -0.003086 | -0.003086 | -1.7974
2019-04-09 14:58 | 1554836290.5557 | 1409 | 0.20952 | 0.000449 | 2e-06 | 99.683 | 100.1215 | 122.477 | 20.6067 | 0.001947 | 0.001914 | -154.771 | 0.23198 | 8.4015 | 0.002666 | 0.003861 | 0.00376 | 4.1e-05 | 0.001098 | 0.28584 | 1.3222 | 2.5475 | 98.51 | 20.11 | 49.398 | 68.0394 | 85.3262 | 67.2368 | 0.011817 | 1.0372 | 0.024174 | 23.3071 | -0.039256 | 0.13801 | 8e-05 | 0.005619 | 0.005619 | 0.98976
2019-04-09 14:58 | 1554836291.5552 | 1410 | 0.20959 | 0.000448 | 3e-06 | 99.6768 | 100.0688 | 122.5838 | 20.6073 | 0.001963 | 0.001868 | -154.748 | 0.23189 | 8.4022 | 0.002656 | 0.00407 | 0.003959 | 4e-05 | 0.001094 | 0.27129 | 1.4042 | 2.5477 | 98.51 | 20.11 | 49.4 | 68.4153 | 85.0733 | 66.1379 | 0.011817 | 1.0405 | 0.024224 | 23.2796 | -0.035112 | 0.13795 | -0.000303 | -0.022154 | -0.022154 | -0.25094
2019-04-09 14:58 | 1554836292.5552 | 1411 | 0.20957 | 0.000445 | 2e-06 | 99.6759 | 100.062 | 122.7953 | 20.6079 | 0.001959 | 0.00182 | -154.7484 | 0.23187 | 8.4019 | 0.002687 | 0.004009 | 0.003817 | 4.2e-05 | 0.001095 | 0.28047 | 1.3965 | 2.5476 | 98.51 | 20.11 | 49.401 | 68.3377 | 85.8208 | 64.9921 | 0.011817 | 1.0441 | 0.024285 | 23.2603 | -0.034577 | 0.13794 | -0.000176 | -0.012969 | -0.012969 | -0.42863
2019-04-09 14:58 | 1554836293.5558 | 1412 | 0.20954 | 0.000444 | 3e-06 | 99.6297 | 100.0895 | 122.6735 | 20.6075 | 0.001958 | 0.001839 | -154.7307 | 0.23118 | 8.4022 | 0.002699 | 0.003916 | 0.00379 | 3.9e-05 | 0.001094 | 0.28783 | 1.3666 | 2.5475 | 98.51 | 20.11 | 49.403 | 68.3102 | 86.1045 | 65.4389 | 0.011818 | 1.0427 | 0.024257 | 23.2641 | -0.04118 | 0.13753 | -1.3e-05 | -0.001104 | -0.001104 | -5.0217
2019-04-09 14:58 | 1554836294.5559 | 1413 | 0.20955 | 0.000442 | 3e-06 | 99.6306 | 99.9315 | 122.6083 | 20.6072 | 0.001978 | 0.001897 | -154.7462 | 0.2317 | 8.4008 | 0.002662 | 0.004003 | 0.003813 | 4.3e-05 | 0.001098 | 0.28871 | 1.3818 | 2.5476 | 98.51 | 20.11 | 49.405 | 68.7901 | 85.2221 | 66.8307 | 0.011818 | 1.0384 | 0.024201 | 23.3057 | -0.026972 | 0.13784 | -7.6e-05 | -0.005663 | -0.005663 | -0.98086
2019-04-09 14:58 | 1554836295.5549 | 1414 | 0.20958 | 0.000442 | 2e-06 | 99.5056 | 100.0083 | 122.3055 | 20.6088 | 0.002018 | 0.001877 | -154.7442 | 0.23154 | 8.4039 | 0.002653 | 0.003903 | 0.003766 | 4e-05 | 0.001103 | 0.2756 | 1.362 | 2.5476 | 98.51 | 20.11 | 49.406 | 69.7548 | 85.0098 | 66.3455 | 0.011819 | 1.0399 | 0.024189 | 23.2603 | -0.045075 | 0.13774 | -0.000215 | -0.015752 | -0.015752 | -0.3524
2019-04-09 14:58 | 1554836296.5549 | 1415 | 0.20953 | 0.000438 | 3e-06 | 99.5109 | 99.9539 | 122.3451 | 20.6084 | 0.001991 | 0.001897 | -154.7601 | 0.23176 | 8.4017 | 0.002637 | 0.003939 | 0.003803 | 4.4e-05 | 0.001104 | 0.2907 | 1.3711 | 2.5477 | 98.51 | 20.11 | 49.406 | 69.087 | 84.6193 | 66.8398 | 0.011819 | 1.0384 | 0.024175 | 23.281 | -0.039728 | 0.13787 | 7.7e-05 | 0.005434 | 0.005434 | 1.0225
2019-04-09 14:58 | 1554836297.554 | 1416 | 0.20953 | 0.000441 | 2e-06 | 99.4803 | 99.9694 | 122.539 | 20.608 | 0.002026 | 0.001878 | -154.767 | 0.23137 | 8.4009 | 0.002635 | 0.003811 | 0.003586 | 4.1e-05 | 0.001102 | 0.31078 | 1.3756 | 2.5475 | 98.51 | 20.11 | 49.405 | 69.9452 | 84.5776 | 66.3892 | 0.011818 | 1.0398 | 0.02421 | 23.284 | -0.043871 | 0.13764 | 2.7e-05 | 0.001844 | 0.001844 | 3.0074
2019-04-09 14:58 | 1554836298.5553 | 1417 | 0.20952 | 0.000444 | 3e-06 | 99.4478 | 99.8923 | 122.5529 | 20.6087 | 0.002001 | 0.001879 | -154.7644 | 0.23172 | 8.403 | 0.00268 | 0.004001 | 0.003795 | 4.4e-05 | 0.001107 | 0.2891 | 1.3325 | 2.5475 | 98.51 | 20.11 | 49.403 | 69.3416 | 85.663 | 66.4014 | 0.011818 | 1.0397 | 0.024211 | 23.2857 | -0.03989 | 0.13785 | 9e-05 | 0.006351 | 0.006351 | 0.87469
2019-04-09 14:58 | 1554836299.5545 | 1418 | 0.20953 | 0.000441 | 3e-06 | 99.4189 | 99.893 | 122.6986 | 20.6092 | 0.001983 | 0.001885 | -154.7635 | 0.23247 | 8.4052 | 0.002676 | 0.003701 | 0.003591 | 4e-05 | 0.0011 | 0.27649 | 1.3462 | 2.5474 | 98.51 | 20.11 | 49.404 | 68.8953 | 85.5637 | 66.5604 | 0.011818 | 1.0393 | 0.02422 | 23.305 | -0.04255 | 0.13829 | 6.3e-05 | 0.004443 | 0.004443 | 1.2543
2019-04-09 14:58 | 1554836300.5551 | 1419 | 0.20952 | 0.000439 | 2e-06 | 99.3656 | 99.7023 | 122.8667 | 20.6094 | 0.001998 | 0.001903 | -154.763 | 0.23231 | 8.4019 | 0.002714 | 0.004278 | 0.004123 | 4.3e-05 | 0.001098 | 0.28622 | 1.345 | 2.5476 | 98.51 | 20.12 | 49.405 | 69.2603 | 86.4653 | 66.9805 | 0.011826 | 1.038 | 0.024221 | 23.3354 | -0.030259 | 0.1382 | 0.00011 | 0.007874 | 0.007874 | 0.70736
2019-04-09 14:58 | 1554836301.5552 | 1420 | 0.20952 | 0.000443 | 3e-06 | 99.309 | 99.6108 | 122.5688 | 20.6093 | 0.00197 | 0.001866 | -154.7688 | 0.23172 | 8.4031 | 0.002649 | 0.004319 | 0.004146 | 4.1e-05 | 0.001104 | 0.29374 | 1.3221 | 2.5474 | 98.51 | 20.12 | 49.405 | 68.6013 | 84.916 | 66.0838 | 0.011826 | 1.0407 | 0.024224 | 23.2763 | -0.027137 | 0.13785 | 0.000124 | 0.008848 | 0.008848 | 0.62788
2019-04-09 14:58 | 1554836302.5552 | 1421 | 0.20953 | 0.000439 | 2e-06 | 99.2614 | 99.8396 | 122.4003 | 20.6085 | 0.001961 | 0.001925 | -154.7771 | 0.23241 | 8.4012 | 0.002633 | 0.004343 | 0.004126 | 4.3e-05 | 0.001102 | 0.29358 | 1.3566 | 2.5476 | 98.51 | 20.12 | 49.407 | 68.3859 | 84.5321 | 67.5154 | 0.011826 | 1.0363 | 0.024156 | 23.3093 | -0.051957 | 0.13826 | 6.8e-05 | 0.004804 | 0.004804 | 1.1599
2019-04-09 14:58 | 1554836303.5551 | 1422 | 0.20952 | 0.000446 | 3e-06 | 99.1968 | 99.5261 | 122.7911 | 20.6098 | 0.001928 | 0.001887 | -154.7814 | 0.23189 | 8.4013 | 0.002708 | 0.004124 | 0.003925 | 4.2e-05 | 0.0011 | 0.27936 | 1.3671 | 2.5476 | 98.51 | 20.12 | 49.411 | 67.5808 | 86.3197 | 66.5906 | 0.011827 | 1.0392 | 0.024228 | 23.3148 | -0.029649 | 0.13795 | 0.000122 | 0.008693 | 0.008693 | 0.63951
2019-04-09 14:58 | 1554836304.5551 | 1423 | 0.20954 | 0.000443 | 3e-06 | 99.1628 | 99.4456 | 122.9488 | 20.6101 | 0.001961 | 0.001846 | -154.7745 | 0.23232 | 8.4016 | 0.002739 | 0.003724 | 0.003555 | 4.2e-05 | 0.001095 | 0.27743 | 1.3643 | 2.5476 | 98.51 | 20.12 | 49.413 | 68.3711 | 87.0875 | 65.6107 | 0.011828 | 1.0422 | 0.024278 | 23.2961 | -0.02547 | 0.13821 | 2.2e-05 | 0.001429 | 0.001429 | 3.8985
2019-04-09 14:58 | 1554836305.5546 | 1424 | 0.20949 | 0.000441 | 2e-06 | 99.1098 | 99.426 | 123.0547 | 20.6085 | 0.00195 | 0.001907 | -154.7865 | 0.23251 | 8.4023 | 0.002624 | 0.004291 | 0.004113 | 4.3e-05 | 0.001098 | 0.28241 | 1.3794 | 2.5476 | 98.51 | 20.12 | 49.421 | 68.1114 | 84.2996 | 67.0801 | 0.01183 | 1.0377 | 0.024236 | 23.3566 | -0.028488 | 0.13832 | 0.000301 | 0.021731 | 0.021731 | 0.25651
2019-04-09 14:58 | 1554836306.5547 | 1425 | 0.20948 | 0.000444 | 3e-06 | 99.1306 | 99.3626 | 122.8007 | 20.6092 | 0.001901 | 0.00195 | -154.7846 | 0.23219 | 8.4022 | 0.002648 | 0.004223 | 0.004039 | 4.4e-05 | 0.001095 | 0.28119 | 1.3261 | 2.5476 | 98.51 | 20.12 | 49.426 | 66.9403 | 84.8986 | 68.1036 | 0.011831 | 1.0346 | 0.024175 | 23.3676 | -0.020909 | 0.13813 | 0.000349 | 0.025122 | 0.025122 | 0.22159
2019-04-09 14:58 | 1554836307.5553 | 1426 | 0.20952 | 0.000445 | 3e-06 | 99.0955 | 99.3816 | 122.7726 | 20.6094 | 0.001998 | 0.001879 | -154.8011 | 0.23284 | 8.4025 | 0.00263 | 0.004624 | 0.004477 | 4.3e-05 | 0.001093 | 0.29358 | 1.3855 | 2.5476 | 98.51 | 20.12 | 49.431 | 69.2666 | 84.4464 | 66.4146 | 0.011832 | 1.0397 | 0.024232 | 23.307 | -0.025783 | 0.13852 | 8.8e-05 | 0.006246 | 0.006246 | 0.89372
2019-04-09 14:58 | 1554836308.5554 | 1427 | 0.20955 | 0.000441 | 2e-06 | 99.0133 | 99.3416 | 122.7411 | 20.6095 | 0.001973 | 0.001906 | -154.7962 | 0.23282 | 8.4034 | 0.002566 | 0.00432 | 0.004145 | 4.3e-05 | 0.00109 | 0.30829 | 1.3224 | 2.5474 | 98.51 | 20.12 | 49.441 | 68.6576 | 82.9112 | 67.0508 | 0.011834 | 1.0378 | 0.024207 | 23.3258 | -0.029612 | 0.1385 | -4.5e-05 | -0.003425 | -0.003425 | -1.6299
2019-04-09 14:58 | 1554836309.5546 | 1428 | 0.20953 | 0.000444 | 2e-06 | 98.9663 | 99.458 | 122.5756 | 20.6082 | 0.002014 | 0.001897 | -154.7911 | 0.23211 | 8.4014 | 0.002629 | 0.00337 | 0.003201 | 4e-05 | 0.001088 | 0.27616 | 1.3342 | 2.5476 | 98.51 | 20.12 | 49.446 | 69.6584 | 84.439 | 66.8449 | 0.011836 | 1.0384 | 0.024198 | 23.3031 | -0.04433 | 0.13808 | 3.4e-05 | 0.002321 | 0.002321 | 2.3977
2019-04-09 14:58 | 1554836310.5558 | 1429 | 0.2095 | 0.000442 | 3e-06 | 98.9031 | 99.3169 | 122.3447 | 20.6093 | 0.001982 | 0.001904 | -154.7878 | 0.23287 | 8.4021 | 0.00273 | 0.003633 | 0.00338 | 4.3e-05 | 0.001086 | 0.30989 | 1.3733 | 2.5476 | 98.51 | 20.12 | 49.452 | 68.8891 | 86.8649 | 67.0052 | 0.011837 | 1.0379 | 0.024169 | 23.2866 | -0.037339 | 0.13853 | 0.000267 | 0.019189 | 0.019189 | 0.29094
2019-04-09 14:58 | 1554836311.5544 | 1430 | 0.20955 | 0.000444 | 3e-06 | 98.8546 | 99.1372 | 122.4944 | 20.6097 | 0.001896 | 0.001876 | -154.7978 | 0.2314 | 8.4028 | 0.002716 | 0.003866 | 0.003741 | 4.3e-05 | 0.001088 | 0.2943 | 1.3793 | 2.5477 | 98.51 | 20.12 | 49.463 | 66.8137 | 86.5178 | 66.3269 | 0.01184 | 1.04 | 0.024208 | 23.2776 | -0.025536 | 0.13766 | -7.5e-05 | -0.005653 | -0.005653 | -0.98137
2019-04-09 14:59 | 1554836312.5545 | 1431 | 0.20953 | 0.000444 | 3e-06 | 98.8623 | 99.3365 | 122.3409 | 20.6101 | 0.001966 | 0.001894 | -154.7877 | 0.23267 | 8.3997 | 0.002698 | 0.005051 | 0.004828 | 4.2e-05 | 0.001086 | 0.28407 | 1.3499 | 2.5475 | 98.51 | 20.12 | 49.468 | 68.5053 | 86.0859 | 66.7705 | 0.011841 | 1.0386 | 0.024177 | 23.2782 | -0.042796 | 0.13841 | 4.6e-05 | 0.003127 | 0.003127 | 1.7839
2019-04-09 14:59 | 1554836313.5535 | 1432 | 0.20954 | 0.000438 | 2e-06 | 98.8391 | 99.1166 | 122.2865 | 20.6095 | 0.001993 | 0.001845 | -154.7922 | 0.23171 | 8.4012 | 0.00267 | 0.003846 | 0.003695 | 4.4e-05 | 0.001089 | 0.30414 | 1.3297 | 2.5476 | 98.51 | 20.12 | 49.478 | 69.1551 | 85.4257 | 65.5931 | 0.011843 | 1.0422 | 0.024214 | 23.2327 | -0.025074 | 0.13784 | -2e-06 | -0.000297 | -0.000297 | -18.6812
2019-04-09 14:59 | 1554836314.5545 | 1433 | 0.20953 | 0.000443 | 3e-06 | 98.8117 | 98.9579 | 121.9844 | 20.6095 | 0.001953 | 0.001877 | -154.776 | 0.23246 | 8.4038 | 0.002673 | 0.003834 | 0.003625 | 4.3e-05 | 0.001087 | 0.29446 | 1.3704 | 2.5477 | 98.51 | 20.12 | 49.482 | 68.1803 | 85.4882 | 66.3594 | 0.011844 | 1.0399 | 0.024156 | 23.2302 | -0.013225 | 0.13829 | 7.2e-05 | 0.005047 | 0.005047 | 1.1043
2019-04-09 14:59 | 1554836315.5536 | 1434 | 0.20957 | 0.00044 | 3e-06 | 98.8049 | 99.1124 | 121.9736 | 20.6102 | 0.001986 | 0.001853 | -154.7787 | 0.23232 | 8.4024 | 0.002597 | 0.004241 | 0.004105 | 4.4e-05 | 0.00109 | 0.28512 | 1.3356 | 2.5476 | 98.51 | 20.12 | 49.489 | 68.9842 | 83.6638 | 65.7718 | 0.011846 | 1.0417 | 0.024176 | 23.209 | -0.027801 | 0.1382 | -0.000183 | -0.01344 | -0.01344 | -0.41441
2019-04-09 14:59 | 1554836316.554 | 1435 | 0.20951 | 0.00044 | 2e-06 | 98.6321 | 99.067 | 122.2735 | 20.61 | 0.001957 | 0.001867 | -154.7811 | 0.2322 | 8.4025 | 0.002608 | 0.003533 | 0.003348 | 4.4e-05 | 0.001094 | 0.29856 | 1.3521 | 2.5475 | 98.51 | 20.12 | 49.499 | 68.2775 | 83.9319 | 66.1297 | 0.011848 | 1.0406 | 0.024193 | 23.2498 | -0.03935 | 0.13814 | 0.000163 | 0.011678 | 0.011678 | 0.47671
2019-04-09 14:59 | 1554836317.555 | 1436 | 0.20954 | 0.000443 | 2e-06 | 98.6563 | 98.9314 | 122.3895 | 20.6105 | 0.001975 | 0.001903 | -154.7785 | 0.23239 | 8.4016 | 0.002605 | 0.004475 | 0.004288 | 4.4e-05 | 0.001101 | 0.26936 | 1.3473 | 2.5476 | 98.51 | 20.12 | 49.503 | 68.7061 | 83.8583 | 66.9806 | 0.011849 | 1.038 | 0.024174 | 23.29 | -0.024912 | 0.13825 | -2.1e-05 | -0.001621 | -0.001621 | -3.4363
2019-04-09 14:59 | 1554836318.5536 | 1437 | 0.20954 | 0.000441 | 2e-06 | 98.5994 | 98.7745 | 122.4153 | 20.6101 | 0.001978 | 0.001857 | -154.774 | 0.23216 | 8.4004 | 0.002679 | 0.003791 | 0.003583 | 4.6e-05 | 0.001098 | 0.28783 | 1.3617 | 2.5476 | 98.51 | 20.12 | 49.508 | 68.7891 | 85.6254 | 65.8676 | 0.01185 | 1.0414 | 0.024217 | 23.2543 | -0.01587 | 0.13811 | -5e-06 | -0.000518 | -0.000518 | -10.7379
2019-04-09 14:59 | 1554836319.5547 | 1438 | 0.20955 | 0.000443 | 3e-06 | 98.5642 | 98.823 | 122.321 | 20.6101 | 0.001953 | 0.001869 | -154.7746 | 0.23213 | 8.4027 | 0.002655 | 0.003783 | 0.003596 | 4e-05 | 0.001101 | 0.29623 | 1.3497 | 2.5476 | 98.51 | 20.12 | 49.511 | 68.1816 | 85.0688 | 66.1625 | 0.011851 | 1.0405 | 0.024197 | 23.2555 | -0.023451 | 0.13809 | -4e-05 | -0.003039 | -0.003039 | -1.8315
2019-04-09 14:59 | 1554836320.5547 | 1439 | 0.20954 | 0.000442 | 2e-06 | 98.4841 | 98.813 | 122.1991 | 20.6105 | 0.001984 | 0.001849 | -154.786 | 0.23261 | 8.4023 | 0.002637 | 0.00401 | 0.00384 | 4e-05 | 0.001102 | 0.27577 | 1.3343 | 2.5475 | 98.51 | 20.12 | 49.513 | 68.9253 | 84.6345 | 65.676 | 0.011852 | 1.042 | 0.024202 | 23.2272 | -0.029821 | 0.13838 | 1.5e-05 | 0.000961 | 0.000961 | 5.8056
2019-04-09 14:59 | 1554836321.5538 | 1440 | 0.20955 | 0.000441 | 2e-06 | 98.5479 | 98.832 | 122.1656 | 20.6109 | 0.001954 | 0.001886 | -154.7931 | 0.23258 | 8.3993 | 0.002585 | 0.004772 | 0.004638 | 4e-05 | 0.001094 | 0.29015 | 1.3489 | 2.5476 | 98.51 | 20.12 | 49.517 | 68.2113 | 83.3665 | 66.577 | 0.011853 | 1.0392 | 0.024167 | 23.2549 | -0.025748 | 0.13836 | -4e-05 | -0.003047 | -0.003047 | -1.83
2019-04-09 14:59 | 1554836322.5548 | 1441 | 0.20954 | 0.000443 | 2e-06 | 98.4591 | 98.8051 | 122.1196 | 20.6111 | 0.001935 | 0.001837 | -154.8043 | 0.23258 | 8.4002 | 0.00266 | 0.003678 | 0.003518 | 4.2e-05 | 0.001086 | 0.28047 | 1.3374 | 2.5475 | 98.51 | 20.12 | 49.519 | 67.745 | 85.1775 | 65.3863 | 0.011853 | 1.0429 | 0.024204 | 23.2097 | -0.031376 | 0.13836 | -6e-06 | -0.000527 | -0.000527 | -10.5812
2019-04-09 14:59 | 1554836323.5549 | 1442 | 0.20953 | 0.000446 | 2e-06 | 98.4502 | 98.8481 | 121.7419 | 20.6098 | 0.001923 | 0.001798 | -154.8055 | 0.23225 | 8.4028 | 0.00265 | 0.003783 | 0.003634 | 4.5e-05 | 0.001084 | 0.27776 | 1.336 | 2.5475 | 98.51 | 20.12 | 49.522 | 67.462 | 84.943 | 64.4569 | 0.011854 | 1.0457 | 0.0242 | 23.142 | -0.036081 | 0.13817 | 5.9e-05 | 0.004189 | 0.004189 | 1.3291
2019-04-09 14:59 | 1554836324.5553 | 1443 | 0.20952 | 0.000442 | 2e-06 | 98.3942 | 98.7187 | 121.6337 | 20.6111 | 0.001965 | 0.001889 | -154.8275 | 0.23316 | 8.4007 | 0.002644 | 0.003336 | 0.003178 | 4e-05 | 0.001089 | 0.30901 | 1.3089 | 2.5475 | 98.51 | 20.12 | 49.527 | 68.4711 | 84.7867 | 66.6426 | 0.011855 | 1.039 | 0.024112 | 23.2064 | -0.029452 | 0.13871 | 0.0001 | 0.007064 | 0.007064 | 0.79129
2019-04-09 14:59 | 1554836325.5553 | 1444 | 0.20955 | 0.000442 | 2e-06 | 98.4615 | 98.6671 | 122.0144 | 20.611 | 0.001987 | 0.001818 | -154.8143 | 0.23163 | 8.402 | 0.002618 | 0.003917 | 0.003742 | 4e-05 | 0.00109 | 0.27527 | 1.3433 | 2.5477 | 98.51 | 20.12 | 49.528 | 69.0124 | 84.1661 | 64.9388 | 0.011855 | 1.0442 | 0.02421 | 23.1844 | -0.018658 | 0.1378 | -7.8e-05 | -0.005791 | -0.005791 | -0.95886
2019-04-09 14:59 | 1554836326.5549 | 1445 | 0.20952 | 0.000443 | 2e-06 | 98.4407 | 98.7907 | 122.1386 | 20.6111 | 0.001945 | 0.00186 | -154.8043 | 0.23248 | 8.4034 | 0.002659 | 0.003652 | 0.003525 | 3.8e-05 | 0.00109 | 0.28711 | 1.3057 | 2.5475 | 98.51 | 20.12 | 49.531 | 68.0009 | 85.1487 | 65.9584 | 0.011856 | 1.0411 | 0.024186 | 23.2311 | -0.031745 | 0.13831 | 0.000106 | 0.007574 | 0.007574 | 0.73594
2019-04-09 14:59 | 1554836327.5553 | 1446 | 0.20957 | 0.000445 | 2e-06 | 98.3951 | 98.7325 | 122.1909 | 20.6114 | 0.00196 | 0.001864 | -154.7762 | 0.23208 | 8.4023 | 0.002631 | 0.004355 | 0.004152 | 4.2e-05 | 0.001092 | 0.29518 | 1.3212 | 2.5476 | 98.51 | 20.12 | 49.531 | 68.363 | 84.4741 | 66.0483 | 0.011856 | 1.0408 | 0.024188 | 23.2392 | -0.030618 | 0.13806 | -0.000175 | -0.012807 | -0.012807 | -0.43445
2019-04-09 14:59 | 1554836328.5542 | 1447 | 0.20954 | 0.000448 | 2e-06 | 98.3778 | 98.7256 | 122.075 | 20.6115 | 0.001952 | 0.001858 | -154.7748 | 0.23258 | 8.4008 | 0.002654 | 0.003478 | 0.003355 | 4.1e-05 | 0.001094 | 0.28655 | 1.3508 | 2.5477 | 98.51 | 20.12 | 49.531 | 68.1556 | 85.0304 | 65.9031 | 0.011856 | 1.0413 | 0.024182 | 23.2232 | -0.031565 | 0.13836 | 2e-05 | 0.001288 | 0.001288 | 4.3278
2019-04-09 14:59 | 1554836329.5529 | 1448 | 0.20955 | 0.000443 | 3e-06 | 98.2974 | 98.5095 | 122.3043 | 20.6103 | 0.001912 | 0.001898 | -154.7715 | 0.23202 | 8.4026 | 0.002618 | 0.003735 | 0.003599 | 4.2e-05 | 0.001103 | 0.27168 | 1.3581 | 2.5476 | 98.51 | 20.12 | 49.53 | 67.2011 | 84.1748 | 66.8666 | 0.011856 | 1.0383 | 0.02417 | 23.278 | -0.019276 | 0.13803 | -8.8e-05 | -0.006549 | -0.006549 | -0.84942
2019-04-09 14:59 | 1554836330.5541 | 1449 | 0.20954 | 0.000446 | 3e-06 | 98.262 | 98.5215 | 122.5817 | 20.6107 | 0.001977 | 0.001905 | -154.7785 | 0.23236 | 8.4012 | 0.002599 | 0.003461 | 0.003263 | 4.4e-05 | 0.0011 | 0.29286 | 1.3096 | 2.5475 | 98.51 | 20.12 | 49.53 | 68.7531 | 83.7072 | 67.0377 | 0.011856 | 1.0378 | 0.024191 | 23.3102 | -0.023594 | 0.13823 | 1.8e-05 | 0.001116 | 0.001116 | 4.9932
2019-04-09 14:59 | 1554836331.5535 | 1450 | 0.20952 | 0.000445 | 2e-06 | 98.2799 | 98.5023 | 122.6843 | 20.6119 | 0.001912 | 0.001836 | -154.7577 | 0.23239 | 8.4001 | 0.002643 | 0.003802 | 0.003651 | 3.9e-05 | 0.001098 | 0.29479 | 1.3273 | 2.5476 | 98.51 | 20.12 | 49.532 | 67.1967 | 84.7608 | 65.3722 | 0.011856 | 1.0429 | 0.024261 | 23.2628 | -0.020218 | 0.13825 | 0.000104 | 0.007413 | 0.007413 | 0.75156
2019-04-09 14:59 | 1554836332.5531 | 1451 | 0.20952 | 0.000446 | 2e-06 | 98.1441 | 98.4228 | 122.5381 | 20.6122 | 0.001969 | 0.001872 | -154.7607 | 0.23193 | 8.4036 | 0.002624 | 0.003919 | 0.003702 | 4.3e-05 | 0.001096 | 0.31564 | 1.3664 | 2.5476 | 98.51 | 20.12 | 49.531 | 68.5635 | 84.3219 | 66.2511 | 0.011856 | 1.0402 | 0.024215 | 23.2791 | -0.025364 | 0.13797 | 0.000102 | 0.00723 | 0.00723 | 0.76902
2019-04-09 14:59 | 1554836333.5535 | 1452 | 0.20955 | 0.000442 | 3e-06 | 98.1521 | 98.2937 | 122.524 | 20.6114 | 0.001992 | 0.001869 | -154.7543 | 0.23269 | 8.4022 | 0.002672 | 0.003452 | 0.003313 | 4.3e-05 | 0.001107 | 0.27793 | 1.3554 | 2.5475 | 98.51 | 20.12 | 49.53 | 69.1274 | 85.4736 | 66.1654 | 0.011856 | 1.0405 | 0.024217 | 23.2749 | -0.012894 | 0.13843 | -3e-05 | -0.002347 | -0.002347 | -2.3769
2019-04-09 14:59 | 1554836334.5546 | 1453 | 0.20951 | 0.000441 | 2e-06 | 98.1521 | 98.3037 | 122.778 | 20.6118 | 0.001932 | 0.001876 | -154.7492 | 0.23308 | 8.4024 | 0.002648 | 0.003903 | 0.003718 | 4.5e-05 | 0.001105 | 0.28783 | 1.3188 | 2.5476 | 98.51 | 20.12 | 49.527 | 67.6832 | 84.884 | 66.3452 | 0.011855 | 1.0399 | 0.024235 | 23.3051 | -0.013803 | 0.13866 | 0.000158 | 0.011321 | 0.011321 | 0.49359
2019-04-09 14:59 | 1554836335.5532 | 1454 | 0.20947 | 0.000443 | 3e-06 | 98.1735 | 98.4045 | 122.6198 | 20.6125 | 0.001992 | 0.001847 | -154.7538 | 0.23301 | 8.4026 | 0.002672 | 0.004321 | 0.004156 | 4.3e-05 | 0.001104 | 0.27151 | 1.3234 | 2.5477 | 98.51 | 20.12 | 49.525 | 69.1285 | 85.4661 | 65.6381 | 0.011855 | 1.0421 | 0.024245 | 23.2658 | -0.021019 | 0.13862 | 0.000429 | 0.031033 | 0.031033 | 0.18001
2019-04-09 14:59 | 1554836336.5542 | 1455 | 0.20945 | 0.000447 | 2e-06 | 98.0866 | 98.326 | 122.4322 | 20.612 | 0.001928 | 0.001877 | -154.7594 | 0.2313 | 8.4004 | 0.002607 | 0.004178 | 0.00397 | 4.3e-05 | 0.001104 | 0.28655 | 1.3387 | 2.5478 | 98.51 | 20.12 | 49.521 | 67.5913 | 83.9105 | 66.3548 | 0.011854 | 1.0399 | 0.024201 | 23.2726 | -0.02181 | 0.1376 | 0.000537 | 0.038764 | 0.038764 | 0.14305
2019-04-09 14:59 | 1554836337.5532 | 1456 | 0.20951 | 0.000443 | 2e-06 | 98.0083 | 98.3942 | 122.3036 | 20.6111 | 0.001919 | 0.001891 | -154.7873 | 0.23206 | 8.4031 | 0.002546 | 0.003675 | 0.003461 | 4.5e-05 | 0.001106 | 0.29142 | 1.3832 | 2.5477 | 98.51 | 20.12 | 49.518 | 67.3744 | 82.4423 | 66.695 | 0.011853 | 1.0388 | 0.024176 | 23.272 | -0.035151 | 0.13805 | 0.000168 | 0.012076 | 0.012076 | 0.46071
2019-04-09 14:59 | 1554836338.5539 | 1457 | 0.20951 | 0.000441 | 2e-06 | 98.0229 | 98.1636 | 122.3982 | 20.6121 | 0.001959 | 0.001903 | -154.79 | 0.23235 | 8.4016 | 0.002604 | 0.004315 | 0.004214 | 4.5e-05 | 0.001104 | 0.2891 | 1.3422 | 2.5475 | 98.51 | 20.12 | 49.517 | 68.3324 | 83.8318 | 66.9932 | 0.011853 | 1.0379 | 0.024175 | 23.2913 | -0.012834 | 0.13822 | 0.000209 | 0.015023 | 0.015023 | 0.37078
2019-04-09 14:59 | 1554836339.5533 | 1458 | 0.2095 | 0.000447 | 2e-06 | 98.0449 | 98.3075 | 122.3815 | 20.6123 | 0.001939 | 0.001885 | -154.7994 | 0.23256 | 8.4001 | 0.002611 | 0.004395 | 0.00415 | 4.5e-05 | 0.001098 | 0.27295 | 1.3578 | 2.5476 | 98.51 | 20.12 | 49.514 | 67.8398 | 84.0068 | 66.5495 | 0.011852 | 1.0393 | 0.024189 | 23.2745 | -0.023922 | 0.13835 | 0.000233 | 0.01676 | 0.01676 | 0.33265
2019-04-09 14:59 | 1554836340.5544 | 1459 | 0.20956 | 0.000442 | 3e-06 | 98.0672 | 98.3185 | 122.2768 | 20.6124 | 0.001918 | 0.001875 | -154.7822 | 0.23239 | 8.3985 | 0.002594 | 0.003869 | 0.003843 | 4.4e-05 | 0.0011 | 0.30055 | 1.3519 | 2.5476 | 98.51 | 20.12 | 49.512 | 67.3414 | 83.5942 | 66.3028 | 0.011851 | 1.04 | 0.024187 | 23.2561 | -0.022887 | 0.13825 | -0.000129 | -0.009482 | -0.009482 | -0.58754
2019-04-09 14:59 | 1554836341.5537 | 1460 | 0.20958 | 0.000442 | 2e-06 | 97.9684 | 98.2338 | 122.2996 | 20.6125 | 0.001935 | 0.001925 | -154.7999 | 0.23317 | 8.4019 | 0.002643 | 0.003844 | 0.00366 | 4.2e-05 | 0.001099 | 0.29214 | 1.3711 | 2.5476 | 98.51 | 20.12 | 49.511 | 67.7635 | 84.7762 | 67.524 | 0.011851 | 1.0363 | 0.024146 | 23.3 | -0.024203 | 0.13871 | -0.00022 | -0.016029 | -0.016029 | -0.34876
2019-04-09 14:59 | 1554836342.5548 | 1461 | 0.20957 | 0.000445 | 3e-06 | 97.9725 | 98.2496 | 122.5604 | 20.6129 | 0.001924 | 0.001858 | -154.7897 | 0.23208 | 8.4015 | 0.002716 | 0.003696 | 0.003498 | 4.5e-05 | 0.001105 | 0.30685 | 1.3321 | 2.5475 | 98.51 | 20.12 | 49.509 | 67.4886 | 86.5327 | 65.9045 | 0.011851 | 1.0413 | 0.02423 | 23.2694 | -0.025264 | 0.13806 | -0.000183 | -0.013422 | -0.013422 | -0.41455
2019-04-09 14:59 | 1554836343.5545 | 1462 | 0.20963 | 0.00045 | 2e-06 | 98.0038 | 98.2469 | 122.977 | 20.6124 | 0.001926 | 0.001844 | -154.7969 | 0.23325 | 8.4008 | 0.002642 | 0.003688 | 0.003472 | 4.5e-05 | 0.001107 | 0.2959 | 1.3688 | 2.5477 | 98.51 | 20.12 | 49.508 | 67.5316 | 84.7567 | 65.56 | 0.01185 | 1.0423 | 0.024283 | 23.297 | -0.022159 | 0.13876 | -0.000545 | -0.039792 | -0.039792 | -0.14053
2019-04-09 14:59 | 1554836344.5525 | 1463 | 0.20955 | 0.000441 | 2e-06 | 97.9017 | 98.2228 | 122.7911 | 20.613 | 0.002019 | 0.001857 | -154.7911 | 0.23156 | 8.4042 | 0.002636 | 0.003909 | 0.003742 | 4.5e-05 | 0.001109 | 0.30901 | 1.3805 | 2.5476 | 98.51 | 20.12 | 49.507 | 69.7731 | 84.6043 | 65.8906 | 0.01185 | 1.0413 | 0.024253 | 23.2908 | -0.029293 | 0.13776 | -3.5e-05 | -0.002682 | -0.002682 | -2.0702
2019-04-09 14:59 | 1554836345.5526 | 1464 | 0.20955 | 0.000446 | 3e-06 | 97.9657 | 98.1588 | 122.6766 | 20.6128 | 0.002005 | 0.001904 | -154.7907 | 0.23256 | 8.4057 | 0.002672 | 0.003831 | 0.003631 | 4.1e-05 | 0.001107 | 0.2855 | 1.3695 | 2.5476 | 98.51 | 20.12 | 49.505 | 69.4408 | 85.4712 | 67.004 | 0.01185 | 1.0379 | 0.024202 | 23.3181 | -0.017614 | 0.13835 | -3.7e-05 | -0.002833 | -0.002833 | -1.9683
2019-04-09 14:59 | 1554836346.554 | 1465 | 0.20957 | 0.000447 | 2e-06 | 97.9684 | 98.123 | 122.6437 | 20.6122 | 0.001969 | 0.001861 | -154.785 | 0.23251 | 8.4008 | 0.002677 | 0.004741 | 0.004524 | 4.6e-05 | 0.001101 | 0.30879 | 1.3384 | 2.5476 | 98.51 | 20.12 | 49.506 | 68.5729 | 85.5955 | 65.9875 | 0.01185 | 1.041 | 0.024235 | 23.2801 | -0.014107 | 0.13832 | -0.000186 | -0.013684 | -0.013684 | -0.40735
2019-04-09 14:59 | 1554836347.5536 | 1466 | 0.20954 | 0.000449 | 2e-06 | 97.9499 | 98.0941 | 122.5397 | 20.6129 | 0.001923 | 0.001853 | -154.7854 | 0.23242 | 8.3997 | 0.002655 | 0.003573 | 0.003386 | 4.2e-05 | 0.001096 | 0.2855 | 1.3237 | 2.5476 | 98.51 | 20.12 | 49.505 | 67.473 | 85.0583 | 65.7845 | 0.01185 | 1.0416 | 0.024232 | 23.2633 | -0.013157 | 0.13827 | -1.5e-05 | -0.001202 | -0.001202 | -4.636
2019-04-09 14:59 | 1554836348.5536 | 1467 | 0.20954 | 0.000447 | 2e-06 | 97.9335 | 98.0862 | 122.3494 | 20.6128 | 0.001893 | 0.001853 | -154.7985 | 0.23256 | 8.4003 | 0.002604 | 0.003953 | 0.003757 | 4.3e-05 | 0.001079 | 0.29679 | 1.3098 | 2.5477 | 98.51 | 20.12 | 49.504 | 66.7363 | 83.8302 | 65.7726 | 0.011849 | 1.0417 | 0.024213 | 23.2448 | -0.013931 | 0.13835 | -2e-05 | -0.001586 | -0.001586 | -3.5158
2019-04-09 14:59 | 1554836349.5531 | 1468 | 0.20958 | 0.000442 | 2e-06 | 97.93 | 98.1464 | 122.2388 | 20.6136 | 0.001988 | 0.001822 | -154.805 | 0.23254 | 8.4004 | 0.002594 | 0.003754 | 0.003648 | 4.2e-05 | 0.001066 | 0.29502 | 1.318 | 2.5476 | 98.51 | 20.12 | 49.502 | 69.0367 | 83.5985 | 65.0358 | 0.011849 | 1.0439 | 0.024229 | 23.209 | -0.019748 | 0.13834 | -0.000252 | -0.018435 | -0.018435 | -0.30241
2019-04-09 14:59 | 1554836350.5541 | 1469 | 0.2096 | 0.000444 | 2e-06 | 97.8463 | 97.9888 | 122.2224 | 20.6131 | 0.001995 | 0.001864 | -154.8073 | 0.23201 | 8.4005 | 0.002608 | 0.003483 | 0.003275 | 4.6e-05 | 0.001061 | 0.29391 | 1.3371 | 2.5477 | 98.51 | 20.12 | 49.501 | 69.1937 | 83.9224 | 66.0563 | 0.011849 | 1.0408 | 0.024191 | 23.2424 | -0.013018 | 0.13802 | -0.000338 | -0.024641 | -0.024641 | -0.22574
2019-04-09 14:59 | 1554836351.5527 | 1470 | 0.20956 | 0.000447 | 2e-06 | 97.9544 | 98.0115 | 121.9959 | 20.613 | 0.001956 | 0.001845 | -154.7927 | 0.23324 | 8.4003 | 0.002619 | 0.004218 | 0.003992 | 4.4e-05 | 0.001066 | 0.2756 | 1.3374 | 2.5475 | 98.51 | 20.12 | 49.499 | 68.2672 | 84.1993 | 65.5886 | 0.011848 | 1.0422 | 0.024185 | 23.2049 | -0.005215 | 0.13875 | -0.000141 | -0.010354 | -0.010354 | -0.54004
2019-04-09 14:59 | 1554836352.5528 | 1471 | 0.20956 | 0.000447 | 2e-06 | 97.8743 | 97.9785 | 122.0585 | 20.6131 | 0.00202 | 0.001848 | -154.79 | 0.232 | 8.4042 | 0.002656 | 0.003735 | 0.003633 | 4.7e-05 | 0.001073 | 0.3115 | 1.2962 | 2.5476 | 98.51 | 20.13 | 49.495 | 69.7941 | 85.0936 | 65.6582 | 0.011855 | 1.042 | 0.024189 | 23.2132 | -0.009517 | 0.13802 | -0.000111 | -0.008183 | -0.008183 | -0.67975
2019-04-09 14:59 | 1554836353.5538 | 1472 | 0.20958 | 0.000445 | 2e-06 | 97.9213 | 98.1488 | 121.7893 | 20.6134 | 0.00197 | 0.00181 | -154.7845 | 0.23268 | 8.4023 | 0.002642 | 0.004174 | 0.003971 | 4.8e-05 | 0.001078 | 0.30127 | 1.3188 | 2.5477 | 98.51 | 20.13 | 49.494 | 68.5911 | 84.7569 | 64.7484 | 0.011854 | 1.0448 | 0.024194 | 23.1565 | -0.020756 | 0.13842 | -0.000219 | -0.01599 | -0.01599 | -0.34887
2019-04-09 14:59 | 1554836354.5542 | 1473 | 0.20958 | 0.000443 | 2e-06 | 97.851 | 98.261 | 121.8548 | 20.6142 | 0.001922 | 0.001836 | -154.7902 | 0.23243 | 8.3996 | 0.002672 | 0.004454 | 0.004291 | 4.9e-05 | 0.001082 | 0.29502 | 1.3622 | 2.5476 | 98.51 | 20.13 | 49.492 | 67.4485 | 85.4591 | 65.3706 | 0.011854 | 1.0429 | 0.024179 | 23.184 | -0.037397 | 0.13827 | -0.000244 | -0.017799 | -0.017799 | -0.31307
2019-04-09 14:59 | 1554836355.5532 | 1474 | 0.20953 | 0.000445 | 2e-06 | 97.8972 | 98.2022 | 121.7518 | 20.6143 | 0.001957 | 0.001905 | -154.7916 | 0.23211 | 8.4025 | 0.002718 | 0.003731 | 0.003599 | 4.5e-05 | 0.001083 | 0.31365 | 1.3673 | 2.5476 | 98.51 | 20.13 | 49.491 | 68.2824 | 86.5723 | 67.0461 | 0.011854 | 1.0378 | 0.024109 | 23.2315 | -0.02782 | 0.13809 | 6.1e-05 | 0.004288 | 0.004288 | 1.2978
2019-04-09 14:59 | 1554836356.5538 | 1475 | 0.20953 | 0.000444 | 2e-06 | 97.8787 | 98.0425 | 122.0184 | 20.6131 | 0.001968 | 0.001862 | -154.7964 | 0.23234 | 8.4024 | 0.002767 | 0.004406 | 0.004281 | 4.4e-05 | 0.001085 | 0.30917 | 1.3509 | 2.5476 | 98.51 | 20.13 | 49.49 | 68.5596 | 87.7549 | 66.0048 | 0.011853 | 1.041 | 0.024172 | 23.2213 | -0.01495 | 0.13822 | 7.2e-05 | 0.005062 | 0.005062 | 1.1004
2019-04-09 14:59 | 1554836357.5529 | 1476 | 0.20956 | 0.000441 | 2e-06 | 97.871 | 98.0793 | 122.1045 | 20.6138 | 0.001993 | 0.001857 | -154.7864 | 0.23288 | 8.4023 | 0.00275 | 0.004285 | 0.004067 | 4.4e-05 | 0.001081 | 0.28423 | 1.3476 | 2.5476 | 98.51 | 20.13 | 49.49 | 69.1538 | 87.3347 | 65.8931 | 0.011853 | 1.0413 | 0.024185 | 23.2257 | -0.019017 | 0.13854 | -0.00013 | -0.009522 | -0.009522 | -0.58638
2019-04-09 14:59 | 1554836358.5531 | 1477 | 0.20955 | 0.000443 | 2e-06 | 97.8734 | 98.2104 | 121.8393 | 20.6149 | 0.002048 | 0.001897 | -154.7704 | 0.2328 | 8.4012 | 0.002741 | 0.004269 | 0.004055 | 4.8e-05 | 0.001079 | 0.30165 | 1.348 | 2.5475 | 98.51 | 20.13 | 49.488 | 70.4666 | 87.1336 | 66.8538 | 0.011853 | 1.0384 | 0.024124 | 23.2333 | -0.030749 | 0.13849 | -3.9e-05 | -0.002982 | -0.002982 | -1.8713
2019-04-09 14:59 | 1554836359.5522 | 1478 | 0.20958 | 0.000441 | 2e-06 | 97.8862 | 98.26 | 121.7846 | 20.614 | 0.002022 | 0.001832 | -154.7555 | 0.23238 | 8.4031 | 0.002718 | 0.003633 | 0.003402 | 4.6e-05 | 0.001076 | 0.28711 | 1.3338 | 2.5477 | 98.51 | 20.13 | 49.486 | 69.8443 | 86.5728 | 65.2907 | 0.011853 | 1.0432 | 0.024175 | 23.1746 | -0.034091 | 0.13824 | -0.000225 | -0.016406 | -0.016406 | -0.33958
2019-04-09 14:59 | 1554836360.552 | 1479 | 0.20951 | 0.000446 | 2e-06 | 97.9675 | 98.2018 | 121.8872 | 20.6131 | 0.001943 | 0.001826 | -154.7617 | 0.23213 | 8.4012 | 0.002727 | 0.004379 | 0.004251 | 4.7e-05 | 0.001076 | 0.26504 | 1.3863 | 2.5477 | 98.51 | 20.13 | 49.484 | 67.9465 | 86.7961 | 65.1337 | 0.011852 | 1.0436 | 0.02419 | 23.179 | -0.02137 | 0.1381 | 0.000197 | 0.014125 | 0.014125 | 0.39401
2019-04-09 14:59 | 1554836361.553 | 1480 | 0.20953 | 0.000445 | 2e-06 | 97.9008 | 98.0889 | 121.8886 | 20.6139 | 0.001989 | 0.001895 | -154.7607 | 0.23292 | 8.4032 | 0.002766 | 0.004356 | 0.004154 | 4.6e-05 | 0.00107 | 0.2808 | 1.3199 | 2.5477 | 98.51 | 20.13 | 49.483 | 69.051 | 87.7377 | 66.7913 | 0.011852 | 1.0385 | 0.024132 | 23.2358 | -0.017175 | 0.13856 | 7.3e-05 | 0.00511 | 0.00511 | 1.0927
2019-04-09 14:59 | 1554836362.5535 | 1481 | 0.20954 | 0.000447 | 2e-06 | 97.9106 | 98.1385 | 121.4741 | 20.6137 | 0.002025 | 0.001887 | -154.7746 | 0.2334 | 8.4012 | 0.002865 | 0.00378 | 0.003686 | 4.9e-05 | 0.001073 | 0.26465 | 1.3506 | 2.5476 | 98.51 | 20.13 | 49.484 | 69.9281 | 90.1177 | 66.5967 | 0.011852 | 1.0391 | 0.024097 | 23.1896 | -0.020795 | 0.13885 | -6e-06 | -0.000511 | -0.000511 | -10.9476
2019-04-09 14:59 | 1554836363.5525 | 1482 | 0.20954 | 0.00044 | 2e-06 | 97.8972 | 98.2748 | 121.472 | 20.613 | 0.001924 | 0.001904 | -154.7808 | 0.23108 | 8.4014 | 0.002749 | 0.003364 | 0.003198 | 4.5e-05 | 0.001076 | 0.28528 | 1.3556 | 2.5476 | 98.51 | 20.13 | 49.484 | 67.4988 | 87.3291 | 67.015 | 0.011852 | 1.0379 | 0.024082 | 23.2037 | -0.034432 | 0.13747 | -1.6e-05 | -0.001321 | -0.001321 | -4.1951
2019-04-09 14:59 | 1554836364.5531 | 1483 | 0.20953 | 0.000442 | 2e-06 | 97.9174 | 98.2266 | 121.836 | 20.6148 | 0.001933 | 0.001848 | -154.7745 | 0.23152 | 8.4043 | 0.002875 | 0.004344 | 0.004088 | 4.8e-05 | 0.00107 | 0.27007 | 1.3181 | 2.5477 | 98.51 | 20.13 | 49.484 | 67.7206 | 90.358 | 65.663 | 0.011852 | 1.042 | 0.024166 | 23.1922 | -0.028196 | 0.13773 | 7.5e-05 | 0.005297 | 0.005297 | 1.0478
2019-04-09 14:59 | 1554836365.5532 | 1484 | 0.20955 | 0.000442 | 3e-06 | 97.9329 | 98.2166 | 121.7623 | 20.6146 | 0.00191 | 0.001865 | -154.7755 | 0.23206 | 8.4021 | 0.002782 | 0.003904 | 0.003774 | 4.5e-05 | 0.001064 | 0.2855 | 1.3638 | 2.5476 | 98.51 | 20.13 | 49.485 | 67.1507 | 88.1099 | 66.0733 | 0.011852 | 1.0407 | 0.024145 | 23.1992 | -0.025872 | 0.13805 | -4.3e-05 | -0.003313 | -0.003313 | -1.6796
2019-04-09 14:59 | 1554836366.5532 | 1485 | 0.20952 | 0.000439 | 3e-06 | 98.0273 | 98.2249 | 121.4825 | 20.615 | 0.001975 | 0.001831 | -154.781 | 0.23158 | 8.3988 | 0.002796 | 0.004101 | 0.00385 | 4.8e-05 | 0.001052 | 0.2907 | 1.3586 | 2.5475 | 98.51 | 20.13 | 49.485 | 68.7122 | 88.4541 | 65.2619 | 0.011852 | 1.0432 | 0.024146 | 23.1448 | -0.018007 | 0.13777 | 0.00015 | 0.010718 | 0.010718 | 0.51802
2019-04-09 14:59 | 1554836367.5522 | 1486 | 0.20953 | 0.00044 | 2e-06 | 98.0458 | 98.3343 | 121.3814 | 20.6158 | 0.001981 | 0.001872 | -154.7919 | 0.23242 | 8.4022 | 0.002812 | 0.004781 | 0.004581 | 4.9e-05 | 0.001056 | 0.29502 | 1.3611 | 2.5477 | 98.51 | 20.13 | 49.485 | 68.8697 | 88.8524 | 66.2376 | 0.011852 | 1.0402 | 0.024101 | 23.1685 | -0.026282 | 0.13827 | 9e-05 | 0.006348 | 0.006348 | 0.87774
2019-04-09 14:59 | 1554836368.5532 | 1487 | 0.20951 | 0.000443 | 2e-06 | 98.0321 | 98.082 | 121.8942 | 20.6147 | 0.001994 | 0.001863 | -154.7978 | 0.23232 | 8.4028 | 0.002701 | 0.004207 | 0.004002 | 4.9e-05 | 0.001058 | 0.30431 | 1.3328 | 2.5476 | 98.51 | 20.13 | 49.484 | 69.1782 | 86.1645 | 66.0227 | 0.011852 | 1.0409 | 0.024159 | 23.2101 | -0.004557 | 0.13821 | 0.000186 | 0.01332 | 0.01332 | 0.41816
2019-04-09 14:59 | 1554836369.5527 | 1488 | 0.20953 | 0.00044 | 2e-06 | 98.0589 | 98.2827 | 121.7858 | 20.6149 | 0.001978 | 0.001898 | -154.8116 | 0.23238 | 8.4018 | 0.002708 | 0.004174 | 0.004017 | 4.9e-05 | 0.001056 | 0.28694 | 1.3599 | 2.5475 | 98.51 | 20.13 | 49.483 | 68.7931 | 86.3263 | 66.8718 | 0.011852 | 1.0383 | 0.024118 | 23.2288 | -0.02039 | 0.13824 | 5.8e-05 | 0.004086 | 0.004086 | 1.3636
2019-04-09 14:59 | 1554836370.5523 | 1489 | 0.20953 | 0.00045 | 2e-06 | 98.0878 | 98.2053 | 121.6649 | 20.6153 | 0.001983 | 0.001829 | -154.7952 | 0.23146 | 8.4006 | 0.002807 | 0.004199 | 0.004041 | 4.8e-05 | 0.001061 | 0.31166 | 1.3634 | 2.5476 | 98.51 | 20.13 | 49.484 | 68.9137 | 88.7099 | 65.2148 | 0.011852 | 1.0434 | 0.024165 | 23.1606 | -0.010706 | 0.13769 | 6.3e-05 | 0.004451 | 0.004451 | 1.2467
2019-04-09 15:00 | 1554836371.5533 | 1490 | 0.2095 | 0.000446 | 2e-06 | 98.0139 | 98.4059 | 121.6004 | 20.615 | 0.001952 | 0.001831 | -154.8184 | 0.23294 | 8.4069 | 0.00288 | 0.004219 | 0.004043 | 4.6e-05 | 0.001065 | 0.29286 | 1.3403 | 2.5475 | 98.51 | 20.13 | 49.484 | 68.1638 | 90.4714 | 65.2455 | 0.011852 | 1.0433 | 0.024158 | 23.1555 | -0.035699 | 0.13858 | 0.000241 | 0.017301 | 0.017301 | 0.32279
2019-04-09 15:00 | 1554836372.5524 | 1491 | 0.20954 | 0.000444 | 2e-06 | 98.0041 | 98.2803 | 122.0987 | 20.6145 | 0.001991 | 0.001849 | -154.7946 | 0.23338 | 8.3996 | 0.002816 | 0.004031 | 0.003843 | 4.7e-05 | 0.001068 | 0.29264 | 1.2873 | 2.5475 | 98.51 | 20.13 | 49.485 | 69.1086 | 88.9434 | 65.6783 | 0.011852 | 1.042 | 0.024192 | 23.2177 | -0.02517 | 0.13884 | -1.2e-05 | -0.001005 | -0.001005 | -5.5672
2019-04-09 15:00 | 1554836373.5524 | 1492 | 0.2095 | 0.000445 | 2e-06 | 98.0791 | 98.3257 | 122.086 | 20.6156 | 0.001961 | 0.001908 | -154.8048 | 0.23236 | 8.4008 | 0.00276 | 0.004336 | 0.004145 | 4.7e-05 | 0.001073 | 0.30021 | 1.3281 | 2.5475 | 98.51 | 20.13 | 49.484 | 68.3913 | 87.6004 | 67.1213 | 0.011852 | 1.0375 | 0.024139 | 23.2659 | -0.022458 | 0.13823 | 0.00026 | 0.01871 | 0.01871 | 0.29775
2019-04-09 15:00 | 1554836374.5534 | 1493 | 0.20952 | 0.000441 | 2e-06 | 98.1039 | 98.4888 | 121.6879 | 20.6154 | 0.001957 | 0.001882 | -154.7973 | 0.23319 | 8.3965 | 0.002754 | 0.004039 | 0.003883 | 4.9e-05 | 0.001081 | 0.29407 | 1.3108 | 2.5475 | 98.51 | 20.13 | 49.482 | 68.2974 | 87.4448 | 66.4791 | 0.011852 | 1.0395 | 0.024123 | 23.206 | -0.035031 | 0.13873 | 0.000121 | 0.008618 | 0.008618 | 0.6487
2019-04-09 15:00 | 1554836375.5515 | 1494 | 0.20952 | 0.000443 | 2e-06 | 97.9844 | 98.3219 | 121.7332 | 20.6157 | 0.001936 | 0.001852 | -154.8062 | 0.23193 | 8.4013 | 0.002764 | 0.004192 | 0.004007 | 4.7e-05 | 0.001082 | 0.29695 | 1.3771 | 2.5477 | 98.51 | 20.14 | 49.48 | 67.7846 | 87.6855 | 65.7708 | 0.011858 | 1.0417 | 0.024152 | 23.1861 | -0.030754 | 0.13798 | 0.000123 | 0.008806 | 0.008806 | 0.63143
2019-04-09 15:00 | 1554836376.5526 | 1495 | 0.20952 | 0.000438 | 2e-06 | 98.1366 | 98.262 | 121.7515 | 20.6161 | 0.0019 | 0.001857 | -154.7875 | 0.23225 | 8.4037 | 0.00278 | 0.003479 | 0.003234 | 4.7e-05 | 0.00108 | 0.30287 | 1.3544 | 2.5476 | 98.51 | 20.14 | 49.479 | 66.922 | 88.0748 | 65.8944 | 0.011858 | 1.0413 | 0.02415 | 23.1921 | -0.011424 | 0.13816 | 0.000156 | 0.011179 | 0.011179 | 0.49806
2019-04-09 15:00 | 1554836377.5515 | 1496 | 0.2095 | 0.000442 | 2e-06 | 98.0565 | 98.2362 | 121.8515 | 20.6166 | 0.001963 | 0.001846 | -154.7949 | 0.23209 | 8.4044 | 0.002786 | 0.004128 | 0.003971 | 4.4e-05 | 0.001079 | 0.28982 | 1.3508 | 2.5476 | 98.51 | 20.14 | 49.475 | 68.4399 | 88.2183 | 65.6132 | 0.011857 | 1.0422 | 0.02417 | 23.192 | -0.016378 | 0.13807 | 0.000234 | 0.01678 | 0.01678 | 0.3316
2019-04-09 15:00 | 1554836378.5516 | 1497 | 0.20952 | 0.000441 | 2e-06 | 98.0696 | 98.3529 | 121.6445 | 20.6164 | 0.00196 | 0.001911 | -154.7982 | 0.23247 | 8.4022 | 0.002696 | 0.004302 | 0.004153 | 4.8e-05 | 0.001081 | 0.27489 | 1.321 | 2.5476 | 98.51 | 20.14 | 49.473 | 68.3678 | 86.0506 | 67.176 | 0.011857 | 1.0374 | 0.024094 | 23.2257 | -0.0258 | 0.1383 | 0.000102 | 0.007215 | 0.007215 | 0.77248
2019-04-09 15:00 | 1554836379.5525 | 1498 | 0.20952 | 0.000443 | 2e-06 | 98.1188 | 98.4104 | 121.7135 | 20.6157 | 0.001989 | 0.001835 | -154.7893 | 0.23187 | 8.4013 | 0.002742 | 0.00405 | 0.003854 | 4.7e-05 | 0.001082 | 0.28407 | 1.3424 | 2.5476 | 98.51 | 20.14 | 49.47 | 69.052 | 87.1531 | 65.3571 | 0.011856 | 1.0429 | 0.024165 | 23.1701 | -0.026545 | 0.13794 | 0.000127 | 0.009108 | 0.009108 | 0.6103
2019-04-09 15:00 | 1554836380.5517 | 1499 | 0.20952 | 0.000438 | 2e-06 | 98.0741 | 98.3161 | 121.794 | 20.6162 | 0.00192 | 0.001859 | -154.8021 | 0.23251 | 8.402 | 0.002668 | 0.003464 | 0.003242 | 4.6e-05 | 0.001087 | 0.29407 | 1.3226 | 2.5475 | 98.51 | 20.14 | 49.467 | 67.3917 | 85.3761 | 65.925 | 0.011855 | 1.0412 | 0.024153 | 23.1972 | -0.022043 | 0.13832 | 0.000117 | 0.008328 | 0.008328 | 0.66939
2019-04-09 15:00 | 1554836381.5516 | 1500 | 0.20954 | 0.000439 | 2e-06 | 98.1399 | 98.3192 | 122.1644 | 20.6154 | 0.002027 | 0.001856 | -154.7962 | 0.23259 | 8.401 | 0.002741 | 0.003661 | 0.003514 | 4.8e-05 | 0.00109 | 0.25984 | 1.3701 | 2.5477 | 98.51 | 20.14 | 49.465 | 69.9742 | 87.123 | 65.8617 | 0.011855 | 1.0414 | 0.024192 | 23.2303 | -0.016324 | 0.13837 | 1.1e-05 | 0.000671 | 0.000671 | 8.3154
2019-04-09 15:00 | 1554836382.5517 | 1501 | 0.20953 | 0.00044 | 2e-06 | 98.0919 | 98.2865 | 121.7276 | 20.6154 | 0.001987 | 0.001893 | -154.8122 | 0.2326 | 8.3987 | 0.002716 | 0.003769 | 0.003619 | 4.8e-05 | 0.001096 | 0.30215 | 1.3941 | 2.5476 | 98.51 | 20.14 | 49.459 | 69.0048 | 86.5399 | 66.7537 | 0.011853 | 1.0387 | 0.024117 | 23.2192 | -0.01772 | 0.13837 | 5e-05 | 0.00346 | 0.00346 | 1.6118
2019-04-09 15:00 | 1554836383.5528 | 1502 | 0.20951 | 0.00044 | 2e-06 | 98.0654 | 98.2789 | 121.4865 | 20.6161 | 0.001981 | 0.001865 | -154.7859 | 0.23236 | 8.4019 | 0.002729 | 0.003712 | 0.003514 | 4.7e-05 | 0.001101 | 0.2959 | 1.4013 | 2.5477 | 98.51 | 20.14 | 49.457 | 68.8644 | 86.8349 | 66.0841 | 0.011853 | 1.0407 | 0.024117 | 23.1733 | -0.019448 | 0.13823 | 0.000214 | 0.015371 | 0.015371 | 0.36243
2019-04-09 15:00 | 1554836384.5518 | 1503 | 0.20952 | 0.000436 | 2e-06 | 98.1578 | 98.2748 | 121.4234 | 20.616 | 0.001979 | 0.001897 | -154.8025 | 0.23258 | 8.4027 | 0.002722 | 0.003971 | 0.003773 | 4.7e-05 | 0.001102 | 0.29015 | 1.4184 | 2.5477 | 98.51 | 20.14 | 49.451 | 68.8262 | 86.6787 | 66.8487 | 0.011852 | 1.0384 | 0.024083 | 23.1934 | -0.010656 | 0.13836 | 0.000129 | 0.00914 | 0.00914 | 0.61008
2019-04-09 15:00 | 1554836385.5518 | 1504 | 0.20954 | 0.000439 | 2e-06 | 98.1613 | 98.5071 | 121.775 | 20.6162 | 0.001938 | 0.001818 | -154.7899 | 0.23271 | 8.401 | 0.002725 | 0.004146 | 0.004048 | 4.9e-05 | 0.001108 | 0.29374 | 1.3908 | 2.5476 | 98.51 | 20.14 | 49.45 | 67.8209 | 86.7397 | 64.9385 | 0.011851 | 1.0442 | 0.024186 | 23.1616 | -0.031449 | 0.13844 | 3e-06 | 8.1e-05 | 8.1e-05 | 68.5188
2019-04-09 15:00 | 1554836386.5532 | 1505 | 0.20952 | 0.000441 | 2e-06 | 98.1777 | 98.5288 | 121.8257 | 20.6161 | 0.001928 | 0.001854 | -154.8305 | 0.23327 | 8.4021 | 0.002675 | 0.004135 | 0.00391 | 4.9e-05 | 0.001105 | 0.29789 | 1.3756 | 2.5475 | 98.51 | 20.14 | 49.447 | 67.5953 | 85.5523 | 65.8102 | 0.011851 | 1.0416 | 0.02416 | 23.1963 | -0.031925 | 0.13877 | 0.000132 | 0.009449 | 0.009449 | 0.59184
2019-04-09 15:00 | 1554836387.5522 | 1506 | 0.2095 | 0.000447 | 2e-06 | 98.251 | 98.5449 | 121.7011 | 20.6159 | 0.00194 | 0.00188 | -154.8122 | 0.23264 | 8.4012 | 0.002688 | 0.004637 | 0.004452 | 4.9e-05 | 0.001105 | 0.30702 | 1.3723 | 2.5477 | 98.51 | 20.14 | 49.445 | 67.8745 | 85.8596 | 66.4428 | 0.01185 | 1.0396 | 0.024125 | 23.206 | -0.026722 | 0.13839 | 0.000226 | 0.016212 | 0.016212 | 0.34402
2019-04-09 15:00 | 1554836388.5522 | 1507 | 0.20952 | 0.000437 | 2e-06 | 98.1846 | 98.2321 | 121.8532 | 20.6158 | 0.002019 | 0.001844 | -154.8054 | 0.23206 | 8.4002 | 0.002744 | 0.00384 | 0.00368 | 4.9e-05 | 0.00111 | 0.2844 | 1.3615 | 2.5475 | 98.51 | 20.14 | 49.444 | 69.7742 | 87.2086 | 65.5819 | 0.01185 | 1.0423 | 0.024171 | 23.1911 | -0.004329 | 0.13805 | 0.000166 | 0.01191 | 0.01191 | 0.46712
2019-04-09 15:00 | 1554836389.5519 | 1508 | 0.2095 | 0.000444 | 2e-06 | 98.2149 | 98.4448 | 122.0151 | 20.6153 | 0.001929 | 0.001874 | -154.7905 | 0.23232 | 8.4026 | 0.002729 | 0.004056 | 0.003921 | 4.8e-05 | 0.001111 | 0.26305 | 1.3603 | 2.5476 | 98.51 | 20.14 | 49.443 | 67.6171 | 86.8435 | 66.294 | 0.01185 | 1.0401 | 0.024162 | 23.2309 | -0.020908 | 0.1382 | 0.000224 | 0.016077 | 0.016077 | 0.34645
2019-04-09 15:00 | 1554836390.551 | 1509 | 0.20952 | 0.000445 | 2e-06 | 98.2754 | 98.6595 | 121.9435 | 20.6172 | 0.002012 | 0.001788 | -154.787 | 0.23387 | 8.3978 | 0.002711 | 0.004774 | 0.004524 | 5.1e-05 | 0.001117 | 0.28926 | 1.3553 | 2.5476 | 98.51 | 20.14 | 49.445 | 69.6195 | 86.4019 | 64.2126 | 0.01185 | 1.0465 | 0.024229 | 23.1527 | -0.034895 | 0.13913 | 0.000153 | 0.010986 | 0.010986 | 0.51035
2019-04-09 15:00 | 1554836391.5519 | 1510 | 0.20952 | 0.000444 | 2e-06 | 98.3013 | 98.4881 | 121.6619 | 20.6158 | 0.001945 | 0.001813 | -154.7726 | 0.23288 | 8.4002 | 0.002724 | 0.004079 | 0.003925 | 5e-05 | 0.001124 | 0.29175 | 1.4163 | 2.5475 | 98.51 | 20.14 | 49.444 | 67.9929 | 86.7174 | 64.8154 | 0.01185 | 1.0446 | 0.024179 | 23.1466 | -0.016986 | 0.13854 | 0.000126 | 0.009041 | 0.009041 | 0.61756
2019-04-09 15:00 | 1554836392.5509 | 1511 | 0.20953 | 0.000442 | 2e-06 | 98.34 | 98.6675 | 122.1559 | 20.6165 | 0.002007 | 0.001853 | -154.7746 | 0.23288 | 8.4011 | 0.002745 | 0.00443 | 0.00418 | 4.9e-05 | 0.001124 | 0.3063 | 1.3622 | 2.5478 | 98.51 | 20.14 | 49.445 | 69.4997 | 87.2257 | 65.7825 | 0.01185 | 1.0416 | 0.024194 | 23.2267 | -0.029733 | 0.13854 | 7e-05 | 0.004989 | 0.004989 | 1.119
2019-04-09 15:00 | 1554836393.551 | 1512 | 0.20951 | 0.000446 | 2e-06 | 98.2766 | 98.505 | 122.447 | 20.6178 | 0.00198 | 0.001873 | -154.7973 | 0.23211 | 8.4022 | 0.002708 | 0.00402 | 0.003885 | 4.8e-05 | 0.001129 | 0.29407 | 1.3868 | 2.5477 | 98.51 | 20.14 | 49.446 | 68.8543 | 86.3454 | 66.2792 | 0.01185 | 1.0401 | 0.024205 | 23.2714 | -0.020766 | 0.13808 | 0.000173 | 0.012416 | 0.012416 | 0.4482
2019-04-09 15:00 | 1554836394.5511 | 1513 | 0.20952 | 0.000447 | 2e-06 | 98.2965 | 98.493 | 122.408 | 20.617 | 0.001859 | 0.001824 | -154.8033 | 0.23292 | 8.4008 | 0.002681 | 0.004244 | 0.003983 | 4.7e-05 | 0.001126 | 0.29839 | 1.3607 | 2.5476 | 98.51 | 20.14 | 49.447 | 65.9282 | 85.7006 | 65.0968 | 0.011851 | 1.0438 | 0.024243 | 23.2272 | -0.017857 | 0.13856 | 0.000137 | 0.009825 | 0.009825 | 0.56835
2019-04-09 15:00 | 1554836395.5515 | 1514 | 0.20953 | 0.000441 | 2e-06 | 98.2355 | 98.6299 | 122.2012 | 20.6171 | 0.001981 | 0.001801 | -154.8128 | 0.23261 | 8.4025 | 0.002726 | 0.003692 | 0.003541 | 4.8e-05 | 0.00113 | 0.30845 | 1.3806 | 2.5476 | 98.51 | 20.14 | 49.447 | 68.8754 | 86.774 | 64.5397 | 0.011851 | 1.0455 | 0.024243 | 23.1884 | -0.035844 | 0.13838 | 4.3e-05 | 0.00301 | 0.00301 | 1.8525
2019-04-09 15:00 | 1554836396.5525 | 1515 | 0.20952 | 0.000446 | 2e-06 | 98.3725 | 98.7163 | 122.3639 | 20.6179 | 0.001981 | 0.001845 | -154.7958 | 0.23236 | 8.4013 | 0.002751 | 0.003789 | 0.0036 | 4.7e-05 | 0.001128 | 0.29966 | 1.3438 | 2.5476 | 98.51 | 20.14 | 49.449 | 68.877 | 87.3815 | 65.588 | 0.011851 | 1.0422 | 0.024221 | 23.2398 | -0.03121 | 0.13823 | 0.000142 | 0.010167 | 0.010167 | 0.54793
2019-04-09 15:00 | 1554836397.5515 | 1516 | 0.20953 | 0.000447 | 2e-06 | 98.3808 | 98.6365 | 122.6636 | 20.6169 | 0.00202 | 0.001874 | -154.8173 | 0.23213 | 8.4035 | 0.002691 | 0.003653 | 0.003422 | 4.8e-05 | 0.001128 | 0.30254 | 1.3994 | 2.5477 | 98.51 | 20.14 | 49.449 | 69.8127 | 85.9308 | 66.2956 | 0.011851 | 1.0401 | 0.024226 | 23.2926 | -0.023215 | 0.1381 | 3.4e-05 | 0.002345 | 0.002345 | 2.3734
2019-04-09 15:00 | 1554836398.5522 | 1517 | 0.20953 | 0.000446 | 2e-06 | 98.4091 | 98.6819 | 122.5641 | 20.6178 | 0.001976 | 0.001814 | -154.8089 | 0.23217 | 8.4045 | 0.002728 | 0.004635 | 0.00452 | 4.6e-05 | 0.001125 | 0.2896 | 1.3816 | 2.5476 | 98.51 | 20.14 | 49.448 | 68.7577 | 86.8147 | 64.8472 | 0.011851 | 1.0445 | 0.024268 | 23.2334 | -0.024762 | 0.13812 | 8.7e-05 | 0.006216 | 0.006216 | 0.89547
2019-04-09 15:00 | 1554836399.5526 | 1518 | 0.20952 | 0.000445 | 2e-06 | 98.4678 | 98.7191 | 122.3062 | 20.6172 | 0.002012 | 0.001875 | -154.8209 | 0.23204 | 8.402 | 0.002694 | 0.004402 | 0.004209 | 4.8e-05 | 0.001127 | 0.28727 | 1.364 | 2.5476 | 98.51 | 20.14 | 49.447 | 69.615 | 86.0117 | 66.3123 | 0.011851 | 1.04 | 0.02419 | 23.2592 | -0.022799 | 0.13804 | 0.000123 | 0.008846 | 0.008846 | 0.62886
2019-04-09 15:00 | 1554836400.5527 | 1519 | 0.20949 | 0.000446 | 2e-06 | 98.4785 | 98.7704 | 122.5148 | 20.6177 | 0.001975 | 0.00184 | -154.8276 | 0.23243 | 8.403 | 0.002694 | 0.004377 | 0.004166 | 4.5e-05 | 0.001125 | 0.27992 | 1.3481 | 2.5476 | 98.51 | 20.15 | 49.445 | 68.7218 | 85.9973 | 65.4652 | 0.011857 | 1.0426 | 0.024241 | 23.25 | -0.026471 | 0.13828 | 0.000298 | 0.021482 | 0.021482 | 0.2594
2019-04-09 15:00 | 1554836401.5524 | 1520 | 0.2095 | 0.000444 | 2e-06 | 98.4818 | 98.7322 | 122.6338 | 20.6182 | 0.001929 | 0.00184 | -154.8554 | 0.23325 | 8.4017 | 0.002727 | 0.004265 | 0.004134 | 4.7e-05 | 0.001124 | 0.27312 | 1.3204 | 2.5476 | 98.51 | 20.15 | 49.443 | 67.6123 | 86.7901 | 65.4729 | 0.011857 | 1.0426 | 0.024252 | 23.2615 | -0.022713 | 0.13876 | 0.000224 | 0.016166 | 0.016166 | 0.34591
2019-04-09 15:00 | 1554836402.5514 | 1521 | 0.20951 | 0.000446 | 3e-06 | 98.5059 | 98.8912 | 122.2235 | 20.6177 | 0.001944 | 0.001875 | -154.8467 | 0.23236 | 8.4028 | 0.002704 | 0.003799 | 0.003558 | 4.8e-05 | 0.001127 | 0.29286 | 1.3292 | 2.5475 | 98.51 | 20.15 | 49.446 | 67.9854 | 86.2455 | 66.319 | 0.011858 | 1.04 | 0.024181 | 23.2516 | -0.034915 | 0.13823 | 0.000186 | 0.013346 | 0.013346 | 0.41741
2019-04-09 15:00 | 1554836403.5528 | 1522 | 0.20948 | 0.000447 | 2e-06 | 98.5922 | 99.0471 | 121.9147 | 20.6182 | 0.001917 | 0.001882 | -154.8464 | 0.23217 | 8.4004 | 0.002692 | 0.00437 | 0.004223 | 4.5e-05 | 0.001126 | 0.31277 | 1.3877 | 2.5476 | 98.51 | 20.15 | 49.45 | 67.3182 | 85.9456 | 66.4826 | 0.011859 | 1.0395 | 0.024145 | 23.2278 | -0.041167 | 0.13812 | 0.000353 | 0.025432 | 0.025432 | 0.21886
2019-04-09 15:00 | 1554836404.5516 | 1523 | 0.20952 | 0.000443 | 3e-06 | 98.5872 | 98.8457 | 122.0656 | 20.6171 | 0.001884 | 0.001832 | -154.8175 | 0.23216 | 8.4025 | 0.002662 | 0.004522 | 0.0043 | 4.5e-05 | 0.001125 | 0.286 | 1.4072 | 2.5476 | 98.51 | 20.15 | 49.455 | 66.5236 | 85.2233 | 65.2774 | 0.01186 | 1.0432 | 0.024203 | 23.2008 | -0.023427 | 0.13811 | 0.000138 | 0.009871 | 0.009871 | 0.56388
2019-04-09 15:00 | 1554836405.5518 | 1524 | 0.20952 | 0.000448 | 2e-06 | 98.6884 | 99.0498 | 122.1606 | 20.6188 | 0.001957 | 0.001849 | -154.8267 | 0.23344 | 8.4009 | 0.002653 | 0.004229 | 0.004115 | 4.6e-05 | 0.00113 | 0.29574 | 1.3662 | 2.5476 | 98.51 | 20.15 | 49.465 | 68.2906 | 85.0101 | 65.6995 | 0.011862 | 1.0419 | 0.024197 | 23.2244 | -0.032694 | 0.13887 | 0.000133 | 0.009534 | 0.009534 | 0.58699
2019-04-09 15:00 | 1554836406.5518 | 1525 | 0.20952 | 0.000446 | 3e-06 | 98.6062 | 98.906 | 122.14 | 20.6189 | 0.002006 | 0.001821 | -154.8304 | 0.23264 | 8.4001 | 0.002701 | 0.003595 | 0.003436 | 4.6e-05 | 0.001129 | 0.29695 | 1.3565 | 2.5476 | 98.51 | 20.15 | 49.471 | 69.4621 | 86.1779 | 65.027 | 0.011864 | 1.044 | 0.024219 | 23.1993 | -0.027146 | 0.1384 | 0.00015 | 0.010731 | 0.010731 | 0.51976
2019-04-09 15:00 | 1554836407.5506 | 1526 | 0.20954 | 0.000445 | 2e-06 | 98.6893 | 99.0247 | 122.0677 | 20.6181 | 0.001918 | 0.001781 | -154.8196 | 0.23351 | 8.4006 | 0.002714 | 0.003917 | 0.003737 | 4.5e-05 | 0.001128 | 0.30397 | 1.3972 | 2.5476 | 98.51 | 20.15 | 49.476 | 67.3414 | 86.4894 | 64.0635 | 0.011865 | 1.0469 | 0.024247 | 23.1594 | -0.030344 | 0.13891 | 2.8e-05 | 0.001909 | 0.001909 | 2.9328
2019-04-09 15:00 | 1554836408.5505 | 1527 | 0.20952 | 0.000444 | 2e-06 | 98.6908 | 99.0147 | 121.9628 | 20.6184 | 0.001995 | 0.001877 | -154.8312 | 0.23312 | 8.4011 | 0.002702 | 0.004011 | 0.003787 | 4.8e-05 | 0.00113 | 0.28368 | 1.3815 | 2.5477 | 98.51 | 20.15 | 49.487 | 69.1943 | 86.1923 | 66.3594 | 0.011867 | 1.0399 | 0.024154 | 23.2281 | -0.029307 | 0.13868 | 0.000122 | 0.008693 | 0.008693 | 0.64289
2019-04-09 15:00 | 1554836409.5516 | 1528 | 0.20952 | 0.000441 | 3e-06 | 98.6712 | 98.9173 | 122.1559 | 20.6181 | 0.002006 | 0.001808 | -154.8355 | 0.23278 | 8.4031 | 0.00272 | 0.003921 | 0.00374 | 4.4e-05 | 0.001126 | 0.30901 | 1.3559 | 2.5478 | 98.51 | 20.15 | 49.492 | 69.4675 | 86.6222 | 64.7014 | 0.011869 | 1.045 | 0.024233 | 23.1897 | -0.022286 | 0.13848 | 0.000163 | 0.011697 | 0.011697 | 0.47711
2019-04-09 15:00 | 1554836410.5506 | 1529 | 0.20952 | 0.000443 | 2e-06 | 98.675 | 98.8633 | 122.055 | 20.6179 | 0.001932 | 0.001875 | -154.8277 | 0.23319 | 8.4031 | 0.002727 | 0.004412 | 0.004207 | 4.4e-05 | 0.001128 | 0.29983 | 1.3653 | 2.5475 | 98.51 | 20.15 | 49.502 | 67.6783 | 86.7905 | 66.3161 | 0.011871 | 1.04 | 0.024165 | 23.2354 | -0.017048 | 0.13872 | 0.000118 | 0.008413 | 0.008413 | 0.66452
2019-04-09 15:00 | 1554836411.5506 | 1530 | 0.20954 | 0.000443 | 2e-06 | 98.6703 | 98.961 | 122.0053 | 20.6186 | 0.002002 | 0.001858 | -154.8034 | 0.23285 | 8.4002 | 0.002699 | 0.003924 | 0.003721 | 4.5e-05 | 0.001128 | 0.27936 | 1.3433 | 2.5476 | 98.51 | 20.15 | 49.508 | 69.3659 | 86.1223 | 65.9149 | 0.011873 | 1.0412 | 0.024174 | 23.217 | -0.026318 | 0.13852 | 2e-05 | 0.001315 | 0.001315 | 4.2453
2019-04-09 15:00 | 1554836412.5517 | 1531 | 0.2095 | 0.000448 | 2e-06 | 98.6357 | 98.8003 | 122.0614 | 20.6191 | 0.00195 | 0.001839 | -154.8119 | 0.2326 | 8.4 | 0.002764 | 0.003768 | 0.003589 | 4.5e-05 | 0.001127 | 0.30215 | 1.3227 | 2.5476 | 98.51 | 20.15 | 49.517 | 68.121 | 87.6874 | 65.4415 | 0.011875 | 1.0427 | 0.024197 | 23.2061 | -0.014912 | 0.13837 | 0.000249 | 0.017945 | 0.017945 | 0.31076
2019-04-09 15:00 | 1554836413.5524 | 1532 | 0.20952 | 0.000443 | 2e-06 | 98.6607 | 98.8333 | 122.278 | 20.6192 | 0.001985 | 0.001847 | -154.7934 | 0.23215 | 8.4025 | 0.002744 | 0.004009 | 0.003804 | 4.8e-05 | 0.001129 | 0.29192 | 1.3669 | 2.5477 | 98.51 | 20.15 | 49.536 | 68.9672 | 87.2174 | 65.64 | 0.011879 | 1.0421 | 0.024211 | 23.2335 | -0.015635 | 0.13811 | 0.000162 | 0.011596 | 0.011596 | 0.47998
2019-04-09 15:00 | 1554836414.5511 | 1533 | 0.2095 | 0.000443 | 3e-06 | 98.689 | 98.9187 | 122.3217 | 20.6183 | 0.001929 | 0.001836 | -154.8149 | 0.23161 | 8.4023 | 0.002758 | 0.004455 | 0.004267 | 4.7e-05 | 0.001132 | 0.29662 | 1.3213 | 2.5475 | 98.51 | 20.15 | 49.546 | 67.6116 | 87.5318 | 65.3752 | 0.011882 | 1.0429 | 0.024225 | 23.2285 | -0.020791 | 0.13778 | 0.000276 | 0.019848 | 0.019848 | 0.27976
2019-04-09 15:00 | 1554836415.5517 | 1534 | 0.20954 | 0.000443 | 2e-06 | 98.6691 | 99.0715 | 122.3909 | 20.6185 | 0.001993 | 0.001847 | -154.7997 | 0.2316 | 8.4034 | 0.002813 | 0.003629 | 0.003515 | 4.5e-05 | 0.001133 | 0.29607 | 1.3259 | 2.5475 | 98.51 | 20.15 | 49.563 | 69.1451 | 88.8685 | 65.6425 | 0.011886 | 1.0421 | 0.024222 | 23.2443 | -0.036406 | 0.13778 | -2e-06 | -0.000303 | -0.000303 | -18.3102
2019-04-09 15:00 | 1554836416.5508 | 1535 | 0.20955 | 0.000446 | 2e-06 | 98.7072 | 99.0667 | 122.5571 | 20.6191 | 0.001917 | 0.00184 | -154.7928 | 0.23225 | 8.4023 | 0.002721 | 0.003971 | 0.003804 | 4.7e-05 | 0.001128 | 0.29015 | 1.3695 | 2.5476 | 98.51 | 20.15 | 49.57 | 67.3285 | 86.6473 | 65.4658 | 0.011887 | 1.0426 | 0.024245 | 23.254 | -0.032516 | 0.13816 | -5.2e-05 | -0.003919 | -0.003919 | -1.4206
2019-04-09 15:00 | 1554836417.5522 | 1536 | 0.20955 | 0.000441 | 2e-06 | 98.7346 | 99.0178 | 122.7165 | 20.6189 | 0.001975 | 0.001861 | -154.7908 | 0.23192 | 8.403 | 0.002748 | 0.004504 | 0.004256 | 4.6e-05 | 0.001133 | 0.29767 | 1.3471 | 2.5475 | 98.51 | 20.15 | 49.577 | 68.7243 | 87.2989 | 65.976 | 0.011889 | 1.041 | 0.024242 | 23.2866 | -0.02562 | 0.13797 | -3.5e-05 | -0.00269 | -0.00269 | -2.0667
2019-04-09 15:00 | 1554836418.5509 | 1537 | 0.20953 | 0.000447 | 2e-06 | 98.7533 | 99.1035 | 122.9615 | 20.6197 | 0.00193 | 0.001808 | -154.8014 | 0.23209 | 8.403 | 0.002792 | 0.003811 | 0.003601 | 4.4e-05 | 0.001131 | 0.30718 | 1.3378 | 2.5476 | 98.51 | 20.15 | 49.591 | 67.6325 | 88.372 | 64.6999 | 0.011892 | 1.045 | 0.024312 | 23.266 | -0.031659 | 0.13807 | 6e-05 | 0.004203 | 0.004203 | 1.3238
2019-04-09 15:00 | 1554836419.5509 | 1538 | 0.20958 | 0.000444 | 2e-06 | 98.7816 | 99.0119 | 122.8409 | 20.6189 | 0.001872 | 0.00184 | -154.8057 | 0.23224 | 8.3996 | 0.002783 | 0.004363 | 0.004216 | 4.5e-05 | 0.001139 | 0.29679 | 1.3474 | 2.5476 | 98.51 | 20.16 | 49.596 | 66.2468 | 88.1541 | 65.4656 | 0.011901 | 1.0426 | 0.024273 | 23.2809 | -0.020831 | 0.13816 | -0.000201 | -0.014754 | -0.014754 | -0.37737
2019-04-09 15:00 | 1554836420.5511 | 1539 | 0.20955 | 0.000446 | 2e-06 | 98.7516 | 99.0574 | 122.7627 | 20.6204 | 0.002057 | 0.001818 | -154.8144 | 0.23176 | 8.4016 | 0.002745 | 0.004459 | 0.004236 | 4.5e-05 | 0.001139 | 0.28479 | 1.3718 | 2.5476 | 98.51 | 20.16 | 49.603 | 70.7019 | 87.2426 | 64.9503 | 0.011903 | 1.0442 | 0.024284 | 23.2558 | -0.027657 | 0.13788 | -3.3e-05 | -0.002539 | -0.002539 | -2.1886
2019-04-09 15:00 | 1554836421.5513 | 1540 | 0.20953 | 0.000447 | 2e-06 | 98.7501 | 98.9338 | 122.7817 | 20.6188 | 0.001954 | 0.001827 | -154.8236 | 0.23216 | 8.402 | 0.002762 | 0.00477 | 0.004555 | 4.5e-05 | 0.001138 | 0.30453 | 1.3129 | 2.5475 | 98.51 | 20.16 | 49.616 | 68.2204 | 87.6488 | 65.1713 | 0.011906 | 1.0435 | 0.024278 | 23.2652 | -0.016629 | 0.13811 | 5.3e-05 | 0.003711 | 0.003711 | 1.5
2019-04-09 15:00 | 1554836422.55 | 1541 | 0.20955 | 0.000444 | 2e-06 | 98.7712 | 99.0708 | 123.2195 | 20.6203 | 0.001955 | 0.001835 | -154.8133 | 0.23283 | 8.4031 | 0.00275 | 0.004348 | 0.004257 | 4.5e-05 | 0.001138 | 0.2975 | 1.3404 | 2.5477 | 98.51 | 20.16 | 49.621 | 68.2423 | 87.3487 | 65.3686 | 0.011907 | 1.0429 | 0.024314 | 23.3134 | -0.027089 | 0.13851 | -6.4e-05 | -0.004788 | -0.004788 | -1.1658
2019-04-09 15:00 | 1554836423.5514 | 1542 | 0.20959 | 0.000448 | 2e-06 | 98.7989 | 99.099 | 123.3465 | 20.6199 | 0.001931 | 0.001813 | -154.8252 | 0.23231 | 8.4042 | 0.002718 | 0.003676 | 0.003537 | 4.5e-05 | 0.001141 | 0.28838 | 1.3947 | 2.5477 | 98.51 | 20.16 | 49.625 | 67.6754 | 86.5777 | 64.8297 | 0.011908 | 1.0446 | 0.024346 | 23.3068 | -0.027129 | 0.1382 | -0.000293 | -0.02153 | -0.02153 | -0.25868
2019-04-09 15:00 | 1554836424.5504 | 1543 | 0.20956 | 0.000444 | 2e-06 | 98.8835 | 99.311 | 123.2364 | 20.6202 | 0.001933 | 0.001798 | -154.8216 | 0.23171 | 8.4012 | 0.002756 | 0.004028 | 0.003882 | 4.4e-05 | 0.00114 | 0.29231 | 1.3726 | 2.5476 | 98.51 | 20.16 | 49.623 | 67.707 | 87.5002 | 64.4739 | 0.011907 | 1.0457 | 0.024348 | 23.2842 | -0.03859 | 0.13784 | -9.8e-05 | -0.007284 | -0.007284 | -0.76261
2019-04-09 15:00 | 1554836425.55 | 1544 | 0.20955 | 0.000443 | 2e-06 | 98.8471 | 99.1768 | 123.0941 | 20.6198 | 0.001956 | 0.001818 | -154.826 | 0.23214 | 8.3996 | 0.002753 | 0.003888 | 0.003686 | 4.6e-05 | 0.001139 | 0.30702 | 1.3408 | 2.5476 | 98.51 | 20.16 | 49.614 | 68.2683 | 87.432 | 64.9584 | 0.011905 | 1.0442 | 0.024316 | 23.2874 | -0.02978 | 0.1381 | -3.5e-05 | -0.00267 | -0.00267 | -2.0846
2019-04-09 15:00 | 1554836426.5511 | 1545 | 0.20954 | 0.000441 | 2e-06 | 98.8817 | 99.2883 | 122.7376 | 20.6202 | 0.001991 | 0.001865 | -154.815 | 0.23239 | 8.4032 | 0.002759 | 0.004223 | 0.004095 | 4.6e-05 | 0.001144 | 0.29551 | 1.3482 | 2.5476 | 98.51 | 20.16 | 49.608 | 69.1141 | 87.5799 | 66.0819 | 0.011904 | 1.0407 | 0.024241 | 23.2923 | -0.036705 | 0.13825 | 1.3e-05 | 0.000842 | 0.000842 | 6.6188
2019-04-09 15:00 | 1554836427.5508 | 1546 | 0.20956 | 0.000443 | 2e-06 | 98.9139 | 99.3406 | 122.1707 | 20.6209 | 0.001938 | 0.001908 | -154.8165 | 0.23167 | 8.4009 | 0.002772 | 0.00418 | 0.003943 | 4.5e-05 | 0.001148 | 0.29966 | 1.3682 | 2.5476 | 98.51 | 20.16 | 49.601 | 67.8395 | 87.8836 | 67.1085 | 0.011902 | 1.0376 | 0.024148 | 23.2735 | -0.038508 | 0.13782 | -8.5e-05 | -0.00631 | -0.00631 | -0.88018
2019-04-09 15:00 | 1554836428.5515 | 1547 | 0.20957 | 0.000448 | 3e-06 | 98.9663 | 99.3375 | 121.9501 | 20.6198 | 0.001904 | 0.001845 | -154.8104 | 0.23263 | 8.3999 | 0.002696 | 0.003851 | 0.00372 | 4.8e-05 | 0.001147 | 0.30055 | 1.3679 | 2.5476 | 98.51 | 20.16 | 49.587 | 67.017 | 86.0578 | 65.5944 | 0.011899 | 1.0422 | 0.02418 | 23.2007 | -0.033491 | 0.13839 | -0.000187 | -0.01371 | -0.01371 | -0.40678
2019-04-09 15:00 | 1554836429.5502 | 1548 | 0.20954 | 0.000448 | 3e-06 | 98.9963 | 99.4143 | 122.0848 | 20.6195 | 0.001872 | 0.001851 | -154.8092 | 0.23271 | 8.4028 | 0.002681 | 0.004126 | 0.003921 | 4.9e-05 | 0.001155 | 0.2959 | 1.3492 | 2.5476 | 98.51 | 20.16 | 49.579 | 66.2505 | 85.6993 | 65.7534 | 0.011897 | 1.0417 | 0.024188 | 23.219 | -0.037681 | 0.13844 | 3.2e-05 | 0.002098 | 0.002098 | 2.6598
2019-04-09 15:00 | 1554836430.5506 | 1549 | 0.20948 | 0.000446 | 2e-06 | 98.9639 | 99.2997 | 122.174 | 20.6207 | 0.001963 | 0.001851 | -154.797 | 0.23328 | 8.4019 | 0.002707 | 0.004358 | 0.004125 | 4.7e-05 | 0.001151 | 0.27599 | 1.3434 | 2.5476 | 98.51 | 20.16 | 49.573 | 68.4398 | 86.3218 | 65.7437 | 0.011895 | 1.0418 | 0.024197 | 23.2271 | -0.030297 | 0.13878 | 0.000366 | 0.026394 | 0.026394 | 0.21189
2019-04-09 15:00 | 1554836431.5517 | 1550 | 0.20949 | 0.000446 | 2e-06 | 99.0166 | 99.3079 | 122.2501 | 20.6204 | 0.001953 | 0.001901 | -154.7952 | 0.23294 | 8.4005 | 0.002738 | 0.003384 | 0.003258 | 4.5e-05 | 0.001146 | 0.30431 | 1.3665 | 2.5477 | 98.51 | 20.15 | 49.558 | 68.2073 | 87.053 | 66.9474 | 0.011885 | 1.0381 | 0.024162 | 23.2756 | -0.026278 | 0.13858 | 0.00032 | 0.022991 | 0.022991 | 0.2429
2019-04-09 15:01 | 1554836432.5517 | 1551 | 0.20952 | 0.000449 | 2e-06 | 99.0875 | 99.3303 | 122.1534 | 20.6207 | 0.001989 | 0.001846 | -154.7991 | 0.23252 | 8.4026 | 0.002683 | 0.003706 | 0.00348 | 4.4e-05 | 0.001141 | 0.28998 | 1.3726 | 2.5476 | 98.51 | 20.15 | 49.55 | 69.0521 | 85.7348 | 65.6273 | 0.011883 | 1.0421 | 0.024199 | 23.2212 | -0.021893 | 0.13833 | 0.000123 | 0.008795 | 0.008795 | 0.63385
2019-04-09 15:01 | 1554836433.5514 | 1552 | 0.20952 | 0.000445 | 2e-06 | 99.0872 | 99.4838 | 121.8907 | 20.6204 | 0.001929 | 0.001845 | -154.8055 | 0.23255 | 8.4006 | 0.002734 | 0.00362 | 0.003418 | 4.6e-05 | 0.001141 | 0.30807 | 1.3741 | 2.5477 | 98.51 | 20.15 | 49.535 | 67.623 | 86.9726 | 65.6012 | 0.011879 | 1.0422 | 0.024174 | 23.1953 | -0.035731 | 0.13834 | 0.000146 | 0.010469 | 0.010469 | 0.53253
2019-04-09 15:01 | 1554836434.5503 | 1553 | 0.20955 | 0.000444 | 2e-06 | 99.1324 | 99.3564 | 122.3001 | 20.6204 | 0.001962 | 0.001793 | -154.7928 | 0.23252 | 8.4015 | 0.002723 | 0.00441 | 0.004236 | 4.3e-05 | 0.001138 | 0.27992 | 1.3544 | 2.5475 | 98.51 | 20.15 | 49.528 | 68.4124 | 86.7111 | 64.3406 | 0.011877 | 1.0461 | 0.02426 | 23.191 | -0.02019 | 0.13833 | -3.1e-05 | -0.002403 | -0.002403 | -2.3195
2019-04-09 15:01 | 1554836435.5505 | 1554 | 0.20955 | 0.000446 | 2e-06 | 99.1122 | 99.4287 | 122.3684 | 20.6216 | 0.002006 | 0.001845 | -154.7991 | 0.23276 | 8.4017 | 0.002668 | 0.004187 | 0.00394 | 4.5e-05 | 0.001139 | 0.30502 | 1.3628 | 2.5476 | 98.51 | 20.15 | 49.52 | 69.4789 | 85.3696 | 65.6094 | 0.011875 | 1.0422 | 0.024221 | 23.241 | -0.028521 | 0.13847 | -5.6e-05 | -0.004227 | -0.004227 | -1.3202
2019-04-09 15:01 | 1554836436.5514 | 1555 | 0.20953 | 0.000446 | 2e-06 | 99.1863 | 99.6734 | 122.4151 | 20.6211 | 0.002003 | 0.001806 | -154.8127 | 0.23314 | 8.4005 | 0.002727 | 0.004631 | 0.004437 | 4.3e-05 | 0.001137 | 0.31492 | 1.3805 | 2.5475 | 98.51 | 20.15 | 49.506 | 69.3981 | 86.7921 | 64.6522 | 0.011872 | 1.0451 | 0.02426 | 23.2126 | -0.043817 | 0.1387 | 7.5e-05 | 0.005295 | 0.005295 | 1.0557
2019-04-09 15:01 | 1554836437.5505 | 1556 | 0.20952 | 0.000445 | 3e-06 | 99.2334 | 99.6242 | 122.7585 | 20.6212 | 0.001954 | 0.001868 | -154.809 | 0.23286 | 8.4015 | 0.002637 | 0.003921 | 0.003773 | 4.2e-05 | 0.001134 | 0.31133 | 1.4105 | 2.5478 | 98.51 | 20.16 | 49.499 | 68.2297 | 84.6404 | 66.164 | 0.011878 | 1.0405 | 0.02424 | 23.2971 | -0.035158 | 0.13853 | 0.000106 | 0.007526 | 0.007526 | 0.74182
2019-04-09 15:01 | 1554836438.5495 | 1557 | 0.20955 | 0.000445 | 3e-06 | 99.237 | 99.7333 | 122.6284 | 20.6205 | 0.001976 | 0.001803 | -154.8183 | 0.23289 | 8.4011 | 0.002782 | 0.003895 | 0.00372 | 4.6e-05 | 0.001131 | 0.29789 | 1.3832 | 2.5476 | 98.51 | 20.16 | 49.486 | 68.7596 | 88.1299 | 64.5991 | 0.011875 | 1.0453 | 0.024283 | 23.231 | -0.044625 | 0.13855 | -4.8e-05 | -0.003676 | -0.003676 | -1.519
2019-04-09 15:01 | 1554836439.5495 | 1558 | 0.20954 | 0.000447 | 3e-06 | 99.3078 | 99.6011 | 122.4395 | 20.6208 | 0.001924 | 0.001819 | -154.8309 | 0.23256 | 8.4049 | 0.002789 | 0.003599 | 0.003488 | 4.3e-05 | 0.001134 | 0.29391 | 1.3504 | 2.5476 | 98.51 | 20.16 | 49.48 | 67.5001 | 88.3028 | 64.9742 | 0.011873 | 1.0441 | 0.024251 | 23.2259 | -0.026379 | 0.13835 | 5e-06 | 0.000185 | 0.000185 | 30.0639
2019-04-09 15:01 | 1554836440.5509 | 1559 | 0.20952 | 0.000446 | 2e-06 | 99.32 | 99.8517 | 122.6181 | 20.6214 | 0.001976 | 0.001849 | -154.8416 | 0.23284 | 8.401 | 0.002723 | 0.003549 | 0.003316 | 4.5e-05 | 0.001133 | 0.2756 | 1.3479 | 2.5475 | 98.51 | 20.16 | 49.474 | 68.74 | 86.6926 | 65.6958 | 0.011872 | 1.0419 | 0.024243 | 23.2677 | -0.047751 | 0.13852 | 0.000133 | 0.009499 | 0.009499 | 0.58768
2019-04-09 15:01 | 1554836441.5496 | 1560 | 0.20955 | 0.000445 | 2e-06 | 99.3218 | 99.6504 | 122.5735 | 20.6235 | 0.001978 | 0.001877 | -154.8257 | 0.23167 | 8.4032 | 0.002684 | 0.00428 | 0.004103 | 4.4e-05 | 0.001133 | 0.3063 | 1.3751 | 2.5476 | 98.51 | 20.16 | 49.462 | 68.799 | 85.7688 | 66.3643 | 0.011869 | 1.0399 | 0.024214 | 23.2864 | -0.029537 | 0.13782 | -2.1e-05 | -0.001649 | -0.001649 | -3.3677
2019-04-09 15:01 | 1554836442.55 | 1561 | 0.20954 | 0.000444 | 2e-06 | 99.4418 | 99.7983 | 122.4998 | 20.6219 | 0.001962 | 0.001863 | -154.8324 | 0.23261 | 8.399 | 0.002642 | 0.003722 | 0.00349 | 4.5e-05 | 0.001138 | 0.29103 | 1.3479 | 2.5475 | 98.51 | 20.16 | 49.457 | 68.4099 | 84.7464 | 66.0453 | 0.011868 | 1.0408 | 0.024219 | 23.2684 | -0.032009 | 0.13838 | 2.7e-05 | 0.001829 | 0.001829 | 3.0491
2019-04-09 15:01 | 1554836443.5497 | 1562 | 0.20953 | 0.000445 | 2e-06 | 99.4064 | 99.926 | 122.8343 | 20.6216 | 0.001951 | 0.001813 | -154.8233 | 0.23236 | 8.4011 | 0.002693 | 0.003849 | 0.003701 | 4.4e-05 | 0.001132 | 0.29679 | 1.3682 | 2.5476 | 98.51 | 20.16 | 49.453 | 68.1491 | 85.9871 | 64.8408 | 0.011867 | 1.0445 | 0.024295 | 23.2588 | -0.046633 | 0.13823 | 4.3e-05 | 0.003003 | 0.003003 | 1.8553
2019-04-09 15:01 | 1554836444.5507 | 1563 | 0.20954 | 0.000447 | 3e-06 | 99.4433 | 99.9057 | 122.4841 | 20.6215 | 0.001942 | 0.001881 | -154.8427 | 0.23305 | 8.4013 | 0.002736 | 0.004158 | 0.003968 | 4.8e-05 | 0.001133 | 0.27848 | 1.3136 | 2.5476 | 98.51 | 20.16 | 49.445 | 67.9376 | 87.0242 | 66.4764 | 0.011865 | 1.0395 | 0.024202 | 23.2817 | -0.041494 | 0.13864 | 3.1e-05 | 0.002123 | 0.002123 | 2.6323
2019-04-09 15:01 | 1554836445.5507 | 1564 | 0.2095 | 0.000449 | 2e-06 | 99.4683 | 99.9969 | 122.2872 | 20.6216 | 0.001923 | 0.001846 | -154.8251 | 0.23284 | 8.4002 | 0.002695 | 0.003854 | 0.003735 | 4.6e-05 | 0.001134 | 0.30613 | 1.3445 | 2.5475 | 98.51 | 20.16 | 49.443 | 67.4869 | 86.0314 | 65.6173 | 0.011864 | 1.0421 | 0.024213 | 23.2336 | -0.047406 | 0.13852 | 0.000244 | 0.01757 | 0.01757 | 0.31772
2019-04-09 15:01 | 1554836446.5497 | 1565 | 0.20952 | 0.000446 | 2e-06 | 99.4666 | 99.7488 | 122.0322 | 20.6216 | 0.001948 | 0.001799 | -154.8153 | 0.23274 | 8.4018 | 0.002764 | 0.0033 | 0.003116 | 4.8e-05 | 0.001131 | 0.30071 | 1.3341 | 2.5476 | 98.51 | 20.16 | 49.44 | 68.0782 | 87.6835 | 64.5027 | 0.011864 | 1.0456 | 0.024227 | 23.1711 | -0.025343 | 0.13846 | 0.000117 | 0.008335 | 0.008335 | 0.66947
2019-04-09 15:01 | 1554836447.5498 | 1566 | 0.20951 | 0.00044 | 3e-06 | 99.4514 | 99.8314 | 121.7637 | 20.6212 | 0.001954 | 0.001888 | -154.8177 | 0.23252 | 8.4 | 0.002703 | 0.004457 | 0.00424 | 5e-05 | 0.001138 | 0.28655 | 1.3906 | 2.5476 | 98.51 | 20.15 | 49.44 | 68.2276 | 86.2114 | 66.6361 | 0.011856 | 1.039 | 0.024125 | 23.2186 | -0.034113 | 0.13833 | 0.000179 | 0.012759 | 0.012759 | 0.43691
2019-04-09 15:01 | 1554836448.5509 | 1567 | 0.20951 | 0.000446 | 2e-06 | 99.4296 | 99.9009 | 121.7318 | 20.6225 | 0.001993 | 0.001792 | -154.7995 | 0.23286 | 8.4003 | 0.002788 | 0.004148 | 0.004032 | 4.5e-05 | 0.00114 | 0.30055 | 1.3665 | 2.5476 | 98.51 | 20.15 | 49.439 | 69.1521 | 88.2756 | 64.3224 | 0.011856 | 1.0461 | 0.024204 | 23.1364 | -0.042295 | 0.13853 | 0.000202 | 0.014522 | 0.014522 | 0.38444
2019-04-09 15:01 | 1554836449.5499 | 1568 | 0.20952 | 0.000441 | 2e-06 | 99.5345 | 99.8796 | 121.9804 | 20.6217 | 0.0019 | 0.001858 | -154.7858 | 0.23299 | 8.401 | 0.002688 | 0.004641 | 0.00442 | 4.9e-05 | 0.001138 | 0.28152 | 1.3926 | 2.5477 | 98.51 | 20.15 | 49.44 | 66.9222 | 85.8715 | 65.9068 | 0.011856 | 1.0413 | 0.024172 | 23.2143 | -0.030959 | 0.1386 | 0.000113 | 0.008048 | 0.008048 | 0.69402
2019-04-09 15:01 | 1554836450.5509 | 1569 | 0.2095 | 0.000446 | 3e-06 | 99.5452 | 99.8682 | 122.0163 | 20.6214 | 0.001962 | 0.001874 | -154.7711 | 0.23231 | 8.4025 | 0.002708 | 0.004347 | 0.004162 | 4.6e-05 | 0.001137 | 0.2912 | 1.3385 | 2.5474 | 98.51 | 20.15 | 49.441 | 68.4025 | 86.349 | 66.3 | 0.011856 | 1.0401 | 0.024162 | 23.2312 | -0.028979 | 0.1382 | 0.000253 | 0.018126 | 0.018126 | 0.30727
2019-04-09 15:01 | 1554836451.5499 | 1570 | 0.20953 | 0.000443 | 3e-06 | 99.5654 | 99.8926 | 122.0395 | 20.622 | 0.001947 | 0.001892 | -154.7904 | 0.23206 | 8.4029 | 0.002767 | 0.004051 | 0.003851 | 4.5e-05 | 0.001141 | 0.29767 | 1.3352 | 2.5476 | 98.51 | 20.15 | 49.442 | 68.0605 | 87.7729 | 66.7202 | 0.011857 | 1.0388 | 0.024149 | 23.2478 | -0.029348 | 0.13805 | 8.8e-05 | 0.006215 | 0.006215 | 0.89516
2019-04-09 15:01 | 1554836452.5499 | 1571 | 0.2095 | 0.000445 | 3e-06 | 99.4931 | 99.947 | 122.1498 | 20.6228 | 0.001964 | 0.001833 | -154.7853 | 0.23323 | 8.4018 | 0.002786 | 0.003805 | 0.003648 | 4.7e-05 | 0.001146 | 0.28335 | 1.3213 | 2.5475 | 98.51 | 20.15 | 49.443 | 68.4536 | 88.225 | 65.3045 | 0.011857 | 1.0431 | 0.02421 | 23.2098 | -0.040718 | 0.13875 | 0.000232 | 0.016689 | 0.016689 | 0.33505
2019-04-09 15:01 | 1554836453.55 | 1572 | 0.20952 | 0.000448 | 3e-06 | 99.438 | 99.8758 | 122.0175 | 20.6228 | 0.001893 | 0.001813 | -154.7893 | 0.2324 | 8.4018 | 0.00277 | 0.003837 | 0.00363 | 4.5e-05 | 0.001148 | 0.2844 | 1.3121 | 2.5475 | 98.51 | 20.15 | 49.446 | 66.7547 | 87.8416 | 64.8357 | 0.011858 | 1.0446 | 0.024214 | 23.1811 | -0.039295 | 0.13825 | 0.00013 | 0.009245 | 0.009245 | 0.60265
2019-04-09 15:01 | 1554836454.5502 | 1573 | 0.20952 | 0.000445 | 3e-06 | 99.4832 | 99.8018 | 122.2522 | 20.6225 | 0.001965 | 0.001863 | -154.7785 | 0.23228 | 8.4017 | 0.002797 | 0.004369 | 0.004248 | 4.6e-05 | 0.00115 | 0.29607 | 1.3029 | 2.5475 | 98.51 | 20.15 | 49.446 | 68.4923 | 88.4889 | 66.0271 | 0.011858 | 1.0409 | 0.024195 | 23.2443 | -0.028595 | 0.13819 | 0.000137 | 0.009773 | 0.009773 | 0.56985
2019-04-09 15:01 | 1554836455.5489 | 1574 | 0.20954 | 0.000442 | 2e-06 | 99.5199 | 99.8858 | 122.536 | 20.6223 | 0.00192 | 0.001871 | -154.7822 | 0.23209 | 8.4035 | 0.002734 | 0.004168 | 0.003883 | 4.5e-05 | 0.001152 | 0.29983 | 1.2984 | 2.5476 | 98.51 | 20.15 | 49.446 | 67.4102 | 86.9735 | 66.2258 | 0.011858 | 1.0403 | 0.024216 | 23.2781 | -0.032825 | 0.13807 | 2.2e-05 | 0.001467 | 0.001467 | 3.7942
2019-04-09 15:01 | 1554836456.5501 | 1575 | 0.20952 | 0.000447 | 2e-06 | 99.5216 | 100.1243 | 123.0892 | 20.623 | 0.001985 | 0.001819 | -154.7785 | 0.23239 | 8.4017 | 0.002732 | 0.003781 | 0.003701 | 4.4e-05 | 0.001154 | 0.26952 | 1.3547 | 2.5477 | 98.51 | 20.15 | 49.445 | 68.9562 | 86.9133 | 64.9841 | 0.011857 | 1.0441 | 0.024315 | 23.2878 | -0.053996 | 0.13825 | 0.000156 | 0.011241 | 0.011241 | 0.49564
2019-04-09 15:01 | 1554836457.5501 | 1576 | 0.20952 | 0.000449 | 3e-06 | 99.5136 | 99.9573 | 123.5385 | 20.6227 | 0.00189 | 0.001778 | -154.7747 | 0.23236 | 8.3999 | 0.002672 | 0.004143 | 0.003916 | 4.5e-05 | 0.001156 | 0.30093 | 1.3286 | 2.5476 | 98.51 | 20.15 | 49.443 | 66.6875 | 85.4784 | 63.9977 | 0.011857 | 1.0472 | 0.024395 | 23.2962 | -0.039795 | 0.13823 | 0.000155 | 0.011164 | 0.011164 | 0.49897
2019-04-09 15:01 | 1554836458.5504 | 1577 | 0.2095 | 0.000446 | 2e-06 | 99.5091 | 99.8534 | 123.9105 | 20.6235 | 0.001986 | 0.001828 | -154.7814 | 0.23244 | 8.4015 | 0.002713 | 0.004134 | 0.004017 | 4.4e-05 | 0.001156 | 0.30237 | 1.3354 | 2.5475 | 98.51 | 20.15 | 49.438 | 68.9901 | 86.4626 | 65.1821 | 0.011856 | 1.0435 | 0.024389 | 23.3722 | -0.030892 | 0.13828 | 0.000234 | 0.016965 | 0.016965 | 0.32848
2019-04-09 15:01 | 1554836459.5501 | 1578 | 0.20951 | 0.000444 | 2e-06 | 99.4856 | 99.9591 | 123.8854 | 20.6238 | 0.001947 | 0.001891 | -154.7843 | 0.23321 | 8.3992 | 0.002686 | 0.003781 | 0.003581 | 4.7e-05 | 0.001159 | 0.29031 | 1.3388 | 2.5477 | 98.51 | 20.16 | 49.434 | 68.0661 | 85.8074 | 66.7018 | 0.011862 | 1.0388 | 0.024332 | 23.4223 | -0.042466 | 0.13874 | 0.000185 | 0.013345 | 0.013345 | 0.41898
2019-04-09 15:01 | 1554836460.5496 | 1579 | 0.2095 | 0.000444 | 3e-06 | 99.452 | 99.9577 | 123.6133 | 20.6237 | 0.001955 | 0.001805 | -154.7997 | 0.23259 | 8.4021 | 0.002676 | 0.004159 | 0.00393 | 4.6e-05 | 0.001162 | 0.29358 | 1.355 | 2.5475 | 98.51 | 20.16 | 49.431 | 68.2494 | 85.5768 | 64.6491 | 0.011861 | 1.0451 | 0.024379 | 23.3258 | -0.045368 | 0.13837 | 0.000277 | 0.020042 | 0.020042 | 0.27822
2019-04-09 15:01 | 1554836461.5496 | 1580 | 0.20952 | 0.000444 | 2e-06 | 99.4871 | 99.8599 | 123.5185 | 20.6231 | 0.001887 | 0.001786 | -154.7863 | 0.23241 | 8.3994 | 0.002761 | 0.003691 | 0.003632 | 4.8e-05 | 0.001161 | 0.31542 | 1.334 | 2.5475 | 98.51 | 20.16 | 49.422 | 66.6208 | 87.6318 | 64.1831 | 0.011859 | 1.0466 | 0.024386 | 23.3008 | -0.033458 | 0.13826 | 0.000111 | 0.007938 | 0.007938 | 0.70191
2019-04-09 15:01 | 1554836462.5497 | 1581 | 0.20958 | 0.000442 | 2e-06 | 99.4871 | 99.8204 | 123.354 | 20.6233 | 0.001993 | 0.001827 | -154.8102 | 0.23297 | 8.401 | 0.002705 | 0.004321 | 0.004049 | 4.6e-05 | 0.001164 | 0.27383 | 1.3605 | 2.5476 | 98.51 | 20.16 | 49.418 | 69.1646 | 86.2687 | 65.1725 | 0.011858 | 1.0435 | 0.024334 | 23.3194 | -0.029912 | 0.1386 | -0.000226 | -0.016606 | -0.016606 | -0.33634
2019-04-09 15:01 | 1554836463.5507 | 1582 | 0.20955 | 0.000449 | 2e-06 | 99.5106 | 99.7202 | 123.5777 | 20.6234 | 0.001976 | 0.001885 | -154.8062 | 0.2317 | 8.4029 | 0.002782 | 0.003814 | 0.00368 | 4.7e-05 | 0.001166 | 0.28744 | 1.3687 | 2.5476 | 98.51 | 20.16 | 49.41 | 68.7451 | 88.1303 | 66.5555 | 0.011856 | 1.0393 | 0.024307 | 23.3881 | -0.018819 | 0.13784 | -6.4e-05 | -0.004799 | -0.004799 | -1.1576
2019-04-09 15:01 | 1554836464.5503 | 1583 | 0.20956 | 0.000446 | 2e-06 | 99.48 | 100.0492 | 123.6279 | 20.6234 | 0.001918 | 0.001836 | -154.8147 | 0.23286 | 8.3993 | 0.002727 | 0.003228 | 0.003097 | 4.6e-05 | 0.001166 | 0.28943 | 1.3839 | 2.5477 | 98.51 | 20.16 | 49.406 | 67.3542 | 86.81 | 65.3816 | 0.011855 | 1.0429 | 0.024354 | 23.3524 | -0.051039 | 0.13853 | -0.000109 | -0.008061 | -0.008061 | -0.69258
2019-04-09 15:01 | 1554836465.6119 | 1584 | 0.20952 | 0.000449 | 3e-06 | 99.5181 | 99.7801 | 123.4307 | 20.6239 | 0.001852 | 0.001837 | -154.7949 | 0.23289 | 8.4029 | 0.002756 | 0.00348 | 0.003336 | 4.4e-05 | 0.001165 | 0.27671 | 1.3343 | 2.5476 | 98.51 | 20.16 | 49.402 | 65.7733 | 87.5077 | 65.4009 | 0.011854 | 1.0428 | 0.024334 | 23.3345 | -0.02352 | 0.13855 | 0.000125 | 0.008964 | 0.008964 | 0.62287
2019-04-09 15:01 | 1554836466.5552 | 1585 | 0.20954 | 0.00045 | 2e-06 | 99.5502 | 99.7997 | 123.2014 | 20.6237 | 0.001949 | 0.00186 | -154.8076 | 0.23217 | 8.4011 | 0.002622 | 0.004451 | 0.004229 | 4.6e-05 | 0.00117 | 0.28512 | 1.3382 | 2.5475 | 98.51 | 20.16 | 49.398 | 68.1034 | 84.2806 | 65.9525 | 0.011854 | 1.0411 | 0.024291 | 23.3318 | -0.022388 | 0.13812 | -1e-06 | -0.000192 | -0.000192 | -28.9452
2019-04-09 15:01 | 1554836467.5561 | 1586 |  | 0.000445 |  | 99.5041 | 99.8338 | 123.0063 | 20.6235 | 0.001864 | 0.001819 | -154.805 | 0.23158 | 8.4048 | 0.002709 | 0.003841 | 0.003593 | 4.4e-05 | 0.001172 | 0.29103 | 1.3093 | 2.5475 | 98.51 | 20.16 | 49.398 | 66.0569 | 86.3753 | 64.986 | 0.011854 | 1.0441 | 0.024307 | 23.2801 | -0.029589 | 0.13777
2019-04-09 15:01 | 1554836468.5561 | 1587 |  | 0.000445 |  | 99.5142 | 100.0217 | 123.0817 | 20.6238 | 0.001942 | 0.001815 | -154.8128 | 0.23236 | 8.404 | 0.002646 | 0.004274 | 0.004124 | 4.8e-05 | 0.001177 | 0.29983 | 1.3444 | 2.5476 | 98.51 | 20.16 | 49.4 | 67.9361 | 84.8592 | 64.8749 | 0.011854 | 1.0444 | 0.024318 | 23.2834 | -0.045499 | 0.13823
2019-04-09 15:01 | 1554836469.5561 | 1588 |  |  |  | 99.5767 | 100.0014 | 122.9418 | 20.6241 | 0.001928 | 0.001822 | -154.8254 | 0.23135 | 8.4029 | 0.002682 | 0.003664 | 0.003486 | 4.5e-05 | 0.001177 | 0.27721 | 1.4054 | 2.5476 | 98.51 | 20.16 | 49.407 | 67.6062 | 85.7311 | 65.053 | 0.011856 | 1.0439 | 0.024298 | 23.2763 | -0.038064 | 0.13763
2019-04-09 15:01 | 1554836470.5561 | 1589 |  |  |  | 99.5797 | 99.905 | 122.8977 | 20.6239 | 0.001932 | 0.001769 | -154.8116 | 0.23181 | 8.4044 | 0.002657 | 0.004013 | 0.003817 | 4.3e-05 | 0.001179 | 0.28982 | 1.3503 | 2.5475 | 98.51 | 20.16 | 49.413 | 67.7058 | 85.1291 | 63.7765 | 0.011857 | 1.0478 | 0.024339 | 23.2281 | -0.029173 | 0.13791
2019-04-09 15:01 | 1554836471.5561 | 1590 |  |  |  | 99.5738 | 100.128 | 122.8852 | 20.6249 | 0.001943 | 0.001837 | -154.8395 | 0.23186 | 8.4007 | 0.002689 | 0.003714 | 0.003508 | 4.2e-05 | 0.001182 | 0.27024 | 1.3541 | 2.5476 | 98.51 | 20.16 | 49.429 | 67.9628 | 85.8777 | 65.4161 | 0.011861 | 1.0428 | 0.024279 | 23.2834 | -0.049652 | 0.13793
2019-04-09 15:01 | 1554836472.5561 | 1591 |  |  |  | 99.6077 | 100.0596 |  | 20.6242 | 0.001893 | 0.001854 | -154.8123 | 0.23181 | 8.4016 | 0.002745 | 0.003673 | 0.003525 | 4.1e-05 | 0.001183 | 0.27688 | 1.3698 | 2.5477 | 98.51 | 20.16 | 49.439 | 66.7505 | 87.2268 | 65.8164 | 0.011863 | 1.0415 |  |  | -0.040483 | 0.13791
2019-04-09 15:01 | 1554836473.5561 | 1592 |  |  |  | 99.5886 | 99.8954 |  | 20.6246 | 0.001942 |  | -154.8253 | 0.23217 | 8.4011 | 0.002634 | 0.004324 | 0.004165 | 4e-05 | 0.001184 | 0.2844 | 1.3814 | 2.5477 | 98.51 | 20.16 | 49.455 | 67.9274 | 84.5624 |  | 0.011867 |  |  |  | -0.027508 | 0.13812
2019-04-09 15:01 | 1554836474.5561 | 1593 |  |  |  | 99.7313 | 100.1002 |  | 20.625 | 0.001895 |  |  | 0.23225 | 8.4026 | 0.002717 | 0.004522 | 0.004324 | 4.4e-05 | 0.001176 | 0.29574 | 1.3991 | 2.5475 | 98.51 | 20.16 | 49.463 | 66.8044 | 86.5656 |  | 0.011869 |  |  |  | -0.033022 | 0.13817
2019-04-09 15:01 | 1554836475.5561 | 1594 |  |  |  | 99.6682 | 100.128 |  | 20.6251 |  |  |  | 0.23249 | 8.4025 | 0.002631 | 0.004127 | 0.003977 | 4.4e-05 | 0.001174 | 0.27007 | 1.3528 | 2.5474 | 98.51 | 20.16 | 49.471 |  | 84.4912 |  | 0.011871 |  |  |  | -0.041176 | 0.13831
2019-04-09 15:01 | 1554836476.5561 | 1595 |  |  |  | 99.6625 | 100.1026 |  | 20.6247 |  |  |  | 0.23224 | 8.3994 | 0.002672 | 0.003917 | 0.003762 | 4.5e-05 | 0.001174 | 0.2891 | 1.339 | 2.5474 | 98.51 | 20.16 | 49.486 |  | 85.4703 |  | 0.011875 |  |  |  | -0.039409 | 0.13816
2019-04-09 15:01 | 1554836477.5561 | 1596 |  |  |  | 99.6845 | 100.1921 |  | 20.6253 |  |  |  | 0.23344 | 8.4008 | 0.002647 | 0.004315 | 0.004081 | 4.4e-05 | 0.001176 | 0.28943 | 1.3547 | 2.5477 | 98.51 | 20.16 | 49.495 |  | 84.8752 |  | 0.011877 |  |  |  | -0.045423 | 0.13887
2019-04-09 15:01 | 1554836478.5561 | 1597 |  |  |  | 99.7739 | 100.339 |  | 20.6251 |  |  |  | 0.23276 | 8.4011 | 0.002664 | 0.003698 | 0.003488 | 4.9e-05 | 0.001174 | 0.28998 | 1.3445 | 2.5475 | 98.51 | 20.16 | 49.501 |  | 85.2808 |  | 0.011878 |  |  |  | -0.05052 | 0.13847
2019-04-09 15:01 | 1554836479.5561

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-04-09 14:35 | 0 | 0 | 0 | 0 | 0 | 52.1 | 0.00378
2019-04-09 14:35 | 1 | 0 | 0 | 0 | 0 | 51.8202 | 0.005139
2019-04-09 14:35 | 2 | 0 | 0 | 0 | 0 | 52.2499 | -0.002035
2019-04-09 14:35 | 3 | 0 | 0 | 0 | 0 | 51.741 | 0.009405
2019-04-09 14:35 | 4 | 0 | 0 | 0 | 0 | 52.2043 | -0.020431
2019-04-09 14:35 | 5 | 0 | 0 | 0 | 0 | 52.0339 | 0.075599
2019-04-09 14:35 | 6 | 0 | 0 | 0 | 0 | 51.8461 | 0.28766
2019-04-09 14:35 | 7 | 0 | 0 | 0 | 0 | 51.7317 | 0.65639
2019-04-09 14:35 | 8 | 0 | 0 | 0 | 0 | 51.7123 | 1.1619
2019-04-09 14:35 | 9 | 0 | 0 | 0 | 0 | 51.5774 | 1.6073
2019-04-09 14:35 | 10 | 0 | 0 | 0 | 0 | 51.4878 | 2.247
2019-04-09 14:35 | 11 | 0 | 0 | 0 | 0 | 51.3569 | 2.6175
2019-04-09 14:35 | 12 | 0 | 0 | 0 | 0 | 51.2394 | 2.9772
2019-04-09 14:35 | 13 | 0 | 0 | 0 | 0 | 51.1525 | 3.297
2019-04-09 14:35 | 14 | 0 | 0 | 0 | 0 | 51.0115 | 3.2565
2019-04-09 14:35 | 15 | 0.15451 | 0 | 0 | 0 | 50.8912 | 3.1821
2019-04-09 14:35 | 16 | 0.13221 | 0 | 0 | 0 | 50.7134 | 3.1464
2019-04-09 14:35 | 17 | 0.13555 | 0 | 0 | 0 | 50.6219 | 3.1093
2019-04-09 14:35 | 18 | 0.12915 | 0 | 0 | 0 | 50.4439 | 3.0876
2019-04-09 14:35 | 19 | 0.13107 | 0 | 0 | 0 | 50.3617 | 3.0611
2019-04-09 14:35 | 20 | 0.13691 | 0 | 0 | 0 | 50.1811 | 3.0995
2019-04-09 14:35 | 21 | 0.12137 | 0 | 0 | 0 | 50.0919 | 3.2269
2019-04-09 14:35 | 22 | 0.15152 | 0 | 0 | 0 | 49.9287 | 3.6553
2019-04-09 14:35 | 23 | 0.12992 | 0 | 0 | 0 | 49.7994 | 4.011
2019-04-09 14:35 | 24 | 0.15884 | 0 | 0 | 0 | 49.6591 | 4.2217
2019-04-09 14:35 | 25 | 0.12947 | 0 | 0 | 0 | 49.4937 | 4.2773
2019-04-09 14:35 | 26 | 0.15042 | 0 | 0 | 0 | 49.3897 | 4.2575
2019-04-09 14:35 | 27 | 0.14515 | 0 | 0 | 0 | 49.198 | 4.2173
2019-04-09 14:35 | 28 | 0.14744 | 0 | 0 | 0 | 49.0984 | 4.1377
2019-04-09 14:35 | 29 | 0.14193 | 0 | 0 | 0 | 48.905 | 4.0937
2019-04-09 14:35 | 30 | 0.13572 | 0 | 0 | 0 | 48.8146 | 4.0128
2019-04-09 14:35 | 31 | 0.14423 | 0 | 0 | 0 | 48.6307 | 3.9022
2019-04-09 14:35 | 32 | 0.11967 | 0 | 0 | 0 | 48.533 | 3.8052
2019-04-09 14:35 | 33 | 0.13978 | 0 | 0 | 0 | 48.3827 | 3.6778
2019-04-09 14:35 | 34 | 0.094434 | 0 | 0 | 0 | 48.2667 | 3.5922
2019-04-09 14:35 | 35 | 0.12343 | 0 | 0 | 0 | 48.1782 | 3.4317
2019-04-09 14:35 | 36 | 0.10198 | 0 | 0 | 0 | 48.0298 | 3.3648
2019-04-09 14:35 | 37 | 0.11308 | 0 | 0 | 0 | 47.968 | 3.2081
2019-04-09 14:35 | 38 | 0.098733 | 0 | 0 | 0 | 47.8092 | 2.973
2019-04-09 14:35 | 39 | 0.083661 | 0 | 0 | 0 | 47.7707 | 2.8281
2019-04-09 14:35 | 40 | 0.094947 | 0.024154 | 2.9679 | 28.3551 | 47.6359 | 2.6922
2019-04-09 14:35 | 41 | 0.079764 | 0.027777 | 3.3672 | 32.0183 | 47.5859 | 2.5539
2019-04-09 14:35 | 42 | 0.094437 | 0.022379 | 2.7278 | 25.189 | 47.4705 | 2.3788
2019-04-09 14:35 | 43 | 0.072209 | 0.028095 | 3.3917 | 30.6015 | 47.4045 | 2.2097
2019-04-09 14:35 | 44 | 0.090331 | 0.021325 | 2.6259 | 23.9037 | 47.3182 | 2.1592
2019-04-09 14:35 | 45 | 0.060412 | 0.030819 | 3.79 | 34.1108 | 47.2334 | 2.0607
2019-04-09 14:35 | 46 | 0.087741 | 0.020239 | 2.5173 | 23.1186 | 47.1854 | 2.0284
2019-04-09 14:35 | 47 | 0.068267 | 0.025076 | 3.1236 | 29.5876 | 47.0672 | 2.0199
2019-04-09 14:35 | 48 | 0.084686 | 0.019692 | 2.4399 | 24.0113 | 47.0412 | 2.0334
2019-04-09 14:35 | 49 | 0.095414 | 0.017039 | 2.1022 | 21.4232 | 46.8989 | 2.0441
2019-04-09 14:35 | 50 | 0.082556 | 0.019438 | 2.3687 | 25.1026 | 46.8554 | 2.0724
2019-04-09 14:36 | 51 | 0.10632 | 0.014667 | 1.7959 | 19.6388 | 46.7264 | 2.0879
2019-04-09 14:36 | 52 | 0.086062 | 0.017876 | 2.1846 | 24.3493 | 46.6522 | 2.0956
2019-04-09 14:36 | 53 | 0.092913 | 0.016336 | 2.0027 | 22.6092 | 46.5492 | 2.1007
2019-04-09 14:36 | 54 | 0.079691 | 0.018833 | 2.3212 | 26.5802 | 46.47 | 2.1182
2019-04-09 14:36 | 55 | 0.10092 | 0.015136 | 1.8291 | 20.9436 | 46.3832 | 2.1137
2019-04-09 14:36 | 56 | 0.062567 | 0.024066 | 2.9371 | 33.801 | 46.2805 | 2.1148
2019-04-09 14:36 | 57 | 0.088313 | 0.01684 | 2.0798 | 23.3393 | 46.2444 | 2.0612
2019-04-09 14:36 | 58 | 0.082408 | 0.018083 | 2.2245 | 24.2339 | 46.1103 | 1.9971
2019-04-09 14:36 | 59 | 0.069571 | 0.020899 | 2.633 | 28.0751 | 46.0814 | 1.9532
2019-04-09 14:36 | 60 | 0.08168 | 0.01755 | 2.231 | 23.6659 | 45.966 | 1.933
2019-04-09 14:36 | 61 | 0.06086 | 0.02327 | 2.9778 | 30.8101 | 45.9249 | 1.8751
2019-04-09 14:36 | 62 | 0.070621 | 0.01988 | 2.5513 | 26.5857 | 45.838 | 1.8775
2019-04-09 14:36 | 63 | 0.058199 | 0.023389 | 3.0523 | 31.2385 | 45.7876 | 1.8181
2019-04-09 14:36 | 64 | 0.077092 | 0.017668 | 2.2654 | 22.8568 | 45.7154 | 1.7621
2019-04-09 14:36 | 65 | 0.051503 | 0.026091 | 3.3431 | 31.7727 | 45.6427 | 1.6364
2019-04-09 14:36 | 66 | 0.068907 | 0.019518 | 2.4489 | 22.0424 | 45.6036 | 1.5189
2019-04-09 14:36 | 67 | 0.055251 | 0.023599 | 2.9947 | 25.969 | 45.5108 | 1.4348
2019-04-09 14:36 | 68 | 0.0718 | 0.018181 | 2.2518 | 18.8314 | 45.487 | 1.3521
2019-04-09 14:36 | 69 | 0.060532 | 0.020916 | 2.611 | 21.912 | 45.3735 | 1.3264
2019-04-09 14:36 | 70 | 0.044214 | 0.027734 | 3.4842 | 28.9181 | 45.3674 | 1.2786
2019-04-09 14:36 | 71 | 0.049308 | 0.024533 | 3.0414 | 25.6219 | 45.2802 | 1.2634
2019-04-09 14:36 | 72 | 0.041518 | 0.028041 | 3.4917 | 28.6219 | 45.2705 | 1.1883
2019-04-09 14:36 | 73 | 0.067468 | 0.017029 | 2.0923 | 16.3706 | 45.1903 | 1.1045
2019-04-09 14:36 | 74 | 0.041511 | 0.026603 | 3.2746 | 25.3874 | 45.1465 | 1.0539
2019-04-09 14:36 | 75 | 0.065483 | 0.016231 | 2.0067 | 15.467 | 45.0972 | 1.0128
2019-04-09 14:36 | 76 | 0.045449 | 0.022969 | 2.7935 | 21.7829 | 45.0244 | 0.99
2019-04-09 14:36 | 77 | 0.061507 | 0.016642 | 1.9942 | 15.5769 | 44.9993 | 0.9581
2019-04-09 14:36 | 78 | 0.042237 | 0.023666 | 2.7965 | 22.4744 | 44.9085 | 0.94924
2019-04-09 14:36 | 79 | 0.056046 | 0.017445 | 2.0378 | 16.7645 | 44.9079 | 0.93958
2019-04-09 14:36 | 80 | 0.052381 | 0.018292 | 2.1159 | 17.7747 | 44.8003 | 0.93106
2019-04-09 14:36 | 81 | 0.033862 | 0.026967 | 3.1722 | 27.2151 | 44.8063 | 0.92157
2019-04-09 14:36 | 82 | 0.056533 | 0.016189 | 1.8541 | 16.3278 | 44.7237 | 0.92306
2019-04-09 14:36 | 83 | 0.036665 | 0.024338 | 2.7929 | 24.8964 | 44.702 | 0.91283
2019-04-09 14:36 | 84 | 0.047907 | 0.018639 | 2.0953 | 19.433 | 44.6445 | 0.93098
2019-04-09 14:36 | 85 | 0.023933 | 0.037446 | 4.1263 | 38.3309 | 44.6125 | 0.91738
2019-04-09 14:36 | 86 | 0.067374 | 0.013276 | 1.44 | 13.6091 | 44.5827 | 0.9169
2019-04-09 14:36 | 87 | 0.033466 | 0.026841 | 2.8723 | 27.3324 | 44.4945 | 0.91471
2019-04-09 14:36 | 88 | 0.033143 | 0.026444 | 2.8621 | 28.2092 | 44.5092 | 0.93495
2019-04-09 14:36 | 89 | 0.037712 | 0.023264 | 2.4887 | 25.088 | 44.4264 | 0.94611
2019-04-09 14:36 | 90 | 0.035502 | 0.024037 | 2.6189 | 26.8357 | 44.435 | 0.95273
2019-04-09 14:36 | 91 | 0.045806 | 0.018163 | 2.0107 | 20.572 | 44.3532 | 0.94232
2019-04-09 14:36 | 92 | 0.023218 | 0.035844 | 3.9369 | 39.6006 | 44.3499 | 0.91943
2019-04-09 14:36 | 93 | 0.051103 | 0.015887 | 1.7833 | 18.389 | 44.2958 | 0.93973
2019-04-09 14:36 | 94 | 0.044129 | 0.017862 | 2.0451 | 20.7129 | 44.2546 | 0.91405
2019-04-09 14:36 | 95 | 0.06096 | 0.012565 | 1.4719 | 14.7552 | 44.204 | 0.89947
2019-04-09 14:36 | 96 | 0.019117 | 0.041168 | 4.6776 | 47.6016 | 44.1452 | 0.91002
2019-04-09 14:36 | 97 | 0.043444 | 0.018615 | 2.0493 | 20.8443 | 44.1515 | 0.90557
2019-04-09 14:36 | 98 | 0.038818 | 0.021485 | 2.2869 | 23.7273 | 44.0636 | 0.92105
2019-04-09 14:36 | 99 | 0.039182 | 0.021773 | 2.2502 | 23.3025 | 44.0729 | 0.91304
2019-04-09 14:36 | 100 | 0.0516 | 0.016535 | 1.7043 | 17.5739 | 43.9825 | 0.90681
2019-04-09 14:36 | 101 | 0.051499 | 0.017018 | 1.6995 | 17.5827 | 43.9724 | 0.9055
2019-04-09 14:36 | 102 | 0.0599 | 0.014626 | 1.4606 | 14.932 | 43.879 | 0.89443
2019-04-09 14:36 | 103 | 0.018678 | 0.046984 | 4.6809 | 48.7302 | 43.8633 | 0.91018
2019-04-09 14:36 | 104 | 0.039344 | 0.022303 | 2.2194 | 23.0483 | 43.8285 | 0.9068
2019-04-09 14:36 | 105 | 0.017682 | 0.052122 | 4.9261 | 50.3363 | 43.7922 | 0.89006
2019-04-09 14:36 | 106 | 0.045305 | 0.020866 | 1.9247 | 19.847 | 43.7823 | 0.89917
2019-04-09 14:36 | 107 | 0.051253 | 0.018427 | 1.6936 | 17.7679 | 43.7058 | 0.91065
2019-04-09 14:36 | 108 | 0.057711 | 0.01636 | 1.4989 | 15.5357 | 43.6809 | 0.89658
2019-04-09 14:36 | 109 | 0.039457 | 0.023364 | 2.188 | 22.126 | 43.5963 | 0.87303
2019-04-09 14:36 | 110 | 0.038436 | 0.023997 | 2.2429 | 22.6295 | 43.5983 | 0.8698
2019-04-09 14:36 | 111 | 0.038312 | 0.023517 | 2.2507 | 23.2089 | 43.5187 | 0.88919
2019-04-09 14:37 | 112 | 0.034724 | 0.025294 | 2.4714 | 26.341 | 43.5219 | 0.91467
2019-04-09 14:37 | 113 | 0.055059 | 0.016007 | 1.564 | 15.6801 | 43.4445 | 0.86333
2019-04-09 14:37 | 114 | 0.02974 | 0.030285 | 2.8857 | 30.0914 | 43.4211 | 0.89492
2019-04-09 14:37 | 115 | 0.061545 | 0.014998 | 1.3984 | 14.6733 | 43.3735 | 0.90306
2019-04-09 14:37 | 116 | 0.035055 | 0.026325 | 2.4555 | 26.1815 | 43.3108 | 0.91778
2019-04-09 14:37 | 117 | 0.026718 | 0.036238 | 3.2446 | 35.0427 | 43.3003 | 0.93626
2019-04-09 14:37 | 118 | 0.017964 | 0.053915 | 4.835 | 52.7495 | 43.2559 | 0.94761
2019-04-09 14:37 | 119 | 0.049571 | 0.019558 | 1.7609 | 18.8486 | 43.255 | 0.93434
2019-04-09 14:37 | 120 | 0.04902 | 0.019791 | 1.7939 | 19.1729 | 43.1639 | 0.93987
2019-04-09 14:37 | 121 | 0.043671 | 0.022152 | 2.0174 | 21.5753 | 43.159 | 0.94221
2019-04-09 14:37 | 122 | 0.04088 | 0.023626 | 2.1626 | 22.4583 | 43.0761 | 0.9181
2019-04-09 14:37 | 123 | 0.038167 | 0.024777 | 2.3253 | 23.7683 | 43.0766 | 0.90716
2019-04-09 14:37 | 124 | 0.057896 | 0.015946 | 1.5315 | 15.6024 | 42.9954 | 0.90332
2019-04-09 14:37 | 125 | 0.025802 | 0.035743 | 3.4316 | 34.3672 | 42.9716 | 0.88674
2019-04-09 14:37 | 126 | 0.056205 | 0.016482 | 1.58 | 16.2194 | 42.9307 | 0.91161
2019-04-09 14:37 | 127 | 0.039999 | 0.023193 | 2.2154 | 22.3623 | 42.8685 | 0.89447
2019-04-09 14:37 | 128 | 0.052846 | 0.017092 | 1.6659 | 16.7484 | 42.8456 | 0.88508
2019-04-09 14:37 | 129 | 0.024759 | 0.036514 | 3.5477 | 35.8294 | 42.7713 | 0.8871
2019-04-09 14:37 | 130 | 0.037167 | 0.023746 | 2.362 | 23.33 | 42.7872 | 0.8671
2019-04-09 14:37 | 131 | 0.041435 | 0.021855 | 2.1101 | 21.0064 | 42.6984 | 0.87039
2019-04-09 14:37 | 132 | 0.028727 | 0.030709 | 3.0293 | 29.5103 | 42.7084 | 0.84774
2019-04-09 14:37 | 133 | 0.032859 | 0.026912 | 2.6463 | 25.7945 | 42.6374 | 0.84758
2019-04-09 14:37 | 134 | 0.028014 | 0.031551 | 3.09 | 30.8351 | 42.6439 | 0.86382
2019-04-09 14:37 | 135 | 0.045524 | 0.019416 | 1.8963 | 18.7861 | 42.5771 | 0.85522
2019-04-09 14:37 | 136 | 0.019308 | 0.04569 | 4.4372 | 43.9291 | 42.5617 | 0.84819
2019-04-09 14:37 | 137 | 0.053 | 0.016644 | 1.6105 | 16.4557 | 42.5259 | 0.87215
2019-04-09 14:37 | 138 | 0.035491 | 0.02484 | 2.3996 | 24.0155 | 42.4662 | 0.85233
2019-04-09 14:37 | 139 | 0.052423 | 0.016846 | 1.6214 | 16.3564 | 42.4491 | 0.85745
2019-04-09 14:37 | 140 | 0.012419 | 0.069333 | 6.8528 | 69.1287 | 42.3736 | 0.85851
2019-04-09 14:37 | 141 | 0.015774 | 0.053149 | 5.3958 | 53.3915 | 42.4143 | 0.84221
2019-04-09 14:37 | 142 | 0.03675 | 0.022848 | 2.3121 | 22.8288 | 42.3367 | 0.83895
2019-04-09 14:37 | 143 | 0.044259 | 0.018947 | 1.9118 | 18.5752 | 42.3436 | 0.82211
2019-04-09 14:37 | 144 | 0.055227 | 0.015213 | 1.5332 | 15.0832 | 42.2485 | 0.833
2019-04-09 14:37 | 145 | 0.03844 | 0.021867 | 2.1901 | 21.3477 | 42.2393 | 0.8206
2019-04-09 14:37 | 146 | 0.055265 | 0.015268 | 1.5223 | 15.1583 | 42.1655 | 0.83773
2019-04-09 14:37 | 147 | 0.013602 | 0.061897 | 6.1611 | 60.5434 | 42.1407 | 0.82353
2019-04-09 14:37 | 148 | 0.043198 | 0.020003 | 1.9328 | 19.0555 | 42.123 | 0.82315
2019-04-09 14:37 | 149 | 0.02443 | 0.034463 | 3.4133 | 33.4638 | 42.0632 | 0.81753
2019-04-09 14:37 | 150 | 0.057542 | 0.014641 | 1.4441 | 14.119 | 42.0636 | 0.81244
2019-04-09 14:37 | 151 | 0.04614 | 0.017766 | 1.7981 | 17.3011 | 41.9582 | 0.79827
2019-04-09 14:37 | 152 | 0.02731 | 0.029997 | 3.0334 | 28.7466 | 41.9739 | 0.78506
2019-04-09 14:37 | 153 | 0.029652 | 0.027694 | 2.7947 | 26.5169 | 41.8991 | 0.78627
2019-04-09 14:37 | 154 | 0.0097531 | 0.08419 | 8.4564 | 79.6422 | 41.9183 | 0.77676
2019-04-09 14:37 | 155 | 0.037033 | 0.022167 | 2.218 | 20.862 | 41.8692 | 0.77258
2019-04-09 14:37 | 156 | 0.023112 | 0.035459 | 3.5195 | 33.9249 | 41.8518 | 0.78407
2019-04-09 14:37 | 157 | 0.064384 | 0.01307 | 1.2577 | 11.9271 | 41.8122 | 0.76791
2019-04-09 14:37 | 158 | 0.031254 | 0.02698 | 2.5817 | 24.6193 | 41.7393 | 0.76945
2019-04-09 14:37 | 159 | 0.042031 | 0.020045 | 1.9091 | 18.3874 | 41.7413 | 0.77284
2019-04-09 14:37 | 160 | 0.027945 | 0.030152 | 2.8554 | 27.1944 | 41.6597 | 0.75996
2019-04-09 14:37 | 161 | 0.038193 | 0.021454 | 2.0793 | 19.9074 | 41.6799 | 0.76033
2019-04-09 14:37 | 162 | 0.047829 | 0.017183 | 1.6617 | 15.5688 | 41.5835 | 0.74464
2019-04-09 14:37 | 163 | 0.0301 | 0.027275 | 2.6238 | 24.422 | 41.5899 | 0.73511
2019-04-09 14:37 | 164 | 0.056379 | 0.014143 | 1.394 | 12.8687 | 41.514 | 0.72552
2019-04-09 14:37 | 165 | 0.041687 | 0.019674 | 1.875 | 16.9776 | 41.4854 | 0.70775
2019-04-09 14:37 | 166 | 0.06613 | 0.012068 | 1.1742 | 11.0917 | 41.4237 | 0.73349
2019-04-09 14:37 | 167 | 0.021448 | 0.037147 | 3.5932 | 32.5094 | 41.3678 | 0.69726
2019-04-09 14:37 | 168 | 0.043136 | 0.018504 | 1.7834 | 17.1229 | 41.3669 | 0.73861
2019-04-09 14:37 | 169 | 0.030424 | 0.02618 | 2.5073 | 22.3962 | 41.2878 | 0.68139
2019-04-09 14:37 | 170 | 0.03835 | 0.02015 | 1.9804 | 18.5066 | 41.3015 | 0.70974
2019-04-09 14:38 | 171 | 0.050948 | 0.014769 | 1.4887 | 13.6506 | 41.2099 | 0.69547
2019-04-09 14:38 | 172 | 0.047788 | 0.016217 | 1.5822 | 15.4471 | 41.2034 | 0.73819
2019-04-09 14:38 | 173 | 0.047566 | 0.015796 | 1.585 | 14.2986 | 41.1149 | 0.68013
2019-04-09 14:38 | 174 | 0.015805 | 0.049086 | 4.7626 | 41.9606 | 41.115 | 0.66319
2019-04-09 14:38 | 175 | 0.030791 | 0.025223 | 2.4438 | 23.4477 | 41.0732 | 0.72197
2019-04-09 14:38 | 176 | 0.011085 | 0.070118 | 6.7811 | 63.2424 | 41.0594 | 0.70105
2019-04-09 14:38 | 177 | 0.04466 | 0.017402 | 1.6767 | 15.8599 | 41.0398 | 0.7083
2019-04-09 14:38 | 178 | 0.022948 | 0.033808 | 3.2559 | 31.6791 | 40.9814 | 0.72696
2019-04-09 14:38 | 179 | 0.047481 | 0.016347 | 1.5706 | 15.0022 | 40.9846 | 0.71231
2019-04-09 14:38 | 180 | 0.045299 | 0.01716 | 1.6472 | 15.9254 | 40.8924 | 0.7214
2019-04-09 14:38 | 181 | 0.025048 | 0.031033 | 2.9747 | 28.8274 | 40.8981 | 0.72207
2019-04-09 14:38 | 182 | 0.043754 | 0.017253 | 1.7033 | 16.0717 | 40.8336 | 0.7032
2019-04-09 14:38 | 183 | 0.042547 | 0.01773 | 1.7503 | 16.0509 | 40.8141 | 0.68291
2019-04-09 14:38 | 184 | 0.063234 | 0.011927 | 1.1725 | 11.1933 | 40.7451 | 0.70779
2019-04-09 14:38 | 185 | 0.023451 | 0.033082 | 3.1489 | 29.1678 | 40.701 | 0.68401
2019-04-09 14:38 | 186 | 0.032051 | 0.024183 | 2.2953 | 21.4523 | 40.688 | 0.68756
2019-04-09 14:38 | 187 | 0.025476 | 0.030503 | 2.8774 | 27.0982 | 40.6385 | 0.69036
2019-04-09 14:38 | 188 | 0.053951 | 0.013979 | 1.3535 | 12.6403 | 40.6298 | 0.68195
2019-04-09 14:38 | 189 | 0.033023 | 0.022881 | 2.2081 | 20.0876 | 40.5415 | 0.66336
2019-04-09 14:38 | 190 | 0.035429 | 0.020648 | 2.0452 | 18.8805 | 40.5592 | 0.66892
2019-04-09 14:38 | 191 | 0.042202 | 0.016836 | 1.7131 | 15.8987 | 40.4698 | 0.67095
2019-04-09 14:38 | 192 | 0.019716 | 0.035984 | 3.6384 | 34.2048 | 40.481 | 0.6744
2019-04-09 14:38 | 193 | 0.035114 | 0.020205 | 2.033 | 19.2253 | 40.4226 | 0.67507
2019-04-09 14:38 | 194 | 0.0070105 | 0.10113 | 10.1564 | 95.7402 | 40.4186 | 0.67119
2019-04-09 14:38 | 195 | 0.05871 | 0.012117 | 1.2138 | 11.5561 | 40.3916 | 0.67845
2019-04-09 14:38 | 196 | 0.037884 | 0.019382 | 1.8782 | 17.7348 | 40.3167 | 0.67186
2019-04-09 14:38 | 197 | 0.039316 | 0.019831 | 1.8083 | 17.4299 | 40.3125 | 0.68528
2019-04-09 14:38 | 198 | 0.028497 | 0.028118 | 2.4941 | 23.736 | 40.2402 | 0.6764
2019-04-09 14:38 | 199 | 0.036764 | 0.022394 | 1.9336 | 18.5279 | 40.2511 | 0.68117
2019-04-09 14:38 | 200 | 0.037271 | 0.021488 | 1.905 | 18.1222 | 40.1679 | 0.67543
2019-04-09 14:38 | 201 | 0.043158 | 0.018599 | 1.6529 | 15.3546 | 40.1749 | 0.66267
2019-04-09 14:38 | 202 | 0.067929 | 0.011825 | 1.0504 | 9.6815 | 40.0781 | 0.65765
2019-04-09 14:38 | 203 | 0.03824 | 0.021043 | 1.8702 | 16.8566 | 40.0508 | 0.6446
2019-04-09 14:38 | 204 | 0.068348 | 0.012063 | 1.0385 | 9.1736 | 39.9897 | 0.627
2019-04-09 14:38 | 205 | 0.036194 | 0.022762 | 1.9519 | 17.0729 | 39.9266 | 0.61794
2019-04-09 14:38 | 206 | 0.064866 | 0.012679 | 1.0812 | 9.4442 | 39.9051 | 0.61261
2019-04-09 14:38 | 207 | 0.039059 | 0.021645 | 1.7893 | 15.9947 | 39.8084 | 0.62474
2019-04-09 14:38 | 208 | 0.035114 | 0.024094 | 1.9798 | 17.6053 | 39.8224 | 0.6182
2019-04-09 14:38 | 209 | 0.047891 | 0.017664 | 1.4456 | 12.8296 | 39.734 | 0.61442
2019-04-09 14:38 | 210 | 0.037701 | 0.022463 | 1.8318 | 16.1235 | 39.7311 | 0.60788
2019-04-09 14:38 | 211 | 0.060434 | 0.014013 | 1.1352 | 10.1825 | 39.6521 | 0.61537
2019-04-09 14:38 | 212 | 0.035713 | 0.023714 | 1.9142 | 16.6542 | 39.6201 | 0.59478
2019-04-09 14:38 | 213 | 0.05807 | 0.014602 | 1.1751 | 10.3283 | 39.5715 | 0.59976
2019-04-09 14:38 | 214 | 0.026424 | 0.032099 | 2.5659 | 22.5443 | 39.5147 | 0.59571
2019-04-09 14:38 | 215 | 0.057759 | 0.014694 | 1.1661 | 10.2066 | 39.5056 | 0.58952
2019-04-09 14:38 | 216 | 0.036287 | 0.023383 | 1.8459 | 16.1153 | 39.41 | 0.58478
2019-04-09 14:38 | 217 | 0.042368 | 0.020019 | 1.5771 | 13.5902 | 39.4273 | 0.57578
2019-04-09 14:38 | 218 | 0.057835 | 0.014645 | 1.1483 | 9.8617 | 39.3231 | 0.57035
2019-04-09 14:38 | 219 | 0.040251 | 0.020505 | 1.6468 | 14.3306 | 39.319 | 0.57682
2019-04-09 14:38 | 220 | 0.050927 | 0.016212 | 1.2986 | 11.072 | 39.2377 | 0.56387
2019-04-09 14:38 | 221 | 0.014045 | 0.057154 | 4.6902 | 40.6233 | 39.2267 | 0.57055
2019-04-09 14:38 | 222 | 0.0401 | 0.020573 | 1.6374 | 14.1219 | 39.1961 | 0.56629
2019-04-09 14:38 | 223 | 0.022872 | 0.036156 | 2.8608 | 25.0967 | 39.1548 | 0.574
2019-04-09 14:38 | 224 | 0.042113 | 0.020123 | 1.5512 | 13.5763 | 39.1434 | 0.57174
2019-04-09 14:38 | 225 | 0.016638 | 0.049623 | 3.9074 | 34.7323 | 39.0796 | 0.57789
2019-04-09 14:38 | 226 | 0.036303 | 0.02272 | 1.7851 | 15.4662 | 39.1004 | 0.56147
2019-04-09 14:38 | 227 | 0.040664 | 0.020308 | 1.5921 | 13.8898 | 39.0102 | 0.56482
2019-04-09 14:38 | 228 | 0.025847 | 0.032008 | 2.5149 | 21.6851 | 39.0235 | 0.56049
2019-04-09 14:38 | 229 | 0.038641 | 0.021987 | 1.6721 | 14.271 | 38.9526 | 0.55145
2019-04-09 14:38 | 230 | 0.029957 | 0.028355 | 2.1564 | 18.4486 | 38.9496 | 0.55266
2019-04-09 14:38 | 231 | 0.069096 | 0.012654 | 0.9346 | 8.2015 | 38.8828 | 0.56669
2019-04-09 14:39 | 232 | 0.046408 | 0.018806 | 1.3822 | 11.9538 | 38.8245 | 0.55476
2019-04-09 14:39 | 233 | 0.071433 | 0.012238 | 0.89746 | 7.6312 | 38.781 | 0.54512
2019-04-09 14:39 | 234 | 0.045193 | 0.018815 | 1.4107 | 11.6566 | 38.6928 | 0.5268
2019-04-09 14:39 | 235 | 0.043779 | 0.019434 | 1.4466 | 12.0833 | 38.6858 | 0.529
2019-04-09 14:39 | 236 | 0.040905 | 0.020291 | 1.5495 | 12.3922 | 38.6046 | 0.5069
2019-04-09 14:39 | 237 | 0.048988 | 0.016935 | 1.2781 | 10.4304 | 38.6012 | 0.51097
2019-04-09 14:39 | 238 | 0.057674 | 0.014362 | 1.076 | 8.7668 | 38.5071 | 0.50562
2019-04-09 14:39 | 239 | 0.030369 | 0.026558 | 2.0195 | 16.5509 | 38.4934 | 0.50263
2019-04-09 14:39 | 240 | 0.061947 | 0.013391 | 0.9862 | 8.2181 | 38.4346 | 0.50908
2019-04-09 14:39 | 241 | 0.015726 | 0.052754 | 3.8604 | 31.9758 | 38.386 | 0.50286
2019-04-09 14:39 | 242 | 0.027671 | 0.030738 | 2.1685 | 17.6004 | 38.391 | 0.48702
2019-04-09 14:39 | 243 | 0.020579 | 0.041395 | 2.9032 | 24.4975 | 38.333 | 0.50414
2019-04-09 14:39 | 244 | 0.04642 | 0.01788 | 1.2826 | 10.8788 | 38.3427 | 0.50499
2019-04-09 14:39 | 245 | 0.046003 | 0.018561 | 1.2926 | 10.6317 | 38.2467 | 0.48909
2019-04-09 14:39 | 246 | 0.016797 | 0.050695 | 3.5219 | 29.9446 | 38.2575 | 0.50296
2019-04-09 14:39 | 247 | 0.030645 | 0.027754 | 1.9179 | 16.382 | 38.2038 | 0.50203
2019-04-09 14:39 | 248 | 0.016861 | 0.05048 | 3.4841 | 29.9044 | 38.2008 | 0.50421
2019-04-09 14:39 | 249 | 0.051317 | 0.016577 | 1.1455 | 9.5524 | 38.1604 | 0.4902
2019-04-09 14:39 | 250 | 0.015176 | 0.056099 | 3.8649 | 33.5313 | 38.113 | 0.50886
2019-04-09 14:39 | 251 | 0.042052 | 0.02027 | 1.3957 | 12.1771 | 38.1174 | 0.51208
2019-04-09 14:39 | 252 | 0.021336 | 0.039972 | 2.7406 | 23.8441 | 38.0391 | 0.50875
2019-04-09 14:39 | 253 | 0.00472 | 0.18022 | 12.3493 | 103.0756 | 38.0738 | 0.48652
2019-04-09 14:39 | 254 | 0.028104 | 0.031091 | 2.0608 | 17.338 | 38.0198 | 0.48727
2019-04-09 14:39 | 255 | 0.038662 | 0.022017 | 1.5014 | 12.7294 | 38.0192 | 0.49215
2019-04-09 14:39 | 256 | 0.068312 | 0.01248 | 0.84844 | 6.974 | 37.9398 | 0.47641
2019-04-09 14:39 | 257 | 0.030554 | 0.028645 | 1.8917 | 15.7007 | 37.8981 | 0.47972
2019-04-09 14:39 | 258 | 0.033225 | 0.026279 | 1.7259 | 12.299 | 37.871 | 0.40863
2019-04-09 14:39 | 259 | 0.011617 | 0.075388 | 4.9392 | 44.7166 | 37.8355 | 0.51947
2019-04-09 14:39 | 260 | 0.037865 | 0.023095 | 1.5066 | 10.9751 | 37.8376 | 0.41557
2019-04-09 14:39 | 261 | 0.015915 | 0.056359 | 3.5735 | 31.0074 | 37.7698 | 0.49347
2019-04-09 14:39 | 262 | 0.01987 | 0.045192 | 2.8584 | 22.9009 | 37.7998 | 0.45505
2019-04-09 14:39 | 263 | 0.050422 | 0.017817 | 1.1249 | 8.774 | 37.7241 | 0.44241
2019-04-09 14:39 | 264 | 0.030362 | 0.02881 | 1.8597 | 15.4174 | 37.7104 | 0.4681
2019-04-09 14:39 | 265 | 0.03161 | 0.02842 | 1.7819 | 15.3605 | 37.6603 | 0.48554
2019-04-09 14:39 | 266 | -0.003166 | -0.2838 | -17.7497 | -145.8506 | 37.6538 | 0.46176
2019-04-09 14:39 | 267 | 0.033874 | 0.026542 | 1.6548 | 13.71 | 37.6503 | 0.46441
2019-04-09 14:39 | 268 | 0.035551 | 0.025227 | 1.5679 | 14.2389 | 37.5924 | 0.50621
2019-04-09 14:39 | 269 | 0.059728 | 0.015423 | 0.93574 | 8.1756 | 37.5757 | 0.48831
2019-04-09 14:39 | 270 | 0.040295 | 0.022852 | 1.3838 | 13.1565 | 37.4831 | 0.53015
2019-04-09 14:39 | 271 | 0.029958 | 0.030747 | 1.8573 | 14.9799 | 37.493 | 0.44877
2019-04-09 14:39 | 272 | 0.057351 | 0.016091 | 0.97319 | 8.3702 | 37.4149 | 0.48004
2019-04-09 14:39 | 273 | 0.035195 | 0.026209 | 1.5852 | 13.9852 | 37.3893 | 0.49221
2019-04-09 14:39 | 274 | 0.032082 | 0.029437 | 1.7347 | 14.8196 | 37.3414 | 0.47543
2019-04-09 14:39 | 275 | 0.014077 | 0.065542 | 3.9439 | 34.8318 | 37.3269 | 0.49033
2019-04-09 14:39 | 276 | 0.053904 | 0.017102 | 1.0304 | 9.5915 | 37.3006 | 0.51702
2019-04-09 14:39 | 277 | 0.030637 | 0.030147 | 1.8117 | 15.1397 | 37.2324 | 0.46383
2019-04-09 14:39 | 278 | 0.047304 | 0.019497 | 1.1672 | 10.3799 | 37.2321 | 0.49102
2019-04-09 14:39 | 279 | 0.025061 | 0.035865 | 2.2008 | 18.5942 | 37.1463 | 0.466
2019-04-09 14:39 | 280 | 0.0085533 | 0.10519 | 6.4758 | 55.5512 | 37.1801 | 0.47514
2019-04-09 14:39 | 281 | 0.045817 | 0.019648 | 1.2026 | 10.5726 | 37.1171 | 0.48441
2019-04-09 14:39 | 282 | 0.024249 | 0.037089 | 2.2716 | 18.002 | 37.1007 | 0.43652
2019-04-09 14:39 | 283 | 0.041715 | 0.021542 | 1.3143 | 11.3757 | 37.0613 | 0.47454
2019-04-09 14:39 | 284 | 0.020537 | 0.044896 | 2.6603 | 21.8392 | 37.025 | 0.44852
2019-04-09 14:39 | 285 | 0.041652 | 0.022131 | 1.3054 | 11.2913 | 37.0118 | 0.47031
2019-04-09 14:39 | 286 | 0.016097 | 0.057243 | 3.3809 | 28.8334 | 36.9513 | 0.46414
2019-04-09 14:39 | 287 | 0.030474 | 0.030267 | 1.7761 | 14.7446 | 36.9713 | 0.44933
2019-04-09 14:39 | 288 | 0.028456 | 0.032425 | 1.889 | 15.9888 | 36.8936 | 0.45497
2019-04-09 14:39 | 289 | 0.017464 | 0.054137 | 3.0788 | 26.4453 | 36.9158 | 0.46185
2019-04-09 14:39 | 290 | 0.051417 | 0.017944 | 1.044 | 8.7991 | 36.8493 | 0.45243
2019-04-09 14:40 | 291 | 0.023303 | 0.039546 | 2.2948 | 19.3154 | 36.8266 | 0.4501
2019-04-09 14:40 | 292 | 0.027936 | 0.033031 | 1.9117 | 16.3459 | 36.7969 | 0.45664
2019-04-09 14:40 | 293 | -0.0023148 | -0.39808 | -22.9013 | -194.6158 | 36.7767 | 0.45049
2019-04-09 14:40 | 294 | 0.042165 | 0.021883 | 1.2547 | 10.6581 | 36.7851 | 0.4494
2019-04-09 14:40 | 295 | 0.032041 | 0.029575 | 1.6499 | 14.0699 | 36.7039 | 0.45081
2019-04-09 14:40 | 296 | 0.026909 | 0.035195 | 1.9542 | 16.5194 | 36.7209 | 0.44452
2019-04-09 14:40 | 297 | 0.035461 | 0.026676 | 1.4762 | 12.1966 | 36.6474 | 0.4325
2019-04-09 14:40 | 298 | 0.016607 | 0.056977 | 3.1529 | 26.6557 | 36.6553 | 0.44266
2019-04-09 14:40 | 299 | 0.041202 | 0.02288 | 1.255 | 10.3394 | 36.6054 | 0.426
2019-04-09 14:40 | 300 | 0.0099562 | 0.09493 | 5.2318 | 44.4576 | 36.5842 | 0.44263
2019-04-09 14:40 | 301 | 0.040262 | 0.023451 | 1.2863 | 10.9536 | 36.5724 | 0.44102
2019-04-09 14:40 | 302 | 0.031119 | 0.030332 | 1.6615 | 13.9474 | 36.5114 | 0.43403
2019-04-09 14:40 | 303 | 0.044798 | 0.020593 | 1.1571 | 9.7358 | 36.5061 | 0.43615
2019-04-09 14:40 | 304 | 0.034054 | 0.027747 | 1.5261 | 12.6588 | 36.4272 | 0.43108
2019-04-09 14:40 | 305 | 0.029465 | 0.032014 | 1.756 | 14.8394 | 36.4367 | 0.43724
2019-04-09 14:40 | 306 | 0.042848 | 0.022028 | 1.2033 | 10.1915 | 36.3646 | 0.43668
2019-04-09 14:40 | 307 | 0.016724 | 0.056457 | 3.0734 | 26.2323 | 36.3591 | 0.4387
2019-04-09 14:40 | 308 | 0.04207 | 0.022461 | 1.2312 | 10.5155 | 36.3202 | 0.44239
2019-04-09 14:40 | 309 | 0.029814 | 0.031723 | 1.7341 | 14.7315 | 36.2821 | 0.43921
2019-04-09 14:40 | 310 | 0.054173 | 0.017031 | 0.94974 | 7.9204 | 36.2538 | 0.42907
2019-04-09 14:40 | 311 | 0.024511 | 0.037745 | 2.102 | 17.2626 | 36.1857 | 0.42313
2019-04-09 14:40 | 312 | 0.02401 | 0.038559 | 2.1415 | 18.2901 | 36.1989 | 0.43915
2019-04-09 14:40 | 313 | 0.026024 | 0.035477 | 1.9703 | 16.7358 | 36.1365 | 0.43554
2019-04-09 14:40 | 314 | 0.019287 | 0.047933 | 2.6473 | 22.4458 | 36.1483 | 0.43292
2019-04-09 14:40 | 315 | 0.033322 | 0.028418 | 1.5343 | 12.6047 | 36.0938 | 0.42001
2019-04-09 14:40 | 316 | 0.011681 | 0.08104 | 4.3511 | 36.8971 | 36.0887 | 0.43098
2019-04-09 14:40 | 317 | 0.047073 | 0.020095 | 1.0781 | 9.3251 | 36.059 | 0.43896
2019-04-09 14:40 | 318 | 0.010345 | 0.089389 | 4.9198 | 42.7512 | 36.0103 | 0.44227
2019-04-09 14:40 | 319 | 0.0057457 | 0.16088 | 8.8178 | 74.5196 | 36.0321 | 0.42817
2019-04-09 14:40 | 320 | -0.001544 | -0.58404 | -32.8795 | -283.7209 | 35.9977 | 0.43806
2019-04-09 14:40 | 321 | 0.02622 | 0.034409 | 1.9344 | 16.7177 | 36.0266 | 0.43834
2019-04-09 14:40 | 322 | 0.045227 | 0.019984 | 1.1195 | 9.5527 | 35.9476 | 0.43203
2019-04-09 14:40 | 323 | 0.020678 | 0.043671 | 2.4448 | 20.4952 | 35.9466 | 0.4238
2019-04-09 14:40 | 324 | 0.02839 | 0.031722 | 1.7734 | 14.9194 | 35.9001 | 0.42355
2019-04-09 14:40 | 325 | 0.0098777 | 0.091378 | 5.0904 | 43.2504 | 35.894 | 0.42722
2019-04-09 14:40 | 326 | 0.05219 | 0.017737 | 0.96261 | 8.2054 | 35.8679 | 0.42824
2019-04-09 14:40 | 327 | 0.0093167 | 0.09941 | 5.3761 | 43.4471 | 35.8076 | 0.40478
2019-04-09 14:40 | 328 | 0.031051 | 0.03053 | 1.606 | 13.3361 | 35.8354 | 0.41411
2019-04-09 14:40 | 329 | 0.047919 | 0.020256 | 1.0364 | 8.4605 | 35.746 | 0.40542
2019-04-09 14:40 | 330 | 0.029643 | 0.032835 | 1.6763 | 13.5672 | 35.7478 | 0.40218
2019-04-09 14:40 | 331 | 0.038624 | 0.025145 | 1.2782 | 10.5973 | 35.6819 | 0.40931
2019-04-09 14:40 | 332 | 0.0078218 | 0.12687 | 6.3254 | 51.9798 | 35.6781 | 0.40658
2019-04-09 14:40 | 333 | 0.044913 | 0.021611 | 1.0954 | 9.1126 | 35.6515 | 0.40927
2019-04-09 14:40 | 334 | 0.031814 | 0.031247 | 1.5467 | 12.7287 | 35.5986 | 0.40495
2019-04-09 14:40 | 335 | 0.035827 | 0.02712 | 1.3757 | 11.2025 | 35.5854 | 0.40135
2019-04-09 14:40 | 336 | 0.017879 | 0.054265 | 2.7466 | 22.2272 | 35.5311 | 0.39739
2019-04-09 14:40 | 337 | 0.037376 | 0.025884 | 1.3063 | 10.3505 | 35.5414 | 0.38686
2019-04-09 14:40 | 338 | 0.037422 | 0.025908 | 1.2981 | 10.6644 | 35.4608 | 0.39908
2019-04-09 14:40 | 339 | 0.016107 | 0.061539 | 3.0066 | 24.8277 | 35.4714 | 0.3999
2019-04-09 14:40 | 340 | 0.04005 | 0.024166 | 1.2064 | 10.1046 | 35.4199 | 0.40469
2019-04-09 14:40 | 341 | -0.016004 | -0.060654 | -3.0256 | -25.1285 | 35.4075 | 0.40216
2019-04-09 14:40 | 342 | 0.02781 | 0.035707 | 1.7343 | 14.569 | 35.4302 | 0.40516
2019-04-09 14:40 | 343 | 0.022999 | 0.043222 | 2.0963 | 18.2288 | 35.3615 | 0.41924
2019-04-09 14:40 | 344 | 0.015144 | 0.065437 | 3.1737 | 26.2648 | 35.3859 | 0.39775
2019-04-09 14:40 | 345 | 0.016292 | 0.060948 | 2.9581 | 24.1383 | 35.3294 | 0.39326
2019-04-09 14:40 | 346 | 0.0064992 | 0.15274 | 7.4298 | 62.7869 | 35.3551 | 0.40807
2019-04-09 14:40 | 347 | 0.026354 | 0.037656 | 1.833 | 15.117 | 35.3101 | 0.39838
2019-04-09 14:40 | 348 | 0.009299 | 0.10672 | 5.1796 | 43.6292 | 35.3098 | 0.40571
2019-04-09 14:40 | 349 | 0.033001 | 0.030755 | 1.4617 | 11.6782 | 35.2836 | 0.38539
2019-04-09 14:40 | 350 | 0.0032182 | 0.32162 | 14.8986 | 118.0499 | 35.2544 | 0.37991
2019-04-09 14:40 | 351 | 0.040313 | 0.025658 | 1.1903 | 10.1365 | 35.2632 | 0.40864
2019-04-09 14:41 | 352 | 0.026195 | 0.039444 | 1.8312 | 16.1949 | 35.1848 | 0.42423
2019-04-09 14:41 | 353 | 0.015501 | 0.066859 | 3.1039 | 26.3751 | 35.2104 | 0.40885
2019-04-09 14:41 | 354 | 0.035159 | 0.029462 | 1.3678 | 11.5272 | 35.1471 | 0.40529
2019-04-09 14:41 | 355 | 0.018366 | 0.056318 | 2.6165 | 22.2145 | 35.1474 | 0.408
2019-04-09 14:41 | 356 | 0.04193 | 0.025224 | 1.147 | 9.6298 | 35.1024 | 0.40378
2019-04-09 14:41 | 357 | 0.019091 | 0.055516 | 2.5225 | 21.0082 | 35.0733 | 0.40106
2019-04-09 14:41 | 358 | 0.03286 | 0.032233 | 1.4678 | 12.2296 | 35.0571 | 0.40186
2019-04-09 14:41 | 359 | 0.011002 | 0.09639 | 4.383 | 35.6648 | 35.0145 | 0.39238
2019-04-09 14:41 | 360 | 0.028898 | 0.036726 | 1.6675 | 13.6017 | 35.0263 | 0.39306
2019-04-09 14:41 | 361 | 0.038363 | 0.028266 | 1.2548 | 10.134 | 34.9584 | 0.38877
2019-04-09 14:41 | 362 | 0.021579 | 0.050123 | 2.2136 | 18.5883 | 34.9556 | 0.40112
2019-04-09 14:41 | 363 | 0.036907 | 0.02944 | 1.2963 | 10.5259 | 34.909 | 0.38848
2019-04-09 14:41 | 364 | 0.011344 | 0.095754 | 4.21 | 35.3041 | 34.8899 | 0.40049
2019-04-09 14:41 | 365 | 0.034742 | 0.030653 | 1.3743 | 11.1703 | 34.8763 | 0.38807
2019-04-09 14:41 | 366 | 0.013532 | 0.080406 | 3.5142 | 29.0157 | 34.8298 | 0.39264
2019-04-09 14:41 | 367 | 0.013751 | 0.079076 | 3.4716 | 27.6512 | 34.8447 | 0.38024
2019-04-09 14:41 | 368 | 0.020953 | 0.051944 | 2.2736 | 18.449 | 34.8001 | 0.38656
2019-04-09 14:41 | 369 | 0.021643 | 0.050251 | 2.1897 | 17.1049 | 34.8036 | 0.3702
2019-04-09 14:41 | 370 | 0.048172 | 0.022593 | 0.98151 | 7.7453 | 34.7522 | 0.3731
2019-04-09 14:41 | 371 | 0.012677 | 0.085637 | 3.7063 | 29.4758 | 34.7211 | 0.37367
2019-04-09 14:41 | 372 | 0.01992 | 0.055629 | 2.3478 | 18.2971 | 34.7182 | 0.36447
2019-04-09 14:41 | 373 | 0.0095737 | 0.11576 | 4.8744 | 38.4271 | 34.6839 | 0.36789
2019-04-09 14:41 | 374 | 0.025184 | 0.044883 | 1.8437 | 14.8603 | 34.6933 | 0.37424
2019-04-09 14:41 | 375 | 0.025206 | 0.044801 | 1.8417 | 14.5199 | 34.6367 | 0.366
2019-04-09 14:41 | 376 | 0.026492 | 0.04257 | 1.7487 | 15.3119 | 34.6427 | 0.40565
2019-04-09 14:41 | 377 | 0.038017 | 0.030239 | 1.2141 | 9.2579 | 34.582 | 0.35196
2019-04-09 14:41 | 378 | 0.015452 | 0.074595 | 2.9903 | 23.9869 | 34.5738 | 0.37063
2019-04-09 14:41 | 379 | 0.040174 | 0.028703 | 1.155 | 8.2217 | 34.5419 | 0.3303
2019-04-09 14:41 | 380 | -0.0028347 | -0.40569 | -16.2878 | -126.7504 | 34.5075 | 0.35929
2019-04-09 14:41 | 381 | 0.018194 | 0.063423 | 2.5463 | 20.7258 | 34.5334 | 0.37709
2019-04-09 14:41 | 382 | 0.032585 | 0.035985 | 1.4213 | 10.6735 | 34.4716 | 0.3478
2019-04-09 14:41 | 383 | 0.029534 | 0.039774 | 1.5638 | 12.6992 | 34.4729 | 0.37506
2019-04-09 14:41 | 384 | 0.022386 | 0.05243 | 2.0597 | 15.4413 | 34.4146 | 0.34567
2019-04-09 14:41 | 385 | -0.0054214 | -0.21224 | -8.4949 | -66.0781 | 34.4319 | 0.35824
2019-04-09 14:41 | 386 | 0.026299 | 0.044686 | 1.7501 | 14.3947 | 34.4124 | 0.37856
2019-04-09 14:41 | 387 | 0.0045602 | 0.25762 | 10.1441 | 78.3104 | 34.39 | 0.35712
2019-04-09 14:41 | 388 | 0.022654 | 0.052806 | 2.0291 | 16.3761 | 34.3953 | 0.37098
2019-04-09 14:41 | 389 | 0.01432 | 0.083687 | 3.2306 | 25.2528 | 34.35 | 0.36161
2019-04-09 14:41 | 390 | 0.020302 | 0.059967 | 2.2599 | 17.7581 | 34.3639 | 0.36053
2019-04-09 14:41 | 391 | 0.020994 | 0.059175 | 2.1895 | 16.9717 | 34.3107 | 0.3563
2019-04-09 14:41 | 392 | 0.010299 | 0.12074 | 4.4605 | 35.8555 | 34.3246 | 0.36928
2019-04-09 14:41 | 393 | 0.010665 | 0.11883 | 4.2982 | 35.5434 | 34.2874 | 0.37907
2019-04-09 14:41 | 394 | 0.011496 | 0.11019 | 4.0167 | 31.8528 | 34.3013 | 0.36618
2019-04-09 14:41 | 395 | 0.069091 | 0.018332 | 0.6636 | 5.3731 | 34.2527 | 0.37124
2019-04-09 14:41 | 396 | 0.031338 | 0.040439 | 1.4639 | 11.9551 | 34.1845 | 0.37465
2019-04-09 14:41 | 397 | 0.03775 | 0.034096 | 1.2116 | 9.9776 | 34.1821 | 0.37665
2019-04-09 14:41 | 398 | 0.025006 | 0.05146 | 1.8386 | 14.4369 | 34.1127 | 0.361
2019-04-09 14:41 | 399 | 0.016303 | 0.079131 | 2.8186 | 22.7777 | 34.1307 | 0.37135
2019-04-09 14:41 | 400 | 0.039876 | 0.032308 | 1.1437 | 9.0348 | 34.0732 | 0.36027
2019-04-09 14:41 | 401 | 0.011692 | 0.11228 | 3.9122 | 31.7259 | 34.062 | 0.37094
2019-04-09 14:41 | 402 | 0.021004 | 0.062397 | 2.1622 | 17.2419 | 34.0419 | 0.36215
2019-04-09 14:41 | 403 | 0.022616 | 0.058009 | 2.0164 | 16.0921 | 34.0207 | 0.36394
2019-04-09 14:41 | 404 | 0.038005 | 0.03456 | 1.192 | 9.9255 | 33.995 | 0.37722
2019-04-09 14:41 | 405 | -0.0036438 | -0.36615 | -12.4146 | -100.0298 | 33.9577 | 0.36449
2019-04-09 14:41 | 406 | -0.0088963 | -0.15011 | -5.0737 | -40.8064 | 33.994 | 0.36303
2019-04-09 14:41 | 407 | 0.0089834 | 0.1518 | 5.0491 | 37.5985 | 33.969 | 0.33776
2019-04-09 14:41 | 408 | 0.012084 | 0.11276 | 3.7478 | 28.1177 | 33.9787 | 0.33979
2019-04-09 14:41 | 409 | 0.026315 | 0.051836 | 1.7106 | 13.0997 | 33.9432 | 0.34472
2019-04-09 14:41 | 410 | 0.0093025 | 0.1465 | 4.8116 | 36.3323 | 33.932 | 0.33798
2019-04-09 14:42 | 411 | 0.055708 | 0.024447 | 0.79649 | 6.0021 | 33.9113 | 0.33436
2019-04-09 14:42 | 412 | 0.035172 | 0.039379 | 1.2589 | 9.4389 | 33.8353 | 0.33199
2019-04-09 14:42 | 413 | 0.034765 | 0.039764 | 1.2651 | 9.8664 | 33.8382 | 0.343
2019-04-09 14:42 | 414 | 0.016859 | 0.083458 | 2.5955 | 20.6919 | 33.7695 | 0.34885
2019-04-09 14:42 | 415 | 0.0053255 | 0.26415 | 8.1813 | 64.4308 | 33.8022 | 0.34313
2019-04-09 14:42 | 416 | 0.036199 | 0.038957 | 1.2007 | 9.6983 | 33.7495 | 0.35107
2019-04-09 14:42 | 417 | 0.010109 | 0.14371 | 4.3015 | 34.7259 | 33.7419 | 0.35104
2019-04-09 14:42 | 418 | 0.025566 | 0.056886 | 1.704 | 13.6237 | 33.7213 | 0.3483
2019-04-09 14:42 | 419 | 0.00065967 | 2.1994 | 65.7752 | 542.3209 | 33.6984 | 0.35775
2019-04-09 14:42 | 420 | 0.029553 | 0.049893 | 1.4704 | 12.1522 | 33.7085 | 0.35913
2019-04-09 14:42 | 421 | 0.022063 | 0.066972 | 1.9672 | 15.8927 | 33.6471 | 0.35064
2019-04-09 14:42 | 422 | 0.014287 | 0.10316 | 3.0303 | 24.4676 | 33.6652 | 0.34957
2019-04-09 14:42 | 423 | 0.013061 | 0.1129 | 3.3026 | 26.7315 | 33.6166 | 0.34914
2019-04-09 14:42 | 424 | 0.019859 | 0.073974 | 2.1604 | 17.5397 | 33.6365 | 0.34832
2019-04-09 14:42 | 425 | 0.043113 | 0.034506 | 0.99081 | 7.9067 | 33.5739 | 0.34088
2019-04-09 14:42 | 426 | 0.010905 | 0.13635 | 3.8989 | 30.9694 | 33.562 | 0.33771
2019-04-09 14:42 | 427 | 0.044109 | 0.034282 | 0.96728 | 7.9588 | 33.5384 | 0.35105
2019-04-09 14:42 | 428 | 0.031186 | 0.048516 | 1.3689 | 10.4965 | 33.4832 | 0.32735
2019-04-09 14:42 | 429 | 0.040645 | 0.037286 | 1.0485 | 8.3865 | 33.4726 | 0.34087
2019-04-09 14:42 | 430 | 0.018327 | 0.083835 | 2.3229 | 18.6286 | 33.4094 | 0.34141
2019-04-09 14:42 | 431 | -0.0064913 | -0.23681 | -6.567 | -52.566 | 33.4358 | 0.34122
2019-04-09 14:42 | 432 | 0.026984 | 0.05699 | 1.5712 | 12.4685 | 33.4091 | 0.33644
2019-04-09 14:42 | 433 | 0.0061913 | 0.25231 | 6.8552 | 55.4767 | 33.3928 | 0.34347
2019-04-09 14:42 | 434 | 0.023538 | 0.06626 | 1.7912 | 14.2666 | 33.3893 | 0.3358
2019-04-09 14:42 | 435 | 0.013355 | 0.11893 | 3.1637 | 24.3385 | 33.351 | 0.32503
2019-04-09 14:42 | 436 | 0.032394 | 0.049638 | 1.2975 | 10.406 | 33.3566 | 0.33709
2019-04-09 14:42 | 437 | 0.027569 | 0.058248 | 1.529 | 11.7321 | 33.2915 | 0.32344
2019-04-09 14:42 | 438 | 0.020778 | 0.078305 | 2.0115 | 16.2794 | 33.3023 | 0.33825
2019-04-09 14:42 | 439 | 0.024098 | 0.067652 | 1.7349 | 13.6768 | 33.248 | 0.32958
2019-04-09 14:42 | 440 | 0.0059418 | 0.27011 | 7.0126 | 57.0187 | 33.2578 | 0.33879
2019-04-09 14:42 | 441 | 0.04358 | 0.037447 | 0.95787 | 7.7993 | 33.2245 | 0.3399
2019-04-09 14:42 | 442 | 0.010926 | 0.14942 | 3.8188 | 30.931 | 33.1855 | 0.33794
2019-04-09 14:42 | 443 | 0.030496 | 0.053524 | 1.3633 | 10.6869 | 33.1924 | 0.32591
2019-04-09 14:42 | 444 | 0.012631 | 0.129 | 3.2969 | 25.4363 | 33.1324 | 0.32128
2019-04-09 14:42 | 445 | 0.0023338 | 0.69984 | 17.7946 | 130.469 | 33.1655 | 0.30448
2019-04-09 14:42 | 446 | 0.016498 | 0.10005 | 2.5092 | 19.1028 | 33.1221 | 0.31517
2019-04-09 14:42 | 447 | 0.00099508 | 1.6644 | 41.7405 | 310.0655 | 33.1381 | 0.30854
2019-04-09 14:42 | 448 | 0.030851 | 0.053751 | 1.3387 | 9.6915 | 33.1106 | 0.299
2019-04-09 14:42 | 449 | 0.015845 | 0.10581 | 2.5929 | 19.0453 | 33.0858 | 0.30177
2019-04-09 14:42 | 450 | 0.028522 | 0.057981 | 1.4341 | 11.2365 | 33.074 | 0.32049
2019-04-09 14:42 | 451 | 0.0090406 | 0.18557 | 4.5159 | 32.1364 | 33.0352 | 0.29053
2019-04-09 14:42 | 452 | 0.024577 | 0.068177 | 1.6562 | 12.5157 | 33.0481 | 0.30759
2019-04-09 14:42 | 453 | 0.036654 | 0.04574 | 1.1092 | 8.0314 | 32.987 | 0.29439
2019-04-09 14:42 | 454 | 0.0087868 | 0.19323 | 4.6155 | 34.5413 | 32.9831 | 0.30351
2019-04-09 14:42 | 455 | 0.048666 | 0.034895 | 0.83204 | 6.5476 | 32.9554 | 0.31864
2019-04-09 14:42 | 456 | 0.01911 | 0.09 | 2.1103 | 17.1369 | 32.9008 | 0.32748
2019-04-09 14:42 | 457 | 0.014524 | 0.1183 | 2.7715 | 22.813 | 32.9125 | 0.33133
2019-04-09 14:42 | 458 | 0.015131 | 0.11519 | 2.6565 | 21.7433 | 32.8696 | 0.329
2019-04-09 14:42 | 459 | 0.024155 | 0.073093 | 1.6625 | 12.5241 | 32.8797 | 0.30252
2019-04-09 14:42 | 460 | 0.041706 | 0.042372 | 0.96548 | 7.3104 | 32.82 | 0.30489
2019-04-09 14:42 | 461 | 0.015272 | 0.1172 | 2.625 | 19.867 | 32.806 | 0.30341
2019-04-09 14:42 | 462 | 0.033766 | 0.053671 | 1.1892 | 8.8744 | 32.7803 | 0.29966
2019-04-09 14:42 | 463 | 0.016886 | 0.10721 | 2.3691 | 18.5301 | 32.7453 | 0.31291
2019-04-09 14:42 | 464 | 0.037112 | 0.049488 | 1.0791 | 8.3394 | 32.7393 | 0.3095
2019-04-09 14:42 | 465 | 0.013476 | 0.13647 | 2.9732 | 23.0621 | 32.6805 | 0.31079
2019-04-09 14:42 | 466 | 0.010208 | 0.17977 | 3.9132 | 30.0177 | 32.7075 | 0.30641
2019-04-09 14:42 | 467 | 0.039963 | 0.045983 | 1 | 8.0844 | 32.6529 | 0.32308
2019-04-09 14:42 | 468 | 0.0082419 | 0.22273 | 4.8136 | 38.2174 | 32.641 | 0.31499
2019-04-09 14:42 | 469 | 0.022151 | 0.084042 | 1.8022 | 14.0032 | 32.6277 | 0.31018
2019-04-09 14:42 | 470 | -0.0048265 | -0.38544 | -8.2508 | -62.7962 | 32.6046 | 0.30309
2019-04-09 14:42 | 471 | 0.011615 | 0.16015 | 3.4221 | 25.73 | 32.628 | 0.29886
2019-04-09 14:43 | 472 | 0.0083963 | 0.22404 | 4.7295 | 36.0229 | 32.5855 | 0.30246
2019-04-09 14:43 | 473 | -0.008852 | -0.20976 | -4.4619 | -34.7473 | 32.614 | 0.30758
2019-04-09 14:43 | 474 | 0.014928 | 0.12586 | 2.6473 | 20.3383 | 32.5938 | 0.30362
2019-04-09 14:43 | 475 | 0.017641 | 0.10798 | 2.2363 | 17.0262 | 32.5878 | 0.30037
2019-04-09 14:43 | 476 | 0.043187 | 0.043936 | 0.90905 | 7.091 | 32.5552 | 0.30623
2019-04-09 14:43 | 477 | -0.0013525 | -1.4253 | -29.2477 | -225.8159 | 32.517 | 0.30542
2019-04-09 14:43 | 478 | 0.010454 | 0.18473 | 3.7636 | 29.1451 | 32.5457 | 0.30469
2019-04-09 14:43 | 479 | 0.015917 | 0.12129 | 2.4733 | 20.11 | 32.496 | 0.32009
2019-04-09 14:43 | 480 | 0.027596 | 0.070733 | 1.4243 | 10.3791 | 32.5123 | 0.28642
2019-04-09 14:43 | 481 | 0.047291 | 0.040831 | 0.83185 | 6.5408 | 32.4403 | 0.30932
2019-04-09 14:43 | 482 | 0.0034553 | 0.56597 | 11.3759 | 86.3486 | 32.4314 | 0.29836
2019-04-09 14:43 | 483 | 0.036495 | 0.053457 | 1.0764 | 8.6096 | 32.4169 | 0.31421
2019-04-09 14:43 | 484 | 0.032141 | 0.060719 | 1.216 | 9.1367 | 32.3662 | 0.29367
2019-04-09 14:43 | 485 | 0.016182 | 0.12063 | 2.4113 | 17.6726 | 32.3563 | 0.28597
2019-04-09 14:43 | 486 | 0.0022202 | 0.8911 | 17.6084 | 128.5228 | 32.3325 | 0.28534
2019-04-09 14:43 | 487 | 0.035357 | 0.055902 | 1.1016 | 8.3073 | 32.341 | 0.29372
2019-04-09 14:43 | 488 | 0.029027 | 0.068893 | 1.341 | 9.9407 | 32.2708 | 0.28855
2019-04-09 14:43 | 489 | 0.0047485 | 0.42024 | 8.1647 | 58.1575 | 32.2873 | 0.27616
2019-04-09 14:43 | 490 | 0.032235 | 0.062669 | 1.2028 | 8.9586 | 32.2508 | 0.28878
2019-04-09 14:43 | 491 | -0.016623 | -0.1215 | -2.3383 | -17.1596 | 32.2382 | 0.28524
2019-04-09 14:43 | 492 | 0.021058 | 0.09707 | 1.8423 | 14.8383 | 32.2652 | 0.31246
2019-04-09 14:43 | 493 | 0.025054 | 0.081033 | 1.5538 | 11.4754 | 32.203 | 0.2875
2019-04-09 14:43 | 494 | 0.0003275 | 6.1806 | 118.1848 | 851.2672 | 32.2214 | 0.27879
2019-04-09 14:43 | 495 | 0.040949 | 0.049984 | 0.94518 | 7.6875 | 32.1882 | 0.31479
2019-04-09 14:43 | 496 | 0.028924 | 0.069994 | 1.3421 | 9.8991 | 32.1509 | 0.28632
2019-04-09 14:43 | 497 | 0.021867 | 0.093566 | 1.7644 | 13.4646 | 32.1307 | 0.29443
2019-04-09 14:43 | 498 | 0.007486 | 0.27392 | 5.1605 | 37.2326 | 32.1084 | 0.27872
2019-04-09 14:43 | 499 | 0.014728 | 0.14075 | 2.6128 | 20.7638 | 32.1108 | 0.30581
2019-04-09 14:43 | 500 | 0.0046387 | 0.45211 | 8.3015 | 65.0652 | 32.0816 | 0.30182
2019-04-09 14:43 | 501 | 0.034921 | 0.060691 | 1.1024 | 8.6207 | 32.0926 | 0.30105
2019-04-09 14:43 | 502 | 0.036635 | 0.057762 | 1.0492 | 7.8397 | 32.0186 | 0.2872
2019-04-09 14:43 | 503 | 0.010703 | 0.19811 | 3.6019 | 26.2715 | 32.0258 | 0.28118
2019-04-09 14:43 | 504 | 0.036982 | 0.058404 | 1.0414 | 7.556 | 31.987 | 0.27943
2019-04-09 14:43 | 505 | -0.01487 | -0.14516 | -2.5837 | -18.7662 | 31.9682 | 0.27905
2019-04-09 14:43 | 506 | 0.0041592 | 0.51861 | 9.1875 | 64.9773 | 32.0013 | 0.27026
2019-04-09 14:43 | 507 | 0.0078502 | 0.2779 | 4.8628 | 36.9201 | 31.9618 | 0.28983
2019-04-09 14:43 | 508 | 0.011732 | 0.18641 | 3.2589 | 23.4576 | 31.9855 | 0.27521
2019-04-09 14:43 | 509 | 0.027644 | 0.079695 | 1.3789 | 10.1769 | 31.9364 | 0.28133
2019-04-09 14:43 | 510 | 0.0029242 | 0.75328 | 12.9965 | 97.0776 | 31.9387 | 0.28387
2019-04-09 14:43 | 511 | 0.033033 | 0.066795 | 1.1546 | 8.5313 | 31.9192 | 0.28181
2019-04-09 14:43 | 512 | 0.013306 | 0.16572 | 2.8485 | 21.546 | 31.8834 | 0.28668
2019-04-09 14:43 | 513 | 0.0090548 | 0.24366 | 4.1802 | 31.1059 | 31.8899 | 0.28166
2019-04-09 14:43 | 514 | 0.0067323 | 0.32684 | 5.5966 | 42.5946 | 31.8641 | 0.28676
2019-04-09 14:43 | 515 | 0.01452 | 0.15234 | 2.611 | 19.5785 | 31.8736 | 0.28429
2019-04-09 14:43 | 516 | 0.03043 | 0.07259 | 1.2383 | 9.0751 | 31.8337 | 0.27616
2019-04-09 14:43 | 517 | 0.0080582 | 0.27033 | 4.6374 | 34.4684 | 31.8208 | 0.27775
2019-04-09 14:43 | 518 | 0.037417 | 0.058195 | 1.005 | 7.3187 | 31.807 | 0.27385
2019-04-09 14:43 | 519 | 0.016478 | 0.13194 | 2.2721 | 16.5269 | 31.7566 | 0.27233
2019-04-09 14:43 | 520 | 0.024842 | 0.088751 | 1.5126 | 11.0516 | 31.7686 | 0.27455
2019-04-09 14:43 | 521 | 0.012353 | 0.1786 | 3.0409 | 22.2451 | 31.7113 | 0.2748
2019-04-09 14:43 | 522 | 0.0018612 | 1.1954 | 20.145 | 147.926 | 31.7427 | 0.27531
2019-04-09 14:43 | 523 | 0.035015 | 0.063532 | 1.0707 | 7.7686 | 31.6982 | 0.27202
2019-04-09 14:43 | 524 | 0.00043617 | 5.1473 | 85.7059 | 632.9324 | 31.6868 | 0.27606
2019-04-09 14:43 | 525 | 0.032496 | 0.068308 | 1.1556 | 8.4127 | 31.6837 | 0.27338
2019-04-09 14:43 | 526 | 0.016642 | 0.13622 | 2.2412 | 16.6887 | 31.6316 | 0.27773
2019-04-09 14:43 | 527 | 0.020573 | 0.1103 | 1.8113 | 13.2064 | 31.6467 | 0.2717
2019-04-09 14:43 | 528 | 0.020683 | 0.10978 | 1.8096 | 13.3477 | 31.5908 | 0.27607
2019-04-09 14:43 | 529 | 0.022477 | 0.10202 | 1.6636 | 12.2837 | 31.6046 | 0.2761
2019-04-09 14:43 | 530 | 0.041678 | 0.054906 | 0.89271 | 6.5405 | 31.5424 | 0.27259
2019-04-09 14:44 | 531 | 0.022502 | 0.10287 | 1.6594 | 12.1842 | 31.5293 | 0.27417
2019-04-09 14:44 | 532 | 0.049106 | 0.047132 | 0.76171 | 5.4443 | 31.4887 | 0.26734
2019-04-09 14:44 | 533 | 0.011751 | 0.19752 | 3.1799 | 23.1633 | 31.4443 | 0.27218
2019-04-09 14:44 | 534 | 0.03519 | 0.065892 | 1.0639 | 7.4181 | 31.4521 | 0.26104
2019-04-09 14:44 | 535 | 0.037756 | 0.061457 | 0.98563 | 7.2513 | 31.378 | 0.27378
2019-04-09 14:44 | 536 | 0.02373 | 0.097665 | 1.5709 | 11.4684 | 31.3805 | 0.27215
2019-04-09 14:44 | 537 | 0.040262 | 0.057485 | 0.92016 | 6.5617 | 31.3247 | 0.26419
2019-04-09 14:44 | 538 | 0.0080722 | 0.28985 | 4.577 | 32.8834 | 31.3101 | 0.26544
2019-04-09 14:44 | 539 | 0.027036 | 0.086559 | 1.3681 | 10.0764 | 31.2981 | 0.27243
2019-04-09 14:44 | 540 | 0.0026407 | 0.89412 | 13.9153 | 102.718 | 31.2649 | 0.27124
2019-04-09 14:44 | 541 | 0.0065602 | 0.36371 | 5.5847 | 42.1732 | 31.287 | 0.27667
2019-04-09 14:44 | 542 | 0.0029087 | 0.82031 | 12.5959 | 88.7465 | 31.2524 | 0.25813
2019-04-09 14:44 | 543 | 0.014824 | 0.16156 | 2.4759 | 17.9969 | 31.2773 | 0.26678
2019-04-09 14:44 | 544 | 0.030227 | 0.079072 | 1.2224 | 8.9188 | 31.2228 | 0.26959
2019-04-09 14:44 | 545 | -0.0046458 | -0.51465 | -7.8948 | -59.5095 | 31.2278 | 0.27647
2019-04-09 14:44 | 546 | 0.023619 | 0.10115 | 1.5517 | 11.4929 | 31.2188 | 0.27145
2019-04-09 14:44 | 547 | 0.0059562 | 0.401 | 6.1334 | 42.5416 | 31.1894 | 0.25339
2019-04-09 14:44 | 548 | 0.027537 | 0.086732 | 1.3292 | 10.0292 | 31.1989 | 0.27617
2019-04-09 14:44 | 549 | 0.020438 | 0.11668 | 1.7794 | 12.6696 | 31.1406 | 0.25895
2019-04-09 14:44 | 550 | 0.0115 | 0.20941 | 3.1604 | 23.2266 | 31.1587 | 0.26712
2019-04-09 14:44 | 551 | 0.026602 | 0.091513 | 1.3588 | 10.3627 | 31.1125 | 0.27567
2019-04-09 14:44 | 552 | 0.0074592 | 0.32555 | 4.8432 | 33.8328 | 31.1123 | 0.25237
2019-04-09 14:44 | 553 | 0.032014 | 0.076711 | 1.1318 | 8.4391 | 31.0889 | 0.27017
2019-04-09 14:44 | 554 | 0.0042136 | 0.58259 | 8.604 | 63.3375 | 31.0593 | 0.26688
2019-04-09 14:44 | 555 | 0.014233 | 0.17406 | 2.5546 | 18.715 | 31.0724 | 0.26638
2019-04-09 14:44 | 556 | 0.021968 | 0.11375 | 1.6526 | 11.9674 | 31.0308 | 0.2629
2019-04-09 14:44 | 557 | 0.011482 | 0.21736 | 3.1331 | 23.9333 | 31.0323 | 0.27481
2019-04-09 14:44 | 558 | 0.028935 | 0.087369 | 1.2542 | 9.5893 | 31.0021 | 0.27747
2019-04-09 14:44 | 559 | 0.017431 | 0.14479 | 2.0826 | 15.2591 | 30.9803 | 0.26598
2019-04-09 14:44 | 560 | 0.03148 | 0.079369 | 1.1474 | 8.4376 | 30.9624 | 0.26562
2019-04-09 14:44 | 561 | 0.017829 | 0.1427 | 2.0282 | 15.2197 | 30.9234 | 0.27135
2019-04-09 14:44 | 562 | 0.01471 | 0.17493 | 2.4594 | 17.9323 | 30.9248 | 0.26379
2019-04-09 14:44 | 563 | 0.0112 | 0.23122 | 3.235 | 25.7111 | 30.8936 | 0.28796
2019-04-09 14:44 | 564 | 0.01497 | 0.17322 | 2.4259 | 18.4778 | 30.9 | 0.27661
2019-04-09 14:44 | 565 | 0.041045 | 0.063006 | 0.8798 | 6.8661 | 30.8595 | 0.28182
2019-04-09 14:44 | 566 | 0.00079025 | 3.2452 | 45.6706 | 361.578 | 30.8324 | 0.28574
2019-04-09 14:44 | 567 | 0.0206 | 0.12459 | 1.7499 | 13.568 | 30.8453 | 0.2795
2019-04-09 14:44 | 568 | 0.018791 | 0.13664 | 1.9344 | 15.2558 | 30.7952 | 0.28667
2019-04-09 14:44 | 569 | 0.0094346 | 0.27216 | 3.8339 | 30.4137 | 30.8096 | 0.28694
2019-04-09 14:44 | 570 | 0.017851 | 0.14273 | 2.0306 | 14.9472 | 30.7724 | 0.26682
2019-04-09 14:44 | 571 | 0.015796 | 0.16141 | 2.2759 | 17.2293 | 30.7754 | 0.27215
2019-04-09 14:44 | 572 | 0.039982 | 0.063525 | 0.90021 | 6.8017 | 30.7361 | 0.27195
2019-04-09 14:44 | 573 | 0.0028033 | 0.9157 | 12.8488 | 93.8946 | 30.7087 | 0.26322
2019-04-09 14:44 | 574 | 0.022549 | 0.11293 | 1.5941 | 12.0802 | 30.7185 | 0.2724
2019-04-09 14:44 | 575 | 0.014618 | 0.17387 | 2.4591 | 18.574 | 30.6689 | 0.27152
2019-04-09 14:44 | 576 | 0.011835 | 0.21668 | 3.0253 | 23.0365 | 30.6882 | 0.27265
2019-04-09 14:44 | 577 | 0.029964 | 0.085283 | 1.1955 | 8.8485 | 30.6411 | 0.26514
2019-04-09 14:44 | 578 | -0.00011808 | -21.6799 | -302.3947 | -2240.3246 | 30.6381 | 0.26454
2019-04-09 14:44 | 579 | 0.034586 | 0.075336 | 1.0336 | 7.4076 | 30.6278 | 0.2562
2019-04-09 14:44 | 580 | 0.014636 | 0.17823 | 2.4452 | 18.8014 | 30.5805 | 0.27517
2019-04-09 14:44 | 581 | 0.012489 | 0.2105 | 2.8545 | 20.9797 | 30.5953 | 0.26201
2019-04-09 14:44 | 582 | 0.013556 | 0.19379 | 2.6279 | 19.7124 | 30.5539 | 0.26723
2019-04-09 14:44 | 583 | 0.024873 | 0.10652 | 1.4292 | 10.4391 | 30.5654 | 0.25965
2019-04-09 14:44 | 584 | 0.03973 | 0.06612 | 0.89393 | 6.5066 | 30.5041 | 0.25851
2019-04-09 14:44 | 585 | 0.013159 | 0.19958 | 2.6902 | 20.7127 | 30.495 | 0.27255
2019-04-09 14:44 | 586 | 0.039175 | 0.066393 | 0.9036 | 6.515 | 30.4669 | 0.25522
2019-04-09 14:44 | 587 | 0.02139 | 0.12182 | 1.653 | 10.8318 | 30.4257 | 0.2317
2019-04-09 14:44 | 588 | 0.026101 | 0.10063 | 1.3632 | 9.7598 | 30.4202 | 0.25474
2019-04-09 14:44 | 589 | 0.0078105 | 0.33383 | 4.4932 | 30.6402 | 30.378 | 0.23931
2019-04-09 14:44 | 590 | 0.016639 | 0.15794 | 2.1267 | 13.3443 | 30.3981 | 0.22203
2019-04-09 14:44 | 591 | 0.033509 | 0.079201 | 1.0531 | 7.6055 | 30.3431 | 0.25485
2019-04-09 14:45 | 592 | 0.0055106 | 0.48138 | 6.3814 | 43.5163 | 30.3409 | 0.2398
2019-04-09 14:45 | 593 | 0.025541 | 0.10456 | 1.3868 | 10.1393 | 30.3224 | 0.25897
2019-04-09 14:45 | 594 | 0.00258 | 1.0369 | 13.726 | 92.6748 | 30.2984 | 0.2391
2019-04-09 14:45 | 595 | 0.017288 | 0.15508 | 2.0487 | 13.714 | 30.309 | 0.23709
2019-04-09 14:45 | 596 | 0.023848 | 0.11263 | 1.4864 | 10.962 | 30.2652 | 0.26143
2019-04-09 14:45 | 597 | 0.016051 | 0.16634 | 2.1908 | 16.25 | 30.2643 | 0.26083
2019-04-09 14:45 | 598 | 0.040545 | 0.066038 | 0.86801 | 6.0785 | 30.2263 | 0.24645
2019-04-09 14:45 | 599 | 0.02769 | 0.096391 | 1.2657 | 9.0523 | 30.1911 | 0.25066
2019-04-09 14:45 | 600 | 0.039508 | 0.068 | 0.88808 | 5.7915 | 30.1667 | 0.22881
2019-04-09 14:45 | 601 | 0.014774 | 0.18359 | 2.3825 | 16.7124 | 30.1195 | 0.24691
2019-04-09 14:45 | 602 | 0.033075 | 0.082119 | 1.0625 | 8.0919 | 30.1274 | 0.26764
2019-04-09 14:45 | 603 | 0.048609 | 0.056258 | 0.7277 | 4.8868 | 30.0541 | 0.23754
2019-04-09 14:45 | 604 | 0.020664 | 0.13241 | 1.7042 | 11.71 | 30.0403 | 0.24197
2019-04-09 14:45 | 605 | 0.030961 | 0.088943 | 1.1422 | 7.4007 | 30.0056 | 0.22913
2019-04-09 14:45 | 606 | -0.0037375 | -0.73775 | -9.4931 | -70.7743 | 29.9873 | 0.26452
2019-04-09 14:45 | 607 | 0.013564 | 0.20451 | 2.618 | 17.7071 | 30.0018 | 0.24017
2019-04-09 14:45 | 608 | 0.0023463 | 1.194 | 15.1467 | 92.2299 | 29.9652 | 0.2164
2019-04-09 14:45 | 609 | 0.01561 | 0.17973 | 2.2663 | 18.1355 | 29.9917 | 0.2831
2019-04-09 14:45 | 610 | 0.025491 | 0.10987 | 1.384 | 9.4256 | 29.9346 | 0.24027
2019-04-09 14:45 | 611 | 0.020485 | 0.13673 | 1.7189 | 12.151 | 29.9445 | 0.24891
2019-04-09 14:45 | 612 | 0.020601 | 0.13484 | 1.7055 | 11.6195 | 29.8939 | 0.23938
2019-04-09 14:45 | 613 | -0.0217 | -0.12806 | -1.618 | -12.3819 | 29.9115 | 0.26869
2019-04-09 14:45 | 614 | 0.024163 | 0.11617 | 1.4619 | 10.311 | 29.9177 | 0.24914
2019-04-09 14:45 | 615 | 0.013321 | 0.21258 | 2.6431 | 24.4478 | 29.8743 | 0.32568
2019-04-09 14:45 | 616 | 0.029799 | 0.094839 | 1.1792 | 3.1141 | 29.8859 | 0.092797
2019-04-09 14:45 | 617 | 0.036566 | 0.077353 | 0.96855 | 2.6669 | 29.8174 | 0.097517
2019-04-09 14:45 | 618 | 0.0095117 | 0.29541 | 3.7212 | 19.5925 | 29.8206 | 0.18636
2019-04-09 14:45 | 619 | 0.021502 | 0.12962 | 1.6511 | 8.1136 | 29.7905 | 0.17446
2019-04-09 14:45 | 620 | -0.0037282 | -0.7527 | -9.5579 | -96.3141 | 29.7845 | 0.35908
2019-04-09 14:45 | 621 | 0.021919 | 0.12711 | 1.6191 | 10.3618 | 29.7864 | 0.22712
2019-04-09 14:45 | 622 | 0.030463 | 0.091541 | 1.1648 | 8.7019 | 29.7433 | 0.26509
2019-04-09 14:45 | 623 | 0.048022 | 0.058085 | 0.74134 | 5.3752 | 29.7246 | 0.25813
2019-04-09 14:45 | 624 | 0.031944 | 0.088694 | 1.1128 | 9.348 | 29.6555 | 0.29861
2019-04-09 14:45 | 625 | 0.01868 | 0.153 | 1.9062 | 13.3031 | 29.66 | 0.2485
2019-04-09 14:45 | 626 | 0.047142 | 0.06117 | 0.75235 | 6.1709 | 29.6093 | 0.29091
2019-04-09 14:45 | 627 | 0.0096683 | 0.30074 | 3.6811 | 26.2115 | 29.5792 | 0.25342
2019-04-09 14:45 | 628 | 0.040585 | 0.072302 | 0.87812 | 7.6049 | 29.5755 | 0.30864
2019-04-09 14:45 | 629 | 0.034121 | 0.086271 | 1.046 | 8.1567 | 29.5055 | 0.27831
2019-04-09 14:45 | 630 | 0.035761 | 0.083103 | 0.99887 | 7.1216 | 29.5063 | 0.25468
2019-04-09 14:45 | 631 | 0.034598 | 0.085912 | 1.0316 | 8.4861 | 29.435 | 0.2936
2019-04-09 14:45 | 632 | 0.015603 | 0.19226 | 2.3048 | 17.3245 | 29.4407 | 0.27032
2019-04-09 14:45 | 633 | 0.03328 | 0.090189 | 1.0801 | 7.9817 | 29.3956 | 0.26563
2019-04-09 14:45 | 634 | 0.023796 | 0.1258 | 1.5066 | 11.2508 | 29.3787 | 0.26772
2019-04-09 14:45 | 635 | 0.059979 | 0.050282 | 0.59647 | 4.7695 | 29.3389 | 0.28607
2019-04-09 14:45 | 636 | 0.034336 | 0.087822 | 1.047 | 7.8227 | 29.2723 | 0.2686
2019-04-09 14:45 | 637 | 0.038542 | 0.079904 | 0.93157 | 7.2017 | 29.2642 | 0.27757
2019-04-09 14:45 | 638 | 0.049891 | 0.063326 | 0.71649 | 5.7958 | 29.1932 | 0.28916
2019-04-09 14:45 | 639 | 0.04115 | 0.079184 | 0.87276 | 7.9087 | 29.1684 | 0.32544
2019-04-09 14:45 | 640 | 0.06483 | 0.051936 | 0.55815 | 7.1268 | 29.1048 | 0.46203
2019-04-09 14:45 | 641 | 0.031535 | 0.10938 | 1.1684 | 19.8922 | 29.0507 | 0.62729
2019-04-09 14:45 | 642 | 0.056775 | 0.061963 | 0.67847 | 12.287 | 29.0291 | 0.69759
2019-04-09 14:45 | 643 | 0.059907 | 0.059785 | 0.68718 | 12.7308 | 28.941 | 0.76267
2019-04-09 14:45 | 644 | 0.067104 | 0.054014 | 0.66322 | 11.6687 | 28.9088 | 0.78302
2019-04-09 14:45 | 645 | 0.07376 | 0.049424 | 0.65346 | 11.1989 | 28.8078 | 0.82603
2019-04-09 14:45 | 646 | 0.045361 | 0.080199 | 1.156 | 19.0081 | 28.7688 | 0.86222
2019-04-09 14:45 | 647 | 0.070056 | 0.052253 | 0.80752 | 13.4547 | 28.7055 | 0.94258
2019-04-09 14:45 | 648 | 0.061819 | 0.059446 | 0.97915 | 16.8384 | 28.635 | 1.0409
2019-04-09 14:45 | 649 | 0.074039 | 0.049514 | 0.87154 | 14.5042 | 28.578 | 1.0739
2019-04-09 14:45 | 650 | 0.065552 | 0.05612 | 1.0497 | 17.6368 | 28.4906 | 1.1561
2019-04-09 14:46 | 651 | 0.094585 | 0.038793 | 0.7729 | 12.7359 | 28.4391 | 1.2046
2019-04-09 14:46 | 652 | 0.087133 | 0.041984 | 0.88767 | 14.396 | 28.3099 | 1.2544
2019-04-09 14:46 | 653 | 0.068735 | 0.052902 | 1.1895 | 18.9894 | 28.2674 | 1.3052
2019-04-09 14:46 | 654 | 0.097209 | 0.037139 | 0.88679 | 13.2489 | 28.1627 | 1.2879
2019-04-09 14:46 | 655 | 0.056331 | 0.064072 | 1.6183 | 25.6361 | 28.0867 | 1.4441
2019-04-09 14:46 | 656 | 0.10003 | 0.036069 | 0.9584 | 15.474 | 28.031 | 1.5478
2019-04-09 14:46 | 657 | 0.12678 | 0.028589 | 0.79674 | 12.8415 | 27.8898 | 1.6281
2019-04-09 14:46 | 658 | 0.11388 | 0.032079 | 0.93378 | 14.8895 | 27.7875 | 1.6956
2019-04-09 14:46 | 659 | 0.11399 | 0.032171 | 0.98016 | 16.2782 | 27.6598 | 1.8555
2019-04-09 14:46 | 660 | 0.11641 | 0.031924 | 1.0288 | 17.5295 | 27.5578 | 2.0406
2019-04-09 14:46 | 661 | 0.14919 | 0.025239 | 0.86894 | 14.8108 | 27.4203 | 2.2097
2019-04-09 14:46 | 662 | 0.14083 | 0.027096 | 0.98593 | 17.0402 | 27.2685 | 2.3997
2019-04-09 14:46 | 663 | 0.15417 | 0.025205 | 0.97406 | 17.1109 | 27.1351 | 2.6379
2019-04-09 14:46 | 664 | 0.14018 | 0.028058 | 1.1544 | 20.9689 | 26.9656 | 2.9394
2019-04-09 14:46 | 665 | 0.15758 | 0.025201 | 1.108 | 18.9272 | 26.8477 | 2.9826
2019-04-09 14:46 | 666 | 0.17078 | 0.023409 | 1.1247 | 18.7756 | 26.6519 | 3.2065
2019-04-09 14:46 | 667 | 0.13659 | 0.029177 | 1.5454 | 24.9157 | 26.5164 | 3.4032
2019-04-09 14:46 | 668 | 0.16827 | 0.023682 | 1.3554 | 21.0877 | 26.3644 | 3.5485
2019-04-09 14:46 | 669 | 0.16554 | 0.024469 | 1.4633 | 22.522 | 26.1864 | 3.7282
2019-04-09 14:46 | 670 | 0.17876 | 0.02309 | 1.434 | 21.2631 | 26.0312 | 3.8009
2019-04-09 14:46 | 671 | 0.15303 | 0.027358 | 1.7198 | 25.6999 | 25.8374 | 3.9327
2019-04-09 14:46 | 672 | 0.16128 | 0.026366 | 1.6777 | 24.7528 | 25.7173 | 3.9922
2019-04-09 14:46 | 673 | 0.18318 | 0.023374 | 1.5197 | 22.6536 | 25.5118 | 4.1497
2019-04-09 14:46 | 674 | 0.15225 | 0.028445 | 1.8876 | 27.8242 | 25.362 | 4.2363
2019-04-09 14:46 | 675 | 0.18782 | 0.023158 | 1.5768 | 22.9188 | 25.193 | 4.3046
2019-04-09 14:46 | 676 | 0.18703 | 0.023303 | 1.6252 | 23.2205 | 24.9937 | 4.3429
2019-04-09 14:46 | 677 | 0.19091 | 0.022844 | 1.6361 | 22.7541 | 24.8194 | 4.3441
2019-04-09 14:46 | 678 | 0.16823 | 0.02603 | 1.902 | 25.7157 | 24.618 | 4.3262
2019-04-09 14:46 | 679 | 0.16909 | 0.02626 | 1.9358 | 25.4689 | 24.4778 | 4.3065
2019-04-09 14:46 | 680 | 0.17186 | 0.02611 | 1.9348 | 24.8584 | 24.2791 | 4.2722
2019-04-09 14:46 | 681 | 0.15551 | 0.02947 | 2.1752 | 27.3774 | 24.1378 | 4.2575
2019-04-09 14:46 | 682 | 0.17732 | 0.026497 | 1.9354 | 23.7439 | 23.9606 | 4.2103
2019-04-09 14:46 | 683 | 0.141 | 0.034152 | 2.4588 | 29.8863 | 23.7954 | 4.214
2019-04-09 14:46 | 684 | 0.15724 | 0.031292 | 2.2176 | 26.1829 | 23.6681 | 4.1169
2019-04-09 14:46 | 685 | 0.13452 | 0.037557 | 2.6028 | 29.2781 | 23.4899 | 3.9385
2019-04-09 14:46 | 686 | 0.094186 | 0.054275 | 3.7337 | 40.6737 | 23.4029 | 3.8309
2019-04-09 14:46 | 687 | 0.11047 | 0.046225 | 3.1795 | 33.5241 | 23.2882 | 3.7033
2019-04-09 14:46 | 688 | 0.10608 | 0.04787 | 3.3035 | 33.2391 | 23.1854 | 3.5261
2019-04-09 14:46 | 689 | 0.11205 | 0.044531 | 3.0909 | 29.39 | 23.0751 | 3.293
2019-04-09 14:46 | 690 | 0.075048 | 0.065283 | 4.5521 | 39.7167 | 22.9708 | 2.9806
2019-04-09 14:46 | 691 | 0.077652 | 0.061515 | 4.2415 | 36.0673 | 22.9163 | 2.8007
2019-04-09 14:46 | 692 | 0.074131 | 0.062492 | 4.1978 | 34.0732 | 22.8157 | 2.5259
2019-04-09 14:46 | 693 | 0.080638 | 0.055112 | 3.6182 | 29.3513 | 22.766 | 2.3668
2019-04-09 14:46 | 694 | 0.077316 | 0.055553 | 3.5621 | 28.6894 | 22.6576 | 2.2181
2019-04-09 14:46 | 695 | 0.034911 | 0.1186 | 7.434 | 59.2482 | 22.6193 | 2.0684
2019-04-09 14:46 | 696 | 0.074752 | 0.053122 | 3.274 | 25.7321 | 22.5697 | 1.9235
2019-04-09 14:46 | 697 | 0.056393 | 0.067585 | 4.0744 | 31.8459 | 22.4816 | 1.7959
2019-04-09 14:46 | 698 | 0.060237 | 0.061274 | 3.601 | 26.8075 | 22.453 | 1.6148
2019-04-09 14:46 | 699 | 0.054382 | 0.065579 | 3.7734 | 28.0055 | 22.3629 | 1.523
2019-04-09 14:46 | 700 | 0.049823 | 0.069061 | 3.892 | 29.1571 | 22.3434 | 1.4527
2019-04-09 14:46 | 701 | 0.078285 | 0.042643 | 2.3445 | 17.8229 | 22.2567 | 1.3953
2019-04-09 14:46 | 702 | 0.039541 | 0.080913 | 4.3934 | 34.4957 | 22.2016 | 1.364
2019-04-09 14:46 | 703 | 0.045838 | 0.067486 | 3.6379 | 27.7783 | 22.1681 | 1.2733
2019-04-09 14:46 | 704 | 0.050207 | 0.059624 | 3.1863 | 24.0899 | 22.1087 | 1.2095
2019-04-09 14:46 | 705 | 0.075535 | 0.03853 | 2.0315 | 15.4501 | 22.0634 | 1.167
2019-04-09 14:46 | 706 | 0.062854 | 0.045346 | 2.3413 | 18.0038 | 21.9667 | 1.1316
2019-04-09 14:46 | 707 | 0.048426 | 0.057937 | 2.9203 | 23.7899 | 21.9387 | 1.152
2019-04-09 14:46 | 708 | 0.054347 | 0.051634 | 2.5023 | 21.0459 | 21.8654 | 1.1438
2019-04-09 14:46 | 709 | 0.035718 | 0.078119 | 3.6865 | 32.7877 | 21.8346 | 1.1711
2019-04-09 14:46 | 710 | 0.056466 | 0.049143 | 2.2671 | 20.867 | 21.7858 | 1.1783
2019-04-09 14:46 | 711 | 0.03418 | 0.081247 | 3.66 | 33.9123 | 21.7299 | 1.1591
2019-04-09 14:47 | 712 | 0.066858 | 0.041238 | 1.8317 | 16.7494 | 21.7057 | 1.1198
2019-04-09 14:47 | 713 | 0.065843 | 0.041644 | 1.8297 | 17.4708 | 21.6038 | 1.1503
2019-04-09 14:47 | 714 | 0.044147 | 0.060631 | 2.6862 | 25.3045 | 21.5791 | 1.1171
2019-04-09 14:47 | 715 | 0.058967 | 0.045119 | 1.9818 | 19.1619 | 21.5077 | 1.1299
2019-04-09 14:47 | 716 | 0.044466 | 0.05896 | 2.5989 | 25.3183 | 21.4666 | 1.1258
2019-04-09 14:47 | 717 | 0.067672 | 0.038058 | 1.6806 | 15.5961 | 21.4112 | 1.0554
2019-04-09 14:47 | 718 | 0.046269 | 0.055543 | 2.4211 | 23.0022 | 21.3405 | 1.0643
2019-04-09 14:47 | 719 | 0.061148 | 0.04131 | 1.806 | 17.2988 | 21.3111 | 1.0578
2019-04-09 14:47 | 720 | 0.061824 | 0.040156 | 1.7598 | 17.2068 | 21.2212 | 1.0638
2019-04-09 14:47 | 721 | 0.047273 | 0.051641 | 2.2729 | 22.2962 | 21.1909 | 1.054
2019-04-09 14:47 | 722 | 0.063299 | 0.038262 | 1.6809 | 16.518 | 21.1206 | 1.0456
2019-04-09 14:47 | 723 | 0.033849 | 0.070825 | 3.1193 | 31.3377 | 21.0737 | 1.0608
2019-04-09 14:47 | 724 | 0.050781 | 0.046831 | 2.0679 | 23.603 | 21.043 | 1.1986
2019-04-09 14:47 | 725 | 0.045961 | 0.051589 | 2.2559 | 22.9133 | 20.9761 | 1.0531
2019-04-09 14:47 | 726 | 0.051135 | 0.04639 | 2.0186 | 21.3235 | 20.9492 | 1.0904
2019-04-09 14:47 | 727 | 0.051518 | 0.046078 | 1.9965 | 21.2642 | 20.8748 | 1.0955
2019-04-09 14:47 | 728 | 0.050625 | 0.046583 | 2.0308 | 21.2295 | 20.8463 | 1.0747
2019-04-09 14:47 | 729 | 0.065248 | 0.036159 | 1.5774 | 16.0706 | 20.7706 | 1.0486
2019-04-09 14:47 | 730 | 0.045092 | 0.052303 | 2.2833 | 25.4469 | 20.723 | 1.1475
2019-04-09 14:47 | 731 | 0.069639 | 0.033553 | 1.4759 | 14.9264 | 20.6719 | 1.0395
2019-04-09 14:47 | 732 | 0.024079 | 0.096424 | 4.2866 | 47.8226 | 20.5984 | 1.1515
2019-04-09 14:47 | 733 | 0.040813 | 0.05674 | 2.5157 | 27.334 | 20.6095 | 1.1156
2019-04-09 14:47 | 734 | 0.073337 | 0.031325 | 1.4034 | 13.8594 | 20.5127 | 1.0164
2019-04-09 14:47 | 735 | 0.048784 | 0.047078 | 2.1049 | 21.3474 | 20.4758 | 1.0414
2019-04-09 14:47 | 736 | 0.059169 | 0.038615 | 1.7436 | 17.6826 | 20.4088 | 1.0463
2019-04-09 14:47 | 737 | 0.037234 | 0.060601 | 2.7544 | 32.3832 | 20.3637 | 1.2058
2019-04-09 14:47 | 738 | 0.055841 | 0.04041 | 1.8367 | 21.7885 | 20.3257 | 1.2167
2019-04-09 14:47 | 739 | 0.043543 | 0.051393 | 2.364 | 25.2096 | 20.2579 | 1.0977
2019-04-09 14:47 | 740 | 0.059943 | 0.037307 | 1.7167 | 17.564 | 20.2326 | 1.0528
2019-04-09 14:47 | 741 | 0.060599 | 0.03684 | 1.6974 | 17.9092 | 20.1415 | 1.0853
2019-04-09 14:47 | 742 | 0.052465 | 0.042577 | 1.963 | 20.0833 | 20.1134 | 1.0537
2019-04-09 14:47 | 743 | 0.073374 | 0.03042 | 1.4044 | 13.6764 | 20.031 | 1.0035
2019-04-09 14:47 | 744 | 0.03794 | 0.058343 | 2.7226 | 28.3407 | 19.9788 | 1.0752
2019-04-09 14:47 | 745 | 0.047833 | 0.045906 | 2.1677 | 26.1435 | 19.9457 | 1.2505
2019-04-09 14:47 | 746 | 0.039528 | 0.055537 | 2.6233 | 29.3931 | 19.8857 | 1.1618
2019-04-09 14:47 | 747 | 0.062381 | 0.034529 | 1.6651 | 19.2276 | 19.8598 | 1.1994
2019-04-09 14:47 | 748 | 0.062478 | 0.034434 | 1.6599 | 16.9389 | 19.7664 | 1.0583
2019-04-09 14:47 | 749 | 0.041819 | 0.050997 | 2.4868 | 28.7183 | 19.7397 | 1.201
2019-04-09 14:47 | 750 | 0.060069 | 0.035413 | 1.7251 | 17.4346 | 19.6746 | 1.0473
2019-04-09 14:47 | 751 | 0.036542 | 0.057708 | 2.842 | 29.231 | 19.6276 | 1.0682
2019-04-09 14:47 | 752 | 0.065442 | 0.032556 | 1.5854 | 17.8009 | 19.5905 | 1.1649
2019-04-09 14:47 | 753 | 0.054201 | 0.038906 | 1.9136 | 19.9179 | 19.5051 | 1.0796
2019-04-09 14:47 | 754 | 0.056653 | 0.036745 | 1.8309 | 19.0908 | 19.4795 | 1.0816
2019-04-09 14:47 | 755 | 0.068034 | 0.030361 | 1.5302 | 17.3177 | 19.3901 | 1.1782
2019-04-09 14:47 | 756 | 0.048849 | 0.041806 | 2.1337 | 24.1914 | 19.3503 | 1.1817
2019-04-09 14:47 | 757 | 0.052569 | 0.038851 | 1.9848 | 19.9628 | 19.288 | 1.0494
2019-04-09 14:47 | 758 | 0.034483 | 0.058626 | 3.0342 | 27.4754 | 19.249 | 0.94745
2019-04-09 14:47 | 759 | 0.056488 | 0.035771 | 1.8538 | 21.3543 | 19.2105 | 1.2063
2019-04-09 14:47 | 760 | 0.038746 | 0.051638 | 2.7148 | 32.1484 | 19.1443 | 1.2456
2019-04-09 14:47 | 761 | 0.042619 | 0.046377 | 2.4723 | 25.3907 | 19.1283 | 1.0821
2019-04-09 14:47 | 762 | 0.05838 | 0.033878 | 1.8136 | 19.4969 | 19.0563 | 1.1382
2019-04-09 14:47 | 763 | 0.044982 | 0.043462 | 2.357 | 28.2924 | 19.0178 | 1.2727
2019-04-09 14:47 | 764 | 0.059547 | 0.032774 | 1.786 | 19.8026 | 18.961 | 1.1792
2019-04-09 14:47 | 765 | 0.037402 | 0.052178 | 2.8507 | 30.711 | 18.906 | 1.1486
2019-04-09 14:47 | 766 | 0.061763 | 0.031675 | 1.7294 | 18.5677 | 18.8763 | 1.1468
2019-04-09 14:47 | 767 | 0.056368 | 0.034727 | 1.8984 | 21.4047 | 18.7887 | 1.2065
2019-04-09 14:47 | 768 | 0.05036 | 0.038393 | 2.1238 | 23.7152 | 18.764 | 1.1943
2019-04-09 14:47 | 769 | 0.060113 | 0.03216 | 1.7796 | 19.6689 | 18.6846 | 1.1824
2019-04-09 14:47 | 770 | 0.048797 | 0.03981 | 2.2113 | 24.267 | 18.6477 | 1.1841
2019-04-09 14:48 | 771 | 0.077057 | 0.024822 | 1.3971 | 15.6721 | 18.579 | 1.2076
2019-04-09 14:48 | 772 | 0.051751 | 0.036982 | 2.0902 | 23.1785 | 18.5049 | 1.1995
2019-04-09 14:48 | 773 | 0.068608 | 0.027841 | 1.5765 | 17.502 | 18.4669 | 1.2008
2019-04-09 14:48 | 774 | 0.063768 | 0.029592 | 1.7008 | 18.407 | 18.372 | 1.1738
2019-04-09 14:48 | 775 | 0.061207 | 0.030835 | 1.7784 | 19.5265 | 18.3391 | 1.1952
2019-04-09 14:48 | 776 | 0.071147 | 0.026875 | 1.5343 | 18.0038 | 18.2472 | 1.2809
2019-04-09 14:48 | 777 | 0.051375 | 0.037093 | 2.1282 | 24.2635 | 18.203 | 1.2465
2019-04-09 14:48 | 778 | 0.065124 | 0.029249 | 1.6823 | 19.1857 | 18.1375 | 1.2494
2019-04-09 14:48 | 779 | 0.050812 | 0.037472 | 2.164 | 24.8033 | 18.0782 | 1.2603
2019-04-09 14:48 | 780 | 0.066898 | 0.028783 | 1.6532 | 18.4636 | 18.0297 | 1.2352
2019-04-09 14:48 | 781 | 0.056261 | 0.034192 | 1.9717 | 22.5457 | 17.9501 | 1.2684
2019-04-09 14:48 | 782 | 0.052405 | 0.036762 | 2.131 | 24.31 | 17.9158 | 1.274
2019-04-09 14:48 | 783 | 0.060995 | 0.031545 | 1.8397 | 21.1448 | 17.8426 | 1.2897
2019-04-09 14:48 | 784 | 0.045679 | 0.042043 | 2.4646 | 28.1181 | 17.7986 | 1.2844
2019-04-09 14:48 | 785 | 0.065275 | 0.029421 | 1.7335 | 19.4271 | 17.7442 | 1.2681
2019-04-09 14:48 | 786 | 0.041581 | 0.046298 | 2.7409 | 31.8991 | 17.6769 | 1.3264
2019-04-09 14:48 | 787 | 0.056767 | 0.033376 | 2.0103 | 22.6439 | 17.6522 | 1.2854
2019-04-09 14:48 | 788 | 0.078376 | 0.024208 | 1.4641 | 16.9396 | 17.5616 | 1.3277
2019-04-09 14:48 | 789 | 0.068528 | 0.02772 | 1.6845 | 19.3483 | 17.5028 | 1.3259
2019-04-09 14:48 | 790 | 0.076331 | 0.024578 | 1.5152 | 17.287 | 17.422 | 1.3195
2019-04-09 14:48 | 791 | 0.047499 | 0.039454 | 2.4451 | 28.0571 | 17.358 | 1.3327
2019-04-09 14:48 | 792 | 0.057481 | 0.032259 | 2.0342 | 22.8931 | 17.3187 | 1.3159
2019-04-09 14:48 | 793 | 0.045757 | 0.040993 | 2.5641 | 29.3406 | 17.2469 | 1.3425
2019-04-09 14:48 | 794 | 0.050558 | 0.036685 | 2.3287 | 26.5328 | 17.2235 | 1.3415
2019-04-09 14:48 | 795 | 0.05983 | 0.031007 | 1.9745 | 22.0449 | 17.1448 | 1.319
2019-04-09 14:48 | 796 | 0.046741 | 0.039705 | 2.5349 | 28.6931 | 17.1085 | 1.3411
2019-04-09 14:48 | 797 | 0.067246 | 0.027586 | 1.7668 | 20.2199 | 17.0449 | 1.3597
2019-04-09 14:48 | 798 | 0.035351 | 0.05254 | 3.3736 | 39.1441 | 16.9849 | 1.3838
2019-04-09 14:48 | 799 | 0.055965 | 0.033172 | 2.1342 | 24.6009 | 16.9629 | 1.3768
2019-04-09 14:48 | 800 | 0.055858 | 0.033153 | 2.1408 | 24.6416 | 16.8768 | 1.3764
2019-04-09 14:48 | 801 | 0.051454 | 0.036029 | 2.3371 | 27.1676 | 16.8524 | 1.3979
2019-04-09 14:48 | 802 | 0.062133 | 0.029826 | 1.944 | 22.6454 | 16.7712 | 1.407
2019-04-09 14:48 | 803 | 0.038545 | 0.048065 | 3.1477 | 36.1011 | 16.7348 | 1.3915
2019-04-09 14:48 | 804 | 0.071254 | 0.026372 | 1.7115 | 19.6646 | 16.682 | 1.4012
2019-04-09 14:48 | 805 | 0.05746 | 0.032622 | 2.1242 | 24.5149 | 16.6019 | 1.4086
2019-04-09 14:48 | 806 | 0.067815 | 0.027613 | 1.8065 | 21.4946 | 16.5628 | 1.4577
2019-04-09 14:48 | 807 | 0.057577 | 0.032858 | 2.1322 | 25.3184 | 16.4706 | 1.4577
2019-04-09 14:48 | 808 | 0.057297 | 0.032631 | 2.1531 | 25.6805 | 16.4453 | 1.4714
2019-04-09 14:48 | 809 | 0.073661 | 0.025406 | 1.6883 | 19.8628 | 16.3527 | 1.4631
2019-04-09 14:48 | 810 | 0.045264 | 0.040777 | 2.7614 | 32.4858 | 16.3073 | 1.4705
2019-04-09 14:48 | 811 | 0.070147 | 0.026383 | 1.7962 | 20.9379 | 16.251 | 1.4687
2019-04-09 14:48 | 812 | 0.058107 | 0.031394 | 2.173 | 25.1671 | 16.1744 | 1.4624
2019-04-09 14:48 | 813 | 0.067334 | 0.027098 | 1.8847 | 21.9555 | 16.1309 | 1.4784
2019-04-09 14:48 | 814 | 0.052768 | 0.034549 | 2.4061 | 27.9897 | 16.0444 | 1.4769
2019-04-09 14:48 | 815 | 0.06332 | 0.029256 | 2.0239 | 23.7166 | 16.0192 | 1.5017
2019-04-09 14:48 | 816 | 0.094722 | 0.019299 | 1.3554 | 15.8086 | 15.9137 | 1.4974
2019-04-09 14:48 | 817 | 0.05877 | 0.031068 | 2.1935 | 25.1755 | 15.8446 | 1.4796
2019-04-09 14:48 | 818 | 0.082647 | 0.02204 | 1.5626 | 17.9338 | 15.784 | 1.4822
2019-04-09 14:48 | 819 | 0.064836 | 0.027787 | 1.9998 | 22.7008 | 15.6876 | 1.4718
2019-04-09 14:48 | 820 | 0.070022 | 0.026097 | 1.8628 | 21.3798 | 15.6494 | 1.4971
2019-04-09 14:48 | 821 | 0.085168 | 0.021422 | 1.5339 | 17.3438 | 15.5457 | 1.4771
2019-04-09 14:48 | 822 | 0.053034 | 0.034389 | 2.4636 | 28.4858 | 15.4895 | 1.5107
2019-04-09 14:48 | 823 | 0.068468 | 0.026331 | 1.9142 | 22.0069 | 15.4298 | 1.5068
2019-04-09 14:48 | 824 | 0.051506 | 0.034946 | 2.5477 | 29.0431 | 15.3588 | 1.4959
2019-04-09 14:48 | 825 | 0.059413 | 0.029903 | 2.2131 | 25.2385 | 15.3217 | 1.4995
2019-04-09 14:48 | 826 | 0.052759 | 0.034116 | 2.4974 | 28.3403 | 15.2427 | 1.4952
2019-04-09 14:48 | 827 | 0.055707 | 0.031926 | 2.3689 | 27.19 | 15.2143 | 1.5147
2019-04-09 14:48 | 828 | 0.053786 | 0.033069 | 2.4568 | 28.8239 | 15.1327 | 1.5503
2019-04-09 14:48 | 829 | 0.033063 | 0.053881 | 4.0172 | 46.482 | 15.1103 | 1.5368
2019-04-09 14:48 | 830 | 0.059763 | 0.029389 | 2.2235 | 25.772 | 15.0563 | 1.5402
2019-04-09 14:49 | 831 | 0.041347 | 0.041998 | 3.2298 | 37.3599 | 14.9996 | 1.5447
2019-04-09 14:49 | 832 | 0.069291 | 0.02506 | 1.9345 | 22.1897 | 14.9637 | 1.5375
2019-04-09 14:49 | 833 | 0.073923 | 0.023442 | 1.8155 | 20.646 | 14.8662 | 1.5262
2019-04-09 14:49 | 834 | 0.059954 | 0.028848 | 2.2392 | 25.2157 | 14.8207 | 1.5118
2019-04-09 14:49 | 835 | 0.066368 | 0.026042 | 2.0239 | 23.1251 | 14.7423 | 1.5348
2019-04-09 14:49 | 836 | 0.051746 | 0.033449 | 2.6015 | 29.641 | 14.692 | 1.5338
2019-04-09 14:49 | 837 | 0.063317 | 0.027661 | 2.1273 | 24.5299 | 14.6331 | 1.5532
2019-04-09 14:49 | 838 | 0.059482 | 0.029428 | 2.2665 | 26.0075 | 14.5681 | 1.547
2019-04-09 14:49 | 839 | 0.072326 | 0.024437 | 1.8645 | 21.6569 | 14.5108 | 1.5664
2019-04-09 14:49 | 840 | 0.070198 | 0.025556 | 1.9286 | 22.6976 | 14.4272 | 1.5933
2019-04-09 14:49 | 841 | 0.052017 | 0.034825 | 2.6069 | 30.5442 | 14.3738 | 1.5888
2019-04-09 14:49 | 842 | 0.070974 | 0.025592 | 1.9213 | 22.8307 | 14.3149 | 1.6204
2019-04-09 14:49 | 843 | 0.064686 | 0.02808 | 2.1149 | 24.8594 | 14.2364 | 1.608
2019-04-09 14:49 | 844 | 0.088773 | 0.020223 | 1.5473 | 18.177 | 14.1796 | 1.6136
2019-04-09 14:49 | 845 | 0.073487 | 0.024454 | 1.878 | 21.95 | 14.068 | 1.613
2019-04-09 14:49 | 846 | 0.059661 | 0.030112 | 2.3237 | 26.8254 | 14.0325 | 1.6004
2019-04-09 14:49 | 847 | 0.080278 | 0.021878 | 1.7341 | 20.289 | 13.9411 | 1.6288
2019-04-09 14:49 | 848 | 0.05873 | 0.030294 | 2.3785 | 27.2294 | 13.8806 | 1.5992
2019-04-09 14:49 | 849 | 0.073213 | 0.024306 | 1.9098 | 21.8488 | 13.8164 | 1.5996
2019-04-09 14:49 | 850 | 0.057884 | 0.030317 | 2.4105 | 27.6862 | 13.7398 | 1.6026
2019-04-09 14:49 | 851 | 0.080096 | 0.021908 | 1.739 | 19.8756 | 13.6929 | 1.592
2019-04-09 14:49 | 852 | 0.069487 | 0.025201 | 1.997 | 23.0322 | 13.5873 | 1.6004
2019-04-09 14:49 | 853 | 0.044853 | 0.039078 | 3.1012 | 35.719 | 13.5575 | 1.6021
2019-04-09 14:49 | 854 | 0.053146 | 0.032845 | 2.6077 | 30.0552 | 13.49 | 1.5973
2019-04-09 14:49 | 855 | 0.047234 | 0.037037 | 2.9455 | 34.3342 | 13.453 | 1.6217
2019-04-09 14:49 | 856 | 0.072287 | 0.024143 | 1.9252 | 22.3835 | 13.389 | 1.618
2019-04-09 14:49 | 857 | 0.059789 | 0.029251 | 2.3331 | 26.996 | 13.3165 | 1.6141
2019-04-09 14:49 | 858 | 0.067984 | 0.025723 | 2.0561 | 23.7461 | 13.2657 | 1.6144
2019-04-09 14:49 | 859 | 0.061654 | 0.027932 | 2.2645 | 26.1114 | 13.1837 | 1.6099
2019-04-09 14:49 | 860 | 0.054256 | 0.031724 | 2.5763 | 29.513 | 13.1425 | 1.6013
2019-04-09 14:49 | 861 | 0.068125 | 0.025307 | 2.0601 | 23.3397 | 13.071 | 1.59
2019-04-09 14:49 | 862 | 0.038435 | 0.045458 | 3.6579 | 41.2642 | 13.0151 | 1.586
2019-04-09 14:49 | 863 | 0.06222 | 0.028117 | 2.2593 | 25.7379 | 12.9821 | 1.6014
2019-04-09 14:49 | 864 | 0.075661 | 0.023385 | 1.8503 | 21.145 | 12.8927 | 1.5999
2019-04-09 14:49 | 865 | 0.058463 | 0.03026 | 2.3913 | 27.4709 | 12.8383 | 1.606
2019-04-09 14:49 | 866 | 0.060656 | 0.029115 | 2.2986 | 26.8612 | 12.7723 | 1.6293
2019-04-09 14:49 | 867 | 0.045116 | 0.039683 | 3.0949 | 35.2814 | 12.7201 | 1.5918
2019-04-09 14:49 | 868 | 0.060329 | 0.029701 | 2.3203 | 26.7985 | 12.6752 | 1.6167
2019-04-09 14:49 | 869 | 0.055473 | 0.031864 | 2.5277 | 29.1769 | 12.6036 | 1.6185
2019-04-09 14:49 | 870 | 0.050157 | 0.035255 | 2.7988 | 32.1051 | 12.5644 | 1.6103
2019-04-09 14:49 | 871 | 0.063327 | 0.027558 | 2.2186 | 24.7863 | 12.4987 | 1.5696
2019-04-09 14:49 | 872 | 0.076685 | 0.022768 | 1.8339 | 20.6355 | 12.4379 | 1.5824
2019-04-09 14:49 | 873 | 0.066099 | 0.026418 | 2.1258 | 23.5341 | 12.352 | 1.5556
2019-04-09 14:49 | 874 | 0.033118 | 0.053474 | 4.243 | 47.2757 | 12.3106 | 1.5657
2019-04-09 14:49 | 875 | 0.052331 | 0.033353 | 2.6705 | 29.8164 | 12.2736 | 1.5603
2019-04-09 14:49 | 876 | 0.056379 | 0.031363 | 2.4767 | 27.842 | 12.2083 | 1.5697
2019-04-09 14:49 | 877 | 0.058647 | 0.03015 | 2.3758 | 26.8789 | 12.1616 | 1.5764
2019-04-09 14:49 | 878 | 0.062613 | 0.028621 | 2.2247 | 25.0595 | 12.0911 | 1.569
2019-04-09 14:49 | 879 | 0.050818 | 0.035251 | 2.7348 | 30.548 | 12.0392 | 1.5524
2019-04-09 14:49 | 880 | 0.080906 | 0.022088 | 1.7116 | 19.4675 | 11.9809 | 1.575
2019-04-09 14:49 | 881 | 0.052493 | 0.033676 | 2.6377 | 29.98 | 11.8903 | 1.5737
2019-04-09 14:49 | 882 | 0.047736 | 0.03698 | 2.8923 | 31.9316 | 11.8707 | 1.5243
2019-04-09 14:49 | 883 | 0.075631 | 0.023416 | 1.8314 | 20.2634 | 11.787 | 1.5325
2019-04-09 14:49 | 884 | 0.06733 | 0.025924 | 2.0468 | 22.5621 | 11.7271 | 1.5191
2019-04-09 14:49 | 885 | 0.077407 | 0 | 0 | 0 | 11.6496 | 1.5142
2019-04-09 14:49 | 886 | 0.054099 | 0 | 0 | 0 | 11.579 | 1.513
2019-04-09 14:49 | 887 | 0.077488 | 0 | 0 | 0 | 11.5312 | 1.5217
2019-04-09 14:49 | 888 | 0.076859 | 0 | 0 | 0 | 11.4295 | 1.5092
2019-04-09 14:49 | 889 | 0.055524 | 0 | 0 | 0 | 11.3828 | 1.5293
2019-04-09 14:49 | 890 | 0.053773 | 0 | 0 | 0 | 11.3139 | 1.5106
2019-04-09 14:50 | 891 | 0.056858 | 0 | 0 | 0 | 11.273 | 1.5543
2019-04-09 14:50 | 892 | 0.078318 | 0 | 0 | 0 | 11.1967 | 1.4475
2019-04-09 14:50 | 893 | 0.048073 | 0 | 0 | 0 | 11.1277 | 1.563
2019-04-09 14:50 | 894 | 0.061112 | 0 | 0 | 0 | 11.0913 | 1.4135
2019-04-09 14:50 | 895 | 0.06973 | 0 | 0 | 0 | 11.0062 | 1.5995
2019-04-09 14:50 | 896 | 0.050373 | 0 | 0 | 0 | 10.9582 | 1.5174
2019-04-09 14:50 | 897 | 0.058468 | 0 | 0 | 0 | 10.9 | 1.6027
2019-04-09 14:50 | 898 | 0.040077 | 0 | 0 | 0 | 10.8459 | 1.5911
2019-04-09 14:50 | 899 | 0.069146 | 0 | 0 | 0 | 10.8097 | 1.381
2019-04-09 14:50 | 900 | 0.056548 | 0 | 0 | 0 | 10.7171 | 1.4735
2019-04-09 14:50 | 901 | 0.035771 | 0 | 0 | 0 | 10.6987 | 1.4838
2019-04-09 14:50 | 902 | 0.055837 | 0 | 0 | 0 | 10.6362 | 1.5868
2019-04-09 14:50 | 903 | 0.049287 | 0 | 0 | 0 | 10.5924 | 1.6052
2019-04-09 14:50 | 904 | 0.051787 | 0 | 0 | 0 | 10.5364 | 1.3726
2019-04-09 14:50 | 905 | 0.040237 | 0 | 0 | 0 | 10.4914 | 1.446
2019-04-09 14:50 | 906 | 0.056613 | 0 | 0 | 0 | 10.4501 | 1.2028
2019-04-09 14:50 | 907 | 0.044361 | 0 | 0 | 0 | 10.3844 | 1.3606
2019-04-09 14:50 | 908 | 0.037901 | 0 | 0 | 0 | 10.3598 | 1.4129
2019-04-09 14:50 | 909 | 0.0604 | 0 | 0 | 0 | 10.302 | 1.2348
2019-04-09 14:50 | 910 | 0.052348 | 0 | 0 | 0 | 10.2454 | 1.4125
2019-04-09 14:50 | 911 | 0.065721 | 0 | 0 | 0 | 10.1935 | 1.3541
2019-04-09 14:50 | 912 | 0.050641 | 0 | 0 | 0 | 10.1201 | 1.4611
2019-04-09 14:50 | 913 | 0.053363 | 0 | 0 | 0 | 10.0886 | 1.2605
2019-04-09 14:50 | 914 | 0.056072 | 0 | 0 | 0 | 10.0137 | 1.3505
2019-04-09 14:50 | 915 | 0.027995 | 0 | 0 | 0 | 9.9825 | 1.2774
2019-04-09 14:50 | 916 | 0.058861 | 0 | 0 | 0 | 9.9446 | 1.3765
2019-04-09 14:50 | 917 | 0.056094 | 0 | 0 | 0 | 9.8714 | 1.3916
2019-04-09 14:50 | 918 | 0.064799 | 0 | 0 | 0 | 9.8313 | 1.2002
2019-04-09 14:50 | 919 | 0.039976 | 0 | 0 | 0 | 9.7494 | 1.3529
2019-04-09 14:50 | 920 | 0.037144 | 0 | 0 | 0 | 9.7462 | 1.2865
2019-04-09 14:50 | 921 | 0.056797 | 0 | 0 | 0 | 9.6702 | 1.2746
2019-04-09 14:50 | 922 | 0.027397 | 0 | 0 | 0 | 9.643 | 1.2718
2019-04-09 14:50 | 923 | 0.045099 | 0 | 0 | 0 | 9.6054 | 1.2443
2019-04-09 14:50 | 924 | 0.038632 | 0 | 0 | 0 | 9.557 | 1.225
2019-04-09 14:50 | 925 | 0.057331 | 0 | 0 | 0 | 9.5233 | 1.2234
2019-04-09 14:50 | 926 | 0.043361 | 0 | 0 | 0 | 9.4498 | 1.2033
2019-04-09 14:50 | 927 | 0.032228 | 0 | 0 | 0 | 9.4361 | 1.1964
2019-04-09 14:50 | 928 | 0.04779 | 0 | 0 | 0 | 9.3793 | 1.2374
2019-04-09 14:50 | 929 | 0.035845 | 0 | 0 | 0 | 9.3461 | 1.1786
2019-04-09 14:50 | 930 | 0.045966 | 0 | 0 | 0 | 9.3034 | 1.2554
2019-04-09 14:50 | 931 | 0.033538 | 0 | 0 | 0 | 9.2587 | 1.1755
2019-04-09 14:50 | 932 | 0.042069 | 0 | 0 | 0 | 9.2317 | 1.1635
2019-04-09 14:50 | 933 | 0.048416 | 0 | 0 | 0 | 9.1751 | 1.1614
2019-04-09 14:50 | 934 | 0.03475 | 0 | 0 | 0 | 9.1392 | 1.176
2019-04-09 14:50 | 935 | 0.049923 | 0 | 0 | 0 | 9.0997 | 1.127
2019-04-09 14:50 | 936 | 0.034215 | 0 | 0 | 0 | 9.0456 | 1.1387
2019-04-09 14:50 | 937 | 0.045405 | 0 | 0 | 0 | 9.0255 | 1.1172
2019-04-09 14:50 | 938 | 0.047124 | 0 | 0 | 0 | 8.9566 | 1.1244
2019-04-09 14:50 | 939 | 0.040537 | 0 | 0 | 0 | 8.9329 | 1.1176
2019-04-09 14:50 | 940 | 0.062572 | 0 | 0 | 0 | 8.87 | 1.1228
2019-04-09 14:50 | 941 | 0.03337 | 0 | 0 | 0 | 8.8189 | 1.0772
2019-04-09 14:50 | 942 | 0.041682 | 0 | 0 | 0 | 8.7954 | 1.0668
2019-04-09 14:50 | 943 | 0.039768 | 0 | 0 | 0 | 8.7369 | 1.0722
2019-04-09 14:50 | 944 | 0.047535 | 0 | 0 | 0 | 8.7139 | 1.0801
2019-04-09 14:50 | 945 | 0.047801 | 0 | 0 | 0 | 8.6438 | 1.0599
2019-04-09 14:50 | 946 | 0.034043 | 0 | 0 | 0 | 8.6214 | 1.0771
2019-04-09 14:50 | 947 | 0.045765 | 0 | 0 | 0 | 8.5705 | 1.0505
2019-04-09 14:50 | 948 | 0.027741 | 0 | 0 | 0 | 8.5357 | 1.0348
2019-04-09 14:50 | 949 | 0.044137 | 0 | 0 | 0 | 8.5077 | 1.0403
2019-04-09 14:50 | 950 | 0.036723 | 0 | 0 | 0 | 8.4523 | 1.0114
2019-04-09 14:50 | 951 | 0.038872 | 0 | 0 | 0 | 8.4323 | 1.0209
2019-04-09 14:51 | 952 | 0.047791 | 0 | 0 | 0 | 8.3733 | 1.0023
2019-04-09 14:51 | 953 | 0.0297 | 0 | 0 | 0 | 8.3428 | 1.0067
2019-04-09 14:51 | 954 | 0.037242 | 0 | 0 | 0 | 8.3091 | 0.9821
2019-04-09 14:51 | 955 | 0.01254 | 0 | 0 | 0 | 8.2746 | 0.98091
2019-04-09 14:51 | 956 | 0.028897 | 0 | 0 | 0 | 8.275 | 0.96751
2019-04-09 14:51 | 957 | 0.027232 | 0 | 0 | 0 | 8.2204 | 0.94528
2019-04-09 14:51 | 958 | 0.015859 | 0 | 0 | 0 | 8.2227 | 0.95998
2019-04-09 14:51 | 959 | 0.03336 | 0 | 0 | 0 | 8.1828 | 0.96768
2019-04-09 14:51 | 960 | 0.012019 | 0 | 0 | 0 | 8.1638 | 0.94345
2019-04-09 14:51 | 961 | 0.031682 | 0 | 0 | 0 | 8.1502 | 0.90315
2019-04-09 14:51 | 962 | 0.022194 | 0 | 0 | 0 | 8.1062 | 0.92155
2019-04-09 14:51 | 963 | 0.030084 | 0 | 0 | 0 | 8.102 | 0.90358
2019-04-09 14:51 | 964 | 0.044446 | 0 | 0 | 0 | 8.0448 | 0.92478
2019-04-09 14:51 | 965 | 0.028539 | 0 | 0 | 0 | 8.0201 | 0.91347
2019-04-09 14:51 | 966 | 0.032198 | 0 | 0 | 0 | 7.9839 | 0.89018
2019-04-09 14:51 | 967 | 0.025573 | 0 | 0 | 0 | 7.957 | 0.90911
2019-04-09 14:51 | 968 | 0.046797 | 0 | 0 | 0 | 7.9269 | 0.88646
2019-04-09 14:51 | 969 | 0.035606 | 0 | 0 | 0 | 7.8708 | 0.89195
2019-04-09 14:51 | 970 | 0.025 | 0 | 0 | 0 | 7.8554 | 0.87253
2019-04-09 14:51 | 971 | 0.050637 | 0 | 0 | 0 | 7.8128 | 0.85221
2019-04-09 14:51 | 972 | 0.035087 | 0 | 0 | 0 | 7.7632 | 0.85625
2019-04-09 14:51 | 973 | 0.032637 | 0 | 0 | 0 | 7.7404 | 0.84153
2019-04-09 14:51 | 974 | 0.030971 | 0 | 0 | 0 | 7.6972 | 0.82907
2019-04-09 14:51 | 975 | 0.037306 | 0 | 0 | 0 | 7.6764 | 0.81782
2019-04-09 14:51 | 976 | 0.045615 | 0 | 0 | 0 | 7.6225 | 0.81347
2019-04-09 14:51 | 977 | 0.028399 | 0 | 0 | 0 | 7.591 | 0.79897
2019-04-09 14:51 | 978 | 0.032355 | 0 | 0 | 0 | 7.5614 | 0.79851
2019-04-09 14:51 | 979 | 0.025487 | 0 | 0 | 0 | 7.5283 | 0.77919
2019-04-09 14:51 | 980 | 0.024141 | 0 | 0 | 0 | 7.5095 | 0.77201
2019-04-09 14:51 | 981 | 0.012736 | 0 | 0 | 0 | 7.4816 | 0.76816
2019-04-09 14:51 | 982 | 0.030648 | 0 | 0 | 0 | 7.4773 | 0.7718
2019-04-09 14:51 | 983 | 0.03815 | 0 | 0 | 0 | 7.4232 | 0.74948
2019-04-09 14:51 | 984 | 0.0022078 | 0 | 0 | 0 | 7.4104 | 0.75175
2019-04-09 14:51 | 985 | 0.033053 | 0 | 0 | 0 | 7.4042 | 0.73745
2019-04-09 14:51 | 986 | 0.026899 | 0 | 0 | 0 | 7.352 | 0.72761
2019-04-09 14:51 | 987 | 0.013717 | 0 | 0 | 0 | 7.3525 | 0.72645
2019-04-09 14:51 | 988 | 0.025229 | 0 | 0 | 0 | 7.319 | 0.71801
2019-04-09 14:51 | 989 | 0.016206 | 0 | 0 | 0 | 7.3053 | 0.73068
2019-04-09 14:51 | 990 | 0.052716 | 0 | 0 | 0 | 7.277 | 0.71711
2019-04-09 14:51 | 991 | 0.040006 | 0 | 0 | 0 | 7.2111 | 0.71302

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062463 | 0.062684 | 0 | 0.2095 | 2019-04-09 11:25 | Group A | 0.1 | 0.22 | Do Nothing | 0.20951 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025185 | 0.025304 | 0 | 0.0261 | 2019-04-09 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044458 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002722 | 0.002734 | 0 | 0.0011 | 2019-04-09 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 2.2964e-06 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.5996 | 183.5796 | 0 | 100 | 2019-04-09 11:25 | None | 0 | 1000 | Do Nothing | 99.8487 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.5348 | 212.1604 | 0 | 100 | 2019-04-09 11:25 | None | 0 | 1000 | Do Nothing | 100.2342 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -14.6891 | 1447 | 0 |  | 2019-04-09 11:13 | None | -10 | 1000 | Do Nothing | 123.4044 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 20.4682 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017953 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0016805 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -154.1311 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.22987 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4017 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.002488 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0038381 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0036742 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 4.7397e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011811 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.2863 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3479 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5476 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 98.5518 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.0832 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 49.806 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.2554 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 80.9043 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 61.491 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.011889 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.055 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.024477 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.2008 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN