## Sheet1
Sections 5.6 - Audible Alarm
Test Data and Report
Manufacturer:
| Model: |  | Serial Number:
| Requirements: | The frequency of an audible alarm shall be within the range of 1000 Hz to 4000 Hz. Where an intermittent alarm is provided, the signal interval shall not exceed 2 s. The A-weighted alarm volume at a distance of 30 cm from the alarm source shall be at least 85 dB(A) and shall not exceed 100 dB(A). If the audible alarm can be shut off or disabled, the instrument shall have a visual alarm and may have a vibration alarm. It shall not be possible to disable or shut off all alarms at the same time.
| Note: | Comments are required when the requirement is not verified.
|  | Ambient Conditions: | °C | %RH |  | mmHg
Test Equipment Used:
Source Data:
Verify
[6 empty] | Yes | No
Is the alarm frequency within 1000 to 4000 Hz?
Where an intermittent alarm is provided, is the interval less than 2 seconds?
Is the alarm volume at a distance of 30 cm within 85 dB(A) and 100 dB(A)?
If the audible alarm can be disabled, does the instrument have a vibration or visual alarm?
Does the instrument have preventive measures for disabling all the alarms?
Comments:
| Completed by: |  |  |  | Date:
| Reviewed by: |  |  |  | Date: