## Disclaimer
Disclaimer | This spreadsheet must be locally validated before use!
2012-12-06 | These spreadsheets are provided as a useful approach for metrology calculations according to the SOP represented here. However, NIST OWM or any other organization involved in their development, collectively and individually, do not warrant this spreadsheets for any specific purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each user agrees to decide if, when and how to use the spreadsheets, does so at his or her sole risk, and is responsible to suitable and applicable verification prior to use. When using the tools provided on the NIST OWM website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST OWM, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST OWM, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST OWM website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## Revisions
Version | Date | By | Revisions
| 2007-02-15 |  | This spreadsheet was originally created by modifying the 3-1 weighing design spreadsheet. It contains the equations from SOP 14 for 1 pt and 2 pt calibrations and for 20 °C and 60 °F reference temperatures. The Glassware reference is with 1 pt calibration to 20 °C. The 5 Gallon Slicker reference is with a 2 pt calibration and 60 °F. Values were compared to the WA volume spreadsheets, intermediate calculations, and to the example in SOP 14 for the 2 L flask.
| 2012-07-19 |  | File was updated to modify all units and symbols consistent with NIST SP 811, SOP 2 (Air Density) 2012 Version from NISTIR 6969, GLP 10 (Water Density) 2012 Version from NISTIR 7383, SOP 14, 2012 Version NISTIR 7383, automatic calculations of uncertainty provided valid process standard deviations are entered. This is set to print on two pages.
| 2013-01-18 |  | Prepared for Volume Metrology Seminar.
| 2014-04-08 | KL | Added Disclaimer, Revisions and Instructions, Software V & V Tab (SAP 10, Form A), Lookup Table and Reference Standards tables and tabs; Edited SOP 14 OPTB Data Sets and SOP 14 OPTB spreadsheet, modified tables to include dropdown list for standards, added Us, revised and modified Laboratory Data, Calibration items and conditions table; Moved uncertainty component description tables; Added Effective Density, Air Density and Water Density Tab and table; Removed uncertainty analysis spreadsheets for Data Set 1 UNC 2 and Data Set 1 UNC 3.
| 2014-09-17 | KL | Reduced listed of acceptable units and descriptions "Lookup Table"; Made a master list of acceptable units and descriptions on "Lookup Table"; Changed all upper case U to lower case u (unless it was a expanded unc); Update all "unit" drop down picks list on UNC Table; Match Uncertainty Budget Table in NISTIR 7383 SOP 14; Hard coded MS1 and MS2 standards on DATA Sets 1 & 2; Fixs formulas in Uncerntainty Intermediate Calcs Table for SOP 14, opt B and SOP 14, opt B Data Set 2.
| 2015-08-20 | GH | Fixed relative percent contributions on Opt B and data sets; fixed reference for accepted s(p) and d.f..

## Instructions
GRAVMETRIC VOLUME CALIBRATION WORKSHEET
FILE NAME: GravCal_SOP14_OptionB_20140917.xlsx
Instructions
This worksheet is designed to calculate volume using the gravimetric method.
Corrections for:
1. The password to 'UNPROTECT' spreadsheet is metrology.
2. The user of this spreadsheet can edit light yellow input cells located on light yellow tabs. All other cells are locked.
3. Fill in "Reference Standard" tables for standards used for SOP 14 OPT B.
4. DO NOT move lookup tables.

## Software V & V
Technical Assessment
This form is completed for each software file used in the laboratory. The form may be added as a tab/worksheet to a file/workbook or may be maintained in a separate electronic/paper file/location in the laboratory.
Lab: |  |  | Software Verification & Validation
Software Description: |  |  | Date:
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense
There are instructions for use
Instructions and data input appear on the visible portion of the first worksheet
Data-entry fields are labeled and color coded (it is recommended to avoid red and green)
The Standard Operating Procedure (SOP) used is clearly specified
The number of digits to be rounded to is specified
The user is warned/notified whenever data-entry fields are left blank
Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations
The software opens at the right location within the file
Unused fields/cells are locked
Rows/columns that the operator need not see are hidden
Unused sheets are removed
Worksheets are named appropriately
B. Mathematical Specification | The correct SOP is used
The formulae and methods chosen from that SOP are specified
Sources and references for formulae are specified
The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty
C. Code review | The formulae in the fields exactly match the SOP
Repeated calculations appropriately reference the correct cells
Calculations, when tested using standard data or reference test data, show appropriate accuracy
Rounding is done at the appropriate locations in the file
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences
Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa
“Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used
E. Component Testing | Each macro used is functional
Each command/button is functional
Combinations of interdependent macros are functional
Plotted graphs are accurate
Worksheets/reports print properly, if needed to
Conditional (color and non-color) formatting is functional
F. Numerical Reference | Look-up tables and lists match the latest calibration report.
Uncertainties match the latest Scope
Values that reference another workbook or spreadsheet are dated
When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct
The evaluation of the embedded data is dated and documented
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty
J. Security | Equation and calculation cells are protected against inadvertent editing
Cells are locked in place; they cannot be moved/dragged
Confidentiality of passwords is appropriate
Files are backed up automatically
Additional back-up is available at alternate facilities
Files on network drives cannot be accidentally deleted

## Lookup Tables
Lists for Drop Downs and Lookups. Do NOT Delete. |  |  |  |  | Master List
|  |  | Acceptable Units | Description | Acceptable Units | Description
Type | Probability | Divisor
|  |  |  | blank |  | blank
A | Normal, 1s | 1 | g | gram | µg | microgram
B | Normal, 2s | 2 | in³ | cubic inch | mg | milligram
| Normal, 3s | 3 | mL | milliliter | g | gram
| Norm@99% | 2.58 | °C | degrees Celsius | kg | kilogram
| Rectangular | 1.7321 | g/cm³ | grams per cubic centimeter | µlb | micro pound
| Triangular | 2.4495 | /°C | per degrees Celsius | lb | pound
| U-Shaped | 1.4142 | mm Hg | millimeter mercury | in | inch
| Rect 1/2 | 3.4641 | % | percent | in² | square inch
|  |  | mg/cm³ | milligrams per cubic centimeter | in³ | cubic inch
[5 empty] | pt | pint
[5 empty] | qt | quart
[5 empty] | gal | gallon
[5 empty] | mL | milliliter
[5 empty] | L | liter
[5 empty] | nm | nanometer
[5 empty] | mm | millimeter
[5 empty] | cm | centimeter
[5 empty] | cm² | square centimeter
[5 empty] | cm³ | cubic centimeter
[5 empty] | m | meter
[5 empty] | m² | square meter
[5 empty] | m³ | cubic meter
[5 empty] | °F | degrees Fahrenheit
[5 empty] | °C | degrees Celsius
[5 empty] | K | Kelvin
[5 empty] | fl dr | fluid drams
[5 empty] | fl oz | Fluid ounce(s)
[5 empty] | ft | foot
[5 empty] | Pa | Pascal
[5 empty] | kPa | Kilopascal
[5 empty] | s | second
[5 empty] | min | minute
[5 empty] | h | hour
[5 empty] | A | Ampere
[5 empty] | C | coulomb
[5 empty] | V | volt
[5 empty] | J | joule
[5 empty] | W | watt
[5 empty] | Hz | hertz
[5 empty] | H | henry
[5 empty] | F | farad
[5 empty] | Ω | ohm
[5 empty] | S | siemens
[5 empty] | Wb | weber
[5 empty] | T | tesla
[5 empty] | A/m | Ampere per meter (magnetic field strength)
[5 empty] | cd | candela
[5 empty] | lm | lumen
[5 empty] | lx | lux
[5 empty] | cd/m² | candela per square meter
[5 empty] | rad | radian
[5 empty] | kg/m³ | kilogram per cubic meter
[5 empty] | g/cm³ | grams per cubic centimeter
[5 empty] | kg/L | kilograms per liter
[5 empty] | m/s | meters per second
[5 empty] | km/h | kilometers per hour
[5 empty] | m/s² | meter per second squared
[5 empty] | m²/s | meter squared per second
[5 empty] | /°C | per degrees Celsius
[5 empty] | mm Hg | millimeter mercury
[5 empty] | % | percent
[5 empty] | mg/cm³ | milligrams per cubic centimeter

## Reference Standards
Mass Standards Used (Ms1) Empty
Mass ID | Nominal (g) | Mass Correction (g) | Expanded Unc from Cal. Rpt. (g) | k factor | us | Density (g/cm3) | Volume (cm3)
Mass Standards Used (Ms2) Filled
Mass ID | Nominal (g) | Mass Correction (g) | Expanded Unc from Cal. Rpt. (g) | k factor | us | Density (g/cm3) | Volume (cm3)

## SOP 14, OptB
GRAVIMETRIC CALIBRATION SOP 14, OPTION B
60 °F Reference Temperature
Laboratory data, calibration item, and conditions:
Vessel Mfg and SN |  | Test Number |  |  |  |  | Test Notes
Nominal Volume (gal) |  | Operator | [8 empty] | 0.0036
Material |  | Date
Cubical Coefficient of Expansion / °F |  | Reference Water Temperature, °F | 60
Note: CCE in /°C | 0 |  | Run 1 Average | Run 2 Average | Uncertainty | Average Environmental
Balance |  | Air Temperature, °C | #DIV/0! | #DIV/0! |  | #DIV/0!
Nominal Load (g) |  | Pressure, mm Hg | #DIV/0! | #DIV/0! |  | #DIV/0!
Standard deviation of the process, from control chart, sp, mL |  | Relative Humidity, % | #DIV/0! | #DIV/0! |  | #DIV/0!
Degrees of Freedom |  | Unc water temperature, °C
Mass standard(s) data (EMPTY):
ID (Note ID; Select for empty Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
Sum and Effective Density
Mass standard(s) data (FILLED):
ID (Note ID; Select for filled Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
Sum and Effective Density
*Mass Correction = True Mass! Density is required for buoyancy corrections.
Measurements:
Run 1 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 |  |  |  |  | Air Temperature, °C
2 | Empty or Drained | O2 |  |  |  |  | Pressure, mm Hg
3 | MS2 | O3 |  |  |  |  | Relative Humidity, %
4 | Filled | O4 |  |  |  |  | Final water temp tw °C:
Run 2 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 |  |  |  |  | Air Temperature, °C
2 | Empty or Drained | O2 |  |  |  |  | Pressure, mm Hg
3 | MS2 | O3 |  |  |  |  | Relative Humidity, %
4 | Filled | O4 |  |  |  |  | Final water temp tw °C:
Final Volume at 60 °F: |  | #DIV/0! | mL | #DIV/0! | gal | #DIV/0! | in3
Reported Uncertainty: |  | #NUM! | mL | #NUM! | gal | #NUM! | in3
Intermediate Calculations:
** Calculated Air Density ** |  |  | ** rounded to 9 digits
Run 1, Average Air Density (g/cm3): |  | #DIV/0! | SOP 2, Option B (NIST IR 6969), 2012
Run 2, AverageAir Density (g/cm3): |  | #DIV/0! | SOP 2, Option B (NIST IR 6969), 2012
** Calculated Water Density ** |  |  | ** rounded to 8 digits
Run 1: Water Density (g/cm3): |  | 0.99984 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
Run 2: Water Density (g/cm3): |  | 0.99984 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
** Calculated Temperature Correction ** |  |  | ** rounded to 6 digits
Run 1: Temperature Correction Factor: |  | 1 | Reference temp 60 °F
Run 2: Temperature Correction Factor: |  | 1 | Reference temp 60 °F
** Calculated Volume at Test Temperature **
Run 1: Volume at Test Temperature |  | #DIV/0! | mL
Run 2: Volume at Test Temperature |  | #DIV/0! | mL
** Calculated Volume at 60 °F Reference**
Run 1: Volume at 60 °F |  | #DIV/0! | mL (Enter in Control Chart if Sc or Lab Std)
Run 2: Volume at 60 °F |  | #DIV/0! | mL (Enter in Control Chart if Sc or Lab Std)
** Control Chart Data and Assessment**
Std Dev of Run 1 and Run 2 |  | #DIV/0! | mL (Enter in Std Dev Chart if no Sc)
Accepted Std Dev (converted to mL) |  | 0 | Incomplete
Measurement Result Units: | mL |  |  | Measurement Range and Parameter:
Instructions: Fill in both "components" and "est unc in meas units" AND | [6 empty] | select "prob distribution" for complete calculations.
Uncertainty Component Description | Symbol | Estimated Uncertainty | Units | d.f. | Est. Unc in Measurement Units (mL) | Type (A, B) | Probability Distribution | Divisor | Std Unc (mL) | Relative Contribution (%) | Explanation/Source/Notes | For effective degrees of freedom calculation.
Mass | us | 0 | g | 100 | 0 | A | Normal, 1s | Finish | Finish | Finish |  | 0
Meniscus Reading | um |  | in³ | 20 | 0 | B | Triangular | Finish | Finish | Finish |  | 0
Accepted sp (Range/1.414) | sp | 0 | mL | 0 | 0 | B | Normal, 1s | Finish | Finish | Finish |  | #DIV/0!
Water Temp | utx | 0 | °C | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Coeff. of Cubical Exp. | uCCE |  | /°C | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
Water Density | upw |  | g/cm³ | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | Uncertainty Intermediate Calcs
Air Temp | uc | 0 | °C | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | 1 | 0.99984
Pressure | up | 0 | mm Hg | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | 1
RH | uRH | 0 | % | 100 | #DIV/0! | B | Normal, 1s | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | 0.99984
Air Density | upa | #DIV/0! | mg/cm³ | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
Mass Density | upm | 0.05 | g/cm³ | 100 | #DIV/0! | B | Rectangular | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0!
[13 empty] | #DIV/0! | #DIV/0!
Instructions: Data Entry is Incomplete!!! | [9 empty] | #DIV/0! |  | #DIV/0!
Min Degrees of Freedom | ν |  |  | 0
Effective Degrees of Freedom | νeff |  |  | #DIV/0!
Combined Uncertainty, uc | [8 empty] | #DIV/0! | Instructions: Assess data entry and values before reporting rounded result. |  | #DIV/0!
Coverage factor, k, uses effective degrees of freedom | [8 empty] | TBD
Expanded Uncertainty, U | [8 empty] | TBD |  |  | #DIV/0!
Expanded Uncertainty, U, Rounded to 2 Significant Digits | [8 empty] | TBD | mL
Note on I35: Determine after weighing, before emptying.
Note on I40: Determine after weighing, before emptying.

## CIPM Air Density 1981_91 (hidden)
Password to unlock: | metrology
Do not delete these cells
Air Density Calculation |  |  |  | They are used to calculate air density.
Barometric Pressure: | 760 | mm Hg |  | 101325 | Pa
Temperature: | 20 | °C |  | 293.15 | °K
Relative Humidity: | 42 | % |  | 42 | %
Calculated Air Density: | 1.2001 | mg/cm3 | 0.0012001 | g/cm3
NOTE: When copying these cells into a spreadsheet, they must be pasted into exactly
the same cells on a new worksheet (i.e., the block from A3 to G7).
Cell references in the equations are "absolute" and fixed to the reference cells in each equation.
Once the cells are pasted, the entire block may be dragged/dropped.

## Water Density (hidden)
| Patterson/Morris (1994), Air-free |  | Jones/Harris (1992), Air-saturated |  | Difference |  | P0 | 999.9736
Water Temp - C | Water Density | [5 empty] | A | 7.0134e-08
5 | 0.99997 | g/m3 | 0.99996 | g/m3 | 4.3703 | ppm | B | 7.9265e-06
6 | 0.99994 | g/m3 | 0.99994 | g/m3 | 5.7643 |  | C | -7.5757e-08
7 | 0.9999 | g/m3 | 0.9999 | g/m3 | 6.1865 |  | D | 7.3149e-10
8 | 0.99985 | g/m3 | 0.99984 | g/m3 | 6.0487 |  | E | -3.5965e-12
9 | 0.99978 | g/m3 | 0.99978 | g/m3 | 5.7462 |  | T0 | 3.9818
10 | 0.9997 | g/m3 | 0.9997 | g/m3 | 6.6588
11 | 0.99961 | g/m3 | 0.9996 | g/m3 | 6.1508
12 | 0.9995 | g/m3 | 0.99949 | g/m3 | 5.5715
13 | 0.99938 | g/m3 | 0.99937 | g/m3 | 6.2557
14 | 0.99925 | g/m3 | 0.99924 | g/m3 | 5.5241
15 | 0.9991 | g/m3 | 0.9991 | g/m3 | 5.6837
16 | 0.99895 | g/m3 | 0.99894 | g/m3 | 5.0284
17 | 0.99878 | g/m3 | 0.99877 | g/m3 | 4.839
18 | 0.9986 | g/m3 | 0.99859 | g/m3 | 5.3844
19 | 0.99841 | g/m3 | 0.9984 | g/m3 | 4.9211
20 | 0.99821 | g/m3 | 0.9982 | g/m3 | 4.6943
21 | 0.99799 | g/m3 | 0.99799 | g/m3 | 4.9383
22 | 0.99777 | g/m3 | 0.99777 | g/m3 | 4.8764
23 | 0.99754 | g/m3 | 0.99753 | g/m3 | 4.722
24 | 0.9973 | g/m3 | 0.99729 | g/m3 | 4.6784
25 | 0.99705 | g/m3 | 0.99704 | g/m3 | 4.94
26 | 0.99678 | g/m3 | 0.99678 | g/m3 | 4.6917
27 | 0.99651 | g/m3 | 0.99651 | g/m3 | 5.1105
28 | 0.99623 | g/m3 | 0.99623 | g/m3 | 4.3651
29 | 0.99595 | g/m3 | 0.99594 | g/m3 | 4.6163
30 | 0.99565 | g/m3 | 0.99564 | g/m3 | 5.0182
31 | 0.99534 | g/m3 | 0.99534 | g/m3 | 4.7179
32 | 0.99503 | g/m3 | 0.99502 | g/m3 | 3.8561
33 | 0.9947 | g/m3 | 0.9947 | g/m3 | 4.5678
34 | 0.99437 | g/m3 | 0.99437 | g/m3 | 3.9823
35 | 0.99403 | g/m3 | 0.99403 | g/m3 | 4.2241
36 | 0.99368 | g/m3 | 0.99368 | g/m3 | 3.4129
37 | 0.99333 | g/m3 | 0.99333 | g/m3 | 3.6644
38 | 0.99297 | g/m3 | 0.99296 | g/m3 | 4.0906
39 | 0.99259 | g/m3 | 0.99259 | g/m3 | 3.7999
40 | 0.99221 | g/m3 | 0.99221 | g/m3 | 3.898
Jones, F.E., and Harris, G. L. ITS-90 Density of Water Formulation for
Volumetric Standards Calibration, NIST Jrnl of Research, 97, 335-340 (1992).

## Effective Density
Effective Density
|  | Mass (g) | Density (g/cm³) | Calculated Volume (cm³)
Weight #1
Weight #2
Weight #3
Weight #4
Weight #5
Weight #6
Weight #7
Weight #8
Weight #9
Weight #10
|  | Total Mass (g) | Calculated Effective Density (g/cm³) | Total Calculated Volume (cm³)
|  | 0 |  | 0
Step 1
Enter Mass Values for
up to 10 weights
Step 2
Enter Densities for
the weights
Step 3
Effective Density

## Air Density
Uncertainty (k=2) of Equation ± µg/cm3 | Bias as compared to CIPM2007 (µg/cm3) | Source | Air Density mg/cm3
34 | #DIV/0! | Lembeck 1974? | 1.14 (summer) 1.15 (winter) (Oklahoma)
2 | #DIV/0! | TN 577 1971 | 0 |  |  |  | Ma | R | Kelvin (T) | Pa | xv | Z
0.35 | #DIV/0! | SOP 2, Option A (F Jones 1978) | 0 |  |  |  | g mol-1 | J mol-1 K-1 | K = 273.15+oC | 101325 Pa = 760 mmHg |  |  | a(0) | a(1) | a(2) | b(0) | b(1) | c(0) | c(1) | d | e
0.026 | #DIV/0! | CIPM 1981 | #DIV/0! |  |  |  | 28.9635 | 8.3144 | 273.15 | 0 | #DIV/0! | #DIV/0! | 1.6242e-06 | -2.8969e-08 | 1.088e-10 | 5.757e-06 | -2.589e-08 | 0.00019297 | -2.285e-06 | 1.73e-11 | -1.034e-08
0.026 | #DIV/0! | SOP 2, Option B, 2003 (CIPM 1981/91) | #DIV/0! |  |  |  | 28.9635 | 8.3145 | 273.15 | 0 | #DIV/0! | #DIV/0! | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09
#DIV/0! | 0 | SOP 2, Option B, 2012 (CIPM 2007) | #DIV/0! |  |  |  | 28.9655 | 8.3145 | 273.15 | 0 | #DIV/0! | #DIV/0! | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09
|  | CIPM 2007 entire formula in one cell | #DIV/0!
Reference Data
20 |  | P (mm Hg) | Temp (oC) | RH (%)
760 | Values
42 | Uncertainty
These Parameters are close to Normal Air (1.2 mg/cm3) using CIPM2007
Equation Components
Equation Components
|  | u(Pressure) | u(Temperature) | u(Relative Humidity) | u(Equation)
|  | 0 | 0 | 0 | #DIV/0!
df
Measurement Equation Inputs | Value
Pressure | 0 | 0 | 0 | 0 | 0
Temperature | 0 | 0 | 0 | 0 | 0
Relative Humidity | 0 | 0 | 0 | 0 | 0
Equation | 0 | 0 | 0 | 0 | #DIV/0!
Result | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | ciui
uc | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | (ciui)2
df | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | <--rel (ciui)2 | #DIV/0!
k | 2 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | ci | Srel (ciui)2
U | #DIV/0! |  | #DIV/0! | Urelative, %
Equation Components
Equation Components
Equation Components
Step 1
Enter Pressure
Step 2
Enter Temperature
Step 3
Enter % Rel Humidity

## H2O Density
Uncertainty (k=2) of Equation ± µg/cm3 | Bias as compared to CIPM (µg/cm3) | Source | Water Density g/cm3 |  | Reference Data Set (°C) | W-B (HB 145, 1986) Air Free (g/cm3) | Jones-Harris (1992) Air Saturated (g/cm3) | Patterson-Morris Air Free (g/cm3) | Tanaka, et al Air Free (g/cm3) | Tanaka, et al Air Saturated (g/cm3)
|  | H. Wagenbreth and W. Blanke, PTB - Mitteilingen 6-71. HB 145, NISTIR 6969 2012 (Air Free) 20 °C | 0.998 202 00 |  | 10 | 0.999 699 | 0.9997 | 0.9997 | 0.9997 | 0.9997
10 | 3.4822 | Jones Harris NIST Jrnl of Res (1992) (Air Saturated) | 0.99985 |  | 15 | 0.999 098 | 0.9991 | 0.9991 | 0.9991 | 0.9991
1 | 1.4408 | GLP 10, NISTIR 7383 (2006) Patterson-Morris (1994) (Air Free) | 0.99984 |  | 20 | 0.998 202 | 0.9982 | 0.99821 | 0.99821 | 0.9982
0.84 | 2.4964 | GLP 10, NISTIR 7383 (2012, 2013) Tanaka, et al (2001) (Air Free) | 0.99984 |  | 25 | 0.997 043 | 0.99704 | 0.99705 | 0.99705 | 0.99705
0.84 | 0 | GLP 10, NISTIR 7383 (2012, 2013) Tanaka, et al (2001) (Air Saturated) | 0.99984
|  | CIPM 2001, Tanaka, et al Note: Air saturated needs to be used in the lab. | 0.99984
Temp (oC)
Values
Uncertainty
Using Tanaka, et al Equation
|  | u(Equation) | u(AFvsAS) | u(Temp)
|  | 8.4e-07 | 4.6164e-06 | 0
| df | 1 | 1 | 1
Measurement Equation Inputs | Value
Equation | 0 | 8.4e-07 | 0 | 0
AFvsAS | 0 | 0 | 4.6164e-06 | 0
Temp | 0 | 0 | 0 | 0
Result | 0.99984 | 0.99984 | 0.99984 | 0.99984
|  | 0 | 0 | 0 | ciui
uc | 0 | 0 | 0 | 0 | (ciui)2
df | 1 | #DIV/0! | #DIV/0! | #DIV/0! | <--rel (ciui)2 | #DIV/0!
k | 2 | 0 | 0 | #DIV/0! | ci | Srel (ciui)2
U | 0 |  | 0% | Urelative, %
Step 1
Enter Temperature

## SOP 14, OptB Data Set 1
GRAVIMETRIC CALIBRATION SOP 14, OPTION B
60 °F Reference Temperature
Laboratory data, calibration item, and conditions:
Vessel Mfg and SN | NA | Test Number | NA |  |  |  | Test Notes
Nominal Volume (gal) | 5 | Operator | NA
Material | SS | Date | 1/20/2013
Cubical Coefficient of Expansion / °F | 2.65e-05 | Reference Water Temperature, °F | 60 |  |  |  |  | Locked Sheet - Data Set 1
Note | 4.77e-05 |  | Run 1 Average | Run 2 Average | Uncertainty | Average Environmental
Balance | NA | Air Temperature, °C | 20.67 | 20.82 | 0.1 | 20.745
Nominal Load (g) | 20000 | Pressure, mm Hg | 752.4 | 752 | 0.5 | 752.2
Standard deviation of the process, from control chart, sp, mL | 1.4 | Relative Humidity, % | 45.47 | 49.2 | 2 | 47.335
Degrees of Freedom | 200 | Unc water temperature, °C |  |  | 0.1
Mass standard(s) data (EMPTY):
ID (Note ID; Select for empty Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
MS1 | 4999.9908 | 0 | 0.00047126 | 1 | 0.00047126 | 8 | 624.9988
Sum and Effective Density | 4999.9908 |  |  |  | 0.00047126 | 8 | 624.9988
Mass standard(s) data (FILLED):
ID (Note ID; Select for filled Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
MS2 | 20000.0205 | 0 | 0.001509 | 1 | 0.001509 | 8 | 2500.0026
Sum and Effective Density | 20000.0205 |  |  |  | 0.001509 | 8 | 2500.0026
*Mass Correction = True Mass! Density is required for buoyancy corrections.
Measurements:
Run 1 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 | 4999.934 |  |  |  | Air Temperature, °C | 20.67 | 20.67
2 | Empty or Drained | O2 | 5088.166 |  |  |  | Pressure, mm Hg | 752.4 | 752.4
3 | MS2 | O3 | 19999.704 |  |  |  | Relative Humidity, % | 45.47 | 45.47
4 | Filled | O4 | 23963.164 |  |  |  | Final water temp tw °C: | 21.1
Run 2 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 | 4999.948 |  |  |  | Air Temperature, °C | 20.82 | 20.82
2 | Empty or Drained | O2 | 5087.91 |  |  |  | Pressure, mm Hg | 752 | 752
3 | MS2 | O3 | 19999.664 |  |  |  | Relative Humidity, % | 49.2 | 49.2
4 | Filled | O4 | 23957.894 |  |  |  | Final water temp tw °C: | 21.1
Final Volume at 60 °F: |  | 18,925.9 | mL | 4.99969 | gal | 1,154.93 | in3
Reported Uncertainty: |  | 3.3 | mL | 0.00087 | gal | 0.20 | in3
Intermediate Calculations:
** Calculated Air Density ** |  |  | ** rounded to 9 digits
Run 1, Average Air Density (g/cm3): |  | 0.0011848 | SOP 2, Option B (NIST IR 6969), 2012
Run 2, AverageAir Density (g/cm3): |  | 0.0011831 | SOP 2, Option B (NIST IR 6969), 2012
** Calculated Water Density ** |  |  | ** rounded to 8 digits
Run 1: Water Density (g/cm3): |  | 0.99797 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
Run 2: Water Density (g/cm3): |  | 0.99797 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
** Calculated Temperature Correction ** |  |  | ** rounded to 6 digits
Run 1: Temperature Correction Factor: |  | 0.99974 | Reference temp 60 °F
Run 2: Temperature Correction Factor: |  | 0.99974 | Reference temp 60 °F
** Calculated Volume at Test Temperature **
Run 1: Volume at Test Temperature |  | 18933.3744 | mL
Run 2: Volume at Test Temperature |  | 18928.3791 | mL
** Calculated Volume at 60 °F Reference**
Run 1: Volume at 60 °F |  | 18928.376 | mL (Enter in Control Chart if Sc or Lab Std)
Run 2: Volume at 60 °F |  | 18923.382 | mL (Enter in Control Chart if Sc or Lab Std)
** Control Chart Data and Assessment**
Std Dev of Run 1 and Run 2 |  | 3.5313 | mL (Enter in Std Dev Chart if no Sc)
Accepted Std Dev (converted to mL) |  | 1.4 | Difference is less than Control Limits, Pass
Measurement Result Units: | in³ |  |  | Measurement Range and Parameter:
Instructions: Fill in both "components" and "est unc in meas units" AND | [6 empty] | select "prob distribution" for complete calculations.
Uncertainty Component Description | Symbol | Estimated Uncertainty | Units | d.f. | Est. Unc in Measurement Units (mL) | Type (A, B) | Probability Distribution | Divisor | Std Unc (mL) | Relative Contribution (%) | Explanation/Source/Notes | For effective degrees of freedom calculation.
Mass | us | 0.0015809 | g | 20 | 0.0015841 | A | Normal, 1s | 1 | 0.0015841 | 0.0001182% |  | 3.1485e-13
Meniscus Reading | um | 1e-15 | mL | 10 | 1.6387e-14 | B | Rectangular | 1.7321 | 9.4611e-15 | 4.217e-27% |  | 8.0124e-58
Accepted sp (Range/1.414) | sp | 1.4 | mL | 10 | 1.4 | B | Normal, 1s | 1 | 1.4 | 92.34% |  | 0.38416
Water Temp | utx | 0.1 | °C | 10 | 0.094564 | B | Rectangular | 1.7321 | 0.054596 | 0.1404% |  | 8.885e-07
Coeff. of Cubical Exp. | uCCE | 1e-06 | /°C | 10 | 0.1136 | B | Rectangular | 1.7321 | 0.065587 | 0.2027% |  | 1.8504e-06
Water Density | upw | 8.3e-07 | g/cm³ | 10 | 0.015765 | B | Rectangular | 1.7321 | 0.0091021 | 0.003903% |  | 6.8639e-10 | Uncertainty Intermediate Calcs
Air Temp | uc | 0.1 | °C | 10 | -0.0030352 | B | Rectangular | 1.7321 | -0.0017524 | 0.0001447% |  | 9.4298e-13 | 0.99973 | 0.99795
Pressure | up | 0.5 | mm Hg | 10 | -0.00013182 | B | Rectangular | 1.7321 | -7.6107e-05 | 2.729e-07% |  | 3.3551e-18 | 0.99973
RH | uRH | 2 | % | 10 | -0.0025212 | B | Rectangular | 1.7321 | -0.0014556 | 9.982e-05% |  | 4.489e-13 | 0.99797
Air Density | upa | 2.6047e-08 | mg/cm³ | 10 | -0.0004331 | B | Normal, 1s | 1 | -0.0004331 | 8.837e-06% |  | 3.5185e-15 | 0.0011835 | 20.845
Viscosity | uv | 3.6039e-05 |  | 10 | 0.68216 | B | Rectangular | 1.7321 | 0.39384 | 7.308% |  | 0.002406 | 0.0011848 | 752.7
Mass Density | upm | 0.05 | g/cm³ | 10 | -0.017419 | B | Normal, 1s | 1.7321 | -0.010057 | 0.004765% |  | 1.023e-09 | 0.0011838 | 49.335
Instructions: Data Entry is Incomplete!!! | [9 empty] | 100% |  | 0.38657
Min Degrees of Freedom | ν |  |  | 10
Effective Degrees of Freedom | νeff |  |  | 11
Combined Uncertainty, uc | [8 empty] | 1.4569 | Instructions: Assess data entry and values before reporting rounded result. |  | 4.5054
Coverage factor, k, uses effective degrees of freedom | [8 empty] | 2.25
Expanded Uncertainty, U | [8 empty] | 3.278 |  |  | 11
Expanded Uncertainty, U, Rounded to 2 Significant Digits | [8 empty] | 3.3 | in³
Note on I35: Determine before emptying flask.
Note on I40: Determine before emptying flask.

## SOP 14, OptB Data Set 2
GRAVIMETRIC CALIBRATION SOP 14, OPTION B
60 °F Reference Temperature
Laboratory data, calibration item, and conditions:
Vessel Mfg and SN | NA | Test Number | NBS Data Set |  |  |  | Test Notes
Nominal Volume (gal) | 5 | Operator | HVO
Material | SS | Date | 3/6/1986
Cubical Coefficient of Expansion / °F | 2.65e-05 | Reference Water Temperature, °F | 60 |  |  |  |  | Locked Sheet - Data Set 2
Note | 4.77e-05 |  | Run 1 Average | Run 2 Average | Uncertainty | Average Environmental
Balance |  | Air Temperature, °C | 25.875 | 25.875 | 0.1 | 25.875
Nominal Load (g) | 24000 | Pressure, mm Hg | 751.06 | 751.06 | 0.5 | 751.06
Standard deviation of the process, from control chart, sp, mL | 1.4 | Relative Humidity, % | 35.15 | 35.15 | 2 | 35.15
Degrees of Freedom | 200 | Unc water temperature, °C |  |  | 0.1
Mass standard(s) data (EMPTY):
ID (Note ID; Select for empty Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
MS1 | 4685 | 0 | 0.00047126 | 1 | 0.00047126 | 8 | 585.625
Sum and Effective Density | 4685 |  |  |  | 0.00047126 | 8 | 585.625
Mass standard(s) data (FILLED):
ID (Note ID; Select for filled Load) | Nominal (g) | Mass Correction* (g) | Expanded Unc from Cal. Rpt. (g) | Unc (k factor) | us | Density (g/cm3) | Volume (cm3)
MS2 | 23554 | 0 | 0.001509 | 1 | 0.001509 | 8 | 2944.25
Sum and Effective Density | 23554 |  |  |  | 0.001509 | 8 | 2944.25
*Mass Correction = True Mass! Density is required for buoyancy corrections.
Measurements:
Run 1 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 | 4685 |  |  |  | Air Temperature, °C | 25.875 | 25.875
2 | Empty or Drained | O2 | 4684.83 |  |  |  | Pressure, mm Hg | 751.06 | 751.06
3 | MS2 | O3 | 23554 |  |  |  | Relative Humidity, % | 35.15 | 35.15
4 | Filled | O4 | 23553.735 |  |  |  | Final water temp tw °C: | 24.8
Run 2 | Measured Item Zero balance before each observation. | Balance Observations, Units: | gram |  |  |  |  | Before | After
1 | MS1 | O1 | 4685 |  |  |  | Air Temperature, °C | 25.875 | 25.875
2 | Empty or Drained | O2 | 4684.83 |  |  |  | Pressure, mm Hg | 751.06 | 751.06
3 | MS2 | O3 | 23554 |  |  |  | Relative Humidity, % | 35.15 | 35.15
4 | Filled | O4 | 23553.735 |  |  |  | Final water temp tw °C: | 24.8
Final Volume at 60 °F: |  | 18,934.8 | mL | 5.00205 | gal | 1,155.47 | in3
Reported Uncertainty: |  | 3.4 | mL | 0.00090 | gal | 0.21 | in3
Intermediate Calculations:
** Calculated Air Density ** |  |  | ** rounded to 9 digits
Run 1, Average Air Density (g/cm3): |  | 0.0011618 | SOP 2, Option B (NIST IR 6969), 2012
Run 2, AverageAir Density (g/cm3): |  | 0.0011618 | SOP 2, Option B (NIST IR 6969), 2012
** Calculated Water Density ** |  |  | ** rounded to 8 digits
Run 1: Water Density (g/cm3): |  | 0.9971 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
Run 2: Water Density (g/cm3): |  | 0.9971 | Tanaka, et al (2012), Air-Saturated (0 °C to 40 °C)
** Calculated Temperature Correction ** |  |  | ** rounded to 6 digits
Run 1: Temperature Correction Factor: |  | 0.99956 | Reference temp 60 °F
Run 2: Temperature Correction Factor: |  | 0.99956 | Reference temp 60 °F
** Calculated Volume at Test Temperature **
Run 1: Volume at Test Temperature |  | 18943.1804 | mL
Run 2: Volume at Test Temperature |  | 18943.1804 | mL
** Calculated Volume at 60 °F Reference**
Run 1: Volume at 60 °F |  | 18934.8265 | mL (Enter in Control Chart if Sc or Lab Std)
Run 2: Volume at 60 °F |  | 18934.8265 | mL (Enter in Control Chart if Sc or Lab Std)
** Control Chart Data and Assessment**
Std Dev of Run 1 and Run 2 |  | 0 | mL (Enter in Std Dev Chart if no Sc)
Accepted Std Dev (converted to mL) |  | 1.4 | Difference is less than Control Limits, Pass
Measurement Result Units: | in³ |  |  | Measurement Range and Parameter:
Instructions: Fill in both "components" and "est unc in meas units" AND | [6 empty] | select "prob distribution" for complete calculations.
Uncertainty Component Description | Symbol | Estimated Uncertainty | Units | d.f. | Est. Unc in Measurement Units (mL) | Type (A, B) | Probability Distribution | Divisor | Std Unc (mL) | Relative Contribution (%) | Explanation/Source/Notes | For effective degrees of freedom calculation.
Mass | us | 0.0015809 | g | 20 | 0.0015855 | A | Normal, 1s | 1 | 0.0015855 | 0.0001046% |  | 3.1596e-13
Meniscus Reading | um | 1e-12 | mL | 10 | 1.6387e-11 | B | Rectangular | 1.7321 | 9.4611e-12 | 3.723e-21% |  | 8.0124e-46
Accepted sp (Range/1.414) | sp | 1.4 | mL | 10 | 1.4 | B | Normal, 1s | 1 | 1.4 | 81.52% |  | 0.38416
Water Temp | utx | 0.1 | °C | 10 | 0.094587 | B | Rectangular | 1.7321 | 0.05461 | 0.124% |  | 8.8938e-07
Coeff. of Cubical Exp. | uCCE | 1e-06 | /°C | 10 | 0.17049 | B | Rectangular | 1.7321 | 0.098432 | 0.403% |  | 9.3873e-06
Water Density | upw | 8.3e-07 | g/cm³ | 10 | 0.015787 | B | Rectangular | 1.7321 | 0.0091146 | 0.003455% |  | 6.9017e-10 | Uncertainty Intermediate Calcs
Air Temp | uc | 0.1 | °C | 10 | -0.00099479 | B | Rectangular | 1.7321 | -0.00057434 | 1.372e-05% |  | 1.0881e-14 | 0.99955 | 0.99707
Pressure | up | 0.5 | mm Hg | 10 | 0.0018403 | B | Rectangular | 1.7321 | 0.0010625 | 4.695e-05% |  | 1.2745e-13 | 0.99955
RH | uRH | 2 | % | 10 | -0.00069428 | B | Rectangular | 1.7321 | -0.00040084 | 6.683e-06% |  | 2.5816e-15 | 0.9971
Air Density | upa | 2.5559e-08 | mg/cm³ | 10 | -0.00042561 | B | Normal, 1s | 1 | -0.00042561 | 7.534e-06% |  | 3.2815e-15 | 0.0011614 | 25.975
Viscosity | uv | 6.0089e-05 |  | 10 | 1.1378 | B | Rectangular | 1.7321 | 0.65689 | 17.95% |  | 0.01862 | 0.0011625 | 751.56
Mass Density | upm | 0.05 | g/cm³ | 10 | -0.017089 | B | Normal, 1s | 1.7321 | -0.0098664 | 0.004049% |  | 9.4763e-10 | 0.0011615 | 37.15
Instructions: Data Entry is Incomplete!!! | [9 empty] | 100% |  | 0.40279
Min Degrees of Freedom | ν |  |  | 10
Effective Degrees of Freedom | νeff |  |  | 14
Combined Uncertainty, uc | [8 empty] | 1.5506 | Instructions: Assess data entry and values before reporting rounded result. |  | 5.781
Coverage factor, k, uses effective degrees of freedom | [8 empty] | 2.2
Expanded Uncertainty, U | [8 empty] | 3.4113 |  |  | 14
Expanded Uncertainty, U, Rounded to 2 Significant Digits | [8 empty] | 3.4 | in³
Note on I35: Determine before emptying flask.
Note on I40: Determine before emptying flask.