## Slide 1: High Flux Backscattering Instrument

## Slide 2: Why Neutrons?
Wavelength of neutrons:
At 2meV, wavelength of
neutrons is about 5Ǻ
typical length scale of interest
in soft matter
Poly(methyl methacrylate)

## Slide 3
Elastic Neutron Scattering
No exchange of energy Ei-Ef=0
Examines change in momentum or direction.
Inelastic Neutron Scattering
Examines both energy and change in momentum or direction.
Pynn, Neutron Scattering: A Primer (1989)
Scattering Concepts
length scales explored:
in terms of Q (=2π/d)

## Slide 4
Scattering function S(Q,ω) {or
S(Q,t)} depends only on your sample!
DSF and ISF are related
by Fourier transform.
Intensity of scattered neutrons
in a given direction θ:
Scattering Quantities
Dynamic Structure Factor
Intermediate Scattering Function

## Slide 5
Measurable quantity in a NS experiment
Coherent and incoherent
coh and inc dynamic structure factors
Scattering Quantities
Dynamic Structure
Factor

## Slide 6
Incoherent intermediate scattering
function relates the motion of a given
nucleus at t=0 and at a later time t;
Incoherent
atom-atom self correlation
Coherent scattering function relates
position of a pair of atoms at different
times
Coherent
pair correlation function
time
integral
S(Q)
PMMA
Scattering Quantities

## Slide 7
Types of Scattering Functions
Dynamic Structure Factors
Intermediate Scattering Functions
Correlation Functions
Fourier
Transform
Inverse
Fourier
Transform
is the probability of finding a particle at position r after time t
if that same particle was at position 0 at time t=0
is the probability of finding a particle at position r after time t
if there was a particle at position 0 at time t=0
Scattering Quantities

## Slide 8
elements Incoherent coherent
σ (H) ~ 82 2
σ (D) ~ 2 5
σ (C) ~ 0 5
σ (O) ~ 0 4
A(H)/B(D)
A(D)/B(H)
Importance of hydrogen?

## Slide 9
Types of Scattering and Motions
Elastic neutron scattering:
is the scattering for which the energy transfer is
identically zero.
Quasi-elastic neutron scattering:
peaks at zero energy transfer, but is broadened
compared to the instrumental resolution. It arises
from diffusive or diffusive-like processes.
Inelastic neutron scattering:
peaks at non-zero energy transfer. This scattering
reflects the vibrational or fast modes of the system.
Elastic
peak
Quasi-elastic
Inelastic
Energy

## Slide 10
Broadening - What Can We Learn?
Fick’s Law
is the solution
Lorentzian function
Intensity
Energy
E=0
Q1
Q2
Q2 > Q1

## Slide 11
Z. Phys. Chem. 224 (2010) 5-32
N-equivalent sites
located on a circle
Broadening - What Can We Learn?

## Slide 12
PVAc at 460K
Time domain Vs frequency domain
NSE
BS
β=0.50

## Slide 13: Why Backscattering?
Basic eq. for resolution:
intrinsic term depends
on crystal properties only
depends only on angular divergence
minimized in BS by choosing θ to be 90 degrees!
NH4ClO4
H→D

## Slide 14: Why HFBS?
i) PST chopper increases effective flux
at sample almost 4 times!
ii) Backscattering with high energy resolution.
iii) Cam-based doppler drive is able to
extend the dynamic window to ±36 μeV.
almost Gaussian

## Slide 15: Choosing your sample
Self motion of PMMA:
have it hydrogenated
A polymer example: poly(methyl methacrylate)
PMMA:
>95%
>70%
[-C-CH2-]
CH3
COO-CH3
[-C-CH2-]
CD3
COO-CD3
α-methyl group motions of PMMA:
[-C-CD2-]
CH3
COO-CD3
>80%
Chain motions of PMMA:
partially deuterated

## Slide 16
Avoid multiple scattering within the sample.
Choose transmission to be about 90%!
Preparing your sample
Scattering depends
on your sample!
Transmission T through the
sample depends on thickness
and total scattering cross section
of your sample.
where
Choose correct thickness:

## Slide 17: Sample environment and cans
use neutron-transparent aluminum in beam.
Depending on your T range, choose
correct sample equipment
correct sample can
correct sealing agent

## Slide 18: Data Reduction
Incoming neutron flux dependence: How to remove it?
The factors to be considered to compare theoretical S(Q,ω) with experimental one:
The neutrons counted in a detector at a given solid angle will depend on:
Double differential cross section
Incoming neutron flux
Sample and beam size
Detector efficiency
Analyzer efficiency
Analyzer area seen by detectors
Scattering from the sample environment
Scattering from the empty can
Monitor normalization
Sample dependent
Vanadium standard
Empty can

## Slide 19: Data Reduction
For complete data reduction:
You have to collect sample data
You have to collect vanadium data for
normalization, detectors efficiency etc.
You have to collect background data
with empty sample can.
You have to collect data for
instrumental resolution.
How to convert measured S(Q,ω)
into absolute units?
Vanadium
standard
measured
true
resolution
Incoherent
scatterer

## Slide 20
Let’s start a backscattering experiment
From where?
Elastic scans
Make the monochromator and analyzers to reflect neutrons
with fixed wave vectors k.
Time scale?
1THz ~ 4.136meV
HWHM for BS: 0.4μeV
time scale ~ 2ns
Doppler frequency = 0

## Slide 21: Example:
Elastic scans can also be used to study phase
diagram of the system.
large hysteresis between cooling and
heating is indicative of undercooling!
Sharp transitions
1st order

## Slide 22
-[CH-CH2
-]
n
O-C
H
3
-
Tg=250K
CH3
rotations
α-relaxation
Elastic Scans
DSC
Example:

## Slide 23: Backscattering Spectroscopy
Need a dynamic window?
Turn ON the Doppler!
maximum vD will determine
the maximum energy transfer
Doppler dynamic
frequency range
15Hz ±11μeV
22Hz ±16μeV
50Hz ±35μeV
Accessible dynamic
range for HFBS
Velocity of
monochromator

## Slide 24: Example
PMMA and PMMA with PEO:
Relaxation time τ is inversely
proportional to the FWHM!
PMMA: Tg=400K
PEO: Tg=220K
PEO acts as a plasticizer
for high Tg component.