## Slide 1
Liquid Chromatography Introduction to Method Development
Lane C. Sander, Ph.D.
Chemical Sciences Division
Material Measurement Laboratory
National Institute of Standards and Technology1
1Contribution of the National Institute of Standards and Technology. Not subject to copyright. Certain commercial equipment, instruments, or materials are identified to specify adequately the experimental procedure. Such identification does not imply recommendation or endorsement by the National Institute of Standards and Technology, nor does it imply that the materials or equipment identified are the best available for the purpose.
Notes: The development of fit-for-purpose analytical methods is an essential component of chemical metrology. This presentation will provide guidance on approaches for establishing robust analytical methods, particularly with applications to liquid chromatography. Because nearly every measurement need represents unique challenges, this guidance is necessarily general in scope.

## Slide 2: Questions to consider before you start
What is the primary objective?
  qualitative analysis
  quantitative analysis
  detection of undesired substances / purity determination
  development of a new method
What are the challenges?
  Which components must be resolved?
  Is sensitivity an issue?
  Interferences?
  What level of accuracy and precision are required?
  Will different sample matrices be analyzed?
  Is analysis time a significant factor?
Notes: Before beginning, you should assess the objectives of the analysis. Depending on the goals, the details of the method used will vary significantly.
If the goals are identification or estimation of components, then a less rigorous method may be suitable.
However, if constituent levels must be determined with a specified level of uncertainty, more rigorous methods will probably be required.
An assessment of compound purity such as is required with reference standards is approached differently than the characterization of complex matrix based samples.
Finally, the development of the method itself may be the primary objective, such as may be required for Reference Measurement Procedures, with clinical applications.
As a part of this needs assessment, the following questions should be considered:
Which compounds are to be determined?
Are the measurements limited by the levels present? Will efforts be required to increase the limits of quantitation?
What interferences might be expected to originate from the reference standards or the sample matrix? How can potential interferences be addressed?
Simply stated, how good do the measurements need to be, to be fit for purpose? This of course will depend on the use of the data. Some user communities my find confidence intervals of 5 to 10% acceptable, but other users may need measurements less than 1%.
Will the method be used for a range of samples or a single example
Finally, is analysis time a significant factor that will further constrain method development?

## Slide 3: Sample Properties:
Matrix
  Bulk properties
  Complexity
  Processing considerations
Analytes
  number of compounds targeted for characterization
  stability, volatility, solubility, molecular weight
  neutral vs ionic
  concentration range
Notes: 1) The sample matrix is a primary consideration in method development. For example, entirely different challenges are presented by particulate based samples compared with tissues. These matrix differences influence the approaches applicable for sample preparation and instrumental analysis.
2) The physical and chemical properties of the analytes are also central to the development of analytical methods. Assess whether or not individual analytes are the primary focus, or if compound classes are of more interest. If a range of compounds is of interest, it may be expedient to report only a subset of the compounds identified, for example, those that are easily quantified. Stability, volatility, and solubility are important considerations in selecting the separation technology. If the analytes are amenable to gas chromatography, this technique may be more a more suitable choice than liquid chromatography due to the inherent high efficiency of the technique. Analyte polarity is another key property that influences method development. In general, polar or ionic compounds are more suited to liquid phase separations whereas nonpolar compounds are often suited to both LC and GC methods. The limit of quantitation is influenced by the efficiency of the separation, but more importantly, it is influenced by the instrumentation used for detection.

## Slide 4: General Considerations
Method Selection
  LC vs GC?
  Detection?
  MS vs MS/MS; mass resolution; mass analyzer technology?
Sample processing
Basis for quantitation
Miscellaneous Considerations
  Sample quantity and availability
  Availability of reference standards and internal standards
  Stability
Notes: Once these basic questions have been answered, an instrumental method can be selected. At this point, it should be possible to select preferred separation and detection approaches. Most methods now utilize mass spectrometry for detection. Additional consideration must be given to the choice of a mass analyzer technology (such as quadrupole or ion trap mass filters). More detail on aspects of mass spectrometry are presented separately.
Sample preparation can be as simple as no preparation, for samples that are directly injected, to extraction and fractionation for two dimensional processing.
When planning an analysis, attention must be given to the basis of quantitation. Three general approaches are common: external standard calibration, calibration with the use of internal standards, and standard addition. External standard calibration requires more skill and care in sample processing, since all manipulations must be quantitative in nature. This approach is most useful for solutions or samples that require minimal processing. The use of an internal standard is recommended for the analysis of complex matrices. With this approach, a known quantity of one or more compounds is added to samples and calibrants, and calculations are based relative to this compound. Losses due to processing are largely compensated, since at least in theory, these losses will affect both the analytes and internal standards equally. With the standard addition approach, known quantities of the analyte are added to the sample to provide calibration and analysis steps. Each of these topics are discussed in detail in a separate presentation on quantitation.
A few practical aspects should also be considered in developing an analytical method. Trace level analyses can sometimes be enhanced by processing larger sample volumes, but this is of course dependent on the availability of the sample. The determination of specific analytes typically requires available reference standards for calibration. If reference standards are not available, calibration is sometimes based on the use of related compounds, such as isomers or other surrogates, but traceability is broken by this approach.

## Slide 5: First Steps…
The best way to start is to carry out a thorough literature search of existing methods
Notes: By far, the best use of initial efforts in developing a method is to perform a thorough literature search of existing methods. Even if none of the methods are suitable, information gained will guide future method development.

## Slide 6: First Steps…
Compare existing methods:
  Will an existing method solve your measurement need?
  Is the method practical and robust? Validated?
  Are there problems with the method?
  What are the reported specifications (LOD, linearity, accuracy, precision)
  Is the method applicable to the target sample matrix?
Evaluate potential method(s)
  Prepare a solution containing the constituents of interest
  Assess resolution and peak shape for target analytes
  Prepare a sample from a typical sample matrix
  Assess interferences
Consider improvements to method
  New technologies (e.g., core shell columns)
  Modifications to conditions
  Use simple solutions of analytes initially, then more complex matrix based samples
Evaluate method performance
  Control samples
Notes: Once candidate methods are identified, these can be evaluated for suitability to the measurement need. Method quality and suitability is often evident, based on such factors as chromatographic peak shape and resolution, the degree of method validation, and method complexity.
After broadly assessing existing methods, the top candidates can be studied in more depth by implementing the instrumental portion of the method with a solution of the target analytes. By starting with a test solution, peak shape and resolution can be easily evaluated, and progress in the assessment is facilitated. If analysis of the test solution is encouraging, a representative complex matrix sample can be extracted and analyzed. At this point, attention should be directed to any indication of unresolved interferences.
It may be possible to easily implement improvements to a published method by using technological advances. Core shell columns have become widely available only in the last few years, and this technology offers significantly enhanced improvements in separation efficiency compared with conventional porous silica, with few disadvantages. The challenge in modifying existing methods through column substitution is in the selection of columns with similar selectivity. For guidelines on how to do this, view the presentation on column selection.
As a final step in the assessment of method suitability, evaluate performance using a control sample, such as a reference material. In the absence of a suitable control, a complex matrix can be spiked with known quantities of the analytes under study, and analyzed. This does not provide all of the evidence that might results from the analysis of a reference material, but it at least will indicate if the system calibration is under control.

## Slide 7: General Considerations
For nonpolar or polar organic molecules, consider a reversed-phase separation
  Most LC separations are carried out on C18 columns
  water/methanol or water/acetonitrile systems preferred
  water/acetonitrile: lower viscosity → lower pressure,
  water/methanol: higher backpressure; different selectivity (hydrogen bonding)
  water/THF: sometimes problematic, peroxide inhibitor has high background, unstable or sloping baseline
Isocratic separations are preferred to gradient separations
Notes: The majority of analyses performed by liquid chromatography are based on reversed-phase separations. Reversed-phase liquid chromatography is applicable to an extremely broad range of measurement needs simply by altering environmental conditions. Most separations are carried out on C18 columns. Hundreds of different C18 columns are available from different manufacturers, and no two columns are identical, but similarities do exist. Knowledge of the details of column properties can greatly assist method development. Aspects of LC column technology are treated in a separate presentation. As another generalization, most methods utilize aqueous mixtures of methanol or acetonitrile. These solvents offer numerous advantages that include UV transparency, low viscosity, moderate toxicity, and easy availability. Differences in selectivity result with the use of different solvents, and these differences can sometimes be attributed to the presence or absence of hydrogen bonding, pi-pi interactions, and of course, solvent polarity. Tetrahydrofuran is specified with some methods, but this solvent can be problematic since additives to inhibit peroxide formation contribute to absorbance detector background within the low UV. Additives and buffers should only be included if essential to the method.
In general, isocratic separations are preferred to gradient separations, when separations can be achieved within 20 to 30 minutes. Gradient separations are indicated for the separation of mixtures with compounds that differ widely in polarity.

## Slide 8: New Method Development
Start as simply as possible
Initial Choices
  Column: 4.6 mm x 150 mm C18 superficially porous “core shell” column
  Mobile phase: water/acetonitrile
  Flow rate: 1 mL/min
  Sample (mixture of reference standards): volume < 10 µL, mass < 0.1 mg
  Additives: none, unless indicated
    Buffer: 25 mmol/L potassium phosphate 2 < pH < 3 (use only if samples ionize: for acidic analytes, work below pKa of samples)
    Other additives: amines, ion-pairing reagents (don’t use initially unless there is a reason)
  Gradient elution: water to acetonitrile over 20 minutes
Evaluate separation
Notes: Start as simply as possible, and add complexity only as needed.
1. If a method must be developed without prior knowledge to direct the development, the following guidelines may be useful.
Select a 4.6 mm x 15 cm C18 column that utilizes core shell technology
Use a water / acetonitrile mobile phase at a flow rate of 1 mL/min.
Prepare a test mixture consisting of the analytes that are to be determined, at a level easily detectable, perhaps 0.1 mg/g
If the target analytes ionize, buffer the mobile phase so that the separation is based on resolution of neutral species. But remember that most columns have an upper pH limit of about 7.5.
Not withstanding the recommendation to use isocratic elution when possible, run the test mixture with gradient elution from water to acetonitrile, over 20 minutes.
2. Careful study of this separation will guide future development, whether or not it is isocratic or gradient elution.

## Slide 9: Initial Conditions
Gradient approach
  Start with a linear gradient; aqueous to organic over 20 minutes
  Adjust program to achieve separation
  Avoid complex multistep programs
Isocratic approach
  Start with strong solvent conditions and work to weaker conditions, or
  Guess based on literature, sample properties, and experience, or
  Estimate starting composition from gradient run
Notes: After carrying out the initial broad gradient run, decide if the separation can be performed by isocratic elution. Isocratic elution is suitable for the determination of analytes with similar polarity. Estimate the composition from the gradient run or based on other means, and make a trial separation by isocratic elution. To save time, it is better to estimate the trial mobile phase composition with stronger solvent conditions and then gradually decrease the strength of the mobile phase in subsequent attempts.

## Slide 10: System Volume
Different from the column dead volume Can be thought of as the chromatograph dead volume
If the mobile phase composition is changed at the chromatograph, how long does it take for the change to occur at the column?
To understand gradient separations (i.e., the gradient composition at the column at any given time) it is necessary to measure the system volume.
Notes: To evaluate the mobile phase composition that is delivered to the column during gradient elution, it is necessary to measure the system volume of the liquid chromatograph. The system volume is different from the column dead volume, and it can be thought of as the chromatograph dead volume.
For example, consider when a change is made in the composition of the mobile phase. How long will it take for the change to occur at the inlet of the column? The system volume is the volume of the tubing and other components prior to the connection with the column. This volume delays changes in mobile phase composition by a measureable, and usually significant, amount.
To understand gradient separations (i.e., the gradient composition at the column at any given time) it is necessary to measure the system volume. The importance of system volume increases with low flow rates, since long periods may be required to change compositions.

## Slide 11: Measuring System Volume
solvent A = methanol or acetonitrile
solvent B = 1 L methanol or acetonitrile + 1 mL acetone
remove the column, connect the lines with a union
run gradient from 100% A to 100% B over 10 minutes at 1 mL/min
detector response as a function of time looks like this:
Important: detector must not be saturated at 100 % B.
Determine the time for 50% response
Calculate gradient time / 2
System volume = A - B
In this example, system volume = 7 – 10/2 = 2 mL
time (min)
Notes: The system volume is a property of the instrumentation, and unless the hardware is modified, the value will always be the same. To determine the system volume, fill one of the solvent containers with a few hundred mL’s of methanol or acetonitrile, and add 1 mL of acetone. The concentration of acetone should not be so great as to saturate the detector when the mixture is eluted. Remove the column and connect the lines with a union. Flush the solvent line with this solution and program a 10 minute gradient from methanol (or acetonitrile) to the dilute acetone solution. Start a run and acquire the detector response data.
The resulting curve will look similar to the figure. In this example, a deflection in the baseline begins at approximately 2 minutes, which indicates that the gradient is delayed by about 2 mL. Similarly, the curve plateaus at about 12 minutes, which is consistent with the 2 mL delay. The system volume can be determined more precisely by measuring the time that corresponds to 50% detector response. This also corresponds to the half way point of the gradient program, at 5 minutes. Subtracting 7 – 5 gives the result of 2 mL.

## Slide 12: System Volume
example:
  4 components
  Gradient 0-100% in 10 min
  1 mL/min
  System volume = 3 mL
1
1
1
0
0
0
2
2
2
3
3
3
4
4
4
5
5
5
6
6
6
7
7
7
8
8
8
9
9
9
10
10
10
gradient
0 -100% in 10 min
40% isocratic
70% isocratic
Retention time (min)
Retention time (min)
Retention time (min)
Retention time (min)
Notes: This slide illustrates the concept of system volume for a hypothetical instrument with a system volume of 3 mL. In the example, a 10 minute gradient from solvent A to solvent B is used. The resulting separation of four compounds is shown at the top right. If the composition of the gradient is incorrectly estimated at 7 minutes to be 70% B, and this value is used isocratically, the solvent strength is much too high to resolve the compounds. See the chromatogram on the bottom, right. But if the system volume is taken into consideration, the composition of the mobile phase at 7 minutes is at 40% of the gradient time, which corresponds to 40% B. A corresponding isocratic separation is shown at the middle, right. This is a better estimate of the appropriate composition for an isocratic method.

## Slide 13: Isocratic vs Gradient Elution
Isocratic elution is most commonly implemented for the analysis of only one or two components
Gradient elution facilitates the separation of components that differ in polarity; primary advantage is reduction in analysis time
Gradient Elution
25% to 100%
aqueous:organic
over 10 minutes
1
0
2
3
4
5
6
7
8
9
10
Isocratic Elution
5
0
10
15
20
90%
80%
70%
60%
Retention time (min)
Retention time (min)
Notes: The following series of separations illustrate the differences between isocratic elution and gradient elution on the spacing of mixture components.
Isocratic elution is most applicable to mixtures of compounds with similar polarity. When large differences in polarity exist among the analytes, either some of the analytes are compressed at the beginning of the separation, or later eluting analytes require excessive time for elution.
Gradient elution is better for mixtures with large differences in polarity. Retention reproducibility is enhanced for isocratic separations, and column re-equilibration steps are not required, since the composition is kept constant.

## Slide 14: Isocratic vs Gradient Elution
Isocratic
  stable system, more rugged
  no equilibration time required between runs
  higher precision (often)
  mixed solvents can be prepared gravimetrically for demanding applications
Gradient
  faster separations of dissimilar components
Notes: Isocratic and gradient elution methods each have advantages and disadvantages. In both cases, proper function of the instrumentation must be maintained for best results.
Isocratic separations may result in a more stable and rugged methods. Since the conditions are constant, no equilibration time is required between runs. This stability often results in higher reproducibility and better precision. For demanding isocratic elution based applications, solvent mixtures can be prepared gravimetrically rather than instrumentally by the liquid chromatograph. Such extreme care is usually not warranted, but this approach is sometimes used for certification of reference materials.
Gradient separations provide faster separations of dissimilar compounds, but lack the advantages of isocratic methods. Under certain conditions, gradient elution may provide altered selectivity and increased peak capacity.

## Slide 15: Separation Goals
Separation time: < 30 min (k’ range 0.5 to 15 isocratic) Often this is compromised to achieve needed resolution
Pressure: < 150 bar desirable Don’t use > 250 bar unless very special circumstances (pump limit 350 bar to 400 bar)
For qualitative analysis complete separation may not be required (example: PTH amino acid analysis for protein sequencing)
Baseline resolution: Rs > 1.5
0.78
0.97
1.57
1.97
A:B = 1:1
A:B=10:1
Notes: This slide lists a few somewhat arbitrary guidelines for an analytical separation. Baseline resolution is achieved for Rs > about 1.5. Integration of components with this degree of resolution is facilitated, and biases from coeluting interferences are minimized.
Rapid methods are always preferred, as long as accuracy and precision goals are maintained. Methods less than about 30 minutes are practical for many applications, but in high throughput environments, separations are often much faster.
Avoid high system backpressures as much as possible. Changes in backpressure occur as filters become restricted, and it is better to have additional pressure capacity so that system shutdown does not occur unexpectedly in the middle of a crucial analysis.
Finally, when considering the development of an analytical method, the analyst should consider possible alternatives to chromatographic separation. Selective detection may provide a different pathway to elimination of interferences.

## Slide 16: Achieving Separation Goals
What if the desired separation cannot be achieved?
  try a different solvent system water/methanol vs. water/ACN vs. water/THF
  try a different column
    monomeric C18 vs polymeric C18 (isomers)
    endcapped vs nonendcapped C18 (polar compounds)
    pentafluorophenyl, phenyl (alternate selectivity)
    C30 (carotenoids)
    PYE column (pi interactions)
    CN, NH2
Notes: When separations are not adequate, a few parameters can be varied to achieve better results. Ultimately, this part of method development is largely trial and error, but experience and knowledge will help to guide the development.
The easiest thing to try is to alter the mobile phase composition. Significant changes in selectivity may result by switching from methanol to acetonitrile, in combination with water. Methanol provides hydrogen bonding functionality, whereas acetonitrile does not. If additional solvent strength is needed, tetrahydrofuran can be used, and this solvent may also offer different selectivity. MTBE is not miscible with water, so the range of compositions is limited, and THF often is packaged to include stabilizers to inhibit peroxide formation, that contribute to the background signal with absorbance detection. Less commonly, ethanol and acetone are used as mobile phase modifiers, and selectivity changes can also be anticipated with their use. It should be noted that ethanol-water mixtures exhibit increased backpressure due to increased viscosity, and acetone exhibits significant UV absorption, which makes its use more suitable for methods with MS detection or detection within the visible spectrum.
Another useful approach is try different columns. A few broad generalizations can be made that can guide selection of columns to be considered.
The designation of C18 column is generic in nature, and is not very useful unless more details are provided since the properties of C18 columns differ widely. Before attempting an independent method development from scratch, consult application notes of the column manufacturers to see if a separation already exits. Also, search the literature for existing methods. One quality of C18 columns does correlate with performance. Columns prepared with polymeric surface modification chemistry are more selective to the molecular shape of the solute than are monomeric columns. Such columns offer advantages for the separation of isomer mixtures like polycyclic aromatic hydrocarbons, and carotenoids.
Column polarity and silanol accessibility are controlled by secondary surface modification reactions called endcapping. It is probably safe to generalize that most C18 columns are endcapped or otherwise deactivated. Basic solutes that exhibit peak tailing may be better separated with columns that are well deactivated. Other compounds may actually depend on column activity for retention and selectivity.
The pentafluorophenyl column is a general purpose reversed phase column that offers different selectivity based on pi-pi interactions with the electron deficient fluorinated phenyl ring. This retention mechanism is sometimes referred to as charge transfer. Some separations are similar to those obtained with a C18 column, but some difficult to separate compound analogs, such as 25(OH)D and 3-epi-25(OH)D are separated at baseline resolution using this column.
The C30 column was developed originally at NIST for applications involving carotenoids isomers. This column is usually prepared with polymeric surface modification, and the column exhibits enhanced shape selectivity.
The pyrene column, or PYE column for short, is a specialty column prepared by surface modification with pyrene. The retention mechanism is probably similar to the PFP column, except that the pyrene ring is electron rich.
Cyano and amino columns are prepared with a propyl chain that is terminated with the functional group. These phases exhibit some reversed phase behavior, but with very limited retention. Selectivity differences are not predictable, but a few very good methods use these columns. The amino column is routinely used in normal phase mode to fractionate environmental contaminants prior to analytical analysis.

## Slide 17: Achieving Separation Goals (continued)
What if the desired separation cannot be achieved?
  Change column temperature
  Better precision results from column temperature control
  Selectivity can change significantly with temperature
    shape recognition
    polar/ionic interaction mechanisms
  Retention decreases with increasing temperature
  This can be useful to reduce run time. If necessary, retention can be adjusted by changing mobile phase conditions
Notes: If changes in mobile phase composition and column selection do not adequately address separation goals, changes in column temperature may provide additional control over selectivity.
Usually, the effect of column temperature on selectivity is modest, but for cases that involve shape selectivity or ionizable solutes, the effect can be quite significant.
Control of column temperature results in better overall precision since retention is made more reproducible. Retention decreases with increasing temperature, but the absolute retention can be compensated by altering the strength of the mobile phase.

## Slide 18: Achieving Separation Goals: Efficiency
What if the desired separation cannot be achieved?
  Use a more efficient column
    Use longer columns
    Use smaller particles
    Use coupled columns
    Use core shell columns
  This is usually a last resort
    higher backpressures
    not as rugged
    only marginal improvement
  To double resolution, need 4x length
  Ultimately, pressure and/or time limits
  the practical column length
Notes: Minor improvements in separations can be achieved by increasing column efficiency.
Resolution is proportional to the square root of the number of theoretical plates, so to improve resolution by a factor of 2, a four fold increase in efficiency is needed.
Efficiency can be increased by using longer columns (such as a 25 cm length column instead of a 10 cm length column), and smaller diameter particulate substrates (for example 3 um particles instead of 5 um particles). Efficiency can also be increased by using two or more columns joined end to end. A relatively recent innovation in column technology is referred to as core shell columns. These columns use microparticulate silica with a solid core and a porous outer layer to improve diffusion processes. With the exception of the latter approach, system pressure is increased with efficiency.
In general, separations should be carried out at pressures as low as possible, to improve method durability, and extend column lifetime. The increased efficiency exhibited by core shell particles is attributed to the reduced pathlength for solute diffusion in and out of the pore network. Core shell columns offer performance advantages over other types of columns, with no significant disadvantages, and this technology can be recommended for the development of new methods.

## Slide 19: Other Approaches
  Consider a more selective detector
    Fluorescence detection
    Absorbance detection (visible wavelength range)
    Mass spectrometry
    Compounds may not need to be resolved if coeluting species have very different responses
  Use multidimensional approaches
    orthogonal chromatographic steps
    sample cleanup
    fractionation
Notes: If the analytes cannot be adequately resolved in the separation domain, selective detection can provide another avenue towards achieving unbiased measurement results.
Virtually all modern analytical methods use some form of selective detection in combination with the chromatographic method. For compounds that exhibit native fluorescence properties, fluorescence detectors offer selective and sensitive detection suitable for complex matrices. Because relatively few analytes exhibit spectra with absorbance in the visible spectrum, when applicable, absorbance detection from 400 nm to 700 nm can also provide selection detection. By far the most generally applicable form of selective detection is mass spectrometry. Most analytes can be ionized under conditions compatible with mass spectrometry. The resulting charged mass fragments are typically characteristic of individual compounds, and provide selective detection even in the presence of coeluting interferences. Additional selectivity results from the use of multiple mass analyzers, such as with triple quadrupole MS.
Finally, multidimensional approaches are useful to simplify the sample prior to the analysis. Solid phase extraction cleanup methods can be developed to either retain either the analytes of interest or the interferences. More discrimination is possible by using an orthogonal separation method to divide the sample into fractions, each of which can be analyzed separately. This approach is commonly used with environmental samples to separate complex extracts into fractions that contain pesticides, polychlorinated biphenyl congeners, and polycyclic aromatic hydrocarbons (or PAHs for short). PAHs can in turn be further divided into fractions based on ring number.
In general, better measurements will result when baseline resolution of analytes is achieved, either through sample cleanup, chromatographic separation, selective detection, or a combination of these steps.

## Slide 20: www.nist.gov