**Preface**

The 2021 revision of Handbook 105-2 includes the following changes.

1. There is no provision for retroactivity. All volumetric standards used for field enforcement of weights and measures need to comply with the specifications and tolerances published in this handbook.
2. Decision rules regarding maximum permissible errors (plus/minus tolerances) were changed. The calibration uncertainty must simply be less than the applicable maximum permissible errors (rather than less than one-third). Statements of conformance to tolerances require the absolute calibration value plus the uncertainty to be within tolerance and the user must consider correction values and uncertainties in use.
3. All reference documents were reviewed and updated with current and applicable references for volumetric calibrations. ASTM International, OIML, ISO, and EURAMET standards and procedures have been reviewed in consideration of international standardization.
4. Formatting was updated for Word and accessibility requirements.

The 1996 revision of Handbook 105-2 included the following changes since it was published in 1971:

1. References to the National Bureau of Standards (NBS) were replaced by the National Institute of Standards and Technology (NIST).
2. Reference to and incorporation of international standards (such as those of the International Organization for Legal Metrology, OIML) and national industry standards (such as those of the American Society for Testing and Materials, ASTM) were made where possible.
3. The addition of references to direct the user to publications that will assist with effective use of field standards.

Additionally, the process for updating the publication was changed to include the following:

1. Conversion of the previous handbook to electronic media to allow future changes to be incorporated in a timelier manner.
2. Organized peer review to ensure incorporation of the latest technology and viewpoints of technical experts.

Note regarding units of measure:

The SI unit of volume is the cubic decimeter (dm3) or the cubic centimeter (cm3). The Twelfth General (International) Conference on Weights and Measures redefined the litre [herein spelled liter] as a "special name for the cubic decimeter," but agreed to permit the continuance of the terms liter (L) and milliliter (mL), except in association with measurements of the highest precision. For volumetric glassware, the difference between the old and new meanings of liter is negligible. Therefore, either mL or cm3 may be marked on flasks and glassware covered in this handbook.

Since commercial applications in the United States use units other than SI or other accepted metric units, this document references other common units in current use.

Certain commercial entities, equipment, or materials may be identified in this document in order to describe an experimental procedure or concept adequately. Such identification is not intended to imply recommendation or endorsement by the National Institute of Standards and Technology, nor is it intended to imply that the entities, materials, or equipment are necessarily the best available for the purpose.

Acknowledgments:

1971: This Handbook was initially written by Blayne C. Keysar of the National Bureau of Standards (now NIST).

1996: Special thanks regarding the 1996 edition were given to Kelleen Moody (Larson), metrologist with the State of Arizona, and to Karl Herken, metrologist with the State of Kansas, for their assistance with review of reference materials, evaluation of comments submitted during peer review, and for typing and editing the document in WordPerfect as the first electronic format. Thanks were also given to numerous metrologists (of both State and industry laboratories) for their technical review of several drafts.

2021: The updated draft was prepared by Georgia L. Harris and reviewed by interested parties. Specific inputs regarding applicable decision rules were submitted by multiple laboratories, with initial feedback on an updated draft provided by Jeremy Nading, metrologist with the State of Oklahoma, Nicholas Santini, metrologist with the State of Michigan, and Micheal Hicks, NIST Office of Weights and Measures

 **Table of Contents**

**Preface i**

Scope 1

Retroactivity 1

Safety Considerations 1

Reference Documents 2

Terminology 2

Specifications 3

Nominal Values 3

Metric 3

Customary (Inch-Pound) 3

Reference Temperature 3

Material and Annealing 4

Physical Properties 4

General 4

Cylindrical Design 4

Construction 4

Base 4

Neck 4

Lines, Graduations and Inscriptions 4

Line Widths 4

Line Orientation 5

Line Construction 5

Graduation Pattern 5

Subdivision Lines 5

Line Color 5

Graduation and Nominal Graduation Lines 5

Graduation and Nominal Graduation Inscriptions 5

Scale Divisions, Metric 5

Scale Divisions, Customary 6

Identification 6

Tolerances (Maximum Permissible Error) 6

Verification Requirements 6

Legal Requirements 6

Traceability 6

Calibration Reports 6

Initial and Periodic Verification 7

Test Methods and References 7

Uncertainties 7

Legal Applications 7

Sources of Variation 8

Abbreviations 8

**Figure 1. 100 ml glass flask.** 10

**Figure 2. 50 ml graduated cylinder.** 10

**Figure 3. ½ pint glass flask.** 11

**Figure 4. 2 fl oz graduated cylinder.** 11

**Notes** 12

**SPECIFICATIONS AND TOLERANCES**

**FOR REFERENCE STANDARD AND FIELD STANDARD**

**WEIGHTS AND MEASURES**

2. Specifications and Tolerances for

Field Standard Measuring Flasks

These specifications and tolerances are recommended as minimum requirements for standards used by State and local weights and measures officials in quantity determination of liquid commodities.

Key Words: field standard measuring flasks; flasks; graduated cylinders; volumetric standards; volumetric specifications and tolerances; weights and measures.

INTRODUCTION

Field standard volumetric flasks and graduated cylinders as described herein are intended to be used by weights and measures officials, manufacturers and distributors of liquid products, research and testing laboratories, and others concerned with accurate measurements of the volume of liquids. Use of these volumetric standards at all appropriate levels of manufacture, distribution, and weights and measures inspection will promote accuracy and uniformity in commerce.

1. Scope
	1. "Field Standard" Classification

This handbook classifies volumetric flasks with graduated necks and graduated cylinders (see Figures 1 to 4) for legal metrology applications as "field standards." Tolerances provided in Tables 2 and 3 are intended to permit use of field standards during normal testing operations as standards having nominal values. Where calibration uncertainties exceed tolerance limits published in this handbook, the reported volumetric values and associated uncertainties must be used and considered in field applications when evaluating packages for compliance. Nominal values may be used only when the absolute value of the volume plus the calibration uncertainty are within maximum permissible errors. For specifications and tolerances for glassware used in laboratory applications, and where smaller tolerances are needed, see the ASTM references listed in Section 2.

* 1. Retroactivity

This handbook applies to all flasks and graduated cylinders for use as field standards (i.e., new and in-use field standard volumetric flasks and graduated cylinders, those intended for replacement of flasks already in use, and to new flasks to be acquired as supplementary standards.)

* 1. Safety Considerations

The accuracy and repeatability of field standards is critically dependent upon cleanliness. Chemicals used in the cleaning process should be evaluated for safety in use and for appropriate disposal methods by reviewing Safety Data Sheets (SDS, formerly Material Safety Data Sheets).

Volumetric glassware should not be emptied by holding onto the neck alone. The bottom of the flask should always be supported to prevent glassware breakage and possible injury.

1. Reference Documents
	1. OIML[[1]](#endnote-2)
		1. R 87, Quantity of product in prepackages, 2016.
		2. R 43, Standard graduated glass flasks for verification officers, 1981.
	2. NIST.[[2]](#endnote-3)
		1. NIST Handbook 44, Specifications, Tolerances, and Other Technical Requirements for Weighing and Measuring Devices, see current edition.
		2. NIST Handbook 133, Checking the Net Contents of Packaged Goods, see current edition.
		3. NISTIR 7383, 2019, Selected Procedures for Volumetric Calibrations.
		4. NIST Special Publication 811, Guide for the Use of the International System of Units (SI), NOTE: The BIPM SI Brochure (9th Edition) was published on May 20, 2019. The NIST SP 330 2019 version has been published and reflects the changes incorporated in the BIPM SI Brochure (9th Edition). This publication, Special Publication 811, has not yet been updated to reflect the changes in the SI that came into effect on May 20, 2019 (World Metrology Day).
	3. ASTM.[[3]](#endnote-4)
		1. E288-10 (2017), Standard Specification for Laboratory Glass Volumetric Flasks.
		2. E438-92 (2018) Standard Specification for Glasses in Laboratory Apparatus.
		3. E542-21, Standard Practice for Gravimetric Calibration of Laboratory Volumetric Instruments.
		4. E694-18, Standard Specification for Laboratory Glass Volumetric Apparatus.
		5. E1272-02(2019) Standard Specification for Laboratory Glass Graduated Cylinders.

1. Terminology

*Borosilicate glass.* A glass of a low cubical coefficient of thermal expansion used for most precision laboratory glassware and known by such trade names[[4]](#endnote-5) as Kimax (KG-33) or Pyrex. See ASTM E 438 for Type I glass specifications.

*Capacity, contained.* The volume of water which the flask contains at the reference temperature when filled to its nominal graduation line and is designated "to contain" or << In >>. The neck graduations of a “to contain” measure represent the volume of liquid in the measure, not the volume of liquid that can be poured from the measure. A “to contain” measure must always be cleaned and dried between successive uses of the measure for purposes of accuracy.

*Capacity, delivered.* The volume of water which the flask delivers at the reference temperature from the specified graduation line when emptied gradually with a 30 s (± 5 s) pour and a 10 s drain while held at a 10 ° to 15 ° angle from vertical. These flasks are designated "to deliver" or << Ex >>. If a flask or graduated cylinder is to be used in a wet condition, it must be calibrated "to deliver." The advantage of using a “to deliver” measure is that the measure does not have to be dried between uses.

*Capacity, nominal.* The nominal capacity of a field standard flask or graduated cylinder is the volume used to designate the flask or cylinder at a given reference temperature (as calibrated to the specified reference temperature of 20 °C).

*Graduation lines.* Numbered lines which extend for at least 3/4 of the flask neck or cylinder circumference.

*Nominal graduation line.* A line extending completely around the flask neck or cylinder circumference (see 4.4.2) that indicates the nominal capacity, and which must be in a contrasting color to the other lines (see 4.5.6).

*Soda-lime glass.* A glass of medium cubical coefficient of thermal expansion often used in field standard flasks and graduated cylinders. See ASTM E 438 for Type II glass specifications.

*Subdivision graduation lines.* Unnumbered intermediate graduation lines between nominal and other graduation lines.

1. Specifications
	1. Nominal Values

A set of field standard flasks and graduated cylinders comprises several flasks. Nominal capacities in the series are chosen in accordance with applications and regulations.

* + 1. Metric

A set of metric field standard flasks consists of a 50 mL graduated cylinder, a 100 mL graduated cylinder or flask, and one each 250 mL, 500 mL, 1 L, 2 L graduated neck flasks. This example does not preclude other sets or nominal sizes.

* + 1. Customary (Inch-Pound)

A set of U.S. customary field standard flasks and graduated cylinders consists of a 2 fl oz graduated cylinder, and one each 1 gill, 0.5 pint, 1 pint, 1 quart, 0.5 gallon, and 1 gallon graduated neck flasks. This example does not preclude other sets or nominal sizes.

* 1. Reference Temperature

The temperature at which the flask or cylinder is intended to contain or deliver a volume equivalent to its nominal capacity, shall be 20 °C (68°F).

Application note: When a product that is normally refrigerated is being tested, a packager is *given the benefit of doubt* in determining volume, unless temperature corrections are made, due to the cubical thermal coefficient of expansion for the glass and for the product; the extent is dependent on whether product is tested at its specified reference temperature, its storage temperature or at a normal indoor environment (i.e., 20 °C.)

* 1. Material and Annealing

A field standard flask or graduated cylinder shall be made of transparent, well annealed clear glass with suitable thermal and chemical properties (such as ASTM Type II, soda-lime or Type I, borosilicate glass). The flask shall be free from chips, cracks, stones, and other visible defects that detract from the appearance or use of the flask. It is particularly important that the graduated portion of the flask or cylinder be free from obvious defects.

* 1. Physical Properties
		1. General

The design shall conform to the general configuration shown in the Figures 1 to 4. The inscriptions and graduations shall be placed in the same relationship to each other and to the position on the flasks as shown.

* + 1. Cylindrical Design

The neck and body of flasks and graduated cylinders must be cylindrical. Any cross section taken in a plane perpendicular to the vertical axis shall be circular.

* + 1. Construction

The shape of a field standard flask or graduated cylinder shall permit complete emptying and thorough cleaning.

* + 1. Base

A standard graduated flask must maintain a stable vertical position without rocking when placed with its base on a flat level surface. Each flask shall be designed with an attached base that is perpendicular to the vertical axis for stability. (A hexagonal base is typically used for maximum stability, but the 1-gallon flask is often designed in such a way that eliminates the need for a base.)

* + 1. Neck

The neck of the flask must be cylindrical. The top edge of the neck shall have a smooth finish and a small flange. The height of the graduated portion of a graduated cylinder shall be at least five times the inside diameter.

* 1. Lines, Graduations and Inscriptions
		1. Line Widths

Graduation and subdivision lines shall be distinct, permanent and of uniform thickness not to exceed 0.3 mm.

* + 1. Line Orientation

Graduation lines shall be perpendicular to the vertical axis of the base of the flask.

* + 1. Line Construction

Graduation lines shall be applied by one of the following methods: etched and filled with a permanent pigment; application of a stain fixed into the glass without etching; or application of an enamel fused onto the glass without etching.

* + 1. Graduation Pattern

There should be no evidence of irregular spacing between graduation lines. The graduation lines shall extend completely around the neck. Due to the difficulty in extending stained or enameled lines completely around the neck, a gap of 4 mm at the closure, or meeting point, is permitted. This gap must be approximately 90 ° from the line of vision when the flask is viewed from the front so as not to interfere with reading a meniscus (see Figures 1 to 4).

* + 1. Subdivision Lines

Subdivision lines shall be uniform and extend at least halfway around the neck.

* + 1. Line Color

If a pigment or enamel is used for graduation lines, the nominal volume line shall be of a contrasting color.

* + 1. Graduation and Nominal Graduation Lines

Graduation lines on flasks with graduated necks shall be marked above and below the nominal line as shown in Volumetric Scale Range in Table 2.

Graduation lines on graduated cylinders shall only be marked beneath the nominal line and shall be marked with Numbered and Minimum Graduations as shown in Table 1.

Subdivision lines on graduated cylinders shall be omitted between the base and the first main graduated line. This will eliminate reading near the base where it is difficult to read and of questionable accuracy. (Striation often occurs in the glass in this area during manufacture when the base is joined to the cylindrical portion.)

* + 1. Graduation and Nominal Graduation Inscriptions

Each nominal capacity line shall be labeled with the appropriate volume and units. The numbers and letters indicating nominal capacity and main graduation capacities shall be placed immediately above the line to which they refer (see Figures 1 to 4 for examples).

* + 1. Scale Divisions, Metric

The scale divisions on a metric graduated cylinder shall be divided into milliliters, and labeling shall so indicate with the appropriate abbreviation: mL. Each subdivision shall comply with Table 2.

* + 1. Scale Divisions, Customary

The scale divisions on a U.S. customary graduated cylinder shall be divided into fluid drams, and labeling shall so indicate with the appropriate abbreviation: fl dr. Each subdivision shall comply with Table 3.

* + 1. Identification

Each field standard flask or graduated cylinder shall be permanently and legibly marked with the following:

1) the manufacturer's name or trademark;

2) serial or identification number;

3) clear identification of "to deliver" or "to contain" use (“to deliver” flasks with a proper *wet down* are typically used for commodity inspection);

4) nominal capacity and appropriate units;

5) reference temperature for calibration; and

6) drain time (e.g., 10 s).

On U.S. customary standard glassware all letters except unit abbreviations are to be in upper case. (See Figures 1 to 4.)

1. Tolerances (Maximum Permissible Error)

The difference between the actual volume and the indicated volume at the prescribed reference temperature (20 °C) shall not be greater than that shown for Tolerance at Nominal or Partial Capacity in Table 2. As noted in Section 8.1 for legal metrology applications, the calibration uncertainty shall be less than the applicable tolerances specified in Table 2 or 3 and the absolute value of the calibrated measurement result plus the calibration uncertainty shall be within the stated tolerance to make conformance statements.

1. Verification Requirements
	1. Legal Requirements

When field standard flasks and graduated cylinders are used for commercial applications they must be inspected for damage prior to each use and prior to calibration. Glassware must be evaluated for conformance to this document and be calibrated by a NIST recognized or accredited laboratory, using appropriate calibration methods such as those shown in Section 7, with calibration certificates provided that include conformity assessment statements, calibration measurement results, and calibration uncertainty values that are less than tolerances listed in this document.

* 1. Metrological Traceability

Field standards shall be calibrated by a recognized or accredited laboratory. Field standard measurement results used for legal metrology shall be traceable to the International System of Units (SI) with associated supporting documentation.

* 1. Calibration Certificates

Acceptable accuracy and traceability to national or international standards shall be documented in a calibration certificate using accepted calibrated methods as shown in Section 7. Calibration values, uncertainties, and tolerance status must be noted on the calibration certificate for the user's evaluation and the certificate must include a conformity assessment statements to ensure that standards fully meet both specifications and tolerances needed to support legal metrology applications. Volumetric standards that do not comply with the specifications and tolerances should not be used for legal weights and measures enforcement activities. Labeled nominal values may be used when the absolute volumetric value plus the calibration uncertainty are within maximum permissible errors; alternatively, the reported calibrated volume and associated uncertainty must be considered by the end user.

* 1. Initial and Periodic Verification

Field standard flasks and graduated cylinders must undergo initial verification for conformance to these specifications and tolerances. Field standards must be inspected prior to use and verified periodically as prescribed by regulation; the frequency of periodic inspection and/or verification depends upon usage but should not exceed 5 years. Glass flasks and graduated cylinders generally do not change capacity values during this period unless damaged. Intermediate comparisons against other standards may be performed occasionally to detect standards in need of recalibration.

1. Test Methods and References

Initial verification to determine whether field standard flasks and graduated cylinders meet applicable tolerances is performed by calibration using accepted volume transfer or gravimetric calibration procedures. The uncertainty of the calibration must be less than the applicable tolerances shown in this publication. If commercial measurements are to be made, there may be additional test/verification requirements (dependent on the jurisdiction in which the field standards will be used). Referenced methods are as follows:

* 1. NISTIR 7383, SOP 14

SOP 14, Standard Operating Procedure for Gravimetric Calibration of Volumetric Ware Using an Electronic Balance, as published in NISTIR 7383, Selected Procedures for Volumetric Calibrations, 2019.

* 1. ASTM E542

Standard Practice for Gravimetric Calibration of Laboratory Volumetric Instruments, 2021.

* 1. ISO 4787

Laboratory glass and plastic ware — Volumetric instruments — Methods for testing of capacity and for use, 2021.

* 1. EURAMET, Calibration Guide

Guidelines on the Determination of Uncertainty in Gravimetric Volume Calibration, EURAMET Calibration Guide No. 19, Version 3.0 (09/2018).

1. Uncertainties
	1. Legal Applications

Uncertainties of the calibration must be evaluated according to the Guide to the Expression of Uncertainties in Measurements,[[5]](#endnote-6) to ensure compliance with recognition and accreditation requirements. The uncertainty for volume calibrations must be less than the tolerances published in this documentary standard. Evaluation of applicable requirements when used to support NIST Handbook 133 package testing or NIST Handbook 44 device testing must be considered by the field officials and service companies as applicable (e.g., especially in situations where the absolute value of the volume plus the calibration uncertainty exceeds the applicable tolerance in this handbook.)

* 1. Sources of Variation

For volumes such as those listed in this handbook, the largest sources of uncertainty are inaccurate reading of the meniscus (see NISTIR 7383, Good Measurement Practice, GMP 3 and ASTM E694 Annex A.1), cleanliness of the container, and proper technique when emptying and draining the flask or cylinder. A 30 s (± 5 s) pour followed by a 10 s drain, with the measure held at between a 10 ° and 15 ° angle from vertical is required during calibration and during application. The uncertainties reported by the laboratory do not reflect the uncertainty in field applications. Field application uncertainties include the same type of factors and are additive to those reported by the laboratory with additional consideration needed for viscosity and opacity of liquids being measured. For microliter volumes measured without a meniscus, the largest source of uncertainty and potential bias may be evaporation; if time of use is limited, evaporation is generally not a concern in weights and measures field applications.

1. Abbreviations

**Table 1. Volume Abbreviations**

|  |  |
| --- | --- |
| Measurement Unit | Abbreviation |
| Cubic Centimeter | cm3 |
| Cubic Decimeter | dm3 |
| Cubic Millimeter | mm3 |
| Fluid Dram | fl dr |
| Fluid Ounce | fl oz |
| Liter | L |
| Milliliter | mL |
| Pint | pt |
| Gallon | gal |

**Table 2. Tolerances (maximum permissible error) for flasks and cylinders (SI)**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Capacity at 20 °CmL | Toleranceat Nominal Capacity± mL | Tolerance at Partial Capacity± mL | Volumetric Scale Range (above and below nominal)mL | NumberingmL | Minimum Graduations(subdivisions)mL |
| 50 cylinder | 0.25 | 0.25 | See 4.5.7 | 5.00 or 10.00 | 1.00 |
| 100 flask | 0.20 | 0.06 | 4.00 | 2.00 | 0.50 |
| 250 | 0.30 | 0.10 | 6.00 | 5.00 | 0.50 |
| 500 | 0.50 | 0.15 | 10.00 | 5.00 | 1.00 |
| 1 000 | 0.80 | 0.22 | 20.00 | 5.00or 10.00 | 1.00 |
| 2 000 | 1.20 | 0.33 | 30.00 | 10.00 | 2.00 |

**Table 3. Tolerances (maximum permissible error) for flasks and cylinders (U.S. Customary)**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Capacity at 20 °C(units)(Metric equivalents provided) | Toleranceat Nominal Capacity± mL | Tolerance at Partial Capacity± mL | Volumetric Scale Range (above and below nominal)fl dr  | Numbering fl dr | Minimum Graduations(subdivisions)fl dr |
| 2 fl oz cylinder(59 mL) | 0.25 | 0.25 | See 4.5.7 | 1.00 or 2.00 | 0.50 |
| 1 Gill flask(118 mL) | 0.20 | 0.10 | 0.50 | 0.50 | 0.25 |
| 0.5 pt(236 mL) | 0.30 | 0.10 | 1.00 | 0.50 | 0.25 |
| 1 pt(473 mL) | 0.40 | 0.15 | 2.00 | 1.00 | 0.50 |
| 1 qt(946 mL) | 0.70 | 0.30 | 4.00 | 2.00 | 1.00 |
| 0.5 gal(1 892 mL) | 1.00 | 0.30 | 6.00 | 2.00 | 1.00 |
| 1 gal(3 785 mL) | 1.20 | 0.30 | 8.00 | 2.00 | 1.00 |

For volumetric measures less than 50 mL, full capacity tolerances do not apply. For these volumetric measures, apply 0.10 mL to individual graduations. For volumes greater than 3 785 mL (1 gal) apply ± 0.02 percent of nominal capacity for tolerances at full capacity and ± 0.3 percent of the minimum graduation for tolerances for individual graduations. For a capacity intermediate between two capacities listed above, the tolerances prescribed for the lower capacity shall be applied.

**Figure 1. 100 ml glass flask.**



**Figure 2. 50 ml graduated cylinder.**



**Figure 3. ½ pint glass flask.**

****

**Figure 4. 2 fl oz graduated cylinder.**



 **Notes**

1. . OIML, L'Organisation Internationale Metrologie Legale, Bureau International De Metrologie Legale, 11 Rue Turgot, 75009 Paris, France. [↑](#endnote-ref-2)
2. . NIST, National Institute of Standards and Technology, Gaithersburg, MD 20899. [↑](#endnote-ref-3)
3. . ASTM International, formerly American Society for Testing and Materials, 100 Barr Harbor Drive, West Conshohocken, PA 19428-2959. [↑](#endnote-ref-4)
4. . Trade names as used in this handbook do not imply recommendation or endorsement by the National Institute of Standards and Technology (NIST). [↑](#endnote-ref-5)
5. . ISO, International Organization for Standardization, Geneva, Switzerland. [↑](#endnote-ref-6)