## Slide 1: How to collect data with a neutron reflectometer
David Hoogerheide
NIST Center for Neutron Research
CHRNS Virtual Neutron School
February 2021
useful
^

## Slide 2: The goal…
…and the tool
J. Dura / NIST
Neutron reflectivity pattern
Real space model
SiO2
Si
air

## Slide 3: The basic idea
Interpretation of reflectivity data is model-based. It is best to know in advance what are the competing models*; then you can prepare your sample and measure your data to maximize your chances of discriminating among the various real-space models.
* this might involve a continuum of structures

## Slide 4: Elements of a reflectivity curve
oscillation period  film thickness
maximum Q  real space resolution
Scattering background  max Q
Statistics  parameter uncertainty

## Slide 5: Experimental considerations
Sample
  Film contrast with substrate
  Film thickness
  Film roughness
  Film uniformity
  Environment
  Sample size
Instrument
  Pre-sample collimation (angular resolution)
  Post-sample collimation (background)
Data collection
  Maximum Q
  Number of Q steps
  Counting time

## Slide 6: Effect of film contrast
When nSLDs match, there is no longer any contrast
Sample design matters!
Recommended contrast:
It’s always worth calculating your expected SLD profile and neutron reflectivity pattern ahead of time

## Slide 7: Effect of film thickness: thick films
Thick films require a tight point spacing
Angular resolution of the measurement starts to matter
Tradeoff with intensity

## Slide 8: Effect of film thickness: thin films
Resolution of the measurement is determined by maximum Q value
High Q measurement takes a long time; reflectivity is very small.
Changes in layer thickness are much more easily detected

## Slide 9: Effect of film thickness: surface roughness
Roughness causes reflectivity to drop off faster
Limits effective Qmax and hence real space resolution
Can also cause artifacts (or offspecular signal!) above about 10 Å rms
It is worth optimizing sample quality for low roughness

## Slide 10: Effect of background
Background limits maximum Q
Comes mainly from interaction of the main beam with the sample and its environment
Background + roughness limits Q range of the measurement

## Slide 11: How to reduce background
Tight post-sample collimation
Use the right materials for substrate and sample environment: avoid hydrogen (plastic)

## Slide 12: How to measure background
Move detector or sample so (offspecular condition)
Interpolate to estimate background under the specular signal (discussed further in data reduction module)
Gains no more than a factor of 10 in reflectivity

## Slide 13: Effect of sample size
Large samples
  constant footprint
  intensity changes with angle
Small samples
  constant beam size
  beam footprint changes with angle

## Slide 14: Mounting the sample
What sample environment do I need?
Do I have the physical hardware/sample mounts to put the sample on the sample table
Talk to your friendly local instrument scientist!

## Slide 15: How to align a reflectometry sample
Eyeball it!
Rough alignment
  Close collimation so detector can handle the direct beam
  Move to a direct beam geometry (sample and detector angles 0)
  Find the sample by scanning sample translation
  Scan sample rotation
  Iterate a couple of times
Fine alignment
  Move to a reflection geometry (sample angle to , detector angle to ).
  Iterate among sample rotation, translation, and tilt. The intensity will be peaked around the optimal values.

## Slide 16: Direct beam alignment
Sample translation
Sample rotation

## Slide 17: Sample alignment video game
From Brian Maranville: https://www.ncnr.nist.gov/instruments/magik/d3-science-v4/demos/reflectometer_alignment_sim_three.html
  Move the sample rotation and translation to direct the beam into the detector.
  In practice, after a rough alignment, this is done by:
    Positioning the detector at an angle
    Optimizing the sample rotation and translation to maximize the count rate in the detector. This is done, along with a sample tilt, in an iterative fashion
    Define the sample rotation angle to be at this position.

## Slide 18: Reflectometry suite @ NCNR

## Slide 19: Monochromatic reflectometers
 MAGIK
Neutron source is a graphite monochromator
  Very good (≈ 1%) wavelength resolution
  Relatively low flux (uses only about 1% of the neutrons in the guide)

## Slide 20: CANDOR polychromatic reflectometer
NIST/Dan Neumann

## Slide 21: CANDOR polychromatic reflectometer
NIST/Dan Neumann
Current speed improvement: 20x

## Slide 22: Questions?
Contact me:
David Hoogerheide
david.hoogerheide@nist.gov
https://www.nist.gov/ncnr/neutron-instruments/reflectometry
Or any of your other friendly instrument scientists!