## Slide 1
Non-Nulling 3D Velocity Testing for Improving the
Uncertainty of CEMS Flow Measurements
(Evolutions in Measurement Technology)
Presented by Aaron Johnson, NIST
Project Team: Iosif Shinder, Jim Filla, Joey Boyd, Vladimir Khromchenko, John Wright, Michael Moldover, and Rodney Bryant
Tom Martz (EPRI), Sylvio Cardoso (EPRI), Richard M. Himes (EPRI)
Source Evaluation Society
September 22, 2021
Notes:   Today I’ll discuss the work that the NIST has done to improve stack emission measurements from coal-fired powerplants
  We have developed what we call “The Non-Nulling Method” to make stack flow measurements faster and more accurate
As I will explain it is important that the NIST method be
  accurate
  Fast
  Robust
In order to be adopted by the electric power industry
  The work has been done
  by the FMG
  NIST staff from other divisions
  In cooperation with EPRI – the Electric power research institute

## Slide 2: Executive Summary: What NIST has done?
Developed Non-Nulling Method to facilitate fast, accurate, and robust flow RATAs
EPA sanctioned pitot probes can be classified into two categories
Probes with limited accuracy, but fast and robust measurements
Probe with good accuracy, but slow measurements and difficult to use
Probe with good accuracy that provides fast, robust measurements
NIST designed a Non-Nulling Hemispherical Probe
Also developed a Non-Nulling Method for Spherical Probe
S-probe (Method 2)
Probe Type 1
Spherical Probe
D = 2.5 cm
D = 2.5 cm
Eliminates need to rotate probe
Hemispherical Probe
D = 2.5 cm
Non-Nulling Probe
D = 2.5 cm
Non-Nulling is robust, provides 1-3% accuracy & is 4x faster than Nulling Method
Probe Type 2
Non-
Nulling Method
D = 2.5 cm
Rotate Probe
P3
P2
P2 = P3
Notes:   Probe has 2 large pressure ports that are not probe to being clogged with particulates or that Point velocity
  Before we jump into the details of the Non-Nulling Method and Stack Testing let me give an executive summary of what we have NIST has accomplished.
  First a few facts
  Contrast
  Nulling can several minutes to find null condition
  Non-Nulling Hemishperical probe measures velocity vector in 5 sec data collection
  Hemispherical has larger pressure ports (6 mm) that
  Based on results of stack testing

## Slide 3: Outline
Explain the Source of Accuracy for the Non-Nulling Method
Review Accuracy of EPA Nulling Method 2F for the Spherical Probe (We assess accuracy of the Non-Nulling Method to EPA Method 2F)
Results of flow RATAs from 4 Coal-Fired Power Plants
Technology Transfer of the Non-Nulling Method to Industry
Conclusions and Future Work

## Slide 4: Source of Accuracy of the Non-Nulling Method
Non-Nulling & Nulling probes calibrated in NIST WIND TUNNEL
  S-Probe
  Spherical & Prism Probes (Nulling)
  Spherical Probe (Non-Nulling)
  Hemispherical probe (Non-Nulling)
Validated accuracy of probes by performing flow RATA in NIST Stack Simulator
Collaborate with EPRI to apply NIST Non-Nulling Method to industrial stacks
Generates smokestack-like flows (i.e., yaw angle 35° & & pitch angle 10°) in a 4 ft Test Section
Facility is equipped with a calibrated reference flow meter with an uncertainty of 0.8 %
Results: Spherical Probe (1 to 2%) & Hemispherical probe (1 to 3%)
NIST WIND TUNNEL
Stack Simulator Facility
EPA’s
Protocols
NIST’s
Protocols
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review
Notes:   With this in mind we have implemented a 3 Tier Strategy to improve the flow RATA and flow CEMSs
  Major

## Slide 5
How Do We Calibrate Non-Nulling Probes?
Calibrations are performed in NIST’s Wind Tunnel
  Air Speeds (Uꝏ) 16 ft/s to 100 ft/s measured with LDA Working Std.
  Grid generated Turbulence Intensities 0 % to 10 %
Automated Traverse System Orients Probe
  Range of pitch angles (α) from -20 ⁰ to 20 ⁰
  Range of yaw angles (ꞵ) from -60 ⁰ to 60 ⁰
Calibration Data
  Requires 500 wind tunnel data points
  Data is used to develop curve fits for Dynamic Press. (Pdyn), Pitch (α), Yaw (ꞵ), and Static Press. (Pstatic)
Who will perform the 500 point non-nulling calibration for each probe?
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review
b

xBF
yBF
LDA
Sensing
Volume
y
x
Notes: Efforts are underway to reduce CO2 emissions from coal fired power plants

## Slide 6
NIST Developed a Universal Curve Fit for the Hemispherical Probe
We procured 6 Hemispherical probes
the 6 probes were manufactured to the same geometrical dimensions
3 different manufacturers, 2 probes made in machine shop, and 4 probes 3D printed
A single universal curve fit characterized the performance of all 6 probes
Dynamic pressure (Pdyn ): ± 2 % to 4 % Accuracy
Yaw Null Angle (βnull): accuracy ± 1º to 2º
Pitch Angle (α): accuracy ± 1º to 2º
Commercial airspeed labs will confirm the accuracy of universal fit at specified pitch, yaw, and velocity
Work is underway to determine geometric tolerances required to ensure probes are interchangeable (i.e., the universal fit applies to all probes within geometric tolerances)
universal fit can be adjusted for probes slightly out of tolerance based on wind tunnel calibrations (preliminary results indicate a 20 to 30 point calibration)
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 7
How do you apply Non-Nulling Method in Stack Test?
Measurements are made with P1 aligned with stack axis (Yaw angle = 0)
(No need to rotate probe to find P2 – P3 = 0 at each traverse point)
Dynamic Pressure, Yaw, Pitch, and Static pressure
are curve fits to wind tunnel data fit as a function of differential pressures
Differential pressures are measured for 5 sec at each traverse point and NIST curve fits are used to determine axial velocity (Note density is NOT determined by Non-Nulling)
Hemispherical
Probe
- Static pressure determined
Differential Pressure Notation:
P12 = P1 – P2
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review
Notes: Cos(Yaw) & Cos(Pitch)

## Slide 8
How do we assess the accuracy of Non-Nulling Method in the field?
Compare Non-Nulling to Method 2F for Spherical Probe
Benefits
Nulling Method 2F is an established protocol and trusted by
industry and EPA
EPRI independently validated the accuracy of the EPA Spherical
Probe Method 2F (1% accuracy) under laboratory conditions (1990’s)
NIST validated the EPA Nulling Method 2F in its stack simulator (1 to 2 % accuracy)
Comparison Nulling vs. Non-Nulling will determine time savings
Challenges
NIST research shows significant errors result if probe is
inadequately nulled, P2 ≠ P3
2) Plugging problems: Lab tests showed that the spherical
probe’s ports “plug” in < 2 min in saturated air loaded with
water droplets (Same test did not plug hemispherical probe)
Plugged Pressure Port
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 9
Determine calibration factors F1 and F2 (measured in wind tunnel at P23 = 0)
1)
No.
Review Method 2F & Assess Magnitude of Nulling Errors
Are nulling errors significant?
Can you “stack” the deck (i.e., determine sign of P23 to yield smaller Pdyn)?
Accuracy of theoretically corrected Pdyn is ± 2% for probe yaw angles within ±10º of βnull
Chart: 
M2F & Theory
Standard M2F
20% Error!
YES!
± 2%
We applied this correction to Method 2F results to correct for nulling errors
Stack Application of Method 2F
Iterate to find βnull yaw null angle (i.e., P23 = 0)
2)
With probe at βnull measure F1 = P45/P12 and use F1 calibration curve to find pitch angle (α)
With probe at βnull measure P12 and quantify F2 using pitch angle (α) and F2 calibration curve
3)
(Method 2F does not specify tolerance P23 = 0)
βprobe - βnull, [deg]
Pdyn / PdynCAL, [%]
Errors in Pdyn as a function of yaw angle
6.6 % spread
(perfect null)
Pdyn / PdynCAL, [%]
Low Pdyn
P23 > 0
How do we fix the problem?
YES!
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 10
-200
-150
-100
-50
0
50
100
150
200
1425
1445
1465
1485
Chart: 
Do the proposed nulling error corrections work in the field?
0.8
0.9
1.0
1.1
1.2
0
100
200
300
400
500
600
time [min]
Non-Nulling gives accurate velocities for all values of P23 providing fast measurement results
P23, [Pa]
54
58
62
66
70
74
0
1
2
3
4
5
6
Iterations to optimize P23=0
1st
2nd
3rd
4th
5th
Time, [sec]
Axial Velocity, [ft/s]
Avg
Actual Nulling Data
M2F Iterations 1 through 4 (P23 ≠ 0)
M2F & Theory
Non-Nulling Algorithm
Method 2F (M2F) Iteration 5
16.5
17.7
18.9
20.1
21.3
22.6
Axial Velocity, [m/s]
P23 = 0
(Ideal Nulling)
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review
Notes:   The correction for nulling errors was based on existing Method 2F & a theoretical correction based on flow around a sphere
  Wind tunnel data showed the correction worked well
  Now we address How well the correction works in the field
  To show the effectiveness of the correction we show a plot of at a single traverse point during a flow RATA where the stack tester attempts to null a spherical probe
  X-aixs is time
  Y-axis P23

## Slide 11: NIST Developed Data Acquisition Box and Software
5 DP transducers
4-way solenoid valves
Reference
Pressure
Manifold
Laptop
Equipment housed in a compact water-proof box
Key Features
  5 Industrial grade differential pressures transducers
  Five 4-way solenoid valves for rapid response, to protect dP sensors during pressure purge
  Weather-proof Laptop
Data Collection & Processing
  LabVIEW data acquisition program
    - Normal sampling rate is 10 Hz for velocity measurements
    - Sample rate is 5 kHz for measuring turbulence
Non-Nulling Algorithm
  Velocity, Yaw, & Pitch calculated using measured dP’s
  Static pressure is also calculated
  Algorithm to troubleshoot for plugging of pressure ports
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 12
Stack testers use high pressure air (100 psi) to periodically purge pitot probe pressure lines
prior testing indicated that purging alone is insufficient to eliminate plugged probes when Nulling Method is implemented using spherical probe
NIST developed software diagnose plugging events
LabVIEW software to identify plugged pressure ports
  Screen shot shows standard deviation of 20 pressure signals (4 spherical probes each with 5 pressure ports) as probes are traversed from point to point
  If stdev falls below empirically determined threshold stack tester is notified
Picture of Plugged/Blocked Pressure Ports
Stdev 20 dP measurements
Empirical threshold to signal plugged data
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 13: Results from 4 Coal-Fired Power Plants
Probe
(Method)
Wet Stack 1
Low
Load
(7 m/s)
High
Load
(16 m/s)
Sphere
(EPA Nulling)
Sphere
(Non-Nulling)
N/A
0.0 %
-2.5 %
-0.2 %
Low
Load
(18 m/s)
High
Load
(27 m/s)
Wet Stack 2
0 %
0 %
0.5 %
0.9 %
0 %
1.2 %
Low
Load
(14 m/s)
High
Load
(20 m/s)
Dry Stack 3
0 %
0 %
0.1 %
0.5 %
0 %
2.6 %
High
Load
(13 m/s)
Wet Stack 4
0* %
1.2 %
* Not corrected for stack tester nulling errors
Taken as
Reference
0* %
0* %
0.9%
Hemisphere
(Non-Nulling)
N/A
Taken as
Reference
Excellent agreement between Spherical Probe Nulling (EPA) and NIST Non-Nulling Method
Spherical probe: Maximum deviation < 1.2 %
Hemispherical probe: Maximum deviation < 2.6 %
Data indicated that Non-Nulling Method was 4x faster than Nulling Method
No plugging events indicated for Hemispherical probe
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 14: Technology Transfer to Industry in Process
Recent test was performed by stack testing company who operated both NIST pressure boxes and software & his own equipment
A Small Business Innovation Research (SBIR) awarded to a stack testing company to develop 1) non-nulling pressure boxes, 2) an automated traverse system, and 3) software to compute velocity
Future testing will use the newly developed industrial non-nulling equipment
Stack tester operating both NIST
and his own equipment
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 15: Conclusions
NIST Non-Nulling Method
At least 4x faster than nulling methods: EPA Nulling Methods 2G and 2F
Hemispherical probe resistant to plugging problems
Validated Non-Nulling Method in NIST’s Stack Simulator and in 4 Coal-Fired Powerplant Stacks
Accuracy of Non-Nulling velocity measurements 1 to 3 % (Compared to Improved Method 2F)
Complete research on interchangeability and calibration of hemispherical probe shapes
Write the draft of the test method for the Non-Nulling Method to propose to the EPA
Work with EPRI, stack industry, and EPA to make Non-Nulling an approved EPA Method
Future Work
Conclusions
Field Results
Accuracy of NN Method
Technology Transfer
Method 2F Review

## Slide 16: Questions?