## Software Technical Assessment
Software Technical Assessment
Software Description | WAMRF-003 (Ver. 7), SOP 31 Scale Plate Calibration Workbook | Approval Date | 2018-09-26
Assessed By | Your Name | Approved By | Your Name
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | All worksheet are appropriately named and easy to understand
| There are instructions for use | Pass | Instructions worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | The workbook opens to the instructions worksheet
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank, light blue for non-blank cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Specified on Uncertainty worksheet
| The number of digits to be rounded to is specified | N/A
| The user is warned/notified whenever data-entry fields are left blank | Pass | In data validation cells only
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry fields on the workbook template are blank. Instructions tell user to break links before saving after a calibration is performed
| The software opens at the right location within the file | Pass | Opens to the Instruction worksheet
| Unused fields/cells are locked | Pass | Unused cells are locked and worksheets are protected
| Rows/columns that the operator need not see are hidden | N/A | No hidden cells or worksheets
| Unused sheets are removed | Pass | No unused worksheets
| Worksheets are named appropriately | Pass | Worksheets are appropriately named
B. Mathematical Specification | The correct SOP is used | Pass | SOP 31
| The formulae and methods chosen from that SOP are specified | Pass | Volume Transfer method
| Sources and references for formulae are specified | Pass | As specified in SOP
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | Pass | For Volume II calibrations
C. Code review | The formulae in the fields exactly match the SOP | Pass | Formulae on calculation worksheet match SOP
| Repeated calculations appropriately reference the correct cells | Pass | Calculation worksheet
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | Pass | Tested using previous example
| Rounding is done at the appropriate locations in the file | Pass | k value rounded to 3 digits
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | Pass | No issues encountered
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | Pass | Formatted correctly throughout workbook
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | Pass | Where appropriate
E. Component Testing | Each macro used is functional | N/A | Mo macros
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | Pass | No issues encountered
| Plotted graphs are accurate | Pass | Graphs on uncertainty worksheet display properly
| Worksheets/reports print properly, if needed to | Pass | Worksheets formatted for printing as necessary
| Conditional (color and non-color) formatting is functional | Pass | On data entry and calculation worksheets
F. Numerical Reference | Look-up tables and lists match the latest calibration report. | Pass | Linked tables are validated and up to date
| Uncertainties match the latest Scope | Pass | Test on uncertainty worksheet will not allow an uncertainty that is better than the latest published scope
| Values that reference another workbook or spreadsheet are dated | Pass | Linked references have been validated and are up to date
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc.) is correct | Pass | All embedded data has been validated
| The evaluation of the embedded data is dated and documented | Pass | Records keep in Software Validation folder on computer hard drive
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | Pass | Evaluated and filed
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | N/A | No hand calculations were done
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Cells are locked and worksheets are password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | State Metrologist and Program Manager only have access
| Files are backed up automatically | Pass | Agency IT backs up servers weekly
| Additional back-up is available at alternate facilities | Pass | Agency IT keeps an additional back up at a remote location
| Files on network drives cannot be accidentally deleted | Pass | Laboratory has its own directory on agency server.

## DCT (WAMRF-003)
Document Changes Tracking Form
Document: | WAMRF-003, SOP 31 Scale Plate Calibration workbook
Date | Page Number | Paragraph Number | Comments | AR Number
2017-03-21 |  |  | Deleted Revision 08 and created Version 1 from new NIST SOP 31 | 253
2017-06-01 |  | SOP 31 worksheet | Inserted updated SOP 31 | 257
2017-07-25 |  |  | Changed Report Number to Certificate Number on Data Entry worksheet Cell E26, changed range name TestNo to CertNo, and some minor reformatting on Uncertainty worksheet. | 264
2017-11-09 |  | Uncertainty worksheet | Corrected uncertainty calculations to match SOP 31, equations 5 and 6 on the Uncertainty Analysis worksheet. | 275
2017-11-09 |  | Uncertainty worksheet | Added SOP 31 equations 5 and 6 under the Specify Equation(s) on the Uncertainty Analysis worksheet.
2017-11-09 |  | Data Entry worksheet | On the Data Entry worksheet changed the word 'points' to intervals' in cell G19 and named cell J19 'Intervals'. | 275
2017-11-17 |  | Data Entry worksheet | Changed conditional formatting for the Nominal Volume Unit, cell A13 on the Data Entry worksheet. | 275
2017-11-17 |  | Data Entry worksheet | Changed cell I30 on the Data Entry worksheet from "Drops" to "Total Drops from Standard" and added a note with an example. | 275
2017-12-05 |  |  | Corrected the ∆sr equations to match equation 3 in SOP 31. | 275
2017-12-05 |  |  | Corrected rounding so enough digits will be displayed. | 275
2017-12-05 |  |  | Labeled mandatory data input for the equations and tests to work properly. | 275
2017-12-05 |  |  | Revised the inputs for the Volumetric Standard used. | 275
2017-12-05 |  |  | Revised the NSCF calculation so it is always visible and corrected rounding for the uncorrected error equation. | 275
2017-12-05 |  |  | Corrected NomVolUnit equations on Uncertainty Analysis worksheet. | 275
2018-09-26 |  |  | Replaced Documentation worksheet with the Document Changes Tracking form | 286
2018-09-26 |  |  | Embedded 2017 version of SOP 31. | 286
2025-03-06 |  |  | Corrected Uncertainty Analysis reference to "gal" from "Nominal Volume (gal)"

## Information
This workbook was developed by Dan Wright, Washington State Department of Agriculture, Weights and Measures Program, Metrology Laboratory. It follows the procedure documented in NIST SOP 31, and the specifications of NIST HB 105-3 and NIST HB 105-4. Georgia Harris and Val Miller, NIST OWM, have evaluated this spreadsheet and provided useful feedback in its development.
Instructions
No. | Description
1 | This workbook must be validated using your internal procedures before use and after any changes or modifications.
2 | The worksheets are password protected. Password: password
3 | This workbook compares the entered data against the specifications of NIST HB 105-3 and NIST HB 105-4.
Range Names
Name | Range
CertNo | ='Data Entry'!$J$1
CustomerID | ='Data Entry'!$L$5
DateReceived | ='Data Entry'!$K$11
Drops | ='Data Entry'!$L$30
Intervals | ='Data Entry'!$L$21
LPG_gal | =Lists!$M$4:$Q$10
LPG_L | =Lists!$M$13:$P$17
NC_Tol | ='Data Entry'!$H$42
NeckUnit.List | =Lists!$A$5:$C$8
NomVol | ='Data Entry'!$D$14
NomVolUnit | ='Data Entry'!$H$14
NSCF | ='Data Entry'!$H$66
NU_Tol | ='Data Entry'!$K$42
Prover_gal | =Lists!$E$4:$K$14
Prover_L | =Lists!$E$17:$I$28
RefTemp | =Lists!$S$3:$S$4
RefTempUnit | ='Data Entry'!$A$27
ScaleGradUnit | ='Data Entry'!$D$15
Spec | ='Data Entry'!$K$13
sr_f4 | ='Data Entry'!$E$46
sr_f6 | ='Data Entry'!$E$48
sr_i | ='Data Entry'!$E$42
Std_df | ='Data Entry'!$H$31
Std_k | ='Data Entry'!$E$31
StdU | ='Data Entry'!$C$31
StdVol | ='Data Entry'!$H$30
TestDate | ='Data Entry'!$D$24

## SOP 31
Note on A1: Unprotect the worksheet, then double click the document to open in Adobe Acrobat

## Data Entry
Data Entry | [7 empty] | Certificate Number:
NOTE: For equations and tests to work properly all entries marked with a ♦ are mandatory.
Customer Information
Company: | [7 empty] | Customer ID:
Address 1: | [7 empty] | Job Order Number:
Address 2: | [7 empty] | POC:
City, St Zip: | [7 empty] | Phone:
Artifact and Measurement Information
Description: | [7 empty] | Date Received:
Manufacture: | [7 empty] | Serial Number:
Model Number: | [7 empty] | Specification ♦:
Nominal Volume ♦: |  |  |  | Nominal Volume Unit ♦:
Scale Graduation Units ♦: |  |  |  | Reference Temperature ♦: |  |  |  |  | Temperature Scale ♦:
|  | Total Neck Scale Plate Volume () |  |  |  | ♦:
FALSE |  | Graduation / Subdivision |  |  |  | ♦:
FALSE |  | Upper Neck Inside Diameter (mm) |  |  |  | ♦:
FALSE |  | Gauge Tube ID (mm) |  |  |  | ♦:
#N/A |  | Graduation / Subdivision Line Width |  |  |  | ♦:
Choose the number of calibration intervals for the test ♦:
The following calculations assume nominal volume graduation is near midpoint of the neck scale
Test Date ♦: |  |  |  |  | Technician Name:
Volumetric Standard Used (2.3.4)
Item ID: |  |  | Description: | [5 empty] | Serial #:
Certificate #: |  |  |  |  | Cal. Date: |  |  |  | Cal. Due ♦:
Note: If Calibration Due Date Cell Turns RED Calibrate the Standard or Choose a Different Standard.
ID Code: |  |  | Delivered Volume @ (gal) |  |  | ♦: |  |  | Total Drops from Std. ♦:
Unc. (gal) | ♦: |  | k factor ♦: |  | df ♦: |  |  | Total Volume (gal):
Neck ID (mm) ♦: |  |  |  | Gauge Tube ID (mm) ♦: |  |  |  | Graduation Line Width (mm) ♦:
Neck Calibration and Uniformity Test Criteria
Neck Scale Tolerance Test (2.4.3)
"Rock" or "Bounce" Vessel after each addition of Volume
Initial Test
Volume Transfers |  |  | Volume Added (in³) | Scale Reading (sr) (in³) | Scale Deviation (Δsr) (in³) |  | Neck Calibration Tolerance ± |  | Scale Uniformity (Δsu) (in³) | Neck Uniformity Tolerance ±
Initial Scale Reading(sri) ♦
1st |  | ♦ | [7 empty] | #VALUE!
2nd |  | ♦
3rd |  | ♦
4th |  | ♦
5th |  | ♦
Final (srf) |  | ♦
Are there any failures in the Neck Uniformity Tolerance Test? ♦
Are there any failures in the Initial Neck Calibration Tolerance Test? ♦
Final Test With The NSCF Applied
Volume Transfers |  |  | Volume Added (in³) | Scale Reading (sr) (in³) | Scale Deviation (Δsr) (in³) |  | Neck Calibration Tolerance ±
Initial Scale Reading(sri)
1st
2nd
3rd
4th
5th
Final (srf)
Are there still failures in the Final Test with the NSCF applied? ♦
Neck Scale Correction Factor Calculation (2.4.3.11)
Neck Scale Correction Factor (NSCF) =
Uncorrected Systematic Error Calculation (4.2)
NOTE: This value is the mean of the scale deviations and may be included in the uncertainty budget for a test measure or prover (if less than the expanded uncertainty when all other components of the test measure or prover calibration uncertainties are combined) and may be included as a rectangular distribution.

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NIST SOP 31, Scale Plate Calibration for Volumetric Field Standards
Specify Equation(s)
Standard Volume at the Reference Temperature (Eqn. 1)
Best Starting Point for the Initial Reading (Eqn. 2)
where
| sri = | Initial scale reading
| NOMprover = | Nominal volume of prover
| VSn = | ‘To Deliver’ volume of water added for each interval of the neck scale calibration
| # calibration points | Total number of intervals making up the neck scale calibration
Accuracy of Neck Scale Plate at Each Interval (Eqn. 3)
where
| Δsrx = | Deviation between volume added and observed scale plate reading
| srx = | number assigned to each increasing neck scale calibration point; x = f for final calibration point
| sri = | Initial scale reading
| xVSn = | ‘To Deliver’ volume of water added for each interval of the neck scale calibration
Neck Scale Correction Factor (Eqn. 4)
where
| NSCF = | Neck scale correction factor
| Vw = | Total volume of water added to neck to increase the water level from sri to srf
| sri = | Initial scale reading
| srf = | Final scale reading
Calibration Uncertainty for Four Calibration Intervals Comprised of One Delivery from the Standard per Interval (Eqn. 5)
Calibration Uncertainty for Four Calibration Intervals Comprised of Two Deliveries from the Standard per Interval (Eqn. 6)
where
| un = | Combined Uncertainty
| uS = | Uncertainty of the standard
| umS = | Ability to read the Meniscus in S
| umX = | Ability to read the Meniscus in X
Estimate of Uncertainty for Reading Meniscus (GMP 3)
where
| π = | The value for pi
| r2 = | The radius of the inside diameter
| h = | The height of the line
Identify Components (4.1)
Uncertainty Component Description |  |  |  |  | Symbol | Source |  |  |  | Typical Distribution
Uncertainty of the standard |  |  |  |  | us | Calibration certificate divided by the coverage factor multiplied by the number of drops |  |  |  | Normal
The ability to read the meniscus in the standard |  |  |  |  | umS | GMP 3, page 4, Uncertainty and Error Analysis in Meniscus Readings |  |  |  | Triangular
The ability to read the meniscus in the unknown |  |  |  |  | umX | GMP 3, page 4, Uncertainty and Error Analysis in Meniscus Readings |  |  |  | Triangular
Uncertainty Calculations (Quantity, Convert, Combined, Expand & Report) (4.1)
Standard Meniscus (GMP 3) |  |  | Neck (mL) |  | umS = | 0 |  | Gauge Tube (mL) |  | umS = | 0
Unknown Meniscus (GMP 3) |  |  | Neck (mL) |  | umX = | 0 |  | Gauge Tube (mL) |  | umX = | 0
Component Description |  | Component Symbol | Estimated Uncertainty | Units | Reported Unc | Type | Distribution | Factor to Normalize | Standard Unc | % Contribution | Degrees of Freedom
Volume Standards |  | us | 0 |  | 0 | B | Normal | 1 | 0 |  | 0
Standard Meniscus |  | umS | 0 |  | 0 | B | Triangular | 2.4495 | 0 |  | 1
Unknown Meniscus |  | umX | 0 |  | 0 | B | Triangular | 2.4495 | 0 |  | 1
Combined U (uc) =
Effective Degrees of Freedom, Welch-Satterthwaite formula, GUM para. G.4.1, equation G2b (veff) =
k factor =
Expanded U (uc*k) =
Expanded U (uc*k) =
Values in
𝑉_𝑆𝑛=(𝑇𝑜𝑡𝑎𝑙 𝑛𝑒𝑐𝑘 𝑠𝑐𝑎𝑙𝑒 𝑝𝑙𝑎𝑡𝑒 𝑣𝑜𝑙𝑢𝑚𝑒)/(# 𝑐𝑎𝑙𝑖𝑏𝑟𝑎𝑡𝑖𝑜𝑛 𝑝𝑜𝑖𝑛𝑡𝑠)
〖𝑠𝑟〗_𝑖=〖𝑁𝑜𝑚〗_𝑝𝑟𝑜𝑣𝑒𝑟−(# 𝑐𝑎𝑙𝑖𝑏𝑟𝑎𝑡𝑖𝑜𝑛 𝑝𝑜𝑖𝑛𝑡𝑠)/2 𝑉_𝑆𝑛
∆〖𝑠𝑟〗_𝑥=〖𝑠𝑟〗_𝑥−(〖𝑠𝑟〗_𝑖+𝑉_𝑆𝑛 )
𝑁𝑆𝐶𝐹=𝑉_𝑤/((〖𝑠𝑟〗_𝑓−〖𝑠𝑟〗_𝑖 ) )
𝑢_𝑛=√((𝑢_𝑆×4)^2+4×(𝑢_𝑚𝑆^2 )+4×(𝑢_𝑚𝑋^2 ) )
𝑢_𝑛=√((𝑢_𝑆×8)^2+8×(𝑢_𝑚𝑆^2 )+8×(𝑢_𝑚𝑋^2 ) )
𝑉𝑜𝑙=𝜋𝑟^2 ℎ

## Lists
Data Validation & Lookup Lists |  |  | NIST HB 105-3 Table | [6 empty] | NIST HB 105-4 Table |  |  |  |  | Reference Temperature
Neck Uniformity Tolerance Vlookup List (NeckUni.List) |  |  | Capacity (gal) | Max Neck ID (mm) | Gauge Tube ID (mm) | Max Subdivision (gal) | Max Subdivision (in³) | Min Neck Vol (gal) | Min Neck Vol (in³) | Capacity (gal) | Max Neck ID (mm) | Max Division (gal) | Max Division (in³) | Total Neck Volume (gal) | Calibration Reference Temperature (ºF)
Scale Graduations | NIST HB 105-3 | NIST HB 105-4 | 1 | 57.15 | 12.7 | 0.001 | 0.25 | 0.015 | 3.5 | 20 | 152.4 | 0.01 | 2 | 0.8 | Calibration Reference Temperature (ºC)
in³ | Maximum Allowed Subdivision (in³) | Maximum Allowed Division (in³) | 5 | 76.2 | 12.7 | 0.002 | 0.5 | 0.075 | 17 | 25 | 152.4 | 0.01 | 2 | 1
gal | Maximum Allowed Subdivision (gal) | Maximum Allowed Division (gal) | 10 | 101.6 | 12.7 | 0.005 | 1 | 0.15 | 35 | 50 | 203.2 | 0.02 | 5 | 2
L | Maximum Allowed Subdivision (L) | Maximum Allowed Division (L) | 25 | 127 | 15.875 | 0.01 | 2 | 0.38 | 88 | 100 | 304.8 | 0.05 | 10 | 4
mL | Maximum Allowed Subdivision (mL) | Maximum Allowed Division (mL) | 50 | 127 | 15.875 | 0.01 | 2 | 0.75 | 173 | 200 | 457.2 | 0.1 | 20 | 8
|  |  | 100 | 177.8 | 15.875 | 0.02 | 5 | 1.5 | 347 | 300 | 457.2 | 0.1 | 20 | 12
|  |  | 200 | 254 | 15.875 | 0.05 | 10 | 3 | 693 | 500 | 457.2 | 0.1 | 20 | 20
|  |  | 500 | 431.8 | 15.875 | 0.1 | 20 | 7.5 | 1733
|  |  | 750 | 431.8 | 15.875 | 0.1 | 20 | 11 | 2610 | Capacity (L) | Max Neck ID (mm) | Max Division (mL) | Total Neck Volume (L)
|  |  | 1000 | 508 | 15.875 | 0.1 | 20 | 10 | 2310 | 100 | 127 | 50 | 4
|  |  | 1500 | 508 | 15.875 | 0.1 | 20 | 15 | 3465 | 200 | 152.4 | 100 | 8
[10 empty] | 500 | 304.8 | 200 | 20
|  |  | Capacity (L) | Max Neck ID (mm) | Gauge Tube ID (mm) | Max Subdivision (mL) | Min Neck Vol (L) |  |  | 1000 | 304.8 | 200 | 40
|  |  | 5 | 57.15 | 12.7 | 5 | 0.075 |  |  | 2000 | 457.2 | 500 | 80
|  |  | 10 | 76.2 | 12.7 | 5 | 0.15
|  |  | 20 | 101.6 | 12.7 | 10 | 0.3
|  |  | 50 | 127 | 12.7 | 25 | 0.75
|  |  | 100 | 127 | 15.875 | 25 | 1.5
|  |  | 200 | 177.8 | 15.875 | 50 | 3
|  |  | 500 | 254 | 15.875 | 100 | 7.5
|  |  | 1000 | 431.8 | 15.875 | 250 | 15
|  |  | 2000 | 431.8 | 15.875 | 500 | 30
|  |  | 3000 | 508 | 15.875 | 500 | 45
|  |  | 3800 | 508 | 15.875 | 500 | 38
|  |  | 5000 | 508 | 15.875 | 500 | 50