## Slide 1
Soxhlet Extractions
Lane C. Sander, Ph.D.
Chemical Sciences Division
Material Measurement Laboratory
National Institute of Standards and Technology1
1Contribution of the National Institute of Standards and Technology. Not subject to copyright. Certain commercial equipment, instruments, or materials are identified to specify adequately the experimental procedure. Such identification does not imply recommendation or endorsement by the National Institute of Standards and Technology, nor does it imply that the materials or equipment identified are the best available for the purpose.
Notes: Soxhlet extraction is one of several sample preparation methods that can be used to transfer targeted analytes from a sample matrix into a solvent, to permit subsequent analysis. This presentation will provide an introduction to the practice of soxhlet extractions, and will demonstrate the technique for a typical sample.

## Slide 2: Overview
Principle of operation: samples are repeatedly extracted with solvent regenerated by distillation through a periodic siphoning process
Advantages
  Classical method
  Automated repeated extraction
  High extraction yields
Disadvantages
  Lengthy process (typically 16 h to 20 h)
  Large solvent volumes
  Possibly sample degradation (oxidation, heat, light)
Notes: With Soxhlet extraction, samples are repeatedly extracted with solvent regenerated by distillation through a periodic siphoning process. Because pure solvent is used for each extraction cycle, the removal of the analytes from the sample matrix is thermodynamically favored compared with simply soaking the sample with a single aliquot of solvent for the same period of time.
There are a number of advantages with soxhlet extraction compared with other approaches. Soxhlet extraction is a classical method, and numerous published methods include this approach as part of the specified analytical procedure. Soxhlet extraction is an automated extraction process that is easier to carry out than combining extracts from successive re-extractions. Soxhlet extraction also often results in high extraction yields.
A few disadvantage can be noted. Extractions are often a lengthy process that can take from 16 to 20 hours unless a longer. Methods use large solvent volumes and there is the possibility of sample degradation though environmental exposure that can include oxidation, thermal degradation, and degradation from exposure to light.

## Slide 3: Apparatus
Notes: This figure illustrates the basic configuration of a Soxhlet extraction apparatus. The equipment consists of four components: the distillation flask and heating source, the Soxhlet extraction glassware, the sample thimble, and the condenser.
The extraction solvent is heated, and solvent vapors travel through the side arm of the Soxhlet glassware to condense in the condenser. The distilled solvent drips into the thimble containing the sample, and over a period of time the thimble fills with solvent. The solvent contacts the sample, and the extract flows out of the bottom of the thimble through a frit to fill the soxhlet chamber. When the level of the solution reaches the top of the siphon tube, the solution siphons into the distillation flask and the process repeats.

## Slide 4: Soxhlet thimbles
50 x 135 mm
275 mL
157 g tare
43 x 116 mm
168 mL
81 g tare
34 x 83 mm
75 mL
34 g tare
Notes: This photograph shows three different sizes of glass thimbles. For sample sizes approx. 1 g or larger, sample masses can be obtained by difference by weighing the thimble before and after addition of the sample. Better accuracy can be obtained for small samples by weighing the sample in a separate foil weighing container before and after transfer to the glass thimble. If this approach is used, it is important to be careful not to lose any sample during the transfer.

## Slide 5: Method Development
Solvent selection
  Polarity and polarizability
  Volatility
  Boiling point
  Pure solvents vs solvent mixtures
    Azeotropes
Extraction time
Notes: The choice of a suitable solvent is the single most important consideration in the development of a Soxhlet extraction method. Potential extraction solvents should provide large partition coefficients for the analyte class under study. The volatility of the solvent should be sufficient for the distillation process, with a boiling point that is compatible with samples and analytes.
It should be noted that for solvent mixtures, the composition of the vapor phase will differ from the composition of the liquid phase unless an azeotropic composition is utilized. Because the composition of a non-azeotropic solvent mixture changes with distillation, the extraction would occur with a variable solvent composition. In practice, this could result in less than optimal extraction conditions and lower extraction recovery.
The kinetics of the extraction favor longer solvent/sample contact times. A common soxhlet time of 16 to 24 hours will provide 50 to 80 extraction cycles and this is usually sufficient for maximum recovery.

## Slide 6: Dispersant
Distributes sample
Dries wet samples
Prevents plug formation
Retains sample within Soxhlet thimble
Types of Dispersants:
  Sodium Sulfate
  Diatomaceous earth (Hydromatrix®)
  Sand
Use of glass wool
Notes: An inert material is commonly used to distribute and dilute the sample, and to adsorb moisture from the sample and thus promote better contact with the extraction solvent. The dispersant and sample are mixed together within the Soxhlet thimble, and the combined matrix acts to prevent plug formation that might otherwise occur with sample alone. Light or fluffy samples are also better retained within the Soxhlet thimble, and in this way, are prevented from migrating through the siphon tube.
*Choice of the dispersant should include consideration of adsorptive interactions that might occur between the analytes and the dispersant matrix. If strong interactions are suspected or are known to occur, then a different dispersant should be selected. Sodium sulfate is applicable to most classes of analytes, whereas diatomaceous earth may retain some compounds, such as brominated flame retardants. Sand is also broadly applicable, and is less likely to be retentive due to the relatively low specific surface area.
For trace level applications, it may also be necessary to pre-clean the dispersant with solvents. The need for pre-cleaning can be evaluated by running an extraction blank (that is, the dispersant without a sample).
*Finally, glass wool should be added to the extraction thimble prior to the addition of any dispersant or samples. This helps to prevent plugging of the glass frit.

## Slide 7: Steps in Sample Processing
Add glass wool and dispersant to thimble
Weigh sample into thimble
Add internal standard to sample
Mix sample and dispersant
Prepare Soxhlet apparatus
  Add solvent and boiling stones to distillation flask
  Place thimble in Soxhlet extraction vessel
  Assemble glassware
Turn on coolant and heater
At the conclusion of the extraction, turn off the heater and allow to cool before further sample processing.
Notes: An overview of the steps in processing a sample by Soxhlet extraction can be provided.
The first step is to add glass wool to the thimble, followed by a layer of the dispersant that has been selected.
The next step is to add the sample to the thimble. The mass of sample can be determined by difference of masses before and after addition to the thimble; or if the sample mass is small relative to the mass of the thimble, then the mass of the sample can be determined by difference from difference with a suitably sized weighing container.
The next step is to add the internal standard to the sample in the thimble
Mix the sample and the dispersant, and an optional step at this point is to add additional dispersant on top of this mixture.
The next step is to prepare the soxhlet apparatus. Add solvent and boiling stones to the boiling flask. Place the thimble together with the sample into the soxhlet apparatus. Use long tongs to insert the thimble. Use care in this step, since it is very easy to break the thimble or soxhlet glassware if the thimble slips unexpectedly. When the thimble is in place, the glassware can be assembled.
Turn the coolant flow on to the condensers, and then turn on the heating mantle to start the extraction process.
At the conclusion of the extraction, turn off the heater and allow the boiling flask to cool before continuing with the sample processing.

## Slide 8
Notes: The following video illustrates these steps for a sample. Begin by assembling the needed materials at the balance on a clean work surface. Transfer an appropriately sized portion of glass wool to the thimble and compact it as much as possible. Next, transfer a few grams of the dispersant on top of the glass wool to a height of perhaps 1 cm. Weigh the thimble containing the glass wool and the initial aliquot of the dispersant. Next, transfer a portion of the sample to the thimble, and weigh again to obtain the quantity transferred. Fill a clean and dry gas tight syringe with a portion of the internal standard solution and weigh. Excess solution on the needle can be removed by wiping, prior to determining this mass. Once the mass of the solution in the syringe is obtained, transfer an aliquot onto the sample in the thimble. The quantity transferred should be planned to be small enough so that no solution will leak through the thimble. This is typically not an issue, if the internal standard solution is 1 mL or less. Weigh the syringe to obtain the quantity transferred. Add an additional portion of the dispersant on top of the sample and mix the contents thoroughly. Optionally, additional dispersant can be added to the sample if desired. If a small mass of the sample is to be extracted, for example, 100 mg or less, it is more accurate to weigh the sample in a foil container, and then transfer the sample to the thimble. The mass transferred can then be determined by difference. Once the sample is prepared in the thimble, assemble the glassware at the location of the Soxhlet extraction setup. Add a few boiling stones to the distillation flask, and then fill the flask with the extraction solvent. It is good practice to fill the container to nearly reach the neck of the flask. This provides insurance that the flask will not go dry during the extraction if losses occur through the ground glass joints. It is usually not necessary to use vacuum grease at the joints, and in fact this may create significant sample contamination and interferences in subsequent analyses. Place the thimble in the Soxhlet glassware using tongs. Use care with this step since it is easy to break the thimble or the Soxhlet glassware if the thimble slips. Assemble the Soxhlet glassware to the distillation flask and the condenser. Turn on the coolant flow, either with a recirculation bath or with chilled water. Next, turn on the heater and adjust the voltage to result in a moderate rate of boiling, perhaps at 70 or 80 % dial setting. Before leaving the apparatus for an extended period, check periodically during the first hour to make sure that the solvent is boiling smoothly, and that there are no obvious vapor leaks. Solvent should readily condense and drip into the extraction thimble. The time between siphon cycles is typically 15 to 20 minutes. This can be adjusted somewhat by raising or lowering the power supplied to the heating mantle from the Variac. The siphon occurs when the level of the solvent reaches the top of the siphon tube. As the extract flows into the distillation flask, the boiling action will moderate briefly, until the boiling point is again regained. At the completion of the extraction, turn off the power to the heating mantle and allow the solution to cool. Ideally, it is useful to time this so the thimble remains nearly filled with the solvent. This seems counter intuitive, but less solvent will need to be evaporated from the extract in subsequent processing steps. Near the end of the extraction, the analyte should not be detectable in this solvent in the sample thimble.

## Slide 9: www.nist.gov
Notes: To learn more about the National Institute of Standards and Technology, visit www.nist.gov.