## Slide 1
Notes: The ability to assess analytical performance is a fundamental component of measurement science. Ultimately, measurement quality depends on the proper function of instrumentation in combination with appropriate methods. Evidence of proper function is typically provided by the instrumentation, chromatograms, and the resulting data generated. Chromatographic peak shape is diagnostic of problems that originate with the instrumentation as well as the method and sample. Once a problem has been clearly identified, steps can be taken to resolve the issue. This presentation will describe common problems than may be encountered, with representative illustrations of the resulting effects on chromatographic performance.

## Slide 2: First Steps…
Sources of Problems:
Mechanical and/or fluid problems
Detector/signal/data system problems
Problems associated with the sample
Problems associated with chromatography
Notes: Troubleshooting is often associated with problems in the instrumentation, but issues may originate from other sources as well. Sample incompatibility with the method (such as solvent mismatch), methods that are not fit for purpose, and column degradation are examples of such issues.
1) With liquid chromatography, fluidic problems often result from mechanical issues that cause erratic flow or leaks.
2) Bubbles in the flow pathways also can contribute to erratic signals; more rarely, issues arise from the electronics
3) For newly implemented methods, problems associated with the sample or the method should not be discounted, particularly if the method has not been validated with an appropriate control.

## Slide 3: Peak Shape…
Ideal: symmetric, Gaussian peak
width at ½ height =
r
Efficiency is indicated by theoretical plate number, N
Notes: Under ideal conditions, chromatographic peak shape can approach a symmetric Gaussian distribution. This quality is desirable, since it results in peaks with the highest separation efficiency. The well defined peak shape also facilitates baseline assignment. Perturbations in the Gaussian profile are easily discerned, and are useful in evaluation of system and method performance.
1) The Gaussian function shown is normalized so that the area under the curve is equal to 1.
At the half height of the peak, the width can be measured to calculate the number of theoretical plates generated,
using the equation at the bottom of this slide. For modern LC applications with a properly configured system, at least 5,000 theoretical plates is typical. This of course depends on the column length and particle diameter, but this value is a reasonable expectation.

## Slide 4
Notes: Non-ideal peak shape may or may not be acceptable in practice. If peak distortion can be attributed to factors that do not bias quantitation, then the non-ideal peak shape may be acceptable. For example, some polar compounds exhibit peak tailing in reversed-phase systems due to competing retention mechanisms with different interaction rates. This should not affect response linearity, but it may make resolution of the targeted analytes more challenging due to loss in effective efficiency. If, however, the non-ideal peak shape can be traced to a coeluting interference, then measurements will be biased by the contribution to peak area. As a general rule, non-ideal peak shape should be investigated and resolved during method development.

## Slide 5
Notes: It is highly useful to analyze a test sample to evaluate system performance prior to beginning analytical measurements. This does not need to be a formal process – any sample can be used, but the analyst should have some prior knowledge about what the result should look like so that the test can provide meaningful results. Probably the best text mixture is a solution that contains the compounds that will be determined with the method to be used; for example, a calibrant solution. If the instrumentation is being used at an earlier point in the process, prior to method development, then a generic test mixture can be used to evaluate mechanical aspects of the instrument. Nonpolar compounds, such as are contained in SRM 869 column performance Test Mixture for LC are best for this purpose. Specialized mixtures can be used to study aspects of chromatographic performance associated with the column. For each of these cases, it is imperative that previously generated results representative of good system performance be available for comparison with the test cases.

## Slide 6
Notes: This slide provides a common example to illustrate the diagnostic qualities of peak shape. For isocratic separations, peak width increases predictably with increasing retention. For gradient separations, peak width typically remains fairly constant, particularly during the gradient elution profile. The appearance of a broad peak, bracketed between narrow peaks, is unusual, and should be investigated. One possible explanation for such an anomaly is carry over from the previous injection. Highly retained compounds may not elute during the method initial run time, and during a subsequent analysis, this component might finally elute, with atypical peak shape. This explanation can be tested by analyzing one or more blanks between the samples – this is always a good idea.

## Slide 7: Sample overload
What happens if too much sample is injected onto the column? Distortion of peak shape!
Two independent issues:
  sample (injection) volume: volume overload
  sample mass applied to the column: mass overload
Injection volume: The nature of the injection solvent (i.e., the sample) influences the maximum injection volume
  Best case: sample is in a weak solvent (i.e., mobile phase composition, or weaker)
  Worst case: sample is in a strong solvent
  (i.e., methylene chloride or hexane)
Notes: 1) The issue of sample overload is another common problem that can occur with any analysis, even if measurements are made at trace levels. When overloading occurs, the resulting distortion in peak shape can compromise the integrity of the measurements.
2) Two types of sample overloading, at least in theory, can occur: overloading related to the volume of sample injected, and overloading due to the mass of the sample applied to the column.
3) Injection volume overloading results when the quantity of solvent in the sample is sufficient to perturb the mobile phase strength and accelerate elution of the analytes on the column. Solvent strength relative to the mobile phase is a key factor. If the solvent strength of the sample is greater than the mobile phase, then the sample is spread out as the sample plug is carried through the column. If however, the solvent strength of the sample is weaker than the mobile phase, then the sample is concentrated as a tight band at the top of the column. Under these conditions, even very large injection volumes may be feasible, without affecting the quality of the separation. This approach is sometimes used to concentrate water samples.

## Slide 8: Column overloading…
Rule of Thumb:
Always use as small an injection volume as possible. Volume can be increased if the sample is in a weak solvent.
Example: Large volumes of water can be injected onto a reversed phase column. The sample is concentrated onto the top of the column by the weak injection solvent, then eluted by the stronger mobile phase.
Notes: It is always a good idea to assess the effect of injection volume on peak shape, as a guide for optimizing injection volume and detection goals.
Better chromatography results with small injections. If the sample is in a weak solvent, larger injection volumes will not have a significant detrimental effect.
1) As an example, large volumes of water can be injected onto a reversed phase column. The sample is concentrated onto the top of the column by the weak injection solvent, then eluted by the stronger mobile phase. With this application, Guassian peak shape is maintained.

## Slide 9: Injection Volume Overload
Solvent strength/volume problems affect early eluting peaks much more than late eluting peaks
Early peaks show fronting
Late peaks are Gaussian (unaffected)
Gradient elution vs. Isocratic elution
Isocratic - more sensitive to solvent volume overload
Gradient - less sensitive; later peaks may be unaffected
Notes: One indicator of injection solvent overload and solvent strength mismatch is evident in characteristic peak shape distortion. Solvent mismatch is indicated by peak fronting that is more prominent for early eluting components than later eluting components.
1) Because gradient elution methods tend to concentrate the sample, gradient methods are somewhat less sensitive to injection volume that are isocratic methods.

## Slide 10: Injection Volume Overload
Column size (inside diameter)
Microbore columns are very sensitive to solvent volume overload.
Reduce injection volume by the ratio of cross sectional areas.
Example: for a 10 µL injection on a 4.6 mm i.d. column:
With a 2 mm i.d. column, the injection volume should be reduced to 2 µL. The flow rate should also be decreased by a factor of 5
Notes: Column dimensions are a critical element in assessing the effect of injection volume on peak shape. The cross sectional area of the column is proportional to the maximum injection volume permitted without overloading the column.
1) Stated another way, if a 10 uL injection provides acceptable results with a 4.6 mm i.d., column, then for a 2 mm column, the injection volume should be reduced by the ratio of the corresponding areas, which in this example is 5.3, resulting in an acceptable injection volume of 2 uL. In addition, to keep the linear velocity the same, then the flow rate should also be reduced by a factor of about 5.

## Slide 11: Injection Volume Overload
concentration
Narrow bands
Resolution remains high
Sample in weak solvent
some
separation
with this
band
Band broadening
Bands will never be narrower than this
Limits resolution
Sample in strong solvent
Notes: This slide graphically illustrates the principle of injection volume overload.
1) When the sample is in a weak solvent and is applied to a chromatographic column, the band is concentrated and focused on the column bed.
2) If the sample is in a strong solvent, the band is not focused, and components in the separation will always be at least as broad as the initial injection.

## Slide 12
Notes: In this illustration, the components of a mixture of priority pollutant PAHs are separated using a gradient method that begins with an initial composition of 50% acetonitrile in water. The sample was evaporated to dryness and dissolved in the starting mobile phase composition. Because there is little or no solvent mismatch, the effect of injection volume on peak shape is minimized.

## Slide 13
Notes: In this example, the 16 component mixture is reconstituted in acetonitrile, which is a strong solvent compared with the starting gradient composition. Peak shape is unaffected for injections less than about 50 uL, but for larger injection volumes, the early eluting compounds demonstrate the characteristic fronting distortion. Later eluting peaks are less affected.

## Slide 14
Notes: A few of the chromatograms are replotted in this figure for clarity. Acetonitrile is the injection solvent, and the gradient elution program begins at a composition of 50% acetonitrile in water. Although this indicates extreme cases, it can be noted that even with the solvent mismatch, a fairly large injection volume of 50 uL results in good peak shape.

## Slide 15
Notes: This figure illustrates a more extreme solvent mismatch, namely, methylene chloride. Not only does this solvent represent high elution strength for reversed-phase separations, but it is also immiscible with 50/50 acetonitrile/water solvent composition. Peak shape degradation is evident at the 10 uL injection volume level for early eluting peaks. It can be further noted that with larger injection volumes, peak shape for even the late eluting components is distorted. Perhaps surprising is the observation that useable results are evident for the 5 uL injection volume, even with the immiscibility property.

## Slide 16
Notes: Methods based on (or at least tolerating) solvent mismatch can sometimes provide useful results. This example shows the analysis of a reformulated gasoline reference material by reversed phase liquid chromatography with isocratic elution, and without a solvent exchange step. The injection size was minimized at 1 uL. Peak shape for the components is good, even though the sample is immiscible with the mobile phase.

## Slide 17: Mass overload
For analytical columns/separations, it is difficult to achieve true mass overload conditions. Usually “mass overload” is really a solvent mismatch.
Rule of thumb: Reversed-phase columns can accept about 1 mg sample per gram packing material. It is difficult to dissolve 1 mg in 10 µL (correspond to ≈10% mass fraction sample concentration).
If the sample is dissolved in the mobile phase or weaker solvent, mass overload is typically not a problem. For stronger (sample) solvents, keep the injection volume low.
Using the above rule of thumb, 1 - 2 µL injections should be OK even for very strong solvents.
Peak width increases with mass injected; resolution of critical pairs may be degraded.
Notes: Mass overload results when nonlinear isotherms occur at very high concentrations. The conditions for mass overload often occurs with preparative chromatography, but it is unlikely to occur for analytical separations.
As a rule of thumb, reversed phase columns can accept about 1 mg of sample per g of the adsorbent packing material. A typical 4.6 mm id by 250 mm length column contains about 3 gram of adsorbent. Mass overload would start to be approached for a 10 uL injection containing a mass of 1 mg, representing a concentration of about 10% mass fraction.
Samples this concentrated are rarely encountered for analytical applications.
This does suggest that even with the most concentrated samples, injection volumes of 1 or 2 uL may provide useable results. If the example in the previous slide for reformulated gasoline is considered in terms of the primary constituents, toluene and benzene, the 1 uL injection represents slightly less than 1 mg mass of sample, and overload conditions are not achieved for the 4.6 mm x 250 mm column.
Overloading conditions potentially affect quantitation by nonlinear response, but overloading also can reduce the resolution of critical solute pairs.

## Slide 18
Notes: This figure illustrates another example of chromatography for a concentrated sample. A saturated solution of fluorene was prepared in 50/50 acetonitrile/water, and separated isocratically in a slightly stronger mobile phase composition. Under these conditions, solvent mismatch does not apply and the injection is concentrated at the inlet of the column. Even with a large injection volume, the primary component peak shows no sign of peak shape distortion.

## Slide 19: Extracolumn effects
Tubing volume should be minimized, but this effect is not as important as some texts would have you believe.
Use 0.010" i.d. tubing between the injector and column and between the column and the detector.
Keep lengths relatively short (in practice, 50 - 100 cm is fine).
Problem becomes more important for small i.d. columns.
Symptom: early peaks tail more than late peaks.
Notes: The importance of minimizing extracolumn volume is frequently stressed in academic texts and in the research literature, in relation to system performance. An examination of the actual influence of extracolumn effects in practice is a useful exercise, since the influence of this property varies considerably with other aspects of the instrumentation.
The extracolumn volume is usually considered to include band broadening contributions from system components after the injector, but not including the column.
Typically, recommendations are to make column connections with very short length, small inside diameter tubing. This advise is OK, but system performance will be indistinguishable with moderate tubing lengths if small i.d. tubing is used. Loss in system performance will become more noticeable with small i.d. columns.

## Slide 20
Notes: The following figure illustrates the influence of extracolumn volume on peak shape for a conventional 4.6 mm x 250 mm 5 um particle diameter column. A 61 m length of 0.010” stainless steel tubing was added after the injector and before the column to assess how this added volume would affect peak shape. This length of tubing increased the extracolumn volume by 3.09 mL. . As expected, the retention times for the three components are each increased by a little over 2 minutes, since 3.09 mL divided by 1.5 mL/min is 2.06 min. Surprisingly, however, very little change in overall peak shape can be noted after the addition of the tubing. This probably can be attributed to the use of the small 0010” i.d. tubing, which helps to prevent band broadening by maintaining a high linear velocity..

## Slide 21
Notes: This figure illustrates the same instrumentation, but operated in the isocratic mode. The two peak profiles are almost identical when they are graphically overlaid, and the efficiency is degraded by only about 25%. For normal applications, the addition of modest tubing length of perhaps 50 cm or less will not practically influence performance, as long as small i.d. tubing is used.

## Slide 22: Retention Time Variability
Possible causes
  pump problem: seals, check valves, leaks
  varying column temperature
  not enough equilibration time in method
  mobile phase composition changing (evaporation)
  pressure too high
Column damage
  Excessive backpressure
  Peak shape degraded: fronting, tailing, peak splitting
  All peaks in the run are affected
Notes: Retention time variability is a very common occurrence in LC applications. Sources of variability often point to problems with the solvent delivery system, that can include the pump seals, check valves, and leaks. Another common source of retention time variability is changes in column temperature. Unless the column is enclosed in an oven or water bath, ambient conditions may fluctuate sufficiently to change retention. Enclosure of the column in a water bath provides precise temperature control, but column ovens are also effective, particularly at elevated temperatures. Other potential sources of variability include inadequate equilibration at the beginning of gradient methods, and changes in mobile phase composition due to evaporation, for solvent reservoirs that contain pre-mixed solvents.
Damage to columns may be indicated by changes in retention. Usually, other aspects of the separation occur as a result of damage, that are evident by increased column backpressure, or peak fronting, tailing or splitting. All peaks in the chromatogram are affected by such mechanical damage.

## Slide 23
Notes: Electronic noise is always present at some level in data systems, but the relative contribution of noise should be low compared with analyte response. The power spectrum for noise from analog detectors, such as absorbance detectors, is different than from mass spectrometry. Sampling rate for conversion from analog to digital can range from 100 hz or higher for some data systems, to only a few hertz for selected ion monitoring programs in MS. Noise can be reduced by replacing any unshielded signal lines with cables that incorporate a grounded mesh shielding. The shielded lines should be only as long as necessary to connect the components; avoid long cable runs.

## Slide 24: Flow-related noise
Most pump problems produce periodic fluctuations
Test if pump-related by turning pump off
If baseline is OK, then problem is flow-related
Possible problem: air bubble in the detector cell
Clear bubble; degas solvents; and/or add backpressure restrictor to outlet line
1 pump stroke
Notes: Almost all liquid chromatographic pump designs are based on technologies that produce periodic flow and pressure fluctuations. These fluctuations can contribute an unwanted signal at the detector that create disturbances in the baseline. In most cases, the disturbance can be easily attributed to flow based on the periodic nature of the noise.
Flow related noise can also be identified by turning off the pump and observing any changes in the detector signal.
A common problem results when an air bubble that is trapped in the cell with absorbance detectors. The bubble moves with the slightest variation in pressure and flow, and creates the characteristic periodic signal. To clear the bubble, flush the cell with isopropanol. If the problem persists, it may be necessary to manually degas the solvents. Excellent results are often obtained by application of a backpressure restrictor to the outlet waste line. Before attempting the latter remedy, make certain that the detector cell can withstand the added backpressure. Typically, 50 or 100 psi backpressure restrictors are safe and effective for this purpose.

## Slide 25
Notes: Baseline drift is characterized by a non periodic and unpredictable change in the baseline over time. Any number of sources may contribute to drift, so elimination of this issue may be time consuming and uncertain. The figure lists many of the problems that can result in baseline drift. After eliminating any mechanical issues related to the pump, it is frequently useful to flush the column with 100% organic modifier at a high flow rate for 10 to 15 minutes. Usually, the baseline will stabilize during this period of time, and research efforts can be continued.

## Slide 26
Notes: Noise spikes are nonperiodic high frequency fluctuations that usually can be attributed to power utilization. If this is deemed to be the case, then an easy solution to the problem is to move the offending instrumentation to a different electrical circuit. Thermostated water baths and other mechanically switched devices can produce periodic spikes, as indicated in this figure.

## Slide 27
Notes: Reproducible shifts in baseline associated with a gradient elution method are also quite common. The effect results when the background absorbance of mixed solvents differs from the individual pure solvents, and is most prevalent for absorbance detection at 190 nm to 210 nm.
Moderate changes in baseline can be tolerated, especially if the magnitude of the shift is small compared with analyte response, and if peak integration is not compromised.

## Slide 28: Excessive Back Pressure
Isolate the source of the excessive back pressure, check and replace any plugged inline filters or tubing
Disconnect tubing at each connection in the system step-wise, starting at the detector outlet Look for significant pressure changes upon disconnection
  Pressure should be very low without column connected.
  Possible sources of plugs: lines, filters, injector loops, rotors, detector cells
The most common source of high back pressure is the precolumn, column and/or column frits
Column frits can be replaced (this may void warrantee) but use care not to disturb the column bed
If this happens often, may need to filter solvents or filter samples (use 0.5 µm syringe filter)
Record normal backpressure values for each column
Notes: It is good practice to record the normal backpressure for each individual column as it is first used, so that comparisons can be made at a later time when excessive backpressure is suspected. For this to be most useful, the conditions should also be recorded, for example, a note could be stored with the column stating: “1100 psi backpressure for 100% methanol at a flow rate of 1 mL at 20 C.
If excessive column backpressure is indicated, several remedies can be attempted.
Initially, remove the line to the inlet of the column and operate the pump. Backpressure should be less than about 10 psi. If larger backpressure is indicated, then a restriction exists in either solvent lines or inline filters, or rarely, in the injector valve. Isolate these components individually, and replace as necessary.
Connect the inlet and outlet lines with a union, thus eliminating the column. Operate the pump. Again, backpressure should be less than 10 psi. If larger backpressure is indicated then a restriction exists in the outlet line or the detector flow cell.
If the excessive backpressure appears to originate in the column, try flushing the column by reversing the inlet and outlet connections. Modern columns can be operated in either direction without damage; back-flushing may help to clear restrictions from inlet and outlet frits.
As a last resort prior to column replacement, it may be possible to remove the inlet and outlet frits internal to the column. This is usually only possible for generic style endfittings, and doing so will void any warrantee that might apply. But frit replacement has a good chance of improving or eliminating pressure restrictions.
If backpressure restrictions are regularly experienced, this may indicate a need to filter solvent or samples using a 0.5 um filter. Particulate matter from other sources, such as wear points from seals or injectors may occasionally be the source of the pressure restriction.

## Slide 29
Notes: Check valves are integral components for most pump designs. The ball and seat parts are highly susceptible to fouling by particulate matter, and when their operation is impeded, pressure delivered by the pump is erratic. To determine proper function, collect the flow pumped through the column over at least a one minute period and compare this volume with the volume based on the specified flow rate. If the collected volume is significantly less than the flowrate would indicate, then the check valves or pump seals are not fully functioning.
To evaluate check valves for proper function, the flow path can be blocked after the pressure transducer, shown as point A in the figure. Under normal operating conditions the pump will shut off as the pressure limit is exceeded. This pressure should be maintained and only gradually decrease with time, if the check valves are functioning properly. A rapid decrease in pressure indicates likely check valve malfunction.

## Slide 30
Notes: Under isocratic conditions, system pressure should be constant within a few psi. If wide variations in pressure are evident, particularly timed with pump action, then this also indicates pump problems. When irregular pressure is noted, the cause of the problem should be identified and remedied. Pressure fluctuations invariably result in retention irreproducibility, and poor measurement precision. The likely sources of the problems are the wear points of the check values and piston seals. These should be replaced as needed.

## Slide 31
Notes: To learn more about the National Institute of Standards and Technology, visit www.nist.gov.