## OWM Disclaimer
Disclaimer
These spreadsheets are provided as a useful approach for metrology calculations according to the SOP represented here. However, NIST OWM or any other organization involved in their development, collectively and individually, do not warrant this spreadsheets for any specific purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each user agrees to decide if, when and how to use the spreadsheets, does so at his or her sole risk, and is responsible to suitable and applicable verification prior to use. When using the tools provided on the NIST OWM website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST OWM, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST OWM, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST OWM website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## Software V&V
Technical Assessment
This form is completed for each software file used in the laboratory. The form may be added as a tab/worksheet to a file/workbook or may be maintained in a separate electronic/paper file/location in the laboratory.
This form is part of the Procedure for Software Verification and Validation from the NIST Office of Weights and Measures.
Codes | Assessment | Pass/Fail | Objective Evidence and Observations
A. Software Inspection | Spreadsheet is clear and makes sense
There are instructions for use
Instructions and data input appear on the visible portion of the first worksheet
Data-entry fields are labeled and color coded (it is recommended to avoid red and green)
The Standard Operating Procedure (SOP) used is clearly specified
The number of digits to be rounded to is specified
The user is warned/notified whenever data-entry fields are left blank
Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations
The software opens at the right location within the file
Unused fields/cells are locked
Rows/columns that the operator need not see are hidden
Unused sheets are removed
Worksheets are named appropriately
B. Mathematical Specification | The correct SOP is used
The formulae and methods chosen from that SOP are specified
Sources and references for formulae are specified
The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty
C. Code review | The formulae in the fields exactly match the SOP
Repeated calculations appropriately reference the correct cells
Calculations, when tested using standard data or reference test data, show appropriate accuracy
Rounding is done at the appropriate locations in the file
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences
Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa
“Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used
E. Component Testing | Each macro used is functional
Each command/button is functional
Combinations of interdependent macros are functional
Plotted graphs are accurate
Worksheets/reports print properly, if needed to
Conditional (color and non-color) formatting is functional
F. Numerical Reference | Look-up tables and lists match the latest calibration report.
Uncertainties match the latest Scope
Values that reference another workbook or spreadsheet are dated
When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct
The evaluation of the embedded data is dated and documented
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty
J. Security | Equation and calculation cells are protected against inadvertent editing
Cells are locked in place; they cannot be moved/dragged
Confidentiality of passwords is appropriate
Files are backed up automatically
Additional back-up is available at alternate facilities
Files on network drives cannot be accidentally deleted

## Change and Rev Control
Owner | NIST Office of Weights and Measures
Tab | Change | Date | Initials
All | Original Files | 2008-04-01 | GH
All | Updated with more uncertainty component columns; added degrees of freedom, with automatic k value and Uncertainty value calculations and rounding to two significant digits. Updated P(n) calculations. | 2018-08-29 | GH

## Laboratory Scope
Lab Name
Date
Measurement Parameter | Range | Uncertainties
Mass, Echelon I, Metric |  | Metric Mass I'!A1
Mass, Echelon I, Avdp |  | Advp Mass I'!A1
Mass, Echelon II, Metric |  | Metric Mass II'!A1
Mass, Echelon II, Avdp |  | Advp Mass II'!A1
Mass, Echelon III, Metric |  | Metric Mass III'!A1
Mass, Echelon III, Avdp |  | Advp Mass III'!A1
Mass, Echelon III, Wheelload Weighers |  | WheelLoad Weigher & Wt Cart'!A1
Mass, Echelon III, Weight Carts |  | WheelLoad Weigher & Wt Cart'!A1
Mass, Echelon III, Railroad Test Cars
Volume, Echelon I, Gravimetric |  | Gravimetric Volume I & SVP'!A1
Volume, Echelon I, Gravimetric SVP |  | Gravimetric Volume I & SVP'!A1
Volume, Echelon II, Volume Transfer |  | Volume Transfer II'!A1
Volume, Echelon II, Volume Transfer, LPG |  | Volume Transfer II LPG'!A1
Length, Metric |  | Length Metric'!A1
Length, US Customary |  | Length US Cust'!A1
Temperature |  | Temperature!A1
Frequency (Tuning Forks) |  | Time & Frequency'!A1
Time (Stopwatches) |  | Time & Frequency'!A1
NOTE: This section and values in this workbook need to match the Recognition Application!
Delete rows in this table and worksheets in this file that are NOT used.

## Uncertainty Budget Template
Measurement Result Units: |  |  |  |  | Measurement Range and Parameter:
Instructions: Fill in both "components" and "est unc in meas units" AND | [6 empty] | select "prob distribution" for complete calculations. | [5 empty] | Lists for Drop Downs and Lookups. Do NOT Delete.
Uncertainty Component Description | Symbol | Estimated Uncertainty | Units | d.f. | Estimated Unc in Measurement Units () | Type (A, B) | Probability Distribution | Divisor | Std Unc () | Relative Contribution (%) | Explanation/Source/Notes | For effective degrees of freedom calculation. |  |  |  | Acceptable Units | Description
[13 empty] | Type | Probability | Divisor
[13 empty] | A | Normal, 1s | 1 |  | blank
[13 empty] | B | Normal, 2s | 2 | µg | microgram
[14 empty] | Normal, 3s | 3 | mg | milligram
[14 empty] | Norm@99% | 2.58 | g | gram
[14 empty] | Rectangular | 1.7321 | kg | kilogram
[14 empty] | Triangular | 2.4495 | µlb | micropound
[14 empty] | U-Shaped | 1.4142 | lb | pound
[14 empty] | Rect 1/2 | 3.4641 | in | inch
[16 empty] | in² | square inch
[16 empty] | in³ | cubic inch
[16 empty] | pt | pint
[16 empty] | qt | quart
[16 empty] | gal | gallon
[16 empty] | mL | milliliter
Instructions: Finish selections or assess resulting values. | [9 empty] | Instructions: Complete data entry. | [5 empty] | L | liter
Min Degrees of Freedom | ν |  |  | Entries incomplete. | [11 empty] | nm | nanometer
Effective Degrees of Freedom | νeff |  |  | TBD | [11 empty] | mm | millimeter
Combined Uncertainty, uc | [8 empty] | 0 | Instructions: Assess data entry and values before reporting rounded result. | [5 empty] | cm | centimeter
Coverage factor, k, uses effective degrees of freedom | [8 empty] | TBD | [6 empty] | cm² | square centimeter
Expanded Uncertainty, U | [8 empty] | TBD | [6 empty] | cm³ | cubic centimeter
Expanded Uncertainty, U, Rounded to 2 Significant Digits | [8 empty] | TBD | No units selected. | [5 empty] | m | meter
[16 empty] | m² | square meter
NOTE: This page is a working section that can be used for calculating effective degrees of freedom. | [15 empty] | m³ | cubic meter
[16 empty] | °F | degrees Fahrenheit
[16 empty] | °C | degrees Celcius
[16 empty] | K | Kelvin
[16 empty] | fl dr | fluid drams
[16 empty] | fl oz | Fluid ounce(s)
[16 empty] | ft | foot
[16 empty] | Pa | Pascal
[16 empty] | kPa | Kilopascal
[16 empty] | s | second
[16 empty] | min | minute
[16 empty] | h | hour
[16 empty] | A | Ampere
[16 empty] | C | coulomb
[16 empty] | V | volt
[16 empty] | J | joule
[16 empty] | W | watt
[16 empty] | Hz | hertz
[16 empty] | H | henry
[16 empty] | F | farad
[16 empty] | Ω | ohm
[16 empty] | S | siemens
[16 empty] | Wb | weber
[16 empty] | T | tesla
[16 empty] | A/m | Ampere per meter (magnetic field strength)
[16 empty] | cd | candela
[16 empty] | lm | lumen
[16 empty] | lx | lux
[16 empty] | cd/m² | candela per square meter
[16 empty] | rad | radian
[16 empty] | kg/m³ | kilogram per cubic meter
[16 empty] | g/cm³ | grams per cubic centimeter
[16 empty] | kg/L | kilograms per liter
[16 empty] | m/s | meters per second
[16 empty] | km/h | kilometers per hour
[16 empty] | m/s² | meter per second squared
[16 empty] | m²/s | meter squared per second
Note on G3: Select Type A or B.
Note on O12: Use if you need to divide the value by 1/2 and the square root of 3.

## Metric Mass I
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 28 |  | Mass Echelon I | [10 empty] | Use ASTM Class 0 or 1, OIML E1 or E2.
|  |  | Specify variable source. | [11 empty] | Class | 0 |  | E1
|  | Std | Process | Density/Buoyancy | Bias | Other | Other | Other | Other
Nominal | Units | us | sp | ubc | ubias | uo | uo | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
100 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on E5: OIML R111 C.3.6-1
Note on F5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Advp Mass I
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 28 |  | Mass Echelon I | [10 empty] | Use ASTM Class 0 or 1, OIML E1 or E2.
|  |  | Specify variable source. | [11 empty] | Class | 0 |  | E1
|  | Std | Process | Density/Buoyancy | Bias | Other | Other | Other | Other
Nominal | Units | us | sp | ubc | ubias | uo | uo | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
1000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
250 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
25 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.05 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.03 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.02 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.01 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.005 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.003 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.002 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.001 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/16 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/32 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/64 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on E5: OIML R111 C.3.6-1
Note on F5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Metric Mass II
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 5 |  | Mass Echelon II | [10 empty] | Use ASTM Class 2 or 3, OIML F1 or F2.Update if not as in template.
|  |  | Specify variable source. | [11 empty] | Class | 2 |  | F2
|  | Std | Process | Density/Buoyancy | Bias | Other | Other | Other | Other
Nominal | Units | us | sp | ubc | ubias | uo | uo | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
1000 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
250 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on E5: OIML R111 C.3.6-1
Note on F5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Advp Mass II
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 5 |  | Mass Echelon II | [10 empty] | Use ASTM Class 2 or 3, OIML F1 or F2.Update if not as in template.
|  |  | Specify variable source. | [11 empty] | Class | 2 |  | F2
|  | Std | Process | Density/Buoyancy | Bias | Other | Other | Other | Other
Nominal | Units | us | sp | ubc | ubias | uo | uo | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
10000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
4000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
250 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
25 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.05 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.03 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.02 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.01 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.005 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.003 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.002 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.001 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/16 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/32 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/64 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on E5: OIML R111 C.3.6-1
Note on F5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Metric Mass III
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 8 |  | Mass Echelon III | [10 empty] | Use 105-1, or OIML M1, M2, M3, or ASTM 4, 5, or 6. Update if not as in template.
|  |  | Specify variable source. | [11 empty] | Class | 105-1 |  | M3
|  | Std | Process | Drift | Sensitivity | Bias | MABC | Other | Other
Nominal | Units | us | sp | udrift max | usens max | ubias | ubc | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
1000 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
250 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | kg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | g | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | mg | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on G5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Advp Mass III
Lab Name |  |  | Procedure |  | Scope |  |  |  |  | Back to Lab Scope
Date |  |  | 8 |  | Mass Echelon III | [10 empty] | Use 105-1, or OIML M1, M2, M3, or ASTM 4, 5, or 6. Update if not as in template.
|  |  | Specify variable source. | [11 empty] | Class | 105-1 |  | M3
|  | Std | Process | Drift | Sensitivity | Bias | MABC | Other | Other
Nominal | Units | us | sp | udrift max | usens max | ubias | ubc | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Tolerance | Pn | Comments (Balance, Modifications, Action Items)
10000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
4000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1000 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
250 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
25 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.5 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.3 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.2 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.1 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.05 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.03 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.02 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.01 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.005 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.003 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.002 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
0.001 | lb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
500 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
300 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
200 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
100 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
50 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
30 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
20 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
10 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
5 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
3 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | µlb | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/2 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/4 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/8 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/16 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/32 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
1/64 | oz | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- |  | ---
Note on C5: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Standard deviation of the process from check standard data over time
Note on G5: Must comply with SOP 29 criteria or corrective action is needed.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## WheelLoad Weigher & Wt Cart
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date |  |  | 33 |  |  |  | Wheel load weighers and weight carts | [7 empty] | Use 105-8
|  |  | Specify variable source. | [10 empty] | Class | 105-8
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Balance | Tolerance | Pn | Comments
6000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Weight Carts
5000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3500 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2500 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
|  |  | Procedure |  |  |  | Scope
Use HB 44, Class IIIL Scales
|  |  | Specify variable source. | [10 empty] | Class | HB 44, IIIL
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc |  | k | U | Units | Instrument | Tolerance | Pn | Comments
40000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Wheel Load Weighers
20000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
15000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
5000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2500 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1500 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1000 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
500 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
250 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | lb | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!
Note on C18: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D18: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M18: Round or Format all expanded uncertainties to two significant digits!

## Metric Mass Tolerance Table
Maximum Permissible Errors for Weights (± in mg)
| OIML (R111-04) | [8 empty] | ASTM (E 617-13) | [9 empty] | NIST (90)
Nominal | E1 | E2 | F1 | F2 | M1 | M1-2 | M2 | M2-3 | M3 | 000 | 00 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | F
5 000 kg |  |  | 25 000 | 80 000 | 250 000 |  | 800 000 |  | 2 500 000 | [6 empty] | 100 000 | 250 000 | 500 000 | 750 000 | 500 000
3 000 kg | [15 empty] | 60 000 | 150 000 | 300 000 | 450 000 | 300 000
2 000 kg |  |  | 10 000 | 30 000 | 100 000 |  | 300 000 |  | 1 000 000 | [6 empty] | 40 000 | 100 000 | 200 000 | 300 000 | 200 000
1 000 kg |  | 1 600 | 5 000 | 16 000 | 50 000 |  | 160 000 |  | 500 000 | [6 empty] | 20 000 | 50 000 | 100 000 | 150 000 | 100 000
500 kg |  | 800 | 2 500 | 8 000 | 25 000 |  | 80 000 |  | 250 000 | [6 empty] | 10 000 | 25 000 | 50 000 | 75 000 | 50 000
300 kg | [15 empty] | 6 000 | 15 000 | 30 000 | 45 000 | 30 000
200 kg |  | 300 | 1 000 | 3 000 | 10 000 |  | 30 000 |  | 100 000 | [6 empty] | 4 000 | 10 000 | 20 000 | 30 000 | 20 000
100 kg |  | 160 | 500 | 1 600 | 5 000 |  | 16 000 |  | 50 000 | [6 empty] | 2 000 | 5 000 | 10 000 | 15 000 | 10 000
50 kg | 25 | 80 | 250 | 800 | 2 500 |  | 8 000 |  | 25 000 |  |  | 63 | 125 | 250 | 500 | 1 000 | 2 500 | 5 000 | 7 500 | 5 000
30 kg | [11 empty] | 38 | 75 | 150 | 300 | 600 | 1 500 | 3 000 | 4 500 | 3 000
25 kg | [11 empty] | 31 | 62 | 125 | 250 | 500 | 1 200 | 2 500 | 4 500 | 2 500
20 kg | 10 | 30 | 100 | 300 | 1 000 |  | 3 000 |  | 10 000 |  |  | 25 | 50 | 100 | 200 | 400 | 1 000 | 2 000 | 3 800 | 2 000
10 kg | 5 | 16 | 50 | 160 | 500 |  | 1 600 |  | 5 000 |  |  | 13 | 25 | 50 | 100 | 200 | 500 | 1 000 | 2 200 | 1 000
5 kg | 2.5 | 8 | 25 | 80 | 250 |  | 800 |  | 2 500 |  |  | 6 | 12 | 25 | 50 | 100 | 250 | 500 | 1 400 | 500
3 kg | [11 empty] | 3.8 | 7.5 | 15 | 30 | 60 | 150 | 300 | 1 000 | 300
2 kg | 1 | 3 | 10 | 30 | 100 |  | 300 |  | 1 000 |  |  | 2.5 | 5 | 10 | 20 | 40 | 100 | 200 | 750 | 200
1 kg | 0.5 | 1.6 | 5 | 16 | 50 |  | 160 |  | 500 |  |  | 1.3 | 2.5 | 5 | 10 | 20 | 50 | 100 | 470 | 100
500 g | 0.25 | 0.8 | 2.5 | 8 | 25 |  | 80 |  | 250 |  |  | 0.6 | 1.2 | 2.5 | 5 | 10 | 30 | 50 | 300 | 70
300 g | [11 empty] | 0.38 | 0.75 | 1.5 | 3 | 6 | 20 | 30 | 210 | 60
200 g | 0.1 | 0.3 | 1 | 3 | 10 |  | 30 |  | 100 |  |  | 0.25 | 0.5 | 1 | 2 | 4 | 15 | 20 | 160 | 40
100 g | 0.05 | 0.16 | 0.5 | 1.6 | 5 |  | 16 |  | 50 |  |  | 0.13 | 0.25 | 0.5 | 1 | 2 | 9 | 10 | 100 | 20
50 g | 0.03 | 0.1 | 0.3 | 1 | 3 |  | 10 |  | 30 |  |  | 0.06 | 0.12 | 0.25 | 0.6 | 1.2 | 5.6 | 7 |  | 10
30 g | [11 empty] | 0.037 | 0.074 | 0.15 | 0.45 | 0.9 | 4 | 5 | 44 | 6
20 g | 0.025 | 0.08 | 0.25 | 0.8 | 2.5 |  | 8 |  | 25 |  |  | 0.037 | 0.074 | 0.1 | 0.35 | 0.7 | 3 | 3 | 33 | 4
10 g | 0.02 | 0.06 | 0.2 | 0.6 | 2 |  | 6 |  | 20 |  |  | 0.025 | 0.05 | 0.074 | 0.25 | 0.5 | 2 | 2 | 21 | 2
5 g | 0.016 | 0.05 | 0.16 | 0.5 | 1.6 |  | 5 |  | 16 |  |  | 0.017 | 0.034 | 0.054 | 0.18 | 0.36 | 1.3 | 2 | 13 | 1.5
3 g | [11 empty] | 0.017 | 0.034 | 0.054 | 0.15 | 0.3 | 0.95 | 2 | 9.4 | 1.3
2 g | 0.012 | 0.04 | 0.12 | 0.4 | 1.2 |  | 4 |  | 12 |  |  | 0.017 | 0.034 | 0.054 | 0.13 | 0.26 | 0.75 | 2 | 7 | 1.1
1 g | 0.01 | 0.03 | 0.1 | 0.3 | 1 |  | 3 |  | 10 |  |  | 0.017 | 0.034 | 0.054 | 0.1 | 0.2 | 0.5 | 2 | 4.5 | 0.9
500 mg | 0.008 | 0.025 | 0.08 | 0.25 | 0.8 |  | 2.5 |  |  |  |  | 0.005 | 0.01 | 0.025 | 0.08 | 0.16 | 0.38 | 1 | 3 | 0.72
300 mg | [11 empty] | 0.005 | 0.01 | 0.025 | 0.07 | 0.14 | 0.3 | 1 | 2.2 | 0.61
200 mg | 0.006 | 0.02 | 0.06 | 0.2 | 0.6 |  | 2 |  |  |  |  | 0.005 | 0.01 | 0.025 | 0.06 | 0.12 | 0.26 | 1 | 1.8 | 0.54
100 mg | 0.005 | 0.016 | 0.05 | 0.16 | 0.5 |  | 1.6 |  |  |  |  | 0.005 | 0.01 | 0.025 | 0.05 | 0.1 | 0.2 | 1 | 1.2 | 0.43
50 mg | 0.004 | 0.012 | 0.04 | 0.12 | 0.4 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.042 | 0.085 | 0.16 | 0.5 | 0.88 | 0.35
30 mg | [11 empty] | 0.005 | 0.01 | 0.014 | 0.038 | 0.075 | 0.14 | 0.5 | 0.68 | 0.3
20 mg | 0.003 | 0.01 | 0.03 | 0.1 | 0.3 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.035 | 0.07 | 0.12 | 0.5 | 0.56 | 0.26
10 mg | 0.003 | 0.008 | 0.025 | 0.08 | 0.25 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.03 | 0.06 | 0.1 | 0.5 | 0.4 | 0.21
5 mg | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.028 | 0.055 | 0.08 | 0.2 |  | 0.17
3 mg | [11 empty] | 0.005 | 0.01 | 0.014 | 0.026 | 0.052 | 0.07 | 0.2 |  | 0.14
2 mg | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.025 | 0.05 | 0.06 | 0.2 |  | 0.12
1 mg | 0.003 | 0.006 | 0.02 | 0.06 | 0.2 | [6 empty] | 0.005 | 0.01 | 0.014 | 0.025 | 0.05 | 0.05 | 0.1 |  | 0.1

## Gravimetric Volume I & SVP
Lab Name |  |  | Procedure |  |  |  |  | Scope |  | Back to Lab Scope
Date |  |  | 14 or 26 |  |  |  |  | Gravimetric Calibrations
|  |  | Specify variable source. | [11 empty] | Class | HB 105-3, 105-7, 105-2, ASTM
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
120 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
25 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
5 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
3 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- | SVP
25 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- | SVP
20 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- | SVP
15 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- | SVP
10 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | --- | SVP
20 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
2 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | Qt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | Pt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1/2 | Pt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on N5: Round or Format all expanded uncertainties to two significant digits!
Note on R5: NOTE: Handbook 105-3,105-4, and 105-7 do not require a 1/3 ratio of uncertainty to tolerance. Handbook 105-2 currently retains a 1/3 requirement!!!

## Volume Transfer II
Lab Name |  |  | Procedure |  |  |  |  | Scope |  | Back to Lab Scope
Date |  |  | 19 |  |  |  |  | Volume Transfer
|  |  | Specify variable source. | [11 empty] | Class | HB 105-3, 105-7, 105-2, ASTM
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
5000 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
3000 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
2000 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1500 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1000 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
500 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
300 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
200 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
25 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
5 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | gal | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
[12 empty] | TBD | TBD |  |  |  | ---
20 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
2 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | L | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
[10 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | Qt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | Pt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
1/2 | Pt | [8 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on N5: Round or Format all expanded uncertainties to two significant digits!

## Glassware Tolerances
TOLERANCES FOR GLASSWARE
| US PHARMACOPEIA |  |  | NNN-C-940B | NNN-B0--789a |  | NNN-P-395C |  | NNN-P-00350b |  | OIML R 4 (Edition 1972) |  | ISO 1042:1998 (E) |  |  |  | OIML R 41 (Edition 1981) | ISO/DIS 385 |  | OIML R 43 (Edition 1981) |  | ISO/DIS 4788 |  |  | OIML R 40 Edition 1981 | ISO/WD 835 |  | OIML D 26 1999 | NBS 602 1963 Ed |  |  |  |  | HB 105-2 1996 Ed |  | ASTM E 237 |  |  | ASTM E 287 |  | ASTM E 288 |  | ASTM E 193 | ASTM E 1293
Capacity (ml) | Chapter 31 (Volumetric Apparatus) |  |  | Table I (Graduated Cylinders) | Table I (Burets) | Table II (Burets) | Table I | Table II | Table I (Pipets) |  | (One-mark Volumetric Flasks) |  | Table 1 (One-mark Narrow-necked Volumetric Flasks) |  | Table 2 (One-mark Wide-necked Volumetric Flasks) |  | (Burette) | Table 1 (Burettes) |  | (Standard Graduated Flasks) |  | Table 1 (Cylinders) |  | Table 23 (Cylinders) | Table II | Table 1 (Pipettes) |  | Table I | Table 9 | Table 10 | Table 11 |  | Table 12 | Table 1 |  | Table 1 (Volumetric Flask) |  | Table 44 (Centrifuge Tube) | Table 1 (Buret) |  | Table 1 (Volumetric Flask) |  | Table 1 | Table 1
| Volumetric Flasks | Transfer Pipets | Burets |  | Class A1 | Class B2 | Class A1 | Class B2 | Class A1 | Class B2 | Class A | Class B | Class A1 | Class B2 | Class A1 | Class B2 |  | Class A1 & AS1 | Class B2 | Type A | Type B | Class A | Class B |  |  | Class A & AS1 | Class B2 | Automatic Pipettes | Volumetric Flasks | Transfer Pipets | Burets | Measuring Pipets | Cylindrical Graduates | Cylinder | Flask | Style I | Style II |  | Class A1 | Class B2 | Class A1 | Class B2 |  | Class A | Class B
0.1 (x 0.01) | [9 empty] | 0.005 | [34 empty] | 0.005
0.1 | [37 empty] | 0.01
0.2 (x 0.01) | [9 empty] | 0.008 | [32 empty] | 0.005
0.2 | [37 empty] | 0.02 | [6 empty] | 0.008
0.3 | [37 empty] | 0.03
0.4 | [37 empty] | 0.05
Above 0.4 | [37 empty] | 0.075
0.5 (x 0.01) | [25 empty] | 0.005 | [16 empty] | 0.01 |  | 0.01
0.5 | [6 empty] | 0.006 | 0.012
1 (x 0.01) | [8 empty] | 0.01 | 0.02 | [7 empty] | 0.006 | 0.01 | [5 empty] | 0.006 | 0.006 | 0.01 | [15 empty] | 0.02
1 (x 0.1) | [8 empty] | 0.01 | 0.02 | [33 empty] | 0.01 | 0.02
1 |  | 0.006 |  |  |  |  | 0.006 | 0.012 |  |  |  |  | 0.025 | 0.05 | [14 empty] | 0.01 | [6 empty] | 0.01 | 0.01
2 (x 0.01) | [17 empty] | 0.01 | 0.02
2 (x 0.02) | [24 empty] | 0.01 | 0.01 | 0.02
2 (x 0.1) | [8 empty] | 0.01 | 0.02 | [33 empty] | 0.01 | 0.02
2 |  | 0.006 |  |  |  |  | 0.006 | 0.012 |  |  |  |  | 0.025 | 0.05 | [15 empty] | 0.006 |  | 0.01 |  |  |  | 0.015 | [6 empty] | 0.04
3 | [6 empty] | 0.01 | 0.02 | [20 empty] | 0.015 | [6 empty] | 0.015 | [6 empty] | 0.06
4 | [6 empty] | 0.01 | 0.02 | [27 empty] | 0.02
5 (x 0.01) |  |  |  |  | 0.01 | [12 empty] | 0.01 | 0.02
5 (x 0.02) | [5 empty] | 0.02 | [11 empty] | 0.01 | 0.02
5 (x 0.05) | [24 empty] | 0.03 | 0.03 | 0.05
5 (x 0.1) |  |  |  | 0.1 |  |  |  |  | 0.02 | 0.04 | [11 empty] | 0.05 | 0.1 | [20 empty] | 0.02 | 0.04
5 (x 0.5) | [23 empty] | 0.2
5 |  | 0.01 |  |  |  |  | 0.01 | 0.02 |  |  | 0.025 | 0.05 | 0.025 | 0.05 | 0.04 | 0.08 | [11 empty] | 0.06 | 0.02 | 0.01 | 0.01 | 0.02 | 0.05 |  |  | 0.02 |  |  |  |  | 0.02 | 0.04
10 (x 0.02) |  |  | 0.02 |  | 0.02 | [12 empty] | 0.02 | 0.05
10 (x 0.05) |  |  |  |  | 0.05 | 0.04 | [10 empty] | 0.02 | 0.03 | 0.05 | [19 empty] | 0.02 | 0.04
10 (x 0.1) |  |  |  | 0.1 |  |  |  |  | 0.03 | 0.06 | [9 empty] | 0.05 |  |  |  |  | 0.05 | 0.05 | 0.1 | [16 empty] | 0.03 | 0.06
10 (x 0.2) |  |  |  | 0.1 | [17 empty] | 0.1 | 0.2
10 (x 1.0) | [23 empty] | 0.3
10 | 0.02 | 0.02 |  |  |  |  | 0.02 | 0.04 |  |  | 0.025 | 0.05 | 0.025 | 0.05 | 0.04 | 0.08 | [11 empty] | 0.08 | 0.02 | 0.02 | 0.02 | 0.03 | 0.08 |  |  | 0.02 |  |  |  |  | 0.02 | 0.04
15 | [6 empty] | 0.03 | 0.06
20 (x 0.1) | [19 empty] | 0.08 | [5 empty] | 0.1 | 0.2
20 | [6 empty] | 0.03 | 0.06 |  |  |  |  | 0.04 | 0.08 | 0.06 | 0.12 | [11 empty] | 0.12
25 (x 0.05) | [16 empty] | 0.03 | 0.03 | 0.05
25 (x 0.10) |  |  | 0.03 |  | 0.03 | 0.06 |  |  | 0.05 | 0.1 | [7 empty] | 0.05 | 0.1 | [6 empty] | 0.1 | [12 empty] | 0.03 | 0.06 |  |  |  | 0.05 | 0.1
25 (x 0.2) |  |  |  | 0.3 | [20 empty] | 0.1 | 0.1 | 0.2
25 (x 0.5) |  |  |  | 0.3 | [17 empty] | 0.25 | 0.5
25 (x 1.0) | [23 empty] | 0.5
25 | 0.03 | 0.03 |  |  |  |  | 0.03 | 0.06 |  |  | 0.04 | 0.08 | 0.04 | 0.08 | 0.06 | 0.12 | [11 empty] | 0.12 | 0.03 |  |  |  | 0.14 |  |  | 0.03 |  |  |  |  | 0.03 | 0.06
30 | [29 empty] | 0.03 | 0.03 | 0.05
40 | [6 empty] | 0.05 | 0.1
50 (x 0.10) |  |  | 0.05 |  | 0.05 | 0.1 | [10 empty] | 0.05 | 0.05 | 0.1 | [19 empty] | 0.05 | 0.1
50 (x 0.2) | [19 empty] | 0.12 | [24 empty] | 0.16
50 (x 1.0) |  |  |  | 0.4 | [17 empty] | 0.5 | 1 | [10 empty] | 0.3
50 (x 2.0) | [23 empty] | 1
50 | 0.05 | 0.05 |  |  |  |  | 0.05 | 0.1 |  |  | 0.06 | 0.12 | 0.06 | 0.12 | 0.1 | 0.2 | [11 empty] | 0.15 | 0.05 | 0.05 | 0.05 | 0.08 | 0.2 | [7 empty] | 0.05 | 0.1
100 (x 0.20) |  |  |  |  | 0.1 | 0.2 | [10 empty] | 0.1 | 0.1 | 0.2 | [19 empty] | 0.1 | 0.2
100 (x 0.5) | [19 empty] | 0.2 | [14 empty] | 0.2
100 (x 1.0) |  |  |  | 0.6 | [17 empty] | 0.5 | 1
100 (x 2.0) | [23 empty] | 1
100 | 0.08 | 0.08 |  |  |  |  | 0.08 | 0.16 |  |  | 0.1 | 0.2 | 0.1 | 0.2 | [13 empty] | 0.2 | 0.08 | 0.08 | 0.1 | 0.15 | 0.35 | [7 empty] | 0.08 | 0.16
200 (x 0.5) | [19 empty] | 0.3
200 | [6 empty] | 0.1 | 0.2 |  |  | 0.15 | 0.3 | 0.15 | 0.3 | [13 empty] | 0.4 | 0.1 | 0.1 | [10 empty] | 0.1 | 0.2
250 (x 0.5) | [19 empty] | 0.3 | [14 empty] | 0.3
250 (x 1.0) | [20 empty] | 0.3
250 (x 2.0) |  |  |  | 1.4 | [17 empty] | 1 | 2
250 (x 5.0) | [23 empty] | 2
250 | 0.12 | [9 empty] | 0.15 | 0.3 | 0.15 | 0.3 | [13 empty] | 0.4 |  |  |  |  | 0.65 | [7 empty] | 0.12 | 0.24
300 | [28 empty] | 0.12
500 (x 1.0) | [19 empty] | 0.5 | [14 empty] | 0.5
500 (x 5.0) |  |  |  | 2.6 | [17 empty] | 2.5 | 5
500 (x 10.0) | [23 empty] | 5
500 | 0.15 | [9 empty] | 0.25 | 0.5 | 0.25 | 0.5 | [13 empty] | 0.5 | 0.15 |  |  |  | 1.1 | [7 empty] | 0.2 | 0.4
1 000 (x 1.0) | [19 empty] | 0.8 | [14 empty] | 0.8
1 000 (x 2.0) | [20 empty] | 0.8
1 000 (x 10.0) |  |  |  | 5 | [17 empty] | 5 | 10
1 000 (x 20.0) | [23 empty] | 10
1 000 (or 1 L) | 0.3 | [9 empty] | 0.4 | 0.8 | 0.4 | 0.8 | 0.6 | 1.2 | [11 empty] | 1 | 0.3 |  |  |  | 2 | [7 empty] | 0.3 | 0.6
2 000 (x 2.0) | [19 empty] | 1.2 | [14 empty] | 1.2
2 000 (x 5.0) | [20 empty] | 1.2
2 000 (x 20.0) |  |  |  | 10 | [17 empty] | 10 | 20
2 000 (x 50.0) | [23 empty] | 20
2 000 (or 2 L) | [10 empty] | 0.6 | 1.2 | 0.6 | 1.2 | [13 empty] | 1 | 0.5 |  |  |  | 3.5 | [7 empty] | 0.5 | 1
Above 2 000 | [28 empty] | 1 part in 4 000
2 500 (or 2.5 L) | [27 empty] | 1.2
5 000 (x 5.0) | [19 empty] | 2.5
5 000 (x 10.0) | [20 empty] | 2.5
5 000 (or 5 L) | [12 empty] | 1.2 | 2.4 | [13 empty] | 2.5
10 000 (x 10.0) | [19 empty] | 5
10 000 (x 20.0) | [20 empty] | 5
10 000 (or 10 L) | [27 empty] | 5
25 000 (x 2.0) | [19 empty] | 1.4
25 000 (x 5.0) | [20 empty] | 1.4
1 Serialized or Nonserialized - precision
2 General purpose
3 Corresponds to accuracy Class B in ISO 384
4 Centrifuge Tube, Conical Bottom, Stoppered & Graduated

## Volume Transfer II LPG
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date |  |  | LPG |  |  |  | 21 and 31
|  |  | Specify variable source. | [10 empty] | Class | HB 105-4
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
200 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
100 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
50 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
30 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
25 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
20 | gal | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | LPG
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!

## Length Metric
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date
|  |  | Specify variable source. | [10 empty] | Class | HB 44
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
90 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Tape or Bench
80 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
70 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
60 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
40 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
5 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
4 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | m | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
90 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Rigid Rule
80 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
70 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
60 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
40 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | cm | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!

## Length US Cust
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date
|  |  | Specify variable source. | [10 empty] | Class | HB 44
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
500 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Tape or Bench
400 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
300 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
200 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
90 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
80 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
70 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
60 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
40 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
9 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Tape or Bench
8 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
7 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
6 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
5 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
4 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | ft | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
96 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Rigid Rule
72 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
60 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
48 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
36 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
24 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
18 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
12 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
6 | in | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!

## Temperature
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date
|  |  | Specify variable source. | [10 empty] | Class | 105-6, HB 44, ASTM
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
200 to 250 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
150 to 200 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 to 150 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 to 100 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 to 50 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
20 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 to 20 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Ice Point | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
0 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
-10 | °C | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
350 to 400 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
300 to 350 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
250 to 300 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
212 to 250 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
212 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
120 to 212 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
68 to 120 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
68 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 to 68 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Ice Point | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
32 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
15 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10 | °F | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!

## Time & Frequency
Lab Name |  |  | Procedure |  |  |  | Scope |  |  | Back to Lab Scope
Date |  |  | Time and Frequency |  |  |  | 24 or 22
|  |  | Specify variable source. | [10 empty] | Class | HB 44, 105-5, SP 960
Nominal | Units | us | sp | uo | uo | uo | uo | uo | uc | d.f. | k | U | Units | Std Used | Tolerance | Pn | Comments
24 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Timing Devices
12 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
6 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1 | h | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
10000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | --- | Frequency, Tuning Forks
9000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
8000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
7000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
6000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
5000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
4000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
3000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
2000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
1000 | Hz | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
100 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
90 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
80 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
70 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
60 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
50 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
40 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
30 | mph | [7 empty] | 0 |  | TBD | TBD |  |  |  | ---
Note on C5: Administrator: Uncertainty of Standard at k=1 from calibration report.
Note on D5: Administrator: Standard deviation of the process from check standard data over time, estimate from range chart, or pooled standard deviations.
Note on M5: Round or Format all expanded uncertainties to two significant digits!