## Documentation
Documentation
This workbook was developed by Dan Wright, Washington State Department of Agriculture Metrology Laboratory, and updated February 2007. It follows the procedure documented in NIST HB 145, SOP 14. NIST/WMD has evaluated this spreadsheet and provided useful feedback in its development.
You can modify this workbook for your own laboratory. However, insure you verify and validate in-accordance-with your quality system requirements.
The workbook and all individual worksheets are password protected. The default password is metrology. After you modify any of the worksheets, be sure to apply a password to all worksheets and the workbook. (Don’t forget your password!)
Please contact me if you need help, have any recommends or find any suspected errors at: dwright@agr.wa.gov
Follow the instructions on the Standards worksheet to input your standards information.

## Report
YOUR
LABORATORY
ADDRESS
AND CONTACT INFORMATION
Report of Calibration #:
Calibration Date |  |  |  | Calibration Due Date
Artifact
Test Item….…: |  |  | Manufacture...:
Volume…........: |  |  | Material……....:
Serial Number.: |  |  | Delivery Type.:
Date Received: |  |  | Specification...:
Submitted By
Name & Address: |  |  | Purchase Order #:
Point of Contact..:
Artifact Calibration Results
The was found � gal @ �F
Calibration Notes
#N/A
#VALUE!
● The artifact described above was calibrated using ASTM Type III or IV Reagent Water.
Calibration Conditions
Technician….….…....: |  |  |  | Procedure: NIST HB 145, SOP 14
Condition of Artifact:
Temperature…..….....: |  | Pressure |  | Humidity:
Water Temperature...:
Laboratory Mass Standard(s) Used
Description | Serial Number | Report Number | Date Calibrated | Date Calibration Due
Laboratory Environmental Standard(s) Used
Description | Serial Number | Report Number | Date Calibrated | Date Calibration Due
Relevant Information
● In-accordance-with ISO/IEC FDIS 17025, General Requirements for the Competence of Testing and Calibration Laboratories, paragraph 5.10.4.4 ‘A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the client. This requirement may be superseded by legal regulations.’
● The results listed in this report relate only to the artifacts described and extent of calibrations performed.
Traceability Statement
The item(s) listed above have been compared to the Standards of the State of YOUT-STATE. The Standards of the State of YOUR STATE are traceable to the National Institute of Standards and Technology (NIST) and are part of a comprehensive measurement assurance program for ensuring continued accuracy and measurement traceability within the level of uncertainty reported by this laboratory. The report of calibration number identified in the title of this report is the unique report number to be used in referencing measurement traceability for the artifact(s) identified in this report only.
Uncertainty Statement
The combined standard uncertainty includes uncertainties reported for the standard, uncertainties associated with the measurement process, uncertainties for any observed deviations from reference values which are less than surveillance limits, and other uncertainties associated with the particular artifact (i.e., reading meniscus, air buoyancy corrections, etc.). A component for viscosity is not included in the uncertainty budget. The combined standard uncertainty is multiplied by k, a coverage factor of 2, to give the expanded uncertainty (which defines an interval with an approximate 95 percent level of confidence). The expanded uncertainty presented in this report is consistent with NIST Technical Note 1297.
Certification Statement
Accredited by the National Institute of Standards and Technology (NIST) through the National Voluntary Laboratory Accreditation Program (NVLAP) for the specified scope of accreditation under lab code YOUR-LAB-CODE. This laboratory meets the requirements of ISO/IEC 17025 and ANSI/NCSL Z540-1. This report may not be used to claim product endorsement by NVLAP or any other government agency, and may not be reproduced, except in full, without written approval from the laboratory.
State Metrologist

## Data Entry
Data Entry | [6 empty] | Report Number:
Customer Information
Name and Address (use Alt+Enter for line breaks) | [6 empty] | PO#
POC
Phone
Artifact Information
Description |  |  |  |  | Date Received
Manufacture |  |  |  |  | Nominal Value (gallon)
Material |  |  |  |  | Serial Number
Cubical Coefficient of Exp. (/ºC) = |  |  |  |  | Artifact Specifications
Artifact Condition |  |  |  |  | Reference Temperature (ºF)
Type of Volume Container (To Deliver - Wet Calibration or To Contain - Dry Calibration)
Was a Calibration Interval Requested or is one Required? | [5 empty] | What is the Interval (months)?
Measurement Information
Date of Calibration |  |  |  |  | Technician
Balance ID Code |  |  | Balance |  |  |  | Balance Unit:
Range Mean from CC (mL) |  |  |  |  | No. of Duplicate Sets from CC
Accredited (e.g. NVLAP, A2LA, etc.) Best Uncertainty for this Calibration (mL)
Measurement Option (A = three observations; B = four observations)
Mass Standard(s) Data
Note: If ID Code cell turns RED the standard calibration is void, calibrate before use.
Mass ID | ID Code | Nominal Mass (g) | True Mass Correction (mg) |  | Us (mg) |  | k-factor | Density (g/cm3) | Uρs (g/cm3) k=1
Environmental Standard(s) Data
Note: If ID Code cell turns RED the standards calibration is void, calibrate before use.
Parameter |  |  | ID Code | Correction |  | Accuracy |  | Unit
Temperature, Water
Temperature, Air
Barometric Pressure
Humidity
Observations
Trial 1
Starting Time and Environmental Conditions |  |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
Water Temperature (ºC)
!!! Zero Balance Before Each Observation !!!
Ending Time and Environmental Conditions |  |  |  |  | Water Temperature (ºC)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)
Trial 2
Starting Time and Environmental Conditions |  |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
Water Temperature (ºC)
!!! Zero Balance Before Each Reading !!!
Ending Time and Environmental Conditions |  |  |  |  | Water Temperature (ºC)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)
Note on A20: M3 - Mettler AT201 S1 - Sartorius CC50000 S4 - Sartorius CC30002 S5 - Sartorius CC5001
Note on B27: A31 - 10 kg W13 - 1 g A32 - 20 kg W14 - 2 g A33 - 30 kg W15 - 3 g W01 - 1 mg W16 - 5 g W02 - 2 mg W17 - 10 g W03 - 3 mg W18 - 20 g W04 - 5 mg W19 - 30 g W05 - 10 mg W20 - 50 g W06 - 20 mg W21 - 100 g W07 - 30 mg W22 - 200 g W08 - 50 mg W23 - 300 g W09 - 100 mg W24 - 500 g W10 - 200 mg W25 - 1 kg W11 - 300 mg W26 - 2 kg W12 - 500 mg W27 - 3 kg W28 - 5 kg
Note on D39: H1 - Humidity Probe SN: V3530001 H2 - Hygrometer SN: Q084229 H3 - Hygrometer SN: Q080537 P1 - Barometer SN: V3530001 P2 - Barometer SN: Y3250001 T1 - Digital Thermometer SN: 86270053 T2 - Temperature Probe SN: V3530001 T3 - Themometer, Liquid-in-Glass SN: 1V9658 T4 - Themometer, Liquid-in-Glass SN: 1V9649

## CIPM Air Density
CIPM Air Density
Air Density Calculations, CIPM 1981/91 Formula, R. S. Davis
Air Density Calculation Run 1 | Do not delete these cells! They are used to calculate air density.
Barometric Pressure (mm Hg)
Temperature (ºC) |  | Pa
Relative Humidity (%) |  | °K
Calculated Air Density (mg/cm3) |  | %
Air Density Calculation Run 2 | Do not delete these cells! They are used to calculate air density.
Barometric Pressure (mm Hg)
Temperature (ºC) |  | Pa
Relative Humidity (%) |  | °K
Calculated Air Density (mg/cm3) |  | %

## Calculations
Calculations
Air Density Determinations
| Run 1 ρa = | g/cm3 | CIPM 1981/1991 Air Density Formula, R.S. Davis
| Run 2 ρa = | g/cm3
Water Density Determinations
| Run 1 ρw = | g/cm3 | Patterson/Morris 1994 Water Density Formula
| Run 2 ρw = | g/cm3
Mass Standards Report Values
| Ms | g | true mass of standard(s)
Volume Determinations
| Trial 1 Vt = | mL | trial 1 & 2 volume at measured temperature
| Trial 2 Vt = | mL
| Trial 1 V @ �F = | mL | trial 1 & 2 volume at reference temperature
| Trial 2 V @ �F = | mL
| V @ �F = | mL | average of trial 1 & 2 at reference temp.
Uncertainty Evaluation
Criteria: The expanded uncertainty must be less than 1/3 of the stated tolerance. For NIST HB 105-3, the expanded uncertainty must be less than 0.07 % of the measured volume.
| 1/3 of Tolerance = | mL
| Expanded Uncertainty = | mL
Tolerance Evaluation
Criteria: The measured volume minus the nominal volume shall not differ more than the maximum permissible error minus the expanded uncertainty.
| Volume Error = | mL
| Tolerance - U (k=2)= | mL
Measurement Control
| Range of Trials = | mL
Results Formatted For Calibration Report
Rounding factor =
V @ �C =
U (k=2) =

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NIST HB 145, SOP 14, Gravimetric Calibration of Volumeteric Ware Using an Electronic Balance
Specify Equation(s)
Measurement Option A |  |  |  | Measurement Option B | [6 empty] | Volume at Reference Temperature
where
where | [10 empty] | Vref = | Volume at reference temperature
Vt = | Result at measured temerature | [9 empty] | α = | Coefficient of cubical expansion of vessel
MS = | Mass standard(s) slightly more than the mass of the filled vessel | [9 empty] | t = | Water temperature
MS1 = | Mass standard(s) slightly more than the mass of the empty/drained vessel | [9 empty] | tref = | Reference temperature
MS2 = | Mass standard(s) slightly more than the mass of the filled vessel
On = | Balance Observations
ρa = | Density of air (CIPM equation)
ρs = | Density of mass standards
ρw = | Density of water (Patterson/Morris equation)
Identify Components Their Uncertainties
Data Values (based on trial 1) | [6 empty] | Reported Uncertainties
| g | (Ms) Mass standard(s) | [5 empty] | g | Uncertainty of mass standard(s), Ms, from calibration report, normal distribution, k =1
| g/cm³ | (ρs) Density of mass standard(s) - Ms | [5 empty] | g/cm³ | Uncertainty of the density of the standard(s), ρs, from the calibration report or from OIML R 111-1: 2004(E) para. B.7.9.3 and Table B.7, uniform distribution
| g | (O1) Balance Obsevation 1 | [5 empty] | g | (sp) Uncertainty associated with the measurement process from the mean range divided by the d*2 value from NIST HB 145, Table 9.1 (value divided by 1.003 mL/g to convert sp from mL to g), normal distribution
| g | (O2) Balance Obsevation 2 | [5 empty] | g | Uncertainty of balance observation 2, included in sp and entered as u(O1 (sp))
| g | (O3) Balance Obsevation 3 | [5 empty] | g | Uncertainty of balance observation 3, included in sp and entered as u(O1 (sp))
| g/cm³ | (ρa) Density of air using the CIPM 1981/1991 Air Density Formula, R.S. Davis | [5 empty] | g/cm³ | Given maximum uncertainty, as long as instrument accuracies for temperature ≤ 0.5 ºC, for barometer ≤ 1.0 mm Hg, and for hygrometer ≤ 10 %, uniform distribution
| g/cm³ | (ρw) Density of water using the Patterson/Morris 1994 Water Density Formula | [5 empty] | g/cm³ | Combined standard uncertainty of the Patterson Morris equation, [J. Res., NIST 97, 335 (1992), 10ppm (0.0000 10 g/cm³)], and possible uncorrected systematic error, k = 1
| /ºC | (α) Cubic Coefficient of Expansion for the vessel material | [5 empty] | /ºC | NBS Report 10 081, Uncertainties in the calibration of large vessels, p. 10 (0.000009 /ºC or 0.000005 /ºF), uniform distribution
| ºC | (tw) Temperature of water in vessel | [5 empty] | ºC | Accuracy for thermometer used for measuring water temperature (NIST HB 145, SOP 14, accuracy must be ≤ 0.1 ºC), uniform distribution
| ºC | (tref) Reference Temperature for this calibration | [5 empty] | ºC | Treated as constant, no uncertainty
Input Table -Measurement Option A (Convert Component Uncertainties For Use In Kragten Spreadsheet)
Input description |  | Variable name, symbol | Value | units | Reported uncer. | Type | Distribution | Factor to normalize | Standard uncer., ui | Rel. ui (%) | df
Mass Standards (matching filled container) |  | Ms |  | g |  | B | uc | 1 | #VALUE! | #VALUE!
Mass Standards Density (matching filled container) |  | ρs |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Mass Standards Balance Observation |  | O1 (sp) |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Empty/Drained Vessel Balance Observation |  | O2 |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Filled Vessel Balance Observation |  | O3 |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Air Density |  | ρa |  | g/cm³ |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Water Density |  | ρw |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Vessel Cubic Coefficient of Expansion |  | α |  | /ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Final Water Temperature |  | tw |  | ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Reference Temperature |  | tref |  | ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Kragten Spreadsheet - Measurement Option A (Calculate Each Component Uncertainty Impact, Combine and Expand)
|  | u(Ms) | u(ρs) | u(O1 (sp)) | u(O2) | u(O3) | u(ρa) | u(ρw) | u(α) | u(tw) | u(tref)
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
df
Measurement Equation Inputs | Value
Ms |  | #VALUE!
ρs |  |  | #VALUE!
O1 (sp) |  |  |  | #VALUE!
O2 |  |  |  |  | #VALUE!
O3 | [5 empty] | #VALUE!
ρa | [6 empty] | #VALUE!
ρw | [7 empty] | #VALUE!
α | [8 empty] | #VALUE!
tw | [9 empty] | #VALUE!
tref | [10 empty] | #VALUE!
Result | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ciui
uc | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | (ciui)2
df | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | <--rel (ciui)2 | #VALUE!
k | 2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ci | ∑ (ciui)2
U | #VALUE! | mL | #VALUE! | (Urelative, %)
Weight of Each Component Uncertainty
Input Table -Measurement Option B (Convert Component Uncertainties For Use In Kragten Spreadsheet)
Input description |  | Variable name, symbol | Value | units | Reported uncer. | Type | Distribution | Factor to normalize | Standard uncer., ui | Rel. ui (%) | df
Mass Standards (matching filled container) |  | Ms1 |  | g |  | B | uc | 1 | #VALUE! | #VALUE!
Mass Standards Density (matching filled container) |  | ρs1 |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Mass Standards (matching filled container) |  | Ms2 |  |  |  | B | U | 1 | #VALUE! | #VALUE!
Mass Standards Density (matching filled container) |  | ρs2 |  |  |  | B | uc | 1 | #VALUE! | #VALUE!
Mass Standards Balance Observation |  | O1 (sp) |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Empty/Drained Vessel Balance Observation |  | O2 |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Filled Vessel Balance Observation |  | O3 |  | g |  | A | Normal | 1 | #VALUE! | #VALUE!
Filled Vessel Balance Observation |  | O4 |  |  |  | A | Normal | 1 | #VALUE! | #VALUE!
Air Density |  | ρa |  | g/cm³ |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Water Density |  | ρw |  | g/cm³ |  | B | uc | 1 | #VALUE! | #VALUE!
Vessel Cubic Coefficient of Expansion |  | α |  | /ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Final Water Temperature |  | tw |  | ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Reference Temperature |  | tref |  | ºC |  | B | Uniform | 1.7321 | #VALUE! | #VALUE!
Kragten Spreadsheet - Measurement Option B (Calculate Each Component Uncertainty Impact, Combine and Expand)
|  | u(Ms1) | u(ρs1) | u(Ms2) | u(ρs2) | u(O1 (sp)) | u(O2) | u(O3) | u(O4) | u(ρa) | u(ρw) | u(α) | u(tw) | u(tref)
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
df
Measurement Equation Inputs | Value
Ms1 |  | #VALUE!
ρs1 |  |  | #VALUE!
Ms2 |  |  |  | #VALUE!
ρs2 |  |  |  |  | #VALUE!
O1 (sp) | [5 empty] | #VALUE!
O2 | [6 empty] | #VALUE!
O3 | [7 empty] | #VALUE!
O4 | [8 empty] | #VALUE!
ρa | [9 empty] | #VALUE!
ρw | [10 empty] | #VALUE!
α | [11 empty] | #VALUE!
tw | [12 empty] | #VALUE!
tref | [13 empty] | #VALUE!
Result | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ciui
uc | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | (ciui)2
df | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | <--rel (ciui)2 | #VALUE!
k | 2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | ci | ∑ (ciui)2
U | #VALUE! | mL | #VALUE! | (Urelative, %)
Weight of Each Component Uncertainty
Determine Expanded Uncertainty For This Calibration
Expand uncertainty for this calibration (k=2) = | [5 empty] | mL
Accredited best uncertainty (e.g. NVLAP, A2LA, etc.) (k=2) = |  |  |  |  | 0 | mL
Uncertainty to be reported for this calibration (k=2) = | [5 empty] | mL

## Standards
Standards
Stds.Table
ID Code | Description | Serial Number | Nominal Value (gram) | True Mass Correction (mg) | Uncertainty (mg) | k factor | Density (g/cm3) | Density Uncertainty k=1 (g/cm3) | Report Number | Date Calibrated | Date Calibration Due
Enviro.Stds Vlookup Table
ID Code | Description | Serial Number | Correction | Accuracy | Units | Report # | Cal Date | Due Date
Balance Vlookup Table
ID Code | Balance | Units
To Change, Add, or Delete Information:
a. Unprotect the worksheet (Select Tools, Protection, Unprotect sheet... enter password and select the OK button).
1. To change a standards information place the curser anywhere in the desired standards list and make any necessary changes.
2. To add a standard, place the curser anywhere in the desired standards list. Enter the new standard on the line with the *.
3. To delete a standard just highlight the row that contains the standard and select Edit, Delete.
4. After entering or deleting standards, click on the arrow in the ID Code block and sort the list in ascending order.
Note: The tables must be sorted by the ID Code in ascending order for the vlookup function to work properly in the formulas!

## Tables & Lists
Tables & Lists |  |  |  |  | Tables & Lists | [8 empty] | Tables & Lists | [6 empty] | Tables & Lists | [5 empty] | Tables & Lists | [6 empty] | Tables & Lists
Rounding Table (Rnd.Table) (Used to round the uncertainty to two significant digits, and rounds the as found/left values to the same number of decimal places that are in the uncertainty.) |  | NISTIR 6969 Table 9.1 (Table_9.1) Use of Range to Estimate Standard Deviation (two replicates) |  | Pick Lists for Data Validation Cells | Volume Conversion Table (VolTable) & Nominal Value Unit (Nom_Val.List) | [8 empty] | Cubical Coefficient of Expansion Table (CCE.Table) & Material List (Materal.List) |  | ASTM E 288-06 Standard Specification For Laboratory Glass Volumetric Flasks |  |  |  |  | ASTM E 1272-02 Standard Specification For Laboratory Glass Graduated Cylinders |  |  | OIML R 43-1981 Standard Graduated Glass Flasks For Verification Officers |  |  | NIST Handbook 105-2 Specifications and Tolerances for Field Standard Measuring Flasks |  |  |  | NIST Handbook 105-3 Specifications and Tolerances for Graduated Neck Type Volumetric Field Standards |  |  | Tolerance Vlookup Table (tolerance)
0 | TableError | 1 | 1.41 | Reference Temperature (ºF) | NomValueUnit | Vref Trial 1 | Vref Trial | Vref | U (k=2) | RndFactor | Tolerance | Symbol | Nominal Value (mL) | Material | cm³/ºC |  | Class A | Class B | Class A "Wide Mouth" | Class B "Wide Mouth" |  | Class A | Class B |  | To Contain | To Deliver | Capacity | Capacity (mL) | Tolerance (mL) | Type | Capacity | Capacity (mL) | Tolerance (mL) | Specification | Tolerance (mL)
1e-12 | 13 | 2 | 1.32 | Reference Temperature (ºC) | Nominal Value (cubic inch) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | in³ | 0 | Glass, Borosilicate (TICA) | 1e-05 | Capacity (mL) | Tolerance (mL) | Tolerance (mL) | Tolerance (mL) | Tolerance (mL) | Capacity (mL) | Tolerance (mL) | Tolerance (mL) | Capacity (mL) | Tolerance (mL) | Tolerance (mL) | 50 mL | 50 | 0.3 | cylinder | 1 gal | 3785.412 | 0.75708 | ASTM E 1272, Class A | #N/A
1e-11 | 12 | 3 | 1.23 |  | Nominal Value (fluid ounce) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | fl oz | 0 | Glass, Borosilicate (TICB) | 1.5e-05 | 5 | 0.02 | 0.04 | 0.08 | 0.16 | 5 | 0.05 | 0.1 | 10 | 0.05 | 0.075 | 2 fl oz | 59.1471 | 0.3 | cylinder | 5 L | 5000 | 1 | ASTM E 1272, Class B | #N/A
1e-10 | 11 | 4 | 1.21 | Option A | Nominal Value (gallon) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | gal | 0 | Glass, Soda-Lime | 2.5e-05 | 10 | 0.02 | 0.04 | 0.08 | 0.16 | 10 | 0.1 | 0.2 | 20 | 0.08 | 0.12 | 100 mL | 100 | 0.2 | flask | 10 L | 10000 | 2 | ASTM E 288, Class A | #N/A
1e-09 | 10 | 5 | 1.19 | Option B | Nominal Value (gill) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | gi | 0 | Plastic, Polycarbonate | 0.00045 | 25 | 0.03 | 0.06 | 0.08 | 0.16 | 25 | 0.17 | 0.34 | 50 | 0.12 | 0.18 | 1 gill | 118.2941 | 0.2 | flask | 5 gal | 18927.06 | 3.7854 | ASTM E 288, Class A "Wide Mouth" | #N/A
1e-08 | 9 | 6 | 1.184 |  | Nominal Value (liquid pint) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | pt | 0 | Plastic, Polypropylene | 0.00024 | 50 | 0.05 | 0.1 | 0.08 | 0.16 | 50 | 0.25 | 0.5 | 100 | 0.2 | 0.3 | 1/2 pint | 236.5882 | 0.3 | flask | 10 gal | 37854.12 | 7.5708 | ASTM E 288, Class B | #N/A
1e-07 | 8 | 7 | 1.178 | To Deliver | Nominal Value (liquid quart) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | qt | 0 | Plastic, Polystyrene | 0.00021 | 100 | 0.08 | 0.16 | 0.1 | 0.2 | 100 | 0.5 | 1 | 200 | 0.3 | 0.4 | 250 mL | 250 | 0.3 | flask | 20 L | 20000 | 4 | ASTM E 288, Class B "Wide Mouth" | #N/A
1e-06 | 7 | 8 | 1.172 | To Contain | Nominal Value (liter) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | L | 0 | Silica, Fused (Quartz) | 1.6e-06 | 200 | 0.1 | 0.2 | 0.2 | 0.4 | 250 | 1 | 2 | 250 | 0.3 | 0.45 | 1 pt | 473.1765 | 0.4 | flask | 50 L | 50000 | 10 | NIST HB 105-2 | #N/A
1e-05 | 6 | 9 | 1.166 |  | Nominal Value (milliliter) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | #N/A | mL | 0 | Steel, Prover, Low Carbon | 3.35e-05 | 250 | 0.12 | 0.24 | 0.2 | 0.4 | 500 | 2 | 4 | 500 | 0.5 | 0.75 | 1 qt | 946.3529 | 0.7 | flask | 25 gal | 94635.3 | 18.9271 | NIST HB 105-3 | 0.1
0.0001 | 5 | 10 | 1.16 | Yes | [9 empty] | Steel, Stainless | 4.77e-05 | 500 | 0.2 | 0.4 |  |  | 1000 | 3 | 6 | 1000 | 0.8 | 1.2 | 1000 mL | 1000 | 0.8 | flask | 100 L | 100000 | 20 | Not Applicable | N/A
0.001 | 4 | 11 | 1.158 | No | [9 empty] | Steel, Terneplate | 3.5e-05 | 1000 | 0.3 | 0.6 |  |  | 2000 | 6 | 12 | 2000 | 1.2 | 1.8 | 1/2 gal | 1892.706 | 1 | flask | 200 L | 200000 | 40 | OIML R 43 | #N/A
0.01 | 3 | 12 | 1.156 | [12 empty] | 2000 | 0.5 | 1 |  |  | 4000 | 14.5 | 29 | 2500 | 1.4 | 2.1 | 2000 mL | 2000 | 1.2 | flask | 100 gal | 378541.2 | 75.7082 | User Specified
0.1 | 2 | 13 | 1.154 | Graduated Cylinder | [19 empty] | 5000 | 2.5 | 3.75 | 1 gal | 3785.412 | 1.2 | flask | 500 L | 500000 | 100
1 | 1 | 14 | 1.152 | Volumetric Flask | [19 empty] | 10000 | 5 | 7.5 |  |  |  |  | 200 gal | 757082.4 | 151.4165
10 | 0 | 15 | 1.15 | Volumetric Container | [26 empty] | 1000 L | 1000000 | 200
100 | -1 | 16 | 1.148 | Slicker Plate Standard | [26 empty] | 500 gal | 1892706 | 378.5412
1000 | -2 | 17 | 1.146 | Prover | [26 empty] | 2000 L | 2000000 | 400
10000 | -3 | 18 | 1.144 | Test Measure | [26 empty] | 3000 L | 3000000 | 600
100000 | -4 | 19 | 1.142 | [27 empty] | 1000 gal | 3785412 | 757.0824
1000000 | -5 | 20 | 1.14 | [27 empty] | 5000 L | 5000000 | 1000
10000000 | -6 | 21 | 1.139 | [27 empty] | 1500 gal | 5678118 | 1135.6236
100000000 | -7 | 22 | 1.138
1000000000 | -8 | 23 | 1.137
10000000000 | -9 | 24 | 1.136
100000000000 | -10 | 25 | 1.135
1000000000000 | -11 | 26 | 1.134
10000000000000 | -12 | 27 | 1.133
100000000000000 | -13 | 28 | 1.132
1000000000000000 | TableError | 29 | 1.131
|  | 30 | 1.13
|  | 31 to infinity | 1.13
Calibration Report Statements
● The delivers the stated volume at the reference temperature from the specified graduation line when it is emptied gradually within a 30 (± 5 seconds) pour and a 10-second drain while held at a 10º to 15º angle from vertical.
● The , when cleaned and dried, contains the stated volume at the reference temperature when filled to its nominal graduation line. The must always be cleaned and dried between successive uses for purposes of accuracy.
● The delivers the stated volume at the reference temperature is when it is drained for a 30 (± 5) second period after cessation of the main flow.