## Slide 1: Single Phase Example
Ferrite Phase from SRM2686a
Nitric acid/methanol extraction
We will be using only a subset of the features provided by Profex/BGMN to run a basic refinement for quantitative analysis; we leave it to you to explore the more advanced refinement features after you have had some experience.

## Slide 2: Introduction to the Rietveld Method, R.A. Young
Provides background on the method, its historical development and subsequent application to laboratory X-ray diffraction
Refining parameters in models for the structure, specimen and instrument effects on the diffraction pattern
Least-squares refinements are carried out until the best fit is obtained between the entire observed powder diffraction pattern and calculated pattern based upon simultaneously refined models for crystal structures and instrument effects
Potential systematic errors in include preferred orientation, background, anisotropic peak broadening, profile shapes, absorption, specimen displacement, specimen transparency, extinction, 2-theta-Zero error, graininess, beam instability
Introduction to the Rietveld Method, R.A. Young in The Rietveld Method, IUCR Monograph on Crystallography #5

## Slide 3: Profex and BGMN*
BGMN Author: Dr. Joerg Bergmann
http://www.bgmn.de
Profex Author: Nicola Dobelin
http://profex.doebelin.org
GPL Open Source, Windows, OS X, Linux; front end for BGMN and for Fullprof 2k
Fundamental Parameters, adaptable to different instruments, multi-platform, robust
Profex Creates a control (input) file
Copies structure models to working folder
Copies the instrument configuration files
(they allow calculation of the instrument contribution to the diffraction peak shapes based upon the X-ray optics)
Allows editing of the working versions of the structure files and saves changes/refined values
Döbelin, N., Kleeberg, R., „Profex: a graphical user interface for the Rietveld refinement program BGMN„, Journal of Applied Crystallography 48 (2015), 1573-1580.
DOI: doi:10.1107/S1600576715014685
* Adapted from Profex tutorials with some modification

## Slide 4: Profex/BGMN
In place of actual powder reference specimens for calibration, crystal structure models allow calculation of phase diffraction patterns. These models include space group, lattice parameters, atomic positions, atomic site occupancies, atomic vibrational parameters.
The code fits a multivariable model of structure, peak profile, background to the observed data

## Slide 5: Initialize Profex
Diffraction pattern display region
Refinement
Progress plot
Plot Options
Select/deselect
Display global parameters and goals

## Slide 6: Profex/BGMN Instrument Configuration
BGMN.de
% BGMN Device Configuration File for Bruker D8 % -------------------------------------------- % Created by Nicola Doebelin, RMS Foundation,
% Switzerland % November 12, 2012 % % Device Configuration: % - Detector: LynxEye % - Radiation: CuKa, Ni-filtered % - Soller Slits: 2.5 degrees % - Divergence Slit: fixed, 0.6 mm % - Anti-Scatter Slit: fixed, 6.76 mm % -Goniometer Radius: 217.5 mm
We will use File=D8-06mm.geq, a modification of the D8 file provided.
Device functions are available for a wide variety of instruments and configurations, though some may require editing and re-saving. Different configurations will require a new device file!

## Slide 7: Profex Configuration
Control file (.SAV)
References model structures from the Cement_Structures Folder
A copy of the structure is stored in the same folder as the data
Instrument profile
Wavelength
Refined Phases
Data File
Set two-theta angles for refinement
(remove the leading “%”)
Zero-Error (really should be zero)
Sample Displacement Parameter
Phase Quantity Calculation

## Slide 8: Profex Configuration
File, refinement details
.LST file: Refinement Summary
refinement statistic measures
refinement global parameter (shift) and phase fraction goals
refinement data, by phase

## Slide 9: Profex Structure Files (.str)
Provided with the installation but may also be transcribed from databases
Element Occupancy Wyckoff Position Fractional Coordinates Thermal Displacement Parameter
E = CA+2(1) Wyckoff=c x=0.0273 y=0.1087 z=0.4920 TDS=0.00971169

## Slide 10: Profex Structure Files (.str)
Provided with the installation but may also be transcribed from databases
Fixed Parameter
A=0.5557
Name Value
Refined Parameter
PARAM=A=0.5557
Refine Name Value
Refined Parameter With Limits
PARAM=A=0.5557_0.551^0.563
Refine Name Value Lower Upper
B1: peak broadening from crystallite size (diffracting domains)
GEWICHT = Scale, Weight
Microabsorption correction, (ifthenelse(ifdef(d),exp(my*d*3/4),1), is not being used

## Slide 11: Sources of Crystal Structure Data
American Mineralogist Crystal Structure Database
http://rruff.geo.arizona.edu/AMS/amcsd.php
Mincryst
http://database.iem.ac.ru/mincryst/
Crystallography Open Database
http://www.crystallography.net/cod/
Cements Crystal Structure Database
http://www.nist.gov/el/building_materials/
inorganic/gsas1.cfm
Commercial:
ICDD
https://icdd.com
ICSD
https://icsd.fiz-karlsruhe.de/
.CIF File: Ferrite

## Slide 12: American Mineralogist Crystal Structure Database
http://rruff.geo.arizona.edu/AMS/amcsd.php
http://jp-minerals.org/vesta/en/
Search by a number of criteria
Get a listing of structure data
Download .CIF file
Import into Profex *
*Use Vesta to view structure and get the Wyckoff positions

## Slide 13: Import CIF Example
Profex->File->Import Structure File
Point toward selected .CIF file
At the same time, open that CIF file using VESTA
  necessary to get the Wyckoff position notation
Add PHASE name: PHASE=Ferrite //
Add phase name to GOAL: GOAL:Ferrite=GEWICHT …
Add atom charge (CA+2, AL+3, FE+3, O-2)
Save to Structures Folder; include extension .str
Profex
Vesta

## Slide 14: Load Data File: File->Insert Scan
Will read numerous formats
- We will use the Bruker .RAW

## Slide 15: Corundum-Silicon File Added
Add/Remove Phases
Display Phases
File and header

## Slide 16: Structure Files
Profex->Preferences
Paths to structure files
Will list according to list order
CementStructures
FlyAsh
InternalStandards
Structures
Once added, select Display to see the copied data, make changes if desired
Add or Remove Phase Windows
Specify Instrument File
Overwrite files in working folder

## Slide 17: Profex: .raw, .sav, .str display
You could decide to refine or fix selected parameters
Typically we will refine lattice parameters, crystallite size (B1), scale (Gewicht), specimen displacement (EPS1). Background is automatically refined
Parameters like thermal (TDS) and site occupancy are best left fixed
Preferred orientation is refined judiciously and it is always best to obtain data with minimal preferred orientation.
A good rule is less is better – do no harm!
Run, Abort Refinement

## Slide 18: Initialize Profex, Insert Scan

## Slide 19: Initialize Profex, Insert Scan

## Slide 20: Phase d&I Search for Phase Identification
Not accounted for with ferrite
?
Keep in mind that this is a selective extraction and some secondary products might occur… but in this case we have a second crystalline phase
Look for a portlandite peak at 18 degrees two-theta, possibly indicating some secondary product resulting from the extraction

## Slide 21: Add Phase
Select one

## Slide 22: Display Structure
For the most part, refined parameters are pre-set and indicated by PARAM
Some parameters are automatic: Background, Scale (gewicht)
Other parameters
Lattice parameters a, b, and c
B1 (Lorentzian line shape related to crystallite size broadening)
Goals: Computed values, crystallite size, adjustments to scale
Preferred orientation: SPHARn, where n=0 (none), 2, 4, 6 Ideally set to 0!
Generally you will not need to initiate or deselect any refined parameters aside from preferred orientation, and only if warranted by peak intensity mismatch and potential for orientation based upon crystal habit or cleavage properties – alite, perhaps ferrite and belite and, for cements, any of the calcium sulfates and calcite
Initiate refinement
Atom positions (x,y,z) and vibrational (thermal) factors (TDS) are not refined partly because the data range is too limited. Site occupancy might be refined, but the initial values for ferrite are based upon earlier refinements, providing typical industrial values.

## Slide 23: Examine Results: Plot, Parameters, LST
Plot key->
Difference plot reflects areas of mismatch. Some displacements reflect a peak shape error while the distinct peaks in the difference pattern reflect a second phase not included in the refinement. Measure the positions and check the ID table or the d&I display. In some cases, peak shifts of one or more peas may reflect a lattice parameter that has reached a limit.
Check the Global Parameters and Goals or LST file window against the lattice parameters and limits; expand them if necessary
Observed and calculated patterns
difference pattern
background
Display legend
Check box to
Turn on/off
Drag window within plot to expand plot or set values through View->Set Plot Range menu item
Compare refined values to parameter limits, adjust limits if justified

## Slide 24: Search for Missing Phase
Guess, or scroll through the possibilities or,
Use the mouse pointer to measure the position and intensity of each peak and use the identification table; add phase, repeat refinement

## Slide 25: Ferrite and Periclase Refinement
Looks much better; low angle mismatch is common, slight intensity mismatch may be related to sampling preferred orientation, or chemical substitution
We could exclude the low angle region by adding the line WMIN=15 to the .SAV tab
Click on an item to highlight it on the plot
This is useful to look at individual phase contributions and to assess which phase may be mismatched with the raw data
Relative phase proportions !

## Slide 26: Preferred Orientation
Since ferrite can exhibit a tabular habit it is possible it could be oriented, skewing the representative orientation ideal and so, the relative peak intensities. Change the zero to 2 and later 4 if necessary after the variable SPHAR and repeat the refinement to see if it improves. Profex runs the refinement and introduces the orientation correction at the last stages, but only if the phase fraction is above a minimum to make it practical.
Fourth order spherical harmonics for the ferrite phase (GEWICHT=SPHAR4) improves the fit
- In an essentially mono-phase powder this is a likely problem, particularly given the bladed habit of the ferrite phase, but keep in mind that it could also be related to sampling or chemical substitution

## Slide 27: Check Parameters

## Slide 28: Look for additional phases
0.303 nm
0.248 nm
the unaccounted peaks are probably from calcite, a secondary product of the extraction