## Slide 1: 2nd NIST Workshop on Cement Materials Characterization July 12–14, 2017 Lecture Room A, Administration Building NIST, Gaithersburg, Maryland LaKesha Perry LaKesha.Perry@nist.gov Materials and Structural Systems Division, Inorganic Materials Group National Institute of Standards and Technology Gaithersburg, MD
ICP-OES for measuring solution composition
Notes: Need pictures

## Slide 2: DISCLAIMERS
Certain commercial equipment, instruments, or materials are identified in this paper to foster understanding. Such identification does not imply recommendation or endorsement by the National Institute of Standards and Technology, nor does it imply that the materials or equipment identified are necessarily the best available for the purpose.
Any opinions expressed solely reflect the author and do not necessarily represent the National Institute of Standards and Technology.

## Slide 3: OVERVIEW
INTRODUCTION
INSTRUMENTATION
PRINCIPLE
APPLICATIONS
RESULTS
CONCLUSION
FUTURE WORK

## Slide 4: INTRODUCTION
What is ICP-OES?
Inductively coupled plasma optical emission spectroscopy (ICP-OES), also referred to as inductively coupled plasma atomic emission spectrometry (ICP-AES) - is a type of emission spectroscopy that uses the inductively coupled plasma to produce excited atoms and ions that emit electromagnetic radiation at wavelengths characteristic of a particular element.

## Slide 5: INTRODUCTION
Elemental analysis with:
Wide elemental coverage
Very low detection limits(ppb /ppm)
Fast analysis times (all elements at once)
Simple spectra
LIMITS OF DETECTION : IN COMPARISION

## Slide 6: INTRODUCTION
Measure almost any element at ppb to ppm levels in almost any material.
Measure all elements in a single analysis.
High sensitivity
Why ICP-OES?
https://goo.gl/images/6FJQKC

## Slide 7: INTRODUCTION
Multiple test methods for elements of interest
Use of harsh chemicals
Time consuming
Not as accurate
Why ICP-OES?
ASTM C114-15 Standard Test Methods for Chemical Analysis of Hydraulic Cement, ASTM International, West Conshohocken, PA, 2015, https://doi.org/10.1520/C0114-15
Standard Test Methods for Chemical Analysis of Hydraulic Cement
(ASTM C114-15)

## Slide 8: INTRODUCTION
Why ICP-OES?
Inductively coupled plasma optical emission spectroscopy
Multiple elements of interest measured at one time.
Typical analysis time for ICP is ~2-3 minutes per sample. This includes flush time, multiple repeats, etc.
Safer chemicals used.
Typical precision, amongst repeats within an analysis, is ~0.5% for accurate and reliable results.
Typical detection limits are ~ 1-10 parts per billion.

## Slide 9: INSTRUMENTATION
ICP Instrument
Argon
Chiller
Peristaltic Pump
Auto Sampler (optional)
Waste Container
Operation Software
Computer system

## Slide 10: INSTRUMENTATION
  Atomic emission spectroscopy measures the intensity of light emitted by atoms or ions of the elements of interest at specific wavelengths;
  Inductively Coupled Plasma spectrometers use emission spectroscopy to detect and quantify elements in a sample;
  ICP-OES uses the argon plasma (6000 K to10000 K) for atomization and excitation of the sample atoms;
  ICP-OES determines concentrations of most elements.

## Slide 11: PRINCIPLES
Sample solution containing elements
Formation of liquid droplets
DESOLVATION
Atomization
Excitation
Emission of Radiation at a Specific Wave Length
Measure the intensity of emitted radiation

## Slide 12: PRINCIPLES
https://goo.gl/images/hoJq9z
Sample typically injected as a liquid
Nebulizer converts liquid sample into an aerosol
Spray Chamber transports aerosol sample to the plasma torch
Plasma torch vaporizes, atomizes, and ionizes aerosol sample
Transfer Optics focus plasma image onto entrance slit of spectrometer
Wavelength dispersive device of the spectrometer isolates proper emission line
Detector and its components measure intensity of the emission line.
Computer software compiles data

## Slide 13: PRINCIPLES
Radial View (Side-on-Plasma):
more suitable for hard matrices (concentrated samples);
alkali metals (Na, K, Li) calibration is more linear;
less spectral interferences;
lower sensitivity (Limit-of-Detection is higher)
Axial View (End-on-Plasma):
more suitable for light matrices;
alkali metals (Na, K, Li) calibration is less linear;
more spectral interferences;
higher sensitivity (Limit-of-Detection is lower)
https://goo.gl/images/jCqKoS

## Slide 14: APPLICATIONS
Clinical Analysis: metals in biological fluids (blood, urine);
Environmental Analysis: trace metals and other elements in waters, soils, plants, composts and sludges;
Pharmaceuticals: traces of catalysts used; traces of poison metals (Cd, Pb etc)
Industry: trace metal analysis in raw materials; noble metals determination.
Forensic science: gunshot powder residue analysis, toxicological examination
https://goo.gl/images/BXXRyS
Fields of Application

## Slide 15: APPLICATION TO CEMENT BINDERS
Quantifying concentrations of readily soluble components in aqueous solutions mixed with cementitious materials by Inductively Coupled Plasma Optical Emission Spectroscopy (ICP-OES)
OBJECTIVE:
https://goo.gl/images/HxSDLu

## Slide 16: CEMENT BINDERS
MATERIAL:
Characterized 3 Class C and 3 Class F fly ashes:
  Each fly ash is designated by a letter-number combination designating Class F or Class C, and an arbitrary index (C1, C2, etc.)
  The bulk oxide compositions of fly ashes measured by X-ray fluorescence (mass %).
Le Saoût, G.; Kocaba, V.; Scrivener, K. Application of the Rietveld Method to the Analysis of Anhydrous Cement. Cem. Concr. Res. 2011, 41 (2), 133–148.

## Slide 17: CEMENT BINDERS
MATERIAL:
Characterized 3 Ordinary Portland Cements (OPC) Type I and Type II:
  Each OPC is designated by OPC and an arbitrary index (OPC1, OPC2, etc.)
Some bulk oxide and phase compositions of Ordinary Portland Cement measured by X-ray diffraction.

## Slide 18: CEMENT BINDERS
PROCEDURE:
Cementitious mix preparation
Shaker-Mixer used for powder blending to ensure quantitative representation of fly ash sample
Obtained (116.6 ± 0.1) g deionized water to (333.0 ± 0.1) g of cementitious powder for a water-solid mass ratio of 0.35
ASTM C1738 certified constant mixer was used to blend water-solid sample for approximately 30 and 60 seconds at 4000 rpm and 10000 rpm, respectively
(25 ± 1) g aliquots of the sample was divided into falcon tubes and kept at constant temperature of 23.0 °C

## Slide 19: CEMENT BINDERS
PROCEDURE:
Pore solution preparation
Centrifugation was conducted at an interval of time for each individual falcon tube sample
The pore solution was extracted using a syringe and filtered using non-pyrogenic NY 0.20 µm membrane
Pore solution for each time interval was diluted using a 5 % nitric acid and deionized water solution for ICP analysis
https://goo.gl/images/benS7y

## Slide 20: CEMENT BINDERS
PROCEDURE:
ICP-OES analysis
ICP-OES was first tuned according to Manufacturer’s Instructional Handbook
Verified parameters and settings for optimum performance (check Hg lamp position, carrier gas flow rates… etc.)
Used a developed test method using the ICP-OES Manufacturer supplied operating software
Selected optimum spectral emission lines for K, Na, S, Si, Al and Ca

## Slide 21: CEMENT BINDERS
PROCEDURE:
ICP-OES analysis
Calibrated the equipment using an in-house solution containing known concentrations of K, Na, S, Si, Al and Ca
The calibration was then verified using a NIST traceable quality control check standard
An autodilutor/autosampler was used for sample introduction
The analysis of a NIST traceable quality control check standard was conducted after every 14 samples to ensure equipment accuracy
Typical calibration curve

## Slide 22: CEMENT BINDERS
PROCEDURE:
Data reporting
Used Manufacturer software to format repot to include both calibration and sample details including:
Method
Date
View (radial and/or axial)
Sample ID
Emission lines
Concentration
Units
Dilution factor

## Slide 23: RESULTS
Fly ash samples
Chart: Ash F1
| 6.0 | 10.0 | 18.0 | 21.0 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0 | 150.0 | 180.0 | 210.0
K | 3.6672 | 4.2718 | 7.9465 | 4.3275 | 4.334 | 4.3449 | 4.401 | 4.5327 | 4.3433 | 4.5412 | 4.6585 | 4.5442
Na | 55.2739 | 61.9754 | 102.3934 | 61.5549 | 61.2475 | 62.9019 | 61.9667 | 63.6182 | 59.8179 | 63.737 | 64.3576 | 62.319
S | 57.5976 | 62.7475 | 103.5605 | 63.1145 | 60.3004 | 59.1486 | 58.4968 | 56.8522 | 51.1482 | 54.2222 | 55.1307 | 54.5725
Ca | 34.4237 | 38.4109 | 61.6315 | 36.8132 | 34.5385 | 33.4107 | 32.6596 | 31.6824 | 29.6389 | 29.7728 | 29.5382 | 30.429
Si | 0.15604 | 0.12273 | 0.24354 | 0.17316 | 0.17971 | 0.21157 | 0.19742 | 0.20832 | 0.19972 | 0.19914 | 0.22002 | 0.2478
Al | 0.30968 | 0.47791 | 0.54707 | 0.35831 | 0.24308 | 0.25888 | 0.19831 | 0.24382 | 0.11566 | 0.10036 | 0.11263 | 0.094225
Chart: Ash C1
| 5.5 | 8.5 | 13.0 | 15.0 | 17.0 | 21.0 | 25.0 | 30.0 | 33.5 | 35.5
K | 8.2406 | 8.2442 | 8.8594 | 9.0816 | 7.9571 | 9.5394 | 10.1012 | 10.3758 | 10.7555 | 10.3544
Na | 120.2912 | 118.7963 | 124.5337 | 124.9193 | 109.2167 | 127.7554 | 133.2245 | 134.2394 | 138.637 | 130.3696
S | 84.8703 | 68.3445 | 62.5791 | 59.5457 | 51.328 | 57.608 | 56.0996 | 56.1838 | 50.7386 | 39.9969
Ca | 56.9115 | 44.5165 | 38.218 | 35.0325 | 29.6796 | 30.8357 | 27.963 | 25.294 | 20.7337 | 13.1435
Si | 0.0097678 | 0.0098509 | 0.014681 | 0.015156 | 0.023001 | 0.026657 | 0.032437 | 0.051248 | 0.055568 | 0.08327
Al | 0.023139 | 0.024004 | 0.02608 | 0.027797 | 0.031627 | 0.029193 | 0.029736 | 0.040213 | 0.033875 | 0.058003

## Slide 24: RESULTS
Fly ash samples
Chart: S
| 6.0 | 10.0 | 18.0 | 21.0 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0 | 150.0 | 180.0 | 210.0
F1 | 57.5976 | 62.7475 | 103.5605 | 63.1145 | 60.3004 | 59.1486 | 58.4968 | 56.8522 | 51.1482 | 54.2222 | 55.1307 | 54.5725
F2 | 60.7277 | 58.8409 | 59.2931 | 56.1807 | 58.5332 | 57.3252 | 56.6495 | 47.0607 | 54.8802 | 57.8211 | 55.5143 | 55.0912
F3 | 60.7277 | 58.8409 | 59.2931 | 56.1807 | 58.5332 | 57.3252 | 56.6495 | 47.0607 | 54.8802 | 57.8211 | 55.5143 | 55.0912
C1 | 84.8703 | 68.3445 | 62.5791 | 59.5457 | 51.328 | 57.608 | 56.0996 | 56.1838 | 50.7386 | 39.9969
C2 | 62.9918 | 61.919 | 60.919 | 60.1258 | 59.3971 | 59.5135 | 59.1538 | 61.5978 | 54.217
C3 | 98.3305 | 98.7588 | 90.9871 | 67.9422
Chart: Ca
| 6.0 | 10.0 | 18.0 | 21.0 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0 | 150.0 | 180.0 | 210.0
F1 | 34.4237 | 38.4109 | 61.6315 | 36.8132 | 34.5385 | 33.4107 | 32.6596 | 31.6824 | 29.6389 | 29.7728 | 29.5382 | 30.429
F2 | 44.3352 | 45.5853 | 44.3577 | 44.6895 | 43.9202 | 42.9554 | 41.3144 | 39.6743 | 37.496 | 36.5662
F3 | 54.1103 | 51.715 | 52.8744 | 51.7258 | 52.0693 | 52.0543 | 51.2659 | 41.9341 | 48.8223 | 50.2587 | 49.4702 | 49.4386
C1 | 56.9115 | 44.5165 | 38.218 | 35.0325 | 29.6796 | 30.8357 | 27.963 | 25.294 | 20.7337 | 13.1435
C2 | 25.1194 | 22.5252 | 20.5241 | 19.1277 | 18.0448 | 17.5782 | 16.1152 | 15.0498 | 11.6672
C3 | 18.3301 | 16.0146 | 13.2342 | 9.3592

## Slide 25: RESULTS
Fly ash samples
Chart: Al
| 6.0 | 10.0 | 18.0 | 21.0 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0 | 150.0 | 180.0 | 210.0
F1 | 0.30968 | 0.47791 | 0.54707 | 0.35831 | 0.24308 | 0.25888 | 0.19831 | 0.24382 | 0.11566 | 0.10036 | 0.11263 | 0.094225
F2 | 0.021731 | 0.027142 | 0.060325 | 0.014776 | 0.015023 | 0.01695 | 0.025672 | 0.024807 | 0.022695 | 0.02246
F3 | 0.021224 | 0.015875 | 0.029625 | 0.015035 | 0.015418 | 0.019124 | 0.014825 | 0.01711 | 0.01359 | 0.024189 | 0.017407 | 0.023621
C1 | 0.023139 | 0.024004 | 0.02608 | 0.027797 | 0.031627 | 0.029193 | 0.029736 | 0.040213 | 0.033875 | 0.058003
C2 | 0.05332 | 0.13165 | 0.05505 | 0.14842 | 0.044154 | 0.059127 | 0.059127 | 0.054605 | 0.080228
C3 | 7.4063 | 8.811 | 7.52 | 6.1252
Chart: Si
| 6.0 | 10.0 | 18.0 | 21.0 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0 | 150.0 | 180.0 | 210.0
F1 | 0.15604 | 0.12273 | 0.24354 | 0.17316 | 0.17971 | 0.21157 | 0.19742 | 0.20832 | 0.19972 | 0.19914 | 0.22002 | 0.2478
F2 | 0.078308 | 0.06212 | 0.05685 | 0.057123 | 0.06339 | 0.078059 | 0.062939 | 0.084445 | 0.076267 | 0.091388
F3 | 0.044519 | 0.053278 | 0.043652 | 0.041492 | 0.033564 | 0.051118 | 0.048068 | 0.059283 | 0.045385 | 0.050358 | 0.042454 | 0.053669
C1 | 0.0097678 | 0.0098509 | 0.014681 | 0.015156 | 0.023001 | 0.026657 | 0.032437 | 0.051248 | 0.055568 | 0.08327
C2 | 0.0134 | 0.022064 | 0.013744 | 0.010753 | 0.0082011 | 0.022408 | 0.022408 | 0.031001 | 0.055663
C3 | 0.028746 | 0.047866 | 0.060767 | 0.041825

## Slide 26: RESULTS
Cement samples
Chart: OPC 1
| 7.5 | 11.66666666666667 | 15.66666666666667 | 21.66666666666667 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0
K | 296.0262 | 323.7566 | 329.8814 | 372.1981 | 374.1484 | 386.3038 | 395.8956 | 384.5279 | 361.249
Na | 121.2322 | 124.3703 | 117.1033 | 126.7429 | 125.3751 | 125.1789 | 128.0197 | 124.6482 | 118.4329
S | 193.4317 | 194.2249 | 183.8592 | 200.3566 | 195.3612 | 193.3607 | 192.3371 | 189.4953 | 179.8676
Ca | 41.2002 | 31.2638 | 24.1011 | 25.177 | 24.235 | 25.2538 | 25.6796 | 23.508 | 21.698
Si | 0.06591 | 0.091665 | 0.092139 | 0.0968 | 0.13757 | 0.080955 | 0.074396 | 0.077972 | 0.081034
Al | 0.055264 | 0.071053 | 0.073116 | 0.069274 | 0.10213 | 0.067791 | 0.075282 | 0.070604 | 0.11604
Chart: OPC 3
| 7.5 | 14.5 | 20.0 | 25.0 | 30.0 | 60.0 | 90.0 | 120.0 | 180.0
K | 391.1692 | 407.7292 | 398.7556 | 408.8844 | 404.9532 | 407.1102 | 398.574 | 405.5142 | 393.7686
Na | 136.0308 | 138.429 | 138.9698 | 140.2319 | 142.741 | 143.1405 | 146.0106 | 147.5127 | 153.5893
S | 260.9824 | 264.2783 | 261.1113 | 262.6576 | 265.7129 | 259.3893 | 261.7958 | 269.9834 | 271.7953
Ca | 19.1855 | 19.5415 | 19.6679 | 20.236 | 20.6368 | 21.6299 | 22.1551 | 21.6041 | 21.1267
Si | 0.15341 | 0.17577 | 0.16322 | 0.12309 | 0.1087 | 0.13296 | 0.13415 | 0.097992 | 0.12556
Al | 0.21601 | 0.1934 | 0.18192 | 0.13692 | 0.20408 | 0.26443 | 0.21318 | 0.22651 | 0.2463

## Slide 27: RESULTS
Cement samples
Chart: Ca
| 7.5 | 11.66666666666667 | 15.66666666666667 | 21.66666666666667 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0
Cement 1 | 41.2002 | 31.2638 | 24.1011 | 25.177 | 24.235 | 25.2538 | 25.6796 | 23.508 | 21.698
Cement 2 | 17.6177 | 17.9891 | 19.4334 | 18.2319 | 17.955 | 19.2525 | 17.6398 | 20.0996 | 19.2446
Cement 3 | 19.1855 | 19.5415 | 19.6679 | 20.236 | 20.6368 | 21.6299 | 22.1551 | 21.6041 | 21.1267
Chart: Si
| 7.5 | 11.66666666666667 | 15.66666666666667 | 21.66666666666667 | 30.0 | 45.0 | 60.0 | 90.0 | 120.0
Cement 1 | 0.06591 | 0.091665 | 0.092139 | 0.0968 | 0.13757 | 0.080955 | 0.074396 | 0.077972 | 0.081034
Cement 2 | 0.049741 | 0.052477 | 0.046032 | 0.047551 | 0.03433 | 0.031054 | 0.02776 | 0.053693 | 0.041759
Cement 3 | 0.15341 | 0.17577 | 0.16322 | 0.12309 | 0.1087 | 0.13296 | 0.13415 | 0.097992 | 0.12556

## Slide 28: CONCLUSION
ICP-OES is an excellent tool for analysis of readily available elements within a pore solution as a function of time
Capable of detecting low limits of Si and Al
Reliable and repeatable measurements verified using NIST traceable certified standards

## Slide 29: FUTURE WORK
Comparison study using X-Ray Fluorescence for verification of ICP method
Analyze pore solutions of cement supplemented with fly ash at multiple levels by mass
Use of data for the development of structure-property models for cementitious materials

## Slide 30
Thank You
LaKesha.Perry@nist.gov